## Sheet1
Filtering (Credit/Age/Class) | Work Type | Case Type (Related Case) | Country | Case Description | Investigation | Corrective Action | SWO Number | RMED FaultCode L1(New) | RMED FaultCode L2(New) | RMED FaultCode L3(New) | RMED FaultCode L4(New) | Level 5 | Case Description Eng | Investigation Eng | Corrective Action Eng
Yes | Dummy | y | DUMM | Dummy | Dummy | Dummy | Dummy1 | [5 empty] | Dummy | Dummy | Dummy
No | Visit | Complaint | RKK | tHbキャル行うとクリーニングになる | ABL800AQC用のアダプター使用していたため、液面までプローブが届かずクリーニングになっていた。 | アンプルダイレクトでの吸引で問題なく実施 | SWO-158049 | User error | Wrong QC Adapter | Wrong QC Adapter | Wrong QC Adapter |  | tHb cal do and get cleaned | Because the adapter for ABL800AQC was used, it had become cleaning without the probe reached the liquid level. | Suction with ampoule direct without any problems
No | 4 | Y | RFRA | Blank | erreur fuite détéctée |  | D1E2D74B2A077F48 | Defect | HW | Inlet module | Inlet Gasket |  | Blank | leaked error
Yes | 4 | Y | RFRA | Blank | Erreur étalonnage pompe |  | D1E2A4432A267409 | Re classified RMED | Double entry |  |  |  | Blank | Pump calibration error
No | Visit | Complaint | RLTD | Agnes Lyczek, Emergency Dept. St Helier Hospital, 0208 2963432 | Analyser requesting inlet gasket holder replacement due to leak detected error. Leak traced to help in Oxi pinch valve tube, tube replaced. On analyser restart, some Oxi parameters failing on quality controls O2Hb, MetHb. Ran tHb calibration - but QC not improved. Replaced faulty hemolyser, performed Oximetry test - passed and tHb calibration - passed. Ran Quality Controls all parameters now passed. Analyser ready for use. Call requested outside of normal working hours (Bank Holiday). Call charges authorised by Craig Wood, Matron ED dept.(Craig.Wood@esth.nhs.uk) Charges: Callout charge (Out of hours) £312.00 On-site labour on of 2 hours £241.10 Parts covered by service contract. Total Charge £553.10 plus VAT Customer to supply purchase order to cover above charges within 7 days. | Replaced split Oxi pinch valve tube. Replaced faulty hemolyser. | SWO-157241 | Defect | HW | Oxi module | Hemolyzer |  | Agnes Lyczek, Emergency Dept. St Helier Hospital, 0208 2963432 | Analyser requesting inlet gasket holder replacement due to leak detected error. Leak traced to help in Oxi pinch valve tube, tube replaced. On analyser restart, some Oxi parameters failing on quality controls O2Hb, MetHb. Ran tHb calibration - but QC not improved. Replaced faulty hemolyser, performed Oximetry test - passed and tHb calibration - passed. Ran Quality Controls all parameters now passed. Analyser ready for use. Call requested outside of normal working hours (Bank Holiday). Call charges authorised by Craig Wood, Matron ED dept.(Craig.Wood@esth.nhs.uk) Charges: Callout charge (Out of hours) £312.00 On-site labour on of 2 hours £241.10 Parts covered by service contract. Total Charge £553.10 plus VAT Customer to supply purchase order to cover above charges within 7 days. | Replaced split Oxi pinch valve tube. Replaced faulty hemolyser.
No | 3 | Y | RESP | ABL90 laboratorio no transmite muestras a omega |  | Se acude, con ayuda de PBM, se chequean todas las conexiones y servicios de Aqure. Limpieza de colas de salida. Equipos empiezan a transmitir OK | D1E302062A16A2E2 | Setup | SW | His/lis | His/lis |  | ABL90 lab does not transmit samples to omega |  | All Aqure connections and services are checked with the help of PBM. Cleaning of outbound queues. Teams start broadcasting OK
No | Visit | Complaint | RCHN | 试剂包漏夜 | 试剂包漏夜 | 更换试剂包后修复 | SWO-160765 | Assisted Replacement | SP | SP | Leaking |  | Reagent bag leak night | Reagent bag leak night | Repair after replacement of reagent kit
No | Visit | Complaint | RCHN | 报1052无法获取试剂包，报5691，报5689错误 | 报1052无法获取试剂包，报5691，报5689错误 | 更换处理包，测试卡，清理内部灰尘后修复 by jesson | SWO-157396 | Cleaned | HW | Storage Module | Sensor, storage wheel | 5689 | Reported 1052 unable to obtain reagent package, reported 5691, reported 5689 errors | Reported 1052 unable to obtain reagent package, reported 5691, reported 5689 errors | Replace the processing pack, test the card, clean the internal dust and repair by jesson
No | Visit | Complaint | RDEU | Datum Uhrzeit falsch | Datum Uhrzeit nicht korrekt | Gerät überprüft, auf der Windowsebene Uhrzeit und Datum korrigiert, Reinstallation nicht möglich, da Verbrauchsmaterial nicht mehr akzeptiert wird, Datum und Uhrzeit zurückgestellt, QC Bereiche eingegeben | SWO-157230 | [5 empty] | Date time incorrect | Date time not correct | Device checked, time and date corrected at Windows level, reinstallation not possible because consumables are no longer accepted, date and time deferred, QC ranges entered
No | Visit | Complaint | RHUN | kalibrációk: OK mérések és QC: tHb szekció nincs mérve, kampós szelep a mintát vissza engedi mérés közben | - | kampós szelep cseréje, teljes OXI teszt, esedékes membrán cserék és tHb kalibráció elvégzése, QC mérés | SWO-157155 | Defect | HW | Oxi module | Pinch valve |  | Calibrations: OK Measurements and QC: tHb section is not measured, Toggle valve allows the specimen to return during measurement | - | Toggle valve replacement, full OXY test, due membrane substitutions and tHb calibration, QC measurement
No | Visit | Complaint | RDEU | 20.12.2016 18:45 Fr. Gassner 051398013278: OXI-Fehlermeldung Neonlamp to high 21.12.2016 Frau Gassner : 051398013278: Oxymetrie Hardware Fehler, 316 | Neon Position Fehler. Rücklieferung und Installation des reparierten Moduls. | OXI Modul 902-833R1966N010 installiert, Leihmodul abgezogen. Lac. QC Fehler beseitigt, fehlenden Dichtring ersetzt, Membran erneuert. OXI Complete Test und Service tHb Kal. i.O. | SWO-157237 | Revisit required | HW | Oxi module | Oxi module |  | 20.12.2016 18:45 Fr. Gassner 051398013278: OXI error message Neonlamp to high 21.12.2016 Ms. Gassner : 051398013278: Oxymetry Hardware Error, 316 | Neon position error. Return and installation of the repaired module. | OXI module 902-833R1966N010 installed, loan module deducted. Lac. QC error eliminated, missing sealing ring replaced, membrane renewed. OXI Complete Test and Service tHb Kal. Ok.
No | Visit | Complaint | RDEU | 02.01.2017, H. Felgentreu, Tel. 030450631168: Halterung für Probennehmer gebrochen | Halterung für Probennehmer abgebrochen | entfernt und ersetzt, Gerät geöffnet und abgebrochene Teile und Proberöhrchen entfernt, Patmessungen durchgeführt, i.O. | SWO-157355 | Defect | HW | Aspiration Module | Tube holder |  | 02.01.2017, H. Felgentreu, Tel. 030450631168: Bracket for sampler broken | Bracket for samplers aborted | removed and replaced, device opened and broken parts and sample tubes removed, pat measurements performed, i.O.
No | 4 | Y | RESP | Error 1258 y 1259 HCT assay conductivity |  | Limpieza del la unidad de Hct con hipoclorito segun tip. Comprobacion con qc y varias muestras. | D1E314862A038D61 | Cleaned | HW | Aspiration Module | Needle Wash Unit | 1258 | Error 1258 y 1259 HCT assay conductivity |  | Cleaning the Hct unit with hypochlorite according to tip. Check with qc and multiple samples.
Yes | Visit | Complaint | RDEU | THb kal erforderlich | Fehlersuche | Thb kal durchgeführt | SWO-157385 | Re classified RMED | ECKHI Email 24.jan-21017 | Oxi module | tHb Cal |  | THb kal required | Troubleshooting | Thb kal performed
No | Visit | Complaint | RLTD | Leak detected | Leakage from tubing near oxi valve | Replaced black pinch valve tubing with new | SWO-157489 | Defect | HW | Oxi module | Pinch valve tube |  | Leak detected | Leakage from tubing near oxi valve | Replaced black pinch valve tubing with new
No | Visit | Complaint | RDAN | Der har været problemer ved skift at væskepakken. Selv om der er skiftet væskepakke, kommer ABL90 igen og ønsker væskepakken skiftet igen. Jeg har talt med Jannie og aftalt at der kommer tekniker 3/1 2017 kl 0930. Jannie har nummer 51680455 | Fandt at inletpakningsholder var defekt og derfor kom der fejl på vakuum test og fejl på væskepakken. | Jeg har skiftet inletpakningsholder. Der er skiftet væskepakke. Alle test under Fluid transport i service er kørt. Der er kørt kal. og QCer. Der er kørt tHb kal. Alt er OK igen. | SWO-157515 | Defect | HW | Inlet module | Inlet Gasket |  | There have been problems when changing the liquid package. Even if the liquid package is changed, ABL90 comes again and wants the liquid package changed again. I have talked to Jannie and agreed that there will be technician 3/1 2017 at 0930. Jannie has number 51680455 | Found that the inlet packet holder was defective and therefore the vacuum test and the defect of the liquid package were faulty. | I've changed the inlet pack holder. A liquid package has been switched. All tests during FLUID transport in service are run. There's a run of the Kal. and QCer. A tHb Kal has been run. Everything is OK again.
No | Visit | Complaint | RDEU | 30.12.2016; Sr. Anke; Tel.: 06722490118; ABL verlangt eine Spülung 30.12.2016; 19:20 Uhr; Sr. Elke; Tel.: 06722490118; Fehler bleibt auch mit dem neuen LP bestehen 03.01.2017, 12:25 Uhr ==> Anrufer Schw. Christina, Tel.: 06722490118 ==> Gerät steht noch immer in der Fehlerbehebung, am 02.01.2017 wollte Techniker kommen, war jedoch noch nicht da und hat sich nicht gemeldet | Gerät meldet Einlassdichtung Wechsel SP wurde bereits gewechselt | Fliesswege gespült ,Dichtung überprüft und erneut eingesetzt. SP erneut eingesetzt Test wiederholt, Aktivitäten kontrolliert. Kalibrierungen und QC durchgeführt. | SWO-157494 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | 30.12.2016; Sr. Anke; Tel.: 06722490118; ABL requires flushing 30.12.2016; 7:20 p.m.; Sr. Elke; Tel.: 06722490118; Error persists with the new LP 03.01.2017, 12:25 o'clock ==> Caller Schw. Christina, Tel.: 06722490118 ==> Device is still in troubleshooting, on 02.01.2017 technicians wanted to come, but was not yet there and did not report | Device reports inlet seal change SP has already been changed | Flow paths rinsed, seal checked and re-inserted. SP re-used test repeatedly, controlling activities. calibrations and QC.
No | Visit | Complaint | RLTD | Settings not registering lac | Customer required settings not in place. | Set up as per request, adult and baby report layouts available. | SWO-157800 | Setup | SW | Patient Layout | Patient Layout |  | Settings not registering lac | Customer required settings not in place. | Set up as per request, adult and baby report layouts available.
No | Visit | Complaint | RLTD | Natalie Vaughan Tel 0208 8695402 The customer has been told to ring in each time about this. 29/12 Inlet probe still stuck - called in to report the fault. | Inlet intermittent stuck out. Tech line raised . | Installed new auto inlet motor assembly as part of tech line case to see if this resolves the issue. Customer will monitor the unit. Part 903-344 not on CRM. | SWO-157452 | Defect | HW | Auto Inlet module | Motor assembly | Grub | Natalie Vaughan Tel 0208 8695402 The customer has been told to ring in each time about this. 29/12 Inlet probe still stuck - called in to report the fault. | Inlet intermittent stuck out. Tech line raised . | Installed new auto inlet motor assembly as part of tech line case to see if this resolves the issue. Customer will monitor the unit. Part 903-344 not on CRM.
No | Visit | Complaint | RCHN | 测试卡信号连接板错误 | 测试卡信号连接板错误，借备件902-972一个 |  | SWO-157393 | [5 empty] | Test card signal connection board error | Test card signal connection board error, borrowed spare parts 902-972 one
No | Visit | Complaint | RCHN | 报错5784,1020，测试卡断裂 | 报错5784,1020，测试卡断裂，借备件636-652 | 5784故障修复后，机器开机后出现错误故障1020，检查机器状态，机器的孵育盘和环境温度没有显示，更换样机上的636-652后修复 by jesson | SWO-157344 | Defect | Consumables | Test Kit | Cartridge | 5784 (Broken) | 5784,1020, test card broken | Error 5784,1020, test card broken, spare part 636-652 | 5784 fault repair, the machine after the start-up error failure 1020, check the status of the machine, the machine incubation disk and ambient temperature is not shown, replace the prototype 636-652 after repair by jesson
No | Visit | Complaint | RCHN | 报1031错误 | 报1031错误 | 自行更换备用光单（R0358N008）修复 by andrew | SWO-157336 | Defect | HW | Process Module | Optical Unit | 1031 | Reported 1031 error | Reported 1031 error | Replace your own backup light sheet (R0358N008) to repair by Andrew
Yes | Visit | Complaint | RCHN | 报1031错误 | 报1031错误 | 自行更换备用光单修复 | SWO-157345 | Re classified RMED | Double entry | Process Module | Optical Unit | 1031 | Reported 1031 error | Reported 1031 error | Replace your own backup light sheet repair
No | Visit | Complaint | RCHN | 光学单元报错1031 | 光学单元报错1031，借备件903-364 | 更换903-364后修复by jesson | SWO-157314 | Defect | HW | Process Module | Optical Unit | 1031 | Optical unit error 1031 | Optical unit error 1031, spare parts 903-364 | Repair by jesson after replacement 903-364
No | Visit | Complaint | RCHN | 报错1031光学强度超标 | 报错1031光学强度超标，借备件903-364 | 使用备用机备件903-364修复 by jesson | SWO-157365 | Defect | HW | Process Module | Optical Unit | 1031 | Reported error 1031 optical strength exceeds the standard | Error 1031 optical strength exceeded the standard, borrowed spare parts 903-364 | Repair by jesson with spare parts 903-364
Yes | 15 | Y | RAME | Z inlet probe bent. |  | Inspected flex Q module and found that it was slightly out of alignment possibly causing the error. | D1E3ED6BA0293B35 | [5 empty] | Z inlet probe bent. |  | Inspected flex Q module and found that it was slightly out of alignment possibly causing the error.
No | 3 | Y | RPOL | Blank | Wymiana ig³y, kalibracja inletu, sprawdzenie kalibracji podajnika, pomiary kontrolne w normie, analizator sprawny technicznie. |  | D1E4343DA0343A65 | Defect | HW | Inlet module | Probe bend |  | Blank | Replacing the wire, calibrating the inlet, checking the calibration of the feeder, control measurements in the standard, the analyser technically efficient.
No | Visit | Complaint | RIND | Received a call from sasoon hospital that all the oximetry parameters are disabled. | Oxi-Module is required to be replaced as suggested by Prakateesh sir. | Replaced the oxi-module sent by Prakatheesh Sir, restarted the machine done the Complete Oximetry Test, after a successful 2 point calibration run 3 samples, results are OK, machine is working fine. Kept the machine under observation for the next 2 days. | SWO-158256 | Revisit required | HW | Oxi module | Oxi module |  | Received a call from sasoon hospital that all the oximetry parameters are disabled. | Oxygen-Module is required to be replaced as suggested by Prakateesh sir. | Replaced the oxi-module sent by Prakatheesh Sir, restarted the machine done the Complete Oximetry Test, after a successful 2 point calibration run 3 samples, results are OK, machine is working fine. Kept the machine under observation for the next 2 days.
No | Visit | Complaint | RDEU | 03.01.2017, 10:40 Uhr ==> Anrufer Herr Lehnert, Tel.: 03391394472 ==> Kalibrierungen werden abgebrochen, Lösungsbehälter ausreichend gefüllt 03.01.2017, 10:55 Uhr ==> Anrufer Herr Lehnert, Tel.: 03391394472 ==> jetzt noch zusätzlich Fehler 0531: upper LS-Error pH/BG-Modul | LS-Fehler im PH/BG-Modul oben | Proteinentfernungsprogramm durchgeführt, LS-Justage, fehlende untere Kappe der PO2-Elektrode ersetzt, Messkammer gereinigt, Kal, QC | SWO-157670 | Defect | HW | pH/BG Module | Valve, air | Membrane | 03.01.2017, 10:40 a.m. ==> Caller Mr. Lehnert, Tel.: 03391394472 ==> Calibrations are cancelled, solution containers are sufficiently filled 03.01.2017, 10:55 a.m. ==> Caller Mr. Lehnert, Tel.: 03391394472 ==> now additional error 0531: upper LS-Error pH/BG-Module | LS errors in PH/BG module above | Protein removal program performed, LS adjustment, missing lower cap of PO2 electrode replaced, measuring chamber cleaned, Kal, QC
No | Visit | Complaint | RDEU | Bereits wenige Tage nach Installation des ABLs / der Lac Elektrode sinkt die Empfindlichkeit unterhalb Grenze | s. case | Standzeit der Lac Membranen aktuell etwa 3 bis 7 Tage. Elektrode sowie Kontaktkappe auf Korrosion untersucht und Reinigungsverfahren mit Hypochlorid und Aqua Dest durchgeführt, beide Dichtringe ersetzt, membraniert, initialisiert, Kalibrationen (Empf. ca. 500pA $mM) und QC Messung, OK | SWO-157530 | Cleaned | Electrode | Electrode | Lac | Cleaned e-tip (378) | Just a few days after installation of the ABL / Lac electrode, the sensitivity below the limit decreases | s. case | Life of the lac membranes currently about 3 to 7 days. Electrode and contact cap tested for corrosion and cleaning procedures performed with hypochloride and aqua dest, both sealing rings replaced, membraneized, initialized, calibrations (emitting approx. 500pA $mM) and QC measurement, OK
Yes | Visit | Complaint | RDEU | Bereits wenige Tage nach Installation des ABLs / der Lac Elektrode sinkt die Empfindlichkeit unterhalb Grenze | Nachtrag Dichtungen | Nachtrag SWO Dichtungen | SWO-157548 | Re classified RMED | Double entry |  |  |  | Just a few days after installation of the ABL / Lac electrode, the sensitivity below the limit decreases | Addendum Seals | Supplement SWO Seals
No | 1 | Y | RAME | Z Customer states Ca parameter dissapearing from panel. |  | Upon arrival, found unit setup using Dynamic Parameter feature which is not needed. | D1E41109A03738C3 | Setup | Patient Layout | Patient Layout | Patient Layout |  | Z Customer states Ca parameter dissapearing from panel. |  | Upon arrival, found unit setup using Dynamic Parameter feature which is not needed.
No | 5 | Y | RAME | P instrument leaking from waste connector, instrument in standby |  | checked data logs for errors, checked flow of solutiuons, checked pumps and tubing, waste tubing disconnected from connector | D1E48093A0218119 | Defect | HW | Basic Unit | Waste tubing |  | P instrument leaking from waste connector, instrument in standby |  | checked data logs for errors, checked flow of solutiuons, checked pumps and tubing, waste tubing disconnected from connector
No | 2 | Y | RAME | L Analyzer not connecting to Radiance |  | Troubleshot and found that after the customer had powered off the analyzer, the static IP settings and defaulted back to DHCP. | D1E5CD44166271EF | Setup | SW | His/lis | His/lis |  | L Analyzer not connecting to Radiance |  | Troubleshot and found that after the customer had powered off the analyzer, the static IP settings and defaulted back to DHCP.
No | 5 | Y | RAME | Z Analyzer will not switch back to power once unplugged |  | When analyzer is unplugged from mains, it switches to battery but when plugged back to mains, the power supply will not switch back to mains. | D1E528C8163E113A | Defect | HW | Power Supply | Power Supply |  | Z Analyzer will not switch back to power once unplugged |  | When analyzer is unplugged from mains, it switches to battery but when plugged back to mains, the power supply will not switch back to mains.
No | 9 | Y | RAME | Z Analyzer will not switch back to power once unplugged. |  | When analyzer is unplugged from mains, it switches to battery but when plugged back to mains, the power supply will not switch back to mains. | D1E52A691657113E | Defect | HW | Power Supply | Power Supply |  | Z Analyzer will not switch back to power once unplugged. |  | When analyzer is unplugged from mains, it switches to battery but when plugged back to mains, the power supply will not switch back to mains.
No | 7 | Y | RAME | Z Battery will not function |  | Checked battery, will not function. | D1E52B1016281141 | Defect | HW | Power Supply | Power Supply |  | Z Battery will not function |  | Checked battery, will not function.
No | Visit | Complaint | RDEU | Es müssen häufig Lösungspack gewechselt werden | Dateien gesichtet, Sichtkontrolle, Test im Servicemenü | Kein Fehler feststellbar, kurzeinweisung, Änderungswünsche Konfiguration aufgenommen und mit Labor besprochen, Pico50 Probennehmer könnten das Problem beheben. | SWO-157881 | No fault found | No fault found | No fault found | Wetability |  | Solution packs often need to be changed | Files spotted, visual inspection, test in the service menu | No error detectable, short instruction, change requests configuration included and discussed with laboratory, Pico50 samplers could fix the problem.
No | Visit | Complaint | RDEU | Proben werden nicht angesaugt, Einlass klemmt. | Sichtkontrolle, Test im Service Menü | Inlet Probe ersetzt da verbogen, Inlet Flow getestet. Kein weiteres, technisches Problem erkennbar. Funktionskontrolle, Probenmessungen. | SWO-157880 | Replaced | HW | Inlet module | Probe bend |  | Samples are not sucked in, inlet clamped. | Visual inspection, test in the service menu | Inlet sample replaces there bent, Inlet Flow tested. No further technical problem recognizable. Functional control, sample measurements.
No | 5 | Y | RPAC | Blank | Inlet Jammed |  | D1E53520161111FC | Adjustment | HW | Auto Inlet module | Inlet drive gear | Grub | Blank | Inlet Jammed
No | Visit | Complaint | RDEU | 03.01.2017; Sr. Stefanie; Tel.: 023815891214; Einlass schließt nicht | Einlass schließt nicht. Uhrzeit ändert sich nach Neustart und muss manuell gestellt werden | Einlass lässt sich elektrisch nicht verstellen. Gerät demontiert und Einlass mechanisch in die richtige Position gebracht. Neustart durchgeführt und Einlass im Service kalibriert. Wegen unterbrochener Spülung, Sensorcassette erneuert. Trotz diverser Kal und QC ist Glu und Lac außer Bereich. Anschließend musste der Lösungspack auch erneuert werden. Diverse Kal und QC erneut durchgeführt. Uhrzeit Gestellt , Date $ Time Einstellungen vorgenommen und abgespeichert. | SWO-158337 | Adjustment | HW | Auto Inlet module | Inlet drive gear |  | 03.01.2017; Sr. Stefanie; Tel.: 023815891214; Entrance does not close | Entrance does not close. Time changes after restart and must be set manually | Inlet cannot be electrically adjusted. device and mechanically placed the inlet in the correct position. restart and inlet calibrated in the service. Due to interrupted flushing, sensor cassette renewed. Despite various Kal and QC, Glu and Lac is out of range. Subsequently, the solution pack had to be renewed. Several Kal and QC performed again. Time Set , Date - Time settings made and saved.
No | Visit | Complaint | RIND | Solution Pack Not connecting but after remove and reconnect showing OK but after Some time again Showing Not Connected. | Solution Pack Not connecting but after remove and reconnect showing OK but after Some time again Showing Not Connected. | Solution Pack Not Inserted properly. Waist Leak Also Found. There is a manufacturing fault in Solution Pack. Door closed properly. and Puted a paper piece to retail. | SWO-158409 | [5 empty] | Solution Pack Not connecting but after remove and reconnect showing OK but after Some time again Showing Not Connected. | Solution Pack Not connecting but after remove and reconnect showing OK but after Some time again Showing Not Connected. | Solution Pack Not Inserted properly. Waist Leak Also Found. There is a manufacturing fault in Solution Pack. Door closed properly. and Puted a paper piece to retail.
No | Visit | Complaint | RCHN | 测试卡参比电极错误 | 测试卡参比电极错误 | 更换测试卡修复 | SWO-157661 | Assisted Replacement | SC | SC | Reference |  | Test card reference electrode error | Test card reference electrode error | Replacement test card repair
No | Visit | Complaint | RCHN | 测试卡定标不过 | 测试卡定标不过 | 更换测试卡修复 | SWO-157663 | Adjustment | SC | SC | Cal |  | Test card calibration is however | Test card calibration is however | Replacement test card repair
No | Visit | Complaint | RCHN | 光单上下传感器无法连接 | 光单上下传感器无法连接，借备件636-550一个 | 更换636-550修复 by jesson | SWO-157359 | Defect | HW | Process Module | Process Unit sensor | 5257 | Light single up and down sensor cannot be connected | Light single up and down sensor can not be connected, borrowed spare parts 636-550 one | Replace 636-550 repair by jesson
No | Visit | Complaint | RCHN | 报5784错误 | 报5784错误，借886-625一个 | 更换配件886-625修复 by jesson | SWO-157644 | Defect | Consumables | Test Kit | Cartridge | 5784 (Broken) | Reported 5784 errors | Reported 5784 error, borrow886-625 one | Replacement accessory 886-625 repair by jesson
No | Visit | Complaint | RCHN | 报错1024，定标两侧都失败 | 报错1024，定标两次都失败,借903-364 | 清洁处理盘灰尘无效，更换903-364后修复 by jesson | SWO-157353 | Defect | HW | Process Module | Optical Unit | 1024 | 1024, bid on both sides failed | 1024, the bid failed twice, borrowed 903-364 | Cleaning the dust of the treatment disc is invalid, repair by jesson after replacing 903-364
Yes | Visit | Complaint | RCHN | 测试卡仓门与面板支撑架断裂 | 测试卡仓门与面板支撑架断裂借备件903-224修复 | 更换903-224修复 by jesson | SWO-157303 | Re classified RMED | OOB |  |  |  | Break silo door and panel support frame | Test card door and panel support frame break borrowed spare part 903-224 repair | Replace 903-224 Fix by Jesson
No | Visit | Complaint | RIND | crea calibration expire. | Check the logs. check the calibration logs. | Replace the Crea consumable & recalibrate system. | SWO-158476 | Assisted Replacement | Consumables | Solution bottle | Crea |  | crea calibration expire. | Check the logs. check the calibration logs. | Replace the Crea consumable & recalibrate system.
No | 3 | Y | RCZE | Blank | nemìøí oxi parametry a selhání v meøení, údajnì vypadlá hadièka potom ok, po kalibraci ale nejde i po výmìnì tìsnìní a odstranìní proteinù, problém z pøedešlého dne |  | D1E781B60261A2B5 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | Blank | Does not measure the Oxi parameters and the failure in the measurement, the purging hose is then OK, but after calibration, it is not possible to change the protein seal and remove the problem from the previous day
No | Visit | Complaint | RDEU | 03.01.2017, Frau Weigand, Tel.: 093319086511. Datenübertragung/ Verbindung verloren an OSM | Überprüfung der Datenübertragung mit ASTM Net --> Übertragung iO. Kontrolle der IP Adressen --> Gateway Adresse Fehlerhaft | Gateway Adresse korrigiert. Übertragung hergestellt, Testmessung mit Datensendung iO. | SWO-157777 | Setup | SW | His/lis | His/lis |  | 03.01.2017, Mrs. Weigand, Tel.: 093319086511. Data transfer/ Connection lost to OSM | Checking data transfer with ASTM Net --> Transfer iO. Control of IP addresses --> Gateway address Incorrect | Gateway address corrected. Transmission, test measurement with data transmission iO.
No | Visit | Complaint | RIND | optical senstivity out of range. | optical senstivity out of range. | cleaned process wheel. cleaned optical unit. measured darkness value. run a test analyzer working fine. | SWO-158465 | Cleaned | HW | Process Module | Process wheel |  | optical senstivity out of range. | optical senstivity out of range. | cleaned process wheel. cleaned optical unit. measured darkness value. run a test analyzer working fine.
No | Visit | Complaint | RHUN | Printer nyitókallantyúja törött. | Printer nyitógombja eltörött. | Printer Unit csere. Kalibrálás, QC mérés. | SWO-158076 | Defect | HW | Printer module | Printer module |  | Printer is broken. | Printer's Open button is broken. | Printer Unit replacement. Calibration, QC measurement.
No | Visit | Complaint | RDEU | 05.01.2017; Sr. Heidi; Tel.: 071131032434; Intervention erfoderlich 05.01.2017, Sr. Heidi, Tel. 071131032434, Probennadel defekt. Kd. hat keine vorrätig. Besuch KDB erforderlich. | SP lässt sich nicht installieren ..keine Lös. erkannt. | Einlass überarbeitet. Dichtung durch Kd. ersetzt. Einlassnadel deformiert. Nadel wieder gerichtet. Einlass getestet ok. Nach einigen Reinstallationen SP endlich ok. Kal 40 Min abgewartet. | SWO-158185 | Replaced | HW | Inlet module | Probe bend |  | 05.01.2017; Sr. Heidi; Tel.: 071131032434; Intervention essential 05.01.2017, Sr. Heidi, tel. 071131032434, sample needle defective. Kd. has none in stock. Visit KDB required. | SP cannot be installed .. no loes. Recognized. | Entrance revised. Seal replaced by Kd. inlet needle deformed. Needle again directed. Inlet tested ok. After some reinstallations SP finally ok. Kal 40 min waited.
Yes | 3 | Y | RFRA | Blank | affichage fenetre remplacement du pack qui apparait lors du passage d'un BG |  | D1E6BF698C50F79F | Re classified RMED | Remote Support |  |  |  | Blank | Window display replacing the pack that appears when passing a BG
No | Visit | Complaint | RDEU | Festplattenspeicher < 15%, trotz mehrmaligem Neustart des Analysators | Fehlersuche --> CF-Karte überprüft --> CF-Karte defekt --> CF-Karte erneuern | Konfiguration gesichert, CF-Karte erneuert, Windows-Software und Analysatorsoftware V3.1MR8 installiert, LP und SC wieder installiert, gesicherte Konfiguration wieder installiert, | SWO-158104 | Defect | HW | DMS_PC | CF Card |  | Disk space of 15%, despite restarting the analyzer several times | Troubleshooting --> CF card checked --> CF card defective --> Renew CF card | Configuration backed up, CF card renewed, Windows software and analyzer software Installed V3.1MR8, LP and SC reinstalled, Backed up configuration reinstalled,
No | Visit | Complaint | RDEU | Lösungen können nicht gepumpt werden | Flüssigkeitstransportfehler | Fehlersuche; Fließwege gespült; div. Programme im Service durchgef.; LP neu initialisiert; Neustart; Startphase abgewartet; Kal. + QC i.O. | SWO-158087 | Assisted Replacement | SP | SP | FT |  | Solutions cannot be pumped | Flüssigkeitstransportfehler | Troubleshooting; Flow paths rinsed; various programs in the service; LP reinitialized; Restart; start-up phase awaited; Kal. + QC in the first place
No | 6 | Y | RFRA | Blank | Après changement cassette pCO2 Lac les calibrations ne passent pas |  | D1E7A90202291069 | No fault found | SC | SC | pCO2 |  | Blank | After changing cassette pCO2 Lake the calibrations do not pass
No | Visit | Complaint | RLTD | Thb set to 3 months, but not showing 3 months to next calibration. | Nick Birch emailed to say that after a thb the next due date did not change to 3 months away, even though it was set to three month interval. | Changed interval date to 2 months, then back to 3, which fixed issue. Reported back to Nick so that he can do the same to the other affected analysers. | SWO-158544 | No fault found | SW | Oxi module | tHb Cal |  | Thb set to 3 months, but not showing 3 months to next calibration. | Nick Birch emailed to say that after a thb the next due date did not change to 3 months away, even though it was set to three month interval. | Changed interval date to 2 months, then back to 3, which fixed issue. Reported back to Nick so that he can do the same to the other affected analysers.
No | 3 | Y | RFRA | Blank | Problèmes po2 et oximétrie (principalement sur l'HBF) |  | D1E6B4318C41C7A1 | Defect | HW | Oxi module | Hemolyzer |  | Blank | Problèmes po2 et oximétrie (principalement sur l'HBF)
No | Visit | Complaint | RIND | Lactate Parameter Failure, Done Hydration Procedure, Placed SC for Cooling in Fridge for 6 Hours, Try to Reinstalled SC But problem Remains, Need to replace SC with New One, PER Sent for Replacement. | Two Point calibration failure, Lactate Parameter Failure. | Done Hydration Procedure, Placed SC for Cooling in Fridge for 6 Hours, Try to Reinstalled SC But problem Remains, Need to replace SC with New One, PER Sent for Replacement. | SWO-158488 | [5 empty] | Lactate Parameter Failure, Done Hydration Procedure, Placed SC for Cooling in Fridge for 6 Hours, Try to Reinstalled SC But problem Remains, Need to replace SC with New One, PER Sent for Replacement. | Two Point calibration failure, Lactate Parameter Failure. | Done Hydration Procedure, Placed SC for Cooling in Fridge for 6 Hours, Try to Reinstalled SC But problem Remains, Need to replace SC with New One, PER Sent for Replacement.
No | Visit | Complaint | RCHN | 插卡底板被腐蚀 | 插卡底板被腐蚀，借备件902-972一个 | 更换902-972后修复 by jesson | SWO-157720 | [5 empty] | The card floor plate is corroded | The card base plate is corroded, with spare parts 902-972 one | Repair by jesson after replacement 902-972
No | Visit | Complaint | RCHN | 测试是卡漏液 | 测试是卡漏液 | 现场更换试剂包修复 | SWO-158272 | [5 empty] | The test is card leakage. | The test is card leakage. | On-site replacement reagent kit repair
No | Visit | Complaint | RCHN | myo定标失败 | myo定标失败，借备件903-364一个，myo定标卡944-214（12231） | 定标卡批号发货错误，已补发，故此单关单处理。 | SWO-158019 | Assisted Replacement | Consumables | Test Kit | Cartridge | Myo | myo calibration failure | myo bid failure, borrowed spare parts 903-364 one, myo calibration card 944-214 (12231) | The bid card lot number shipping error, has been reissued, so this single clearance processing.
Yes | Visit | Complaint | RCHN | myo定标失败 | myo定标失败，借备件903-364一个，myo定标卡944-214（12231） |  | SWO-158047 | Re classified RMED | Double entry |  |  |  | myo calibration failure | myo bid failure, borrowed spare parts 903-364 one, myo calibration card 944-214 (12231)
No | 60 | Y | RESP | Calibración abortada con frecuencia y error de deriva en el pH e iones. |  | Comprobar estado de los tubos, bombas y válvulas, cambiar tubos de bombas, calibrar y probar con controles. | D1E6BCF18C05CA94 | Defect | HW | pH/BG Module | Pump tubes |  | Frequently aborted calibration and drift error in pH and ions. |  | Check the status of tubes, pumps and valves, change pump tubes, calibrate and test with controls.
No | Visit | Complaint | RDAN | Jan melder at ABL'en ofte har drift på elektrolyterne ved midnats kalibreringen. | Afprøvet og undersøgt. | Elektrodeconnector udskiftet, slutafprøvet. SEJ sender ny referenceelektrode. | SWO-158140 | Assisted Replacement | Electrode | Electrode | Reference |  | Jan reports that the ABL often has operation on the electrolytes at the midnight calibration. | Tested and investigated. | Elektrodeconnector replaced, final tested. COOL sends new reference electrode.
No | Visit | Complaint | RDEU | 03.01.2017, 10:40 Uhr ==> Anrufer Herr Lehnert, Tel.: 03391394472 ==> Kalibrierungen werden abgebrochen, Lösungsbehälter ausreichend gefüllt 03.01.2017, 10:55 Uhr ==> Anrufer Herr Lehnert, Tel.: 03391394472 ==> jetzt noch zusätzlich Fehler 0531: upper LS-Error pH/BG-Modul 04.01.2017 13:30 Herr Schwarz 03391394637 meldet, dass der Fehler immer noch besteht. | sehr oft Error 500: Spüllösung vom pH/BG Modul erkannt, wenn nicht erwartet | Ventile von den Messmodulen kontrolliert -> beide Air Ventile hingen unten fest und Membran vom pH $BG Modul durchstoßen beide Ventile ausgetauscht LS Justage + Pumpenkal | SWO-158071 | Defect | HW | pH/BG Module | Valve, air |  | 03.01.2017, 10:40 a.m. ==> Caller Mr. Lehnert, Tel.: 03391394472 ==> Calibrations are cancelled, solution containers are sufficiently filled 03.01.2017, 10:55 a.m. ==> Caller Mr. Lehnert, Tel.: 03391394472 ==> now additional error 0531: upper LS-Error pH/BG-Module 04.01.2017 13:30 Mr. Schwarz 03391394637 reports that the error still exists. | very often Error 500: Rinsing solution detected by pH/BG module when not expected | Valves controlled by the measuring modules -> both air valves hung tightly at the bottom and penetrate membrane from the pH $BG module both valves replaced LS Adjustment + Pump Scale
Yes | Visit | Complaint | RDAN | Jan melder at ABL'en ofte har drift på elektrolyterne ved midnatskalibreringen. | Afprøvet og undersøgt. | Elektrodeconnector udskiftet, slutafprøvet. SEJ sender ny referenceelektrode. | SWO-158139 | Re classified RMED | Double entry |  |  |  | Jan reports that the ABL often operates on the electrolysis at the Midnatal calibration. | Tested and investigated. | Elektrodeconnector replaced, final tested. COOL sends new reference electrode.
Yes | Visit | Complaint | RDEU | Seit Installation des Analysators gibt es konstant Schwankungen hoch und runter im pCO2 Ergebnis L3 | Erneut bzw. immer noch zickzack-Linie bei pCO2 QC L3 | Regelmäßig QC Ausreißer nach unten bei pCO2 in Level drei. Erneute Anfrage an RDEU gestartet. Ggf. erneuter Datenextrakt und erneuter Tausch der Elektrode (zuletzt ersetzt am 26.08.2016). | SWO-158149 | Re classified RMED | Courtesy / Follow Up |  |  |  | Since the installation of the analyzer, there are constant fluctuations up and down in the pCO2 result L3 | Re-line or still zigzag line at pCO2 QC L3 | Regularly QC outliers down at pCO2 in level three. Re-enquiry to RDEU started. If necessary, re-data extract and re-exchange of the electrode (last replaced on 26.08.2016).
No | 9 | Y | RAME | 1 PH will not pass calibration |  | Looked in datalogs and found that status is reading 6.672, Sensitivity is reading 98.3 and Status 0 is reading 7.175 | D1E68E118C37A1AA | User error | Electrode | Electrode | Reference | Expired | 1 PH will not pass calibration |  | Looked in datalogs and found that status is reading 6.672, Sensitivity is reading 98.3 and Status 0 is reading 7.175
No | Visit | Complaint | RIND | pco2 value above reference range. | pco2 control value above reference range. senstivity of co2 was low. | cleaned chamber. run protien removal. shaked co2 electrode. analyzer working fine. | SWO-158463 | Cleaned | Fluidic pathway | Fluidic pathway | pCO2 |  | pco2 value above reference range. | pco2 control value above reference range. senstivity of co2 was low. | cleaned chamber. run protien removal. shaked co2 electrode. analyzer working fine.
No | Visit | Complaint | RDAN | Jan melder at ABL'en ikke kan kalibrere CO2 + O2. | Afprøvet og undersøgt. Lavt tryk i gas 1. Gas 2 flasken er en gas 1! | Ny gas 1+2 monteret, slutafprøvet. | SWO-158133 | User error | Consumables | Gas empty | Gas empty |  | Jan reports that the ABL cannot calibrate CO2 + O2. | Tested and investigated. Low pressure in gas 1. Gas 2 bottle is a gas 1! | New gas 1 + 2 mounted, final tested.
No | Visit | Complaint | RDAN | Marianne Mertz melder at CoHb er skredet efter, at ABL'en var på -R-. | Afprøvet og undersøgt. | Service tHb og alm. tHb kalibreret, slutafprøvet. | SWO-158136 | Adjustment | HW | Oxi module | tHb cal |  | Marianne Mertz reports that CoHb has progressed after the ABL was on-R-. | Tested and investigated. | Service tHb og alm. tHb kalibreret, slutafprøvet.
No | Visit | Complaint | RDAN | Marianne Mertz melder at CoHb er skredet efter, at ABL'en var på -R-. | Afprøvet og undersøgt. | Service tHb og alm. tHb kalibreret, slutafprøvet. | SWO-158135 | Adjustment | HW | Oxi module | tHb cal |  | Marianne Mertz reports that CoHb has progressed after the ABL was on-R-. | Tested and investigated. | Service tHb og alm. tHb kalibreret, slutafprøvet.
No | Visit | Complaint | RDEU | 05.01.2017, H. Sabih, Tel. 02151324630: Leckage erkannt | Einlasssonde verbogen. | Messkammer und Einlassdichtung gereinigt, Verbogene Einlasssonde gerichtet. Gerät mehrfach kalibriert. | SWO-158111 | Defect | HW | Inlet module | Probe bend |  | 05.01.2017, H. Sabih, tel. 02151324630: Leakage detected | Einlasssonde verbogen. | Measuring chamber and inlet seal cleaned, bent inlet probe directed. calibrated several times.
No | 3 | Y | RAME | Z Failed inlet calibration;error code 0675-position calculation error |  | Found inlet motor and friction detector not assembled correctly, loose | D1E66B5D8C3DA3ED | Defect | HW | Inlet module | Friction detector unit | Screws not tightened | Z Failed inlet calibration;error code 0675-position calculation error |  | Found inlet motor and friction detector not assembled correctly, loose
No | 3 | Y | RFRA | Blank | Erreur Température avant séchage |  | D1E6E9988C17B64B | Defect | HW | Process Module | PCB, Cup dryer |  | Blank | Temperature error before drying
Yes | 6 | Y | RAME | P New user bent inlet probe |  | Confirmed reported issue by phone with Dr. Kim | D1EB93A2644C04BE | Re classified RMED |  |  |  |  | P New user bent inlet probe |  | Confirmed reported issue by phone with Dr. Kim
No | 6 | Y | RAME | Z The analyzer keeps saying to replaced the solution pack. |  | Confirmed problem. Preesure test errors. | D1E827DB0219B59A | Defect | HW | Inlet module | Inlet Gasket |  | Z The analyzer keeps saying to replaced the solution pack. |  | Confirmed problem. Preesure test errors.
No | Visit | Complaint | RDEU | Leckage erkannt | Sichtkontrolle | Bluteingabedichtung wurde nicht korrekt angebracht. Dichtung ersetzt, alle test im Service Menü erfolgreich Neustart, Kalibration und QC-Kontrolle nicht erfogreich da Sensor länger nicht konditioniert wurde. Daher Sensor ersetzt. | SWO-158521 | User error | HW | Inlet module | Inlet Gasket |  | Leakage detected | Visual inspection | Blood entry seal was not applied correctly. Seal replaced, all test in the service menu successfully restart, calibration and QC control not achieved because sensor has not been conditioned for longer. Therefore, sensor replaced.
Yes | 2 | Y | RFRA | Blank | Erreur de rendu de la bili |  | D1E7EDD90222AB0B | Re classified RMED | Remote Support |  |  |  | Blank | Error in the rendering of the bili
No | Visit | Complaint | RCHN | LAC定标失败，D/E漂移错误 | LAC定标失败，D/E漂移错误 | 更换测试卡后修复 | SWO-159794 | [5 empty] | LAC calibration failure, D/E drift error | LAC calibration failure, D/E drift error | Repair after replacing test card
No | Visit | Complaint | RCHN | lac定标不过 | lac定标不过 | 更换测试卡修复 | SWO-158325 | [5 empty] | lac calibration is not | lac calibration is not | Replacement test card repair
No | Visit | Complaint | RCHN | PO2定标不过 | PO2定标不过 | 更换测试卡修复 | SWO-158315 | [5 empty] | PO2 calibration is not the | PO2 calibration is not the | Replacement test card repair
No | Visit | Complaint | RCHN | 除PO2分压之外其他参数都定标不过 | 除PO2分压之外其他参数都定标不过 | 更换测试卡修复 | SWO-158284 | [5 empty] | Other parameters except PO2 sub-pressure are calibrated but | Other parameters except PO2 sub-pressure are calibrated but | Replacement test card repair
No | Visit | Complaint | RCHN | CO2灵敏度低，LAC状态值高，灵敏度低，泵定标出错 | CO2灵敏度低，LAC状态值高，灵敏度低，泵定标出错 | 更换测试卡及试剂包后修复 | SWO-160179 | Assisted Replacement | SC | SC | pCO2 | Sensitivity | Low CO2 sensitivity, high LAC status value, low sensitivity, pump calibration error | Low CO2 sensitivity, high LAC status value, low sensitivity, pump calibration error | Repair after replacing test cards and reagent kits
No | Visit | Complaint | RCHN | 不均匀冲洗液 | 不均匀冲洗液 | 更换试剂包修复 | SWO-158349 | Assisted Replacement | SP | SP | FT |  | Uneven flushing fluid | Uneven flushing fluid | Replacement reagent kit repair
No | Visit | Complaint | RCHN | 无法读取芯片数据 | 无法读取芯片数据 | 更换试剂包后修复 | SWO-158222 | Assisted Replacement | SP | SP | Chip |  | Can't read chip data | Can't read chip data | Repair after replacement of reagent kit
No | Visit | Complaint | RCHN | 光单上下传感器错误，报错5784,5257 | 光单上下传感器错误，报错5784,5257，借备件636-550一个 | 更换备件636-550修复 | SWO-158218 | Defect | Consumables | Test Kit | Cartridge | 5784 (Broken) | Light single up and down sensor error, reported error 5784,5257 | Light single up and down sensor error, reported error 5784,5257, borrowed spare parts 636-550 one | Replacement spare parts 636-550 repair
No | Visit | Complaint | RIND | PO2 status shows yellow some time. | Check and found PO2 status shows yellow and it's sensitivity is higher than the limit | Check PO2 electrode and clean with brush and clean all the chamber,clean inlet gasket and waste line and run Protein removal and 2-point calibration which is Ok.Run QC which is OK and PO2 status is also OK and problem has been resolved. | SWO-158470 | Cleaned | Fluidic pathway | Fluidic pathway | pO2 |  | PO2 status shows yellow some time. | Check and found PO2 status shows yellow and it's sensitivity is higher than the limit | Check PO2 electrode and clean with brush and clean all the chamber,clean inlet gasket and waste line and run Protein removal and 2-point calibration which is Ok.Run QC which is OK and PO2 status is also OK and problem has been resolved.
No | Visit | Complaint | RCHN | 打印机错误 | 打印机错误，借备件910-275一个 | 更换910-275修复 by jesson | SWO-158278 | Defect | HW | Printer module | Printer module |  | Printer error | Printer error, borrowed spare parts 910-275 one | Replace 910-275 Repair by jesson
No | Visit | Complaint | RDEU | 06.01.2017, Herr Tauber, Tel.: 05281992212: Kalibrationsfehler bei Ca 06.01.2017, 11:28 Uhr ==> Anrufer Her Tauber, Tel.: 05281992212 ==> Kalibrierfehler Ca besteht weiterhin (Kal.-staus 26 mmol/L) pCO2 Empf. außer Bereich | pCO2, K, pH Staus, Drift und Empf. Fehler Bei Ref.- und Ca-Elektrode ist die Elektrodenspitze blank poliert (Silberchlorid-Beschichtung fehlt) | Verstopften Belüftungsstutzen gespült Ref.- und Ca- Leihelektroden eingebaut Dichtring der pH-Elektrode ersetzt Probelauf | SWO-158450 | Assisted Replacement | Electrode | Membrane | Reference |  | 06.01.2017, Mr. Tauber, Tel.: 05281992212: Calibration error at Ca 06.01.2017, 11:28 ==> Caller Her Tauber, Tel.: 05281992212 ==> Calibration error Ca persists (Cal.-staus 26 mmol/L) pCO2 Mpf. out of range | pCO2, K, pH congestion, drift and empf. Error For Ref. and Ca electrode, the electrode tip is polished (silver chloride coating is missing) | Clogged ventilation nozzle flushed Ref.- and Ca- rental electrodes installed Sealing ring of the pH electrode replaced Rehearsal
No | 7 | Y | RAME | V Inlet cal / position errors. Sample type is defaulting to proficiency test. |  | 0676 inlet sensitivity out of range errors in activity log. | D1E7BE47022BE682 | Defect | HW | Inlet module | Friction detector unit |  | V Inlet cal / position errors. Sample type is defaulting to proficiency test. |  | 0676 inlet sensitivity out of range errors in activity log.
No | Visit | Complaint | RDEU | pCO2 Kal Empf. im laufenden Betrieb konstant weit unterhalb von 85% | s. case | Fehlersuche: Keine Gerinnsel o.Ä. feststellbar; pCO2 Elektrode mit Cleaner gereinigt und vorerst mit der selben Hülse membraniert. Mehrere Kal, Empf. abfallend von 92% Richtung Untergrenze => Elektrode ersetzt, Initialisierung, Kal und QC Messung, OK | SWO-158383 | Assisted Replacement | Electrode | Electrode | pCO2 |  | pCO2 Kal Empf. in operation constantly well below 85% | s. case | Troubleshooting: No clots or similar can be detected; pCO2 electrode cleaned with cleaner and membraned with the same sleeve for the time being. Multiple Kal, Empf. sloping from 92% towards the lower limit => electrode replaced, initialization, cal and QC measurement, OK
No | Visit | Complaint | RKK | Bil QC Lv3 測定値が範囲外 | Bil QC Lv3 測定値が範囲外 tHbCAL行うが解消せず | ヘモライザーユニット交換 OXIコンプリートテスト確認 コンタミ除去・エアライン・測定チャンバー清掃 Glu・Lac・Refメンブラン交換 CAL・tHbCAL・StHbCAL・QC測定データ確認 | SWO-158453 | Defect | HW | Oxi module | Hemolyzer |  | Bil QC Lv3 Measurements out of range | Bil QC Lv3 Measurements out of range tHbCAL do but do not eliminate | Hemolyzer unit replacement OXI Complete Test Confirmation Contamination removal, airline measurement chamber cleaning Glu, Lac, Ref, Membrane Replacement Cal, tHbCAL, StHbCAL, QC measurement data confirmation
Yes | 2 | Y | RPAC | Blank |  |  | D1E7BCA6024271C9 | Re classified RMED | Delivery |  |  |  | Blank
No | Visit | Complaint | RDEU | Inlet undicht | Inlet undicht, pO2 Nullpunkt Fehler | Inletmodul demontiert und geprüft, Abfallweg geprüft, Ersatzmodul installiert, Nullpunkt Kalibration durchgeführt, Funktionskontrolle ok | SWO-158392 | Defect | HW | Inlet module | Inlet module |  | Inlet leaks | Inlet leak, pO2 zero point error | Inlet module dismantled and tested, waste path tested, replacement module installed, zero calibration performed, function control ok
No | Visit | Complaint | RIND | SC temperature not stable. | analyzer displaying SC temperature not stable. | removed casssette. reinstallled cassette and solution pack. cassette has installed successfully. | SWO-158460 | [5 empty] | SC temperature not stable. | analyzer displaying SC temperature not stable. | removed casssette. reinstallled cassette and solution pack. cassette has installed successfully.
No | Visit | Complaint | RIND | Touch screen was not working. | Touch screen was not working | remove the touch panel and cleaned it properly,then upgraded software but still the touch screen was not working. | SWO-160079 | [5 empty] | Touch screen was not working. | Touch screen was not working | remove the touch panel and cleaned it properly,then upgraded software but still the touch screen was not working.
Yes | Visit | Complaint | RIND | Touch screen was not working. | Touch screen was not working | replace the touch panel with the new one and it was working properly | SWO-160080 | Re classified RMED | Double entry | same as day before |  |  | Touch screen was not working. | Touch screen was not working | replace the touch panel with the new one and it was working properly
No | Visit | Complaint | RHUN | Tűtartó, oldatcserét kér a készülék, vasárnap éjszaka óta nincs Astrup Nem állt helyre, szervízkiszállást kérnek | Készülék nem méréskész,többször tűtartó cserét kér | Visula Flow Check rendben, Tűtartó kimosása, Felszívótű ellenőrzése, Oldatpakk újrainstallálása, Szervíztesztek, Indítás, Kalibrációk, QC mérés | SWO-158694 | Assisted Replacement | SP | SP | FT |  | Needle holder, Request for Solution Exchange, no Astrup since Sunday night Not recovered, service is requested | The appliance is not ready to be measured, repeatedly requesting | The Visula Flow Check is OK, you can wash the needle holder, check the suction needle, reinstall Solution Pack, service tests, start, calibration, QC measurement
No | 6 | Y | RAME | P New User bent sample inlet probe |  | Confirmed reported issue by visual inspection. | D1EB9528644F04C0 | Replaced | HW | Inlet module | Probe bend |  | P New User bent sample inlet probe |  | Confirmed reported issue by visual inspection.
No | Visit | Complaint | RDEU | 09.01.2017, Fr. Weimer, Tel. 09117580992753: Leckage erkannt | Fehlersuche --> Einlassnadel und -dichtung und Lösungspack überprüft --> Alles i.O.; Analysator im Service überprüft (Programme des Fluid Transport) | Analysator im Service überprüft (Adjust Manifold, Pressure Test, Clot detect $remove, Pumpenkal., SP Check, LS Justage), Startphase, System und Funktion i.O. | SWO-158678 | Adjustment | SW | Service pgm | Manifold |  | 09.01.2017, Ms. Weimer, Tel. 09117580992753: Leakage detected | Troubleshooting --> Inlet needle and seal and solution pack checked --> Everything in the art; Analyzer checked in service (fluid transport programs) | Analyzer checked in service (Adjust Manifold, Pressure Test, Clot detect $remove, Pump Scale, SP Check, LS Adjustment), Start Phase, System and Function i.O.
No | Visit | Complaint | RDEU | Glu. und Lac. Kal. Fehler | Kal. Daten analysiert . Glu und Lac Status zu hoch | Sensorkassette und Lösungspack erneuert. Defekte SC und LP wird reklamiert. | SWO-159045 | Assisted Replacement | SC | SC | Glu/Lac |  | Glu. and Lac. Kal. Error | Kal. Data analyzed . Glu and Lac status too high | Sensor cassette and solution pack. Defective SC and LP are complained about.
No | Visit | Complaint | RCHN | 所有参数没有终点，超出响应范围 | 所有参数没有终点，超出响应范围 | 更换测试卡 | SWO-159173 | [5 empty] | All parameters have no end point, out of response range | All parameters have no end point, out of response range | Replace the test card
No | Visit | Complaint | RCHN | CA定标不过 | CA定标不过 | 更换测试卡修复 | SWO-158702 | [5 empty] | CA calibration is not | CA calibration is not | Replacement test card repair
No | Visit | Complaint | RCHN | 参比电极错误 | 参比电极错误 | 更换测试卡修复 | SWO-159801 | Assisted Replacement | SC | SC | Reference |  | Reference electrode error | Reference electrode error | Replacement test card repair
No | Visit | Complaint | RCHN | 进样口问题 | 进样口问题，借备件903-139一个 |  | SWO-158594 | Defect | HW | Inlet module | Inlet module |  | Sample problems | Sample problem, borrow edipart 903-139 one
No | Visit | Complaint | RCHN | CO2定标失败 | CO2定标失败 | 更换测试卡修复 | SWO-158561 | Assisted Replacement | SC | SC | pCO2 | Calibration | CO2 calibration failure | CO2 calibration failure | Replacement test card repair
No | Visit | Complaint | RCHN | 外置条码扫描器不能扫描定标卡条码 | 外置条码扫描器不能扫描定标卡条码,借902-924一个 |  | SWO-158602 | Defect | HW | Analyzer Control Module | Barcode reader, display |  | External barcode scanner cannot scan calibration card barcodes | External barcode scanner can not scan the calibration card bar code, borrow 902-924 one
No | Visit | Complaint | RCHN | 机器报错1054PCT定标失败 | 机器报错1054PCT定标失败，借备件944-394（12323）一个 | 更换备件944-394修复(by tim) | SWO-158549 | Assisted Replacement | Consumables | Test Kit | Cartridge | PCT CAL | Machine misstated 1054PCT calibration failure | Machine misstated 1054PCT calibration failure, borrowed spare parts 944-394 (12323) one | Replacement spare parts 944-394 repair (by tim)
Yes | Visit | Complaint | RCHN | 机器报错1054PCT定标失败 | 定标报1021，定标底物溢出，借备件903-364和903-236各一个 | 更换备件944-394修复 | SWO-158550 | Re classified RMED | Double entry |  |  |  | Machine misstated 1054PCT calibration failure | Bid 1021, bid substrate overflow, spare parts 903-364 and 903-236 each one | Replacement parts 944-394 repair
No | Visit | Complaint | RCHN | 机器报1263，HCT测量出错 | 机器报1263，HCT测量出错，借备件903-329一个 |  | SWO-158626 | Defect | HW | Aspiration Module | Needle Wash Unit |  | Machine report 1263, HCT measurement error | Machine report 1263, HCT measurement error, borrowed spare parts 903-329 one
No | Visit | Complaint | RDEU | 06.01.2017, 10:40 Uhr ==> Anrufer Herr Thiele, Tel.: 036458543434 ==> Probe bei Ansaugung vorzeitig im Oxi-Modul erkannt | Probe bei Ansaugung vorzeitig im Oxi-Modul erkannt LS Oximetriefehler | Fehlersuche Wartung Ventilblock BG und EL div. Pumpenschläuche erneuert, Rollenkränze überprüft, LS Justage, Pumpenkal ,Proteinentfernung tHb kal , Lac membraniert | SWO-158777 | Defect | HW | pH/BG Module | Pump tubes |  | 06.01.2017, 10:40 a.m. ==> Caller Mr. Thiele, Tel.: 036458543434 ==> Sample in suction detected prematurely in the Oxi module | Sample detected prematurely in Oxi module during suction LS Oximetry Error | Troubleshooting Maintenance valve block BG and EL div. Pump hoses renewed, roller wreaths checked, LS Adjustment, Pump Scale ,Protein Removal tHb kal , Lac membraneed
Yes | 10 | Y | RPAC | Blank | Analyser indicating OXI Hardware error at times. |  | D1EB78D764581EC8 | [5 empty] | Blank | Analyser indicating OXI Hardware error at times.
No | Visit | Complaint | RDEU | Probentransportfehler | Einlassnadelhalterung gebrochen | Inletmodul geprüft, Probe Holder ausgetauscht, Funktionskontrolle ok | SWO-159054 | Defect | HW | Inlet module | Probe arm |  | Probentransportfehler | Inlet needle holder broken | Inlet module tested, sample holder replaced, function control ok
No | Visit | Complaint | RCHN | O2测量有问题 | O2测量有问题 | 更换O2电极修复 | SWO-160260 | Assisted Replacement | Electrode | Electrode | pO2 |  | There is a problem with the O2 measurement | There is a problem with the O2 measurement | Replacement of O2 electrode repair
No | Visit | Complaint | RDEU | Schwarzer Bildschirm, keine Aktion möglich Fehler trat bereits mehrfach auf | Schwarzer Bildschirm, keine Aktion möglich Reklamation/ Garantie zu SWO | Datensicherung, SW erneuert (XPE, Applikation), Rücksicherung Daten, Kal u. QC i.O. | SWO-158731 | Defect | HW | DMS_PC | Touch screen |  | Black screen, no action possible Errors have occurred several times | Black screen, no action possible Complaint/guarantee to SWO | Data backup, SW renewed (XPE, application), backup data, Kal & QC i.O.
No | Visit | Complaint | RDEU | Kunde berichtet, dass wieder ein "schwarzer Bilkdschirm" aufgetreten ist und nur mit Neustart behoben werden konnte. Kunde bittet um Rep., da Gerät neu ist (Installation 08/2016) | Häufig schwarzer Bildschirm , keine Aktion möglich | Datensicherung, SW erneuert (XPE, Applikation), Rücksicherung Daten, Kal u. QC i.O. | SWO-158730 | Defect | HW | DMS_PC | Touch screen |  |  | Frequently black screen, no action possible | Data backup, SW renewed (XPE, application), backup data, Kal & QC i.O.
No | Visit | Complaint | RDEU | 09.01.2017, Kundin, Tel.: 09117580992308: Kalibrationsfehler bei Lactat 09.01.2017, Frau Bänkert, Tel.: 09117580992308 : Weiterhin Kalibrationsfehler bei Laktat | Fehlersuche (Messung instabil bei pCO2 und Lac) --> Messkammer, Elektrode überprüft | Reinigung der Messkammern und der Glasmembranfläche, System und Funktion i.O. | SWO-158806 | Assisted Replacement | Electrode | Electrode | Lac |  | 09.01.2017, Customer, Tel.: 09117580992308: Calibration error in lactate 09.01.2017, Mrs. Bänkert, Tel.: 09117580992308 : Further calibration error in lactate | Troubleshooting (measurement unstable for pCO2 and Lac) --> Measuring chamber, electrode checked | Cleaning of the measuring chambers and the glass membrane surface, system and function in the
Yes | Visit | Complaint | RCHN | 打印机卡纸 | 打印机卡纸,借910-275一个 |  | SWO-158630 | Re classified RMED | OOB |  |  |  | Printer paper jam | Printer paper jam, borrow 910-275 one
No | 2 | Y | RAME | Z Customer called with autochack errors and OXI errors |  | Investigated to find 743 - Autochack Breaker error, 247 - Spectrometer A $D converter failed ,269 - Hemolyzer phase regulation failed. | D1EBBFAA643100B0 | Defect | HW | Auto Check Module | Breaker Motor |  | Z Customer called with autochack errors and OXI errors |  | Investigated to find 743 - Autochack Breaker error, 247 - Spectrometer A $D converter failed ,269 - Hemolyzer phase regulation failed.
Yes | Visit | Complaint | RDEU | Fehler in kal ph K Na Ca Cl Lac | Fehlersuche | SC erneuert | SWO-159017 | Re classified RMED | ECKHI Email 24.jan-21017 | SC | Reference |  | Error in kal ph K Na Ca Cl Lac | Troubleshooting | SC renewed
No | Visit | Complaint | RDEU | 10.01.2017 10:50Uhr Frau Benz 062181024322: QC-Verletzung bei PO2 im LevelB. | QC-Verletzung pO2 Level B (tief) | Analysator überprüft (kein Fehler feststellbar) - Erneuerung Lösungspack notwendig - Funktionsüberprüfung etc. ohne Beanstandung ! | SWO-159050 | Assisted Replacement | SP | SP | pO2 |  | Jan 10, 2017 10:50AM Ms Benz 062181024322: QC violation at PO2 in LevelB. | QC violation pO2 Level B (deep) | Analyzer checked (no error detectable) - renewal solution pack necessary - function check etc. without complaint !
No | 3 | Y | RPOL | Blank | Uszkodzony zasilacz i eletronika w analizatorze. Wymiana zasilacza, jednostki komputetra i p³ytki sekcji mokrej. Konoieczna wymiana zasliania akumulatorowego. Analizator sprawny technicznie |  | D1ECDC6CDA19B9AA | Defect | HW | Power Supply | Power Supply |  | Blank | Faulty power supply and eletronics in the analyser. Replacing the PSU, Komputetra unit and the wet section. Battery replacement. Technically efficient analyzer
No | Visit | Complaint | RDEU | 09.01.2017; Fr. Horner; Tel.: 081613993; QC Messfehler alle Level 09.01.2017 16:00 Fr.Horner 081613993: QC-Fehler Level B bei pO2 | QC Fehler pO2 Level B zu niedrig. Kunde hat LP und SC bereits getauscht, Fehler bleibt bestehen. Kein Material mehr vorhanden. STK laut Frau Kostova durchführen (noch kein Protokoll vorhanden) | LP erneuert - pO2 Fehler nicht mehr vorhanden. STK durchgeführt - ok. Nach Hochfahren und Kal Kalibrierfehler Glu vorhanden. SC Tausch und Reinitialisierung brachte keine Besserung. Glu nach Rücksprache mit Kunden deaktiviert. Kunde beobachtet, ob sich Glu über Nacht einläuft. Herr Dr. Herden gibt Mittwoch Rückmeldung an Techniker. Kalibration und QC's vorgenommen(alle Parameter außer Glu in Ordnung) Kunden SC und LP für Lager überlassen. Sollte Fehler noch vorhanden sein wird ein Reparaturersatzgerät geliefert, da E-Teil nicht im Bestand. | SWO-159195 | Assisted Replacement | SP | SP | pO2 |  | 09.01.2017; Fr. Horner; Tel.: 081613993; QC Messfehler alle Level 09.01.2017 16:00 Fr.Horner 081613993: QC-Fehler Level B bei pO2 | QC error pO2 Level B too low. Customer has already exchanged LP and SC, errors remain. No more material available. STK according to Ms. Kostova perform (no protocol yet) | LP renewed - pO2 error no longer exists. STK performed - ok. After ramp-up and kal calibration error Glu present. SC exchange and reinitialization brought no improvement. Glu deactivated after consultation with customers. customer observes if Glu enters overnight. Dr. Herden gives feedback to technicians on Wednesday. Calibration and QC's (all parameters except Glu in order) customers SC and LP for warehouses. If there are still errors, a repair replacement device is delivered, as E-part is not in stock.
No | Visit | Complaint | RDEU | Lösungen können nicht gepumpt werden | Lösungen können nicht gepumpt werden | div. Test`s im Service durchgef.; SC + LP gew.; Startphase abgewartet; Kal. i.O.; QC-Level A + C außer Bereich -> wegen zu kalter Lösung im LP (war im Kühlschrank); weitere Kal.; QC i.O. | SWO-159044 | Assisted Replacement | SP | SP | FT |  | Solutions cannot be pumped | Solutions cannot be pumped | div. Test it in the service. SC + LP. start-up phase awaited; Kal. i.O.; QC-Level A + C out of range -> because of too cold solution in the LP (was in the refrigerator); more Kal.; QC in the first
No | Visit | Complaint | RSAF | Inlet probe bent. | Checked data logs and inspected all hardware. | Replaced inlet gasket and inlet probe. | SWO-159248 | Replaced | HW | Auto Inlet module | Probe bend |  | Inlet probe bent. | Checked data logs and inspected all hardware. | Replaced inlet gasket and inlet probe.
No | 6 | Y | RPAC | Blank | Sampling issues occuring on the analyser. |  | D1ECEA8CDA58DE3B | Defect | HW | Auto Inlet module | Inlet Gasket |  | Blank | Sampling issues occuring on the analyser.
No | 10 | Y | RAME | Z Customer found a leak in the analyzer. |  | Found that the sensor cassette probe was clotted and potentially the pumps. | D1ECEC4EDA2E92C0 | [5 empty] | Z Customer found a leak in the analyzer. |  | Found that the sensor cassette probe was clotted and potentially the pumps.
No | Visit | Complaint | RCHN | pco2灵敏度超出范围，血样模块有问题 | pco2灵敏度超出范围，血样模块有问题，借备件902-980一个 |  | SWO-158897 | Defect | HW | Oxi module | Hemolyzer |  | pco2 sensitivity out of range, blood module problem | pco2 sensitivity out of range, blood module problem, spare part 902-980 one
No | Visit | Complaint | RCHN | 打印机打印缺失 | 打印机打印缺失,借910-286一个 |  | SWO-158904 | Defect | HW | Printer module | Printer module |  | Printer printing missing | Printer printing missing, borrow 910-286 one
No | Visit | Complaint | RDEU | Fehler taucht immer wieder auf trotz Ersatz des ElMet Pumpenschlauchs | S. Case | Beim Eintreffen nach Pumpenkal o.g. Fehler. ElMet Modul ausgebaut und beide Ventile sowie Pumpe überprüft, keine sichtbaren Fehler feststellbar. Beide Ventile ersetzt, weitere Arbeiten im Rahmen de HJW (s. anderer SWO). | SWO-159078 | Defect | HW | El/Met Module | Pump tubes |  | Error keeps popping up despite replacement of ElMet pump hose | S. Case | On arrival after pump scale above o.g. error. ElMet module removed and both valves and pump checked, no visible errors detectable. Both valves replaced, further work within the framework of HJW (see other SWO).
No | Visit | Complaint | RDEU | 10.01.2017 Frau Domberg 05115322529 meldet, dass ein Gerät eine Leckage erkennt. 10.01.2017 Frau Domberg 05115322529 wünscht einen Vorortservice | Undichtigkeit im Bereich des Elyte Modul's | Gerät demontiert und gereinigt, Elektroden geprüft, Funktionskontrolle ok | SWO-159107 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | 10.01.2017 Ms. Domberg 05115322529 reports that a device detects a leakage. 10.01.2017 Mrs Domberg 05115322529 wants a suburban service | Leakinence in the area of the Elyte module | Device dismantled and cleaned, electrodes tested, function control ok
No | 2 | Y | RAME | Z Cannot power on the analyzer. |  | Unplug and connect to another power outlet and still same issue. Primary fuses in power supply unit keeps blowing. | D1ECE863DA04FFD3 | Defect | HW | Power supply | Power supply |  | Z Cannot power on the analyzer. |  | Unplug and connect to another power outlet and still same issue. Primary fuses in power supply unit keeps blowing.
No | Visit | Complaint | RLTD | lauren spratling - 07788 27 60 94 | Customer reported, solution pack not accessible (currently propped up with tape). Analyser not scanning and accepting new test cartridges. Inventory doesn't show all cartridges in analyser. | Opened analyser and removed a large amount of dust from around the barcode scanner, removed tape from under solution pack and removed and refitted, no issues found. Cleaned around solution pack enclosure. Removed cartridges and tested scanner, all seems ok. Error rotate storage wheel error on data logs, operated all functions in service, no issues found. Customer ran sample with no issues. | SWO-159156 | Cleaned | HW | Storage Module | Barcode reader, Storage wheel |  | Lauren spratling - 07788 27 60 94 | Customer reported, solution pack not accessible (currently propped up with tape). Analyser not scanning and accepting new test cartridges. Inventory doesn't show all cartridges in analyser. | Opened analyser and removed a large amount of dust from around the barcode scanner, removed tape from under solution pack and removed and refitted, no issues found. Cleaned around solution pack enclosure. Removed cartridges and tested scanner, all seems ok. Error rotate storage wheel error on data logs, operated all functions in service, no issues found. Customer ran sample with no issues.
No | Visit | Complaint | RDEU | Das forderft das erneut Einsetzen des Solutions packs | Rinse Fail im Solutions pack, Solution pack Check mehrmals durchgeführt und Visual Flow, das Gerät hat nicht gepumpt . | Erneuerung der Sonde, Einlassdichtungseinheit, Solution pack | SWO-159465 | Assisted Replacement | SP | SP | FT | 1176 | This requires the re-insertion of the Solutions pack | Rinse Fail in the Solutions pack, Solution pack Check performed several times and Visual Flow, the device has not pumped. | Renewal of the probe, inlet seal unit, solution pack
No | Visit | Complaint | RDEU | Kund bemängelt lange StartupTime bei Lac | lange Startup Time bei Lac nach SC wechsle am 04.01.17 | SC wurde am 04.01.17 06:49Uhr ersetzt, Lac QC Level A fehlerfrei 04.01.17 23:04Uhr, Datalogs & SD erstellt und an TS geschickt | SWO-159383 | Assisted Replacement | SC | SC | Lac |  | Customer criticizes long StartupTime at Lac | long Startup Time at Lac to SC change on 04.01.17 | SC was replaced on 04.01.17 06:49, Lac QC Level A error-free 04.01.17 23:04, Datalogs & SD created and sent to TS
No | Visit | Complaint | RDEU | Kunde berichte von sporadischen Ca Sensor Ansprache Fehler bei Patientenmessung | Kunde bemängelt das sporadische Auftreten von Ca Sensoransprachefehler bei der kapillaren Patientenmessung, QC & Kal sind fehlerfrei | Kunde wird die SC am 14.01.17 (fällig) ersetzten und sich ggf. melden wenn der Fehler weiter besteht, | SWO-159380 | Assisted Replacement | SC | Sc | Ca++ | LEH mail | Customer reports of sporadic Ca sensor response error in patient measurement | Customer criticizes the sporadic occurrence of Ca sensor contact error in capillary patient measurement, QC & Kal are error-free | Customer will replace the SC on 14.01.17 (due) and, if necessary, report it if the error persists,
No | Visit | Complaint | RDEU | Kunde bemängelt Glu Startup Time > 4h | Glu Startup Time größer 4 h | SC wurde am 27.12.16 14:08 Uhr ersetzt, Glu QC Level A fehlerfrei am 27.12.16 20:51 Uhr, datalogs & SD erstellt und an TS geschickt | SWO-159442 | Assisted Replacement | SC | SC | Glu |  | Customer criticizes Glu Startup Time > 4h | Glu Startup Time greater 4 h | SC was replaced on 27.12.16 14:08, Glu QC Level A error-free on 27.12.16 20:51, datalogs & SD created and sent to TS
No | Visit | Complaint | RDEU | SC Startup Time größer als 4h | Kunde bemängelt eine Startup Time >4h bei Lac | SC wurde am 03.01.17 7:24Uhr ersetzt, Lac Level A fehlerfrei am 04.01.17 07:15Uhr, Datalogs und SD erstellt und an TS geschickt | SWO-159430 | Assisted Replacement | SC | SC | Lac |  | SC Startup Time larger than 4h | Customer complains of a startup time >4h at Lac | SC was replaced on 03.01.17 7:24o, Lac Level A error-free on 04.01.17 07:15, Datalogs and SD created and sent to TS
No | 5 | Y | RPNZ | Blank | Sensor cassette replacement due to internal reference electrode error. |  | D1EE242D5045F2DA | Assisted Replacement | SC | SC | Reference |  | Blank | Sensor cassette replacement due to internal reference electrode error.
No | Visit | Complaint | RDEU | Kunde bemängelt eine Lac Startup Time von größer 4h | Lac Startup Time >4h, bei Eintreffen Lac QC werte im Level A & C außer Bereich | SC wurde am 11.1.17 07.30Uhr ersetzt, bei Anwesenheit Lac QC Werte in Level A & C fehlerhaft, Gerät mehrfach kalibriert, QC wiederholt, Lac QC Wert in Level A dann i.O., in Level C wert auf oberer Grenze ->info an Med-Technik, Datalogs & SD erstellt und an TS geschickt | SWO-159438 | Assisted Replacement | SC | SC | Lac |  | Customer criticizes a Lac Startup Time of greater 4h | Lac Startup Time >4h, on arrival Lac QC values in level A & C out of range | SC was replaced on 11.1.17 07.30o, in presence Lac QC values in level A & C faulty, device repeatedly calibrated, QC repeated, Lac QC value in level A then i.O., in level C value on upper limit ->info created to Med-Technique, Datalogs & SD and sent to TS
No | Visit | Complaint | RDEU | Kunde bemängelt eine Lac Startup time von größer 4h | Lac Startup Time >4h | SC wurde am 10.01.17 13.20 ersetzt, Lac QC Level A fehlerfrei 10.01.17 19:39 Uhr, datalogs & SD erstellt und an TS geschickt | SWO-159435 | Assisted Replacement | SC | SC | Lac |  | Customer complains of a Lac Startup time of greater 4h | Lac Startup Time >4h | SC was replaced on 10.01.17 13.20, Lac QC Level A error-free 10.01.17 19:39, datalogs & SD created and sent to TS
No | Visit | Complaint | RCHN | 接口板烧掉了 | 接口板烧掉了，借备件902-972一个 |  | SWO-159225 | [5 empty] | The interface board's burned. | The interface board was burned, and the spare parts 902-972 were one
No | Visit | Complaint | RCHN | 温度错误 | 温度错误，借备件902-972一个 |  | SWO-159186 | [5 empty] | Temperature error | Temperature error, borrowed spare parts 902-972 one
No | Visit | Complaint | RIND | Data Base Full. | Data Base Full. MSG coming to remove data Base. | All data Base download and deleted. Analyzer Serviced and calibrated properly. Handed over in a fine working condition. | SWO-159548 | [5 empty] | Data Base Full. | Data Base Full. MSG coming to remove data Base. | All data Base download and deleted. Analyzer Serviced and calibrated properly. Handed over in a fine working condition.
No | Visit | Complaint | RCHN | 报5769和5785错误 | 报5769和5785错误 | 更换备用机上的580-421后修复 | SWO-160155 | Defect | HW | Storage Module | Stepper Motor, punch out |  | reported 5769 and 5785 errors | reported 5769 and 5785 errors | Repair after replacing 580-421 on the standby
No | 13 | Y | RPOL | Blank | El pO2 - wymiana |  | D1F9581E763D7799 | Assisted Replacement | Electrode | Electrode | pO2 |  | Blank | El pO2-Replacement
No | Visit | Complaint | RDEU | Seit Installation des Analysators gibt es konstant Schwankungen hoch und runter im pCO2 Ergebnis L3 | Weiterhin zickzack Linie bei L3 pCO2 Ergebnissen | Erneut einen Datenextrakt an RDEU Service gesendet und die Elektrode ersetzt. Gerät komplett überprüft, sichtkontrolliert und Fließverhalten getestet, keine Unregelmäßigkeiten feststellbar. Kalibrationen, Funktionstest, per Remote eine L3 Messung ausgelöst, OK | SWO-159318 | Assisted Replacement | Electrode | Electrode | pCO2 | Grey zone | Since the installation of the analyzer, there are constant fluctuations up and down in the pCO2 result L3 | Continue zigzag line at L3 pCO2 results | Once again, a data extract is sent to RDEU Service and the electrode is replaced. Device completely checked, visually controlled and flow behavior tested, no irregularities detectable. Calibrations, function test, remotely triggered an L3 measurement, OK
No | Visit | Complaint | RDEU | Kein Neustart möglich, Meldung "Einlassklappen entfernt" obwohl beide montiert sind | s. case | Neustarttaste nicht anwählbar beim Eintreffen, obwohl beide Klappen richtig sitzen. Klappen bewegt, ohne Änderung der Situation. Klappenposition im Servicemenü überprüft, alle Varianten werden korrekt angezeigt. Neustart nun möglich aber mit abnormalem Startbild => Kaltstart ausgelöst, Startverfahren und Initialisierung. Bitte morgen erneut beobachten! | SWO-159426 | Unknown | HW | Inlet module | PCB, FLAP | Revisit required?? |  | s. case | Restart button cannot be selected on arrival, even though both flaps are seated correctly. flaps moved without changing the situation. Flap position checked in the service menu, all variants are displayed correctly. Restart now possible but with abnormal start screen => Cold start triggered, start procedure and initialization. Please watch again tomorrow!
Yes | 2 | Y | RPAC | Blank |  |  | D1EE09A6504F3B4F | Re classified RMED | Installation |  |  |  | Blank
No | 6 | Y | RAME | P Instrument leaking Blood from blank bottle next to rinse solution, perform tHb cal |  | checked data logs for errors, checked flow of waste solutions, checked for clotts, | D1EE9D295061FCA9 | Defect | Hw | Basic Unit | Waste tubing |  | P Instrument leaking Blood from blank bottle next to rinse solution, perform tHb cal |  | checked data logs for errors, checked flow of waste solutions, checked for clotts,
No | Visit | Complaint | RDEU | 09.01.2017/14:00, Herr Esche: Lac. Kal. Fehler nach Membranierung 10.01.2017, 8.15 Uhr Herr Esche, Tel.: 089857915954 : Weiterhin Kalibrationsfehler bei Laktat | Fehlersuche vor Ort | Nullpunktkalibrierung durchgeführt. Lac Elektrode neu membraniert. Mehrere Kalibrierungen durchgeführt. QCs für alle Level durchgeführt. ABL wieder Messbereit. | SWO-159339 | Adjustment | HW | El/Met Module | Lac | Zero cal/adj | 09.01.2017/14:00, Mr Esche: Lac. Kal. Errors after membraneing 10.01.2017, 8.15 a.m. Mr. Esche, Tel.: 089857915954 : Further calibration error in lactate | Troubleshooting on site | zero-point calibration. Lac electrode re-membraneed. Multiple calibrations performed. QCs for all levels. ABL ready to measure again.
No | Visit | Complaint | RCHN | 打印机卡纸 | 打印机卡纸，借备件910-275 |  | SWO-159204 | Defect | HW | Printer Module | Printer Module |  | Printer paper jam | Printer paper jam, spare parts 910-275
No | Visit | Complaint | RIND | Customer complaint that no print is coming after Performing the Sample. | Checked and found that the printer not working | reset the USB Cable, now working fine | SWO-159862 | Defect | SW | Analyzer Control Module | Off/On | Printer Module | Customer complaint that no print is coming after Performing the Sample. | Checked and found that the printer not working | reset the USB Cable, now working fine
Yes | Visit | Complaint | RDEU | Das Gerät meldet verdacht auf Gerinnsel in Sensorkassette und dass der Solution pack nicht ordentlich installiert wurde | Das gerät meldete laut Analysator, dass der Solution pack Lot :. XT-50 nicht ordentlich installiert wurde, obwohl die letzte Änderung am 31/12/16 statt gefunden hat. Meldete auch Gerinnsel in Sensorkassette | Erneuerung des Solutions packs gegen konditionierte Ware lot:.XV-01 | SWO-159766 | Re classified RMED | ECKHI Email 24.jan-21017 | SP | Chip |  | The device reports suspected clots in sensor cartridge and that the solution pack has not been installed properly | According to the analyzer, the device reported that the solution pack Lot:. XT-50 was not installed properly, although the last change took place on 31/12/16. Also reported clots in sensor cassette | Renewal of the Solution pack against conditioned goods lot:. XV-01
No | 25 | Y | RAME | 3 The pO2 measures low the first time external QC Level 1 is measured. |  | The analyzer works better than originally, but the pO2 still measures low on Level 1 external QC the first time it is measured. | D1EF5916C6236C18 | [5 empty] | 3 The pO2 measures low the first time external QC Level 1 is measured. |  | The analyzer works better than originally, but the pO2 still measures low on Level 1 external QC the first time it is measured.
No | Visit | Complaint | RCHN | LAC定标失败 | LAC定标失败 | 更换测试卡 | SWO-159521 | Assisted Replacement | SC | SC | Lac |  | LAC calibration failure | LAC calibration failure | Replace the test card
No | Visit | Complaint | RCHN | LAC定标不过 | LAC定标不过 | 更换测试卡后修复 | SWO-161720 | Assisted Replacement | SC | SC | Lac |  | LAC calibration is not | LAC calibration is not | Repair after replacing test card
No | 5 | Y | RPAC | Blank | Waste Bottle has overflowed twice, waste detecter does not appear to be functioning correctly. |  | D1EF50E6C60FF059 | Defect | HW | Basic Unit | Waste detector |  | Blank | Waste Bottle has overflowed twice, waste detecter does not appear to be functioning correctly.
No | Visit | Complaint | RDEU | 06.01.2017, Frau Horstmann, Tel.: AQC Fehler pCO2 Level 3 | PCo2 in QC Level 2 und 3 immer änder unteren Grenze bzw. zeitweise zu niedrig | E-Lyte Modul eingesetzt und Gerät entsprechend konfiguriert, Kal und QC durchgeführt | SWO-159753 | Assisted Replacement | Electrode | Electrode | pCO2 | 8X0 / El/Met installed | 06.01.2017, Mrs. Horstmann, Tel.: AQC Error pCO2 Level 3 | PCo2 in QC Level 2 and 3 always change lower limit or at times too low | E-Lyte module inserted and device configured accordingly, Kal and QC performed
Yes | 7 | Y | RFRA | Blank |  |  | D1EFA6D8C605CA40 | Re classified RMED | Collect data |  |  |  | Blank
No | Visit | Complaint | RDEU | 12.01.2017, Pfl. Roland, Tel.: 021663942114: Kalibrationsfehler bei pCO2 12.01.2017 17:30 Pfl. Roland 021663942114: nach Kal jetzt QC Fehler | QC Lvl 1 pCO2 Fehler (Messwert zu hoch: 29,8mmHg (25,1-28,9)) und QC Lvl 3 pH und pCO2 Fehler (pH 7,463 (7,470-7,510) pCO2 22,0 mmHg (18,0-21,0)). Fehler bestehen seit Erneuerung des LP. Anfang Januar Glu Fehlerhaft in der QC. | Nach Erneuerung des LP (lot YB-50) besteht ein Problem in der QC Lvl 1 und Lvl 3 mit pH und pCO2. LP (lot YB-50) ausgetauscht - genau das gleiche Fehlerbild. LP mit einer anderen lot Nummer aus dem Kreißsaal organisiert - Fehler behoben. Der Kunde hat eine Lieferung mit 4 LP bekommen alle mit der gleichen Lot. Die verbleibenden 2 LP nehme ich mit da diese vermutlich das gleiche Bild zeigen werden. Der Kunde bestellt neue LP. Kalibrierung und QC ok - Ampel steht auf grün. Datensicherung durchgeführt zur Auswertung der QC Probleme Anfang Januar (Ableitstrom erkannt 964 - Kunde tauschte das LP zur Problembehebung. | SWO-160067 | Assisted Replacement | SP | SP | pCO2 |  | 12.01.2017, Pfl. Roland, Tel.: 021663942114: Kalibrationsfehler bei pCO2 12.01.2017 17:30 Pfl. Roland 021663942114: nach Kal jetzt QC Fehler | QC Lvl 1 pCO2 error (measurement too high: 29.8mmHg (25.1-28.9)) and QC Lvl 3 pH and pCO2 error (pH 7.463 (7,470-7,510) pCO2 22.0 mmHg (18.0-21.0)). Errors have existed since renewal of the LP. In early January, Glu erroneous in the QC. | After renewal of the LP (lot YB-50) there is a problem in the QC Lvl 1 and Lvl 3 with pH and pCO2. LP (lot YB-50) - exactly the same error image. LP organized with another lot number from the Kreißsaal - bug fixed. The customer has received a delivery with 4 LP all with the same lot. I take the remaining 2 Lp with me as they will probably show the same picture. The customer orders new LP. Calibration and QC ok - traffic light is on green. Data backup performed to evaluate QC problems at the beginning of January (Leakestrom detected 964 - customer exchanged the LP for troubleshooting.
Yes | 5 | Y | RAME | O tHb and Met Hb near limits of CAP Survey material. |  | Perform inspection and performance check. | D1F0AE9B3C463A45 | Re classified RMED | Courtesy |  |  |  | O tHb and Met Hb near limits of CAP Survey material. |  | Perform inspection and performance check.
No | Visit | Complaint | RDEU | pCO2-QC-Fehler | Medizintechnik (Herr Söllner und Herr Griebel) versuchen schon den ganzen Freitag 13.1.17 den pH und pCO2-QC-Fehler mit Tausch von SC und LP in den Griff zu bekommen --> Fehlersuche --> Überprüfung der Dateiaktivitäten --> Kal.-Ergebnisse i.O. --> QC-Ergebnisse (pH nach unten und pCO2 nach oben raus) --> LP (YB05) verursacht den Fehler --> LP erneuern | LP (XT06) eingebaut und (YB05) ausgebaut, System und Funktion i.O. | SWO-160053 | Assisted Replacement | SP | SP | pCO2 |  | pCO2-QC-Fehler | Medical technology (Mr. Söllner and Mr. Griebel) are already trying to get a grip on the pH and pCO2-QC errors all Friday 13.1.17 by exchanging SC and LP --> Troubleshooting --> Checking file activity --> Cal.-Results usually --> QC results (pH down and pCO2 up) --> LP (YB05) causes the error --> Renew LP | LP (XT06) built-in and (YB05) removed, system and function in the
No | Visit | Complaint | RDEU | 13.01.2017, Frau Zeitler, Tel.: 095150326311: LP wechseln | Fehlersuche --> Dateiaktivitäten überprüft --> QC-Probleme bei pH und pCO2 --> LP - Problem (YB05) --> Spülung des Analysators, Überprüfung im Service | LP (XT06) eingebaut und (YB05) ausgebaut, Clot detect $remove, Spülung des Analysators, System und Funktion i.O. | SWO-160014 | Assisted Replacement | SP | SP | pCO2 |  | 13.01.2017, Mrs. Zeitler, Tel.: 095150326311: Change LP | Troubleshooting --> File Activities Checks --> QC Problems with pH and pCO2 --> LP - Problem (YB05) --> Flushing of the analyzer, verification in the service | LP (XT06) installed and (YB05) removed, Clot detect $remove, flushing of the analyzer, system and function in the order
No | Visit | Complaint | RDEU | Analysator zeigt mehrere Fehlermeldungen. Betriebsmittel (SC + LP) wurden bereits getauscht, ohne Erfolg. | Erneut Intervention erforderlich --> Fehlersuche --> Einlassnadel, Einlassdichtung und Schlauchwege überprüft --> Gerinnsel im System, vermutlich durch Probennehmer oder falsche Handhabung verursacht | Analysator gespült und im Service überprüft (Pressure test, Clot detect $remove), Startphase, System und Funktion i.O. | SWO-160052 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | Analyzer displays several error messages. Equipment (SC + LP) has already been exchanged, to no avail. | Re-intervention required --> Troubleshooting --> Inlet needle, inlet seal and hoseways checked --> clots in the system, probably caused by samplers or incorrect handling | Analyzer rinsed and checked in the service (Pressure test, Clot detect $remove), start phase, system and function in the
Yes | Visit | Complaint | RDEU | 13.01.2017, Frau Grimm, Tel.: 0610391261756: immer wieder SC wechseln |  |  | SWO-160064 | Re classified RMED |  |  |  |  | 13.01.2017, Mrs. Grimm, Tel.: 0610391261756: change SC again and again
No | Visit | Complaint | RLTD | Known psu issues wont switch to battery, replacement required when known good PSUs are available. | PSU unit needs replacing | Replaced PSU. Tested and checked. All working ok | SWO-159936 | Defect | HW | Power Supply | Power Supply |  | Known psu issues wont switch to battery, replacement required when known good PSUs are available. | PSU unit needs replacing | Replaced PSU. Tested and checked. All working ok
No | Visit | Complaint | RCHN | 测试卡漏液导致接口板腐蚀 | 测试卡漏液导致接口板腐蚀，借备件902-972一个 |  | SWO-159843 | [5 empty] | Test card leak causes interface plate corrosion | Test card leakage causes interface plate corrosion, spare part 902-972 one
No | Visit | Complaint | RCHN | SIS/LIS启动选项选不了 | SIS/LIS启动选项选不了,借367-689一个 |  | SWO-159842 | [5 empty] | SIS/LIS startup option not selected | SIS/LIS start option can't be selected, borrow 367-689 one
No | Visit | Complaint | RIND | SP not connected Msg coming again and again and not installing. | SP not connected. | Found SP not setting properly need to push in by external efforts then only smart chip connecting then again open. Solution pack issue as waist also coming out. SP not setting properly and not initialysing when try to reset and install. Need to replace SP. SP replaced and calibrated. analyser working fine. | SWO-160121 | [5 empty] | SP not connected Msg coming again and again and not installing. | SP not connected. | Found SP not setting properly need to push in by external efforts then only smart chip connecting then again open. Solution pack issue as waist also coming out. SP not setting properly and not initialysing when try to reset and install. Need to replace SP. SP replaced and calibrated. analyser working fine.
No | 6 | Y | RAME | The PCO2 would not pass Level 1 of the External QC. |  | Checked the data logs and verified the problem. Ordered and replaced the signal data board and sensor connector board. This did not solve the problem. | D1F0B2C83C0C5AB4 | [5 empty] | The PCO2 would not pass Level 1 of the External QC. |  | Checked the data logs and verified the problem. Ordered and replaced the signal data board and sensor connector board. This did not solve the problem.
No | Visit | Complaint | RCHN | 多参数定标不过 | 多参数定标不过 | 更换测试卡 | SWO-161365 | Assisted Replacement | SC | SC | Reference |  | Multi-parameter calibration but | Multi-parameter calibration but | Replace the test card
No | 28 | Y | RESP | Error en la calibración de la bomba del módulo pH gases. |  | Comprobar bombas y válvulas, cambiarlos los tubos de las bombas, calibrar y probar. | D1F0C2863C13CAC9 | Defect | HW | pH/BG Module | Pump tubes |  | Error in the pump calibration of the pH gas module. |  | Check pumps and valves, replace pump tubes, calibrate and test.
No | 9 | Y | RAME | Z Analyzer has a blank screen. |  | Customer rebooted analyzer several times but did not help. Checked power supply. | D1F088943C054D8E | Defect | HW | DMS_PC | CPU |  | Z Analyzer has a blank screen. |  | Customer rebooted analyzer several times but did not help. Checked power supply.
No | Visit | Complaint | RDEU | 13.01.2017/12:20Uhr Frau Keil 0626183311: Das Gerät meldet seit mehreren Tagen Pumpen-Kal-Fehler 793 | Beim Service zeifte der Analysator keine Fehlermeldung an, war grün und messbereit. Seit Anfang Januar sporadisch pH/BG Pumpen-Asymmetriefehler . Weitere viel häufigere Fehlermeldung Autocheck Carousel failure . Trotz der Meldung werden die Qualitätskontrollen richtig und ohne Fehler durchgeführt. Überprüfung der Rollenkränze der Pumpen, Funktion in Ordnung, keine Behinderung. Überprüfung der LS Werte im Service, keine Auffälligkeit. | Tausch der Pumpenschläuche. Durchführung einer LS Justage und einer Pumpenkalibration. Ein Lösungspumpenschlauch Funktion defekt. Tausch des Flachbandkabels im Autocheck Modul. Funktion nach Inbetriebnahme in Ordnung, Fehlermeldung erscheint nicht mehr unter Aktivitäten. | SWO-160075 | Defect | HW | pH/BG Module | Pump tubes |  | 13.01.2017/12:20pm Woman Wedge 0626183311: The device has been reporting pump-kal error for several days 793 | During the service, the analyzer did not display an error message, was green and ready to measure. Since the beginning of January sporadically pH/BG pump asymmetry error . More much more common error message Autocheck Carousel failure . Despite the notification, the quality controls are carried out correctly and without errors. Checking the roller wreaths of the pumps, function in order, no obstruction. Checking the LS values in the service, no conspicuousness. | Replacement of the pump hoses. Performing an LS adjustment and pump calibration. A solution pump hose function defective. Replacement of the ribbon cable in the Autocheck module. Function in order after commissioning, error message no longer appears under Activities.
No | 2 | Y | RAME | G Customer report Crea not calibrating. |  | Checked data logs. Tried to reset zero calibration. Reading no signal from electrodes. Tried electrode replacement for fault isolation. Order MetII module for replacement. | D1F06D363C51B1D2 | Defect | HW | El/Met Module | Crea |  | G Customer report Crea not calibrating. |  | Checked data logs. Tried to reset zero calibration. Reading no signal from electrodes. Tried electrode replacement for fault isolation. Order MetII module for replacement.
No | Visit | Complaint | RHUN | a készülék szivárgási hibát ír, oldat cserét, és átmosást kér | a készülékben lévő oldat nincs teljesen rögzítve, az oldat cserét az eredetileg használt oldat elfogyása miatt jelezte a készülék, oldat csere közben 1134-es hibakód jelezve (oldat chip írása/olvasása során hiba történt) | oldat csere, készülék újra indítása, csere készülék beüzemelése, kalibrációk, QC mérés | SWO-160119 | User error | SP | SP Activation | SP Activation |  | The device writes a leak defect, a solution replacement, and is asked to | The solution in the device is not fully attached, The solution was changed due to the exhaustion of the solution originally used, with error code 1134 when replacing the solution (solution chip read/write error occurred) | replacement, device re-starting, Installation of the replacement apparatus, calibrations, QC measurement
No | Visit | Complaint | RCHN | 测试卡报错：全部参数定标失败；试剂包报错：PH K NA CA CL定标错误 | 测试卡报错：全部参数定标失败；试剂包报错：PH K NA CA CL定标错误 | 连续更换一张测试卡和一个试剂包 | SWO-160156 | [5 empty] | Test card error: All parameter calibration failed; | Test card error: All parameter calibration failed; | Replace one test card and one reagent kit in a row
No | Visit | Complaint | RIND | customer complaint for sample aborted | error no-1999 and sample aborted problem. found cup stuck on the incubator wheel. also found blockage in the wash block | removed the cup from incubator wheel and cleaned the wash block and cleaned all tubing, now working fine. | SWO-160111 | Defect | Consumables | SP | Solution Pack | 1999 | customer complaint for sample aborted | error no-1999 and sample aborted problem. found cup stuck on the incubator wheel. also found blockage in the wash block | removed the cup from incubator wheel and cleaned the wash block and cleaned all tubing, now working fine.
No | Visit | Complaint | RDEU | Nach dem Austausch der Lösungspack wurde Leckage erkannt. | Frau Heidrich hat bereits die Einlassdichtung mit Halter, Lösungspack und Sensorkassette ausgetauscht. | Fehlersuchverfahren in Serviceprogramm. Oxi-Schlauchventil ausgetauscht. Pumpenkal und LS- Justage. Neutart. | SWO-160210 | Defect | HW | Oxi module | Pinch valve |  | After replacing the solution pack, leakage was detected. | Ms. Heidrich has already replaced the inlet seal with holder, solution pack and sensor cassette. | Error search procedure in service program. Oxi hose valve. Pump scale and LS adjustment. Neutart.
No | Visit | Complaint | RCHN | lac敏感数值偏低 | lac敏感数值偏低 | 更换测试卡修复 | SWO-160219 | [5 empty] | Low lac sensitivity value | Low lac sensitivity value | Replacement test card repair
No | 8 | Y | RAME | Z Printer not printing. |  | Checked printer - no problem found. | D1F55D4914121CC8 | No fault found | No fault found | No fault found | No fault found |  | Z Printer not printing. |  | Checked printer - no problem found.
No | Visit | Complaint | RDEU | 16.01.2017/13:40Uhr Herr Wolf 06213832465: Das Gerät meldet pO2 und pCO2-Fehler | Analysator zeigt bei Messungen mit pH-Wert von etwa 7,1 keine pCO2(T) und pO2(T) -Werte Kalibration und QC ohne Beanstandung ! | Analysator überprüft (kein Fehler feststellbar) - evt. Rücksprache bei neuen Erkenntnissen | SWO-160388 | No fault found | No fault found | No fault found | pCO2 |  | 16.01.2017/13:40H Mr Wolf 06213832465: The device reports pO2 and pCO2 error | Analyzer shows no pCO2(T) and pO2(T) values for measurements with pH of about 7.1 Calibration and QC without complaint ! | Analyzer checked (no error detectable) - evt. Consultation with new findings
No | Visit | Complaint | RDEU | pH und pCO2 QC-Fehler, trotz Erneuerung des SP | Besorgung von Lösungspacks, bei verschiedenen Kunden, da keine mehr im Besitz der GS19 --> Fahrt zum Klinikum A, danach Fahrt zum Klinikum B, danach Fahrt zum Klinikum am Bruderwald, Bamberg Fehlersuche --> Lösungspack überprüft --> Lösungspack verursacht die QC-Probleme bei pH und pCO2 (SP Lot YB50) | Lösungspack (XT06) eingebaut und (YB50) ausgebaut, System und Funktion i.O. --> Datalogs und SD gezogen | SWO-160386 | Assisted Replacement | SP | SP | pCO2 |  | pH and pCO2 QC errors, despite renewal of SP | Procurement of solution packs, from different customers, as no longer owned by GS19 --> Drive to Klinikum A, then drive to Klinikum B, then drive to the Klinikum am Bruderwald, Bamberg Troubleshooting --> Solution Pack Checks --> Solution Pack Causes QC Problems with pH and pCO2 (SP Lot YB50) | Solution pack (XT06) built-in and (YB50) removed, system and function in the --> Datalogs and SD pulled
Yes | Visit | Complaint | RSAF | Nonhlanhla phoned me with regards to the Abl90 Flex plus not switching on. I advised customer to bypass the UPS and see if analyzer starts up. Customer came back to me and informed me that the UPS is the reason why the analyzer does not start up. | Tested UPS in Unit and not switching on. | Took UPS unit into workshop for further investigations. | SWO-160308 | Re classified RMED | UPS fault |  |  |  | Nonhlanhla phoned me with regards to the Abl90 Flex plus not switching on. I advised customer to bypass the UPS and see if analyzer starts up. Customer came back to me and informed me that the UPS is the reason why the analyzer does not start up. | Tested UPS in Unit and not switching on. | Took UPS unit into workshop for further investigations.
No | Visit | Complaint | RDEU | 15.01.2017/16:30Uhr Floian Rüttelross auf Mailbox 0062415015777: Dies ist der tritte Anruf. Alle Wartungsmaßnahmen waren erfolglos 16.01.2017/07:50Uhr Rückruf der Station: Frau Frenzel 0062415016015. Das Problem besteht weiterhin, ein Servcie ist notwendig. | QC Verletzungen pH Level C (leicht zu tief) pCO2 Level A und C RiliBÄK-Verletzung (erhöht) Kalibration ohne Beanstandung | Anwender hat bereits zweimal SP erneuert (LOT YB-50) - Analysator gespült und überprüft (kein Fehler feststellbar) - Erneuerung Anschlussdichtung - nochmalige Erneuerung SP (LOT XV05) - Analysator zeigt Kalibrierfehler pCO2 (Kalibrierdrift 2 außer Toleranz) - mehrfache Kalibration durchgeführt und QC gestartet - Analysator ohne Beanstandung ! | SWO-160349 | Assisted Replacement | SP | SP | pCO2 |  | 15.01.2017/16:30h Floian Rüttelross on Mailbox 0062415015777: This is the kicked call. All maintenance measures were unsuccessful 16.01.2017/07:50h Call back of the station: Ms. Frenzel 0062415016015. The problem persists, a Servcie is necessary. | QC Violations pH Level C (slightly too deep) pCO2 Level A and C RiliBÄK Violation (increased) Calibration without complaint | User has already renewed SP twice (LOT YB-50) - Analyzer flushed and checked (no error detectable) - Renewal connection seal - Renewed renewal SP (LOT XV05) - Analyzer shows calibration error pCO2 (calibration drift 2 except tolerance) - multiple Calibration performed and QC started - analyzer without complaint !
No | Visit | Complaint | RCHN | 液体无法循环 | 液体无法循环,借备件903-099一个 |  | SWO-160150 | [5 empty] | The liquid cannot be circulated. | Liquid cannot be recycled, spare parts 903-099 one
No | Visit | Complaint | RKK | 温度エラー頻発の為、装置点検作業実施。 | 温度エラー頻発 | ゼロカセット、目視等で異常確認 | SWO-161747 | [5 empty] | Equipment inspection work was carried out due to frequent temperature errors. | Frequent temperature errors | Check abnormalities by zero cassette, visual, etc.
No | Visit | Complaint | RCHN | 最大暗视野测试检查失败 | 最大暗视野测试检查失败,借903-364一个 |  | SWO-160181 | Defect | HW | Process Module | Optical Unit | 1035 | Maximum dark field test check failed | Maximum dark field test check failed, borrowing 903-364 one
No | Visit | Complaint | RDEU | Multiparameter fehler bei Kalibration | Reklamationsdaten | Getät überprüft, Reklamationsdaten ausgedruckt, Datensicherung und QC durchgeführt | SWO-160479 | [5 empty] | Multiparameter fehler bei Kalibration | Reklamationsdaten | Done checked, complaint data printed out, data backup and QC performed
No | 8 | Y | RFRA | Blank | Derrive Ca++/CQ PCO2 hors gamme (trop fort)/erreur 478 derrive PCO2/Fuite par joint d'intro. |  | D1F460BE9E2CCEC4 | Revisit required | Electrode | Electrode | pCO2 |  | Blank | Derrive Ca/CQ PCO2 off-range (too strong)/error 478 derrive PCO2/Leak by intro seal.
No | Visit | Complaint | RDEU | 06.01.2017, Herr Tauber, Tel.: 05281992212: Kalibrationsfehler bei Ca 06.01.2017, 11:28 Uhr ==> Anrufer Her Tauber, Tel.: 05281992212 ==> Kalibrierfehler Ca besteht weiterhin (Kal.-staus 26 mmol/L) pCO2 Empf. außer Bereich | Ca- und Referenz- Elektrode ausgetauscht | Ca- und Ref.- Elektode ersetzt, Probelauf, Kal. Und QC ok | SWO-160423 | Assisted Replacement | Electrode | Electrode | Reference |  | 06.01.2017, Mr. Tauber, Tel.: 05281992212: Calibration error at Ca 06.01.2017, 11:28 ==> Caller Her Tauber, Tel.: 05281992212 ==> Calibration error Ca persists (Cal.-staus 26 mmol/L) pCO2 Mpf. out of range | Ca- and reference electrode replaced | Ca- and Ref.- Elektode replaced, test run, Cal. And QC ok
No | 9 | Y | RAME | 2 PC02 level 1 and 2 dropping out low and failing QC |  | Replaced the PC02 membrane and the dirift was 18 mm low on cal 2 and 10 mm low on cal 1. after 2 point passed and drift was corrected the QC came right up - before membrane values are at L1=64, L2=38.9, L3=20.9 $ $ After replacement values are at L1=66.8, L2=40.5, L3=22.9 this puts level 2 and 3 at a higher value than instrument #1 and less than 1 mm difference between both units on level 1. | D1F47ADF9E09E867 | Assisted Replacement | Electrode | Membrane | pCO2 |  | 2 PC02 level 1 and 2 dropping out low and failing QC |  | Replaced the PC02 membrane and the dirift was 18 mm low on cal 2 and 10 mm low on cal 1. after 2 point passed and drift was corrected the QC came right up - before membrane values are at L1=64, L2=38.9, L3=20.9 $ $ After replacement values are at L1=66.8, L2=40.5, L3=22.9 this puts level 2 and 3 at a higher value than instrument #1 and less than 1 mm difference between both units on level 1.
Yes | 3 | Y | RPAC | Blank | Operator and reference manuals. |  | D1F450C79E561547 | Re classified RMED | Logistical upSWO |  |  |  | Blank | Operator and reference manuals.
No | Visit | Complaint | RDEU | pCO2 Kal. Fehler | Kal. Daten analysiert. Nach längerer Ruhestellung pCO2 Kal Fehler Empf. 45% | Membrane D788 Lot. R0759 undicht, Flüssigkeit tritt an Membrane aus. Elektrode erneut membraniert , weiterhin Empf. Fehler 50%. Mehrere Kal. durchgeführt . pCO2 Elektrode erneuert und mit gleicher Membran eingesetzt. Alle Werte o.k. | SWO-160482 | Assisted Replacement | Electrode | Membrane | pCO2 |  | pCO2 Kal. Error | Kal. analyzed data. After prolonged resting position pCO2 Kal error empf. 45% | Membrane D788 Lot. R0759 leaks, liquid leaks to membrane. Electrode re-membraneed, continues empf. Error 50%. Several Kal. Performed. pCO2 electrode renewed and used with the same membrane. All values ok.
No | 1 | Y | RFRA | Blank | Tension neon too low |  | D1F48F4B9E067876 | Defect | HW | Oxi module | PCB, Oxi module |  | Blank | Tension neon too low
Yes | Visit | Complaint | RDEU | pCo 2 Fehler in QC c | Fehlersuche | QC durchgeführt | SWO-160582 | Re classified RMED | ECKHI Email 24.jan-21017 | Unknown | pCO2 | QC Level C | pCo 2 Fehler in QC c | Troubleshooting | QC performed
No | Visit | Complaint | RDEU | Meldungen keine Probe während Ansaugung erkannt kommt häufig bei den Kapillar Messungen | Häufig Probenabbrüche bei Kapillar Messung | Jahreswartung lt. Vertrag und RDEU Vorgaben durchgeführt, thb Kal durchgeführt , Kunde verwendet Kabe BK140 Kapillären zur Messung -> Kunden auf Einhaltung des Opl Vertages hingewiesen | SWO-160812 | User error | User error | User error | Capillary tubes | Kabe BK140 | Messages no sample detected during suction is common in capillary measurements | Frequent sample interruptions during capillary measurement | Annual maintenance carried out according to contract and RDEU specifications, thb Kal carried out, customer uses Kabe BK140 capillary for measurement -> Customers advised to comply with the Opl Contract
No | 1 | Y | RCZE | Blank | Error 1020, 1028, 1030, 1034 |  | D1F6EAF38A068709 | Revisit required | HW | Oxi module | Oxi module | Data investigation | Blank | Error 1020, 1028, 1030, 1034
No | Visit | Complaint | RIND | electrolytes values high compering with other instrument and sometimes QC values also out of range. | found problem with sensor board. | cleaned the sensor board and signal data board, did zeroing for sensor connector board, did calibration and QC all values found with in the range instrument kept on observation for one day | SWO-162925 | [5 empty] | electrolytes values high compering with other instrument and sometimes QC values also out of range. | found problem with sensor board. | cleaned the sensor board and signal data board, did zeroing for sensor connector board, did calibration and QC all values found with in the range instrument kept on observation for one day
No | Visit | Complaint | RDEU | 16.01.2017/10:50Uhr Frau Sager 03948599852: Das Gerät zeigte im laufendem Betrieb pH und pCO2-Kal-Fehler | Nach Neuinstallation eines Lösungspack war dieser undicht | Analysator beim Eintreffen dicht, Gerät gereinigt, Funktionskontrolle undichter Lösungspack: Lot 229816 | SWO-160578 | [5 empty] | 16.01.2017/10:50pm Mrs. Sager 03948599852: The device showed pH and pCO2-Kal errors during operation | After reinstalling a solution pack, this leaked | Analyzer on arrival dense, device cleaned, functional control of leaky solution pack: Lot 229816
No | Visit | Complaint | RCHN | LAC定标不通过 | LAC定标不通过 | 连续更换两张测试卡修复 | SWO-161724 | [5 empty] | LAC calibration does not pass | LAC calibration does not pass | Replace two test cards in a row for repair
No | Visit | Complaint | RCHN | HCT测量出错 | HCT测量出错,借903-329一个 |  | SWO-160463 | Cleaned | HW | Aspiration Module | Needle Wash Unit | 1267 | HCT measurement error | HCT measurement error, borrow 903-329 one
No | Visit | Complaint | RKK | ｲﾝﾚｯﾄﾎﾟｼﾞｼｮﾝが不定期で悪くなりｶﾞｽｹｯﾄから液漏れが頻発しﾘｰｸ電流発生。ﾌﾘｸｼｮﾝﾃﾞｨﾃｸﾀｰﾕﾆｯﾄ交換。 | ｲﾝﾚｯﾄﾎﾟｼﾞｼｮﾝが不定期で悪くなりｶﾞｽｹｯﾄから液漏れが頻発しﾘｰｸ電流発生。 | ﾌﾘｸｼｮﾝﾃﾞｨﾃｸﾀｰﾕﾆｯﾄ交換。 | SWO-160540 | Defect | HW | Inlet module | Friction detector unit |  | The inlet position deteriorates irregularly and leakage current occurs frequently from the gasket. Friction detector unit replacement. | The inlet position deteriorates irregularly and leakage current occurs frequently from the gasket. | Friction detector unit replacement.
No | Visit | Complaint | RDEU | Keine Datenübertragung | Keine Datenübertragung | Netzwerk Verbindung geprüft, Hub neu gestartet, IT Hotline informiert | SWO-161450 | Setup | SW | His/lis | His/lis |  | No data transfer | No data transfer | Network connection checked, hub restarted, IT hotline informed
No | Visit | Complaint | RIND | Touch screen was not working | Touch screen was not working | Cleaned the touch screen properly,reconnect the touch screen connector and then switch on the instrument and touch screen was working properly. | SWO-161586 | Defect | HW | DMS_PC | Touch screen |  | Touch screen was not working | Touch screen was not working | Cleaned the touch screen properly,reconnect the touch screen connector and then switch on the instrument and touch screen was working properly.
No | Visit | Complaint | RIND | O2 CO2 values not acceptable. | Parameter values not Ok. | QC setup done. QC ampul installed. As Lactate not calibrating ask for new lactate mambrane and QC sheduled for next visit. | SWO-163119 | Revisit required | HW | Auto Check Module | QC Ampul | Installed | O2 CO2 values not acceptable. | Parameter values not Ok. | QC setup done. QC ampul installed. As Lactate not calibrating ask for new lactate mambrane and QC sheduled for next visit.
No | Visit | Complaint | RDEU | PCO2 Status und Empfindlichkeitsfehler | PCO2 Status und Empfindlichkeitsfehler nach Wartungsmassnahmen, Glasbruch | pCO2 Elektrode (E788 XP-03) geprüft und ausgetauscht, Funktionskontrolle ok | SWO-161448 | Assisted Replacement | Electrode | Electrode | pCO2 | Glass breakage | PCO2 Status and Sensitivity Error | PCO2 status and sensitivity errors after maintenance measures, glass breakage | pCO2 electrode (E788 XP-03) tested and replaced, function control ok
No | 3 | Y | RAME | Z AutoCheck not scanning all the ampoules |  | Perform scan in service not reading all the ampoules | D1F5D20E14012E12 | Defect | HW | Auto Check Module | FPC Cable |  | Z AutoCheck not scanning all the ampoules |  | Perform scan in service not reading all the ampoules
No | 3 | Y | RAME | 9 INTERMITTENT LAC CAL FAILURE. |  | CONFIRMED ISSUE WHEN ONSITE. | D1F5A0D3140480C0 | No fault found | Electrode | Membrane | Lac |  | 9 INTERMITTENT LAC CAL FAILURE. |  | CONFIRMED ISSUE WHEN ONSITE.
No | Visit | Complaint | RDEU | 17.01.2017/13:20Uhr Frau Göttisch 05617398320: Das Gerät akzeptiert die neuen Cartrigde nicht. 17.01.2017; 14:15 Uhr; Fr. Bender; Tel.: 05617398322; die Kal kann nicht gestartet werden | Kalibration startet nicht | Der Barcode wurde bereits am Gerät erkannt. Sehr wahrscheinlich wurde die Meldung übersehen, dass zu wenig Kapazitäten in der Pufferkassette waren, um eine Kalibration zu starten. Im Servicemenu zeigte das Gerät keine Fehler mehr. Pufferkassette erneuert, danach Kalibration durchgeführt. Anwender startet danach eine Qualitätskontrolle. Gerät überprüft. | SWO-160794 | Unknown | Not Cat. | Not Cat. | Not Cat. |  | 17.01.2017/13:20pm Ms. Göttisch 05617398320: The device does not accept the new Cartrigde. 17.01.2017; 2:15 p.m.; Ms. Bender; Tel.: 05617398322; the Kal cannot be started | Calibration does not start | The barcode has already been detected on the device. Most likely, the message that there was too little capacity in the buffer cartridge to start a calibration was overlooked. In the service menu, the device no longer showed any errors. Buffer cassette renewed, then calibrated. Users then start quality control. device checked.
No | Visit | Complaint | RDEU | Glu Lac QC Fehler |  |  | SWO-160997 | Assisted Replacement | SC | SC | Glu/Lac |  | Glu Lac QC Fehler
No | Visit | Complaint | RDEU | 18.01.2017, 12:10 Uhr ==> Tel.: 04421892330 ==> Benutzerintervention erforderlich ==> Fehler 1342: Leckage erkannt 18.01.2017/12:30Uhr Herr Beewen 04421892330: Trotz Wartungsmaßnahmen weiterhin Transportfehler | Intervention erforderlich, immer wider LP Tausch oder durchspülen vom Gerät wird angezeigt, Verschmutzung ist in der SC | manuell gespült, Qwetschschlauch getauscht, SC getauscht, Kaltstartphase, manuelle Kalibrierungen durchgeführt, QC,s bei PH und pCO2 nicht bestanden, LP getauscht, Kal und QC IO | SWO-161063 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | 18.01.2017, 12:10 pm ==> Tel.: 04421892330 ==> User intervention required ==> Error 1342: Leakage detected 18.01.2017/12:30pm Mr Beewen 04421892330: Despite maintenance measures, transport errors continue | Intervention required, always against LP exchange or flushing through the device is displayed, Pollution is in the SC | manually rinsed, qwetsch hose exchanged, SC exchanged, cold start phase, manual calibrations performed, QC,s failed at PH and pCO2, LP exchanged, Kal and QC IO
No | Visit | Complaint | RDEU | Analysator zeigt mehrere Fehlermeldungen. Betriebsmittel (SC + LP) wurden bereits getauscht, ohne Erfolg. | erneut Intervention erforderlich --> siehe SWO-160052 und 155090 Fehlersuche --> Quetschventil und Hämolysator im Service überprüft --> Quetschventil schaltet gut; Clot detect/remove i.O. --> vorsichtshalber Hämolysator erneuern | Hämolysator (R0698N009) und Schlauch am Quetschventil erneuert, System und Funktion i.O. --> SD gezogen | SWO-160978 | Defect | HW | Oxi module | Hemolyzer |  | Analyzer displays several error messages. Equipment (SC + LP) has already been exchanged, to no avail. | intervention required again --> see SWO-160052 and 155090 Troubleshooting --> Squeeze valve and hemolyser checked in service --> Squeeze valve switches well; Clot detect/remove i.O. --> renew hemolyser as a precaution | Hemolyser (R0698N009) and hose on the crushing valve renewed, system and function in the order --> SD pulled
No | Visit | Complaint | RCHN | 报错1302温度不稳定，偏高。接口板坏了 | 报错1302温度不稳定，偏高。接口板坏了,借902-972一个 |  | SWO-160817 | [5 empty] | The error 1302 temperature is unstable and high. The interface board is broken. | The error 1302 temperature is unstable and high. The interface board is broken, borrow 902-972 one
No | Visit | Complaint | RIND | electrolytes values high compering with other instrument and sometimes QC values also out of range. | NA | compared the patient results and QC results with other instrument found all values with in the range. | SWO-162927 | [5 empty] | electrolytes values high compering with other instrument and sometimes QC values also out of range. | In | compared the patient results and QC results with other instrument found all values with in the range.
No | Visit | Complaint | RLTD | Email received from Vitor Roldao | Fault is intermittent, described by customer that during sleep study without any user interacting with the monitor a mouse cursor appears on the screen, then shortly after the screen goes black. Can be reset by power on/off button. Screen also observed to freeze but stay present again needs resetting by power off/on button. Downloaded system report and system dump to send to Radiometer Basel for investigation. Tested monitor, but no fault observed during visit. Monitor removed to Radiometer Limited workshop for investigation/repair. | Unable to explain cause of error, monitor removed for investigation, (Sensor & Mains lead left with customer). | SWO-161014 | [5 empty] | Email received from Vitor Roldao | Fault is intermittent, described by customer that during sleep study without any user interacting with the monitor a mouse cursor appears on the screen, then shortly after the screen goes black. Can be reset by power on/off button. Screen also observed to freeze but stay present again needs resetting by power off/on button. Downloaded system report and system dump to send to Radiometer Basel for investigation. Tested monitor, but no fault observed during visit. Monitor removed to Radiometer Limited workshop for investigation/repair. | Unable to explain cause of error, monitor removed for investigation, (Sensor & Mains lead left with customer).
No | Visit | Complaint | RLTD | Email from Vitor Roldao. | Monitor had 71% gas left Friday pm, by Monday Morning was at 36%. Checked that calibration chamber gasket was correctly fitted - found to be ok. Cleaned chamber and surface of electrode. Advised customer to ensure electrode face is cleaned of contact gel after use, before placing back into calibration chamber. Took system log and service dump for evaluation. Checked fitting of calibration chamber gasket on the 11 other customer TCM5s. | System log & service dump to be evaluated. | SWO-161157 | [5 empty] | Email from Vitor Roldao. | Monitor had 71% gas left Friday pm, by Monday Morning was at 36%. Checked that calibration chamber gasket was correctly fitted - found to be ok. Cleaned chamber and surface of electrode. Advised customer to ensure electrode face is cleaned of contact gel after use, before placing back into calibration chamber. Took system log and service dump for evaluation. Checked fitting of calibration chamber gasket on the 11 other customer TCM5s. | System log & service dump to be evaluated.
No | 23 | Y | RFRA | Blank |  |  | D1F6F9D18A19F131 | Cleaned | HW | Oxi module | Hemolyzer |  | Blank
No | Visit | Complaint | RDEU | 06.01.2017, 10:40 Uhr ==> Anrufer Herr Thiele, Tel.: 036458543434 ==> Probe bei Ansaugung vorzeitig im Oxi-Modul erkannt 09.01.2017/10:30, Herr Thiele: weiterhin das Problem --> Probe vorzeitig im OXI LS erkannt 17.01.2017, Herr Kutschbach, Tel.: 036458543435; : Weiterhin Probe vorzeitig im Oxi LS erkannt | Probe vorzeitig im Oxi LS erkannt tritt hauptsächlich bei AC Messung auf | Fehlersuche Wartung Fließwege, Einlaß, Luftansaugstutzen Rollenkranz vonLösungspumpe gewartet Ventilwartung Oximodul ( scheint etwas zu haken) LS justage , Pumpenkalibrierung Austausch div. Schläuche ABL meldet nach Neustart bei Testmessung OXI Hardwarefehler, aber keine Meldung unter Systemmeldung mehrfach Neustart durchgeführt Testmessungen | SWO-160872 | Defect | HW | Fluidic manifold | Pump tubes | Probably revisit | 06.01.2017, 10:40 a.m. ==> Caller Mr. Thiele, Tel.: 036458543434 ==> Sample in suction detected prematurely in the Oxi module 09.01.2017/10:30, Mr. Thiele: continue the problem --> Sample detected prematurely in OXI LS 17.01.2017, Mr. Kutschbach, Tel.: 036458543435; : Further sample detected prematurely in Oxi LS | Sample detected prematurely in Oxi LS occurs mainly with AC measurement | Troubleshooting Maintenance flow paths, inlet, air intake nozzle Roller wreath serviced by solution pump Valve maintenance Oximodul (seems to hook a bit) LS adjustment , pump calibration Replacement of div. hoses ABL reports OXI hardware failure after restart on test measurement, but no message under system message Restarted multiple times Test measurements
No | 35 | Y | RAME | 1 PH would intermittently fail QC and calibrations. |  | Customer stating that PH electrode and measuring chamber would fill up with blood numerous times. | D1F6FBC58A620C76 | Defect | Electrode | Electrode | Sealing ring | Sealing ring, pH | 1 PH would intermittently fail QC and calibrations. |  | Customer stating that PH electrode and measuring chamber would fill up with blood numerous times.
No | 20 | Y | RAME | 1 PH would intermittently fail QC and calibrations. |  | Customer stating that PH electrode and measuring chamber would fill up with blood numerous times. | D1F6D9E98A570C6B | Defect | Electrode | Electrode | Sealing ring | Sealing ring, pH | 1 PH would intermittently fail QC and calibrations. |  | Customer stating that PH electrode and measuring chamber would fill up with blood numerous times.
No | 5 | Y | RCZE | Blank | Chyba 251 |  | D1F6ECEE8A4C870B | Adjustment | HW | Oxi module | Lamp unit | Halogen | Blank | I guess 251
No | Visit | Complaint | RKK | pCO2 QC測定値が変動する | pCO2 QC測定値が変動する | 装置引上げの為、代替機設置・動作確認 | SWO-161040 | Unknown | Electrode | Electrode | pCO2 |  | PCO2 QC readings fluctuate | PCO2 QC readings fluctuate | Installation and operation confirmation of alternative machines for equipment pulling up
No | Visit | Complaint | RKK | pCO2・Cl QC測定値が変動する | pCO2・Cl QC測定値が変動する | 修理完了装置入替 CAL・QC測定データ確認 | SWO-161030 | Defect | HW | Inlet module | Friction detector unit | pCO2 and CL issues as well | PCO2 and Cl QC readings fluctuate | PCO2 and Cl QC readings fluctuate | Replacement of repair completion equipment Cal/QC measurement data confirmation
No | 4 | Y | RAME | P Inlet Leak. |  | Fluid leaking from inlet. Suspect is worn pump tubing or clogged inlet gasket. | D1FC47FB625007AE | Defect | HW | Fluidic manifold | Pump tubes |  | P Inlet Leak. |  | Fluid leaking from inlet. Suspect is worn pump tubing or clogged inlet gasket.
Yes | 3 | Y | RCZE | Blank | posunuté meze na kontrolách pCO2 na horní hranici, je v limitu ale Nekl požaduje upravit na støed |  | D1F7ED230016CED7 | Re classified RMED | Planned re install module |  |  |  | Blank | The shifted limits on pCO2 checks on the upper limit, but the limit is not required to adjust to the middle
Yes | Visit | Complaint | RDEU | 10.01.2017 Frau Bartmann 06142881562 meldet, dass es immer wieder QXI-QC Abweichungen gibt. 16.01.2017/11:40Uhr Frau Bartmann 06142881562: Weiterhin QC Ergebnisse bei CO2Hb und O2Hb grenzwertig. |  |  | SWO-160965 | Re classified RMED |  |  |  |  | 10.01.2017 Ms. Bartmann 06142881562 reports that there are always QXI-QC deviations. 16.01.2017/11:40pm Ms Bartmann 06142881562: Further QC results for CO2Hb and O2Hb borderline.
No | Visit | Complaint | RIND | O2 CO2 values not acceptable. | QC to run. | Analyser Calibrate properly. QC run and found all values well with in range. Inform to managment and Dr. Indrajeet for same. Dr convensed and confirm to do all test on analyser. Analyser Handed over in a fine working condition. | SWO-163120 | Assisted Replacement | HW | Auto Check Module | QC Ampul | Installed | O2 CO2 values not acceptable. | QC to run. | Analyser Calibrate properly. QC run and found all values well with in range. Inform to managment and Dr. Indrajeet for same. Dr convensed and confirm to do all test on analyser. Analyser Handed over in a fine working condition.
No | 9 | Y | RAME | Z Internal bar code scanne rwill no light up(Intermittent) |  | Tested scanner. Will not light ujp | D1F6E41F8A08BB25 | Defect | HW | Basic Unit | Barcode reader, Basic Unit |  | Z Internal bar code scans rwill no light up (Intermittent) |  | Tested scanner. Will not light ujp
No | 3 | Y | RPNZ | Blank | Unstable pH measurments |  | D1F7744E8A0CC4AB | Cleaned | HW | pH/BG Module | Measuring chamber |  | Blank | Unstable pH measurments
No | 3 | Y | RAME | Z Customer called with OXY errors and failing QC | Changed the complaint visit tab from N to Y. | Complete OXI test passed Lamp adjustment pass | D1F6E5398A16CA4E | Adjustment | HW | Oxi module | Lamp unit |  | Z Customer called with OXY errors and failing QC | Changed the complaint visit tab from N to Y. | Complete OXY test passed Lamp adjustment Pass
No | Visit | Complaint | RLTD | Results | Discrepant result- Radiometer medical recommended action change of parts & Main Visit | Changed all parts on machine that were outlined by the Techline investigated by Radiometer RMED. 6 month visit was due but main service visit carried out. All parts of main kit changed. Bank tests and LQC ran on board. Site to Re-calibrate and run LQC all parameters. Site to monitor. | SWO-160920 | Unknown | Techline | Techline | Techline |  | Results | Discrepant result- Radiometer medical recommended action change of parts & Main Visit | Changed all parts on machine that were outlined by the Techline investigated by Radiometer RMED. 6 month visit was due but main service visit carried out. All parts of main kit changed. Bank tests and LQC ran on board. Site to Re-calibrate and run LQC all parameters. Site to monitor.
No | 8 | Y | RFRA | Blank | Erreur calibration PCT lot 12499 (blocage mécanique durant la calibration ) |  | D1F6CAD88A5518B6 | Defect | Consumables | Test Kit | Cartridge, PCT | Broken | Blank | PCT lot 12499 calibration error (mechanical blockage during calibration)
Yes | Visit | Complaint | RDEU | Anwender haben teile, die der Anwender tauschen kann bereits getauscht, auch Betriebsmittel, ohne Erfolg. Analysator will weiterhin Teile getauscht haben. | Es wurden visual Flow check überprüft, und die ganze Reihe des Fluistransports und der Fehler bleibt unverändert | Erneuerung der Sensorkassette und Solution pack | SWO-161315 | Re classified RMED | ECKHI Email 24.jan-21017 | SP | FT | Leakage | Users have parts that the user can exchange already, even equipment, to no avail. Analyzer wants to continue to exchange parts. | Visual flow check has been checked and the whole series of fluis transport and the error remains unchanged | Renewal of the sensor cartridge and solution pack
Yes | 1 | Y | RAME | Z Install analyzer. |  | Had analyzer shipped to my home. I turned analyzer on and installed SC and SP. It continued to give the ambient temp error. I calibrated the thermisters in service and the hemolyzer would not go above 31.12C. I also had total blockage in waste line. Flush that out. I removed hemolyzer and flushed and checked all connections. Analyzer would start initializing and then go to ambient temp error. I spoke with Andy and will ship this one back. | D1F82B390046E1B2 | Re classified RMED | OOB |  |  |  | Z Install analyzer. |  | Had analyzer shipped to my home. I turned analyzer on and installed SC and SP. It continued to give the ambient temp error. I calibrated the thermisters in service and the hemolyzer would not go above 31.12C. I also had total blockage in waste line. Flush that out. I removed hemolyzer and flushed and checked all connections. Analyzer would start initializing and then go to ambient temp error. I spoke with Andy and will ship this one back.
No | Visit | Complaint | RDEU | nach SC Wechesl Glu/Lac QC Fehler, danach mehrfacher LP Wechsel | nach SC wechsle Glu/Lac QC Fehler, am nächsten Tag fordert Gerät zum LP Gerät ist bei Anwesenheit grün und messbereit | SC wurde routinemäßig am 17.01.17 7:33Uhr ersetzt ->QC Fehler bei Glu $Lac bis 18.01.17 7:12 Uhr ->Gerät empfiehlt LP wechsle LP wurde ersetzt DataLogs & SD erstellt und an TS geschickt | SWO-161248 | Assisted Replacement | SC | SC | Glu/Lac | Start up | after SC Wechesl Glu/Lac QC error, then multiple LP change | after SC change Glu/Lac QC error, the next day prompts device to LP Device is green and ready for measurement in presence | SC was routinely replaced on 17.01.17 7:33o ->QC error with Glu $Lac to 18.01.17 7:12 o'clock ->Device recommends LP change LP has been replaced DataLogs & SD created and sent to TS
No | Visit | Complaint | RSAF | Customer called to say her analyzer does not want to correct even after doing everything that it instructs her to do. | Check data logs and saw error started shortly after the replacement of n sensor cassette on the 12/01/2017 | Done trouble shooting in service mode to try and correct error but had to replace sensor cassette and error went away immediately. All calibrations ok and Qc's ok. | SWO-161415 | Assisted Replacement | SC | SC | Cl |  | Customer called to say her analyzer does not want to correct even after doing everything that it instructs her to do. | Check data logs and saw error started shortly after the replacement of n sensor cassette on the 12/01/2017 | Done trouble shooting in service mode to try and correct error but had to replace sensor cassette and error went away immediately. All calibrations ok and Qc's ok.
No | Visit | Complaint | RCHN | LAC定标不过，HCT,NA,N/E未计算 | LAC定标不过，HCT,NA,N/E未计算 | 更换2张测试卡后修复 | SWO-161332 | [5 empty] | LAC calibration, HCT, NA, N/E uncalculated | LAC calibration, HCT, NA, N/E uncalculated | Repair after replacing 2 test cards
No | Visit | Complaint | RCHN | 光谱不匹配，浑浊度高 | 光谱不匹配，浑浊度高，借备件902-980和903-197各一个 | 更换902-980和903-197修复(By Ran Gao) | SWO-161055 | Defect | HW | Oxi module | Hemolyzer |  | Spectrum mismatch, high turbidity | Spectrum mismatch, high turbidity, one for spare parts 902-980 and 903-197 each | Replacement 902-980 and 903-197 repairs (By Ran Gao)
Yes | Visit | Complaint | RCHN | 1061定标中止 | 1061定标中止,借944-220（12325）一个 | 工程师报错机器序列号，此单取消 （By Ran Gao） | SWO-161071 | Re classified RMED | AQT issue |  |  |  | 1061 Calibration abort | 1061 calibration suspension, loan 944-220 (12325) one | The engineer incorrectly reported the machine serial number and this order was cancelled (By Ran Gao)
No | Visit | Complaint | RCHN | PO2质控失败 0375 | PO2质控失败 0375 | 更换测试卡正常。 | SWO-161755 | Assisted Replacement | SC | SC | pO2 | 375 | PO2 QC Failure 0375 | PO2 QC Failure 0375 | The replacement test card is OK.
Yes | Visit | Complaint | RCHN | 光学强度超出范围 | 光学强度超出范围,借903-364一个 | 更换备件903-364后修复(By Ran Gao) | SWO-161033 | Defect | HW | Process Module | Optical Unit | 1031 | Optical strength out of range | Optical strength out of range, borrowed 903-364 a | Replacement of spare parts 903-364 after repair (By Ran Gao)
No | Visit | Complaint | RCHN | 1061定标中止 | 1061定标中止,借944-220（12325）一个 |  | SWO-161077 | Cleaned | HW | Storage Module | Process wheel | 1061 | 1061 Calibration abort | 1061 calibration suspension, loan 944-220 (12325) one
No | Visit | Complaint | RDEU | 19.01.2017/11:30Uhr Herr Viehoff 028412002770: pH-Fehler bei Kal 19.01.2017, 13.05 Uhr , Herr Viehoff: 028412002770: pH Kalibrationsfehler bleibt bestehen. | Umsichtigkeit Messkammer pH / BG | Messkammer gereinigt, Ref Elektrode geprüft, Dichtungsring an der pH Elektrode erneuert. | SWO-161290 | Defect | Electrode | Electrode | Sealing ring | Sealing ring, pH | 19.01.2017/11:30pm Mr Viehoff 028412002770: pH error at Kal 19.01.2017, 13.05 , Mr. Viehoff: 028412002770: pH Calibration error remains. | Umsichtigkeit Messkammer pH / BG | Measuring chamber cleaned, ref electrode tested, sealing ring renewed at pH electrode.
No | 4 | Y | RFRA | Blank | L'écran se bloque de temps en tps, lorsque les tubes sont insérés, puis OFF/ON pour redémarrer l'AQT |  | D1F8266E004A0D56 | [5 empty] | Blank | The screen hangs from time to time, when the tubes are inserted, then OFF/ON to restart the AQT
No | Visit | Complaint | RDEU | 19.01.2017, Frau Gollheim, Tel.: 03616541582 : AQC Fehler bei pCO2 Level A+C | QC Fehler | neues LP mit anderer LOT geliefert und initialisieret, alle QC Level bestanden | SWO-161368 | Assisted Replacement | SP | SP | pCO2 |  | 19.01.2017, Mrs. Gollheim, Tel.: 03616541582 : AQC error at pCO2 Level A+C | QC Fehler | new LP with other LOT supplied and initialized, all QC levels passed
No | Visit | Complaint | RDEU | 20.01.2017/10:00Uhr Herr Schneiter 04421891559: QC-Verletzung bei pH und CO2 20.01.2017; 14:15 Uhr; Hr. Schneiter; 04421891559; Spülung des Fließwegs funktioniert nicht | System steht in Erneuerung vom LP, Lösungspack nicht komplett arretiert | Lösungspack richtig arretiert, Kaltstartphase, Kalibrierung und alle drei QC,s IO System messbereit | SWO-161592 | Assisted Replacement | SP | SP | pCO2 |  | 20.01.2017/10:00uhr Mr Schneiter 04421891559: QC violation in pH and CO2 20.01.2017; 2:15 p.m.; Mr. Schneiter; 04421891559; Flushing of the flow path does not work | System is in renewal of the LP, solution pack not completely locked | Solution pack properly locked, cold start phase, calibration and all three QC,s IO system ready for measurement
No | Visit | Complaint | RBNL | Erreur PO2 | Mail du client : La callibration sort en erreur. Nous avons déjà remplacer une fois la casette des capteurs mais toujours le même souci. Savez vous nous aider svp. Sens. 91% Etat -25.2 mmHg | En ouvrant la fenêtre du SC je constate que la partie supérieur bouge un peu vers la gauche. Après démontage je vois que la partie supérieur du module Sensor Interface n'est pas bien fixe et donc ne se positionne pas toujours à la bonne place dans la Sensor Cassette. Après serrage de 2 visses , ajustage des contacts au dessus de la cassette, et remontage du module j'ai redémarré l'ABL et j'ai passé plusieurs calibrations (test de pO2 fait dans programme de service) sans erreur. Les 3 contrôles sont bien passées également.. | SWO-161510 | Adjustment | HW | Basic Unit | Sensor bay |  | PO2 error | Customer mail: Callibration is a mistake. We already replace once the box of sensors but still the same concern. Know you please help us. meaning. 91% Status -25.2 mmHg | When I open the SC window I notice that the top part moves a little to the left. After dismantling I see that the upper part of the Sensor Interface module is not well fixed and therefore does not always position itself in the right place in the Sensor Cassette. After tightening 2 screws, adjusting the contacts above the cassette, and reassembling the module I restarted the ABL and I passed several calibrations (pO2 test done in service program) without error. All three checks are well passed as well.
No | Visit | Complaint | RSAF | Leak currents detected during isolation test. | Stripped analyzer and inspected for leaks. Checked all contact points. Checked sample path. | Cleaned all electrical contact points. Reinitiliased analyzer. | SWO-161795 | Cleaned | HW | Basic Unit | Solution Pack PCB Id |  | Leak currents detected during isolation test. | Stripped analyzer and inspected for leaks. Checked all contact points. Checked sample path. | Cleaned all electrical contact points. Reinitiliased analyzer.
No | 2 | Y | RPOL | Blank | Sprawdzenie aparatu. Wpisanie zakresów referencyjnych. Analizator sprawny. |  | D1F958057622F00D | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | Blank | Check the camera. Enter reference ranges. Functional Analyzer.
No | Visit | Complaint | RDEU | Nach SC tausch sind nur die Blutgase und pH grün, Oxymetrie und Elektrolyte sind rot in Parameterleiste hinterlegt. Analysator gibt keine Info an warum. | Oximetrieparameter rot Keine Elektrolyte in Verwendung | Gerät mit gelber Ampel => LP nur noch 23 Zyklen Kal und QC ohne Fehler => Oximetrie nicht aktiviert Durchführung Service Test => Fehler bei Lampenstatus (LED Max Intensität zu hoch) und bei Hämolysatorstatus (Kavitation Dämpfen zu tief) Gerät zerlegt und Hemolyzer optisch überprüft => QC Austritt an oberer Schlauchverbindung und LS => Hemolyzer gereinigt und gespült Geräteneustart => Flüssigkeitssensorkalibrierung i.O. Oximetrieparameter aktiviert => anschließende Kal und QC i.O. Uhrzeit gestellt Analysator trennt und verbindet sich jede Minute fortlaufend mit Radiance | SWO-161480 | [5 empty] | After SC exchange, only the blood gases and pH are green, oxymetry and electrolytes are stored red in parameter bar. Analyzer does not specify why. | Oximetry parameters red No electrolytes in use | Device with yellow traffic light => LP only 23 cycles Kal and QC without error => Oximetry not activated Execution Service Test => Error in lamp status (LED Max intensity too high) and hemolyser status (cavitation vapours too low) Disassembled device and optically checked Hemolyzer => QC outlet to upper hose connection and LS => Hemolyzer cleaned and rinsed Device restart => Liquid sensor calibration in the o. Oximetry parameter activated => subsequent Cal and QC in the following Time set Analyzer separates and connects to Radiance continuously every minute
No | Visit | Complaint | RIND | continuously two sensor cassette and two solution got failed and its showing verification check failure | checked the machine and find out sensor connector board having rust in some ports . | connected zeroing cassette and zeroing done and installed solution pack and sensor cassette again its showing verification check failure .we need to replace sensor cassette connector board and signal data board for this isuee. | SWO-162943 | [5 empty] | continuously two sensor cassette and two solution got failed and its showing verification check failure | checked the machine and find out sensor connector board having rust in some ports . | connected zeroing cassette and zeroing done and installed solution pack and sensor cassette again its showing verification check failure .we need to replace sensor cassette connector board and signal data board for this isuee.
No | Visit | Complaint | RLTD | Oxi pinch valve failed during training. | Three month technical training completed. Whilst training the user and demonstrating a tHb calibration, the Oxi pinch valve failed. | New Oxi pinch valve fitted to the analyser. | SWO-161565 | Defect | HW | Oxi module | Pinch valve |  | Oxy Pinch valve failed during training. | Three month technical training completed. Whilst training the user and demonstrating a tHb calibration, the Oxi pinch valve failed. | New Oxi pinch valve fitted to the analyser.
No | Visit | Complaint | RDEU | 06.01.2017, 10:40 Uhr ==> Anrufer Herr Thiele, Tel.: 036458543434 ==> Probe bei Ansaugung vorzeitig im Oxi-Modul erkannt 09.01.2017/10:30, Herr Thiele: weiterhin das Problem --> Probe vorzeitig im OXI LS erkannt 17.01.2017, Herr Kutschbach, Tel.: 036458543435; : Weiterhin Probe vorzeitig im Oxi LS erkannt 19.01.2017; Hr. Kutschbach; Tel.: 036458543435; der Fehler besteht weiterhin | Probe bei Ansaugung vorzeitig im Oxi-LS erkannt Fehler tritt immer wieder beiMessung , aber vor allem bei QC auf; nicht permanent | Austausch Meßkammern neue Ventile und Ventilgehäuse am BG und EL Modul Dichtring pH El. erneuert Elektroden membraniert Einlaufverfahren | SWO-161518 | Defect | Electrode | Electrode | Sealing ring | Sealing ring, pH | 06.01.2017, 10:40 a.m. ==> Caller Mr. Thiele, Tel.: 036458543434 ==> Sample in suction detected prematurely in the Oxi module 09.01.2017/10:30, Mr. Thiele: continue the problem --> Sample detected prematurely in OXI LS 17.01.2017, Mr. Kutschbach, Tel.: 036458543435; : Further sample detected prematurely in Oxi LS 19.01.2017; Mr. Kutschbach; Tel.: 036458543435; the error still exists | Sample detected prematurely in Oxi-LS during suction Errors occur again and again in measurement , but especially with QC; not permanent | Replacement of measuring chambers new valves and valve housings on the BG and EL modules Sealing ring pH El. Renewed Electrodes membraneed Entry procedure
No | Visit | Complaint | RDEU | Auswahlliste für Layout leer - keine Möglichkeit Plichtfeld für Personal ID einzurichten | Auswahlliste für Layout leer | Wiederherstellung der Layouts nicht möglich, Konfiguration von weiterem Gerät im Hause geholt und eingespielt, Layouts angepasst, fällige Membran ersetzt, mehrere Kal, QC, Patienten ID Variable eingerichtet | SWO-161606 | Setup | SW | Patient Layout | Patient Layout |  | Selection list for layout empty - no way to set up plicht field for Personal ID | Selection list for layout empty | Restoration of layouts not possible, configuration of further device fetched in the house and recorded, layouts adapted, overdue membrane replaced, multiple Kal, QC, patient ID variable set up
No | Visit | Complaint | RDEU | Anwender hat viel zeit in BGA in den letzten Tagen verbracht. Zur Zeit sind Driften bei Lac und Cl Elektrode. Anwender sehr unzufrieden. | Fehlersuche vor Ort. | Bei Ankunft war Analysator Messbereit. Im Archiv war zu sehen dass die Kals seit geraumer Zeit abgebrochen wurden weil erst Spüllösung leer war und dann Kal 2 Lösung. Diese wurden jedoch in der Früh von Frau Schirp getauscht. Danach war die Kal wieder i.O. Frau Guth und Frau Schirp darüber informiert warum so viele Kals abgebrochen wurden. Trotzdem die Ventile aus PH BG und ELE MET Modul ausgebaut, wegen den ab und zu vorherrschenden Driften bei Lac. Beide Luftventile waren fest. Beide Ventile gangbar gemacht. Luftkanal gespült. LS Justage und Pumpenkal durchgeführt. Mehrere Kalibrierungen durchgeführt --> i.O. Keine fehlerhaften Driften mehr vorhanden. QCs für alle Level durchgeführt. Fragen von Personal ( Frau Guth, Schirp )zu Kalibrierung, Driften und Lösungstausch beantwortet und am ABL Menüpunkte zur Problemlösung aufgezeigt. ABL wieder Messbereit. Zwei Patientenmessungen durchführen lassen. --> diese waren i.O. | SWO-161509 | Cleaned | HW | pH/BG Module | Valve, air |  | User has spent a lot of time in BGA in the last few days. Currently, drifts are at Lac and Cl electrode. Users very dissatisfied. | Troubleshooting on site. | On arrival, analyzer was ready to measure. In the archive it could be seen that the kals were canceled for quite some time because first the flushing solution was empty and then Kal 2 solution. However, these were exchanged in the morning by Mrs Schirp. Afterwards, the Kal was again in the I.O. Mrs. Guth and Ms. Schirp informed about why so many kals were canceled. Nevertheless, the valves from PH BG and ELE MET module were removed, because of the from time to time prevailing drifts at Lac. Both air valves were fixed. Both valves made workable. Air duct flushed. LS adjustment and pump scale. Several calibrations performed --> in the end No more faulty drifts. QCs for all levels. Questions answered by staff ( Ms. Guth, Schirp ) about calibration, drifting and solution exchange and shown on the ABL menu items for problem solving. ABL ready to measure again. Have two patient measurements carried out. --> these were in the end.
No | 4 | Y | RAME | Z Analyzer stuck in hold mode. Leak detected message. |  | Examined inle assembly. Found bent probe. | D1F95A51764008A6 | Defect | HW | Inlet module | Probe bend |  | Z Analyzer stuck in hold mode. Leak detected message. |  | Examined inle assembly. Found bent probe.
No | Visit | Complaint | RKK | テストカートリッジのツメが破損で器械停止 | 確認 | サービスでは取り出せないので、分解してカートリッジを取り出す。下まで分解できないので、近くを確認したが破損のツメは見つからない。動作確認 | SWO-161782 | Defect | Consumables | Test Kit | Cartridge | Broken | The test cartridge's jamming is damaged and the instrument stops. | Check | Because it cannot be taken out by the service, it disassembles it and the cartridge is taken out. Because it is not possible to disassemble it to the bottom, the swallow of the damage is not found though it confirmed near. Check operation
No | Visit | Complaint | RDEU | 19.01.2017/08:40Uhr Frau Kügle 08252946323: Das Probenröhrchen wird nicht mehr gehalten, weil eine Kralle abgebrochen ist. | Tube Holder gebrochen. | Tube Holder erneuert. Abgebrochenes Teil von Tube Holder aus gerät entfernt. Nadel ausgebaut und überprüft - ok. LQC durchgeführt. Fehler im Zuge der Wartung beseitigt. | SWO-161500 | Defect | HW | Aspiration Module | Tube holder |  | 19.01.2017/08:40pm Mrs. Kügle 08252946323: The sample tube is no longer held because a claw is broken off. | Tube Holder broken. | Tube Holder renewed. Aborted part of Tube Holder from device removed. Needle removed and checked - ok. LQC. Fixed errors during maintenance.
No | Visit | Complaint | RDEU | Analysator bricht die Proben ab mit Meldung 1067. 20.01.2017, Frau Beier, Tel.: 023234995368: erneut Probenabrüche | Datenspeicher überprüft, der AQT hatte am 17.1. Meldung 1999 am 18.1. 5669 und heute 5680 und 5670 | System auf Cups überprüft, Motorspindeln gereinigt und geschmiert. Servicedump gesichert, Funktion z.Zt. OK. | SWO-161880 |  |  |  |  | 1067 | Analyzer cancels the samples with message 1067. 20.01.2017, Mrs. Beier, Tel.: 023234995368: probenprobenabruch | data storage, the AQT had on 17.1. Report 1999 on 18.1. 5669 and today 5680 and 5670 | System checked for cups, motor spindles cleaned and lubricated. Service dump secured, function currently OK.
No | Visit | Complaint | RDEU | 19.01.2017/15:45Uhr Frau Bender 05617398320: Das Gerät wurde wegen einem Gerätefehler neu gestartet. Danach besteht das Gerät die Initialsierung nicht und meldet Fehler 2167 | Messungen brechen ab. Calibrate13-Fehler, auch nach Ein-Ausschalten. | Gerät im Servicemenu überprüft. Cupsensor Threshold bei 18. servicedump gezogen und auswerten lassen. Cupsensor gereinigt und Patientenmessungen durchgeführt. Sollte Fehler wieder auftreten, muss der Sensor erneuert werden. E-Teil bestellt. | SWO-162081 | Revisit required | HW | Incubator Module | Sensor, Cup |  | 19.01.2017/15:45pm Ms. Bender 05617398320: The device was restarted due to a device error. After that, the device does not have the initialization and reports error 2167 | Measurements break off. Calibrate13 error, even after power-on. | device in the service menu. Cupsensor Threshold at 18th servicedump and have it evaluated. cup sensor and patient measurements carried out. If errors occur again, the sensor must be replaced. E-part.
No | Visit | Complaint | RCHN | 报错1024，底板底物外漏 | 报错1024，底板底物外漏，借备件903-364,TNI定标卡944-212（12162），ckmb定标卡944-216（12179），myo定标卡944-214（12338），D二聚体定标卡944-220（12124）各一个 | 更换903-364,重新定标TNI定标卡944-212（12162），ckmb定标卡944-216（12179），myo定标卡944-214（12338），D二聚体定标卡944-220（12124）后修复 （By Ran Gao） | SWO-161379 | Defect | HW | Process Module | Optical Unit | 1024 | Error 1024, leakage of the bottom plate | Error 1024, leakage of the bottom plate, spare part 903-364, TNI calibration card 944-212 (12162), ckmb calibration card 944-216 (12179), myo calibration card 944-214 (12338), D dipolymer calibration card 944-220 (12124) each | Replace 903-364, re-order TNI calibration card 944-212 (12162), ckmb calibration card 944-216 (12179), myo calibration card 944-214 (12338), D two polyp calibration card 944-220 (12124) after repair (By Ran Gao)
No | Visit | Complaint | RCHN | 调节屏幕角度时，机器自动重启 | 调节屏幕角度时，机器自动重启，借备件636-571，910-298各一个 | 更换636-571和910-298修复 by andrew | SWO-161679 | Defect | HW | Analyzer Control Module | FPC Cable |  | When adjusting the screen angle, the machine restarts automatically | When adjusting the screen angle, the machine automatically restarts, with spare parts 636-571,910-298 each | Replace 636-571 and 910-298 Repair by Andrew
No | Visit | Complaint | RDEU | LP erneuern da LP leer | Fehlersuche | LP erneuert | SWO-161823 | User Error | SP | SP | Empty |  | LP renew da LP empty | Troubleshooting | LP renewed
Yes | Visit | Complaint | RDEU | 23.01.2017; Pfl. Domenik; Tel.: 021663942111; Gerätekonfiguration stimmt nach Überprüfung nicht mehr | Ausdruck ist nicht mehr einseitig. | Patientenlayout stimmt nicht mehr - Falsche Vorgabe gesetzt. Das richtige Layout wieder als Standard eingestellt. | SWO-162000 | Re classified RMED | Phone support |  |  |  | 23.01.2017; Pfl. Domenik; Tel.: 021663942111; Device configuration no longer corrects after verification | Expression is no longer one-sided. | Patient layout is no longer correct - wrong default set. The correct layout is set as the default again.
No | Visit | Complaint | RDEU | Interner Komunikationsfehker zwischen CPU und AU Unit. Gerät bleibt Offline. | Gerät bootet ohne Fehlermeldungen, bleibt jedoch im Offlinemodus hängen Touchscreen ohne Funktion. Gerät bootet sich selbstständig nach einigen Minuten durch. | Da sich durch den Fehler der Touchscreen nicht bedienen lässt , testweise neuen Monitor eingebaut. Kein Erfolg! Sicherungen auf AU Board getestet. o.k. Steckerverbindungen getestet o.k. Gerät zur Reparatur in GS17 genommen. | SWO-162695 | Revisit required | analyzer removed | Power Supply | Power Supply | See UpSWO 25/1 | Internal communication feud between CPU and AU unit. Device remains offline. | Device boots without error messages, but remains in offline mode hanging touchscreen without function. Device boots itself after a few minutes. | Since the error does not allow the touchscreen to be operated, new monitor built-in on a test basis. No success! Fuses tested on AU board. Ok. Plug connections tested o.k. Device taken for repair in GS17.
No | 7 | Y | RFRA | Blank |  |  | D2024E4FB0540019 | Assisted Replacement | SP | SP | FT |  | Blank
No | Visit | Complaint | RDEU | 23.01.2017, Frau Wingerter, Tel.: 062181024322: Einlass schließt nicht richtig und AQC Fehler bei pCO2 | Bei Eintreffen Analysator nur noch mit Systemmeldung Inkonsistente Datenbank | Analysator überprüft - Einlass demontiert und Anschlussdichtung erneuert - Rückmontage und Neustart durchgeführt - Funktionsüberprüfung etc. ohne Beanstandung ! Bemerkung: Bei telefonischer Unterstützung am Wochenende schloss Einlass bei SP-Wechsel erst nach Reset des Analysators ! | SWO-161836 | [5 empty] | 23.01.2017, Mrs Wingerter, Tel.: 062181024322: Inlet does not close properly and AQC error with pCO2 | On arrival Analyzer only with system message Inconsistent database | Analyzer checked - inlet disassembled and connection seal renewed - reassembly and restart performed - function check etc. without complaint ! Note: With telephone support at the weekend, admission for SP change only closed after the reset of the analyzer !
Yes | Visit | Complaint | RBNL | Roma.ini aangepast | Op verzoek van klant, Roma.ini file heb aangepast zodat i.p.v. het patiëntennummer, het labnummer bij de resultaten te zien is. | Aanpassing Roma.ini | SWO-161975 | Re classified RMED | Changing Roma.ini on customer request |  |  |  | Roma. INI Modified | At the customer's request, Roma. ini file has been modified so that instead of the patient number, the lab number is shown in the results. | Adaptation Roma. ini
No | Visit | Complaint | RDEU | pH-Werte QC Level C zu niedrig pCO2-Werte bei QC Level A+C zu hoch | SP90 LOT YB-50 installiert | Kein Lösungspack einer anderen Charge vor Ort. | SWO-161993 | Assisted Replacement | SP | SP | pH |  | pH values QC Level C too low pCO2 values too high at QC Level A+C | SP90 LOT YB-50 installed | No solution pack of another batch on site.
Yes | Visit | Complaint | RDEU | pH-Werte QC Level C zu niedrig pCO2-Werte bei QC Level A+C zu hoch | Ersatzlösung-Pack installieren | SP90 LOT YB-51 installiert, QC ok | SWO-161994 | Re classified RMED | Same day same problem |  |  |  | pH values QC Level C too low pCO2 values too high at QC Level A+C | Install Replacement Solution Pack | SP90 LOT YB-51 installiert, QC ok
Yes | Visit | Complaint | RBNL | Roma.ini aanpassing | Op verzoek van klant, Roma.ini file heb aangepast zodat i.p.v. het patiëntennummer, het labnummer bij de resultaten te zien is. | Aanpassing Roma.ini | SWO-161972 | Re classified RMED | Changing Roma.ini on customer request |  |  |  | Roma. INI adaptation | At the customer's request, Roma. ini file has been modified so that instead of the patient number, the lab number is shown in the results. | Adaptation Roma. ini
No | Visit | Complaint | RCHN | 报1024错误代码 | 报1024错误代码，借903-364和903-236各一个 |  | SWO-161722 | Defect | HW | Process Module | Optical Unit | 1024 | Reported 1024 error code | Reported 1024 error code, borrow 903-364 and 903-236 each one
No | 6 | Y | RESP | Inlet wheel no gira correctamente. |  | Desmontado inlet wheel. Comprobados sensores y motores. | D1FD311CD84402E2 | [5 empty] | Inlet wheel does not rotate properly. |  | Disassembled wheel inlet. Tested sensors and motors.
No | 8 | Y | RESP | Aguja doblada. Probable causa suciedad en el sensor. |  | Sustituida aguja y sensor. Realizado ajuste de la aguja. | D1FD262ED85802DC | [5 empty] | Needle bent. Likely cause dirt on the sensor. |  | Replaced needle and sensor. Made needle adjustment.
No | Visit | Complaint | RDEU | Es werden keine AC's gemessen | Es werden keine AC's gemessen | AC Modul geprüft und demontiert, Flachbandkabel im AC Modul ausgetauscht, Funktionskontrolle ok | SWO-162173 | Defect | HW | Auto Check Module | Flat cable |  | No AC's are measured | No AC's are measured | AC module tested and disassembled, ribbon cable replaced in AC module, function control ok
No | Visit | Complaint | RDEU | 23.01.2017 10:55 Frau Silke Rosenow Tel.: 03055185041 Fehler: ABL tropft aus Einlass | während Spülung tritt Lösung am Einlaß aus | Fehlersuche, Reinigung, Einlaßdichtung entfernt, Schlauch angebracht und durchgespült, Abfallwege getestet und gespült, alles wieder montiert, LS Justage, Pumpenkal, Kal, QC | SWO-162167 | Revisit required | HW | Inlet module | Leakage |  | 23.01.2017 10:55 Mrs Silke Rosenow Tel.: 03055185041 Error: ABL drips from inlet | during rinsing solution at the inlet | Troubleshooting, cleaning, inlet seal removed, hose attached and flushed, waste paths tested and rinsed, everything reassembled, LS adjustment, pump scale, kal, QC
No | Visit | Complaint | RDEU | pCO2 Ergebnisse in L3 Ergebnissen zickzackmäßig an/unter der Untergrenze | s. case | Nach Rücksprache mit RDEU Service und Remote-Aufschaltung den Analysator aufgesucht, überprüft und ServiceDump sowie Datenextrakt gezogen und an EvaRa weitergeleitet zwecks Sendung an RMED und Auswertung. Zurzeit Ampel grün, keine Fehler. | SWO-161902 | [5 empty] | pCO2 results in L3 results zigzag at/below the lower limit | s. case | After consultation with RDEU Service and remote connection, the analyzer is visited, checked and ServiceDump and data extract pulled and forwarded to EvaRa for transmission to RMED and evaluation. Currently traffic light green, no errors.
No | Visit | Complaint | RDEU | 18.01.2017, Herr Döhler, Tel.: 017615323439: Betriebsmittelplan erscheint Mittags 12 Uhr | Konfiguration angepasst | Konfiguration angepasst | SWO-162172 | Setup | SW | Configuration | Configuration | Shuld not be an upSWO | 18.01.2017, Mr. Döhler, Tel.: 017615323439: Equipment plan is published at noon noon | Configuration adapted | Configuration adapted
Yes | Visit | Complaint | RDEU | Aufarbeitung und vorgezogene Neugeräteaufstellung wegen defektem Altgerät | Aufarbeitung Neugerät vorgezogen wegen defektem Altgerät. | Installation und Konfiguration vorbereitet. Aufstellung in Abteilung MEI durch Jona Lünenbürger. | SWO-162690 | Re classified RMED | HW | New Analyzer |  |  | Refurbishment and early installation of new equipment due to defective old equipment | Refurbishment of new device preferred defective old equipment. | installation and configuration. Installation in Department MEI by Jona Lünenbürger.
No | Visit | Complaint | RIND | Analyzer showing error on the screen, after restarting Analyzer problem not solved need to visit at the site | checked and found that the cartridge blocked the storage wheel. | removed the cartridge and inserted new one now working satisfactorily. also done PCT Calibration. | SWO-161814 | Defect | Consumables | Test Kit | Cartridge, Broken | 5689 | Analyzer showing error on the screen, after restarting Analyzer problem not solved need to visit at the site | checked and found that the cartridge blocked the storage wheel. | removed the cartridge and inserted new one now working satisfactorily. also done PCT Calibration.
No | Visit | Complaint | RDEU | 19.01.2017/15:45Uhr Frau Bender 05617398320: Das Gerät wurde wegen einem Gerätefehler neu gestartet. Danach besteht das Gerät die Initialsierung nicht und meldet Fehler 2167 | Gerät zur Reparatur abgeholt. | Gerät am 20.01.17 zur Reparatur abends abgeholt. Cup-Sensor erneuert und kalibriert. Blutmessungen und Myoglobin-Kalibration durchgeführt. Fehler trat nicht mehr auf. Gerät am 23.01.17 wieder beim Kunden aufgestellt und mit Qualitätakontrolle und weiteren Blutmessungen überprüft. | SWO-162080 | Revisit required | HW | Incubator Module | Sensor, Cup |  | 19.01.2017/15:45pm Ms. Bender 05617398320: The device was restarted due to a device error. After that, the device does not have the initialization and reports error 2167 | device for repair. | Device picked up on 20.01.17 for repair in the evening. Cup sensor renewed and calibrated. blood measurements and myoglobin calibration. Errors no longer occurred. Device placed again at the customer on 23.01.17 and checked with quality control and further blood measurements.
No | Visit | Complaint | RBNL | Bij het aanbieden van ene buis gebeurt er niets. Kan alleen buis verwijderen. Gisteren (donderdag zou er een storing geweest zijn waarbij instrument uitgezet is geweest en daarna weer opgestart. Foutmelding 3553: incubatorwheelmove_31 en 1010: systeemreinigingsprocedure is niet met succes voltooid. | Instrument meetklaar. In hisorie foutmelingen 3553 en 1010. 19-1-2017 15:03 fout 6732 storagedoor obstructed, kwartier later is wel de TNT cassette vervangen. Probeer vanmorgen een bepaling aan te vragen maar instrument herkent buis niet, geen reactie na plaatsen buis. Na verwijderen insert en bewegen vinger voor scanner is het wel weer mogelijk een bepaling te starten en met succes af te ronden. | Scanner gedemonteerd, kabelcontacten en lensjes schoongemaakt met alkohol. Serviceinstellingen gecontroleerd, OK. Nogmaals bepaling aangevraagd, scant direct buisje en barcode. Systeemreiniging en -controle met succes uitgevoerd. | SWO-161799 | [5 empty] | When offering one tube nothing happens. Can only remove tube. Yesterday (Thursday there would have been a malfunction where instrument was turned off and then restarted. Error message 3553: INCUBATORWHEELMOVE_31 and 1010: System cleaning procedure has not been completed successfully. | Instrument measurement ready. In Hisorie Errors 3553 and 1010. 19-1-2017 15:03 error 6732 storage by obstructed, fifteen minutes later the TNT cassette was replaced. Try to request a provision this morning but instrument does not recognize tube, no reaction after post tube. After removing insert and moving finger for scanner it is possible to start a provision and finalize it successfully again. | Scanner disassembled, cable contacts and Lenes cleaned with Alkohol. Service settings checked, OK. Once again provision requested, scans direct tube and barcode. System cleaning and control performed successfully.
No | Visit | Complaint | RDEU | 19.01.2017; Fr. DOmberg; Tel.: 05115322259; ABL ist mit dem pO2 2x durch den Ringversuch gefallen | PO2 Ringversuchs Ergebnisse unterschiedlich zu anderen ABL 90 | Mit dem Anwender (Frau Domberg) Ringversuche (RfB) durchgeführt und ausgewertet | SWO-162303 | No fault found | Parameter | Parameter | Analyzer comparison |  | 19.01.2017; Ms. DOmberg; Tel.: 05115322259; ABL has fallen with the pO2 2x through the ring test | PO2 ring test results different from other ABL 90 | Performed and evaluated with the user (Mrs. Domberg) ring tests (RfB)
No | Visit | Complaint | RLTD | Oxi failing QC 3. Suzi Knightly 02082962282 | Checked QC data log - analyser had been failing level 2 QC HbF low intermittently. Performed tHb calibration. Repeated QC - now more mid range. | performed tHb calibration | SWO-162286 | Adjustment | HW | Oxi module | tHb Cal |  | Oxi failing QC 3. Suzi Knightly 02082962282 | Checked QC data log - analyser had been failing level 2 QC HbF low intermittently. Performed tHb calibration. Repeated QC - now more mid range. | performed tHb calibration
No | Visit | Complaint | RLTD | Please call Jane Camileri on 0208 296 2282 | On arrival analyser in intervention required, error 1184 Leak Detected. Oximetry QC, tHb calibration and leak errors traced to a hole which had formed in the oximetry valve tube. Replaced faulty tube. Tested fluid flow - now ok. Performed tHb calibration - passed. Performed Quality Controls - passed. Analyser ready for use. | Replaced faulty Oxi valve tube. | SWO-162206 | Defect | HW | Oxi module | Pinch valve tube |  | Please call Jane Camileri on 0208 296 2282 | On arrival analyser in intervention required, error 1184 Leak Detected. Oximetry QC, tHb calibration and leak errors traced to a hole which had formed in the oximetry valve tube. Replaced faulty tube. Tested fluid flow - now ok. Performed tHb calibration - passed. Performed Quality Controls - passed. Analyser ready for use. | Replaced Faulty oxy valve tube.
No | Visit | Complaint | RDEU | 24.01.2017, Herr Söllner, Tel.: 095150315242: Temperaturfehler SC 1066 24.01.2017, 09.15 Uhr, Herr Söllner, 095150315242: Temperaturfehler bei SC und Systemtemperaturfehler, Vakuum u. Fließfehler | Fehlersuche --> Analysator im Service überprüft (Sensorkassettenfach angesteuert, Thermistor Calibration (Temperaturen bei SC außer Bereich), usw.) --> Sensorkassettenfach öffnet/schließt nicht richtig und Thermistor-Fehler --> Flachbandkabel sitzt nicht richtig; Spülung des Analysators --> Verbindungen der Leitungen/Kabel überprüft | Breites Flachbandkabel oberhalb des Sensorkassettenfaches wieder richtig eingebaut, Servicetest (Thermistorkalibration und Ansteuerung des Sensorkassettenfaches) --> beide Funktionen i.O., Startphase, System und Funktion i.O. | SWO-162177 | Defect | HW | Sensor Compartment | FPC Cable | 1066 | 24.01.2017, Mr. Söllner, Tel.: 095150315242: Temperature error SC 1066 24.01.2017, 09.15, Mr. Söllner, 095150315242: Temperature error at SC and system temperature error, vacuum and flow error | Troubleshooting --> Analyzer checked in service (sensor cassette compartment controlled, Thermistor Calibration (temperatures at SC out of range), etc.) --> Sensor cartridge compartment opens/closes not properly and thermistor error --> Ribbon cable does not sit properly; Rinsing the analyzer --> Connections of cables/cables checked | Wide ribbon cable above the sensor cassette compartment reinstalled correctly, service test (thermistor calibration and control of the sensor cassette compartment) --> both functions i.O., start phase, system and function i.O.
No | Visit | Complaint | RDEU | 16.01.2017/14:50Uhr Frau Defaber 00223274230: Ein Techniker soll bei der Netzwerkanbindung anwesend sein. | Conworksanbindung vorbereitet. Aber Conworx wird vorerst nicht genutzt. | Zusammen mit Herrn Lück die Netzwerkeinstellungen im Haus und am Gerät vorgenommen. QC Daten werden nun übertragen. Fehler im LP, Glu Level 1 fällt raus. Es ist kein anderes LP im Haus vorhanden. Neues LP ist bestellt. | SWO-162405 | Setup | SW | His/lis | His/lis |  | 16.01.2017/14:50pm Ms. Defaber 00223274230: A technician should be present at the network connection. | Conworks connection. But Conworx is not used for the time being. | Together with Mr. Lück, the network settings were made in the house and on the device. QC data is now transferred. Error in LP, Glu Level 1 falls out. There is no other LP in the house. New LP has been ordered.
No | 29 | Y | RESP | No pasa el test de fugas. |  | Comprobar tubos, cambiar la aguja toma muestras, calibrar y probar con muestras y controles. | D1FE56144E46945A | Defect | HW | Basic Unit | Pre Heater tube |  | It doesn't pass the leak test. |  | Check tubes, change the needle takes samples, calibrate and test with samples and controls.
No | 3 | Y | RPOL | Blank | Wykryto przeciek w systemie. Brak uszczelki pod membran¹ Cl. Wymiana membrany. Kalibracje. Analizator sprawny. |  | D1FE73E74E14C819 | Assisted Replacement | Electrode | Membrane | Cl |  | Blank | A leak has been detected in the system. No membrane GASKET ¹ CL. Replacing the membrane. Calibrations. Functional Analyzer.
No | 5 | Y | RAME | P Liquid Sensor Error. |  | Customer had inlet liquid sensor error and oxi liquid sensor errors. | D2001AE7C43B75A0 | Cleaned | HW | Basic Unit | Fluidic pathway |  | P Liquid Sensor Error. |  | Customer had inlet liquid sensor error and oxi liquid sensor errors.
No | Visit | Complaint | RDEU | 23.01.2017 10:55 Frau Silke Rosenow Tel.: 03055185041 Fehler: ABL tropft aus Einlass | Ursache für Ausfall gestern überprüfen, pH Kal gescheitert, QC bei div Parametern gescheitert | Fehlersuche, QC abgebrochen dadurch alle Parameter gesperrt, Kalwerte durchgesehen - alle Parameter membraniert, mehrere Kal, QC | SWO-162166 | Revisit required | HW | Inlet module | Leakage | 2'nd visit | 23.01.2017 10:55 Mrs Silke Rosenow Tel.: 03055185041 Error: ABL drips from inlet | Cause of failure yesterday check, pH Kal failed, QC failed at div parameters | Troubleshooting, QC aborted thereby locking all parameters, looking through cal values - all parameters membraned, several cal, QC
No | Visit | Complaint | RDEU | 24.01.2017 14:09 Frau Rosner Tel.: 03658282610 Fehler: tHb-Kalfehler in Kal1 und manueller Kal | THb-Kalibrierfehler Luftkanal gespült | Fehlersuche Proteinentfernung, Pumpenkalibrierung, LS-Justage Oximentrietest im Serviceprogramm durchgeführt Einlassdichtung erneuert Mehrmals 2-Pkt.-Kalibrierung und THb-Kalibrierung durchgeführt Fehler besteht bei QC-Messung ==> Ersatzteil muss bestellt werden | SWO-162311 | Cleaned | HW | Fluidic pathway | Fluidic pathway |  | 24.01.2017 14:09 Ms Rosner Tel.: 03658282610 Error: tHb-Kalfehler in Kal1 and manual Kal | THb calibration error Air duct flushed | Troubleshooting Protein removal, pump calibration, LS-adjustment Oximentrie test carried out in the service program Inlet seal renewed Several 2-pcs calibration and THb calibration performed Error exists with QC measurement ==> Spare part must be ordered
No | Visit | Complaint | RDEU | ABL startet sporadisch mehrmals am Tag neu - Meldung "Netz ein" und eine kurze Init.Phase von Ca. 1,5 Minuten 13.10.2016 Herr Gelfert 09117580992880 kann den Temperaturfehler jetzt nicht mehr tollerieren. | Sporadische Abschaltung und Neustart des ABL | Tausch des AU Boards, I-Botton getauscht, Konfiguration, Roma.db und Roma.ini gesichert und nach Tausch wieder eingespielt, System Keys aufgespielt, Software aufgespielt, Update Type ausgeführt, System- und Funktionskontrolle in Ordnung | SWO-162425 | Defect | HW | Analyzer Control Module | PCB, Analyzer control |  |  | Sporadic shutdown and restart of the ABL | Exchange of the AU board, I-Botton exchanged, configuration, Roma.db and Roma.ini backed up and replayed after exchange, system keys played up, software played up, Update Type executed, system and function control in order
No | 5 | Y | RAME | F Oxi hardware problem. |  | Intermittent Oxi hardware problem on QC with error 0579. No errors when arrived. Ran Complete OXI test to isolate. Lamp unit test fails with Neon intensity too low, 34%. need to be >50%. | D1FE52AB4E322F7C | Defect | HW | Oxi module | Lamp unit | Neon | F Oxy hardware problem. |  | Intermittent Oxi hardware problem on QC with error 0579. No errors when arrived. Ran Complete OXI test to isolate. Lamp unit test fails with Neon intensity too low, 34%. Need to be > 50%.
No | Visit | Complaint | RIND | complaint about the optical sensitivity out of range in TropI | cleaned the process wheel but not solved foiund that the dark count of optical unit is above 50. so need to replace the optical unit | replaced the optical unit, now working fine and test are running smoothly. | SWO-162277 | Defect | HW | Process Module | Optical Unit | 1029 | complaint about the optical sensitivity out of range in TropI | cleaned the process wheel but not solved foiund that the dark count of optical unit is above 50. so need to replace the optical unit | replaced the optical unit, now working fine and test are running smoothly.
No | Visit | Complaint | RIND | Test not activated after cartridge inserted inside the analyser. Cartridge not scanned FTC_106 | checked and found that the alignment of cartridge barcode need to be align. also seen error FTC_106 in the activity log. | aligned the barcode scanner and check complete wet section as well as the liquid transport. the needle wash unit got blockage due to which it is showing error , cleaned it properly. now analyser working fine. | SWO-162273 | Adjustment | HW | Storage Module | Barcode reader, Storage wheel |  | Test not activated after cartridge inserted inside the analyser. Cartridge not scanned FTC_106 | checked and found that the alignment of cartridge barcode need to be align. also seen error FTC_106 in the activity log. | aligned the barcode scanner and check complete wet section as well as the liquid transport. the needle wash unit got blockage due to which it is showing error , cleaned it properly. now analyser working fine.
Yes | Visit | Complaint | RDEU | Solution pack leer | Solution pack ist leer | Erneuerung des Solutions pack | SWO-162559 | Re classified RMED | ECKHI Email 24.jan-21017 |  |  |  | Solution pack leer | Solution pack is empty | Renewal of the Solutions pack
No | Visit | Complaint | RKK | ABL90 リーク検出・圧力テストフローエラー等、エラー多数にて装置稼働せず。 | インレットプローブが完全に詰まっていた。またプローブの湾曲も確認。内部流路の汚れもひどい。 | インレットプローブ及びガスケット交換。内部流路洗浄。各流体系テスト項目チェック。 Cal $QC $測定：OK | SWO-162587 | Replaced | HW | Inlet module | Probe clot |  | ABL90 leak detection and pressure test flow error, etc., the device does not operate with a large number of errors. | The inlet probe was completely jammed. Also, the curvature of the probe was confirmed. The dirt of the internal flow path is also terrible. | Inlet probes and gasket replacements. Internal flow path cleaning. Check each flow system test item. Cal $QC $Measurement: OK
No | 10 | Y | RAME | Z Customer called the Analyzer was failing Oximetry in blood samples |  | Blood samples with error 581 Spectrum mismatch. | D1FFB319C4479EF0 | Defect | HW | Oxi module | Hemolyzer |  | Z Customer called the Analyzer was failing Oximetry in blood samples |  | Blood samples with error 581 Spectrum mismatch.
No | Visit | Complaint | RIND | IN the instrument shows Replace sensor cassette | Check and found new sensor cassette Calibration is ok but QC is not OK | In new sensor cassette in Qc level B & C Lac level is above the range.Run fluid transport and take data log for further inspection. | SWO-162858 | Assisted Replacement | SC | SC | Reference | 1361 | IN the instrument shows Replace sensor cassette | Check and found new sensor cassette Calibration is ok but QC is not OK | In new sensor cassette in Qc level B & C Lac level is above the range. Run fluid transport and take data log for further inspection.
No | Visit | Complaint | RDEU | Touchscreen lässt sich nicht bedienen und Interner Komunikationsfehker zwischen CPU und AU Unit. Gerät bleibt Offline. | ABL90 wieder auf Station installiert. | Geräteaufstellung in OP nach Reparatur in GS17. | SWO-162691 | Defect | HW | Power Supply | Power Supply |  | Touchscreen cannot be operated and Internal communication feud between CPU and AU unit. Device remains offline. | ABL90 reinstalled on station. | Equipment installation in op after repair in GS17.
No | 2 | Y | RAME | P user called stating that analyzer keeps asking to replace sp |  | found that it states unable to pump solutions | D1FFDC00C4440177 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | P user called stating that analyzer keeps asking to replace sp |  | found that it states unable to pump solutions
No | Visit | Complaint | RDEU | 25.01.2017/08:00Uhr Herr Kotz 071164893256: Das Gerät meldet regelmäßig Ableitstromfehler | Lösungspack mit Ableitstromfehler. | Lösungspack aus GS17 erneuert. | SWO-162692 | Assisted Replacement | SP | SP | Leak current |  | 25.01.2017/08:00uhr Mr. Kotz 071164893256: The device regularly reports leak-current faults | Solution pack with leak-current fault. | Solution pack from GS17 renewed.
No | 9 | Y | RAME | G After changing SP and SC multiple times electolytes, pH and COOX calirbation & QC are out of control. |  | Lot XV-15 SP caused calibration errors driving QC out of range. Instrument had sat in maintenance hold for several days without SC being cycled and needs replacement. | D1FFD13CC46384A5 | Assisted Replacement | SP | SP | FT |  | G After changing SP and SC multiple times electolytes, pH and COOX calirbation & QC are out of control. |  | Lot XV-15 SP caused calibration errors driving QC out of range. Instrument had sat in maintenance hold for several days without SC being cycled and needs replacement.
No | Visit | Complaint | RKK | 勝手にｲﾝﾚｯﾄがOPEN状態。ｲﾝﾚｯﾄPCB、ｲﾝﾚｯﾄﾓｼﾞｭｰﾙ交換。 | シリンジポジションでインレットがOPENのまま停止。 | ｲﾝﾚｯﾄ：PCB・ﾓｼﾞｭｰﾙ交換 | SWO-162316 | Defect | HW | Auto Inlet module | Inlet Position Cal. |  | The inlet is in the state without permission. Inlet PCB, inlet module replacement. | The inlet is stopped as OPEN in the syringe position. | Inlet: PCB module replacement
No | Visit | Complaint | RLTD | Mark Jones 07818 490128 | Failing Pressure tests on vacuum value. | Customer had followed guided trouble shooting and replaced inlet gasket and solution pack. I replaced inlet gasket, still failed. Took inlet from adjacent analyser, which passed. Tried another gasket, which failed. No more gaskets on site. Travelled to Glenfield to collect a gasket from their spare analyser. Now passing vacuum test but failing cthb on drift. Analyser now fully functional. | SWO-162481 | Defect | HW | Auto Inlet module | Inlet Gasket |  | Mark Jones 07818 490128 | Failing Pressure tests on vacuum value. | Customer had followed guided trouble shooting and replaced inlet gasket and solution pack. I replaced inlet gasket, still failed. Took inlet from adjacent analyser, which passed. Tried another gasket, which failed. No more gaskets on site. Travelled to Glenfield to collect a gasket from their spare analyser. Now passing vacuum test but failing cthb on drift. Analyser now fully functional.
No | 2 | Y | RCZE | Blank | Oxi problem, vid predchadzajuci swo |  | D1FFA82BC40D41ED | Defect | HW | Basic Unit | Wetsection cotrol PCB |  | Blank | Oxi problem, vid predchadzajuci swo
No | Visit | Complaint | RCHN | PCO2定标不过 | PCO2定标不过,借903-334一个 |  | SWO-162336 | [5 empty] | PCO2 calibration is not the | PCO2 calibration, however, borrows 903-334 a
No | Visit | Complaint | RDEU | 24.01.2017 13:00 Fr. Daub 06233317220: QC-Probleme pH, pCO2 24.01.2017 13:50 Frau Daub Tel.: 06233317218 Fehlermeldung: Fließwege verstopft (Kundin war sehr erregt) | QC Verletzungen und sehr häufiger Betriebsmitteltausch | Jahreswartung wird vorgezogen (siehe SWO) | SWO-162423 | [5 empty] | 24.01.2017 13:00 Fr. Daub 06233317220: QC problems pH, pCO2 24.01.2017 13:50 Ms. Daub Tel.: 06233317218 Error message: Flow paths clogged (customer was very excited) | QC Injuries and very frequent equipment exchange | Annual maintenance is brought forward (see SWO)
No | Visit | Complaint | RCHN | 报1024错误代码 | 报错1024，之前借的光单903-364有问题，现重新借903-364，TNI定标卡944-212（12269）个一个 |  | SWO-162385 | Defect | HW | Process Module | Optical Unit | 1024 | Reported 1024 error code | 1024, previously borrowed light sheet 903-364 problem, now re-borrow903-364, TNI calibration card 944-212 (12269) one
No | Visit | Complaint | RCHN | 无法读取测试卡 | 无法读取测试卡，借910-304一个 |  | SWO-162330 | Defect | HW | Storage Module | Barcode reader, Storage wheel |  | The test card could not be read | Failed to read test card, borrow 910-304 one
No | Visit | Complaint | RLTD | Tiana 02082164934 | Not running QC | Cleaned dust from unit and test ok. | SWO-162454 | Cleaned | HW | Auto Check Module | Barcode Reader |  | Tiana 02082164934 | Not running QC | Cleaned dust from unit and test ok.
No | Visit | Complaint | RDEU | 25.01.2017; Fr. Keil; Tel.: 0626183311; AC Karussell hängt fest | Autocheck Karussell lässt sich nicht manuell aus dem Modul entfernen. Ausbau notwendig. | Ausbau des Autocheck Moduls, auseinander bauen. Motor von Karussell Träger gelöst. Autocheck Karussell mit Antriebsachse verklebt, gelöst und gründlich gereinigt. Autocheck Modul zusammengebaut, Funktion in Ordnung. | SWO-162563 | Cleaned | HW | Auto Check Module | caroussel |  | 25.01.2017; Ms. Keil; Tel.: 0626183311; AC carousel hangs tightly | Autocheck Carousel cannot be manually removed from the module. Expansion necessary. | Expansion of the Autocheck module, disassemble. Motor detached from carousel carrier. Autocheck carousel with drive axle glued, dissolved and thoroughly cleaned. Autocheck module assembled, function fine.
No | 6 | Y | RAME | P WASTE NIPPLE CLOGGED |  | CLOG REQUIRES REMOVAL AND FLUSHHING | D20165C73A54066A | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | P WASTE NIPPLE CLOGGED |  | CLOG REQUIRES REMOVAL AND FLUSHHING
No | Visit | Complaint | RDEU | 09.01.2017/14:00, Herr Esche: Lac. Kal. Fehler nach Membranierung 10.01.2017, 8.15 Uhr Herr Esche, Tel.: 089857915954 : Weiterhin Kalibrationsfehler bei Laktat | erneut sporadisch Lac Driften | Proteinentfernung durchgeführt - Messkammer gereinigt Lactat Elektrode erneuert Lac Vorverstärkernullpunkt kalibriert Luftweg im Fluidic Modul gespült Ventile im pH $BG sowie El.Met. Modul überprüft - ok Gerät mehrfach kalibriert, Drift bei Lactat 0,0 QC Messungen durchgeführt | SWO-162549 | Assisted Replacement | Electrode | Electrode | Lac |  | 09.01.2017/14:00, Mr Esche: Lac. Kal. Errors after membraneing 10.01.2017, 8.15 a.m. Mr. Esche, Tel.: 089857915954 : Further calibration error in lactate | again sporadic Lac Driften | Protein removal performed - measuring chamber cleaned Lactate electrode renewed Lac preamplifier zero point calibrated Airway flushed in fluidic module Valves in pH $BG and El.Met. Module checked - ok Device calibrated multiple times, drift at lactate 0.0 QC measurements performed
No | Visit | Complaint | RDEU | Inlet undicht | Inletmodul undicht | Original Inlet Modul zurück gebaut, Modulventil V1-V4 demontiert und überholt, Funktionskontrolle ok | SWO-162776 | Unknown | HW | Inlet module | Leakage |  | Inlet leaks | Inlet module leaks | Original Inlet module built back, module valve V1-V4 disassembled and overhauled, function control ok
No | Visit | Complaint | RKK | AQTのインレットポート右扉の動作不良が発生。自動では完全に閉じないため、エラーとなる。代替えのフロントパネルに入れ替え、持ち帰る。 | AQTのインレットポート右扉の動作不良が発生。自動では完全に閉じないため、エラーとなる. | 代替えのフロントパネルに入れ替え、持ち帰る。 | SWO-163128 | Defect | HW | Cover and Chassis | Inlet Door |  | AQT's inlet port right door malfunctioned. It is an error because it does not close completely automatically. Replace with an alternate front panel and take it home. | AQT's inlet port right door malfunctioned. Automatic ally will not close completely and will result in an error. | Replace with an alternate front panel and take it home.
Yes | Visit | Complaint | RDEU | Kein Solutionpack lieferbar, Gerät steht | No | Solutionpack geliefert und eingesetzt, Kalibration, QC- Kontrolle | SWO-162705 | Re classified RMED | Logistical upSWO |  |  |  | No solution pack available, device is available | No | Solutionpack delivered and used, calibration, QC control
No | 4 | Y | RAME | F Spectrum error on blood samples |  | Intermittent Spectrum 581 erros on blood samples. | D201161A3A379EFA | Defect | HW | Oxi module | Hemolyzer |  | F Spectrum error on blood samples |  | Intermittent Spectrum 581 erros on blood samples.
No | 6 | Y | RAME | V They said the printer and touch screen would not work locked up |  | The touch screen was out of cal | D200D5D43A3913C7 | Adjustment | Touch screen | DMS_PC | Touch screen | Cal | V They said the printer and touch screen would not work locked up |  | The touch screen was out of cal
No | Visit | Complaint | RSCH | Blank | Instabilité divers paramètres lors de la mesure sang complet | Nettoyage de la section fluidique e.o Modification du setup e.o Calibration e.o contrôle de qualité e.o | SWO-162704 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | Blank | Instability various parameters when measuring full blood | Cleaning the fluidsection e.o Changing the setup e.o Calibration e.o quality control e.o
No | 6 | Y | RAME | 2 PCO2 High on S7430 control after solution pack and sensor replacement |  | checked data logs for errors, checked QC log, Red qc upper range is 69 instrument reading 71-74mmHg | D20152713A0AFD02 | [5 empty] | 2 PCO2 High on S7430 control after solution pack and sensor replacement |  | checked data logs for errors, checked QC log, Red qc upper range is 69 instrument reading 71-74mmHg
No | Visit | Complaint | RIND | Gone to provide service as the analyzer was unable to read the barcode of the consumable so cleaned the barcode but still was unable to read the code so ordered for new barcode. And now waiting for the part. | Found problem with barcode so ordered for new barcode |  | SWO-162843 | Revisit required | HW | Basic unit | Barcode reader, Basic Unit |  | Gone to provide service as the analyzer was unable to read the barcode of the consumable so cleaned the barcode but still was unable to read the code so ordered for new barcode. And now waiting for the part. | Found problem with barcode so ordered for new barcode
No | Visit | Complaint | RKK | FLEXモジュールの動作不良。2番と3番のスロットにおいてシリンジを置いても認識しない状況。画面上のリアクションはなかった。また、シリンジの2度置きによるものでもなかった。 |  |  | SWO-162589 | Unknown | HW | FlexQ Module | Barcode Reader |  | Flex module malfunctioned. A situation where the syringe is not recognized even if the syringe is placed in slots 2 and 3. There was no reaction on the screen. In addition, it was not due to the place of the syringe twice.
No | Visit | Complaint | RDEU | 26.01.2017, Frau Domberg, tel.. 05115322529 : Leckage erkannt | Leckage erkannt | Gerät geprüft, Messkammer geprüft und gereinigt, Funktionskontrolle ok | SWO-162809 | Cleaned | HW | pH/BG Module | Measuring chamber |  | 26.01.2017, Mrs. Domberg, tel.. 05115322529 : Leakage detected | Leakage detected | Device tested, measuring chamber tested and cleaned, function control ok
No | Visit | Complaint | RDEU | pCO2 Ergebnisse in AC Level 3 springen zickzackartig und regelmäßig aus der Untergrenze mit regelmäßig nötigen QC-Zweitmessungen | Diesmal pCO2 Werte oberhalb RiLiBÄK Grenze | pCO2 Wert in AC5+ L3 zwei mal zu hoch gemessen mit zwischenzeitlicher Reinigung $ Kal 2; Messkammern überprüft und leicht blutige Ablagerungen vor Ref. entfernt. pCO2 Elektrode gereinigt und neu membraniert, Kontakte geputzt. Kalibrationen, AC5+ Messung => OK | SWO-162659 | Cleaned | HW | Fluidic pathway | Fluidic pathway | pCO2 | pCO2 results in AC Level 3 jump zigzag-like and regularly from the lower limit with regularly required QC second measurements | This time pCO2 values above RiLiBÄK limit | pCO2 value in AC5+ L3 measured twice too high with intermediate cleaning . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . measuring chambers and slightly bloody deposits before ref. away. pCO2 electrode cleaned and re-membraneed, contacts cleaned. Calibrations, AC5+ measurement => OK
No | 4 | Y | RAME | W DMS not connected to AU error. |  | This happened after a reboot. All reboots after that would not resolve issue. | D200F80B3A0B4FC6 | Defect | SW | Win 7 | Re-Install |  | W DMS not connected to AU error. |  | This happened after a reboot. All reboots after that would not resolve issue.
No | Visit | Complaint | RDEU | 27.01.2017 8:00 H. Putz 03696137433: LP aus Expresslieferung Fehlerhaft? | Kalibrierwerte nach Erneuerung SP Bei Lc und Glu fehlerhaft | Fehlersuche SP ist bereits seit 24 h installiert, mehrere Kal sind gelaufen, Statuswerte zu hoch Reinitialisierung alter SP,keine Besserung Austausch SP, Fehler bleibt, StatuswerteGlu, Lac zu hoch Wechsel SC SP und SC zur Reklamationsmeldungen an RDeu | SWO-163031 | Defect | SC | SC | Glu/Lac |  | 27.01.2017 8:00 H. Putz 03696137433: LP from Express Delivery Faulty? | Calibration values after renewal SP Faulty at Lc and Glu | Troubleshooting SP has been installed since 24 h, several Kal have run, Status values too high Reinitialization of old SP, no improvement Exchange SP, error remains, status valuesGlu, Lac too high Change SC SP and SC to report complaints to RDeu
No | 19 | Y | RAME | Z Customer reporting problem during Sensor Cassette replacement |  | While customer was replacing SC analyzer appears to have froze, also customer logon failed to work. Suspect customer rushed SC installation and did not wait for prompt for next steps causing analyzer to stall. Also customer tried to insatll an expired SC. | D201913FB01C27A4 | User error | SC installation | Expired SC | Expired SC |  | Z Customer reporting problem during Sensor Cassette replacement |  | While customer was replacing SC analyzer appears to have froze, also customer logon failed to work. Suspect customer rushed SC installation and did not wait for prompt for next steps causing analyzer to stall. Also customer tried to insatll an expired SC.
No | Visit | Complaint | RDEU | 27.01.2017 Frau Glöß 0302506428 meldet einen Leckagefehler. | Gerät steht im Interventionsmodus | Fehlersuche, Einlaßdichtung+ LP wurde bereits ersetzt, Abfallwege gespült - speziell Hemolysator, Einlaß gespült, Servicetests durchgeführt, Vakuumfehler behoben, Neustartphase abgewartet, QC | SWO-163083 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | 27.01.2017 Ms. Glöß 0302506428 reports a leakage error. | Device is in intervention mode | Troubleshooting, inlet seal+ LP has already been replaced, waste paths flushed - especially hemolysator, inlet flushed, service tests performed, vacuum error fixed, restart phase awaited, QC
Yes | 22 | Y | RAME | Z Customer reporting error 734 - General WSM exception. Also mentioned other errors but wasn't able to accurately describe them during our call. |  | Apparently this user came in this morning to the analyzer displaying the error which occured overnight. No calls overnight from the overnight shift reporting any issues. | D201ED43B04BE665 | Re classified RMED | Phone support |  |  |  | Z Customer reporting error 734 - General WSM exception. Also mentioned other errors but wasn't able to accurately describe them during our call. |  | Apparently this user came in this morning to the analyzer displaying the error which occured overnight. No calls overnight from the overnight shift reporting any issues.
Yes | Visit | Complaint | RDEU | keine Lieferung von LP erhalten | keine Lieferung von LP erhalten Gerät steht | LP geliefert und eingesetzt Startphase System IO | SWO-163010 | Re classified RMED | Logistical upSWO |  |  |  | No delivery from LP | no delivery of LP received device is available | LP supplied and used start-up phase system IO
Yes | Visit | Complaint | RDEU | keine Lieferung von LP erhalten | keine Lieferung von LP pünktlich erhalten | LP geliefert | SWO-163012 | Re classified RMED | Logistical upSWO |  |  |  | No delivery from LP | No delivery of LP received on time | LP supplied
No | Visit | Complaint | RLTD | Mark Jones 07818 490128 | Analyser failing on all oximetry parameters. | Ran complete oximetry tests all passing. Replaced pinch valve tube which solved the issue. Passed calibration and QCs straight away. | SWO-163124 | Defect | HW | Oxi module | Pinch valve tube |  | Mark Jones 07818 490128 | Analyser failing on all oximetry parameters. | Ran complete oximetry tests all passing. Replaced pinch valve tube which solved the issue. Passed calibration and QCs straight away.
No | Visit | Complaint | RCHN | PCO2定标不过 | PCO2定标不过 | 因前一个SWO借件发货有问题 ，重开一个。 | SWO-163161 | [5 empty] | PCO2 calibration is not the | PCO2 calibration is not the | Re-open one due to a problem with the previous SWO debit shipment.
No | 5 | Y | RAME | Z No display |  | Telephone call indicated symptoms of failed CF card. Arrived and observed sound at power up, but no display | D2026453B0034415 | [5 empty] | Z No display |  | Telephone call indicated symptoms of failed CF card. Arrived and observed sound at power up, but no display
No | Visit | Complaint | RDEU | Gas2 Druck zu niedrig trotz neuer Gasflasche Auftragsnummer A303/17-002087 | Gas2 Druck zu niedrig trotz neuer Gasflasche , Gas1 normal | Fehlersuche, Gas2 Regelung getestet, Drucktest negativ, Ventil durchgespült, Anschluß erneuert, Drucktest wiederholt - i.O. , Regelung getestet - i.O., Neustart, Kal, QC | SWO-163086 | Cleaned | HW | Basic Unit | Gas 2, valve |  | Gas2 pressure too low despite new gas cylinder Number of mandate A303/17-002087 | Gas2 pressure too low despite new gas cylinder, Gas1 normal | Troubleshooting, Gas2 control tested, pressure test negative, valve flushed through, connection renewed, pressure test repeated - i.O., control tested - i.O., restart, Kal, QC
No | 9 | Y | RPAC | Blank | Aborting QC |  | D20220F1B0283D08 | Defect | HW | Inlet module | Probe bend |  | Blank | Aborting QC
No | Visit | Complaint | RDEU | Nach HJW konstant Einlass-Kalibrierfehler | s. case | Fehler aufgetreten nach HJW => neue Einlassdichtung benetzt, ohne Erfolg. Servicemenü aufgerufen und mehrfach Einlass kalibriert, o.E. Nadelantrieb manuell bewegt, o.E. Nadel überprüft, sie ist vollständig gerade und ohne Verschleißmarken. Weitere neue Einlassdichtung eingesetzt, o.E. Einlassnadel ersetzt, nun sofort Einlass-Kal bestanden => erste Einlassdichtung (aus W-Kit) eingesetzt, ebenfalls problemlos. Weiteres Verfahren s. HJW-Case | SWO-163044 | Defect | HW | Inlet module | Needle |  | After HJW constant inlet calibration error | s. case | Error occurred after HJW => new inlet seal wetted, to no avail. Service menu called and repeatedly calibrated inlet, o.E. needle drive manually moved, o.E. needle checked, it is completely straight and without wear marks. Further new inlet seal inserted, o.E. inlet needle replaced, now immediately inlet-scale passed => first inlet seal (made of W-Kit) inserted, also without problems. Further procedure see HJW-Case
No | Visit | Complaint | RDEU | 24.01.2017 14:09 Frau Rosner Tel.: 03658282610 Fehler: tHb-Kalfehler in Kal1 und manueller Kal | Fehler 0581: Oxi spectrum mismatch bei THb-Kalibrierung und QC-Messungen | OXI-Modul ausgetauscht S $N entnommenes OXI-Modul: 902-833R2034N002 THb-Kalibrierung | SWO-163056 | Defect | HW | Oxi module | Spectrometer | Spectrum mismatch | 24.01.2017 14:09 Ms Rosner Tel.: 03658282610 Error: tHb-Kalfehler in Kal1 and manual Kal | Error 0581: Oxi spectrum mismatch for THb calibration and QC measurements | OXI module replaced S $N oxI module taken from the us: 902-833R2034N002 THb calibration
No | Visit | Complaint | RDEU | Laktatfehler, kurze Standzeit | Laktatfehler, kurze Standzeit | Daten überprüft, Messkammer gereinigt und neue Membranen bestellen lassen Gerät kalibriert und mit QC überprüft | SWO-163117 | Assisted Replacement | Electrode | Membrane | Lac | Short lifetime | Lactate error, short service life | Lactate error, short service life | Data checked, measuring chamber cleaned and new membranes ordered Device calibrated and checked with QC
No | Visit | Complaint | RSAF | Inlet module not detected. | Tested all inlet hardware. | Replaced wet section control PCB. Calibrated inlet module. | SWO-163259 | Defect | HW |  |  |  | Inlet module not detected. | Tested all inlet hardware. | Replaced wet section control PCB. Calibrated inlet module.
No | Visit | Complaint | RCHN | 检测失败 0298，0581 | 检测失败 0298，0581，借902-980一个。 |  | SWO-163138 | Defect | HW | Oxi module | Hemolyzer |  | Detection failure 0298,0581 | Detection failure 0298,0581, borrowing 902-980 one.
No | Visit | Complaint | RCHN | PCO2定标失败 1316，0089，0378 | PCO2定标失败 1316，0089，0378 | 更换测试卡正常。 | SWO-163139 | [5 empty] | PCO2 calibration failure 1316,0089,0378 | PCO2 calibration failure 1316,0089,0378 | The replacement test card is OK.
No | Visit | Complaint | RCHN | TNI结果偏高 | TNI结果偏高 | 更换903-119一个正常。 by andrew | SWO-163165 | Defect | HW | Process Module | Cup wash unit |  | TNI results are high | TNI results are high | Replace the 903-119 one normal. by andrew
No | 12 | Y | RESP | Resultado discrepante. Falso positivo |  | sustituida process Wheel y filtro LS2. Pasados QCs y test de blancos. Se abre techline | D203ACB12625FB7A | Defect | HW | Process Module | Process wheel |  | Dissenting result. False positive |  | replaced process Wheel and LS2 filter. Past QCs and white tests. Techline opens
No | Visit | Complaint | RIND | TNT Calibration Failed, After Analyzer Showing Message Optical intensity out of range, Calibration Adjustment Analyte Outliers check failed,Optical movement Failure, Cleaned Process Wheel, Tried to reinstalled Optical Unit, Cleaned Analyzer, Reinstalled Motor, Tried to Resolve, Problem Found with Optical Unit, process wheel and Motor, need spares for Testing, waiting for spares. | TNT Calibration Failure. Optical Unit Movement Failure, Optical intensity out of range, Calibration Adjustment Analyte Outliers check failed. | Cleaned Process Wheel, Tried to reinstalled Optical Unit, Cleaned Analyzer, Reinstalled Motor, Tried to Resolve, Problem Found with Optical Unit, process wheel and Motor, need spares for Testing, waiting for spares. | SWO-163133 | Revisit required | HW | Process Module | Optical Unit |  | TNT Calibration Failed, After Analyzer Showing Message Optical intensity out of range, Calibration Adjustment Analyte Outliers check failed,Optical movement Failure, Cleaned Process Wheel, Tried to reinstalled Optical Unit, Cleaned Analyzer, Reinstalled Motor, Tried to Resolve, Problem Found with Optical Unit, process wheel and Motor, need spares for Testing, waiting for spares. | TNT Calibration Failure. Optical Unit Movement Failure, Optical intensity out of range, Calibration Adjustment Analyte Outliers check failed. | Cleaned Process Wheel, Tried to reinstalled Optical Unit, Cleaned Analyzer, Reinstalled Motor, Tried to Resolve, Problem Found with Optical Unit, process wheel and Motor, need spares for Testing, waiting for spares.
No | Visit | Complaint | RKK | ＳＰ不良クレーム | SP不良にて測定不能 | ＳＰ交換 | SWO-163320 | Assisted Replacement | SP | SP | SP |  | SP Fault Claims | Sp failure and measurement cannot be | SP Exchange
Yes | 18 | Y | RPAC | Blank | Solution pack |  | D207299B88293C0D | Assisted Replacement | SP | SP | Leak current |  | Blank | Solution pack
Yes | 6 | Y | RAME | P Flow selector errors |  | 1177 flow selector errors. Customer replaced solution pack and got the same result. Got a couple of 1177 errors while checking after maintenance performed. Unit was moved and had discharged over the week end. | D2062B9612223188 | Defect | HW | Basic Unit | PCB, Flow Selector |  | P Flow selector errors |  | 1177 flow selector errors. Customer replaced solution pack and got the same result. Got a couple of 1177 errors while checking after maintenance performed. Unit was moved and had discharged over the week end.
Yes | 12 | Y | RPAC | Blank | Sensor Cassette failure pO2 |  | D207154A881122BF | Re classified RMED | Logistical upSWO |  |  |  | Blank | Sensor Cassette failure pO2
No | Visit | Complaint | RDEU | 27.01.2017/16:00Uhr Frau Ütmü 062213629165 oder 015147127161: Das Gerät wurde im Oktober 2016 geliefert und war zwei Tage in Betrieb. Da das Gerät vor Ort nicht reparierbar war wurde es ausgetauscht. Das Gerät kam am Vortag wieder zurück und war seit dem nicht meßbereit. Das Gerät meldet einen Kalibrationsfehler bei Glu 1225. | Glucose Kalibration gescheitert. Analysator zeigt Meldung Mehr Zeit für weitere Hydrierung zulassen oder Sensorkassette erneuern. SC ist aber schon über ein Tag im Analysator. Bereits die zweite SC mit gleichem Fehler innerhalb kurzer Zeit. | Nullpunktkalibrierung mit Zeroing Kassette durchgeführt. Werte sind innerhalb Grenzwerte nach Manuel. Erneute Kalibration mit Qualitätskontrolle brauchte kein Erfolg. Leihgerät aus Geschäftsstelle mitgenommen zum Test, ob Betriebsmittel im Leihgerät sich gleich verhalten. LP im Anwender Gerät zu wenig Messzyklen. Anwender haben keinen Ersatz LP. Messzyklen im Leihgerät zu wenig um zu tauschen. Leihgerät wieder mitgenommen. Anwender wird den LP tauschen, wenn geliefert und die SC. Anwender meldet sich, wenn weiterhin Qualitätskontroll Grenzverletung vorhanden. | SWO-163400 | [5 empty] | 27.01.2017/16:00h Mrs. Ütmü 062213629165 or 015147127161: The device was delivered in October 2016 and was in operation for two days. Since the device could not be repaired on site, it was replaced. The device returned the day before and had not been ready for measurement since then. The device reports a calibration error with Glu 1225. | Glucose calibration failed. Analyzer displays message Allow more time for further hydrogenation or renew sensor cartridge. BUT SC has been in the analyzer for over a day. Already the second SC with the same error in a short time. | Zero-point calibration performed with zeroing cassette. Values are within limits according to Manuel. Renewing calibration with quality control did not require success. Loaner taken from the office for testing whether equipment in the rental device behaves the same. LP in the user device too few measuring cycles. Users do not have a replacement LP. measuring cycles in the rental device too little to exchange. borrower. User will exchange the LP when delivered and the SC. User reports if quality control border vulnerability is still present.
No | 5 | Y | RAME | Z Leaking/Blood clot |  | sample errors (0328) and leaking at inlet capillary. | D20665BD1227A4F3 | Defect | HW | Inlet module | Inlet Gasket |  | Z Leaking/Blood clot |  | sample errors (0328) and leaking at inlet capillary.
No | Visit | Complaint | RDEU | Bildschirm friert häufig ein. | Bildschirm friert häufig ein. | Monitor erneuert. Kalibriert und Töne konfiguriert. | SWO-163367 | Defect | HW | DMS_PC | Touch screen |  | Screen often freezes. | Screen often freezes. | Monitor renewed. Calibrated and configured totones.
No | 3 | Y | RAME | Z Customer reported AutoCheck mechanical errors. |  | AutoCheck unit errors intermittently during auto QC runs. Found breaker arm not returning to full return position. | D20648A61202ABB4 | Adjustment | HW | Auto Check Module | Breaker Motor |  | Z Customer reported AutoCheck mechanical errors. |  | AutoCheck unit errors intermittently during auto QC runs. Found breaker arm not returning to full return position.
No | Visit | Complaint | RIND | TNT Calibration Failed, After Analyzer Showing Message Optical intensity out of range, Calibration Adjustment Analyte Outliers check failed,Optical movement Failure, Cleaned Process Wheel, Tried to reinstalled Optical Unit, Cleaned Analyzer, Reinstalled Motor, Tried to Resolve, Problem Found with Optical Unit, process wheel and Motor, need spares for Testing, waiting for spares. | For Replacement for Optical Unit. | Replaced Optical Unit with New One, Cleaned Process wheel, Reinstalled Process Unit, Problem Found In Stepper Motor while Movement for the Optical Process, Ordered Stepper Motor for Replacement. | SWO-164636 | Defect | HW | Process Module | Optical Unit | 1031 | TNT Calibration Failed, After Analyzer Showing Message Optical intensity out of range, Calibration Adjustment Analyte Outliers check failed,Optical movement Failure, Cleaned Process Wheel, Tried to reinstalled Optical Unit, Cleaned Analyzer, Reinstalled Motor, Tried to Resolve, Problem Found with Optical Unit, process wheel and Motor, need spares for Testing, waiting for spares. | For Replacement for Optical Unit. | Replaced Optical Unit with New One, Cleaned Process wheel, Reinstalled Process Unit, Problem Found In Stepper Motor while Movement for the Optical Process, Ordered Stepper Motor for Replacement.
No | Visit | Complaint | RKOR | Sensor cassette error | . | Sensor cassette replace. Gasket module clean. Solution pack connecter clean. Calibration & Control check. | SWO-163490 | Assisted Replacement | SC | SC | SC |  | Sensor cassette error | . | Sensor cassette replace. Gasket module clean. Solution pack connecter clean. Calibration & Control check.
No | Visit | Complaint | RDEU | 30.01.2017/12:30Uhr Frau Watmunt 03957752405: QC-Fehler bei pH und pCO2 bei Level A und C | YB - 50 LP in Verwendung QC Lvl C pCO2 und pH Fehler - > reklamiert | LP ersetzt - QC i.O. ein LP YB-50 XL zurückgenommen, da er sehr wahrscheinlich nicht funktionieren wird | SWO-163607 | Assisted Replacement | SP | SP | pCO2 | YB50 | 30.01.2017/12:30Uhr Frau Watmunt 03957752405: QC-Fehler bei pH und pCO2 bei Level A und C | YB - 50 LP in use QC Lvl C pCO2 and pH error - > complained | LP replaced - QC in the world an LP YB-50 XL is taken back as it will most likely not work
No | 2 | Y | RFRA | Blank | Erreur Inlet |  | D208A5ABFE12AD42 | Defect | HW | Auto Inlet module | Auto Inlet module |  | Blank | Inlet error
No | Visit | Complaint | RLTD | After rebooting during training the analyser is just saying Offline in the top left hand corner. | While training was being given on replacing consumables the analyser had to be rebooted and after that it came up saying offline and it was unable to log on to it. I checked the wet section control PCB and there were no obvious problems. The sample mixer & battery modules were disconnected to see if that helped but it didn't. | It looks like the wet section is not communicating with PC module, I will return with a new PC Module, Power Supply & Wet Section Control PCB to see which part is at fault. | SWO-163706 | Revisit required | SW | Offline | Offline |  | After rebooting during training the analyser is just saying Offline in the top left hand corner. | While training was being given on replacing consumables the analyser had to be rebooted and after that it came up saying offline and it was unable to log on to it. I checked the wet section control PCB and there were no obvious problems. The sample mixer & battery modules were disconnected to see if that helped but it didn't. | It looks like the wet section is not communicating with PC module, I will return with a new PC Module, Power Supply & Wet Section Control PCB to see which part is at fault.
No | Visit | Complaint | RKOR | SP reading error, over flow | . | Over flow check. Waste manifold line leak. Clean & check. I will replace it the next time a problem. Data log clean. | SWO-163489 | [5 empty] | SP reading error, over flow | . | Over flow check. Waste manifold line leak. Clean & check. I will replace it the next time a problem. Data log clean.
No | 7 | Y | RAME | Z blank screen |  | Troubleshot to bad CPU which was ordered | D2078D0E885CC6A9 | [5 empty] | Z blank screen |  | Troubleshot to bad CPU which was ordered
No | 3 | Y | RESP | SE HA DESCONECTADO DE AQURE |  | CHEQUEO DE CONEXIONES Y CONFIGURACIÓN DEL EQUIPO. TRAS ABRIR INCIDENCIA EN INFORMÁTICA DE HOSPITAL, SE SOLUCIONA JUNTO CON PERSONAL DE IT | D20753C988168D95 | Setup | His/lis | His/lis | His/lis |  | HAS BEEN DISCONNECTED FROM AQURE |  | CHECK OF CONNECTIONS AND EQUIPMENT SETTINGS. AFTER OPENING INCIDENT IN HOSPITAL INFORMATION TECHNOLOGY, IT IS SOLVED TOGETHER WITH IT STAFF
No | Visit | Complaint | RDEU | kritischer Systemfehler | System zeigt kritischen Systemfehler und ist nicht Messbereit Button zum löschen der letzten Daten ist gewählt aber trotzdem sperrt das System des Weiteren ist der ABL nicht mehr mit liveconnect verbunden | alle Daten auf ein USB Stick verschoben und Neustart durch geführt, System wieder messbereit Daten Ingo Sprenger zur Kontrolle geschickt | SWO-163691 | [5 empty] | critical system error | System shows critical system errors and is not ready to measure Button to delete the last data is selected but still the system locks furthermore, the ABL is no longer connected to liveconnect | all data moved to a USB stick and rebooted, system ready to be measured again Data Sent to Ingo Sprenger for control
No | Visit | Complaint | RIND | crea calibration expire. | check the logs. | replace the crea reagent. calibrate system multiple time.calibration pass & qc crea pass. | SWO-164608 | Assisted Replacement | Consumables | Solution bottle | Crea |  | crea calibration expire. | check the logs. | replace the crea reagent. calibrate system multiple time.calibration pass & qc crea pass.
Yes | 39 | Y | RESP | Cliente informa electrodo roto de pH |  | Comprobación del estado del analizador, telefónicamente cliente informa que el electrodo de pH esta roto. Se envia electrodo nuevo para su reposición. | D20EB7664C1479F9 | Re classified RMED | Telph support |  |  |  | Customer reports broken pH electrode |  | Checking the status of the analyzer, telephone client informs that the pH electrode is broken. New electrode is sent for replacement.
Yes | Visit | Complaint | RDEU | 31.01.2017; Fr. Johnson; Tel.: 061157715490; Schlauch undicht/abgesprungen |  |  | SWO-163676 | Re classified RMED | Telph support |  |  |  | 31.01.2017; Ms. Johnson; Tel.: 061157715490; Hose leaking/bounced
No | Visit | Complaint | RKK | ネオンランプ電圧が低い | オキシ項目 ・・・・ | オキシモジュール交換 | SWO-163746 | Defect | HW | Oxi module | Lamp unit | Neon low | Low neon lamp voltage | Oxy item ... | Oxy module replacement
No | Visit | Complaint | RDEU | OSM bei QC Messungen, vier mal in Folge ohne äußeren Einfluss | s. case | Ohne äußeren Einfluss vier QC Messungen mit OSM => Proteinentfernung, Kal und tHb-Kal => OSM. OxiCompleteTest im Servicemenü, alle Werte nahezu mittig und OK. Mehrere Kalibrationen und erneute tHb-Kal, OK. Endkontrolle: 3 AC5+ Messungen => Alle OK | SWO-163789 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway | tHb | OSM in QC measurements, four times in a row without external influence | s. case | Without external influence four QC measurements with OSM => protein removal, kal and tHb-kal => OSM. OxiCompleteTest in the service menu, all values almost central and OK. Multiple calibrations and re-calibration tHb-Kal, OK. Final control: 3 AC5+ measurements => All OK
No | Visit | Complaint | RDEU | 31.01.2017/09:00Uhr Herr Döhler 017615323439: QC-Verletzung im Level3 bei tHb | QC-Verletzung im Level3 bei tHb, tHb Kal. notwendig. | Reinigung, Proteinentfernung, 1.Kal. und tHb Kal. durchgeführt. tHb Test und QCs i.O. | SWO-163689 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway | tHb | 31.01.2017/09:00am Mr Döhler 017615323439: QC violation at level 3 at tHb | QC-Verletzung im Level3 bei tHb, tHb Kal. notwendig. | Reinigung, Proteinentfernung, 1.Kal. und tHb Kal. durchgeführt. tHb Test und QCs i.O.
No | 3 | Y | RAME | F OXI hardware error. |  | Neon lamp intensity and neon voltage are out of range. | D20747F0884663D9 | Defect | HW | Oxi module | Lamp unit |  | F OXI hardware error. |  | Neon lamp intensity and neon voltage are out of range.
No | Visit | Complaint | RDAN | Ingen forbindelse til AU | Software genindlæst AU board forsøgt skiftet Kabel til AU board forsøgt skiftet | pH $BG modul PCB skiftet | SWO-163289 | Defect | HW | pH/BG Module | PCB, ph/BG Module | Afledt fejl på Win | No connection to AU | Software Reloaded AU board tried the shift AU Board Cable tried the switch | pH $BG modul PCB skiftet
No | Visit | Complaint | RDEU | Kundin hat bereits das Gerät gespült und den Einlass gereinigt. | Das gerät lässt sich wider Spülen, noch saugen, die graue dichtung ist komplett verstopft und die sonde ist verbogen. Durch Visual Flow Check war zu sehen, dass das Gerät keine Lösung durch kaufen lässt. |  | SWO-164038 | Replaced | HW | Inlet module | Probe bend |  | Customer has already rinsed the device and cleaned the inlet. | This device can be resisted, still sucking, the grey seal is completely clogged and the probe is bent. Visual Flow Check showed that the device doesn't let a solution through.
No | Visit | Complaint | RDEU | 01.02.2017/11:20Uhr Frau Mai 02065972808: Nach dem LP-Tausch geht das Gerät in den Fehlerbehebungsmodus. Alle Maßnahmen waren erfolglos. | Vaccumfehler | Anwender hat LP bereits erneuert. Startphase im Serviceprogramm wiederholt, Funktion OK. Initialisierung durchgeführt, OK. | SWO-164073 | No fault found | SP | SP | Vaccum error |  | 01.02.2017/11:20pm Mrs May 02065972808: After the LP exchange the device goes into troubleshooting mode. All measures were unsuccessful. | Vaccumfehler | User has already renewed LP. Repeated start-up phase in service program, function OK. Initialization performed, OK.
No | Visit | Complaint | RLTD | LC IT connection | LC IT connection set up. | LC IT connection set up. Local IT changed IT settings on system. Static IP address assigned to ABL90. Set up and tested. Connection set up and working. | SWO-166659 | Setup | His/lis | His/lis | His/lis |  | LC IT connection | LC IT connection set up. | LC IT connection set up. Local IT changed IT settings on system. Static IP address assigned to ABL90. Set up and tested. Connection set up and working.
No | Visit | Complaint | RLTD | Ranzir on 01865 221596 | QCs still failing on Oximetery parameters. The solution pack has been changed and a tHb calibration has been run. | Replaced the hemolyser and ran a tHb calibration. Also while on site I cleaned up the old ABL827s from ED Resus & EAU and then took them back to Radiometer. | SWO-163981 | Defect | HW | Oxi module | Hemolyzer |  | Ranzir on 01865 221596 | QCs still failing on Oximetery parameters. The solution pack has been changed and a tHb calibration has been run. | Replaced the hemolyser and ran a tHb calibration. Also while on site I cleaned up the old ABL827s from ED Resus & EAU and then took them back to Radiometer.
No | Visit | Complaint | RSAF | After system check analyzer enter intervention required | 3 SP replaced with leak currents still detected. Same errors as original installation. UPS was disconnected - ruled out, Mains power supply from Ward UPS suspected. | Changed power cables and moved plug from UPS socket to normal power socket. | SWO-164196 | Unknown | HW | SP | Leak current |  | After system check analyzer enter intervention required | 3 SP replaced with leak currents still detected. Same errors as original installation. UPS was disconnected - ruled out, Mains power supply from Ward UPS suspected. | Changed power cables and moved plug from UPS socket to normal power socket.
No | Visit | Complaint | RDEU | 31.01.2017, Pfl Matthias: AQC Fehler bei pCO2 in Level a+C 31.01.2017 Fr. Preikschas 023815891214: Problem bleibt weiterhin bestehen | - QC Fehler Level A,C - PH / pCO2 - inkonsistente Datenbank | Gerät geprüft. Wegen inkonsistente Datenbank Fehler Neustart durchgeführt. Angelieferten Lösungspack nach Startphase installiert. Neuer Lösungspack bringt nun pCO2 Kalfehler! Diverse Kal und QC durchgeführt. | SWO-164072 | Assisted Replacement | SP | SP | pCO2 |  | 31.01.2017, Pfl Matthias: AQC error at pCO2 in level a+C 31.01.2017 Ms. Preikschas 023815891214: Problem remains | - QC Error Level A,C - PH / pCO2 - inconsistent database | device checked. Restart due to inconsistent database error. Delivered solution pack after start-up phase. New solution pack now brings pCO2 cal error! Various Kal and QC performed.
No | 26 | Y | RESP | Imprecisión en medidas de thb |  | sustitución de hemolizador completo, calibración de Cthb, prueba de varias muestras analizadas seguidas y comprobando la repetibilidad. | D2088C7CFE372358 | [5 empty] | Inaccuracy in thb measures |  | complete hemolizer replacement, Cthb calibration, testing of several samples analyzed in a row and checking repeatability.
No | 11 | Y | RESP | ERRORES VARIOS NEUMATICA |  | SE SUSTITUYEN DIFERENTES PIEZAS Y TUBOS Y REACTIVOS DE DIFERENTES LOTES, PARA SUBSANAR ERRORES NO SIENDO EFECTIVIO. DESMONTAJE DE MODULO DE COOXIMETRIA PARA LIMPIEZA DE VALVULAS DE GANCHO Y HEMOLIZADOR, EN EL QUYE SE DETECTA UNA OBSTRUCCION, SE VUELVEN A REALIZAR PRUEBAS SIENDO CORRECTAS. CALIBRADO Y VERIFICADO CON Q.C. | D208BE27FE4833B1 | Cleaned | Fluidic pathway | Oxi module | Hemolyzer |  | VARIOUS NEUMATIC ERRORS |  | DIFFERENT PARTS AND TUBES AND REACTIVES OF DIFFERENT LOTS ARE SUBSTITUTE, TO CHANGE ERRORS NOT BEING EFFECTIVE. DISASSEMBLY OF COOXIMETRIA MODULE FOR CLEANING OF GANCHO AND HEMOLIZER VALVE, IN WHICH AN OBSTRUCTION IS DETECTED, TESTS ARE RETURNED BEING CORRECT. CALIBRATED AND VERIFIED WITH Q.C.
No | 4 | Y | RAME | Z Display intermittently freerzing. |  | Examined display assembly. Verified IP settings for NetOpt remote support connection. | D20AFCC7EA5D2C4A | Defect | HW | DMS_PC | Touch screen |  | Z Display intermittently freerzing. |  | Examined display assembly. Verified IP settings for NetOpt remote support connection.
No | 6 | Y | RAME | P Inlet leaking /sample errors |  | Sample errors (328) and leaking at capillary | D208E590FE46C1B5 | Defect | HW | Inlet module | Inlet Gasket |  | P Inlet leaking /sample errors |  | Sample errors (328) and leaking at capillary
No | Visit | Complaint | RIND | Sample error. | Sample measurement error. | Found problem with pH $BG module air Flat membrane problem, replaced the same restarted instrument calibrated results are OK, performed the sample found results are OK found instrument is functional. | SWO-164729 | Defect | HW | pH/BG Module | Valve, air |  | Sample error. | Sample measurement error. | Found problem with pH $BG module air Flat membrane problem, replaced the same restarted instrument calibrated results are OK, performed the sample found results are OK found instrument is functional.
No | Visit | Complaint | RDEU | Bildschirm reagiert nur langsam o. garnicht | Reinigungsarbeiten, Fehlersuche | z.Zt. Kein Fehler feststellbar, Prüfung Festplattenkapazität LW C: , Prüfung Konfiguration, Kontrolle Schnittstelle, Kontrolle Firewall, Reinigung Bildschirm, Probelauf Hinweis, dass bitte keine alkoholischen Desinfektions- und Reinigungsmittel für den Bildschirm verwendet werden !!! | SWO-163812 | Defect | HW | DMS_PC | Touch screen |  | Screen reacts slowly o.e. not at all | Cleaning work, troubleshooting | Currently. No error detectable, checking disk capacity LW C: , check configuration, control interface, control firewall, cleaning screen, test run Please note that no alcoholic disinfectants and cleaning agents are used for the screen !!!
No | Visit | Complaint | RDEU | pCO2 Ergebnisse in L3 Ergebnissen zickzackmäßig an/unter der Untergrenze | s. case | Nach Rücksprache mit RDEU und RMED den pH Dichtring ersetzt durch 834-618 (834-682). pH Messkammer sowie Elektrode gereinigt, Kalibration, Endkontrolle: AC5+ L3 Messung => Alle Werte OK, pCO2 nahezu mittig. Gerät bzw. L3 Messungen werden weiter beobachtet | SWO-163921 | Defect | Electrode | Electrode | Sealing ring | Sealing ring, pH | pCO2 results in L3 results zigzag at/below the lower limit | s. case | After consultation with RDEU and RMED, the pH sealing ring is replaced by 834-618 (834-682). pH measuring chamber and electrode cleaned, calibration, final inspection: AC5+ L3 Measurement => All values OK, pCO2 almost central. Device or L3 measurements are further observed
No | Visit | Complaint | RDEU | Herr Moritz meldet einen regelmäßig "einfrierenden" Touchscreen, der dann nicht mehr bedienbar ist | s. case | Erneut Probleme mit eingefrorenem Touchscreen laut Anwender, Herrn Otten, Monitoreingaben seien nicht mehr möglich. Beim Eintreffen aktuell keine Probleme feststellbar => neuen Monitor mit geänderter Firmware bestellt; wird nach Lieferung ersetzt. | SWO-163978 | Defect | HW | DMS_PC | Touch screen |  |  | s. case | Once again, according to the user, Mr Otten, problems with a frozen touchscreen are no longer possible. No problems can be detected on arrival at the moment => new monitor with changed firmware ordered; will be replaced after delivery.
No | Visit | Complaint | RDEU | Einlassnadel nach oben gebogen | s. case | Einlassnadel nach oben geknickt => im Rahmen der Tätigkeiten der FAN915-318 gerichtet. | SWO-163989 | Defect | HW | Inlet module | Probe bend |  | Inlet needle bent upwards | s. case | Inlet pin bent upwards => within the scope of the activities of fan915-318.
No | 7 | Y | RAME | F Error codes 0316 and 257 neon lamp intensity too low |  | Fails OXI test for neon lamp | D2089D85FE36E966 | Defect | HW | Oxi module | PCB, Oxi module |  | F Error codes 0316 and 257 neon lamp intensity too low |  | Fails OXI test for neon lamp
No | Visit | Complaint | RDEU | 03.01.2017, Frau Weigand, Tel.: 093319086511. Datenübertragung/ Verbindung verloren an OSM 30.01.2017 10:30 Fr. Weigand 093319086512: Datenübertragung nicht mehr möglich; vermutlich geiche Problematik wie am 03.01.2017 | Überprüfung der Einstellungen und IP Adressen | IP Adressen korrigiert. Speicherbefehl mit Batch Datei ausgeführt. Datenübertragung iO. | SWO-163950 | Setup | SW | His/lis | His/lis |  | 03.01.2017, Mrs. Weigand, Tel.: 093319086511. Data transfer/ Connection lost to OSM 01/30/2017 10:30 Fr. Weigand 093319086512: Data transfer no longer possible; probably a problem like on 03.01.2017 | Checking settings and IP addresses | IP addresses corrected. Memory command executed with batch file. Data transfer iO.
No | 4 | Y | RFRA | Blank | Erreur 5784 |  | D20B3ADCEA3DE2E8 | Defect | Consumables | Test Kit | Cartridge | Broken, nose | Blank | Error 5784
No | Visit | Complaint | RDEU | Kundin hat bereits das Gerät gespült und den Einlass gereinigt. | Fehlersuche durchgeführt. Diverse Verbrauchsmaterialien aus Kundenbestand erneuert. Hemolyzer erneuert. LS-Sensoren überprüft. | Hemolyzer erneuert. | SWO-164375 | Defect | HW | Oxi module | Hemolyzer |  | Customer has already rinsed the device and cleaned the inlet. | Troubleshooting. Various consumables from customer stock renewed. Hemolyzer. LS sensors. | Hemolyzer.
No | Visit | Complaint | RDEU | Kunde hat bereits 2 mal die Sensor Kassette und das LP erneuert, Fehler bleibt bestehen. | pO2 und pCO2 QC Fehler Level A Datei Kal. Datei QC Datei Betriebsmittel | Gerät im Service überprüft, Kunde hat in den letzten Tagen 2x SC90 und einmal LP90 getauscht, versucht LP90 an einem anderen ABL einzusetzen - ging nach Rücksprache Arzt nicht, Tausch des LP90, nach Startphase und QC alles wieder grün, System- und Funktionskontrolle in Ordnung | SWO-164306 | Assisted Replacement | SP | SP | pCO2 and pO2 |  | Customer has already renewed the sensor cassette and the LP 2 times, error remains. | pO2 and pCO2 QC Error Level A File Kal. File QC File Equipment | Device checked in service, customer has exchanged 2x SC90 and once LP90 in the last days, tried to use LP90 on another ABL - did not go after consultation doctor, exchange of the LP90, after start phase and QC everything back green, system and function control in order
No | 2 | Y | RFRA | Blank | CQ HbF hors gamme (trop faible) depuis mois de décembre |  | D20A091B74130AF6 | Defect | HW | Oxi module | Hemolyzer |  | Blank | CQ HbF hors gamme (trop faible) depuis mois de décembre
No | 3 | Y | RAME | P user called stating that keeps getting unable to pump solution. |  | found that the flow value is coming up at 16000. | D209F36974424041 | Defect | HW | Basic Unit | Wetsection control PCB |  | P user called stating that keeps getting unable to pump solution. |  | found that the flow value is coming up at 16000.
No | Visit | Complaint | RBNL | Instrument blijft op batterijvoeding staan als er weer netvoeding is ingeplugd. Schakelt niet over. Instrument volledig uitzetten en weer opstarten noodzakelijk. | Blijft op batterijvoeding. Defecte power supply 910-285P0020N682. | Nieuwe powersupply 910-285R0022N469 gemonteerd. | SWO-164096 | Defect | HW | Power Supply | Power Supply |  | Instrument remains on battery power when the power supply is plugged in again. Does not switch. Fully shutdown and restart necessary. | Remains on battery power. Defective power supply 910-285P0020N682. | New Powersupply 910-285R0022N469 mounted.
No | Visit | Complaint | RLTD | After rebooting during training the analyser is just saying Offline in the top left hand corner. | Replaced the Wet Section Control PCB & The CPU unit but neither of these solved the problem so the originals were re-fitted. Replaced the power supply which got the analyser past saying Offline and not being able to log on but when it booted up it didn't recognise that the inlet module was fitted. Tried replacing the Wet Section Control PCB again to see if it would recognise the inlet module but it didn't so the original PCB was re-fitted again. | As I was unable to get the analyser up and running it was decided to take this analyser back to Radiometer and return with a new one to replace it. | SWO-164367 | Revisit required | analyzer removed | analyzer removed | analyzer removed |  | After rebooting during training the analyser is just saying Offline in the top left hand corner. | Replaced the Wet Section Control PCB & The CPU unit but neither of these solved the problem so the originals were re-fitted. Replaced the power supply which got the analyser past saying Offline and not being able to log on but when it booted up it didn't recognise that the inlet module was fitted. Tried replacing the Wet Section Control PCB again to see if it would recognise the inlet module but it didn't so the original PCB was re-fitted again. | As I was unable to get the analyser up and running it was decided to take this analyser back to Radiometer and return with a new one to replace it.
No | Visit | Complaint | RDEU | Gestern LP Wechsel Seit dem Probleme mit diversen Parametern Kein weiteres LP vorrätig | Nach gestrigen LP Wechsel Probleme mit diversen Parametern | Beim Eintreffen nur pH und pCO2 grün => Rest rot => Kal und QC aber i.O. => bis auf pH und pCO2 alle Parameter deaktiviert Weitere Überprüfung und Reparatur im Rahmen der fälligen Jahreswartung | SWO-164218 | [5 empty] | Yesterday LP Change Since the problem with various parameters No more LP in stock | After yesterday's LP change problems with various parameters | On arrival only pH and pCO2 green => rest red => Kal and QC but i.O. => all parameters deactivated except for pH and pCO2 Further inspection and repair as part of the due annual maintenance
No | 5 | Y | RESP | Error de corriente de fugas |  | Limpieza del sistema con agua. Chequeo por servicio. Sustitucion de sp Cal, qc | D209E960742CCC70 | Assisted Replacement | SP | SP | Leak current |  | Leakage current error |  | Cleaning the system with water. Check for service. Substitution of sp Cal, qc
No | Visit | Complaint | RCHN | TNI值偏高，冲洗单元有问题 | TNI值偏高，冲洗单元有问题 | 更换备件903-119后修复 by jesson | SWO-164350 | Defect | HW | Process Module | Cup wash unit |  | TNI value is high, there is a problem with the flushing unit | TNI value is high, there is a problem with the flushing unit | Repair by jesson after replacement of spare parts 903-119
No | Visit | Complaint | RMAY | Check data log - All QC High bias for PCO2 and Lactate. Perform routine maintenance. Check measuring chamber for clot. Run protein removal Run calibration. | Check electrode and membrane. Perform routine maintenance. | Change Lactate and PCO2 membrane. | SWO-164725 | Assisted Replacement | Electrode | Membrane | Lac | and pCO2 | Check data log - All QC High bias for PCO2 and Lactate. Perform routine maintenance. Check measuring chamber for clot. Run protein removal Run calibration. | Check electrode and membrane. Perform routine maintenance. | Change Lactate and PCO2 membrane.
No | Visit | Complaint | RLTD | Baneta Yelda Tel. 0208 2164934 | Printer missing lines. | Replaced printer head and test ok. | SWO-164201 | Defect | HW | Printer Module | Printer head |  | Baneta Yelda Tel. 0208 2164934 | Printer missing lines. | Replaced printer head and test ok.
No | Visit | Complaint | RIND | Auto check Scanner error. | Auto check Scanner error. | Removed the Auto check module, Cleaned the analyzer and Scanner restarted instrument found it was scanning the all the ampules found instrument is functional, | SWO-164730 | Cleaned | HW | Auto Check Module | Barcode reader, Autocheck Module |  | Auto check Scanner error. | Auto check Scanner error. | Removed the Auto check module, Cleaned the analyzer and Scanner restarted instrument found it was scanning the all the ampules found instrument is functional,
No | Visit | Complaint | RDEU | 01.02.2017, Frau Bek (Labor), Tel.: 07318040 : Bildschirm defekt (blue) | Fehlersuche vor Ort |  | SWO-164144 | Defect | HW | DMS_PC | CPU |  | 01.02.2017, Mrs. Bek (Laboratory), Tel.: 07318040 : Screen defective (blue) | Troubleshooting on site
No | 2 | Y | RFRA | Blank | Erreur de spectre sur CQ niveau 3 |  | D209B1C1745A0A0E | Adjustment | HW | Oxi module | Lamp unit |  | Blank | Spectrum Error on QC Level 3
No | Visit | Complaint | RDEU | 01.02.2017, Frau Java, Tel.: 017631658717: Messungen instabil bei pH und pO2 02.02.2017/12:00Uhr Frau Java 017631658717 Rückruf: 022147887890: Oxifehler 581, pH instabil und pO2 instabil bei Blutmessung. | Bei jeder (Schweine-) Blutmessung Fehler 581 OxiSpectrumMismatch | Mehrere Proteinentfernungen, Reinigung, Kal sowie OxiCompleteTest durchgeführt, AC5+ Messung OK, erste Blutmessung erneut mit Fehler 581 => OxiModul ersetzt, Kal, OCT, tHb-Kal, AC5+ Messung, Blutmessung => weiterhin 581 => Originalmodul wieder eingebaut, abgeglichen, kalibriert, 3 Level QC kontrolliert, alle OK. | SWO-164296 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | 01.02.2017, Mrs. Java, Tel.: 017631658717: Measurements unstable at pH and pO2 02.02.2017/12:00h Ms Java 017631658717 Recall: 022147887890: Oxi error 581, pH unstable and pO2 unstable in blood measurement. | For each (pig) blood measurement error 581 OxiSpectrumMismatch | Several protein removals, cleaning, cal and OxiCompleteTest performed, AC5+ measurement OK, first blood measurement replaced again with error 581 => OxiModul, Kal, OCT, tHb-Kal, AC5+ measurement, blood measurement => still 581 => original module re-installed, matched, calibrated, 3 level QC controlled, all OK.
No | Visit | Complaint | RDEU | 09.01.2017/14:00, Herr Esche: Lac. Kal. Fehler nach Membranierung 10.01.2017, 8.15 Uhr Herr Esche, Tel.: 089857915954 : Weiterhin Kalibrationsfehler bei Laktat | diverse Driftfehler | Erneuerungsintervalle für Membranen auf Anweisung von TSP (BorWe) angelegt Glu, Lac und Ref. Membran monatlich alle anderen 3 monatlich | SWO-164297 | Assisted Replacement | Electrode | Membrane | Lac |  | 09.01.2017/14:00, Mr Esche: Lac. Kal. Errors after membraneing 10.01.2017, 8.15 a.m. Mr. Esche, Tel.: 089857915954 : Further calibration error in lactate | different Driftfehler | Renewal intervals for membranes created on the instruction of TSP (BorWe) Glu, Lac and Ref. Membrane monthly all other 3 monthly
No | Visit | Complaint | RSCH | Folgereparatur SWO-154647 | Folgereparatur SWO-154647 | Display wurde durch René ersetzt wie abgemacht (Folgereparatur SWO-154647) | SWO-164089 | Defect | HW | DMS_PC | Touch screen |  | Folgereparatur SWO-154647 | Folgereparatur SWO-154647 | Display was replaced by René as done (Follow-up repair SWO-154647)
No | 2 | Y | RPNZ | Blank | Set up and complete validation for install into NICU site. |  | D209D9A87406AC90 | Defect | HW | Basic Unit | PCB, Analyzer control |  | Blank | Set up and complete validation for install into NICU site.
No | 2 | Y | RAME | P Leaking at the Y adapter. |  | Confirmed problem. | D209C131742D550C | Cleaned | Fluidic pathway | Fluidic manifold | Y-piece | Clot/Leakage | P Leaking at the Y adapter. |  | Confirmed problem.
Yes | 6 | Y | RFRA | Blank | Support tube échantillon cassé |  | D209E0577405B204 | [5 empty] | Blank | Broken sample tube support
Yes | Visit | Complaint | RDEU | 02.02.2017, Fr. Zafiriadou, Tel. 0694071878: Zähler für Nadel Ersatz hat den Warnwert erreicht | Erreichte Nadelstiche 9535. Analysator sonst ohne Fehlermeldung. Daten gesichert. | Ausbau der alten Nadel und Einbau einer neuen Nadel. Überprüfung der Beweglichkeit der Nadel im Needleblock, Funktion in Ordnung. Qualitätskontroll Messungen durchgeführt. | SWO-164322 | Re classified RMED |  |  |  |  | 02.02.2017, Ms. Zafiriadou, tel. 0694071878: Meter for needle replacement has reached the warning value | Achieved needle stitches 9535. Analyzer otherwise without error message. Data backed up. | Removal of the old needle and installation of a new needle. Checking the mobility of the needle in the needle block, function in order. Quality control measurements carried out.
No | Visit | Complaint | RDEU | Analysator bricht die Proben ab mit Meldung 1067. 20.01.2017, Frau Beier, Tel.: 023234995368: erneut Probenabrüche 02.02.2017/13:50Uhr Frau Beier 023234995290: Häufig Probenabruch mit Fehler 1067 | Anwender hat Cup auf dem Abfallbehälter gefunden. | AQT zur Zeit betriebsbereit, SD gesichert. Abfallbehälter XV-58 | SWO-164365 | Defect | Consumables | SP | Cup Stack | 1067 | Analyzer cancels the samples with message 1067. 20.01.2017, Mrs. Beier, Tel.: 023234995368: probenprobenabruch 02.02.2017/13:50pm Mrs Beier 023234995290: Common sample break with error 1067 | User has found Cup on the waste container. | AQT currently operational, SD secured. Waste container XV-58
No | Visit | Complaint | RDEU | an unserem Troponin-Testgerät AQT 90 flex ist der Scanner für die Patienten-Barcodes defekt --> Kein Scanner"Strich" sichtbar, es lässt sich ebenfalls kein Barcode scannen. | Scanner für die Patienten-Barcodes defekt --> Kein Scanner Strich sichtbar, es lässt sich ebenfalls kein Barcode scannen. | Verbindung Barcodescanner überprüft. Geräteneustart. Scannerfunktion wieder i.O. | SWO-164136 | Defect | HW | Aspiration Module | Barcode reader, Inlet |  |  | Scanner for the patient barcodes defective --> No scanner stroke visible, also no barcode can be scanned. | Connection barcode scanner checked. Device restart. Scanner function again in the
No | Visit | Complaint | RLTD | Denise on 0208 9499061 speak to charmiane, - | Removed sensor interface pcb and then removed sensor pin assembly. Cleaned with alcohol contacts on sensor interface pcb and both ends of all contact pins. Checked all contact pins are had free moving spring action - all ok. Re-assembled checking alignment of pins and tested analyser operation. Glucose parameter after allowing time to stabilise is now passing calibrations. Performed quality controls all 3 levels passed for all parameters. Analyser ready for use. | Cleaned sensor contact pins. | SWO-164917 | Cleaned | HW | Basic Unit | SC Contacts |  | Denise on 0208 9499061 speak to charmiane, - | Removed sensor interface pcb and then removed sensor pin assembly. Cleaned with alcohol contacts on sensor interface pcb and both ends of all contact pins. Checked all contact pins are had free moving spring action - all ok. Re-assembled checking alignment of pins and tested analyser operation. Glucose parameter after allowing time to stabilise is now passing calibrations. Performed quality controls all 3 levels passed for all parameters. Analyser ready for use. | Cleaned sensor contact pins.
No | 8 | Y | RAME | Z Analyzer powered down on it's own and will not turn on. |  | Verified. | D20B3CD1EA0C9264 | Revisit required | HW | Battery module | Battery |  | Z Analyzer powered down on it's own and will not turn on. |  | Verified.
No | Visit | Complaint | RDEU | Anwender berichten, Analysator zeigt Meldung Tausch des Lösungspacks an, obwohl LP genug Messzyklen hat. Ampelstatus ist auch grün. | Analysator zeigt bei Service keine Fehlermeldung an. Meldung erscheint sporadisch und wird nicht aufgezeichnet unter Aktivitäten. | Anwender haben Analysator neu gestartet bei Erscheinen der Meldung. Danach ist Analysator wieder ohne Meldung. Kein weiterer Fehler festgestellt. Datalogs und Service Dump gezogen zur weiteren Auswertung. | SWO-164536 | Defect | SW | SW | SP exchange |  | Users report, analyzer displays message exchange of solution pack, even though LP has enough measurement cycles. Traffic light status is also green. | Analyzer does not display an error message to Service. Message appears sporadically and is not recorded under Activities. | Users have restarted Analyzer when the message appears. After that, the analyzer is again without a message. No further error detected. Datalogs and Service Dump drawn for further evaluation.
No | Visit | Complaint | RDEU | Analysator gibt Meldung raus, dass Lösungspack leer ist und getauscht werden soll. Meldung kann man wegdrücken oder Analysator neu starten, dann ist Meldung weg, kommt aber wieder. | Im Haus unterwegs zu weiteren Service. | Analysator zeigte beim Service die Meldung nicht. Überprüfung des Analysators, keine Meldung zum Lösungspack Fehler gefunden. Datalogs und Service Dump gezogen zur Auswertung. pO2 Kalibrierfehler, SC erneuert. | SWO-164537 | Defect | SW | SW | SP exchange |  | Analyzer issues a message that solution pack is empty and should be replaced. Message can be pushed away or analyzer can be restarted, then message is gone, but comes back. | In the house on the way to further service. | The analyzer did not show the message to the service. Checking the analyzer, no message found about the solution pack error. Datalogs and Service Dump drawn for evaluation. pO2 calibration error, SC renewed.
No | 6 | Y | RPAC | Blank | inlet issues |  | D20B2233EA39CC49 | Adjustment | HW | Auto Inlet module | Inlet drive gear |  | Blank | inlet issues
No | Visit | Complaint | RDEU | 02.02.2017, OA Dr. Stiller, Tel.: 05112893835: Kalibrations-und AQC Fehler bei pCO2 02.02.2017, 8.00 Uhr, Dr. Pape, 05112893835: Kalibrationen ok, Messungen werden z.T. abgebrochen | Messungen Abgebrochen, Billi bei einigen Messungen nicht berichtet. Einlassdichtung an der Nadel undicht. | Einlassdichtung erneuert, Fluid Transport geprüft, tHb Kal. durchgeführt. | SWO-164834 | Defect | HW | Auto Inlet module | Inlet Gasket |  | 02.02.2017, OA Dr. Stiller, Tel.: 05112893835: Calibration and AQC error in pCO2 02.02.2017, 8.00 a.m., Dr. Pape, 05112893835: Calibrations ok, measurements are partly canceled | Measurements Cancelled, Billi did not report on some measurements. inlet seal on the needle leaks. | Inlet seal renewed, fluid transport tested, tHb Kal. Performed.
No | Visit | Complaint | RDEU | 02.02.2017 12:00 H. Kapell 03696168088: Kunde wünscht Austauschgerät | QC Fehler K, pOs im Lev.1 Kunde wünscht Austauschgerät | Überprüfung SC, SP $ beides wurde durch Kunden bereits erneuert, SP reinitialisiert nach Einlaufprozedur QC und Kal i.O. | SWO-164398 | [5 empty] | 02.02.2017 12:00 H. Kapell 03696168088: Customer wants replacement device | QC Error K, pOs in Lev.1 Customer wants replacement device | Review SC, SP both have already been renewed by customers, SP reinitialized after the inlet procedure QC and Kal i.O.
No | 11 | Y | RAME | W Analyzer shows "Software Exception Error" |  | Troubleshot to a bad CF card | D20EB4F04C6210D2 | [7 empty] | Troubleshot to a bad CF card
No | Visit | Complaint | RCHN | LAC定标失败 | LAC定标失败 | 更换测试卡后修复 | SWO-168133 | [5 empty] | LAC calibration failure | LAC calibration failure | Repair after replacing test card
No | Visit | Complaint | RCHN | LAC定标失败 | LAC定标失败 | 更换测试卡正常。 | SWO-164352 | Assisted Replacement | SC | SC | Lac |  | LAC calibration failure | LAC calibration failure | The replacement test card is OK.
No | Visit | Complaint | RCHN | GLU LAC定标状态超出限度 | GLU LAC定标状态超出限度 | 更换测试卡 | SWO-164819 | Assisted Replacement | SC | SC | Glu/Lac |  | GLU LAC calibration status exceeds limit | GLU LAC calibration status exceeds limit | Replace the test card
Yes | 4 | Y | RCZE | Blank | tece u vstupu i po výmìnì tìsnìní |  | D20AB7F8EA32B97A | [5 empty] | Blank | At the inlet and after the seal is
No | Visit | Complaint | RDEU | Gerät tropft aus Einlass | Gerät, Blut tropft aus Einlass Einlassanschluss unter Inletgasket verstopft und verbogen | ABL äußerlich gereinigt Proteinentfernung durchgeführt Tubing nipple kit sowie Abfallpumpenschläuche erneuert Gerät dekontaminiert, Inlet Gasket aus Kundenbestand erneuert ABL kalibriert und überprüft | SWO-164453 | Cleaned | Fluidic pathway | Inlet module | Fluidic pathway |  | Device drips from inlet | Device, blood drips from inlet Inlet connection under Inletgasket clogged and bent | ABL externally cleaned Protein removal performed Tubing nipple kit as well as waste pump hoses renewed Device decontaminated, Inlet Gasket renewed from customer base ABL calibrated and checked
No | Visit | Complaint | RDEU | pH Messung instabil, sporadisch inhomogene Probe | pH Messung instabil, sporadisch inhomogene Probe | Proteinentfernung durchgeführt pH Elektrode gereinigt, Inletgasket aus Kundenbestand erneuert Schläuche zwischen Einlasssonde und pH $BG Modul erneuert Gerät kalibriert und überprüft, Blutmessungen durchgeführt | SWO-164454 | Cleaned | Fluidic pathway | pH/BG Module | Fluidic pathway |  | pH measurement unstable, sporadic inhomogeneous sample | pH measurement unstable, sporadic inhomogeneous sample | Protein removal performed pH electrode cleaned, inlet gasket renewed from customer base Hoses between inlet probe and pH $BG module renewed Device calibrated and checked, blood measurements carried out
No | Visit | Complaint | RDEU | 03.02.2017 12:30 H. Soest 020172384285: Deckel AQC-Modul fährt ohne Abschaltung auf/zu | Beim Schließen des Deckels des AQC-Moduls stoppt der Motor nicht bei geschlossener Position. Nach kurzer Zeit fährt der Deckel wieder runter. | Überprüfung der Deckelsensoren über das Servicemenü: Sensoren arbeiten unzuverlässig, vor allem der untere. --> Erneuerung der entsprechenden Autocheck Interface PCB incl. Sensoren. Funktionskontrolle: I.O. -Garantie- | SWO-164571 | Defect | HW | Auto Check Module | PCB, Autocheck |  | 03.02.2017 12:30 H. Soest 020172384285: Lid AQC module goes on/to without shutdown | When closing the lid of the AQC module, the motor does not stop when the position is closed. After a short time, the lid goes down again. | Checking the lid sensors via the service menu: Sensors work unreliably, especially the lower one. --> Renewal of the corresponding Autocheck Interface PCB incl. sensors. Functional control: I.O. -Guarantee
No | Visit | Complaint | RLTD | ranvir - 01865 2215 96, | Waste line blocked | Removed waste line and waste bottle connector and flushed as blocked removed blood clot. Replaced waste line and test ok. | SWO-164482 | Cleaned | Fluidic pathway | Basic Unit | Fluidic pathway |  | ranvir - 01865 2215 96, | Waste line blocked | Removed waste line and waste bottle connector and flushed as blocked removed blood clot. Replaced waste line and test ok.
No | Visit | Complaint | RDEU | Inlet Error |  | Mehrfach Inlet Kal in Service Programm durchgeführt. Ok. Nach Neustart wieder Inlet Kal Error. Einlassdichtung ersetzt. | SWO-164527 | Defect | HW | Inlet module | Inlet gasket |  | Inlet Error |  | Multiple Inlet Kal in Service Program performed. Ok. After rebooting again Inlet Kal Error. replaced inlet seal.
No | Visit | Complaint | RDEU | 03.02.2017, Frau Balling, Tel.: 093319086511: AQT läuft aus | Überprüfung der Schlauchverbindungen | Schlauch Nr 6 steckt nicht auf dem Manifold. Schlauch wieder angesteckt. Funktionstest iO. | SWO-164553 | Defect | Consumables | SP | Valve, blocked |  | 02/03/2017, Mrs Balling, Tel.: 093319086511: AQT expires | Checking the hose connections | Tube No 6 is not on the Manifold. hose replugged. Functional test iO.
No | 31 | Y | RPOL | Blank | Wymiana taœmy po³¹czeniowej pomiêdzy jednostk¹ CPU a ekranem dotykowym. Analizator sprawny technicznie |  | D20AFF6DEA1C40A7 | Defect | HW | Analyzer Control Module | Cable, UIM |  | Blank | Replacing the connection volume between the CPU unit and the touch screen. Technically efficient analyzer
No | Visit | Complaint | RDEU | Probenabbrüche 1067 | Probenabbrüche 1067, 1020 | Umgebungstemperatur angeglichen, Probennnehmerkennung gereinigt und kalibriert, Probenmessungen o.k. | SWO-164474 | Setup | HW | Analyzer | Ambient Temp | 1020 | Probenabbrüche 1067 | Probenabbrüche 1067, 1020 | Ambient temperature, sample detection cleaned and calibrated, sample measurements ok.
No | 4 | Y | RFRA | Blank |  |  | D20ADBCFEA26746F | Defect | HW | Process Module | Tubes |  | Blank
No | Visit | Complaint | RCHN | PCO2,CA,GLU,PH定标失败 | PCO2,CA,GLU,PH定标失败 | 更换测试卡正常。 | SWO-164641 | [5 empty] | PCO2, CA, GLU, PH calibration failure | PCO2, CA, GLU, PH calibration failure | The replacement test card is OK.
No | Visit | Complaint | RCHN | CA,GLU定标不过 | CA,GLU定标不过 | 更换测试卡后修复 | SWO-165225 | [5 empty] | CA, GLU calibration but | CA, GLU calibration but | Repair after replacing test card
No | Visit | Complaint | RCHN | PCO2定标失败 | PCO2定标失败 | 更换测试卡后修复 | SWO-164637 | Assisted Replacement | SC | SC | pCO2 |  | PCO2 calibration failure | PCO2 calibration failure | Repair after replacing test card
No | Visit | Complaint | RCHN | 报1024错误代码 | 光单依然存在质量问题，报错1031，需要重新再借一个903-364 |  | SWO-164634 | Defect | HW | Process Module | Optical Unit | 1031 | Reported 1024 error code | Light list still has quality problems, reported error 1031, need to borrow another 903-364
No | Visit | Complaint | RCHN | 液路故障，FCT-107 | 液路故障，FCT-107，借备件903-209,902-918各一个 |  | SWO-164646 | Defect | HW | Process Module | PCB, Inlet |  | Liquid road fault, FCT-107 | Liquid road fault, FCT-107, spare parts 903-209, 902-918 each
No | Visit | Complaint | RDEU | pO2 Fehler bei QC Level B. Messung zu tief! | Überprüfung der Kal, QC und Systemdateien | Lösungspack erneuert und Ersatz bereitgestellt. Funktionskontrolle in Ordnung. | SWO-164738 | Assisted Replacement | SP | SP | pO2 |  | pO2 error at QC Level B. Measurement too deep! | Checking Kal, QC and system files | solution pack renewed and replacement provided. Functional control in order.
No | Visit | Complaint | RCHN | 测试卡初始化失败 | 测试卡初始化失败，借902-972一个 | 更换备件902-972后修复(by tim) | SWO-164631 | [5 empty] | Test card initialization failed | Test card initialization failed, borrowing 902-972 one | Repair after replacement of spare parts 902-972 (by tim)
No | Visit | Complaint | RCHN | K,NA,CA,CL定标失败 | K,NA,CA,CL定标失败，借903-183，903-158各一个。 | 更换903-183后定标通过，组装时发现903-158上的4个弹片弹起的程度明显不一致，更换后修复。by jesson | SWO-164401 | Defect | HW | Basic Unit | PCB, Sensor interface |  | K, NA, CA, CL calibration failed | K, NA, CA, CL calibration failed, borrowing 903-183, 903-158 each. | After the replacement 903-183 after the calibration pass, assembly found that the 903-158 on the four shrapnel up the degree of significantinin, after replacement repair. by jesson
No | Visit | Complaint | RCHN | 处理盘错误 | 处理盘错误，借备件636-550一个 | 更换636-550后修复 by jesson | SWO-164615 | Defect | HW | Process Module | Process Unit sensor | 5784 | Handling disk errors | Processing tray error, borrowed spare parts 636-550 one | Repair by jesson after replacement 636-550
No | Visit | Complaint | RCHN | 报5784错误 | 报5784错误，借636-550一个 |  | SWO-164712 | Defect | HW | Process Module | Process Unit sensor | 5784 | Reported 5784 errors | Reported 5784 error, borrow 636-550 one
No | 8 | Y | RAME | Z The ABL90 would not power on when the battery was connected. |  | Reconnected the battery and verifed the problem. After the battery was replaced the ABL90 still would not power on. | D20F401D4C07A7AC | Unknown | HW | Power Supply | Power Supply |  | Z The ABL90 would not power on when the battery was connected. |  | Reconnected the battery and verifed the problem. After the battery was replaced the ABL90 still would not power on.
No | Visit | Complaint | RDEU | Kal. Und QC-Fehler bei Glu und Lac | Fehlersuche --> SC wurde gegen 9 Uhr getauscht --> deswegen noch Kal. Und QC-Fehler bei Glu und Lac | Mehrere Kalibrierungen und QC-MESSUNGEN wiederholt, System und Funktion i.O. | SWO-165015 | Assisted Replacement | SC | SC | Glu/Lac | Start up | Kal. Und QC-Fehler bei Glu und Lac | Troubleshooting --> SC was exchanged at 9 o'clock --> therefore Kal. And QC errors in Glu and Lac | Multiple calibrations and QC MEASUREMENTS repeated, system and function in the
Yes | Visit | Complaint | RDEU | Keine LPs im Haus vorhanden | LPs ausgeliefert. | Kreißsaal hat ein LP erhalten, Intensiv hat zwei LPs erhalten, Notaufnahme hat ein PL erhalten. In de Notaufnahme noch mehrere Fragen zum ABL beantwortet und die Fehlerbehebung durchgesprochen. | SWO-165196 | Re classified RMED | Logistical upSWO |  |  |  | No LPs available in the house | LPs. | Kreißsaal has received an LP, Intensively received two LPs, Emergency room has received a PL. In the emergency room several questions about the ABL answered and the troubleshooting was answered.
No | Visit | Complaint | RDEU | Intervention erforderlich - Meldung nach SP Tausch | Einlassdichtungen ersetzt, LS Pack ersetzt, LS Pack mit einem anderen Gerät gerauscht. Der LS Pack ist I.O. Verbindungsschlauch zwischen Oximetrie und SC ersetzt, nach wie vor Leckage Fehler. Abfallschlauch ersetzt. | Startphase QC und Kalibration. | SWO-164924 | Revisit required | HW | Leakage | Leakage |  | Intervention required - message after SP exchange | Replaced inlet seals, LS Pack replaced, LS Pack roughed with another device. The LS Pack is I.O. connection hose replaced between oximetry and SC, still leakage error. replaced by the waste hose. | Start phase QC and calibration.
No | Visit | Complaint | RDEU | 06.02.2017 Herr Wagner 065194783730 meldet das sein ABL auf Grund eines Fehlers ausgeschaltet wurde. | Siehe Cas vom 05.02.2017 | WE Dienst | SWO-165159 | Revisit required | analyzer removed | analyzer removed | analyzer removed |  | 06.02.2017 Mr. Wagner 065194783730 reports that his ABL was switched off due to an error. | See Cas of 05.02.2017 | WE Service
No | Visit | Complaint | RDEU | Ca. Messabweichungen. | Herr Wagner Berichtet, dass Ca. massive Messabweichungen hat. Er hat mir Befunde von einer Probe gezeigt. Das erste Ergebnis war 1,06 das Zweite drei Minuten später 0,9 und das Dritte noch mal drei Minuten später 0,6. | Ich habe auf der Station B2 Ca. aktiviert beide ABL´s ein kalibriert und QC´s gemessen. Danach haben wir vier Vergleichsmessungen durchgeführt. Das Ergebnis bei Ca. lag bei beiden ABL´s bei vier Messungen zwischen 1,28 und 1,29. Das Ca. auf der St. B2 habe ich nach den Vergleichsmessungen deaktiviert. | SWO-165167 | No fault found | Performance | Performance | Ca |  | Approximate measurement deviations. | Mr. Wagner reports that approx. has massive measurement deviations. He showed me findings from a sample. The first result was 1.06 the second three minutes later 0.9 and the third another three minutes later 0.6. | I have activated on station B2 Ca. both ABL's one calibrated and MEASURED QC's. After that, we carried out four comparative measurements. The result for Approx. was four measurements between 1.28 and 1.29 for both ABLs. The approx. on the St. B2 I deactivated after the comparison measurements.
Yes | 4 | Y | RFRA | Blank | Analyseur non disponible lors du passage du BG après remplacement du pack de solution |  | D20EE0C04C50B98C | [5 empty] | Blank | Analyzer not available when the BG passes after replacing the solution pack
No | Visit | Complaint | RBNL | De barcode reader van ABL 1 serienr. N0025 leest af en toe een verkeerde barcode. Kunnen jullie hier iets aan doen? Gr. Marlies Met vriendelijke groet, M.E.Schuil-Meijers | Het probleem beschreven door de gebruiker; De barcode reader van ABL 1 serienr. N0025 leest af en toe een verkeerde barcode. Kunnen jullie hier iets aan doen? M.E.Schuil-Meijers ABL90 leest soms verkeerde labnummer in. Dit is meerdere malen voorgekomen. Bij voorbeeld; Bloedmonsterbarcode met het labnummer: 7020306971 door de ABL90 is geidentificeerd als; 7010306971, of Bloedmonsterbarcode met het labnummer: 7020231571 door de ABL90 is geidentificeerd als; 7020251571. Het barcodetype is interleaved 2 of 5. | Zodoende de barcodescanner unit heb vervangen vervolgens heb gecontroleerd. Tevens de beveiligingsopties en maken van de back-up, in beide ABL90 toestellen heb uitgelegd $aangepast. | SWO-165075 | User error | Barcode Quality | Barcode Quality | Barcode Quality |  | The barcode reader of ABL 1 serienr. N0025 occasionally reads a wrong barcode. Can you do something about this? Gr. Marlies Sincerely M. E. Hides-Meijers | The problem described by the user; The barcode reader of ABL 1 serienr. N0025 occasionally reads a wrong barcode. Can you do something about this? M. E. Hides-Meijers ABL90 sometimes reads wrong lab number. This has occurred several times. For example; Blood sample barcode with the lab number: 7020306971 has been identified by the ABL90 as; 7010306971, or Blood sample barcode with the lab number: 7020231571 has been identified by the ABL90 as; 7020251571. The barcode type is interleaved 2 or 5. | So the barcode scanner unit got replaced then checked. Also explained the security options and create the backup, in both ABL90 devices $aangepast.
No | Visit | Complaint | RCHN | PO2偏高 | PO2偏高,借902-972一个，902-982一个 |  | SWO-164880 | [5 empty] | PO2 high | PO2 high, borrow 902-972 one, 902-982 one
No | Visit | Complaint | RCHN | LAC定标不过，偏高 | LAC定标不过，偏高 | 更换测试卡 | SWO-164824 | [5 empty] | LAC calibration is not, high | LAC calibration is not, high | Replace the test card
Yes | 2 | Y | RFRA | Blank | Initialisation SC80 impossible. |  | D20EDA304C3E9A86 | [5 empty] | Blank | SC80 initialization impossible.
Yes | 5 | Y | RAME | P Unit leaking from sensor cassette |  | Checked lines for crimps. Checked waste pump. Found problem to be defective solution pack | D20EDABD4C2770A0 | [5 empty] | P Unit leaking from sensor cassette |  | Checked lines for crimps. Checked waste pump. Found problem to be defective solution pack
No | 4 | Y | RAME | P Suspect Clot in Analyzer |  | Found analyzer inoperable upon arriving, customer knew that it was clotted from the blood that was run. | D21B45EAE8565126 | [5 empty] | P Suspect Clot in Analyzer |  | Found analyzer inoperable upon arriving, customer knew that it was clotted from the blood that was run.
No | 11 | Y | RAME | W Analyzer shows "Software Exception Error". |  | Troubleshot to a bad CF card | D20EB8344C1210D5 | [7 empty] | Troubleshot to a bad CF card
No | Visit | Complaint | RCHN | 电极灵敏度低 | 电极灵敏度低，借备件945-737 |  | SWO-164883 | [5 empty] | Low electrode sensitivity | Low electrode sensitivity, spare part 945-737
No | Visit | Complaint | RCHN | 定标底物外漏，检查失败 | 定标底物外漏，检查失败 | 更换903-364修复(by zhenjie) | SWO-164938 | Defect | HW | Process Module | Optical Unit | 1024 | The calibration substrate leak, check failed | The calibration substrate leak, check failed | Replacement 903-364 repair (by zhenjie)
No | Visit | Complaint | RCHN | 报错5487 | 报错5487，借备件580-087,636-550各一个 |  | SWO-164785 | Defect | HW | Process Module | Process Unit sensor | 5784 | 5487 error | Error 5487, spare parts 580-087, 636-550 each
No | 5 | Y | RESP | Error 1223 y 1257. No mide correctamente el hct |  | Limpiado céluda de hematocrito. Pasados QCs y muestsras para comprobar el estado del equipo. | D20EA9BA4C13C2DF | Cleaned | HW | Aspiration Module | Needle Wash Unit | 1223 | Error 1223 and 1257. Doesn't measure hct correctly |  | Cleaned with hematocrit. Passed QCs and standards to check the status of the equipment.
No | Visit | Complaint | RIND | Oxy-metery parameter error. | Check the system. check the logs. | Reset the Oxy-metry module & Recalibrate system. Run oxy test & passed. calibrate system & qc pass. | SWO-166257 | Adjustment | HW | Oxi module | tHb Cal |  | Oxy-metery parameter error. | Check the system. check the logs. | Reset the Oxy-metry module & Recalibrate system. Run oxy test & passed. calibrate system & qc pass.
No | 13 | Y | RAME | P PH/BG liquid sensor error. |  | Patti said that everything was out. It was givomng a liquid sensor error. I asked her to check the data logs and found that she was getting pH $BG module detected rinse when not expected. This is the tiop of the air vavle broken off. That is what I found. Patti had to leave so I will sign verbal. | D20ED6E34C08B6BF | Defect | HW | pH/BG Module | Valve, air |  | P PH/BG liquid sensor error. |  | Patti said that everything was out. It was givomng a liquid sensor error. I asked her to check the data logs and found that she was getting pH $BG module detected rinse when not expected. This is the tiop of the air vavle broken off. That is what I found. Patti had to leave so I will sign verbal.
No | Visit | Complaint | RKK | OXIｴﾗｰ：Halogen Lamp intensity high | halogenn Lammpu intensity high | Halogen Lamp 調整 3400 | SWO-165102 | Adjustment | HW | Oxi module | Lamp unit | Halogen | OXIｴﾗｰ：Halogen Lamp intensity high | halogenn Lammpu intensity high | Halogen Lamp Adjustment 3400
No | 5 | Y | RAME | Z Crea Sens 2 is too high. |  | Confirmed problem upon arrival. | D20F2A504C50169D | Adjustment | HW | Met II | Crea | Zero Current | Z Crea Sens 2 is too high. |  | Confirmed problem upon arrival.
No | Visit | Complaint | RDEU | 06.02.2017 11:00 Sr. Pia 021028512912: Glu/Lac Probleme | Empfl. Glu. und Lakt. sind zu niedrig, nach einer Dekontamination am Freitag sind beide Membranen def. | keine Lactatmembranen auf der Station deshalb keine Membranierung möglich, Glukoseelektrode membraniert und Gerät kalibriert, | SWO-164971 | Assisted Replacement | Electrode | Membrane | Glu | potentiel revisit (Lac) | 02/06/2017 11:00 Sr. Pia 021028512912: Glu/Lac Problems | Empfl. Glu. and lact. are too low, after decontamination on Friday both membranes are def. | no lactate membranes on the station therefore no membraneing is possible, glucose electrode membranes and device calibrated,
No | Visit | Complaint | RDEU | Fr. Poloni 0211/4092367 meldet Einlasskalibrierfehler | Einl.kalibrierfehler, Einl.positionierfehler, | Einl.dichtung gereinigt, Gerät kalibriert, div. Konfigurationseinstellungen durchgeführt, | SWO-164944 | Adjustment | HW | Inlet module | Probe cal |  | Ms. Poloni 0211/4092367 reports inlet calibration error | Einl.kalibrierfehler, Einl.positionierfehler, | Sealing cleaned, device calibrated, various configuration settings performed,
No | 11 | Y | RAME | F Neon # too low on 2/4, 0316 and 0258. |  | Neon # too low on 2 $4, 0316 and 0258. Also tHb on Calibration is reading 613. Also, when adjusting lamp, value is not stable. | D20EB0AA4C5A987F | Defect | HW | Oxi module | Lamp unit |  | F Neon # too low on 2/4, 0316 and 0258. |  | Neon # too low on 2 $4, 0316 and 0258. Also tHb on Calibration is reading 613. Also, when adjusting lamp, value is not stable.
No | Visit | Complaint | RIND | TNT Calibration Failed, After Analyzer Showing Message Optical intensity out of range, Calibration Adjustment Analyte Outliers check failed,Optical movement Failure, Cleaned Process Wheel, Tried to reinstalled Optical Unit, Cleaned Analyzer, Reinstalled Motor, Tried to Resolve, Problem Found with Optical Unit, process wheel and Motor, need spares for Testing, waiting for spares. | Replacement of Optical Unit, Flat Cable. | Replaced and Re Installed New Optical Unit and Flat Cable with New one,perform system cleaning, Problem Resolve. | SWO-166269 | Defect | HW | Process Module | Optical Unit |  | TNT Calibration Failed, After Analyzer Showing Message Optical intensity out of range, Calibration Adjustment Analyte Outliers check failed,Optical movement Failure, Cleaned Process Wheel, Tried to reinstalled Optical Unit, Cleaned Analyzer, Reinstalled Motor, Tried to Resolve, Problem Found with Optical Unit, process wheel and Motor, need spares for Testing, waiting for spares. | Replacement of Optical Unit, Flat Cable. | Replaced and Re Installed New Optical Unit and Flat Cable with New one,perform system cleaning, Problem Resolve.
Yes | 5 | Y | RAME | L Analyzer reporting normal ranges on random samples and not on others. | Changed type of service from Techline to complaint. | Customer noticed this at the end result of LIS report. Looking further Radiance was also showing a variance between what was reported. Upon further inspection results are coming from the analyzer like that. Machine setup verified. Sample input information verified as the same. | D20FF93CC23D9D49 | [5 empty] | L Analyzer reporting normal ranges on random samples and not on others. | Changed type of service from Techline to complaint. | Customer noticed this at the end result of LIS report. Looking further Radiance was also showing a variance between what was reported. Upon further inspection results are coming from the analyzer like that. Machine setup verified. Sample input information verified as the same.
No | Visit | Complaint | RDEU | Sr. Elke (0234/5092085) meldet diverse Stati Fehler 07.02.2017/08:20Uhr Sr.Elke 02345092085: Alle Wartungsmaßnahmen ohne Erfolg. Das Gerät zeigt ständig Fehlermeldungen an. | Anwender hat SC um 7:00 ersetzt, Glu und Lac bei QC noch außerhalb | Datenspeicher überprüft, Kal und QC wiederholt, z. Zt. OK. | SWO-165391 | Assisted Replacement | SC | SC | Glu/Lac | Start up | Sr. Elke (0234/5092085) reports various Stati errors 07.02.2017/08:20uhr Sr.Elke 02345092085: All maintenance measures without success. The device constantly displays error messages. | User replaced SC at 7:00, Glu and Lac at QC still outside | Data store checked, Kal and QC repeated, currently OK.
Yes | Visit | Complaint | RDEU | Es wurden keine sp geliefert, 3 Packs waren falsch beschriftet und abgelaufen | No | 4 Lösungspacks geliefert | SWO-165534 | Re classified RMED | Logistical upSWO |  |  |  | No sp were delivered, 3 packs were incorrectly labeled and expired | No | 4 solution packs delivered
No | Visit | Complaint | RDEU | 06.02.2017 9:30 Fr. Meckenstock 02114092520: Kundenbeschwerde wegen mangelnder Betriebsbereitschaft |  | Gerät mit warmen Wasser 10- 15 mal durchgespült. Einlasssonde durch Kundeneigene ersetzt. Kal und QC sind OK. Absprache mit Kunden getroffen, dass ein ABL90 LP als Ersatz geliefert wird. Aktuelle Lot ist die YH50. | SWO-165536 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | 06.02.2017 9:30 Fr. Meckenstock 02114092520: Customer complaint due to lack of operational readiness |  | Device rinsed with warm water 10-15 times. inlet probe replaced by customer-owned. Kal and QC are OK. agreement with customers that an ABL90 LP will be delivered as a replacement. The current lot is the YH50.
Yes | Visit | Complaint | RDEU | 30.01.2017/12:30Uhr Frau Watmunt 03957752405: QC-Fehler bei pH und pCO2 bei Level A und C | . | LP zum Kunden gebracht, da kein neuer vorrätig ist bzw. Bestellung nicht kam | SWO-165519 | Re classified RMED | Logistical upSWO |  |  |  | 30.01.2017/12:30Uhr Frau Watmunt 03957752405: QC-Fehler bei pH und pCO2 bei Level A und C | . | LP brought to the customer as no new stock is available or order did not arrive
Yes | 6 | Y | RPAC | Blank | Sensor cassette failure (1361) - Internal reference electrode error in sensor cassette |  | D21007B4C22C7672 | Re classified RMED | Ewen mail | Thursday, February 09, 2017 12:56 AM | Reference |  | Blank | Sensor cassette failure (1361) - Internal reference electrode error in sensor cassette
Yes | 3 | Y | RPAC | Blank | Glucose Calibration failure |  | D20FE80CC2057665 | Re classified RMED | Ewen mail | Thursday, February 09, 2017 12:56 AM | Glu | LEH Mail | Blank | Glucose Calibration failure
No | Visit | Complaint | RCHN | 温度错误 | 温度错误，借备件902-972一个 |  | SWO-165118 | [5 empty] | Temperature error | Temperature error, borrowed spare parts 902-972 one
No | 9 | Y | RAME | Z Humming noise from rear of analyzer |  | Found defective cooling on EMI computer. | D2102C23C2284343 | [5 empty] | Z Humming noise from rear of analyzer |  | Found defective cooling on EMI computer.
No | Visit | Complaint | RIND | Temperature Malfunction. | Temperature Malfunction. | Found that Sensor cassette got leaked the solution, do to that found T M error, replaced the SC but same error present, found connector problem, need to replace with new SBC ASAP. failed SC details 300 $30, Lot 230569, Sn, 1520247. | SWO-166197 | [5 empty] | Temperature Malfunction. | Temperature Malfunction. | Found that Sensor cassette got leaked the solution, do to that found T M error, replaced the SC but same error present, found connector problem, need to replace with new SBC ASAP. failed SC details 300 $30, Lot 230569, Sn, 1520247.
Yes | Visit | Complaint | RIND | Temperature Malfunction. | Temperature Malfunction. | Taken the sensor connector board from standby analyzer replaced the same, performed the zeroing, installed the new SC found instrument is functional. | SWO-166199 | Re classified RMED | Double entry |  |  |  | Temperature Malfunction. | Temperature Malfunction. | Taken the sensor connector board from standby analyzer replaced the same, performed the zeroing, installed the new SC found instrument is functional.
No | Visit | Complaint | RCHN | 报5784错误 | 报5784错误 | 清除测试卡角内的碎片，更换新测试卡后修复 | SWO-165469 | Defect | Consumables | Test Kit | Cartridge | 5784 (Broken) | Reported 5784 errors | Reported 5784 errors | Remove debris from the test card angle and repair it after replacing the new test card
No | Visit | Complaint | RIND | Glucose parameter calibration error. | Check the calibration logs. | Replace the membrane & recalibrate system. | SWO-166256 | Assisted Replacement | Electrode | Membrane | Glu |  | Glucose parameter calibration error. | Check the calibration logs. | Replace the membrane & recalibrate system.
No | Visit | Complaint | RDEU | 06.02.2017 Herr Thiele 036458543434 meldet, dass die pCO2 Elektrode vermutlich defekt ist. | Fehlersuche Sporatische Kal.-driftfehler THb | pCO2-Elektrode erneuert (alt: DK-02) Diverse Pumpenschläuche erneuert | SWO-165414 | Assisted Replacement | Electrode | Electrode | pCO2 |  | 06.02.2017 Mr. Thiele 036458543434 reports that the pCO2 electrode is probably defective. | Fehlersuche Sporatische Kal.-driftfehler THb | pCO2 electrode renewed (old: DK-02) Various pump hoses renewed
No | Visit | Complaint | RIND | instrument showing printer error. | printer not rotating properly, | replaced the printer head with new printer head now printer working. | SWO-166268 | Defect | HW | Printer Module | Printer head |  | instrument showing printer error. | printer not rotating properly, | replaced the printer head with new printer head now printer working.
No | Visit | Complaint | RDEU | 31.01.2017; Fr. Johnson; Tel.: 061157715490; Schlauch undicht/abgesprungen | Schlauch erneut abgesprungen | Alle Fliesswege erneuert , Probe Messungen beobachtet. Kein Fehler feststellbar | SWO-165289 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | 31.01.2017; Ms. Johnson; Tel.: 061157715490; Hose leaking/bounced | Hose bounced off again | All flow paths renewed, sample measurements observed. No error detectable
No | 17 | Y | RAME | Z AutoCheck lid keeps opening and closing |  | Found sensor dirty | D20FE02AC24D73CD | Cleaned | HW | Auto Check Module | Sensor |  | Z AutoCheck lid keeps opening and closing |  | Found sensor dirty
No | Visit | Complaint | RDEU | 07.02.2017, Fr. Dieburg, Tel. 017615322529: Gerät in rot und HALT | Oximetrie Fehler, Application Error NHSTW32.exe | Oxi Modul geprüft, AC's wiederholt ok, SW geprüft, Viren Scan negativ, Funktionskontrolle ok | SWO-165374 | Cleaned | Fluidic pathway | Oxi module | Oxi module, Cleaned |  | 07.02.2017, Fr. Dieburg, Tel. 017615322529: Device in red and HALT | Oximetrie Fehler, Application Error NHSTW32.exe | Oxi module checked, AC's repeated ok, SW checked, virus scan negative, function control ok
No | 4 | Y | RPNZ | Blank | Screen Locked - Disconnect AU |  | D222BF81AC127EB4 | Defect | HW | DMS_PC | Touch screen |  | Blank | Screen Locked - Disconnect AU
No | 9 | Y | RFRA | Blank | erreur 1020 : température incubateur hors gamme |  | D210019EC2052542 | Defect | HW | Incubator Module | PCB, Incubator | 1020 | Blank | Error 1020: out-of-range incubator temperature
No | Visit | Complaint | RDEU | 30.01.2017/13:19Uhr Frau Hellberger 06044615776: Das Gerät hat am Morgen den Fehler 3412 angezeigt und wollte einen Neustart. 06.02.2017, Frau Hof, Tel. 06044615906: Das Gerät hat wieder einen Neustart verlangt mit der Meldung 1999, UNDEFINED_LOOKUP_STRING | Überprüfung nach mehreren Fehlern | Fehlerursache beschränkt sich auf Verstopfung der Abfallstrecke und Needle Wash Unit. Reinigungsprozeduren der Pumpen 1&2 und Pumpen 4&5 durchgeführt. Pumpenkopf P5 ausgetauscht. Aufgrund des Fehler 1999 wurde das Antriebsrad des Incubatorwheels erneuert. Initialisierungsphase. Test im Servicemenü ohne Beanstandung. Blankmessungen durchgeführt und mittels Tool analysiert. Alle 5 Testmessungen unter 1600, i.O QC in allen Parametern ohne Beanstandung Sporadische Fehlermeldung BNP Sensor kaputt oder nicht vorhanden. Sensor wird bestellt und nach Erhalt eingebaut. | SWO-165407 | Cleaned | HW | Aspiration Module | Needle Wash Unit | 1999 | 30.01.2017/13:19Pm Ms Hellberger 06044615776: The device displayed error 3412 in the morning and wanted a restart. 06.02.2017, Mrs. Hof, Tel. 06044615906: The device has requested a restart with the message 1999, UNDEFINED_LOOKUP_STRING | Checking for multiple errors | The cause of the error is limited to blockage of the waste section and Needle Wash Unit. cleaning procedures of pumps 1&2 and pumps 4&5. Pump head P5. Due to the fault in 1999, the drive wheel of the Incubatorwheel was renewed. Initialization phase. Test in the service menu without complaint. blank measurements and analyzed using a tool. All 5 test measurements under 1600, i.O QC in all parameters without complaint Sporadic error message BNP sensor broken or non-existent. Sensor is ordered and installed upon receipt.
No | Visit | Complaint | RLTD | mohammed, - 01895 828 757 | Glucose keeps failing sensitivity even with a good SC fitted. | Cleaned sensor pins on interface PCB on power up now has temp error top plate out of range. Unable to resolve issue parts required. | SWO-165708 | Revisit required | HW | Basic Unit | PCB, Sensor interface |  | mohammed, - 01895 828 757 | Glucose keeps failing sensitivity even with a good SC fitted. | Cleaned sensor pins on interface PCB on power up now has temp error top plate out of range. Unable to resolve issue parts required.
No | 1 | Y | RFRA | Blank | blocage du dispositif suite à bouchage |  | D2113BD0380DD6D8 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | Blank | locking the device as a result of plugging
No | 4 | Y | RAME | P FREQUENT FLUIDIC BLOCKAGE ERRORS. |  | CONFIRMED ISSUE WHEN ONSITE. | D2117DF13841E546 | [5 empty] | P FREQUENT FLUIDIC BLOCKAGE ERRORS. |  | CONFIRMED ISSUE WHEN ONSITE.
No | Visit | Complaint | RCHN | 打印机无法工作 | 打印机无法工作 | 更换备用机上的903-001后修复(by tim) | SWO-165493 | [5 empty] | The printer does not work | The printer does not work | Repair after replacing 903-001 on the standby (by tim)
Yes | 11 | Y | RAME | 6 CA not calibrated on analyzer. Customer states that the analzyer is failing system cycles for muliple parameters. |  | Spoke with Bobbie and found that the CA failed to calibrate days after SC being replaced. Arrived onsite and verifed the failure. | D2117BA53808375A | [5 empty] | 6 CA not calibrated on analyzer. Customer states that the analzyer is failing system cycles for muliple parameters. |  | Spoke with Bobbie and found that the CA failed to calibrate days after SC being replaced. Arrived onsite and verifed the failure.
No | Visit | Complaint | RDEU | 07.02.2017 Frau Hell 03696167004 meldet, dass der Parameter Kalium nicht funktioniert. | Gerät häufig mit Fehlern belastet Kunden wünscht Austauschgerät | Kundengerät zur Überprüfung in GS Aufstellung Reparaturaustauschgerät LP und SC installiert tHb-Kal i.O. Pat-Datenbank gelöscht | SWO-165622 | [5 empty] | 07.02.2017 Ms. Hell 03696167004 reports that the potassium parameter does not work. | Device frequently burdened with errors Customer wants replacement device | Customer device for verification in GS Installation of repair replacement device LP and SC installed tHb-Kal i.O. Pat database deleted
No | 2 | Y | RAME | F Error codes 292 Turbidity too high, 581 oxi spectrum mismatch, 688 tHb too low for oximeter calculation on multiple patient samples. When the same sample is run on the analyzer next to it, results are without any errors. |  | Ran service test to check hemolyzer. Failed Hemolyzer Fres, Hemolyzer peak voltage, Hemolyzer frequency, Cavitation absorb. | D2115E983862B7BD | [5 empty] | F Error codes 292 Turbidity too high, 581 oxi spectrum mismatch, 688 tHb too low for oximeter calculation on multiple patient samples. When the same sample is run on the analyzer next to it, results are without any errors. |  | Ran service test to check hemolyzer. Failed Hemolyzer Fres, Hemolyzer peak voltage, Hemolyzer frequency, Cavitation absorb.
No | Visit | Complaint | RCHN | 电源坏了，无法开机 | 电源坏了，无法开机 ，借902-970一个。 | 更换902-970后修复(by tim) | SWO-165224 | [5 empty] | The power supply is out of order and can't turn it on. | The power supply is out of order and can't turn it on. , borrow edith of 902-970. | Repair after replacement 902-970 (by tim)
No | Visit | Complaint | RCHN | 无法读取测试卡芯片信息 | 无法读取测试卡芯片信息 | 更换测试卡正常。 | SWO-168131 | Assisted Replacement | SC | SC | chip not read |  | Failed to read test card chip information | Failed to read test card chip information | The replacement test card is OK.
No | Visit | Complaint | RCHN | 报错1999 | 报错1999，借902-978一个 |  | SWO-165564 | Cleaned | HW | Aspiration Module | Cavo Pump | 1999 | Error 1999 | Misquoted 1999, loan 902-978 one
No | Visit | Complaint | RCHN | 液路故障 3412 | 液路故障 3412，借903-329，636-561各一个。 |  | SWO-165548 | Cleaned | HW | Aspiration Module | Needle Wash Unit | 3412 | Liquid Road Fault 3412 | Liquid Road Fault 3412, borrow 903-329, 636-561 each.
No | 17 | Y | RAME | V Customer reports that the AutoCheck module cover continuously opens and closes |  | Analyzer has been powered down and was likely down overnight. AutoCheck cover $door is observed to open and close continuously. A grinding sound is also heard. Upon further investigation, fluid is observed to have entered into the AutoCheck module. | D21156D5381C251B | Defect | HW | Auto Check Module | PCB, Autocheck | Fluid damage | V Customer reports that the AutoCheck module cover continuously opens and closes |  | Analyzer has been powered down and was likely down overnight. AutoCheck cover $door is observed to open and close continuously. A grinding sound is also heard. Upon further investigation, fluid is observed to have entered into the AutoCheck module.
No | Visit | Complaint | RDEU | OSM bei QC Messungen, vier mal in Folge ohne äußeren Einfluss | Jetzt regelmäßig hohe Driften bei Hb Nullpunkt-Kal | Regelmäßig hohe Driften bei Hb Kal => per Remotezugriff diverse Aktivitäten durchgeführt, Drift nun gegen 0. Bei fortwährenden Problemen muss ggf. noch das Ocimodul ersetzt werden! | SWO-165600 | Defect | HW | Oxi module | Oxi module, Replaced |  | OSM in QC measurements, four times in a row without external influence | Now regularly high drifts at Hb zero point-kal | Regularly high drifts at Hb Kal => performed via remote access various activities, drift now against 0. In case of continuous problems, the Ocimodul may need to be replaced!
No | 4 | Y | RAME | 7 cl- need tho be changed often as quality controls run high after days of changing membrane |  | discussed instrument data: drew up the plan to clean and remembrane all electrodes. instrument study for 2 weeks no maintenance to be performed .noted salts in reference chamber and crescent shaped debris removed from cl- chamber | D2114C1B382F3DD3 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | 7 cl- need tho be changed often as quality controls run high after days of changing membrane |  | discussed instrument data: drew up the plan to clean and remembrane all electrodes. instrument study for 2 weeks no maintenance to be performed .noted salts in reference chamber and crescent shaped debris removed from cl- chamber
No | Visit | Complaint | RDEU | 07.02.2017 Herr Rawer 03518566313 meldet einen OXI-Fehler. | Datenbank geprüft, Oxi-complete-Test | Beim Eintreffen Fehlermeldung 316,258,247,248 --> ABL aus $ein --> Fehler 0258,0316 --> Oxi-complete-Test durchgeführt --> fehlerhaft Ersatzteile werden bestellt Station und MedTech informiert | SWO-165571 | Assisted Replacement | Electrode | Electrode | pCO2 |  | 02/07/2017 Mr Rawer 03518566313 reports an OXI error. | Database checked, Oxi-complete test | On arrival error message 316,258,247,248 --> ABL from $ein --> Error 0258,0316 --> Oxi-complete test performed --> faulty Spare parts are ordered Station and MedTech informed
No | Visit | Complaint | RIND | customer complaint for sample aborted | 2nd time customer complaint for sample aborted. | no action required user error. Performed the system clean procedure. | SWO-166064 | User error | User error | User error | User error |  | customer complaint for sample aborted | 2nd time customer complaint for sample aborted. | no action required user error. Performed the system clean procedure.
No | Visit | Complaint | RDEU | 30.01.2017/13:19Uhr Frau Hellberger 06044615776: Das Gerät hat am Morgen den Fehler 3412 angezeigt und wollte einen Neustart. 06.02.2017, Frau Hof, Tel. 06044615906: Das Gerät hat wieder einen Neustart verlangt mit der Meldung 1999, UNDEFINED_LOOKUP_STRING | BNP Sensor nicht erkannt oder defekt | Analysator geöffnet. BNP Temperatursensor an Cup Wash Unit ausgetauscht. Initialisierungsphase. Servicetests QC ohne Beanstandung Eingabe der Anwender zur QC Prüfung | SWO-165701 | Defect | HW | Process Module | BNP Sensor |  | 30.01.2017/13:19Pm Ms Hellberger 06044615776: The device displayed error 3412 in the morning and wanted a restart. 06.02.2017, Mrs. Hof, Tel. 06044615906: The device has requested a restart with the message 1999, UNDEFINED_LOOKUP_STRING | BNP sensor not detected or defective | Analyzer open. BNP temperature sensor to Cup Wash Unit. Initialization phase. Service tests QC without complaint Entering the users for QC examination
Yes | 4 | Y | RAME | Z Unit being sent in due to error Cal error Sensitivity that will not clear. |  | Cal error occurred on one of the sensor during sensor calibration. When inspect and replaced membrane, noticed one layer of the membrane still stuck on the sensor surface. | D2118764383301E5 | Cleaned | HW | Electrode | Membrane |  | Z Unit being sent in due to error Cal error Sensitivity that will not clear. |  | Cal error occurred on one of the sensor during sensor calibration. When inspect and replaced membrane, noticed one layer of the membrane still stuck on the sensor surface.
No | 4 | Y | RAME | Z Monitor is reading high on pCO2 compaired to the other analyzers at his location. Customer swapped electrodes with other monitors onsite and verified the error stayed with the monitor and not the electrode. |  | Ran a performance test and was able to recreate a high pCO2 measurement. | D21167F83839D9BB | Defect | HW | Tosca module |  |  | Z Monitor is reading high on pCO2 compaired to the other analyzers at his location. Customer swapped electrodes with other monitors onsite and verified the error stayed with the monitor and not the electrode. |  | Ran a performance test and was able to recreate a high pCO2 measurement.
Yes | Visit | Complaint | RDEU | SP90 liefern | SP90 liefern | SP90 geliefert | SWO-166022 | Re classified RMED | Logistical upSWO |  |  |  | Deliver SP90 | Deliver SP90 | SP90 delivered
No | Visit | Complaint | RLTD | mohammed, - 01895 828 757 | Top plate temp out of range | Replaced sensor interface PCB and test ok Calibration and QC ok. | SWO-166005 | Defect | HW | Basic Unit | PCB, Sensor interface |  | mohammed, - 01895 828 757 | Top plate temp out of range | Replaced sensor interface PCB and test ok Calibration and QC ok.
No | Visit | Complaint | RCHN | 泵轮断裂 | 泵轮断裂 | 更换924-738一个正常。(bt tim) | SWO-165837 | [5 empty] | Pump wheel break | Pump wheel break | Replace 924-738 one normal. (bt tim)
No | Visit | Complaint | RCHN | 断开交流电后机器停机 | 断开交流电后机器停机，借备件431-018C一个 |  | SWO-165861 | [5 empty] | Machine shutdown after AC is disconnected | Machine shutdown after AC disconnect, spare part 431-018C one
No | Visit | Complaint | RCHN | GLU定标错误 | GLU定标错误 | 更换测试卡 | SWO-166729 | Assisted Replacement | SC | SC | Glu | 0377 | GLU calibration error | GLU calibration error | Replace the test card
No | 4 | Y | RESP | NO CALIBRA LA pC02. NO PASA LOS CONTROLES EN pC02. |  | REVISIÓN DE HISTÓRICOS. LIMPIEZA BAHIA DE ELECTRODOS. REALIZAR TEST TÉCNICOS. REEMPLAZAR CASSETTE DE ELECTRODOS DEFECTUOSO. REALIZAR TEST TÉCNICOS. CALIBRACIONES Y CONTROLES. | D212AD33AE274CC8 | Assisted Replacement | SC | SC | pCO2 |  | DO NOT CALIBRATE THE pC02. DO NOT PASS THE CONTROLS IN pC02. |  | HISTORICAL REVIEW. CLEANING BAHIA DE ELECTRODOS. PERFORM TECHNICAL TESTS. REPLACEMENT DEFECT ELECTRODE CASSETTE. PERFORM TECHNICAL TESTS. CALIBRATIONS AND CONTROLS.
No | Visit | Complaint | RCHN | 报3410，FTC-101 | 报3410，FTC-101，借580-098一个 |  | SWO-165898 | Defect | HW | Process Module | Optical Unit | 3410 | 3410, FTC-101 | 3410, FTC-101, 580-098
No | 61 | Y | RESP | Medidas inestables en pH e iones. |  | Comprobar electrodo de referencia quitar un anillo sellador sobrante, limpiar y tensar los contactos eléctricos del electrodo. Revisión normal del equipo con limpieza de la cámara de medida y salida del desperdicio, comprobación del módulo oxi, realización de operaciones de mantenimiento y prueba con controles. | D212B3BBAE26AD2A | User error | Electrode | Electrode | Sealing ring | Double sealing ring | Unstable measures in pH and ions. |  | Check reference electrode remove a leftover sealing ring, clean and tighten the electrical contacts of the electrode. Normal equipment check with measuring chamber cleaning and waste output, oxy module check, maintenance operations and testing with controls.
No | Visit | Complaint | RKK | ｌａｃ応答不良 | ｌａｃ応答不良 | ｌａｃ電極交換 Ｃａｌ２ ＱＣ ＯＫ | SWO-167356 | Assisted Replacement | Electrode | Electrode | Lac | 376 | Lac response bad | Lac response bad | Lac electrode replacement Cal2 QC OK
No | Visit | Complaint | RDEU | OSM bei QC Messungen, vier mal in Folge ohne äußeren Einfluss | OxiSpectrumMismatch again... | Fehlersuche: Erneut Spectrum Mismatch Fehler bei QC Messungen => Oxi-Modul ersetzt, zweimal Kal1, OxiCompleteTest durchgeführt. Neustart, Kalibrationen, tHb-Kal, Funktionstest, QC Messungen OK | SWO-165964 | Defect | HW | Oxi module | Spectrometer |  | OSM in QC measurements, four times in a row without external influence | OxiSpectrumMismatch again... | Troubleshooting: Again Spectrum Mismatch Error in QC Measurements => Oxi module replaced, twice Kal1, OxiCompleteTest performed. Restart, calibrations, tHb-Kal, function test, QC measurements OK
Yes | Visit | Complaint | RDEU | Na Werte zu hoch bei QC | s. case | Aufschaltung per Remotezugriff: Bei allen QC Leveln Na immer wieder oberhalb Grenze(n). Elektrode wurde bereits membraniert, keine Änderung. Ca sporadisch (selten) ebenfalls betroffen. Einsatz ab $am Montag erforderlich! | SWO-165986 | Re classified RMED | Remote Support |  |  |  | Na values too high at QC | s. case | Remote access: At all QC levels, Na repeatedly above limit(s). Electrode has already been membraneed, no change. Ca sporadically (rarely) also affected. Use from $am Monday required!
No | 4 | Y | RPNZ | Blank | DMS Not Connecting |  | D222CEC1AC317EBF | Defect | HW | DMS_PC | Power Supply |  | Blank | DMS Not Connecting
No | 6 | Y | RPOL | Blank | Kalibracja tHb, wgranie nowych operatorów, testy kontrolne. Analizator sprawny. |  | D213CFF02401E17C | Adjustment | HW | Oxi module | tHb Cal |  | Blank | THb calibration, new operator uploads, control tests. Functional Analyzer.
Yes | 19 | Y | RAME | P Numerous fluid flow errors. |  | Remote connection is active. Contacted customer, accessed analyzer, view data logs, accessed service diagnostics, determined the fluid flow issues most likely caused by a faulty new SC just installed. | D216AE59101B21C2 | [5 empty] | P Numerous fluid flow errors. |  | Remote connection is active. Contacted customer, accessed analyzer, view data logs, accessed service diagnostics, determined the fluid flow issues most likely caused by a faulty new SC just installed.
No | 6 | Y | RAME | E Error 1079 sensor impedance error & 1070 sensor response error for pH,K+, Na+, Ca+, Cl-, also error 1406 - analyzer not horizontal. |  | Reviewed data logs, calibration failing, observed error 1406 (not horizontal) ocuring from 2 $8 to 2 $9 intermittently, no errors after 2 $9 @ 8:15AM. Sensor response errors began 2 $9 @ 00:52. SC had been replace approximately a day earlier. Customer tried a second SC with the same issues. Learned while onsite another RT lead flushed the analyzer on 2 $9 in the AM and calibration began passing. Suspect a possible unknown material passing through analyzer may have caused reported problem. | D213EF612444E088 | No fault found | No fault found | No fault found | No fault found | Cleaned Fluidic P. | E Error 1079 sensor impedance error & 1070 sensor response error for pH,K+, Na+, Ca+, Cl-, also error 1406 - analyzer not horizontal. |  | Reviewed data logs, calibration failing, observed error 1406 (not horizontal) ocuring from 2 $8 to 2 $9 intermittently, no errors after 2 $9 @ 8:15AM. Sensor response errors began 2 $9 @ 00:52. SC had been replace approximately a day earlier. Customer tried a second SC with the same issues. Learned while onsite another RT lead flushed the analyzer on 2 $9 in the AM and calibration began passing. Suspect a possible unknown material passing through analyzer may have caused reported problem.
No | 10 | Y | RAME | w les than 15 percent disk soace |  | error1307 | D213DC26243F9D13 | Adjustment | SW | DMS_PC | Disc Space | temp files deleted | w les than 15 percent disk soace |  | error1307
No | 4 | Y | RAME | w less than 15 percent disk space |  | error 1307 | D213DE31242C9D17 | Adjustment | SW | DMS_PC | Disc Space | temp files deleted | w less than 15 percent disk space |  | error 1307
No | Visit | Complaint | RLTD | Auto inlet does not raise fully. | Auto inlet does not raise fully. | Removed cover and adjusted inlet connection. Swapped back and removed spare analyser. Shut down spare analyser and returned to Glenfield workshop. Analyser sensor cassette had expired, replaced and ensured analyser passed all tests. | SWO-166208 | Adjustment | HW | Auto Inlet module | Inlet Position Cal. |  | Auto inlet does not raise fully. | Auto inlet does not raise fully. | Removed cover and adjusted inlet connection. Swapped back and removed spare analyser. Shut down spare analyser and returned to Glenfield workshop. Analyser sensor cassette had expired, replaced and ensured analyser passed all tests.
No | Visit | Complaint | RCHN | lac定标失败 | lac定标失败 | 更换测试卡后修复 | SWO-166092 | [5 empty] | lac calibration failed | lac calibration failed | Repair after replacing test card
No | 10 | Y | RAME | 3 PO2 fails level 1 QC |  | Checked all data logs for any previous issues and previous issues and found the level 1 QC fails PO2 when ran the first time. | D213AF9B2459C92A | [5 empty] | 3 PO2 fails level 1 QC |  | Checked all data logs for any previous issues and previous issues and found the level 1 QC fails PO2 when ran the first time.
No | Visit | Complaint | RCHN | 报4584，进样门开闭不正常 | 报4584，进样门开闭不正常，借903-028一个 |  | SWO-166109 | Defect | HW | Cover and Chassis | Inlet Door | 4584 | Reported 4584, into the sample door open and closed abnormal | Reported 4584, into the sample door open and closed abnormal, borrow 903-028 a
No | Visit | Complaint | RCHN | 干燥前温度超范围 | 干燥前温度超范围，借903-210一个 |  | SWO-166099 | Defect | HW | Process Module | PCB, Cup dryer | 1068 | Overrange temperature before drying | Over-range temperature before drying, borrow 903-210 one
No | Visit | Complaint | RCHN | 仪器开不了机，显示没有电源 | 仪器开不了机，显示没有电源,借910-343一个 | 更换备用机上的903-175后修复(by tim) | SWO-165789 | Defect | HW | Analyzer Control Module | Power Supply |  | The instrument won't turn on and shows no power | The instrument can't open, shows no power, borrows 910-343 one | Post-replacement 903-175 on standby (by tim)
No | 65 | Y | RPAC | Blank | Oximetry |  | D213F5D1245DBDAC | Defect | HW | DMS_PC | Power Supply |  | Blank | Oximetry
No | 10 | Y | RAME | 2 Pco2 trending lower than other unit (820) on QC |  | checked upper valves and replaced flat valve membrane and measuring chamber. | D21410A3245282E3 | Defect | HW | pH/BG Module | Valve, air |  | 2 Pco2 trending lower than other unit (820) on QC |  | checked upper valves and replaced flat valve membrane and measuring chamber.
No | 8 | Y | RPAC | Blank | Inlet leaking |  | D2175EAF8622BC68 | Defect | HW | Inlet module | Inlet Gasket | Leaking | Blank | Inlet leaking
No | 6 | Y | RPAC | Blank | Neon Lamp Intensity Too Low errors all the time. |  | D213EFFB242D10DD | Defect | HW | Oxi module | Lamp unit |  | Blank | Neon Lamp Intensity Too Low errors all the time.
No | 11 | Y | RAME | Z nipple on fluidic block |  | found solution nipple broken | D213D1C524281F52 | Defect | HW | Inlet module | Nipple broken |  | Z nipple on fluidic block |  | found solution nipple broken
No | Visit | Complaint | RKK | 電解質ﾓｼﾞｭｰﾙ下部LSｴﾗｰと、オキシLSｴﾗｰが同時に出ていたが、タンパク除去等の作業でオキシLSｴﾗｰが残った。 | ｻｰﾋﾞｽﾃｽﾄにおけるLS値が異常。 | 各ﾎﾟﾝﾌﾟﾁｭｰﾌﾞ、エアライン洗浄にて良好。 | SWO-166394 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | And electrolyte module lower LS error, but oxy LS error was out at the same time, oxyLS error remained in the work such as protein removal. | The LS value in the service test is abnormal. | Good for each pump tube and airline cleaning.
No | Visit | Complaint | RKK | シリンジ、キャピラリー共に吸わない。 | シリンジ、キャピラリー共に吸わない | ｲﾝﾚｯﾄﾌﾟﾛｰﾌﾞ、ｲﾝﾚｯﾄｶﾞｽｹｯﾄ交換。FlexQﾓｼﾞｭｰﾙの位置調整。 | SWO-166392 | Defect | HW | Inlet module | Inlet Gasket |  | Syringes and capillaries do not suck both. | Syringes and capillaries don't smoke both | Inlet probe, inlet gasket replacement. FlexQ module positioning.
No | Visit | Complaint | RIND | CKMB Results failure, Analyzer Showing error 1029 Optical Sensitivity low, Perform System cleaning and Run sample, Problem Remains, System Check failed for optical Unit, Need to replace Optical Unit, Ordered placed for testing. | System Check failed for optical Unit, Need to replace Optical Unit. CKMB results Failure. | CKMB Results failure, Analyzer Showing error 1029 Optical Sensitivity low, Perform System cleaning and Run sample, Problem Remains, order for Optical Unit for testion. | SWO-166271 | Revisit required | HW | Process Module | Optical Unit |  | CKMB Results failure, Analyzer Showing error 1029 Optical Sensitivity low, Perform System cleaning and Run sample, Problem Remains, System Check failed for optical Unit, Need to replace Optical Unit, Ordered placed for testing. | System Check failed for optical Unit, Need to replace Optical Unit. CKMB results Failure. | CKMB Results failure, Analyzer Showing error 1029 Optical Sensitivity low, Perform System cleaning and Run sample, Problem Remains, order for Optical Unit for testion.
No | Visit | Complaint | RDEU | Fehler nach dem einlegen einer Kassette | Gerät geöffnet, Sichtkontrolle, Funktionskontrolle, Systemüberprüfung. | Eine TnI Testkassette mit abgebrochenem Mitnehmer geborgen. Storagewheel ausgebaut und erfolglos nach dem Bruchstück gesucht. Evtl. War die Kassette vor dem Einlegen durch Sturz beschädigt. Gerät zusammengebaut, Funktionskontrolle | SWO-166296 | Defect | Consumables | Test Kit | Cartridge | Broken, nose | Error after inserting a cassette | Device opened, visual inspection, function control, system verification. | A TnI test cassette with aborted taker recovered. Storagewheel expanded and searched unsuccessfully for the fragment. Possibly the cassette was damaged by fall before inserting it. Device assembled, function control
Yes | 7 | Y | RPNZ | Blank | Sensor cassette replacement 1361: internal reference electrode error |  | D216B827105782BA | Re classified RMED | Telp support /PER |  |  |  | Blank | Sensor cassette replacement 1361: internal reference electrode error
No | Visit | Complaint | RSCH | Blank | Message erreur (équipement ABL90 Flex en mode de dépannage : remplacement introducteur en permanence) Demande d'intervention à la demande de Mme Ulrike Maier (sage femme) | Nettoyage de la section fluidique (suspicion caillot de sang dans la section fluidique) e.o Calibration e.o contrôle de qualité e.o | SWO-166325 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | Blank | Error message (ABL90 Flex equipment in troubleshooting mode: continuous introductory replacement) Request for intervention at the request of Ms. Ulrike Maier (midwife) | Cleaning the fluid section (suspicion blood clot in the fluid section) e.o Calibration e.o quality control e.o
No | Visit | Complaint | RDEU | Es können keine Patienten gescannt werden. | Der Barcode Scanner ist ohne Funktion. | Den Scanner getauscht. Danach Kal,QC durchgeführt. Scanner Funktion geprüft ok. | SWO-166542 | Defect | HW | Basic Unit | Barcode reader, display |  | Patients cannot be scanned. | The barcode scanner is without function. | Replace the scanner. Then Kal,QC performed. Scanner function checked ok.
No | Visit | Complaint | RKK | リーク検出 | 現地の装置確認 | パーツキット交換 | SWO-166677 | Defect | HW | Oxi module | Pinch valve |  | Leak detection | Local equipment check | Parts kit replacement
No | 8 | Y | RAME | P Replace Solution Pack. |  | Instrument was asking customer to replace solution pack multiple times. Customer has already replaced the inlet gasket holder. Found bent inlet probe. | D21F163B4A5C552B | Replaced | HW | Inlet module | Probe bend |  | P Replace Solution Pack. |  | Instrument was asking customer to replace solution pack multiple times. Customer has already replaced the inlet gasket holder. Found bent inlet probe.
No | Visit | Complaint | RLTD | Please call Katrina Gregg, 0131 537 1664/1665 Customer called back at 01:12 on 11Feb2017. Craig Stout | Black tubing between oximetry and sensor bay leaking | Replaced tubing using non stock spare, fitted original sc from spare analyser | SWO-166594 | Defect | HW | Oxi module | Pinch valve tube |  | Please call Katrina Gregg, 0131 537 1664/1665 Customer called back at 01:12 on 11Feb2017. Craig Stout | Black tubing between oximetry and sensor bay leaking | Replaced tubing using non stock spare, fitted original sc from spare analyser
No | Visit | Complaint | RDEU | 10.02.2017, 13:07 Uhr ==> Anrufer Schw. Katrin, Tel.: 03731787523 ==> SP muss erneuert werden, jedoch können Chipdaten des neuen SP nicht gelesen werden | Lieferung eines LP | neu eingesetzter LP Lot YB-53 meldet Chipdaten können nicht gelesen werden, Keine LP mehr vorrätig, bei eintreffen ist Gerät eingeschaltet und meldet einen inkonsistenten Datenbank Fehler -> LP eingesetzt , Startphase abgewartet, pCO2 & Glu $Lac Kal Fehler -> SC sollte ersetzt werden ( fällig am 15.02) keine neue vorrätig -.>Schw. Kathrin kommt heute nachmittag | SWO-166553 | Assisted Replacement | SP | SP | chip not read |  | 10.02.2017, 13:07 ==> Caller Schw. Katrin, Tel.: 03731787523 ==> SP must be renewed, but chip data of the new SP cannot be read | Delivery of an LP | newly inserted LP Lot YB-53 reports chip data cannot be read, No more LP in stock, when it arrives, device is switched on and reports an inconsistent database error -> LP inserted, start phase waited, pCO2 & Glu $Lac Kal error -> SC should be replaced (due on 15.02) no new in stock -.>Schw. Kathrin is coming this afternoon
No | 2 | Y | RFRA | Blank | Pb de fuite |  | D217B2C48635D009 | Replaced | HW | Auto Inlet module | Probe bend |  | Blank | Leakpb
No | Visit | Complaint | RCHN | LAC定标失败 | LAC定标失败 | 更换测试卡后修复 | SWO-168143 | [5 empty] | LAC calibration failure | LAC calibration failure | Repair after replacing test card
No | Visit | Complaint | RCHN | 接口板有锈蚀 | 接口板有锈蚀,借902-972一个 |  | SWO-166499 | [5 empty] | The interface plate is rusted. | Interface plate has rust, borrow 902-972 one
Yes | 7 | Y | RAME | P Unable to perform system cycle due to blockage |  | Error 2106 possible system blockage | D2178B27865373D8 | [5 empty] | P Unable to perform system cycle due to blockage |  | Error 2106 possible system blockage
No | 1 | Y | RCZE | Blank | z pøístroje použit CPU ESM-945 na opravu v Ústí n.O. 30.12.2016 |  | D2175743864AA02F | [5 empty] | Blank | The device uses the CPU ESM-945 to repair in Ústí n.O. 30.12.2016
No | Visit | Complaint | RCHN | CO2 GLU LAC定标不过 | CO2 GLU LAC定标不过 | 更换测试卡 | SWO-166726 | Assisted Replacement | SC | SC | Glu/Lac |  | CO2 GLU LAC calibration is not | CO2 GLU LAC calibration is not | Replace the test card
No | Visit | Complaint | RCHN | 开机无反应 | 开机无反应，借备件903-175 | 更换备件903-175后修复(By Ran) | SWO-166274 | Defect | HW | Analyzer Control Module | Power Supply |  | No reaction on | No reaction to boot, spare parts 903-175 | Replacement of spare parts 903-175 after repair (By Ran)
No | 77 | Y | RESP | ERRORES DE HARDWARE OXI |  | SE INTENTA RESTITUIR EL SISTEMA REINICIANDO EL EQUIPO EN VARIAS OCASIONES NO SIENDO ESTO EFEECTIVO. SE SUSTITUYE PLACA DE CONTROL OXI, SE EFECTUAN AJUSTES Y CALIBRACION DE tHb. CALIBRADO Y VERIFICADO CON Q.C. | D2178699861093BA | Defect | HW | Oxi module | PCB, Oxi module |  | OXI HARDWARE ERRORS |  | ATTEMPTS TO REST THE SYSTEM BY RESTARTING THE EQUIPMENT IN SEVERAL OCCASIONS NOT BEING THIS EFEECTIVE. OXI CONTROL PLACE IS SUBSTITUTE, tHb ADJUSTMENTS AND CALIBRATION ARE EFFECTED. CALIBRATED AND VERIFIED WITH Q.C.
No | 8 | Y | RPAC | Blank | Leaking from inlet |  | D2179B968638C041 | Defect | HW | Inlet module | Nipple broken |  | Blank | Leaking from inlet
No | Visit | Complaint | RDEU | 13.02.2017/08.00Uhr Frau 091318535184: Die Patientenprobe wird eingezogen, aber vorzeitig abgebrochen und nicht gemessen. Aus dem Einlass tropft Flüssigkeit. | Fehlersuche --> Einlassnadel und Einlassdichtung, sowie Messkammern und Elektroden überprüft --> alles in Ordnung | Messkammer gesäubert, mehrere Blutmessungen durchgeführt, System und Funktion i.O. | SWO-166694 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | 13.02.2017/08.00H Woman 091318535184: The patient sample is collected, but prematurely aborted and not measured. Liquid drips from the inlet. | Troubleshooting --> Inlet needle and inlet seal, as well as measuring chambers and electrodes checked --> everything in order | Measuring chamber cleaned, several blood measurements carried out, system and function in the
No | Visit | Complaint | RDEU | Oxi-LS Fehler und nach LS-Justage inlet LS-Fehler | s. case | Lufteinlass komplett verstopft => freigespült. Weitere Tätigkeiten im Rahmen der HJW (gesonderter SWO). | SWO-166549 | Cleaned | Fluidic pathway | Fluidic pathway | Air Intake |  | Oxi-LS error and after LS-Adjustment inlet LS error | s. case | Air inlet completely clogged => rinsed free. Other activities under the HJW (separate SWO).
No | Visit | Complaint | RDEU | Na Werte zu hoch bei QC | s. case | Alle in Frage kommenden Komponenten überprüft, Elektrolyt-Elektroden membraniert, Messkammern gereinigt. Ventiltest im Servicemenü, OK. Neustart, Kalibrationen, Endkontrolle: QC Messung, OK. Bei weiteren Ausreißern geht ein Datenextrakt zur Überprüfung an RDSEU | SWO-166615 | Assisted Replacement | Electrode | Membrane | Na+ |  | Na values too high at QC | s. case | All eligible components checked, electrolyte electrodes membraneed, measuring chambers cleaned. Valve test in the service menu, OK. Restart, calibrations, final control: QC measurement, OK. For further outliers, a data extract goes to RDSEU for review
No | 7 | Y | RAME | P Leak detected. |  | Found seal rings missing from pO2 electrode and Glucose electrode. | D217A4F186619B7E | User error | Electrode | Electrode | Sealing ring | Sealing ring, pO2 | P Leak detected. |  | Found seal rings missing from pO2 electrode and Glucose electrode.
No | 7 | Y | RAME | P instrument leaking from inlet and electrode chamber, pump tube maint |  | checked flow of solutions, checked waste lines and tubing, checked electrode chamber, | D218105D860E4858 | Defect | HW | pH/BG Module | Measuring chamber, Leak |  | P instrument leaking from inlet and electrode chamber, pump tube maint |  | checked flow of solutions, checked waste lines and tubing, checked electrode chamber,
No | Visit | Complaint | RDEU | 13.02.2017, Frau Domberg, Tel.: 05115322529: AQC Ampullen werden nicht gelesen | Es werden kein AC Ampullen erkannt | AC Modul demontiert und gereinigt, Funktionskontrolle ok | SWO-166899 | Cleaned | HW | Auto Check Module | AutoCheck Module |  | 13.02.2017, Mrs. Domberg, Tel.: 05115322529: AQC ampoules are not read | No AC ampoules are detected | AC module disassembled and cleaned, function control ok
No | Visit | Complaint | RDEU | Fr. Poloni 0211/4092367 meldet Einlasskalibrierfehler | Einlasssonde defekt. | Einlasssonde erneuert. | SWO-166662 | Defect | HW | Inlet module | Probe bend |  | Ms. Poloni 0211/4092367 reports inlet calibration error | Inlet probe defective. | Inlet probe renewed.
No | Visit | Complaint | RDEU | 07.02.2017 Herr Rawer 03518566313 meldet einen OXI-Fehler. | Siehe SWO | Oximodul erneuert --> Einlasssonde verbogen, gerichtet Zusatz der Lösungen war nach Wechsel nicht aktiviert--> Fehler Crea --> berichtigt Finalkey eingespielt | SWO-166458 | Defect | HW | Oxi module | Oxi module, Replaced |  | 02/07/2017 Mr Rawer 03518566313 reports an OXI error. | See SWO | Oximodul renewed --> inlet probe bent, directed Addition of solutions was not activated after change--> Error Crea --> corrected Finalkey recorded
Yes | 6 | Y | RAME | Z To collect data logs for techline case. |  | While looking at analyzer Chuck and I found inconsistencies in the setup of this analyzer and the one next to it. | D218BA24FC1F167C | [5 empty] | Z To collect data logs for techline case. |  | While looking at analyzer Chuck and I found inconsistencies in the setup of this analyzer and the one next to it.
No | Visit | Complaint | RKOR | tHb sample data error | . | CBC sample data compare test. User Pre analytical error. User training. | SWO-167043 | User error | Sample | Oxi module | tHb Cal |  | tHb sample data error | . | CBC sample data compare test. User Pre analytical error. User training.
Yes | Visit | Complaint | RDEU | Hotline - kein LP mehr vorhanden | Hotline - kein LP mehr vorhanden | LP installiert und Startphase abgewartet, neue Barcodetype zur Benutzer ID getestet, QC abgewartet | SWO-167060 | Re classified RMED | Logistical upSWO |  |  |  | Hotline - no lp available anymore | Hotline - no lp available anymore | LP installed and start-up phase awaited, new barcode type tested to user ID, QC awaited
Yes | Visit | Complaint | RDEU | Lösungen können nicht gepumpt werden | Anruf Frau Wenzel Tel 062155012534 | Kundin erklärt wie sie das Gerät manuell spült, Kundin geraten auch zusätzlich die Einlassdichtung zu erneuern. Sie wird sich wieder melden falls weitere Hilfen ötig ist. | SWO-166973 | Re classified RMED | Phone support |  |  |  | Solutions cannot be pumped | Call Mrs. Wenzel Tel 062155012534 | Customer explains how she rinses the device manually, customer also advises to renew the inlet seal. She will come back if further aid is available.
No | 11 | Y | RAME | W unit kept going into the boot screen |  | Unit would be in home screen then would reboot, customer had turned unit off and restarted the analyzer this morning | D218EA81FC199F8B | Defect | SW | Roma.ini | Reboot |  | W unit kept going into the boot screen |  | Unit would be in home screen then would reboot, customer had turned unit off and restarted the analyzer this morning
No | Visit | Complaint | RDEU | 13.02.2017 19:30 H. Paqwereau 071419996572: pCO2 Kal-Fehler 14.02.2017, Fr. Graef-Gärtner, Tel. 071419996572: pCO2 und Cl Kal Fehler nach erneutem LP und SC Wechsel | Kal. Daten und QC Analysiert. Reiner pCO2 Kal. Fehler -> QC in Kontrollgrenzen . CL Fehler wegen pCo2 Kal. Fehler. | Diverse Tubings und Einlassdichtung erneuert. System und SC gespült. Kein Erfolg. SC Lot. 573 eingesetzt. Initalisierungsfehler braucht sehr lange bis alle Parameter o.k. sind. | SWO-167051 | Assisted Replacement | SC | SC | pCO2 |  | Feb 13, 2017 7:30 PM H. Paqwereau 071419996572: pCO2 Kal Error 14.02.2017, Ms. Graef-Gärtner, tel. 071419996572: pCO2 and Cl Kal error after re-LP and SC change | Kal. Data and QC Analyzed. Pure pCO2 Kal. Error -> QC in control limits . CL error due to pCo2 Kal. Error. | Various tubings and inlet seal renewed. system and SC. No success. SC Lot. 573. Initalization errors take a long time until all parameters are OK.
No | Visit | Complaint | RLTD | The display screen has a pink hue to it. | The analyser was not reverting from battery to mains power when the mains was re-established. This is due to a tolerance issue on the signal that swaps between battery and mains supply. This issue has been corrected on new power supplies. | The power supply was replaced with a new one which has had the issue corrected. | SWO-167012 | Defect | HW | Power Supply | Power Supply |  | The display screen has a pink hue to it. | The analyser was not reverting from battery to mains power when the mains was re-established. This is due to a tolerance issue on the signal that swaps between battery and mains supply. This issue has been corrected on new power supplies. | The power supply was replaced with a new one which has had the issue corrected.
No | Visit | Complaint | RCHN | 气泡探测定标失败，漂移错误 | 气泡探测定标失败，漂移错误 | 更换测试卡和试剂包后修复 | SWO-167296 | [5 empty] | Bubble detection calibration failed, drift error | Bubble detection calibration failed, drift error | Repair after replacing test cards and reagent kits
No | Visit | Complaint | RCHN | 泵轮断裂 | 泵轮断裂，借924-738一个。 |  | SWO-166688 | [5 empty] | Pump wheel break | The pump wheel breaks, borrowing 924-738 one.
No | 3 | Y | RPNZ | Blank | QC failure |  | D218C02FFC60C408 | [5 empty] | Blank | QC failure
No | Visit | Complaint | RCHN | PO2定标超范围 | PO2定标超范围 |  | SWO-166734 | Assisted Replacement | SC | SC | pO2 |  | PO2 Calibration Over range | PO2 Calibration Over range
No | Visit | Complaint | RDEU | 07.02.2017, Fr. Dieburg, Tel. 017615322529: Gerät in rot und HALT | Analysator steht | Software XP DMS910 und Analysator SW 6.16 neu installiert, Konfiguration geladen, Funktionskontrolle ok | SWO-167258 | Defect | SW | Re install | Re install |  | 07.02.2017, Fr. Dieburg, Tel. 017615322529: Device in red and HALT | Analyzer is | Software XP DMS910 and analyzer SW 6.16 reinstalled, configuration loaded, function control ok
No | 6 | Y | RAME | V Barcode Not Reading. |  | Customer could not scan badges to log on to instrument. Recycled power with no change. Could not find operators in system. Not sure how they got deleted and they were missing. Customer is using Win7 operating system. | D21B9C4AE804DD93 | Setup | SW | Win7 | Re install |  | V Barcode Not Reading. |  | Customer could not scan badges to log on to instrument. Recycled power with no change. Could not find operators in system. Not sure how they got deleted and they were missing. Customer is using Win7 operating system.
No | 4 | Y | RPAC | Blank | Poor QC performance for pCO2. |  | D218CBCEFC27C91C | Cleaned | Electrode | Electrode | pCO2 |  | Blank | Poor QC performance for pCO2.
No | Visit | Complaint | RDEU | Lac muss alle 14 Tage neu mebraniert werden, da QC- bzw. Kal-Fehler | Lac muss alle 14 Tage membraniert werden (31.01.2017 und 08.02.2017, jeweils Lot 3433) | Lac nach Anweisung RMED membraniert, Servicedump gezogen, Kal u. QC i.O. | SWO-166878 | Assisted Replacement | Electrode | Membrane | Lac |  | Lac muss alle 14 Tage neu mebraniert werden, da QC- bzw. Kal-Fehler | Lac must be membraneed every 14 days (31.01.2017 and 08.02.2017, each lot 3433) | Lac according to instruction RMED membraneed, service dump drawn, Kal and QC i.O.
No | 3 | Y | RAME | F 316 neon number too low, 258 neon intensity too low |  | checked data logs for errors, checked lamp unit and hemolyzer, | D21948BDFC434862 | Defect | HW | Oxi module | PCB, Oxi module |  | F 316 neon number too low, 258 neon intensity too low |  | checked data logs for errors, checked lamp unit and hemolyzer,
No | Visit | Complaint | RKK | ＴＳＴ修理完了器返却。インキュベーターＰＣＢ交換 |  |  | SWO-166893 | Defect | HW | Incubator Module | PCB, Incubator |  | Return tst repair completion machine. Incubator PCB Replacement
No | 3 | Y | RFRA | Blank | Pb détection tube échantillon |  | D21900A7FC45472D | Cleaned | HW | Aspiration Module | Sensor, Vacutainer |  | Blank | Pb sample tube detection
No | Visit | Complaint | RDEU | Das gerät meldet Drucküberprüfung Vakuumfehler und Fehler bei der Pumpekalibrierung erkannt | Die Einlassdichtungseinheit wurde getestet und es sind Luft bei der visuellen Checkflow zu sehen. | ELD wurde erneuert und die Luftblasen sind verschwunden und somit die Fehlermeldung. | SWO-167064 | Defect | HW | Inlet module | Inlet Gasket |  | The device reports pressure check vacuum error and error detected during pump calibration | The inlet seal unit has been tested and there is air at the visual checkflow. | ELD has been renewed and the air bubbles have disappeared and thus the error message.
Yes | Visit | Complaint | RDEU | System Temperatur außer Bereich | Fehlermeldung System Temperatur außer Bereich. Parkplatz suche. | Analysator zerlegt und Fehlersuche durchgeführt. Fließstrecke untersucht. Kein Fehler gefunden Analysator zusammengesetzt und wieder hochgefahren. Fehlermeldung war nicht mehr vorhanden. | SWO-167145 | Re classified RMED | Same day upSWO |  |  |  | System temperature out of range | Error message System temperature out range. Find parking. | Analyzer disassembled and troubleshooting performed. flow line. No error found analyzer assembled and restarted. Error message no longer existed.
No | Visit | Complaint | RDEU | Einlassdichtung verbraucht. | Fehlerbehebung wird vom Gerät angefordert. | Einlassdichtung erneuert. Fehler beseitigt. Einlaufverfahren durchgeführt. | SWO-167734 | Defect | HW | Auto Inlet module | Inlet Gasket |  | inlet seal is used up. | Troubleshooting is requested by the device. | Inlet seal renewed. Fixed a bug. run-in procedures.
No | 3 | Y | RCZE | Blank |  |  | D21A32E5720C6890 | Defect | HW | Basic Unit | Wet sec control PCB |  | Blank
No | 12 | Y | RAME | 3 The PO2 was showing N/E |  | I checked the logs and saw a few N but then it came in on the next run | D219F43C724D4B7C | [5 empty] | 3 The PO2 was showing N/E |  | I checked the logs and saw a few N but then it came in on the next run
No | Visit | Complaint | RKK | tHbキャリブレーションエラー多発。OXIモジュール引上げ。 | tHbキャルエラー多発 | モジュール引上げ。 | SWO-168072 | Defect | HW | Oxi module | Oxi module, Replaced | tHb | Frequent tHb calibration errors. OXI module raise. | frequent tHb car error | Module pull up.
Yes | 19 | Y | RPAC | Blank | Semi Annual Service |  | D21A00B8722B32D6 | Re classified RMED | PM |  |  |  | Blank | Semi Annual Service
Yes | 19 | Y | RAME | Z inlet clip broken |  | inlet probe clip broken | D21A1D4F72247888 | [5 empty] | Z inlet clip broken |  | inlet probe clip broken
No | Visit | Complaint | RDEU | 13.02.2017; Hr. Odins; Tel.: 03855205818; FlexQ Positionierungsfehler | Kundengespräch und Fehlersuche durchgeführt. FlexQ-Modul erneuert. Justagen, Kalibrationen und Qualitätskontrollen in Ordnung. | FlexQ-Modul erneuert. | SWO-167178 | Defect | HW | FlexQ Module | FlexQ Module |  | 13.02.2017; Mr. Odins; Tel.: 03855205818; FlexQ Positioning Error | customer interview and troubleshooting. FlexQ module is renewed. adjustments, calibrations and quality controls in order. | FlexQ module is renewed.
No | Visit | Complaint | RLTD | Lisa Tyler 01865 221596 | Trying another display had not worked and the analyser is still experiencing problems with the display freezing and strange things happening when buttons are pressed on the screen. | A complete software reinstallation was done to see if this resolves the problem. If the problems still exist after this then the PC module will be replaced. Also while on-site reinstalled the Live Connect certificate on SCBU, ED Majors, | SWO-167484 | Defect | HW | DMS_PC | Touch screen |  | Lisa Tyler 01865 221596 | Trying another display had not worked and the analyser is still experiencing problems with the display freezing and strange things happening when buttons are pressed on the screen. | A complete software reinstallation was done to see if this resolves the problem. If the problems still exist after this then the PC module will be replaced. Also while on-site reinstalled the Live Connect certificate on SCBU, ED Majors,
No | 2 | Y | RAME | Z Customer called with a black screen on the analyzer. |  | Checked the power supply, monitor and CPU. | D21A041F72467C8E | Defect | HW | DMS_PC | CPU |  | Z Customer called with a black screen on the analyzer. |  | Checked the power supply, monitor and CPU.
No | Visit | Complaint | RDEU | 14.02.2017 8:10 Sr. Juliane 045518011151: Systemreinigung endet nicht | System überprüft und Kundengespräch durchgeführt. Kein Fehler feststellbar. | System überprüft und Kundengespräch durchgeführt. Kein Fehler feststellbar. | SWO-167242 | Defect | SW | Analyzer Control Module | Off/On |  | Feb 14, 2017 8:10 Sr. Juliane 045518011151: System cleaning does not end | system and customer call performed. No error detectable. | system and customer call performed. No error detectable.
No | Visit | Complaint | RLTD | Please call Lisa on 07796 694 479 | Investigated lack of disk space - due to build up of archived data logs. Showed customer how to delete archived data logs to free disk space. Analyser now at 44% free disk space. | Removed archived data logs to free disk space. | SWO-167670 | Adjustment | SW | DMS_PC | Disc Space | temp files deleted | Please call Lisa on 07796 694 479 | Investigated lack of disk space - due to build up of archived data logs. Showed customer how to delete archived data logs to free disk space. Analyser now at 44% free disk space. | Removed archived data logs to free disk space.
No | Visit | Complaint | RDEU | Frau Rohlfing meldet 02363/1082704 Leckage erkannt | Diverse Probleme mit Lösungen, Glucose außerhalbBereich | Datenspeicher ausgewertet, Fehler treten teilweise nach Probenaspiration gescheitert auf, mit den Anwendern im Labor über Präanalytik gesprochen. SC aus Kundenbestand ersetzt. Reklamation an RDEU gemailt | SWO-167743 | Unknown | Unknown | Unknown | Unknown |  | Ms Rohlfing reports 02363/1082704 Leak detected | Various problems with solutions, glucose outside range | Data storage evaluated, errors occur partly after sample aspiration failed, talked to the users in the laboratory about preanalysis. SC replaced from customer base. Complaint to RDEU emailed
No | Visit | Complaint | RCHN | 条码扫描失败 | 条码扫描失败，借902-992一个。 | 清洁扫描口后修复,but don't change the parts(By Ran) | SWO-166770 | [5 empty] | Barcode scanning failed | Barcode scanning failed, borrowing 902-992 one. | Repair after cleaning the scan, but don't change the parts (By Ran)
No | Visit | Complaint | RDEU | 15.02.2017/13:40Uhr Frau Kutschka 03455294154: Trotz Zubehörwechsel weiterhin Fehler Luftblasen erkannt | Interner Lösungsschlauch hat sich vom Anschlussstutzen gelöst | Fehlersuche, Befestigung Lösungsschlauch, Installation Lösungspack, Funktionskontrolle | SWO-167409 | [5 empty] | 15.02.2017/13:40pm Mrs. Kutschka 03455294154: Despite accessory changes still error air bubbles detected | Internal solution hose has detached from the connection nozzle | Troubleshooting, fixing solution hose, installation solution pack, function control
No | 10 | Y | RESP | ERRORES DE SL, BOMBAS Y FLUIDICA |  | SE EFECTUAN PRUEBAS CON DIFERENTES REACTIVOS Y LIMPIEZA DE TODO EL SISTEMA NO RESOLVIENDO EL PROBLEMA, SE DESMONTA PARCIALMENTE EL EQUIPO Y SE SUSTITUYE TUBO DE HEMOLIZADOR BAHIA, TRAS DIFERENTES PRUEBAS Y AJUSTES QUEDA TODO SOLUCIONADO. CALIBRADO Y VERFIFICADO CON Q.C. | D21B3A94E83BB106 | Defect | HW | Oxi module | Pinch valve tube |  | SL, PUMP AND FLUIDICA ERRORS |  | TESTS ARE EFFECTED WITH DIFFERENT REACTIVOS AND CLEANING OF THE ENTIRE SYSTEM DO NOT RESOLVE THE PROBLEM, THE EQUIPMENT IS PARTIALLY DISASSEMBLED AND BAHIA HEMOLIZER TUBE IS SUBSTITUTE, AFTER DIFFERENT TESTS AND ADJUSTMENTS IS ALL SOLVED. CALIBRATED AND VERFIFIED WITH Q.C.
No | Visit | Complaint | RCHN | 参比电极错误 | 检测雷度质控液水平2，PCO2结果偏高 | 更换测试卡 | SWO-167784 | Assisted Replacement | SC | SC | Reference | 1361 | Reference electrode error | Detection of lightning quality control liquid level 2, PCO2 results are high | Replace the test card
No | Visit | Complaint | RCHN | 温度超范围 | 温度超范围 | 更换测试卡后修复 | SWO-167405 | Assisted Replacement | SC | SC | Temperature | 1066 | Temperature over range | Temperature over range | Repair after replacing test card
No | Visit | Complaint | RSAF | Calcium and electrolytes failing calibrations.(Status out of limits) | Checked data logs. Visually inspected reference latch and electrode. | Replaced reference electrode. Cleaned measuring chamber and electrode latch. | SWO-169509 | Assisted Replacement | Electrode | Electrode | Reference |  | Calcium and electrolytes failing calibrations. (Status out of limits) | Checked data logs. Visually inspected reference latch and electrode. | Replaced reference electrode. Cleaned measuring chamber and electrode latch.
No | 24 | Y | RCZE | Blank |  |  | D21C7D635E5E5E56 | Assisted Replacement | Electrode | Electrode | Ca |  | Blank
No | 2 | Y | RAME | Z Touch display intermittently freezing. |  | Examined display cable and assembly. | D21B5758E84CEF5A | Defect | HW | DMS_PC | Touch screen |  | Z Touch display intermittently freezing. |  | Examined display cable and assembly.
No | 9 | Y | RFRA | Blank | Derrive Ca++/erreur 478 derrive PCO2 |  | D21B5166E8171ED7 | Defect | HW | Basic Unit | Gas module |  | Blank | Derrive Ca/error 478 derrive PCO2
Yes | 5 | Y | RFRA | Blank | SUPPORT TUBE CASSE |  | D21B4ADFE85EE920 | [5 empty] | Blank | TUBE HOLDER BREAKS
No | Visit | Complaint | RDEU | 15.02.2017, H. Benischke, Tel. 0646177446: Gerät ohne Funktion | Barcodescanner ohne Funktion | Barcodescanner im Servicemenu nicht ansprechbar. Main-Board bestellt. | SWO-168255 | Revisit required | HW | Analyzer Control Module | PCB, Main Board |  | 15.02.2017, H. Benischke, Tel. 0646177446: Device without function | Barcode scanner without function | barcode scanner in the service menu is not responsive. Main-Board.
No | Visit | Complaint | RDEU | Das Gerät fordert die Erneuerung der Sensorkassette, aufgrund eines Kalibrierfehler bei Laktat Parameter | Es wird ein Fehler der Ware vermutet | Erneuerung der Sensorkassette gegen konditionierte Ware. Die ausgetauschte Ware wird getestet | SWO-167953 | Assisted Replacement | SC | SC | Lac | 1310 | The device requires the renewal of the sensor cartridge due to a calibration error in lactate parameters | An error of the goods is suspected | Renewal of the sensor cassette against conditioned goods. The exchanged goods are tested
No | Visit | Complaint | RLTD | Annette on 0300 3001761 | Analyser failing to power on, loud long beep at regular intervals, screen completely blank. Disconnected battery unit, and wet section from the power supply - no difference. Fault is with either CPU unit or Power supply, neither part available at time of visit. Return visit to be made on Monday 20th February, with required parts to complete repair. | Unable to resolve during this visit. | SWO-167960 | Revisit required | HW | Power Supply | Power Supply |  | Annette on 0300 3001761 | Analyser failing to power on, loud long beep at regular intervals, screen completely blank. Disconnected battery unit, and wet section from the power supply - no difference. Fault is with either CPU unit or Power supply, neither part available at time of visit. Return visit to be made on Monday 20th February, with required parts to complete repair. | Unable to resolve during this visit.
No | Visit | Complaint | RDEU | Touchscreen lässt sich nicht bedienen und Interner Komunikationsfehker zwischen CPU und AU Unit. Gerät bleibt Offline. | Netzteil erneuert siehe SWO 162694 Touchscreen ohne Funktion und Gerät bleibt im Offlinemodus | Netzteil erneuert. | SWO-168189 | Defect | HW | Power Supply | Power Supply |  | Touchscreen cannot be operated and Internal communication feud between CPU and AU unit. Device remains offline. | Power supply renewed see SWO 162694 Touchscreen without function and device remains in offline mode | power supply is renewed.
Yes | 5 | Y | RAME | Z Unable to load solution pack. |  | Customer had resolved issue upon arrival. Examined solution pack loading mechanism. | D21C9FFE5E134583 | [5 empty] | Z Unable to load solution pack. |  | Customer had resolved issue upon arrival. Examined solution pack loading mechanism.
No | Visit | Complaint | RDEU | 06.02.2017 9:30 Fr. Meckenstock 02114092520: Kundenbeschwerde wegen mangelnder Betriebsbereitschaft 08.02.2017, Fr. Bayrak, Tel. 02114092620: Kunde wartet auf Service!!! 0802.2017 Herr Jatho 0211409-0 Klärung dieser Beschwerde | OXI Bügelventil verunreinigt. | Jw Kit eingebaut, Gerät forderte anschließend keine neuen LP an. OXI Bügelventil demontiert und gereinigt, Sensor Kassette erneuert. Gerät ausgeliefert. | SWO-167914 | Defect | HW | Oxi module | Pinch valve |  | 06.02.2017 9:30 Fr. Meckenstock 02114092520: Customer complaint due to lack of operational readiness 08.02.2017, Fr. Bayrak, tel. 02114092620: Customer is waiting for service!!! 0802.2017 Mr Jatho 0211409-0 Clarification of this complaint | OXI ironing valve contaminated. | Jw Kit installed, device then did not requested a new LP. OXI ironing valve disassembled and cleaned, sensor cassette renewed. device.
No | 4 | Y | RAME | Z Customer reports the printer is not working. |  | Staff was on phone with tech support and the issue could not be fixed.. | D21C63715E190CB7 | No fault found | HW | Printer Module | Printer Module |  | Z Customer reports the printer is not working. |  | Staff was on phone with tech support and the issue could not be fixed..
No | Visit | Complaint | RLTD | lisa - 0208 216 4934 | Glucose failing cal status and sensitivity | Removed interface PCB and removed and cleaned contact pins. Refitted and test ok Glucose passing calibration with good sensitivity and status. Run QC all passed ok. | SWO-167963 | Cleaned | HW | Sensor Compartment | Contacts |  | lisa - 0208 216 4934 | Glucose failing cal status and sensitivity | Removed interface PCB and removed and cleaned contact pins. Refitted and test ok Glucose passing calibration with good sensitivity and status. Run QC all passed ok.
No | 1 | Y | RPOL | Blank | Reinstalacja Windowsa na nowej karcie CF. Sprawdzenie analizatora. Kalibracje. Aparat sprawny |  | D21C7D9C5E2C1CEC | Defect | HW | CF Card | CF Card |  | Blank | Reinstalling windows in the New CF tab. Check the analyzer. Calibrations. Efficient camera
No | Visit | Complaint | RDEU | Einlassnadel verbogen | Sichtkontrolle | Kurzeinweisung, neue Bestellisten verteilt, Einlsassnadel ausgerichtet, Zählerstände und Lagerbestände aller Geräte im Haus überprüft, insg nur 3 SP90 im Haus, Funktionskontrolle | SWO-167906 | Replaced | HW | Auto Inlet module | Probe bend |  | Einlassnadel verbogen | Visual inspection | Short briefing, new orderists distributed, single-seat needle aligned, meter readings and stock of all appliances in the house checked, in total only 3 SP90 in the house, function control
No | Visit | Complaint | RDEU | 15.02.2017/13:40Uhr Frau Kutschka 03455294154: Trotz Zubehörwechsel weiterhin Fehler Luftblasen erkannt | Interne Schlauchverbindung ok, Sensorkassettte schadhaft 945-719 Lot 230927 | Fehlersuche, Installation neue Sensorkassette, Funktionskontrolle | SWO-167807 | [5 empty] | 15.02.2017/13:40pm Mrs. Kutschka 03455294154: Despite accessory changes still error air bubbles detected | Internal hose connection ok, sensor box defective 945-719 lot 230927 | Troubleshooting, installation of new sensor cartridge, function control
No | Visit | Complaint | RCHN | LAC,GLU定标失败 | LAC,GLU定标失败 | 更换测试卡正常。 | SWO-168124 | Assisted Replacement | SC | SC | Glu/Lac |  | LAC, GLU bid failed | LAC, GLU bid failed | The replacement test card is OK.
No | Visit | Complaint | RCHN | 液路传感器定标失败 | 液路传感器定标失败，借备件903-129,903-199各一个 |  | SWO-167700 | Defect | HW | Basic Unit | Cassette retainer | 529 | Failed to calibrate the liquid-road sensor | Failed to calibrate the liquid-road sensor, one for spare parts 903-129,903-199 each
No | Visit | Complaint | RCHN | 清洁液无法读取 | 清洁液无法读取，借备件905-953一个 |  | SWO-167712 | Revisit required | HW | Aspiration Module | PCB, Vacu vision |  | Cleaning fluid cannot be read | Cleaning liquid can not be read, borrowed spare parts 905-953 one
No | Visit | Complaint | RCHN | 定标卡条码扫描器无法扫描 | 定标卡条码扫描器无法扫描，借902-924一个 |  | SWO-167726 | Defect | HW | Analyzer Control Module | Barcode reader, display |  | Calibration card barcode scanner cannot be scanned | Calibration card barcode scanner can not be scanned, borrow 902-924 one
No | Visit | Complaint | RCHN | LIS系统无法连接 | LIS系统无法连接 | 自行更换备用机件910-299后修复 | SWO-167413 | Defect | HW | Analyzer Control Module | DMS_PC |  | LIS system cannot be connected | LIS system cannot be connected | Repair after replacing spare parts 910-299
Yes | Visit | Complaint | RSAF | Calcium and electrolytes failing calibrations.(Status out of limits) | Checked data logs. | Removed new reference electrode. Installed customer unit. | SWO-169516 | Re classified RMED | Courtesy/Follow up |  |  |  | Calcium and electrolytes failing calibrations. (Status out of limits) | Checked data logs. | Removed new reference electrode. Installed customer unit.
Yes | 12 | Y | RESP | Manutenção preventiva e erro de posicionamento da entrada. |  | Susbtituição da sonda , descontaminação do circuito, limpeza das camaras de medida, substituição das membranas, verificação do estado de funcionamento, estabilização e passagem do QC. | D21C9EAE5E3ECDB4 | Re classified RMED | PM |  |  |  | Preventive maintenance and positioning error of the input. |  | Replacement of the probe, decontamination of the circuit, cleaning of the measurement chambers, replacing the membranes, checking the state of operation, stabilization and passage of the QC.
No | Visit | Complaint | RDEU | 16.02.2017/10:17Uhr Frau Rabisch 093423031810: Bei einer Probenmessung verlangte das Gerät einen Neustart und hat die Pat-Probe verworfen. Nach dem Neustart meldet das Gerät Fehler 4794 | Pipettenfehler | Analysator geöffnet. Probenröhrchen im Analysator. Nadel stark verbogen. Austausch der Nadeleinheit. Prüfung im Servicemenü. Inlet Door öffnet auf der linken Seite zeitweise nicht. Fehlende Sicherungsringe an Motorwelle. Übergangsweise mit Schlauchstück fest fixiert. Sicherungsringe werden bestellt und eingebaut. Initialisierungsphase Rücksetzen des Nadelzählers QC durch Labor | SWO-167910 | Defect | HW | Aspiration Module | Bend needle |  | 16.02.2017/10:17uhr Ms. Rabisch 093423031810: During a sample measurement, the device demanded a restart and discarded the Pat sample. After reboot, the device reports error 4794 | Pipettenfehler | Analyzer open. Sample tubes in the analyzer. Needle strongly bent. Replacement of the needle unit. Check in the service menu. Inlet Door does not open temporarily on the left. Missing locking rings on motor shaft. Temporarily fixed with hose piece. Fuse rings are ordered and installed. Initialization phase Reset the needle counter QC by Laboratory
No | Visit | Complaint | RDEU | 15.02.2017, H. Benischke, Tel. 0646177446: Gerät ohne Funktion | Mainboard und Scanner müssen erneuert werden. | Mainboard und Scannereinheit erneuert. Software neu installiert. Dateien zurückgespielt. Qualitätskontrollen durchgeführt. Gerät überprüft. | SWO-168254 | Defect | HW | Analyzer Control Module | PCB, Main Board | 9098 | 15.02.2017, H. Benischke, Tel. 0646177446: Device without function | The motherboard and scanner need to be renewed. | mainboard and scanner unit. Software reinstalled. Files played back. quality controls carried out. device checked.
No | Visit | Complaint | RDEU | 16.01.2017/09:30Uhr Frau Tzschoppe 0502192105722: Das Gerät meldet kein Zugriff auf Pufferkassette | Datei Betriebsmittel geprüft. | Das besagte Lösungspack nicht in den Daten gefunden. In den Erneuerungen vom letzten Jahr ein XR-52 LP gefunden, Verfallsdatum 14.12.16. Kunde hat einen verfallenen LP installiert. | SWO-167983 | User error | Consumables | Consumables | SP expired |  | 16.01.2017/09:30pm Ms Tzschoppe 0502192105722: The device does not report access to buffer cartridge | File equipment checked. | The said solution pack not found in the data. In the renewals of last year found an XR-52 LP, expiry date 14.12.16. Customer has installed a forfeited LP.
No | Visit | Complaint | RCHN | 干燥前温度超出范围1068 | 干燥前温度超出范围1068,借903-210，903-118各一个 |  | SWO-167985 | Defect | HW | Process Module | PCB, Cup dryer | 1068 | Temperature out of range 1068 before drying | Pre-drying temperature over range 1068, borrow 903-210, 903-118 each
No | Visit | Complaint | RIND | In sample results Glu & Lac shows very low and Na+ shows abnormal.In calibration shows calibration unstable. | Check and found in calibration & sample Na+.K+.Ca2+.Glu & Lac results are abnormal | In calibration Na+,K+ shows calibration unstable.In system message shows El $met upper liquid sensor error.Check and clean liquid sensor,calibrate which is Ok except Na+calibration is Ok.But in sample Glu & Lac result shows very low.Check measuring valve and found the problem in air valve membrane which is need to be replace to resolve the problem. | SWO-170050 | Revisit required | HW | El/Met Module | Valve, air |  | In sample results Glu & Lac shows very low and Na+ shows abnormal. In calibration shows calibration unstable. | Check and found in calibration & sample Na+. K+. Ca2+. Glu & Lac results are abnormal | In calibration Na+,K+ shows calibration unstable. In system message shows El $met upper liquid sensor error. Check and clean liquid sensor,calibrate which is Ok except Na+calibration is Ok.But in sample Glu & Lac result shows very low. Check measuring valve and found the problem in air valve membrane which is need to be replace to resolve the problem.
No | Visit | Complaint | RDEU | Leckage erkannt, Einlassnadel, Einlassdichtung und Lösungspack wurde bereits durch den Kunden erneuert. Wochenenddienst Freigabe durch Frau Gerhardt und den Oberarzt. | Leckage erkannt, Einlassnadel, Einlassdichtung und Lösungspack wurde bereits durch den Kunden erneuert. Wochenenddienst Freigabe durch Frau Gerhardt und den Oberärztin Dr. Königs-El Amrawi | Drucküberprüfung Vakuum Fehler, Druck Überprüfung wird nicht bestanden. Fließwege gereinigt - keine Verstopfung. LP entfernt und das Gerät neu gestartet - Fehler besteht weiterhin. Fließwege auseinander gebaut, separat gespült und gereinigt. Drucktests im Serviceprogramm durchgeführt - Fehler besteht weiterhin. Einlassdichtung verschlissen - zieht Luft mit rein. Einlassdichtung nicht vor Ort und nicht im Auto. Reparatur nicht möglich. Ersatzteil wird morgen organisiert und eingebaut. | SWO-168129 | Revisit required | HW | Auto Inlet module | Inlet Gasket |  | Leakage detected, inlet needle, inlet seal and solution pack has already been renewed by the customer. Weekend service clearance by Ms. Gerhardt and the chief physician. | Leakage detected, inlet needle, inlet seal and solution pack has already been renewed by the customer. Weekend service released by Ms. Gerhardt and the senior physician Dr. Königs-El Amrawi | Pressure verification vacuum error, pressure check is not passed. Flow paths cleaned - no constipation. REMOVED LP and the device restarted - error still exists. Flow paths disassembled, rinsed separately and cleaned. Print tests performed in the service program - errors persist. Inlet seal worn out - draws air with it. Inlet seal not on site and not in the car. Repair not possible. Spare part will be organized and installed tomorrow.
No | Visit | Complaint | RCHN | 多次定标失败 | 多次定标失败,借903-334，902-972各一个 |  | SWO-168019 | [5 empty] | Multiple bid failures | Multiple bid failures, one each borrowed 903-334,902-972
No | Visit | Complaint | RCHN | LAC定标失败 | LAC定标失败 | 更换测试卡后修复 | SWO-168094 | [5 empty] | LAC calibration failure | LAC calibration failure | Repair after replacing test card
No | Visit | Complaint | RCHN | 964探测到有漏电流 | 964探测到有漏电流,借902-980，903-183各一个 |  | SWO-168018 | Revisit required | HW | Basic Unit | PCB, Sensor interface | 964 | 964 detected leakage current | 964 detected leakage current, borrowed 902-980, 903-183 each one
No | Visit | Complaint | RCHN | 0679 tnt错误 | 0679 tnt错误,借903-644一个 |  | SWO-168026 | Defect | HW | Basic Unit | Wet sec control PCB |  | 0679 tnt error | 0679 tnt error, borrow 903-644 one
No | Visit | Complaint | RCHN | 清洁液无法读取 | 清洁液无法读取，借备件905-953一个不能修复，追加借件903-186一个 | 更换905-953和903-186修复(By Ran) For this SWO, we didn't find 903-186 in 905-953 Kit. | SWO-167981 | Defect | HW | Aspiration Module | PCB, Vacu vision |  | Cleaning fluid cannot be read | Cleaning liquid can not be read, spare parts 905-953 one can not be repaired, additional borrowing 903-186 one | 更换905-953和903-186修复(By Ran) For this SWO, we didn't find 903-186 in 905-953 Kit.
No | Visit | Complaint | RSCH | Blank | Chargé configuration 6.13 sur version 6.16 (incompatibilité) | Réinstallation complète du software sytème et XP (main board 910) e.o Réinstallation de la section fluidique (analyzer control PCB) e.o Introduction complète manuelle de la configuration e.o Calibration e.o contrôle de qualité e.o | SWO-169549 | Defect | HW | Basic Unit | PCB, Analyzer control |  | Blank | Charge configuration 6.13 on version 6.16 (incompatibility) | Complete system and XP (main board 910) e.o software reinstallation Reinstallation of the fluidic section (analyzer control PCB) e.o Full manual introduction of the e.o configuration Calibration e.o quality control e.o
No | Visit | Complaint | RDEU | 20.02.2017 10:20Uhr Frau Marx 05542504816: Weiterhin pO2-Fehler | pO2-Qc Fehler | Lösungspack erneuert. Danach alles QC in Ordnung. Gerät überprüft. | SWO-169305 | Assisted Replacement | SP | SP | pO2 |  | 20.02.2017 10:20pm Ms Marx 05542504816: Continued pO2 errors | pO2-Qc Fehler | Solution pack renewed. After that, all QC fine. device checked.
No | Visit | Complaint | RLTD | Annette on 0300 3001761 | Analyser has a blank screen due to faulty CPU unit. | Replaced CPU and fitted original CF card. Powered on unit and CPU is working correctly with the customer's original CF card. | SWO-169304 | Defect | HW | DMS_PC | CPU |  | Annette on 0300 3001761 | Analyser has a blank screen due to faulty CPU unit. | Replaced CPU and fitted original CF card. Powered on unit and CPU is working correctly with the customer's original CF card.
No | Visit | Complaint | RDEU | 20.02.2017; Fr. Breitmar; Tel.: 021039080821; ZBV Verfallsdatum überschritten | SC intern verfallen, zweite vorhandene SC am 16.02. abgelaufen (946-009R0549). Keine weitere SC vor Ort. LP im Gerät noch bis 08.03. haltbar, zwei weitere LPs vor Ort abgelaufen (beide 944-197XT02). | Sensorkassette mitgebracht und installiert. Lösungspack zur Reserve ebenfalls vor Ort gelassen. Die Kundin reklamiert, dass nicht klar genug auf die kurze Haltbarkeit der Sensorkassetten und Lösungspacks hingewiesen wurde und erwartet daher kulanterweise kostenlosen Ersatz für 1 ST Sensorkassette und 1 ST Lösungspacks. Die Haltbarkeiten der beiden Produkte wurden besprochen. | SWO-168325 | User error | SC Expired | SC Expired | SC Expired |  | 20.02.2017; Ms. Breitmar; Tel.: 021039080821; ZBV expiry date exceeded | SC internally forfeited, second existing SC expired on 16.02. (946-009R0549). No more SC on site. LP in the device still durable until 08.03. | sensor cartridge and installed. Solution pack to the reserve also left on site. The customer complains that the short shelf life of the sensor cartridges and solution packs has not been clearly enough and therefore expects free replacements for 1 ST sensor cassette and 1 ST solution packs. The durability of the two products was discussed.
No | Visit | Complaint | RDEU | Intervention erforderlich - Meldung nach SP Tausch 20.02.2017; Hr. Heeb; Tel.: 065194783735; Intervention erforderlich 20.02.2017 08.10Herr Heeb: 065194783735: Weiterhin Fehlermeldung | Undichtigkeit im Messsystem | Oxi- Schlauchventil ausgetauscht.eine kompletteJahreswartung durchgeführt. | SWO-169372 | Defect | HW | Oxi module | Pinch valve tube |  | Intervention required - message after SP exchange 20.02.2017; Mr. Heeb; Tel.: 065194783735; Intervention needed 20.02.2017 08.10Mr Heeb: 065194783735: Continue error message | Leakage in the measuring system | Oxi hose valve replaced.a complete annual maintenance carried out.
No | Visit | Complaint | RDEU | Leckage erkannt, Einlassnadel, Einlassdichtung und Lösungspack wurde bereits durch den Kunden erneuert. Wochenenddienst Freigabe durch Frau Gerhardt und den Oberarzt. | Leckage erkannt. | Einlassdichtung undicht - erneuert. Gerät ohne LP. Es gestartet. Sensorkassette muss wieder einlaufen aufgrund der langen Standzeit. Empfindlichkeit von pCO2 kommt nicht in den Sollbereich. Sensorkassette erneuert. Neue Sensorkassette muss einlaufen, ich melde mich telefonisch und frage den Status ab. | SWO-168279 | Defect | HW | Auto Inlet module | Inlet Gasket | Leaking | Leakage detected, inlet needle, inlet seal and solution pack has already been renewed by the customer. Weekend service clearance by Ms. Gerhardt and the chief physician. | Leakage detected. | Inlet seal leaks - renewed. Device without LP. It started. Sensor cartridge must re-enter due to the long service life. Sensitivity of pCO2 does not reach the target range. Sensor cartridge is renewed. New sensor cartridge has to come in, I log in by phone and check the status.
No | 2 | Y | RFRA | Blank | Er calibration pO2 |  | D2205171C051BCC9 | Assisted Replacement | SP | SP | pO2 |  | Blank | He's not a good father
No | Visit | Complaint | RCHN | 温度错误 | 温度错误，借902-972和903-334各一个 |  | SWO-168104 | [5 empty] | Temperature error | Temperature error, borrow one each for 902-972 and 903-334
Yes | 16 | Y | RAME | V SC Smartchip Corrupt |  | Verified SC Smartchip Corrupt error. | D2207ED3C013F0E4 | [5 empty] | V SC Smartchip Corrupt |  | Verified SC Smartchip Corrupt error.
No | Visit | Complaint | RIND | failure of Ca parameter. | Continuous failure of Ca parameter. | Activated Ca parameter. Given two point calibration. Observed that sensor board was faulty. Need to replace new sensor Board. | SWO-169888 | [5 empty] | failure of Ca parameter. | Continuous failure of Ca parameter. | Activated Ca parameter. Given two point calibration. Observed that sensor board was faulty. Need to replace new sensor Board.
No | 17 | Y | RAME | W CUSTOMER IS NOTICING THEIR QC RESULTS ARE NOT AUTOMATICALLY UPLOADED TO THEIR MIDDLEWARE AFTER RECENT DATABASE CONSOLIDATION. REQUEST ASSISTANCE TO CHECK INTO SYSTEM SETTINGS AS PATIENT RESULTS WERE UPLOADING PROPERLY EXCEPT FOR QC. |  |  | D22080F0C0412180 | [5 empty] | W CUSTOMER IS NOTICING THEIR QC RESULTS ARE NOT AUTOMATICALLY UPLOADED TO THEIR MIDDLEWARE AFTER RECENT DATABASE CONSOLIDATION. REQUEST ASSISTANCE TO CHECK INTO SYSTEM SETTINGS AS PATIENT RESULTS WERE UPLOADING PROPERLY EXCEPT FOR QC.
Yes | 13 | Y | RESP | Error de base de datos, status code:2. |  | Error en la base de datos se realizan reinicios, se reinstala la versión del programa borrando la base de datos, chequeo de funcionamiento general, instalación completa de consumibles, calibración general, test conntroles de calidad he introdución de muestra, funcionando correctamente. | D22041B6C01FC151 | [5 empty] | Database error, status code:2. |  | Database errors are performed restarts, program version is reinstalled by erasing database, general operation check, complete installation of consumables, general calibration, quality conntroles test I sample introduction, working properly.
No | Visit | Complaint | RCHN | 964探测到有漏电流 | 964探测到有漏电流，借902-183，902-980各一个，因902-183没货，需要从北京发货。 |  | SWO-168053 | Revisit required | HW | Basic Unit | PCB, Sensor interface | 964 | 964 detected leakage current | 964 detected leakage current, borrow 902-183, 902-980 each one, because 902-183 is out of stock, need to ship from Beijing.
No | Visit | Complaint | RCHN | PO2定标超范围 | PO2定标超范围 ，借903-183一个。 |  | SWO-168079 | Revisit required | HW | Basic Unit | PCB, Sensor interface |  | PO2 Calibration Over range | PO2 Calibration Over range , borrowed 903-183.
No | Visit | Complaint | RCHN | O2定标不过 | O2定标不过 | 更换测试卡后修复 | SWO-168083 | Assisted Replacement | SC | SC | pO2 |  | O2 calibration but | O2 calibration but | Repair after replacing test card
No | Visit | Complaint | RCHN | LAC,CL,校准不通过，定标漂移 | LAC,CL,校准不通过，定标漂移 | 更换测试卡后修复 | SWO-168082 | Assisted Replacement | SC | SC | Lac | 210 | LAC, CL, calibration does not pass, calibration drift | LAC, CL, calibration does not pass, calibration drift | Repair after replacing test card
No | Visit | Complaint | RLTD | lorna mcdougal - 0141 451 6070 | Gas bottle stuck in chamber, will require strip down to remove Sensor reading high | Removed sensor 92-67029 to workshop for investigation Demo monitor returned to workshop and customer monitor returned | SWO-168247 | [5 empty] | lorna mcdougal - 0141 451 6070 | Gas bottle stuck in chamber, will require strip down to remove Sensor reading high | Removed sensor 92-67029 to workshop for investigation Demo monitor returned to workshop and customer monitor returned
No | Visit | Complaint | RCHN | 定标失败 | 定标失败，借CKMB定标卡944-216（12386）一张 |  | SWO-168063 | Adjustment | Consumables | Test Kit | CKMB | Calibration | Bid failed | Bid failure, borrow CKMB calibration card 944-216 (12386) one
No | Visit | Complaint | RCHN | 处理盘错误 | 2月4日报修处理盘错误，2月5日修复。现又发生同类问题，故再次借636-550一个 |  | SWO-168158 | Defect | HW | Process Module | Process Unit sensor | 5257 | Handling disk errors | February 4 Daily Repair Processing Disc Error, February 5 repair. Now there is a similar problem, so again borrow 636-550 one
No | Visit | Complaint | RCHN | 报错1031 | 报错1031。借903-364一个 |  | SWO-168109 | Defect | HW | Process Module | Optical Unit | 1031 | 1031 | 1031. Borrow 903-364 one
No | Visit | Complaint | RCHN | 报1068错误 | 报1068错误，借903-210一个 |  | SWO-168126 | Defect | HW | Process Module | PCB, Cup dryer | 1068 | Reported 1068 error | Reported 1068 error, borrow 903-210 one
No | 3 | Y | RFRA | Blank |  |  | D2207C9DC029C2CD | Defect | HW | Inlet module | Inlet module |  | Blank
No | Visit | Complaint | RIND | In sample results Glu & Lac shows very low and Na+ shows abnormal.In calibration shows calibration unstable. | Problem found in the air valve membrane which is cut. | Replace the air valve membrane by another instrument for temporary solution.Run LS adjust & Pump Calibration and run 2-point calibration which is Ok.then run OC which is also OK and run some sample which result is satisfactory. | SWO-170054 | Defect | HW | El/Met Module | Valve, air |  | In sample results Glu & Lac shows very low and Na+ shows abnormal. In calibration shows calibration unstable. | Problem found in the air valve membrane which is cut. | Replace the air valve membrane by another instrument for temporary solution. Run LS adjust & Pump Calibration and run 2-point calibration which is Ok.then run OC which is also OK and run some sample which result is satisfactory.
No | Visit | Complaint | RDEU | Seit Wochen immer wieder deutliche Abweichungen bei MetHb und HHB Messungen trotz mehrfacher tHb und diversen Reinigungsverfahren | s. case | Sporadische Aussetzer bzw. Grenzverletzungen bei MetHb und HHb => Proteinentfernung, Kal, OxiCompleteTest, zweimal hintereinander Hemolyzer Phase Regulation Failure => Hemolyzerunit ersetzt, komplett abgeglichen, Service-tHb-Kal, Kal, Funktionstest und QC Messung, OK. | SWO-169349 | Defect | HW | Oxi module | Hemolyzer |  | For weeks again and again significant deviations in MetHb and HHB measurements despite multiple tHb and various cleaning methods | s. case | Sporadic dropouts or boundary injuries in MetHb and HHb => Protein Removal, Kal, OxiCompleteTest, twice in a row Hemolyzer Phase Regulation Failure => Hemolyzerunit replaced, completely matched, Service-tHb-Kal, Kal, Function Test and QC Measurement, OK.
No | Visit | Complaint | RDEU | Einlassnadel nach oben gebogen | s. case | Einlassnadel durchgängig nach oben verbogen => gerade gerichtet. | SWO-169355 | Defect | HW | Inlet module | Probe bend |  | Inlet needle bent upwards | s. case | Inlet pin bent consistently upwards => straight enlongated.
No | Visit | Complaint | RDEU | Grüne LED in FlexQ Position 3 ohne Funktion | s. case | Grüne LED in FlexQ Position 3 ohne Funktion. Ersatzteil bestellt; wird eingebaut sofort nach Lieferung. Drei Dichtringe von FAN915-318 nachgetragen. | SWO-169357 | Revisit required | HW | FlexQ Module | Position-LED | Known issue AMF | Green LED in FlexQ Position 3 without function | s. case | Green LED in FlexQ position 3 without function. spare part ordered; will be installed immediately after delivery. Three sealing rings of FAN915-318.
No | 2 | Y | RPOL | Blank | Aparat nie aspirowa³ kontroli z automatycznego podajnika ze wzglêdu na pogiêt¹ ig³ê. Wyprostowano ig³ê. Kalibracja inletu. Analizator sprawny. |  | D220E9AFC0332789 | Defect | HW | Inlet module | Probe bend |  | Blank | The camera does not aspirator control from the automatic feeder due to the bending. The needle is straighsigned. Calibrating the inlet. Functional Analyzer.
No | 3 | Y | RAME | F Upon serviceman Installing unit, unit giving Neon lamp Intensity too low. | Changed the compalint visit tab from N to Y and changed type of service from install to complaint. | Upon arrival, verified above issue with unit Install. Found unit Exhibiting Neon Lamp low Intensity. Verified Failure in Oximetry service test. | D2217879362FF541 | Defect | HW | Oxi module | Spectrometer |  | F Upon serviceman Installing unit, unit giving Neon lamp Intensity too low. | Changed the compalint visit tab from N to Y and changed type of service from install to complaint. | Upon arrival, verified above issue with unit Install. Found unit Exhibiting Neon Lamp low Intensity. Verified Failure in Oximetry service test.
No | Visit | Complaint | RDEU | 17.02.2017; Hr. Rautmann; Tel.: 0405259503951; Unlock_17 | Kundengespräch und Fehlersuche durchgeführt. Fehlerhafte Parameterkassette mit abgebrochener Kunststoffführung entnommen und ersetzt. Justagen und Kalibrationen durchgeführt. Systemüberprüfung und Funktionstests in Ordnung. | Fehlerhafte Parameterkassette mit abgebrochener Kunststoffführung entnommen und ersetzt. Justagen und Kalibrationen durchgeführt. Systemüberprüfung und Funktionstests in Ordnung. | SWO-168169 | Defect | HW | Storage Module | Cartridge guide |  | 17.02.2017; Hr. Rautmann; Tel.: 0405259503951; Unlock_17 | customer interview and troubleshooting. Incorrect parameter cassette with broken plastic guide removed and replaced. adjustments and calibrations. System verification and functional tests are in order. | Incorrect parameter cassette with broken plastic guide removed and replaced. adjustments and calibrations. System verification and functional tests are in order.
No | Visit | Complaint | RBNL | Storagewheel kan niet in safe positie komen | Storagewheel loopt vast doordat er van één cartridge een lipje afgebroken is. Defecte cartridge verwijderd en een cuptransport check uitgevoerd: OK. | Defecte cartridge verwijderd. | SWO-168317 | Defect | Consumables | Test Kit | Cartridge | 5689 (Broken) | Storagewheel cannot enter safe position | Storagewheel crashes because a single cartridge has broken down a tab. Defective cartridge removed and a cuptransport check performed: OK. | Defective cartridge removed.
Yes | 3 | Y | RFRA | Blank |  |  | D221DF83364ED610 | Re classified RMED | PM |  |  |  | Blank
No | Visit | Complaint | RDEU | 21.02.2017; Hr. Möller; Tel.: 05622997240; inhomogene Spüllösung | Pumpenkalibrationsfehler. Lösungsfehler. | Einlassdichtung erneuert. Gerät im Servicemenü überprüft. Pumpenfehler war erst nach einer Liquidsensorkalibrierung wieder in Ordnung. Alle Test fehlerfrei, danach Neustart und Kalibrierung. Anwender hat die falschen Kapillaren bestellt, die nicht für Elektrolytmessungen geeignet sind. Gerät überprüft | SWO-170025 | Assisted Replacement | SP | SP | FT |  | 21.02.2017; Mr. Möller; Tel.: 05622997240; inhomogeneous rinsing solution | Pumpenkalibrationsfehler. Lösungsfehler. | Inlet seal renewed. device in the service menu. Pump fault was only correct after a liquid sensor calibration. All test error-free, then restart and calibration. User has ordered the wrong capillaries that are not suitable for electrolyte measurements. Device checked
No | Visit | Complaint | RDEU | 20.02.2017; Hr. Barchfeld; Tel.: 04088908780 -621; ABL hat die IP Adresse verloren | IP Adresse hat das Gerät verloren | IP Adresse neu konfiguriert und auf Sd Karte gespeichert | SWO-169540 | Setup | His/lis | His/lis | His/lis |  | 20.02.2017; Mr. Barchfeld; Tel.: 04088908780 -621; ABL has lost its IP address | IP address has lost the device | IP address reconfigured and stored on Sd card
No | Visit | Complaint | RDEU | Kein Fernzugriff möglich | Kein Fernzugriff möglich | NetOP Einstellungen geprüft, Software Windows und V3.1 MR8 neu installiert, Funktionskontrolle fehlerhaft, Kontaktaufnahme mit der IT Abteilung Herr Molau, es wurde kein Port für den Analysator angelegt, Funktionskontrolle ok | SWO-170245 | Setup | Remote Support | Remote Support | Remote Support |  | No remote access possible | No remote access possible | NetOP settings checked, software Windows and V3.1 MR8 reinstalled, function control faulty, contact with the IT department Mr. Molau, no port was created for the analyzer, function control ok
No | Visit | Complaint | RCHN | 报错1031 | 报错1031，借903-364一个 |  | SWO-169408 | Defect | HW | Process Module | Optical Unit | 1031 | 1031 | 1031, loan 903-364 one
No | Visit | Complaint | RCHN | 1068干燥前温度超出范围 | 1068干燥前温度超出范围，借备件903-210一个 |  | SWO-169433 | Defect | HW | Process Module | PCB, Cup dryer | 1068 | 1068 temperature out of range before drying | 1068 pre-drying temperature out of range, spare parts 903-210 one
Yes | Visit | Complaint | RCHN | PCT定标失败 | PCT定标失败，借定标卡944-3942张（12383） |  | SWO-169486 | Re classified RMED | PCT not released | PCT not released | PCT not released | upSWO date 21.02 | PCT calibration failed | Failure of PCT calibration, 944-3942 cards (12383)
No | Visit | Complaint | RDEU | Datum, Uhrzeit | Uhrzeit, Datum falsch | Gerät überprüft, Patch eingespielt, Datum und Uhrzeit korrigiert, verfallenes Verbrauchsmaterial ersetzt, QC außerhalb Grenzen, da das Lot keine ABL9 Grenzwerte hat | SWO-169649 | [5 empty] | Date, time | Time, date incorrect | device checked, patch checked, date and time corrected, expired consumable sumer, QC outside limits as the lot has no ABL9 limits
Yes | 7 | Y | RPAC | Blank | FAN915-318R5 |  | D221BFB8365F5FE4 | Re classified RMED | FAN |  |  |  | Blank | FAN915-318R5
No | Visit | Complaint | RDEU | Mitteilung vom WE-Dienst: Lac hält ca. 14 Tage | Lac hält ca. 14 Tage | Lac gem. Anweisung gereinigt und neu membraniert, akz. Kal abgewartet, Kal u. QC i.O. Servicedump | SWO-169879 | Assisted Replacement | Electrode | Membrane | Lac |  | Announcement from the WE service: Lac lasts about 14 days | Lac lasts about 14 days | Lac according to instructions cleaned and re-membraneed, akz. Kal waited, Kal & QC i.O. Service dump
No | Visit | Complaint | RDEU | AQC-Modul mit ratternden Geräuschen beim drehen des Karussells. | Ratternde Geräusche des AQC Moduls beim drehen des Karussells. -Garantie- | QC Flachbandkabel 14 polig erneuert, kein Erfolg, gleiches Verhalten. Vermutlich Motor oder Motoransteuerung defekt. Kabel zurückgetauscht. Vorübergehend AQC Modul aus Demogerät GS17 eingebaut bis Kundenmodul repariert werden kann. | SWO-170676 | Revisit required | HW | Auto Check Module | Auto Check Module |  | AQC module with rattling noises when rotating the carousel. | Rattling noises of the AQC module while rotating the carousel. -Guarantee | QC ribbon cable 14 poles renewed, no success, same behavior. Probably engine or motor control defective. cable replaced. Temporarily AQC module from demo device GS17 installed until customer module can be repaired.
No | Visit | Complaint | RIND | low senstivity of Ca and high control value of pco2. | senstivity of Ca was low. control value of pco2 was high. | cleaned measuring chamber. run protien removal. given two point calibration. run control. | SWO-170657 | Cleaned | Fluidic pathway | Fluidic pathway | Ca |  | low senstivity of Ca and high control value of pco2. | senstivity of Ca was low. control value of pco2 was high. | cleaned measuring chamber. run protien removal. given two point calibration. run control.
No | 4 | Y | RPNZ | Blank | No AU Connection |  | D2231B35AC14F975 | Defect | HW | Basic Unit | PCB, Analyzer control |  | Blank | No AU Connection
No | 2 | Y | RAME | 9 Lactate drifting problems |  | Customer has tried a couple of Electrode without any good results $Also, different membranes | D221AA8B3635308F | Defect | HW | El/Met Module | PCB, El/Met |  | 9 Lactate drifting problems |  | Customer has tried a couple of Electrode without any good results $Also, different membranes
No | Visit | Complaint | RBNL | Vals positieve HbF detectie | Sinds enkele weken valt het op dat er zeer frequent vals positieve analyses worden bepaald voor de HbF. Ook de MetHb resultaten zijn niet betrouwbaar. Er werd reeds een tHb kalibratie uitgevoerd en een Co-ox servicetest. Deze waren OK en deze resultaten werden doorgestuurd naar de fabrikant voor verder onderzoek (nog geen definitieve oorzaak gevonden). Ondertussen werd er lokaal besloten om preventief de hemolyser te vervangen. | Hemolyser (R0596N0002) vervangen. tHb kalibratie en Co-ox test OK CalQC : OK Er werden eveneens een 8-tal stalen bepaald Voorlopig afwachten of de problemen hiermee opgelost werden. | SWO-170303 | [5 empty] | False positive HbF detection | For several weeks, it has been said that very frequent false positive analyses are being determined for the HbF. Also the MetHb results are not reliable. A tHb calibration has already been performed and a Co-ox service test. These were OK and these results were forwarded to the manufacturer for further investigation (no definitive cause yet found). Meanwhile, it was decided locally to replace the Hemolyser preventively. | Hemolyser (R0596N0002) replaced. tHb Calibration and Co-ox test OK CalQC: OK 8 samples were also determined For now, wait and see if the problems were resolved.
No | Visit | Complaint | RIND | SC Not Initializing, Cleaned SC and Placed for Cooling and Reinstalled, Na and Hct Parameter Not calibrate after Trying Re installation for Multiple times but problem remains, Replaced SC with New one, Problem resolved. | SC Not Initializing. | Cleaned SC and Placed for Cooling and Reinstalled, Na and Hct Parameter Not calibrate after Trying Re installation for Multiple times but problem remains, Replaced SC with New one, Problem resolved. | SWO-170656 | [5 empty] | SC Not Initializing, Cleaned SC and Placed for Cooling and Reinstalled, Na and Hct Parameter Not calibrate after Trying Re installation for Multiple times but problem remains, Replaced SC with New one, Problem resolved. | SC Not Initializing. | Cleaned SC and Placed for Cooling and Reinstalled, Na and Hct Parameter Not calibrate after Trying Re installation for Multiple times but problem remains, Replaced SC with New one, Problem resolved.
No | 4 | Y | RAME | Z Oxi spectrum mismatch, error 581. |  | Checked the data logs and verified the problem. | D222C1EDAC16BB08 | [5 empty] | Z Oxi spectrum mismatch, error 581. |  | Checked the data logs and verified the problem.
No | Visit | Complaint | RCHN | LAC,GLU超出范围 | LAC,GLU超出范围 | 更换测试卡后修复 | SWO-169816 | [5 empty] | LAC, GLU out of range | LAC, GLU out of range | Repair after replacing test card
No | 3 | Y | RESP | Fuga de liquidos. |  | Desmontagem do equipamento, verificação de fuga na P4, substituição do tubos , verificação do funcionamento, tudo OK | D2231B4EAC0A9521 | Defect | HW | Aspiration Module | Tubes |  | Liquid leak. |  | Disassembly of the equipment, leakage check on P4, replacing the tubes, checking the operation, everything OK
No | Visit | Complaint | RDEU | Einlassnadel offenbar ohne äußeren Einfluss bananenkrumm und gleichmäßig nach oben gebogen | s. case | Einlassnadel erneut und offenbar ohne äußere Einwirkung bananenkrumm. Nadel gerichtet. | SWO-169945 | Defect | HW | Inlet module | Probe bend |  | Inlet needle apparently without external influence banana crumb and evenly bent upwards | s. case | Inlet needle again and apparently without external influence banana-krumm. Needle directed.
No | Visit | Complaint | RDEU | Beim Herausziehen der K-Elektrode verbleibt die Hülse in der Messkammer | s. case | Oberer Gummi-Dichtring des K Messplatzes offenbar ab Werk nicht in Position, dadurch Blockage der Elektrode. Messkammer zerlegt und Ring neu positioniert. | SWO-169947 | User error | Electrode | Electrode | K+ |  | When pulling out the K-electrode, the sleeve remains in the measuring chamber | s. case | Upper rubber sealing ring of the K measuring station apparently not in position ex works, thus blocking the electrode. measuring chamber disassembled and ring repositioned.
No | Visit | Complaint | RDEU | Immer wieder pCO2 sowie Ca zu hoch gemessen in L3 | s. case | Fehlersuche: Fließwege überprüft, Nadel gerichtet (s. a. SWO), Ca sowie Referenz membraniert, pH Dichtung ersetzt (s.a. SWO), mechanische Blockade im K Messplatz beseitigt (s.a. SWO), alle Messkammern gereinigt. Nach Neustart und Kal sowie tHb-Kal Lac Empf. unterhalb Grenze und pCO2 an Grenze => beide membraniert. Mehrere Kalibrationen, Werte stabil und Driften gegen 0, QC Messung L3 => pH und E`lyte satt außerhalb Bereich, weitere Kal 2 und Folgemessung => keine Änderung. Referenz erneut membraniert, Kal und QC Messung => Alle Werte nun fast mittig. | SWO-169954 | Assisted Replacement | Electrode | Membrane | Reference |  | Again and again pCO2 and Ca too high measured in L3 | s. case | Troubleshooting: Flow paths checked, needle directed (see a. SWO), Ca and reference membraneed, pH seal replaced (see a. SWO), mechanical blockage in the K measuring station eliminated (see s.a. SWO), all measuring chambers cleaned. After restart and Kal as well as tHb-Kal Lac Empf. below limit and pCO2 at limit => both membraneed. Several calibrations, values stable and drift against 0, QC measurement L3 => pH and e'lyte enough outside range, further scale 2 and follow-up measurement => no change. Reference re-logged, Kal and QC measurement => All values now almost central.
No | Visit | Complaint | RDEU | Erste Kal nach Routinestart immer mit Na Driftgrenzverletzung | s. case | Start der Kalibrier- $ Reinigungsroutine um 06:00 Uhr; erste Kal gegen 06:15 Uhr immer mit Na Driftgrenzverletzung. Datenextrakt gezogen zwecks Auswertung durch RDEU $ RMED | SWO-170056 | Unknown | Unknown | Unknown | Unknown |  | First Kal after routine start always with Na drift limit injury | s. case | Start of the calibration cleaning routine at 06:00 a.m.; first Kal around 06:15 always with Na drift limit injury. Data extract pulled for evaluation by RDEU - RMED
No | 4 | Y | RAME | F 316 neon number too low, 258 neon intensity too low |  | checked data logs for errors, checked hemolyzer and lamp unit, performed complete oxi test | D2234ED9AC0DAB5D | Defect | HW | Oxi module | Lamp unit |  | F 316 neon number too low, 258 neon intensity too low |  | checked data logs for errors, checked hemolyzer and lamp unit, performed complete oxi test
No | Visit | Complaint | RDEU | 21.02.2017 Herr Malter 03518561328 meldet einen Konditionierungfehler im El/Met-Modul. | Datenbank und Ventile geprüft | Ventile geprüft und gereinigt --> Neustart--> Kal i.O | SWO-169877 | Cleaned | HW | El/Met Module | Valve, air |  | 21.02.2017 Mr. Malter 03518561328 reports a conditioning error in the El/Met module. | Database and valves checked | Valves tested and cleaned --> Restart--> Kal i.O
No | 2 | Y | RAME | 2 3 Po2 and PCo2 errors during scheduled cal 2 every 8 hours. Linearity off on level 1 when running VK kit. |  | Found that only scheduled Cal 2's questioning. Cal 1's ok. Suspected Gas 2 valve problem. | D222C686AC22BEFB | Defect | HW | Basic Unit | Gas module |  | 2 3 Po2 and PCo2 errors during scheduled cal 2 every 8 hours. Linearity off on level 1 when running VK kit. |  | Found that only scheduled Cal 2's questioning. Cal 1's ok. Suspected Gas 2 valve problem.
No | Visit | Complaint | RDEU | 08.02.2017, Fr. Domberg, Tel. 05115322529: pH und Lyte bestehen Kal und QC nicht | QC Lvl.C Fehler und Kal. Empfindlichkeit Fehler, nach Installation des 944-197.XV-03. Kunde konnte den Pack erfolgreich wechseln, Fehler beseitigt. | Lösungspack zur Überprüfung abgezogen, Reklamationen erstellt. Mit den Anwendern über Lagerung und Frostindikator gesprochen. Lösungspack in der GS Werkstatt am ABL90+ geprüft, Fehler reproduzierbar. Die QC Lvl.C hat tHb und Met. bei Null, pH, pCO2 und E-Lyte. entsprechen den Kal.2 Lsg. Konzentrationen, Kal.2 Lsg. macht Empf. Fehler. Kal.2 Lsg. QC Lvl.C Lsg. müssen vertauscht sein. | SWO-170195 | Assisted Replacement | SP | SP | Pouch Mixup | XV-03 | 08.02.2017, Ms. Domberg, Tel. 05115322529: pH and Lyte do not exist Kal and QC | QC Lvl.C Error and Kal. Sensitivity error, after installing the 944-197.XV-03. Customer was able to change the pack successfully, fixed bugs. | Solution pack subtracted for review, complaints created. Talked to users about storage and frost indicator. Solution pack tested in the GS workshop on the ABL90+, errors reproducible. The QC Lvl.C has tHb and Met. at zero, pH, pCO2 and E-Lyte. correspond to the Kal.2 Lsg. concentrations, Kal.2 Lsg. makes Empf. Error. Kal.2 Lsg. QC Lvl.C Lsg. must be reversed.
Yes | 6 | Y | RPAC | Blank | Sensor cassette failure (1361) - Internal reference electrode error in sensor cassette |  | D2241DBA2214A13D | Assisted Replacement | SC | SC | Reference |  | Blank | Sensor cassette failure (1361) - Internal reference electrode error in sensor cassette
No | Visit | Complaint | RKK | インレットが閉じない | インレットが閉じない | インレットの位置調整 | SWO-170764 | Adjustment | HW | Auto Inlet module | Inlet Position Cal. |  | Inlet does not close | Inlet does not close | Positioning inlets
No | Visit | Complaint | RCHN | PO2定标失败 | PO2定标失败 | 更换测试卡正常。 | SWO-170121 | [5 empty] | PO2 calibration failure | PO2 calibration failure | The replacement test card is OK.
No | Visit | Complaint | RLTD | Contact Tara Goddard | Tested monitor on arrival - on first power on, TCM software failed to boot up all the way, left with grey screen. On/off button tested,no fault observed during visit. Monitor removed for investigation at Radiometer Ltd. workshop. Loan TCM5 monitor provided serial no. 500R0043N001. | Monitor removed for further investigation. | SWO-170625 | [5 empty] | Contact Tara Goddard | Tested monitor on arrival - on first power on, TCM software failed to boot up all the way, left with grey screen. On/off button tested,no fault observed during visit. Monitor removed for investigation at Radiometer Ltd. workshop. Loan TCM5 monitor provided serial no. 500R0043N001. | Monitor removed for further investigation.
No | Visit | Complaint | RCHN | Reflective sensor 功率值为0，1997 | Reflective sensor 功率值为0，1997,借636-539一个。 | 更换636-539一个正常。 by andrew | SWO-169773 | Defect | HW | Incubator Module | Sensor, Cup | 1997 | Reflective sensor power value is 0,1997 | The Reflective sensor power value is 0, 1997, and borrows 636-539 one. | Replace 636-539 one normal. by andrew
No | Visit | Complaint | RCHN | BNP温度传感器损坏 1235 | BNP温度传感器损坏 1235 | 更换905-908一个正常。by andrew | SWO-170346 | Defect | HW | Process Module | Sensor, BNP | 1235 | BNP temperature sensor damage 1235 | BNP temperature sensor damage 1235 | Replace 905-908 one normal. by andrew
No | Visit | Complaint | RCHN | 仪器蓝屏导致定标数据都不存在 | 仪器蓝屏导致定标数据都不存在，借944-258（12175）、944-216（12278）、944-220（12388）各一张 |  | SWO-170117 | Defect | SW | Analyzer Control Module | Off/On |  | Instrument blue screen causes calibration data to be non-existent | Instrument blue screen results in the non-existence of calibration data, with one each of the 944-258 (12175), 944-216 (12278), 944-220 (12388)
No | 4 | Y | RAME | G Customer report inlet calibration errors after power on reset for oximetry errors. |  | Found unit in hold status with inlet calibration error. tHB calibration failing with a 314 ( Oxi wavelength deviation too high) error. Found lamp intensity too high. | D224541522237EF1 | Defect | HW | Inlet module | Inlet Gasket |  | G Customer report inlet calibration errors after power on reset for oximetry errors. |  | Found unit in hold status with inlet calibration error. tHB calibration failing with a 314 ( Oxi wavelength deviation too high) error. Found lamp intensity too high.
No | Visit | Complaint | RDEU | Pumpenkal Fehler | Pumpenkal Fehler | Pumpenschläuche aus Kundenbestand erneuert. | SWO-170374 | Defect | HW | Fluidic manifold | Pump tubes | Why double | Pump scale error | Pump scale error | Pump hoses from customer base.
No | Visit | Complaint | RDEU | Pumpenkalibrierung Fehlermeldung | Fehlermeldung analysiert. | Pumpenschlauch erneuert. | SWO-170375 | Defect | HW | Fluidic manifold | Pump tubes | Why double | Pump calibration Error message | Error message analyzed. | Pump hose renewed.
Yes | 3 | Y | RAME | C Customer concerned about metHb precision on analyzer. |  | Found the calibration ok on analyzer. QC levels out. Ran QC and all levels in and OK. Ran a precision run for tHb calibration in service and found precise for both new analyzers. Ran blood sample x7 on both analzyers for a comparision of MetHb and all values commuicated to specialist Keith Clark and all looks good. | D22423E6221D60A9 | Re classified RMED | courtesy call |  |  |  | C Customer concerned about metHb precision on analyzer. |  | Found the calibration ok on analyzer. QC levels out. Ran QC and all levels in and OK. Ran a precision run for tHb calibration in service and found precise for both new analyzers. Ran blood sample x7 on both analzyers for a comparision of MetHb and all values commuicated to specialist Keith Clark and all looks good.
No | 6 | Y | RFRA | Blank | Message récupération des données...veuillez patienter de manière récurrente + erreur 4008 : LOCKINLET 24 |  | D22438262246CCC2 | Defect | SW | Analyzer Control Module | Off/On |  | Blank | Message recovering data... please wait repeatedly - error 4008: LOCKINLET 24
Yes | Visit | Complaint | RDEU | Lieferung Solution pack 903-006 |  |  | SWO-170280 | Re classified RMED | Logistical upSWO |  |  |  | Delivery Solution pack 903-006
No | Visit | Complaint | RLTD | Please call Sam or Mary-June 0208 296 4709 | On arrival analyser in Intervention required screen - message Leak detected 1342. Checked data logs - pressure test vacuum errors 1063 several times since 00:01 today. Checked fluid pathway by flushing - fluid pathway clear. Checked inlet - Inlet probe slightly bent in two directions. Carefully straightened inlet probe and refitted. Tested fluid transport in service - all tests passed. Restarted analyser and verified operation - all now ok. | Corrected bends in inlet probe to make it perfectly straight. | SWO-170626 | Defect | HW | Inlet module | Probe bend |  | Please call Sam or Mary-June 0208 296 4709 | On arrival analyser in Intervention required screen - message Leak detected 1342. Checked data logs - pressure test vacuum errors 1063 several times since 00:01 today. Checked fluid pathway by flushing - fluid pathway clear. Checked inlet - Inlet probe slightly bent in two directions. Carefully straightened inlet probe and refitted. Tested fluid transport in service - all tests passed. Restarted analyser and verified operation - all now ok. | Corrected bends in inlet probe to make it perfectly straight.
No | Visit | Complaint | RCHN | PO2定标失败 | PO2定标失败 | 更换测试卡正常。 | SWO-170357 | Assisted Replacement | SC | SC | pO2 |  | PO2 calibration failure | PO2 calibration failure | The replacement test card is OK.
No | Visit | Complaint | RCHN | 测试卡参比液泄漏 | 测试卡参比液泄漏 | 更换测试卡正常。 | SWO-170813 | Assisted Replacement | SC | SC | Reference | Leaking | Test card reference fluid leak | Test card reference fluid leak | The replacement test card is OK.
No | 56 | Y | RESP | Errores en sensibilidad de lactato. |  | Tras envío de Electrodo de Lactato el pasado 22 Feb y tras chequear equipo y ver que se ha solucioinado, se vuelve a producir el mismo problema. Chequeo de electrodo, sustitución de membrana, limpieza a fondo de preamplficador, Limpieza general de equipo y resto de preamplificadores | D225AD80983E4B69 | Assisted Replacement | Electrode | Membrane | Lac |  | Errors in lactate sensitivity. |  | After sending Lactato Electrode on 22 Feb and after checking equipment and seeing that it has been solucioined, the same problem occurs again. Electrode check, membrane replacement, in-depth cleaning of the preamprifer, General cleaning of equipment and other preamps
No | 62 | Y | RESP | errores en cal de lactato y cloro |  | Tras recibir electroodo de lactato sustituyo éste, remembranado, limpieza de camara de medida, limpieza de preamplificador con restos de oxido, al estar una de las lenguetas suelta se pega y fija, limpieza de preamplificadores, limpieza de camara de medida con bastantes restos secos de liquido en pocillo ref, preamplificadoor de Ref con muelle roto, limpieza de nipples de entrada, Ajuste de lámpara halógena tras fallo de oxi, Impresora rota pegado de pieza negra de plastico partida. instalación de Licencia Net-op sin éxito | D225AFDB98394B6D | Assisted Replacement | Electrode | Membrane | Lac |  | errors in lactate lime and chlorine |  | After receiving electroode of lactate I replace this, remembrane, cleaning chamber of measurement, cleaning of preamp with oxide residues, being one of the loose tab sticks and fixed, cleaning of preamps, cleaning of measuring camera with quite a few Dry remnants of liquid in poe ref, Ref preamplificadoor with broken spring, cleaning of inlet nipples, Adjustment of halogen lamp after failure of oxi, Broken printer glued black part part part part part part part split. unsuccessful Net-op License installation
Yes | 8 | Y | RPAC | Blank | Drift on Electrolytes after cleaning |  | D22539089835EB5E | Re classified RMED | Phone support |  |  |  | Blank | Drift on Electrolytes after cleaning
No | Visit | Complaint | RKK | ｲﾝﾚｯﾄﾌﾟﾛｰﾌﾞが曲がっている | ｲﾝﾚｯﾄﾌﾟﾛｰﾌﾞが曲がっている | ｲﾝﾚｯﾄﾌﾟﾛｰﾌﾞ交換 | SWO-170441 | Defect | HW | Inlet module | Probe bend |  | Inlet probe is bent | Inlet probe is bent | Inlet probe replacement
No | Visit | Complaint | RDEU | Grüne LED in FlexQ Position 3 ohne Funktion | s. case | Beim Eintreffen einwandfreie Funktion aller drei Sampler Pos. LEDs. LED manuell bewegt, Schlitten bewegt, keinerlei Auffälligkeiten oder Aussetzer mehr feststellbar => neue PCB vorerst NICHT eingebaut. | SWO-170505 | Defect | HW | FlexQ Module | PCB, FlexQ | via short | Green LED in FlexQ Position 3 without function | s. case | On arrival all three sampler pos. Leds. LED manually moved, slide moved, no more abnormalities or dropouts detectable => new PCB NOT installed for the time being.
No | Visit | Complaint | RCHN | 测试卡右翼缺失 2872 | 测试卡右翼缺失 2872 | 更换测试卡正常。 | SWO-171704 | Defect | Consumables | Test Kit | Cartridge | Broken, wings | Test card right wing missing 2872 | Test card right wing missing 2872 | The replacement test card is OK.
Yes | Visit | Complaint | RCHN | 结果打问号 1029 | 结果打问号 1029 ，借903-364一个。 |  | SWO-170665 | Defect | HW | Process Module | Optical Unit | 1029 | Results Question Mark 1029 | Results Question Mark 1029 , borrow edith of 903-364.
Yes | Visit | Complaint | RCHN | PCO2定标失败 | PCO2定标失败 | 更换测试卡 | SWO-171698 | Re classified RMED | OOB |  |  |  | PCO2 calibration failure | PCO2 calibration failure | Replace the test card
No | Visit | Complaint | RCHN | 系统检查失败 1031 | 系统检查失败 1031，借903-364一个。 |  | SWO-170704 | Defect | HW | Process Module | Optical Unit | 1031 | System check failed 1031 | System check failed 1031, borrowing 903-364 one.
No | Visit | Complaint | RIND | Gas pressure offset voltage out of range. | Gas2 pressure low, | replace the pressure transducer with unit chasis,check the leakage test for gas 2 and found the flow regulator valve for gas2 was not working,need to replace the the valve. | SWO-171314 | Defect | HW | Gas module | Gas module |  | Gas pressure offset voltage out of range. | Gas2 pressure low, | replace the pressure transducer with unit chasis,check the leakage test for gas 2 and found the flow regulator valve for gas2 was not working,need to replace the the valve.
No | 14 | Y | RPAC | Blank | Sensor cassette failed Glu/Lact |  | D22A9A7E705F69D2 | Assisted Replacement | SC | SC | Glu/Lac |  | Blank | Sensor cassette failed Glu/Lact
Yes | 5 | Y | RPOL | Blank | Konserwacja roczna, kalibracja ctHb w normie, pomiary techniczne w normie, analizator sprawny technicznie, przegl wa¿ny do: 02/2018r. |  | D2292EF8FA61DCD1 | [5 empty] | Blank | Annual maintenance, ctHb calibration in standard, technical measurements in standard, technically efficient analyser, browser up to: 02/2018r.
Yes | 14 | Y | RPAC | Blank | Solution pack failure - Liquid transport failed |  | D2292661FA416792 | Re classified RMED | PER |  |  |  | Blank | Solution pack failure - Liquid transport failed
Yes | 13 | Y | RPAC | Blank | Sensor cassette failure pO2 |  | D22918C7FA25678F | Re classified RMED | PER |  |  |  | Blank | Sensor cassette failure pO2
No | Visit | Complaint | RSCH | Gerät blockiert | Gerät blockiert | Allgemeine Kontrolle ok. Sensorkassette und Lösungspack entfernt. System durchgespült ok. System mit Sensorkassette durchgepült ok. Alle Tests im Servicememü durchgeführt ok. Neustart ok. Warten 55 min für Initialisierung | SWO-170957 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | Device blocked | Device blocked | General control ok. Sensor cartridge and solution pack. System flushed ok. System with sensor cassette put through ok. All tests performed in the service memü ok. Restart ok. Wait 55 min for initialization
Yes | Visit | Complaint | RDEU | Lösungspack leer und kein Ersatz vorrätig | Lösungspack erneuert | Lösungspack erneuert und installiert. | SWO-171092 | Re classified RMED | Logistical upSWO |  |  |  | Solution pack empty and no replacement in stock | Solution pack renewed | Solution pack renewed and installed.
Yes | Visit | Complaint | RDEU | Lösungspack leer und kein Vorrat. | Lösungspack geliefert und installiert. | Fehlenden Lösungspack geliefert und installiert. SC aus Kundenbestand erneuert. | SWO-171093 | Re classified RMED | Logistical upSWO |  |  |  | Solution pack empty and no stock. | solution pack is delivered and installed. | Missing solution pack delivered and installed. SC from customer base.
Yes | Visit | Complaint | RCHN | LAC定标失败 | LAC定标失败 | 更换测试卡正常。 | SWO-170824 | [5 empty] | LAC calibration failure | LAC calibration failure | The replacement test card is OK.
No | 3 | Y | RESP | 1068 - Temperatura fora do intervalo antes da secagem |  | Analise do registo de dados e mensagens do sistema. Calibração do Cup Dryer , impossivel calibrar. Substituição da cabeça das bombas, novo tentativa de calibração do Cup Dryer sem sucesso, Abertura do equipamento, limpeza e verificação da placa da unidade de secagem dos copos, limpeza da unidade de lavagem. inicialização. Realização de 2 verificações de sistema com resultados OK. Realização de dois controlos de quailidade OK. | D22951A1FA36E77A | Defect | HW | Process Module | Compressor | 1068 | 1068-temperature outside the range before drying |  | Analyze data logging and system messages. Calibration of Cup Dryer, impossible to calibrate. Replacement of the pump head, new attempt to calibrate the Cup Dryer without success, opening the equipment, cleaning and checking the plate of the drying unit of the cups, cleaning the washing unit. Startup. Realization of 2 system checks with results OK. Realization of two FINE quaility checks.
No | 1 | Y | RFRA | Blank | cache usb lateral- boot up problem and I2cissue |  | D2291052FA32B1CF | [5 empty] | Blank | cache usb lateral- boot up problem and I2cissue
No | Visit | Complaint | RIND | Rinse error. | Check the system. Check the logs. check the activity logs. check the pump calibration ,LS calibration & both calibration pass. | Check the Inlet gasket. Check the tubing .Check the upper liquid sensor & lower inlet section. Check the rinse & air valve both valve faulty. Replace the air & rinse valve. calibrate system multiple time. Calibration & qc pass. | SWO-172375 | Defect | HW | pH/BG Module | Valve, air |  | Rinse error. | Check the system. Check the logs. check the activity logs. check the pump calibration ,LS calibration & both calibration pass. | Check the Inlet gasket. Check the tubing . Check the upper liquid sensor & lower inlet section. Check the rinse & air valve both valve faulty. Replace the air & rinse valve. calibrate system multiple time. Calibration & qc pass.
No | 4 | Y | RFRA | Blank |  |  | D22912C9FA51FB94 | Defect | HW | Inlet module | Inlet module |  | Blank
No | Visit | Complaint | RKK | データのバックアップが出来ない。アプリケーションエラー発生。 | データのバックアップが出来ない | ＣＰＵボード不良の為交換。 | SWO-171697 | Defect | HW | DMS_PC | CPU |  | I can't back up my data. An application error occurred. | I can't back up my data | Replacement due to cpu board failure.
No | 8 | Y | RAME | P Excessive leaking at inlet |  | Inlet area leaking on cals, QC and blood gases. David worked on analyzer Friday. | D22965B4FA36913D | Defect | HW | Inlet module | Inlet Gasket |  | P Excessive leaking at inlet |  | Inlet area leaking on cals, QC and blood gases. David worked on analyzer Friday.
No | 5 | Y | RAME | 3 Customer reported by email zero current errors on 2 point calibrations. |  | Checked data logs for errors. Zero current errors on pO2. | D2292B10FA090F14 | Adjustment | HW | pH/BG Module | Zero current |  | 3 Customer reported by email zero current errors on 2 point calibrations. |  | Checked data logs for errors. Zero current errors on pO2.
No | 3 | Y | RAME | H Customer reported inhomogeneous errors during QC runs. |  | Checked data logs for errors. Problem is intermittent occurring on QC most of the time. . Checked fluid path. | D2294587FA600F19 | Defect | Electrode | Electrode | Sealing ring | Sealing ring, pH | H Customer reported inhomogeneous errors during QC runs. |  | Checked data logs for errors. Problem is intermittent occurring on QC most of the time. . Checked fluid path.
No | Visit | Complaint | RDEU | Probenhalterung abgebrochen | Probenhalterung abgebrochen | Analysator geöffnet - 3 Probenehmer entfernt - Funktionsüberprüfung etc. ohne Beanstandung | SWO-170931 | Defect | HW | Aspiration Module | Tube holder |  | Sample holder aborted | Sample holder aborted | Analyzer open - 3 samplers removed - function check etc. without complaint
No | Visit | Complaint | RHUN | új SC csere után a fenti hibák vannak, borzasztó státusz és érzékenység értékekkel | - | SC csere után a kalibráció stabilizálódik | SWO-171212 | Assisted Replacement | SC | SC | Reference |  | After a new SC replacement, the above errors are, terrible status and sensitivity values | - | After SC replacement, the calibration stabilates
No | Visit | Complaint | RDEU | Kalibrationsfehler | Kalibrationsfehler trotz Wechsel der Verbrauchsmaterialien | Gerät überprüft, Fließwege durchgespült, Lösungspack ersetzt, Sensor erneuert, Kal und QC o.k. Eingabe Fixe Zielwerte korregiert | SWO-171401 | Assisted Replacement | SC | SC | pCO2 & Ca |  | Kalibrationsfehler | Calibration error despite change of consumables | Device checked, flow paths flushed through, solution pack replaced, sensor renewed, Kal and QC ok. Input Fixed target values corrected
No | 23 | Y | RAME | Z Customer reporting display is dark, no actions possible, mixer light and inlet handle LEDs are lit. |  | Performed voltage checks, no issue with voltages, suspect CPU failed. | D22AD7D17029F57A | Defect | HW | DMS_PC | CPU |  | Z Customer reporting display is dark, no actions possible, mixer light and inlet handle LEDs are lit. |  | Performed voltage checks, no issue with voltages, suspect CPU failed.
No | Visit | Complaint | RDEU | sO2 in QC Level B zu niedrig | Gerät überprüft, | tHb Kal durchgeführt, alle drei QCs durchgeführt | SWO-171387 | Adjustment | HW | Oxi module | tHb Cal |  | sO2 in QC Level B too low | Device checked, | tHb Kal, all three QCs performed
Yes | 6 | Y | RPAC | Blank | Sensor cassete and Solution pack failure after backup file installed |  | D22AA66C70130B04 | Re classified RMED | FSE error |  |  |  | Blank | Sensor cassete and Solution pack failure after backup file installed
No | Visit | Complaint | RDEU | Verstopfung im Gerät, läßt sich nicht durchspülen | Verstopfung im Gerät, läßt sich nicht durchspülen ABL stand aufgrund fehlendem SP länger still | Fehlersuche, div. Schläuche erneuert, Einlaßdichtung neu Einlaßnadel neu SC erneuert, da nach wie vorFehlermeldung und Fließweg nicht frei Überprüfung Fluid Transport im ServiceProgramm Hemolyser verstopft, läßt sich nicht freispülen Einlaufverfahren fehlerhaft | SWO-171281 | Revisit required | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | Constipation in the device, cannot be flushed | Constipation in the device, cannot be flushed ABL stood still for longer due to lack of SP | Troubleshooting div. Hoses renewed, inlet seal new Inlet needle new SC is renewed, as the error message and flow path are still not free Review Fluid Transport in the ServiceProgram Hemolyser clogged, cannot be rinsed free Inlet method faulty
No | 10 | Y | RESP | El equipo se queda bloequeado cuando lleva más de dos horas sin analizar una muestra. |  | Recargado software. Comprobado buen funcionamiento del equipo y pasadas muestras | D22A9F3C7012091D | [5 empty] | The equipment is locked out when it hasn't been scanned for more than two hours. |  | Reloaded software. Proven smooth equipment performance and past samples
No | Visit | Complaint | RIND | Instrument is frequently shut down by itself. | Instrument is frequently shut down by itself. | Waited observed the few pointes, it was temperature issue or power failure, explained the same to Biomedical eng, suggested to provide the UPS point and room temperature need to maintain 18 to 22 C, it will be under observation next weel. | SWO-172389 | [5 empty] | Instrument is frequently shut down by itself. | Instrument is frequently shut down by itself. | Waited observed the few pointes, it was temperature issue or power failure, explained the same to Biomedical eng, suggested to provide the UPS point and room temperature need to maintain 18 to 22 C, it will be under observation next weel.
No | 11 | Y | RAME | 3 The PO2 would not pass after the sensor cassette was replaced. |  | Checked the data logs and verified the problem. | D22A67CC70143CB1 | [5 empty] | 3 The PO2 would not pass after the sensor cassette was replaced. |  | Checked the data logs and verified the problem.
No | Visit | Complaint | RCHN | 参比电极错误 | 参比电极错误 | 更换测试卡 | SWO-171090 | Assisted Replacement | SC | SC | Reference | 1361 | Reference electrode error | Reference electrode error | Replace the test card
No | Visit | Complaint | RCHN | 气体定标时检测到液体 | 气体定标时检测到液体 | 更换834-222修复 by jesson | SWO-171076 | Defect | HW | pH/BG Module | Valve, air | Membrane | Liquid detected during gas calibration | Liquid detected during gas calibration | Replace 834-222 Repair by jesson
No | Visit | Complaint | RIND | Chloride Sensitivity Low | Chloride Electrode Membrane Need To Be Change | Replaced The Chloride Membrane, Machine Working Ok | SWO-172357 | Assisted Replacement | Electrode | Membrane | Cl | Sensitivity | Chloride Sensitivity Low | Chloride Electrode Membrane Need To Be Change | Replaced The Chloride Membrane, Machine Working Ok
No | Visit | Complaint | RDEU | KAL-Abbruch | KAL-Abbruch Oberer pH/BG LS keine Spüllösung 509 | Analysator heruntergefahren - Überprüfung etc. Messventile - Erneuerung Luftmembran Ventil pH $BG Modul - Einstellung Hallogenlampe Oximetrie -Funktionsüberprüfung etc. | SWO-171241 | Defect | HW | pH/BG Module | Valve, air |  | KAL Cancellation | KAL Cancellation Upper pH/BG LS no rinsing solution 509 | Analyzer shut down - check etc. Measuring valves - Renewal air membrane valve pH $BG module - Setting Hallogen lamp oximetry - function check etc.
No | Visit | Complaint | RDEU | Pumpenasymetriefehler | Reinigungsarbeiten, Fehlersuche | PumpenKal im Service, probehalber Modul- und Lösungspumpenschläuche ersetzt, Einlassdichtung und Einlaßsonde ersetzt, PumpenKal, Pumpenschläuche zurück getauscht, PumpenKal ok, QK ok | SWO-171559 | Defect | HW | Fluidic manifold | Pump tubes |  | Pumpenasymetriefehler | Cleaning work, troubleshooting | PumpsKal in service, test-tested module and solution pump hoses replaced, inlet seal and inlet probe replaced, pumpsKal, pump hoses replaced, pumpsKal ok, QK ok
No | 2 | Y | RCZE | Blank |  |  | D22A56E1702A4CD0 | Defect | HW | pH/BG Module | Measuring chamber |  | Blank
Yes | 11 | Y | RFRA | Blank | investigation suite à problème de mesure de D-DIMMER |  | D22AAF7B70448858 | Re classified RMED | HW | Aspiration Module | Needle assembly | This is revisit w/o complaint | Blank | Investigation suite à problème de mesure de D-DIMMER
No | Visit | Complaint | RDEU | Kalibrationsfehler | pCO2 in Level A außerhalb Grenzen, bzw grenzwertig | Gerät überprüft, Level A und C grenzwertig, Lösungspack ersetzt | SWO-171521 | Assisted Replacement | SP | SP | pCO2 | RDEU GEMBA | Kalibrationsfehler | pCO2 in level A outside limits, or borderline | Device checked, level A and C borderline, solution pack replaced
Yes | 13 | Y | RAME | Customer complains time on analyzer loses sync with LIS and results fail to transfer over. |  | Upon arrival, verified customers stated complaint and discovered analyzer time was a minute faster and failing to sync with LIS. | D22BF0B1E6563B8D | [5 empty] | Customer complains time on analyzer loses sync with LIS and results fail to transfer over. |  | Upon arrival, verified customers stated complaint and discovered analyzer time was a minute faster and failing to sync with LIS.
Yes | 12 | Y | RAME | Z Customer complains time on analyzer loses sync with LIS and results fail to transfer over. |  | Upon arrival, verified customers stated complaint and discovered analyzer time was a minute faster and failing to sync with LIS. | D22BF43EE6243B90 | [5 empty] | Z Customer complains time on analyzer loses sync with LIS and results fail to transfer over. |  | Upon arrival, verified customers stated complaint and discovered analyzer time was a minute faster and failing to sync with LIS.
Yes | 13 | Y | RAME | Z Customer complains time on analyzer loses sync with LIS and results fail to transfer over. |  | Upon arrival, verified customers stated complaint and discovered analyzer time was a minute faster and failing to sync with LIS. | D22BF4FEE6153B92 | [5 empty] | Z Customer complains time on analyzer loses sync with LIS and results fail to transfer over. |  | Upon arrival, verified customers stated complaint and discovered analyzer time was a minute faster and failing to sync with LIS.
Yes | 12 | Y | RAME | Z Customer complains time on analyzer loses sync with LIS and results fail to transfer over. |  | Upon arrival, verified customers stated complaint and discovered analyzer time was a minute faster and failing to sync with LIS. | D22BF5B8E6043B95 | [5 empty] | Z Customer complains time on analyzer loses sync with LIS and results fail to transfer over. |  | Upon arrival, verified customers stated complaint and discovered analyzer time was a minute faster and failing to sync with LIS.
Yes | 3 | Y | RAME | F Oxi values not coming in on QC levels. |  | Examined data logs. Examined oxi module. | D22B9E3EE65330AF | [5 empty] | F Oxi values not coming in on QC levels. |  | Examined data logs. Examined Oxi module.
No | 4 | Y | RAME | G - MULTI |  | Jay called to advise they are having longstanding issues with Glu $Lac constantly falling out of calibration and failing QC | D22BCC9BE604CDC5 | Defect | HW | Basic Unit | PCB, Sensor interface |  | G - MULTI |  | Jay called to advise they are having longstanding issues with Glu $Lac constantly falling out of calibration and failing QC
Yes | 9 | Y | RPAC | Blank |  |  | D22D14815C613E66 | [5 empty] | Blank
No | 2 | Y | RFRA | Blank | Erreur d'etalonnage de la pompe |  | D22B7B14E654DA08 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | Blank | Pump stalling error
Yes | Visit | Complaint | RSAF | Sr Naransamy phoned saying that the analyzer is saying operator intervention is required. and that the other analyzer has not been working for two days now. | no more activities left on solution pack. | Tried to locate solution pack from out laying hospitals but nobody had to borrow. Used the non working analyzers solution pack to get the machine functional. | SWO-171452 | Re classified RMED | Logistical upSWO |  |  |  | Sr Naransamy phoned saying that the analyzer is saying operator intervention is required. and that the other analyzer has not been working for two days now. | no more activities left on solution pack. | Tried to locate solution pack from out laying hospitals but nobody had to borrow. Used the non working analyzers solution pack to get the machine functional.
No | Visit | Complaint | RDEU | Verstopfung im Gerät, läßt sich nicht durchspülen | Verstopfung im Gerät, läßt sich nicht durchspülen | Hemolyser verstopft Hemolyser Unit erneuert tHb Kal Einlaufverfahren | SWO-171470 | Defect | HW | Oxi module | Hemolyzer | Clot | Constipation in the device, cannot be flushed | Constipation in the device, cannot be flushed | Hemolyser clogged Hemolyser Unit renewed tHb Kal Entry procedure
No | Visit | Complaint | RCHN | LAC定标失败 | LAC定标失败 | 更换测试卡 | SWO-171699 | [5 empty] | LAC calibration failure | LAC calibration failure | Replace the test card
No | 11 | Y | RAME | 3 PO2 would not pass after sensor cssette was replaced. |  | Checked data logs and verified problem. | D22B913FE64B2DA4 | [5 empty] | 3 PO2 would not pass after sensor cssette was replaced. |  | Checked data logs and verified problem.
Yes | Visit | Complaint | RCHN | GLU,LAC定标失败 | GLU,LAC定标失败 | 更换测试卡后修复 | SWO-177949 | Assisted Replacement | SC | SC | Glu/Lac |  | GLU, LAC calibration failed | GLU, LAC calibration failed | Repair after replacing test card
No | Visit | Complaint | RKK | モニターエラー (Calibration chamber leakage after calibration) | Calibration chamber leakage after calibration センサーシールパーツの歪み | センサーシールパーツ交換 | SWO-171344 | [5 empty] | Monitor error (Calibration chamber leakage after calibration) | Calibration chamber leakage after calibration Distortion of sensor seal parts | Sensor seal parts replacement
No | Visit | Complaint | RCHN | 报1024错误 | 报1024错误，借903-364一个 | 更换903-364修复 by jesson | SWO-171137 | Defect | HW | Process Module | Optical Unit | 1024 | Reported 1024 error | Reported 1024 error, borrow903-364 one | Replace 903-364 Repair by jesson
No | Visit | Complaint | RDEU | unterschiedliche Messergebnisse bei Pat.Messung | Patientenergebnis sehr unterschiedlich bei Wiederholungsmessungen zB 1Messung CO2 60 O2 41 2.Messung pCO2 50,1 O2 49. nicht ein Ergebnis ist bei einer Wiederholung so wie das erste!!! Vergleichsmessungen am ABL5 vom Kunden immer gleich nur ganz kleine Unterschiede | SD an Kopenhagen weitergeleitet | SWO-171756 | [5 empty] | different measurement results for Pat.Measurement | Patient outcome very different for repeat measurements e.g. 1Measurement CO2 60 O2 41 2.Measurement pCO2 50.1 O2 49. Not a result is in a repetition like the first!!! Comparative measurements on the ABL5 from the customer always only very small differences | SD forwarded to Copenhagen
Yes | 3 | Y | RPAC | Blank | PREPARE MACHINE FOR TRARALGON |  | D22BD5B9E62CF611 | Re classified RMED | DEMO Analyzer |  |  |  | Blank | PREPARE MACHINE FOR TRARALGON
No | Visit | Complaint | RDEU | 01.03.2017; Hr. Zwoch; Tel.: 089889220895; Touchscreen ohne Funktion 01.03.2017; Hr. Zwoch; Tel.: 089889220895; auch nach dem Neustart gibt es Probleme mit dem Touchscreen | Touchscreen reagiert in der Mitte sowie unten rechts nicht | Geräteneustart durchgeführt Touchscreen kalibriert Gerät kalibriert und überprüft | SWO-171568 | Defect | HW | DMS_PC | Touch screen |  | 01.03.2017; Mr. Zwoch; Tel.: 089889220895; Touchscreen without function 01.03.2017; Mr. Zwoch; Tel.: 089889220895; even after the restart there are problems with the touchscreen | Touchscreen does not respond in the middle and bottom right | Device restart performed Touchscreen calibrated Device calibrated and checked
No | Visit | Complaint | RDEU | Kein Ausdruck möglich | Kein Ausdruck möglich, Rollenlösehebel gebrochen | Computereinheit demontiert, Druckkopf ausgetauscht, Auto Druckfunktion deaktiviert, Funktionskontrolle ok, Med Technik H. Wahrlich über die Berechnung informiert | SWO-171590 | Defect | HW | Printer Module | Printer head |  | No expression possible | No printing possible, roller release lever broken | Computer unit dismantled, print head replaced, car print function deactivated, function control ok, Med Technique H. Truly informed about the calculation
No | Visit | Complaint | RDEU | Der vor wenigen Tagen eingesetzte Dichtring 834-682 (=834-618) ist UNDICHT | s. case | Der am 26.01.2017 im Rahmen der FAN915-318 eingesetzte Dichtring 834-682 $ 834-618 ist bereits UNDICHT und es tropft aus dem Überlauf der Messkammer; beim Entnehmen der Elektrode verblieb der Dichtring in der Messkammer. Wurde von Frau Klubert neu positioniert. | SWO-171552 | Defect | Electrode | Electrode | Sealing ring | Sealing ring, pH | The sealing ring 834-682 (=834-618) used a few days ago is UNDICHT | s. case | The sealing ring 834-682 x 834-618, which was used on January 26, 2017 as part of the FAN915-318, is already UNDICHT and it drips out of the overflow of the measuring chamber; when the electrode was removed, the sealing ring remained in the measuring chamber. Was repositioned by Ms. Klubert.
No | Visit | Complaint | RIND | printer of analyser was not working. | printer was not taking rolls properly. | required new printer head. | SWO-171520 | Defect | HW | Printer Module | Printer head |  | printer of analyser was not working. | printer was not taking rolls properly. | required new printer head.
Yes | 1 | Y | RPAC | Blank | Replaced pH o-ring |  | D22CD33E5C563D7B | Re classified RMED | FAN |  |  |  | Blank | Replaced pH o-ring
No | Visit | Complaint | RDEU | PH instabil bei Messung | Fehlersuche vor Ort | Bei eintreffen war der Analysator messbereit. Frau Neiswirth hat mir mitgeteilt dass pH bei der Messung sporadisch instabil ist. PH Elektrode gereinigt und unteren pH Dichtungsring getauscht. Mehrere Kalibrierungen durchgeführt. QCs für alle Messungen durchgeführt. ABL wieder messbereit. Mehrere Patientenmessungen durchgeführt. Keine instabile Messung mehr feststellbar. | SWO-171658 | Defect | Electrode | Electrode | Sealing ring | Sealing ring, pH | PH unstable when measured | Troubleshooting on site | When the analyzer arrived, it was ready to measure. Mrs Neiswirth told me that pH is sporadically unstable in the measurement. PH electrode cleaned and lower pH sealing ring replaced. Multiple calibrations performed. QCs for all measurements. ABL ready to be measured again. Several patient measurements have been carried out. No more unstable measurement can be detected.
No | 3 | Y | RAME | 1 Customer reported pH parameter unstable on blood samples. | Changed the complaint visit tab from N to Y. | Checked and cleaned all measuring chamber wells in both modules. Found leakage in Na well. Checked all tubing between measuring chambers. | D22BDFCAE6012BD9 | Assisted Replacement | Electrode | Electrode | pH |  | 1 Customer reported pH parameter unstable on blood samples. | Changed the complaint visit tab from N to Y. | Checked and cleaned all measuring chamber wells in both modules. Found leakage in Na well. Checked all tubing between measuring chambers.
No | Visit | Complaint | RCHN | 打印机无法打印 | 打印机无法打印,借903-001一个 |  | SWO-171692 | [5 empty] | The printer can't print | Printer can't print, borrow 903-001 one
No | Visit | Complaint | RDEU | Keine Anbindung möglich. | Keine lokalen Fehler festzustellen. | Gerät an einen den Standort ZNA AQT90Flex ( 838R0096N010) transportiert und dort angebunden. Es konnte keine Verbindung zum Radiance mit dem ABL80COOX aufbauen. Verbindungstest zwischen Notebook und ABL80COOX mit ASTM.Net durchgeführt. Verbindung war OK. Qualitätskontrollen und Patientenbefunde konnten vom Gerät zum Laptop übertragen werden. Gerät wurde gemäß Kundenangaben wieder konfiguriert (IP: 10.96.11.54, Subnet Mask: 255.255.255.0, GT: 10.96.0.200). Verbindung zum Radiance konnte weiterhin nicht aufgebaut werden. | SWO-172293 | [5 empty] | No connection possible. | No local errors to detect. | device to a location ZNA AQT90Flex ( 838R0096N010) and connected there. It was unable to connect to the Radiance with the ABL80COOX. Connection test between notebook and ABL80COOX with ASTM.Net performed. Connection was OK. Quality controls and patient findings could be transferred from the device to the laptop. Device has been reconfigured according to customer specifications (IP: 10.96.11.54, Subnet Mask: 255.255.255.0, GT: 10.96.0.200). The connection to the radiance could still not be established.
No | Visit | Complaint | RIND | SC80 Installation Failure, leakage found from SC80, Sensor Board Get Defective By Leakage from SC,Placed Order for Connector Board and PER Sent for replacement | SC80 not Initializing. | SC80 Installation Failure, leakage found from SC80, Sensor Board Get Defective By Leakage from SC,Placed Order for Connector Board and PER Sent for replacement | SWO-172371 | [5 empty] | SC80 Installation Failure, leakage found from SC80, Sensor Board Get Defective By Leakage from SC,Placed Order for Connector Board and PER Sent for replacement | SC80 not Initializing. | SC80 Installation Failure, leakage found from SC80, Sensor Board Get Defective By Leakage from SC,Placed Order for Connector Board and PER Sent for replacement
No | 7 | Y | RAME | Z Analyzer dropped and display is not functioning. |  | Analyzer had physical damage to the display. | D22D24635C5EFE1B | [5 empty] | Z Analyzer dropped and display is not functioning. |  | Analyzer had physical damage to the display.
No | 2 | Y | RFRA | Blank | pO2 hors tolérance |  | D22D295F5C0557A1 | [5 empty] | Blank | pO2 out of tolerance
No | Visit | Complaint | RCHN | 血氧参数无值 | 血氧参数无值，借903-197和902-980各一个 | 更换903-197修复 by jesson | SWO-171172 | Defect | HW | Oxi module | Spectrometer | 1034 | Blood oxygen parameters are worthless | Blood oxygen parameters are worthless, borrow 903-197 and 902-980 each one | Replace 903-197 Fix by Jesson
No | Visit | Complaint | RCHN | GLU，LAC定标失败 | GLU，LAC定标失败 | 更换测试卡正常。 | SWO-172027 | Assisted Replacement | SC | SC | Glu/Lac | 1378 | GLU, LAC calibration failed | GLU, LAC calibration failed | The replacement test card is OK.
No | Visit | Complaint | RCHN | 液体传感器失败 | 液体传感器失败 | 更换试剂包后修复 | SWO-177966 | Assisted Replacement | SP | SP | FT | 1268 | Liquid sensor failed | Liquid sensor failed | Repair after replacement of reagent kit
No | Visit | Complaint | RCHN | 传感器损坏 5784，5257 | 传感器损坏 5784，5257 ，借636-550一个 |  | SWO-171700 | Defect | HW | Process Module | Process Unit sensor |  | Sensor damage 5784,5257 | Sensor damage 5784,5257 , borrow 636-550 one
No | Visit | Complaint | RCHN | 光学强度超出范围 | 光学强度超出范围,借903-364一个 | 更换903-364修复 by jesson | SWO-171694 | Defect | HW | Process Module | Optical Unit | 1031 | Optical strength out of range | Optical strength out of range, borrowed 903-364 a | Replace 903-364 Repair by jesson
No | Visit | Complaint | RCHN | 屏幕下方扫描器扫不出，测试卡扫描器故障 | 屏幕下方扫描器扫不出，测试卡扫描器故障，借910-304和902-924各一个 |  | SWO-171781 | Defect | HW | Analyzer Control Module | Barcode reader, display |  | The scanner below the screen can't be scanned, the test card scanner is faulty | The scanner below the screen cannot be scanned, the test card scanner is faulty, and one each is borrowed from 910-304 and 902-924
No | Visit | Complaint | RCHN | 干燥前温度超范围 1068 | 干燥前温度超范围 1068，借903-210,905-908各一个。 |  | SWO-171769 | Defect | HW | Process Module | PCB, Cup dryer | 1068 | Over-range temperature before drying 1068 | The pre-drying temperature exceeds the range of 1068, with a loan of 903-210, 905-908 each.
No | Visit | Complaint | RSCH | Reparatur | Bricht Messung seit Membranenwartung Glu ab Kein vorderes Luftsegment im Einlass-Flüssigkeitsfühler im Zeitrahmen Gerinnsel in pH Messkammer | Einlassdichtung und Nadel überprüft i.o Gerinnsel in pH entfernt i.o O-Ring pH ersetzt i.o Messkammer pH $BG-Modul ausgebaut, ersetzt da Risse im pO2 Zylinder und Messkammer(Nicht die Fehlerursach) i.o Ventile pH $BG-Modul ausgebaut und gereinigt i.o Nachfüllung, LS justage und Pumpen Kal. i.o Drei Testmessungen mit Blut i.o Danach 2. Kal. i.o Testmessung mit FlexQ i.o Erneut 2. Kal. i.o Testmessung mit FlexQ i.o QC Level 3 i.o Danach erneut 2. Kal. Na instabil Membranenwartung Na durchgeführt i.o Kal. und QC i.o Testmessung i.o Messung Kunde i.o | SWO-171970 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | Repair | Cancels measurement since membrane maintenance Glu No front air segment in the inlet liquid sensor in the time frame Clots in pH measuring chamber | Inlet seal and needle checked i.o Clots removed in pH i.o O-ring pH replaces i.o measuring chamber pH $BG module removed, replaces cracks in the pO2 cylinder and measuring chamber (Not the error error) i.o Valves pH $BG module removed and cleaned i.o Refilling, LS adjustment and pumps Kal. Ok Three test measurements with blood i.o After that 2nd Kal. Ok Test measurement with FlexQ i.o Again 2. Ok Test measurement with FlexQ i.o QC Level 3 i.o Then again 2nd Kal. Na unstable Membrane maintenance Na carried out i.o Kal. and QC i.o Test measurement i.o Measurement customer i.o
No | Visit | Complaint | RIND | crea calibration expire. | Check the logs. check the calibration logs. | Replace the Crea reagent & calibrate system. | SWO-172340 | Assisted Replacement | Consumables | Solution bottle | Crea Solution |  | crea calibration expire. | Check the logs. check the calibration logs. | Replace the Crea reagent & calibrate system.
No | Visit | Complaint | RCHN | THB定标不出值 | THB定标不出值，借902-458一个 | 更换902-458修复 by jesson | SWO-171430 | Defect | HW | Oxi module | Lamp unit |  | THB calibration not value | THB bid is not worth, borrow 902-458 a | Replace 902-458 Repair by Jesson
No | Visit | Complaint | RDEU | PO2 Nullstromfehler | PO2 Nullstromfehler | PO2 Elektrode geprüft und ausgetauscht, Funktionskontrolle ok | SWO-172076 | Assisted Replacement | Electrode | Electrode | pO2 |  | PO2 Nullstromfehler | PO2 Nullstromfehler | PO2 electrode tested and replaced, function control ok
No | 5 | Y | RAME | H Inhomogenous sample error |  | Examined inlet module and QC logs. Found bent inlet probe. | D22CFD3D5C28C493 | Defect | HW | Inlet module | Probe bend |  | H Inhomogenous sample error |  | Examined inlet module and QC logs. Found bent inlet probe.
No | Visit | Complaint | RDAN | Claus Boel melder at ABL'en konsekvent har Neon lamp intensity too low . | Afprøvet og undersøgt. | Neon lamp udskiftet, slutafprøvet. | SWO-171963 | Defect | HW | Oxi module | Lamp unit | Neon Lamp | Claus Boel reports that the ABL consistently has Neon lamp intensity too low. | Tested and investigated. | Neon lamp replaced, final tested.
No | 5 | Y | RAME | F Customer reported unit failing oximetry on QC. |  | Found error code 314 (Oxi wavelength deviation high) in QC log. Failed tHB calibration in service with status 23 (Wavelength error). Not failing complete oximetry tests. | D22CE4FA5C0CED7F | Defect | HW | Oxi module | Spectrometer |  | F Customer reported unit failing oximetry on QC. |  | Found error code 314 (Oxi wavelength deviation high) in QC log. Failed tHB calibration in service with status 23 (Wavelength error). Not failing complete oximetry tests.
No | Visit | Complaint | RKK | インレットキャルエラーが頻発し、ホールド画面から回復できず。 | インレットモジュール入れ替え | インレットモジュール入れ替え | SWO-171681 | Defect | HW | Inlet module | Inlet module |  | Inlet care error occurs frequently and cannot be recovered from the hold screen. | Replacing inlet modules | Replacing inlet modules
No | 5 | Y | RAME | Z Customer is unable to pass oxy measurement on level 2 and 3 external QC due to error 581 oxy spectrum mismatch errors. |  | Arrived onsite and could not get oxy mearsurements with consistant resluts even though sample entering hemolyzer. Error 581. | D22D036C5C032815 | Defect | HW | Oxi module | Hemolyzer |  | Z Customer is unable to pass oxy measurement on level 2 and 3 external QC due to error 581 oxy spectrum mismatch errors. |  | Arrived onsite and could not get oxy mearsurements with consistant resluts even though sample entering hemolyzer. Error 581.
No | 4 | Y | RAME | 2 PCO2 sensitivity errors and drift errors on P02. Error 689 gas pressure votlage out of range. |  | Arrived onsite and the PC02 and the P02 are not calibrating. | D22D14725C2D281C | Defect | HW | pH/BG Module | Tubing | grey | 2 PCO2 sensitivity errors and drift errors on P02. Error 689 gas pressure votlage out of range. |  | Arrived onsite and the PC02 and the P02 are not calibrating.
No | 3 | Y | RAME | W QC is running at wrong time |  | Verified time is set correctly in Windows and on the analyzer setup. Also pum calibration error. | D22D1A155C48E02C | Setup | SW | Time setting | Time setting |  | W QC is running at wrong time |  | Verified time is set correctly in Windows and on the analyzer setup. Also pum calibration error.
No | Visit | Complaint | RIND | Satesafe22 Error, After Checking found that Cartridge were Stuck in Storage wheel, Fixed the Problem, Perform Set safe passed Completely, Cleaned analyzer Internally and Externally, Problem Resolve. | Satesafe22 Error, After Checking found that Cartridge were Stuck in Storage wheel. | Fixed the Problem, Perform Set safe passed Completely, Cleaned analyzer Internally and Externally, Problem Resolve. | SWO-172368 | Cleaned | HW | Storage Module | Storage wheel |  | Satesafe22 Error, After Checking found that Cartridge were Stuck in Storage wheel, Fixed the Problem, Perform Set safe passed Completely, Cleaned analyzer Internally and Externally, Problem Resolve. | Satesafe22 Error, After Checking found that Cartridge were Stuck in Storage wheel. | Fixed the Problem, Perform Set safe passed Completely, Cleaned analyzer Internally and Externally, Problem Resolve.
No | Visit | Complaint | RDEU | 20.02.2017; Fr. Jagusch; Tel.: 022612886440; Messabbrüche | Erneut sporadisch aber wiederkehrend Probenabbrüche | Fehlersuche: Dateien überprüft, unter Aktivitäten diverse Meldungen über mechanische Probleme des Storagewheels sowie des Processwheels und weitere ähnliche Fehlermeldungen. Beim Öffnen des Analysators fiel ein Testcup unbekannter Herkunft aus dem Gerät; im Servicemenü safe positions angefahren, OK. Sämtliche in Frage kommende Komponenten sichtkontrolliert und per Serviceprogramm aktiviert, OK. Weder übermäßige Staubbelastung noch mechanische Probleme erkennbar. Im Flows Programm insgesamt 16 Testcups in Folge transportiert, alle einwandfrei absolviert. Montage, Neustart, Endkontrolle: LQC Messung OK | SWO-172002 | Defect | SW | Analyzer Control Module | Off/On |  | 20.02.2017; Ms. Jagusch; Tel.: 022612886440; Measurement interruptions | Again sporadic but recurring rehearsal breaks | Troubleshooting: Checks files, under Activities various messages about mechanical problems of the storage wheel as well as the process wheel and other similar error messages. When the analyzer was opened, a test cup of unknown origin fell from the device; in the safe positions service menu, OK. All eligible components are visually controlled and activated by a service program, OK. No excessive dust exposure or mechanical problems can be detected. Transported in the Flows program a total of 16 test cups in a row, all completed flawlessly. Assembly, restart, final control: LQC measurement OK
No | 4 | Y | RFRA | Blank | ERR IMPRIMANTE |  | D22E5AD2D21183DE | Defect | HW | Printer Module | Printer Module |  | Blank | ERR PRINTER
Yes | Visit | Complaint | RCHN | LAC定标失败 | LAC定标失败 | 更换测试卡正常。 | SWO-172113 | [5 empty] | LAC calibration failure | LAC calibration failure | The replacement test card is OK.
No | Visit | Complaint | RIND | Air Detection calibration Failure, Sc not initialising, Cleaned machine Internally and Externally, Done Zeroing, Reinstalled SC80 and Replaced SP with New one, problem Resolve. | Air Detection calibration Failure, Sc not initialising. | Cleaned machine Internally and Externally, Done Zeroing, Reinstalled SC80 and Replaced SP with New one, problem Resolve. | SWO-172372 | [5 empty] | Air Detection calibration Failure, Sc not initialising, Cleaned machine Internally and Externally, Done Zeroing, Reinstalled SC80 and Replaced SP with New one, problem Resolve. | Air Detection calibration Failure, Sc not initialising. | Cleaned machine Internally and Externally, Done Zeroing, Reinstalled SC80 and Replaced SP with New one, problem Resolve.
No | Visit | Complaint | RCHN | 干燥前温度超范围 1068 | 干燥前温度超范围 1068 ，借903-210一个。 |  | SWO-172051 | Defect | HW | Process Module | PCB, Cup dryer | 1068 | Over-range temperature before drying 1068 | Over-range temperature before drying 1068 , borrow edith of 903-210.
No | Visit | Complaint | RCHN | 定标卡扫描失败 | 定标卡扫描失败，借902-924一个。 |  | SWO-172025 | Defect | HW | Analyzer Control Module | Barcode reader, display |  | Bid card scan failed | The bid card scan failed, borrowing 902-924 one.
No | Visit | Complaint | RKK | 納品時から電解質項目のQCが外れる。 PH、Ca、Naのﾄﾞﾘﾌﾄがでかい。 | 納品時から電解質項目のQCが外れる。 PH、Ca、Naのﾄﾞﾘﾌﾄがでかい。 | Ref電極交換。 PHｼｰﾘﾝｸﾞﾘﾝｸﾞ張り付きあり。 | SWO-172530 | Assisted Replacement | Electrode | Membrane | Reference |  | QC electrolyte item is removed from the time of delivery. Ph, Ca, Na drift is big. | QC electrolyte item is removed from the time of delivery. Ph, Ca, Na drift is big. | Ref electrode replacement. There is with PH sealing ring.
No | Visit | Complaint | RDEU | pCO2 bei QC L2 außer Bereich trotz neuer Membran | Level2 wird häufig im RiliBÄK Bereich verletzt, Membranierung brachte keinen Erfolg | Messkammer und Elektrode gereinigt, erneut membraniert - Membran defekt- wiederholt, mehrere Kal, QC | SWO-172249 | Assisted Replacement | Electrode | Membrane | pCO2 |  | pCO2 at QC L2 out of range despite new membrane | Level2 is often violated in the RiliBÄK area, membraneing did not bring success | Measuring chamber and electrode cleaned, re-membraneed - membrane defective- repeated, several kal, QC
No | Visit | Complaint | RDEU | 02.03.2017 16:20Uhr Herr Bärwolf 03055184118: Der Fehler Oxi-LS-Justagefehler erscheint regelmäßig trotz Wartungsmaßnahme. | Oxi LS Fehler | Sensor entfernt und Anschluss gereinigt, Sensor gereinigt, LS Servicejustage, Pumpenkal, Werte normal, pH instabil, mehrere Kal, QC | SWO-172250 | Defect | HW | Oxi module | LS |  | 02.03.2017 16:20o Mr. Bärwolf 03055184118: The Error Oxi-LS-Adjustment Error appears regularly despite maintenance measures. | Oxi LS Error | Sensor removed and connection cleaned, sensor cleaned, LS service adjustment, pump scale, values normal, pH unstable, multiple kal, QC
No | 9 | Y | RAME | V Touch screen calibration. |  | Upon arrival, several parameters were red, reference membrane needed to be changed, and the waste container was full. | D23B1DAF6E089936 | Adjustment | HW | DMS_PC | Touch screen | Calibration | V Touch screen calibration. |  | Upon arrival, several parameters were red, reference membrane needed to be changed, and the waste container was full.
No | Visit | Complaint | RIND | Lactate Sensitivity Low | Lactate Electrode Sensitivity Low | Replaced Lactate Membrane , Machine Working Ok | SWO-172361 | Assisted Replacement | Electrode | Membrane | Lac | Sensitivity | Lactate Sensitivity Low | Lactate Electrode Sensitivity Low | Replaced Lactate Membrane , Machine Working Ok
Yes | 5 | Y | RAME | L To gather data log files |  | To gather log files for the investigation of QC not cross over correctly to LIS. | D22E719ED20C8447 | Re classified RMED | Collect data |  |  |  | L To gather data log files |  | To gather log files for the investigation of QC not cross over correctly to LIS.
No | 2 | Y | RAME | G Ca++ an K+ are out of calibration and move instrument. |  | Ca++ & K+ are out of status limits. Cal. 1solution, Ca+ electrode, K+ electrode and reference had been changed by customer. All PCB's check good. Replacement reference electrode used by customer was bad too. | D22DEB6FD21AA422 | Assisted Replacement | Electrode | Electrode | Reference |  | G Ca++ an K+ are out of calibration and move instrument. |  | Ca++ & K+ are out of status limits. Cal. 1solution, Ca+ electrode, K+ electrode and reference had been changed by customer. All PCB's check good. Replacement reference electrode used by customer was bad too.
No | 6 | Y | RPNZ | Blank | pCO2 failing again |  | D231CB5F344394C7 | Assisted Replacement | Electrode | Electrode | pCO2 |  | Blank | pCO2 failing again
No | 2 | Y | RAME | Z Touch screen not working. |  | Touch screen will not calibrate. | D22E475BD22EDB10 | Defect | HW | Touch Screen | Touch Screen |  | Z Touch screen not working. |  | Touch screen will not calibrate.
No | Visit | Complaint | RCHN | 血红蛋白没得的值偏差比较大，HCT没得的值偏低 | 血红蛋白测得的值偏差比较大，HCT测得的值偏低,借902-982一个 | 测量仓有血迹，冲洗并做了零点校准后修复 by jesson | SWO-172030 | [5 empty] | The value deviation of hemoglobin is relatively large, and the value not obtained by HCT is low. | The value deviation measured by hemoglobin is relatively large, the value measured by HCT is low, and 902-982 is a | The measuring compartment has blood, rinsed and made a zero-point calibration to repair by jesson
No | Visit | Complaint | RIND | Leakage In Inlet | Clot In Inlet Because Of That Leakage | Cleaned The Inlet By Using Hypo chloride , Machine Working Ok | SWO-172360 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | Leakage In Inlet | Clot In Inlet Because Of That Leakage | Cleaned The Inlet By Using Hypo chloride , Machine Working Ok
No | Visit | Complaint | RDEU | Gerät meldet Probleme mit der Luftblasen erkennung | Neue Initialisierung der SC erforderlich. | Kunde hatte kein LP vor Ort und keine Sensor Kassette. LP und Sensor Kassette in der GS geholt. Kontakte an der SC gereinigt, Fliesswege durchgespült. SC initialisiert, Kal und QC sind OK. Neue SC und LP beim Kunden gelassen. | SWO-172410 | [5 empty] | Device reports problems with air bubble detection | New initialization of sc required. | Customer had no LP on site and no sensor cassette. LP and sensor cassette in the GS. Contacts cleaned at the SC, flow paths rinsed through. SC initialized, Kal and QC are OK. New SC and LP left at the customer'
Yes | Visit | Complaint | RCHN | LAC定标不过 | LAC定标不过 | 更换测试卡后修复 | SWO-172975 | [5 empty] | LAC calibration is not | LAC calibration is not | Repair after replacing test card
No | Visit | Complaint | RKOR | Sensor cassette clot error | . | Sensor cassette clot clean. Liquid flow line clean. Calibration failed. Sensor cassette replace. Calibration completed. | SWO-172479 | [5 empty] | Sensor cassette clot error | . | Sensor cassette clot clean. Liquid flow line clean. Calibration failed. Sensor cassette replace. Calibration completed.
No | Visit | Complaint | RDEU | Kal. Error->keine Spüllösung, Flüssigkeit während Gaskal. erkannt, Pumpenkal. Error | Wochenendeinsatz laut Frau Dr. Ganzlmeier dringend erforderlich Ventilmembran (Airventil im pH/BG Modul) erneuert | Fehlersuche, alle vom Anwender durchgeführten, möglichen Maßnahmen brachten keinen Erfolg, Kal.- $und Aktivitätendateien überprüft, auf Serviceprogrammebene Fluidicmodulventile angesteuert ok, bei Messmodulen vor allem bei der Pumpenkal. Luftblasenbildung im Ventilblock bei pH $BG Modul, Ventilmembran (Airventil im pH $BG Modul) defekt->erneuert, Neustart, Pumpenkal. und LS Justage ok, Kalibration, QC und Testmessung ok | SWO-172411 | Defect | HW | pH/BG Module | Valve, air |  | Kal. Error->no rinsing solution, liquid during gas scale. detected, pump scale. Error | Weekend assignment urgently needed according to Dr. Ganzlmeier Valve membrane (air valve in pH/BG module) renewed | Troubleshooting, all possible measures carried out by the user did not succeed, Kal.- $und activity files checked, controlled at the service program level fluidic module valves ok, with measuring modules especially with the pump scale. Air bubble formation in the valve block at pH $BG module, valve membrane (air valve in pH $BG module) defective->renewed, restart, pump scale. and LS Adjustment ok, calibration, QC and test measurement ok
No | Visit | Complaint | RDEU | 06.03.2017 10:00Uhr Frau Eurich 062181024322: Das Gerät meldet trotz Wartungsmaßnahmen weiterhin LP erneuern. | Lösungen können nicht gepumpt werden | Allgemeine Überprüfung - Analysator gespült - Erneuerung Einlassdichtung - Serviceprogramme durchgeführt - Funktionsüberprüfung etc. ohne Beanstandung | SWO-172569 | Defect | HW | Inlet module | Inlet Gasket |  | 06.03.2017 10:00am Ms Eurich 062181024322: The device continues to renew LP despite maintenance measures. | Solutions cannot be pumped | General inspection - Analyzer rinsed - Renewal inlet seal - Service programs carried out - Functional check etc. without complaint
No | Visit | Complaint | RDEU | Kal-Fehler in pH und Elyten nach Kassettentausch. | Neue Kassette defekt. | Referenzlösung in der Kassette war ausgelaufen 946-009 Lot R0579. Neue Kassette aus Schwalmstadt geholt und eingesetzt. Gerät überprüft. Vier Lösungspacks abgezogen. | SWO-174116 | Assisted Replacement | SC | SC | Reference | Leaking | Kal errors in pH and elytes after cassette replacement. | New cassette defective. | Reference solution in the cassette had expired 946-009 Lot R0579. New cassette from Schwalmstadt fetched and installed. device checked. Four solution packs subtracted.
No | Visit | Complaint | RDEU | Drucküberprüfung Vakuumfehler | Gerät überprüft, | Lösungspack richtig aktiviert, Einlassdichtung erneuert, Gerät neu gestartet, Kaltstartphase abgewartet, tHb Kal, Kal und QCs durchgeführt | SWO-172905 | User error | SP | Activation | Activation |  | Drucküberprüfung Vakuumfehler | Device checked, | Solution pack correctly activated, inlet seal renewed, device restarted, cold start phase awaited, tHb Kal, Kal and QCs performed
No | Visit | Complaint | RDEU | Externer Scanner ohne Funktion. Kunde wollte Scanner anschließen, funktioniert aber nicht. | Externer Scanner ohne Funktion. Nacharbeit zu Auslieferung letzte Woche, da Einstellungen nicht übernommen wurden. | Einstellungen in Roma.ini nochmals eingegeben. Nach Neustart Scanner nochmals ab und angesteckt Scanner funktioniert jetzt. Kalibration und QC's vorgenommen. | SWO-172609 | Setup | SW | Roma.ini | Ext. Barcode reader |  | External scanner without function. Customer wanted to connect scanner, but did not work. | External scanner without function. Rework on delivery last week because settings have not been applied. | settings in Roma.ini. After rebooting scanner again and plugged in Scanner now works. calibration and QC's.
No | Visit | Complaint | RDEU | Vakuum Fehler | Einlassnadel verbogen | Gerät geprüft, Einlassnadel (aus Kundenbestand) ersetzt, Funktionskontrolle ok | SWO-172862 | Replaced | HW | Auto Inlet module | Probe bend |  | Vacuum error | Einlassnadel verbogen | Device checked, inlet pin (from customer stock) replaced, function control ok
Yes | Visit | Complaint | RBNL | Bij tijd en wijle zijn er problemen dat de mixer niet start als er een spuit wordt ingelegd. De verlichting gaat daarbij ook flikkeren. | Mixen gaat nu goed. Overname van batterijstatus naar netvoeding ook goed. | Preventief mixerunit vervangen. | SWO-172619 | Re classified RMED | HW | Basic Unit | Sample Mixer | Agreed WR 2017-03-30 | At times, there are problems that the mixer does not start when a syringe is inlaid. The lighting is also flickering. | Mixes now goes well. Take charge of battery status to power supply also good. | Replace the preventive mixer unit.
Yes | Visit | Complaint | RBNL | Bij tijd en wijle zijn er problemen dat de mixer niet start als er een spuit wordt ingelegd. De verlichting gaat daarbij ook flikkeren. | Mixen gaat nu goed. Correcte overname van baterrijstatus naar netvoeding. | Preventief mixer vervangen. | SWO-172628 | Re classified RMED | HW | Basic Unit | Sample Mixer | Agreed WR 2017-03-30 | At times, there are problems that the mixer does not start when a syringe is inlaid. The lighting is also flickering. | Mixes now goes well. Correct acquisition of baterdrive status to power supply. | Replace the preventive mixer.
No | Visit | Complaint | RBNL | Bij tijd en wijle zijn er problemen dat de mixer niet start als er een spuit wordt ingelegd. De verlichting gaat daarbij ook flikkeren. | Nu geen mix problemen. Ook correcte overname van batterijstatus naar netstatus. | Preventief mixer vervangen. | SWO-172632 | Defect | HW | Basic Unit | Sample Mixer | Preventive | At times, there are problems that the mixer does not start when a syringe is inlaid. The lighting is also flickering. | Now no mix problems. Also correct battery status acquisition to Netstatus. | Replace the preventive mixer.
Yes | Visit | Complaint | RBNL | Bij tijd en wijle zijn er problemen dat de mixer niet start als er een spuit wordt ingelegd. De verlichting gaat daarbij ook flikkeren. | Correcte mixing en overname van batterijstatus naar netspanning. | Preventief mixer vervangen. | SWO-172642 | Re classified RMED | HW | Basic Unit | Sample Mixer | Agreed WR 2017-03-30 | At times, there are problems that the mixer does not start when a syringe is inlaid. The lighting is also flickering. | Correct mixing and take-over of battery status to mains voltage. | Replace the preventive mixer.
Yes | Visit | Complaint | RBNL | Bij tijd en wijle zijn er problemen dat de mixer niet start als er een spuit wordt ingelegd. De verlichting gaat daarbij ook flikkeren. | Nu geen problemen, ook niet bij overname van batterijstatus naar netvoeding. | Preventief mixer vervangen. | SWO-172776 | Re classified RMED | HW | Basic Unit | Sample Mixer | Agreed WR 2017-03-30 | At times, there are problems that the mixer does not start when a syringe is inlaid. The lighting is also flickering. | No problems, even when battery status is transferred to mains power. | Replace the preventive mixer.
No | Visit | Complaint | RSCH | sO2, HbF, HHb im QC ausser Bereich | sO2, HbF, HHb im QC ausser Bereich | Allgemeine Kontrolle ok. Lösungspack wurde ohne Erfolg ersetzt ok. Hemolyzer ers ok. Schläuche von JW Kit ersetzt ok. Thb Kal. I.o. System i.o Kal. + QC i.o. | SWO-172597 | Defect | HW | Basic Unit | PM kit |  | sO2, HbF, HHb in QC out of range | sO2, HbF, HHb in QC out of range | General control ok. Solution pack has been replaced ok without success. Hemolyzer ers ok. Hoses replaced by JW Kit ok. Thb Kal. Ok. System i.o Kal. + QC i.o.
Yes | Visit | Complaint | RCHN | LAC定标不过 | LAC定标不过 | 更换测试卡后修复 | SWO-172974 | [5 empty] | LAC calibration is not | LAC calibration is not | Repair after replacing test card
Yes | Visit | Complaint | RCHN | LAC定标不过 | LAC定标不过 | 更换测试卡后修复 | SWO-172930 | [5 empty] | LAC calibration is not | LAC calibration is not | Repair after replacing test card
No | 2 | Y | RPAC | Blank | Analyser cleaned |  | D231FAFD3401DCEB | [5 empty] | Blank | Analyser cleaned
No | Visit | Complaint | RIND | Two Point Calibration Failure, SC not Initialising, tried to reinstall SC, Replaced SC with New one, Problem Resolve. | SC80 Initialisation Failure. | Two Point Calibration Failure, SC not Initialising, tried to reinstall SC, Replaced SC with New one, Problem Resolve. | SWO-175135 | [5 empty] | Two Point Calibration Failure, SC not Initialising, tried to reinstall SC, Replaced SC with New one, Problem Resolve. | SC80 Initialisation Failure. | Two Point Calibration Failure, SC not Initialising, tried to reinstall SC, Replaced SC with New one, Problem Resolve.
Yes | 2 | Y | RPAC | Blank | Set up patient layout for Jason. |  | D231FC303448DCEE | Re classified RMED | Customer training |  |  |  | Blank | Set up patient layout for Jason.
No | 6 | Y | RAME | Z external QC quick keys will not appear on the main screen |  | verified all quick keys load to the main screen except external QC | D23215DC345CA3AE | [5 empty] | Z external QC quick keys will not appear on the main screen |  | verified all quick keys load to the main screen except external QC
No | Visit | Complaint | RCHN | 频繁更换测试卡 | 频繁更换测试卡，借备件903-135一个 | 由于定标不通过导致的更换测试卡频率较高，怀疑机器存在问题，借用903-135检查机器的放大电路和电极接触是否存在问题。目前未发现问题 | SWO-172355 | Assisted Replacement | SC | SC | Unknown |  | Frequent change of test cards | Frequent replacement of test cards, spare parts 903-135 one | Due to the high frequency of replacement test cards due to non-passing the calibration, a problem is suspected of the machine, and 903-135 is used to check the machine's amplification circuit and electrode contact for problems. No problems have been identified at this time
Yes | Visit | Complaint | RCHN | GLU定标不过 | GLU定标不过 | 更换测试卡后修复 | SWO-172937 | [5 empty] | GLU calibration is not | GLU calibration is not | Repair after replacing test card
No | Visit | Complaint | RIND | Oxy-metery parameter error. | check the logs. check the tHb calibration. | uninterrupted shutdown causing A $D converter fail. shutdown system for 15-20 min & reset the Oxi module. calibrate system calibration pass. guide for take precaution for power supply. | SWO-175140 | Power off/on | HW | Oxi module | Off/On |  | Oxy-metery parameter error. | check the logs. check the tHb calibration. | uninterrupted shutdown causing A $D converter fail. shutdown system for 15-20 min & reset the Oxi module. calibrate system calibration pass. guide for take precaution for power supply.
No | Visit | Complaint | RIND | Lactate frequently fails with it's sensitivity | Check calibration status and calibration data log and found Lac is failing frequently. | Clean all the chambers,check Lac membrane status and run cleaning then 2-point calibration which is OK.Take data log for further inspection to resolve the Lac issue. | SWO-172914 | Cleaned | HW | El/Met Module | Fluidic pathway |  | Lactate frequently fails with it's sensitivity | Check calibration status and calibration data log and found Lac is failing frequently. | Clean all the chambers,check Lac membrane status and run cleaning then 2-point calibration which is OK. Take data log for further inspection to resolve the Lac issue.
No | Visit | Complaint | RDEU | pCO2 und K Empfindlichkeitsfehler | Undichtigkeiten in der Messkammer | Elektrode demontiert und geprüft, Messkammern gereinigt, pO2 $pCO2 $Ca. $K $Na. Elektroden neu membraniert, Funktionskontrolle ok | SWO-172976 | Assisted Replacement | Electrode | Membrane | pCO2 |  | pCO2 and K Sensitivity errors | Leaks in the measuring chamber | Electrode dismantled and tested, measuring chambers cleaned, pO2 $pCO2 $Ca. $K $Na. Electrodes re-membraneed, function control ok
No | 5 | Y | RFRA | Blank | CV important sur les CQ |  | D2321259340F3E2D | Cleaned | HW | pH/BG Module | Fluidic pathway |  | Blank | Important CV on CQ
No | Visit | Complaint | RIND | Leaking from inlet gasket | Leaking from inlet gasket found problem in waste connector was block | Cleaned the waste block & checked equipment is functional | SWO-175107 | Cleaned | HW | Inlet module | Fluidic pathway |  | Leaking from inlet gasket | Leaking from inlet gasket found problem in waste connector was block | Cleaned the waste block & checked equipment is functional
Yes | Visit | Complaint | RDEU | Das Gerät sollte von 825 mit Flex Q auf ABL800 ohne Flex Q abgerüstet werden | Option Key installiert | Option Key installiert | SWO-172904 | Re classified RMED | Upgrade option key |  |  | RDEU GEMBA | The device should be upgraded from 825 with Flex Q to ABL800 without Flex Q | Option Key installed | Option Key installed
Yes | Visit | Complaint | RDEU | Cl aufrüsten | Cl. Elektrode eingebaut | Cl. - Elektrode eingebaut und Optionkey installiert. | SWO-172902 | Re classified RMED | Upgrade option key |  |  | RDEU GEMBA | Cl upgrade | Cl. Electrode installed | Cl. - Electrode installed and optionkey installed.
No | 32 | Y | RPOL | Blank | Czyszczenie bearing wheel, ususniêcie z pok³adu uszkodzonej kasetki D-dimer Analizator sprawny technicznie. |  | D2320B8D3400D13B | Defect | Consumables | Test Kit | Cartridge, D-Dimer | Broken | Blank | Cleaning the bearing wheel, removing the damaged D-dimer cassette Analyzer technically efficient.
No | Visit | Complaint | RIND | Sample aborting again and again. | Sample aborting. | Gone through service program and found inlet wheel and Storage wheel failed. Storage wheel blocking by a broken holder. Re-fix Storage wheel after service and got fixed. Analyser serviced and calibrated properly. Sample done and found values well. Analyser Handed over in afine working condition. | SWO-174399 | Defect | HW | Storage Module | Storage wheel |  | Sample aborting again and again. | Sample aborting. | Gone through service program and found inlet wheel and Storage wheel failed. Storage wheel blocking by a broken holder. Re-fix Storage wheel after service and got fixed. Analyser serviced and calibrated properly. Sample done and found values well. Analyser Handed over in afine working condition.
No | 4 | Y | RPOL | Blank | Wymiana ig³y oraz uszczelki wlotu. Analizator sprawny technicznie. |  | D2335A20AA3F658D | Replaced | HW | Inlet module | Probe |  | Blank | Replacing the needle and inlet seals. Technically efficient analyzer.
No | Visit | Complaint | RDEU | Frau Weimar Tel: 09117580992753 | Inhomogene Kal. 3 Lösung | ABL war bei Eintreffen wieder messbereit, ABL im Service überprüft, System- und Funktionskontrolle in Ordnung | SWO-173046 | No fault found | No fault found | No fault found | FT |  | Mrs Weimar Tel: 09117580992753 | Inhomogeneous Kal. 3 Solution | ABL was ready to be measured again on arrival, ABL checked in service, system and function control in order
No | 12 | Y | RAME | R. 1178 pum calibration error, 1324 inhomogeneous rins solution, 1111 inhomogeneous air, 1062, pressure test pressure error, 1086,1085,1087,1084,1083 ihomogeneous calibations solutions. |  | Found sensor cassette fully clogged. Unable to get solution through it. | D2334227AA21BD11 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | R. 1178 pum calibration error, 1324 inhomogeneous rins solution, 1111 inhomogeneous air, 1062, pressure test pressure error, 1086,1085,1087,1084,1083 ihomogeneous calibations solutions. |  | Found sensor cassette fully clogged. Unable to get solution through it.
No | 3 | Y | RAME | D Customer says COHGB does not correlate between analyzers |  | Had customer run a tHb calibration. Results were good. System cycle and QC's are all within range. Blood results continue to run significantly higher than ABL 80 OSM right next to this analyzer. | D2337FC2AA02CCED | [5 empty] | D Customer says COHGB does not correlate between analyzers |  | Had customer run a tHb calibration. Results were good. System cycle and QC's are all within range. Blood results continue to run significantly higher than ABL 80 OSM right next to this analyzer.
No | Visit | Complaint | RCHN | 定标错误 1320，1318，1319 | 定标错误 1320，1318，1319 | 更换测试卡 | SWO-174226 | Assisted Replacement | SC | SC | Multi |  | Calibration error 1320,1318,1319 | Calibration error 1320,1318,1319 | Replace the test card
No | Visit | Complaint | RCHN | CKMB和MYO定标失败 | CKMB和MYO定标失败，借944-216（12278）和944-214（12231）各一个 |  | SWO-172981 | Defect | HW | Incubator Module | Punch up | 5297 | CKMB and MYO bid failed | CKMB and MYO failed to bid, borrowing one each of 944-216 (12278) and 944-214 (12231)
No | 22 | Y | RESP | Fallo de calibración en el potasio de forma diaria. |  | kit de tubos. Susituida cámara de medida del pH. limpieza cámara del medida, circuitos neumáticos, hemolizador, filtro ventilador, limpieza exterior. Comprobación y ajuste de las bombas, válvulas, entrada, sensores de líquidos, oximetro, autocheck, impresora, electrodos, gas, temperatura. Descontaminación. Cal 2 ptos | D2331D94AA386EB4 | Cleaned | HW | El/Met Module | Measuring chamber |  | Calibration failure in potassium on a daily basis. |  | tube kit. Susituted pH measuring chamber. Cleaning measurement camera, pneumatic circuits, hemolizer, fan filter, exterior cleaning. Checking and adjusting pumps, valves, inlet, liquid sensors, oximeter, autocheck, printer, electrodes, gas, temperature. Decontamination. Cal 2 ptos
Yes | 21 | Y | RAME | D COHb is not comparing on patient samples |  | COHb is running lower on this analyzer on blood samples; it is also running low on QC 5.7 peer 2 is 6.0, Travel from San Rafael | D2336232AA0B02A5 | [5 empty] | D COHb is not comparing on patient samples |  | COHb is running lower on this analyzer on blood samples; it is also running low on QC 5.7 peer 2 is 6.0, Travel from San Rafael
No | Visit | Complaint | RDEU | 07.03.2017 11:30Uhr Frau Domberg 05115322529: Der Touchsreen reagiert nicht oder nur schlecht auf Tastendruck. | Bildschirm reagiert schlecht , PH instabil | Bildschirm geprüft, Leihbildschirm installiert, BG Messkammer gereinigt, Ref. Elektrode neu membraniert, Dichtring an der PH Elektrode ersetzt,Funktionskontrolle ok | SWO-174686 | Defect | HW | DMS_PC | Touch screen |  | 07.03.2017 11:30am Ms. Domberg 05115322529: The touchsreen does not react or only badly at the touch of a button. | Screen reacts badly, PH unstable | Screen checked, rental screen installed, BG measuring chamber cleaned, ref. Electrode re-membraneed, sealing ring replaced at THE PH electrode,function control ok
No | 7 | Y | RAME | Z Display Screen intermittently unresponsive. |  | issue while entering patient data. | D23888C3824B8056 | Defect | HW | DMS_PC | Touch screen |  | Z Display Screen intermittently unresponsive. |  | issue while entering patient data.
No | 11 | Y | RAME | 2 pCO2 issues |  | pCO2 low on level 1 quality control. pCO2 also drifts on calibration intermittently. David and I have tried new pCO2 electrodes also. This is an ABL820 at software version 6.16. | D2331C26AA0B1E81 | Assisted Replacement | Electrode | Membrane | pCO2 |  | 2 pCO2 issues |  | pCO2 low on level 1 quality control. pCO2 also drifts on calibration intermittently. David and I have tried new pCO2 electrodes also. This is an ABL820 at software version 6.16.
No | Visit | Complaint | RIND | Sample aborting again and again. | Analyser is on observation. | Observation done. and found analyser working well. Handed over in a fine working condition. | SWO-174401 | No fault found | Courtesy | Courtesy | Courtesy |  | Sample aborting again and again. | Analyser is on observation. | Observation done. and found analyser working well. Handed over in a fine working condition.
No | Visit | Complaint | RDEU | QK-Röhrchen werden beim einsetzen in AQT nicht erkannt Fehlende Daten QK in CobasIT1000 | Reinigungsarbeiten, Fehlersuche | Fehlende QK-Daten in der Konfiguration angelegt, Schnittstellenkonfiguration geprüft, es ist davon auszugehen, dass ein Anwenderfehler vorliegt -> da fehlende QK mglw. als Patient gemessen wurden Zusätzlich sind sehr viele Verbindungsabbrüche und entsprechende Wiederherstellungen der Verbindung vom AQT zu CobasIT dokumentiert! Hinweis an Kunden über Anlage von QK !! | SWO-174188 | Setup | His/lis | His/lis | His/lis |  | QK tubes are not detected when used in AQT Missing data QK in CobasIT1000 | Cleaning work, troubleshooting | Missing QK data created in the configuration, interface configuration checked, it is to be assumed that there is a user error -> since missing QK mglw. measured as a patient In addition, a lot of disconnections and corresponding restores of the connection from AQT to CobasIT are documented! Note to customers about qK plant !!
No | Visit | Complaint | RDEU | 07.03.2017; Hr. Klein (Med. Technik); Tel.: 089889220962; Beschwerde per Fax - ABL nicht messbereit, er kann keine weiteren Angaben machen. 07.03.2017; Hr. Spangenberg (Geräteverantwortliche); Tel.: 089889220410; ABL ist wegen Temperaturfehler ausgeschaltet. Herr Spangenberg hat den ABL wieder eingeschaltet, Rückruf in 1 - 1,5 Std., dann auch Klärung wegen Hotline. 07.03.2017; 13:05 Uhr; Hr. Spangenberg; Tel.: 089889220410; Der ABL ist seit 1,5 Stunden in der Startphase von 55Min. | Fehler: Temperatur Subtrat der Sensorkassette außer Bereich. Rollenkranz rastet nicht korrekt auf Achse ein. | Abfallschlauch und Ventilschlauch überprüft - ok. Rollenkranz erneuert (aus Neugeräte ausgebaut, da nicht im Bestand). SC und LP aus Kundenbestand erneuert. Kalibrationen und QC's vorgenommen. | SWO-174422 | Unknown | Unknown | Unknown | Unknown |  | 07.03.2017; Mr. Klein (Med. Technik); Tel.: 089889220962; Complaint by fax - ABL is not ready for measurement, he cannot provide any further information. 07.03.2017; Mr. Spangenberg (device manager); Tel.: 089889220410; ABL is switched off due to temperature errors. Mr. Spangenberg has switched on the ABL again, recall in 1 - 1.5 hours, then also clarification due to hotline. 07.03.2017; 1:05 p.m.; Mr. Spangenberg; Tel.: 089889220410; The ABL has been in the 55min start phase for 1.5 hours. | Error: Temperature substrate of the sensor cartridge out of range. Roller wreath does not snap correctly to the axis. | Waste hose and valve hose checked - ok. Roller wreath renewed (removed from new equipment, as not in stock). SC and LP from customer base. calibrations and QC's.
No | 4 | Y | RPOL | Blank | Wymiana wygiêtej ig³y aspiracyjnej. Testy kontrolne. Analizator sprawny. |  | D234AE952027F2E9 | Replaced | HW | Inlet module | Probe |  | Blank | Replacement of bent aspiration. Control tests. Functional Analyzer.
Yes | 1 | Y | RCZE | Blank | nenabíhá, hlásí chybu - systemtray MFC Application has encountered a problem and needs to close- Po dohodì odvoz do office. Pøístroj z dìtského pracovištì onko |  | D234D0F620410E28 | Re classified RMED | DEMO Analyzer |  |  |  | Blank | Error-Systemtray MFC Application has encountered a problem and needs to close-after an agreement to transfer to office. The device from the
No | Visit | Complaint | RCHN | 接口板坏了，所有参数定标失败 | 接口板坏了，所有参数定标失败，借902-972一个。 |  | SWO-174218 | [5 empty] | The interface board is broken and all parameters fail to be calibrated | The interface board is broken, all parameters have failed to mark, borrowing 902-972 one.
No | Visit | Complaint | RIND | Temp error. | Temp error, | Temp error, during SC installation, found SC was not connected to analyzer, removed the SC reinstalled the SC found Air detection calibration failed, inactivated the Air detection, recalibrated SC found all the parameter was working, found instrument is functional. | SWO-181350 | [5 empty] | Temp error. | Temp error, | Temp error, during SC installation, found SC was not connected to analyzer, removed the SC reinstalled the SC found Air detection calibration failed, inactivated the Air detection, recalibrated SC found all the parameter was working, found instrument is functional.
No | Visit | Complaint | RCHN | 系统开机无法启动 | 系统开机无法启动,借910-342一个 |  | SWO-174302 | Defect | HW | DMS_PC | DMS_PC |  | System boot cannot be started | System boot can not start, borrow 910-342 one
No | Visit | Complaint | RCHN | LAC响应错误 | LAC响应错误 | 更换测试卡 | SWO-174208 | Assisted Replacement | SC | SC | Lac | Response | LAC response error | LAC response error | Replace the test card
Yes | Visit | Complaint | RCHN | PCO2定标错误，PCO2 LAC质控错误 | PCO2定标错误，PCO2 LAC质控错误 | 连续更换一个测试卡和一个试剂包修复 | SWO-176054 | [5 empty] | PCO2 calibration error, PCO2 LAC quality control error | PCO2 calibration error, PCO2 LAC quality control error | Replace one test card and one reagent kit repair in a row
No | Visit | Complaint | RCHN | 参比电极错误 | 参比电极错误 | 更换测试卡后修复 | SWO-177952 | Assisted Replacement | SC | SC | Reference | 1361 | Reference electrode error | Reference electrode error | Repair after replacing test card
No | Visit | Complaint | RCHN | 测试卡扫描失败 | 测试卡扫描失败,借910-304一个。 |  | SWO-174238 | Defect | HW | Storage Module | Barcode reader, Storage wheel |  | Test card scan failed | The test card scan failed, borrowing 910-304 one.
No | Visit | Complaint | RCHN | MYO定标失败 | MYO定标失败,借944-214（12338）一个 |  | SWO-174230 | Assisted Replacement | Consumables | Test Kit | Cartridge | Myo | MYO calibration failure | MYO bid failure, borrow 944-214 (12338) one
No | Visit | Complaint | RCHN | P3泵不转 3412 | P3泵不转 3412，借903-112，903-198各一个。 |  | SWO-174268 | Defect | HW | Analyzer Control Module | PCB, Main Board | 3412 | P3 pump does not turn 3412 | P3 pumps do not turn 3412, borrow 903-112, 903-198 each.
Yes | 11 | Y | RFRA | Blank | Fuite site d'introduction + derive Ca2+ |  | D2348A2F2019D135 | [5 empty] | Blank | Leaked introductory site ' derive Ca2'
No | Visit | Complaint | RDEU | Bildschirm friert häufig ein. | Bildschirm friert häufig ein. | Bildschirm aus Altgerät eingesetzt. Kalibriert und Töne konfiguriert. | SWO-174697 | Defect | HW | DMS_PC | Touch screen |  | Screen often freezes. | Screen often freezes. | screen made of old equipment. Calibrated and configured totones.
No | Visit | Complaint | RDEU | Kunde musste Lac-Membran nach 14 Tagen wieder ersetzen Info durch Ralf Ehlert | Lac muss 14tägig wegen QC-Fehler membraniert werden D7077, Lot 3361 | Servicedump gezogen, Lac gereinigt und neu membraniert, akz. Kal abgewartet, Kal und QC i.O. | SWO-174544 | Assisted Replacement | Electrode | Membrane | Lac | RDEU GEMBA | Customer had to replace Lac membrane after 14 days Info by Ralf Ehlert | Lac must be membraned 14 days due to QC error D7077, Lot 3361 | Service dump pulled, Lac cleaned and re-membraneed, akz. Kal waited, Kal and QC i.O.
No | Visit | Complaint | RDEU | 08.03.2017 Frau Klubert 02214785283 meldet, dass ein OXI-Fehler vorliegt. | Service Menü, Dateien gesichtet, sporadisch neon number too low und oxy spectrum mismatch | Oxy Modul ersetzt, oxy complete test, Kalibrationen, Validierung der tHb Kalibration, QC Kontrolle | SWO-174516 | Defect | HW | Oxi module | Oxi module, Replaced | Neon/Spectrometer | 08.03.2017 Ms Klubert 02214785283 reports that there is an OXI error. | Service menu, files spotted, sporadic neon number too low and oxy spectrum mismatch | Oxy module replaced, oxy complete test, calibrations, validation of tHb calibration, QC control
No | 7 | Y | RAME | F Oxi hardware error, Neon lamp voltage too high. |  | Oxi hardware error, Neon lamp voltage too high. | D2347CD32039DB26 | Defect | HW | Oxi module | Lamp unit | Neon | F Oxi hardware error, Neon lamp voltage too high. |  | Oxi hardware error, Neon lamp voltage too high.
No | Visit | Complaint | RDEU | Lac muss alle 14 Tage neu mebraniert werden, da QC- bzw. Kal-Fehler | Lac instabil in Kal u. QC nach Membranierung Glu | Servicdump gezogen, mehrfach kalibriert, Lac neu membraniert, akz. Kal abgewartet, Kal u. QC i.O. | SWO-174545 | Assisted Replacement | Electrode | Membrane | Lac | RDEU GEMBA | Lac muss alle 14 Tage neu mebraniert werden, da QC- bzw. Kal-Fehler | Lac unstable in Kal & QC after membraneation Glu | Servicdump pulled, multiple calibrated, Lac re-membraneed, akz. Kal waited, Kal & QC i.O.
No | 2 | Y | RFRA | Blank | Consommation importante bouteille gaz 2 |  | D2345D65203C8EE0 | Defect | HW | Gas module | Gas2 |  | Blank | Significant gas cylinder consumption 2
No | 4 | Y | RPOL | Blank | Aparat zg³asza³ problem FTC 106, zrobiono procedurê odbia³czania, wymiana pompy P5, testy przep³ywowe, analizator sprawny. |  | D23475BE20521762 | Defect | HW | Aspiration Module | P5 | FTC 106 | Blank | Camera reported problem FTC 106, a procedure for reprocessing, P5 pump replacement, overrun tests, analyser functional.
No | Visit | Complaint | RDEU | 08.03.2017, Frau Scheipl, Tel.: 01711189713: Probenhalterung (Rädchen) im Einlass abgebrochen, Probe liegt im AQT. | Tausch der Probenhalterung | Probenhalterung getauscht weil gebrochen. Systemüberprüfung durchgeführt. Mehrere Systemreinigungen durchgeführt. QC Level 2 Multicheck durchgeführt --> i.O. AQT wieder messbereit. | SWO-174355 | Defect | HW | Aspiration Module | Tube holder |  | 08.03.2017, Mrs. Scheipl, Tel.: 01711189713: Sample holder (wheel) aborted in the inlet, sample is in the AQT. | Replacement of the sample holder | Sample holder replaced because broken. system verification. Several system cleanings performed. QC Level 2 Multicheck performed --> i.O. AQT ready to measure again.
No | 3 | Y | RPOL | Blank | Testy analizatora - ok. Kalibracja CRP - 12521. Analizator sprawny technicznie. |  | D2347F3D205D53E4 | Assisted Replacement | Consumables | Test Kit | CRP |  | Blank | Analyzer tests-OK. Calibrating CRP-12521. Technically efficient analyzer.
No | 6 | Y | RAME | F user called stating that oxi parameters COHB SO2 and METhb |  | found that they are failing internal and external qc for OXI parameters | D235CCFF96025412 | Adjustment | HW | Oxi module | tHb Cal |  | F user called stating that oxi parameters COHB SO2 and METhb |  | found that they are failing internal and external qc for OXI parameters
Yes | 7 | Y | RAME | G SC Failing Multiple Parameters. |  | Sensors cassettes were failing multiple parameters on calibration. Replacements were ordered through tech support. | D23768150C384EFB | Re classified RMED | PER action |  |  |  | G SC Failing Multiple Parameters. |  | Sensors cassettes were failing multiple parameters on calibration. Replacements were ordered through tech support.
No | Visit | Complaint | RDAN | ADG melder, at ABL 90'eren ikke sætter netikonet på efter, at have kørt på batteri. | Afprøvet og undersøgt. Fejlen bekræftet. | Strømforsyning udskiftet, slutafprøvet. | SWO-174673 | Defect | HW | Power Supply | Power Supply |  | ADG reports that the ABL 90 's do not put the Neti icon on after having run on battery. | Tested and investigated. Error confirmed. | Power supply replaced, final tested.
No | Visit | Complaint | RIND | Solution & Blood leak. | Solution & Blood leak from waste. | Problem with waste drain, cleaned the waste drain, restarted instrument found instrument is functional. | SWO-181352 | [5 empty] | Solution & Blood leak. | Solution & Blood leak from waste. | Problem with waste drain, cleaned the waste drain, restarted instrument found instrument is functional.
No | Visit | Complaint | RCHN | 除PO2之外定标均失败 | 除PO2之外定标均失败 | 更换测试卡正常。 | SWO-174557 | [5 empty] | Bids failed except PO2 | Bids failed except PO2 | The replacement test card is OK.
No | 2 | Y | RAME | Z PCO2 not passing internal qc level 2. |  | Calibrations passing. Tried new sensor cassette and new solution pack; neither resolved the issue. | D23653EC9642C89E | [5 empty] | Z PCO2 not passing internal qc level 2. |  | Calibrations passing. Tried new sensor cassette and new solution pack; neither resolved the issue.
No | Visit | Complaint | RCHN | 血氧单元硬件错误 1034 | 血氧单元硬件错误 1034 ，借902-980一个。 |  | SWO-174584 | Defect | HW | Oxi module | Hemolyzer | 1034 | Blood oxygen unit hardware error 1034 | Blood oxygen unit hardware error 1034 , borrow 902-980 one.
No | Visit | Complaint | RHUN | Nyomtató hiba, Glu, Lac, Cl pontatlan | Nyomtató hiba, Glu, Lac, Cl pontatlan | Nyomtató cseréje, Glu, Lac membrán cserék, Mérőkamra tisztítása, Oktatás (QC mérések), Kalibrálás, QC mérés | SWO-174665 | Defect | HW | Printer Module | Printer Module |  | Printer error, Glu, LAC, Cl inaccurate | Printer error, Glu, LAC, Cl inaccurate | Printer replacement, Glu, Lac membrane replacements, measuring chamber cleaning, training (QC measurements), calibration, QC measurement
Yes | 9 | Y | RPAC | Blank | Sample errors |  | D235B78196293E49 | [5 empty] | Blank | Sample errors
No | 2 | Y | RFRA | Blank | Erreur S8377 |  | D235CD32962D823A | Setup | Replacement schedule | Replacement schedule | Replacement schedule |  | Blank | Error S8377
No | Visit | Complaint | RIND | Analyzer Showing Error Sate safe_22, Cartidge not placed Properly, Remove and reinstalled Cartridge, passed setsafe test, Restart the Analyzer, problem resolve. | Analyzer Not working. | Analyzer Showing Error Sate safe_22, Cartidge not placed Properly, Remove and reinstalled Cartridge, passed setsafe test, Restart the Analyzer, problem resolve. | SWO-175144 | User error | Consumables | Test Kit | Cartridge | Set_Safe_22 | Analyzer Showing Error Sate safe_22, Cartidge not placed Properly, Remove and reinstalled Cartridge, passed setsafe test, Restart the Analyzer, problem resolve. | Analyzer Not working. | Analyzer Showing Error Sate safe_22, Cartidge not placed Properly, Remove and reinstalled Cartridge, passed setsafe test, Restart the Analyzer, problem resolve.
No | Visit | Complaint | RDEU | 09.03.20167 16:00 H. Weingärtler 09191610391: Interventionsmeldung/Verfallsdatum 09.03.2017 17.00 Uhr, Dr. Weingärtner Tel.: 09191610334: Leckage erkannt, Lp,SC sind schon gewechselt 09.03.2017, 18.10 Uhr, Hebamme, Tel: 09191610334: Weiterhin Intervention erforderlich-jetzt Einlassdichtung wechseln | Fehlersuche --> Lösungspack, Sensorkassette und Einlassmodul überprüft --> Einlassdichtung verschlissen und Einlasszwischendichtung verschmutzt | Einlassdichtung (Kundenbestand) und Einlasszwischendichtung (Kundenbestand) erneuert, System und Funktion i.O. | SWO-174964 | Defect | HW | Auto Inlet module | Connector Gasket |  | 09.03.20167 16:00 H. Weingärtler 09191610391: Intervention notification/expiry date 09.03.2017 17.00, Dr. Weingärtner Tel.: 09191610334: Leak age detected, Lp,SC have already changed 09.03.2017, 6.10 pm, midwife, Tel: 09191610334: Further intervention required-now change inlet seal | Troubleshooting --> Solution pack, sensor cassette and inlet module checked --> Inlet seal worn and inlet intermediate seal dirty | Inlet seal (customer stock) and inlet intermediate seal (customer stock) renewed, system and function as it is
Yes | Visit | Complaint | RDEU | Kunde hat seine Bestellung an LP nicht bekommen | Kunde kann nicht messen, kein LP geliefert | LP aus GS besorgt und geliefert, | SWO-175113 | Re classified RMED | Logistical upSWO |  |  |  | Customer did not receive his order to LP | Customer can't measure, no LP delivered | LP from GS,
Yes | Visit | Complaint | RDEU | Kunde hat seine Bestellung an LP nicht bekommen | Kein LP geliefert, Kunde kann nicht messen | Lieferung u. Installation LP, Kal u. QC i.O. | SWO-175114 | Re classified RMED | Logistical upSWO |  |  |  | Customer did not receive his order to LP | No LP delivered, customer can't measure | Lieferung u. Installation LP, Kal u. QC i.O.
No | Visit | Complaint | RSCH | Fehler 1034 | Problemi ossimetrici. | Sostituito emolizzatore, apparecchio completamente ricalibrato e testato. | SWO-175257 | Defect | HW | Oxi module | Hemolyzer | 1034 | Error 1034 | Oxymetric problems. | Replaced molizer, fully recalibrated and tested device.
No | Visit | Complaint | RCHN | 定标失败 | 定标失败，借PCT定标卡944-394（12383），NT-proBNP944-258（11925）各2张 |  | SWO-174860 | Assisted Replacement | Consumables | Test Kit | Cartridge | PCT, NT-proBNP | Bid failed | Failed to bid, 2 each by PCT bid card 944-394 (12383), NT-proBNP944-258 (11925)
No | Visit | Complaint | RCHN | 定标底物外露检查失败 1024 | 定标底物外露检查失败 1024 | 更换备用光学单元903-364（903-364R0231N008）正常。 | SWO-175212 | Defect | HW | Process Module | Optical Unit | 1024 | Calibration substrate exposure check failed 1024 | Calibration substrate exposure check failed 1024 | Replacing the backup optical unit 903-364 (903-364R0231N008) is normal.
No | Visit | Complaint | RCHN | 报1068错误 | 报1068错误，借848-037一个 |  | SWO-174888 | Defect | HW | Process Module | Membrane Pump | 1068 | Reported 1068 error | Reported 1068 error, borrow848-037 a
No | Visit | Complaint | RCHN | 存储盘下面的加热板坏了 | 存储盘下面的加热板坏了,借902-478，903-138各一个 | 903-138后修复 | SWO-174566 | Defect | HW | Incubator Module | PCB, Incubator |  | The heating plate under the storage tray is broken. | The heating plate under the storage tray is broken, one each is 902-478, 903-138 | 903-138 Post-Fix
No | 37 | Y | RESP | Resultados discrepantes para el pH y oximetría |  |  | D237008F0C292611 | [5 empty] | Dissenting results for pH and oximetry
No | Visit | Complaint | RDEU | tHb und Glu. QC Fehler | tHb und Glu. QC Fehle. Glu. Dichtring fehlt. | Glu. Dichtring und Membran ersetzt. Reinigung, Proteinentfernung und 1.Kal., tHb Kal. durchgeführt, Kal. tHb Test und QCs i.O. 2x 834-618 Dichtring als Ersatz vor Ort gelassen. | SWO-175365 | Cleaned | HW | Oxi module | Oxi module, Cleaned | tHb | tHb und Glu. QC Fehler | tHb and Glu. QC misguided. Glu. Sealing ring is missing. | Glu. sealing ring and membrane. Cleansing, protein removal and 1.Cal., tHb Kal. carried out, Kal. tHb test and QCs in the o. 2x 834-618 sealing ring as replacement left on site.
No | Visit | Complaint | RDEU | 01.03.2017; Hr. Zwoch; Tel.: 089889220895; Touchscreen ohne Funktion 01.03.2017; Hr. Zwoch; Tel.: 089889220895; auch nach dem Neustart gibt es Probleme mit dem Touchscreen | Touchscreen reagiert in der Mitte sowie unten rechts nicht | Touchscreen erneuert und kalibriert Gerät überprüft | SWO-175073 | Defect | HW | DMS_PC | Touch screen |  | 01.03.2017; Mr. Zwoch; Tel.: 089889220895; Touchscreen without function 01.03.2017; Mr. Zwoch; Tel.: 089889220895; even after the restart there are problems with the touchscreen | Touchscreen does not respond in the middle and bottom right | Touchscreen renewed and calibrated Device checked
No | Visit | Complaint | RDEU | Nullstromfehler O2 Elektrolytfehler | O2 Elektrode und Ref Elektrode fehlerhaft. Ca Elektrode defekt. | O2elektrode erneuert. Referenz Elektrode aus kundenbestand erneuert. Mehrfach kalibriert. Ca lässt sich nicht einkalibrieren. Ca Elektrode erneuert. Elektrode muss noch einkalibriert werden. Station wird fals erforderlich sich am Montag erneut melden. | SWO-175096 | Assisted Replacement | Electrode | Electrode | pO2 |  | Nullstromfehler O2 Elektrolytfehler | O2 electrode and ref electrode faulty. Ca electrode defective. | O2electrode renewed. Reference electrode renewed from customer stock. Multiple calibrated. Ca cannot be calibrated. Ca electrode renewed. Electrode has yet to be calibrated. Station will report fals required again on Monday.
No | 6 | Y | RPNZ | Blank | AutoCheck No Go/Replace AutoCheck Module/AutoCheck Ampoule Reader |  | D23AB5A16E42CB04 | Defect | HW | Auto Check Module | Auto Check Module |  | Blank | AutoCheck No Go/Replace AutoCheck Module/AutoCheck Ampoule Reader
No | Visit | Complaint | RDEU | 02.03.2017 Herr Stark 02114092297 meldet eine erhöhte Empf. bei den Parametern pH & Ca 09.03.2017 16:30Uhr Sr.Birgit 02114092297: Weiterhin pH und Ca-fehler | pH und Ca Empf zu hoch, Knickstelle an der Einlassnadel, Nadel wurde gebogen | Elektrodenkappe Referenzelektrode gereinigt, Meßkammer und pH Elektrode gesäubert, O-Ring an der pH Elektrodenkappe und Einlassnadel ersetzt. Funktion z.Zt. OK. | SWO-175255 | Defect | HW | Inlet module | Probe bend |  | 02.03.2017 Mr. Stark 02114092297 reports an increased empf. parameters pH & Ca 09.03.2017 16:30h Sr.Birgit 02114092297: Continue pH and Ca-error | pH and Ca Empf too high, buckling at the inlet needle, needle has been bent | Electrode cap reference electrode cleaned, measuring chamber and pH electrode cleaned, O-ring replaced at the pH electrode cap and inlet needle. Function currently OK.
No | Visit | Complaint | RBNL | Storagewheel draait niet meer. | AQT staat in error, storagewheel is ook via het servicemenu niet werkzaam te krijgen. Na demontage blijkt dat er van een cartridge een lipje is afgebroken, waardoor deze niet op zijn plek zit en het storagewheel blokkeert. Na verwijderen defecte cartridge werkt AQT weer naar behoren. Aansluitend het halfjaarlijks onderhoud uitgevoerd, zie SWO-175017 | Defecte cartridge verwijdert. | SWO-175007 | Defect | Consumables | Test Kit | Cartridge | Broken | Storagewheel is no longer running. | AQT is in error, Storagewheel is also not able to get active via the service menu. After disassembly, it appears that a cartridge has broken down a tab, leaving it in place and blocking the Storagewheel. After uninstalling defective cartridge AQT will work properly again. Followed by the half-yearly maintenance, see SWO-175017 | Defective cartridge.
Yes | Visit | Complaint | RCHN | LAC定标失败 漂移错误 | LAC定标失败 漂移错误 | 更换测试卡正常。 | SWO-175221 | [5 empty] | LAC calibration failure Drift error | LAC calibration failure Drift error | The replacement test card is OK.
No | Visit | Complaint | RIND | Temperator not stable (High) for new sensor cassette | Checked the data log found Temperator not stable (High) for new cassette tried to install another 2 new SC it was showing verfication check failed cross checked signal data board found working confirmed signal data board is defective | Need to replace signal data board | SWO-175986 | [5 empty] | Temperator not stable (High) for new sensor cassette | Checked the data log found Temperator not stable (High) for new cassette tried to install another 2 new SC it was showing verfication check failed cross checked signal data board found working confirmed signal data board is defective | Need to replace signal data board
No | Visit | Complaint | RDEU | Flüss. kommt aus Einlass |  | Mit Fr. Hintzen die Einlassdichtung erneuert. Nadel überprüft und die beiden Schläuche gespült. Einlass jetzt trocken. | SWO-175190 | Defect | HW | Inlet module | Inlet Gasket |  | River. comes from admission |  | With Fr. Hintzen the inlet seal is renewed. needle and rinsed the two hoses. Entrance now dry.
No | Visit | Complaint | RLTD | Contact Penny Clark - 0131 5371664 | Analysser touch screen unresponsive | Rebooted analyser, still same Removed from department to replace display with new Spare analyser put in place New screen fitted and analyser put back in ICU New inlet clip fitted on both analysers | SWO-175753 | Defect | HW | DMS_PC | Touch screen |  | Contact Penny Clark - 0131 5371664 | Analysser touch screen unresponsive | Rebooted analyser, still same Removed from department to replace display with new Spare analyser put in place New screen fitted and analyser put back in ICU New inlet clip fitted on both analysers
No | Visit | Complaint | RDEU | Nach SP Tausch wird Fehlersuche angezeigt. | Fehler nach LP Wechsel. | Bei Ankunft kein Fehler mehr vorhanden. LP und Einlass nochmals überprüft. Keine Auffälligkeiten. Kal. , QC und tHb Kal. durchgeführt. | SWO-175994 | No fault found | No fault found | No fault found | No fault found |  | After SP swap, troubleshooting is displayed. | Error after LP change. | There is no more error on arrival. LP and inlet checked again. No abnormalities. Kal. , QC and tHb Kal. Performed.
No | Visit | Complaint | RCHN | 温度错误 | 温度错误 | 更换测试卡正常。 | SWO-175327 | [5 empty] | Temperature error | Temperature error | The replacement test card is OK.
No | Visit | Complaint | RCHN | 感应板锈了，固定感应板螺丝松了 | 感应板锈了，固定感应板螺丝松了，借924-738和902-972各一个 |  | SWO-175286 | [5 empty] | The sensing plate is rusty and the fixing sensor screw is loose. | The sensing plate is rusty, the fixed sensor screw is loose, and one each is 924-738 and 902-972
No | Visit | Complaint | RCHN | LAC定标失败 | LAC定标失败 | 更换测试卡后修复 | SWO-175705 | [5 empty] | LAC calibration failure | LAC calibration failure | Repair after replacing test card
No | 3 | Y | RPAC | Blank |  |  | D23FBCA646501D98 | [5 empty] | Blank
No | Visit | Complaint | RCHN | 报1265错误 | 报1265错误，借903-329一个 | 更换903-329一个正常。(By Ran) | SWO-175149 | Defect | HW | Aspiration Module | Needle Wash Unit |  | Reported 1265 errors | Reported 1265 error, borrowed 903-329 one | Replace the 903-329 one normal. (By Ran)
No | 11 | Y | RAME | 2 Pco2 level 1 running on low side of the peer 2 range |  | Reviewed logs and same as last visit no QC has failed sense 1 $16 $2017 | D23AC2616E24DDA8 | Revisit required | Electrode | Electrode | pCO2 | Technline | 2 Pco2 level 1 running on low side of the peer 2 range |  | Reviewed logs and same as last visit no QC has failed sense 1 $16 $2017
No | Visit | Complaint | RDEU | 13.03.2017; Hr. Kolbe; Tel.: 0821400165694; Probenproblem; Probe vorzeitig im El/Met und oder Oxy Flüssigkeitsfüher | Probenproblem, Elektrolyt- und Oximetrie-QC Fehler beide Luftventile im pH/BG sowie El. Met kleben fest Ventilmembran von Luftventil im El. Met Modul gerissen | Proteinentfernung durchgeführt - Messkammer gereinigt Ventile gängig gemacht, Ventilmembranen der Luftventile erneuert Gerät kalibriert und überprüft | SWO-175573 | Defect | HW | El/Met Module | Valve, air | Membrane | 13.03.2017; Mr. Kolbe; Tel.: 0821400165694; Sample problem; Sample prematurely in El/Met and or Oxy liquid feeder | Sample problem, electrolyte and oximetry QC error both air valves in pH/BG and El. Met stick tightly Valve membrane of air valve in The El. Met module ripped | Protein removal performed - measuring chamber cleaned Valves made common, valve membranes of the air valves renewed Device calibrated and checked
No | 9 | Y | RAME | V Chloride Electrode Connector Broken |  | Found the electrode connector missing the spring and could not find the spring. | D243CEF0A8283E1E | User error | HW | pH/BG Module | Electrode connector |  | V Chloride Electrode Connector Broken |  | Found the electrode connector missing the spring and could not find the spring.
No | 1 | Y | RFRA | Blank |  |  | D23AA0C06E5AE5F3 | Defect | HW | Cover and Chassis | Front Cover |  | Blank
No | Visit | Complaint | RDEU | 13.03.2017, Frau Schulz, Tel.: 045518021894: Fehler 4794 Nedle Detektion failed, am Wochenende ist eine Probe falsch herum geladen worden, seitdem gibt es Probleme. Der AQT wurde schon mehrmals neu gestartet. | Kundengespräch und Fehlersuche durchgeführt. Nadel total verbogen. Nadel bestellt. | Nadel bestellt. | SWO-175754 | Revisit required | HW | Aspiration Module | Bend needle | 4794 | 13.03.2017, Mrs Schulz, Tel.: 045518021894: Error 4794 Nedle Detection failed, at the weekend a sample was loaded wrongly around, since then there have been problems. The AQT has been restarted several times. | customer interview and troubleshooting. Needle totally bent. Needle ordered. | Needle ordered.
No | 5 | Y | RPOL | Blank | Zg³oszono czêste wymianu pakietów odczynnikowych. Kompletna wymiana owê¿owania, ponowna wymiana wygiêtej ig³y, za³o¿enie Inlet Clip, wymiana uszczelki wlotu. Oczyszczenie analizatora, kalibracja tHb. Testy kontrolne. Analizator sprawny. |  | D23C10A3E4215C05 | Defect | HW | Inlet module | Probe bend |  | Blank | The reported replacement of reagent packets. Complete replacement of the bearing, re-replacing the curved needle, the Inlet Clip, replacing the inlet gasket. Cleaning the analyser, calibrating the tHb. Control tests. Functional Analyzer.
No | 1 | Y | RPAC | Blank | Communication error window |  | D23BF727E45B6F4F | Defect | SW | Windows | Communication | V3.1 MR5 | Blank | Communication error window
No | Visit | Complaint | RDEU | 14.03.17, 8:30 Uhr, Herr Sahrhage, 052177279550 oder 551, 1342 Leckage erkannt, | Leckage erkannt 1342 | Gerät gespült, defekten Ventilschlauch ersetzt, checkout and adjustment, Gerät kalibriert und mit QCs überprüft,Ok | SWO-175985 | Defect | HW | Oxi module | Pinch valve tube |  | 14.03.17, 8:30, Mr. Sahrhage, 052177279550 or 551, 1342 Leakage detected, | Leakage detected 1342 | Device rinsed, defective valve hose replaced, checkout and adjustment, device calibrated and checked with QCs,Ok
No | Visit | Complaint | RCHN | LAC定标失败 | LAC定标失败 | 更换测试卡正常。 | SWO-176254 | [5 empty] | LAC calibration failure | LAC calibration failure | The replacement test card is OK.
Yes | Visit | Complaint | RCHN | LAC定标失败 | LAC定标失败 | 更换测试卡正常。 | SWO-175658 | [5 empty] | LAC calibration failure | LAC calibration failure | The replacement test card is OK.
No | Visit | Complaint | RCHN | 血氧参数无值 | 血氧参数无数值，借备件903-199一个 |  | SWO-175643 | Defect | HW | Oxi module | Spectrometer | 1034 | Blood oxygen parameters are worthless | There is no value for blood oxygen parameters, one spare part 903-199
No | 3 | Y | RESP | Ha perdido la conexión con Radiance y no envía los resultados al SIL |  | Chequeo de conexiones, cable y de configuración observando que la puerta de enlace ha cambiado y no es la correcta, se configura conecta a Radiance, configuración de vista de registro de pacientes ya que quieren que se vea la IDMuestra. prueba de envío de muestras, pcalibración y pasado qc. | D23C32B7E461AFE7 | Setup | His/lis | Roma.ini | Radiance |  | You have lost the connection to Radiance and do not send the results to the SIL |  | Checking connections, cable, and configuration by observing that the gateway has changed and is not correct, it is configured connect to Radiance, patient log view settings since they want the SampleID to be seen. sample submission test, precalibration and past qc.
Yes | Visit | Complaint | RCHN | PO2，LAC定标不过 | PO2，LAC定标不过 | 更换测试卡及试剂包后修复 | SWO-175690 | [5 empty] | PO2, LAC calibration but | PO2, LAC calibration but | Repair after replacing test cards and reagent kits
No | 66 | Y | RPAC | Blank | Corelations of the analyser with ICU & ED, not the best for MettHb and COtHB. |  | D23BF5A6E40EABDE | Defect | HW | Oxi module | Oxi module, Replaced |  | Blank | Corelations of the analyser with ICU & ED, not the best for MettHb and COtHB.
No | Visit | Complaint | RIND | Paper feed not working | Checked the printer head and found gear assembly broken | Replaced printer head from a working ABL800, printer is working now. | SWO-175693 | Defect | HW | Printer Module | Printer head |  | Paper feed not working | Checked the printer head and found gear assembly broken | Replaced printer head from a working ABL800, printer is working now.
No | 12 | Y | RPOL | Blank | Wymiana hemolizatora. Kalibracja hemoglobiny od strony serwisu. Testy serwisowe - ok. Analizator sprawny technicznie, |  | D23C16C3E4488855 | Defect | HW | Oxi module | Hemolyzer |  | Blank | Replacement of the Haemolyser. Calibration of hemoglobin from the site. Service Tests-OK. Technically efficient analyzer,
No | Visit | Complaint | RDAN | Jan Pedersen melder, at ABL'en har alle de oxifejl der findes. | Afprøvet og undersøgt. Spectrometeret og temperatur reguleringen virker! Resten fejler! Skiftet neon lamp og PCB'et uden resultat. | Hemolyzer udskiftet, opjusteret, service tHb kalibreret, slutafprøvet. | SWO-175772 | Defect | HW | Oxi module | Hemolyzer |  | Jan Pedersen reports that the ABL has all the oxidityerrors that exist. | Tested and investigated. The spectrometer and temperature regulation works! The rest fails! Changed the Neon lamp and the PCB without result. | Hemolyzer replaced, upwards adjusted, service tHb calibrated, final tested.
No | Visit | Complaint | RLTD | Lisa 01865 221596 | customer reporting that the analyser intermittently has problems inputting data using the on screen keyboard. | checked over the analyser and but experienced no issues whilst investigating on site. decided to run a screen calibration within the windows control panel. advised customer to monitor and report back if problem persists. | SWO-175937 | Defect | HW | DMS_PC | Touch screen |  | Lisa 01865 221596 | customer reporting that the analyser intermittently has problems inputting data using the on screen keyboard. | checked over the analyser and but experienced no issues whilst investigating on site. decided to run a screen calibration within the windows control panel. advised customer to monitor and report back if problem persists.
No | Visit | Complaint | RLTD | Lisa 01865 221596 | Customer reporting that they have been experiencing leaks coming from the inlet of the analyser during rinse after a sample has been run. This happens on both samples and QCs that are run. Customer had tried to replace the inlet gasket and had run various flush and tubing checks. | Arrived on site and attended analyser. found that the analyser had some qc staining on the auto check cover which indicated it may have been leaking from the inlet gasket after the QC run that morning. opened the analyser and ran a qc then monitored the rinse afterwards. no leak coming from the gasket. checked the tubings for blockages or perforations. found none. flushed the waste lines and soaked with protein remover to break down any possible clots or blockages. checked for debris being flushed out of the waste bottle nipple onto the tissue placed below it. none removed. checked the inlet by removing the gasket and inspecting the seal. this looked in good condition. inspected the inlet probe and the blue inlet probe tubing. these both looked good. calibrated the inlet probe as a precaution. refitted all parts and ran a number of rinses in a row to check the rinse procedure and monitor the inlet for leakage. none found and all looked to be in working order. tried clearing the blue inlet nipple on the fluidic module to ensure this was clear. upon rinsing, there was no fluid drawn in to the nipple. inspected the valve operation. this did not seem correct. checked in service programs for any flagged failures. none shown. powered down the device and removed $inspected and cleaned all valves for signs of perforation or blockage. refitted all parts and powered up the analyser. set all settings to default and retested the blue inlet nipple. this now took liquid in and rinsed as it should. suspect small blockage in waste valve. calibrated and qcd analyser. customer will monitor. | SWO-175940 | Cleaned | Fluidic pathway | Inlet module | Inlet module |  | Lisa 01865 221596 | Customer reporting that they have been experiencing leaks coming from the inlet of the analyser during rinse after a sample has been run. This happens on both samples and QCs that are run. Customer had tried to replace the inlet gasket and had run various flush and tubing checks. | Arrived on site and attended analyser. found that the analyser had some qc staining on the auto check cover which indicated it may have been leaking from the inlet gasket after the QC run that morning. opened the analyser and ran a qc then monitored the rinse afterwards. no leak coming from the gasket. checked the tubings for blockages or perforations. found none. flushed the waste lines and soaked with protein remover to break down any possible clots or blockages. checked for debris being flushed out of the waste bottle nipple onto the tissue placed below it. none removed. checked the inlet by removing the gasket and inspecting the seal. this looked in good condition. inspected the inlet probe and the blue inlet probe tubing. these both looked good. calibrated the inlet probe as a precaution. refitted all parts and ran a number of rinses in a row to check the rinse procedure and monitor the inlet for leakage. none found and all looked to be in working order. tried clearing the blue inlet nipple on the fluidic module to ensure this was clear. upon rinsing, there was no fluid drawn in to the nipple. inspected the valve operation. this did not seem correct. checked in service programs for any flagged failures. none shown. powered down the device and removed $inspected and cleaned all valves for signs of perforation or blockage. refitted all parts and powered up the analyser. set all settings to default and retested the blue inlet nipple. this now took liquid in and rinsed as it should. suspect small blockage in waste valve. calibrated and qcd analyser. customer will monitor.
No | Visit | Complaint | RDEU | 13.03.2017, Frau Schulz, Tel.: 045518021894: Fehler 4794 Nedle Detektion failed, am Wochenende ist eine Probe falsch herum geladen worden, seitdem gibt es Probleme. Der AQT wurde schon mehrmals neu gestartet. | Kundengespräch durchgeführt. Nadel erneuert und justiert. Systemüberprüfung durchgeführt. | Nadel erneuert und justiert. | SWO-175755 | Defect | HW | Aspiration Module | Bend needle |  | 13.03.2017, Mrs Schulz, Tel.: 045518021894: Error 4794 Nedle Detection failed, at the weekend a sample was loaded wrongly around, since then there have been problems. The AQT has been restarted several times. | customer interview. Needle renewed and adjusted. system verification. | Needle renewed and adjusted.
No | Visit | Complaint | RCHN | LAC定标失败 | LAC定标失败 | 连续更换2张测试卡后修复 | SWO-177455 | [5 empty] | LAC calibration failure | LAC calibration failure | Repair after 2 consecutive test cards
No | Visit | Complaint | RCHN | 不均匀的冲洗液 | 不均匀的冲洗液 | 更换试剂包后修复 | SWO-179664 | Assisted Replacement | SP | SP | FT | 1081 | Uneven flushing fluid | Uneven flushing fluid | Repair after replacement of reagent kit
Yes | 12 | Y | RPOL | Blank | Diagnoza problemu, testy serwisowe, oczyszczenie modu³u OXI, testy kontrolne. Analizator sprawny. |  | D23D80A05A185C0F | [5 empty] | Blank | Problem diagnosis, service tests, OXI module cleaning, control tests. Functional Analyzer.
No | Visit | Complaint | RDEU | Kunde musste Lac-Membran nach 14 Tagen wieder ersetzen Info durch Ralf Ehlert | Lac muss alle 14 Tage ersetzt werden | Lac neu membraniert mit D7077, Lot R3473, akz. Kal abgewartet, Kal u. QC i.O. | SWO-176103 | Assisted Replacement | Electrode | Membrane | Lac | Sensitivity | Customer had to replace Lac membrane after 14 days Info by Ralf Ehlert | Lac must be replaced every 14 days | Lac re-membraneed with D7077, lot R3473, akz. Kal waited, Kal & QC i.O.
No | Visit | Complaint | RDEU | Lac muss alle 14 Tage neu mebraniert werden, da QC- bzw. Kal-Fehler | Lac muss 14tägig membraniert werden | Lac neu membraniert (Lot R 3473), akz. Kal abgewartet, Kal u. QC i.O. | SWO-176154 | Assisted Replacement | Electrode | Membrane | Lac | Sensitivity | Lac muss alle 14 Tage neu mebraniert werden, da QC- bzw. Kal-Fehler | Lac must be membraned 14 days a day | Lac re-membraneed (Lot R 3473), akz. Kal waited, Kal & QC i.O.
No | Visit | Complaint | RDEU | 14.03.17, 13:45 Uhr, Herr Kurscheid, 02241187361, Starke Driften bei fast allen Parametern und Anbindungsprobleme | Gerät kann keine Daten an die Herzlungenmaschine senden. Nach jeder Messung wird eine Kalibration durchgeführt. Lac und Glu Kal Fehler. | Haken bei Kalibration nach jeder Messung raus genommen. Glu und Lac neu membraniert. Mehrfach Kalibriert. Optionkeys installiert. Fehlersuche an der Netzwerkverbindung durchgeführt, fehler der Kis Verbindung nicht beheben können. | SWO-176341 | Setup | His/lis | His/lis | His/lis |  | 14.03.17, 13:45, Mr. Kurscheid, 02241187361, Strong drifts with almost all parameters and connection problems | Device cannot send data to the heartlung machine. After each measurement, a calibration is performed. Lac and Glu Kal error. | Hook scalding taken out after each measurement. Glu and Lac. Multiple calibrated. Optionkeys installed. Troubleshooting the network connection, the Kis connection fails.
No | Visit | Complaint | RDEU | CPU Unit Einbau | CPU Unit Einbau | CPU Unit in der GS eingebaut. Startphase abgewartet, NT-pro BNP kalibriert. | SWO-177501 | Defect | HW | Analyzer Control Module | CPU |  | CPU Unit Installation | CPU Unit Installation | CPU unit built into the GS. start-up phase, NT-pro BNP calibrated.
No | Visit | Complaint | RLTD | Intermittent Inhomogeneous solution errors since installation on 09/03/17. | Analyser going into troubleshooting due to various Inhomogeneous solution errors. Fault traced to incorrectly fitted inlet connector gasket. Reseated this gasket and fluid transport errors resolved. Then encountered errors with pO2 QC level A reading low and level C reading high - resolved by reseating Oximetry pinch valve tube at Sensor Bay. Also had Lactate calibration errors - resolved by cleaning contact pins with alcohol. Analyser verified after restart - now passing calibration and quality controls on all parameters. | Reseated inlet connector gasket and Oximetry pinch valve tube. Cleaned Sensor Cassette bay contact pins. | SWO-177697 | Defect | HW | Oxi module | Pinch valve tube/blood |  | Intermittent Inhomogeneous solution errors since installation on 09/03/17. | Analyser going into troubleshooting due to various Inhomogeneous solution errors. Fault traced to incorrectly fitted inlet connector gasket. Reseated this gasket and fluid transport errors resolved. Then encountered errors with pO2 QC level A reading low and level C reading high - resolved by reseating Oximetry pinch valve tube at Sensor Bay. Also had Lactate calibration errors - resolved by cleaning contact pins with alcohol. Analyser verified after restart - now passing calibration and quality controls on all parameters. | Reseated inlet connector gasket and Oximetry pinch valve tube. Cleaned Sensor Cassette bay contact pins.
No | Visit | Complaint | RLTD | Failing pO2 QC level A low and Level C high | Issue resolved by replacing the Oximetry pinch valve tube to improve seal to the Sensor Cassette bay. Verified operation on restart - all ok. | Replaced Oximetry pinch valve tube. | SWO-177698 | Defect | HW | Oxi module | Pinch valve tube/blood |  | Failing pO2 QC level A low and Level C high | Issue resolved by replacing the Oximetry pinch valve tube to improve seal to the Sensor Cassette bay. Verified operation on restart - all ok. | Replaced Oximetry pinch valve tube.
No | Visit | Complaint | RLTD | PO2 failing quality controls, level A reading low, level C reading high. | Issue resolved by reseating the Oximetry pinch valve tube at the Sensor Cassette bay connection. Verified operation on restart - now passing pO2 on quality controls, in middle of control limits. | Reseated Oximetry pinch valve tube. | SWO-177696 | Defect | HW | Oxi module | Pinch valve tube/blood |  | PO2 failing quality controls, level A reading low, level C reading high. | Issue resolved by reseating the Oximetry pinch valve tube at the Sensor Cassette bay connection. Verified operation on restart - now passing pO2 on quality controls, in middle of control limits. | Reseated Oximetry pinch valve tube.
No | Visit | Complaint | RDEU | Keine Oximetrie-Werte | Keine Oximetrie-Werte | Gerät aus dem OP abgezogen, Hemolyzer erneuert, checkout and adjustment, Gerät kalibriert, mit QCs überprüft und div. Blutmessungen durchgeführt,OK. defekt: 908-552R1841N019 | SWO-177427 | Defect | HW | Oxi module | Hemolyzer |  | No oximetry values | No oximetry values | Device removed from the operating room, Hemolyzer renewed, checkout and adjustment, device calibrated, checked with QCs and various blood measurements performed,OK. defective: 908-552R1841N019
No | Visit | Complaint | RCHN | PO2测量不准 | PO2测量不准,借902-972，903-334各一个 | 更换902-972一个正常。(By Ran) | SWO-175727 | [5 empty] | PO2 measurement is not allowed | PO2 is not measured, borrow 902-972, 903-334 each | Replace 902-972 one normal. (By Ran)
No | Visit | Complaint | RCHN | 接口板腐蚀 | 接口板腐蚀 | 更换备用机件902-972后修复(By Ran) | SWO-177482 | [5 empty] | Interface plate corrosion | Interface plate corrosion | Repair after replacement of spare parts 902-972 (By Ran)
Yes | Visit | Complaint | RCHN | 多参数定标不过，气泡探测定标失败 | 多参数定标不过，气泡探测定标失败 | 更换测试卡及换试剂包后修复 | SWO-177459 | [5 empty] | Multi-parameter calibration, but bubble detection calibration failed | Multi-parameter calibration, but bubble detection calibration failed | Repair after replacing test cards and reagent kits
No | Visit | Complaint | RIND | Leakage from SC80 while installation, Sensor connector get Defective, clean Sensor Board and Tried to reinstall, Problem Remains, Need Sensor Board, SC80 PER need to sent. waiting for Spare. | Leackage for SC80 | Leakage from SC80 while installation, Sensor connector get Defective, clean Sensor Board and Tried to reinstall, Problem Remains, Need Sensor Board, SC80 PER need to sent. waiting for Spare. | SWO-177820 | [5 empty] | Leakage from SC80 while installation, Sensor connector get Defective, clean Sensor Board and Tried to reinstall, Problem Remains, Need Sensor Board, SC80 PER need to sent. waiting for Spare. | Leackage for SC80 | Leakage from SC80 while installation, Sensor connector get Defective, clean Sensor Board and Tried to reinstall, Problem Remains, Need Sensor Board, SC80 PER need to sent. waiting for Spare.
No | Visit | Complaint | RCHN | 测试卡温度超范围 1064，1065，1066 | 测试卡温度超范围 1064，1065，1066，借903-183一个 | 更换903-183一个正常。(By Ran) | SWO-176019 | Defect | HW | Basic Unit | PCB, Sensor interface |  | Test card temperature over range 1064,1065,1066 | Test card temperature over range 1064,1065,1066, borrowed 903-183 one | Replace 903-183 one normal. (By Ran)
No | Visit | Complaint | RCHN | 血氧单元氖灯光强 1020 | 血氧单元氖灯光强 1020 ，借 903-165，903-197，902-980各一个。 | 更换903-197后修复(By Ran) | SWO-177249 | Defect | HW | Oxi module | Spectrometer | 1020 | Blood oxygen unit xenon light strong 1020 | Blood oxygen unit xenon light strong 1020 , borrow 903-165, 903-197, 902-980 each. | Post-replacement 903-197 repair (By Ran)
No | Visit | Complaint | RIND | Patient printout not clearly visible | Checked the machine Patient printout not clearly visible found printer head defective | Need to replace printer head | SWO-176284 | Revisit required | HW | Printer Module | Printer head |  | Patient printout not clearly visible | Checked the machine Patient printout not clearly visible found printer head defective | Need to replace printer head
No | Visit | Complaint | RIND | Analyser showing the error message DMS not connected to analyzing unit. | Showing the error message DMS Not connected to analyzing unit. | Checked the connection between DMS PC and Analyzer Control PCB. After reconnecting the cable analyser found working fine . But repeating the same error intermittently. Has to check with new 636- 506 :Power $Communication Cable, Analyzer Control | SWO-177299 | Revisit required | HW | Basic Unit | PCB, Analyzer control |  | Analyser showing the error message DMS not connected to analyzing unit. | Showing the error message DMS Not connected to analyzing unit. | Checked the connection between DMS PC and Analyzer Control PCB. After reconnecting the cable analyser found working fine . But repeating the same error intermittently. Has to check with new 636- 506 :Power $Communication Cable, Analyzer Control
No | 9 | Y | RAME | 1 pH issue |  | 5 Blood samples taken with low pH readings. No patient injury occurred as stated by blood gas coordinator. Possible cause could have been blood clot. Follwoing the 5 consecutive blood samples,on the next calibration pH drift occurred. Data logs show no further calibration errors since event. | D23EB60CD0487930 | Cleaned | Fluidic pathway | pH/BG Module | Measuring chamber, Clot |  | 1 pH issue |  | 5 Blood samples taken with low pH readings. No patient injury occurred as stated by blood gas coordinator. Possible cause could have been blood clot. Follwoing the 5 consecutive blood samples,on the next calibration pH drift occurred. Data logs show no further calibration errors since event.
No | 3 | Y | RAME | 2 pCO2 QC results are running low in peer comparison. |  | QC results for all levels have been 2 mmHg lower since instal and trended down begining 12-18-16. There have been many rinse errors from pH $BG upper LS sensor. There is evidence of a leak at top of module pump tubing. Gas 2 is also the original since install. | D23EA37FD00FAF60 | Defect | HW | pH/BG Module | Pump tubes |  | 2 pCO2 QC results are running low in peer comparison. |  | QC results for all levels have been 2 mmHg lower since instal and trended down begining 12-18-16. There have been many rinse errors from pH $BG upper LS sensor. There is evidence of a leak at top of module pump tubing. Gas 2 is also the original since install.
No | 2 | Y | RAME | P Customer reported by email, error code 0793, pH/BG asymmetrical pump error. |  | Customer had already replaced pump tubing on pH $BG module. Checked pump rotor assy. Pump calibration showed calibration error. | D23E9841D04368D5 | Defect | HW | pH/BG Module | Pump tubes |  | P Customer reported by email, error code 0793, pH/BG asymmetrical pump error. |  | Customer had already replaced pump tubing on pH $BG module. Checked pump rotor assy. Pump calibration showed calibration error.
No | Visit | Complaint | RDEU | 15.03.2017, Frau Asters, tel: 0714410196086: AQT reagiert nicht mehr. Nach Einlegen einer Probe wurde der Barcode nicht vollständig gelesen und der AQT blieb in diesem Bild hängen. | Gerät stürzt vereinzelt ab. Friert ein. | Servicedump gezogen. Daten zur weiteren Auswertung nach Willich geschickt. | SWO-177404 | Revisit required | SW | Aspiration Module | Barcode reader, Inlet |  | 15.03.2017, Ms. Asters, tel: 0714410196086: AQT stops responding. After inserting a sample, the barcode was not fully read and the AQT got stuck in this image. | Device crashes sporadically. Freezes. | Service dump. data sent to Willich for further evaluation.
No | 2 | Y | RPOL | Blank | Pomoc u¿ytkownikowi w konfiguracji przesy³u danych pacjenta. Analizator sprawny. |  | D23E50AAD0301FCA | Setup | His/lis | His/lis | His/lis |  | Blank | User assistance in configuring the patient data transmission. Functional Analyzer.
No | Visit | Complaint | RDEU | Gerät fordert immer wieder SC Erneuerung an. Glu Probleme. | SC defekt. | Gerät überprüft und prophylaktisch 10-15 durchgespült. Halterung der Einlassdichtung seitlich gebrochen. Einlassdichtung mit Halterung durch Kundeneigene ersetzt. LP und SC erneuert, Gerät kalibriert, SC erneuert. Absprache mit Kundin getroffen, dass eine telefonische Rücksprache nach Initialisierung der SC erfolgen wird. | SWO-177774 | Assisted Replacement | SC | SC | Glu |  | Device repeatedly requests SC renewal. Glu problems. | SC defective. | device and rinsed prophylactic 10-15. Bracket of the inlet seal broken sideways. Inlet seal with bracket replaced by customer's own. LP and SC renewed, device calibrated, SC renewed. Agreement with customer that a telephone consultation will take place after initialization of the SC.
No | Visit | Complaint | RDEU | nach der 5-6 Messung friert der Bildschirm ein, Probe wird nicht ausgedruckt, erst nach einem Neustart kann das Gerät wieder benutzt werden und der Patientenausdruck kann gedruckt werden, |  | ich habe die Sotware noch einmal aufgespielt, | SWO-177733 | Defect | SW | SW | Freeze |  | after the 5-6 measurement the screen freezes, sample is not printed, only after a restart the device can be used again and the patient printout can be printed, |  | I have replayed the Sotware,
Yes | 6 | Y | RPAC | Blank | Sensor cassette failure (1361) - Internal reference electrode error in sensor cassette. |  | D24036594647345F |  | PER |  |  | HEAD LINE = "Consumable Failure" = PER | Blank | Sensor cassette failure (1361) - Internal reference electrode error in sensor cassette.
Yes | 1 | Y | RPAC | Blank | Provided S Box with new analyser |  | D244B8861E5D4519 | Re classified RMED | Logistical upSWO |  |  |  | Blank | Provided S Box with new analyser
No | Visit | Complaint | RLTD | PO2 failing quality controls, Level A reading low, Level C reading high. | Replaced oximetry pinch valve tube and reseated Sensor Cassette in Sensor Bay. Further issues with pH and Glucose failing intermittently. Flushed and reseated Sensor Cassette again. Now all parameters are working correctly. | Replaced Oximetry pinch valve tube. Flushed and reseated Sensor Cassette. | SWO-177699 | Defect | HW | Oxi module | Pinch valve tube/blood |  | PO2 failing quality controls, Level A reading low, Level C reading high. | Replaced oximetry pinch valve tube and reseated Sensor Cassette in Sensor Bay. Further issues with pH and Glucose failing intermittently. Flushed and reseated Sensor Cassette again. Now all parameters are working correctly. | Replaced Oximetry pinch valve tube. Flushed and reseated Sensor Cassette.
No | Visit | Complaint | RCHN | PH,CO2定标失败 | PH,CO2定标失败 | 更换试剂包和测试卡后修复 | SWO-177878 | [5 empty] | PH, CO2 calibration failure | PH, CO2 calibration failure | Repair after replacement of reagent kits and test cards
No | 10 | Y | RAME | 3 pO2 failing system cycles. |  | Found P02 disabled and failing system cycles on P02. Customer has replaced two SC and one SP and P02 is still not calibrating. | D2489904800E4BB4 | [5 empty] | 3 pO2 failing system cycles. |  | Found P02 disabled and failing system cycles on P02. Customer has replaced two SC and one SP and P02 is still not calibrating.
No | Visit | Complaint | RDEU | 17.03.2017, Herr Schlei, Tel.: 07819321752, : Herr Schlei berichtet, Fließwege verstopf-im Service Menue die Fleißwege überprüft-ok, reinitialisierung der SC nicht möglich, neue SC läßt sich nicht einsetzten-bzw initialisieren. | Temperatur instabil | neue SC eingesetzt, Initialisierungsabbruch nach 29%, Zero Check durchgeführt und rekalibriert, SC erneut initialisiert, Kal. und QC i.O. | SWO-177633 | [5 empty] | 17.03.2017, Mr. Schlei, Tel.: 07819321752, : Mr. Schlei reports, flow paths clogged-in the service menu checks the diligent paths-ok, reinitialization of the SC cannot be possible, new SC cannot be used or initialized. | Temperature unstable | new SC used, initialization abort after 29%, zero check performed and recalibrated, SC reinitialized, Kal. and QC in the first place
Yes | Visit | Complaint | RCHN | LAC定标、质控均失败 | LAC定标、质控均失败 | 更换测试卡后修复 | SWO-177955 | [5 empty] | LAC calibration, quality control are failed | LAC calibration, quality control are failed | Repair after replacing test card
No | Visit | Complaint | RCHN | 定标失败 1024 | 定标失败 1024 ，借903-364一个。 | 更换903-364后修复 | SWO-176334 | Defect | HW | Process Module | Optical Unit | 1024 | Bid Failure 1024 | Bid Failure 1024 , borrow edith of 903-364. | Repair after replacement 903-364
No | Visit | Complaint | RCHN | 定标失败 | 定标失败，借944-258（12301）一张 | 重新定标后修复(By Ran) | SWO-177480 | Adjustment | Consumables | Test Kit | Cartridge | 5784 (Broken) | Bid failed | Failed to bid, borrow 944-258 (12301) | Repair after re-calibration (By Ran)
No | Visit | Complaint | RCHN | 进样仓左侧门损坏 | 进样仓左侧门损坏 ，借886-638一个。 | 更换886-638一个正常。by jesson | SWO-178261 | Defect | HW | Cover and Chassis | Inlet Door |  | Damage to the left door of the sample silo | Damage to the left door of the sample silo , borrow edith 886-638. | Replace 886-638 one normal. by jesson
No | Visit | Complaint | RIND | crea calibration expire. | Check the system. check the logs. | Replace the Crea reagent & calibrate system. | SWO-177841 | Assisted Replacement | Electrode | Membrane | Crea Solution |  | crea calibration expire. | Check the system. check the logs. | Replace the Crea reagent & calibrate system.
No | Visit | Complaint | RDEU | Kunde hat gestern nach Gasfehler defekte Ventilmembran des Gaseinlaßventils ersetzt, weiterhin Gasfehler vorhanden | Reinigungsarbeiten, Fehlersuche | Membranventile der Messmodule kontrolliert, Verstopfung am Gaseinlaßventil beseitigt, Spülfehler, Kontrolle Spülzyklus, Lufteinlaßventil am ElMetModul ersetzt, Inbetriebnahme, korrekte Kal abgewartet, QK ok | SWO-177861 | Defect | HW | El/Met Module | Valve, air |  | Customer replaced faulty valve membrane of the gas inlet valve yesterday after gas fault, gas fault still present | Cleaning work, troubleshooting | Diaphragm valves of the measuring modules checked, clogging on the gas inlet valve eliminated, flushing error, control flushing cycle, air inlet valve replaced at ElMetmodule, commissioning, correct calving waited, QK ok
No | Visit | Complaint | RKK | pCO2 QC測定値が変動する | pCO2 QC測定値が変動する | 修理完了装置返却 | SWO-177622 | Assisted Replacement | Electrode | Membrane | pCO2 |  | PCO2 QC readings fluctuate | PCO2 QC readings fluctuate | Return repair completion equipment
No | Visit | Complaint | RDEU | 16.03.2017; Hr. Bärwald; Tel.: 03055184118; IP Adresse muss geändert | Gerät konnte noch nicht auf Conworx umgestellt werden wegen Schnittstelle zu anderem IT System, Anbindung an Conworx | Anbindung gemäß Vorgaben durchgeführt, Kontrolle der Verbindung, Kal und QC | SWO-177663 | Setup | His/lis | His/lis | His/lis |  | 16.03.2017; Mr. Bärwald; Tel.: 03055184118; IP address must be changed | Device could not yet be switched to Conworx due to interface to other IT system, connection to Conworx | Connection carried out according to specifications, control of the connection, Kal and QC
No | 3 | Y | RAME | F Customer reported oximetry errors during calibrations. |  | Unit getting error code 0316: Neon Number Too Low. Tested oximetry module. Failing Neon position too low. Tried adjusted lamp and running tHB calibration. This did not correct issue. | D240029E46473A1D | Defect | HW | Oxi module | Spectrometer | Lamp unit ? | F Customer reported oximetry errors during calibrations. |  | Unit getting error code 0316: Neon Number Too Low. Tested oximetry module. Failing Neon position too low. Tried adjusted lamp and running tHB calibration. This did not correct issue.
No | Visit | Complaint | RDEU | 16.03.2017 10:50Uhr Frau Anika Hochhäuser 084216019434: Die Kontrolle bei NT-ProBNP ist immer zu niedrig. | LQC Fehler bei NT-ProBNP | Neue Charge von NT-ProBNP kalibriert, Systemüberprüfung gestartet, Systemüberprüfungen alle in Ordnung, ServiceDump gezogen, nach ServiceDump Software Version 8.8.16 erneut installiert, Startphase abgewartet, eine weitere Systemüberprüfung gestartet - alles in Ordnung, neue LQC Charge eingelesen, LQC MCC2 Level 2 - Lot 06162 gemessen, TnT und NT-ProBNP war erfolgreich, System- und Funktionskontrolle in Ordnung | SWO-177727 | Assisted Replacement | Consumables | Test Kit | NT-Pro BNP |  | Mar 16, 2017 10:50AM Ms Anika Hochhäuser 084216019434: Control at NT-ProBNP is always too low. | LQC Fehler bei NT-ProBNP | New batch of NT-ProBNP calibrated, system check started, system checks all in order, ServiceDump pulled, reinstalled after ServiceDump software version 8.8.16, start phase waited, another system check started - all right, new LQC batch read in, LQC MCC2 Level 2 - Lot 06162 measured, TnT and NT-ProBNP was successful, system and function control in order
No | 4 | Y | RFRA | Blank | couvercle amovible incubateur wheel manquant |  | D23F9CA1460A3E48 | Defect | HW | Incubator Module | Incubator Lid |  | Blank | removable incubator cover missing wheel
No | Visit | Complaint | RCHN | 测试卡参比电极错误 | 测试卡参比电极错误 | 更换测试卡后修复 | SWO-177815 | Assisted Replacement | SC | SC | Reference |  | Test card reference electrode error | Test card reference electrode error | Repair after replacing test card
No | Visit | Complaint | RCHN | LAC,PCO2定标不过 | LAC,PCO2定标不过 | 更换测试卡后修复 | SWO-178268 | Assisted Replacement | SC | SC | Lac | Sensitivity | LAC, PCO2 calibration but | LAC, PCO2 calibration but | Repair after replacing test card
No | Visit | Complaint | RCHN | 舱门无法开启关闭 | 舱门无法开启关闭，借备件925-016一个 | 更换925-016后修复(By Ran) | SWO-177817 | Defect | HW | Cover and Chassis | Inlet Door |  | The hatch cannot be opened and closed | The hatch cannot be opened and closed, spare parts 925-016 one | Repair after replacement 925-016 (By Ran)
No | 7 | Y | RAME | Z QC aborting. |  | Autocheck aborting QC without user message. Activity log shows error 0742, Autocheck Carrosel failure - couldn't move. | D2420EAA323F35F3 | Defect | HW | Auto Check Module | FPC Cable |  | Z QC aborting. |  | Autocheck aborting QC without user message. Activity log shows error 0742, Autocheck Carrosel failure - couldn't move.
No | Visit | Complaint | RDEU | 20.03.17 16:00 Herr Kortz Tel.071164893256 , TouchScreen lässt sich nicht mehr richtig bedienen . | Touch Screen lässt sich ungenau bedienen. | Touch Kal. im Controll Panel durchgeführt. Funktionen wieder normal. | SWO-178837 | Adjustment | HW | DMS_PC | Touch screen |  | 20.03.17 16:00 Mr. Kortz Tel.071164893256, TouchScreen can no longer be operated properly. | Touch screen can be operated inaccurately. | Touch Kal. performed in the Controll Panel. functions back to normal.
No | Visit | Complaint | RCHN | LAC定标错误 | LAC定标错误，借备件903-334一个 | 更换备件903-334后修复 by andrew | SWO-177834 | [5 empty] | LAC calibration error | LAC calibration error, borrowed spare parts 903-334 one | Repair by Andrew after replacing spare parts 903-334
No | Visit | Complaint | RIND | Machine is not getting on | Machine is not getting on cross checked power supply board found working confirmed power supply board is defective | Need to replace power supply board | SWO-179690 | [5 empty] | Machine is not getting on | Machine is not getting on cross checked power supply board found working confirmed power supply board is defective | Need to replace power supply board
No | 4 | Y | RPNZ | Blank | pH QC low. |  | D2440037A843B5F4 | [5 empty] | Blank | pH QC low.
No | 12 | Y | RAME | Z Solution pack door not staying fully closed. |  | Paula said the solution pack door would come back a little when she latches it closed. She believes it is not pushing the SP in tight enough. Replaced door hinge and it is tighter. Will also claim the COOX zero cassette to get off my inventory. | D2435D0FA85FFB31 | [5 empty] | Z Solution pack door not staying fully closed. |  | Paula said the solution pack door would come back a little when she latches it closed. She believes it is not pushing the SP in tight enough. Replaced door hinge and it is tighter. Will also claim the COOX zero cassette to get off my inventory.
No | Visit | Complaint | RCHN | O2定标失败 | O2定标失败，借945-605一个 |  | SWO-178015 | [5 empty] | O2 calibration failure | O2 bid failure, borrow 945-605 one
No | Visit | Complaint | RCHN | 血氧参数无值 | 仍旧出现间歇性故障，借备件903-664,910-285各一个 |  | SWO-177869 | Defect | HW | Basic Unit | Cassette retainer | 1034 | Blood oxygen parameters are worthless | Intermittent failure continues, with spare parts 903-664,910-285 each
No | Visit | Complaint | RCHN | 测试卡参比电极错误 | 测试卡参比电极错误 | 更换测试卡后修复 | SWO-177891 | Assisted Replacement | SC | SC | Reference | 1361 | Test card reference electrode error | Test card reference electrode error | Repair after replacing test card
No | Visit | Complaint | RCHN | 定标中止 | 定标中止，借定标卡944-214（12231），944-216（12278）各一张 | 重新定标正常。 by andrew | SWO-177850 | Assisted Replacement | Consumables | Test Kit | Cartridge, Myo | 5297 | Bid abort | Bid suspension, borrowing bid card 944-214 (12231), 944-216 (12278) each | Recalibration is normal. by andrew
No | Visit | Complaint | RCHN | 液路控制板漏液，4.5号马达被腐蚀 | 液路控制板漏液，4.5号马达被腐蚀，借580-096及924-591各一个 |  | SWO-177910 | Defect | HW | Aspiration Module | Wet Section, Manifold |  | Liquid circuit control plate leakage, Motor 4.5 corrosion | Liquid circuit control plate leakage, 4.5 motor corrosion, borrow 580-096 and 924-591 each one
No | Visit | Complaint | RCHN | 干燥前温度超范围 1068 | 干燥前温度超范围 1068，借903-210，848-037各一个。 |  | SWO-178037 | Defect | HW | Process Module | PCB, Cup dryer | 1068 | Over-range temperature before drying 1068 | The pre-drying temperature exceeds the range of 1068, borrowing one each at 903-210, 848-037.
No | 13 | Y | RPOL | Blank | Regulacja lampy halogenowej |  | D24D8AF458503B45 | Adjustment | HW | Oxi module | Lamp unit |  | Blank | Halogen lamp Adjustment
No | Visit | Complaint | RCHN | 进样处废液漏液 | 进样处废液漏液，借905-847一个 |  | SWO-177976 | Defect | HW | Inlet module | Nipple kit |  | Waste fluid leak at the sample | Waste fluid leak at the sample, borrow 905-847 one
No | Visit | Complaint | RIND | Lactate Values Below Reference Range In QC Level 1 | Lactate Values for OC Level 1 Below Reference Range For Last 3 QC Test | Cleaned The Lactate Chamber & Lactate Electrode Tip, Run The 2 Point Calibration. After Passing 2 Point Calibration Again Run QC Level 1 ,In Result Lactate Values Are In Range . Machine Working Fine. | SWO-179545 | Cleaned | Fluidic pathway | El/Met Module | Measuring chamber | Lac | Lactate Values Below Reference Range In QC Level 1 | Lactate Values for OC Level 1 Below Reference Range For Last 3 QC Test | Cleaned The Lactate Chamber & Lactate Electrode Tip, Run The 2 Point Calibration. After Passing 2 Point Calibration Again Run QC Level 1 ,In Result Lactate Values Are In Range . Machine Working Fine.
No | Visit | Complaint | RDEU | pH Messkammer läuft voll | pH Messkammer läuft voll | Messkammer im pH $BG Modul gereinigt, Tausch des pH Dichtrings, mehrere Kal. durchgeführt, Proteinentfernung durchgeführt, System- und Funktionskontrolle in Ordnung | SWO-178079 | Cleaned | Fluidic pathway | pH/BG Module | Measuring chamber |  | pH measuring chamber runs full | pH measuring chamber runs full | Measuring chamber in the pH $BG module cleaned, replacement of the pH sealing ring, several cal. performed, protein removal performed, system and function control in order
No | Visit | Complaint | RDEU | 20.03.2017, Frau Langgut, Tel.: 09117580992378: häufige Kalibrationsfehler bei Laktat-häufig neu membraniert. 20.03.2017; 12:15 Uhr; Fr.Langgut; Tel.: 09117580992378; weiterhin Lac Kal Fehler | Lac Kal. und QC Fehler | Lac Elektrode gründlich mit Proteinentferner gereinigt, Membrane erneut getauscht - sie schien undicht zu sein, Messkammer gereinigt, Glasfläche der Elektrode gründlich gereinigt, Einlassnadel und Einlassdichtung aus Kundenbestand getauscht - Einlassnadel war stark verbogen (Knick), Lufteinlass gespült, mehrere Kal. durchgeführt, System- und Funktionskontrolle in Ordnung | SWO-178154 | Assisted Replacement | Electrode | Membrane | Lac |  | 20.03.2017, Mrs. Langgut, Tel.: 09117580992378: common calibration errors in lactate-frequently re-membraneized. 20.03.2017; 12:15 p.m.; Fr.Langgut; Tel.: 09117580992378; Continue Lac Kal Error | Lac Kal. und QC Fehler | Lac electrode thoroughly cleaned with protein remover, membrane replaced - it seemed to be leaking, measuring chamber cleaned, glass surface of the electrode thoroughly cleaned, inlet needle and inlet seal replaced from customer stock - inlet needle was strong bent (kink), air inlet rinsed, several cal. performed, system and function control in order
No | Visit | Complaint | RDEU | Ca+ QC Lebel 3 Probleme | Ca+ Membran defekt. | Messkammer gereinigt. Ref Elektrode membraniert, Ca+ Elektrode Membraniert. Gerät kalibriert und QC durchgeführt. | SWO-178795 | Cleaned | Fluidic pathway | El/Met Module | Measuring chamber |  | CA QC Lebel 3 Probleme | Ca+ membrane defective. | measuring chamber. Ref electrode membranes, Ca+ electrode membraneed. calibrated and QC performed.
No | 5 | Y | RAME | 2 PO2 and PCO2 fails Qc |  | Customer remembraned electrodes and still the same issue. | D243BF53A821FF63 | Cleaned | Fluidic pathway | pH/BG Module | Measuring chamber |  | 2 PO2 and PCO2 fails Qc |  | Customer remembraned electrodes and still the same issue.
No | 6 | Y | RAME | P Leakage atb inlet module |  | Replaced inlet nipples and inlet gasket. | D243C300A84DFF68 | Defect | HW | Inlet module | Inlet Gasket |  | P Leakage atb inlet module |  | Replaced inlet nipples and inlet gasket.
No | Visit | Complaint | RIND | Analyzer In Error Mode, Showing Error 3050 Corroptxml, Restart The Analyzer, Done Auto Refresh for Incubation Temperature, Analyzer took 20 mins to stabilize incubation temperature, problem resolve, Analyzer working ok. | Analysis Disable, Error 3050 Corroptxml. | after Checking found Analyzer , Showing Error 3050 Corroptxml, Restart The Analyzer, Done Auto Refresh for Incubation Temperature, Analyzer took 20 mins to stabilize incubation temperature, problem resolve, Analyzer working ok. | SWO-179534 | Defect | SW | Analyzer Control Module | Off/On | 3050 | Analyzer In Error Mode, Showing Error 3050 Corroptxml, Restart The Analyzer, Done Auto Refresh for Incubation Temperature, Analyzer took 20 mins to stabilize incubation temperature, problem resolve, Analyzer working ok. | Analysis Disable, Error 3050 Corroptxml. | after Checking found Analyzer , Showing Error 3050 Corroptxml, Restart The Analyzer, Done Auto Refresh for Incubation Temperature, Analyzer took 20 mins to stabilize incubation temperature, problem resolve, Analyzer working ok.
No | 6 | Y | RCZE | Blank |  |  | D24427E9A84A7BF8 | User error | User error | User error | User error |  | Blank
No | Visit | Complaint | RLTD | Not connected to time server . | Not connected to time server. | Time server IP address added to Internet time server setting and test connection ok. | SWO-178501 | Setup | SW | Time setting | Time setting |  | Not connected to time server . | Not connected to time server. | Time server IP address added to Internet time server setting and test connection ok.
No | Visit | Complaint | RSAF | Inlet faulty | Inlet bent gasket cracked | Replaced inlet probe and inlet gasket w. Holder. Fluid transport tests ok. Start up tests completed. PCO2 sensitivity out of range.(rising with every calibration-customer to monitor and replace if necessary) | SWO-180249 | Replaced | HW | Auto Inlet module | Probe bend |  | Inlet faulty | Inlet bent gasket cracked | Replaced inlet probe and inlet gasket w. Holder. Fluid transport tests ok. Start up tests completed. PCO2 sensitivity out of range. (rising with every calibration-customer to monitor and replace if necessary)
No | Visit | Complaint | RIND | Analyzers Touch Not working, After checking found that Touch Panal Get Defective, Replaced Touch Panel with New One, And Reinstalled SC and SP, Problem Resolve. | Touch Panel not Working. | After checking found that Touch Panal Get Defective, Replaced Touch Panel with New One, And Reinstalled SC and SP, Problem Resolve. | SWO-179536 | [5 empty] | Analyzers Touch Not working, After checking found that Touch Panal Get Defective, Replaced Touch Panel with New One, And Reinstalled SC and SP, Problem Resolve. | Touch Panel not Working. | After checking found that Touch Panal Get Defective, Replaced Touch Panel with New One, And Reinstalled SC and SP, Problem Resolve.
No | 6 | Y | RFRA | Blank | Recupération données discordant pO2 |  | D244ADB01E0D1D31 | [5 empty] | Blank | Recovering data discordant pO2
No | Visit | Complaint | RCHN | 开机死机 | 开机死机，借910-342一个 |  | SWO-178220 | Defect | HW | DMS_PC | CPU |  | Power on on | Power on, borrow 910-342 one
No | Visit | Complaint | RCHN | LAC,GLU定标失败 0210 | LAC,GLU定标失败 0210 | 更换测试卡正常。 | SWO-179175 | Assisted Replacement | SC | SC | Glu/Lac | 210 | LAC, GLU calibration failure 0210 | LAC, GLU calibration failure 0210 | The replacement test card is OK.
No | Visit | Complaint | RCHN | 测试卡参比电极错误 | 测试卡参比电极错误 | 更换测试卡 | SWO-178920 | Assisted Replacement | SC | SC | Reference | 1361 | Test card reference electrode error | Test card reference electrode error | Replace the test card
Yes | Visit | Complaint | RCHN | 报错1024 | 报错1024，借903-364一个 | 更换903-364一个正常。 by andrew | SWO-177867 | Defect | HW | Process Module | Optical Unit | 1024 | 1024 error | 1024, loan 903-364 one | Replace 903-364 one normal. by andrew
No | 4 | Y | RAME | V RECURRING ZERO CURRENT ERROR E408 FOR GLU/LAC AND CAL DRIFT OUT OF RANGE FOR LAC E376. |  | CONFIRMED ISSUE WHEN ONSITE. EI $MET MODULE WAS ALREADY REPLACED BY HJP ON FEB 21 SO ISSUE COULD BE CAUSED BY MINOR RINSE SOLUTION MIXED W CAL SOLUTION IN RIGHT FLUIDIC MODULE. | D244C6B71E580287 | Defect | HW | El/Met Module | Valve |  | V RECURRING ZERO CURRENT ERROR E408 FOR GLU/LAC AND CAL DRIFT OUT OF RANGE FOR LAC E376. |  | CONFIRMED ISSUE WHEN ONSITE. EI $MET MODULE WAS ALREADY REPLACED BY HJP ON FEB 21 SO ISSUE COULD BE CAUSED BY MINOR RINSE SOLUTION MIXED W CAL SOLUTION IN RIGHT FLUIDIC MODULE.
No | 6 | Y | RAME | F Customer reported calibrations failing oximetry. |  | Checked calibration log for errors. Unit failing with a :0251 Halogen Lamp intensity too high. | D244C5D01E3E8D37 | Adjustment | HW | Oxi module | Lamp unit | Halogen | F Customer reported calibrations failing oximetry. |  | Checked calibration log for errors. Unit failing with a :0251 Halogen Lamp intensity too high.
No | 6 | Y | RFRA | Blank | 1223/1264/1265 : PAS DE DETERMINATION Hct 1218 : ECHANTILLON INSUFFISANT |  | D244BF211E4DCFFA | Defect | HW | Aspiration Module | P4 | 1218 | Blank | 1223/1264/1265: NOT DETERMINATION Hct 1218: ECHANTILLON
No | Visit | Complaint | RBNL | Na een systeem reiniging is een crp calibratie afgebroken. Na uit- en aanzetten is een TNT calibratie wel gelukt. | Systeemcontrole gaat goed. Zie in activiteiten overzicht melding: undefined_ lookup_ string46.-1. Zie in failure.xml bestand: iets met pipette. Dit kan wijzen op een mechanisch defect. Aangezien volgende ingezette TNT calibratie wel goed gaat waarschijnlijk een onprettige samenloop van omstandigheden. Ben na een half uur weer opgeroepen CRP test staat weer in het rood, melding calibratie noodzakelijk. Zie in activiteiten overzicht dat om 16:05 een actieve calibratie aanpassing werd uitgeschakeld (0898). Er moet opnieuw gecalibreerd worden. | Vanmorgen binnengekomen nieuwe CRP kit (12521) met succes gecalibreerd. | SWO-178436 | Defect | Consumables | Test Kit | Cartridge, TnT | Broken | After a system cleaning, a CRP calibration has been aborted. After turning off and on, a TNT calibration did succeed. | System check is going well. See in Activity Overview notification: Undefined_ lookup_ string46.-1. See in failure. xml file: something with pipette. This may indicate a mechanical defect. Since the next bet TNT calibration is going well probably an unpleasant confluence of circumstances. Am after half an hour again called CRP Test state again in red, notification calibration necessary. See in activity Overview that at 16:05 an active calibration adjustment was turned off (0898). It has to be calibrated again. | This morning, new CRP Kit (12521) was successfully calibrated.
No | Visit | Complaint | RDEU | Akku hält nicht mehr lange | Gerät geht bei 20% Akkukapazität aus | Akku defekt, Ersatzteil bestellt | SWO-178664 | Revisit required | HW | Battery module | Battery |  | Battery doesn't last long anymore | Device runs out at 20% battery capacity | Battery defective, spare part ordered
No | Visit | Complaint | RDEU | 22.03.2017 09:40Uhr Frau Strobel 021318886152: Das Personal auf der Station verlangt einen Service-Kit der nicht vorhanden wäre. | Fehler 1178: Fehler Pumpenkalibrierung. | Fluid Transport-Serviceprogramme durchgeführt. Durch Pumpenkalibration Fehler beseitigt. Nach Kaltstart Fehler 1396, 1340 sowie Kalibrierfehler Glu & pCO2. Sensorkassettentausch empfohlen und nach erneuter gescheiterter Kalibration durchgeführt. Einlassdichtung und Anschlussdichtung erneuert und mit dem Kunden Vorgehensweise besprochen. Einlaufphase der Sensorkassette beobachtet. Funktion fehlerfrei. | SWO-178951 | Adjustment | HW | Service | Pump Calibration |  | 22.03.2017 09:40Hrs Ms Strobel 021318886152: The staff at the station requires a service kit that would not be available. | Error 1178: Error pump calibration. | Fluid Transport service programs. Eliminated by pump calibration errors. After cold start error 1396, 1340 as well as calibration error Glu & pCO2. sensor cartridge replacement and performed after another failed calibration. Inlet seal and connection seal renewed and discussed with the customer's procedure. inlet phase of the sensor cartridge. Function error-free.
No | Visit | Complaint | RDEU | 22.03.2017 12:00 Fr. Essel 0221890718954: Radiance-Anbindung gestört | Netzwerkeinstellungen verloren. | Fehlersuche Netzwerkeinstellungen vorgenommen. 3 Geräte im Haus stillgelegt. | SWO-178969 | Setup | His/lis | His/lis | Radiance |  | 22.03.2017 12:00 Fr. Essel 0221890718954: Radiance connection disrupted | Network settings lost. | Troubleshooting network settings. 3 appliances in the house shut down.
No | 8 | Y | RPNZ | Blank | Sensor cassette replacement due to internal reference electrode error |  | D248DB5B80580DFD | Assisted Replacement | SC | SC | Reference |  | Blank | Sensor cassette replacement due to internal reference electrode error
No | Visit | Complaint | RCHN | LAC定标失败 | LAC定标失败 | 更换测试卡正常。 | SWO-178933 | [5 empty] | LAC calibration failure | LAC calibration failure | The replacement test card is OK.
No | Visit | Complaint | RCHN | LAC定标失败 | LAC定标失败 | 更换测试卡 | SWO-179180 | [5 empty] | LAC calibration failure | LAC calibration failure | Replace the test card
No | Visit | Complaint | RCHN | GLU定标失败 | GLU定标失败 | 更换GLU电极膜正常。 | SWO-179173 | Assisted Replacement | Electrode | Membrane | Glu |  | GLU calibration failed | GLU calibration failed | Replace the GLU electrode film normally.
No | Visit | Complaint | RDEU | 22.03.2017, Herr Esche, Tel.: 089857915954: kap. Messungen nicht möglich, er hat die Einlassdichtung gewechselt, Proteinentfernung und LS Justage durchgeführt | Fehlersuche vor Ort | Bei eintreffen war die Nadel verbogen. Deswegen Nadel und Einlassdichtung getauscht. ausKundenbestand entnommen. Ventile und Ventilsitze gereinigt aus pH bG und Ele Met Modul. Pumpenschlscueh getauscht. Mehrere Kalibrierungen durchgeführt. tHb Kal durchgeführt. --> i.O. Fcuv.: 0,9933 Kal Verifizierung durchgeführt--> i.O. Kappiläre Testmessung mit QC Lösung durchgeführt. --> Lösung wird ohne Probleme angesaugt. QCs für alle Level durchgeführt. ABL wieder messbereit. | SWO-178870 | Defect | HW | Inlet module | Probe bend |  | 22.03.2017, Mr. Esche, Tel.: 089857915954: kap. Measurements not possible, he has changed the inlet seal, performed protein removal and LS adjustment | Troubleshooting on site | When it arrived, the needle was bent. Therefore, needle and inlet seal are exchanged. taken from customer stock. Valves and valve seats cleaned from pH bG and Ele Met module. Pump scueh. Multiple calibrations performed. tHb Kal. --> i.O. Fcuv.: 0.9933 Kal Verification performed--> i.O. Kappilular test measurement performed with QC solution. --> Solution is sucked in without any problems. QCs for all levels. ABL ready to be measured again.
No | 7 | Y | RAME | W HELP OPTION - ALPHABETICALLY INDEX FAILS TO RESPOND BY LINKING TOPICS TO RESPECTIVE SUBJECTS OF INTEREST. |  | CONFIRMED ISSUE WHEN ONSITE. | D2467113944B029E | Setup | His/lis | His/lis | His/lis |  | W HELP OPTION - ALPHABETICALLY INDEX FAILS TO RESPOND BY LINKING TOPICS TO RESPECTIVE SUBJECTS OF INTEREST. |  | CONFIRMED ISSUE WHEN ONSITE.
Yes | 7 | Y | RAME | K Customer needed option key added for pleura fluid. |  | Pleura fluid not installed. | D246042294343C69 | Re classified RMED | User request |  |  |  | K Customer needed option key added for pleura fluid. |  | Pleura fluid not installed.
Yes | 6 | Y | RAME | K Customer needed pleura key installed on analyzer. |  | Pleura key not installed on analyzer. | D2460956942F3C6D | Re classified RMED | User request |  |  |  | K Customer needed pleura key installed on analyzer. |  | Pleura key not installed on analyzer.
No | 2 | Y | RFRA | Blank |  |  | D2461B2694494788 | Defect | HW | Storage Module | Barcode reader, Storage wheel |  | Blank
No | 2 | Y | RFRA | Blank | Erreur 5669 ROTATEINLETWHEEL_24 |  | D2460D5E941F2206 | No fault found | No fault found | No fault found | No fault found | 5669 | Blank | Error 5669 ROTATEINLETWHEEL-24
No | Visit | Complaint | RDEU | Der Analysator meldet leckage | Im Flüssigkeitstransport sind Luftblasen zu sehen. Fehler erkannt, die Einlass-Sonde ist verbogen. | Einlass-Sonde wurde begradigt und das gerät im option Perform pressurtest mehrmals kalibriert | SWO-179145 | Replaced | HW | Inlet module | Probe bend |  | The analyzer reports leakage | Air bubbles can be seen in the fluid transport. Error detected, the inlet probe is bent. | Inlet probe has been straightened and this is calibrated several times in the option Perform pressurtest
No | Visit | Complaint | RDEU | 23.03.2017 09:40Uhr Martin Zerwagner 070712987226: Am Morgen wurde an einem Patienten 6g/dl tHb gemessen. Im Labor wurde am Abend 11g/dl tHb gemessen. | Gerät überprüft. Vergleichsmessungen mit ABL 8xx HTG 11,1 zu 11,1 g/dl identisch. Pat Akte Vorbefunde zeigen kein unnormales Verhalten . Vergleichsmessungen mit Zentrallabor 8,1 zu 8,5 ok. Präanalytik mit Stationsarzt besprochen. Evtl. Sediment Messung. |  | SWO-179423 | User error | Pre analytical | Pre analytical | Sample mix |  | 23.03.2017 09:40uhr Martin Zerwagner 070712987226: In the morning 6g/dl tHb was measured on a patient. In the laboratory, 11g/dl tHb was measured in the evening. | device checked. Comparison measurements identical to ABL 8xx HTG 11.1 to 11.1 g/dl. Pat file preliminary findings show no abnormal behavior. Comparative measurements with central laboratory 8.1 to 8.5 ok. Pre-analysis discussed with ward doctor. Possibly sediment measurement.
No | 5 | Y | RAME | 5 NA++ |  | Sodium correlation high value value compared to ABBOT architect | D2473B050A59E780 | No fault found | Perfomance | Perfomance | Na |  | 5 NA + + |  | Sodium correlation high value value compared to ABBOT architect
No | Visit | Complaint | RDEU | 20.03.2017; Dr. Dieckmann; Tel.: 0602132433200; keine Radiance-Verbindung 21.03.2017 Pfl.Thomas 06021324723: keine Radiance-Verbindung | Keine Radianceverbindung Überprüfung Ethernet | IT Connection Ethernet überprüft. Keine feste IP und Netzwerkdaten gespeichert, Stand DHCP. Eingabe der Netzwerkdaten. Speicherung auf Windowsebene. Verbindung hergestellt. Neustart. Keine Verbindung nach Neustart. Fehlermeldung Ethernetadapter. Ethernet Adapter speichert Daten nicht ab. Überprüfung und ändern der Etherneteinstellungen. Erneute Eingabe der Berbindungsdaten. Speicherung auf Windowsebene. Nicht erfolgreich nach Neustart. Deinstallation der Netzwerkkarte. Neuinstallation nach Neustart. Eingabe der Verbindungsdaten und Speicherung. Netzwerkverbindung nach Neustart nicht hergestellt. Nochmals manuell hergestellt. Verbindung wird nach einem Neustart wieder verloren gehen! CPU inkl. Netzwerkkarte wird bestellt und nach Erhalt eigebaut, keine Speicherung der Anbindungsdaten möglich. | SWO-179058 | Revisit required | HW | DMS_PC | CPU | His/Lis | 20.03.2017; Dr. Dieckmann; Tel.: 0602132433200; No radiance connection 21.03.2017 Pfl.Thomas 06021324723: no radiance connection | No radiance connection Checking Ethernet | IT Connection Ethernet. No fixed IP and network data stored, status DHCP. Entering the network data. Windows-level storage. Connection established. Restart. No connection after reboot. Error message Ethernet adapter. Ethernet Adapter does not store data. Check and change Ethernet settings. Re-enter the binding data. Windows-level storage. Not successful after reboot. Uninstall the network card. Reinstall after reboot. Enter the connection data and save it. Network connection not established after restart. Manufactured manually again. Connection will be lost again after a restart! CPU incl. network card is ordered and built after receipt, no storage of the connection data possible.
No | 3 | Y | RFRA | Blank | problème de rendu du paramètre bilirubine |  | D247954C0A3B5F7C | Defect | HW | Oxi module | Hemolyzer |  | Blank | rendering problem of the bilirubin setting
No | Visit | Complaint | RLTD | Following user training the analyser started asking for the inlet gasket to be replaced. | I had come in to check the inlet module on ITU1 when Marriann asked me to have a look at this one as it was asking for the inlet gasket to be replaced after she had given some training. | First of all the solution pack was removed and replaced but this didn't fix the problem. The inlet module was removed and a blood clot found at the end of the probe, this was flushed out and the inlet gasket flushed but this still didn't cure the problem. A new inlet gasket was fitted and following this the analyser did pass its pressure test. | SWO-179301 | Defect | HW | Auto Inlet module | Inlet Gasket |  | Following user training the analyser started asking for the inlet gasket to be replaced. | I had come in to check the inlet module on ITU1 when Marriann asked me to have a look at this one as it was asking for the inlet gasket to be replaced after she had given some training. | First of all the solution pack was removed and replaced but this didn't fix the problem. The inlet module was removed and a blood clot found at the end of the probe, this was flushed out and the inlet gasket flushed but this still didn't cure the problem. A new inlet gasket was fitted and following this the analyser did pass its pressure test.
No | 8 | Y | RFRA | Blank | blocage de la baie +accroche ponctuelle du site d'introduction |  | D2474FC10A185F77 | Defect | HW | Sample mixer | Sample mixer | LEH Mail | Blank | blocking the bay - spot hook of the introductory site
No | Visit | Complaint | RCHN | 接口板腐蚀 | 接口板腐蚀,借902-972一个 | 拆备用机上的902-972一个更换后正常。 by andrew | SWO-178582 | [5 empty] | Interface plate corrosion | Interface plate corrosion, borrow 902-972 one | Dismantling the 902-972 on the spare machine is normal after replacement. by andrew
Yes | Visit | Complaint | RCHN | LAC定标失败 | LAC定标失败 | 更换测试卡正常。 | SWO-178932 | [5 empty] | LAC calibration failure | LAC calibration failure | The replacement test card is OK.
No | Visit | Complaint | RCHN | 报1035错误 | 报1035错误，借903-364一个 | 更换903-364一个正常。by andrew | SWO-178293 | Defect | HW | Process Module | Optical Unit | 1035 | Reported 1035 error | Reported 1035 error, borrow 903-364 one | Replace 903-364 one normal. by andrew
No | Visit | Complaint | RIND | Machine Not Initialising | Due To Electrical Failure Machine Shut Down Suddenly , After Switching On Machine Not Initialising | After Sudden Power Failure , Machine Started Again & Gives The Leak Detected Error. Shut Down The Machine By Standard Method, Checked All The Connections, Again Started The Machine, Machine Initialised & Passed 2 Point Calibration. Run The Sample , Machine Working Ok | SWO-179547 | User error | Mains Power Failure | Mains Power Failure | Mains Power Failure |  | Machine Not Initialising | Due To Electrical Failure Machine Shut Down Suddenly , After Switching On Machine Not Initialising | After Sudden Power Failure , Machine Started Again & Gives The Leak Detected Error. Shut Down The Machine By Standard Method, Checked All The Connections, Again Started The Machine, Machine Initialised & Passed 2 Point Calibration. Run The Sample , Machine Working Ok
Yes | 4 | Y | RPAC | Blank | Reference parameters drifting |  | D247503B0A1FC50E | Re classified RMED | Phone support |  |  |  | Blank | Reference parameters drifting
No | Visit | Complaint | RDEU | 22.03.17 11:45 Frau Possen Tel. 02151324078 , Kalibrationsfehler bei PH und den Elektrolyten 22.03.2017 12:15 Pfl. Oliver 02151324079: Kal-Fehler bei pH und Elektrolyten 22.03.2017 15:28 Frau Passens 021513245070, Frau Susan 02151324078: Das Gerät funktionierte für kurze Zeit. Aktuell zeigt das Gerät wieder den bekannten Fehler | Kal Fehler bei pH und allen Elektrolyten. Status und Empfindlichkeit außerhalb der Kal Grenzen. | Messkammer gereinigt und Ref Kontakte nachgebogen - keine Verbesserung, Ref Elektrode erneuert - keine Verbesserung, Ref Elektrode zurück getauscht, Ref Connector erneuert - Empfindlichkeit und Status aller Elektrolyte und pH OK, mehrfach Kalibriert und alle 3 QC Lvl durchgeführt. | SWO-179121 | Assisted Replacement | Electrode | Electrode | Reference |  | 22.03.17 11:45 Mrs Possen Tel. 02151324078 , Calibration error in PH and electrolytes 22.03.2017 12:15 Pfl. Oliver 02151324079: Kal error in pH and electrolytes 22.03.2017 15:28 Ms. Passens 021513245070, Ms. Susan 02151324078: The device worked for a short time. Currently, the device again shows the known error | Kal faults with pH and all electrolytes. Status and sensitivity outside the Cal limits. | Measuring chamber cleaned and ref contacts bent - no improvement, ref electrode renewed - no improvement, ref electrode replaced, ref connector renewed - sensitivity and status of all electrolytes and pH OK, multiple calibrated and all 3 QC Lvl Performed.
No | Visit | Complaint | RIND | Lactate Sensitivity Low | Lactate Sensitivity Low & Lactate Membrane Needs To Be Change | Lactate Membrane Need To Be Replace, Replaced The Lactate Membrane, Machine Passed 2 Point Calibration. Run The Sample, Machine Working Fine | SWO-179551 | Assisted Replacement | Electrode | Membrane | Lac |  | Lactate Sensitivity Low | Lactate Sensitivity Low & Lactate Membrane Needs To Be Change | Lactate Membrane Need To Be Replace, Replaced The Lactate Membrane, Machine Passed 2 Point Calibration. Run The Sample, Machine Working Fine
No | 10 | Y | RAME | H Inhomogenous sample errors on random QC samples. |  | Verified customer's complaint. | D24761760A0B6D7E | Adjustment | HW | Inlet module | Probe cal. |  | H Inhomogenous sample errors on random QC samples. |  | Verified customer's complaint.
No | 7 | Y | RPNZ | Blank | AutoCheck Replacement |  | D24EE66FCE632D60 | Revisit required | HW | Auto Check Module | Auto Check Module |  | Blank | AutoCheck Replacement
No | 3 | Y | RAME | K CUSTOMER STATED THEY HAD NOT RECEIVED LINEARITY REPORT FOLLOWING INSTALLATION |  | UPON ARRIVED DISCOVED THE REPORT FOR BOTH ANALYZERS HAD BEEN RECEIVED AFTER INSTALLATION. MISIDENTIFIED IN EMAIL. CUSTOMER WAS ABLE TO PRINT AND REVIEW THE REPORTS. ALSO STATED THEY NEEDED HELP WITH CONNECTION FOR HIS $LIS SETTINGS | D247D0EC0A1DBA7C | Setup | His/lis | His/lis | His/lis |  | K CUSTOMER STATED THEY HAD NOT RECEIVED LINEARITY REPORT FOLLOWING INSTALLATION |  | UPON ARRIVED DISCOVED THE REPORT FOR BOTH ANALYZERS HAD BEEN RECEIVED AFTER INSTALLATION. MISIDENTIFIED IN EMAIL. CUSTOMER WAS ABLE TO PRINT AND REVIEW THE REPORTS. ALSO STATED THEY NEEDED HELP WITH CONNECTION FOR HIS $LIS SETTINGS
No | Visit | Complaint | RDEU | Akku hält nicht mehr lange | Batterie Option defekt | Austausch der Batterie Option, alter Akku 911-311P0072N001 | SWO-179382 | Defect | HW | Battery module | Battery |  | Battery doesn't last long anymore | Battery option defective | Battery replacement option, old battery 911-311P0072N001
No | 9 | Y | RAME | G Numerous and various issues with pO2 and pCO2. | Changed the complaint visit tab from N to Y and changed the type of service from routine to complaint. | Four sensor cassettes had to be installed to get one that would work. | D24888FA803575D2 | Defect | HW | Basic Unit | PCB, Sensor interface |  | G Numerous and various issues with pO2 and pCO2. | Changed the complaint visit tab from N to Y and changed the type of service from routine to complaint. | Four sensor cassettes had to be installed to get one that would work.
No | Visit | Complaint | RDEU | 23.03.17 16:35 Frau Freund Tel. 0221890715069 , Ca . jede 3 Messung wird abgebrochen mit Fehler 1190 | Fehler 1190 sporadisch aber wiederkehrend | Einlass Positionierung im Servicemenü überprüft, OK. INLET POSITION DETECTOR PCB 903-337 ERSETZT (aus Demogerät 092R0104N005), Neustart, Test der Funktionen im Servicemenü, OK, Endkontrolle: QC Messung OK | SWO-179287 | Defect | HW | Auto Inlet module | PCB, Inlet Position |  | 23.03.17 16:35 Mrs. Friend Tel. 0221890715069 , ca . every 3 measurement is aborted with error 1190 | Error 1190 sporadically but recurrent | Inlet positioning checked in the service menu, OK. INLET POSITION DETECTOR PCB 903-337 REPLACEMENT (from demo device 092R0104N005), restart, test of functions in the service menu, OK, final control: QC measurement OK
No | Visit | Complaint | RIND | QC Results out of Range, Cleaned SC80, Cleaned Connector perform Zeroing, Reinstalled SC, Run QC Problem Remains,Need to replace Sensor Board, TO Placed for Spare. | QC Failure. | QC Results out of Range, Cleaned SC80, Cleaned Connector perform Zeroing, Reinstalled SC, Run QC Problem Remains,Need to replace Sensor Board, TO Placed for Spare. | SWO-179539 | [5 empty] | QC Results out of Range, Cleaned SC80, Cleaned Connector perform Zeroing, Reinstalled SC, Run QC Problem Remains,Need to replace Sensor Board, TO Placed for Spare. | QC Failure. | QC Results out of Range, Cleaned SC80, Cleaned Connector perform Zeroing, Reinstalled SC, Run QC Problem Remains,Need to replace Sensor Board, TO Placed for Spare.
No | Visit | Complaint | RCHN | 触摸屏无反应 | 触摸屏无反应，借903-239一个。 | 更换903-239修复 by andrew | SWO-178972 | [5 empty] | Touch screen does not respond | Touch screen does not respond, borrow 903-239 one. | Replace 903-239 Repair by Andrew
No | Visit | Complaint | RCHN | 压力测试真空错误 1063 | 压力测试真空错误 1063 | 更换试剂包正常。 | SWO-179227 | Assisted Replacement | SP | SP | FT | 1063 | Pressure test vacuum error 1063 | Pressure test vacuum error 1063 | The replacement kit is OK.
No | Visit | Complaint | RCHN | 进样口模块有故障 | 进样口模块有故障,借903-139一个 | 更换903-139一个正常。 by andrew | SWO-178994 | Defect | HW | Inlet module | Inlet module |  | There is a problem with the inlet module | The sample module is faulty, with a 903-139 | Replace the 903-139 one normal. by andrew
No | 6 | Y | RFRA | Blank | CV important sur les CQ |  | D24885DF803C5111 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | Blank | Important CV on CQ
No | 6 | Y | RPNZ | Blank | El/Met pump asymmetry errors. |  | D2547362A634E69C | Defect | HW | El/Met Module | Pump tubes |  | Blank | El/Met pump asymmetry errors.
No | Visit | Complaint | RCHN | 流量选择器出错 1177 | 流量选择器出错 1177，借903-185，580-417各一个。 | 更换903-185，580-417修复 by andrew | SWO-179242 | Defect | HW | Wet section | PCB, Flow Selector | 1177 | Traffic Selector Error 1177 | The flow selector made an error of 1177, borrowing one each at 903-185, 580-417. | Replace 903-185,580-417 Fix by Andrew
No | Visit | Complaint | RCHN | 测试卡参比电极错误 | 测试卡参比电极错误 | 更换测试卡 | SWO-179549 | Assisted Replacement | SC | SC | Reference | 1361 | Test card reference electrode error | Test card reference electrode error | Replace the test card
No | Visit | Complaint | RCHN | 不通电 | 不通电，借903-175，903-198各一个 | 更换903-175和903-198后，仍旧未修复，追加借件910-330一个 by andrew | SWO-179570 | Revisit required | HW | Analyzer Control Module | PCB, Main Board |  | No power | No power, borrow 903-175, 903-198 each | After replacing 903-175 and 903-198, still unrepaired, added borrowing 910-330 a by andrew
No | Visit | Complaint | RIND | AQT non functional. Sample stuck in side. | AQT ERROR Massage. | Cup Stuck inside incubator wheel. Cup Removed from incubator wheel. Analyzer serviced and calibrated properly. Handed over in a fine working condition. one test run to check. working fine. | SWO-179569 | Defect | Consumables | SP | Cup Stack |  | AQT non functional. Sample stuck in side. | AQT ERROR Massage. | Cup Stuck inside incubator wheel. Cup Removed from incubator wheel. Analyzer serviced and calibrated properly. Handed over in a fine working condition. one test run to check. working fine.
No | Visit | Complaint | RCHN | 机器忽然断电，开机后无法进入系统 | 机器忽然断电，开机后无法进入系统,借933-303一个 |  | SWO-179603 | [5 empty] | The machine suddenly lost power and was unable to enter the system after power ing-up | The machine suddenly power down, after the power can not enter the system, borrow 933-303 a
No | Visit | Complaint | RCHN | 进样门的条码扫描器无法扫描 | 进样门的条码扫描器无法扫描,借903-186一个 |  | SWO-179614 | Defect | HW | Aspiration Module | Barcode reader, Inlet |  | The barcode scanner at the inlet door cannot be scanned | The barcode scanner for the inlet door cannot be scanned, and 903-186 is a
No | Visit | Complaint | RCHN | 不通电 | 更换903-175和903-198后，仍旧未修复，追加借件910-343(原货号是910-330)一个 |  | SWO-179601 | Defect | HW | Analyzer Control Module | PCB, Main Board |  | No power | After replacing 903-175 and 903-198, still unrepaired, additional borrowing 910-343 (original number 910-330) one
Yes | Visit | Complaint | RLTD | Ranzir on 01865 221596 |  |  | SWO-180028 | Re classified RMED | Phone call, FSE dispatch |  |  |  | Ranzir on 01865 221596
No | Visit | Complaint | RDEU | 24.03.2017, Pfl. David, Tel.: 070712985091: Probenaspiration gescheitert bei allen Proben, Pfl. David hat noch keine Probe messen können. Der ABL steht seit heute Vorort. 24.03.2017 16.11Uhr Pfl David 0070712985091: Das Gerät empfiehlt den Tech.Service zu rufen. | Gerät überprüft verlangt immer Lös Pack oder Spülen. Vakuumfehler erscheint immer wieder | Gerät überprüft. Lös Pack scheint ok. Hemolysator und Tubings überprüft, ok. Einlassdichtung ersetzt. Ohne Erfolg ,erneute Einlassdichtung, dann Sol Pack Check mehrfach ok. Pressuretest jetzt auch ok. Initialisierung abgewartet. Nach mehrfacher Kal und Kalibration und QC wieder ok. Blutmessungen durchgeführt. | SWO-179941 | Defect | HW | Inlet module | Inlet Gasket |  | 24.03.2017, Pfl. David, Tel.: 070712985091: Sample aspiration failed in all samples, Pfl. David has not yet been able to measure a sample. The ABL has been on the ground since today. 24.03.2017 16.11Pm Pfl David 0070712985091: The device recommends calling the Tech.Service. | Device checked always requires loess pack or rinsing. Vacuum error appears again and again | device checked. Lös Pack seems ok. Hemolysator and tubings checked, ok. replaced inlet seal. Without success, new inlet seal, then Sol Pack Check several times ok. Pressuretest now ok. initialization awaited. After multiple cal and calibration and QC ok again. blood measurements.
No | 2 | Y | RAME | Z Will not print. Error 1297 printer is offline. |  | Checked printer connections. | D24C92D1E22F81BA | Defect | HW | Printer Module | Printer Module |  | Z Will not print. Error 1297 printer is offline. |  | Checked printer connections.
No | Visit | Complaint | RSAF | Trouble shooting done telephonically but needed to have consumables replaced and customer does not have spare stock on hand. | Service mode entered and done tests, all passed. | Firstly replaced solution pack but only corrected glu error, Replaced sensor cassette and all ok. | SWO-179988 | Assisted Replacement | SC | SC | Glu |  | Trouble shooting done telephonically but needed to have consumables replaced and customer does not have spare stock on hand. | Service mode entered and done tests, all passed. | Firstly replaced solution pack but only corrected glu error, Replaced sensor cassette and all ok.
No | Visit | Complaint | RBNL | Er zijn wat problemen met het mengen van monsters mbv de ingebouwde mixer. Start soms niet bij inleggen van een spuit. | Er is al een mixer-unit gewisseld, probleem nog niet geheel verdwenen. | Preventief voedings-unit vervangen. | SWO-179815 | Defect | HW | Power Supply | Power Supply |  | There are some problems with mixing samples MBV the built-in mixer. Sometimes do not Start when a syringe is put in. | A mixer unit has already been changed, the problem has not completely disappeared. | Preventive power unit replacement.
Yes | Visit | Complaint | RBNL | Problemen met mixen. Tevens gaat de analyzer direct uit als het netsnoer wordt uitgetrokken. | Tijdens demonstraties start de mixer soms niet als een spuit wordt opgelegd. Analyzer valt direct uit als het netsnoer losgetrokken wordt. | Voeding vervangen en een nieuwe accu geplaatst. | SWO-179864 | Re classified RMED | HW | Power Supply | Power Supply | Agreed WR 2017-03-30 | Problems with mixing. The analyzer also goes out immediately when the power cord is pulled out. | During demonstrations, the mixer sometimes does not start when a syringe is imposed. Analyzer is immediately removed when the power cord is pulled out. | Replace the power supply and put a new battery.
Yes | Visit | Complaint | RBNL | Bij tijd en wijle zijn er problemen dat de mixer niet start als er een spuit wordt ingelegd. De verlichting gaat daarbij ook flikkeren. | Mixing problematiek nog steeds niet naar tevredenheid klant. Bij opstarten vloeistof transport problemen, bij ingang sensorcassette verstopt. | Preventief powersupply vervangen. Na weken door kunnnn spuiten met demiwater. | SWO-179831 | Re classified RMED | HW | Power Supply | Power Supply | Agreed WR 2017-03-30 | At times, there are problems that the mixer does not start when a syringe is inlaid. The lighting is also flickering. | Mixing problems still fail to customer satisfaction. When booting fluid transport problems, at input sensor cassette clogged. | Preventive powersupply replacement. After weeks by kunnnn spraying with demi water.
No | Visit | Complaint | RCHN | 测试卡温度 | 测试卡温度不稳定，借备件902-972一个 |  | SWO-179743 | [5 empty] | Test card temperature | Test card temperature is not stable, borrowed spare parts 902-972 one
No | 15 | Y | RAME | P the SP leaking |  | The customer said the SP was leaking | D24C8325E2098B9E | [5 empty] | P the SP leaking |  | The customer said the SP was leaking
No | Visit | Complaint | RIND | O2 Values out of range, QC Failure, Cleaned Connector and Interior Dust and Reinstalled. Problem Remains, Replaced Sensor Cassette with New one. | QC Failure, O2 Values on higher Side. | O2 Values out of range, QC Failure, Cleaned Connector and Interior Dust and Reinstalled. Problem Remains, Replaced Sensor Cassette with New one. | SWO-181286 | [5 empty] | O2 Values out of range, QC Failure, Cleaned Connector and Interior Dust and Reinstalled. Problem Remains, Replaced Sensor Cassette with New one. | QC Failure, O2 Values on higher Side. | O2 Values out of range, QC Failure, Cleaned Connector and Interior Dust and Reinstalled. Problem Remains, Replaced Sensor Cassette with New one.
No | Visit | Complaint | RCHN | 温度出错 | 温度出错 | 更换测试卡后修复 | SWO-180000 | [5 empty] | Temperature error | Temperature error | Repair after replacing test card
No | 7 | Y | RAME | F OXI would not cal |  | the OXI wauld not cal and was turned off | D24C58DEE201AA21 | [5 empty] | F OXI would not cal |  | the OXI wauld not cal and was turned off
No | 9 | Y | RAME | P LEAKAGE BY WASTE PORT OF FRONT PLATE, OVERFLOWED TO COUNTER TOP. |  | CONFIRMED ISSUE WHEN ONSITE. CUSTOMER HAS ALREADY TRIED DIFFERENT SENSOR CASSETTE BUT LEAK CONTINUED. | D24C63D6E2141C50 | [5 empty] | P LEAKAGE BY WASTE PORT OF FRONT PLATE, OVERFLOWED TO COUNTER TOP. |  | CONFIRMED ISSUE WHEN ONSITE. CUSTOMER HAS ALREADY TRIED DIFFERENT SENSOR CASSETTE BUT LEAK CONTINUED.
No | 5 | Y | RESP | Han cambiado el SP y les sigue pidiendo cambio No se pueden bombear soluciones |  | Sustitución de SP cogido de otro servicio Sustitución de Sonda, gasket y junta de conexión Inicio completo | D24CC909E24A25A3 | Defect | HW | Inlet module | Inlet Gasket |  | They have changed the SP and keep asking them for change You can't pump solutions |  | Replacement of SP caught from other service Probe replacement, gasket and connection joint Full start
No | Visit | Complaint | RCHN | SO2,THB值过高 | SO2,THB值过高，借902-980和903-197一个 |  | SWO-179720 | Defect | HW | Oxi module | Hemolyzer |  | SO2, THB value is too high | SO2, THB value is too high, borrow 902-980 and 903-197 one
No | Visit | Complaint | RCHN | TNI定标中止 | TNI定标中止，借备件903-364一个 |  | SWO-179747 | Defect | HW | Process Module | Optical Unit | 1024 | TNI calibration abort | TNI calibration suspended, borrowed spare parts 903-364 one
No | Visit | Complaint | RCHN | 开机无法进入系统，蓝屏 | 开机无法进入系统，蓝屏 | 自行更换备用件910-274后修复 by jesson | SWO-179773 | Defect | HW | Power Supply | Power Supply |  | Power on can not enter the system, blue screen | Power on can not enter the system, blue screen | Repair by jesson after replacing spare parts 910-274
No | Visit | Complaint | RIND | Lactate Value Low in QC Level 1 | Lactate Value Low In QC Level 1 | Cleaned the Lactate Electrode Tip & Measuring Chamber , Machine Passed 2 Point Calibration & QC Level 1. If Problem Persist Again Then Lactate Membrane Need To Be Change | SWO-181279 | Cleaned | Fluidic pathway | El/Met Module | Measuring chamber | Lac | Lactate Value Low in QC Level 1 | Lactate Value Low In QC Level 1 | Cleaned the Lactate Electrode Tip & Measuring Chamber , Machine Passed 2 Point Calibration & QC Level 1. If Problem Persist Again Then Lactate Membrane Need To Be Change
No | 14 | Y | RAME | 0 tHb out of limits on verification material. |  | S7940 and S7950 levels of the verification material VK-R7 lot #459 is low out of limits. | D24C8F4BE22230E3 | Adjustment | HW | Oxi module | Oxi Complete |  | 0 tHb out of limits on verification material. |  | S7940 and S7950 levels of the verification material VK-R7 lot #459 is low out of limits.
No | 10 | Y | RAME | 0 tHb out of limits on verification material. |  | tHb is borderline low on S7940 and S7950 levels. | D24C90DFE26130E4 | Adjustment | HW | Oxi module | Oxi Complete |  | 0 tHb out of limits on verification material. |  | tHb is borderline low on S7940 and S7950 levels.
No | Visit | Complaint | RDEU | Patlayout Auswahlliste leer - keine Anpassung möglich | Konfiguration kann icht angepasst werden - Layout leer, Referenzwerte werden teils, teils nicht ans Zielsystem übertragen | Fehlersuche, Layout+Messmodi und Referentwerte Standart wiederhergestellt, Layout und Refwerte neu eingegeben, Messmodi angepasst, mehrere Test durchgeführt, Werte verglichen - i.O., Konfiguration gespeichert | SWO-180037 | Setup | SW | Patient layout | Patient layout |  | Patlayout selection list empty - no customization possible | Configuration can be adjusted - layout empty, reference values are partly, partly not transferred to the target system | Troubleshooting, Layout+Measurement Modes and Referent values Standset restored, layout and ref values re-entered, measurement modes adjusted, multiple tests performed, values compared - i.O., configuration saved
No | 2 | Y | RAME | 3 , Customer having trouble getting p02 and electrolytes to pass calibrations |  | Had a zero current calibration error. Customer remembraned and cleaned chambers | D24C60AEE2481163 | Defect | Electrode | Electrode | Sealing ring | Sealing ring, pO2 | 3 , Customer having trouble getting p02 and electrolytes to pass calibrations |  | Had a zero current calibration error. Customer remembraned and cleaned chambers
No | 8 | Y | RAME | Z Unit stopped sending data to Sunquest after performing an update type after removing Oxi module. |  | Unit is connected serially via a lantronix box. Not connected to Radiance. Unit operating normally otherwise. | D24C5C9CE20A30E0 | Setup | His/lis | His/lis | His/lis |  | Z Unit stopped sending data to Sunquest after performing an update type after removing Oxi module. |  | Unit is connected serially via a lantronix box. Not connected to Radiance. Unit operating normally otherwise.
No | Visit | Complaint | RDEU | 24.03.2017 Frau Hessberger 0602132434553 meldet, dasss sie seit einigen Tagen fehler bei den Parametern K Na Ca haben. | Kalibrationsfehler bei Na&Ca nach Membranierung | Durchsicht de Kalibrationen. Kalibrationen seit Membranierung der Referenz fehlerhaft. Messkammer gereinigt und Ref neu membraniert. Kalibrationen mehrfach durchgeführt. QC in allen Bereichen nachgeholt, ohne Beanstandung | SWO-179851 | Assisted Replacement | Electrode | Membrane | Reference |  | 24.03.2017 Ms. Hessberger 0602132434553 reports that they have had errors with the parameters K Na Ca for several days. | Calibration error at Na&Ca after membraneization | Review of the calibrations. calibrations since membraneization of the reference faulty. measuring chamber cleaned and ref re-membraneed. calibrations performed several times. QC caught up in all areas, without complaint
No | 4 | Y | RAME | F Customer called with oxi calibration errors. |  | Investigation found error 316 (neon number low) 643 (DMS not connected) and 238 OXI hardware errors. | D24C6D3FE20F7BEE | Defect | HW | Oxi module | PCB, Oxi module |  | F Customer called with Oxi calibration errors. |  | Investigation found error 316 (neon number low) 643 (DMS not connected) and 238 OXI hardware errors.
Yes | 7 | Y | RFRA | Blank | changement de lot Myo regularisation stock |  | D24C71F8E2155DA5 | Re classified RMED | Logistical upSWO |  |  |  | Blank | Myo lot change regularization stock
No | Visit | Complaint | RLTD | amanda 0131 242 1330 | Analyser fluid path blocked completely at hemolyser causing intervention required | Flushed analyser vigourously and carried out several clot removed procedures, eventually got large clot of blood to release through inlet gasket. | SWO-180673 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway | Big CLOT :) | amanda 0131 242 1330 | Analyser fluid path blocked completely at hemolyser causing intervention required | Flushed analyser vigourously and carried out several clot removed procedures, eventually got large clot of blood to release through inlet gasket.
No | Visit | Complaint | RLTD | Anne-Marie Curtis 0131 2427908 | Analyser in troubleshooting needed state, asking for inlet gasket to be replaced Failing pressure test in service program Found leakage from valve tubing | Replaced inlet probe and gasket, still same Replaced tubing for valve Pressure test ok | SWO-180380 | Defect | HW | Oxi module | Pinch valve tube |  | Anne-Marie Curtis 0131 2427908 | Analyser in troubleshooting needed state, asking for inlet gasket to be replaced Failing pressure test in service program Found leakage from valve tubing | Replaced inlet probe and gasket, still same Replaced tubing for valve Pressure test ok
No | 9 | Y | RAME | Y Stuck in troubleshooting loop and the customer has replaced SC, SP inlet gasket, and inlet connector. The analyzer hasn't been used for patiant testing yet. |  | Powered the analyzer on and installed a new SC and SP and was able to re-create the troubleshooting loop issue. | D24D9FA55855B84B | Defect | HW | Wet section | Connector Gasket |  | Y Stuck in troubleshooting loop and the customer has replaced SC, SP inlet gasket, and inlet connector. The analyzer hasn't been used for patiant testing yet. |  | Powered the analyzer on and installed a new SC and SP and was able to re-create the troubleshooting loop issue.
No | Visit | Complaint | RDEU | 20.03.2017; Dr. Dieckmann; Tel.: 0602132433200; keine Radiance-Verbindung 21.03.2017 Pfl.Thomas 06021324723: keine Radiance-Verbindung | Austausch CPU | Datensicherung durchgeführt. Analysator geöffnet. Computerunit ausgetauscht, Installation der Kundenspezifischen Speicherkart. Neustart und Initialisierungsphase. Eintragen der Netzwerkverbindungsdaten. Speicherung auf Windowsebene. Fehlermeldung des Ethernetadapters mit Austausch beseitigt. Neustart und Initialisierungsphase. Anbindungsdaten wurden nicht gespeichert. Erneutes herstellen der Verbindung und Speicherung. Anbindung wird nach Neustart wieder verloren gehen. (Notstromtest am 01.04.2017) Austausch der CF Karte notwendig. Dr. Pietsch wünscht keine weitere Blockierung des Analysators! Blanko CF Karte aus neuer CPU wird vorkonfiguriert und nach Terminabsprache mit Dr.Pietsch eingebaut. | SWO-180178 | Defect | HW | DMS_PC | CPU | CF Card ? | 20.03.2017; Dr. Dieckmann; Tel.: 0602132433200; No radiance connection 21.03.2017 Pfl.Thomas 06021324723: no radiance connection | Replacement CPU | data backup. Analyzer open. Computer unit, installation of customer-specific memory cart. Restart and initialization phase. Enter the network connection data. Windows-level storage. Eliminates the error message of the Ethernet adapter with replacement. Restart and initialization phase. Connection data was not saved. Reconnect and save. Connection will be lost again after restart. (Emergency power test on 01.04.2017) Replacement of the CF card is necessary. Dr. Pietsch does not want any further blocking of the analyzer! Blank CF card made of new CPU is preconfigured and installed upon appointment with Dr.Pietsch.
No | 4 | Y | RAME | P Intervention required. |  | Customer came in this morning and analyzer was in intervention screen requesting a flush. Error code 1350. Seems to start after midnight after clot detect. Happens every day. | D24DC579583389FF | Defect | HW | Basic Unit | PM kit |  | P Intervention required. |  | Customer came in this morning and analyzer was in intervention screen requesting a flush. Error code 1350. Seems to start after midnight after clot detect. Happens every day.
Yes | 9 | Y | RPAC | Blank | multiple consumable failures |  | D24E017058317DE2 |  | PER |  |  |  | Blank | multiple consumable failures
No | 3 | Y | RFRA | Blank | ERR QC VALEURS CO-OX APRS CHANGEMENT DE LA CASSETTE |  | D24DC9B158167EDF | Defect | HW | Oxi module | Hemolyzer |  | Blank | ERR QC VALEURS CO-OX APRS CHANGE OF THE CASSETTE
No | Visit | Complaint | RCHN | 机器忽然断电，开机后无法进入系统 | 机器忽然断电，开机后无法进入系统，借备件903-381,933-306各一个 |  | SWO-180093 | [5 empty] | The machine suddenly lost power and was unable to enter the system after power ing-up | The machine suddenly powered off, can not enter the system after power-up, borrowed spare parts 903-381,933-306 each one
No | Visit | Complaint | RCHN | LAC定标失败 | LAC定标失败 | 更换测试卡后修复 | SWO-180161 | [5 empty] | LAC calibration failure | LAC calibration failure | Repair after replacing test card
No | Visit | Complaint | RCHN | 无法读取试剂包芯片 | 无法读取试剂包芯片，借备件903-644,903-158,903-159.636-555各一个 |  | SWO-180132 | Defect | HW | Basic Unit | PCB, SP ID |  | Unable to read reagent package chip | Unable to read reagent package chips, spare parts 903-644, 903-158, 903-159.636-555 each
No | Visit | Complaint | RCHN | 报错4548 | 报错4548，借备件886-638一个 |  | SWO-180050 | Defect | HW | Cover and Chassis | Inlet Door | 4548 | 4548 error | 4548, spare parts 886-638 one
No | Visit | Complaint | RIND | Autocheck scanner not working properly | Checked and found QC ampoules are not being scanned in autocheck module, scratches found on the golden plate in the scan engine. | Cleaned and aligned laser position, it was not possible to adjust the value 191-193 and adjusted at 196. Only two QC ampoules were scanned. | SWO-180690 | Cleaned | HW | Auto Check Module | Barcode reader, Autocheck Module |  | Autocheck scanner not working properly | Checked and found QC ampoules are not being scanned in autocheck module, scratches found on the golden plate in the scan engine. | Cleaned and aligned laser position, it was not possible to adjust the value 191-193 and adjusted at 196. Only two QC ampoules were scanned.
No | 5 | Y | RCZE | Blank |  |  | D24DC5EE58534690 | Assisted Replacement | Electrode | Electrode | pO2 |  | Blank
No | Visit | Complaint | RDEU | 20.03.17 11:10 Herr Hilwerth Tel.. 017615322524 , Kalibratzionsfehler bei pCo2 | PCO2 Elektrode gebrochen | pCO2 Elektrode (E788 XK-01) geprüft und (aus Kundenbestand) ausgetauscht, Funktionskontrolle ok | SWO-180199 | Assisted Replacement | Electrode | Electrode | pCO2 | Broken | 20.03.17 11:10 Mr. Hilwerth Tel. 017615322524 , Calibration error in pCo2 | PCO2 electrode broken | pCO2 electrode (E788 XK-01) tested and replaced (from customer stock), function control ok
No | 5 | Y | RAME | 2 pCO2 QC results are running low compared to peer. |  | pCO2 QC results trended low compared to peer since 12-18 reference membrane change. The results plot shows increases over time with different consumable repalcement but results have still been lower than original history. | D24DD3C75801F428 | Defect | Electrode | Electrode | Sealing ring | Sealing ring, pH | 2 pCO2 QC results are running low compared to peer. |  | pCO2 QC results trended low compared to peer since 12-18 reference membrane change. The results plot shows increases over time with different consumable repalcement but results have still been lower than original history.
No | Visit | Complaint | RDEU | Anpassung des Ausdrucks | Anpassung des Ausdrucks | Anpassung des Ausdrucks | SWO-180155 | Setup | SW | Printer Layout | Printer Layout |  | Adaptation of the expression | Adaptation of the expression | Adaptation of the expression
Yes | 9 | Y | RFRA | Blank | Lenteur |  | D24DCA0558003EE8 | Re classified RMED | SWU |  |  |  | Blank | Slowness
Yes | 18 | Y | RPAC | Blank | Analyser unable to pump solutions |  | D24EE02ACE2CB203 | [5 empty] | Blank | Analyser unable to pump solutions
No | Visit | Complaint | RDEU | 20.03.2017; Dr. Dieckmann; Tel.: 0602132433200; keine Radiance-Verbindung 21.03.2017 Pfl.Thomas 06021324723: keine Radiance-Verbindung | Einsetzen der konfigurierten Speicherkarte | Speicherkarte (mitgeliefert in CPU) in GS vorinstalliert (Windows, Applikationssoftware und Kundenkonfig vom 27.03.17) Öffnen des Analysators und einsetzen der CF Karte. Neustart und Initialisierungsphase. Neu einlesen der Betriebsmittel. Initialisierung 50Min. Kal und QC ohne Beanstandung Dauerhafte Speicherung der Radianceverbindung erfolgreich. | SWO-180403 | Defect | HW | DMS_PC | CF Card |  | 20.03.2017; Dr. Dieckmann; Tel.: 0602132433200; No radiance connection 21.03.2017 Pfl.Thomas 06021324723: no radiance connection | Inserting the configured memory card | Memory card (supplied in CPU) pre-installed in GS (Windows, application software and customer confectionery from 27.03.17) Open the analyzer and insert the CF card. Restart and initialization phase. Re-reading the equipment. Initialization 50min. Kal and QC without complaint Permanent storage of the radiance connection successfully.
No | Visit | Complaint | RDEU | 22.03.2017 11:00 H. Kruse 0211904361249: Kalkulationsfehler aufgrund zu niedriger Hb-Werte | Gerät verstopft. | Gerät zerlegt, Hemolyzer mehrfach durchgespült. tHB Kal durchgeführt. Mehrfach blutige Messungen durchgeführt. | SWO-180756 | [5 empty] | 22.03.2017 11:00 H. Kruse 0211904361249: Calculation error due to too low Hb values | device clogged. | device disassembled, Hemolyzer rinsed several times. tHB Kal. Several bloody measurements performed.
No | Visit | Complaint | RCHN | LAC定标失败 | LAC定标失败 | 更换测试卡正常。 | SWO-180410 | [5 empty] | LAC calibration failure | LAC calibration failure | The replacement test card is OK.
No | Visit | Complaint | RCHN | 报错4794 | 报错4794，借备件903-171,636-549各一个 |  | SWO-180296 | Defect | HW | Aspiration Module | Bend needle | 4794 | 4794 error | Error 4794, one spare part 903-171,636-549 each
No | Visit | Complaint | RCHN | 进样盘底座上面的轮子的卡子断了，样本固定不了 | 进样盘底座上面的轮子的卡子断了，样本固定不了,借903-584一个 |  | SWO-180319 | Defect | HW | Aspiration Module | Tube holder |  | The clip of the wheel above the base of the intake disc is broken and the sample cannot be fixed | The clip of the wheel above the base of the intake disc is broken and the sample cannot be fixed, and the sample is not fixed, and the 903-584 is a
No | 14 | Y | RAME | 0 tHb is low out of limits on verification material. |  | S7950 is out of limits low. S7940 is out or borderline low. | D24EEE86CE602038 | Defect | HW | Oxi module | Hemolyzer |  | 0 tHb is low out of limits on verification material. |  | S7950 is out of limits low. S7940 is out or borderline low.
No | Visit | Complaint | RDEU | Touchscreen ohne Funktion; Medizintechnik war bereits vor Ort und hat Monitor-Neustart ausgelöst! | TS ohne Funktion bei Bedienung. Die Medizintechnik war bereits vor Ort und hat Monitor-Neustart ausgelöst; kurze Zeit später erneute Fehlermeldung durch die Anwender | TS-Eingabe überprüft, bei meinem Eintreffen waren alle Funktionen wieder einwandfrei, aktuell keine Fehler feststellbar! Kal. und QC Messung im Rahmen der FAN915-318 durchgeführt. | SWO-180488 | Defect | HW | DMS_PC | Touch screen |  | Touchscreen without function; Medical technology was already on site and triggered monitor restart! | TS without function when operating. Medical technology was already on site and triggered monitor restart; a short time later, the user's error message | TS input checked, when I arrived all functions were flawless again, currently no errors can be detected! Kal. and QC measurement within the scope of the FAN915-318.
No | 8 | Y | RAME | W demo option key for lac expiring |  | found that the lac key did not install right at installation of analyzers. | D24EE3B5CE4BE4E0 | Setup | Demo key exp | Demo Key Exp | Demo key expired |  | W demo option key for lac expiring |  | found that the lac key did not install right at installation of analyzers.
No | Visit | Complaint | RDEU | 29.03.2017 Frau Hankel 03948370265 berichtet, dass in der letzten Woche auch einen Pumpenkall-Fehler gab der von einem Mitarbeiter behoben wurde. | Airventilmembran im pH/BG Modul defekt | Ventile überprüft und gereinigt defekte Airventilmembran im pH $BG Modul ersetzt Kal- und QC | SWO-180465 | Defect | HW | pH/BG Module | Valve, air |  | 29.03.2017 Ms. Hankel 03948370265 reports that last week there was also a pump-kall error that was fixed by an employee. | Airvalve membrane defective in pH/BG module | Valves checked and cleaned defective air valve membrane in pH $BG module replaced Kal- and QC
No | 6 | Y | RAME | P Customer reported oxi liquid sensor error after unit power up restart. |  | Air line needs flushed. Also noticed halogen lamp voltage too high on calibrations intermittently. | D24F6B15CE2D06CF | Cleaned | Fluidic pathway | Oxi module | Oxi module, Cleaned |  | P Customer reported oxi liquid sensor error after unit power up restart. |  | Air line needs flushed. Also noticed halogen lamp voltage too high on calibrations intermittently.
No | Visit | Complaint | RDEU | 29.03.2017, Frau Schäfer, tel.: 091318543855: Autocheck Mechanik Fehler, Frau Schäfer hat die Chargen geprüft, Die Autocheck Türe geht auf, Karussell bleibt liegen. | AutoCheck Ampullen werden nicht gemessen | AutoCheck Modul im Service überprüft, Ampullen werden eingelesen, sie werden konditioniert, die Ampullen werden aufgebrochen, das Karussell dreht sich, der Deckel (Lid) öffnet sich, der Becher worin das AutoCheck-Karussell gelagert ist hebt sich bei Ansteuerung nicht an, zum Test das 14-polige Flachbandkabel getauscht - brachte keine Besserung, erforderliche Ersatzteile werden bestellt... AutoCheck Ampullen werden manuell gemessen | SWO-180458 | Revisit required | HW | Autocheck module | Autocheck module |  | 29.03.2017, Mrs. Schäfer, tel.: 091318543855: Autocheck mechanics error, Mrs. Schäfer has checked the batches, The car check door opens, carousel remains. | AutoCheck ampoules are not measured | AutoCheck module checked in the service, ampoules are read in, they are conditioned, the ampoules are broken open, the carousel turns, the lid opens, the cup in which the AutoCheck carousel is stored does not lift at the control, to the Test the 14-pin ribbon cable replaced - brought no improvement, required spare parts are ordered... AutoCheck ampoules are measured manually
No | 2 | Y | RCZE | Blank | neon position failed - Fneon: 2.645 ( limit do 2.000 ) |  | D2506D904402D23E | Defect | HW | Oxi module | Spectrometer | LEH , why | Blank | neon position failed - Fneon: 2.645 ( limit do 2.000 )
No | 2 | Y | RCZE | Blank | neon position failed, Fneon: 2.178 ( limot 2.000 ), dále chyba v hemolyzeru - Impedance Shape : Phaseregulation error |  | D2506FAD4449D242 | Defect | HW | Oxi module | Spectrometer | LEH , why | Blank | neon position failed, Fneon: 2.178 ( limot 2.000 ), dále chyba v hemolyzeru - Impedance Shape : Phaseregulation error
No | Visit | Complaint | RDEU | Parameterprobleme pH und Elektrolyte Sensorkassette und Lösungspack wurden getauscht | Sensorfehler, Gerät steht nicht waagerecht | Gerät überprüft, Standort kontrolliert, Sensor erneuert, tHb Kal durchgeführt, Lösungspack ausgetauscht | SWO-180786 | Assisted Replacement | SC | SC | Reference |  | Parameter problems pH and electrolytsensor cassette and solution pack have been replaced | Sensor error, device is not horizontal | Device checked, location checked, sensor renewed, tHb Kal performed, solution pack replaced
No | Visit | Complaint | RDEU | 30.03.2017 11:13Uhr Frau Knabe 03661462543: Seit gestern Abend meldet das Gerät div. Fehler: Flow-Selector, Vaccumfehler, Leckagefehler, Flüssigkeit nicht gepumpt usw. | Fehlersuche | Probenweg gespült Verschleißteile und Schläuche sowie Quetschventil erneuert | SWO-180835 | Defect | HW | Basic Unit | PM kit |  | 30.03.2017 11:13am Mrs. Knabe 03661462543: Since last night the device reports various errors: Flow-Selector, Vaccum error, leakage error, liquid not pumped etc. | Troubleshooting | Sample path rinsed Wear parts and hoses as well as crushing valve renewed
No | Visit | Complaint | RLTD | Please call Baneta Yelda on 0208 216 4934 | Time on all analyser jumping back an hour when connected to net work time server. | Changed settings in RMED Dst managed service from manual to automatic as advised by RMED. All changes saved to CM card ok. Test time sync with time server ok. All 11 analysers have been done. | SWO-180875 | Setup | SW | Time setting | Time setting |  | Please call Baneta Yelda on 0208 216 4934 | Time on all analyser jumping back an hour when connected to net work time server. | Changed settings in RMED Dst managed service from manual to automatic as advised by RMED. All changes saved to CM card ok. Test time sync with time server ok. All 11 analysers have been done.
No | Visit | Complaint | RSAF | Failing calibrations and Qc's on p02 and giving inhomogeneous samples/rinse |  |  | SWO-180866 | Assisted Replacement | SP | SP | FT |  | Failing calibrations and Qc's on p02 and giving inhomogeneous samples/rinse
Yes | Visit | Complaint | RDEU | LP wird benötigt | Absprache mit Frau Essers | 2 LPs auf die E2 gebracht als Reserve für das Haus | SWO-180717 | Re classified RMED | Logistical upSWO |  |  |  | LP is required | Arrangement with Ms. Essers | 2 LPs brought to the E2 as a reserve for the house
No | 3 | Y | RFRA | Blank | Erreur 375 : valeur état pO2 hors gamme suite au remplacement de la cassette de capteur |  | D2505CED445902DE | Defect | HW | Basic Unit | PCB, Sensor interface |  | Blank | Error 375: off-range pO2 status value following sensor cassette replacement
No | Visit | Complaint | RCHN | 黑屏 | 黑屏，借903-381一个，933-215暂时缺货不借 | 更换903-381后修复by jesson | SWO-180434 | [5 empty] | Black screen | Black screen, borrow 903-381 one, 933-215 temporarily out of stock | Repair by jesson after replacement 903-381
No | 4 | Y | RAME | D Customer says COHGB is running higher on this analyzer compared to 313195 OSM right next to it. |  | Checked data logs. Ran comparison bloods obtained from another customer site. COHGB consistently ran 1.0 higher on this analyzer compared to the OSM right next to it. | D250808144329889 | [5 empty] | D Customer says COHGB is running higher on this analyzer compared to 313195 OSM right next to it. |  | Checked data logs. Ran comparison bloods obtained from another customer site. COHGB consistently ran 1.0 higher on this analyzer compared to the OSM right next to it.
No | 17 | Y | RAME | Z Soultion pack leaking inside the solution pack cavity. |  | Found the molded door hinge latch broken. It was not keeping the solution pack tight against the waste luer. When I got it back up it failed the Glu low on level 1 external QC. Put in a new sensor cassette and it gave a sensor circuit error. Replaced it with another SC and all came in. | D2502CA844590CCF | [5 empty] | Z Soultion pack leaking inside the solution pack cavity. |  | Found the molded door hinge latch broken. It was not keeping the solution pack tight against the waste luer. When I got it back up it failed the Glu low on level 1 external QC. Put in a new sensor cassette and it gave a sensor circuit error. Replaced it with another SC and all came in.
No | Visit | Complaint | RCHN | 报错5784，测试卡断裂，造成卡盘 | 报错5784，测试卡断裂，造成卡盘 | 更换测试卡后修复 | SWO-180944 | Adjustment | Consumables | Test Kit | Cartridge | 5784 (Broken) | 5784, test card break, resulting in card disk | 5784, test card break, resulting in card disk | Repair after replacing test card
No | Visit | Complaint | RIND | Lactate Sensitivity Low | Lactate Sensitivity Low , Membrane Need To Be Change | Replaced Lactate membrane With New One ,Machine Passed 2 Point Calibration & QC . Machine Working Fine | SWO-181280 | Assisted Replacement | Electrode | Membrane | Lac | Sensitivity | Lactate Sensitivity Low | Lactate Sensitivity Low , Membrane Need To Be Change | Replaced Lactate membrane With New One ,Machine Passed 2 Point Calibration & QC . Machine Working Fine
No | 14 | Y | RAME | 0 tHb out of limits on verification material. |  | 3 different lot #'s of verification run low out of limits on s7950 solution. Unit is operating normally otherwise with all QC near mean values. | D2501DEF441B344B | Defect | HW | Fluidic manifold | Pump tubes |  | 0 tHb out of limits on verification material. |  | 3 different lot #'s of verification run low out of limits on s7950 solution. Unit is operating normally otherwise with all QC near mean values.
No | 10 | Y | RAME | 0 tHb results on S7950 verification results are low out of range. |  | # lots of material tested. An occasional result came in on the low limit. Unit operating normally otherwise with no cal or QC problems. | D2503D44443C122B | Defect | HW | Fluidic manifold | Pump tubes |  | 0 tHb results on S7950 verification results are low out of range. |  | # lots of material tested. An occasional result came in on the low limit. Unit operating normally otherwise with no cal or QC problems.
No | Visit | Complaint | RDEU | 30.03.2017 11:20Uhr Herr Rittermann 02086974019: Bei den Vergleichsmessungen mit einem weiteren ABL und dem Labor ist aufgefallen, dass es bei den Elektrolyten große Unterschiede gibt. | Patientenmessungen sind in den E'lyten nicht mit anderen BgA-Geräten vergleichbar, E'lyte werden niedriger gemessen, pH wird höher gemessen, der Kunde hat die Ref.elektrode vor 2 Tagen mit Lot.R1977 bezogen, seit dem gibt es ein Problem. | ich habe die Ref.elektrode noch einmal neu bezogen und das Gerät kalibriert, Probemessungen durchgeführt und mit dem Gerät auf der Intensivst. verglichen, | SWO-180824 | Assisted Replacement | Electrode | Membrane | Reference |  | 30.03.2017 11:20am Mr. Rittermann 02086974019: In the comparative measurements with another ABL and the laboratory it has been noticed that there are big differences in the electrolytes. | Patient measurements are not comparable in the E'lytes with other BgA devices, e'lytes are measured lower, pH is measured higher, the customer obtained the ref.electrode 2 days ago with Lot.R1977, since then there has been a problem. | I have re-received the ref.electrode and calibrated the device, carried out test measurements and used the device on the intensive test. Compared
No | 4 | Y | RPNZ | Blank | High bias on sodium results. |  | D2500E2C441B7748 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | Blank | High bias on sodium results.
No | 7 | Y | RAME | F Customer reported neon lamp errors. |  | Calibration results showing errors 0258: Neon Lamp Intensity Too Low. Ran complete oximetry test with Neon Intensity failing at 24%. | D25038434415A4AE | Defect | HW | Oxi module | Lamp unit |  | F Customer reported neon lamp errors. |  | Calibration results showing errors 0258: Neon Lamp Intensity Too Low. Ran complete oximetry test with Neon Intensity failing at 24%.
No | Visit | Complaint | RDEU | Lösungen können nicht gepumpt werden | Gerinsel im Einlassmodul | Gerät mehrfach gespült und Proteinentfernungen durchgeführt. | SWO-181037 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | Solutions cannot be pumped | Gerinsel in the inlet module | device rinsed several times and carried out protein removals.
No | Visit | Complaint | RDEU | 29.03.2017, Frau Schäfer, tel.: 091318543855: Autocheck Mechanik Fehler, Frau Schäfer hat die Chargen geprüft, Die Autocheck Türe geht auf, Karussell bleibt liegen. | AutoCheck Mechanik Fehler Karussell hebt sich bei Messung nicht mehr an Keine mechanische Blockade festzustellen 14-poliges Flachbandkabel Ok | Tausch des AutoCheck Moduls, AutoCheck Modul im Service überprüft, System- und Funktionskontrolle in Ordnung | SWO-180693 | Defect | HW | Auto Check Module | Auto Check Module |  | 29.03.2017, Mrs. Schäfer, tel.: 091318543855: Autocheck mechanics error, Mrs. Schäfer has checked the batches, The car check door opens, carousel remains. | AutoCheck Mechanics Error Carousel no longer rises when measured No mechanical blockage to detect 14-pin ribbon cable Ok | Replacement of the AutoCheck module, AutoCheck module checked in service, system and function control in order
Yes | Visit | Complaint | RDEU | tHb kal. lässt sich nicht durchführen. | tHb Kal. lässt sich nicht durchführen. Seriennummer laut Guido Bohn ( Zentrale ) überprüft. | tHb lässt sich nich Kalibrieren, da die Oxiwerte abgewählt sind. Nur pH, pCO2, pO2 , K, Lac. Seriennummer stimmt mit CRM und Liste von Guido Bohn. In den nächsten Tagen wird mit Oberärztin Fr. Manal geklärt ob die Konfiguration so gewollt ist. | SWO-181370 | Re classified RMED | OXI de-selected, could not do tHb Cal |  |  |  | tHb kal. cannot be performed. | tHb Kal. cannot be performed. Serial number according to Guido Bohn ( Headquarters ) checked. | tHb cannot be calibrated because the oxi values have been deselected. Only pH, pCO2, pO2 , K, Lac. Serial number matches CRM and list of Guido Bohn. In the next few days, Chief Physician Fr. Manal will clarify whether the configuration is so intentional.
No | Visit | Complaint | RDEU | Trotz Time Server Anbindung stimmt Uhrzeit nicht, geht 1h nach | Kontrolle und Anpassung Time Servereinstellung | Nach Haken bei Daylight saving rein $raus stimmt Uhrzeit mit Update Now nicht mehr, Haken gestetzt, Timer Sever Test Uhrzeit okay | SWO-181086 | Setup | SW | Time setting | Time setting |  | Despite Time Server connection, time is not right, 1h goes up | Control and customization Time Server setting | After hooking at Daylight saving in $raus time with Update Now no longer corrects, hook stung, Timer Sever test time okay
No | Visit | Complaint | RBNL | ABL90 met nr 16-005-0431 (op de trolley) ziet een onderdeel bij het reagens pack los. Zou je het losse plaatje willen vastzetten, zodat wij de reagens cassette er weer in kunnen plaatsen? (is duidelijk zichtbaar) | Hoe?? | Wasteconnector gedemonteerd om lipje weer op zijn plaats te schuiven. | SWO-181119 |  | LEH: Is this Connector gasket??? |  |  |  | ABL90 with Nr 16-005-0431 (on the trolley), a component is detached from the reagent pack. Would you like to pin the loose picture so that we can put the reagent cassette back in? (is clearly visible) | How?? | The Wasteconnector disassembled to slide the tab back into place.
No | Visit | Complaint | RBNL | ABL90 16-005-0435 is de accu kortgeleden vervangen, maar als wij bij deze ABL de stoomstekker er uit trekken, dan gaat het niet over op de accu, maar gaat het apparaat uit. Zou je dit willen verhelpen | Status geeft 0% aan. Afgelopen maandag al een nieuwe voeding geplaatst i.v.m. mixing problemen. | Accu gereset. Accu spanning blijft 0%, laad niet op. Ook na een half uur nog steeds 0%. Zal een batterij bestellen. | SWO-181078 | Revisit required | HW | Battery module | Battery |  | ABL90 16-005-0435 The battery has recently been replaced, but if we pull out the steam plug in this ABL, it will not be on the battery, but the device will turn off. Would you like to remedy this | Status indicates 0%. Last Monday a new PSU was placed due to mixing problems. | Battery Reset. Battery voltage remains 0%, do not charge. Even after half an hour still 0%. Will order a battery.
No | Visit | Complaint | RKK | 検体の汚れでtHb,PCO2,CLドリフトエラー発生 | tHｂ、PCO2、CLドリフトエラー | 回路洗浄 | SWO-181381 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway | 1316 | THb, PCO2, CL drift error occurs due to specimen stains | tHb, PCO2, CL drift error | Circuit cleaning
No | 3 | Y | RFRA | Blank | ERREUR QC sur l'oxymétrie. |  | D25151E5BA18F54A | Defect | HW | Oxi module | Spectrometer |  | Blank | QC ERROR on oxymetry.
No | 4 | Y | RCZE | Blank | Po príchode chyby pH, minerály a Lac |  | D2517126BA03E525 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway | SC Clot | Blank | Upon arrival of pH defects, minerals and LAC
No | Visit | Complaint | RCHN | PCO2,HCT未计算，K,NA,CA,CL偏低 | PCO2,HCT未计算，K,NA,CA,CL偏低，借902-972一个。 |  | SWO-181063 | [5 empty] | PCO2,HCT未计算，K,NA,CA,CL偏低 | PCO2, HCT uncalculated, K, NA, CA, CL low, borrowed 902-972 one.
No | Visit | Complaint | RCHN | 报1005未连接分析仪 | 报1005未连接分析仪，借903-228一个 | 更换903-228一个正常。by jesson | SWO-180684 | [5 empty] | Reported 1005 unconnected analyzer | Reported 1005 unconnected analyzer, borrowed 903-228 one | Replace 903-228 one normal. by jesson
No | Visit | Complaint | RCHN | 电源板电压太低，启动不了机器 | 电源板电压太低，启动不了机器，借910-246一个 | 更换910-246后修复 by jesson | SWO-180682 | [5 empty] | The supply board voltage is too low to start the machine | The supply board voltage is too low to start the machine, borrow 910-246 one | Repair by jesson after replacement 910-246
No | Visit | Complaint | RCHN | 血氧硬件错误 | 血氧硬件错误，借备件902-980一个 |  | SWO-180924 | Defect | HW | Oxi module | Hemolyzer | 1034 | Blood oxygen hardware error | Blood oxygen hardware error, borrowed spare parts 902-980 one
No | Visit | Complaint | RCHN | 报错1999.22.-1 | 报错1999.22.-1，借备件903-030一个 | 更换903-030后修复 by jesson | SWO-180603 | Defect | HW | Storage Module | PCB, Cartridge Lock | 199.22.-1 | Error 1999.22.-1 | Error 1999.22.-1, borrowed spare parts 903-030 one | Repair by jesson after replacement 903-030
No | Visit | Complaint | RIND | crea calibration expire. | Check the logs. Check the calibration. | Replace the crea regent. Calibrate system multiple time & calibration pass. | SWO-181209 | Assisted Replacement | Consumables | Solution bottle | Crea expired |  | crea calibration expire. | Check the logs. Check the calibration. | Replace the crea regent. Calibrate system multiple time & calibration pass.
No | Visit | Complaint | RBNL | Problème imprimante | Imprimante défecteueuse. | Remplacement de l'imprimante et essai : OK | SWO-181197 | Defect | HW | Printer Module | Printer head |  | Printer problem | Defective printer. | Printer replacement and test: OK
No | 9 | Y | RCZE | Blank |  |  | D2512E7DBA11E520 | Cleaned | Fluidic pathway | Oxi module | Fluidic pathway |  | Blank
No | 6 | Y | RCZE | Blank |  |  | D2512F7DBA07E523 | Cleaned | Fluidic pathway | Oxi module | Fluidic pathway |  | Blank
No | Visit | Complaint | RKK | pCO2 QC測定値 変動 | pCO2 QC測定値 変動 | pCO2電極・電極コネクター交換 Clメンブラン交換 LS調整実施 CAL・QC測定データ確認 | SWO-181341 | Assisted Replacement | Electrode | Membrane | pCO2 |  | pCO2 QC readings Variation | pCO2 QC readings Variation | Exchange of pCO2 electrodes and electrode connectors Cl-Menbran Replacement Ls adjustment implemented Cal/QC measurement data confirmation
No | Visit | Complaint | RKK | pCO2 QC測定値 変動 | pCO2 QC測定値 変動 | pCO2電極・電極コネクター交換 Clメンブラン交換 LS調整実施 CAL・QC測定データ確認 | SWO-181342 | Assisted Replacement | Electrode | Membrane | pCO2 |  | pCO2 QC readings Variation | pCO2 QC readings Variation | Exchange of pCO2 electrodes and electrode connectors Cl-Menbran Replacement Ls adjustment implemented Cal/QC measurement data confirmation
No | 11 | Y | RAME | 1 pH sensitivity errors. Unit also has k, Na and Ca sensitivity errors along with Na and Ca status errors. |  | Reference has been remembraned X3 by customer with lot #2022. pH sens. 111.9, K 105.4, Na 87.9 and Ca 122. | D251A9CEBA4E1D36 | Assisted Replacement | Electrode | Electrode | Reference |  | 1 pH sensitivity errors. Unit also has k, Na and Ca sensitivity errors along with Na and Ca status errors. |  | Reference has been remembraned X3 by customer with lot #2022. pH sens. 111.9, K 105.4, Na 87.9 and Ca 122.
No | Visit | Complaint | RDEU | Glu auch nach Membranwechsel immer Kalfehler, Lac nach div Kal nun auch fehlerhaf, Wechsel ohne Erfolg | Kaldatei durchgesehen, auch pCO2 ist jetzt fehlerhaft | Fehlersuche, Glu und Lac Membran neu - Elektroden mit Cleaner behandelt, pCO2 wurde eben erst membraniert, Messkammern gereinigt, mehrfach kalibriert, fehlerhafte QC wiederholt | SWO-181396 | Assisted Replacement | Electrode | Membrane | Glu |  | Glu even after membrane changes always cal error, Lac after div Kal now also errorhaf, change without success | Cal file reviewed, even pCO2 is now faulty | Troubleshooting, Glu and Lac membrane new - electrodes treated with cleaner, pCO2 has just been membraneed, measuring chambers cleaned, multiple calibrated, faulty QC repeated
No | 10 | Y | RAME | Z Error 374: Inhomogenous sample at pH/BG module. |  | Examined data logs.This is a repetitive random issue. Examined inlet and pH $BG module. | D25182BFBA349E5F | Defect | HW | Inlet module | Nipple kit |  | Z Error 374: Inhomogenous sample at pH/BG module. |  | Examined data logs. This is a repetitive random issue. Examined inlet and pH $BG module.
No | Visit | Complaint | RCHN | LAC偏高，漂红 | LAC偏高，漂红 | 更换测试卡正常。 | SWO-181282 | [5 empty] | LAC high, red | LAC high, red | The replacement test card is OK.
No | Visit | Complaint | RDEU | 31.03.2017; Sr. Alexandra; Tel.: 08213167351; Rad am Probenhalter ist abgebrochen | Probehalber gebrochen. Nadel ist in Ordnung. Abgebrochenes Teil wurde vom Kunden bereits entsorgt, da es nicht ins Gerät viel. Einsatz wurde von Sr Alexandra Pospisil ngefordert. Sie wurde daraufhin gewiesen ,dass der Einsatz kostenpflichtig ist. | 2 Probenröhrchen aus Gerät entfernt. Tube Holder erneuert. Systemüberprüfung ok. 1x CRP Testmessungen zu Funktionsüberprüfung durchgeführt - ok. Messung wurde ohne Fehler durchgeführt. LQC konnte nicht durchgeführt werden, da Labor geschlossen war. | SWO-181389 | Defect | HW | Aspiration Module | Tube holder |  | 31.03.2017; Sr. Alexandra; Tel.: 08213167351; Wheel on the sample holder is broken off | Broken for trial. Needle is fine. Cancelled part has already been disposed of by the customer as it does not enter the device much. Mission was requested by Sr Alexandra Pospisil. She was then instructed that the use is chargeable. | Removed 2 sample tubes from the device. Tube Holder renewed. System check ok. 1x CRP test measurements for function verification performed - ok. Measurement was carried out without errors. LQC could not be performed because Labor was closed.
Yes | Visit | Complaint | RDEU | LP wird dringend benötigt. | Gerät kann nicht mehr messen. | Kunden darüber aufgeklärt das der Einsatz berechnet wird. LP ausgeliefert. | SWO-181354 | Re classified RMED | Logistical upSWO |  |  |  | LP is urgently needed. | Device can no longer measure. | informed customers that the use will be charged. LP delivered.
No | 10 | Y | RPNZ | Blank | Daylight savings did not change automatically |  | D253C6CFA63E7C11 | Setup | SW | Time setting | Time setting | Summer time | Blank | Daylight savings did not change automatically
No | 6 | Y | RPNZ | Blank | Daylight savings time did not change automatically |  | D253B3AFA60FEA15 | Setup | SW | Time setting | Time setting |  | Blank | Daylight savings time did not change automatically
No | 11 | Y | RPNZ | Blank | Daylight savings did not change automatically |  | D253BC46A61A7C0D | Setup | SW | Time setting | Time setting |  | Blank | Daylight savings did not change automatically
No | Visit | Complaint | RLTD | Karen Shaw - 0131 536 0206 | Analyser in intervention required screen with leak detected error Found inlet probe bent and gasket faulty | Fitted new probe and gasket, fitted new inlet clip which was missing Carried out all fluid transport check in service program ok | SWO-181623 | Replaced | HW | Inlet module | Probe bend |  | Karen Shaw - 0131 536 0206 | Analyser in intervention required screen with leak detected error Found inlet probe bent and gasket faulty | Fitted new probe and gasket, fitted new inlet clip which was missing Carried out all fluid transport check in service program ok
No | Visit | Complaint | RDEU | Hr. Handlooser, Tel.: 060329394123, Touchscreen falsch kalibriert. | Touchscreen Kalibration | Touchscreen Kalibration konnte nicht durchgeführt werden. Bereiche der linken Displayhälfte werden nicht erkannt. Analysator neu gestartet. Weiterhin fehlerhaft. Touchscreen wird bestellt und nach Erhalt eingebaut | SWO-181626 | Revisit required | HW | DMS_PC | Touch screen |  | Mr. Handlooser, Tel.: 060329394123, Touchscreen incorrectly calibrated. | Touchscreen Calibration | Touchscreen calibration could not be performed. Areas of the left half of the display are not recognized. Analyzer restarted. Still faulty. Touchscreen is ordered and installed upon receipt
No | 2 | Y | RAME | Z Customer called with a white screen on the analyzer. |  | Marge informed me that the analyzer fell off the table and the white screen appeared after start up. | D255271B1C2EF8AF | User error | HW | Analyzer dropped | Analyzer dropped |  | Z Customer called with a white screen on the analyzer. |  | Marge informed me that the analyzer fell off the table and the white screen appeared after start up.
No | Visit | Complaint | RSAF | Complaining again that the analyzer is failing Qc's only this time on pO2 for Qc's level A and B | Analyzer giving Qc errors since 02/04/2017 on level A and C only for pO2 parameter. | Entered service program and done all tests but showed no errors, replaced solution pack as analyzer requested but soon after replacing solution pack required replacement of sensor cassette as well. | SWO-181821 | Assisted Replacement | SC | SC | pO2 |  | Complaining again that the analyzer is failing Qc's only this time on pO2 for Qc's level A and B | Analyzer giving Qc errors since 02/04/2017 on level A and C only for pO2 parameter. | Entered service program and done all tests but showed no errors, replaced solution pack as analyzer requested but soon after replacing solution pack required replacement of sensor cassette as well.
No | Visit | Complaint | RDEU | Neu eingesetzte SC90 hat nach 48h pCO2-Kalibrier- u. QC-Fehler Kunde hat keine SC90 mehr verfügbar | Neue SC hat pCO2-Kalibrier- u. QC-Fehler, Kunde hat durch SC-Fehler keine SC90 mehr verfügbar | Neue SC90 installiert, Credit Claim Form erstellt, Startphase abgewartet, Kal u. QC i.O. Defekte SC90: Lot R0595 - Verfallsdatum erreicht (Verfallsdatum 2017-08-07, vorher nicht benutzt!), Lot: 1009-191 - pCO2 Kalibrier- u. QC-Fehler | SWO-181741 | Assisted Replacement | SC | SC | pCO2 |  | Newly installed SC90 has failed after 48h pCO2 calibration and QC errors Customer no longer has SC90 available | New SC has pCO2 calibration &i.e. error, customer no longer has SC90 available due to SC error | New SC90 installed, Credit Claim Form created, start-up phase awaited, Kal & QC i.O. Defective SC90: Lot R0595 - Expiration date reached (expiry date 2017-08-07, not used before!), Lot: 1009-191 - pCO2 Calibration and QC error
Yes | 2 | Y | RPAC | Blank | Used for training |  | D2555C421C2A4D24 | Re classified RMED |  |  |  |  | Blank | Used for training
No | Visit | Complaint | RDEU | Einlass ist ausgefahren und schliest sich nicht mehr. | Funktionskontrolle Einlassnadel verbogen | Einlasssystem zerlegt, gereinigt, montiert und geprüft, verbogene Einlassnadel ausgerichtet. Neustart, abwarten der Kalibration und QC Kontrolle oxymetrie ausser Bereich. Hemolyzer mit proteinentferner und Wasser gespült, clot detect $remove und weitere Programme im Service durchgeführt, mehrere Kalibration erforderlich wg thb drift, QC- Kontrolle | SWO-181673 | Replaced | HW | Auto Inlet module | Probe bend |  | Entrance is extended and does not close anymore. | Funktionskontrolle Einlassnadel verbogen | Inlet system disassembled, cleaned, assembled and tested, bent inlet needle aligned. Restart, wait for calibration and QC control oxymetry off range. Hemolyzer with protein remover and water rinsed, clot detect $remove and other programs performed in the service, multiple calibration required wg thb drift, QC control
No | Visit | Complaint | RDEU | 03.04.2017, Herr Lüttke, Tel.: 02941672240: Lösungen können nicht gepumpt werden. Herr Lüttke ist eigentlich von der Intensivstation. | Lösung QC2 inhomogen | Anwender in Fehlersuche eingewiesen, neues defektes SP ersetzt, SC mit Glucosefehler auch erneuert, Gerät gespült, Kal und QCs | SWO-182033 | Assisted Replacement | SP | SP | FT |  | 04/03/2017, Mr. Lüttke, Tel.: 02941672240: Solutions cannot be pumped. Mr Lüttke is actually from the intensive care unit. | Solution QC2 inhomogeneous | User instructed in troubleshooting, new defective SP replaced, SC with glucose error also renewed, device flushed, Kal and QCs
No | Visit | Complaint | RDEU | eventl. Verstopfung, keine Kalibr.möglich, alle Parameter in rot, | Fehler 2108 Possible fluid pathway blockage . Sensorkassette nicht initialisiert. | Fehler im Zuge der Jahreswartung beseitigt. | SWO-181465 | [5 empty] | eventl. constipation, no calibration possible, all parameters in red, | Error 2108 Possible fluid pathway blockage . sensor cartridge not initialized. | Fixed errors during annual maintenance.
Yes | 13 | Y | RAME | D COHGB are not correlating with OSM 313195. Customers whole blood control is running 1.0 higher than OSM 313195 as well as patient samples. |  | Checked data logs. Ran service test. Failed cavitation absorption. | D25542AA1C11331B | [5 empty] | D COHGB are not correlating with OSM 313195. Customers whole blood control is running 1.0 higher than OSM 313195 as well as patient samples. |  | Checked data logs. Ran service test. Failed cavitation absorption.
No | 4 | Y | RPNZ | Blank | pH failing QC below range. |  | D25855117E366305 | [5 empty] | Blank | pH failing QC below range.
No | 5 | Y | RAME | 0 tHb calibration below zero tHb and being questioned on system cycles All oxi parameters are red |  | Tried a reboot and tHb calibration-No change | D25523431C51FEE8 | [5 empty] | 0 tHb calibration below zero tHb and being questioned on system cycles All oxi parameters are red |  | Tried a reboot and tHb calibration-No change
No | 1 | Y | RPNZ | Blank | Informed it was not working, no details of errors. |  | D255A1C31C0CE1D2 | [5 empty] | Blank | Informed it was not working, no details of errors.
Yes | 75 | Y | RESP | Baja pieza olvidada en parte de reacondicionado. |  | Baja pantalla | D25514811C523413 | Re classified RMED | Run 15 |  |  |  | Low piece forgotten in part of reconditioning. |  | Low screen
No | 4 | Y | RCZE | Blank |  |  | D256B327923B46B2 | Cleaned | HW | El/Met Module | Valve, air |  | Blank
No | 11 | Y | RAME | Z Sample errors. Not getting any tHb results on patient samples. Unable to calibrate. |  | Examined data logs. Found Oxi Hardware failure codes. | D255490C1C539F18 | Adjustment | HW | Oxi module | Lamp unit |  | Z Sample errors. Not getting any tHb results on patient samples. Unable to calibrate. |  | Examined data logs. Found Oxi Hardware failure codes.
Yes | 11 | Y | RPNZ | Blank | Queried ABE/pH results |  | D25574A81C31E1D0 | Re classified RMED | Phone support |  |  |  | Blank | Queried ABE/pH results
No | Visit | Complaint | RDAN | Inlet kal. error | Inlet tjek | Inlet modul skiftet | SWO-181512 | Defect | HW | Inlet module | Inlet module |  | Inlet kal. error | Inlet tjek | Inlet Module Shift
No | Visit | Complaint | RKK | キャル残り6分40秒で中止される。0500予定外にpH/BGモジュールでリンス液検出のエラー。チューブ、バルブ、チャンバー、異常なし。pH/BG上部液センサーが常にFull表示される。液センサー交換し改善。 | キャル残り6分40秒で中止される。0500予定外にpH/BGモジュールでリンス液検出のエラー。チューブ、バルブ、チャンバー、異常なし。pH/BG上部液センサーが常にFull表示される。 | 液センサー交換 | SWO-181719 | Defect | HW | pH/BG Module | LS Sensor |  | The car is aborted in 6 minutes and 40 seconds remaining. 0500 Unscheduled rinse fluid detection error in pH/BG module. Tubes, valves, chambers, no abnormalities. PH / BG upper liquid sensor is always full display. Liquid sensor replacement and improvement. | The car is aborted in 6 minutes and 40 seconds remaining. 0500 Unscheduled rinse fluid detection error in pH/BG module. Tubes, valves, chambers, no abnormalities. PH / BG upper liquid sensor is always full display. | Liquid sensor replacement
Yes | 1 | Y | RCZE | Blank |  |  | D256B1C8922546AA | Re classified RMED | UPS Faulty |  |  |  | Blank
No | 33 | Y | RPOL | Blank | B³¹d kalibracji pompy P3. Pompa P3 nie pracowa³a - wymieniono ca³y modu³ sekcji mokrej. Analizator sprawny technicznie. |  | D25549D01C2370CC | Defect | HW | Process Module | Pump 3 |  | Blank | Calibration error of P3 pump. The P3 pump has not worked-the same wet section module is listed. Technically efficient analyzer.
No | Visit | Complaint | RDEU | SO2 bei QC Level 2 an der oberen Grenze | SO2 bei QC Level 2 an der oberen Grenze | tHb-Kal durchgeführt, Kal, QC | SWO-181900 | Adjustment | HW | Oxi module | tHb Cal |  | SO2 at QC Level 2 at the upper limit | SO2 at QC Level 2 at the upper limit | tHb-Kal durchgeführt, Kal, QC
No | Visit | Complaint | RDEU | 22.03.2017 11:00 H. Kruse 0211904361249: Kalkulationsfehler aufgrund zu niedriger Hb-Werte 30.03.11:20 Hr. Kruse Fehler ist wieder aufgtreten 03.04.2017; 17:15 Uhr; Sr. Marion; der Fehler ist wieder aufgetreten | Kalkulationsfehler aufgrund zu niedriger Hb-Werte. | Systemüberprüfung durchgeführt - keine Fehlermeldungen. Fehler tritt nur sporadisch auf. Oximetrie Ventil ausgebaut und gereinigt, Schläuche erneuert. QA Zyklus durchgeführt. Blutige Messungen durchgeführt. Keine Fehler aufgetreten. | SWO-182103 | [5 empty] | 22.03.2017 11:00 H. Kruse 0211904361249: Calculation error due to too low Hb values 30.03.11:20 Mr. Kruse error is re-emerging 03.04.2017; 5:15 p.m.; Sr. Marion; the error has occurred again | Calculation error due to too low Hb values. | System check performed - no error messages. Error occurs only sporadically. Oximetry valve removed and cleaned, hoses renewed. QA cycle. Bloody measurements carried out. No errors occurred.
No | 1 | Y | RPOL | Blank | Po w³¹czeniu analizator skalibrowa³ siê poprawnie. Podczas testu suchego pomiaru aparat pokaza³ znacznie wiêksze wyniki od 0. Wyczyszczono czujnik i ponownie wymieniono membranê. Po tej operacji na suchym pomiarze czujnik pokaza³ wartoœæ ,0 co jest wyniekim poprawnym. W analizatorze skoñczy³ siê gaz kalibracyjny. Konieczna wymiana gazu. Zaleca siê obserwacjê analizatora. W przypadku popwtarzaj¹cych siê problemów proszê o niezw³oczny kontakt z firm¹ Radiometer. |  | D2565937921CB360 | [5 empty] | Blank | After switching on, the analyser calibrator is correctly calibrated. During the dry measurement test, the camera showed significantly more results from 0. The sensor is cleaned and the membrane is replaced again. After this operation, the sensor's dry measurement showed the value of,0 which is correct. The calibration gas is in the analyser. Gas replacement required. It is recommended to observe the analyser. In the event of a problem, please contact the Radiometer for a non-contacting company.
No | Visit | Complaint | RDEU | Bei Patientenmessungen regelmäßig Fehler 877 FlexQ Positionierungsfehler | s. case | Modul im Servicemenü überprüft, Funktion OK. Modul ausgebaut und Mechanik getestet, keine Fehler feststellbar. Montage, ausgiebiger Test aller Funktionen im Servicemenü, keine Fehler feststellbar. Vorbeugend Datenextrakt gezogen. Neustart, Kal und QC Messung OK | SWO-181964 | Adjustment | HW | FlexQ Module | FlexQ Module |  | Error 877 FlexQ positioning error on patient measurements | s. case | Module checked in the service menu, function OK. Module removed and mechanics tested, no faults detectable. Installation, extensive testing of all functions in the service menu, no errors detectable. Preventive data extract drawn. Restart, Kal and QC measurement OK
No | 4 | Y | RAME | V Instrument will not print |  | checked printer and printer voltage | D256E05D92414D62 | Defect | HW | Printer Module | Printer head |  | V Instrument will not print |  | checked printer and printer voltage
No | Visit | Complaint | RDEU | 0351/8566313 Hr. Helm QC Fehler bei Level 1 und 3 Gluc und Crea und PCo2 31.03.2017 16:45 Hr. Helm Cre Kal Fehler nach Membran Tausch + PCO2 Qc Fehler 03.04.2017 Herr Helm 03518566313 meldet, dass einige Parameter die QC nicht bestehen. | Lac, pCO2, Crea-Kal/ QC-Fehler | Crea-Elektroden membraniert; pCO2 gereinigt; Lac konnte nicht membraniert werden, da keine Membranen vorhanden | SWO-182022 | User error | Electrode | Membrane | Crea expired |  | 0351/8566313 Mr. Helm QC Error at Level 1 and 3 Gluc and Crea and PCo2 31.03.2017 16:45 Mr. Helm Cre Kal Error after Membrane Swap + PCO2 Qc Error 03.04.2017 Mr Helm 03518566313 reports that some parameters do not pass the QC. | Lac, pCO2, Crea-Kal/ QC-Fehler | Crea electrodes membraneed; pCO2 cleaned; Lac could not be membraned because there are no membranes
No | Visit | Complaint | RDEU | Hr. Handlooser, Tel.: 060329394123, Touchscreen falsch kalibriert. | Austausch Display (Touchscreen defekt/Lieferverfügbarkeit des Displays) | Analysator geöffnet und CPU ausgebaut. Austausch des Displays, einsetzen des Barcodescanners. Initialisierungsphase Touchscreen kalibriert und gespeichert. Kal und QC ohne Beanstandung | SWO-182245 | Defect | HW | DMS_PC | Touch screen |  | Mr. Handlooser, Tel.: 060329394123, Touchscreen incorrectly calibrated. | Replacement display (touchscreen defective/delivery availability of the display) | Analyzer opened and CPU expanded. Replacement of the display, inserting the barcode scanner. Initialization phase touchscreen calibrated and stored. Kal and QC without complaint
No | Visit | Complaint | RDEU | 05.04.2017 12:15 POC-Beauftragte 01622349235: Patientenlayout; nur Ca im Ausdruck | Nach Arbeiten und Korrekturen durch Kunden im Patientenlayout sowie in der Parameterkonfiguration werden diverse Parameter nicht mehr berichtet | Vorhergehende Konfiguration des letzte Woche ersetzten ABL90 zurückgesichert, Probelauf ok, Berichtbaren Messbereich für Ca nach unten auf 0,1mmol $L korrigiert, Testmessung ok | SWO-182371 | Setup | SW | Print layout | Print layout |  | 04/05/2017 12:15 POC Representative 01622349235: Patient Layout; only Ca in print | After work and corrections by customers in the patient layout as well as in the parameter configuration, various parameters are no longer reported | Previous configuration of last week's replaced ABL90 backsecured, test run ok, reportable measuring range for Ca down to 0.1mmol $L corrected, test measurement ok
No | 17 | Y | RAME | P Possible Blockage | Changed the complaint visit tab from N to Y. | Able to reproduce error. | D257735E08068B3D | [5 empty] | P Possible Blockage | Changed the complaint visit tab from N to Y. | Able to reproduce error.
No | Visit | Complaint | RIND | Sample problem. Wrong reportings. | Sample problem. | Analyser showing Sample problem while reporting and give all with ? mark. Inlet gasket and sample probe cleaned. Lac and Glu membrane changed. Analyser serviced and calibrated properly. Handed over in a fine working condition. | SWO-182411 | Assisted Replacement | Electrode | Membrane | Glu |  | Sample problem. Wrong reportings. | Sample problem. | Analyser showing Sample problem while reporting and give all with ? mark. Inlet gasket and sample probe cleaned. Lac and Glu membrane changed. Analyser serviced and calibrated properly. Handed over in a fine working condition.
No | Visit | Complaint | RKK | pCO2 QC測定値 変動 | pCO2 QC 変動 | インレットホルダー交換 インレットモジュール調整 タンパク除去、LS調整・ポンプCAL確認 CAL・QC測定データ確認 | SWO-184176 | Cleaned | Fluidic pathway | pH/BG Module | pCO2 |  | pCO2 QC readings Variation | pCO2 QC Variation | Inlet holder replacement Inlet module adjustment Protein removal, LS adjustment and pump CAL confirmation Cal/QC measurement data confirmation
No | 6 | Y | RCZE | Blank | závada na oximteru - Hemolyzer phase regulation failed, mìøení možné. ve všech výsledcích + obèas v QC. |  | D25A6424F434010F | Defect | HW | Oxi module | Hemolyzer | phase err | Blank | Failure on Oximteru-Hemolyzer phase regulation failed, measuring possible. In all results + time in QC.
No | Visit | Complaint | RIND | sample error and pH measurement unstable. | found problem with pH Electrode, found clots at pH electrode. | cleaned the pH electrode and did calibration found all results with in the range. kept instrument observation. | SWO-183229 | Cleaned | Electrode | Electrode | pH |  | sample error and pH measurement unstable. | found problem with pH Electrode, found clots at pH electrode. | cleaned the pH electrode and did calibration found all results with in the range. kept instrument observation.
No | 12 | Y | RAME | 3 PO2 zero on 2pt Cals sometimes falls below 0mmHg and cal is not excepted |  | Reviewed data logs ,and on 2Pt cals the PO2zero falls below 0 | D257BFF50835BEBA | Adjustment | SW | pH/BG Module | Zero current |  | 3 PO2 zero on 2pt Cals sometimes falls below 0mmHg and cal is not excepted |  | Reviewed data logs ,and on 2Pt cals the PO2zero falls below 0
No | 4 | Y | RAME | V Printer printing extra bold print and has a burning smell |  | checked for possible areas of printer burning and printer conections | D25821730862945B | Defect | HW | Printer Module | Printer head |  | V Printer printing extra bold print and has a burning smell |  | checked for possible areas of printer burning and printer conections
No | Visit | Complaint | RKK | インレットの注入口から、液が垂れる。 | 機器状況確認 | プローブホルダー交換作業一式 | SWO-187017 | Defect | HW | Inlet module | Probe holder |  | From the inlet of the inlet, the liquid hangs down. | Equipment status check | Probe holder replacement work set
No | Visit | Complaint | RKK | インレットの注入口から液が垂れる | 機器状況確認 | プローブホルダー交換作業一式 | SWO-187023 | Defect | HW | Inlet module | Probe holder |  | The liquid hangs from the inlet of the inlet. | Equipment status check | Probe holder replacement work set
No | Visit | Complaint | RDEU | Printer druckt sporadisch nicht mehr | Drucker reagiert sporadisch nicht auf seine Druckaufträge. | Testweise neuen kompl. Drucker eingesetzt. Drucker druckt nicht , Papiervorschub o.k. Streckverbinder bis zur CPU kontrolliert. Keine Auffälligkeiten.! Danach einmalige Funktionalität des Druckers , dann wieder keine Reaktion auf Druckauftrag. Eventuell Flachbandkabel und $ oder CPU- Board defekt. Teile müssen bestellt werden. | SWO-182518 | Revisit required | HW | DMS_PC | CPU |  | Printer stops printing sporadically | Printer does not respond sporadically to its print jobs. | Test-wise new compl. printer. Printer does not print , paper feed ok. extend connectors up to the CPU. No abnormalities.! After that, one-time functionality of the printer , then again no reaction to print job. Possibly ribbon cable and a faulty or CPU board. Parts must be ordered.
No | 5 | Y | RFRA | Blank | MACHINE BOUCHEE/CAILLOT DETECTE |  | D25909A17E1DE7A8 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | Blank | MACHINE BOUCHEE/CAILLOT DETECTE
No | 6 | Y | RAME | Z Windows errors. | Changed the complaint visit tab from N to Y and changed the type of service from routine to complaint. | Corrupt files in windows. | D25917C77E1B20C8 | Defect | HW | DMS_PC | CF Card |  | Z Windows errors. | Changed the complaint visit tab from N to Y and changed the type of service from routine to complaint. | Corrupt files in windows.
No | 4 | Y | RAME | P Customer called with leak detected errors. |  | Investigation found 1184 (leak detected) and 1063 (Pressure test vacuum error) | D258D02C7E55C9E0 | Defect | HW | Inlet module | Inlet Gasket |  | P Customer called with leak detected errors. |  | Investigation found 1184 (leak detected) and 1063 (Pressure test vacuum error)
No | 8 | Y | RAME | 9 LAC and GLU cal is out. Sensor response errors- both parameters are --- out on calibration / zero response from sensor cassette. Sensitivity is also 0. SC installed on arrival is lot 1005. |  | Multiple SC's and SP's replaced, instrument has been in and out, will replace sensor interconnect PCB to test for defective Preamp. Tried a new SC lot 1005-198 and same error ( note same lot as the one on analyzer. it too had Glu and Lac errors | D25969AB7E1AF22A | Defect | HW | Basic Unit | PCB, Sensor interface | Glu/Lac | 9 LAC and GLU cal is out. Sensor response errors- both parameters are --- out on calibration / zero response from sensor cassette. Sensitivity is also 0. SC installed on arrival is lot 1005. |  | Multiple SC's and SP's replaced, instrument has been in and out, will replace sensor interconnect PCB to test for defective Preamp. Tried a new SC lot 1005-198 and same error ( note same lot as the one on analyzer. it too had Glu and Lac errors
No | Visit | Complaint | RDEU | Einlaß schließt nur nach dem erneuern der Einlaßdichtung | Einlaß soll häufig nur nach dem erneuern der Einlaßdichtung schließen | Gerät überprüft, Einlaß im Sevice kontrolliert, Anschlussdichtung getauscht, vermeintlich defekte Dichtung eingebaut, Funktion o.k. | SWO-182793 | Revisit required | HW | Auto Inlet module | Auto Inlet module | Gasket? | Admission closes only after renewing the inlet seal | Inlet should often close only after renewing the inlet seal | Device checked, inlet checked in the Sevice, connection seal exchanged, supposedly defective seal installed, function ok.
No | Visit | Complaint | RDEU | 03.04.2017 16:00 Sr. Reinhild 017222000947: Sommerzeit einstellen | FAN | Zeiteinstellung durchgeführt | SWO-182792 | Setup | SW | Time setting | Time setting | Summer time | 04/03/2017 16:00 Sr. Reinhild 017222000947: Set daylight saving time | FAN | Time setting performed
No | Visit | Complaint | RLTD | Contact Stephanie Sandy | Printer door will not close positively. When paper torn off door sometimes springs open. Also paper has been jamming intermittently or printout quality poor. Replaced faulty printer unit. Tested operation - now ok. | Replaced faulty printer unit. | SWO-182589 | Defect | HW | DMS_PC | Printer Module |  | Contact Stephanie Sandy | Printer door will not close positively. When paper torn off door sometimes springs open. Also paper has been jamming intermittently or printout quality poor. Replaced faulty printer unit. Tested operation - now ok. | Replaced faulty printer unit.
No | Visit | Complaint | RCHN | PH,CA,K,NA,CL定标错误 | PH,CA,K,NA,CL定标错误 | 更换测试卡后修复 | SWO-182762 | Assisted Replacement | SC | SC | Reference |  | PH,CA,K,NA,CL定标错误 | PH,CA,K,NA,CL定标错误 | Repair after replacing test card
No | Visit | Complaint | RCHN | GLU响应错误 | GLU响应错误，借903-183一个。 | 备件没有使用，清理了一下后修复(by tim) | SWO-182146 | Cleaned | HW | Basic Unit | SC Contacts |  | GLU response error | GLU responded to the error, borrowing 903-183 one. | Spare parts are not in use, cleanup after repair (by tim)
No | Visit | Complaint | RHUN | tö bb OXI hiba egyszerre, a készülék kétszeri újra indítása sem szűnteti meg a hibákat | - | OXI PCB csere, SW frissítés, halogén lámpa intenzítás beállítása, teljes OXI teszt, tHb kalibráció, 2 pontos kalibráció, QC mérés | SWO-182695 | Defect | HW | Oxi module | PCB, Oxi module |  | If the device is not restarted twice, the machine will no | - | OXY PCB replacement, SW update, halogen lamp intensity adjustment, full OXY test, tHb calibration, 2 point calibration, QC measurement
No | Visit | Complaint | RCHN | NA,K,CL质控偏低 | NA,K,CL质控偏低，借945-603一个 | 更换942-058后修复(by tim) | SWO-182053 | Assisted Replacement | Electrode | Membrane | Reference |  | NA, K, CL quality control low | NA, K, CL low quality control, borrow 945-603 one | Repair after replacement 942-058 (by tim)
No | 6 | Y | RCZE | Blank | po vèerejší výmìnì hemolyzéru kontorola výsledkù, nová chyba oxi mismatch, podezøení na hemolyzér - nový od Marka |  | D25A6646F4330114 | Defect | HW | Oxi module | Hemolyzer |  | Blank | A new error of Oxi mismatch, suspicion of Hemolyzer-new from Mark, and the most recent change in the
Yes | 3 | Y | RAME | 1 Ph out on calibrations. |  | Troubleshot with customer. Could not fix. | D258F47C7E061A3F | Re classified RMED | Phone support |  |  |  | 1 Ph out on calibrations. |  | Troubleshot with customer. Could not fix.
No | Visit | Complaint | RIND | sample error and pH measurement unstable. | found lactate calibration error | replaced the lactate membrane, did the calibration and samples all results found with in the range. | SWO-183232 | Assisted Replacement | Electrode | Membrane | Lac |  | sample error and pH measurement unstable. | found lactate calibration error | replaced the lactate membrane, did the calibration and samples all results found with in the range.
No | Visit | Complaint | RIND | Sample Error In QC Level 3 | All Parameters Out In QC Level 3 & Machine Showing Sample Error | Cleaned The Inlet Gasket , Calibrate Inlet , Pump & Liquid Sensor , Perform Protein Removal , Machine Passed 2 Point Calibration & Run The QC Level 3 Successfully. | SWO-183100 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | Sample Error In QC Level 3 | All Parameters Out In QC Level 3 & Machine Showing Sample Error | Cleaned The Inlet Gasket , Calibrate Inlet , Pump & Liquid Sensor , Perform Protein Removal , Machine Passed 2 Point Calibration & Run The QC Level 3 Successfully.
Yes | Visit | Complaint | RDEU | 05.04.2017 Frau Portmann 023234995291 hat sich noch einmal gemeldet. Es bleibt der Fehler 1005 Warngrenze für Nadel erreicht. | Nadelstiche bei 9500 | Nadel ersetzt während Rep, | SWO-182743 | Re classified RMED | Needle replacement Warning |  |  |  | 04/05/2017 Ms. Portmann 023234995291 has contacted us again. The error 1005 warning limit for needle is reached. | Needle stitches at 9500 | Needle replaced during Rep,
No | Visit | Complaint | RDEU | 05.04.2017 Frau Portmann 023234995291 meldet, dass die Proben abgebrochen werden. | Zeitweise Proben abgebrochen 1065 und 1067 + ungenügende Probe Im Speicher Meldung 2877 Cartrige _40 und 5684 Rotateprocesswheel_40 | Fleißstrecke mit Hypochlorid gespült, Sensoren Storagewheel, Cartrige Bewegung, Proccesswheel , Inkubatorwheel, Punch down und Nadel gereinigt, Motoren gefettet, Systemüberprüfung Systemreinigung durchgeführt, während Überprüfung QC Meldung 4549 Openclose_3 - Stecker Sensor Punch down Sitz korrigiert, Sensoren und Antrieb Inlet Door gereinigt, QC durchgeführt Blindmessungen durchgeführt i.O , Anwender beobachtet und meldet sich gegebenenfalls wieder | SWO-182742 | Cleaned | HW | Aspiration Module | Needle Wash Unit | 1067 | 05.04.2017 Ms. Portmann 023234995291 reports that the samples are being aborted. | Intermittent samples aborted 1065 and 1067 + insufficient sample In memory message 2877 Cartrige _40 and 5684 Rotateprocess_40 | Hard track with hypochloride rinsed, sensors storagewheel, cartrige movement, proccesswheel, incubator wheel, punch down and needle cleaned, motors greased, system verification system cleaning performed while checking QC message 4549 Openclose_3 - plug Sensor punch down seat corrected, sensors and drive Inlet Door cleaned, QC performed blind measurements carried out i.O , user observed and reports again if necessary
No | Visit | Complaint | RDEU | Hr. Giese, Tel.: 071717013732, möchte einen Zebra Drucker installieren, fragt wie der Treiber installiert werden. | Information über Etikettendrucker an Herrn Gieße Med. Technischer Leiter von Dräger TGM | Herrn Gieße über die augenblickliche Situation zur Installation des Etikettendruckers von Zebra unterrichtet. Über den Sachverhalt aufgeklärt, dass derzeit ( laut offizieller Information vom letzten SEKO Meeting ) Radiometer keine Unterstützung und Verantwortung für einen externen Drucker und dessen Installation mittels Druckertreiber übernimmt. Die Situation wird derzeit mit TS und Koppenhagen erörtert und dann entschieden wie weiterhin in dieser Frage gehandelt werden darf im Hinblick bezüglich auf des Medizinproduktegesetz. Herr Gieße versteht die Situation und erwartet kontinuierlich über neue Erkenntnisse unterrichtet zu werden. Herr Gieße steht im Kontakt mit dem Geräteverantwortlichen Funktionsleiter ZNA Herr Miehe. | SWO-183062 | Revisit required | SW | Analyzer Control Module | Printer Driver |  | Mr. Giese, tel.: 071717013732, wants to install a Zebra printer, asks how the driver will be installed. | Information about label printers to Mr. Gieße Med. Technical Director of Dräger TGM | Mr. Gieße is informed about the current situation with the installation of Zebra's label printer. Informed about the fact that currently ( according to official information from the last SEKO meeting ) Radiometer assumes no support and responsibility for an external printer and its installation by means of a printer driver. The situation is currently being discussed with TS and Koppenhagen and then decided how to continue to act on this issue with regard to the Medical Devices Act. Mr. Gieße understands the situation and expects to be continuously informed about new findings. Mr. Gieße is in contact with the device manager function manager ZNA Mr. Miehe.
Yes | 9 | Y | RAME | Blank |  |  | D25D3BA3E02A35F1 | Re classified RMED | No input |  |  |  | Blank
No | Visit | Complaint | RDEU | 07.04.2017, Frau Salomon, Tel.: 06213834524: Kal Fehler bei pH und pCO2; empf: pH 52,5; pCO2:.... Frau Salomon hat eine man. Spülung durchgeführt | Nach 2. SC-Wechsel Lac und Glu fehlerhaft | Überprüfung, KAL und QC | SWO-183037 | Assisted Replacement | SC | SC | pCO2/pH |  | 07.04.2017, Mrs. Salomon, Tel.: 06213834524: Kal error in pH and pCO2; empf: pH 52.5; pCO2:.... Mrs Salomon has one. Rinse performed | After 2nd SC change Lac and Glu faulty | Review, KAL and QC
No | Visit | Complaint | RDEU | 07.04.2017 Herr Kotz 071164893256 ABL erkennt eine Leckage. F1342 | Gerät meldet Leckage erkannt. | Inlet Gasket with Holder erneuert, Einlassdichtungsgummi hat sich beim erneuern durch Kunde nach vorne raus gedrückt. Gerät und Sensorkassette gespült. Lösungspack neu installiert. Einlaufverfahren durchgeführt, alles wieder gut. | SWO-183052 | Defect | HW | Inlet module | Inlet Gasket |  | 04/07/2017 Mr. Kotz 071164893256 ABL detects a leakage. F1342 | Device reports leakage detected. | Inlet Gasket with Holder renewed, inlet seal rubber has pushed itself forward when renewed by customer. device and sensor cartridge. Solution pack reinstalled. Run-in procedures carried out, all well again.
No | Visit | Complaint | RDEU | Einlaß schließt nur nach dem Erneuern der Einlaßdichtung 07.04.2017 Herr Metje 023815891214 meldet, das der fehler weiterhin besteht. Herr M. wünscht einen Vorortservice. | Intervention erforderlich, Vacuumfehler, FTC 1189 | Gerät überprüft, JW durchgeführt | SWO-183042 | Defect | HW | Auto Inlet module | Inlet Gasket |  | Admission closes only after renewing the inlet seal 07.04.2017 Mr Metje 023815891214 reports that the error still exists. Mr. M. wants a suburban service. | Intervention required, Vacuum error, FTC 1189 | Device checked, JW performed
No | 5 | Y | RPAC | Blank | Daylight savings setup |  | D25A3B9DF4424325 | Setup | SW | Time setting | Time setting | Daylight | Blank | Daylight savings setup
Yes | Visit | Complaint | RBNL | ABL90 16-005-0435 is de accu kortgeleden vervangen, maar als wij bij deze ABL de stoomstekker er uit trekken, dan gaat het niet over op de accu, maar gaat het apparaat uit. Zou je dit willen verhelpen | Batterij nog steeds 0% opgeladen. | Nieuwe batterij geplaatst. Nu wel energie 89% --> 100%. Ook normale schakeling tussen battery voeding en netspanning. | SWO-182939 | Re classified RMED | Same day upSWO | Battery module | Battery |  | ABL90 16-005-0435 The battery has recently been replaced, but if we pull out the steam plug in this ABL, it will not be on the battery, but the device will turn off. Would you like to remedy this | Battery still 0% charged. | New battery inserted. Now energy is 89%--> 100%. Also normal switching between battery power and mains voltage.
No | Visit | Complaint | RBNL | Er is al een antal malen de foutmelding lek gedetecteerd voorgekomen. Schoonmaken inletgasket, vervangen gasket, doorspuiten demiwater en of vervangen reagenspack helpt voor een korte periode. | Komt dagelijk voor na systeemcheck maar ook routine (DEMO) gebruik. Vacuum te hoog. Fout 1184 en 1342. Andere inletgasket (nieuw) en druk op gasket uitoefenen geen soulaas. | Inlet mechanisme gedemonteerd en gecontroleerd en weer gemonteerd (spring wind). Alle tubings vervangen. Analyzer start direct zonder problemen op. Vijf bloedmonsters bepaald. | SWO-182885 | Defect | HW | Auto Inlet module | Inlet Gasket |  | There is already an Antal times the error message leak detected occurred. Cleaning inletgasket, replacing gasket, by spraying demi water and or replacing reagent pack helps for a short period of time. | Occurs daily after system check but also routine (DEMO) use. Vacuum too high. Error 1184 and 1342. Other Inletgasket (new) and press gasket do not exercise any soulaas. | Inlet mechanism disassembled and controlled and reassembled (Springwind). Replace All tubings. Analyzer starts immediately without any problems. Five blood samples determined.
No | Visit | Complaint | RBNL | Analyzer schakeld niet over van batterij status naar netspanning als een netsnoer wordt aangesloten. | Tijdens reaparatie lek gedetecteerd, geconstateerd dat de status niet overgaat van batterijvoeding naar netspanning. Uit- en aanschakelen is nodig. | Nieuwe voedingsuinit gemonteerd. Nu wel overgang batterij status naar netspanning. | SWO-182954 | Defect | HW | Power Supply | Power Supply |  | Analyzer does not switch from battery status to mains voltage when a power cord is connected. | During reaparation leak detected, it was found that the status does not pass from battery power to mains voltage. Switching off and on is required. | New feed suinit mounted. Now it does transition battery status to mains voltage.
No | Visit | Complaint | RCHN | LAC定标失败 | LAC定标失败，借903-334一个。 | 更换903-334后修复(by tim) | SWO-182189 | [5 empty] | LAC calibration failure | LAC calibration failed, borrowing 903-334 one. | Repair after replacement 903-334 (by tim)
No | Visit | Complaint | RCHN | 打印机坏了 | 打印机坏了，借902-985一个 | 更换902-985后修复(by tim) | SWO-182116 | [5 empty] | The printer's broken. | The printer is broken, borrow 902-985 one. | Repair after replacement 902-985 (by tim)
No | Visit | Complaint | RIND | Received a complaint from Nimai Children Hospital that after the failure of previously installed SC, the analyzer is displaying 'Sensor Circuit Error' for the newly installed cassette. | Need to check the status of the sensor connector board as well as the zeroing values. | Done zeroing calibration, all the values are fine, replaced the sensor cassette, after calibration, machine is working fine. | SWO-184741 | [5 empty] | Received a complaint from Nimai Children Hospital that after the failure of previously installed SC, the analyzer is displaying 'Sensor Circuit Error' for the newly installed cassette. | Need to check the status of the sensor connector board as well as the zeroing values. | Done zeroing calibration, all the values are fine, replaced the sensor cassette, after calibration, machine is working fine.
No | 23 | Y | RAME | 2 Pco2 fails system cycles after a couple of weeks of the SC been installed |  | Normally, the Pco2 fails system cycle QC $Level 2 low after a couple of weeks of the SC been installed. I had checked the Zero cal and it was fine. Customer tried another SC and it also failed. | D25A3472F45E96AA | [5 empty] | 2 Pco2 fails system cycles after a couple of weeks of the SC been installed |  | Normally, the Pco2 fails system cycle QC $Level 2 low after a couple of weeks of the SC been installed. I had checked the Zero cal and it was fine. Customer tried another SC and it also failed.
Yes | Visit | Complaint | RCHN | 1031定标错误 | 1031定标错误，借903-364一个 | 更换903-364后修复(by tim) | SWO-182436 | Defect | HW | Process Module | Optical Unit | 1031 | 1031 Calibration Error | 1031 calibration error, borrow 903-364 one | Repair after replacement 903-364 (by tim)
No | 6 | Y | RCZE | Blank | po kontrole stále chyba, opakuje se u cca 20% odbìrù ( 3 stejní pacienti ), ostatní OK, sledování odbìrù a mìøení na místì, prùchod krve do oximetru ve vèech pøípadech bez problému, OXI ventil tìsní, krev v tubièce bez pohybu |  | D25A67FCF4050119 | User error | Sample volume | Sample volume | Sample volume |  | Blank | After the inspection, the error is repeated at approximately 20% of the selection (3 same patients), other OK, monitoring of the selection and measurement at the place, blood flow to the oximeter in all cases without problem, OXI valve of the body, blood in the Tubière without movement
No | Visit | Complaint | RDEU | Hb und SO2 Messwerte sind sehr oft in den QC's fehlerhaft und es bestehen oft Driften in den Kalibrierungen. | Hb und SO2 QC Ergebnisse außerhalb der Grenzen. Außerdem bestehen oft driften in den Kalibrierungen. | Lampenspannung überprüft - Ergebnisse OK. Es wurde schon öfters eine Proteinentfernung und eine tHb Kalibrierung durchgeführt. Der Fehler tritt dennoch weiterhin auf. Oximetrie Modul erneuert und reklamiert. Startphase durchgeführt, Lampenspannung überprüft, Kalibriert, service tHb kal und tHb kal durchgeführt. QC's durchgeführt. | SWO-183009 | Defect | HW | Oxi module | Oxi module, Replaced |  | Hb and SO2 readings are very often faulty in the QC's and there are often drifts in the calibrations. | Hb and SO2 QC results outside the limits. In addition, there are often drifts in the calibrations. | Lamp voltage checked - results OK. A protein removal and tHb calibration has been performed several times. However, the error still occurs. Oximetry module renewed and complained. Start phase carried out, lamp voltage checked, calibrated, service tHb kal and tHb kal carried out. QC's.
No | Visit | Complaint | RDEU | 05.04.2017; Fr. Fitzner; Tel.: 06963015479; pCO2 Kal Fehler | Wartungsmassnahme am Analysator. Empfindlichkeit der pCO2 Elektrode schwankt um den unteren Grenzwert von 85%. Anwender haben mehrmals die Elektrode in den letzten Tagen membraniert. | Tausch der pCO2 Elektrode. Nach mehrmaliger Kalibration liegt die Empfindlichkeit bei über 95%, und der Staus liegt etwas über 20 mmHg, Werte in Ordnung. | SWO-183026 | Assisted Replacement | Electrode | Electrode | pCO2 |  | 05.04.2017; Ms. Fitzner; Tel.: 06963015479; pCO2 Kal Error | Maintenance measure on the analyzer. Sensitivity of the pCO2 electrode fluctuates around the lower limit of 85%. Users have membranized the electrode several times in the last few days. | Replacement of the pCO2 electrode. After several calibrations, the sensitivity is over 95%, and the traffic jams are just over 20 mmHg, values are fine.
No | Visit | Complaint | RDEU | Frontaler Barcode-Scanner hat sporadische Aussetzer. | Frontaler Barcode-Scanner hat sporadische Aussetzer. Im Flex-Q-Modul wurden ebenfalls sporadisch Probleme mit dem Barcode-Scanner beobachtet. | Frontalen Barcodescanner gereinigt und justiert. Barcodescanner Flex-Q-Modul gereinigt. Funktionskontrolle beider Scanner i.O. | SWO-182948 | Cleaned | HW | Basic Unit | Barcode reader, Front |  | Frontal barcode scanner has sporadic dropouts. | Frontal barcode scanner has sporadic dropouts. The Flex-Q module also observed sporadic problems with the barcode scanner. | Frontal barcode scanner cleaned and adjusted. barcode scanner Flex-Q module. Functional control of both scanners in the
No | Visit | Complaint | RDEU | Printer druckt sporadisch nicht mehr | Neuer Drucker ohne Funktion. | Kabelverbindungen getestet, kein Erfolg. PC Modul getestet. Bei abgesteckten Anschlüssen 1 und 3 Funktioniert der Drucker, werden diese wieder angesteckt ( blauer und schwarzer Stecker ) keine Reaktion des Druckers. CPU Unit complete DMS910 erneuert. Danach wieder vollständige Funktion gegeben. | SWO-183051 | Defect | HW | DMS_PC | CPU |  | Printer stops printing sporadically | New printer without function. | Cable connections tested, no success. PC module tested. With plugged-in connectors 1 and 3 If the printer works, these are plugged back on (blue and black plug) no reaction of the printer. CPU Unit complete DMS910 renewed. Then full function is given again.
No | 5 | Y | RESP | Error 24 en la process wheel |  | Revision de mantenimiento. Quitar cubeta de la process wheel Realizacion de test de servicio. Realizacion de limpieza. Comprobacion con qc. | D25CA3A2E043793A | Defect | Consumables | SP | Waste lid |  | Error 24 en la process wheel |  | Maintenance review. Remove bucket from the process wheel Performing service test. Cleaning. Check with qc.
No | Visit | Complaint | RDEU | 06.04.2017, Frau Jagusch, Tel.: 022612886440: Fehler 5684 Rotateprocesswheel 56684; | s. case | Beim Eintreffen Gerät prinzipiell betriebsbereit (Testkapazität niedrig). AQT geöffnet und mechanisch beschädigtes Cup entfernt; Anwender hat zuvor weitere 4 Cups entfernt bei Entnahme der Reagenzienkassette. Alle mechanischen Komponenten überprüft, keine Fehler oder Fehlfunktionen feststellbar. Fehler wurde wahrscheinlich verursacht durch fehlerhaften Abtransport verwendeter Cups und deren Blockade in Reag. Kassette. Mit aktueller, gestern neu eingesetzter Reag. Kassette Cup Transport von 10 Cups erfolgreich absolviert. Neustart, Initialisierung, Systemüberprüfung OK QC Messung wird später vom Anwender erledigt! | SWO-182883 | Defect | Consumables | SP | Waste lid |  | 04/06/2017, Mrs Jagusch, Tel.: 022612886440: Error 5684 Rotateprocesswheel 56684; | s. case | When arriving device is basically ready for operation (test capacity low). AQT opened and mechanically damaged cup removed; User has previously removed a further 4 cups when removing the reagent cartridge. All mechanical components checked, no errors or malfunctions detectable. Error was probably caused by incorrect removal of used cups and their blocking in Reag. Cassette. With current, yesterday newly installed Reag. Cassette Cup Transport of 10 Cups successfully completed. Restart, Initialization, System Check OK QC measurement is done later by the user!
Yes | 2 | Y | RPAC | Blank | Damaged sensor cassette |  | D25BE7706A3FE6B8 | Re classified RMED | Phone call |  |  |  | Blank | Damaged sensor cassette
No | Visit | Complaint | RCHN | 气泡探测定标失败 | 气泡探测定标失败，借备件902-972一个 | 更换902-972后修复(tim) | SWO-182835 | [5 empty] | Bubble probe calibration failed | Bubble detection calibration failed, with spare parts 902-972 one | Repair after replacement 902-972 (tim)
No | Visit | Complaint | RCHN | 多参数定标飘红 | 多参数定标飘红 | 更换测试卡后修复 | SWO-183129 | [5 empty] | Multi-parameter calibration floating red | Multi-parameter calibration floating red | Repair after replacing test card
No | Visit | Complaint | RCHN | LAC,GLU定标失败 | LAC,GLU定标失败 | 更换测试卡后修复 | SWO-183075 | Assisted Replacement | SC | SC | Glu/Lac |  | LAC, GLU bid failed | LAC, GLU bid failed | Repair after replacing test card
No | Visit | Complaint | RIND | Lact calibration error | Checked the data logs found Lactate calibration sensitivity out of range | Replaced Lactate membrane | SWO-184551 | Assisted Replacement | Electrode | Membrane | Lac |  | Lact calibration error | Checked the data logs found Lactate calibration sensitivity out of range | Replaced Lactate membrane
No | Visit | Complaint | RCHN | 测试卡漏液 | 测试卡漏液，借备件902-972一个 |  | SWO-183148 | [5 empty] | Test card leakage | Test card leak, spare part 902-972 one
No | Visit | Complaint | RCHN | 软件版本不匹配 | 软件版本不匹配，借903-198一个 | 更换903-198后修复(by tim) | SWO-182906 | Defect | HW | Analyzer Control Module | PCB, Main Board |  | Software version does not match | Software version does not match, borrow 903-198 a | Repair after replacement 903-198 (by tim)
No | Visit | Complaint | RCHN | 定标失败5140 | 定标失败5140，借定标卡944-220（12325）2张 |  | SWO-183143 | Defect | HW | Incubator Module | Punch up | 5140 | Bid failure 5140 | Bid failure 5140, debit bid card 944-220 (12325) 2
No | Visit | Complaint | RCHN | 定标卡条码无法扫描 | 定标卡条码无法扫描，借备件902-924一个 | 更换902-924后修复 | SWO-183099 | Defect | HW | Analyzer Control Module | Barcode reader, display | BR002 | Bid card barcode cannot be scanned | Calibration card barcode can not be scanned, borrowed spare parts 902-924 one | Repair after replacement 902-924
No | 5 | Y | RAME | Z CUSTOMER STATING THAT UNIT IS NOT TRANSMITTING CORRECTLY TO RALS. |  | VERIFIED UNIT NOT TRANSMITTING. | D25DFC49560D57FA | Setup | His/lis | His/lis | His/lis | ASTM file | Z CUSTOMER STATING THAT UNIT IS NOT TRANSMITTING CORRECTLY TO RALS. |  | VERIFIED UNIT NOT TRANSMITTING.
Yes | 3 | Y | RPOL | Blank | Prawid³owo podpiêto UPS-a. Sprawdzenie aparatu. Analizator sprawny. |  | D25DC8275638CC7B | [5 empty] | Blank | The UPS is properly signed. Check the camera. Functional Analyzer.
No | Visit | Complaint | RDEU | Zeitanzeige nach Sommerzeitumstellung nicht korrekt. | Zeitversatz um eine Stunde | Zeitzone auf GTB (Athen) gesetzt. Zeit mit Zeitserver synchronisiert, Daten auf CF-Karte gesichert. | SWO-183237 | Setup | SW | Time setting | Time setting | Summer time | Time display after daylight saving time change not correct. | Time offset by one hour | time zone set to GTB (Athens). Time synchronized with time server, data backed up to CF card.
No | Visit | Complaint | RSAF | Low pO2 values on Qc's even after solution pack and sensor cassette as been replaced. | failing Qc's on pO2 parameter, | Entered service program and done fluid transport test and failed on pressure test. Investigated and straightened bent probe and reran test and passed, | SWO-184297 | Unknown | HW | Auto Inlet module | Probe bend | ext. O2 Leaking | Low pO2 values on Qc's even after solution pack and sensor cassette as been replaced. | failing Qc's on pO2 parameter, | Entered service program and done fluid transport test and failed on pressure test. Investigated and straightened bent probe and reran test and passed,
No | Visit | Complaint | RDEU | 10.04.2017; Fr. Schmitt; Tel.: 065168342215; Probe verloren | Einlassdichtung bleibt in Kapillar Position hängen. | Einlasssonde und Dichtung ersetzt. | SWO-183622 | Defect | HW | Auto Inlet module | Inlet Position Cal. | Revisit - days (MRM) | 10.04.2017; Ms. Schmitt; Tel.: 065168342215; Sample lost | Inlet seal remains stuck in capillary position. | inlet probe and seal.
No | Visit | Complaint | RCHN | 试剂包安装失败，废液口不通 | 试剂包安装失败，废液口不通 | 更换试剂包后修复 | SWO-184153 | [5 empty] | Reagent pack installation failed, waste liquid port is not accessible | Reagent pack installation failed, waste liquid port is not accessible | Repair after replacement of reagent kit
Yes | 15 | Y | RAME | D COHB OUT OF BLOOD COMPARISON RANGE 1.1TO2.1 , RESULT 4.8 |  | COMPLETED CUST. QC RUN AND CUSTOMER RUN WITH ALL RESULTS OUTSIDE NORMALS. EVALUATION OF SYSTEM RESULTS COMPLETE. | D25E476656378A6D | [5 empty] | D COHB OUT OF BLOOD COMPARISON RANGE 1.1TO2.1 , RESULT 4.8 |  | COMPLETED CUST. QC RUN AND CUSTOMER RUN WITH ALL RESULTS OUTSIDE NORMALS. EVALUATION OF SYSTEM RESULTS COMPLETE.
No | 23 | Y | RAME | 2 Pco2 and P02 seem to go out every time a SC/SP is replaced. |  | Customer had a major leak so he had to replace the SP $Then Po2 was out on level 8302 too high 165. | D25E33EC565696B6 | [5 empty] | 2 Pco2 and P02 seem to go out every time a SC/SP is replaced. |  | Customer had a major leak so he had to replace the SP $Then Po2 was out on level 8302 too high 165.
No | Visit | Complaint | RCHN | 测试卡电路开路，接口板腐蚀 | 测试卡电路开路，接口板腐蚀 | 自行更换备用机上的902-972后修复（By Ran） | SWO-183213 | [5 empty] | Test card circuit open, interface plate corrosion | Test card circuit open, interface plate corrosion | Post-902-972 repair on standby (By Ran)
No | Visit | Complaint | RCHN | 光学强度超范围 1031 | 光学强度超范围 1031，借903-364一个。。 |  | SWO-183202 | Defect | HW | Process Module | Optical Unit | 1031 | Optical strength over range 1031 | Optical strength over range 1031, borrowed 903-364 one.
No | 3 | Y | RFRA | Blank | Erreur calibration pCO2 : sensibilité = 65% |  | D25DEB1D56260548 | Assisted Replacement | Consumables | Gas cylinder | Gas2 |  | Blank | PCO2 calibration error: sensitivity - 65%
No | Visit | Complaint | RDEU | Gerät geht nicht in Betriebsmittel -Halt nach Abziehen von Flaschen | Flaschenerkennung ist aktiv, LEDs gehen aus und an. Erkennung ist auch im Serviceprogramm aktiv und korrekt, trotzdem kein Betriebsmittel - Halt nach Abzug der Flaschen. | Fehlersuche: Fluidic- sowie AutoCheck Modul ausgebaut und CANBUS Kabelführung zu beiden Bottle Illum. PCBs und zum Fluidic Modul überprüft. Am linken Fl.Mod. Steckplatz war ein Kontaktpin seitlich abgebogen => Pin gerichtet, Stecker mehrfach testweise gesteckt und überprüft, OK. Alles wieder montiert und sämtliche Funktionen überprüft, Alles OK. Kalibrationen, QC Kontrolle, OK. | SWO-183501 | Defect | HW | Basic Unit | PCB, Illumination |  | Device does not go into equipment -stop after withdrawing bottles | Bottle detection is active, LEDs go out and on. Detection is also active and correct in the service program, yet no equipment - hold after deduction of the bottles. | Troubleshooting: Fluidic and AutoCheck module removed and CANBUS cable routing to both Bottle Illum. PCBs and to the Fluidic module. On the left Fl.Mod. Slot was a contact pin bent sideways => Pin directed, plug plug plug plug repeatedly plugged and checked, OK. Everything reassembled and all functions checked, All OK. Calibrations, QC control, OK.
No | 5 | Y | RCZE | Blank |  |  | D25EFDCDCC0FF9E9 | Defect | HW | El/Met Module | Valve |  | Blank
No | 7 | Y | RFRA | Blank | 8130: ERREUR FLUIDIQUE |  | D25E2822560A4A8E | Cleaned | HW | Aspiration Module | Needle Wash Unit | 8130 | Blank | 8130: FLUIDIC ERROR
No | Visit | Complaint | RKK | SC交換後、pCO2がQCA/B/Cの３濃度で範囲上限値の2倍以上の値で測定される。SP交換後も変化ない。 | SC不良 | SC交換。 SC交換後はQC3レベルとも範囲内に収まる。 | SWO-184155 | Assisted Replacement | SC | SC | pCO2 |  | After SC exchange, pCO 2 is measured at a value of more than twice the range upper limit at 3 concentrations of QCA / B / C. It does not change even after sp exchange. | SC bad | SC replacement. After SC exchange, both QC3 levels fit within the range.
No | Visit | Complaint | RDEU | Diverse Erneuerungsmaßnahmen und Elektrodenprobleme | Glukose Reaktionsfehler, Statusfehler Glu Lac Sensorkassette auskristallisiert | JW durchgeführt, Sensorkassette ausgetauscht | SWO-183920 | Assisted Replacement | SC | SC | Glu |  | Various renewal measures and electrode problems | Glucose Reaction Errors, Status Error Glu Lac Sensor cassette crystallized | JW performed, sensor cartridge replaced
No | Visit | Complaint | RDEU | 11.04.2017 13.00Uhr Herr Ginschel 095150315247: Das Einlassmodul wurde vom Personal abgebrochen | Fehlersuche --> Halterung am Einlassarm abgebrochen --> durch Fehlbedienung vom Anwender verursacht | Einlassarm (R1102N010) eingebaut, System und Funktion i.O. | SWO-183833 | Defect | HW | Inlet module | Inlet Handle | Broken | 11.04.2017 13.00Hr Mr Ginschel 095150315247: The inlet module was cancelled by the staff | Troubleshooting --> Bracket on the inlet arm broken --> caused by incorrect operation by the user | Inlet arm (R1102N010) built-in, system and function in the
No | 15 | Y | RAME | P Liquid Sensor Calibration Errors. |  | The customer replaced the printer paper and then had general printer errors. The customer shut down the instrument but had liquid sensor errors after restarting the instrument. Checked all internal connections and found printer power connection was loose which may have caused the instrument electronics to malfunction. | D25F7E0CCC386F62 | Adjustment | HW | Printer Module | Cable | Loose | P Liquid Sensor Calibration Errors. |  | The customer replaced the printer paper and then had general printer errors. The customer shut down the instrument but had liquid sensor errors after restarting the instrument. Checked all internal connections and found printer power connection was loose which may have caused the instrument electronics to malfunction.
Yes | 5 | Y | RPAC | Blank | Sensor cassette failure (1361) - Internal reference electrode error in sensor cassette. |  | D25F8630CC03900B | [5 empty] | Blank | Sensor cassette failure (1361) - Internal reference electrode error in sensor cassette.
No | Visit | Complaint | RSAF | Can't get the inlet module to lower to the closed position. | Inlet stuck in syringe mode position. | manually turned curved wheel assembly to the closed $resting position. Tested inlet in service mode and functioned well in all positions. | SWO-184318 | Defect | HW | Auto Inlet module | Inlet Position Cal. |  | Can't get the inlet module to lower to the closed position. | Inlet stuck in syringe mode position. | manually turned curved wheel assembly to the closed $resting position. Tested inlet in service mode and functioned well in all positions.
No | 2 | Y | RAME | P Customer reported Possible blockage error message intermittently during QC testing. |  | Checked data logs for errors. Possible blockage error message intermittently. | D25F7B12CC14578C | [5 empty] | P Customer reported Possible blockage error message intermittently during QC testing. |  | Checked data logs for errors. Possible blockage error message intermittently.
No | 2 | Y | RAME | P Customer reported Possible blockage error message intermittently during QC testing. |  | Checked data logs for errors. Intermittent Possible blockage message during QC runs. | D25F7721CC0E5787 | [5 empty] | P Customer reported Possible blockage error message intermittently during QC testing. |  | Checked data logs for errors. Intermittent Possible blockage message during QC runs.
Yes | 15 | Y | RAME | D COHB CONTROL OUT HIGH CONTINUED |  | PREVIOUS SWO ESTABLISHED HEMOLYZER AND SPECTROMETER NECESSARY TO RESTORE RESULTS | D25F291FCC4FCEF1 | [5 empty] | D COHB CONTROL OUT HIGH CONTINUED |  | PREVIOUS SWO ESTABLISHED HEMOLYZER AND SPECTROMETER NECESSARY TO RESTORE RESULTS
Yes | 7 | Y | RAME | D COHB CONTROL OUT HIGH CONTINUED FROM SWO BABV3JAA7 ON ALTERNET INSTRUMENT |  | SAME AS SWO BABV3JAA7 FOR COHB HIGH ON CONTROL BLOOD | D25F2A90CC5CCEF6 | [5 empty] | D COHB CONTROL OUT HIGH CONTINUED FROM SWO BABV3JAA7 ON ALTERNET INSTRUMENT |  | SAME AS SWO BABV3JAA7 FOR COHB HIGH ON CONTROL BLOOD
No | Visit | Complaint | RIND | QC Failure for O2 and CL, Replaced Signal Data Board, Sensor Board and perform zeroing, Passed QC and two ppoint calibration report, Working ok, problem resolve. | QC Failure. | QC Failure for O2 and CL, Replaced Signal Data Board, Sensor Board and perform zeroing, Passed QC and two point calibration report, Working ok, problem resolve. | SWO-184534 | [5 empty] | QC Failure for O2 and CL, Replaced Signal Data Board, Sensor Board and perform zeroing, Passed QC and two ppoint calibration report, Working ok, problem resolve. | QC Failure. | QC Failure for O2 and CL, Replaced Signal Data Board, Sensor Board and perform zeroing, Passed QC and two point calibration report, Working ok, problem resolve.
No | 2 | Y | RAME | P Customer states fluid blockage occuring in unit. |  | Upon arrival confirmed fluid blockage errors. | D25F3A1BCC2FFEAF | [5 empty] | P Customer states fluid blockage occuring in unit. |  | Upon arrival confirmed fluid blockage errors.
No | Visit | Complaint | RCHN | GLU定标失败 | GLU定标失败 | 更换测试卡后修复 | SWO-183631 | Assisted Replacement | SC | SC | Glu |  | GLU calibration failed | GLU calibration failed | Repair after replacing test card
No | Visit | Complaint | RCHN | 孵育盘感应器坏 5685 | 孵育盘感应器坏 5685，借903-138一个。 | 更换903-138后修复(By Ran) | SWO-183378 | Defect | HW | Incubator Module | PCB, Incubator | 5685 | Incubation plate sensor bad 5685 | Incubation disk sensor bad 5685, borrow 903-138 one. | Repair after replacement 903-138 (By Ran)
No | 5 | Y | RPNZ | Blank | K calibration failures |  | D260E430424DDBE0 | Assisted Replacement | Electrode | Electrode | K+ |  | Blank | K calibration failures
Yes | 10 | Y | RAME | p rinse leaking from front |  | found tubing crimped #5 line, hole in solution pump tubing,glu missing o ring, ph not sealing, had impeded probe,rinse errors. blood sample tested with clotted instrument | D25F4B32CC0F05A8 | [5 empty] | p rinse leaking from front |  | found tubing crimped #5 line, hole in solution pump tubing,glu missing o ring, ph not sealing, had impeded probe,rinse errors. blood sample tested with clotted instrument
No | 13 | Y | RPNZ | Blank | Investigate pCO2 QC low, pH issues and bilirubin results |  | D260EA0F423DDBE2 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway | pCO2 | Blank | Investigate pCO2 QC low, pH issues and bilirubin results
No | 6 | Y | RAME | F Customer called with 581 oxi errors |  | Checked all logs to verify past performance and found 581 errors on QC and patient samples. Last THB cal 1.0277 | D25F2390CC4C70B2 | Defect | HW | Oxi module | Hemolyzer |  | F Customer called with 581 oxi errors |  | Checked all logs to verify past performance and found 581 errors on QC and patient samples. Last THB cal 1.0277
No | Visit | Complaint | RDEU | Analysator hat in der Vergangenheit mehrere Fehlermeldungen angezeigt, wie z. B. Probenabbruch, 1067, oder 5684 rotate process Wheel und 1223 + 1265 Hct Zelle nicht leer. | Überprüfung Analysator auf Fehlermeldung im System und unter Aktivitäten. Überprüfung auch unter Patientenmessungen. Mehrmals Probenabbruch mit Meldung 1067, und Einmal mit Meldung Hct Fehler. | Proteinreinigung nach Service Manuel durchgeführt, Systemreinigung durchgeführt. Reinigung des Washblocks, der Pipette und des Pipetteblocks. Reinigung des vorderen Filters. Reinigung der Cup Wash unit. Schläuche Pumpen P2 und P4 unterschiedlich lang und Abrieb an Achsen. Pumpenköpfe getauscht. Blankmessungen durchgeführt, Messwerte innerhalb Grenzwerte, in Ordnung. Flowwerte überprüft, sehr gut, Werte liegen zwischen 1100 und 1200. | SWO-183837 | Cleaned | HW | Aspiration Module | Needle Wash Unit | 1223 | Analyzer has displayed several error messages in the past, such as sample cancellation, 1067, or 5684 rotate process wheel and 1223 + 1265 Hct cell not empty. | Check analyzer for error message in the system and under Activities. Check also under patient measurements. Multiple sample abort with message 1067, and Once with message Hct error. | Protein cleaning carried out according to Service Manuel, system cleaning performed. Cleaning of the washblock, the pipette and the pipette block. Cleaning the front filter. Cleaning the Cup Wash unit. Hoses pumps P2 and P4 different length and abrasion on axes. pump heads. Blank measurements performed, measurements within limit values, fine. Flow values checked, very well, values are between 1100 and 1200.
No | Visit | Complaint | RDEU | Set Save 22, Rotate Process Wheel 40 | Reinigungsarbeiten, Fehlersuche | Prozeßrad nicht drehbar, optische Einheit demontiert, Verklemmtes Testcup entfernt, Montage, Neustart ok, SafePositions ok, Startphase, Systemüberprüfung ok, QK TnI ok | SWO-183887 | Defect | Consumables | SP | Cup Stack |  | Set Save 22, Rotate Process Wheel 40 | Cleaning work, troubleshooting | Process wheel not rotatable, optical unit disassembled, clamped test cup removed, assembly, restart ok, SafePositions ok, start phase, system check ok, QK TnI ok
No | 8 | Y | RAME | Z Customer reports intermittent printer errors |  | Checked all logs to verify performance and found activities when staff had printer issues and restarted and $or opened and closed the printer lid. | D260A6D34209BDF5 | Defect | HW | Printer Module | Printer Module |  | Z Customer reports intermittent printer errors |  | Checked all logs to verify performance and found activities when staff had printer issues and restarted and $or opened and closed the printer lid.
No | 6 | Y | RAME | 3 pO2 keeps failing qualicheck qc. |  | Kathy is running qc studies and, pO2 is intermittently failing on each one of the external qc. She replaced the sensor cassette (Part #944-005, Lot #966) on 4 $8 $17 for the same problem. | D260BA6642513143 | Defect | HW | Basic Unit | PCB, Sensor interface |  | 3 pO2 keeps failing qualicheck qc. |  | Kathy is running qc studies and, pO2 is intermittently failing on each one of the external qc. She replaced the sensor cassette (Part #944-005, Lot #966) on 4 $8 $17 for the same problem.
No | Visit | Complaint | RDEU | 10.04.2017; Fr. Schmitt; Tel.: 065168342215; Probe verloren | Einlassmodul ersetzt. | Startphase QC und Kalibration. | SWO-183915 | Defect | HW | Auto Inlet module | Auto Inlet module |  | 10.04.2017; Ms. Schmitt; Tel.: 065168342215; Sample lost | inlet module. | Start phase QC and calibration.
Yes | Visit | Complaint | RDEU | 10.04.2017; Fr. Schmitt; Tel.: 065168342215; Probe verloren | E-Teil eingetragen | E-Teil eingetragen | SWO-187361 | Re classified RMED | Double entry |  |  |  | 10.04.2017; Ms. Schmitt; Tel.: 065168342215; Sample lost | E-part entered | E-part entered
No | 23 | Y | RAME | 2 Pco2 internal/external QC problems |  | Normally the level one QC for Pco2 runs high on both units. When internal Qc fails it is either the Pco2 low on 8302 or the Po2 high on 8303 internal Qc. Ex: Pco2 68.6 too low or P02 92 too high. This happens with many different lots of both SC $SP. The room temperatures run between 66-68 F-No major air draft noticed. On this unit I have replaced Sensor connector $ Analog $ Spine $Manifold with no change in its operation. | D2605FF9422C7EAF | [5 empty] | 2 Pco2 internal/external QC problems |  | Normally the level one QC for Pco2 runs high on both units. When internal Qc fails it is either the Pco2 low on 8302 or the Po2 high on 8303 internal Qc. Ex: Pco2 68.6 too low or P02 92 too high. This happens with many different lots of both SC $SP. The room temperatures run between 66-68 F-No major air draft noticed. On this unit I have replaced Sensor connector $ Analog $ Spine $Manifold with no change in its operation.
Yes | Visit | Complaint | RCHN | GLU和LAC定标失败 | GLU和LAC定标失败 | 更换测试卡后修复 | SWO-184017 | [5 empty] | GLU and LAC bid failed | GLU and LAC bid failed | Repair after replacing test card
Yes | Visit | Complaint | RCHN | 测试卡参比电极错误 | 测试卡参比电极错误 | 更换测试卡后修复 | SWO-184154 | [5 empty] | Test card reference electrode error | Test card reference electrode error | Repair after replacing test card
No | Visit | Complaint | RCHN | 测试卡参比电极错误 1361 | 测试卡参比电极错误 1361 | 更换测试卡正常。 | SWO-186650 | Assisted Replacement | SC | SC | Reference | 1361 | Test card reference electrode error 1361 | Test card reference electrode error 1361 | The replacement test card is OK.
Yes | Visit | Complaint | RCHN | 参比电极错误，GLU和LAC定标失败 | 参比电极错误，GLU和LAC定标失败 | 连续更换2张测试卡后修复 | SWO-184015 | [5 empty] | Reference electrode error, GLU and LAC calibration failed | Reference electrode error, GLU and LAC calibration failed | Repair after 2 consecutive test cards
No | Visit | Complaint | RKK | pCO2 QC測定値 変動 | pCO2 QC 測定値変動 | 配管チューブ交換(リンスライン) タンパク除去、LS調整・ポンプCAL確認 エアライン清掃 CAL GAS2交換 CAL・QC測定データ確認 | SWO-184147 | Assisted Replacement | Consumables | Gas cylinder | Gas2 |  | pCO2 QC readings Variation | pCO2 QC Measurement variation | Pipe tube replacement (rinse line) Protein removal, LS adjustment and pump CAL confirmation Airline Cleaning CAL GAS2 replacement Cal/QC measurement data confirmation
No | Visit | Complaint | RKK | pCO2 QC測定値 変動 | pCO2 QC 測定値 変動 | 配管チューブ交換(リンスライン) コンタミ除去、LS調整・ポンプCAL確認 | SWO-184161 | Assisted Replacement | Consumables | Gas cylinder | Gas2 |  | pCO2 QC readings Variation | pCO2 QC Readings Variation | Pipe tube replacement (rinse line) Contamination removal, LS adjustment and pump CAL confirmation
Yes | Visit | Complaint | RDEU | 12.04.2017 09:50Uhr Herr Alsbaki 06142881033: Oximetriefehler bei der Kalibration 251 |  |  | SWO-184094 | Re classified RMED | PM |  |  |  | 12.04.2017 09:50H Mr Alsbaki 06142881033: Oximetry error in calibration 251
No | 6 | Y | RAME | 9 Intermittent Lactate drift during cals. |  | Checked data cal logs for drift errors, found several lactate cal errors. Found a crack in the ca++ measuring well. Checked grounding strap for modules. Following power up all calibration cleared following the third cal. QC and samples were run, the ca++ results for the samples read lower when compared to the other analyzer.Performed a cal2, the sens on ca++ went from 99% to 86% and lactate drift was at .3. After several cals wasn't able to clear errors untill cleaning connectors. | D266FF80901B58B8 | Defect | HW | El/Met Module | Measuring chamber, Lac | Cracked | 9 Intermittent Lactate drift during cals. |  | Checked data cal logs for drift errors, found several lactate cal errors. Found a crack in the ca++ measuring well. Checked grounding strap for modules. Following power up all calibration cleared following the third cal. QC and samples were run, the ca++ results for the samples read lower when compared to the other analyzer. Performed a cal2, the sens on ca++ went from 99% to 86% and lactate drift was at .3. After several cals wasn't able to clear errors untill cleaning connectors.
No | 2 | Y | RAME | 0 tHb |  | Neon Intensity too low. | D260A5D54210CD73 | Defect | HW | Oxi module | Lamp unit | Neon | 0 tHb |  | Neon Intensity too low.
No | Visit | Complaint | RDEU | 13.04.2017 9:15 Fr. Kirchner 0201650568154: Uhrzeit ist verstellt | Gerät steht noch auf Winterzeit. Ebenso der AQT S/N 838R0418N014. | Der Time-Server stellt nicht automatisch von Winter- auf Sommerzeit um, daher entsprechende Einstellung im Betriebssystem vorgenommen. Nach Geräteneustart bleibt die korrekte Zeit hinterlegt. Einstellung analog am AQT S $N 838R0418N014 vorgenommen. | SWO-184446 | Setup | SW | Time setting | Time setting | Summer time | 13.04.2017 9:15 Ms. Kirchner 0201650568154: Time is adjusted | device is still on winter time. Also the AQT S/N 838R0418N014. | The Time server does not automatically switch from winter to daylight saving time, so the appropriate setting is made in the operating system. After the device is restarted, the correct time is stored. Setting analogously to the AQT S $N 838R0418N014.
No | Visit | Complaint | RDEU | Leckage erkannt | LP erneuert, Analysator gespült, Visual Flow Test und Backflush Prozedur durchgeführt. Hemolysator undicht! | Hemolysator erneuert, Ventilschlauch getauscht, tHb Kalibration, Systemkontrolle. | SWO-184343 | Revisit required | HW | Oxi module | Hemolyzer | Leakage | Leakage detected | LP renewed, analyzer flushed, Visual Flow test and backflush procedure performed. Hemolysator leaks! | Hemolysator renewed, valve hose exchanged, tHb calibration, system control.
No | Visit | Complaint | RDEU | Leckage erkannt | Inhomogene Kal 3 Lösung | LP erneuert. Startphase, Kal, QC und Blutmessung iO. | SWO-184415 | Assisted Replacement | SP | SP | FT | Leakage | Leakage detected | Inhomogeneous Kal 3 Solution | LP renewed. Start phase, Kal, QC and blood measurement iO.
No | 15 | Y | RAME | P Liquid Sensor Calibration Error. |  | Liquid sensors 1 and 2 are failing between the sensor cassette which are part of the bay unit. The values for LS1 and LS2 are 245 high out of range which could indicate and electrical malfunction of the sensors. | D261A402B8046F75 | Defect | HW | Basic Unit | Sensor Cassette Bay |  | P Liquid Sensor Calibration Error. |  | Liquid sensors 1 and 2 are failing between the sensor cassette which are part of the bay unit. The values for LS1 and LS2 are 245 high out of range which could indicate and electrical malfunction of the sensors.
No | 3 | Y | RAME | P Fluid transport error. |  | Found the analyzer is asking for back-flush. It is suspecting a clot. | D261EC01B85457BB | Defect | HW | Basic Unit | Sensor Cassette Bay |  | P Fluid transport error. |  | Found the analyzer is asking for back-flush. It is suspecting a clot.
Yes | Visit | Complaint | RDEU | Alle Lyte bei QK fehlerhaft, Kal ok | Reinigungsarbeiten, Fehlersuche | Kal, QK fehlerhaft, SP90 vor zwei Tagen neu, Sensorkasette ersetzt, Startphase | SWO-184368 | Re classified RMED | SC | SC | Reference |  | Alle Lyte bei QK fehlerhaft, Kal ok | Cleaning work, troubleshooting | Kal, QK faulty, SP90 new two days ago, sensor calet replaced, start-up phase
No | Visit | Complaint | RDEU | Alle Lyte bei QK fehlerhaft, Kal ok | Siehe SWO 184368 | Siehe SWO 184368 Beide Unterschriften zu SWO 184368 & zu SWO 184404 | SWO-184367 | Assisted Replacement | SC | SC | Reference |  | Alle Lyte bei QK fehlerhaft, Kal ok | See SWO 184368 | See SWO 184368 Both signatures for SWO 184368 & SWO 184404
No | Visit | Complaint | RDEU | 12.04.17 , 12:25 Herr Nasirie Tel. 0401818879728 , Messabbrüche und diverse Fehler | Diverse Messabbrüche. | Im Datei Archive nachgeschaut. Die Fehlermeldungen sind keine Probe erkannt. Den Analysator im Service Menü durchgecheckt. Keine Meldungen vorhanden. | SWO-184258 | No fault found | No fault found | No fault found | No fault found |  | 12.04.17 , 12:25 Mr. Nasirie Tel. 040181879728 , Discontinuations and various errors | Several Messabbriche. | Looked in the Archive file. The error messages are not detected as a sample. Checked the analyzer in the service menu. There are no messages.
No | Visit | Complaint | RDEU | Lösungspack verweigert Installation | Lösungspack verweigert Installation. Inlet Gasket , Inlet Probe und Connector für Kundenbestand überlassen. Bitte berechnen! | Mit MED. Techniker Herr Leistner Lösungspack erneut installiert. Lösungspack war nicht vollständig nach unten gedrückt. Einlassdichtung streifte an LP und blockierte die korrekte Positionierung. Einlassdichtung gewechselt und gleichen LP wieder installiert. Gerät durchgespült, tHb Kal. fällig und durchgeführt . Reklamationsverfahren erklärt. | SWO-184422 | User error | SP | SP Activation | SP Activation |  | Solution pack denies installation | Solution pack denies installation. Inlet Gasket , Inlet Probe and Connector for customer inventory. Please calculate! | With MED. Technician Mr. Leistner solution pack reinstalled. Solution pack was not fully pushed down. Inlet seal swiped at LP and blocked the correct positioning. inlet seal changed and reinstalled the same LP. device flushed through, tHb Kal. due and carried out . complaint procedure.
No | Visit | Complaint | RCHN | 接口板被腐蚀 | 接口板被腐蚀，借备件902-972一个 | 更换902-972一个正常。(By Ran) | SWO-183938 | [5 empty] | Interface plate is corroded | Interface board is corroded, spare parts 902-972 one | Replace 902-972 one normal. (By Ran)
No | Visit | Complaint | RCHN | 测试卡漏液 | 测试卡漏液，借902-972一个 | 更换902-972后修复(By Ran) | SWO-183864 | [5 empty] | Test card leakage | Test card leak, borrow 902-972 one | Repair after replacement 902-972 (By Ran)
Yes | Visit | Complaint | RIND | Temp-D and Temp-B error. | Sensor Cassette Temp-D and Temp-B error. | Found Temp A, B, D errors, Sensor cassette got leaked solution, sensor connector got damaged, replaced the standby Sensor connector PCB, performed the Zeroing, installed the new SC found all the parameters are working, performed the test found results are Ok, instrument working satisfactorily. Defective SC details: Lot: 234692, Sn: 1562492. New SC details: Lot: 235032, Sn: 1566297 | SWO-184179 | [5 empty] | Temp-D and Temp-B error. | Sensor Cassette Temp-D and Temp-B error. | Found Temp A, B, D errors, Sensor cassette got leaked solution, sensor connector got damaged, replaced the standby Sensor connector PCB, performed the Zeroing, installed the new SC found all the parameters are working, performed the test found results are Ok, instrument working satisfactorily. Defective SC details: Lot: 234692, Sn: 1562492. New SC details: Lot: 235032, Sn: 1566297
No | Visit | Complaint | RCHN | 不均匀的定标3溶液1084 | 不均匀的定标3溶液1084 | 更换试剂包后修复 | SWO-184470 | Assisted Replacement | SP | SP | FT | 1084 | Uneven calibration 3 solution 1084 | Uneven calibration 3 solution 1084 | Repair after replacement of reagent kit
No | Visit | Complaint | RCHN | 进样口试管支架断裂 | 进样口试管支架断裂 ，借903-584一个。 | 更换备用机备件903-584后修复(By Ran) | SWO-184207 | Defect | HW | Aspiration Module | Tube holder |  | The test tube support in the inlet is broken | The test tube support in the inlet is broken , borrow edith of 903-584. | Repair after replacement of spare parts 903-584 (By Ran)
No | Visit | Complaint | RIND | crea calibration expire. | Check the system. check the logs. | Replace the Crea reagent & calibrate the system. | SWO-184571 | Assisted Replacement | Consumables | Consumables | Crea expired |  | crea calibration expire. | Check the system. check the logs. | Replace the Crea reagent & calibrate the system.
Yes | 12 | Y | RFRA | Blank | Fuite site d'introduction |  | D261B080B83C525C | [5 empty] | Blank | Leaked introductory site
No | 4 | Y | RCZE | Blank | hodnoty na Lac v QC mimo požadavek OKB, v našem limitu ale tìsnì u spodní hranice |  | D261E9BBB83EA2FC | Assisted Replacement | Electrode | Membrane | Lac | Sensitivity | Blank | Values to Lac in QC outside the OKB requirement, but in our limit, at the lower limit
No | 13 | Y | RPNZ | Blank | Grey Radiometer ABL800 Series screen. |  | D261F3D1B85DD03F | Defect | HW | DMS_PC | Touch screen |  | Blank | Grey Radiometer ABL800 Series screen.
No | Visit | Complaint | RSCH | Reparatur | Seit 10.04.2017 Lac QC Level 3 zu tief Über Nacht hatte Gerät 2. Kal Abgebrochen Keine Spüllösung während Kal. pH-Modul nicht innert Zeitrahmen gefüllt | Allgemeine Kontrolle i.o Membranenwartung Lac gleiche Lot R3306 durchgeführt i.o Kal. i.o QC Level 3 i.o Erneut Kal. i.o Alle QC i.o Erneut Kal. Empfindlichkeit schwankt QC Level 3 sinkend Lac Elektrode ersetzt mit neuer Lot Membrane aus Thalwil Alte Lac Elektrode YH03 = Garantiersatz i.o Kal. und QC i.o Letzte 2. Kal. wurde abgebrochen Gleicher Fehler wie in der Nacht Ventile pH $BG-Modul ausgebaut und gereinigt i.o Spüllösung ersetzt i.o Schlausch Spüllösung ersetzt i.o Zwei 2. Kal. i.o | SWO-184384 | Cleaned | HW | pH/BG Module | Valve |  | Repair | Since 10.04.2017 Lac QC Level 3 too deep Overnight had device 2nd Kal Aborted No rinsing solution during Kal. pH module not filled within timeframe | General control i.o Membrane Maintenance Lac Same Lot R3306 Performed i.o Kal. Ok QC Level 3 i.o Once again Kal. Ok All QC i.o Once again Kal. Sensitivity fluctuates QC Level 3 sinking Lac electrode replaced with new solder membrane from Thalwil Old Lac Electrode YH03 = Warranty Kit i.o Kal. and QC i.o Last 2nd Kal. has been canceled Same mistake as at night Valves pH $BG module removed and cleaned i.o Rinsing solution replaces i.o Schlausch rinsing solution replaces i.o Two 2nd Kal. Ok
No | Visit | Complaint | RKK | ラディアンスとの接続が出来ないので引上げ、代替器械セット | 確認 | 機械修理後セットアップ、ラディアンスとの接続確認、測定確認 | SWO-184453 | Setup | His/lis | His/lis | His/lis |  | Because it is not possible to connect with radiance, it raises, and an alternative instrument set | Check | Machine repair setup, connection confirmation with radiance, measurement confirmation
No | Visit | Complaint | RDEU | Set Save 22, Rotate Process Wheel 40 13.04.2017; Sr. Gabi; Tel.: 03303522870; wieder Fehlermeldung: Setsafe_22 | Fehler Sétsave 22 | Reinigung, Funktionen von Storage-, Inkubator- und Prozesswheel kontrolliert, Testcups aus Prozessrad und Inkubatorrad entfernt, Startphase, Systemtest ok, TestQK, Setsave 22, erneut Testcup entfernt sowie Solutionpack kontrolliert und ersetzt, Startphase, Systemtest ok, TestQK ok | SWO-184404 | Defect | Consumables | SP | Cup Stack | Revisit | Set Save 22, Rotate Process Wheel 40 13.04.2017; Sr. Gabi; Tel.: 03303522870; wieder Fehlermeldung: Setsafe_22 | Bug Sétsave 22 | Cleaning, functions of storage, incubator and process wheel controlled, test cups removed from process wheel and incubator wheel, start phase, system test ok, TestQK, Setsave 22, re-removed test cup as well as solutionpack controlled and replaced, start phase, system test ok , TestQK ok
No | 15 | Y | RPAC | Blank | Computer failure |  | D262D5332E332E4D | Revisit required | HW | DMS_PC | CPU |  | Blank | Computer failure
No | 3 | Y | RPOL | Blank | Na dwóch kasetach identyczne problemy z parametrami. Po przep³ukaniu w koñcu zaskoczy³. |  | D268181F06340DC1 | Unknown | Unknown | Unknown | Unknown |  | Blank | On two cartridges identical problems with parameters. After the disc has been blown, it will surprise.
No | 2 | Y | RPOL | Blank | Wymiana odczynników ze wzglêdu na b³¹d w dacie analizatora. Testy kontrolne. Analizator sprawny. |  | D2715BF54041D70C | User error | SP Expired | SP Expired | SP Expired |  | Blank | Replacement of reagents due to error on the date of the analyser. Control tests. Functional Analyzer.
No | 2 | Y | RPOL | Blank | W pozycji kapilary aparat nie odbija³ uszczelki po odsuniêciu kapilary. W efekcie aparat domaga³ siê usuniêcia próbki i po chwili porzuca³ pomiar próbki. Wymiana ig³y i uszczelki wlotu. Szkolenie u¿ytkownika. Analizator sprawny. |  | D262F4622E0CE2F3 | Defect | HW | Auto Inlet module | Inlet Position Cal. |  | Blank | In the capillary position, the camera does not bounce the seals after removing the capillaries. As a result, the camera has requested that the sample be removed and the sample is discarded after a while. Replacing the inlet and gasket. User training. Functional Analyzer.
No | Visit | Complaint | RCHN | 报错 5257 | 报错 5257，借636-550一个。 | 自行更换备用机备件636-550后修复(By Ran) | SWO-184325 | Defect | HW | Process Module | Process Unit sensor | 5257 | Error 5257 | 5257, borrow 636-550 one. | Self-replacement of spare parts 636-550 after repair (By Ran)
No | Visit | Complaint | RCHN | 报错1024，1025 | 报错1024，1025 | 更换903-364正常。(By Ran) | SWO-184478 | Defect | HW | Process Module | Optical Unit | 1024 | Error 1024,1025 | Error 1024,1025 | Replacement of 903-364 is normal. (By Ran)
No | Visit | Complaint | RCHN | HCT不能被正确测量 | HCT不能被正确测量 ，借903-329一个。 | 更换903-329一个正常。(By Ran) | SWO-184292 | Defect | HW | Aspiration Module | Needle Wash Unit |  | HCT cannot be measured correctly | HCT cannot be measured correctly , borrow edith of 903-329. | Replace the 903-329 one normal. (By Ran)
No | Visit | Complaint | RKK | 測定中のpH応答不良 | 測定中のpH応答不良 | pH電極交換 | SWO-184455 | Assisted Replacement | Electrode | Electrode | pH |  | Poor pH response during measurement | Poor pH response during measurement | pH electrode replacement
No | 2 | Y | RFRA | Blank | Erreur 0090 : Valeur MMC1 en dessous de la gamme pour Nt ProBNP |  | D262E0822E4A03E6 | Defect | HW | Process Module | Optical Unit | 1024 | Blank | Error 0090: MMC1 value below the range for Nt ProBNP
No | Visit | Complaint | RDEU | Leckage erkannt | Wochenendeinsatz laut Frau Paul zwingend erforderlich Nacharbeit zu SWO-184415 | Fehlersuche, laut Frau Paul QC grenzwertig seit letzten Serviceeinsatz und jetzt außer Bereich bei pO2 Level C und erneut Leckage Fehler->Überprüfung des internen Flüssigkeitstransportweges von Abfall bis Einlass, Sensor Bay geöffnet und LP herausgenommen->erneutes Einsetzen des LP sehr schwergängig->mittlere Kunststoffführung des LP am Analysator sitzt nicht korrekt->diese Führung wieder festgezogen, Leckagetest bestanden Startphase bestanden->QC bei pO2 in allen Leveln Zielwert | SWO-184685 | Adjustment | HW | Basic Unit | SP Guide |  | Leakage detected | Weekend assignment mandatory according to Ms Paul Rework on SWO-184415 | Troubleshooting, according to Ms Paul QC borderline since last service use and now out of range at pO2 Level C and again leakage error->checking of the internal liquid transport path from waste to inlet, Sensor Bay opened and LP removed->again Inserting the LP very difficult-to-use->medium plastic guide of the LP on the analyzer sits incorrectly->this guide reattached, leakage test passed Start phase passed->QC at pO2 at all levels target value
No | 6 | Y | RAME | 3 pO2 calibration error. |  | Upon arrival, found the analyzer presenting error 964 Leak current in relation to Solution Pack detected. Also found, pO2 calibrations failing with error 1002- pO2 dark data is out of range. | D264398EA45392B0 | Defect | HW | Basic Unit | PCB, Sensor interface |  | 3 pO2 calibration error. |  | Upon arrival, found the analyzer presenting error 964 Leak current in relation to Solution Pack detected. Also found, pO2 calibrations failing with error 1002- pO2 dark data is out of range.
No | Visit | Complaint | RSAF | Failing Qc on po2 level B even after replacement of Solution pack. | Done complete service test on consumables. | Replaced bent probe. Analyzer calibrated and passed all Qc's | SWO-184844 | Unknown | HW | Auto Inlet module | Probe bend | ext. O2 Leaking | Failing Qc on po2 level B even after replacement of Solution pack. | Done complete service test on consumables. | Replaced bent probe. Analyzer calibrated and passed all Qc's
No | Visit | Complaint | RDEU | Einlass nicht selbständig geschlossen, seit dem hängt das gerät in der Erneuerungsschleife. | Analysator am Vortag schon wieder Messbereit, Lac. QC Fehler nach Startphase, Kal. und QC seit 20 Uhr grün, bei Testmessung mit Erwachsenen Blut werden Probenfehler berichtet. Kundin besteht auf Überprüfung des Gerätes am Samstag. Fr. Dr. Gehrken ( diensthabende Ärztin ) hat einen Wochenendeinsatz beauftragt. Kunde hat angeblich kein ZVBM mehr auf Lager. | ZVBM aus Kl. Detmold geholt. Analysator geprüft, keinen Fehler festgestellt. Kapillar Modus standardmäßig auf 65ul gestellt, Kundin hat bei Probemessung die 65ul Kapillare nicht komplett gefüllt. Testmessung im 45ul Modus durchgeführt, Messung i.O. 45ul als Standart Modus eingestellt. Ersatz SC90 noch im Kühlschrank, Lösungspack und Einlassdichtung als Ersatz vor Ort gelassen. | SWO-184768 | User error | Sample volume | Sample volume | Sample volume |  | Entrance not closed independently, since then the device hangs in the renewal loop. | Analyzer the day before already ready to measure again, Lac. QC error after start phase, Kal. and QC since 8 p.m. green, in test measurement with adult blood sample errors are reported. Customer insists on checking the device on Saturday. Ms. Dr. Gehrken (on duty doctor ) has commissioned a weekend assignment. Customer reportedly no longer has a ZVBM in stock. | ZVBM from Kl. Detmold. Analyzer checked, no error detected. Capillary mode set to 65ul by default, customer did not completely fill the 65ul capillary when sampled. Test measurement performed in 45ul mode, measurement i.O. 45ul set as standard mode. Replacement SC90 still left in the refrigerator, solution pack and inlet seal as a replacement on site.
No | Visit | Complaint | RCHN | LAC结果偏高 | LAC结果偏高 | 更换测试卡 | SWO-184635 | [5 empty] | LAC results are high | LAC results are high | Replace the test card
No | Visit | Complaint | RCHN | 光学强度超范围 1031 | 光学强度超范围 1031，借903-364一个。 | 更换903-364后修复(By Ran) | SWO-184481 | Defect | HW | Process Module | Optical Unit | 1031 | Optical strength over range 1031 | Optical strength over range 1031, borrowed 903-364 one. | Repair after replacement 903-364 (By Ran)
No | Visit | Complaint | RLTD | Leak detected. | On arrival the analyser was in operator intervention required and the analyser was reporting that a leak was detected, the analyser was suggesting that Radiometer should be contacted. On arrival to the analyser it was not possible to open the inlet, it appeared as if the inlet gasket was being trapped on the solution pack. The analyser failed calibrations and QC on start-up on pCO2, Ca, Glu and Lac. The calibrations sensitivity and status values were very poor and not close to being within limits. Looking through the data logs the analyser has not been working since the 13th April. This is a chargeable call. Labour, 1 hour - £129.00 VAT - £25.80 Total - £154.80 | The inlet gasket was lifted slightly and this then allowed the inlet to open. The inlet gasket was pressed along the probe, (as if introducing a sample) and an audible click was heard as the gasket sat home. The inlet now opened and closed acoordingly. The fluid transport tests were ran in the service program and all passed. The analyser was allowed to restart. The Sensor cassette was replaced due to the calibration and QC failures especially as solution has not been pumped through the analyser since the 13th April. The SC was left to start-up and stabilise. | SWO-184607 | Adjustment | HW | Inlet module | Inlet Gasket |  | Leak detected. | On arrival the analyser was in operator intervention required and the analyser was reporting that a leak was detected, the analyser was suggesting that Radiometer should be contacted. On arrival to the analyser it was not possible to open the inlet, it appeared as if the inlet gasket was being trapped on the solution pack. The analyser failed calibrations and QC on start-up on pCO2, Ca, Glu and Lac. The calibrations sensitivity and status values were very poor and not close to being within limits. Looking through the data logs the analyser has not been working since the 13th April. This is a chargeable call. Labour, 1 hour - £129.00 VAT - £25.80 Total - £154.80 | The inlet gasket was lifted slightly and this then allowed the inlet to open. The inlet gasket was pressed along the probe, (as if introducing a sample) and an audible click was heard as the gasket sat home. The inlet now opened and closed acoordingly. The fluid transport tests were ran in the service program and all passed. The analyser was allowed to restart. The Sensor cassette was replaced due to the calibration and QC failures especially as solution has not been pumped through the analyser since the 13th April. The SC was left to start-up and stabilise.
No | Visit | Complaint | RLTD | Analyser stuck in OIR loop, no inlet gasket available. | On arrival the analyser the analyser was in operator intervention required and the analyser was requesting the inlet gasket to be replaced, the user was unable to locate a spare inlet gasket. This is a chargeable call as it is out of hours. Call out - £317.00 Labour, 0.5 hour - £64.50 924-816 Inlet gasket - £99.40 Sub total of £480.90 VAT - £96.18 Total - £577.08 | Inlet gasket removed and inspected, the gasket appeared to be deformed. New inlet gasket fitted to the analyser and the analyser restarted passing calibrations and QC. | SWO-184608 | User error | HW | Inlet module | Inlet Gasket |  | Analyser stuck in OIR loop, no inlet gasket available. | On arrival the analyser the analyser was in operator intervention required and the analyser was requesting the inlet gasket to be replaced, the user was unable to locate a spare inlet gasket. This is a chargeable call as it is out of hours. Call out - £317.00 Labour, 0.5 hour - £64.50 924-816 Inlet gasket - £99.40 Sub total of £480.90 VAT - £96.18 Total - £577.08 | Inlet gasket removed and inspected, the gasket appeared to be deformed. New inlet gasket fitted to the analyser and the analyser restarted passing calibrations and QC.
No | Visit | Complaint | RCHN | LAC定标错误，值偏高 | LAC定标错误，值偏高 | 更换测试卡后修复 | SWO-184618 | [5 empty] | LAC calibration error, high value | LAC calibration error, high value | Repair after replacing test card
No | Visit | Complaint | RCHN | LAC定标失败 | LAC定标失败 | 更换测试卡 | SWO-184636 | Assisted Replacement | SC | SC | Lac |  | LAC calibration failure | LAC calibration failure | Replace the test card
No | Visit | Complaint | RCHN | 报错4794 | 报错4794，借备件636-549一个 | 更换636-549一个正常。 by andrew | SWO-184603 | Defect | HW | Aspiration Module | Needle Sensor | 4794 | 4794 error | 4794, borrowed spare parts 636-549 one | Replace 636-549 one normal. by andrew
No | 15 | Y | RAME | 0 tHb out of limits on verification material. |  | Cals, QC and tHb check are all at or near mean values. | D266E46E9055D1C3 | Setup | Sample volume | Sample volume | Sample volume |  | 0 tHb out of limits on verification material. |  | Cals, QC and tHb check are all at or near mean values.
No | 12 | Y | RAME | 0 tHb out of limits on level 3 of verification material. |  | QC and tHb calibration and cal are all on mean values. | D266B0E3905AD1BF | Setup | Sample volume | Sample volume | Sample volume |  | 0 tHb out of limits on level 3 of verification material. |  | QC and tHb calibration and cal are all on mean values.
No | 6 | Y | RAME | F Customer reported oximetry errors on QC. |  | Checked data logs. Found 581 failures and high MetHb values on QC. Calibration also failing with tHB drifting intermittently. | D266CC849024CE00 | Adjustment | HW | Oxi module | Lamp unit | Halogen | F Customer reported oximetry errors on QC. |  | Checked data logs. Found 581 failures and high MetHb values on QC. Calibration also failing with tHB drifting intermittently.
No | Visit | Complaint | RIND | after checking found that Touch screen not working, Done touch screen calibration, Restart the analyzer, problem resolve. | Touch screen not working. | Done touch screen calibration, Restart the analyzer, problem resolve. | SWO-186894 | Defect | HW | DMS_PC | Touch screen |  | after checking found that Touch screen not working, Done touch screen calibration, Restart the analyzer, problem resolve. | Touch screen not working. | Done touch screen calibration, Restart the analyzer, problem resolve.
No | Visit | Complaint | RIND | Analyzer Showing error calibrate39, After checking found that measuring cup is stuck in incubator wheel, remove incubator wheel and discard used cup from wheel, cleaned interior dust and calibrate incubator wheel, process wheel and punch out, problem resolve. | Analyzer Showing error calibrate39, After checking found that measuring cup is stuck in incubator wheel. | Remove incubator wheel and discard used cup from wheel, cleaned interior dust and calibrate incubator wheel, process wheel and punch out, problem resolve. | SWO-186895 | Cleaned | HW | Incubator Module | Cup Sensor | calibrate39 | Analyzer Showing error calibrate39, After checking found that measuring cup is stuck in incubator wheel, remove incubator wheel and discard used cup from wheel, cleaned interior dust and calibrate incubator wheel, process wheel and punch out, problem resolve. | Analyzer Showing error calibrate39, After checking found that measuring cup is stuck in incubator wheel. | Remove incubator wheel and discard used cup from wheel, cleaned interior dust and calibrate incubator wheel, process wheel and punch out, problem resolve.
No | Visit | Complaint | RIND | Cartridge not installing and stuck. | Cartridge stuck and not installing. | Found storage wheel stuck and not moving. Analyser opened and cleaned for storage wheel, Punch-out sensor and incubatore wheel. Analyser serviced and calibrated properly. handed over in afine working condition. | SWO-184922 | Cleaned | HW | Storage Module | Storage wheel sensor |  | Cartridge not installing and stuck. | Cartridge stuck and not installing. | Found storage wheel stuck and not moving. Analyser opened and cleaned for storage wheel, Punch-out sensor and incubatore wheel. Analyser serviced and calibrated properly. handed over in afine working condition.
Yes | 4 | Y | RAME | Z - MISC |  | ANNUAL PM DUE | D267F97F063DF378 | [5 empty] | Z - MISC |  | ANNUAL PM DUE
Yes | 6 | Y | RPAC | Blank | Sensor cassette failure (1361) - Internal reference electrode error in sensor cassette. |  | D2688B25060CEA3E | [5 empty] | Blank | Sensor cassette failure (1361) - Internal reference electrode error in sensor cassette.
No | Visit | Complaint | RLTD | The analyser is in the OIR loop and will not pass when retested. Please call Stephen Gardener - 01162585530 | Leak detected. | Split in pinch valve tubing, replaced tubing, now ok. | SWO-184793 | Defect | HW | Oxi module | Pinch valve tube |  | The analyser is in the OIR loop and will not pass when retested. Please call Stephen Gardener - 01162585530 | Leak detected. | Split in pinch valve tubing, replaced tubing, now ok.
No | Visit | Complaint | RSAF | Keeps on saying intervention required after all trouble shooting was done and new solution pack installed and telephonic assistance done. | Last worked on the 17th at 11pm when solution pack ran empty. | Correctly inserted inter connecting Gasket and straightened slightly bent probe. | SWO-185159 | Defect | HW | Auto Inlet module | Connector Gasket |  | Keeps on saying intervention required after all trouble shooting was done and new solution pack installed and telephonic assistance done. | Last worked on the 17th at 11pm when solution pack ran empty. | Correctly inserted inter connecting Gasket and straightened slightly bent probe.
No | 19 | Y | RAME | 0 THB failing system cycle. The sensitivity is 0.00. |  | The tHb sensitivity is failing causing internal QC to be Not Calculated. | D2681F13064DC0EF | [5 empty] | 0 THB failing system cycle. The sensitivity is 0.00. |  | The tHb sensitivity is failing causing internal QC to be Not Calculated.
No | Visit | Complaint | RCHN | 温度不稳定 | 温度不稳定，借备件903-334一个 |  | SWO-184729 | [5 empty] | Temperature instability | Temperature instability, spare parts 903-334 one
No | Visit | Complaint | RIND | Temperature Malfunction. | Temperature malfunction during Cassette installation. | Replace the Sensor Cassette with the new one,Electrolyte parameter got passed but PO2,PCO2,pH failed,Run the Zeroing check ,O2 was not stabilized,Then replace the sensor board,again run the zeroing check ,it got passed successfully,then install the sensor cassette successfully. | SWO-185463 | [5 empty] | Temperature Malfunction. | Temperature malfunction during Cassette installation. | Replace the Sensor Cassette with the new one,Electrolyte parameter got passed but PO2,PCO2,pH failed,Run the Zeroing check ,O2 was not stabilized,Then replace the sensor board,again run the zeroing check ,it got passed successfully,then install the sensor cassette successfully.
No | Visit | Complaint | RCHN | 试剂包漏液 | 试剂包漏液 | 更换试剂包后修复 | SWO-185015 | Assisted Replacement | SP | SP | Leaking |  | Reagent bag leak | Reagent bag leak | Repair after replacement of reagent kit
No | Visit | Complaint | RCHN | 光学敏感性超范围 1031 | 光学敏感性超范围 1031 ，借903-364一个。 | 更换903-364一个正常。 by andrew | SWO-184633 | Defect | HW | Process Module | Optical Unit | 1031 | Optical sensitivity over range 1031 | Optical sensitivity over range 1031 , borrow edith of 903-364. | Replace 903-364 one normal. by andrew
No | Visit | Complaint | RCHN | 软件不匹配 1999，0002 | 软件不匹配 1999，0002,借903-198，903-215各一个。 | 更换903-198及903-215后修复 by andrew | SWO-184616 | Defect | HW | Analyzer Control Module | PCB, Main Board | 1999 | Software mismatch 1999,0002 | Software mismatch 1999,0002, borrowing 903-198,903-215 each. | Repair by Andrew after replacing 903-198 and 903-215
No | 5 | Y | RAME | Z Autocheck aborting. Carousel lifting to inlet position and spinning and then immmediatley aborting. |  | Checked carousel for debris. Checked to make sure customer had proper lot numbers installed etc.. | D267DE9606224D7B | Defect | HW | Autocheck module | FPC Cable |  | Z Autocheck aborting. Carousel lifting to inlet position and spinning and then immmediatley aborting. |  | Checked carousel for debris. Checked to make sure customer had proper lot numbers installed etc..
Yes | Visit | Complaint | RKK | ウィルススキャン | ウィルススキャン | ウィルススキャン | SWO-188701 | Re classified RMED | Virus scan |  |  |  | Virus Scan | Virus Scan | Virus Scan
No | Visit | Complaint | RDEU | Akku lädt sich nicht am Netz | Akku ausgetauscht, keine Besserung, Netzteil über kreuz getauscht mit Labor-Gerät, Fehler liegt am Netzteil | Netzteil wird bestellt | SWO-185285 | Revisit required | HW | Basic Unit | Power Supply |  | Battery does not charge on the mains | Battery replaced, no improvement, power supply cross-exchanged with laboratory device, fault lies on the power supply | Power supply is ordered
No | Visit | Complaint | RDEU | fortlaufende Spülung unterbrochen | Spülung unterbrochen | Fehlersuche; Fließwege gespült; Einlaßdichtung und -halter gew.; Kal. + QC i.O. | SWO-185185 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | continuous rinsing interrupted | Rinse interrupted | Troubleshooting; Flow paths rinsed; inlet seal and holder. Kal. + QC in the first place
No | Visit | Complaint | RSAF | Inlet stuck in open position. | Inlet stuck in open position , no movement . Message , remove foreign object from the inlet | Remove inlet module. Remove probe and holder. Check for obstruction .Assemble inlet module, Fit inlet module and perform inlet calibration and analyzer calibration. All OK. | SWO-185479 | Defect | HW | Auto Inlet module | Auto Inlet module | User error? | Inlet stuck in open position. | Inlet stuck in open position , no movement . Message , remove foreign object from the inlet | Remove inlet module. Remove probe and holder. Check for obstruction . Assemble inlet module, Fit inlet module and perform inlet calibration and analyzer calibration. All OK.
No | Visit | Complaint | RSAF | Failing Qc level B on pO2 only. | Failing on pO2 Qc level B only | Replaced all tubing's and inlet gasket but still fails Qc, Noticed $heard oxygen leaking from the cubical next to the analyzer and the entrance to the cubical in less than a meter away, located leak and solved the issue of oxygen leaking. Left analyzer to stabilize and informed ward sister in charge and lab tech to check on the analyzer if it corrected the pO2 Qc issue, spoke to Paul the tech from lab on 19 $04 $2017 at 10H00 to move analyzer to were there is no oxygen contamination in the air. Awaiting his feedback with regards to the move if it corrected the pO2 issue. | SWO-185221 | User error | Inviroment | Inviroment | ext. O2 Leaking | ext. O2 Leaking | Failing Qc level B on pO2 only. | Failing on pO2 Qc level B only | Replaced all tubing's and inlet gasket but still fails Qc, Noticed $heard oxygen leaking from the cubical next to the analyzer and the entrance to the cubical in less than a meter away, located leak and solved the issue of oxygen leaking. Left analyzer to stabilize and informed ward sister in charge and lab tech to check on the analyzer if it corrected the pO2 Qc issue, spoke to Paul the tech from lab on 19 $04 $2017 at 10H00 to move analyzer to were there is no oxygen contamination in the air. Awaiting his feedback with regards to the move if it corrected the pO2 issue.
No | 20 | Y | RESP | Valores de Na baixos Manutenção Preventiva |  | Verificação dos registos de dados. Passagem de 2 niveis de QC com resultados OK. Realização de gasimetria, o resultado de Na continua baixo. Execução de limpeza conforme procedimento. Verificação da impressora. Calibração e verificação com QC. | D2694AF67C21D4B8 | [5 empty] | Low Preventive Maintenance values |  | Verification of data records. Passage of 2 levels of QC with results OK. Achievement of Gasimetry, the result of Na is still low. Cleaning execution as per procedure. Check the printer. Calibration and verification with QC.
No | Visit | Complaint | RCHN | 接口板腐蚀 | 接口板腐蚀，借备件902-972一个 | 更换902-972一个正常。 by andrew | SWO-184658 | [5 empty] | Interface plate corrosion | Interface plate corrosion, spare parts 902-972 one | Replace 902-972 one normal. by andrew
No | Visit | Complaint | RCHN | 离子不稳定，感应板有腐蚀 | 离子不稳定，感应板有腐蚀，借备件902-972一个 |  | SWO-185042 | [5 empty] | Ions are unstable and the sensing plate is corroded | Ion instability, sensing plate corrosion, spare part 902-972 one
No | Visit | Complaint | RCHN | 试剂包废液口堵塞 | 试剂包废液口堵塞 | 更换试剂包后修复 | SWO-185361 | [5 empty] | Reagent pack waste mouth blockage | Reagent pack waste mouth blockage | Repair after replacement of reagent kit
No | 1 | Y | RESP | El equipo está con la pantalla en negro, se escucha un fuerte pitido y no trabaja. |  | (18-4-2016) Verificación del fallo (se cambia la F.A, comprobación de la pantalla, bateria y módulo PC (siendo este la causa de la avería). Se solicita un módulo nuevo. (19-4-2016) Desmontaje y sustitución del módulo PC, verificación del fungible reutilizable, calibración y verifación con ciclos de calidad. | D26947277C4B1A7B | [5 empty] | The computer is with the screen black, you hear a loud beep and it does not work. |  | (18-4-2016) Fault check (changes F.A., screen check, battery and PC module (this being the cause of the breakdown). A new module is requested. (19-4-2016) Disassembly and replacement of the PC module, reusable fungible verification, calibration and verification with quality cycles.
No | Visit | Complaint | RCHN | 无法开机 | 无法开机，借备件910-342一个 | 更换910-342一个正常。by andrew | SWO-184707 | Defect | HW | DMS_PC | CPU |  | Cannot power on | Cannot power on, borrow spare parts 910-342 one | Replace 910-342 one normal. by andrew
No | Visit | Complaint | RCHN | 漏电流 | 漏电流，借备件903-581一个 | 更改903-581后修复 by andrew | SWO-184662 | Defect | HW | Basic Unit | PCB, Sensor interface | 0964 | Leak current | Leak current, spare parts 903-581 one | Fix by andrew after changing 903-581
No | Visit | Complaint | RCHN | 报错3414 | 报错3414，定标中止，借924-591和TNI定标卡944-212（12602） | 更换备件924-591重新定标后修复 by andrew | SWO-184671 | Defect | HW | Aspiration Module | Wet Section, Manifold | 3414 | Error 3414 | Error 3414, bid suspension, loan 924-591 and TNI bid card 944-212 (12602) | Repair by andrew after replacement of spare parts 924-591 re-calibration
No | Visit | Complaint | RKK | 患者検体測定時のｐH電極応答不良 | ケース詳細の通り | pH電極交換 | SWO-188681 | Assisted Replacement | Electrode | Electrode | pH |  | Poor pH electrode response during patient specimen measurement | As the case details | pH electrode replacement
No | Visit | Complaint | RDEU | Einlass tropft erheblich bei Spülung, jeweils im letzten Zyklus | s. case | Einlass tropft jeweils in letzter Sequenz der Spülzyklen erheblich. Sonde während Spülung untersucht, sie fährt offenbar etwas weiter heraus als üblich; Nadel hat längenmittig einen zurückgebogenen Knick => Nadel und Einlassdichtung ersetzt, mehrere Kalibrationen und Spülzyklen beobachtet, kein weiterer Flüssigkeitsaustritt feststellbar. Noch installierte Cl Elektrode durch Dummy ersetzt. Endkontrolle: QC Messung OK | SWO-185168 | Defect | HW | Inlet module | Probe bend |  | Inlet drips considerably during rinse, in each case in the last cycle | s. case | Inlet drips significantly in the last sequence of the rinsing cycles. Probe examined during flushing, it apparently travels a little further than usual; Needle has replaced a bent crease => needle and inlet seal in the middle of the length, observed several calibrations and rinsing cycles, no further liquid leak can be detected. Still installed Cl electrode replaced by dummy. Final control: QC measurement OK
No | Visit | Complaint | RDEU | Regelmäßige Na Driften in beide Richtungen; sporadisch Ca ebenfalls betroffen | s.case | Alle relevanten Komponenten untersucht, Ventile im Service überprüft, Alles OK. Einlassnadel leicht gebogen - gerichtet. Messkammer auf Dichtheit geprüft, OK. pH Elektrode und Referenz ausgebaut, überprüft und klebenden pH Dichtring ersetzt. ServiceDump erstellt sowie Datenextrakt gezogen zur Weiterleitung an RDEU. Kalibrationen und QC Messung OK | SWO-185192 | Defect | Electrode | Electrode | Sealing ring | Sealing ring, pH | Regular Na drifting in both directions; sporadically Ca also affected | s.case | All relevant components examined, valves checked in service, all OK. Inlet needle slightly bent - directed. Measuring chamber tested for tightness, OK. pH electrode and reference removed, checked and adhesive pH sealing ring replaced. ServiceDump created and data extract drawn for forwarding to RDEU. Calibrations and QC measurement OK
No | 6 | Y | RAME | 3 REPEATITIVE PO2 CAL ERROR ON 0.1mmHg LEVEL SINCE MARCH 18 |  | CONFIRMED ISSUE WHEN ONSITE. | D2696F207C4C43F1 | Adjustment | HW | pH/BG Module | pO2 | Zero adjustment | 3 REPEATITIVE PO2 CAL ERROR ON 0.1mmHg LEVEL SINCE MARCH 18 |  | CONFIRMED ISSUE WHEN ONSITE.
No | 2 | Y | RCZE | Blank |  |  | D26BACDB6819FA1E | Defect | Electrode | Electrode | Sealing ring | Sealing ring, pH | Blank
No | Visit | Complaint | RDEU | Kunde hat wiederholt Kal.fehler bei pCO2 | Wiederholt Kal.fehler bei pCO2, Empfindlichkeit teilw. über 100% | Messkammer, Konnektor überprüft, Elektrode neu membraniert, gesäubert und einkalibriert, Empf. bleibt wiederholt zu hoch, testweise gebrauchte Elektrode membraniert und eingesetzt, Verlauf wird beobachtet | SWO-185266 | Assisted Replacement | Electrode | Membrane | pCO2 |  | Customer has repeated Kal.error at pCO2 | Repeats Kal.error at pCO2, sensitivity partiallyw. over 100% | Measuring chamber, connector checked, electrode re-membraneed, cleaned and calibrated, empf. remains repeatedly too high, test-used electrode membranes and used, course is observed
No | Visit | Complaint | RDEU | 19.04.2017; Hr. Weinmann, Tel.: 07361553302; Leckage erkannt 19.04.2017 9:30 H. Weinmann 07361550: Problem besteht weiterhin | Lösungspack lässt sich nicht installieren. Anschließende Jahreswartung wird vorgezogen. | Einlassdichtung mit Halter defekt. Schwarze Dichtung hat sich nach vorn herausgeschossen. Garantie. | SWO-186620 | Defect | HW | Inlet module | Inlet Gasket |  | 19.04.2017; Mr. Weinmann, Tel.: 07361553302; Leakage detected 19.04.2017 9:30 H. Weinmann 07361550: Problem persists | Solution pack cannot be installed. Subsequent annual maintenance is preferred. | Inlet seal with holder defective. Black poetry has shot forward. Guarantee.
No | 9 | Y | RAME | Z Decontamination. |  | Decontamination and Planned Maintenance. | D26ABB96F252A18E | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | Z Decontamination. |  | Decontamination and Planned Maintenance.
No | Visit | Complaint | RDEU | 08.03.2017; Hr. Hoppe; Tel.: 016093980307; IP Adresse muss am ABL eingegeben werden | IP konfiguriert | IP und Netzwerk konfiguriert | SWO-185582 | Setup | His/lis | His/lis | His/lis |  | 08.03.2017; Mr. Hoppe; Tel.: 016093980307; IP address must be entered on the ABL | IP configured | IP and network configured
No | 5 | Y | RCZE | Blank | Na výsledkoch sa nezobrazuju výsledky oximetrie. |  | D26AC2F6F237EB8B | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | Blank | The results do not show the result of Oximetry.
No | Visit | Complaint | RCHN | 接口板腐蚀 | 接口板腐蚀，借备件902-972一个 |  | SWO-185320 | [5 empty] | Interface plate corrosion | Interface plate corrosion, spare parts 902-972 one
No | Visit | Complaint | RKK | QCにてClが外れる件で、シグナルデータボード、センサーボード交換実施。ゼロチェック実施。様子見。 | QCにてClが外れる件で、シグナルデータボード、センサーボード交換実施。ゼロチェック実施。様子見。 | QCにてClが外れる件で、シグナルデータボード、センサーボード交換実施。ゼロチェック実施。様子見。 | SWO-185292 | [5 empty] | In the case where cl is removed in QC, the signal data board and the sensor board exchange were carried out. Zero check ed. Wait. | In the case where cl is removed in QC, the signal data board and the sensor board exchange were carried out. Zero check ed. Wait. | In the case where cl is removed in QC, the signal data board and the sensor board exchange were carried out. Zero check ed. Wait.
Yes | 2 | Y | RPAC | Blank |  |  | D26A93D7F25DB96A | Re classified RMED | setup new analyzer |  |  |  | Blank
Yes | 2 | Y | RPAC | Blank |  |  | D26A9398F22DB969 | Re classified RMED | setup new analyzer |  |  |  | Blank
No | 49 | Y | RESP | SE QUEDA EN PANTALLA DE INICIO STARTING UP REVISIÓN PREVENTIVA |  | SUSTITUCIÓN DE HDD SATA CONFIGURACIÓN COMPLETA REVISIÓN PREVENTIVA SEGÚN PROCEDIMIENTO | D26AD5A2F206919F | Defect | HW | DMS_PC | HDD |  | STAYS ON START SCREEN STARTING UP PREVENTIVE REVIEW |  | SATA HDD SUBSTITUTE FULL CONFIGURATION PREEMPTIVE REVIEW AS PROCEDURE
Yes | 57 | Y | RESP | Revisión preventiva |  | Revisión preventiva según procedimiento | D26A8908F23B9198 | Re classified RMED | PM |  |  |  | Preventive review |  | Preventive review as procedure
No | 7 | Y | RPOL | Blank | Diagnoza problemu, wymiana wentylatora w module AutoCheck, testy kontrolne, konieczna dalsza naprawa. |  | D26BC404682ACD30 | Revisit required | HW | Auto Check Module | Fan |  | Blank | Diagnosis of the problem, replacement of the fan in AutoCheck module, control tests, further repairs needed.
No | Visit | Complaint | RDEU | Beide Sicherungsringe der Antriebsräder am Inlet door fehlen. Übergangsweise Schlauchstücke zur Festigung aufgesetzt. Sicherungsringe werden bestellt | Wartungsmassnahme am Analysator. | Einbau der Sicherungsringe an den Antriebsrädchen der Inlet Doors. Bestellnummer der Sicherungsringe:048-290. Sie sind nicht in CRM enthalten. | SWO-185574 | Defect | HW | Cover and Chassis | Inlet Door |  | Both locking rings of the drive wheels on the Inlet door are missing. Temporarily attached hose pieces for consolidation. Fuse rings are ordered | Maintenance measure on the analyzer. | Installation of the safety rings on the drive wheel of the Inlet Doors. Order number of the fuse rings:048-290. They are not included in CRM.
No | 10 | Y | RFRA | Blank | Erreur d'intégration des résultats |  | D26AD054F21D80EB | Setup | His/lis | His/lis | His/lis |  | Blank | Results integration error
No | Visit | Complaint | RBNL | Analyzer geeft foutmelding 4794: Pipette sensor error, possibly the needle have been bent and is hitting against the sensor. Exchange needle ensure correct mounting. | De naald is U-vormig rondgebogen. Na vervangen naald diverse instellingen aangepast met de gebruiker. Alle controles berpaald. Vijf bloedmonsters met alle vier testen bepaald, vergelijk Chemie en AQT Hengelo. Instrument naar SEH verplaatst, systeem controle OK. | Nieuwe naald gemonteerd en afgesteld. Teller gereset. Systeemcontrole, OK. | SWO-185409 | Defect | HW | Aspiration Module | Bend needle |  | Analyzer geeft foutmelding 4794: Pipette sensor error, possibly the needle have been bent and is hitting against the sensor. Exchange needle ensure correct mounting. | The needle is U-shaped bent around. After replacing needle various settings adjusted with the user. All controls are Berpaald. Five blood samples with all four tests determined, compare Chemie and AQT Hengelo. Instrument moved to SEH, System check OK. | New needle mounted and adjusted. Counter reset. System Check, OK.
No | Visit | Complaint | RDEU | 19.04.2017; Fr. Eschmann; Tel.: 020885083200; keine Aqure Verbindung | AQT ist nicht verbunden. IP adresse ist falsch. Serveradr. ist falsch, | neue Adressen eingegeben, Adressen abgespeichert, Datenübertragung kontrolliert, | SWO-185545 | Setup | His/lis | His/lis | His/lis |  | 19.04.2017; Ms. Eschmann; Tel.: 020885083200; No Aqure connection | AQT is not connected. IP address is incorrect. Serveradr. is wrong, | new addresses, stored addresses, data transmission controlled,
Yes | 5 | Y | RPOL | Blank | Przegl¹d roczny. Wymiana wê¿yków. Czyszczenie. Kalibracja tHb. Analizator sprawny |  | D26BC2FD6812BCB2 | [5 empty] | Blank | Annual overview. Replacing the parts. Cleaning. THb Calibration. Functional analyzer
Yes | Visit | Complaint | RDEU | Kal und QCs werden nicht bestanden | QC Werte PCO2 sind viel zu hoch. | Glaskörper gereinigt und LP erneuert. Ohne Erfolg. SC erneuert. | SWO-186659 | Re classified RMED | SC | SC | pCO2 |  | Kal and QCs fail | QC values PCO2 are far too high. | Vitreous cleaned and LP renewed. Without success. SC renewed.
No | Visit | Complaint | RDEU | Kal und QCs werden nicht bestanden | Nach einsetzen der neuen SC Einsatz im Josefhospital, danach wieder vor Ort um zu überprüfen ob die neue SC den Fehler behoben hat. Kalibriert und Backlog Fehler bei PCO2, Na, Cl, Glu und Lac. PCO2 QC Fehler trotz Kal Fehler wieder im Norm Bereich. | Mehrfach kalibriert und QC Kontrollen durchgeführt. Alle Werte wieder grün. Gerät ist einsatzbereit. | SWO-186701 | Assisted Replacement | SC | SC | pCO2 |  | Kal and QCs fail | After starting the new SC deployment in the Josefhospital, then back on site to check if the new SC has fixed the error. Calibrated and backlog errors at PCO2, Na, Cl, Glu and Lac. PCO2 QC error despite Kal error back in the standard range. | Multiple calibrated and QC checks performed. All values again green. Device is ready for use.
No | 21 | Y | RAME | W Screen lock up and rebooting does not help |  | Rebooted anlyzer and does not help | D26BD4A2685FCF68 | Defect | SW | SW | Freeze |  | W Screen lock up and rebooting does not help |  | Rebooted anlyzer and does not help
No | Visit | Complaint | RDEU | 20.4.2017 Herr Grabler 09117580995142 meldet, dass er immer wieder aufgefordert wird das LP zu Tauschen. 21.04.2017 Frau Calow 09117580992603 wünscht einen Rückruf vor dem Service. | Spülung des Analysators | Fehlersuche, Lösungen können nicht gepumpt werden, ELD, Einlasssonde und Anschlussdichtung überprüft ok, SC und LP ausgeworfen, von Abfall zu linken Anschluss der SC gespült ok, SC gespült ok, von rechten Anschluss der SC zu Einlass gespült (schwergängig->nach mehrmaligem Spülen ok), nochmals komplette Spülung des Analysators durchgeführt, auf Serviceprogrammebene werden jetzt alle Testprogramme bestanden Startphase bestanden, System-und Funktionskontrolle i.O. | SWO-186855 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | 20.4.2017 Mr. Grabler 09117580995142 reports that he is repeatedly asked to exchange the LP. 21.04.2017 Ms. Calow 09117580992603 wants a recall before the service. | Rinsing the analyzer | Troubleshooting, solutions cannot be pumped, ELD, inlet probe and connection seal checked ok, SC and LP ejected, from waste to left connection the SC flushed ok, SC flushed ok, from right connection of the SC to inlet flushed (difficult-to-go->after repeated Rinse ok), again complete flushing of the analyzer is carried out, at the service program level all test programs are now passed start-up phase passed, system and function control i.O.
No | 6 | Y | RAME | 3 pO2 failing qualicheck qc. | Changed the complaint visit tab from N to Y. | The customer is running qc study and, pO2 is failing on all qc levels. Recommended the customer to run waste blood every day and every time during the qc study so the sensor cassette is conditioned at all times for qc. Customer is requesting my assistance with this task. i will speak to Don Masters about the customer's request. | D26BFD3E6832D4B9 | Defect | HW | Basic Unit | PCB, Sensor interface |  | 3 pO2 failing qualicheck qc. | Changed the complaint visit tab from N to Y. | The customer is running qc study and, pO2 is failing on all qc levels. Recommended the customer to run waste blood every day and every time during the qc study so the sensor cassette is conditioned at all times for qc. Customer is requesting my assistance with this task. i will speak to Don Masters about the customer's request.
No | Visit | Complaint | RLTD | When in intervention mode, unable to select consumables or enter service. | When in intervention mode, unable to select consumables or enter service. | Reloaded software now ok. | SWO-186796 | Defect | SW | RMED SW | RMED SW | Service tip available ??? | When in intervention mode, unable to select consumables or enter service. | When in intervention mode, unable to select consumables or enter service. | Reloaded software now ok.
No | Visit | Complaint | RDEU | Einlassdichtung wechseln wird vom Gerät angefordert. | Einlassdichtung wechseln -Garantie- | Gerät verlangt immer wieder Einlassdichtungswechsel. Mit Kunden das Wechseln der Einlassdichtung durchgeführt , Aktion bestätigt und Wartungsintervalle neu eingestellt. Systemtest durchgeführt, alles o.k. | SWO-186836 | Assisted Replacement | HW | Auto Inlet module | Inlet Gasket |  | Change inlet seal is requested by the device. | Changing inlet seal -Guarantee | device repeatedly requires inlet seal changes. With customers changing the inlet seal performed, action confirmed and maintenance intervals re-set. System test performed, all ok.
No | 2 | Y | RCZE | Blank | opakovaná chyba v citlivosti Lac. Provedeny 3 výmìny SC. |  | D26BE601683AF5B1 | Revisit required | SC | SC | Lac | Sensitivity | Blank | Repeated error in the Lac sensitivity. 3 Replacement SC.
No | Visit | Complaint | RKK | SC交換後より温度エラーにて初期化されない、との連絡。 | センサーボード腐食。 | センサーボード交換。SC交換。 | SWO-186689 | [5 empty] | It is not initialized by the temperature error after sc exchange. | Sensor board corrosion. | Sensor board replacement. SC replacement.
No | Visit | Complaint | RDEU | Demokey wird am 21.4.2017 ablaufen | bestellten Demooptionkey installieren | Klärung des Standortes des Gerätes nach Umzug Neubau, Klärung Zugang zum Gerät wg. Neubau, DemoOptionKey (gültig bis 21.7.2017) installiert, KAL & QK ok | SWO-186780 | Setup | Demo key exp | Demo key exp | Demo key exp |  | Demokey will expire on April 21, 2017 | Install ordered demo option key | Clarification of the location of the device after moving new building, clarification access to the device wg. New building, DemoOptionKey (valid until 21.7.2017) installed, KAL & QK ok
No | Visit | Complaint | RKK | QCレベル4ｐH電解質異常発生。Caはレンジオーバー。その他項目は上限、下限ぎりぎり。 | QC症状確認 | 比較電極交換。 | SWO-186880 | Assisted Replacement | Electrode | Electrode | Reference |  | QC level 4pH electrolyte abnormality occurs. Ca is over range. Other items are the upper limit and the lower limit. | QC symptom check | Comparison electrode replacement.
No | 7 | Y | RPOL | Blank | Wymiana modu³u AutoCheck, testy kontrolne. Analizator sprawny. |  | D26BCB9C6838CD33 | Defect | HW | Autocheck module | Autocheck module |  | Blank | AutoCheck module replacement, control tests. Functional Analyzer.
No | Visit | Complaint | RKK | 装置測定時に開始ボタンが押せない状態が頻発した為、再インストール作業を実施して様子見を行う。 | 状態確認 | 測定ソフトウェア再インストール | SWO-188425 | Defect | SW | RMED SW | Start button |  | Since the state in which the start button can not be pressed at the time of device measurement is frequently, it performs a state view by carrying out the reinstall operation. | Status Check | Re-install measurement software
No | Visit | Complaint | RKK | 装置測定時に開始ボタンが押せない状態が頻発した為、再インストール作業を実施して様子見を行う。 | 状態確認 | 測定ソフトウェア再インストール実施 | SWO-188418 | Defect | SW | RMED SW | Start button |  | Since the state in which the start button can not be pressed at the time of device measurement is frequently, it performs a state view by carrying out the reinstall operation. | Status Check | Reinstallation of measurement software
No | Visit | Complaint | RDEU | 18.04.2017, Frau Dill, Tel.: 071351011402 Herr Rank: Touchscreen reagiert verzögert bis gar nicht. Der ABL wurde neu gestartet und der Touchscreen gereinigt 20.04.2017, Herr Rank, Tel.: 0713510144403: Touchscreen reagiert weiterhin sehr verzögert oder `friert`ein. | Touchscreen reagiert zeitweise nicht. -Garantie- | Monitor komplett erneuert. Einlaufverfahren durchgeführt. | SWO-186839 | Defect | HW | DMS_PC | Touch screen |  | 18.04.2017, Mrs. Dill, Tel.: 071351011402 Mr. Rank: Touchscreen reacts delayed to not at all. The ABL has been restarted and the touchscreen cleaned 20.04.2017, Mr. Rank, Tel.: 0713510144403: Touchscreen continues to react very delayed or 'freezes'. | Touchscreen does not respond at times. -Guarantee | Monitor completely renewed. run-in procedures.
No | 8 | Y | RAME | F Customer is seeing downward trend on oxy resulst. |  | Found a pin hole in one of the tubes. | D26BD4DA6853A080 | Defect | HW | Oxi module | Tubings |  | F Customer is seeing downward trend on oxy resulst. |  | Found a pin hole in one of the tubes.
No | Visit | Complaint | RIND | FTC 103, 104, 105 Error. | FTC 103, 104, 105 error | Performed the service programs, wet section problem with P4 and P5 was failed, replaced the wash block and cleaned the filter, restarted instrument but wash block was not responding, found problem with inlet PCB need to replace with new. Waiting for spares. | SWO-186931 | Defect | HW | Aspiration Module | PCB, Inlet | FTC 103 | FTC 103, 104, 105 Error. | FTC 103, 104, 105 error | Performed the service programs, wet section problem with P4 and P5 was failed, replaced the wash block and cleaned the filter, restarted instrument but wash block was not responding, found problem with inlet PCB need to replace with new. Waiting for spares.
No | 2 | Y | RCZE | Blank | stále problémy s Lac i po vyèištìní, kontrole kanálù, výmìnì RP a SC |  | D26F50E4CA627EE9 | Replaced | Analyzer replaced | Analyzer replaced | Lac | Sensitivity | Blank | Still problems with Lac even after the check out, channel control, RP and SC
No | Visit | Complaint | RCHN | 无法开机 | 无法开机，借903-381一个。 | 更换903-381一个正常。by andrew | SWO-185609 | [5 empty] | Cannot power on | Cannot power on, borrow 903-381 one. | Replace 903-381 one normal. by andrew
No | Visit | Complaint | RCHN | 打印机无法打印 | 打印机无法打印，借备件903-001一个 | 更换903-001一个正常。 by andrew | SWO-185421 | [5 empty] | The printer can't print | Printer cannot be printed, spare part 903-001 one | Replace 903-001 one normal. by andrew
No | Visit | Complaint | RCHN | 定标失败 1054 | 定标失败 1054，借944-220（12568）一张。 | 重新定标正常。by andrew | SWO-186621 | Assisted Replacement | Consumables | Test Kit | D-Dimer | 1054 | Bid failure 1054 | Bid failure 1054, borrow 944-220 (12568) one. | Recalibration is normal. by andrew
No | Visit | Complaint | RCHN | 血氧单元可能有凝块，无法抽吸溶液 1352，1350，1343 | 血氧单元可能有凝块，无法抽吸溶液 1352，1350，1343，借903-199，902-980各一个。 | 更换903-199后修复(by tim) | SWO-186908 | Defect | HW | Oxi module | Hemolyzer | Clot | Blood oxygen units may have clots that cannot be pumped out of solution 1352,1350,1343 | Blood oxygen units may have clots that cannot pump solution siphons 1352,1350,1343, borrowing 903-199,980 each. | Repair after replacement 903-199 (by tim)
No | Visit | Complaint | RCHN | 加热温度超范围 1068 | 加热温度超范围 1068，借903-210一个。 | 更换903-210一个正常。by andrew | SWO-186886 | Defect | HW | Process Module | PCB, Cup dryer | 1068 | Heating temperature over range 1068 | The heating temperature exceeds the range of 1068, with a loan of 903-210. | Replace the 903-210 one normal. by andrew
No | 2 | Y | RAME | V Bar code reader working intermittently. |  | Verified. | D26FA2E2CA594A37 | Defect | HW | Basic Unit | Barcode reader, display |  | V Bar code reader working intermittently. |  | Verified.
No | 3 | Y | RAME | V Bar code reader working intermittently. |  | Verified. | D26FA489CA4E4A39 | Defect | HW | Basic Unit | Barcode reader, display |  | V Bar code reader working intermittently. |  | Verified.
No | Visit | Complaint | RDEU | 24.04.2017 Sr. Christiane / Pfl. Jan 065194783750 meldet, dass der ABL einen Leckagefehler zeigt. | Vacuum value liegt bei 0,8 [0,25-0,75] | Oxi - Schlauch ausgetauscht. Einlassdichtung aus dem Kundenbestand ausgetauscht. Neustart. | SWO-187548 | Defect | HW | Oxi module | Pinch valve tube |  | 24.04.2017 Sr. Christiane / Pfl. Jan 065194783750 reports that the ABL is showing a leakage error. | Vacuum value is 0.8 [0.25-0.75] | Oxi - Hose replaced. Inlet seal from the customer stock. Restart.
Yes | 10 | Y | RAME | Blank |  | Phone support. Customer reporting leaking from to solution packs. | D26F9C9CCA0E3B33 | [5 empty] | Blank |  | Phone support. Customer reporting leaking from to solution packs.
Yes | Visit | Complaint | RCHN | LAC定标失败 | LAC定标失败 | 更换测试卡 | SWO-187642 | [5 empty] | LAC calibration failure | LAC calibration failure | Replace the test card
No | Visit | Complaint | RIND | Parameters not calibrating. | Lac not calibrating. | Try to calibrate. Reinstall the membrane and try again. But not calibrating lactate parameter. Need to replace Lactate membrane. Asked from indore will be expecting tomorrow. | SWO-188576 | Assisted Replacement | Electrode | Membrane | Lac |  | Parameters not calibrating. | Lac not calibrating. | Try to calibrate. Reinstall the membrane and try again. But not calibrating lactate parameter. Need to replace Lactate membrane. Asked from indore will be expecting tomorrow.
No | 6 | Y | RAME | 1 Customer reported pH unstable error messages on patient samples. |  | Research data logs. Checked inlet flow. | D2702597CA62B21D | Cleaned | Fluidic pathway | pH/BG Module | Measuring chamber |  | 1 Customer reported pH unstable error messages on patient samples. |  | Research data logs. Checked inlet flow.
No | Visit | Complaint | RSCH | Reep Lactat | Lactat nach wenigen Tagen QC Probleme | Lichtfilter El $Met ist nicht eingebaut. Provisorisch abgedunkelt. Lichtfilter bei RMED bestellt. | SWO-187180 | Defect | HW | El/Met Module | Kapton filter | missing | Reep Lactat | Lactate after a few days QC problems | Light filter El $Met is not installed. Provisionally darkened. light filter ordered from RMED.
No | Visit | Complaint | RDEU | Seit einiger Zeit ist ABL nicht Messbereit | Keine Pump Kal.......LS Justage.......kein Pressure.Test wird bestanden. | Einlassdichtung überprüft.ist neu.. Nadel nicht Verbogen. Gerät überprüft gespült. Stenose $ Gerinnsel freigespült. Pump..LS ...Pressure Test wird bestanden. Initialisierung. Gerät überprüft arbeitet, aber PCO2 Glu.Lac sind durch stillstand nicht ok. Neue SC ist nicht vorhanden, kein Wechsel möglich. | SWO-187290 | Revisit required | SC | SC | Glu/Lac |  | For some time, ABL has not been ready for measurement | No Pump Kal....... LS Adjustment....... no Pressure.Test is passed. | Inlet seal checked.is new.. Needle not bent. Device checked flushed. Stenosis - clots rinsed free. Pump.. Ls... Pressure test is passed. Initialization. Device checked works, but PCO2 Glu.Lac are not ok by standstill. New SC is not available, no change is possible.
Yes | 15 | Y | RAME | F OXI TURBIDITY ERROS ON PATIENT SAMPLES ONLY |  | RAN SEVERAL SAMPLES WITH REOCCURRING RESULTS | D2722C1AB61A2BC4 | [5 empty] | F OXI TURBIDITY ERRORS ON PATIENT SAMPLES ONLY |  | RAN SEVERAL SAMPLES WITH REOCCURRING RESULTS
No | Visit | Complaint | RBNL | Vloeistoftransport error | Diverse vloeistoftransport errors. Bij nader onderzoek blijkt dat de afvoer van Inlet naar SC verstopt te zijn. De verstopping heb verwijderd vervolgens in het service menu heb gecontroleerd. Hierna start de ABL90 zonder probleem op. | De afvoer Inlet naar SC is verstopt. De blokkage heb verwijderd, OK. | SWO-187480 | Cleaned | Fluidic pathway | Fluidic pathway | Clot |  | Vloeistoftransport error | Various liquid transport errors. Upon further examination it appears that the drain from Inlet to SC is clogged. Have removed the blockage then checked in the service menu. After this, the ABL90 starts up without problem. | The drain Inlet to SC is clogged. Have removed the block, OK.
No | 3 | Y | RPOL | Blank | Wymieniono modu³ inletu (903-338), podl¹czono aparat do NetOp-a,diagnostyka podzespo³ów i testy dzia³ania. Analizator sprawny technicznie. |  | D270DB09400FE37D | Defect | HW | Auto Inlet module | Auto Inlet module |  | Blank | The module (903-338) has been replaced, the NetOp-a camera, the subsystem diagnostics, and the operation tests have been completed. Technically efficient analyzer.
Yes | 20 | Y | RAME | D COHGB is too high on whole blood control |  | Checked data logs. Found results with high COHGB. Reran whole blood control, it is still running 4.5. The range is 2.1 to 3.9 | D27119CE4045F04F | [5 empty] | D COHGB is too high on whole blood control |  | Checked data logs. Found results with high COHGB. Reran whole blood control, it is still running 4.5. The range is 2.1 to 3.9
No | 17 | Y | RAME | Z Molded door hinge assembly for solution pack is broken. |  | The latch that holds the SP door hinge assemby closed is broken. | D2709FA0401AC107 | [5 empty] | Z Molded door hinge assembly for solution pack is broken. |  | The latch that holds the SP door hinge assemby closed is broken.
No | Visit | Complaint | RCHN | 报错 5784 | 报错 5784 | 更换测试卡 | SWO-187394 | Defect | Consumables | Test Kit | Cartridge | 5784 | Error 5784 | Error 5784 | Replace the test card
No | Visit | Complaint | RIND | Parameters not calibrating. | Lactate membrane to replace. | Lactate membrane replaced. Analyser serviced and calibrated properly. Handed over in a fine working condition. Electrolyte and Crea Glucose profile created as requested by staff. | SWO-188581 | Assisted Replacement | Electrode | Membrane | Lac |  | Parameters not calibrating. | Lactate membrane to replace. | Lactate membrane replaced. Analyser serviced and calibrated properly. Handed over in a fine working condition. Electrolyte and Crea Glucose profile created as requested by staff.
No | Visit | Complaint | RIND | Sample and rinse leakage from inlet gasket. | Checked inlet gasket and found that there was clot inside gasket. | Cleaned gasket with hypo chloride solution and also cleaned waste path. Then given pump calibration. Unit is working ok. | SWO-188196 | Cleaned | HW | Inlet module | Inlet Gasket |  | Sample and rinse leakage from inlet gasket. | Checked inlet gasket and found that there was clot inside gasket. | Cleaned gasket with hypo chloride solution and also cleaned waste path. Then given pump calibration. Unit is working ok.
No | Visit | Complaint | RKK | リーク検出 | リーク検出 | インレットプローブ曲がり修正 エアライン清掃 LS調整実施 CAL・QC測定データ確認 | SWO-187947 | Defect | HW | Inlet module | Probe bend | 663 | Leak detection | Leak detection | Inlet probe bending fix Airline Cleaning Ls adjustment implemented Cal/QC measurement data confirmation
No | 5 | Y | RAME | V Autocheck aborting QC. |  | No sample detected errors. Also some QC logging as unknowns. | D270CFF1403222DB | Defect | HW | Autocheck module | FPC Cable |  | V Autocheck aborting QC. |  | No sample detected errors. Also some QC logging as unknowns.
No | 4 | Y | RAME | F oxi error 581 |  | oxi spectrum mismatch on samples | D270C8FA400FF6E4 | Defect | HW | Oxi module | Hemolyzer |  | F oxi error 581 |  | oxi spectrum mismatch on samples
No | Visit | Complaint | RDEU | 24.04.2017; Fr. Dohmes; Tel.: 06946086125; QC Messfehler - Fehler vorhanden in Messung mit LQC-Lösung in Position | Überprüfung Einsatz mit AleSc | ServiceDump erstellt und analysiert. Letzte 5Flows unter 1600. Analysator geöffnet. Pumpe P2 und P4 ausgetauscht. Schläuche verschließen. Verunreinigte Cup Wash Unit bei Halbjahreswartung gereinigt. QC durchgefallen. Parameter neu kalibriert. Weiterhin fehlerhafte QC. Kunde wird erneut angefahren Im Anschluss Halbjahreswartung | SWO-187574 | Performance | HW | Process Module | Tubings |  | 24.04.2017; Ms. Dohmes; Tel.: 06946086125; QC measurement error - error present in measurement with LQC solution in position | Review Use with AleSc | ServiceDump creates and analyzes. Last 5Flows under 1600. Analyzer open. pump P2 and P4. Hoses. Contaminated Cup Wash Unit cleaned at half-year maintenance. QC. parameter recalibrated. Further faulty QC. Customer is approached again Following half-year maintenance
No | Visit | Complaint | RDEU | 24.04.2017; Sr. Tina; Tel.: 01702950263; Probenaspiration gescheitert 24.04.2017, Frau Dr. Paulick, Tel.: 016093853036: weiterhin Probenaspiration gescheitert | Quetschventil geprüft | Ventil schwergängig, erneuert Testmessung durchgeführt i.O | SWO-187720 | Defect | HW | Oxi module | Pinch valve |  | 24.04.2017; Sr. Tina; Tel.: 01702950263; Sample aspiration failed 24.04.2017, Dr. Paulick, Tel.: 016093853036: continued rehearsal aspiration failed | Squeeze valve tested | Valve heavy-duty, renewed Test measurement carried out in i.O
Yes | 15 | Y | RAME | F TURBIDITY ERRORS ON PAT SAMPLES |  | REPLACED ANDF TESTED WITH NEW ANNUAL TUBING, AND SPECTROMETER. TRY WET SECTION PCB | D2722F2FB6152BCB | [5 empty] | F TURBIDITY ERRORS ON PAT SAMPLES |  | REPLACED ANDF TESTED WITH NEW ANNUAL TUBING, AND SPECTROMETER. TRY WET SECTION PCB
No | Visit | Complaint | RLTD | Unable to pump solutions, changed SP and flushed still not working. Steve Wells 07527 185464 | Analyser in intervention required status Checked consumables and found SP not properly activated and aligned on analyser | Removed SP and activated and re inserted into analyser Carried out fluid transport checks within service program to get analyser back in ready mode | SWO-187916 | User error | SP | SP Activation | SP Activation |  | Unable to pump solutions, changed SP and flushed still not working. Steve Wells 07527 185464 | Analyser in intervention required status Checked consumables and found SP not properly activated and aligned on analyser | Removed SP and activated and re inserted into analyser Carried out fluid transport checks within service program to get analyser back in ready mode
No | Visit | Complaint | RDEU | Akku lädt sich nicht am Netz | Akku nicht defekt, wie Tausch über Kreuz ergeben hatte | power supply erneuert, Teile zurücckgetauscht | SWO-187650 | Defect | HW | Basic Unit | Power Supply |  | Battery does not charge on the mains | Battery not defective, as exchange over cross had revealed | power supply renewed, parts exchanged
No | Visit | Complaint | RDEU | Die. Fehlermeldung | SC wurde heute Nacht ersetzt, Gerät zeigt im laufenden Betreib Temp. Hardware Fehler, Lac QC fehlerhaft | Gerät zerlegt, Austausch aller Schlauchverbindungen, Quetschventil im Oxy Unit geprüft i.O., Sensor Bay Unit, Hemolyzer Unit & Einslassnadel gespült, Befestigung an unterteil des Oxy Unit abgebrochen ->ersetzt div. Serviceprogramme durchgeführt (Fluidtransport, kpl. Oxy & pO2 Test) fehlerfrei, Kal & QC | SWO-187812 | Defect | HW | Oxi module | Hemolyzer |  | The. Error message | SC was replaced tonight, device shows in running operation Temp. Hardware error, Lac QC faulty | Device disassembled, replacement of all hose connections, squeezing valve in the Oxy Unit tested i.O., Sensor Bay Unit, Hemolyzer Unit & inlet needle flushed, attachment to lower part of the Oxy Unit aborted ->replaced various service programs (fluid transport, cpl. Oxy & pO2 Test) error-free, Kal & QC
No | 5 | Y | RAME | z black screen |  | lights and alarms black screen | D272263FB638865B | Revisit required | HW | DMS_PC | CPU |  | z black screen |  | lights and alarms black screen
No | Visit | Complaint | RDEU | Das Gerät akzeptiert nur die Radiometer Vorgabe als Standard. Änderungen werden nicht akzeptiert. | Das Gerät akzeptiert nur die Radiometer Vorgabe als Standard. Änderungen werden nicht akzeptiert. Der Kunde wünscht, das verschiedene Einstelllungen geändert werden. | Datensicherung durchgeführt, Software neu installiert, Datensicherung wieder eingespielt, Patientenlayout nach Kundenwunsch angepasst, Referenzbereiche eingetragen, 1 std Startphase durchgeführt, Einstellungen nach Kundenwunsch geändert. | SWO-187826 | Defect | SW | Patient layout | Patient layout |  | The device accepts only the radiometer default as the default. Changes will not be accepted. | The device accepts only the radiometer default as the default. Changes will not be accepted. The customer wants different settings to be changed. | Data backup performed, software reinstalled, data backup re-recorded, patient layout adapted to customer requirements, reference areas entered, 1 std start phase performed, settings changed according to customer request.
No | 4 | Y | RFRA | Blank | affiche fuite detectée |  | D2722DB5B65E2F5D | Defect | HW | Auto Inlet module | Inlet Gasket |  | Blank | detected leaked poster
Yes | 10 | Y | RAME | D COHGB is running high on customers whole blood control |  | Checked data logs, observed the out of range values. Customers whole blood control is running 4.1. | D272512DB607791B | [5 empty] | D COHGB is running high on customers whole blood control |  | Checked data logs, observed the out of range values. Customers whole blood control is running 4.1.
No | 9 | Y | RAME | K Analyzer asking for Remote support license. |  | The analyzer could not find the remote support key. | D272368AB661C10E | [5 empty] | K Analyzer asking for Remote support license. |  | The analyzer could not find the remote support key.
No | 20 | Y | RAME | 2 The PCO2 was not passing level 1 of external QC. |  | Checked the data logs and verrified the problem. The customer used 24 ampuoles of level 1 QC over past few days in trying to get the analyzer working. | D271FCFDB632A3B4 | [5 empty] | 2 The PCO2 was not passing level 1 of external QC. |  | Checked the data logs and verrified the problem. The customer used 24 ampuoles of level 1 QC over past few days in trying to get the analyzer working.
No | Visit | Complaint | RCHN | 报错4587 | 报错5766，借备件903-198一个 | 更换903-198后修复(by tin) | SWO-187058 | Defect | HW | Analyzer Control Module | PCB, Main Board | 5766 | 4587 error | 5766, borrowed spare parts 903-198 one | Repair after replacement 903-198 (by tin)
No | Visit | Complaint | RCHN | 血氧模块未激活 0682 | 血氧模块未激活 0682，借902-833一个。 | 更换902-833后修复(by tim) | SWO-187279 | Defect | HW | Oxi module | Oxi module, Replaced | 0682 | Blood oxygen module not activated 0682 | The blood oxygen module is not activated 0682, borrowing 902-833 one. | Repair after replacement 902-833 (by tim)
No | Visit | Complaint | RCHN | 条码扫描器无法扫描 | 条码扫描器无法扫描，借备件902-822一个 | 更换902-822一个正常。(by tim) | SWO-187324 | Defect | HW | Basic Unit | Barcode reader, Basic Unit |  | Barcode scanner cannot be scanned | Barcode scanner cannot be scanned, borrowed spare parts 902-822 one | Replace 902-822 one normal. (by tim)
No | 4 | Y | RAME | P Fluid flow issues |  | pH eelctrode chamber has fluid build up intermittently. | D2723B50B6428B15 | Defect | Electrode | Electrode | Sealing ring | Sealing ring, pH | P Fluid flow issues |  | pH eelctrode chamber has fluid build up intermittently.
No | Visit | Complaint | RIND | All Sample measurements are Positive. | All Sample measurements are Positive. | Found All Sample measurements are Positive, Cleaned the process wheel, and cup wash unit but problem was remaining same need to replace with process wheel cup wash unit and OU. | SWO-187793 | Cleaned | HW | Process Module | Process wheel |  | All Sample measurements are Positive. | All Sample measurements are Positive. | Found All Sample measurements are Positive, Cleaned the process wheel, and cup wash unit but problem was remaining same need to replace with process wheel cup wash unit and OU.
No | Visit | Complaint | RDEU | Frau Vogelsberger meldet einen Vorfall eines hochgradig positiven TnI beim Patienten. Analysator meldet negativ | Überprüfung laut Anweisung | Abarbeitung des Data Collection Guide Analysatortechnisch. ServiceDump analysiert, Flow ohne Beanstandung Cup Wash Unit versehen mit leichter Kratzspur, CWU zerlegt und gereinigt, freier Lauf und einwandfreie Funktion. Nadelkolben gereinigt, leichte Verzögerung im zurückgleiten. Nadel ungewöhnlich stumpf (Counter 3000 pierces), erneuert und Counter resetet. Weitere Überprüfungen des DCG ohne Beanstandung. Blankmessungen und Flowanalyse i.O. DataLogs gesichert. Vorfall: Kist am 20.04.2017 12:47 | SWO-187879 | Cleaned | HW | Aspiration Module | Needle assembly |  | Ms Vogelsberger reports an incident of a highly positive TnI in the patient. Analyzer reports negative | Verification according to instruction | Processing of the Data Collection Guide analytical technical. ServiceDump analyzed, flow without complaint Cup Wash Unit provided with light scratch track, CWU disassembled and cleaned, free running and flawless function. Needle piston cleaned, slight delay in sliding back. Needle unusually blunt (Counter 3000 pierces), renewed and counter resets. Further checks of the DCG without complaint. Blank measurements and flow analysis in the first place DataLogs. Incident: Kist on 20.04.2017 12:47
No | 6 | Y | RFRA | Blank | AQT 90 arrété : support de tube cassé |  | D27267EBB6442F61 | Defect | HW | Aspiration Module | Tube holder |  | Blank | AQT 90 stopped: broken tube holder
Yes | Visit | Complaint | RDEU | 24.04.2017; Fr. Dohmes; Tel.: 06946086125; QC Messfehler - Fehler vorhanden in Messung mit LQC-Lösung in Position | Die Qualitätskontrolle bei TnT ist im Level rot und gelb seit Anfang des Jahres am oberen Grenzwert. Auch bei D-Dimer sind die Qualitätskontroll Werte nach oben erhöht. Überprüfung Messwerte bei den letzten TnT Kalibrationen, in Ordnung. | Durchführung einer Jahreswartung Maßnahme. Zusätzlich Erneuerung des Pumpenkopfes P2 und P5. Durchführung von Blankmessungen. Messwerte in Ordnung innerhalb der Grenzwerte. Überprüfung der Flowwerte, Werte liegen nun unter dem Wert von 1300. Sind nun niedriger, vorher bei über 1400. Qualitätskontrollen in Ordnung, zwischen oberem Grenzwert und Zielwert. Service Dump gezogen zur Ausarbeitung an Technischen Support. Entscheidung, ob noch Hardware getauscht wird. | SWO-187911 | Re classified RMED | Double entry |  |  |  | 24.04.2017; Ms. Dohmes; Tel.: 06946086125; QC measurement error - error present in measurement with LQC solution in position | The quality control at TnT has been at the upper limit since the beginning of the year in the red and yellow level. Even with D-Dimer, the quality control values are increased upwards. Check measurements at the last TnT calibrations, all right. | Implementation of an annual maintenance measure. In addition, renewal of the pump head P2 and P5. Execution of blank measurements. Measured values in order within the limit values. Checking the flow values, values are now below the value of 1300. Are now lower, previously at over 1400. Quality controls are fine, between the upper limit value and the target value. Service Dump pulled to technical support for elaboration. decide whether to replace hardware.
No | Visit | Complaint | RLTD | Heather Fife 01315371898 | Analyser in windows screen when I arrived Rebooted analyser and failed due to memory error Possible database corruption | Backed up setup files in explorer Installed fresh database Loaded setup files | SWO-188204 | Defect | SW | Windows | Windows |  | Heather Fife 01315371898 | Analyser in windows screen when I arrived Rebooted analyser and failed due to memory error Possible database corruption | Backed up setup files in explorer Installed fresh database Loaded setup files
No | 5 | Y | RAME | z black screen |  | at power up speaker beeps barcode active, lights on | D27351282C111658 | Unknown | HW | DMS_PC | CPU |  | z black screen |  | at power up speaker beeps barcode active, lights on
No | Visit | Complaint | RCHN | 无法开机 | 无法开机，借933-148一个。 | 更换933-148后修复(by tim) | SWO-187638 | [5 empty] | Cannot power on | Cannot power on, borrow 933-148 one. | Repair after replacement 933-148 (by tim)
No | 2 | Y | RPAC | Blank | Fluid blockage error. Leaking solution pack |  | D2738F952C5E009B | [5 empty] | Blank | Fluid blockage error. Leaking solution pack
No | Visit | Complaint | RCHN | 测试卡无法安装 | 测试卡无法安装，借备件902-972一个 | 更换902-972后修复(by tim) | SWO-187430 | [5 empty] | The test card could not be installed | Test card cannot be installed, spare parts 902-972 one | Repair after replacement 902-972 (by tim)
No | Visit | Complaint | RCHN | 测试卡舱门无法打开 | 测试卡舱门无法打开，借580-083一个 | 更换580-083后修复(by tim) | SWO-187748 | Defect | HW | Basic Unit | SC Door |  | Test card door cannot be opened | Test card door cannot be opened, borrow 580-083 one | Repair after replacement 580-083 (by tim)
No | Visit | Complaint | RIND | crea calibration expire. | Crea calibration error. | Replace the Crea reagent. | SWO-188573 | User error | Consumables | Consumables | Crea expired |  | crea calibration expire. | Crea calibration error. | Replace the Crea reagent.
No | Visit | Complaint | RIND | Sample Problem was shown on ABL800 analyzer at Sasoon Hospital, Pune. | Need to check the inlet casket for any clots. | Removed the inlet casket, dipped it in sodium hypochloride solution in order to remove the blood, it was found to be blocked. Run protein removal after connecting the inlet casket. Told the technicians about the protein removal process. Run 2 point calibration, run few samples machine is found working fine. | SWO-188676 | Cleaned | HW | Inlet module | Inlet Gasket | Clot | Sample Problem was shown on ABL800 analyzer at Sasoon Hospital, Pune. | Need to check the inlet casket for any clots. | Removed the inlet casket, dipped it in sodium hypochloride solution in order to remove the blood, it was found to be blocked. Run protein removal after connecting the inlet casket. Told the technicians about the protein removal process. Run 2 point calibration, run few samples machine is found working fine.
No | 4 | Y | RAME | 2 Qc not coming in. |  | pCO2 high on low QC compared to peer group. | D273665C2C10D536 | Defect | Electrode | Electrode | Sealing ring | Sealing ring, pH | 2 Qc not coming in. |  | pCO2 high on low QC compared to peer group.
No | Visit | Complaint | RIND | Lamp intensity low. | Lamp intensity low. | Found Lamp intensity low, performed the complete Oximetry test, found Lamp adjustment was not OK, performed the Lamp adjustment performed the calibration found ctHb values are OK, performed the test found ctHb and SO2 values are OK. Found Instrument working satisfactorily. | SWO-188542 | Adjustment | HW | Oxi module | Lamp unit | Low intensity | Lamp intensity low. | Lamp intensity low. | Found Lamp intensity low, performed the complete Oximetry test, found Lamp adjustment was not OK, performed the Lamp adjustment performed the calibration found ctHb values are OK, performed the test found ctHb and SO2 values are OK. Found Instrument working satisfactorily.
No | 13 | Y | RAME | H Failing QC. Error code 374: Inhomogeneous sample at pH/BG. Sample errors |  | Examined QC data logs. | D2735F392C0EE404 | Defect | HW | Inlet module | Inlet module |  | H Failing QC. Error code 374: Inhomogeneous sample at pH/BG. Sample errors |  | Examined QC data logs.
No | Visit | Complaint | RLTD | Evangeline 01372 735279 (DDL) | Inspected inlet probe and gasket - inlet probe slightly bent, inlet gasket split. Replaced inlet probe (Radiometer stock) and inlet gasket (customer stock). Ran all fluid transport tests in service - all passed. Also cleaned contacts for solution pack with alcohol to improve contact to solution pack. Restarted analyser and verified operation. - all now ok. | Replaced bent inlet probe and damaged $split inlet gasket. | SWO-188537 | Defect | HW | Inlet module | Inlet Gasket |  | Evangeline 01372 735279 (DDL) | Inspected inlet probe and gasket - inlet probe slightly bent, inlet gasket split. Replaced inlet probe (Radiometer stock) and inlet gasket (customer stock). Ran all fluid transport tests in service - all passed. Also cleaned contacts for solution pack with alcohol to improve contact to solution pack. Restarted analyser and verified operation. - all now ok. | Replaced bent inlet probe and damaged $split inlet gasket.
No | Visit | Complaint | RLTD | sam patel 07989 278754 | Printer missing brass bush from print roller. | Replaced printer. | SWO-188465 | Defect | HW | Printer Module | Printer |  | sam patel 07989 278754 | Printer missing brass bush from print roller. | Replaced printer.
No | Visit | Complaint | RDEU | Zweimal Fehler 1190 Einlaß geöffnet in zwei Tagen | Messung abgebrochen - Einlaß geöffnet | Clutch Schrauben nachgezogen, Probelauf ok | SWO-188443 | Adjustment | HW | Auto Inlet module | Inlet drive gear |  | Twice error 1190 inlet opened in two days | Measurement aborted - inlet open | Clutch screws tightened, test run ok
No | 3 | Y | RAME | P Customer reported leakage from front of analyzer. |  | QC leaking from front of unit during system cycles. | D2752CEFA203CBC8 | [5 empty] | P Customer reported leakage from front of analyzer. |  | QC leaking from front of unit during system cycles.
No | Visit | Complaint | RCHN | 报错5297 | 报错5297，借备件903-371一个 | 更换903-371后修复(by tim) | SWO-187955 | Defect | HW | Incubator Module | Punch up |  | 5297 error | 5297, spare parts 903-371 one | Repair after replacement 903-371 (by tim)
No | 62 | Y | RESP | Error 689 tensión de offset de presión fuera de rango. |  | Comprobar módulo de gases, limpiar transductor y tubos capilares de paso de gases, calibrar y probar con controles. | D2749480A24847DB | Cleaned | HW | Gas module | Gas module | Replace Transducer expected | Error 689 out-of-range pressure offset voltage. |  | Check gas module, clean transducer and gas passing capillary tubes, calibrate and test with controls.
No | Visit | Complaint | RDEU | Seit Wochen immer wieder deutliche Abweichungen bei MetHb und HHB Messungen trotz mehrfacher tHb und diversen Reinigungsverfahren | s. case | ABL überprüft, fließtechnisch alle Funktionen OK, keine Auffälligkeiten. Weiterhin Driften hoch und runter bei Kal und Mismatch Meldung bei QC Messung => Oximodul ersetzt, Abgleich, OxiCompleteTest, Kalibrationen, Endkontrolle: drei QC Messungen OK | SWO-188483 | Defect | HW | Oxi module | Oxi module, Replaced | MetHb and HHb | For weeks again and again significant deviations in MetHb and HHB measurements despite multiple tHb and various cleaning methods | s. case | ABL checked, flowed all functions OK, no abnormalities. Further drifting up and down at Kal and mismatch message at QC measurement => Oximodul replaced, matching, OxiCompleteTest, calibrations, final control: three QC measurements OK
No | Visit | Complaint | RDEU | 27.04.2017; Fr. Domberg; Tel.: 05115322529; Lac QC Messfehler Level 3 | Spüllösungsrückstände in der Messkammer | Elyte Messkammer geprüft und gereinigt, Funktionskontrolle ok | SWO-188474 | Cleaned | HW | El/Met Module | Measuring chamber | Residue | 27.04.2017; Ms. Domberg; Tel.: 05115322529; Lac QC Measurement Error Level 3 | Rinsing solution residues in the measuring chamber | Elyte measuring chamber tested and cleaned, function control ok
No | Visit | Complaint | RSCH | Reparatur | QC pCO2 seit 06.04.2017 sporadisch zu tief Kunde hatte um 04:22 Membranenwartung durchgeführt Danach Empfindlichkeit sporadisch über 100% Sporadisch drift und Schwankung bei Status QC Lot Level 4 neue Lot nicht eingelesen | Allgemeine Kontrolle i.o pCO2 Elektrode ersetzt und Membranenwartung durchgeführt i.o Membrane aus Thalwil Danach Status instabil Erneut Membranenwartung durchgeführt i.o Garantieaustausch pCO2 Elektrode XT04 Kal. und QC i.o | SWO-188383 | Assisted Replacement | Electrode | Electrode | pCO2 | Sensitivity high | Repair | QC pCO2 sporadically too deep since 06.04.2017 Customer had performed membrane maintenance at 04:22 After that sensitivity sporadically over 100% Sporadic drift and fluctuation in status QC Lot Level 4 new lot not read in | General control i.o pCO2 electrode replaced and membrane maintenance carried out i.o Membrane from Thalwil After that status unstable Again membrane maintenance carried out i.o Warranty replacement pCO2 electrode XT04 Kal. and QC i.o
No | Visit | Complaint | RDEU | 26.04.2017 16:40Uhr Herr Kamp 06946086123: Das Gerät wurde nach der Wartung wieder auf seinen Platz gestellt. Nach dem Hochfahren meldet das Gerät Fehler 1051/1052 und 3412 | Ftc_101 und Ftc_103 nach Reagenzienwechsel Keine Lösung an P1 und P3 zu gegebener Zeit Zeitliche Überlagerungen des Problems während auslesen des Storagewheel | Analysator geöffnet und eingeschaltet. ServiceDump analysiert, i.O. Pumpentest P1-P5 durchgeführt, Änderungen des Value an LS und Htc Zelle ohne Beanstandung. Schlauchlage optimiert, leichte knicke im Schlauchsystem. Metallklammern gelockert. Initialisierungsphase 5 Blankmessungen durchgeführt, ServiceDump erstellt und analysiert. Flow zwischen 1249-1309 Analysator auf Messplatz installiert und gestartet. Analysator messbereit MCC2 9162 | SWO-188461 | Adjustment | HW | Aspiration Module | Tubings |  | 26.04.2017 16:40Mr Kamp 06946086123: The device was put back in place after maintenance. After booting up, the device reports errors 1051/1052 and 3412 | Ftc_101 and Ftc_103 after reagent change No solution to P1 and P3 in due time Time overlays of the problem while reading out the storagewheel | Analyzer opened and turned on. ServiceDump analyzed, i.O. Pump test P1-P5 performed, changes in value to LS and Htc cell without complaint. Hose position optimized, slight bends in the hose system. Metal clamps loosened. Initialization phase 5 blank measurements performed, ServiceDump created and analyzed. Flow between 1249-1309 Analyzer installed and started on measuring station. Analyzer ready for measurement MCC2 9162
No | Visit | Complaint | RCHN | 软件升级失败，定标数据丢失 | 软件升级失败，定标数据丢失，借944-216（12452）和944-214（12482）各一张 | 重新定标后修复(By Ran) | SWO-187959 | Defect | SW | RMED SW | SW - Upgrade | LEH SW Version | Failed software upgrade, bid data loss | Failed software upgrade, bid data loss, borrowing 944-216 (12452) and 944-214 (12482) each | Repair after re-calibration (By Ran)
No | Visit | Complaint | RCHN | 无法开机 | 无法开机，借910-285一个 |  | SWO-188637 | Defect | HW | Power Supply | Power Supply |  | Cannot power on | Cannot power on, borrow 910-285 one
No | Visit | Complaint | RCHN | 报错1031错误 | 报错1031错误，借903-364一个 |  | SWO-188658 | Defect | HW | Process Module | Optical Unit | 1031 | 1031 error | 1031 error, borrow 903-364 one
No | Visit | Complaint | RCHN | 进样盘上的滚轮坏了 | 进样盘上的滚轮坏了，借903-584一个 | 更换903-584后修复(by tim) | SWO-188604 | Defect | HW | Aspiration Module | Tube holder |  | The roller on the sample plate is broken. | The roller on the sample plate is broken, borrow 903-584. | Repair after replacement 903-584 (by tim)
Yes | 23 | Y | RAME | 3 PO2 will not pass calibration. |  | Customer reported that the P02 would not pass cals or QC, they have replaced the SC and SP twice as well as flushed the analyzer. | D28DDD4CDA53C699 | [5 empty] | 3 PO2 will not pass calibration. |  | Customer reported that the P02 would not pass cals or QC, they have replaced the SC and SP twice as well as flushed the analyzer.
No | 1 | Y | RAME | Z MISC |  | ISSUES WITH BARCODE SCANNING WHILE SETTING UP INTERFACE WITH OMNITECH | D2785714045A3288 | Setup | SW | Barcode reader | Barcode reader, display |  | Z MISC |  | ISSUES WITH BARCODE SCANNING WHILE SETTING UP INTERFACE WITH OMNITECH
Yes | 2 | Y | RPAC | Blank | Manual |  | D2814D943E4FE80F | Re classified RMED | Delivering Manual |  |  |  | Blank | Manual
No | Visit | Complaint | RCHN | pco2定标失败 | pco2定标失败，借903-581 |  | SWO-188728 | Defect | HW | Basic Unit | PCB, Sensor interface |  | pco2 calibration failure | pco2 bid failure, borrowed 903-581
Yes | Visit | Complaint | RKK | 北海道医療センター手術室機を移設。BML機。 オーバーホール済み。 | ケース詳細の通り | ABL715⇒ABL800Basic（825GL）へ入れ替え | SWO-194986 | Re classified RMED | Installation |  |  |  | Relocated the operating room machine of Hokkaido Medical Center. BML machine. Overhauled. | As the case details | Replaced by ABL715⇒ABL800Basic(825GL)
Yes | 7 | Y | RAME | P Leaking at inlet |  | Checked all tubings and inlet area | D27904197A0AD1C6 | [5 empty] | P Leaking at inlet |  | Checked all tubings and inlet area
No | 11 | Y | RPNZ | Blank | Low CO2 on QC |  | D2825BCCB41308A4 | Cleaned | HW | pH/BG Module | Measuring chamber | pCO2 | Blank | Low CO2 on QC
No | 11 | Y | RAME | F COHb, METHb QC is out of range intermittently and not comparing to peer. Intermittent OXI spectrum mismatch errors and out of range blood peer comparisons. pO2 QC L-3 is having high results. |  | tHb 1 point calibration value was below specifications. QC ranges are have been changed multiple times to accommodate QC result trend or high CV's. Sample mixing from tubes is not optimum for blood sample comparison and may contribute to spectrum mis-match errors when not mixed properly. pO2 QC results are most likely from running QC manually accross the two instruments as L-3 will be most affected by room air contamination. | D278ACB90435DBAE | Adjustment | HW | Oxi module | Lamp unit |  | F COHb, METHb QC is out of range intermittently and not comparing to peer. Intermittent OXI spectrum mismatch errors and out of range blood peer comparisons. pO2 QC L-3 is having high results. |  | tHb 1 point calibration value was below specifications. QC ranges are have been changed multiple times to accommodate QC result trend or high CV's. Sample mixing from tubes is not optimum for blood sample comparison and may contribute to spectrum mis-match errors when not mixed properly. pO2 QC results are most likely from running QC manually accross the two instruments as L-3 will be most affected by room air contamination.
No | 3 | Y | RAME | Z Customer reported by text touch screen not responding correctly. |  | Touch positions not working correctly. | D278B2790402AFEA | Defect | HW | DMS_PC | Touch screen |  | Z Customer reported by text touch screen not responding correctly. |  | Touch positions not working correctly.
Yes | 7 | Y | RPOL | Blank | Aparat zabrudzony krwi¹ w œrodu. Zalany zosta³ czujnik po³o¿enia ig³y co doprowadzi³o do jej po³amania. Testy kontrolne. Analizator sprawny. |  | D2829FB9B4468E34 | [5 empty] | Blank | The camera is contaminated with blood ¹ in the family. The sensor for the location of the device has been floated and has been Control tests. Functional Analyzer.
Yes | Visit | Complaint | RDEU | Materiallieferung | Materiallieferung von SC90 946-008 (R0516), da Kunde falsch bestellt hat --> Aufgrund von Kundenverschulden (falsches Material bestellt) muss Material geliefert werden -->Frau Leidenberger weiß über die Berechnung der Lieferung Bescheid | Materiallieferung SC90 946-008 (R0516) | SWO-189163 | Re classified RMED | Logistical upSWO |  |  |  | Materiallieferung | Material delivery of SC90 946-008 (R0516) because customer ordered incorrectly --> Due to customer fault (wrong material ordered) material must be delivered -->Mrs. Leidenberger knows about the calculation of the delivery | Materiallieferung SC90 946-008 (R0516)
No | Visit | Complaint | RCHN | LAC定标失败 | LAC定标失败 | 更换测试卡 | SWO-188890 | [5 empty] | LAC calibration failure | LAC calibration failure | Replace the test card
No | Visit | Complaint | RKOR | Solution pack recognition error | i Button error | i Button replace. Solution pack replace. 2-point calibration check. | SWO-195187 | [5 empty] | Solution pack recognition error | i Button error | i Button replace. Solution pack replace. 2-point calibration check.
Yes | Visit | Complaint | RCHN | PO2定标失败 | PO2定标失败 | 更换测试卡后修复 | SWO-188915 | [5 empty] | PO2 calibration failure | PO2 calibration failure | Repair after replacing test card
No | Visit | Complaint | RCHN | 报1031错误 | 报1031错误，借903-364一个 | 更换903-364后修复 by jesson | SWO-188729 | Defect | HW | Process Module | Optical Unit | 1031 | Reported 1031 error | Reported 1031 error, borrow 903-364 one | Repair by jesson after replacing 903-364
No | Visit | Complaint | RDEU | 02.05.2017 Sr. Anja 02214784039 meldert, dass der ABL einen Spülfehler hat. Sie wünscht einen Vorortservice. | s. case | Erneut ist die Membrane des BG Modul Gas-Ventils mittig ausgestanzt => Ventil komplett erneuert. Startverfahren etc., Kal und QC Messung OK | SWO-189061 | Defect | HW | pH/BG Module | Valve, air | Membrane | 02.05.2017 Sr. Anja 02214784039 notifys that the ABL has a flushing error. She wants a suburban service. | s. case | Once again, the membrane of the BG Module gas valve is punched out in the middle => valve is completely renewed. Start method etc., Kal and QC measurement OK
No | 15 | Y | RAME | H Inhomogeneous sample detected at pH/BG module. |  | Examined data logs. | D279C70E7A31C66C | Defect | HW | Inlet module | Inlet module | clot? | H Inhomogeneous sample detected at pH/BG module. |  | Examined data logs.
No | Visit | Complaint | RIND | Error In QC Level 1 | Minor Blood Clot In PH Causes Error In QC | Cleaned The PH Chamber, Run The Cleaning Mode, Machine Passed 2 Point Calibration & QC Level 1. Machine Working Fine | SWO-190127 | Cleaned | HW | pH/BG Module | Measuring chamber | Clot | Error In QC Level 1 | Minor Blood Clot In PH Causes Error In QC | Cleaned The PH Chamber, Run The Cleaning Mode, Machine Passed 2 Point Calibration & QC Level 1. Machine Working Fine
No | Visit | Complaint | RIND | touch screen not working properly. | touch not activating sometimes | replaced the touch screen with new, did calibration for touch screen ,found touch working properly.instrument kept under observation. | SWO-188987 | Defect | HW | DMS_PC | Touch screen |  | touch screen not working properly. | touch not activating sometimes | replaced the touch screen with new, did calibration for touch screen ,found touch working properly.instrument kept under observation.
No | Visit | Complaint | RSCH | Rep Kond Fehler | 0219: Konditionierfehler im El/Met Modul Gerät rinnt bei Spülung aus Kapillar Einlass | Kein Fehler festgestellt Friction Detector Unit ausgetauscht Einlass Nadel und Einlass Dichtung ausgetauscht Ventile überprüft Kapillare Testmessung durchgeführt Funktionskontrolle i.o | SWO-188956 | Revisit required | HW | Inlet module | Inlet Gasket |  | Rep Kond Error | 0219: Conditioning error in the El/Met module Device runs during flushing from capillary inlet | No error detected Detector Unit replaced Inlet needle and inlet seal replaced Valves checked Capillary test measurement performed Functional control i.o
No | Visit | Complaint | RSCH | Rep Kond Fehler | friction detector | Friction Detector zu SWO188956 | SWO-189153 | Defect | HW | Inlet module | Friction detector unit |  | Rep Kond Error | friction detector | Detector Friction to SWO188956
No | Visit | Complaint | RIND | All Sample measurements are Positive. | All Sample measurements are positive. | Found and replaced with Process wheel and OU, Cup wash, Calibrated instrument found Calibration results are Ok, performed the test results are OK. Found instrument is working satisfactorily. | SWO-190116 | Defect | HW | Process Module | Process wheel |  | All Sample measurements are Positive. | All Sample measurements are positive. | Found and replaced with Process wheel and OU, Cup wash, Calibrated instrument found Calibration results are Ok, performed the test results are OK. Found instrument is working satisfactorily.
No | Visit | Complaint | RKOR | Storage wheel error | Storage wheel error | Storage wheel check. reagent cartrige check. Service mode check. analyzer reset. | SWO-195190 | Adjustment | HW | Storage Module | Storage wheel | 5689 | Storage wheel error | Storage wheel error | Storage wheel check. reagent cartrige check. Service mode check. analyzer reset.
No | 17 | Y | RFRA | Blank | lecture des codes barres impossible |  | D27ACFE9F05AD6A6 | Defect | HW | Barcode reader | Barcode reader, display |  | Blank | reading barcodes impossible
No | Visit | Complaint | RDEU | 03.05.2017 11:30Uhr Herr Antrest 095150315241: Die Uhrzeit wurde korrigiert, aber nach kurzer Zeit ist die Uhrzeit wieder um eine Stunde verstellt. | Fehlersuche --> Uhrzeit eigestellt und Verfahren für die automatische Zeiteinstellung angewandt | Uhrzeit eigestellt und Verfahren für die automatische Zeiteinstellung angewandt, System und Funktion i.O. | SWO-189528 | Setup | SW | Time setting | Time setting |  | 05/03/2017 11:30AM Mr. Antrest 095150315241: The time has been corrected, but after a short time the time is adjusted again by one hour. | Troubleshooting --> Time set and automatic time setting procedures applied | time and procedures for automatic time setting applied, system and function in the order
Yes | 13 | Y | RAME | Z D ECONTAMINATION AND P LANNED M AINTENANCE . |  | SCHEDULED D ECONTAMINATION. | D27B0020F03BA6CD | Re classified RMED | PM |  |  |  | Z D ECONTAMINATION AND P LANNED M AINTENANCE . |  | SCHEDULED D ECONTAMINATION.
No | 4 | Y | RFRA | Blank | Plus de tactile en bas à droite |  | D27ACD80F05DCA51 | Adjustment | HW | DMS_PC | Touch screen |  | Blank | More touch at the bottom right
No | Visit | Complaint | RLTD | Bent Probe | The inlet probe got bent over the weekend. | Replaced the inlet probe with a new one. | SWO-189414 | Replaced | HW | Auto Inlet module | Probe bend |  | Bent Probe | The inlet probe got bent over the weekend. | Replaced the inlet probe with a new one.
No | Visit | Complaint | RDEU | 24.04.2017, Sr. Anja, tel.: 006021323661 : ständiger SC und LP Wechsel, Sr. Anja bemängelt, dass durch die ganzen Wechsel, der ABL oft nicht zur Verfügung steht. Sie bittet um eine Überprüfung. | Jahreswartung vorbeugend durchgeführt. Alle Service Tests i.O. Kal und QC i.O | Präventiv Einweisung durchgeführt. | SWO-189814 | Defect | HW | PM Kit | PM Kit |  | 24.04.2017, Sr. Anja, tel.: 006021323661 : permanent SC and LP change, Sr. Anja complains that due to all the changes, the ABL is often not available. It asks for a review. | Annual maintenance carried out preventively. All service tests in kal and QC i.O | Preventive instruction.
Yes | Visit | Complaint | RBNL | Roma.ini aangepast | Op verzoek, de Roma.ini file heb aangepast zodat je het labnummer bij dataoverzicht kan zien. | Roma.ini heb aangepast. | SWO-189326 | Re classified RMED | SW | His/lis | His/lis | FSE : On request of the client, Roma.ini file has to be adjusted in order to change the patient ID to AccesionNumber. | Roma. INI Modified | On request, the Roma. ini file has been modified so you can see the lab number in data overview. | Roma. INI have adjusted.
Yes | Visit | Complaint | RBNL | Roma.ini aangepast | Op verzoek, de Roma.ini file heb aangepast zodat je het labnummer bij dataoverzicht kan zien. | Roma.ini heb aangepast. | SWO-189327 | Re classified RMED | SW | His/lis | His/lis | See SWO-189326 | Roma. INI Modified | On request, the Roma. ini file has been modified so you can see the lab number in data overview. | Roma. INI have adjusted.
No | 4 | Y | RFRA | Blank | Pas de sensibilité Glu et Lac |  | D27B1188F0055958 | Revisit required | Unknown | Unknown | Unknown | Techline | Blank | No Glu and Lake sensitivity
No | Visit | Complaint | RDEU | 02.05.2017 14:30Uhr Ütmü 062213629165: Die Uhrzeit verstellt sich regelmäßig um eine Stunde. Das Datum geht um einen Tag falsch. | Fr. Ütmü, Tel. 015147127161, Datum und Uhrzeit verstellen sich weiterhin. | Fr. Ütmü hat wie gebeten SC und LP erneuert, zuvor das Datum+Uhrzeit richtig gestellt. Die Uhrzeit verstellt sich weiterhin nach Neustart um 1 Std. Vermutlich Sommer Winterzeit Problem. Das Datum ändert sich wohl im Betrieb auf Morgen. Fr. Ütmü bittet um einen Service Einsatz. | SWO-189365 | [5 empty] | 02.05.2017 14:30H Ütmü 062213629165: The time is regularly adjusted by one hour. The date is one day wrong. | Ms. Ütmü, tel. 015147127161, date and time are still changed. | Ms. Ütmü has renewed SC and LP as requested, previously set the date+time correctly. The time still adjusts by 1 hour after reboot. Probably summer winter time problem. The date probably changes to tomorrow. Ms. Ütmü asks for a service.
Yes | Visit | Complaint | RIND | SC Calibration error. | SC calibration error. | Found SC 80 air detection failure, Inactivated the Air detection reinstalled the SC but problem was remaining same, replaced with new SC 80 found instrument is functional. | SWO-190119 | [5 empty] | SC Calibration error. | SC calibration error. | Found SC 80 air detection failure, Inactivated the Air detection reinstalled the SC but problem was remaining same, replaced with new SC 80 found instrument is functional.
Yes | 3 | Y | RPAC | Blank | Solution pack leaking. |  | D27BFF906642EC8B | [5 empty] | Blank | Solution pack leaking.
Yes | Visit | Complaint | RCHN | LAC,GLU,PCO2定标失败 | LAC,GLU,PCO2定标失败 | 连续更换试剂包和测试卡后修复 | SWO-189520 | [5 empty] | LAC, GLU, PCO2 bid failed | LAC, GLU, PCO2 bid failed | Repair after continuous replacement of reagent kits and test cards
No | Visit | Complaint | RCHN | CO2定标不过 | CO2定标不过 | 更换测试卡后修复 | SWO-189492 | Assisted Replacement | SC | SC | pCO2 |  | CO2 calibration is not the | CO2 calibration is not the | Repair after replacing test card
No | Visit | Complaint | RCHN | 报错 1031，1025，1024 | 报错 1031，1025，1024，借903364，944-220（12325）一张。 | 更换903-364，重新定标。by jesson | SWO-188848 | Defect | HW | Process Module | Optical Unit | 1031 | Error 1031,1025,1024 | 1031,1025,1024, borrow 903364,944-220 (12325). | Replace the 903-364 and re-mark the bid. by jesson
No | Visit | Complaint | RDEU | Anpassung der Konfiguration für das Cobra PDMS erforderlich. | Konfg angepasst. | Geräte gemäß Absprache mit Frau Bolk und Dr. Jedzek - Zader konfiguriert und getestet. | SWO-189400 | Setup | SW | His/lis | His/lis |  | Configuration adjustment required for the Cobra PDMS. | Konfg adjusted. | Devices configured and tested in consultation with Ms. Bolk and Dr. Jedzek - Zader.
No | 15 | Y | RAME | V Analyzer Shutting Down When Entering Demographics. |  | Customer reported that after scanning the accession number during a sample then pressing enter on the external keyboard would cause the analyzer to shut down. | D28DE495DA2EC6A6 | Defect | HW | Power Supply | Power Supply | Same day | V Analyzer Shutting Down When Entering Demographics. |  | Customer reported that after scanning the accession number during a sample then pressing enter on the external keyboard would cause the analyzer to shut down.
No | 15 | Y | RAME | v Flex-Q Positions 2-3 Not Working |  | Performed flex Q calibration in service mode, found position 1 was calibrated OK but positions 2 and 3 were off and needed adjusting. | D28DE06DDA1FC69E | Adjustment | HW | FlexQ Module | Positioning | Same day | v Flex-Q Positions 2-3 Not Working |  | Performed flex Q calibration in service mode, found position 1 was calibrated OK but positions 2 and 3 were off and needed adjusting.
No | Visit | Complaint | RDEU | Fr. Domberg, Tel.: 05115322529, pH und E-Lyt. Kal. Fehler. | PH/Elyte Empfindlichkeitsfehler | BG Messkammer geprüft, demontiert und gereinigt, Funktionskontrolle ok | SWO-189688 | Cleaned | HW | El/Met Module | Measuring chamber |  | Ms. Domberg, Tel.: 05115322529, pH and E-Lyt. Kal. Error. | PH/approx. Empfindlichkeitsfehler | BG measuring chamber tested, dismantled and cleaned, function control ok
No | 3 | Y | RAME | P Customer reports that QC's are intermittently aborted |  | 0331 errors are present, indicating no sample detected. Ampoule breaker appears to break ampoules. However, some are broken out of alignment or access to a different ampoule than the one broken is observed. | D27AE573F02DC117 | Defect | HW | Autocheck module | FPC Cable |  | P Customer reports that QC's are intermittently aborted |  | 0331 errors are present, indicating no sample detected. Ampoule breaker appears to break ampoules. However, some are broken out of alignment or access to a different ampoule than the one broken is observed.
No | 3 | Y | RAME | P - FLUID FLOW |  | Client advised of failing QCs for inhomogeneous soltiions intermittently. | D27ADB60F0516A4D | Adjustment | HW | Autocheck module | Carroussel |  | P - FLUID FLOW |  | Client advised of failing QCs for inhomogeneous soltiions intermittently.
No | Visit | Complaint | RDEU | Fr. Weber, Tel.: 06174201222, LQC Barcodes lassen sich nicht einlesen, Monitor BCR verstaubt. | Fr. Weber, Tel.: 06174201222, LQC Barcodes lassen sich nicht einlesen, Monitor BCR verstaubt. | Mit Fr. Weber das Scannerfenster gereinigt, der erste Barcode wurde erkannt, beim zweiten bleibt der Fehler, zweiten Barcode von Hand eingeben lassen. Kundin bittet um Service Einsatz am Montag oder Mittwoch nächste Woche. | SWO-189341 | Cleaned | HW | Analyzer Control Module | Barcode reader, display |  | Ms. Weber, Tel.: 06174201222, LQC barcodes cannot be read in, Monitor BCR dusted. | Ms. Weber, Tel.: 06174201222, LQC barcodes cannot be read in, Monitor BCR dusted. | With Ms. Weber cleaned the scanner window, the first barcode was detected, the second one remains the error, the second barcode is entered by hand. Customer asks for service on Monday or Wednesday next week.
No | 4 | Y | RFRA | Blank | Pas de lecture code barre |  | D27AD44EF0077B2B | Cleaned | HW | Storage Module | Barcode reader, Storage wheel |  | Blank | No barcode playback
No | Visit | Complaint | RDEU | QC fehlerhaft in Level B u. C bei tHb, Hb-Derivaten und tBil | QC-Fehler Level B u. C bei tHb, Hb-Derivaten, tBil | Reinigung Fließwege, Service-Programme ausgeführt, tHb-Kal wiederholt, Kal i.O., QC i.O. | SWO-189950 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway | tHb | QC faulty in Level B and C for tHb, Hb derivatives and tBil | QC-Fehler Level B u. C bei tHb, Hb-Derivaten, tBil | Cleaning flow paths, service programs executed, tHb-Kal repeated, Kal i.O., QC i.O.
No | 5 | Y | RAME | P error 1343 unable to pump solutions. |  | Reviewed circumstances leading to reported issue. Customer indicates a SP was recently replaced, having difficulty removing it to re-install. Suspect SP was not activated completely as cause of issue. | D28086803E5B8768 | User error | SP | SP Activation | SP Activation |  | P error 1343 unable to pump solutions. |  | Reviewed circumstances leading to reported issue. Customer indicates a SP was recently replaced, having difficulty removing it to re-install. Suspect SP was not activated completely as cause of issue.
No | Visit | Complaint | RDEU | 03.05.2017 11:30Uhr Herr Antrest 095150315241: Die Uhrzeit wurde korrigiert, aber nach kurzer Zeit ist die Uhrzeit wieder um eine Stunde verstellt. | Fehlersuche --> Einstellungen auf Standard zurückgestellt, da die Analysatoren an einem Zeitserver hängen!!!! | Einstellungen auf Standard zurückgestellt, da die Analysatoren an einem Zeitserver hängen!!!! Anschließend Analysator mehrfach neu gestartet, um zu überprüfen, dass die Einstellungen passen System und Funktion i.O. | SWO-189663 | Setup | SW | Time setting | Time setting |  | 05/03/2017 11:30AM Mr. Antrest 095150315241: The time has been corrected, but after a short time the time is adjusted again by one hour. | Troubleshooting --> Settings reset to default because the analyzers are attached to a time server!!!! | Settings are set back to default because the analyzers are attached to a time server!!!! Then analyzer restarted several times to verify that the settings fit System and function in the
No | Visit | Complaint | RDEU | 03.05.2017 11:40Uhr Frau Bergk 06032944607: Das Gerät meldet einen Gerätefehler und ist nicht einsatzbereit. | DMS nicht mit Analysatoreinheit verbunden | Datensicherung auf allen Ebenen durchgeführt. Roma wurde beendet und nicht neu gestartet. Überprüfung der Module und Neustart des Analysators. Initialisierungsphase 1Stunde. Kalibrierfehler pCO2, Glu und Lac Fällige tHb Kal durchgeführt Kal und QC in diversen Parametern nicht i.O. SC wurde getauscht und reklamiert. | SWO-189667 | Defect | SW | Power off/on | Power off/on |  | 05/03/2017 11:40AM Ms. Bergk 06032944607: The device reports a device error and is not ready for use. | DMS not connected to analyzer unit | data backup at all levels. Roma has been terminated and not restarted. Check the modules and restart the analyzer. Initialization phase 1hour. Calibration error pCO2, Glu and Lac Due tHb Kal performed Kal and QC in various parameters not in the SC has been exchanged and complained about.
No | Visit | Complaint | RSCH | Blank | Erreur configuration dans le fichier patient (impossible de consulter la colonne de travail) garantie | Copie Roma Ini de l'équipement ABL90 + du Labo e.o Test informatique ok e.o Calibration e.o contrôle de qualité e.o | SWO-189611 | Setup | SW | Roma.ini | Roma.ini |  | Blank | Error configuration in the patient file (impossible to view the work column) guaranteed | Roma Ini copy of the ABL90 equipment - from the Lab e.o Computer test ok e.o Calibration e.o quality control e.o
Yes | Visit | Complaint | RBNL | Roma.ini aanpassing | Op verzoek, de Roma.ini file heb aangepast zodat je het labnummer bij dataoverzicht kan zien. | Roma.ini heb aangepast. | SWO-189720 | Re classified RMED | SW | His/lis | His/lis | See SWO-189326 | Roma. INI adaptation | On request, the Roma. ini file has been modified so you can see the lab number in data overview. | Roma. INI have adjusted.
Yes | 5 | Y | RFRA | Blank |  |  | D27CABB86630FFD2 | Re classified RMED | Logistical upSWO | Supply SC |  |  | Blank
No | Visit | Complaint | RCHN | 除PO2外定标不过，CA定标不过 | 除PO2外定标不过，CA定标不过 | 更换测试卡后修复 | SWO-189755 | [5 empty] | Except for PO2, however, the CA calibration is not | Except for PO2, however, the CA calibration is not | Repair after replacing test card
No | Visit | Complaint | RCHN | 接口板腐蚀 | 接口板腐蚀，借902-972一个。 |  | SWO-189554 | [5 empty] | Interface plate corrosion | Interface plate corrosion, borrow 902-972 one.
No | Visit | Complaint | RCHN | 进样口模块弹簧老化 | 进样口模块弹簧老化，借备件903-139一个 |  | SWO-189484 | Defect | HW | Inlet module | Inlet module |  | Inlet module spring aging | Sample module spring aging, spare parts 903-139 one
No | 6 | Y | RESP | INDICA CAL3 NO HOMOGENIA |  | REVISIÓN DE HISTÓRICOS. REALIZAR PURGAS DE LAVADO DEL CIRCUITO DE FLUÍDICA CON NaCl (10%) Y AGUA DESTILADA. LIMPIEZA DE LA SONDA DE ENTRADA. CALIBRACIONES Y CONTROLES | D27C1366660C68A2 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | INDICA CAL3 NO HOMOGENIA |  | HISTORICAL REVIEW. MAKE WASH PURGES OF THE FLUDIC CIRCUIT WITH NaCl (10%) AND DESTILATED WATER. CLEANING THE INPUT PROBE. CALIBRATIONS AND CONTROLS
No | Visit | Complaint | RCHN | 测试卡条码扫描器故障 | 测试卡条码扫描器故障，借910-304一个 | 更换910-304一个正常。by jesson | SWO-189162 | Defect | HW | Storage Module | Barcode reader, Storage wheel | BR001 | Test card barcode scanner failure | Test card barcode scanner failure, borrow 910-304 one | Replace 910-304 one normal. by jesson
No | 9 | Y | RAME | U Autocheck not working. |  | QC samples not always aspirating, random inlet errors. | D27C9CC466422867 | Defect | HW | Inlet module | Probe |  | U Autocheck not working. |  | QC samples not always aspirating, random inlet errors.
No | Visit | Complaint | RDEU | 03.05.2017 10:00Uhr Herr Ruf Med-Technik 0711310382232: Die Station soll wegen einem pH-Fehler angerufen werden. Keine Fehlerdetails bekannt. 03.05.2017 10:10Uhr Frau Neumann 07113103 2512 oder 82500: pH-Fehler bei der QC 03.05.2017 11:10Uhr Frau Neumann 0711310382500: Nach der Proteinentfernung weiterhin 03.05.2017 15:00 Frau Neumann 0711310382500 weiterhin pH Fehler. Vorortservice gewünscht | pH Wert wird zu tief gemessen bei Q.C Level1 | Gerät überprüft. Empf.Status unauffällig. Messkammer BG Modul gereinigt. Blut an pH Elektrode entfernt. Einlass durchgespült. Air Valve V7 gespült. Nach mehreren Kal pH ok. Na Grenzwertig hoch bei Q.C. nach erneuter Kal und Q.C alles ok. | SWO-189725 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | 03.05.2017 10:00o Mr. Ruf Med-Technik 0711310382232: The station should be called because of a pH error. No error details known. 05/03/2017 10:10AM Ms Neumann 07113103 2512 or 82500: pH error in QC 05/03/2017 11:10Am Ms Neumann 0711310382500: After protein removal continue 05/03/2017 15:00 Ms Neumann 0711310382500 continue pH error. On-site service desired | pH value is measured too low at Q.C Level1 | device checked. Empf.Status inconspicuous. Measuring chamber BG module cleaned. Blood removed from pH electrode. Inlet flushed through. Air Valve V7 rinsed. After several cal pH ok. Well high at Q.C. after re-cal and Q.C all ok.
No | Visit | Complaint | RDAN | Apparatet står på 3041. USB porten foran på appatet virker ikke. Der forsøges med ny USB nøgler om det gør en forskel. | Afprøvning | USB port testet med andre usb stick og der er godt nok forlænget ventetid, men det kan kan lade sig gøre at benytte porten. | SWO-189965 | No fault found | SW | DMS_PC | USB | Slow comm. | The device is 3041. The USB port on the front of the app does not work. Attempts are made with new USB keys if it makes a difference. | Testing | USB Port tested with other USB stick and there is well enough prolonged latency, but it can be possible to use the port.
No | Visit | Complaint | RDAN | Cl ligger alt for højt på S7755 og falder ud 50 % af gangene. | Se på QC plot | Udskiftning af Cl elektroden | SWO-189976 | Assisted Replacement | Electrode | Electrode | Cl |  | Cl is too high on S7755 and falls out 50% of the hallways. | Look at QC plot | Replacing the Cl electrode
No | 2 | Y | RAME | Z Customer reported from courtesy call, unit is getting DMS disconnected messages. |  | Checked data logs for errors. Unit getting errors code 643 DMS disconnected from AU. | D27C33D16636954C | Defect | HW | Basic Unit | PCB, Analyzer control |  | Z Customer reported from courtesy call, unit is getting DMS disconnected messages. |  | Checked data logs for errors. Unit getting errors code 643 DMS disconnected from AU.
No | 11 | Y | RFRA | Blank | Lenteur constaté par le client pour lancer les analyses |  | D27BF31F666172A2 | Unknown | Unknown | Unknown | Unknown | Techline | Blank | Slowness found by the customer to start the analyses
No | 7 | Y | RFRA | Blank | AQT 90 arrété : support de tube cassé |  | D27C2511662B044D | Defect | HW | Aspiration Module | Tube holder |  | Blank | AQT 90 stopped: broken tube holder
Yes | 5 | Y | RAME | Z Calibration failed - Under warranty. Royln Robinsin 804-2974464 Install known good sensor but still failed. | Changed the complaint visit tab from N to Y. | Found Comm, PCB was defective. | D27C3EBC66201E6A | [5 empty] | Z Calibration failed - Under warranty. Royln Robinsin 804-2974464 Install known good sensor but still failed. | Changed the complaint visit tab from N to Y. | Found Comm, PCB was defective.
No | Visit | Complaint | RKK | キャリブレーション不良 リンス溶液に気泡混入 | ｷｬﾘﾌﾞﾚｰｼｮﾝ不良 | ガスケット洗浄→回路洗浄 | SWO-190212 | Assisted Replacement | SP | SP | FT | 1081 | Defective calibration, bubbles mixed in the rinse solution | Poor calibration | Gasket cleaning → Circuit cleaning
No | Visit | Complaint | RDEU | 04.05.2017 Herr Kratz 0602132434705 meldet vierschiedene Kal und QC Fehler. | Gerät geprüft. Es wurde mehrmals versucht, Sensorkassette mit überschrittenem Verfallsdatum zu installieren. | Nach Erneuerung der Sensorkassette, hat sich das Gerät einkalibriert. Alle Service Tests durchgeführt. Kal. und QCs sind i.O | SWO-190494 | User error | SC | SC | SC expired |  | 04.05.2017 Mr. Kratz 0602132434705 reports four-different Kal and QC errors. | device checked. Several attempts were made to install the sensor cartridge with the expiry date exceeded. | After renewing the sensor cartridge, the device has calibrated itself. All service tests performed. Kal. and QCs are i.O.
Yes | Visit | Complaint | RDEU | 04.05.2017 Herr Kratz 0602132434705 meldet vierschiedene Kal und QC Fehler. | Gerät überprüft; Laut Aktivitäten, wurde mehrmals versucht verfallene Sensorkassette zu installieren. Alle Service Tests durchgeführt. Kal und QCs sind i.O | Einweisung durchgeführt. | SWO-190911 | Re classified RMED | Double entry |  |  |  | 04.05.2017 Mr. Kratz 0602132434705 reports four-different Kal and QC errors. | device checked; According to activities, several attempts were made to install dilapidated sensor cartridge. All service tests performed. Kal and QCs are i.O | instruction carried out.
No | Visit | Complaint | RDEU | 05.05.2017 09:30Uhr Martina Krause 01702245229: Das Gerät ist seit dem 4.5. um 4Uhr nicht mehr am Netz. Ein weiteres Gerät hat sich am 29.4. um 4Uhr vom Netz getrennt. Das dritte Gerät hat sich am 2.5 um 4Uhr vom Netz getrennt. | ABL90 und 2 weitere nicht mit Radiance verbunden. | Konfiguration aller 3 Geräte überprüft, neu eingegeben und abgespeichert. Radiance Verbindung wieder verfügbar und Daten werden übermittelt. Status der Gerät OK. | SWO-190146 | Setup | SW | His/lis | His/lis |  | 05.05.2017 09:30h Martina Krause 01702245229: The device has not been on the grid since 4am. Another device disconnected from the network at 4am on 29.4. The third device disconnected from the mains at 4am on 2.5am. | ABL90 and 2 others not connected to Radiance. | Configuration of all 3 devices checked, re-entered and saved. Radiance connection is available again and data is transmitted. Status of the device OK.
No | Visit | Complaint | RDEU | D | Fehlersuche --> Dateiarchive (QC, Kal und Betriebsmittel) überprüft --> pCO2 in den 3 QCs außer Bereich (zu hoch) --> SC defekt --> keine SC mehr vor Ort (Frau Fraundörfer organisiert sich eine SC auf einer anderen Station) | SC wird vom Kunden erneuert und erhält von mir (TobEh) gegen 12 Uhr einen Rückruf!!!! | SWO-189861 | Assisted Replacement | SC | SC | pCO2 |  | D | Troubleshooting --> File archives (QC, Kal and equipment) checked --> pCO2 in the 3 QCs out of range (too high) --> SC defective --> no more SC on site (Mrs. Fraundörfer organizes a SC on another station) | SC is renewed by the customer and receives a call back from me (TobEh) around 12 o'clock!!!!
No | 2 | Y | RCZE | Blank | po výmìnì RP a SC nefunkèní, netìsnost, i po opakované reinstalaci a èištìní |  | D2822741B431F144 | Defect | HW | Auto Inlet module | Inlet Gasket |  | Blank | After the RP and SC are not working, it does not function, even after repeated reinstalling and
No | Visit | Complaint | RDEU | 05.05.2017 10:45Uhr Frau Seilbach 03445721761: Der Ausdruck ist auf eine Seite gedränkt und druckt kleine Buchstaben. | Drucker defekt | Austausch des defekten Druckers aus Kundengerät Kal- und QC | SWO-190045 | [5 empty] | 05.05.2017 10:45am Ms. Seilbach 03445721761: The printout is dimped on one side and prints small letters. | Printer defective | Replacing the defective printer from customer device Kal- and QC
No | 4 | Y | RAME | Z ABL is making noise and battery life is zero. |  | Computer unit fan is making excessive noise and battery has no life at all. | D27D7061DC27D518 | [5 empty] | Z ABL is making noise and battery life is zero. |  | Computer unit fan is making excessive noise and battery has no life at all.
No | Visit | Complaint | RCHN | 接口板腐蚀 | 接口板腐蚀，借902-972一个 |  | SWO-189877 | [5 empty] | Interface plate corrosion | Interface plate corrosion, borrow 902-972 one
No | Visit | Complaint | RCHN | LAC定标不过 | LAC定标不过 | 更换测试卡后修复 | SWO-189866 | [5 empty] | LAC calibration is not | LAC calibration is not | Repair after replacing test card
No | Visit | Complaint | RCHN | 光单北京噪音检测超范围 | 光单背景噪音检测超范围，借备件903-364一个 | 更换903-364后未修复，要重新借903-236一个 by jesson | SWO-189218 | Revisit required | HW | Process Module | Optical Unit | 1029 | Light single Beijing noise detection over range | Light single background noise detection over range, spare parts 903-364 one | Replace 903-364 after not repaired, to borrow 903-236 a by jesson
Yes | Visit | Complaint | RCHN | 报错1031 | 报错1031，借903-364一个 |  | SWO-189858 | Re classified RMED | HW | Process Module | Optical Unit | 1031 | 1031 | 1031, loan 903-364 one
No | Visit | Complaint | RCHN | 报错1031 | 报错1031，借备件903-364一个 | 更换903-364一个正常。by jesson | SWO-189563 | Defect | HW | Process Module | Optical Unit | 1031 | 1031 | Error 1031, borrowed spare parts 903-364 one | Replace 903-364 one normal. by jesson
No | Visit | Complaint | RCHN | 报错1031 | 报错1031，借备件903-364一个 | 更换903-364后修复 | SWO-189546 | Defect | HW | Process Module | Optical Unit | 1031 | 1031 | Error 1031, borrowed spare parts 903-364 one | Repair after replacement 903-364
No | Visit | Complaint | RCHN | 报错1031 | 更换光单后，定标数据丢失，借BNP定标卡944-258（12398）一张 |  | SWO-189971 | Defect | HW | Process Module | Optical Unit | 1031 | 1031 | After replacing the light sheet, the calibration data loss, by BNP calibration card 944-258 (12398) one
No | Visit | Complaint | RCHN | 触屏不灵敏 | 触屏不灵敏，借910-320一个。 |  | SWO-189846 | Defect | HW | DMS_PC | Touch screen |  | The touch screen is not sensitive | The touch screen is not sensitive, borrow 910-320 one.
Yes | Visit | Complaint | RCHN | PU板故障 | AU板故障，借备件902-847一个 | 更换902-847后修复 by jesson | SWO-189513 | Re classified RMED | OOB |  |  |  | PU board failure | AU board failure, borrowed spare parts 902-847 one | Repair by jesson after replacement 902-847
No | Visit | Complaint | RDAN | ABL835 754R2358N0003 Ustabil måling på Glucose. Besked fra Pernille: Ustabile målinger på mange pt. målinger. Jeg kan ikke se noget unormalt, og det hjælper ikke at skifte membran. Kalib og QC er ikke mere ustabilt end på de andre ABL800. Hun har tjekket fodpakning og om der er snask i målekammeret. | Nul strøm tjekket | Målekammer renset Glu elektrode skiftet | SWO-189972 | Cleaned | HW | El/Met Module | Measuring chamber |  | ABL835 754R2358N0003 Unstable measurement of Glucose. Message from Pernille: Unstable measurements on many pt. measurements. I do not see anything abnormal and it does not help to change the membrane. Kalib and QC are no more unstable than on the other ABL800. She has checked the foot pack and whether there is a snash in the measuring chamber. | Zero Power Check | Measuring chamber cleaned The Glu electrode switch
No | Visit | Complaint | RDEU | 0351/8566313 Hr. Helm QC Fehler bei Level 1 und 3 Gluc und Crea und PCo2 31.03.2017 16:45 Hr. Helm Cre Kal Fehler nach Membran Tausch + PCO2 Qc Fehler 03.04.2017 Herr Helm 03518566313 meldet, dass einige Parameter die QC nicht bestehen. | Datenbank und Betriebsmittel geprüft | Kal2-Lösung niedriger Stand --> Kal-Fehler bei Empf. 2und Empf. 3, alle Lösungen erneuert, Sitz der Elektroden geprüft --> Kal. --> Kal- Fehler Empf.3 --> Neumembranierung nicht möglich, da keine Membranen auf Station verfügbar sind --> Station bestellt Material und führt die Membranierung in Verbindung mit einer Einweisung ( Radiometer ) durch MedTech wird informiert | SWO-189906 | Assisted Replacement | Electrode | Membrane | Glu |  | 0351/8566313 Mr. Helm QC Error at Level 1 and 3 Gluc and Crea and PCo2 31.03.2017 16:45 Mr. Helm Cre Kal Error after Membrane Swap + PCO2 Qc Error 03.04.2017 Mr Helm 03518566313 reports that some parameters do not pass the QC. | Database and equipment checked | Kal2 solution low-level --> Kal error at Empf. 2and Empf. 3, all solutions renewed, seat of electrodes tested --> Kal. --> Kal- Error Empf.3 --> Re-membraneing not possible, as no membranes are available on station --> Station orders material and performs the membraneing in conjunction with a briefing ( radiometer ) MedTech is informed
No | 3 | Y | RAME | F 0581 Oxi spectrum mismatching when running qc. |  | Performed complete oxi test. All results in range. Problem seems to be intermittent. | D27D69ACDC121963 | Defect | HW | Oxi module | Lamp unit |  | F 0581 Oxi spectrum mismatching when running qc. |  | Performed Complete oxy test. ALL results in range. Problem seems to be intermittent.
No | Visit | Complaint | RDEU | 04.05.2017, Herr Kurscheidt, Tel.: 02241187361: AQC Fehler Oximetrie Level 1 04.05.2017, Herr Kurscheidt, tel.: 02241187361: weiterhin Fehler Oximetrie Level 1 04.05.2017 15:50Uhr Herr Kurscheidt 02241187361: Oximetriefehler Im Level1 und bei tHb-Kal | Oxy Modul defekt | Fehlersuche durchgeführt. Neues Oxymodul eingebaut. Keine Fehler mehr vorhanden. | SWO-189993 | Defect | HW | Oxi module | Oxi module, Replaced |  | 05/04/2017, Mr Kurscheidt, Tel.: 02241187361: AQC Error Oximetry Level 1 04.05.2017, Mr. Kurscheidt, tel.: 02241187361: continue error Oximetry Level 1 04.05.2017 15:50H Mr Kurscheidt 02241187361: Oximetry error in Level 1 and tHb-Kal | Oxy module defective | Troubleshooting. New oxymodule built-in. No more errors.
No | 4 | Y | RCZE | Blank |  |  | D27D90D2DC5EED6B | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | Blank
No | Visit | Complaint | RCHN | 装卡后显示无更多测试 | 装卡后显示无更多测试 | 更换测试卡后修复 | SWO-190099 | Assisted Replacement | SC | SC | Chip coding |  | No more tests displayed after card loading | No more tests displayed after card loading | Repair after replacing test card
No | Visit | Complaint | RIND | Not switching ON. | Not Switching ON. No activity found from Analyser. | Try to replace Power Supply unit- but not switching ON. Try to replace Analyser control PCB but Same problem found. CPU unit cleaned and try to reconnect but not switching ON. A burn of IC found upon CPU unit. Need to replace CPU unit. Sending TO for this will tack 1 or 2 week. | SWO-190161 | Revisit required | HW | DMS_PC | CPU | Return module to EJNSC | Not switching ON. | Not Switching ON. No activity found from Analyser. | Try to replace Power Supply unit- but not switching ON. Try to replace Analyser control PCB but Same problem found. CPU unit cleaned and try to reconnect but not switching ON. A burn of IC found upon CPU unit. Need to replace CPU unit. Sending TO for this will tack 1 or 2 week.
No | Visit | Complaint | RKK | 測定時エラーコード1029：光学感度範囲外の為、測定不可 | 測定時エラーコード1029：光学感度範囲外の為、測定不可 | 代替機を設置し、機器引上げ。 | SWO-193231 | Defect | HW | Process Module | Optical Unit | 1029 | Measurement error code 1029: Not measurable because it is outside the optical sensitivity range | Measurement error code 1029: Not measurable because it is outside the optical sensitivity range | An alternative machine is installed and the equipment is raised.
No | Visit | Complaint | RCHN | LAC响应错误 | LAC响应错误，借备件903-183一个 |  | SWO-190164 | Defect | HW | Basic Unit | PCB, Sensor interface |  | LAC response error | LAC response error, borrowed spare parts 903-183 one
No | Visit | Complaint | RCHN | LAC定标失败 | LAC定标失败 | 更换测试卡后修复 | SWO-190205 | Assisted Replacement | SC | SC | Lac |  | LAC calibration failure | LAC calibration failure | Repair after replacing test card
No | Visit | Complaint | RCHN | 光单北京噪音检测超范围 | 之前借903-364，更换903-364后未修复，现重新借903-236一个 backgroud check failure | 更换903-236后修复(By Ran) | SWO-189837 | Defect | HW | Process Module | Optical Unit | 1029 | Light single Beijing noise detection over range | Previously borrowed 903-364, replaced 903-364 after not repaired, now borrow 903-236 one backgroud check failure | Repair after replacement 903-236 (By Ran)
No | Visit | Complaint | RCHN | 报错1999 | 报错1999，借备件636-541一个，903-329一个 | 更换636-541及903-329后修复 | SWO-190117 | Revisit required | HW | Aspiration Module | Needle Wash Unit | 1999 | Error 1999 | Error 1999, borrowed spare parts 636-541 one, 903-329 one | Repair after replacement 636-541 and 903-329
No | Visit | Complaint | RCHN | 报错1999 | 之前借备件636-541，现加借备件903-329一个 | 追加借用备件 903-329 | SWO-190121 | Defect | HW | Aspiration Module | Needle Wash Unit | 1999 | Error 1999 | Previously borrowed spare parts 636-541, now with spare parts 903-329 one | Additional borrowed spare parts 903-329
No | 2 | Y | RAME | DOA Sensor Cassette |  | Non linear pH, pCO2 and p02. Level 3 reading high and Level 2 readings low. | D28148F73E097C0D | Assisted Replacement | SC | SC | BG | VRK | DOA Sensor Cassette |  | Non linear pH, pCO2 and p02. Level 3 reading high and Level 2 readings low.
No | Visit | Complaint | RDEU | Lösungen können nicht gepumpt werden. | LP defekt. | Inbetriebnahme für Problauf durchgeführt. Neues LP (Lot: YI-61 defekt. LP erneuert (Lot: YI-52). | SWO-190466 | Assisted Replacement | SP | SP | FT |  | Solutions cannot be pumped. | LP defective. | commissioning for Problauf. New LP (Lot: YI-61 defective. LP renewed (lot: YI-52).
Yes | 9 | Y | RPAC | Blank | Sensor cassette failure (1361) - Internal reference electrode error in sensor cassette |  | D281889B3E0D16AD | [5 empty] | Blank | Sensor cassette failure (1361) - Internal reference electrode error in sensor cassette
Yes | 4 | Y | RPAC | Blank | Cord |  | D29301A4B261F127 | Re classified RMED | Logistical upSWO |  |  |  | Blank | Cord
No | Visit | Complaint | RDEU | 05.05.2017 09:50Uhr Sr.Jenny 06021323668: Der Techniker hat bereits die Uhrzeit geprüft und eingestellt. Die Uhrzeit auf dem Display stimmt, aber auf dem Ausdruck und im Laborsystem stimmt die Uhrzeit um eine Stunde nicht. | Zeitumstellung | Zeitumstellung nach RMED Anleitung durchgeführt. Gerät zum Testen 2 x neu gestartet. Alles i.O | SWO-191013 | Setup | SW | Time setting | Time setting |  | 05.05.2017 09:50h Sr.Jenny 06021323668: The technician has already checked and set the time. The time on the display is correct, but on the printout and in the laboratory system, the time is not correct by one hour. | Time change | time change according to RMED instructions. Device to test 2 x restarted. Everything i.O.
Yes | 1 | Y | RPAC | Blank | SC failed to calibrate Lactate after 4 days. Repalced and working fine |  | D28E043ADA3A5025 | [5 empty] | Blank | SC failed to calibrate Lactate after 4 days. Repalced and working fine
No | Visit | Complaint | RCHN | 无法读取试剂包芯片信息 | 无法读取试剂包芯片信息 借903228一个 |  | SWO-190368 | [5 empty] | Unable to read reagent package chip information | Unable to read reagent package chip information Borrow 903228 one
No | Visit | Complaint | RDEU | 02.05.2017 14:30Uhr Fr. Ütmü 062213629165: Die Uhrzeit verstellt sich regelmäßig um eine Stunde. Das Datum geht um einen Tag falsch. 03.05.2017 14:30Uhr Fr. Ütmü, Tel. 015147127161, Datum und Uhrzeit verstellen sich weiterhin. | Uhrzeit deutlich verstellt. Analysator wird regelmäßig, jeden Tag, vom Anwender ausgeschaltet. Beim Neustart stellt sich die Uhrzeit immer eine Stunde vor. Überprüfung der Uhrzeit im auf Betriebssystem Ebene. Einstellungen stimmen. | Korrekturen die im Bereich Date und Time gemacht und gespeichert wurden, werden beibehalten. Auch eine kurzzeitige Ausschaltung verändert das Datum nicht. die Sommerzeit deaktiviert. Falls weiterhin Zeitveränderung vorhanden, erneute Anfahrt. | SWO-190538 | [5 empty] | 02.05.2017 14:30 H Fr. Ütmü 062213629165: The time is regularly adjusted by one hour. The date is one day wrong. 03.05.2017 14:30H Fr. Ütmü, tel. 015147127161, date and time are still changed. | time is clearly adjusted. Analyzer is turned off regularly, every day, by the user. When restarting, the time is always set to one hour. Checking the time in the operating system level. Settings are correct. | Corrections made and saved in the Date and Time area are retained. Even a short-term shutdown does not change the date. daylight saving time is deactivated. If there is still time change, re-enter.
No | Visit | Complaint | RCHN | 报1031错误 | 报1031错误，借903-364、944-258（12749）、944-394（12450）、944-220（12568）各一个 |  | SWO-190260 | Defect | HW | Process Module | Optical Unit | 1031 | Reported 1031 error | 1031, borrow 903-364, 944-258 (12749), 944-394 (12450), 944-220 (12568) each
No | Visit | Complaint | RCHN | 试管无法读取 | 试管无法读取，借备件905-953及636-551各一个 | 更换备件905-953及636-551后修复 | SWO-190179 | Defect | HW | Aspiration Module | Barcode reader, Inlet |  | Test tube cannot be read | Test tube can not be read, spare parts 905-953 and 636-551 each one | Replacement of spare parts 905-953 and 636-551 after repair
No | Visit | Complaint | RCHN | PH,代谢物模块测量仓警报灯不亮，血氧单元钨灯超范围，液体电磁阀生锈 | PH,代谢物模块测量仓警报灯不亮，血氧单元钨灯超范围，液体电磁阀生锈 | 自行更换备用件902-458，902-844两个，902-700后修复(By Ran) | SWO-190232 | Defect | HW | Oxi module | Lamp unit | USED Analyzer | PH, metabolite module measuring warehouse alarm light is not lit, blood oxygen unit tungsten lamp super range, liquid solenoid valve rust | PH, metabolite module measuring warehouse alarm light is not lit, blood oxygen unit tungsten lamp super range, liquid solenoid valve rust | Replace spare parts 902-458, 902-844 two, 902-700 post-fix (By Ran)
No | Visit | Complaint | RKK | サンプル移動エラー多発 | 機器状況確認 | ELモジュール部、フラットメンブランバルブ用メンブラン破断の為、交換。 | SWO-190574 | Defect | HW | El/Met Module | El/Met Module | Sample transport | Frequent sample movement errors | Equipment status check | Replacement for EL module part, membrane breakage for flat membrane valve.
No | Visit | Complaint | RIND | Calibration Error In Lactate, Glucose & Sodium | Suspected Membranes Might be Defective. | Replaced The Membranes But Still Calibration Error In Lactate, Glucose & Sodium. Further Testing Will be Done After Replacing Reference Membrane | SWO-193688 | Unknown | HW | El/Met Module | Tubings | Email | Calibration Error In Lactate, Glucose & Sodium | Suspected Membranes Might be Defective. | Replaced The Membranes But Still Calibration Error In Lactate, Glucose & Sodium. Further Testing Will be Done After Replacing Reference Membrane
No | Visit | Complaint | RIND | Leaking from inlet gasket & sample problem | Checked the machine found blood clot in inlet gasket & waste was block | Cleaned the inlet gasket & waste tube performed protien removal 2 point calibration done sample processed found ok | SWO-190897 | Cleaned | HW | Inlet module | Inlet Gasket |  | Leaking from inlet gasket & sample problem | Checked the machine found blood clot in inlet gasket & waste was block | Cleaned the inlet gasket & waste tube performed protien removal 2 point calibration done sample processed found ok
No | Visit | Complaint | RDEU | ^08.05.2017 09:30 Uhr He Berlinghoff Tel.: 062213967926 Fehlermeldung: pCO2-Kalibrierfehler | Trotz mehrfacher Membranierung der pCO2 Elektrode durch Anwender , bleibt die Empfindlichkeit stark außer Grenzen. | Erneuerung der pCO2 Elektrode, Reinigung der Messkammer und der Elektrodenkontakte. Reinigung auch der pH Messkammer. Werte nach mehrmaligen Kalibration wieder in Ordnung. | SWO-190537 | Assisted Replacement | Electrode | Electrode | pCO2 |  | 05/08/2017 09:30 He Berlinghoff Tel.: 062213967926 Error message: pCO2 calibration error | Despite multiple membraneing of the pCO2 electrode by users, the sensitivity remains very beyond limits. | Renewal of the pCO2 electrode, cleaning of the measuring chamber and the electrode contacts. Cleaning also of the pH measuring chamber. Values are fine again after several calibrations.
No | Visit | Complaint | RIND | Printer problem. | Printer Head roller problem. | Found head roller and lock problem need to replace with new, suggested to customer we can’t give new printer head, took old ABL800 printer and replaced the same restarted instrument found printer was working satisfactorily | SWO-191787 | Defect | HW | Printer Module | Printer head |  | Printer problem. | Printer Head roller problem. | Found head roller and lock problem need to replace with new, suggested to customer we can’t give new printer head, took old ABL800 printer and replaced the same restarted instrument found printer was working satisfactorily
No | 12 | Y | RAME | Z NO SAMPLES DETECTED WHEN PERFORMING QC AUTOMATICALLY. NEEDLE OR PROBE NOT EXTENDED INTO QC HOUSING FOR ASPIRATION. PATIENT (BLOOD) SAMPLES WILL ASPIRATE, JUST NOT QC SAMPLES. |  | CONFIRMED ISSUE WHEN ONSITE. 14-POLE FLEX CABLE WAS NOT INSTALLED PROPERLY. | D2810F943E3D4A55 | Defect | HW | Autocheck module | FPC Cable |  | Z NO SAMPLES DETECTED WHEN PERFORMING QC AUTOMATICALLY. NEEDLE OR PROBE NOT EXTENDED INTO QC HOUSING FOR ASPIRATION. PATIENT (BLOOD) SAMPLES WILL ASPIRATE, JUST NOT QC SAMPLES. |  | CONFIRMED ISSUE WHEN ONSITE. 14-POLE FLEX CABLE WAS NOT INSTALLED PROPERLY.
No | Visit | Complaint | RDEU | Fr. Weber, Tel.: 06174201222, LQC Barcodes lassen sich nicht einlesen, Monitor BCR verstaubt. | Barcodes werden nur sporadisch gelesen über Scanner am Display | Kein kontinuierlicher Strahl des Scanners, flackert. Kompletten Scanner ausgetauscht. Analysator eingeschaltet. Initialisierungsphase. Barcodes der QC Chargen eingelesen und korrigiert. Volle Funktion des Scanners. QC wird durch Frau Vogelsberger gemessen. | SWO-190495 | Defect | HW | Analyzer Control Module | Barcode reader, display |  | Ms. Weber, Tel.: 06174201222, LQC barcodes cannot be read in, Monitor BCR dusted. | Barcodes are only read sporadically via scanner on the display | No continuous beam of the scanner, flickers. Complete scanner replaced. analyzer switched on. Initialization phase. barcodes of the QC batches are read and corrected. Full function of the scanner. QC is measured by Ms Vogelsberger.
No | Visit | Complaint | RDEU | 08.05.2017 09:15Uhr Frau Eigen 0212992942257: Der Bildschirm ist schwarz, der Scanner am Bildschirm funktioniert nicht. | Display Kabel ersetzt. Display Flackert nach wie vor. Display bestellt. | Startphase | SWO-190646 | Defect | HW | Analyzer Control Module | Cable, Display |  | 08.05.2017 09:15pm Mrs. Own 0212992942257: The screen is black, the scanner on the screen is not working. | Display cable replaced. Display still flickers. display ordered. | Start
No | Visit | Complaint | RDEU | 08.05.2017 10:20Uhr Frau Ganjic 03026383308: Das Gerät bricht die Messung ab mit dem Fehler 1068 Temperaturfehler | sporadisch Meldung: Temperatur außer Bereich vor Trocknungsschritt | Fehlersuche; div. Anwendungen im Serviceprogr. durchgef.; Gerät aus; Blower + Filter ausgebaut, gereinigt u. eingebaut; Gerät ein; Startphase abgewartet; QC i.O. | SWO-190443 | Cleaned | HW | Process Module | Filter, Cup dryer | 1068 | 05/08/2017 10:20AM Ms Ganjic 03026383308: The device cancels the measurement with error 1068 Temperature error | sporadic message: Temperature out of range before drying step | Troubleshooting; various applications in the service progr. through it; device off; Blower + filter removed, cleaned and installed; device; start-up phase awaited; QC in the first
No | Visit | Complaint | RDEU | 20.03.17 16:00 Herr Kortz Tel.071164893256 , TouchScreen lässt sich nicht mehr richtig bedienen . 08.05.2017 09:40Uhr Frau Husbeck 071164893016: Der TouchScrenn lässt sich nur schwer bedienen, weil die Tasten seitlich versetz regaieren. | Touchscreen lässt sich nicht präzise bedienen. Einlassdichtung klemmt beim Spritzen - rückzug. | Touchscreen Kalibrierung durchgeführt. EWFMGR C:- commit durchgeführt. Einlasssonde ist verbogen, Sonde und Einlassdichtung aus Kundenbestand erneuert. | SWO-191256 | Adjustment | HW | DMS_PC | Touch screen |  | 20.03.17 16:00 Mr. Kortz Tel.071164893256, TouchScreen can no longer be operated properly. 08.05.2017 09:40uhr Ms. Husbeck 071164893016: The TouchScrenn is difficult to operate because the buttons are re-energizing on the side. | Touchscreen is not precisely operated. Inlet seal clamps during spraying - retraction. | Touchscreen calibration performed. EWFMGR C:- commit performed. Inlet probe is bent, probe and inlet seal is renewed from customer stock.
Yes | Visit | Complaint | RDEU | LP soll erneuert werden | Lösungspack erneuern | Gerät manuell gespült LP erneuert Gerät überprüft | SWO-190805 | Re classified RMED | DEMO Setup |  |  |  | LP to be renewed | Renew ingenusolution pack | Device manually rinsed LP renewed Device checked
No | 7 | Y | RAME | Z Printer won't print. |  | Printer won't print. Error code 1298, Printer lid open. Opening and closing lid doesn't work. Rebooting does not resolve issue. | D282545DB400D1AE | Defect | HW | Printer Module | Printer Module |  | Z Printer won't print. |  | Printer won't print. Error code 1298, Printer lid open. Opening and closing lid doesn't work. Rebooting does not resolve issue.
No | Visit | Complaint | RDEU | 09.05.2017; Fr. Wertenbroich; Tel.: 0213152953154; Messvorgang dauert 10 Minuten | Kapillar Proben werden direkt abgebrochen. | Die Einlassnadel ist verbogen - Einlassnadel erneuert. Testmessungen durchgeführt - Kapillar Proben werden wieder gemessen. Messverfahren mit BGA Probenehmer nochmal erklärt. Analysator Status grün. | SWO-191022 | Replaced | HW | Inlet module | Probe bend |  | 09.05.2017; Ms. Wertenbroich; Tel.: 0213152953154; Measuring takes 10 minutes | Capillary samples are aborted directly. | The inlet needle is bent - inlet needle renewed. Test measurements carried out - capillary samples are measured again. Measurement method with BGA sampler explained again. Analyzer status green.
No | 12 | Y | RESP | El equipo se queda con la pantalla congelada cuando lleva dos horas de inactividad sin pasar muestras |  | Comprobado conexionado. Sustituida sensord boar (placa grande frontal) por la del equipo de laboratorio con el cable incluido para comprobación. | D2826CEEB433B7A0 | [5 empty] | The computer is left with the screen frozen when it takes two hours of inactivity without passing samples |  | Proven connection. Replaced sensord boar (large front plate) with that of laboratory equipment with the included cable for testing.
No | Visit | Complaint | RCHN | LAC定标失败 | LAC定标失败 | 更换测试卡正常。 | SWO-190878 | [5 empty] | LAC calibration failure | LAC calibration failure | The replacement test card is OK.
No | Visit | Complaint | RCHN | GLU,LAC定标失败 | GLU,LAC定标失败 | 更换测试卡正常。 | SWO-191198 | Assisted Replacement | SC | SC | Glu/Lac |  | GLU, LAC calibration failed | GLU, LAC calibration failed | The replacement test card is OK.
Yes | 2 | Y | RFRA | Blank | 92-67103 Calibrations répétées et stabilisation interrompu |  | D2825A3BB407F4D6 | Re classified RMED | Double entry |  |  |  | Blank | 92-67103 Repeated calibrations and interrupted stabilization
No | 2 | Y | RFRA | Blank | 92-56727 Erreur capteur [05100010] |  | D28257D6B438F4D2 | [5 empty] | Blank | 92-56727 Sensor Error [05100010]
Yes | 8 | Y | RAME | w status code 2 |  | rebooting brings you back to same code 2 error | D2826A17B400AA20 | [5 empty] | w status code 2 |  | rebooting brings you back to same code 2 error
Yes | 2 | Y | RPAC | Blank | Software upgrade |  | D2825E0CB423AFAD | Re classified RMED | SWU |  |  |  | Blank | Software upgrade
Yes | 2 | Y | RPAC | Blank | Pre-Delivery |  | D2825D89B43CAFAB | Re classified RMED | Upgrade |  |  |  | Blank | Pre-Delivery
Yes | 27 | Y | RESP | Erro de leitura dos códigos de barras e manutenção |  | Desmontagem do equipamento, limpeza interna e externa, limpeza dos sensores, da agulha, verificação do estado da bombas, verificação do funcionamento, tudo ok. | D282AFCAB45DAC04 | Re classified RMED | PM done | is scheduled in acc w prev. |  |  | Error reading bar codes and maintenance |  | Disassembly of the equipment, internal and external cleaning, cleaning of the sensors, the needle, checking the state of the pumps, checking the operation, everything OK.
No | Visit | Complaint | RBNL | AQT geeft errormelding 1029 (Optical sensitivity out of range) | AQT geeft de foutmelding 1029. zowel op metingen als bij de systemcheck (solid std afwijking ligt op -65,1) Gehandeld vlgs de troubleshooting tree voor error 1029. reference cup schoongemaakt en system check uitgevoerd, errormelding is niet weg. Daarom de Optical unit vervangen, en de [calibrate factory reference] uitgevoerd, deze is oke. System check uitgevoerd deze is nu oke. Alle parameters opnieuw gekalibreerd, deze zijn goed behalve de TNT, hierop errormelding 1031. Dit betekent dat er een nieuwe Optical Unit geinstalleerd moet worden. Deze is besteld, 11-5-17 vervolgbezoek ingepland. | Optical unit vervangen. | SWO-190753 | Defect | HW | Process Module | Optical Unit | 1029 | AQT geeft errormelding 1029 (Optical sensitivity out of range) | AQT Displays the error message 1029. Both on measurements and at the system check (solid STD deviation is on-65.1) acted ACC the troubleshooting tree for error 1029. Reference Cup cleaned and system check performed, error message is not gone. Therefore, the Optical unit replaced, and the [Calibrate Factory reference] performed, this is okay. System Check performed this is okay now. All parameters recalibrated, these are good except the TNT, this error message 1031. This means that a new Optical Unit needs to be installed. This has been ordered, 11-5-17 follow-up visit scheduled. | Optical Unit.
No | Visit | Complaint | RDEU | 08.05.2017 09:15Uhr Frau Eigen 0212992942257: Der Bildschirm ist schwarz, der Scanner am Bildschirm funktioniert nicht. | Display ersetzt. | QC Gemessen. | SWO-190666 | Defect | HW | Analyzer Control Module | Display |  | 08.05.2017 09:15pm Mrs. Own 0212992942257: The screen is black, the scanner on the screen is not working. | Display replaced. | QC Measured.
No | 3 | Y | RFRA | Blank | 1067:MESURES ANNULEES |  | D282AD86B40D19BE | Defect | Consumables | SP | Cup Stack | 1067 | Blank | 1067:UNDOMEASURES
No | 4 | Y | RAME | V Printer roller broken |  | Printer roller broken on one end. | D283EDC82A286D07 | Defect | HW | Printer Module | Printer Module |  | V Printer roller broken |  | Printer roller broken on one end.
No | Visit | Complaint | RDEU | 05.05.2017 09:50Uhr Sr.Jenny 06021323668: Der Techniker hat bereits die Uhrzeit geprüft und eingestellt. Die Uhrzeit auf dem Display stimmt, aber auf dem Ausdruck und im Laborsystem stimmt die Uhrzeit um eine Stunde nicht. | Zeitumstellung | Zeitumstellung nach RMED Anleitung durchgeführt. Gerät zum Test 2 x neu gestartet. Alles ok. | SWO-190967 | Setup | SW | Time setting | Time setting |  | 05.05.2017 09:50h Sr.Jenny 06021323668: The technician has already checked and set the time. The time on the display is correct, but on the printout and in the laboratory system, the time is not correct by one hour. | Time change | time change according to RMED instructions. Device to test 2 x restarted. All ok.
No | Visit | Complaint | RCHN | 监测过程中提示更换电极膜 | 监测过程中提示更换电极膜，借945-737一个。 | 更换945-737一个正常。(By Ran) | SWO-190909 | [5 empty] | Prompt to replace the electrode film during monitoring | During the monitoring process prompt to replace the electrode film, borrow 945-737 one. | Replace the 945-737 one normal. (By Ran)
No | Visit | Complaint | RCHN | LAC定标失败 | LAC定标失败，借903-581一个 | 更换903-581后修复(By Ran) | SWO-190640 | Defect | HW | Basic Unit | PCB, Sensor interface |  | LAC calibration failure | LAC bid failed, borrow 903-581 one | Repair after replacement 903-581 (By Ran)
Yes | 10 | Y | RPAC | Blank | Annual Service |  | D28397EF2A116DEC | [5 empty] | Blank | Annual Service
No | Visit | Complaint | RDEU | Regelmäßige Na Driften in beide Richtungen; sporadisch Ca ebenfalls betroffen | Weiterhin Na Drift zu hoch | Referenz Elektrode ersetzt und weiteren Datenextrakt gezogen. Neustart, Kalibrationen und QC Messung OK | SWO-190947 | Assisted Replacement | Electrode | Membrane | Reference |  | Regular Na drifting in both directions; sporadically Ca also affected | Continue Na Drift too high | Reference electrode replaced and further data extract pulled. Restart, calibrations and QC measurement OK
No | Visit | Complaint | RCHN | 血氧模块断开，电流有问题 | 血氧模块断开，电流有问题，借备件902-847及910-274各一个 | 无更换备件，检查到是液路到试剂下的灯，中间连接的排线破损导致短路，故处理了下排线后修复(By Ran) no spare part changed. Just the cable was broken.Fix the cable, then fix the problem. | SWO-190168 | Defect | HW | Basic Unit | Cable |  | Blood oxygen module disconnect, current problem | Blood oxygen module disconnect, current problem, spare parts 902-847 and 910-274 each one | No replacement spare parts, check is the liquid road to the reagent under the lamp, the middle connection of the line break caused by a short circuit, so the lower row after repair (By Ran) no spare part part changed. Just the cable was broken. Fix the cable, then fix the problem.
No | 7 | Y | RPNZ | Blank | Unable to get pH to pass cal on status and sens after clot removal |  | D28AE766EE131DBF | Cleaned | HW | pH/BG Module | Measuring chamber, Clot |  | Blank | Unable to get pH to pass cal on status and sens after clot removal
No | 4 | Y | RAME | F Customer reported by direct call, oximetry errors on calibrations. |  | Checked logs for errors, found 0316 Neon number too low and 0258 Neon intensity too low . Ran complete oximetry diagnostics and had failures on lamp unit. | D283B6EE2A54973B | Defect | HW | Oxi module | Lamp unit |  | F Customer reported by direct call, oximetry errors on calibrations. |  | Checked logs for errors, found 0316 Neon number too low and 0258 Neon intensity too low . Ran complete oximetry diagnostics and had failures on lamp unit.
Yes | Visit | Complaint | RDEU | Kal-Kassette TnT Lot 12771 | TnT-Kal-Kassette liefern | Kalkassette geliefert | SWO-191058 | Re classified RMED | Logistical upSWO |  |  |  | Kal-Kassette TnT Lot 12771 | Supply TnT-Kal cassette | Calf cassette supplied
No | Visit | Complaint | RIND | Calibration failed with Optical intensity out of range error | Cal data Checked | to replace Optical unit | SWO-190991 | Defect | HW | Process Module | Optical Unit | 1031 | Calibration failed with Optical intensity out of range error | Cal data Checked | to replace Optical unit
Yes | Visit | Complaint | RDEU | 10.05.2017 Frau Eschmann 020885083200 meldet, dass Teste nicht erkannt und auf die Reagenzien cass. nicht zugegriffen werden kann. | Gerät steht im Fehler, auf Reag.cassette kann nicht zugegriffen werden, Datum ist auf 01.01.2006 zurückgesprungen, | Gerät neu eingestellt, Schacht der Reag.cassette von alten Cups gereinigt, neue Reag.cassette eingesetzt, Gerät neu gestartet, | SWO-191153 | [5 empty] | 10.05.2017 Ms. Eschmann 020885083200 reports that tests are not detected and on the reagents cass. inaccessible. | Device is in error, Reag.cassette cannot be accessed, date has jumped back to 01.01.2006, | Device re-adjusted, shaft of Reag.cassette cleaned of old cups, new Reag.cassette inserted, device restarted,
No | Visit | Complaint | RDEU | 08.05.2017; Hr. Völker; Tel.: 015732231046; Kal und QC Messfehler 10.05.2017; Herr Völker, Tel.: 030434307030: weiterhin Kal und QC Fehler bei K und Laktat | Reinigungsarbeiten, Fehlersuche | Sensorkassette ersetzt, Startphase abgewartet, weitere Korrekte Kal abgewartet, QK ok, weitere Ansprachefehler des pCO2-Sensors -> Info an Kunden diesen Parameter zu beobachten und ggf. Rückruf an mich | SWO-191308 | Assisted Replacement | SC | SC | pCO2 |  | 08.05.2017; Mr. Völker; Tel.: 015732231046; Kal and QC measurement error 10.05.2017; Mr Völker, tel.: 030434307030: continue Kal and QC errors in K and lactate | Cleaning work, troubleshooting | Sensor cassette replaced, start phase awaited, further correct cal waited, QK ok, further speech errors of the pCO2 sensor -> Info to customers to observe this parameter and, if necessary, call back to me
No | 22 | Y | RAME | V Machine not working POC Lisa RRT |  | System will not power on. | D284F421A027715A | Defect | HW | Power Supply | Power Supply |  | V Machine not working POC Lisa RRT |  | System will not power on.
No | Visit | Complaint | RCHN | 试剂包漏夜导致接口板腐蚀 | 试剂包漏夜导致接口板腐蚀，借902-972一个 | 更换902-972一个正常。(By Ran) | SWO-190915 | [5 empty] | Reagent package leakage night causes interface plate corrosion | Reagent package leakage night causes interface plate corrosion, borrow 902-972 one | Replace 902-972 one normal. (By Ran)
No | Visit | Complaint | RCHN | 仪器无法启动 | 仪器无法启动，借902-970一个 | 更换902-970一个正常。(By Ran) | SWO-190687 | [5 empty] | The instrument does not start | The instrument does not start, borrow 902-970 one | Replace the 902-970 one normal. (By Ran)
No | Visit | Complaint | RCHN | PH,PO2定标漂移 | PH,PO2定标漂移，借备件902-972一个 | 更换902-972后修复(By Ran) | SWO-190667 | [5 empty] | PH, PO2 CALIBRATION DRIFT | PH, PO2 calibration drift, spare part 902-972 one | Repair after replacement 902-972 (By Ran)
No | Visit | Complaint | RCHN | 白屏 | 白屏，借903-365一个 | 更换903-365后修复(By Ran) | SWO-190732 | [5 empty] | White screen | White screen, borrow 903-365 one | Repair after replacement 903-365 (By Ran)
No | Visit | Complaint | RCHN | MYO定标失败 | MYO定标失败，借定标卡944-214（12481）一张 | 重新定标正常。(By Ran) Recal, OK. | SWO-190619 | Assisted Replacement | Consumables | Test Kit | Cartridge, Myo |  | MYO calibration failure | MYO calibration failure, borrowing bid card 944-214 (12481) one | Recalibration is normal. (By Ran) Recal, OK.
No | Visit | Complaint | RCHN | TNI定标失败 | TNI定标失败，借TNI定标卡944-212 2张（12798） |  | SWO-191248 | Assisted Replacement | Consumables | Test Kit | Cartridge, TNI |  | TNI calibration failed | TNI bid failure, borrow TNI calibration card 944-212 2 (12798)
No | 6 | Y | RPNZ | Blank | Printer Head Faulty |  | D2929DC1B21EDE5F | Defect | HW | Printer Module | Printer head |  | Blank | Printer Head Faulty
No | Visit | Complaint | RSAF | Not scanning all the ampules in the carousel. | Not scanning all the ampules in the carousel | Stripped and cleaned the entire AQC module scanner and mirror | SWO-193851 | Cleaned | HW | Autocheck module | Barcode Scanner |  | Not scanning all the ampules in the carousel. | Not scanning all the ampules in the carousel | Stripped and cleaned the entire AQC module scanner and mirror
Yes | 10 | Y | RCZE | Blank |  |  | D284E5FAA00C257C | [5 empty] | Blank
No | Visit | Complaint | RKK | 装置引き上げ | 元々はサンプル移動エラー対応。 上記ｴﾗｰ対応後、QC測定にてｴﾗｰ発生。 | 装置引き上げ | SWO-191438 | Unknown | Unknown | Unknown | Unknown |  | Equipment pulling up | Originally supported sample move error. After the above error response, an error occurs in QC measurement. | Equipment pulling up
No | Visit | Complaint | RDEU | 11.05.2017 10:00 Fr. Benkert 09117580992308: Kal-Fehler pCO2n 11.05.2017 10:50 Fr. Benkert 09117580992308: Kal-fehlr pCO2 | pCO2 Kal. und QC Fehler Glu und Lac des Öfteren Kal. Fehler | Proteinentfernung durchgeführt, Lufteinlass gespült, Datei Archiv Kal. und QC überprüft - Kal. und QC Fehler auch bei Glu, Lac und Na angezeigt, pCO2 nochmal neu membraniert - ohne Erfolg - Status zu gering, pCO2 Elektrode E788 getauscht, im Zuge des Einsatzes Glu, Lac und Ref. neu membraniert, bei Glu und Lac fehlt der untere Dichtring - Dichtring ersetzt, alle Messkammern gereinigt, Elektrodenkontakte und Connectoren mit Glasfaserpinsel gereinigt, Einlassnadel und Einlassdichtung aus Kundenbestand getauscht - Nadel war verbogen, LS Justage und Pumpen Kalibration durchgeführt, Nullstrom Kal. für pO2, Glu und Lac durchgeführt, System- und Funktionskontrolle in Ordnung | SWO-191570 | User error | Fluidic pathway | Fluidic pathway | Maintenance Parts |  | May 11, 2017 10:00 Fr. Benkert 09117580992308: Kal Error pCO2n May 11, 2017 10:50 Fr. Benkert 09117580992308: Kal-fehlr pCO2 | pCO2 Kal. and QC error Glu and Lac often Kal. Error | Protein removal performed, air inlet flushed, file archive Kal. and QC checked - Kal. and QC error also indicated at Glu, Lac and Na, pCO2 re-membraneed - without success - status too low, pCO2 electrode E788 exchanged, in the course of using Glu, Lac and Ref. newly membraneed, at Glu and Lac the lower sealing ring is missing - sealing ring replaced, all measuring chambers cleaned, electrode contacts and connectors cleaned with glass fiber brush, inlet needle and inlet seal replaced from customer stock - needle was bent, LS adjustment and Pumps calibration carried out, zero current cal. performed for pO2, Glu and Lac, system and function control in order
Yes | Visit | Complaint | RKK | 納品後、インレットCALエラー、リーク検出、Ca感度エラー、オキシメトリーハードウェアエラー多発 | 修理完了機搬入 | 装置返却作業 | SWO-191503 | Re classified RMED | OOB, although 27 days |  |  |  | After delivery, inlet CAL error, leak detection, Ca sensitivity error, oxymetry hardware error frequent | The machine was repaired and moved in. | Equipment return work
No | 6 | Y | RAME | Fails PO2 calibrations. |  | Customer replaced a new PO2 electrode and membrane. PO2 fails calibrations frequently. | D284E91FA0417FFF | Defect | HW | pH/BG Module | Measuring chamber, pO2 |  | Fails PO2 calibrations. |  | Customer replaced a new PO2 electrode and membrane. PO2 fails calibrations frequently.
No | 7 | Y | RAME | 6 CA is out of customer ranges for level 3 QC. |  | CA is out of customer range for Level3 QC. Replaced E Met electrode pump tubing, cleaned CA and Reference electrodes chambers and contacts. | D28501F9A00CBFC5 | Cleaned | HW | El/Met Module | Measuring chamber |  | 6 CA is out of customer ranges for level 3 QC. |  | CA is out of customer range for Level3 QC. Replaced E Met electrode pump tubing, cleaned CA and Reference electrodes chambers and contacts.
No | Visit | Complaint | RDEU | 10.05.2017; Fr. Hof; Tel.: 06044615906; Einlass tropft | Probenmaterial fließt über Einlassdichtung nicht komplett in die Abfallstrecke | Abfallstrecke, Abfallschläuche, Anschlussstücke der Einlassdichtung, Verschlauchung von Inlet zu Abfallpumpe und Einlassdichtung erneuert. Inlet- und LS Justage. Pumpen Kal ohne Beanstandung Kalibrationen durchgeführt. QC Messung und Testmessung ohne Beanstandung Verwendete Probennehmer sind Pico50 | SWO-191412 | Defect | HW | Basic Unit | Waste tubes |  | 10.05.2017; Ms. Hof; Tel.: 06044615906; Inlet drips | Sample material does not flow completely into the waste section via inlet seal | Waste section, waste hoses, connecting pieces of the inlet seal, hoseing from Inlet to waste pump and inlet seal renewed. Inlet and LS adjustment. Pumps Kal without complaint calibrations. QC Measurement and test measurement without complaint Samplers used are Pico50
No | Visit | Complaint | RDEU | 0351/8566313 Hr. Helm QC Fehler bei Level 1 und 3 Gluc und Crea und PCo2 31.03.2017 16:45 Hr. Helm Cre Kal Fehler nach Membran Tausch + PCO2 Qc Fehler 03.04.2017 Herr Helm 03518566313 meldet, dass einige Parameter die QC nicht bestehen. | Membranen und Lösungen geprüft | Alle Lösungen erneuert, Membraniert, offene Fragen geklärt, Schulung durchgeführt ( Geräteverantwortlicher ) Crea in Empf. 3 mit Fehlermeldung --> ABL war eine Woche in Ruhestellung--> erneuter Besuch am 12.05.2017 Display ohne Funktion --> nach Neustart wieder i.O --> Ersatzteil wird bestellt | SWO-191415 | Assisted Replacement | Consumables | Solution bottle | Crea |  | 0351/8566313 Mr. Helm QC Error at Level 1 and 3 Gluc and Crea and PCo2 31.03.2017 16:45 Mr. Helm Cre Kal Error after Membrane Swap + PCO2 Qc Error 03.04.2017 Mr Helm 03518566313 reports that some parameters do not pass the QC. | Membranes and solutions tested | All solutions renewed, membraned, open questions clarified, training carried out ( device manager ) Crea in Empf. 3 with error message --> ABL was a week in retirement--> another visit on 12.05.2017 Display without function --> after restart i.O --> Spare part is ordered
No | Visit | Complaint | RIND | Analyzer not Initializing it is showing error type 7 with message AU_BARCODE_SETU P_ERR | Checked and found that problem with vision control PCB. | need to replace vision control PCB and Cable. | SWO-192555 | Defect | HW | Aspiration Module | PCB, Vacu vision |  | Analyzer not Initializing it is showing error type 7 with message AU_BARCODE_SETU P_ERR | Checked and found that problem with vision control PCB. | need to replace vision control PCB and Cable.
No | Visit | Complaint | RBNL | AQT geeft errormelding 1029 (Optical sensitivity out of range) | Vervolg op bezoek van 9-5-2017. Bij de kalibratie van de TnT verscheen foutmelding 1031, dus de geplaatste nieuwe Optical Unit is niet goed. In overleg met Jan Guldberg uit CPH besloten zowel de Optical Unit als het Processwheel te vervangen. [calibrate factory reference]uitgevoerd, deze is oke. TnT opnieuw gekalibreerd, deze wordt afgebroken. AQT gecontroleerd op fouten en checkouts uitgevoerd, alles is goed. 5blanks op de DDIM gemeten, deze zijn goed. TnT weer gekalibreerd, nu is deze ook goed. Hierna DDIM en CRP gekalibreerd. Nogmaals de [calibrate factory reference] uitgevoerd, deze is wederom oke. QC's gemeten deze zijn allen goed. AQT is nu klaar voor gebruik. | Optical Unit en Processwheel vervangen. | SWO-191352 | Defect | HW | Process Module | Optical Unit | 1031 | AQT geeft errormelding 1029 (Optical sensitivity out of range) | Continuation of visit of 9-5-2017. When calibration of the TnT appeared error message 1031, so the placed new Optical Unit is not good. In consultation with Jan Guldberg from CPH, both the Optical Unit and the Processwheel were decided to replace. [Calibrate Factory reference] performed, this is okay. TnT is re-calibrated, it is aborted. AQT checked for errors and checkouts performed, all is well. 5blanks on the DDIM measured, these are good. TnT calibrated again, now this is good too. Hereinafter DDIM and CRP calibrated. Once again the [calibrate Factory reference] performed, this is again okay. Qcs measured these are all good. AQT is now ready to use. | Optical Unit and Processwheel replacement.
No | Visit | Complaint | RKK | 測定時エラーコード1029：光学感度範囲外の為、測定不可 | 測定時エラーコード1029：光学感度範囲外の為、測定不可 | 機器返却実施。 | SWO-194774 | Defect | HW | Process Module | Optical Unit | 1029 | Measurement error code 1029: Not measurable because it is outside the optical sensitivity range | Measurement error code 1029: Not measurable because it is outside the optical sensitivity range | Equipment return ed.
No | Visit | Complaint | RDEU | 09.05.2017; Sr. Christina; Tel.: 035126561426; pCO2 QC Messfehler Level C 09.05.2017 Frau Keller 035126561426: Nach Tausch des neuen LPs zeigt das Gerät einen QC-fehler LevelB bei pO2 09.05.2017 Frau Susanne Keller 035126561426: Alle Wartungsmaßnahmen ohne Erfolg | Überprüfung Datenbank | Datenbank des ABl geprüft --> i.O Ausdruck CC | SWO-191642 | Assisted Replacement | SP | SP | pO2 |  | 09.05.2017; Sr. Christina; Tel.: 035126561426; pCO2 QC Measurement Error Level C 09.05.2017 Mrs. Keller 035126561426: After replacing the new LPs, the device shows a QC error LevelB at pO2 05/09/2017 Mrs Susanne Keller 035126561426: All maintenance measures without success | Review Database | Database of OJ checked --> i.O Expression CC
No | Visit | Complaint | RDEU | 09.05.2017 08:25Uhr Ledbuhr 02863209937: QC-Fehler bei Level B und C 10.05.2017 09:30 Uhr, Frau Ledbuhr Tel.: 02863209936 Fehlermeldung: tHb-QC-Fehler bei Level B und C | THB Messung zu niedrig bei QCS | Kal., THB kal, QCs durchgeführt, OK | SWO-191802 | Adjustment | HW | Oxi module | tHb Cal |  | 09.05.2017 08:25Uhr Ledbuhr 02863209937: QC-Fehler bei Level B und C 10.05.2017 09:30 Uhr, Frau Ledbuhr Tel.: 02863209936 Fehlermeldung: tHb-QC-Fehler bei Level B und C | THB measurement too low at QCS | Kal., THB kal, QCs durchgeführt, OK
Yes | 3 | Y | RPOL | Blank |  |  | D2861D1A1642831C | Re classified RMED | No fact |  |  |  | Blank
No | Visit | Complaint | RDEU | 26.04.2017, Herr Michels, tel.: 070716063755: Scanner funktioniert nicht richtig 27.04.2017, Herr Michels, Tel.: 070716063755: Scanner weiterhin ohne Funktion | Scanner Konfiguration in Roma..ini. Überprüft.ok. Scanner geöffnet und gereinigt. War aber sauber. | Testscan ohne Erfolg..Scanner funktioniert liest aber keine Daten. SW neu installieren erfolgreich. Scanner List jetzt auch Pat . Etiketten. Initialisierung abgewartet.QC durchgeführt und in Ordnung. | SWO-191719 | Defect | SW | RMED SW | Barcode reader, display |  | 26.04.2017, Mr. Michels, tel.: 070716063755: Scanner does not work properly 27.04.2017, Mr. Michels, Tel.: 070716063755: Scanner still without function | Scanner configuration in Roma.. Ini. Checked.ok. Scanner opened and cleaned. But was clean. | Test scan without success.. Scanner works but does not read data. Reinstall SW successfully. Scanner List now also Pat . Labels. initialization awaited. QC performed and in order.
No | Visit | Complaint | RDEU | 09.05.2017, Frau Klingenberg, tel.: 0582421242: AQC Fehler pO2 in Level A und C 11.05.2017; Fr. Weißbach; Tel.: 0582421699; wieder pO2 QC Messfehler | System überprüft => LP defekt und Bügelventil schwergänig | LP erneuert und Bügelventil gangbar gemacht $ Kalibration und QC durchgeführt | SWO-191590 | Defect | HW | Oxi module | Pinch Valve |  | 05/09/2017, Mrs. Klingenberg, tel.: 0582421242: AQC Error pO2 in Level A and C 11.05.2017; Ms. Weißbach; Tel.: 0582421699; again pO2 QC measurement error | System checked => LP defective and ironing valve heavy-duty | LP renewed and ironing valve made workable . . . . . . . . . . . . . . . . . . . . . . . . . . . . .
No | Visit | Complaint | RSCH | Blank | Erreur vide pression : code 1063 malgré le remplacement du pack de solutions | Remplacé le kit de maintenance annuel e.o Nettoyage système fluidique e.o Calibration e.o contrôle de qualité e.o | SWO-191730 | Defect | HW | Basic Unit | PM Kit |  | Blank | Empty pressure error: code 1063 despite replacement of solution pack | Replaced the annual maintenance kit e.o Cleaning fluid system e.o Calibration e.o quality control e.o
No | 3 | Y | RFRA | Blank | ERREUR VIDE TEST DE PRESSION |  | D28A19A2782F785B | Defect | HW | Auto Inlet module | Inlet Gasket |  | Blank | EMPTY PRESSURE TEST ERROR
No | Visit | Complaint | RIND | Analysis disabled. Insufficient solution pack cycle. New solution pack VC failed Message. | Analysis disabled. Insufficient solution pack cycle. New solution pack VC failed Message. | Sensor Cassette flushed and solution pack reset. Cleaned both. Reinstall SP and SC. Analyser serviced and calibrated properly. Handed over in a fine working condition. | SWO-191801 | [5 empty] | Analysis disabled. Insufficient solution pack cycle. New solution pack VC failed Message. | Analysis disabled. Insufficient solution pack cycle. New solution pack VC failed Message. | Sensor Cassette flushed and solution pack reset. Cleaned both. Reinstall SP and SC. Analyser serviced and calibrated properly. Handed over in a fine working condition.
No | 65 | Y | RESP | SGD OO CONECTADO A UNIDAD DE ANALISIS |  | CHEQUEO DE ANALYZER CONTROL PCB, TENSIONES DE LA POWER SUPPLY Y VENTILADOR INLET SIENDO TODOO OK Y PERSISTIENDO EL PROBLEMA. SE DESMONTA CADA UNO DE LOS 4 MICRO FUSIBLES Y SE LIMPIAN LOS CONTACTOS DE ÉSTOS. REINICIO DE EQUIPO Y SOLUCIÓN DEL TEMA. | D28632481661D62A | Defect | HW | Basic Unit | PCB, Analyzer control |  | SGD OO CONNECTED TO ANALISIS UNIT |  | ANALYZER CONTROL PCB CHECK, POWER SUPPLY TENSIONS AND INLET FAN BEING ALL OK AND PERSISTING THE PROBLEM. EACH OF THE 4 MICRO FUSIBLES IS DISMOUNTED AND THE CONTACTS OF THESE ARE CLEANED. RESTARTING OF EQUIPMENT AND SOLUTION OF THE THEME.
No | Visit | Complaint | RKK | pCO2 QC測定値 変動 | 5/2ユーザー作業にて、Glu・Lac・Ref・Clメンブラン交換後からQC pCO2測定値が変動した様子 | Clメンブラン取り付け直し ｐHシーリングリング取り付け直し タンパク除去 エアライン・測定チャンバー清掃 LS調整・ポンプCAL確認 CAL・QC測定データ確認 | SWO-191877 | Cleaned | HW | pH/BG Module | Measuring chamber |  | pCO2 QC readings Variation | Qc pCO2 readings fluctuated after the replacement of Glu, Lac, Ref, and Clmenbran in 5/2 user work | Cl Membrane Reinstallation Reseating the pH sealing ring Protein removal Airline measurement chamber cleaning LS adjustment and pump CAL confirmation Cal/QC measurement data confirmation
No | Visit | Complaint | RKK | pH 応答不良 多発 | pH 応答不良 多発 | pH電極交換 エアライン・測定チャンバー清掃 LS調整・ポンプCAL確認 タンパク除去、Glu・Lac・Refメンブラン交換 CAL・QC測定データ確認 | SWO-191876 | Assisted Replacement | Electrode | Electrode | pH |  | pH Bad Response Multiple | pH Bad Response Multiple | pH electrode replacement Airline measurement chamber cleaning LS adjustment and pump CAL confirmation Protein removal, Glu Lac, Ref membrane replacement Cal/QC measurement data confirmation
No | Visit | Complaint | RHUN | Lámpa hiba | Oxi modul hibaüzenetek: Neon number too low , Neon intenzity too low | Neon lámpa csere után megszűnt a hibaüzenet. Oxi tesztek, kalibrálás, QC mérés. | SWO-191717 | Defect | HW | Oxi module | Lamp unit | Neon | Lamp error | Oxy Hibaüzenetek mode: Neon number too low, Neon intenzity too low | After a Neon lamp replacement, the error message is resolved. Oxy tests, calibration, QC measurement.
No | Visit | Complaint | RDEU | QC Fehler | QC Fehler bei pCO und pO2 Kal. Werte sind Ok QC bei pCO2 und pO2 sind zu hoch | pCO2 und pO2 neu membraniert, Nullstrom Kal. für pO2, Glu, Lac und Crea durchgeführt, Messkammern gereinigt, Lufteinlass gespült, Einlassnadel überprüft - ist gerade, System- und Funktionskontrolle in Ordnung | SWO-191649 | Assisted Replacement | Electrode | Membrane | pCO2 |  | QC Fehler | QC Fehler bei pCO und pO2 Kal. Werte sind Ok QC bei pCO2 und pO2 sind zu hoch | pCO2 and pO2 re-membraneed, zero current Cal. for pO2, Glu, Lac and Crea performed, measuring chambers cleaned, air inlet flushed, inlet needle checked - is straight, system and function control in order
No | Visit | Complaint | RDEU | 0351/8566313 Hr. Helm QC Fehler bei Level 1 und 3 Gluc und Crea und PCo2 31.03.2017 16:45 Hr. Helm Cre Kal Fehler nach Membran Tausch + PCO2 Qc Fehler 03.04.2017 Herr Helm 03518566313 meldet, dass einige Parameter die QC nicht bestehen. | Siehe SWO | Elektroden über Kreuz getauscht --> Fehler 959 erscheint nicht --> beide Elektroden erneuert und neu membraniert --> Kal und QC durchgeführt i.O | SWO-191641 | User error | Electrode | Electrode | Crea mixed | Email | 0351/8566313 Mr. Helm QC Error at Level 1 and 3 Gluc and Crea and PCo2 31.03.2017 16:45 Mr. Helm Cre Kal Error after Membrane Swap + PCO2 Qc Error 03.04.2017 Mr Helm 03518566313 reports that some parameters do not pass the QC. | See SWO | Electrodes exchanged over cross --> Error 959 does not appear --> both electrodes renewed and re-membraneed --> Kal and QC performed i.O
No | 13 | Y | RAME | F Customer getting variance in qc results for oximetery. |  | Had Keith Clark on site and looked through the resulst for qc. Found too great of varience in results. Ordered new spectrometer with fiber optic cable and replaced. Powered up and had neon voltage too low errors on spectrometer. DOA. Replaced oxi complete with old analyzer from 693N004 with customers permission and started up fine. Matched lamp to new unite. tHb at 555 pA. Other analyzer at 560 pA. Ran all levels of controls. Level 2 FCOHb not passing at 3.2. Analyzer range set to 2.0-2.8. Radometer range 1-5 with mean at 3. Lot 147. All others pass. Spoke with operator and no manager on duty. Rouella said not to adjust ranges and to make sure the qc passes. Not able to pass level 2. Ordered oxi complete for monday delivery. Analyzer oxi down. | D286596216363552 | Revisit required | HW | Oxi module | Oxi module |  | F Customer getting variance in qc results for oximetery. |  | Had Keith Clark on site and looked through the resulst for qc. Found too great of varience in results. Ordered new spectrometer with fiber optic cable and replaced. Powered up and had neon voltage too low errors on spectrometer. DOA. Replaced oxi complete with old analyzer from 693N004 with customers permission and started up fine. Matched lamp to new unite. tHb at 555 pA. Other analyzer at 560 pA. Ran all levels of controls. Level 2 FCOHb not passing at 3.2. Analyzer range set to 2.0-2.8. Radometer range 1-5 with mean at 3. Lot 147. All others pass. Spoke with operator and no manager on duty. Rouella said not to adjust ranges and to make sure the qc passes. Not able to pass level 2. Ordered oxi complete for monday delivery. Analyzer oxi down.
No | Visit | Complaint | RDEU | 29.03.2017, Frau Schäfer, tel.: 091318543855: Autocheck Mechanik Fehler, Frau Schäfer hat die Chargen geprüft, Die Autocheck Türe geht auf, Karussell bleibt liegen. 31.03.2017 9:00 Fr. Schaefer 091318543855 QC werden nicht automatisch gemessen sondern nur wenn jemand vorm Gerät Fortsetzen drückt | AutoCheck Mechanik Fehler | Proteinentfernung durchgeführt, neues AutoCheck Modul (R2709N001) gegen Leihmodul ausgetauscht, AutoCheck Modul im Service überprüft, mehrere Kal. durchgeführt, tHb Kal. durchgeführt, QC Messungen gestartet, AutoCheck Modul konditioniert Ampullen wie gewohnt, System- und Funktionskontrolle in Ordnung | SWO-191691 | Defect | HW | Autocheck module | Autocheck module |  | 29.03.2017, Mrs. Schäfer, tel.: 091318543855: Autocheck mechanics error, Mrs. Schäfer has checked the batches, The car check door opens, carousel remains. 31.03.2017 9:00 Ms. Schaefer 091318543855 QC are not measured automatically but only when someone presses continue in front of the device | AutoCheck Mechanics Error | Protein removal performed, new AutoCheck module (R2709N001) replaced with rental module, AutoCheck module checked in service, several cal. tHb Kal. performed, QC measurements started, AutoCheck module conditioned ampoules as usual, system and function control in order
Yes | Visit | Complaint | RCHN | NA,HCT定标失败 | NA,HCT定标失败 | 更换测试卡 | SWO-191816 | [5 empty] | NA, HCT calibration failed | NA, HCT calibration failed | Replace the test card
No | Visit | Complaint | RCHN | 试剂包支架损坏 | 试剂包支架损坏，借886-776一个。 |  | SWO-191790 | Defect | HW | Basic Unit | Actuator |  | Damage to reagent package stent | The reagent package bracket is damaged, borrow 886-776 one.
No | Visit | Complaint | RCHN | 数据传输失败 | 数据传输失败，借910-342一个。 | 更换910-342后修复 by andrew | SWO-191191 | Defect | HW | DMS_PC | USER INTERFACE MODULE |  | Data transfer failed | Data transfer failed, borrowing 910-342 one. | Repair by andrew after replacing 910-342
No | Visit | Complaint | RCHN | PCO2质控超统计范围 | PCO2质控超统计范围 | 更换测试卡正常。 | SWO-192939 | Assisted Replacement | SC | SC | pCO2 |  | PCO2 QUALITY CONTROL OVER-STATISTICAL RANGE | PCO2 QUALITY CONTROL OVER-STATISTICAL RANGE | The replacement test card is OK.
No | Visit | Complaint | RCHN | 测试卡初始化失败 | 测试卡初始化失败 | 更换测试卡 | SWO-192675 | [5 empty] | Test card initialization failed | Test card initialization failed | Replace the test card
No | Visit | Complaint | RDEU | Kunde hat nach SC Wechsle pCO2 Sensor Ansprache Fehelr | nach SC Wechsle pCO2 Kal Empf bei 79% und Sensor Ansprache Fehler QC Fehler in Level A bei pCO2, Cl, Glu Gerät hat in letzter Zeit lt. Datenbank häufig Kal Fehler bei verschieden SC & LP | Kal wiederholt i.O. fehlerhafte QC Level A wiederholt Kal Fehler bei pCO2 während QC Messung SC Fach im Serviceprog. geöffnet, es scheint es seine die Pins nicht alle auf einer höhe, Sensor Interface PCB bestellt ->info an Kunden rep. am Mittwoch | SWO-192057 | Defect | HW | Basic Unit | PCB, Sensor interface |  | Customer has according to SC Wechsle pCO2 sensor address Fehelr | according to SC Change pCO2 Kal Empf at 79% and Sensor Address Error QC Error in Level A at pCO2, Cl, Glu Device has lately according to database frequently Kal errors at different SC & LP | Kal repeats I.O. faulty QC Level A repeats Kal error in pCO2 during QC measurement SC Fach in the service prog. open, it seems its the pins not all ordered on a height, sensor interface PCB ->info to customers rep. on Wednesday
No | Visit | Complaint | RDEU | Tel. Service Einlass erneuert | Gerät meldet wieder Einlassdichtung wechseln und dann LP wechseln. -Garantie- | ABL90 durchgespült und Teile im Umfang eines Jahreskits erneuert. Einlasssonde zur Bevorratung übergeben. ( Einlasssonde berechnen ) | SWO-192234 | Defect | HW | Inlet module | Inlet Gasket |  | Tel. Service Entrance renewed | Device reports again changing inlet seal and then changing LP. -Guarantee | ABL90 rinsed and parts in the scope of an annual kit renewed. inlet probe for stockpiling. ( Calculate inlet probe )
No | 4 | Y | RFRA | Blank | problème de rendu du paramètre bilirubine |  | D289E54A78621B8E | Defect | HW | Oxi module | Spectrometer |  | Blank | rendering problem of the bilirubin setting
No | Visit | Complaint | RDEU | 12.05.2017; Fr. Theinert; Tel.: 06221182546; Abfallweg ist verstopft | Abflusstrichter übergelaufen. Bei Service Beginn laut Anwender ist Abfluss besser. Ursache evtl. auch Rollenkranz der Förderpumpe. | Ausbau der Flüsdigkeits-Verschlauchung und Erneuerung durch einen Jahreswartungs-Kit. Abfalltrichter gereinigt. Flüssigkeit läuft gut ab, wird gut abgesaugt. | SWO-192153 | [5 empty] | 12.05.2017; Ms. Theinert; Tel.: 06221182546; Waste path is clogged | Drain hopper overflowed. At service start according to user, outflow is better. Cause possibly also roller ring of the pump. | Removal of fluid hose and renewal by an annual maintenance kit. Waste hopper cleaned. Liquid runs well, is well extracted.
No | Visit | Complaint | RIND | Received a call from Joshi Hospital, Pune that the machine is throwing Temperature Error even if no sensor Cassette is connected. | Need to check the Sensor Connector Board for any damage, as technicians have reported that leakage from the back of newly installed SC was observed. | Found that the Sensor connector board is damaged, tried installing new SC, same error was observed, the machine is displaying SC temp not stable even if no SC is connected. Required Sensor Connector Board as well as Signal Data Board for testing purpose. | SWO-192040 | [5 empty] | Received a call from Joshi Hospital, Pune that the machine is throwing Temperature Error even if no sensor Cassette is connected. | Need to check the Sensor Connector Board for any damage, as technicians have reported that leakage from the back of newly installed SC was observed. | Found that the Sensor connector board is damaged, tried installing new SC, same error was observed, the machine is displaying SC temp not stable even if no SC is connected. Required Sensor Connector Board as well as Signal Data Board for testing purpose.
No | Visit | Complaint | RCHN | 无法读取或写入试剂包芯片信息 | 无法读取或写入试剂包芯片信息，借903-158一个 |  | SWO-191962 | Defect | HW | Basic Unit | Solution Pack ID PCB | 1134 | Unable to read or write package chip information | Unable to read or write reagent package chip information, borrow 903-158 one
No | Visit | Complaint | RCHN | 定标底物外露检查失败 1024 | 定标底物外露检查失败 1024 ，用备用光单903-364R0156n004 |  | SWO-191854 | Defect | HW | Process Module | Optical Unit | 1024 | Calibration substrate exposure check failed 1024 | Calibration substrate exposure check failed 1024 , with spare light sheet 903-364R0156n004
Yes | 11 | Y | RPAC | Blank | Ref electrode-Na, Cl and Ca drifting after cleaning |  | D28A10EE781E3946 | [5 empty] | Blank | Ref electrode-Na, Cl and Ca drifting after cleaning
No | Visit | Complaint | RDEU | Touchscreen lässt sich teilweise nicht bedienen--> nach Neustart wieder i.O | Touchscreen geprüft | Display-unit defekt ( keine Bedienung möglich $ Touchscreen funktioniert nicht ) --> Display unit erneuert | SWO-191984 | Defect | HW | DMS_PC | Touch screen |  | Touchscreen can partially not be operated--> after reboot i.O | Touchscreen checked | Display-unit defective (no operation possible - Touchscreen does not work ) --> Display unit renewed
No | 14 | Y | RAME | F FCOHb running high on analyzer. Return visit to replace the oximetry module due to DOA spectrometer. |  |  | D28A5D767848BF27 | Defect | HW | Oxi module | Oxi module, Replaced | FCOHb | F FCOHb running high on analyzer. Return visit to replace the oximetry module due to DOA spectrometer.
No | Visit | Complaint | RDEU | 21.04.2017; Fr. Gansauer; Tel.: 05861838001; Testcatridges werden nicht richtig erkannt | System überprüft => Scanner falsch eingestellt | Scanner neu justiert Fax 05861838099 | SWO-192168 | Adjustment | HW | Storage Module | Barcode reader, Storage wheel |  | 21.04.2017; Ms. Gansauer; Tel.: 05861838001; Testcatridges are not detected correctly | System checked => Scanner set incorrectly | Scanner re-adjusted Fax 05861838099
No | Visit | Complaint | RDEU | 11.05.2017, Frau Fadel, Tel.: 03026383308: Anmeldung nur über Manager möglich; | keine Anwenderanmeldung mgl. | Fehlersuche; div. Sicherheitseinst. angepaßt; Anwender klärt Zugriffsrechte mit Labor Limbach | SWO-192002 | User error | SW | User access | User access |  | 11.05.2017, Mrs. Fadel, Tel.: 03026383308: Registration only possible via manager; | no user login mgl. | Troubleshooting; div. Security one. adapted; User clarifies access rights with Limbach laboratory
Yes | 6 | Y | RPOL | Blank |  |  | D28B67D7EE15A2FF | [5 empty] | Blank
Yes | 6 | Y | RPOL | Blank | Awaria klapy wejsciowej kasety sensorowej. Awaria usuniêta, wymiana SC90 LOT : R0621. Analizator sprwny technicznie |  | D28B67D8EE31DBE2 | [5 empty] | Blank | Sensor cassette input flap failure. Fault removed, replacement SC90 LOT: R0621. Technical Analyzer
No | 10 | Y | RAME | O user called stating that Failed Proficiency Test for THB |  | checked out the report and found that they failed 3 out of 5 for THB | D28B4B86EE470983 | Adjustment | HW | Oxi module | tHb Cal |  | O user called stating that Failed Proficiency Test for THB |  | checked out the report and found that they failed 3 out of 5 for THB
No | 22 | Y | RAME | V Leak detected error 1184 |  | Pinch valve not opening $closing. | D28B37C0EE54E22C | Defect | HW | Oxi module | Pinch valve | 1184 | V Leak detected error 1184 |  | Pinch valve not opening $closing.
No | Visit | Complaint | RIND | Lac parameter is failed in SC. | Cleaned Sensor cassette with distilled water. Also checked for sample clot. | Kept in fridge for 30 minutes and reinstalled sensor cassette. But still it is showing Lac sensitivity low. Need to change SC. | SWO-193205 | [5 empty] | Lac parameter is failed in SC. | Cleaned Sensor cassette with distilled water. Also checked for sample clot. | Kept in fridge for 30 minutes and reinstalled sensor cassette. But still it is showing Lac sensitivity low. Need to change SC.
No | 3 | Y | RAME | P Possible blockage |  | pink fluid all over counter. Paged FSE | D28B0498EE3DD0A9 | [5 empty] | P Possible blockage |  | pink fluid all over counter. Paged FSE
No | Visit | Complaint | RCHN | 无法开机 | 无法开机，借910-285一个 |  | SWO-192188 | Defect | HW | Power Supply | Power Supply |  | Cannot power on | Cannot power on, borrow 910-285 one
No | Visit | Complaint | RCHN | CKMB定标失败 | CKMB定标失败，借CKMB定标卡944-216（12758）一张 |  | SWO-192282 | Defect | Consumables | Test Kit | Cartridge | 5784 | CKMB calibration failed | CKMB calibration failure, debit CKMB calibration card 944-216 (12758) one
No | Visit | Complaint | RDEU | 15.05.2017 16:30Uhr Frau Patenheim 04421891871: Wartungsmaßnahmen greifen nicht, weiterhin Touchscreen funktioniert schlecht. | Touchscreen teilweise ohne Funktion | nach Neustart Display wieder in Funktion, vermutlich CPU Unit defekt, ET bestellt | SWO-192451 | Defect | HW | DMS_PC | Touch screen |  | 15.05.2017 16:30Pm Ms Patenheim 04421891871: Maintenance measures do not work, touchscreen still works badly. | Touchscreen partially without function | after restart display back in function, presumably CPU unit defective, ET ordered
No | 10 | Y | RAME | Z Customer called with touch screen issues. Remote session could not resolve and customer requested on site service. |  | Touch screen is not out of calibration but is not responding correctly. The number 6 and letter G are populating when they should not be and scroll arrows are not working. PO2 zero errors found when checking the logs for performance. | D28B3F7DEE5374B8 | Adjustment | HW | DMS_PC | Touch screen |  | Z Customer called with touch screen issues. Remote session could not resolve and customer requested on site service. |  | Touch screen is not out of calibration but is not responding correctly. The number 6 and letter G are populating when they should not be and scroll arrows are not working. PO2 zero errors found when checking the logs for performance.
No | 6 | Y | RFRA | Blank | Valeurs THB à ... |  | D28D9053DA427BDC | Revisit required | HW | Oxi module | tHb Cal | Air in Fluidic Pathway | Blank | THB values to ...
No | 6 | Y | RAME | 2 PCO2 fails level 1&2 QC |  | Fails PCO2 on level 1&2 QC on the lower range of packet insert. PCO2 electrode was remembraned and sill the same issue. | D28B751CEE45CE1C | Assisted Replacement | Electrode | Membrane | pCO2 | pH | 2 PCO2 fails level 1&2 QC |  | Fails PCO2 on level 1&2 QC on the lower range of packet insert. PCO2 electrode was remembraned and sill the same issue.
Yes | Visit | Complaint | RDEU | 15.05.2017; Hr. Kurscheid, Tel.: 02241187361; Parameter Glu fehlt bei dem neuen ABL | Elektrode fehlt | Glu installiert. Mehrfavh kalibriert. | SWO-192449 | Re classified RMED | Glu missing |  |  |  | 15.05.2017; Mr. Kurscheid, Tel.: 02241187361; Parameter Glu is missing from the new ABL | Electrode is missing | Glu installed. Multifavh calibrated.
No | Visit | Complaint | RDEU | Wiederholt Qualitätskontroll Grenzverletzung bei Parameter TnT. Unterstützung und Prüfung durch Techniker Kollege aus CPH. | Techniker Kollege aus CPH hat Optische Einheit überprüft, ob zentral gemessen wird im Testcup. Überprüfung über spezielles Tool und Testcup. Messung nicht zentral im Testcup. Daten gezogen und Service Dump. | Tausch der Optischen Einheit und des Process Wheels. Überprüfung ob Zentral gemessen wird. Werte besser wie vorher. Neukalibration der beiden Parameter TnT und D- Dimer. Testkits aus der Geschäftsstelle vorher geholt. Qualitätskontrolle außer Grenzen und höher gemessen als vorher. Einbau der alten Optischen Einheit und Ausbau der neuen. Erneute Kalibration der beiden Parameter. Qualitätskontroll Werte wieder in Ordnung. | SWO-192506 | Adjustment | HW | Process Module | Process wheel |  | Repeated quality control boundary violation at parameter TnT. Support and verification by technician colleague from CPH. | Technician colleague from CPH has checked optical unit whether centrally measured in the test cup. Check via special tool and test cup. Measurement not central in the Test Cup. Data dragged and service dump. | Replacement of the Optical Unit and the Process Wheel. Check whether central is measured. Values better than before. Recalibration of the two parameters TnT and D- Dimer. Test kits from the office beforehand. Quality control out of bounds and measured higher than before. Installation of the old optical unit and extension of the new one. Re-calibration of the two parameters. Quality control values back in order.
No | Visit | Complaint | RDEU | 24.04.2017; Fr. Dohmes; Tel.: 06946086125; QC Messfehler - Fehler vorhanden in Messung mit LQC-Lösung in Position | Überprüfung des Analysators durch Technik Kollege aus Kopenhagen. Überprüfung der Optischen Einheit, ob sie zentral im Testcup misst. Arbeiten am AQT anbei. | Techniker Kollege hat noch Testmessungen durchgeführt um den Zustand der Optischen Einheit zu prüfen. Tausch von Hardware heute aus Zeitgründen nicht mehr möglich. | SWO-192507 | Revisit required | HW | Process Module | Optical Unit | 211 | 24.04.2017; Ms. Dohmes; Tel.: 06946086125; QC measurement error - error present in measurement with LQC solution in position | Review of the analyzer by technical colleague from Copenhagen. Check the optical unit to see if it measures centrally in the Test Cup. work on the AQT. | Technician colleague has still carried out test measurements to check the condition of the optical unit. Replacement of hardware is no longer possible due to time constraints.
No | Visit | Complaint | RDEU | TnI Werte eines Patienten weichen um ca. 0,11 bis 0,17 nach unten ab Messung 1: 15.05.17-10.12 Uhr AQT = 0,017, Labor = 0,131 ng/ml Messung 2: 15.05.17-15.45 Uhr AQT = 0,024, Labor = 0,196 ng/ml | Systemcheck Service Dump und Datensicherung | Blutbilder gereinigt, Schlauchsystem mit Hypochlorid gespült, Sensoren gereinigt, Piercingtool und Wash-Unit gereinigt, Processwheel gereinigt 5 Blankmessungen - ok Photocounts ok 3 mal Level 1 LQC TnI - ok Level 2 und Level 3 LQC - TnI ok | SWO-192533 | Cleaned | HW | Process Module | Process wheel |  | TnI values of a patient deviate by approximately 0.11 to 0.17 downwards Measurement 1: 15.05.17-10.12 O'clock AQT = 0.017, laboratory = 0.131 ng/ml Measurement 2: 15.05.17-15.45 O'clock AQT = 0.024, laboratory = 0.196 ng/ml | System check Service Dump and Data Backup | Blood images cleaned, hose system flushed with hypochloride, sensors cleaned, piercing tool and wash unit cleaned, processwheel cleaned 5 blank measurements - ok Photocounts ok 3 times Level 1 LQC TnI - ok Level 2 and Level 3 LQC - TnI ok
No | Visit | Complaint | RDEU | Systemtemperatur außer Bereich | Top Temperatur 39,46 Bottom temperature 28,13 Substrate temperature 30,16 Dabei ist die Sensorkassette beschädigt | Ein und ausschalten des Anlysators | SWO-192835 | Assisted Replacement | SC | SC | Thermistor |  | System temperature out of range | Top Temperature 39.46 Bottom temperature 28.13 Substrate temperature 30.16 The sensor cartridge is damaged | Switching the analyzer on and off
No | Visit | Complaint | RLTD | AMELIA DAVIES 07500 225 828 | Analyser failing with leak detected. Failing vacuum pressure test value too high. Fault traced to small hole in Oxi pinch valve tube. Pinch tube replaced as part of service visit also performed today. | Replace Oxi pinch valve tube. | SWO-192829 | Defect | HW | Oxi module | Pinch valve tube | 1342 | AMELIA DAVIES 07500 225 828 | Analyser failing with leak detected. Failing vacuum pressure test value too high. Fault traced to small hole in Oxi pinch valve tube. Pinch tube replaced as part of service visit also performed today. | Replace Oxygen Pinch valve tube.
No | Visit | Complaint | RDEU | EvB Potsdam Labor meldet, dass der Pat. ID Barcode nicht erkannt wird, Auftrags ID Barcode wird erkannt | Pat Barcodeswerden nicht erkannt, verwendeter Barcodetyp ist nicht bekannt | Bei eintreffen ist Radianceverbindung nicht aktiv-> Haken für Kommunikationsverbindung entfernt $gesetzt -> keine Verbessung -> Haus EDV nicht erreichbar bemängelte Barcodes eingescannt-> werden nicht erkannt, Druck des Barcodes auf dem Etikettenrand, externe Scanner angeschlossen -> Barcodes werden nicht erkannt -> Info an Station das die Qualität der Barcoddruck zu prüfen ist-> Rücksprache mit Hr. Gebhard in EvB Potsdam Gerät neu gestartet, Netzwerkabel geprüft-> Verbindung in Netzwerkdose lose -> befestigt ( Wastebehälter steht davor und drückt dagegen)-> Info an Stat.leitung das Kabel durch Haus EDV ersetzt werden soll, Pat ID Code per Hand eingegeben -> Daten kommen im Soaris an Labor: Telefonat mit Fr. Roland: Problem der Qualität der Barcodeetiketten ist bekannt und soll behoben werden | SWO-192731 | User error | HW | Cable | Cable |  | EvB Potsdam Labor reports that the Pat. ID barcode is not recognized, order ID barcode is detected | Pat barcodes are not recognized, used barcode type is not known | When ireads arrive, radiance connection is not active-> hook for communication connection removed $gesetzt -> no improvement -> house computer not reachable criticized barcodes scanned-> are not recognized, printing of the barcode on the label edge, external scanners are connected -> barcodes are not recognized -> Info to station that the quality of barcod printing is to be checked-> Consultation with Mr. Gebhard in EvB Potsdam Device restarted, network-abel checked-> connection in network socket loose -> fixed (waste container stands in front of it and presses against it)-> Info to Stat.leitung the cable should be replaced by house COMPUTER, Pat ID code entered by hand -> Data arrive in the Soaris Laboratory: Phone call with Ms. Roland: Problem of the quality of the barcode labels is known and should be fixed
No | Visit | Complaint | RDEU | 03.05.2017 11:30Uhr Herr Antrest 095150315241: Die Uhrzeit wurde korrigiert, aber nach kurzer Zeit ist die Uhrzeit wieder um eine Stunde verstellt. 17.05.2017; Hr. Söllner (Med. Technik), Tel.: 095150315242; die Uhrzeiz verstellt sich nach wie vor. | Fehlersuche --> Einstellungen im Control Panel (Date and Time) überprüft | Date and Time Einstellungen auf Windows-Ebene angepasst, Speicherbefehl ausgeführt, 2 Neustart des Analysators durchgeführt --> keine Zeitverstellung mehr feststellbar, System und Funktion i.O. | SWO-192664 | Setup | SW | Time setting | Time setting |  | 05/03/2017 11:30AM Mr. Antrest 095150315241: The time has been corrected, but after a short time the time is adjusted again by one hour. 17.05.2017; Mr. Söllner (Med. Technik), tel.: 095150315242; the clock line is still disfigured. | Troubleshooting --> Checks settings in the Control Panel (Date and Time) | Date and Time settings adapted at the Windows level, memory command executed, 2 restart of the analyzer performed --> no time adjustment can be determined anymore, system and function i.O.
No | Visit | Complaint | RDEU | 17.05.2017; Hr. Söllner; Tel.: 095150315242; Uhrzeit verstellt sich immer | Fehlersuche --> Einstellungen im Control Panel (Date and Time) überprüft | Date and Time Einstellungen auf Windows-Ebene angepasst, Speicherbefehl ausgeführt, 2 Neustart des Analysators durchgeführt --> keine Zeitverstellung mehr feststellbar, System und Funktion i.O. | SWO-192663 | Setup | SW | Time setting | Time setting |  | 17.05.2017; Mr. Söllner; Tel.: 095150315242; Time is always changed | Troubleshooting --> Checks settings in the Control Panel (Date and Time) | Date and Time settings adapted at the Windows level, memory command executed, 2 restart of the analyzer performed --> no time adjustment can be determined anymore, system and function i.O.
No | 22 | Y | RAME | V Blank display. |  | Power supply LEDs not all on. Battery voltage higher than warning threshold. | D28C7D8B6401C76F | Defect | HW | Power Supply | Power Supply |  | V Blank display. |  | Power supply LEDs not all on. Battery voltage higher than warning threshold.
No | Visit | Complaint | RDEU | 16.05.2017; Fr. Krause-Reinbrodt (Labor Stein), Hr. Klein (ITS); Tel.: 0213152953888; keine Netzwerkverbindung vorhanden | Keine Netzwerkverbindung vorhanden. | Netzwerkeinstellungen, Speicherbefehl und die Startphase durchgeführt. | SWO-192822 | Setup | SW | His/lis | His/lis |  | 16.05.2017; Ms. Krause-Reinbrodt (Laboratory Stone), Mr. Klein (ITS); Tel.: 0213152953888; There is no network connection | There is no network connection. | network settings, memory command, and the boot phase.
No | Visit | Complaint | RDEU | ABL steht in Fehlerbehebung | Analysator stand in Fehlerbehebung | Fließwege gespült, Startphase durchlaufen. Ventilfunktion im Service überprüft, Ventil leichtgängig. Mehrere Testmessungen durchgeführt, Funktion OK. | SWO-192920 | Cleaned | HW | Oxi module | Fluidic pathway |  | ABL is in troubleshooting | Analyzer was in troubleshooting | Flow paths rinsed, go through the start phase. Valve function checked in service, valve smooth. Several test measurements performed, function OK.
No | Visit | Complaint | RDEU | Mixer Tray gebrochen | Riss/ Kunststoffbruch im Mixer Tray | Mixer Tray erneuert. | SWO-192787 | Defect | HW | Basic Unit | Mixer Cover |  | Mixer Tray broken | Crack/ plastic break in the mixer tray | Mixer Tray renewed.
Yes | Visit | Complaint | RCHN | NA,HCT,LAC定标失败 | NA,HCT,LAC定标失败 | 连续更换2张测试卡后修复 | SWO-193278 | [5 empty] | NA, HCT, LAC calibration failed | NA, HCT, LAC calibration failed | Repair after 2 consecutive test cards
No | 17 | Y | RAME | I SC temperature errors on several SC. |  | Ran zero calibration test and passed, checked sensor connector board for damage and none was found. | D28C78566460B0F7 | [5 empty] | I SC temperature errors on several SC. |  | Ran zero calibration test and passed, checked sensor connector board for damage and none was found.
No | Visit | Complaint | RCHN | LAC定标失败 | LAC定标失败 | 更换测试卡 | SWO-192599 | [5 empty] | LAC calibration failure | LAC calibration failure | Replace the test card
No | Visit | Complaint | RCHN | 触摸屏按键触摸不准 | 触摸屏按键触摸不准 | 重新校准触摸屏后修复 by andrew | SWO-192556 | Adjustment | HW | DMS_PC | Touch screen |  | Touch touch button is not touched | Touch touch button is not touched | Fix by andrew after recalibrating the touch screen
No | Visit | Complaint | RCHN | LAC定标不过 | LAC定标不过 | 更换测试卡后修复 | SWO-195590 | Assisted Replacement | SC | SC | Lac |  | LAC calibration is not | LAC calibration is not | Repair after replacing test card
No | Visit | Complaint | RCHN | 测试卡底部温度超范围 | 测试卡底部温度超范围 | 更换测试卡后修复 | SWO-195591 | Assisted Replacement | SC | SC | Thermistor | 1065 | Test card bottom temperature over range | Test card bottom temperature over range | Repair after replacing test card
No | Visit | Complaint | RCHN | 样本量不足，定标卡扫描失败 1218 | 样本量不足，定标卡扫描失败 1218，借902-924，924-621各一个。 |  | SWO-192607 | Revisit required | HW | Analyzer Control Module | Barcode reader, display | 1218 | Insufficient sample size, bid card scan failed 1218 | The sample size is insufficient, the bid card scan failed 1218, borrowing 902-924, 924-621 each.
No | Visit | Complaint | RCHN | 定标底物外露检查失败 1024，1025 | 定标底物外露检查失败 1024，1025，借903-364一个。 |  | SWO-192547 | Defect | HW | Process Module | Optical Unit | 1024 | Calibration substrate exposure check failed 1024,1025 | The calibration substrate exposure check failed 1024,1025, borrowed 903-364 one.
No | 66 | Y | RESP | ERROR 0643...ABL/DMS PC DISCONNECTED FROM AU VUELVE A SALTAR EL MISMO ERROR DE LA SEMANA PASADA |  | TRAS RECIBIR MATERIAL SE PROCEDE A CHECHEO DE MÓDULO PC, MEDICIÓN DE TENSIONES DE LA POWER SUPPLY, CABLES DE CONEXIÓN PC-AU, ANALYZER CONTROL PCB, FUSIBLES, Y CONEXIONES Y CABLES EN GENERAL DESCONECTANADO Y CONECTANDO TODOS DE NUEVO, NO SE OBSERVA NADA ANÓMALO. SE SUSTITUYEN LOS 4 FUSIBLES Y SE VERIFICAN Y REFUERZAN LOS CONTACTOS EN LA ANALYZER CONTROL PCB, SE VERIFICA CONTINUIDAD EN CONTACTOS, SE SUSTITUYE CABLE DE COMUNICACIÓN DE PC UNIT-ANALIZER CONTROL.. | D28C768E6405D63F | Revisit required | HW | Basic Unit | PCB, Analyzer control | Email | ERROR 0643...ABL/DMS PC DISCONNECTED FROM AU RETURNS THE SAME ERROR OF LAST WEEK |  | AFTER RECEIVING MATERIAL, PC MODULE REGULATION, POWER SUPPLY TENSION MEASURE, PC-AU CONNECTION CABLES, ANALYZER CONTROL PCB, FUSIBLE, AND GENERAL CONNECTIONS AND CABLES ARE AVAILABLE AND CONNECTED ALL NEW, NOTHING IS NOT OBSERVED. THE 4 FUSIBLES ARE SUBSTITUTED AND THE CONTACTS ARE CHECKED AND REFURBED IN THE ANALYZER CONTROL PCB, CONTINUITY IS VERIFIED IN CONTACTS, PC UNIT-ANALIZER CONTROL COMMUNICATION CABLE IS SUBSTITUTE.
No | Visit | Complaint | RHUN | elhajlottt a tű, ideiglenesen visszahajtogatták, de csere kell | - | felszívó tű cseréje, bemeneti egység tisztítása, LS és pumpa kalibrációk, QC mérés | SWO-192657 | Defect | HW | Inlet module | Probe bend |  | Bent over the needle, temporarily refolded, but replacement should | - | Suction needle replacement, input unit cleaning, LS and pump calibrations, QC measurement
No | Visit | Complaint | RDEU | 15.05.2017 12:20Uhr Sr.Anett 030450665122: Nach der Probeeingabe ertönt normaler Weise ein Ton zum Warnen, dass die Einlassklappe geschlossen werden soll. Dieser Ton fehlt | Datensicherung durchgeführt, erneut Systemeinstellungen überprüft - keine Fehler gefunden | Betriebssystem neu installiert, Analysatorsoftware installiert, Backup zurückgespielt, keine Sound hörbar, Lautsprecher ausgebaut - Kabel an Stecker angelötet, wieder zusammengesetzt, Neustart abgewartet, mehrere Kal, Sound eingestellt, QC | SWO-192700 | Defect | HW | DMS_PC | Loudspeaker |  | 15.05.2017 12:20h Sr.Anett 030450665122: After the test entry, a sound sounds normally to warn that the inlet flap should be closed. This tone is missing | Data backup performed, system settings checked again - no errors found | Operating system reinstalled, analyzer software installed, backup backplayed, no sound audible, speakers removed - cable soldered to plug, reassembled, restart awaited, multiple kal, sound set, QC
No | Visit | Complaint | RKK | pH 検体測定時、応答不良 | pH 検体測定時、応答不良 | pH電極交換 タンパク除去 エアライン・測定チャンバー清掃 LS調整・ポンプCAL確認 CAL・QC測定データ確認 | SWO-193961 | Assisted Replacement | Electrode | Electrode | pH | email | pH When measuring specimens, poor response | pH When measuring specimens, poor response | pH electrode replacement Protein removal Airline measurement chamber cleaning LS adjustment and pump CAL confirmation Cal/QC measurement data confirmation
No | Visit | Complaint | RKK | pH 応答不良 多発 | pH 応答不良 | pH $BGベースプレート交換 タンパク除去 エアライン・測定チャンバー清掃 LS調整・ポンプCAL確認 CAL・QC測定データ確認 | SWO-193952 | Assisted Replacement | Electrode | Electrode | pH | Email | pH Bad Response Multiple | pH Bad response | pH $BG base plate replacement Protein removal Airline measurement chamber cleaning LS adjustment and pump CAL confirmation Cal/QC measurement data confirmation
No | Visit | Complaint | RDEU | Kunde muss in kurzen Intervallen Lac neu membranieren | Abnehmende Lac Empf. | Lac nach Service Tipp gesäubert und mit Proteinentferner behandelt, Elektrode gespült und neu bezogen (LOT 3488) mehrfach 1 Punkt-Kalibrierung durchgeführt, Empfindlichkeit deutlich höher als zuvor aber Empf. springt z.Z. noch, Kunde beobachtet weiteres Verhalten | SWO-192722 | Assisted Replacement | Electrode | Membrane | Lac | Sensitivity | Customer must re-membrane Lac at short intervals | Decreasing Lac Empf. | Lac cleaned according to service tip and treated with protein remover, electrode rinsed and re-sourced (LOT 3488) performed several times 1 point calibration, sensitivity significantly higher than before but empf. is still jumping at the time, customer is watching further behavior
No | Visit | Complaint | RDEU | 24.04.2017; Fr. Dohmes; Tel.: 06946086125; QC Messfehler - Fehler vorhanden in Messung mit LQC-Lösung in Position | Am Analysator auch eine Wartungsmassnahme durchgeführt. | Ausbau der Optischen Einheit und des Process Wheels und Einbau einer neuen Optischen Einheit und eines neuen Procees Wheels. Alte Kalibrationen gelöscht und neu kalibriert. Kalibration Werte in Ordnung, Qualitätskontrolle wieder zielwertig. | SWO-192761 | Revisit required | HW | Process Module | Optical Unit | 211 | 24.04.2017; Ms. Dohmes; Tel.: 06946086125; QC measurement error - error present in measurement with LQC solution in position | A maintenance measure was also carried out on the analyzer. | Expansion of the optical unit and the process wheel and installation of a new optical unit and a new Procees Wheel. Old calibrations deleted and recalibrated. Calibration values in order, quality control again targeted.
Yes | 6 | Y | RAME | 9 Lac trending high |  | Verified Lac, level $slot 3 Quality Control Plot. Carbon dust in dual valve housing. | D28DE34CDA4D28D7 | Re classified RMED | ABL8XX entry |  |  |  | 9 Lac trending high |  | Verified Lac, level $slot 3 Quality Control Plot. Carbon dust in dual valve housing.
Yes | Visit | Complaint | RDEU | Meldung Lac gesperrt, QC Werte Lac außer Meßbereich | Meldung Lac gesperrt über Radiance (ABL ohne Radiancanbindung !), QC Werte Lac außer Meßbereich Fehler trat nach SW upgrade und JW auf, siehe auch SWO 192793 | heutige Überprüfung ergibt, Sperrung Lac nicht mehr vorhanden ( nach Neustart?) QC Werte Lac nach wie vor alle außer Meßbereich Kal, Reinitialisierung LP Anderes LP eingesetzt, anschl. wieder Lac gesperrt, QC Werte nach wie vor außerhalb Meßbereich | SWO-192911 | Re classified RMED | SWU | FSE not following Service Tip | Disabled electrolyte | Lac | Message Lac blocked, QC values Lac out of measurement range | Message Lac blocked via radiance (ABL without radiancane binding !), QC values Lac out of measurement range Error occurred after SW upgrade and JW, see also SWO 192793 | Today's review reveals blocking Lac no longer exists (after reboot?) QC Values Lac still all except measurement range Kal, reinitialization LP Other LP used, followed by again Lac locked, QC values still outside measuring range
No | 6 | Y | RAME | Z System keeps shutting off by itself. |  | Inspection of the system revealed the power supply is not providing continuous power to the system. | D28D276BDA47879E | Defect | HW | Basic Unit | Power Supply |  | Z System keeps shutting off by itself. |  | Inspection of the system revealed the power supply is not providing continuous power to the system.
No | Visit | Complaint | RDEU | 17.05.2017, Herr Viernich, Tel.: 021596795102: Patientenetiketten können nicht eingescannt werden | Patientenetiketten können nicht eingescannt werden - Interleaved 2 aus 5. | Barcode Einstellungen für ITS 2 aus 5 hinzugefügt. Änderungen im System gespeichert. Startphase durchgeführt und mit dem Anwender eine Testmessung durchgeführt. | SWO-193057 | Setup | SW | Roma.ini | Barcode |  | 17.05.2017, Mr. Viernich, Tel.: 021596795102: Patient labels cannot be scanned | Patient labels cannot be scanned - Interleaved 2 out of 5. | Added barcode settings for ITS 2 out of 5. changes stored in the system. phase and a test measurement carried out with the user.
No | 4 | Y | RPOL | Blank | Sprawdzenie kabla RS232, reinstalacja windows na analizatorze wraz z oprogramowaniem 3.3. Stwierdzono uszkodzenie komputera w analizatorze. Nowe czêœci zamienne zamówiono. Analizator sprawny w celach wykonywania oznaczeñ , brak transmisji danych LIS/HIS. |  | D28EFB70505F56BF | Revisit required | HW | DMS_PC | CPU |  | Blank | Check the RS232 cable, reinstall Windows on the analyzer with the 3.3 software. The computer has been damaged in the analyzer. New replacement parts are ordered. Analyser is efficient for the purposes of the mark-up, no data transmission LIS/HIS.
No | Visit | Complaint | RBNL | Error 1227: De correctie voor bicarbonaat bevat fouten uit PH, PCO2 | Sinds vanochtend verscheen error 1227: De correctie voor bicarbonaat bevat fouten uit PH, PCO2 . Daarbij ook de kalibratie/QC's worden afgekeurd. Het probleem begon nadat er vanochtend een nieuwe SC/RP werd geplaatst. Na telefonisch overleg werd besloten om SC opnieuw te vervangen, zonder resultaat. De inletgasket heb vervangen, zonder resultaat. Vanaf Inlet tot waste alle onderdelen heb gecontroleerd, geen afwijking heb geconstateerd. De Bayunit en de contactelementen heb gereinigd, hierna kalibratie/QC's zijn OK. | Bayunit, contactelemenren heb gereinigd. | SWO-193201 | Cleaned | HW | Basic Unit | Sensor Bay Unit |  | Error 1227: The correction for bicarbonate contains errors from PH, PCO2 | Since this morning, error 1227 appeared: The correction for bicarbonate contains errors from PH, PCO2. Calibration/Qcs are also rejected. The problem started after a new SC/RP was placed this morning. After telephone consultation it was decided to replace SC again, with no result. The inletgasket have replaced without result. From Inlet to waste all parts have checked, no deviation have been noted. The Bayunit and the contact elements have been cleaned, hereafter calibration/Qcs are OK. | Bayunit, Contactelemenren have cleaned.
No | Visit | Complaint | RDEU | 17.05.2017 Frau Janik 0211904361209 meldet, dass der Datenspeicher voll ist. | QC Datenspeicher voll, | Daten auf einen USB Stick gesichert und am Gerät gelöscht, | SWO-193026 | [5 empty] | 17.05.2017 Ms. Janik 0211904361209 reports that the data store is full. | QC data storage full, | data backed up to a USB stick and deleted from the device,
No | Visit | Complaint | RCHN | 全参数定标不过 | 全参数定标不过，借备件903-334一个 | 更换903-334后修复 by andrew | SWO-192680 | [5 empty] | Full parameter calibration is not the | Full parameter calibration, however, with spare parts 903-334 one | Repair by andrew after replacing 903-334
No | 16 | Y | RAME | 8 Paula called with the Glu failing low on level 1 external QC but the analzyer did not flag it out. |  | Paula ran level 1 external QC on May 17th at 8:35 am. The Glu result was 25. The range on the analyzer is (26-38). The analyzer stayed green and did not flat the result as being out low. They did not catch it at the time because it was green. She was looking at some results and found the issue. | D28DB816DA16F8A0 | [5 empty] | 8 Paula called with the Glu failing low on level 1 external QC but the analzyer did not flag it out. |  | Paula ran level 1 external QC on May 17th at 8:35 am. The Glu result was 25. The range on the analyzer is (26-38). The analyzer stayed green and did not flat the result as being out low. They did not catch it at the time because it was green. She was looking at some results and found the issue.
Yes | Visit | Complaint | RCHN | 样本量不足，定标卡扫描失败 1218 | 样本量不足，定标卡扫描失败 1218，借902-924，924-621各一个(924-621在DMS无产品信息，向丹麦订货后新建信息。故两单发货） |  | SWO-192924 | Re classified RMED | Logistical upSWO |  |  | 1218 | Insufficient sample size, bid card scan failed 1218 | Insufficient sample size, bid card scanning failure 1218, borrow 902-924,924-621 each (924-621 in DMS no product information, after ordering to Denmark after the new information.) So two orders are shipped)
Yes | Visit | Complaint | RCHN | 定标失败 | 定标失败，借944-394（12450）、944-220（12325）、944-214（12841）、944-216（12452）、944-258（12150）各一张 | 自行更换备用光单903-364，重新定标后修复 by andrew | SWO-192309 | Re classified RMED | Stock recovery |  |  |  | Bid failed | Failed to bid, borrow 944-394 (12450), 944-220 (12325), 944-214 (12841), 944-216 (12452), 944-258 (12150) each | Replace the spare light sheet 903-364 yourself and fix by Andrew after re-calibration
No | 39 | Y | RESP | Fallo en cooximetria por detección de aire |  | Kit de tubos. Revisada neumática, calibración y QCS. Todo ok. | D28DAE08DA0512F2 | [5 empty] | Air detection cooxymetry failure |  | Tube kit. Pneumatic review, calibration and QCS. It's all right.
No | Visit | Complaint | RIND | crea calibration expire. | Check the calibration. | Replace the crea reagent & calibrate the system. | SWO-193592 | Assisted Replacement | Consumables | Solution bottle | Crea |  | crea calibration expire. | Check the calibration. | Replace the crea reagent & calibrate the system.
No | Visit | Complaint | RCHN | 血氧模块氖灯电压过低，温度不对 | 血氧模块氖灯电压过低，温度不对，借备件902-833一个 | 更换902-833一个正常。 by andrew | SWO-192630 | Defect | HW | Oxi module | Lamp unit | Neon | The oxygen module lamp voltage is too low, the temperature is not correct | Blood oxygen module xenon lamp voltage is too low, temperature is not correct, borrow edipart 902-833 one | Replace 902-833 one normal. by andrew
No | Visit | Complaint | RDEU | Pumpen Asymetriefehler | Pumpen Asymetriefehler | Pumpenkalibration bringt auf dauer keine Verbesserung. Pumpenschläuche der Elektroden Module aus dem Kundenbestand ausgetauscht. | SWO-193161 | Defect | HW | pH/BG Module | Pump tubes |  | Pumps Asymetry Errors | Pumps Asymetry Errors | Pump calibration does not bring any improvement in the long run. Pump hoses of the electrode modules are replaced from the customer's inventory.
No | Visit | Complaint | RDEU | 18.05.17, 09:15, Frau Belke, 0211919481650, Autocheck Temepraturfehler 18.05.2017, 12.42 Uhr Frau Belke: tel.: 0211919481650 Autocheck TemperaturFehler | Temperaturfehler im QC Modul, Lüfter ist def. | Lüfter im Modul ersetzt, Gerät kalibriert und QC durchgeführt, | SWO-193066 | Defect | HW | Autocheck module | Fan | 763 | 18.05.17, 09:15, Mrs. Belke, 0211919481650, Autocheck Temepratur error 18.05.2017, 12.42 pm Mrs. Belke: tel.: 0211919481650 Autocheck TemperatureError | Temperature error in QC module, fan is def. | replaced in the module, device calibrated and QC performed,
No | 8 | Y | RAME | Z - pO2 failed 2 2-point Calibrations for zero current. |  | Customer had done several rinses and the problem resolved itself before FSE onsite. | D28DC367DA234A79 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | Z - pO2 failed 2 2-point Calibrations for zero current. |  | Customer had done several rinses and the problem resolved itself before FSE onsite.
No | Visit | Complaint | RDEU | 18.05.2017; Fr. Meier; Tel.: 02114092367; pO2 Nullstromfehler bei Kalibrationen, Flüssigkeit hinter der Einlassdichtung | pO2 Nullstromfehler bei der Kalibrierung. | Reparatur im Rahmen der Halbjahreswartung durchgeführt. | SWO-193327 | Adjustment | HW | pH/BG Module | Zero current |  | 18.05.2017; Ms. Meier; Tel.: 02114092367; pO2 Zero current fault during calibrations, liquid behind the inlet seal | pO2 Zero current error during calibration. | repair carried out as part of half-year maintenance.
No | 4 | Y | RAME | F Halogen lamp intensity too high |  | Found lamp intensity on the edge of out. Sometimes in, sometime out. | D28DAC92DA068FE8 | Adjustment | HW | Oxi module | Lamp unit |  | F Halogen lamp intensity too high |  | Found lamp intensity on the edge of out. Sometimes in, sometime out.
Yes | Visit | Complaint | RDEU | Meldung Lac gesperrt, QC Werte Lac außer Meßbereich Fehler trat nachJW und SW update auf | Meldung Lac gesperrt, QC Werte Lac außer Meßbereich Fehler trat nachJW und SW update auf | komplette Radiometer SW neu installiert, Haken bei Abfrage komplett rausgenommen, ABL neu konfiguriert Service Dump gezogen- Weiterleitung an TS Ursache Sperrung Lac ist , wenn K, Na oder Ca deaktiviert werden, wird Lac automatisch gesperrt, beim Kundengerät jetzt Ca zugeschaltet | SWO-193392 | Re classified RMED | SWU | FSE not following Service Tip | Disabled electrolyte | Lac | Message Lac blocked, QC values Lac out of measurement range Error occurred after JW and SW update | Message Lac blocked, QC values Lac out of measurement range Error occurred after JW and SW update | complete radiometer SW reinstalled, hook completely removed when query, ABL reconfigured Service Dump Pulled- Forwarding to TS Cause blocking Lac is when K, Na or Ca are deactivated, Lac is automatically locked, with customer device now Ca switched on
No | Visit | Complaint | RDEU | 19.05.2017 14:00Uhr Herr Müller 0362891950245: Seit der letzten Softwarpflege funktioniert der Scanner nicht. | Scanner funktioniert Barcode der Patienten werden. nicht erkannt | muß Änderung in der roma.ini vorgenommen werden ABL erkennt Tastatur nicht, andere Tastatur nicht vorhanden Kunde informiert Problem muß noch behoben werden | SWO-193563 | Revisit required | SW | Roma.ini | Roma.ini |  | 19.05.2017 14:00h Mr. Müller 0362891950245: The scanner has not worked since the last softwar maintenance. | Scanner works barcode of the patients. not recognized | must be changed in the roma.ini ABL does not recognize keyboard, other keyboard does not exist Customer informed Problem still needs to be fixed
No | Visit | Complaint | RKOR | Sample data error | Hb sample data error | Gasket hold clean. Gasket module clean. tHb calibration work. Sensor cassette clean. Calibration check QC check. Sample test. Activity log check. | SWO-195224 | Adjustment | HW | Oxi module | tHb Cal |  | Sample data error | Hb sample data error | Gasket hold clean. Gasket module clean. tHb calibration work. Sensor cassette clean. Calibration check QC check. Sample test. Activity log check.
No | 4 | Y | RPOL | Blank | Wymiana komputera wraz z kart¹ CF. Instalacja oprogramowania 3.3 Analizator sprawny technicznie, transmisja danych LIS/HIS prawid³owa. Ze wzglêdu na wewnêtrzne procedury szpitalne dotycz¹ce ochrony danch osobowych, karta pamiêci znajduj¹ca siê w uszkodzonym komputerze o numerze KC2GF-303 (* 1650153*23* ) zosta³a przekazana kierownikowi laboratorium. |  | D28EFC97506156C0 | Defect | HW | DMS_PC | CPU |  | Blank | Replacing the computer with the CF card. Installing the 3.3 Software Analyzer technically efficient, data transmission LIS/HIS correctly. Due to the internal hospital procedures for the protection of personal data, the memory card in the damaged computer number KC2GF-303 (* 1650153 * 23) has been transferred to the laboratory manager.
No | Visit | Complaint | RCHN | 温度-B，偏高 1302 | 温度-B，偏高 1302，借903-334一个。 | 更换903-334后修复 by andrew | SWO-192595 | [5 empty] | Temperature-B, high 1302 | Temperature-B, high 1302, borrowed 903-334 one. | Repair by andrew after replacing 903-334
No | Visit | Complaint | RCHN | 接口板腐蚀 | 接口板腐蚀，借件902-972一个 | 更换902-972后修复 by andrew | SWO-192940 | [5 empty] | Interface plate corrosion | Interface plate corrosion, borrowing 902-972 one | Repair by andrew after replacement 902-972
No | Visit | Complaint | RCHN | 所有参数定标失败 | 所有参数定标失败 | 更换测试卡后修复 | SWO-193517 | Assisted Replacement | SC | SC | Unknown |  | Failed to calibrate all parameters | Failed to calibrate all parameters | Repair after replacing test card
No | Visit | Complaint | RKK | QC測定時にサンプル移動エラーが出る。 | 検体痩躯亭でも10件に1件程度エラーが出ていた。 | Met2ﾓｼﾞｭｰﾙｴｱﾊﾞﾙﾌﾞがｴﾗｰ出していた。また、PHﾓｼﾞｭｰﾙﾘﾝｽ側ﾊﾞﾙﾌﾞの動きもいまいちのため交換。 | SWO-194751 | Defect | HW | Met2 Module | Valve |  | Sample movement error occurs during QC measurement. | There was an error about the case in the specimen singing-out. | The Met2 module air valve was out of error. In addition, the movement of the PH module rinse side valve is also bad, so we replaced it.
No | Visit | Complaint | RKK | QCでNaが範囲外 | QCでNaが範囲外 | 回路洗浄実施（次亜塩素酸溶液） 全メンブラン交換 Na電極交換 Cal2 QC OK | SWO-193265 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | Na out of range in QC | Na out of range in QC | Circuit cleaning (hypochlorite solution) All Membrane Replacement Na electrode replacement Cal2 QC OK
No | Visit | Complaint | RDEU | 19.05.2017 12:00 H. Schaikowski 017615322584: Kal-Fehler pCO2 | Elektrodenspitze gebrochen | pCO2 Elektrode (E788 JT04) geprüft und ersetzt, Funktionskontrolle ok | SWO-193533 | Assisted Replacement | Electrode | Electrode | pCO2 | Revisit? | May 19, 2017 12:00 PM H. Schaikowski 017615322584: Kal error pCO2 | Electrode tip broken | pCO2 electrode (E788 JT04) tested and replaced, function control ok
Yes | Visit | Complaint | RDEU | 19.05.2017 12:00 H. Schaikowski 017615322584: Kal-Fehler pCO2 | pCO2 Elektrode (E788 YB02) keine Status und Empfindlichkeitswerte | pCO2 Elektrode geprüft und ausgetauscht | SWO-193534 | Re classified RMED | Double entry |  |  |  | May 19, 2017 12:00 PM H. Schaikowski 017615322584: Kal error pCO2 | pCO2 electrode (E788 YB02) no status and sensitivity values | pCO2 electrode tested and replaced
No | 4 | Y | RFRA | Blank | erreur 3412 |  | D28F3470501F9D19 | Defect | Hw | Aspiration Module | Needle Wash Unit | 3412 | Blank | error 3412
No | Visit | Complaint | RCHN | 定标失败 | 定标失败，借TNI定标卡944-212（12602）一张 | 重新定标后修复 by andrew | SWO-193339 | Assisted Replacement | Consumables | Test Kit | Cartridge | TnI | Bid failed | Failed to bid, borrow TNI bid card 944-212 (12602) one | Fix by andrew after re-calibration
No | Visit | Complaint | RCHN | 零点校准失败 | 零点校准失败，借903-334一个 |  | SWO-193666 | [5 empty] | Zero calibration failed | Zero calibration failure, borrow 903-334 one
No | Visit | Complaint | RCHN | 报错1035 | 报错1035 | 更换库存机903-364后修复 by andrew | SWO-193678 | Defect | HW | Process Module | Optical Unit | 1035 | 1035 error | 1035 error | Repair by andrew after replacing stock machine 903-364
No | Visit | Complaint | RCHN | 定标底物外露检查失败 1024，1025 | 再次定标底物外露检查失败 1024，借903-236一个 | 更换903-364，903-236各一个正常。 by jesson | SWO-193602 | Defect | HW | Process Module | Optical Unit | 1024 | Calibration substrate exposure check failed 1024,1025 | Re-calibration substrate exposure check failed 1024, borrow 903-236 one | Replace 903-364, 903-236 each normal. by jesson
No | Visit | Complaint | RIND | tHb status blank in calibration and sample measurement | Checked and found Halogen lamp intensity too high | Performed Lamp unit adjustment, 2 point calibration performed successfully. | SWO-193698 | Adjustment | HW | Oxi module | Lamp unit |  | tHb status blank in calibration and sample measurement | Checked and found Halogen lamp intensity too high | Performed Lamp unit adjustment, 2 point calibration performed successfully.
Yes | 6 | Y | RFRA | Blank | FUITE DETECTEE |  | D292DF1AB222C2B4 | [5 empty] | Blank | DETECTE FUITE
No | Visit | Complaint | RDEU | 19.05.2017 14:00Uhr Herr Müller 0362891950245: Seit der letzten Softwarpflege funktioniert der Scanner nicht. | Barcode des Anwenders wird nicht erkannt | Änderung in der roma.ini vorgenommen Neustart und Test (i.O.) | SWO-193993 | Setup | SW | Roma.ini | Barcode reader, display |  | 19.05.2017 14:00h Mr. Müller 0362891950245: The scanner has not worked since the last softwar maintenance. | User's barcode is not recognized | Change made in the roma.ini Restart and Test (i.O.)
No | Visit | Complaint | RLTD | Customer querying bilirubin results | Checked tHb calibration ok Complete oximetry check ok All calibration and qc for bilirubin ok | None required Analyser working within specification | SWO-194081 | No fault found | Performance | Oxi module | ctBil |  | Customer querying bilirubin results | Checked tHb calibration ok Complete oximetry check ok All calibration and qc for bilirubin ok | None required Analyser working within specification
Yes | 13 | Y | RPAC | Blank | NA QC Level B failure |  | D29324A6B22E80E1 | [5 empty] | Blank | NA QC Level B failure
No | 4 | Y | RFRA | Blank | SOL NON HOMOGENES |  | D2931E66B217DAFE | Defect | HW | Auto Inlet module | Inlet Gasket |  | Blank | SOL NON HOMOGENES
No | Visit | Complaint | RCHN | 开机程序不运行 | 开机程序不运行 借备件933-306一个 |  | SWO-193864 | [5 empty] | The power-on program does not run | Power-on program does not run borrowing spare parts 933-306 one
Yes | Visit | Complaint | RCHN | LAC,PCO2,CA,LAC定标失败 | PCO2,CA,LAC定标失败 | 更换测试卡正常。 | SWO-194177 | [5 empty] | LAC, PCO2, CA, LAC calibration failed | PCO2, CA, LAC calibration failed | The replacement test card is OK.
No | Visit | Complaint | RCHN | 血氧单元不匹配 0581 | 血氧单元不匹配 0581 ，借902-980，903-197各一个。 |  | SWO-193880 | Defect | HW | Oxi module | Hemolyzer | 581 | Blood oxygen unit mismatch 0581 | Blood oxygen unit mismatch 0581 , borrow 902-980, 903-197 each.
No | Visit | Complaint | RIND | Not switching ON. | CPU Unit to replace. | CPU Unit replaced. Analyzer serviced and calibrated properly. Handed over in fine working condition. Analyzer is on observation till next day. | SWO-195452 | Defect | HW | DMS_PC | CPU |  | Not switching ON. | CPU Unit to replace. | CPU Unit replaced. Analyzer serviced and calibrated properly. Handed over in fine working condition. Analyzer is on observation till next day.
No | Visit | Complaint | RKK | Neon Lamp intesity too low | Neon Lamp intesity too low | モジュール引き上げ | SWO-193796 | Defect | HW | Oxi module | Oxi module, Replaced | Neon | Neon Lamp intesity too low | Neon Lamp intesity too low | Module pull up
No | Visit | Complaint | RIND | No printout | Gear assembly in printer head found dismantled | Reassembled the gears, 2 point calibration performed successfully. | SWO-194848 | Defect | HW | Printer Module | Printer head |  | No printout | Gear assembly in printer head found dismantled | Reassembled the gears, 2 point calibration performed successfully.
Yes | 3 | Y | RPAC | Blank | Compensate 1 box of Lactate membranes |  | D293363DB24480E7 | [5 empty] | Blank | Compensate 1 box of Lactate membranes
No | Visit | Complaint | RDAN | Dorthe Skovsende melder, at ABL'en per. har Neon number too low og Neon lamp intensity too low. | Afprøvet og undersøgt. | OXI PCB udskiftet, service tHb og alm. tHb kalibreret, slutafprøvet. | SWO-194087 | Defect | HW | Oxi module | PCB, Oxi module |  | Dorthe Skovsende melder, at ABL'en per. har Neon number too low og Neon lamp intensity too low. | Tested and investigated. | OXI PCB udskiftet, service tHb og alm. tHb kalibreret, slutafprøvet.
Yes | 2 | Y | RAME | Z: C USTOMER REQUESTS R ADIANCE CHANGED TO AQURE; T IME CHANGED TO 24H F ORMAT AND I.T C ONNECTION ESTABLISHED |  | C USTOMER REQUESTS R ADIANCE CHANGED TO AQURE; T IME CHANGED TO 24H F ORMAT AND I.T C ONNECTION ESTABLISHED | D29423BB28396F12 | Re classified RMED | Planned by Customer request |  |  |  | Z: C USTOMER REQUESTS R ADIANCE CHANGED TO AQURE; T IME CHANGED TO 24H F ORMAT AND I.T C ONNECTION ESTABLISHED |  | C USTOMER REQUESTS R ADIANCE CHANGED TO AQURE; T IME CHANGED TO 24H F ORMAT AND I.T C ONNECTION ESTABLISHED
Yes | 2 | Y | RPAC | Blank |  |  | D2953ED39E36F284 | Re classified RMED | Installation |  |  |  | Blank
No | Visit | Complaint | RLTD | Issue raised with engineer on arrival at site by Stephanie Sandy. | Solution pack QC lot not consistent with Installed solution pack - Solution pack installation error. As solution pack almost empty, new solution pack fitted. Error message still present. Backed up database and upgraded software to v3.3. Analyser restarted after software upgrade, error now resolved. Verified operation of analyser - all ok. | Upgraded and rebooted software. | SWO-194579 | Defect | SW | RMED SW | Data base | 1405 | Issue raised with engineer on arrival at site by Stephanie Sandy. | Solution pack QC lot not consistent with Installed solution pack - Solution pack installation error. As solution pack almost empty, new solution pack fitted. Error message still present. Backed up database and upgraded software to v3.3. Analyser restarted after software upgrade, error now resolved. Verified operation of analyser - all ok. | Upgraded and rebooted software.
No | Visit | Complaint | RLTD | Stephaine 07863330762 | Analyser failing clot detection test, value too high for Sensor Cassette. Fault traced to inlet probe which was bent at the back end that connects into the inlet connector gasket. Fitted new inlet probe and clot detection test now passes. Also noted persistent Inhomogeneous QC2 error in solution pack test in service, so replaced solution pack. Lactate sensor failed with response error, fitted new Sensor Cassette. Verified operation on restart - all ok. | Replaced bent inlet probe. | SWO-194581 | Replaced | HW | Auto Inlet module | Probe bend |  | Stephaine 07863330762 | Analyser failing clot detection test, value too high for Sensor Cassette. Fault traced to inlet probe which was bent at the back end that connects into the inlet connector gasket. Fitted new inlet probe and clot detection test now passes. Also noted persistent Inhomogeneous QC2 error in solution pack test in service, so replaced solution pack. Lactate sensor failed with response error, fitted new Sensor Cassette. Verified operation on restart - all ok. | Replaced bent inlet probe.
No | Visit | Complaint | RCHN | BNP温度传感器损坏 1235 | BNP温度传感器损坏 1235 | 更换905-908一个正常。by jesson | SWO-194292 | Defect | HW | Process Module | Sensor, BNP |  | BNP temperature sensor damage 1235 | BNP temperature sensor damage 1235 | Replace 905-908 one normal. by jesson
No | Visit | Complaint | RCHN | 报错1031，1029，试管架的轮子掉了 | 报错1031，1029，试管架的轮子掉了，借903-364，903-584各一个。 |  | SWO-194367 | Defect | HW | Process Module | Optical Unit | 1031 | The wheel of the test tube rack was lost at 1031,1029. | The wheels of the test tube rack were lost at 1031,1029, borrowing one each at 903-364, 903-584.
No | Visit | Complaint | RCHN | 测试卡一侧导轨脱落 | 测试卡一侧导轨脱落 | 测试卡复位后修复 | SWO-194807 | Defect | Consumables | Test Kit | Cartridge | Broken wings | One side rail of the test card is off | One side rail of the test card is off | Repair after test card reset
Yes | Visit | Complaint | RIND | Not switching ON. | General Visit. | Analyser checked and found working fine. Analyzer calibrated and found all parameters well with in range. Solutions replaced. | SWO-195453 | Re classified RMED | Follow Up |  |  |  | Not switching ON. | General Visit. | Analyser checked and found working fine. Analyzer calibrated and found all parameters well with in range. Solutions replaced.
No | 4 | Y | RPAC | Blank | Blank and Black screen |  | D29523CF9E3CB519 | Revisit required | HW | DMS_PC | DMS_PC |  | Blank | Blank and Black screen
No | Visit | Complaint | RDEU | 23.5.2017 10:40 Uhr Herr Wilkicz Tel.: 08988922435, Barcode für Lösungen wird immer wieder gefordert | Gerät bei eintreffen Meßbereit | Lösungen abgezogen und eingespannt Kein Fehler feststellbar, bei QC Überprüfung O2 instabil pO2 neu membraniert Gerät kalibriert und überprüft | SWO-194424 | Cleaned | HW | Basic Unit | Barcode reader, Basic Unit |  | 23.5.2017 10:40 Mr. Wilkicz Tel.: 08988922435, barcode for solutions is demanded again and again | Device ready to be measured when it arrives | Solutions removed and clamped No error detectable, in QC review O2 unstable pO2 re-membraneed Device calibrated and checked
No | 6 | Y | RAME | P Fluid Flow issues |  | Checked logs. Checked all valves. Checked pump rotors. Checked waste lines. | D29407D9286284B5 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | P Fluid Flow issues |  | Checked logs. Checked all valves. Checked pump rotors. Checked waste lines.
No | 3 | Y | RAME | F Customer reports oxy hardware errors |  | 0251, halogen lamp intensity too high and 0579, not possible to measure errors are present. Halogen lamp intensity is observed to be higher than the 3100-3900 range at >4000. | D2940285285F5925 | Adjustment | HW | Oxi module | Lamp unit | Halogen | F Customer reports oxy hardware errors |  | 0251, halogen lamp intensity too high and 0579, not possible to measure errors are present. Halogen lamp intensity is observed to be higher than the 3100-3900 range at >4000.
No | 4 | Y | RAME | Z Touch screen not working. |  | The touch screen is out of calibration. Also, there is a bubble under the screen and unable to perform calibration. | D293DA48281F0407 | [5 empty] | Z Touch screen not working. |  | The touch screen is out of calibration. Also, there is a bubble under the screen and unable to perform calibration.
No | Visit | Complaint | RDEU | Benutzerintervention erforderlich, nach Tausch des LP und der SC | Anruf von der Medizintechnik Herr Wolfgang Ginschel erhalten, er bittet darum, dass ich mir das Gerät ansehe, da die Medizintechnik selber wenig Zeit dazu hat!!!! Fehlersuche --> Enlassnadel, Einlassdichtung, Lösungspack, Sensorkassette überprüft --> Einlassnadel und Einlassdichtung i.O.; Lösungspack (LP) und Sensorkassette (SC) auch in Ordnung | Lösungspack vom Analysator auswerfen lassen, selben Lösungspack nochmals installiert, Überprüfung des Analysators im Service (Adjust Manifold, Pressuretest, LS-Justage, Pumpenkal., SP Check, Clot detect $remove); Analysator gespült, System und Funktion i.O. --> aktuelle SC (435-112) bei vorheriger SC(425-189 --> installiert am 19.05.17) ebenfalls QC-Fehler bei pCO2, welcher sich erst am nächsten Tag (20.05.17) bemerkbar machte --> Kal. und QC-Fehler bei pCO2 --> keine SC mehr vor Ort --> FSE (TobEh) organisiert sich eine SC und wird diese am Freitag einbauen | SWO-194827 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | user intervention required, after replacing the LP and SC | Received a call from the medical technology Mr. Wolfgang Ginschel, he asks that I look at the device, because the medical technology itself has little time to do so!!!! Troubleshooting --> Discharge needle, inlet seal, solution pack, sensor cassette checked --> Inlet needle and inlet seal in the case; Solution pack (LP) and sensor cartridge (SC) also fine | Let the analyzer eject the solution pack, reinstall the same solution pack, check the analyzer in service (Adjust Manifold, Pressuretest, LS-Adjustment, Pump Calca, SP Check, Clot detect $remove); Analyzer flushed, system and function i.O. --> current SC (435-112) with previous SC(425-189 --> installed on 19.05.17) also QC error with pCO2, which only became noticeable the next day (20.05.17) --> Kal. and QC error at pCO2 --> no more SC on site --> FSE (TobEh) organizes a SC and will install it on Friday
Yes | 2 | Y | RPAC | Blank | Sensor cassette failure Lac and Na |  | D2953F189E05F285 | [5 empty] | Blank | Sensor cassette failure Lac and Na
No | Visit | Complaint | RCHN | 固定样本处断裂 | 固定样本处断裂，借924-738一个 |  | SWO-194607 | [5 empty] | Break at fixed sample | Break at fixed sample, borrow 924-738 one
No | Visit | Complaint | RCHN | 提示试剂包门栓没关 | 提示试剂包门栓没关，借902-975一个 |  | SWO-194769 | [5 empty] | Prompt reagent pack door bolt not closed | Prompt reagent package door bolt not closed, borrow 902-975 one
No | Visit | Complaint | RDEU | Meßunterschiede zum ABL 8XX Intensiv | Meßunterschiede | Gerät überprüft, Vergleichsmessungen durchgeführt, tHb Kal und Zero Abgleich durchgeführt, Konstanzmessungen o.k. | SWO-195233 | [5 empty] | Measurement differences to ABL 8XX Intensive | Meßunterschiede | device checked, comparative measurements carried out, tHb Kal and Zero matching performed, Constance measurements ok
No | Visit | Complaint | RCHN | PO2定标失败 | PO2定标失败，借902-972和903-334各一个 |  | SWO-194749 | [5 empty] | PO2 calibration failure | PO2 calibration failure, borrow 902-972 and 903-334 each one
No | Visit | Complaint | RCHN | 传感器加热器重复定标失败 | 传感器加热器重复定标失败 | 更换测试卡后修复 | SWO-194606 | Assisted Replacement | SC | SC | Thermistor |  | Sensor heater repeat calibration failed | Sensor heater repeat calibration failed | Repair after replacing test card
No | Visit | Complaint | RCHN | 测试卡参比电极错误 | 测试卡参比电极错误 | 更换测试卡后修复 | SWO-194630 | Assisted Replacement | SC | SC | Reference | 1361 | Test card reference electrode error | Test card reference electrode error | Repair after replacing test card
No | Visit | Complaint | RCHN | PO2，THB质控失败，定标失败 1531，0090 | PO2，THB质控失败，定标失败 1531，0090 | 更换两个试剂包正常。 | SWO-194987 | Assisted Replacement | SP | SP | FT | 1531 | PO2, THB QC failed, bid failure 1531,0090 | PO2, THB QC failed, bid failure 1531,0090 | Replacing both reagent kits is normal.
No | Visit | Complaint | RCHN | 仪器死机，无原因重启 | 仪器死机，无原因重启，借910-342一个 | 更换910-342一个正常。by jesson | SWO-194291 | Defect | HW | DMS_PC | CPU |  | Instrument crashes, no reason to restart | Instrument crash, no reason to restart, borrow 910-342 one | Replace 910-342 one normal. by jesson
No | Visit | Complaint | RCHN | 报5784 | 报5784 | 更换测试卡后修复 | SWO-195084 | Defect | Consumables | Test Kit | Cartridge | 5784 | Reported 5784 | Reported 5784 | Repair after replacing test card
Yes | 8 | Y | RCZE | Blank |  |  | D29594BD9E287ED9 | [5 empty] | Blank
Yes | 8 | Y | RCZE | Blank |  |  | D2955D639E297825 | [5 empty] | Blank
No | 4 | Y | RPAC | Blank | Blank and black screen |  | D2952B959E63B51D | Defect | HW | DMS_PC | CPU |  | Blank | Blank and black screen
No | Visit | Complaint | RDEU | 24.05.2017; Hr. Schmidt; Tel.: 022147897564; Flüssigkeit währned Gaskal erkannt | Messkammer ist voller Flüssigkeit | pCO2 neu Membraniert, Messkammer gereinigt, LS Justage und Pumpenkal durchgeführt. Mehrfach kalibriert. | SWO-195287 | Assisted Replacement | Electrode | Membrane | pCO2 |  | 24.05.2017; Mr. Schmidt; Tel.: 022147897564; Liquid detected during Gaskal | Measuring chamber is full of liquid | pCO2 newly membraneed, measuring chamber cleaned, LS adjustment and pump scale performed. Multiple calibrated.
No | Visit | Complaint | RDEU | 24.05.2017 10:50Uhr Herr Schmitt 022147897564: Aus dem Einlass tropft Flüssigkeit | Fehlersuche durchgeführt. Wahrscheinlich liegt der Fehler im Einlassmodul | Friction Detection Unit bestellt | SWO-195302 | Revisit required | HW | Inlet module | Friction detector unit |  | May 24, 2017 10:50am Mr Schmitt 022147897564: Liquid drips out of the inlet | Troubleshooting. Probably the error is in the inlet module | Friction Detection Unit ordered
No | Visit | Complaint | RCHN | 血红素不出值 | 血红素不出值，借902-833一个。 | 更换902-833后修复 by jesson | SWO-194256 | Defect | HW | Oxi module | Oxi module, Replaced | Hemolyzer | Hemoglobin is not worth it. | Hemoglobin is not worth, borrow 902-833 one. | Repair by jesson after replacement 902-833
No | 2 | Y | RAME | P Customer reports analyzer leak. |  | QC soution is observed leaking in and around the analyzer. The pH sealing ring is also observed to be missing from the pH electrode. | D29553CD9E219A1A | Defect | Electrode | Electrode | Sealing ring | Sealing ring, pH | P Customer reports analyzer leak. |  | QC soution is observed leaking in and around the analyzer. The pH sealing ring is also observed to be missing from the pH electrode.
No | Visit | Complaint | RBNL | Inletdeur blijft open | Tijdens het aanbieden van een bloedbuis, de inletdeur blijft open staan met de melding Verwijder de buis terwijl er geen buis in zit. Aan/uit zetten van de AQT lost het probleem ook niet op. Bij nader onderzoek blijkt dat een klein stukje papier op de barcodescanner is geplakt. Zodoende deze verwijderd vervolgens de scanner heb schoongemaakt vervolgens in het service menu heb gecontroleerd, OK. | De barcodescanner van de Inletmodule heb gereinigd. | SWO-194961 | Cleaned | HW | Aspiration Module | Sensor, Vacutainer |  | Inletdoor remains open | While offering a blood tube, the inletdoor remains open with the notification remove the tube while there is no tube in it. Setting the AQT on/off does not solve the problem either. On closer examination it appears that a small piece of paper is stuck on the barcode scanner. So this removed the scanner then got cleaned then in the Service menu checked, OK. | The barcode scanner of the Inletmodule has been cleaned.
No | Visit | Complaint | RDEU | 23.05.2017 13:45 Fr. Kraus 0726460237: Setsafefehler 5784 | Analysator zeigt Fehlermeldung 5784 Setsafe 22 . Überprüfung des Analysators im Service. Alle Positionen gehen auf Setsafe, außer Storage Wheel. Eine Testkassette blockiert das Drehen des Storage Wheels. Anwender hatten vor Fehler des Analysators erst Kalibrationskassette eingelegt. | Eine Kunststoff-Nippel an der Testkassette isst abgebrochen, somit konnte der Analysator die Testkassette nicht mehr ins Storage Wheel zurückziehen. Testkassette herausgenommen und Bereich des Storage Wheels überprüft, ob Nippel vorhanden und nicht Rotation blockiert. Kalibration gestartet. Kalibration on Ordnung. | SWO-194962 | Defect | Consumables | Test Kit | Cartridge | Broken tooth | 23.05.2017 13:45 Fr. Kraus 0726460237: Setsafe error 5784 | Analyzer displays error message 5784 Setsafe 22 . Checking the analyzer in the service. All positions go to Setsafe, except Storage Wheel. A test cartridge blocks the storage wheel from rotating. Users had only inserted calibration cassettes before the analyzer's errors. | A plastic nipple on the test cartridge eats aborted, so the analyzer could no longer retract the test cartridge into the storage wheel. Test cartridge removed and area of the storage wheel checks whether nipples exist and does not block rotation. calibration started. Calibration on order.
Yes | Visit | Complaint | RDEU | Sensor Kassette defekt. Keine SC mehr vorrätig. | SC defekt. | Sensor Kassette in Bottrop geholt und zum Kunden geliefert. Kunde hatte noch eine SC dieser er in der Zwischenzeit eingesetzt hatte. Gerät kalibriert und QC durchgeführt. Problematik ist entstanden, weil der Kunde den ABL90 in Kombination mit nicht heparinisierten Spritzen verwendet. Mit dem Kunden die Problematik besprochen. Kunde wird in Zukunft Pico50 Pn verwenden. | SWO-195378 | Re classified RMED | Logistical upSWO |  |  |  | Sensor cassette defective. No more SC in stock. | SC defective. | Sensor cassette in Bottrop and delivered to the customer. Customer still had a SC which he had used in the meantime. calibrated and QC performed. Problems arose because the customer uses the ABL90 in combination with non-heparinized syringes. Discussed the issue with the customer. Customer will use Pico50 Pn in the future.
Yes | Visit | Complaint | RDEU | Sensor Kassette defekt. Keine SC mehr vorrätig. |  | Anhang zum Notdienst. Lieferung von 2 Sensor Kasetten und 1 LP. | SWO-195379 | Re classified RMED | See SWO-195378 |  |  |  | Sensor cassette defective. No more SC in stock. |  | Appendix to the emergency service. Delivery of 2 sensor sets and 1 LP.
No | Visit | Complaint | RCHN | PCT定标失败，光学强度超范围1031 | PCT定标失败，光学强度超范围1031 | 更换903-364一个正常. by jesson | SWO-195012 | Defect | HW | Process Module | Optical Unit | 1031 | PCT calibration failure, optical strength over range 1031 | PCT calibration failure, optical strength over range 1031 | Replace 903-364 one normal. by jesson
No | Visit | Complaint | RCHN | 报错1031，1029，试管架的轮子掉了 | 报错1031，1029，试管架的轮子掉了，借903-364一个。 | 更换903-364,903-584各一个正常。by jesson | SWO-195029 | Defect | HW | Process Module | Optical Unit | 1031 | The wheel of the test tube rack was lost at 1031,1029. | Error 1031,1029, the test tube frame wheel fell off, borrowing 903-364 one. | Replace 903-364, 903-584 each normal. by jesson
No | Visit | Complaint | RCHN | 排线断了 | 排线断了，借备件636-652一个 |  | SWO-195066 | Defect | HW | Incubator Module | Cable |  | The line is broken. | The line is broken, with spare parts 636-652 one
No | Visit | Complaint | RKK | Lacキャリブレーションエラー(感度133.0pA/mM)。Lacメンブラン交換。QC測定。 | Lacキャリブレーションエラー(感度133.0pA/mM) | Lacメンブラン交換。QC測定。 | SWO-195169 | Assisted Replacement | Electrode | Membrane | Lac | Sensitivity | Lac calibration error (sensitivity 133.0pA/mM). Lac Membrane replacement. QC measurement. | Lac calibration error (sensitivity 133.0pA/mM) | Lac Membrane replacement. QC measurement.
No | Visit | Complaint | RIND | In calibration Po2 sensitivity is above the range. | Check analyzer status and calibration result and found PO2 status is yellow and it's sensitivity is higher than the range. | Clean all the chambers ,clean PO2 electrode and done 2-point calibration which is OK but near the high limit.Run QC which is also OK. | SWO-195588 | Assisted Replacement | Electrode | Membrane | pCO2 |  | In calibration Po2 sensitivity is above the range. | Check analyzer status and calibration result and found PO2 status is yellow and it's sensitivity is higher than the range. | Clean all the chambers ,clean PO2 electrode and done 2-point calibration which is OK but near the high limit. Run QC which is also OK.
No | 8 | Y | RAME | P 0653 sample in pH/BG module detected prematurley. |  | Error happening on patient samples. Minor calibration drift recently. | D296AABA14309E2F | Defect | Electrode | Electrode | Sealing ring | Sealing ring, pH | P 0653 sample in pH/BG module detected prematurley. |  | Error happening on patient samples. Minor calibration drift recently.
No | 3 | Y | RCZE | Blank | neon position failed - Fneon: 2.645 ( limit do 2.000 ) |  | D296E8A114637A09 | Defect | HW | Oxi module | Oxi module, Replaced | Neon | Blank | neon position failed - Fneon: 2.645 ( limit do 2.000 )
No | 3 | Y | RCZE | Blank | neon position failed, Fneon: 2.178 ( limit 2.000 ), dále chyba v hemolyzeru - Impedance Shape : Phaseregulation error |  | D296EB2014587A0E | Defect | HW | Oxi module | Spectrometer | LEH | Blank | neon position failed, Fneon: 2.178 ( limit 2.000 ), dále chyba v hemolyzeru - Impedance Shape : Phaseregulation error
No | Visit | Complaint | RDEU | Benutzerintervention erforderlich, nach Tausch des LP und der SC | Fehlersuche --> pCO2 zu hoch --> QC-Fehler in allen Leveln Dateiarchive (Kal., QC und Betriebsmittel) überprüft --> SC muss erneuert werden | SC (R0539) eingebaut und (R435) ausgebaut, System und Funktion i.O. | SWO-195310 | Assisted Replacement | SC | SC | pCO2 |  | user intervention required, after replacing the LP and SC | Troubleshooting --> pCO2 too high --> QC errors in all levels file archives (cal., QC and equipment) checked --> SC needs to be renewed | SC (R0539) built-in and (R435) removed, system and function in the
No | Visit | Complaint | RLTD | Leak detected and vacuum error. Sue - 07954 576051 | Checked service programs. | Performed Manual flush Replaced gasket and ran fluid transport checks. Analyser is now passing. Initialised, calibrated and has passed all QC's. | SWO-195466 | Defect | HW | Auto Inlet module | Inlet Gasket |  | Leak detected and vacuum error. Sue - 07954 576051 | Checked service programs. | Performed Manual flush Replaced gasket and ran fluid transport checks. Analyser is now passing. Initialised, calibrated and has passed all QC's.
No | Visit | Complaint | RDEU | 26.05.2017 14:45 Uhr; Frau Winterle; Tel.: 0091285043201: Einlass bleibt in Spritzenposition und kracht | Einlass schließt nicht und rattert/kracht beim Versuch zu schließen | Einlass entfernt und Stellrad bewegt - lässt sich ohne Probleme bewegen und man spürt keine Blockade oder unrunde Drehung, ABL beginnt ohne Einlassmodul die Startphase von 10 Min., nach 10 Min. ABL mittels Reset-Knopf ausgeschaltet, im ausgeschalteten Zustand Einlassmodul wieder angesetzt und ABL neu gestartet, Startphase von 50 Min. hat begonnen, nach Startphase mehrere Testmessungen in Spritzen- und Kapillarmodus durchgeführt - Einlass öffnet und schließt ganz normal und ohne krachenden Geräusche, System- und Funktionskontrolle in Ordnung | SWO-195534 | Adjustment | SW | Auto Inlet module | Off/On | Inlet Open | May 26, 2017 2:45 PM; Mrs Winterle; Tel.: 0091285043201: Inlet remains in syringe position and crashes | Entrance does not close and rattles/crashes when trying to close | Inlet removed and adjusting wheel moved - can be moved without problems and you can not feel any blockage or uncircular rotation, ABL starts the start phase of 10 min. without inlet module, after 10 min. ABL is switched off by means of a reset button, when the inlet module is switched off again started and ABL restarted, start phase of 50 min. has begun, after start phase several test measurements in syringe and capillary mode - inlet opens and closes normally and without crashing noise, system and function control in order
No | Visit | Complaint | RCHN | 感应板腐蚀 | 感应板腐蚀，借备件902-972一个 | 更换902-972一个正常。by jesson | SWO-195049 | [5 empty] | Induction plate corrosion | Induction plate corrosion, spare parts 902-972 one | Replace 902-972 one normal. by jesson
No | Visit | Complaint | RCHN | 触摸屏触摸位置偏移 | 触摸屏触摸位置偏移 ，借910-324一个。 | 更换910-324一个正常。by jesson | SWO-194997 | Defect | HW | DMS_PC | Touch screen |  | Touch touch position offset | Touch touch position offset , borrow edith of 910-324. | Replace 910-324 one normal. by jesson
No | Visit | Complaint | RCHN | 光学强度超范围 1031 | 光学强度超范围 1031，借903-364，944-212（12296）一张。 | 更换903-364一个，重新定标正常。by jesson | SWO-195041 | Defect | HW | Process Module | Optical Unit | 1031 | Optical strength over range 1031 | Optical strength over range 1031, borrowed 903-364, 944-212 (12296) one. | Replace the 903-364 one and re-label normally. by jesson
No | Visit | Complaint | RCHN | 报1031错误 | 报1031错误 | 清洁处理盘，并系统保养 | SWO-195501 | Defect | HW | Process Module | Optical Unit | 1031 | Reported 1031 error | Reported 1031 error | Clean the processing disc and maintain the system
No | Visit | Complaint | RDEU | Laktat Ansprachefehler und QC außer Bereich | Systemstatusdatei, Elektrodensitz kontrolliert | Lak. Membran gewechselt (alte Membran noch auf der Elektrodenspitze) , Probelauf ok | SWO-195438 | Assisted Replacement | Electrode | Membrane | Lac |  | Lactate speech error and QC out of range | System status file, electrode seat checked | Lak. Membrane changed (old membrane still on the electrode tip) , test run ok
No | Visit | Complaint | RDEU | 26.05.2017 Herr Schraven 015174669714 meldet, das der ABL schon kurz nach der Aufstellung unglaubwürdige Werte zeigt. Eine Überprüfung vor Ort wird gewünscht. | pH Messkammer verunreinigt. | pH Messkammer gereinigt, Gerät kalibriert und QC durchgeführt. | SWO-195388 | Cleaned | Fluidic pathway | pH/BG Module | Measuring chamber |  | 26.05.2017 Mr. Schraven 015174669714 reports that the ABL shows implausible values shortly after the installation. An on-site check is requested. | pH measuring chamber contaminated. | pH measuring chamber cleaned, device calibrated and QC performed.
No | Visit | Complaint | RKK | クロット検出 | クロット検出 | 回路洗浄、ガスケット・溶液パック排液ドレイン掃除 | SWO-195625 | Cleaned | HW | Auto Inlet module | Gasket - Clot |  | Clot detection | Clot detection | Circuit cleaning, gasket and solution pack drain drain cleaning
Yes | Visit | Complaint | RCHN | 小风扇不转了 | 小风扇不转了，借636-520一个 | 更换636-520后修复 by jesson | SWO-195199 | [5 empty] | The little fan doesn't turn. | The little fan doesn't turn, borrow 636-520 one | Repair by jesson after replacement 636-520
No | Visit | Complaint | RCHN | 报错1031 | 报错1031。借备件903-364，BNP定标卡944-258（12413）各一个 | 更换光单，重新定标 | SWO-195500 | Defect | HW | Process Module | Optical Unit | 1031 | 1031 | 1031. Borrowed spare parts 903-364, BNP calibration card 944-258 (12413) each | Replace the light sheet and re-label
No | Visit | Complaint | RCHN | PCT定标失败，光学强度超范围1031 | PCT定标失败，借PCT定标卡944-394（12383）一张 |  | SWO-195541 | Defect | HW | Process Module | Optical Unit | 1031 | PCT calibration failure, optical strength over range 1031 | Failure of PCT calibration, one by pct calibration card 944-394 (12383)
No | Visit | Complaint | RCHN | 软件版本不匹配 0001 | 软件版本不匹配 0001 ，借903-198一个。 | 更换903-198后修复(by tim) | SWO-195074 | Defect | HW | Analyzer Control Module | PCB, Main Board | 0001 | Software version mismatch 0001 | Software version mismatch 0001 , borrowing 903-198. | Repair after replacement 903-198 (by tim)
No | Visit | Complaint | RCHN | 光学强度超范围 1031 | 光学强度超范围 1031，借903-364一个。 | 更换903-364后修复 by jesson | SWO-195262 | Defect | HW | Process Module | Optical Unit | 1031 | Optical strength over range 1031 | Optical strength over range 1031, borrowed 903-364 one. | Repair by jesson after replacing 903-364
No | Visit | Complaint | RCHN | 液路错误 3412 | 液路错误 3412，借903-329，924-504各一个。 | 更换备件903-329及924-504后修复(by tim) | SWO-195079 | Defect | HW | Aspiration Module | Needle Wash Unit | 3412 | Liquid Road Error 3412 | Liquid Road Error 3412, borrow 903-329, 924-504 each. | Replacement of spare parts 903-329 and 924-504 after repair (by tim)
Yes | Visit | Complaint | RCHN | 打印机无法打印 | 打印机无法打印，借910-275一个。 | 更换910-275后修复 by jesson | SWO-195263 | [5 empty] | The printer can't print | The printer cannot print, borrow 910-275 one. | Repair by jesson after replacement 910-275
No | Visit | Complaint | RIND | In calibration Po2 sensitivity is above the range. | After 2-3 calibration again Po2 sensitivity increases and shows above than 60. | Again clean all chambers, again scan both Gases, run zero current calibration and done 2-point calibration but problem remains same.Replace PO2 electrode but not effected much more.Again check and found PCo2 membrane is leaking. Replace po2 & pco2 membrane and done calibration but sensitivity is more than 40.Replace electrode with new membrane and done calibration and problem has been resolved. | SWO-195592 | Assisted Replacement | Electrode | Membrane | pO2 | Sensitivity | In calibration Po2 sensitivity is above the range. | After 2-3 calibration again Po2 sensitivity increases and shows above than 60. | Again clean all chambers, again scan both Gases, run zero current calibration and done 2-point calibration but problem remains same. Replace PO2 electrode but not effected much more. Again check and found PCo2 membrane is leaking. Replace po2 & pco2 membrane and done calibration but sensitivity is more than 40.Replace electrode with new membrane and done calibration and problem has been resolved.
No | 6 | Y | RAME | Z Customer states Inlet clip broken. Probe not secure. |  | Customer states Inlet clip broken. Probe not secure. | D298F0BF003001FE | Defect | HW | Inlet module | Inlet clip |  | Z Customer states Inlet clip broken. Probe not secure. |  | Customer states Inlet clip broken. Probe not secure.
No | 9 | Y | RAME | P 0653 Sample in pH/BG upper liquid sensor detected prematurely. | Changed the complaint visit tab from 'blank' to Y. | Intermittently occurring on patient samples and QC. | D298F070003C1B01 | Defect | HW | pH/BG Module | Pump tubes | 653 | P 0653 Sample in pH/BG upper liquid sensor detected prematurely. | Changed the complaint visit tab from 'blank' to Y. | Intermittently occurring on patient samples and QC.
No | Visit | Complaint | RCHN | 触摸屏无反应 | 触摸屏无反应，借备件903-239一个 | 备件没有使用，重装软件后修复 by jesson | SWO-195503 | [5 empty] | Touch screen does not respond | Touch screen non-reactive, spare parts 903-239 one | Spare parts not used, reload software to repair by jesson
No | 3 | Y | RPNZ | Blank | pH QC failures on level 1 and 4. |  | D29AD86676072D41 | [5 empty] | Blank | pH QC failures on level 1 and 4.
No | Visit | Complaint | RCHN | 液体传感器错误 | 液体传感器错误，借903-165和903-199各一个 | 更换903-199后修复 by jesson | SWO-195277 | Defect | HW | Basic Unit | Cassette retainer |  | Liquid sensor error | Liquid sensor error, one each for 903-165 and 903-199 | Repair by jesson after replacement 903-199
No | Visit | Complaint | RSCH | Blank | Impossible de redémarrer le pack de solutions (divers messages) | Remplacé le rotor péristaltique e.o (blocage) e.o Calibration e.o contrôle de qualité e.o | SWO-195882 | Defect | HW | Basic Unit | Pump |  | Blank | Can't restart the solution pack (various messages) | Replaced the peristaltic rotor e.o (blocking) e.o Calibration e.o quality control e.o
No | Visit | Complaint | RDEU | 29.05.2017 09:20Uhr Herr Raxingern08517205132: Das Gerät meldet Flüssigkeitstransportfehler und Intervention erforderlich. | Herr Draxinger hat angerufen, da er wusste, dass ich wegen des Ausfalls des ABL auf der Intensivstation auf dem Weg bin. Inhomogene Lösungen. Kundin hat LP vor meiner Ankunft eingelegt. | Kundin hatte LP bereits getauscht. ABL war in 55 minütiger Startphase. Kalibrationen und QC's vorgenommen. | SWO-195720 | Assisted Replacement | SP | SP | FT |  | 29.05.2017 09:20Uhr Mr. Raxingern08517205132: The device reports fluid transport errors and intervention required. | Mr Draxinger called because he knew I was on my way because of the failure of the ABL in the intensive care unit. Inhomogeneous solutions. Customer inserted LP before my arrival. | Customer had already exchanged LP. ABL was in 55 minutes of launch. calibrations and QC's.
No | Visit | Complaint | RCHN | LAC定标不过 | LAC定标不过 | 更换测试卡后修复 | SWO-201252 | [5 empty] | LAC calibration is not | LAC calibration is not | Repair after replacing test card
No | Visit | Complaint | RKK | SC交換するとVCエラーやO2のエラーで初期化できない。今回も社内の装置では問題なく初期化できたSCを現地で取り付けるとO2が＃＃＃になってしまう。納品後1年以内でエラー出すぎなので一度しっかり見てほしいと依頼あり。 | SC交換時毎回何かしらのエラー | SC数回交換。 | SWO-195920 | [5 empty] | SC replacement cannot be initialized with VC error or O2 error. O2 is ###になってしまう when the SC which was able to be initialized without a problem is installed in the locale this time withthe equipment in the company. Because there is an error too much within one year after delivery, there is a request that I want you to see it firmly once. | Any errors every time you replace SC | SC several times replaced.
No | Visit | Complaint | RKK | QCでNaが範囲外 | QCで電解質の値が測定毎にふらつく | 比較電極交換 Cal2 QC OK | SWO-195660 | Assisted Replacement | Electrode | Electrode | Na |  | Na out of range in QC | The value of the electrolyte is wobbly every measurement in QC | Comparative electrode replacement Cal2 QC OK
No | Visit | Complaint | RDAN | Lars Boje melder, at den ikke har kørt QC siden den 24/5. | Afprøvet og undersøgt. Hammer kører ikke. Med ny motor virker det ej heller! |  | SWO-196060 | Revisit required | HW | Autocheck module | Autocheck module |  | Lars Boje reports that it has not run QC since 24/5. | Tested and investigated. Hammer is not running. With new engine it does not work either!
Yes | 1 | Y | RFRA | Blank | FUITE SUITE DEMARRAGE |  | D29BB5C9EC395029 | [5 empty] | Blank | LEAKING AFTER STARTING
No | 4 | Y | RFRA | Blank | Plus d'écran, suite intervention SAV M.Martina |  | D29B9943EC05D2CE | Replaced | Analyzer | Analyzer replaced | Analyzer replaced |  | Blank | More screen, following intervention SAV M.Martina
No | Visit | Complaint | RDEU | 29.05.2017 17:30 Fr. Sammer 09405182390: Backlogfehler nach mehrfachem SC-Wechsel | Vor Ankunft Backlog Fehler. pO2 QC Fehler Level C zu niedrig. Kunde hatte bereits 3x SC und 1x LP erneuert. | Abfallschlauch und Ventilschlauch überprüft - kein Flüssigkeitsaustritt. Abfallschlauch am Waste Connector nachgesetzt. ABL manuell gespült. Pinch Valve im Service überprüft - schaltet einwandfrei. Oxi Complete Test ok. pO2 Complete und Pressure Test ok. tHb Kalibration ausgeführt (Fcuv = 101,2%). Einlassdichtung erneuert. LP aus Kundenbestand erneuert. | SWO-196017 | Cleaned | Fluidic pathway | Fluidic pathway | Wetability |  | 29.05.2017 17:30 Fr. Sammer 09405182390: Backlog error after multiple SC change | Before arrival backlog error. pO2 QC error level C too low. Customer had already renewed 3x SC and 1x LP. | Waste hose and valve hose checked - no liquid leakage. waste hose on the Waste Connector. ABL manually rinsed. Pinch Valve checked in service - switches flawlessly. Oxi Complete Test ok. pO2 Complete and Pressure Test ok. tHb calibration (Fcuv = 101.2%). Inlet seal renewed. LP from customer base.
Yes | 3 | Y | RCZE | Blank | po poslední návštìvì objednali nové tìsnìní, v krabièce na plus bylo ale na flex |  | D29D3A3D621B79BA | Re classified RMED | Logistical upSWO |  |  |  | Blank | After the last visit, they ordered a new seal, but in the box the plus was on the flex
No | Visit | Complaint | RLTD | Battery not holding charge | Battery faulty after software upgrade to 3.3. | Replaced battery and test ok. | SWO-196427 | Defect | HW | Battery module | Battery |  | Battery not holding charge | Battery faulty after software upgrade to 3.3. | Replaced battery and test ok.
No | Visit | Complaint | RIND | Electrolyte values not satisfactory. | doubt on parameters. | Preseason study done on QC and two different patient samples and found all acceptable. Analyser Serviced and calibrated properly. Handed over in a fine working condition. Staff Satisfied with results. | SWO-196587 | [5 empty] | Electrolyte values not satisfactory. | doubt on parameters. | Preseason study done on QC and two different patient samples and found all acceptable. Analyser Serviced and calibrated properly. Handed over in a fine working condition. Staff Satisfied with results.
Yes | Visit | Complaint | RIND | Lac Failure. | Lac Failure | Found calibration values Lac sensitivity was low, tri to calibrate but problem was remining same, need to replace with new Sensor Cassette and confirm. | SWO-196832 | [5 empty] | Lac Failure. | Lac Failure | Found calibration values Lac sensitivity was low, tri to calibrate but problem was remining same, need to replace with new Sensor Cassette and confirm.
No | Visit | Complaint | RDEU | Parameter gesperrt weil angeblich QC Level 3 fehlerhaft gemessen wurde, obwohl der keine QC 3 hat | System überprüft => Parameter wurden wegen fehlerhafter QC Level 3 gesperrt .Kunde hat diesen aber nicht. Es wurde ein QC gemessen wo Luft erkannt wurde . | Level 1 gelöscht und als Level 3 angelegt . Dann QC gemessen , System hat sich entsperrt, QC wieder richtig konfiguriert, Kalibration und QC durchgeführt | SWO-196026 | [5 empty] | Parameters blocked because qc level 3 was allegedly measured incorrectly even though the one does not have a QC 3 | System checked => Parameters have been locked due to incorrect QC Level 3 . But customer does not have this. A QC was measured where air was detected. | Level 1 deleted and created as Level 3 . Then QC measured, system has been unlocked, QC reconfigured correctly, calibration and QC performed
No | Visit | Complaint | RIND | Autocheck ampoules not scanned | Laser scan engine checked and found defective | Replaced laser scan unit from a working ABL800, now error is rectified | SWO-197040 | Defect | HW | Autocheck module | Barcode reader, Autocheck Module |  | Autocheck ampoules not scanned | Laser scan engine checked and found defective | Replaced laser scan unit from a working ABL800, now error is rectified
No | Visit | Complaint | RLTD | Charles - 02476 968 510 | Remove and inspect autocheck module, found that the impeller from the Autocheck fan had fallen off. | Thoroughly dust out autocheck module, clean mirror on laser scanning unit, fit new Autocheck fan. On power on and calibration had leak currents detected, carried out extensive tests to pinpoint cause of leak currents, found nothing conclusive but still had leak currents. Stripped and cleaned pH chamber and seals. Perform 2 x 1 point calibration, and run all 4 levels of QC to clear errors. | SWO-196808 | Defect | HW | Autocheck module | FAN | OOB | Charles - 02476 968 510 | Remove and inspect autocheck module, found that the impeller from the Autocheck fan had fallen off. | Thoroughly dust out autocheck module, clean mirror on laser scanning unit, fit new Autocheck fan. On power on and calibration had leak currents detected, carried out extensive tests to pinpoint cause of leak currents, found nothing conclusive but still had leak currents. Stripped and cleaned pH chamber and seals. Perform 2 x 1 point calibration, and run all 4 levels of QC to clear errors.
No | Visit | Complaint | RDEU | Sporadisch jede 2. bis 3. Probe Oxy Hardware Error wegen neon Lamp | Erneut Oxi-Probleme; Modul wurde bereits am 28.04. mit SWO-188483 ersetzt | Fehlersuche: Nach Reinigung der Fließwege Oxi-Complete-Test absolviert, Neon Intensity too low, Neo Position failure => Modul ersetzt und Garantie beantragt. Abgleich und OCT, OK. Kalibrationen, tHb-Kal, QC Messungen OK | SWO-196232 | Defect | HW | Oxi module | Oxi module, Replaced | Neon lamp | Sporadic every 2nd to 3rd sample Oxy Hardware Error due to neon lamp | Again Oxi problems; Module has already been replaced with SWO-188483 on 28.04. | Troubleshooting: After cleaning the flow paths Completed Oxi-Complete-Test, Neon Intensity too low, Neo Position failure => Module replaced and warranty requested. Matching and OCT, OK. Calibrations, tHb-Kal, QC measurements OK
No | Visit | Complaint | RHUN | 2 x leállva találták a gépet, utána jöttek az OXI hibák. | Áramkimaradás után az OXI paramétereket hibásan méri. | Oximetriás tesztek hibásak. Hemolizátor csere után a tesztek megfelelőek. Kalibrálás, tHb kalibrálás, QC mérés. 4-5 méréssel tesztelve a készülék. | SWO-196595 | Defect | HW | Oxi module | Hemolyzer |  | 2 x They found the machine stopped, after which the OXY defects came. | After a power failure, the OXY parameters are incorrectly measured. | Oximetry tests are incorrect. After a haemolytic replacement, the tests are appropriate. Calibration, tHb calibration, QC measurement. 4-5 is tested by measuring the device.
No | 6 | Y | RAME | P Customer complains the unit is stuck on hold with a lead detected error present and haven't been able to resolve it. |  | Upon arrival the customers stated complaint was verified. Checked the fluidics module for leaks and any worn tubing with all tubing looking ok. Checked all tubing on the inlet module and discovered the tubing on the inlet heater was loose and leaking slightly, Also, the inlet gasket looked to have been past it's replacment date. | D29CC88062227CEF | Defect | HW | Inlet module | Tubings |  | P Customer complains the unit is stuck on hold with a lead detected error present and haven't been able to resolve it. |  | Upon arrival the customers stated complaint was verified. Checked the fluidics module for leaks and any worn tubing with all tubing looking ok. Checked all tubing on the inlet module and discovered the tubing on the inlet heater was loose and leaking slightly, Also, the inlet gasket looked to have been past it's replacment date.
No | 9 | Y | RAME | P 0653 sample in pH/BG module detected prematurley. |  | checked data logs for errors, checkec valves and liquid sensors, checked electrodes and electrode chambers, checked electrode membranes, | D29D54E9623F83FC | Defect | HW | pH/BG Module | Pump tubes | Email | P 0653 sample in pH/BG module detected prematurley. |  | checked data logs for errors, checkec valves and liquid sensors, checked electrodes and electrode chambers, checked electrode membranes,
No | 16 | Y | RAME | pH out on blood samples |  | Error code 476 & 293 on blood samples. All calibrations and quality controls are in. | D29D0FE662380F39 | Assisted Replacement | Electrode | Electrode | pH | 476 | pH out on blood samples |  | Error code 476 & 293 on blood samples. All calibrations and quality controls are in.
No | 15 | Y | RAME | Z QC sample error. |  | Viewed QC logs - QC intermittently getting sample errors. 14 pole cable in auto check crimped. | D29CC3E2621E0421 | Defect | HW | Autocheck module | FPC Cable |  | Z QC sample error. |  | Viewed QC logs - QC intermittently getting sample errors. 14 pole cable in auto check crimped.
No | 4 | Y | RAME | 3 Customer reported 2 point calibrations failing with zero current errors on pO2 intermittently. |  | Checked data logs for errors. Zero current errors intermittently. All other checks ok. | D29CEC1D622142DA | Adjustment | HW | pH/BG Module | Zero current |  | 3 Customer reported 2 point calibrations failing with zero current errors on pO2 intermittently. |  | Checked data logs for errors. Zero current errors intermittently. All other checks ok.
No | 2 | Y | RFRA | Blank | Fuite sous joint d'intro |  | D29CDFFD62185F9D | Defect | HW | Inlet module | Nipple kit |  | Blank | Leak under intro seal
No | Visit | Complaint | RDEU | Anwender müssen in letzter Zeit sehr häufig membranieren. | Kal. Fehler bei Metaboliten (Driften), Kunde hat Lac Elektrode ersetzt | Nach SW Update keine Auffälligkeiten bei Lac mehr feststellbar, Glu Ansprachefehler, Glu neu membraniert, Messkammer, Pumpenschläuche und Ventile überprüft, Glu einkalibriert (Lac weiterhin unauffällig), Kal. Und QC ohne Beanstandung | SWO-196203 | Assisted Replacement | Electrode | Membrane | Glu |  | Users have had to membrane very frequently lately. | Kal. Error in Metabolites (Drifting), Customer Has Replaced Lac Electrode | After SW Update no more abnormalities can be detected at Lac, Glu speech error, Glu re-membraneed, measuring chamber, pump hoses and valves checked, Glu calibrated (Lac still inconspicuous), Kal. And QC without complaint
No | Visit | Complaint | RSAF | Customer said that the Abl90 Flex does not want to aspirate samples. | Found analyzer green. Ran water as sample and aspirated perfectly, asked sisters for blood sample to observe the way they introduce the sample. | Done data log download as requested. Set the analyzers time to correct time as it was an hour ahead. Ran 5x samples and all ok. | SWO-196693 | No fault found | No fault found | No fault found | No fault found |  | Customer said that the Abl90 Flex does not want to aspirate samples. | Found analyzer green. Ran water as sample and aspirated perfectly, asked sisters for blood sample to observe the way they introduce the sample. | Done data log download as requested. Set the analyzers time to correct time as it was an hour ahead. Ran 5x samples and all ok.
No | Visit | Complaint | RDEU | seit 29.05. keine Verbindung zu Aqure | Verbindungsdaten überprüft | die Netzwerkkonfiguration auf Windowsebene war nicht mehr vorhanden, erneut eingegeben und auf CF Karte gesichert, Neustart abgewartet, Verbindung i.O. | SWO-196795 | Setup | His/lis | His/lis | His/lis |  | since 29.05. no connection to Aqure | Connection data checked | the network configuration at the Windows level was no longer present, re-entered and backed up to CF card, restart awaited, connection i.O.
No | Visit | Complaint | RDEU | Halterung für Probenmischer gebrochen | Halterung für Probenmischer gebrochen | Sampler Tray getauscht Gerät überprüft | SWO-196621 | Defect | HW | Sampler Mixer | Mixer Cover |  | Bracket for sample mixer broken | Bracket for sample mixer broken | Sampler Tray exchanged Device checked
No | Visit | Complaint | RDEU | 30.05.2017; Sr. Silke; Tel.: 03696167049; mögliche Verstopfung der Fließwege erkannt | mögliche Verstopfung der Fließwege erkannt - häufige Meldung bei Probenmessung | Überprüfung Fließwege Austausch Schlauchsatz und Pumpen tHb Kal LP erneuert ( Fälligkeit lt. Plan) | SWO-196536 | [5 empty] | 30.05.2017; Sr. Silke; Tel.: 03696167049; possible blockage of the flow paths detected | possible blockage of the flow paths detected - frequent notification during sample measurement | Review flow paths Replacement hose set and pumps tHb Kal LP renewed (due date according to plan)
Yes | 11 | Y | RAME | I Temp errors 1305 and 1302. |  | Zero check in service not stable. One pin on SC connector is black. | D29DF369D8237A54 | [5 empty] | I Temp errors 1305 and 1302. |  | Zero check in service not stable. One pin on SC connector is black.
No | Visit | Complaint | RCHN | 定标卡无法扫描 | 定标卡无法扫描，借备件902-924一个 | 更换902-924一个正常。(by tim) | SWO-195945 | Defect | HW | Analyzer Control Module | Barcode reader, display | BR002 | Bid card cannot be scanned | The bid card cannot be scanned, one spare part 902-924 | Replace the 902-924 one normal. (by tim)
No | Visit | Complaint | RCHN | 定标失败1024 | 定标失败1024，借903-364,902-924各一个 | 更换903-364和902-924后修复(By Ran) | SWO-196500 | Defect | HW | Process Module | Optical Unit | 1024 | Bid Failure 1024 | Bid failure 1024, borrow 903-364, 902-924 each | Replacement 903-364 and 902-924 post-repair (By Ran)
Yes | 2 | Y | RPAC | Blank | After cleaning program the Na, Cl and Ca are drifting |  | D29DE963D807FBFE | [5 empty] | Blank | After cleaning program the Na, Cl and Ca are drifting
No | Visit | Complaint | RDEU | 26.05.2017 Herr Schraven 015174669714 meldet, das der ABL schon kurz nach der Aufstellung unglaubwürdige Werte zeigt. Eine Überprüfung vor Ort wird gewünscht. | pH in der Kalibrierung und in der QC instabil, außerhalb der Grenzen. Um 03:58 wurde eine Kapilläre BGA gemessen welche durch ein Proben Problem abgebrochen wurde. | Einlass gereinigt und Fließwege gespült, Messkammer gereinigt, pH Dichtring erneuert, zwei Kalibrierungen und QC's durchgeführt. | SWO-196926 | Cleaned | Fluidic pathway | pH/BG Module | Measuring chamber |  | 26.05.2017 Mr. Schraven 015174669714 reports that the ABL shows implausible values shortly after the installation. An on-site check is requested. | pH in calibration and in the QC unstable, outside the limits. At 03:58 a capillary BGA was measured which was aborted by a sample problem. | Inlet cleaned and flow paths rinsed, measuring chamber cleaned, pH sealing ring renewed, two calibrations and QC's performed.
No | 4 | Y | RFRA | Blank | Résultats discordant D-dimer sur résultats patient |  | D29E19B4D858889A | Defect | HW | Aspiration Module | Needle | Techline | Blank | Discordant D-dimer results on patient outcomes
No | Visit | Complaint | RDEU | 30.05.2017 12:30Uhr Herr Heilig 071717011236: Display ist dunkel, der Barcodeleser leuchtet. | Funktionstest durchgeführt | Ersatzmonitor aus Demogerät testweise eingebaut. Kein Erfolg. Steckverbindungen geprüft, keine Auffälligkeiten. CPU Unit ausgebaut und von anderen Komponenten getrennt um Startverhalten der CPU zu testen. Keine Reaktion, kein Booten. CPU Unit voraussichtlich defekt. Monitor zurückgebaut. CPU Unit bestellt. Gerät mitgenommen um in GS17 zu reparieren wenn CPU Unit geliefert wird. Zubehör und ext. Drucker vom Kunden mitgenommen. | SWO-197335 | Revisit required | HW | Analyzer Control Module | CPU |  | May 30, 2017 12:30pm Mr. Holy 071717011236: Display is dark, the barcode reader lights up. | Functional test performed | Replacement monitor from demo device installed on a test basis. No success. Connectors checked, no abnormalities. CPU Unit and disconnected from other components to test THE START behavior of the CPU. No reaction, no booting. CPU unit is expected to be defective. Monitor. CPU Unit. device to repair in GS17 when CPU Unit is supplied. Accessories and ext. printer taken by the customer.
No | Visit | Complaint | RDEU | 31.05.2017 Sr. Waldtraud 0817175680 meldet, dass ein Akku - Fehler angezeigt wird. | Akku defekt - Fehler 1559 | defekten Akku (910-311P0129N008) ausgebaut neuen Akku (910-311P0137N064) eingebaut Gerät kalibriert und überprüft | SWO-196785 | Defect | HW | Battery module | Battery | 1559 | 31.05.2017 Sr. Waldtraud 0817175680 reports that a battery error is displayed. | Battery defective - Error 1559 | defective battery (910-311P0129N008) removed new battery (910-311P0137N064) built-in Device calibrated and checked
No | Visit | Complaint | RDEU | 31.05.2017 Anwender beobachtet, dass die pH Empfindlichkeit nicht stabil ist. Sie schwankt zwischen 90 und 113%. 01.06.2017 Frau Laut 0987485308 meldet, das der pH-Fehler bestehen bleibt. die SC wurde ersetzt. Ein Vorortservice wird gewünscht. | Fehlersuche --> Dateiarchive (Kal., QC, Aktivitäten) überprüft --> Kalibrierfehler pH (pH 1070: Sensor Ansprachefehler) --> SC (R0545) defekt | SC (946-008 R0539) installiert und (946-008 R0545) deinstalliert, System und Funktion i.O. | SWO-196994 | Assisted Replacement | SC | SC | pH | Wetability | 31.05.2017 Users observed that pH sensitivity is not stable. It varies between 90 and 113%. 01.06.2017 Ms. Laut 0987485308 reports that the pH error remains. the SC has been replaced. A suburban service is desired. | Troubleshooting --> File archives (Cal., QC, Activities) checked --> Calibration error pH (pH 1070: Sensor contact error) --> SC (R0545) defective | SC (946-008 R0539) installed and uninstalled (946-008 R0545), system and function I.O.
No | Visit | Complaint | RDEU | Häufig Probenabbruch bei MBU | Häufig Probenabbruch bei MBU Häufig Erneuerung Lösungspack notwendig | 2 mal SP reklamiert - Erläuterung etc. Analysator gespült und überprüft - Erneuerung Nadel (gekrümmt) und Einlassdichtung - Funktionsüberprüfung etc. | SWO-196876 | Replaced | HW | Auto Inlet module | Probe bend |  | Frequent sample cancellation at MBU | Frequent sample cancellation at MBU Frequently renewal solution pack necessary | 2 times SP complained - explanation etc. Analyzer rinsed and checked - renewal needle (curved) and inlet seal - function check etc.
No | Visit | Complaint | RBNL | Inlet blijft open | Nog steeds afgebroken capillaire bloedmonsters door de ABL90. De inletgasket zit soms op de capillaire positie vast. | De Inletmodule compleet heb vervangen vervolgens heb gecontroleerd, OK. | SWO-196913 | User error | HW | Auto Inlet module | Broken Capillary |  | Inlet remains open | Still degraded capillary blood samples by the ABL90. The inletgasket is sometimes attached to the capillary position. | The Inletmodule has completely replaced then checked, OK.
No | 4 | Y | RCZE | Blank |  |  | D29F27914E43728D | Assisted Replacement | SC | SC | Lac | Techline | Blank
No | Visit | Complaint | RCHN | 接口板槽锈了 | 接口板槽锈了，借902-972一个 | 更换902-972一个正常。(by tim ) | SWO-196800 | [5 empty] | The interface board groove is rusty. | The interface board groove is rusted, borrowed 902-972 one | Replace 902-972 one normal. (by tim)
No | Visit | Complaint | RCHN | LAC定标失败 | LAC定标失败 | 更换测试卡正常。 | SWO-196840 | [5 empty] | LAC calibration failure | LAC calibration failure | The replacement test card is OK.
No | Visit | Complaint | RCHN | PO2,CL质控失败 | PO2,CL质控失败 | 更换测试卡正常。 | SWO-196859 | [5 empty] | PO2, CL QC FAILED | PO2, CL QC FAILED | The replacement test card is OK.
No | Visit | Complaint | RCHN | 温度不稳定 | 温度不稳定，借备件902-972一个 | 更换902-972后修复(by tim) | SWO-196285 | [5 empty] | Temperature instability | Temperature instability, spare parts 902-972 one | Repair after replacement 902-972 (by tim)
No | 3 | Y | RAME | 2 pCO2 not passing system cycle. |  | Examined sensor cassette and solution pack. Customer has replaced SC and SP. | D2A0DA0DC453E1F1 | [5 empty] | 2 pCO2 not passing system cycle. |  | Examined sensor cassette and solution pack. Customer has replaced SC and SP.
No | Visit | Complaint | RCHN | ph传感器阻抗错误 | ph传感器阻抗错误 | 更换测试卡后修复 | SWO-198088 | Assisted Replacement | SC | SC | pH | 1079 | ph sensor impedance error | ph sensor impedance error | Repair after replacing test card
No | 2 | Y | RESP | HA ECHADO HUMO NO ENCIENDE |  | SUSTITUCIÓN DE FUENTE DE ALIMENTACIÓN | D29FC91D4E1B2584 | Defect | HW | Power Supply | Power Supply | LEH Mail to Alex | HE'S SMOKED, HE DOESN'T TURN ON. |  | REPLACING THE POWER SUPPLY
No | Visit | Complaint | RCHN | 测试卡扫描失败 | 测试卡扫描失败，借910-304一个。 | 更换910-304后修复(by tim) | SWO-196408 | Defect | HW | Storage Module | Barcode reader, Storage wheel |  | Test card scan failed | The test card scan failed, borrowing 910-304 one. | Repair after replacement 910-304 (by tim)
No | Visit | Complaint | RCHN | CKMB定标失败 1025 | CKMB定标失败 1025 ，借944-216（12923）一张。 | 重新定标后修复(by tim) | SWO-196526 | Defect | HW | Process Module | Optical Unit | 1024 | CKMB Calibration Failure 1025 | CKMB Calibration Failure 1025 , borrow 944-216 (12923). | Repair after re-calibration (by tim)
No | 67 | Y | RPAC | Blank | Ongoing issues with COtHB and MettHb, regarding correlations. Software Upgrade analyser to Version V6.17. |  | D29F7EFD4E374E3A | Revisit required | HW | Oxi module | Unknown | COtHb & MetHb | Blank | Ongoing issues with COtHB and MettHb, regarding correlations. Software Upgrade analyser to Version V6.17.
No | Visit | Complaint | RKK | 血液やQCでHbFたまに多く出る |  |  | SWO-196868 | Unknown | Unknown | Unknown | Unknown | LEH Mail | HbF sometimes comes out in blood and QC
No | Visit | Complaint | RDEU | Monitor lässt sich nicht bedienen, nach Neustart funktioniert es wieder | Monitor lässt sich nicht bedienen | Datensicherung, Monitor überprüft (HWParam 4.10.007), Firmware installiert | SWO-196953 | Defect | HW | DMS_PC | Touch screen |  | Monitor cannot be operated, after reboot it will work again | Monitor cannot be operated | Data backup, monitor checked (HWParam 4.10.007), firmware installed
Yes | 5 | Y | RPAC | Blank | Rreference parameters failing 3am cal |  | D29F284A4E07DAB4 | Re classified RMED | Logistical upSWO |  |  |  | Blank | Rreference parameters failing 3am cal
No | 7 | Y | RPNZ | Blank | Calibration failures on reference related parameters. Ca QC running high. |  | D29F51F54E33A178 | Cleaned | Fluidic pathway | pH/BG Module | Measuring chamber |  | Blank | Calibration failures on reference related parameters. Ca QC running high.
No | Visit | Complaint | RIND | Oximetry results not coming. | Oxymetry Hardware problem and temprature problem Msg Present. | Drane and remove the Oxy Module. After Half hour again refix and calibrate, Cleaning, Protean removal, Oxy complete test and tHb Calibration done and pass. Analyser seriviced and calibrated properly. Handed over in a fine working condition. | SWO-196881 | Cleaned | HW | Oxi module | Oxi module, Cleaned |  | Oximetry results not coming. | Oxymetry Hardware problem and temprature problem Msg Present. | Drane and remove the Oxy Module. After Half hour again refix and calibrate, Cleaning, Protean removal, Oxy complete test and tHb Calibration done and pass. Analyser seriviced and calibrated properly. Handed over in a fine working condition.
No | Visit | Complaint | RDEU | 1.06.2017 08:40Uhr Herr Viehoff 025263005237: Kal-Fehler bei pCO2 | PH Status und Driftfehler bei Kal, PCo2 Empf bei <10%, Fehler bei Proben Ansaugung | PH Elektrode untere Dichtung fehlt komplett, Blut über Überlauf PH Modul ausgelaufen, Blutreste entfernt Dichtung PH Elektrode ersetzt, PH Fehler seit gestern, PCO2 Fehler seit membranieren heute morgen 7:00, pCO2 Membran ersetzt Kontakte der Elektrode gereinigt, Empf bleibt < 10%, PCo2 Elektrode ersetzt, Kal und QC durchgeführt | SWO-197033 | Assisted Replacement | Electrode | Electrode | pCO2 |  | 1.06.2017 08:40Am Mr Viehoff 025263005237: Kal error at pCO2 | PH Status and drift error at Kal, PCo2 empf at 10%, error in samples suction | PH electrode lower seal is missing completely, blood spilled over overflow PH module, blood residue removed seal PH electrode replaced, PH fault since yesterday, PCO2 fault since membraneing this morning 7:00, pCO2 membrane replaced contacts of the electrode cleaned, mpf remains 10%, PCo2 electrode replaced, Kal and QC performed
Yes | 13 | Y | RFRA | Blank | Lenteur Software |  | D29F82994E635B15 | Re classified RMED | SWU |  |  |  | Blank | Slow Software
No | Visit | Complaint | RDEU | Lösungen können nicht gepumpt werden | - Einlaßdichtung fehlerhaft - mehrere LP def. | neue Einlaßdichtung eingesetzt; LP installiert; 1x LP hinterlegt; Kal. + QC i.O. | SWO-197367 | Defect | HW | Inlet module | Inlet Gasket |  | Solutions cannot be pumped | - Inlet seal faulty - multiple LP def. | new inlet seal inserted; LP installed; 1x LP deposited; Kal. + QC in the first place
No | Visit | Complaint | RLTD | Please call Girlre on 0191445 3195 | Sensor bay not closing properly, possible obstruction Removed sensor bay module and found closed sensor being obstructed by sensor flag | Realigned sensor flag so that it did not obstruct closing of bay | SWO-197380 | Adjustment | HW | Basic Unit | Sensor door |  | Please call Girlre on 0191445 3195 | Sensor bay not closing properly, possible obstruction Removed sensor bay module and found closed sensor being obstructed by sensor flag | Realigned sensor flag so that it did not obstruct closing of bay
No | Visit | Complaint | RSAF | Replacemet screen, inlet gasket the Solution pack | Inlet gasket warn out and never replaced | Replaced inlet gasket and start up tests ok. | SWO-197645 | Defect | HW | Auto Inlet module | Inlet Gasket |  | Replacemet screen, inlet gasket the Solution pack | Inlet gasket warn out and never replaced | Replaced inlet gasket and start up tests ok.
No | Visit | Complaint | RDEU | 01.06.2017 Herr Baldini 08517205301 meldet tHb - Fehler fHb, sO2, tHb | Oximetrie QC Fehler trotz durchgeführter tHb Kalibration | ABL90 manuell gespült Hemolyzer sowie Ventilschlauch erneuert, tHb kalibriert danach Glucose bei QC Level C zu tief, SCi + LP aus Kundenbestand erneuert nach SCi Tausch Glu im Level C um 2 mg $dL zu hoch SCi reinstalliert Gerät kalibriert und überprüft | SWO-197194 | Defect | HW | Oxi module | Hemolyzer |  | 01.06.2017 Herr Baldini 08517205301 meldet tHb - Fehler fHb, sO2, tHb | Oximetrie QC Fehler trotz durchgeführter tHb Kalibration | ABL90 manually rinsed Hemolyzer and valve hose renewed, tHb calibrated glucose at QC Level C too deep, SCi + LP renewed from customer base after SCi exchange glu at level C reinstalled by 2 mg $dL too high SCi Device calibrated and checked
Yes | 4 | Y | RAME | 0 Customer called to request assistance running tHb calibration and how to properly verify the result. |  | Customer was properly running both test and validating the results, however they were under the impression that there was a script that was supposed to walk them through both parts. | D2A0DDA5C40B2B06 | Re classified RMED | Customer training |  |  |  | 0 Customer called to request assistance running tHb calibration and how to properly verify the result. |  | Customer was properly running both test and validating the results, however they were under the impression that there was a script that was supposed to walk them through both parts.
No | Visit | Complaint | RDEU | Anwender müssen in letzter Zeit sehr häufig membranieren. | Wiederholte Driftfehler bei Metaboliten und Elektrolyten | Elektrolyt-Ersatzmodul aus GS eingesetzt, SW angepasst und mehrfach einkalibriert, Kal. und QC i.O., Kunde wird weiteres Verhalten beobachten | SWO-197799 | Defect | HW | El/Met Module | El/Met Module | Electrolyte and Met | Users have had to membrane very frequently lately. | Repeated drift errors in metabolites and electrolytes | Electrolyte replacement module from GS used, SW adapted and multiple calibrated, Kal. and QC i.O., customer will observe further behavior
No | Visit | Complaint | RDEU | Fehler 5684, 5784 | Reagenzienkassette verklemmt | Gerät überprüft, Cups entfernt, defekte Reagenzienkassette ersetzt, Messung o.k. | SWO-197314 | Defect | Consumables | Solution Pack | Cup Stack |  | Error 5684, 5784 | Reagent cassette jammed | Device checked, cups removed, defective reagent cartridge replaced, measurement ok
No | Visit | Complaint | RDEU | 30.05.2017 12:30Uhr Herr Heilig 071717011236: Display ist dunkel, der Barcodeleser leuchtet. | Geräteaufstellung beim Kunde nach Reparatur in GS17 | Einlaufverfahren Systemtest und QC durchgeführt. | SWO-197771 | Defect | HW | Analyzer Control Module | CPU |  | May 30, 2017 12:30pm Mr. Holy 071717011236: Display is dark, the barcode reader lights up. | Device installation at the customer after repair in GS17 | inlet procedure system test and QC.
Yes | Visit | Complaint | RDEU | Netzwerkanbindung an H+S | Netzwerkkonfiguration nach Kundenvorgabe durchgeführt. | Netzwerkeinstellungen an AQT eingegeben. Verbindung zu H+S wurde hergestellt. Med. Technik Herr Babel , Stationsleiter Herr Miehe und die IT wurden informiert und überzeugten sich vor Ort über die Funktionalität der Anbindung. | SWO-197773 | Re classified RMED |  |  |  |  | Netzwerkanbindung an H+S | network configuration according to customer specifications. | network settings to AQT. Connection to H+S has been established. Med. Technik Mr. Babel , station manager Mr. Miehe and the IT were informed and convinced themselves on site about the functionality of the connection.
No | Visit | Complaint | RDEU | 02.06.2017 08:50Uhr Herr Schwarz 015127648326: Alle Proben werden mit den Fehlermeldungen 1067, 1236 und 2872 abgebrochen | Probemmessung, Monovette ist aus Halter gefallen | Monovette aus Wet-Section entfernt Probelauf | SWO-197387 | User error | User error | User error | User error | Vacutainer fallen from holder | 06/02/2017 08:50Am Mr Schwarz 015127648326: All samples are cancelled with error messages 1067, 1236 and 2872 | Sample measurement, monovette has fallen from holder | Removed monovette from Wet-Section Rehearsal
No | Visit | Complaint | RDEU | Leckage erkannt. | Fehlersuche Verbindungsschlauch der Oxi- Modul und Sensorkassette undicht | Verbindungsschlauch getauscht. Einlassdichtung gereinigt. Alle Prozeduren in Service durchgeführt. Neustart 1 Std. | SWO-197810 | Defect | HW | Oxi module | Pinch valve tube |  | Leakage detected. | Troubleshooting Connection hose of the Oxi module and sensor cassette leaks | Connection hose is replaced. Inlet seal cleaned. All procedures performed in service. Restart 1 hr.
No | Visit | Complaint | RCHN | lac值低 | lac值低 | 更换测试卡后修复 | SWO-197694 | [5 empty] | Low lac value | Low lac value | Repair after replacing test card
No | Visit | Complaint | RCHN | 扫码器不扫码 | 扫码器不扫码 借902-924一个 | 更换902-924一个正常。(by tim) | SWO-197122 | Defect | HW | Basic Unit | Barcode reader, display |  | The code sweeper does not sweep the code | The codesweeper doesn't sweep the code, borrows 902-924 one. | Replace the 902-924 one normal. (by tim)
No | Visit | Complaint | RCHN | PO2质控不稳定 | PO2质控不稳定，借902-972，903-334各一个。 | 更换903-334，902-972各一个正常。(by tim) | SWO-197454 | [5 empty] | PO2 QC instability | PO2 QC is unstable, borrowing 902-972, 903-334 each. | Replace 903-334, 902-972 each normal. (by tim)
Yes | Visit | Complaint | RCHN | 系统温度超范围 加热器重复定标失败 1221，1326 | 系统温度超范围 加热器重复定标失败 1221，1326 | 更换测试卡正常。 | SWO-197494 | [5 empty] | System temperature overrange Heater repeat calibration failure 1221,1326 | System temperature overrange Heater repeat calibration failure 1221,1326 | The replacement test card is OK.
No | 20 | Y | RAME | I Temp errors |  | Customer said he received temperature errors 1065 and 1066. Prior to my arrival the errors cleared themselves and the analyzer was now working properly. Inspected SC pins and found all ok. Ran temperature calibration in the Service menu. Customer said they tried flushing the unit but no water came out of the inlet probe. | D2A9DF2EFE5BE238 | Cleaned | Fluidic pathway | Fluidic pathway | Clot | 1066 | I Temp errors |  | Customer said he received temperature errors 1065 and 1066. Prior to my arrival the errors cleared themselves and the analyzer was now working properly. Inspected SC pins and found all ok. Ran temperature calibration in the Service menu. Customer said they tried flushing the unit but no water came out of the inlet probe.
No | Visit | Complaint | RCHN | 接口板腐蚀 | 接口板腐蚀，借902-972一个 |  | SWO-197594 | [5 empty] | Interface plate corrosion | Interface plate corrosion, borrow 902-972 one
No | Visit | Complaint | RKK | ブランクキャル失敗を連発 | ブランクキャル失敗を連発 | 代替え器設置 | SWO-197646 | [5 empty] | Blank car failure spree | Blank car failure spree | Installation of an alternative device
No | Visit | Complaint | RIND | SC Temp- B problem. | Sensor board to replace. | Sensor board replaced. Zeroing calibration done and found heater not functioning. Need to replace Signal Data board. Asked from Pune. will be coming till Wednesday. | SWO-199081 | [5 empty] | SC Temp- B problem. | Sensor board to replace. | Sensor board replaced. Zeroing calibration done and found heater not functioning. Need to replace Signal Data board. Asked from Pune. will be coming till Wednesday.
Yes | Visit | Complaint | RCHN | 传感器响应错误 | 传感器响应错误 | 更换试剂包正常。 | SWO-202723 | [5 empty] | Sensor response error | Sensor response error | The replacement kit is OK.
No | Visit | Complaint | RCHN | 定标失败 1031 | 定标失败 1031，借944-214（12481）一张。 | 重新定标后修复(By Ran) recal, OK | SWO-197447 | Defect | HW | Process Module | Optical Unit | 1031 | Bid Failure 1031 | Bid failure 1031, borrow 944-214 (12481) one. | Re-calibration repair (By Ran) recal, OK
No | 3 | Y | RAME | 2 Customer states unit running low on QC. |  | Upon arrival, found unit had been running low on Peer 2. The past four days unit running well within Peer 2 ranges. Will install version 6.17 in unit. | D2A44760260E676B | Setup | Training | Training | Training |  | 2 Customer states unit running low on QC. |  | Upon arrival, found unit had been running low on Peer 2. The past four days unit running well within Peer 2 ranges. Will install version 6.17 in unit.
No | 5 | Y | RAME | F Customer reported by direct text oximetry errors on calibrations. Error code 0251, Halogen Lamp intensity too high. |  | Tested oximetry system. Found halogen lamp above 3800. | D2A49606262B62E4 | Adjustment | HW | Oxi module | Lamp unit | Halogen | F Customer reported by direct text oximetry errors on calibrations. Error code 0251, Halogen Lamp intensity too high. |  | Tested oximetry system. Found halogen lamp above 3800.
No | Visit | Complaint | RDEU | 20.03.17 16:00 Herr Kortz Tel.071164893256 , TouchScreen lässt sich nicht mehr richtig bedienen . 08.05.2017 09:40Uhr Frau Husbeck 071164893016: Der TouchScrenn lässt sich nur schwer bedienen, weil die Tasten seitlich versetz regaieren. | Touchscreen lässt sich nicht richtig bedienen. | Touch regiert vor allem unten links nicht korrekt. Verschobene Touchfunktion trotz mehrmaliger Touch Kalibration. ( siehe vorherige SWOs ) Monitor erneuert. ewf commit durchgeführt. | SWO-198005 | Defect | HW | DMS_PC | Touch screen |  | 20.03.17 16:00 Mr. Kortz Tel.071164893256, TouchScreen can no longer be operated properly. 08.05.2017 09:40uhr Ms. Husbeck 071164893016: The TouchScrenn is difficult to operate because the buttons are re-energizing on the side. | Touchscreen is not operated properly. | Touch does not govern correctly, especially at the bottom left. Shifted touch function despite several touch calibrations. ( see previous SWOs ) Monitor renewed. ewf commit.
Yes | Visit | Complaint | RDEU | Kein LP vor Ort | Gerät überprüft und neu gestartet | Kaltstartphase abgewartet und LP installiert, Kal und QCs durchgeführt | SWO-198048 | Re classified RMED | Logistical upSWO |  |  |  | No LP on site | Device checked and restarted | Cold start phase waited and LP installed, Kal and QCs performed
No | Visit | Complaint | RDEU | Einlassdichtung defekt --> Intervention erforderlich | Fehlersuche --> Einlassdichtung abegbrochen | Einlassdichtung erneuert, System und Funktion i.O. | SWO-198116 | Defect | HW | Inlet module | Inlet Gasket |  | Inlet seal defective --> Intervention required | Troubleshooting --> Inlet seal abegbrochen | Inlet seal renewed, system and function in the
No | Visit | Complaint | RDEU | Probenmischerbeleuchtung flackert, obwohl keine Spritze eingelegt | Probenmischerbeleuchtung flackert, obwohl keine Spritze eingelegt | Sample Mixer Unit erneuert Gerät kalibriert und überprüft | SWO-197766 | Defect | HW | Sampler Mixer | Sampler Mixer |  | Sample mixer lighting flickers even though no syringe is inserted | Sample mixer lighting flickers even though no syringe is inserted | Sample Mixer Unit renewed Device calibrated and checked
No | Visit | Complaint | RDEU | Benutzer-ID kommt bei Synlab nicht an. | Benutzer-ID kommt bei Synlab nicht an. | Konfiguration überprüft, in Ordnung. Trance mit astmnet erstellt und Herrn Handl, Synlab zugeschickt. Benutzer-ID auf Trance, aber nicht in Conworx sichtbar. Rückruf Herr Korte, Siemens. Laut Herrn Korte ist Benutzer-ID in der Datenbank sichtbar, wird aber nicht an Synlab weitergeleitet. Herr Korte übernimmt die weitere Fehlersuche. | SWO-198122 | No fault found | His/lis | His/lis | His/lis |  | User ID does not arrive at Synlab. | User ID does not arrive at Synlab. | Configuration checked, all right. Trance created with astmnet and sent to Mr. Handl, Synlab. User ID on trance, but not visible in Conworx. Recall Mr. Korte, Siemens. According to Mr. Korte, user ID is visible in the database, but is not forwarded to Synlab. Mr Korte will continue to search for the bug.
No | 4 | Y | RFRA | Blank | Erreurs de lecture de pack+ Base de données corrompue |  | D2A5911B9C2B140F | Defect | HW | Basic Unit | Solution Pack ID PCB |  | Blank | Pack-to-pack reading errors Corrupted Database
Yes | 1 | Y | RPAC | Blank | Sensor response error |  | D2A5AF569C546EEC | [5 empty] | Blank | Sensor response error
No | Visit | Complaint | RDEU | Analysator ist erst seit ein paar Tagen bei Anwender. gedruckt wird das Feld Patienten ID, aber nicht die gemessenen Werte. | Kein Ausdruck | Gerät konfiguriert | SWO-197929 | [5 empty] | Analyzer has only been with users for a few days. The Patient ID field is printed, but not the measured values. | No expression | Device configured
No | 2 | Y | RESP | TU0153; FAN 915-353; 915-361 Error en la cal de lactato. |  | Sustitucion de la sensor interface pcb. Actualizacion soft 3.3 Cal, qc | D2A5E6FC9C577ED7 | Defect | HW | Basic Unit | PCB, Sensor interface |  | TU0153; FAN 915-353; 915-361 Lactate cal error. |  | Replacement of the pcb interface sensor. Soft update 3.3 Cal, qc
No | Visit | Complaint | RCHN | tni定标不过 | tni定标不过,借一个903-364 |  | SWO-197687 | Defect | HW | Process Module | Optical Unit | 1031 | tni calibration but | tni calibration, however, borrow a 903-364
No | Visit | Complaint | RCHN | 报错1031 | 报错1031，借903-364一个 | 没有故障，工程师报错仪器序列号了。(By Ran) wrong SN. CASE closed. | SWO-197888 | Defect | HW | Process Module | Optical Unit | 1031 | 1031 | 1031, loan 903-364 one | There was no fault and the engineer reported the wrong instrument serial number. (By Ran) wrong SN. CASE closed.
No | Visit | Complaint | RCHN | 3414-FTC-101 | 3414-FTC-101 借备件903-119一个 | 更换903-119备件一个后修复 | SWO-197593 | Defect | HW | Process Module | Cup wash unit |  | 3414-FTC-101 | 3414-FTC-101 Spare parts 903-119 one | Replace 903-119 spare parts with one post-fix
No | Visit | Complaint | RCHN | 报修1031 | 报修1031，借备件903-364一个 |  | SWO-197899 | Defect | HW | Process Module | Optical Unit | 1031 | Repair 1031 | Repair 1031, spare parts 903-364 one
No | Visit | Complaint | RCHN | TNI定标中止 | TNI定标中止，借944-212（12862）2张 |  | SWO-197688 | Assisted Replacement | Consumables | Test Kit | Cartridge, Tni |  | TNI calibration abort | TNI calibration suspension, borrowing 944-212 (12862) 2 sheets
No | Visit | Complaint | RDEU | 06.06.2017, 09:20 Uhr ==> Anrufer Fr. Mentrup, Tel.: 0252299441 ==> Fehler 0793: Pumpenasymetriefehler pH/BG-Modul, Fehler ist nach Pumpenkalibration weg und tritt aber wieder auf | Pumpenasym. Fehler | Fehler wurde durch fällige Wartung behoben | SWO-198654 | Defect | HW | pH/BG Module | Pump tubes | Assymetri | 06.06.2017, 09:20 ==> Caller Ms. Mentrup, Tel.: 0252299441 ==> Error 0793: Pump asymetry error pH/BG module, error is gone after pump calibration and reappears | Pumpasym. Error | Error fixed due maintenance
No | Visit | Complaint | RCHN | 无法开机 | 无法开机，借备件910-332一个 | 910-332有质量问题，重新开单借910-274、910-332和910-321各一个 | SWO-197588 | Revisit required | HW | DMS_PC | CPU | Used analyzer | Cannot power on | Cannot power on, borrow spare parts 910-332 one | 910-332 has quality issues, re-billing borrowing 910-274, 910-332 and 910-321 each
No | Visit | Complaint | RCHN | LAC定标失败 | LAC定标失败 | 更换LAC电极膜正常。 | SWO-197764 | Assisted Replacement | Electrode | Membrane | Lac |  | LAC calibration failure | LAC calibration failure | Replacing the LAC electrode film is normal.
No | Visit | Complaint | RCHN | 零点电流错误 0408 | 零点电流错误 0408 ，借902-829一个。 | 更换电滤板备件一个 | SWO-197695 | Defect | HW | El/Met Module | PCB, El/Met | 408, LEH contact | Zero current error 0408 | Zero current error 0408 , borrowed 902-829 one. | Replace one of the panel spare parts
No | Visit | Complaint | RHUN | Szennyeselszívó cső eldugult. | Műszer folyik. | Szennyeselszívó cső eldugult, cső csere után a folyás megszűnt. | SWO-197991 | Defect | HW | Inlet module | Tubings | Clot | Dirt suction tube is blocked. | Instrument is underway. | Dirty suction tube is blocked, after pipe replacement the discharge is discontinued.
No | 3 | Y | RPAC | Blank | Windows corrupted when customer trying to obtain MAC address |  | D2A5CB9D9C006233 | Setup | SW | His/lis | His/lis |  | Blank | Windows corrupted when customer trying to obtain MAC address
No | Visit | Complaint | RDEU | 06.06.2017 11:30Uhr Frau Eilers-Gladrow 05315950: Hb wird nicht gemessen. Hardwarefehler bei Kal | System Status Datei überprüft - Halogenlamp Voltage too high | Halogenlamp justiert, Kal, QC | SWO-198059 | Adjustment | HW | Oxi module | Lamp unit | Halogen | 06.06.2017 11:30pm Mrs Eilers-Gladrow 05315950: Hb is not measured. Hardware failure at Kal | System Status File Checked - Halogenlamp Voltage Too High | Halogenlamp justiert, Kal, QC
Yes | Visit | Complaint | RLTD | Annette Tel 353 61462119 or 462168 | Adding operators issue | Set both machines back to centralised user use. Site can now push operators. | SWO-198263 | Re classified RMED | His/Lis settup |  |  |  | Annette Tel 353 61462119 or 462168 | Adding operators issue | Set both machines back to centralised user use. Site can now push operators.
No | Visit | Complaint | RDEU | 06.06.2017, 12:50 Uhr ==> Anrufer Schw. Simone, Tel.: 022133085058 ==> Analysator geht nicht aus der Fehlerbehebungsroutine, Kunde hat schon 2x SP erneuert und besitzt keine neuen Einlassdichtungshalter 06.06.2017, 15:15 Uhr ==> Anrufer Schw. Manuela, Tel.: 022133085058 ==> Fehler besteht weiterhin, Kunde wünscht den Einsatz eines Technikers, kein Einlassdichtungshalter vor Ort vorhanden | s. case | Fehlersuche: Einlassdichtung ersetzt - weiterhin Fehler. System in beide Richtungen gespült, ohne Erfolg. SP ersetzt, ohne Erfolg. Alle Fließwege auf Dichtheit überprüft, Funktion und Leichtgängigkeit des Bügelventils geprüft, ok. Einlassnadel ersetzt, Fluidic Programme im Servicemenü absolviert, weiterhin fehlerhafter Flüssigkeitstransport, diverse Lösungen werden nicht gefördert. Jahreswartung mit Wartungskit vorgezogen (s. gesonderter case) und Hemolyzer auf Durchlässigkeit überprüft. Neustart, nach etwa 5 min wieder in der geführten Fehlerbehebung => SC ersetzt und erneut sämtliche bezogenen Serviceprogramme mehrfach wiederholt, nun (mit neuer SC) regelmäßig Vakuumfehler => SC zurück getauscht. Fluidic Programme nun ok, Neustart, Startphase, Kal und QC Messungen OK. | SWO-198446 | Revisit required | SC | SC | FT | Vacum | 06.06.2017, 12:50 o'clock ==> Caller Schw. Simone, Tel.: 022133085058 ==> Analyzer does not go out of the troubleshooting routine, customer has already renewed 2x SP and has no new inlet seal holders 06.06.2017, 15:15 o'clock ==> Caller Schw. Manuela, Tel.: 022133085058 ==> Error persists, customer wishes to use a technician, no inlet seal holder available on site | s. case | Troubleshooting: Inlet seal replaced - still errors. System flushed in both directions, to no avail. SP replaced, to no avail. All flow paths checked for leak, function and smoothness of the ironing valve checked, ok. Replacement of the inlet needle, fluidic programs completed in the service menu, further faulty fluid transport, various solutions are not promoted. Annual maintenance with maintenance kit preferred (see separate case) and Hemolyzer checked for permeability. Restart, replaced after about 5 min in the guided troubleshooting => SC and again repeated all service programs several times, now (with new SC) regularly replaced vacuum errors => SC. Fluidic programs now ok, restart, start phase, Kal and QC measurements OK.
No | 5 | Y | RCZE | Blank | po vèerejší výmìnì SC není zkalibrována Glu ani po 24 hod., hodnoty ... |  | D2A843C688363DBA | Revisit required | HW | Basic Unit | PCB, Sensor interface | Techline RCZE 781396503 | Blank | After the last change of the SC is not calibrated Glu even after 24 hours, the values...
No | Visit | Complaint | RDEU | 06.06.2017, Frau Niebius, Tel: 03445721750: Flüssigkeit läuft unten am Gerät aus | Gerät läuft aus | Störung im Rahmen der Wartung beseitigt | SWO-198185 | [5 empty] | 06.06.2017, Mrs. Niebius, Tel: 03445721750: Liquid runs out at the bottom of the device | Device expires | Trouble-free maintenance
Yes | 16 | Y | RAME | F ENTIRE OXY PANEL PARAMETERS WERE DISABLED DUE TO AIR DETECTED IN FLUIDIC PATHLINE, E1218. |  | CONFIRMED ISSUE WHEN ONSITE. | D2A6E5F31232B253 | [5 empty] | F ENTIRE OXY PANEL PARAMETERS WERE DISABLED DUE TO AIR DETECTED IN FLUIDIC PATHLINE, E1218. |  | CONFIRMED ISSUE WHEN ONSITE.
No | Visit | Complaint | RIND | SC Temp- B problem. | Signal Data Board to replace. | Signal Data board replaced. Zeroing calibration done. SC and SP installed. Sample analysis done. Analyser serviced and calibrated properly. Handed over in a fine working condition. | SWO-199087 | [5 empty] | SC Temp- B problem. | Signal Data Board to replace. | Signal Data board replaced. Zeroing calibration done. SC and SP installed. Sample analysis done. Analyser serviced and calibrated properly. Handed over in a fine working condition.
No | Visit | Complaint | RCHN | 报错1031 | 报错1031，借备件903-364一个 |  | SWO-198105 | Defect | HW | Process Module | Optical Unit | 1031 | 1031 | Error 1031, borrowed spare parts 903-364 one
No | Visit | Complaint | RCHN | 报错1031 | 报错1031，借备件903-364一个 |  | SWO-198081 | Defect | HW | Process Module | Optical Unit | 1031 | 1031 | Error 1031, borrowed spare parts 903-364 one
No | Visit | Complaint | RCHN | 报1024错误 | 报1024错误，借903-364一个 |  | SWO-198200 | Defect | HW | Process Module | Optical Unit | 1031 | Reported 1024 error | Reported 1024 error, borrow903-364 one
No | Visit | Complaint | RCHN | 打印机损坏 | 打印机损坏，借910-275一个 |  | SWO-198228 | Defect | HW | Printer Module | Printer Module |  | Damaged printer | Printer damage, borrow 910-275 one
No | Visit | Complaint | RCHN | 无法开机 | 无法开机，借的910-332有质量问题，重新借910-274、910-332和910-321各一个 | 更换910-332后修复(By Ran) | SWO-197706 | Defect | HW | DMS_PC | HDD |  | Cannot power on | Cannot be switched on, the loan 910-332 has quality problems, re-borrow 910-274, 910-332 and 910-321 each | Repair after replacement 910-332 (By Ran)
Yes | 2 | Y | RCZE | Blank |  |  | D2A6E1DF1246AE3E | Re classified RMED | Preventive maintenance |  |  |  | Blank
No | Visit | Complaint | RDEU | Anwender müssen in letzter Zeit sehr häufig membranieren. | Driftprobleme bei Glu | Glu und Messkammer gereinigt, Glu neu membraniert und einkalibriert, Kal. Und QC i.O. | SWO-198310 | Cleaned | HW | El/Met Module | Measuring chamber |  | Users have had to membrane very frequently lately. | Drift problems with Glu | Glu and measuring chamber cleaned, Glu re-membraneed and calibrated, Kal. And QC i.O.
No | 4 | Y | RFRA | Blank | Pas d'ouverture de la bay unit lors du remplacement de la cassette |  | D2A95B5EFE2AEBFA | Defect | HW | Basic Unit | SC Door |  | Blank | No opening of the bay unit when replacing the cassette
No | Visit | Complaint | RDEU | pO2 Meldung und Intervention erforderlich | Leak voltage liegt bei 367,387 mV | Erneuerung des Solutions pack gegen konditioniert Ware | SWO-198657 | Assisted Replacement | SP | SP | Leak current | 1331 | pO2 Notification and intervention required | Leak voltage is 367,387 mV | Renewal of the solution pack against conditioned goods
No | Visit | Complaint | RDEU | Gerät bootet nicht richtig. ABL bleibt im Offlinemodus hängen, keine AU Verbindung. | Fehlersuche mit Teilen aus Demogerät. | PC Unit getauscht, kein Erfolg. AU PCB Getauscht , kein Erfolg. Netzteil getauscht, kein Erfog. Monitor getauscht, Kein Erfolg. Alle getauschten Module in Demogerät eingebaut. alles o.k. Im Kundengerät Einlassmodul Einheit getauscht, kein Erfolg. Mischer Unit für sPico getauscht, alles o.k. . Gerät bootet wieder normal. Mischer Unit defekt. Alle Module wieder original zurückgebaut. MischerUnit von AU PCB getrennt. Gerät bootet normal . Mischer Unit muß bestellt werden. Einlaufverfahren durchgeführt. | SWO-198687 | Revisit required | HW | Sampler Mixer | Sampler Mixer |  | Device does not boot properly. ABL is stuck in offline mode, no AU connection. | Troubleshooting with parts from demo device. | PC Unit exchanged, no success. AU PCB Exchanged, no success. Power supply replaced, no erfog. Monitor swapped, No success. All replaced modules built into demo device. everything ok. In the customer device inlet module unit exchanged, no success. Mixer unit for sPico exchanged, all OK. . Device boots back to normal. Mixer unit defective. All modules restored to their original. separating MixerUnit from AU PCB. Device boots normally. Mixer unit must be ordered. run-in procedures.
No | Visit | Complaint | RDEU | 06.06.2017, 12:50 Uhr ==> Anrufer Schw. Simone, Tel.: 022133085058 ==> Analysator geht nicht aus der Fehlerbehebungsroutine, Kunde hat schon 2x SP erneuert und besitzt keine neuen Einlassdichtungshalter 06.06.2017, 15:15 Uhr ==> Anrufer Schw. Manuela, Tel.: 022133085058 ==> Fehler besteht weiterhin, Kunde wünscht den Einsatz eines Technikers, kein Einlassdichtungshalter vor Ort vorhanden 07.06.2017 19:30 Sr Simone 022133081750 Fehler bleibt bestehen | s. case | ABL hat gestern Abend und heute morgen wieder den selben Fehler gezeigt; das SP wurde anwenderseitig 2 x neu initialisiert - beim Eintreffen grüne Ampel. Im Servicemenü Pressure Test aufgerufen, ok. Einlassdichtung erneuert => Vakuumfehler, permanent reproduzierbar; zuvor eingesetzte (gestern erneuerte) Einlassdichtung eingesetzt => Vakuumfehler, permanent reproduzierbar; weitere (3. neue) Einlassdichtung eingesetzt => keine weiteren Fehlermeldungen, mehrfach sämtliche Service-Testprogramme absolviert, alle fehlerfrei! Mehrere Blutmessungen Spritze $Kapillare abgewartet, alle OK. ServiceDump gezogen, QC Messung OK | SWO-198674 | Defect | HW | Inlet module | Inlet Gasket |  | 06.06.2017, 12:50 o'clock ==> Caller Schw. Simone, Tel.: 022133085058 ==> Analyzer does not go out of the troubleshooting routine, customer has already renewed 2x SP and has no new inlet seal holders 06.06.2017, 15:15 o'clock ==> Caller Schw. Manuela, Tel.: 022133085058 ==> Error persists, customer wishes to use a technician, no inlet seal holder available on site 06/07/2017 19:30 Sr Simone 022133081750 Error remains | s. case | ABL showed the same mistake again last night and this morning; the SP has been reinitialized 2 times on the user side - when green traffic lights arrive. Called in the pressure test service menu, ok. Inlet seal renewed => Vacuum fault, permanently reproducible; previously used (yesterday renewed) inlet seal => vacuum fault, permanently reproducible; further (3rd new) inlet seal inserted => no further error messages, completed several times all service test programs, all error-free! Several blood measurements syringe $Kapillare waited, all OK. ServiceDump pulled, QC measurement OK
No | Visit | Complaint | RDEU | 08.06.2017 Frau Wingerter 062181024320 meldet eine Leckage. 08.06.2017 10:00 Fr. Wingerter: Status von Complete auf Waiting FSE umgestellt (FSE wurde ausgelöst) | Analysator steht in Intervention erforderlich mit der Meldung: Leckage erkannt. Überprüfung des Analysators im Service. Flow Test negativ. | Fließwege mit Spülwerkzeug durchgespült. Fließwege wieder frei, verstopft durch Gerinnsel. Verwendung von Sarstedt Probenmehmern. Nach Neustart und erneuter Initialisierungsphase ist Analysator funktionsbereit. | SWO-198841 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway | 1184 | 06/08/2017 Ms Wingerter 062181024320 reports a leak. 06/08/2017 10:00 Fr. Wingerter: Status changed from Complete to Waiting FSE (FSE was triggered) | Analyzer is required in intervention with the message: Leakage detected. Checking the analyzer in the service. Flow test negative. | Flow paths rinsed with rinsing tool. Flow paths free again, clogged by clots. Use of Sarstedt sample companies. After rebooting and reinitialization, Analyser is ready for use.
No | Visit | Complaint | RDEU | 07.06.2017 08:45Uhr Sr.Barbara 03381781926: Wegen einem Lac-Fehler wurde die SC getauscht. Aktuell wieder Lac-Fehler 08.06.2017 8:30 Sr.Barbara 03381781926: Nach wie vor LAC-Fehler | Sensorfach überprüft, dabei festgestellt, dass sich Ablagerungen / Auskristallisierungen auf den goldenen Pins befinden zudem sind 2 Pins kürzer | SC und LP zum Test ersetzt, aber Problem dass Empf. bei Lac = 0 ist, bleibt bestehen es ist bereits die 3. SC mit gleichem Fehlerbild vermutlich Sensor Interface Board defekt - muss bestellt werden Gerät wird zur Reparatur mitgenommen | SWO-198847 | Revisit required | HW | Basic Unit | PCB, Sensor interface | Lac | 06/07/2017 08:45H Sr.Barbara 03381781926: Sc was replaced due to a Lac error. Currently again Lac error 06/08/2017 8:30 Sr.Barbara 03381781926: Lac Errors Still | Sensor compartment checked, found that deposits / crystallizations are located on the golden pins in addition, 2 pins are shorter | SC and LP replaced for testing, but problem that Empf. at Lac = 0, remains it is already the 3rd SC with the same error image presumably sensor interface board defective - must be ordered Device is taken for repair
Yes | 5 | Y | RCZE | Blank |  |  | D2A83EC6882642F5 | [5 empty] | Blank
No | Visit | Complaint | RLTD | Unable to access patient logs from Aqcure remote access | Patient logs could not be accessed from internal remote support. On entering the service menu and internal remote support it was was found that patient log access was disabled. | The disable patient record view box was unchecked. | SWO-198836 | Setup | SW | His/lis | Remote support |  | Unable to access patient logs from Aqcure remote access | Patient logs could not be accessed from internal remote support. On entering the service menu and internal remote support it was was found that patient log access was disabled. | The disable patient record view box was unchecked.
No | Visit | Complaint | RBNL | Afgebroken bloedmonsters | Er zijn problemen met de capillaire metingen bij de ABL90. Tijdens de meting (meestal capillaire metingen), de inletgasket blijft vast zitten en daardoor de meting wordt afgebroken. | De Inletmodule heb vervangen vervolgens heb gecontroleerd. | SWO-198828 | Defect | HW | Auto Inlet module | Auto Inlet module |  | Broken down blood samples | There are problems with the capillary measurements at the ABL90. During the measurement (usually capillary measurements), the inletgasket remains stuck and therefore the measurement is broken down. | The Inletmodule have replaced then checked.
Yes | Visit | Complaint | RLTD | AQURE connection and configuration. | Checked Analyser connection to AQURE as it is showing as not connected. | Reconfigured Analyser, checked and tested. Setup Analyser to appear on AQURE and allocated it to Newham - Rainbow Ward. | SWO-198849 | Re classified RMED | Installation | His/lis | Aqure |  | AQURE connection and configuration. | Checked Analyser connection to AQURE as it is showing as not connected. | Reconfigured Analyser, checked and tested. Setup Analyser to appear on AQURE and allocated it to Newham - Rainbow Ward.
No | 5 | Y | RCZE | Blank | instalována ovìøená SC a RP ( bylo vyzkoušeno a instalováno na pøístroji R0116N008 ), hodnoty pøed odebráním ze stroje Glu stav 972, citl. 191 a Lac stav 947, citl. 351 = OK |  | D2A8453C884C3DBD | Defect | HW | Basic Unit | PCB, Sensor interface | Techline RCZE 781396503 | Blank | Installed SC and RP (tested and installed on the device R0116N008), the values before removal from the machine Glu 972 state, Citl. 191 and Lac status 947, Citl. 351 = OK
No | 17 | Y | RAME | Z Leaking from solution pack. |  | The solution pack door is broken again and not pushing the SP in all the way. The analyzer is failing Glu. Paula will let it run and then run the QC later. She said it needs to hydrate. | D2A81FC98808EDAB | [5 empty] | Z Leaking from solution pack. |  | The solution pack door is broken again and not pushing the SP in all the way. The analyzer is failing Glu. Paula will let it run and then run the QC later. She said it needs to hydrate.
No | 12 | Y | RAME | Z Corrupt system file. |  | Paula had a possible fluid blockage. I suggested turning it off, remove solution pack and sensor cassette and then turn back on. When it came back up it showed corrupt system file and corrupt parameter file. It gavee all parametes accross the top and said no reportable parameters. Reloaded SW did not solve issue. Replaced SC and it solved the issue. | D2A81919882BEDA8 | [5 empty] | Z Corrupt system file. |  | Paula had a possible fluid blockage. I suggested turning it off, remove solution pack and sensor cassette and then turn back on. When it came back up it showed corrupt system file and corrupt parameter file. It gavee all parametes accross the top and said no reportable parameters. Reloaded SW did not solve issue. Replaced SC and it solved the issue.
No | Visit | Complaint | RIND | SC Temp- B problem. | Analyser to observe . | Observation of analyser done. And found working fine. Analyser calibrated and serviced properly. Handed over in a fine working condition. | SWO-199089 | [5 empty] | SC Temp- B problem. | Analyser to observe . | Observation of analyser done. And found working fine. Analyser calibrated and serviced properly. Handed over in a fine working condition.
No | Visit | Complaint | RDEU | 08.06.2017, Frau Dammann, Tel.: 04441853685: Fehlermeldung bei jeder Messung; | Fehler 581 Oxiparameter werden nicht gemessen , mögliche Verstopfung | Gerät geprüft , Diverse Messungen mit Hypochlorid durchgeführt. Sensorcassette erneuert (Verfallsdatum erreicht) Gerät demontiert , Quetschventil und Fliesswege überprüft. Hemolyzer erneut gespült. Anschließend kal , thb kal und QC durchgeführt. Blutige Messung auch ok. | SWO-199148 | [5 empty] | 08.06.2017, Mrs. Dammann, Tel.: 04441853685: Error message for each measurement; | Error 581 Oxi parameters are not measured, possible constipation | Device tested, Various measurements performed with hypochloride. Sensor cassette renewed (expiry date reached) Device dismantled, squeezing valve and flow paths checked. Hemolyzer rinsed again. Subsequently kal , thb kal and QC performed. Bloody measurement also ok.
No | Visit | Complaint | RCHN | 002软件版本错误，1024定标底部外泄 | 002软件版本错误，1024定标底部外泄 借903-364和636-564备件各一个 | 更换903-364备件后修复（By Ran） | SWO-198212 | Defect | HW | Process Module | Optical Unit | 1024 | 002 software version error, 1024 calibration bottom leak | 002 software version error, 1024 calibration bottom leak borrowed 903-364 and 636-564 spare parts one each | Repair after replacement of 903-364 spare parts (By Ran)
No | Visit | Complaint | RCHN | 报错1031 | 报错1031，借903-364一个，Tni944-212（12945）和PCT944-394（12546），D二聚体定标卡944-220，（12691和12568）各一张 |  | SWO-198467 | Defect | HW | Process Module | Optical Unit | 1031 | 1031 | 1031, loan 903-364 one, Tni944-212 (12945) and PCT944-394 (12546), D dipolypation calibration card 944-220, (12691 and 12568) each
No | Visit | Complaint | RCHN | 报错1031 | 报错1031，借944-258（12413),944-214(12338),944-394(12546)各一张。 |  | SWO-198486 | Defect | HW | Process Module | Optical Unit | 1031 | 1031 | 1031, borrow 944-258 (12413), 944-214 (12338), 944-394 (12546) each.
No | Visit | Complaint | RCHN | PCT定标失败 | PCT定标失败，借944-394（12450）一张。 |  | SWO-198490 | Assisted Replacement | Consumables | Test Kit | Cartridge, PCT |  | PCT calibration failed | The PCT bid failed, borrowing 944-394 (12450).
No | Visit | Complaint | RCHN | 报错5784 | 报错5784 | 更换测试卡后修复 | SWO-198672 | Defect | Consumables | Test Kit | Cartridge | 5784 | 5784 | 5784 | Repair after replacing test card
Yes | Visit | Complaint | RCHN | 报错1031 | 报错1031，借备件903-364一个 | 更换903-364备件后修复(By Ran) | SWO-198082 | Re classified RMED | OOB |  |  | 1031 | 1031 | Error 1031, borrowed spare parts 903-364 one | Repair after replacement of 903-364 spare parts (By Ran)
Yes | 8 | Y | RPAC | Blank | Crea Issues |  | D2A85664881C0918 | [5 empty] | Blank | Crea Issues
No | Visit | Complaint | RDEU | 07.06.2017, Herr Reimer, Tel.: 02151324907: AQC Fehler Level 3 alle Parameter seit dem Wochenende | In den letzten Tagen sind bei den Hb Parametern viele Kalibrierdriften und Fehlerhafte QC Ergebnisse aufgetreten. | Servicedump erstellt, Oximetrie-Modul ausgetauscht und Reklamatiert. Service tHb und tHb kal durchgeführt, Kalibriert und QC's durchgeführt. | SWO-198724 | Defect | HW | Oxi module | Oxi module, Replaced |  | 06/07/2017, Mr. Reimer, Tel.: 02151324907: AQC Error Level 3 all parameters since the weekend | In recent days, the Hb parameters have experienced many calibration drifts and faulty QC results. | Service dump created, oximetry module replaced and advertised. Service tHb and tHb kal performed, calibrated and QC's performed.
Yes | 20 | Y | RAME | Z Customer reports that sensor cassette(SC) bay will not open |  | Customer attempted to replace the SC but the SC bay would not open. Although, the gears could be heard turning. Analyzer was just PM'd 6 $8 $2017. During the PM, the bay door was closed beyond the point that would allow the analyzer to normally open it. | D2A9A3B5FE09D60E | [5 empty] | Z Customer reports that sensor cassette(SC) bay will not open |  | Customer attempted to replace the SC but the SC bay would not open. Although, the gears could be heard turning. Analyzer was just PM'd 6 $8 $2017. During the PM, the bay door was closed beyond the point that would allow the analyzer to normally open it.
No | Visit | Complaint | RKK | 1543と1084を頻発する | 状態確認 | 部品交換 | SWO-199318 | Defect | HW | Inlet module | Inlet Gasket |  | Frequently 1543 and 1084 | Status Check | Parts replacement
No | Visit | Complaint | RDEU | 07.06.2017 08:45Uhr Sr.Barbara 03381781926: Wegen einem Lac-Fehler wurde die SC getauscht. Aktuell wieder Lac-Fehler 08.06.2017 8:30 Sr.Barbara 03381781926: Nach wie vor LAC-Fehler | . | Gerät wieder zum Kunden gebracht und funktionstüchtig übergeben | SWO-199155 | Defect | HW | Basic Unit | PCB, Sensor interface | Lac | 06/07/2017 08:45H Sr.Barbara 03381781926: Sc was replaced due to a Lac error. Currently again Lac error 06/08/2017 8:30 Sr.Barbara 03381781926: Lac Errors Still | . | Device brought back to the customer and handed over functionally
No | 2 | Y | RAME | W System disk error |  | Upon boot up analyzer would give system disk error, insert system disk message. | D2A96915FE06DEDD | Defect | HW | DMS_PC | CF Card |  | W System disk error |  | Upon boot up analyzer would give system disk error, insert system disk message.
Yes | 3 | Y | RPAC | Blank | Failed pH and Electrolytes. |  | D2A956B4FE0BF215 | [5 empty] | Blank | Failed pH and Electrolytes.
No | Visit | Complaint | RDEU | Scanner defekt | Scanner springt nicht an | Gerät und Scanner kontrolliert, Scanner erneuert | SWO-199018 | Defect | HW | Basic Unit | Barcode reader, display |  | Scanner defective | Scanner does not start | Device and scanner checked, scanner renewed
No | Visit | Complaint | RDEU | 09.06.2017, Sr. Marion, Tel.: 0211904361241: mögliche Verstopfung der Fließwege | Fließwege im Serviceprogramm überprüft | SC gespült, Fließwege gereinigt, LS Justage durch reboot ausgelöst, Funktion OK. | SWO-199107 | [5 empty] | 09.06.2017, Sr. Marion, Tel.: 0211904361241: possible blockage of the flow paths | Flow paths checked in the service program | SC flushed, flow paths cleaned, LS adjustment triggered by reboot, function OK.
No | Visit | Complaint | RCHN | 1304温度不稳定 | 1304温度不稳定,需要借件902-972一个 | 更换902-972后修复(By Ran) | SWO-198579 | [5 empty] | 1304 Temperature Instability | 1304 Temperature is unstable and requires a loan of 902-972 | Repair after replacement 902-972 (By Ran)
Yes | Visit | Complaint | RCHN | LAC,GLU定标失败 | LAC,GLU定标失败 | 连续更换2张测试卡后修复 | SWO-199201 | [5 empty] | LAC, GLU bid failed | LAC, GLU bid failed | Repair after 2 consecutive test cards
No | Visit | Complaint | RCHN | 报1031错误 | 报1031错误，借903-364一个、944-216（12758）一张 | 更换903-364，重新定标正常。 by andrew | SWO-198968 | Defect | HW | Process Module | Optical Unit | 1031 | Reported 1031 error | Reported 1031 error, borrow 903-364 one, 944-216 (12758) one | Replace the 903-364 and re-mark normally. by andrew
No | Visit | Complaint | RCHN | 进样口和测试卡舱门无法关闭 | 进样口和测试卡舱门无法关闭，借903-198和636-551各一个 | 更换903-198，636-551后故障依旧，加借903-117，903-028各一个。 by andrew | SWO-199029 | Revisit required | HW | Analyzer Control Module | PCB, Main Board |  | The sample and test card doors cannot be closed | The sample entrance and test card door cannot be closed, with one each of the 903-198 and 636-551 | After replacing 903-198, 636-551, the fault remains, plus 903-117, 903-028 each one. by andrew
No | Visit | Complaint | RKK | キャル時、Hb値のドリフトエラーが多発 | キャル時のドリフト、QC測定時のHb測定値エラー多発 | 代替えモジュール設置でOxiモジュール引き上げ作業実施 | SWO-202120 | Defect | HW | Oxi module | Oxi module, Replaced |  | Frequent drift errors of Hb values at the time of | Drift during cal, Hb measurement error frequent lye at QC measurement | Oxi module pulling work implemented by installing an alternative module
No | Visit | Complaint | RDAN | Blank Layout fix før SW 6.17 | NA | Software opdateret til ver. 6.17 Setup rettet | SWO-199381 | Defect | SW | SW | Blank Layout |  | Blank Layout fix før SW 6.17 | In | Software updated for Ver. 6.17 Setup fixed
No | 9 | Y | RAME | P Customer reported inhomogeneous errors during QC samples. |  | Checked data logs. Found error code 0374: Inhomogeneous sample at pH $BG module during QC runs. Some fluid leakage from front of analyzer. Customer already replaced inlet probe and inlet gasket. | D2A989B9FE3F625D | Adjustment | HW | Inlet module | Probe cal. | and Tubings | P Customer reported inhomogeneous errors during QC samples. |  | Checked data logs. Found error code 0374: Inhomogeneous sample at pH $BG module during QC runs. Some fluid leakage from front of analyzer. Customer already replaced inlet probe and inlet gasket.
No | Visit | Complaint | RDEU | Lösungen können nicht mehr gepumpt werden | nachdem der LP heute nach gewechselt wurde, erscheint die Meldung: Lösungen können nicht gepumpt werden am Tel. festgestellt, dass der Pressure Test nicht i.O. ist, auch nach Wechsel des Einlasshalters bleibt das Problem bestehen ein weiterer Einlasshalter oder LP ist nicht vorhanden Kunde Ärztin Frau Dr. Hedayati beauftragte kostenpflichtigen WE Service | Einlasshalter gewechselt, dann Pressure Test i.O. Service Programme durchlaufen lassen i.O. Startup abgewartet - Gerät stand seit heute Nacht um 02:00Uhr Anmerkung: Bitte in Zukunft für ausreichend Material sorgen. Es müssen zu jeder Zeit mind. 2 LP, 2 SC und 2 Einlasshalter auf Vorrat gehalten werden. | SWO-199269 | Defect | HW | Inlet module | Inlet Gasket |  | Solutions can no longer be pumped | after the LP has been changed today, the message appears: Solutions cannot be pumped on the phone it was determined that the pressure test is not o.o., even after changing the inlet holder the problem remains another inlet holder or LP is not available Customer doctor Dr. Hedayati commissioned paid WE service | Inlet holder changed, then Pressure Test i.O. Let service programs go through as a. Startup waited - device stood since tonight at 02:00o Note: Please ensure sufficient material in the future. At least 2 LP, 2 SC and 2 inlet holders must be kept in stock at all times.
No | Visit | Complaint | RCHN | 报错1031 | 报错1031，借903-364一个 | 更换903-364后修复(By Ran) | SWO-198894 | Defect | HW | Process Module | Optical Unit | 1031 | 1031 | 1031, loan 903-364 one | Repair after replacement 903-364 (By Ran)
No | Visit | Complaint | RCHN | 温度强度超出范围，定标底部外露 | 温度强度超出范围，定标底部外露，需要借件903-364一个 | 更换903-364后修复(By Ran) | SWO-198895 | Defect | HW | Process Module | Optical Unit |  | Temperature intensity out of range, calibration bottom exposed | Temperature intensity out of range, calibration bottom exposed, need to borrow 903-364 a | Repair after replacement 903-364 (By Ran)
No | Visit | Complaint | RDEU | Alle gültigen Kalibrationen nicht mehr vorhanden und AQT verlangt immer Neustart 1999 undefiniert lookup String 46.1 , 3264 Environment | System überprüft | Lichtschranken gereinigt, optische Einheit gereinigt, Neustart durchgeführt , Testmessungen und QC durchgeführt Wahrscheinlich technischer Defekt ,sporadisch immer diverse Fehlermeldungen Timeouts verschiedenster Module | SWO-199215 | Defect | HW | HW | Temperature |  | All valid calibrations no longer exist and AQT always requires restart 1999 Undefined lookup String 46.1 , 3264 Environment | System checked | Light barriers cleaned, optical unit cleaned, restart performed, test measurements and QC performed Probably technical defect ,sporadically always various error messages timeouts of various modules
No | Visit | Complaint | RCHN | 测试卡中层、底层温度超范围 | 测试卡中层、底层温度超范围,借903-165一个。 |  | SWO-199236 | Defect | HW | Basic Unit | PCB, Wet section Control |  | Test card mid- and bottom-temperature over range | Test card mid- and bottom-temperature over-range, borrowing 903-165 one.
No | Visit | Complaint | RCHN | 进样口和测试卡舱门无法关闭 | 进样口和测试卡舱门无法关闭 借903-117，903-028各一个。 |  | SWO-199244 | Defect | HW | Analyzer Control Module | PCB, Main Board |  | The sample and test card doors cannot be closed | The sample and test card doors cannot be closed Borrow 903-117, 903-028 each.
No | Visit | Complaint | RDEU | Parameter nicht messbereit, Fragen zu Bestellung und Lieferzeiten, Fragen zu Probennehmern und Preanalytik | IQC-Fehler bei pCO2 nach SC-Tausch | Reinigung Fließwege, Einlassdichtungshalter erneuert, Fluid Transport im Service durchgeführt, QC wiederholt - pCO2 zu hoch, pH in Level B zu tief, SC90 ersetzt, | SWO-199703 | Assisted Replacement | SC | SC | pCO2 | Rillibäk? | Parameters not ready for measurement, questions about ordering and delivery times, questions about samplers and preanalysis | IQC error on pCO2 after SC exchange | Cleaning flow paths, inlet seal holder renewed, fluid transport carried out in service, QC repeated - pCO2 too high, pH in level B too deep, SC90 replaced,
No | Visit | Complaint | RDEU | nach Umsetzung des Gerätes auf einen anderen Tisch im Labor Glu mit Sensorfehler, kein Status | SC sollte erneuert werden | erneut kalibriert ohne Ergebnis, anschließend die SC herausgenommen, gespült und wieder eingesetzt,Kal i.O., QC i.O. (keine Messung die Ursache!) | SWO-199744 | User error | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | after conversion of the device to another table in the laboratory Glu with sensor error, no status | SC should be renewed | recalibrated without result, then the SC removed, rinsed and reinserted,Kal i.O., QC i.O. (no measurement the cause!)
No | Visit | Complaint | RKK | SC交換するとVCエラーやO2のエラーで初期化できない。今回も社内の装置では問題なく初期化できたSCを現地で取り付けるとO2が＃＃＃になってしまう。納品後1年以内でエラー出すぎなので一度しっかり見てほしいと依頼あり。 | 気泡エラーが頻発する | センサーボード故障により現象発生。 センサーボード交換を実施しました。 | SWO-199768 | [5 empty] | SC replacement cannot be initialized with VC error or O2 error. O2 is ###になってしまう when the SC which was able to be initialized without a problem is installed in the locale this time withthe equipment in the company. Because there is an error too much within one year after delivery, there is a request that I want you to see it firmly once. | Frequent bubble errors | A phenomenon occurs due to sensor board failure. The sensor board was replaced.
No | Visit | Complaint | RIND | Cassette is failing continuously. | Checked and found that the zeroing not going properly. | done the zeroing and change the sensor board on standby purpose, need to replace the sensor board. | SWO-201041 | [5 empty] | Cassette is failing continuously. | Checked and found that the zeroing not going properly. | done the zeroing and change the sensor board on standby purpose, need to replace the sensor board.
No | 6 | Y | RAME | Z: Oximetry failing on system cycle. |  | Oximetry failing on system cycle | D2AD250E605D3542 | [5 empty] | Z: Oximetry failing on system cycle. |  | Oximetry failing on system cycle
No | 8 | Y | RPNZ | Blank | 0722, 0374 errors ? on QC results intermittently. Inhomogeneous sample at pH/BG module |  | D2AD3D4F60199792 | Defect | HW | Inlet module | Inlet Gasket |  | Blank | 0722, 0374 errors ? on QC results intermittently. Inhomogeneous sample at pH/BG module
No | 8 | Y | RPNZ | Blank | 0722, 0374 errors ? on QC results intermittently. Inhomogeneous sample at pH/BG module |  | D2AD3D4F60199792 | Defect | HW | Inlet module | Inlet Gasket | 722 | Blank | 0722, 0374 errors ? on QC results intermittently. Inhomogeneous sample at pH/BG module
No | 18 | Y | RPNZ | Blank | pCO2 Drifts and Bilirubin Incorrect |  | D2B218C2384A9FB0 | Defect | HW | Oxi module | Hemolyzer |  | Blank | pCO2 Drifts and Bilirubin Incorrect
No | 10 | Y | RAME | H Customer reported inhomogeneous errors during patient and QC samples. |  | Examine data logs. Check all inlet flow path. | D2AD55D56060C12B | Defect | HW | Inlet module | Nipple kit |  | H Customer reported inhomogeneous errors during patient and QC samples. |  | Examine data logs. Check all inlet flow path.
No | 5 | Y | RCZE | Blank |  |  | D2AE337BD630AE50 | Defect | HW | Gas module | Gas Module |  | Blank
Yes | 13 | Y | RFRA | Blank | investigation suite à problème de mesure de troponine + valeurs de blanc trop hautes |  | D2AD454D6006A5BA | [5 empty] | Blank | investigation following troponin measurement problem - white values too high
No | Visit | Complaint | RDEU | Gerät bootet nicht richtig. ABL bleibt im Offlinemodus hängen, keine AU Verbindung. | Gerät bleibt in Offline Modus wegen defekter Sample Mixer Unit | Sample Mixer Unit erneuert. Einlaufverfahren durchgeführt. Alle Funktionen wieder o.k. | SWO-200205 | Defect | HW | Basic Unit | Sampler Mixer |  | Device does not boot properly. ABL is stuck in offline mode, no AU connection. | Device remains in offline mode due to defective Sample Mixer Unit | Sample Mixer Unit renewed. run-in procedures. All functions again ok.
No | 5 | Y | RAME | Z BATTERY BACKUP FAILURE |  | TECHLINE SUGGESTS POWER SUPPLY | D2AE88CED613E5E2 | Defect | HW | Power Supply | Power Supply | Techline | Z BATTERY BACKUP FAILURE |  | TECHLINE SUGGESTS POWER SUPPLY
No | Visit | Complaint | RDEU | ABL geh nach einer abgebrochenen Messung in die Fehlerbehebung. Manuelles Spülen und Tausch des LP waren ohne Erfolg. | Visual Flow Check, Systemspülung, Überprüfung der Kabelsteckverbindungen. | Einlaßdichtung erneuert. Systemneustart. Kal, QC und Blutmessung iO. | SWO-200122 | Defect | HW | Inlet module | Inlet Gasket |  | ABL goes into troubleshooting after an aborted measurement. Manual rinsing and replacing of the LP was unsuccessful. | Visual Flow Check, system flushing, checking of cable connectors. | Inlet seal renewed. System restart. Kal, QC and blood measurement iO.
Yes | 3 | Y | RPAC | Blank | Training |  | D2AE8567D647CFD8 | Re classified RMED | Training |  |  |  | Blank | Training
No | Visit | Complaint | RBNL | Na het aanbieden van een bloedmonster gaat de inlet niet terug. Ook in servicemenu geen beweging in inlet, alleen ratelend geluid. | Inlet staat wel boven reagenspack maar de gasket is niet op nippel van pack aangesloten, staat er iets boven. In servicemenu geen beweging aan inlethouder, wel ratelend geluid. Zonder inlethouder wel ronddraaien van geleidingskrans. | Na verwijderen inlethouder start instrument wel op met transportfoutemeldingen. Kan na plaatsen inlethouder geheel via servicemenu goed laten functioneren. Instrument opnieuw op laten starten, kalibreren en controles meten, OK. | SWO-199960 | Adjustment | HW | Auto Inlet module | Inlet Position Cal. | LEH | After offering a blood sample, the inlet does not return. Also in service menu no movement in inlet, only rattling sound. | Inlet is above reagent pack but the gasket is not attached to the nipple of pack, something is above it. In Service Menu No movement to the holder, but rattling sound. Without a holder, you can rotate the guide wreath. | After removing the entry-holder start instrument with transport foulesthings. Can function properly after placing the holder in a service menu. Instrument to restart, calibrate and measure controls, OK.
Yes | Visit | Complaint | RBNL | Afgebroken bloedmonsters | Op verzoek van RMED, Service Dump file heb gekopieerd om samen met de Inletmodule naar RMED te verzenden. Need your help, please return the Inlet Module, and service dump to me. ANALYZER_TYPE : ABL90 FLEX PLUS(092R0101N0016) // | SD-file heb gekopieerd voor RMED. | SWO-200528 | Re classified RMED | To collect Auto inlet module |  |  |  | Broken down blood samples | At RMED's request, Service Dump file has been copied to send to RMED along with the Inletmodule. Need your help, please return the Inlet Module, and service dump to me. ANALYZER_TYPE: ABL90 FLEX PLUS (092R0101N0016)// | SD-File copied for RMED.
No | Visit | Complaint | RKK | キャピラリー測定時ガスケットの戻りが遅い | キャピラリー測定時ガスケットの戻りが遅い | ホルダつきインレットガスケットの戻りが鈍い(キャピラリー測定時)。 ホルダ付インレットガスケット、インレットプローブ交換。 インレットモジュール分解点検実施。 別途ご依頼のレポートレイアウト変更。 | SWO-199933 | Adjustment | HW | Auto Inlet module | Cap - Spring |  | Gasket return slowduring capillary measurement | Gasket return slowduring capillary measurement | The return of the inlet gasket with the holder is dull (at capillary measurement). Inlet gasket with holder, inlet probe replacement. Inlet module disassembly inspection was carried out. The report layout change of the request separately.
No | Visit | Complaint | RKK | ブランクキャル失敗を連発 | ブランクキャル失敗を連発 | 再現性なし。 ヘモライザーとピンチバルブ交換実施。 | SWO-199765 | [5 empty] | Blank car failure spree | Blank car failure spree | No reproducibility. Hemolyzer and pinch valve replacement were implemented.
No | Visit | Complaint | RIND | Analyzer not calibrating. | Parameters not calibrating. | Try to reinstall and reset SC and SP one by one. But not calibrated. SC replaced. PER returned for replacement. | SWO-200864 | [5 empty] | Analyzer not calibrating. | Parameters not calibrating. | Try to reinstall and reset SC and SP one by one. But not calibrated. SC replaced. PER returned for replacement.
No | Visit | Complaint | RCHN | CO2定标不过 | CO2定标不过 | 更换测试卡后修复 | SWO-201891 | Assisted Replacement | SC | SC | pCO2 |  | CO2 calibration is not the | CO2 calibration is not the | Repair after replacing test card
No | Visit | Complaint | RCHN | 样本显示温度错误 | 样本显示温度错误，借903-210、636-561和848-037各一个 | 更换903-210、636-561和848-037各一个正常。 by andrew | SWO-199393 | Defect | HW | Process Module | PCB, Cup dryer | 1070 | Sample shows temperature error | Samples show temperature error, with one each in 903-210, 636-561 and 848-037 | Replace the 903-210, 636-561, and 848-037 each normal. by andrew
Yes | 7 | Y | RPNZ | Blank | pH failures, calibration sens and QC low |  | D2AEEFF6D60C97A1 | Re classified RMED | Phone support |  |  |  | Blank | pH failures, calibration sens and QC low
No | Visit | Complaint | RDEU | Parameter O2HB fällt aus dem Ringversuch | Parameter O2HB fällt aus dem Ringversuch | Reinigung, Kal.1, Servicekalibration, Gerät kalibriert und mit QCs überprüft, Ok | SWO-200149 | Adjustment | HW | Oxi module | Oxi module |  | Parameter O2HB falls out of the ring test | Parameter O2HB falls out of the ring test | Cleaning, Kal.1, service calibration, device calibrated and checked with QCs, Ok
No | 18 | Y | RPNZ | Blank | Unstable pH measurements |  | D2AEEEBAD644979F | Revisit required | Electrode | Electrode | pH |  | Blank | Unstable pH measurements
No | Visit | Complaint | RDEU | Alle gültigen Kalibrationen nicht mehr vorhanden und AQT verlangt immer Neustart 1999 undefiniert lookup String 46.1 , 3264 Environment Schwester Nicole gesprochen, beauftragt hat Ute Exter | System überprüft => Main Board defekt , Kalibration ungültig | Main Board erneuert , AQT neu kalibriert , keine QC durch geführt da noch diverse Parameter kalibriert werden müßen, Kunde mißt nach Durchführung die QC Frau Gansau informiert | SWO-200146 | Defect | HW | Analyzer Control Module | PCB, Main Board | 3264 | All valid calibrations no longer exist and AQT always requires restart 1999 Undefined lookup String 46.1 , 3264 Environment Sister Nicole spoke, commissioned Ute Exter | System checked => Main board defective , calibration invalid | Main board renewed, AQT recalibrated, no QC performed as various parameters still need to be calibrated, customer measures after carrying out the QC Informed Ms Gansau
No | 1 | Y | RFRA | Blank | Pas de réponse de l'écran tactile de façon intermittente |  | D2AE8B5DD6181A28 | [5 empty] | Blank | No response from touchscreen intermittently
No | 10 | Y | RAME | 3 pCO2 errors on internal and external QC. Customer has tried three different sensor cassettes. |  | Confirm when on site. | D2AFE38E4C3F4A90 | Defect | HW | Basic Unit | PCB, Sensor interface |  | 3 pCO2 errors on internal and external QC. Customer has tried three different sensor cassettes. |  | Confirm when on site.
No | Visit | Complaint | RDEU | 14.06.2017, 13:40 Uhr ==> Anrufer Schw. Rita, Tel.: 0221890713652 ==> Benutzerintervention erforderlich | Leckage erkannt. | Einlassdichtung aus JW Satz eingebaut. Im Servicemenü Programme mehrfach wiederholt bis alle Fehler beseitigt waren. | SWO-200584 | Defect | HW | Inlet module | Inlet Gasket |  | 14.06.2017, 13:40 ==> Caller Schw. Rita, Tel.: 0221890713652 ==> User intervention required | Leakage detected. | Inlet seal built from JW set. Repeated several times in the service menu programs until all errors have been fixed.
No | 6 | Y | RAME | Z Analyzer dead, no power |  | the analyzer showed no life | D2AF78374C3995E4 | Revisit required | HW | Power Supply | Power Supply |  | Z Analyzer dead, no power |  | the analyzer showed no life
No | Visit | Complaint | RDEU | Zahlenwerte passen nicht zu den Einheiten im Befund | Dateien müssen beim ABL92 angepasst werden | Dateien eingespielt, Gerät neu gestartet | SWO-200482 | Setup | SW | RMED | Parameter setup |  | Numerical values do not match the units in the findings | Files need to be customized on ABL92 | Files played in, device restarted
No | Visit | Complaint | RCHN | PO2,HCT定标失败 | PO2,HCT定标失败，借902-972和903-334各一个 | 更换902-972后修复 by andrew | SWO-199954 | [5 empty] | PO2, HCT calibration failed | PO2, HCT calibration failure, borrowing 902-972 and 903-334 each one | Repair by andrew after replacement 902-972
No | Visit | Complaint | RCHN | 温度错误 | 温度错误，借903-183（被替换成903-581）一个 | 更换903-183备件后修复 by andrew | SWO-199814 | Defect | HW | Basic Unit | PCB, Sensor interface | 1326 | Temperature error | Temperature error, borrow 903-183 (replaced by 903-581) one | Repair by andrew after replacement of 903-183 spare parts
No | Visit | Complaint | RCHN | 测试卡参比电极错误 1361 | 测试卡参比电极错误 1361 | 更换测试卡 | SWO-200190 | Assisted Replacement | SC | SC | Reference |  | Test card reference electrode error 1361 | Test card reference electrode error 1361 | Replace the test card
Yes | 13 | Y | RPAC | Blank | Replace the pH Black Tip seal |  | D2AFD67A4C23E4F5 | [5 empty] | Blank | Replace the pH Black Tip seal
Yes | 12 | Y | RPAC | Blank | Replace the pH Black tip seal |  | D2AFD8D34C02E4FA | [5 empty] | Blank | Replace the pH Black tip seal
No | Visit | Complaint | RDEU | 13.06.2017, 09:25 Uhr ==> Anrufer Fr. Dill, Tel.: 07131492115 ==> Ref.-Elektrode wurde aufgrund Ca-Kalibrierfehler neu membraniert, jetzt besteht Kalibrierfehler pH und Elektrolyte | pH und Elyte bei Kal. außerhalb Statusgrenzen. | Undichte Messkammer gereinigt , Airkanal und Wastekanal. gespült , Ref.Elektrode erneuert. | SWO-200526 | Assisted Replacement | Electrode | Electrode | Reference |  | 13.06.2017, 09:25 ==> Caller Ms. Dill, Tel.: 07131492115 ==> Ref.-Electrode has been re-logged due to Ca-calibration errors, now there is calibration error pH and electrolytes | pH and Elyte at Kal. outside state limits. | Leaky measuring chamber cleaned, air channel and waste channel. rinsed, Ref.electrode renewed.
No | 3 | Y | RAME | 6 Ca sensitivity errors. |  | Sensitivity high above upper limit. Problem occurred while unit was being validated for method comparisons. | D2AFAADB4C3C51E0 | Defect | HW | El/Met Module | Pump tubes |  | 6 Ca sensitivity errors. |  | Sensitivity high above upper limit. Problem occurred while unit was being validated for method comparisons.
No | 1 | Y | RFRA | Blank | Echec de calibration |  | D2AF84DE4C048EED | [5 empty] | Blank | Calibration failure
No | 3 | Y | RFRA | Blank | PAS D'OUVERTURE DE LA BAY UNIT POUR REMPLACER LA SC |  | D2B14070C218C217 | Adjustment | HW | Basic Unit | SC Door |  | Blank | NOT THE BAY UNIT TO REMPLACER THE SC
No | Visit | Complaint | RDEU | Tritt häufig Spülfehler auf Gerät gerade ausgeliefert | nach Auslieferung häufiges Auftreten von Spülfehler Meldung 1340 | Austausch div Verschleißteile Ventil überprüft Überprüfung Fließwege | SWO-200822 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway | New analyzer | Frequently occurs flushing error Device currently delivered | after delivery frequent occurrence of flushing errors Message 1340 | Replacement div consumables Valve checked Review flow paths
No | Visit | Complaint | RCHN | lac定标不过 | lac定标不过,需要借件903-334一个 | 更换903-334及测试卡后修复 by andrew | SWO-199815 | [5 empty] | lac calibration is not | lac calibration, however, need to borrow 903-334 a | Repair by andrew after replacing 903-334 and test card
No | Visit | Complaint | RKK | ファンの音が大きい。クーリングファン交換。 | ファンの音が大きい。 | クーリングファン交換。 | SWO-204314 | [5 empty] | The fan is loud. Cooling fan replacement. | The fan is loud. | Cooling fan replacement.
No | Visit | Complaint | RCHN | PCO2,PH定标不出结果 | PCO2,PH定标不出结果，借903-334一个。 | 更换903-334后修复 by andrew | SWO-200328 | [5 empty] | PCO2, PH calibration results | PCO2, PH calibration results, borrow 903-334 one. | Repair by andrew after replacing 903-334
No | Visit | Complaint | RCHN | 进样口垫片损坏 | 进样口垫片损坏 | 更换924-816后修复 by jesson | SWO-200998 | Defect | HW | Basic Unit | Inlet Gasket |  | Damage to the inlet gasket | Damage to the inlet gasket | Repair by after replacement 924-816 jesson
No | Visit | Complaint | RCHN | 无法读取芯片数据 | 无法读取芯片数据 | 更换测试卡后修复 | SWO-200810 | Assisted Replacement | SC | SC | Chip coding |  | Can't read chip data | Can't read chip data | Repair after replacing test card
No | Visit | Complaint | RCHN | OXI液路传感器故障 0537 | OXI液路传感器故障 0537,借903-125一个。 | 更换903-125备件后修复 by andrew | SWO-200327 | Defect | HW | Oxi module | PCB, Oxi module | 537 | OXI liquid circuit sensor fault 0537 | OXI liquid circuit sensor fault 0537, borrow edgy 903-125 one. | Repair by andrew after replacement of 903-125 spare parts
No | 8 | Y | RPNZ | Blank | Blue Screen. Error code 2 |  | D2B2236F38509FB8 | Defect | SW | Windows | Windows | Error code 2 | Blank | Blue Screen. Error code 2
No | Visit | Complaint | RIND | Sample aborted. | Check the patient result ,found last few samples aborted,check the activity log and found inlet does not filled properly during rinse. | Clean the inlet gasket properly,clean the waste pump tube,flush the inlet probe,run the inlet cali,liquid sensor cal & pump cal,after successful calibration run few sample,got the satisfactory result. | SWO-201264 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | Sample aborted. | Check the patient result ,found last few samples aborted,check the activity log and found inlet does not filled properly during rinse. | Clean the inlet gasket properly,clean the waste pump tube,flush the inlet probe,run the inlet cali,liquid sensor cal & pump cal,after successful calibration run few sample,got the satisfactory result.
Yes | 9 | Y | RAME | 6 CA drifts |  | Checked all data logs for QC issues and found CA falls below the customer set ranges but are within the Radiometer packet insert ranges. Cleaned Reference and Ca electrode connector. | D2B0F225C20395D9 | Re classified RMED | Within RMED Specs |  |  |  | 6 CA drifts |  | Checked all data logs for QC issues and found CA falls below the customer set ranges but are within the Radiometer packet insert ranges. Cleaned Reference and Ca electrode connector.
No | 6 | Y | RAME | F temp error of 270 Hemolyzer temp error |  | looked through the logs and error started on June 8 intermittently 8 to 10 times a day | D2B10E0DC21117E4 | No fault found | HW | Oxi module | Thermistor | 270 | F temp error of 270 Hemolyzer temp error |  | looked through the logs and error started on June 8 intermittently 8 to 10 times a day
No | 4 | Y | RCZE | Blank |  |  | D2B253F0380C154C | Defect | HW | FlexQ Module | FlexQ Cover |  | Blank
No | Visit | Complaint | RDEU | D-Dimer lässt sich nicht Kalibrieren | D-Dimer lässt sich nicht Kalibrieren | Gerät geprüft, Verbrauchsmaterial (LP und Kassetten) ersetzt, Funktionskontrolle ok | SWO-200803 | Defect | Consumables | Test Kit | Cartridge | D-Dimer | D-Dimer cannot be calibrated | D-Dimer cannot be calibrated | Device checked, consumables (LP and cassettes) replaced, function control ok
Yes | 9 | Y | RFRA | Blank | valeurs de blanc trop hautes |  | D2B0DA42C20666D6 | [5 empty] | Blank | white values too high
No | 5 | Y | RFRA | Blank | Cupules de blancs élevées sur les cal |  | D2B0F8CDC2152F39 | Defect | HW | Process Module | Optical Unit |  | Blank | White cups raised on cal
Yes | 4 | Y | RFRA | Blank | Erreur MCC1 et MCC2 hors gamme |  | D2B21C1C38257601 | Re classified RMED | See next upSWO | Test Kit | TnT | Calibration | Blank | MCC1 and MCC2 error off-range
No | 4 | Y | RFRA | Blank | ERREUR DE LQC |  | D2B0F8FBC24152D2 | Defect | HW | Process Module | Optical Unit |  | Blank | LQC ERROR
No | Visit | Complaint | RDEU | Externer Drucker druckt auf mehren Seiten | Externer Drucker druckt auf mehren Seiten - Falsche Schriftgrösse | Fehlersuche : Druckertreiber überprüft, sämtliche Einstellungen weg. Treiber Einstellungen korrigiert, 1x Kopie Treiber gelöscht. Patres Dateien im Odin und Language Ordner kopiert. Einstellungen mit Commit Befehl gespeichert. Startphase abgewartet, Testdrucke gemacht. | SWO-201200 | Setup | SW | External Printer | External Printer |  | External printer prints on several pages | External printer prints on several pages - Incorrect font size | Troubleshooting: Printer driver checks, all settings gone. Driver settings corrected, 1x copy driver deleted. Patres files copied to the Odin and Language folder. Settings saved with commit command. start-up phase waited, test prints made.
No | 8 | Y | RAME | Z Acqure Not Connecting. |  | Verified CPU and wireless usb adapter were malfunctioning. | D2B271B73822CCBC | Defect | HW | DMS_PC | CPU |  | Z Acqure Not Connecting. |  | Verified CPU and wireless usb adapter were malfunctioning.
No | Visit | Complaint | RDEU | 16.06.2017, Herr/Frau Tiator, tel: 05112893838: Intervention erforderlich; der ABL zeigt seit einigen Tagen immer wieder Interventionsmeldungen. Es wurde mehrmals man. gespült, Einlassdichtung, LP und SC wurden getauscht. | K./Na. Empfindlichkeitsfehler , Verstopfung im Hemolyser | Gerät demontiert, Schlauchsystem gespült, SC 90 neu installiert, Funktionskontrolle ok | SWO-201569 | Cleaned | Fluidic pathway | Oxi module | Hemolyzer |  | 16.06.2017, Mr/Mrs Tiator, tel: 05112893838: Intervention required; the ABL has been showing intervention reports over and over again for several days. It became several times man. rinsed, inlet seal, LP and SC were replaced. | K./Na. Sensitivity error , constipation in the Hemolyser | Device dismantled, hose system flushed, SC 90 reinstalled, function control ok
No | 2 | Y | RESP | HAN PUESTO VARIOS SP SALTA INTERVENCIÓN REQUERIDA |  | KIT ANUAL REVISIÓN PREVENTIVA | D2B220873853DCC6 | Defect | HW | PM Kit | PM Kit |  | HAVE PUT SEVERAL SP SKIPS INTERVENTION REQUIRED |  | ANNUAL KIT PREVENTIVE REVIEW
No | Visit | Complaint | RCHN | 定标卡无法扫码 | 定标卡无法扫码，需要借件902-924一个 | 更换902-924后修复 by jesson | SWO-200701 | Defect | HW | Analyzer Control Module | Barcode reader, display |  | The bid card cannot be scanned | The bid card cannot be scanned, need to borrow 902-924 one | Repair by jesson after replacement 902-924
No | Visit | Complaint | RCHN | 报错5766 | 报错5766，借备件903-198一个 | 更换903-198后修复 by andrew | SWO-200727 | Defect | HW | Analyzer Control Module | PCB, Main Board | 5766 | 5766 error | 5766, borrowed spare parts 903-198 one | Repair by andrew after replacement 903-198
No | 2 | Y | RCZE | Blank | Pristroj pise - volejte technika |  | D2B21E27380DAE58 | [5 empty] | Blank | Pristroj Pisa-Call technology
No | Visit | Complaint | RDEU | Oberer LS-Fehler El/ Met-Modul | Oberer LS-Fehler El/ Met-Modul, oft Lac-Probleme berichtet | Aufgef. Verschleißteile ersetzt, dazu Analysator runtergefahren, Flüssigkeit war in Sensor gelaufen, Analysator wieder hochgefahren , LS-Justage im Serviceprogramm i.O., Nachfüllprogramm, Lac-Elektrode nach Radiometer Verfahren gereinigt und mit neuer Lot (3526) membraniertKal, QC | SWO-201208 | Adjustment | HW | El/Met Module | LS |  | Upper LS error El/ Met module | Upper LS error El/ Met module, often Lac problems reported | Aufgef. Replacement of wear parts, with analyser shut down, liquid had run into sensor, analyzer restarted, LS-Adjustment in the service program i.O., refill program, Lac electrode cleaned according to radiometer procedure and membraned with new solder (3526)Kal Qc
No | 3 | Y | RAME | F OXI DRIFT AND SPECTRUM MISMATCH |  | CLEANED AND DECONTAMINATED HEMOLYZER WITH NO RESULTS. | D2B2C7873855B0C6 | Cleaned | HW | Oxi module | Oxi module, Cleaned | Revisit? | F OXI DRIFT AND SPECTRUM MISMATCH |  | CLEANED AND DECONTAMINATED HEMOLYZER WITH NO RESULTS.
No | 5 | Y | RCZE | Blank |  |  | D2B25454384D1550 | Defect | HW | FlexQ Module | FlexQ Cover | LEH Email | Blank
No | Visit | Complaint | RDEU | 16.06.2017, Herr/Frau Tiator, tel: 05112893838: Intervention erforderlich; der ABL zeigt seit einigen Tagen immer wieder Interventionsmeldungen. Es wurde mehrmals man. gespült, Einlassdichtung, LP und SC wurden getauscht. | pCO2 Kal. Backlog-Fehler | Gerät demontiert, Ventileinheit demontiert und gereinigt, Ventilschlauch ersetzt, SC 90 erneuert, Funktionskontrolle ok | SWO-201568 | Defect | HW | PM Kit | PM Kit |  | 16.06.2017, Mr/Mrs Tiator, tel: 05112893838: Intervention required; the ABL has been showing intervention reports over and over again for several days. It became several times man. rinsed, inlet seal, LP and SC were replaced. | pCO2 Kal. Backlog-Fehler | Device dismantled, valve unit disassembled and cleaned, valve hose replaced, SC 90 renewed, function control ok
No | Visit | Complaint | RCHN | 不均匀的定标溶液 | 不均匀的定标溶液 | 更换试剂包后修复 | SWO-201893 | Assisted Replacement | SP | SP | FT |  | Uneven calibration solution | Uneven calibration solution | Repair after replacement of reagent kit
No | Visit | Complaint | RCHN | 定标失败 | 定标失败，借903-364一个 | 更换新机器后修复 by andrew | SWO-200908 | Defect | HW | Process Module | Optical Unit | 1059 | Bid failed | Bid failure, borrow 903-364 one | Fix by andrew after replacing a new machine
Yes | Visit | Complaint | RCHN | 定标失败 | 定标失败,加借944-220（12568）一张 | 更换新机器后修复 by andrew | SWO-200921 | Re classified RMED | Double entry |  |  |  | Bid failed | Failed to bid, plus 944-220 (12568) | Fix by andrew after replacing a new machine
No | Visit | Complaint | RCHN | BNP定标不过 | 借BNP定标卡944-258（13034）一张 |  | SWO-201259 | Assisted Replacement | Consumables | Test Kit | BNP Calibration |  | BNP calibration is not | By BNP Bid Card 944-258 (13034) One
No | Visit | Complaint | RCHN | 探测到漏电流 | 探测到漏电流，902-813 | 更换902-813后修复 by andrew | SWO-200990 | Defect | HW | Inlet module | PCB, Flap Detector |  | Leak current detected | Leak current detected, 902-813 | Repair by andrew after replacement 902-813
Yes | 5 | Y | RAME | P Customer reporting leak out of inlet gasket port for probe. |  | Confirmed reported issue with customer. | D2B31562AE407867 | Re classified RMED | Phone support |  |  |  | P Customer reporting leak out of inlet gasket port for probe. |  | Confirmed reported issue with customer.
No | 32 | Y | RAME | Z Customer reporting SC Refence error - replace SC |  | Confirmed issue. Someone prior to calling for assistance had begun the replacement process but stopped, with cassette bay door open. | D2B483812446F077 | Assisted Replacement | SC | SC | Reference |  | Z Customer reporting SC Refence error - replace SC |  | Confirmed issue. Someone prior to calling for assistance had begun the replacement process but stopped, with cassette bay door open.
No | Visit | Complaint | RCHN | CO2质控失败，GLU，PCO2定标失败1070，0210 | CO2质控失败，GLU，PCO2定标失败1070，0210 | 更换测试卡,试剂包正常。 | SWO-201334 | Assisted Replacement | SC | SC | Glu |  | CO2 QC Failure, GLU, PCO2 Calibration Failure 1070,0210 | CO2 QC Failure, GLU, PCO2 Calibration Failure 1070,0210 | Replace the test card and the reagent package is OK.
No | Visit | Complaint | RCHN | MYO定标失败 | 定标失败，借MYO定标卡944-214（13027）一张 | 重新定标正常。(by tim) | SWO-201265 | Assisted Replacement | Consumables | Test Kit | Myo Calibration |  | MYO calibration failure | Failed to bid, borrow MYO bid card 944-214 (13027) one | Recalibration is normal. (by tim)
No | Visit | Complaint | RCHN | 定标失败 | 定标失败，借BNP定标卡944-258（12301）一张 | 重新定标后修复 by andrew | SWO-200932 | Assisted Replacement | Consumables | Test Kit | BNP Calibration |  | Bid failed | Failed to bid, borrow BNP bid card 944-258 (12301) one | Fix by andrew after re-calibration
No | Visit | Complaint | RDEU | 16.06.2017, Frau Jahnke, Tel.: 03044030700: Intervention erforderlich - LP wechseln. Frau Jahnke berichtet, dass ständig das LP gewechselt werden müsste. | bei Eintreffen war der Analysator Messbereit | Aktivitäten durchgesehen - Problem heute Nacht war, dass ein Ableitstrom im LP erkannt wurde, danach wurde er gewechselt Ableitstrom aktuell überprüft war bei -37mV und damit i.O. Oxi Schlauch, Einkassdichtungen überprüft i.O. $ Serviceprogramme durchlaufen lassen i.O. | SWO-201565 | Assisted Replacement | SP | SP | Leak current |  | 16.06.2017, Mrs. Jahnke, Tel.: 03044030700: Intervention required - change LP. Ms. Jahnke reports that the LP needs to be changed all the time. | on arrival, the analyzer was ready to measure | Activities reviewed - Problem tonight was that a leakage stream was detected in the LP, after which it was changed Discharge current was currently checked at -37mV and thus i.O. Oxi hose, cash seals checked in i.O. • Let service programs go through i.O.
No | Visit | Complaint | RDEU | 16.06.2017, 11:45 Uhr ==> Anrufer Herr Murabagic, Tel.: 023815891214 ==> Eingabe über Touchscreen funktioniert nicht richtig, ca. 1cm versetzt zur Taste 19.06.2017 08:30 Uhr Anruf Herr Debußmann Tel.: 023815891214 Eingabe am Touchscreen nur über Maus möglich | Touchsreen reagiert nicht Flüssigkeitsschaden | Gerät überprüft, Touchsreen erneuert | SWO-202066 | Defect | HW | DMS_PC | Touch screen |  | 16.06.2017, 11:45 a.m. ==> Caller Mr. Murabagic, Tel.: 023815891214 ==> Entering via touchscreen does not work properly, approx. 1cm offset to the button 19.06.2017 08:30 Call Mr. Debußmann Tel.: 023815891214 Entry on the touchscreen only via mouse possible | Touchsreen does not respond Liquid damage | Device checked, touchsreen renewed
Yes | Visit | Complaint | RDEU | 19.06.2017; 10:44 Uhr; H. Wagner; Tel. 062415014815: pH durchgestrichen; Kal und QC Fehler 19.06.2017, 11.30 Uhr, Herr Wagner, Tel.: 062415014815: keine neue SC mehr vor Ort | KAL Fehler pH | Allgemeine Überprüfung - Erneuerung Sensorkassette - Funktionsüberprüfung etc. | SWO-201766 | Re classified RMED | Logistical upSWO |  |  |  | 19.06.2017; 10:44 a.m.; H. Wagner; Tel. 062415014815: pH crossed out; Kal and QC error 19.06.2017, 11.30 a.m., Mr. Wagner, Tel.: 062415014815: no new SC on site | KAL Error pH | General inspection - Renewal sensor cassette - Function check etc.
No | Visit | Complaint | RLTD | Fault phoned in by David Ruaux (RLTD) found when on site for user training. | Immediate issue with QC out of range for oximetry was resolved by running a tHb calibration. On arrival checked fluidic service values - all ok. Remobved oximetry module and physically checked - all ok. Fitted new oxi pinch valve tube as precaution. Re-assembled and tested - all ok. | Replaced Oxi pinch valve tube. No actual fault found. | SWO-201802 | Adjustment | HW | Oxi module | tHb Cal |  | Fault phoned in by David Ruaux (RLTD) found when on site for user training. | Immediate issue with QC out of range for oximetry was resolved by running a tHb calibration. On arrival checked fluidic service values - all ok. Remobved oximetry module and physically checked - all ok. Fitted new oxi pinch valve tube as precaution. Re-assembled and tested - all ok. | Replaced Oxi pinch valve tube. No current fault found.
No | 5 | Y | RPNZ | Blank | 6.71 level 4 range 6.72-6.82. Level 1 6.93 (6.93-7.01) |  | D2B604DB9A4A8A0E | [5 empty] | Blank | 6.71 level 4 range 6.72-6.82. Level 1 6.93 (6.93-7.01)
No | 27 | Y | RAME | 2 PCO2 was high on level 1 external QC. |  | Checked the data logs. | D2B5E0859A1075FA | [5 empty] | 2 PCO2 was high on level 1 external QC. |  | Checked the data logs.
No | 9 | Y | RAME | I temp error 270 hemolyzer temp error |  | error 270 in the manual instructs to replace hemolyzer | D2B5FF529A5C2FCB | Defect | HW | Oxi module | Hemolyzer | 270 | I temp error 270 hemolyzer temp error |  | error 270 in the manual instructs to replace hemolyzer
Yes | Visit | Complaint | RDAN | Godmorgen ABL800 nr 2 her på KBA kommer med flg. meddelelse: 0600: Demoperiode udløber snart Med venlig hilsen Pernille Pflug-Christensen Kemi-Immunkemi og POCT specialist Mobil: 21380995 pernille.pflug-christensen@regionh.dk Bispebjerg og Frederiksberg Hospital Klinisk Biokemisk Afdeling Bispebjerg Bakke 23, indgang 6 2400 København NV | Demonøgler fra ikke slået fra. | Demonøgler er nu slået fra og fejl 0600 Demoperiode udløber snart skulle gerne være væk i morgen. | SWO-201690 | Re classified RMED |  |  |  |  | Good morning ABL800 nr 2 Here at KBA comes with FLG. Message: 0600: Demo period expires soon Sincerely, Pernille Pflug-Christensen Chemistry-immune Chemistry and POCT specialist Mobile: 21380995 pernille.pflug-christensen@regionh.dk Bispebjerg and Frederiksberg Hospital Clinical Biochemical Department Bispebjerg Bakke 23, interior entrance 6 2400 Copenhagen NV | Demo keys from not muted. | Demo keys are now turned off and Error 0600 Demo period expires soon should be like to be gone tomorrow.
No | 2 | Y | RAME | F Out of box failure for all oximetry on Cal Ver level III. |  | All Cals and QC passed but Cal Ver level III would not pass, would get oxi spectrum mismatch error code 0581, 0688, 0094. Suspect bad Hemolyzer. | D2B64A459A50D7A2 | Defect | HW | Oxi module | Hemolyzer | Spectrum mismatch | F Out of box failure for all oximetry on Cal Ver level III. |  | All Cals and QC passed but Cal Ver level III would not pass, would get oxi spectrum mismatch error code 0581, 0688, 0094. Suspect bad Hemolyzer.
No | 4 | Y | RAME | F Out of box failure for all oximetry on Cal Ver level III. |  | All Cals and QC passed but Cal Ver level III would not pass, would get oxi spectrum mismatch error code 0581, 0688, 0094. Suspect bad Hemolyzer. | D2B739FA1049ED55 | Revisit required | HW | Oxi module | Hemolyzer |  | F Out of box failure for all oximetry on Cal Ver level III. |  | All Cals and QC passed but Cal Ver level III would not pass, would get oxi spectrum mismatch error code 0581, 0688, 0094. Suspect bad Hemolyzer.
No | 15 | Y | RESP | Error 3414 FTC 105. |  | Comprobar bombas y válvulas rellenar los tubos de paso del reactivo, calibrar y probar con un blanco. Revisión del equipo y puesta en marcha despues de estar un largo tiempo parado. | D2B5EF739A08F999 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway | 3414 | Error 3414 FTC 105. |  | Check pumps and valves fill the reagent passage tubes, calibrate and test with a white. Equipment review and commissioning after being stopped for a long time.
No | 9 | Y | RFRA | Blank | 3412: FTC 103 |  | D2B651039A5E9A99 | Defect | HW | Aspiration Module | Needle Wash Unit | 3412 | Blank | 3412: FTC 103
No | 5 | Y | RFRA | Blank | erreur 3412 |  | D2B62CE79A149D7E | Defect | HW | Aspiration Module | Needle Wash Unit | 3412 | Blank | error 3412
No | Visit | Complaint | RDEU | 09.06.2017, Frau Eschmann, Tel.: 020885083200: ABL meldet SC Tausch; die abendlich Kontrolle sei ständig fehlerhaft. 16.06.2017 8:15 Fr. Eschmann 020885083200: QC-C (20:00 Uhr) pO2 Wert nach wie vor zu hoch 20.06.2017 9:30 Fr. Eschmann weiterhin Probleme mit pO2 trotz Erneuerung | pO2 wird im Level C etwas zu hoch gemessen, Das LP wurde am 05.06.2017 ersetzt, seit dem besteht das Problem, LP def. | Fließwege gespült, neues LP eingelegt, Gerät kalibriert und QC durchgeführt, | SWO-202226 | Assisted Replacement | SP | SP | pO2 QC C High |  | 09.06.2017, Mrs. Eschmann, Tel.: 020885083200: ABL reports SC Tausch; the evening check is constantly flawed. 16.06.2017 8:15 Ms. Eschmann 020885083200: QC-C (20:00) pO2 Value still too high 20.06.2017 9:30 Ms. Eschmann continues to have problems with pO2 despite renewal | pO2 is measured a little too high in level C, the LP was replaced on 05.06.2017, since then the problem exists, LP def. | Flow paths rinsed, new LP inserted, device calibrated and QC performed,
No | 7 | Y | RAME | 2 pCO2 & lactate calibration & QC errors. |  | pCO2 calibration was in but QC results were out of range. Lactate calibration sensitivity was low and controls out of range. Customer has changed several SC's & SP's but each time the failed parameters change for different reasons. | D2B6FF85105CA17F | Assisted Replacement | SC | SC | Lac | and pCO2 | 2 pCO2 & lactate calibration & QC errors. |  | pCO2 calibration was in but QC results were out of range. Lactate calibration sensitivity was low and controls out of range. Customer has changed several SC's & SP's but each time the failed parameters change for different reasons.
Yes | Visit | Complaint | RCHN | CL N/E没有终点，温度-F | CL N/E没有终点，温度-F | 连续更换2张测试卡后修复 | SWO-202075 | [5 empty] | CL N/E has no end point, temperature-F | CL N/E has no end point, temperature-F | Repair after 2 consecutive test cards
No | Visit | Complaint | RIND | Instrument rebooting very often. | Instrument rebooting very often. | Instrument was rebooting, check the CF care it was OK, need to replace with new CF card and check awaiting for Spares. | SWO-202541 | [5 empty] | Instrument rebooting very often. | Instrument rebooting very often. | Instrument was rebooting, check the CF care it was OK, need to replace with new CF card and check awaiting for Spares.
No | Visit | Complaint | RCHN | CA定标失败 | CA定标失败，借902-972,903-334各一个 |  | SWO-201880 | [5 empty] | CA calibration failure | CA calibration failure, borrow 902-972, 903-334 each one
No | Visit | Complaint | RIND | pH pCO2 and Lac Calibration error. | pH pCO2 and Lac Calibration error. | Sharing. Check the instrument found blood got clotted inside the Sensor cassette, flushed the SC with water reinstalled the SC found SC was working instrument is working satisfactorily. | SWO-202561 | [5 empty] | pH pCO2 and Lac Calibration error. | pH pCO2 and Lac Calibration error. | Sharing. Check the instrument found blood got clotted inside the Sensor cassette, flushed the SC with water reinstalled the SC found SC was working instrument is working satisfactorily.
No | 4 | Y | RPNZ | Blank | pH QC failures below range, level 1 and 4. |  | D2B789DD104F8A18 | [5 empty] | Blank | pH QC failures below range, level 1 and 4.
No | Visit | Complaint | RIND | SC Temp B, Not Initializing Msg coming. | Sensor Board Replace. | Sensor Board Replaced. SC and SP installed. Zeroing calibration done. QC check and found not passing. SP replaced but same problem found. SC replaced and found some parameters not calibrating. Need to replace both SC and SP. Analyser Serviced and Calibrated properly. Handed over in a fine working condition. | SWO-203337 | [5 empty] | SC Temp B, Not Initializing Msg coming. | Sensor Board Replace. | Sensor Board Replaced. SC and SP installed. Zeroing calibration done. QC check and found not passing. SP replaced but same problem found. SC replaced and found some parameters not calibrating. Need to replace both SC and SP. Analyser Serviced and Calibrated properly. Handed over in a fine working condition.
No | Visit | Complaint | RCHN | LAC定标失败 | LAC定标失败，借903-183一个。 | 更换903-183后修复(by tim) | SWO-201419 | Defect | HW | Basic Unit | PCB, Sensor interface |  | LAC calibration failure | LAC calibration failed, borrowing 903-183 one. | Repair after replacement 903-183 (by tim)
No | 7 | Y | RPNZ | Blank | Oxi Sample Problems |  | D2BB02517212A0EE | Defect | HW | Oxi module | Hemolyzer |  | Blank | Oxygen Sample Problems
No | 7 | Y | RAME | Z Customer changing Gas1 every 7-10 days |  | checked replacement logs, Gas1 changed on 5 $29, 6 $5, 6 $12, 6 $19, | D2B7BD3610248F03 | Revisit required | HW | Gas module | Gas valve |  | Z Customer changing Gas1 every 7-10 days |  | checked replacement logs, Gas1 changed on 5 $29, 6 $5, 6 $12, 6 $19,
No | Visit | Complaint | RDAN | OXI lampe intensitet for høj | OXI test | Lampe justeret tHb kal. udført | SWO-202229 | Adjustment | HW | Oxi module | Lamp unit | Halogen | OXI Lamp intensity too high | OXYGEN Test | Lampe justeret tHb kal. udført
No | 4 | Y | RAME | F Out of box failure for all oximetry on Cal Ver level III |  | All Cals and QC passed but Cal Ver level III would not pass, would get oxi spectrum mismatch error code 0581, 0688, 0094. Suspect bad Hemolyzer. | D2B73D45100549DC | Defect | HW | Oxi module | Hemolyzer |  | F Out of box failure for all oximetry on Cal Ver level III |  | All Cals and QC passed but Cal Ver level III would not pass, would get oxi spectrum mismatch error code 0581, 0688, 0094. Suspect bad Hemolyzer.
No | Visit | Complaint | RDEU | Trotz Time Server Anbindung stimmt Uhrzeit nicht, geht 1h nach |  | Einstellung Timeserver deaktiviert | SWO-202413 | Setup | SW | Time setting | Time setting |  | Despite Time Server connection, time is not right, 1h goes up |  | Setting Timeserver disabled
No | Visit | Complaint | RSAF | Customer phoned and complained about a burning smell coming from the analyzer. | Got to Nicu and investigated where the burning smell was coming from and located it coming from the UPS system. | Installed new UPS and removed the faulty UPS | SWO-204074 | No fault found | HW | UPS | UPS |  | Customer phoned and complained about a burning smell coming from the analyzer. | Got to Nicu and investigated where the burning smell was coming from and located it coming from the UPS system. | Installed new UPS and removed the faulty UPS
No | Visit | Complaint | RDAN | Anne Marie melder, at ABL'en måler for højt på QC C. Den beder selv om en ny væskepakke. Den har fået to uden, at det virker og den beder stadigvæk om en ny væskepakke. | Afprøvet. | Sensorcassette udskiftet, startet op og slutafprøvet. | SWO-202621 | Assisted Replacement | SC | SC | Lac |  | Anne Marie reports that the ABL is measuring too high on QC C. It asks for a new liquid package. It has got two without it working and it still asks for a new liquid package. | Tested. | Sensorcassette replaced, booted and finished tested.
No | Visit | Complaint | RDEU | Frau Steinbach 0651/68342215 pO2 QC Fehler Lvl.2 19.06.2017 15:25 Uhr Frau Steinbach Tel.: 065168342215 -> Kundin hatSC ersetzt, Fehler bleibt 20.06.2017 13:10Uhr Frau Schmitt 065168342215: Auch nach LP und SC-Tausch weiterhin QC-Verletzung | ABL hat den QC Fehler seit dem der LS-Pack am 18.06 ersetzt wurde. Einen Ls-Pack mit anderem LOT eingesetzt. | Startphase QC und Kalibration. | SWO-202481 | Assisted Replacement | SP | SP | pO2 |  | Ms Steinbach 0651/68342215 pO2 QC Error Lvl.2 19.06.2017 15:25 Ms Steinbach Tel.: 065168342215 -> Customer has replacedSC, error remains 20.06.2017 13:10pm Ms Schmitt 065168342215: Even after LP and SC exchange still QC violation | ABL has had the QC error since the LS-Pack was replaced on 18.06. Use an Ls-Pack with other LOTS. | Start phase QC and calibration.
No | Visit | Complaint | RCHN | LAC定标失败，漂移超范围 | LAC定标失败，漂移超范围 | 更换测试卡 | SWO-203542 | Assisted Replacement | SC | SC | Lac |  | LAC calibration failure, drift over range | LAC calibration failure, drift over range | Replace the test card
No | Visit | Complaint | RCHN | 测试卡无法扫描 | 测试卡无法扫描，借备件910-304一个 |  | SWO-202339 | Defect | HW | Storage Module | Barcode reader, Storage wheel | BR001 | Test card can't be scanned | Test card can not be scanned, borrowed spare parts 910-304 one
No | Visit | Complaint | RCHN | 报错3412 | 报错3412,需要借件903-329一个 | 更换903-329备件一个(by tim) | SWO-201918 | Defect | HW | Aspiration Module | Needle Wash Unit | 3412 | Error 3412 | 3412, need to borrow 903-329 one | Replace 903-329 spare parts one (by tim)
No | Visit | Complaint | RCHN | 测试卡卡盘 | 测试卡卡盘 | 取出测试卡后清理修复(by tim) | SWO-202449 | Defect | Consumables | Test Kit | Cartridge | Broken | Testing the kaka plate | Testing the kaka plate | Clean up the repair after removing the test card (by tim)
No | Visit | Complaint | RCHN | 报错1204，系统检测失败 | 报错1204，系统检测失败 | 更换903-364后修复 | SWO-202445 | Defect | HW | Process Module | Optical Unit | 1204 | Error 1204, system detection failure | Error 1204, system detection failure | Repair after replacement 903-364
No | 3 | Y | RAME | F 0376 Calibration drift 1 out of range. |  | Drift in the range of 20 out. Tried complete oxi test and lamp unit adjust. Problem goes away and comes back. | D2B8645E8644B4D3 | Defect | HW | Oxi module | PCB, Oxi module |  | F 0376 Calibration drift 1 out of range. |  | Drift in the range of 20 out. Tried complete oxi test and lamp unit adjust. Problem goes away and comes back.
No | Visit | Complaint | RDEU | 21.06.2017, 08:45 Uhr, Fr. Rathmer, Tel. 025639126985: Meldung: Probenmessung abgebrochen | Fehler 1067 Alle Proben werden abgebrochen | Gerät überprüft, keine weiteren Fehlermeldungen vorhanden, Daten gelöscht, Filter gereinigt, Probenmessungen o.k. | SWO-202918 | No fault found | No fault found | No fault found | No fault found | 1067 | 21.06.2017, 08:45, Ms. Rathmer, Tel. 025639126985: Message: Sample measurement aborted | Error 1067 All samples are cancelled | Device checked, no further error messages available, data deleted, filter cleaned, sample measurements ok.
No | Visit | Complaint | RDEU | Kassetten werden nicht angenommen. | Kassetten werden nicht angenommen. | Mangelhaft justierten Barcodeleser ausgerichtet. | SWO-202678 | Adjustment | HW | Storage Module | Barcode reader, Storage wheel |  | Cassettes are not accepted. | Cassettes are not accepted. | Poorly adjusted barcode reader aligned.
No | Visit | Complaint | RDEU | Lekage Fehler | ABL neu gestartet und LS Pack neu installiert, alle Programme im Service Menü durchgeführt>>I.O | Startphase QC und Kalibration. | SWO-203031 | Assisted Replacement | SP | SP | Leakage |  | Lekage Error | ABL restarted and LS Pack reinstalled, all programs performed in the service menu>>I.O | Start phase QC and calibration.
No | Visit | Complaint | RBNL | Analyzer geeft lekgedetecteerd aan. Al doorgespoten, inletgasket vervangen, nieuwe reagenspack en sensorcassette geplaatst maar de fout blijft. | Door gebruiker al doorgespoten met demiwater inletgasket, reagenspack en sensor vervangen. Lijkt eerst niet te willen pompen, ook niet uit bakje demiwater, na doorspuiten wel. Vacuumcheck te hoog, 0,92. Na vervangen alle tubings blijft vacuum te hoog. Bij controleren actuator-rod ook vacuum te hoog. Verwisselen van de complete inlethouder en reagenspack blijft het vacuum te hoog. Zie wel dat bij verwijderen sensorcassette het onderste deel van de bayunit nat is. Andere sensor cassette, die waarschijnlijk enkele dagen op tafel heeft gestaat, geplaatst, status gevoeligheid fouten GLU en LAC. | Sensor cassette vervangen. | SWO-202937 | Assisted Replacement | SC | SC | Clot |  | Analyzer indicates leak detected. Already sprayed, Inletgasket replaced, new reagent pack and sensor cassette inserted but the error remains. | Replaced by user already sprayed with demi water inletgasket, reagent pack and sensor. Seems to first not want to pump, even out of Dish demi water, after by spraying though. Vacuumcheck too high, 0.92. After replacing all tubings, vacuum remains too high. When checking actuator-Rod also vacuum too high. Changing the complete holder and reagent pack keeps the vacuum too high. Do see that when removing sensor cassette the lower part of the bayunit is wet. Other sensor cassette, which probably has been placed on the table for several days, status sensitivity errors GLU and LAC. | Replace the Sensor cassette.
Yes | Visit | Complaint | RLTD | Contact Jenny 01276 526543 | Lactate failing calibration - sense is ... Or low at 61%, but then .... Again. Need to check contact pins. | Service visit required. Passed toColin Maunder. | SWO-202836 | Re classified RMED | Phone support |  |  |  | Contact Jenny 01276 526543 | Lactate failing calibration - sense is ... Or low at 61%, but then .... Again. Need to check contact pins. | Service visit required. Passed toColin Maunder.
No | Visit | Complaint | RLTD | Contact Jenny 01276 526543 | Checked data logs for fault history. The customer has already replaced the sensor cassette to try to resolve the issue, but Lactate is still failing calibration. Checked sensor readings in service programs. | Cleaned gasket and flushed cassette. Cleaned sensor cassette contact pins. Removed solution pack and checked it was activated correctly. Clicks could be heard when pressure was applied to the top. Replaced sensor cassette. Analyser is now passing calibrations. I will call POCT 23 $06 am to see if QC is now passing. Both cassettes have been removed and will claimed for. | SWO-203438 | Assisted Replacement | SC | SC | Lac |  | Contact Jenny 01276 526543 | Checked data logs for fault history. The customer has already replaced the sensor cassette to try to resolve the issue, but Lactate is still failing calibration. Checked sensor readings in service programs. | Cleaned gasket and flushed cassette. Cleaned sensor cassette contact pins. Removed solution pack and checked it was activated correctly. Clicks could be heard when pressure was applied to the top. Replaced sensor cassette. Analyser is now passing calibrations. I will call POCT 23 $06 am to see if QC is now passing. Both cassettes have been removed and will claimed for.
No | Visit | Complaint | RCHN | 接口板螺丝打滑 | 接口板螺丝打滑，需要借件886-674备件一个 | 备件接错，现需要重新借备件924-738，故此单先关单(by tim) | SWO-202853 | [5 empty] | Interface board screw slips | Interface board screw slip, need to borrow 886-674 spare parts one | Spare parts are wronganded and need to be re-borrowed for spare parts 924-738, so the single first clearance (by tim)
No | Visit | Complaint | RCHN | 测试卡断裂 | 测试卡断裂 | 更换测试卡后修复 | SWO-202868 | Defect | Consumables | Test Kit | Cartridge | Broken | Test card break | Test card break | Repair after replacing test card
No | Visit | Complaint | RCHN | hct值不能备份 | hct值不能备份,需要借件903-329 | 更换903-329后修复(by tim) | SWO-202451 | Defect | HW | Aspiration Module | Needle Wash Unit |  | hct value cannot be backed up | Hct value cannot be backed up, need s903-329 | Repair after replacement 903-329 (by tim)
No | 7 | Y | RPAC | Blank | Touch screen intermittently not working that well |  | D2B9A65CFC08F92C | Adjustment | HW | DMS_PC | Touch screen |  | Blank | Touch screen intermittently not working that well
No | 11 | Y | RAME | Z Customer having intermittent errors on analyzer. Screen going blank and roma.ini not responding. |  | Remotely connected to analyzer and found that the Roma.ini file and AU board stop communicating or not responding. Replaced the roma.ini. and still happening. Reloaded SW 6.16v and roma.ini up and running but customer went to run sample and touch screen is not working results not crossing over. Arrived onsite and found that the touch screen cable was loose and not properly attached. Found that the USB port was broken and TCP $IP was not responding. Results not crossing over to cerner. | D2B9F7DCFC01E971 | Defect | HW | DMS_PC | CPU |  | Z Customer having intermittent errors on analyzer. Screen going blank and roma.ini not responding. |  | Remotely connected to analyzer and found that the Roma.ini file and AU board stop communicating or not responding. Replaced the roma.ini. and still happening. Reloaded SW 6.16v and roma.ini up and running but customer went to run sample and touch screen is not working results not crossing over. Arrived onsite and found that the touch screen cable was loose and not properly attached. Found that the USB port was broken and TCP $IP was not responding. Results not crossing over to cerner.
No | 4 | Y | RAME | 1 Customer reporting random PH unstable errors while running samples. |  | Calibrations all ok. No Ph errors reported. Nothing physical showing. | D2B9C12AFC5EF14E | Defect | HW | pH/BG Module | Valve, air |  | 1 Customer reporting random PH unstable errors while running samples. |  | Calibrations all ok. No Ph errors reported. Nothing physical showing.
No | Visit | Complaint | RDAN | Neon lamp intensity &number too low | OXI test | OXI modul skiftet ud med et fra de nye ABL'ere Hjemtaget til reparation | SWO-203384 | Defect | HW | Oxi module | Oxi module, Replaced | Neon lamp | Neon lamp intensity &number too low | OXYGEN Test | OXI module replaced with one from the new Abls Home taken for repair
No | Visit | Complaint | RLTD | Thaira Tel 01372 735735 ext (6105) | Replaced faulty touch screen. Tested operation of new touch screen - all ok. Verified analyser operation -all ok. | Replaced touch screen. | SWO-203457 | Defect | HW | DMS_PC | Touch screen |  | Thaira Tel 01372 735735 ext (6105) | Replaced faulty touch screen. Tested operation of new touch screen - all ok. Verified analyser operation -all ok. | Replaced touch screen.
No | Visit | Complaint | RLTD | Simon Donnelly Tel 0203 2994446 Bleep 411 | Analyser had been found by technician with the inlet stuck open and in the replacement guide for inlet gasket. After replacing the gasket there were multiple liquid transport errors, Inhomogeneous solutions, pressure test errors etc. Customer continued through guided troubleshooting and replaced solution pack, flushed the sample pathway - still faulty. Then at 16:28 just prior to my arrival test again was tried - analyser started up ok. Checked fluid transport service program - all values well within range. However, pCO2 failing calibration with 1070 response error when QC level A is run, ok on other QCs. Swapped Solution Pack - no change. Replaced Sensor Cassette - waited for start-up to complete and pCO2 still failing just when QC level A is run. As no further Sensor Cassettes available advised staff on unit and left it to stabilize. May need a second new Sensor Cassette to resolve pCO2 QC/Calibration error. | Fluid transport errors - resolved by customer actions, but unknown which action taken was the solution. Subsequent pCO2 errors - possibly needs new Sensor Cassette of a different lot. | SWO-203484 | No fault found | SC | SC | pCO2 |  | Simon Donnelly Tel 0203 2994446 Bleep 411 | Analyser had been found by technician with the inlet stuck open and in the replacement guide for inlet gasket. After replacing the gasket there were multiple liquid transport errors, Inhomogeneous solutions, pressure test errors etc. Customer continued through guided troubleshooting and replaced solution pack, flushed the sample pathway - still faulty. Then at 16:28 just prior to my arrival test again was tried - analyser started up ok. Checked fluid transport service program - all values well within range. However, pCO2 failing calibration with 1070 response error when QC level A is run, ok on other QCs. Swapped Solution Pack - no change. Replaced Sensor Cassette - waited for start-up to complete and pCO2 still failing just when QC level A is run. As no further Sensor Cassettes available advised staff on unit and left it to stabilize. May need a second new Sensor Cassette to resolve pCO2 QC/Calibration error. | Fluid transport errors - resolved by customer actions, but unknown which action taken was the solution. Subsequent pCO2 errors - possibly needs new Sensor Cassette of a different lot.
No | Visit | Complaint | RLTD | Stephanie 07863330762 | Checked display, removed, cleaned and replaced. There is fluid trapped in between the panels that I am unable to remove. The issue is on the bottom left of the panel and is making the touch screen panel unresponsive. | Swapped phase 0 Analyser display unit with faulty ED Analyser. Touch screen panel needs to be replaced, due to liquid damage. Analyser has passed calibration and QC. | SWO-203635 | Revisit required | HW | DMS_PC | Touch screen |  | Stephanie 07863330762 | Checked display, removed, cleaned and replaced. There is fluid trapped in between the panels that I am unable to remove. The issue is on the bottom left of the panel and is making the touch screen panel unresponsive. | Swapped phase 0 Analyser display unit with faulty ED Analyser. Touch screen panel needs to be replaced, due to liquid damage. Analyser has passed calibration and QC.
No | 1 | Y | RPAC | Blank | Multiple temperature errors |  | D2BB16BF72086C9F | Defect | HW | Basic Unit | PCB, Sensor interface |  | Blank | Multiple temperature errors
Yes | Visit | Complaint | RCHN | 报错1024 | 报错1024，需要借件903-364备件一个 | 更换903-364后修复(by tim) | SWO-202446 | Re classified RMED | OOB |  |  | 1024 | 1024 error | 1024, need to borrow 903-364 spare parts one | Repair after replacement 903-364 (by tim)
No | Visit | Complaint | RCHN | BNP定标失败，操作老师操作失误导致定标数据失活 | BNP定标失败，操作老师操作失误导致定标数据失活，借944-258*12749一张 | 重新定标正常。(by tim) | SWO-202851 | User error | Consumables | Test Kit | BNP Calibration |  | BNP calibration failure, operation teacher error led to calibration data inactivation | BNP calibration failure, operation teacher error led to the calibration data inactivation, borrow 944-258 x 12749 a | Recalibration is normal. (by tim)
No | Visit | Complaint | RLTD | Zoe Barnfield Tel. 01707 666 366 Ask for her in ER | Bar code scanner not scanning ampoules . | Removed auto check module cleaned dust from unit. Refitted and test ok. | SWO-203494 | Cleaned | HW | Autocheck module | Barcode reader, Autocheck Module |  | Zoe Barnfield Tel. 01707 666 366 Ask for her in ER | Bar code scanner not scanning ampoules . | Removed auto check module cleaned dust from unit. Refitted and test ok.
No | 12 | Y | RPNZ | Blank | Sensor cassette replacement. |  | D2BE8784D42A1338 | Assisted Replacement | SC | SC | pCO2 |  | Blank | Sensor cassette replacement.
No | Visit | Complaint | RCHN | 接口板螺丝打滑 | 接口板螺丝打滑，加借备件924-738一个 | 更换924-738一个正常。(by tim) | SWO-203144 | [5 empty] | Interface board screw slips | Interface board screw slip, plus spare parts 924-738 one | Replace 924-738 one normal. (by tim)
No | Visit | Complaint | RCHN | 废液探测器错误，血氧打问号0581 | 废液探测器错误，血氧打问号0581，借902-496，902-833各一个。 |  | SWO-203518 | Defect | HW | Basic Unit | Waste detector |  | Waste detector error, blood oxygen test question mark 0581 | Waste detector error, blood oxygen question mark 0581, borrow 902-496, 902-833 each.
No | Visit | Complaint | RCHN | 分析仪未连接0734，1223，1065，1069 | 分析仪未连接0734，1223，1065，1069，借903-165一个。 |  | SWO-203540 | Defect | HW | Basic Unit | PCB, Wet section Control |  | Analyzer not connected to 0734,1223,1065,1069 | The analyzer is not connected to 0734, 1223, 1065, 1069, and borrows 903-165.
No | Visit | Complaint | RCHN | 光学强度超出范围 | 光学强度超出范围,需要借件903-364备件一个，需要借件五张定标卡 944-220 批次号：12806、944-212 批次号：12757、944-216 批次号：12758、 944-214 批次号 13027（这个四张没货），944-258 批次号：13034 （有货） | 更换903-364，重新定标正常。(by tim) | SWO-203248 | Defect | HW | Process Module | Optical Unit | 1031 | Optical strength out of range | Optical strength beyond range, need to borrow a 903-364 spare parts one, need to borrow five bid cards 944-220 batch number: 12806, 944-212 batch number: 12757, 944-216 batch number: 12758, 944-214 batch number 13027 (this four no items), 944-258 Batch No.: 13034 (Available) | Replace the 903-364 and re-mark normally. (by tim)
Yes | Visit | Complaint | RDEU | Kunde beachtete schon länger, dass thb QC Werte sich in allen 3 Leveln an der oberen Grenzenbefinden | bei eintreffen thb im Level 3 fehlerhaft, andere QC Werte liegen nur 0,1 mmol/L unter oberem Grenzwert thb Kal fällig, keine Kalibrator vorrätig LP wurde mehrfach ersetzt | Thb Kal durchgeführt, Werte im Level 3 0,1 unter Oberem Grenzwerte, in Level 2 0,2 mmol $L kpl. Oxy Test durchgeführt i.O. Hemolyzer Unit bestellt -> Info an Kunden | SWO-203859 | Re classified RMED | HW | Oxi module | Hemolyzer |  | Customer has long noticed that thb QC values are at the upper limit in all 3 levels | if thb arrives in level 3 faulty, other QC values are only 0.1 mmol/L below upper limit value thb Kal due, no calibrator in stock LP has been replaced several times | Thb Kal performed, values in level 3 0.1 below upper limit values, in level 2 0.2 mmol $L Compl. Oxy test performed in the first place Hemolyzer Unit ordered -> Info to customers
No | 4 | Y | RAME | Z user called stating that screen is black when powering on. |  | found that the analyzer will not power on the screen at all. | D2BEE006D42B0FEC | [5 empty] | Z user called stating that screen is black when powering on. |  | found that the analyzer will not power on the screen at all.
Yes | 9 | Y | RPAC | Blank | Analyser having fluidic issues. |  | D2C17F55C03B4976 | [5 empty] | Blank | Analyser having fluidic issues.
No | Visit | Complaint | RDEU | 26.06.2017, 10:08 Uhr, Fr. Dreyer, Tel. 051181152811: Ableitstrom im Anschluss LP erkannt | Ventilschlauch undicht | Schlauchsystem geprüft, Schlauch ausgetauscht siehe SWO Jahreswartung | SWO-204395 | Defect | HW | Oxi module | Pinch valve tube |  | 26.06.2017, 10:08 a.m., Ms. Dreyer, tel. 051181152811: Discharge current detected in the following LP | Valve hose leaks | Hose system tested, hose replaced see SWO annual maintenance
No | Visit | Complaint | RLTD | Stephanie 07863330762 | Removed faulty touch screen panel | Replaced touch screen panel on the phase 0 Analyser and then swapped back to original displays. Both Analysers are ready for use. | SWO-204501 | User error | HW | DMS_PC | Touch screen |  | Stephanie 07863330762 | Removed faulty touch screen panel | Replaced touch screen panel on the phase 0 Analyser and then swapped back to original displays. Both Analysers are ready for use.
Yes | 1 | Y | RPAC | Blank | pCO2 sesnor cassette failure |  | D2BF0EF3D407053D | [5 empty] | Blank | pCO2 sesnor cassette failure
No | 3 | Y | RFRA | Blank | ERREURS CAL PO2/PCO2 |  | D2BF02A0D45C77AF | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | Blank | ERRORCAL PO2/PCO2
No | Visit | Complaint | RCHN | 测试卡漏液导致接口板腐蚀 | 测试卡漏液导致接口板腐蚀，借902-972一个 |  | SWO-203765 | [5 empty] | Test card leak causes interface plate corrosion | Test card leakage causes interface plate corrosion, borrows 902-972 one
No | Visit | Complaint | RCHN | 定标失败 | 定标失败，借902-778一个。 |  | SWO-203710 | [5 empty] | Bid failed | The bid failed, borrowing 902-778 one.
No | Visit | Complaint | RCHN | 定标失败 | 定标失败，借902-778一个。 |  | SWO-203713 | [5 empty] | Bid failed | The bid failed, borrowing 902-778 one.
No | Visit | Complaint | RCHN | 定标失败 | 定标失败，借902-778一个。 |  | SWO-203709 | [5 empty] | Bid failed | The bid failed, borrowing 902-778 one.
No | Visit | Complaint | RCHN | 测试卡下层温度超范围，系统温度超范围，无法读取或者写入传感器芯片 1066，1221，1123 | 测试卡下层温度超范围，系统温度超范围，无法读取或者写入传感器芯片 1066，1221，1123 | 更换三张测试卡正常。 | SWO-203966 | Assisted Replacement | SC | SC | Thermistor |  | The test card underlet temperature is over range, the system temperature is over range, cannot read or write to the sensor chip 1066,1221,1123 | The test card underlet temperature is over range, the system temperature is over range, cannot read or write to the sensor chip 1066,1221,1123 | Replacing the three test cards is OK.
Yes | Visit | Complaint | RCHN | GLU,LAC定标不过 | GLU,LAC定标不过 | 连续更换2张测试卡后修复 | SWO-203794 | [5 empty] | GLU, LAC calibration but | GLU, LAC calibration but | Repair after 2 consecutive test cards
No | Visit | Complaint | RCHN | 报1999.61.22错误 | 报1999.61.22错误，借944-394（12450）一个 |  | SWO-203665 | Defect | HW | Aspiration Module | Needle Wash Unit | 1999 | Reported 1999.61.22 | Reported 1999.61.22 error, loan 944-394 (12450) one
No | Visit | Complaint | RDEU | 21.06.2017 10:15Uhr Herr Kohl 03441860814: Die Parameterleiste ist oft auf gelb, weil die Kalibration abgebrochen wird. | Immer wieder abgebrochene Kal | Überprüfung Dateiarchiv => immer wieder sporadisch Fehler 0367 pH $BG Modul nicht zeitgerecht gefüllt und 0878 keine Spüllösung während Kal Überprüfung, Ausbau und Reinigung alles Messventile => Membran vom pH $BG Airventil defekt => Austausch Mehrmals Kal ohne Abbrüche durchgeführt | SWO-203825 | Defect | HW | pH/BG Module | Valve, air | Membrane | 21.06.2017 10:15Am Mr. Kohl 03441860814: The parameter bar is often yellow because the calibration is aborted. | Time and again aborted Kal | Checking file archive => again and again sporadically error 0367 pH $BG module not filled on time and 0878 no flushing solution during Kal Inspection, removal and cleaning of all measuring valves => membrane from pH $BG air valve defective => Replacement Several times Kal performed without aborts
No | 8 | Y | RAME | V Analyzer is not booting up. |  | The analyzer keeps shutting itself off and, then it won't startup windows thus it would not load the AU software. | D2BEA516D44A4027 | Defect | HW | DMS_PC | Power Supply |  | V Analyzer is not booting up. |  | The analyzer keeps shutting itself off and, then it won't startup windows thus it would not load the AU software.
No | 6 | Y | RAME | 1 Ph unstable on random patient results. |  | Continued work from last visit same problem. | D2BF07D5D4121CC2 | Assisted Replacement | Electrode | Electrode | pH |  | 1 Ph unstable on random patient results. |  | Continued work from last visit same problem.
No | Visit | Complaint | RLTD | mori - 01865 221596 Lisa Tyler called back 23/06/2017, new electrode has not solved the issue. 01865 221596 | Failing pO2 and pCO2 QCs. | Inspected chamber and electrodes, all ok. Checked valves in service mode, checked gas flow in service, both ok. Ph chamber filling with fluid and not draining. Ph $bg air valve stuck, unable to pass air through. Stripped and cleaned valve which was very dirty. | SWO-203843 | Cleaned | HW | pH/BG Module | Valve, air |  | mori - 01865 221596 Lisa Tyler called back 23/06/2017, new electrode has not solved the issue. 01865 221596 | Failing pO2 and pCO2 QCs. | Inspected chamber and electrodes, all ok. Checked valves in service mode, checked gas flow in service, both ok. Ph chamber filling with fluid and not draining. Ph $bg air valve stuck, unable to pass air through. Stripped and cleaned valve which was very dirty.
No | Visit | Complaint | RDEU | Gerät meldet unter Status und System Meldungen: Neon Lamp Voltage too high | OXI Modul defekt. | Gerät überprüft. OXI Modul PCB erneuert, Fehler bleibt bestehen. Absprache mit Kundin getroffen,dass ein Ersatzteil ( OXI Modul ) bestellt wird. | SWO-205189 | Revisit required | HW | Oxi module | Oxi module |  | Device reports under Status and System Messages: Neon Lamp Voltage too high | OXY module Puncture. | device checked. OXI module PCB renewed, error remains. agreement with customer,that a spare part ( OXI module ) is ordered.
No | 5 | Y | RAME | d cohb cv too high on L2 cohb 13% |  | checked data | D2BF0F5CD41CCB8A | Defect | HW | Oxi module | Hemolyzer |  | d cohb cv too high on L2 cohb 13% |  | checked data
No | 2 | Y | RAME | W SCREEN FREEZING AND SLOW |  | checked opertaion screen very slow | D2BF1213D42CCB91 | Defect | HW | DMS_PC | Touch screen | Firmware | W SCREEN FREEZING AND SLOW |  | checked opertaion screen very slow
No | 2 | Y | RCZE | Blank | chyba 1067 pocas merania vzorky, reštart nepomaha |  | D2BE96BCD42B8F5B | Defect | HW | Storage Module | PCB, Cartridge Lock | 1067 | Blank | Error 1067 during sample measurement, restart Nepomaha
No | 23 | Y | RPOL | Blank | Wymieniono skaner, pomoc w konfiguracji raportów pacjenta. Testy analizatora. Aparat Sprawny. |  | D2BFEF374A5095CF | Defect | HW | Basic Unit | Barcode reader, display |  | Blank | Listed scanner, help in configuring patient reports. Analyzer tests. Camera functional.
No | Visit | Complaint | RLTD | Ranvir Tel 01865 221596 | On arrival the battery was indicating 0%. Within the service program the battery was indicating 0. The battery reset button was pressed, no change to battery. | Battery unit replaced. The LED indicator is indicating that the battery is charging and the battery percentage is raising. The battery values are updating in service. | SWO-204243 | Defect | HW | Battery module | Battery |  | Ranvir Tel 01865 221596 | On arrival the battery was indicating 0%. Within the service program the battery was indicating 0. The battery reset button was pressed, no change to battery. | Battery unit replaced. The LED indicator is indicating that the battery is charging and the battery percentage is raising. The battery values are updating in service.
No | Visit | Complaint | RDEU | 27.06.2017 Frau Villalobos 02014551832 meldet, dass die Einlassdichtung mit Halter abgebrochen ist. | Einlassdichtung in der Halterung abgebrochen. Rest der Dichtung lässt sich nicht ohne Weiteres entfernen. | Einlass demontiert und den Rest der Dichtung entfernt. Einlassnadel und Dichtung erneuert (aus Kundenbestand) und Gerät neu gestartet (Kaltstart). Funktionskontrolle: I.O. | SWO-204362 | Defect | HW | Inlet module | Inlet gasket | Broken | 27.06.2017 Mrs Villalobos 02014551832 reports that the inlet seal with holder has been broken. | Inlet seal in the bracket is broken. The rest of the seal cannot be easily removed. | inlet and remove scan the rest of the seal. Inlet needle and seal renewed (from customer stock) and device restarted (cold start). Functional control: I.O.
No | 6 | Y | RAME | Z No power |  | The analyzer had no power | D2BFCD444A44C32B | Defect | HW | Power Supply | Power Supply |  | Z No power |  | The analyzer had no power
No | 10 | Y | RAME | 3 PO2 Failing External QC. |  | PO2 was failing external qc during 20 day qc study. Level 1 was low, Levels 2 and 3 were high. Suspect is silicone contamination or malfunctioning sensor cassettes. Customer is concerned about instrument performance and would like instrument replaced if problem is not resolved by FSE. | D2C014BB4A0A4B8B | Assisted Replacement | SC | SC | Wetability |  | 3 PO2 Failing External QC. |  | PO2 was failing external qc during 20 day qc study. Level 1 was low, Levels 2 and 3 were high. Suspect is silicone contamination or malfunctioning sensor cassettes. Customer is concerned about instrument performance and would like instrument replaced if problem is not resolved by FSE.
No | 10 | Y | RFRA | Blank | blocage récurrent de la porte de baie cassette lors du changement de cassette |  | D2C013714A61D74C | Defect | HW | Basic Unit | Detector plate |  | Blank | recurring blocking of the cassette bay door when changing cassette
No | Visit | Complaint | RCHN | LAC定标失败 | LAC定标失败 | 更换测试卡后修复 | SWO-204847 | [5 empty] | LAC calibration failure | LAC calibration failure | Repair after replacing test card
No | Visit | Complaint | RLTD | Heather Canavan on 0141 4521324 | Monitor stuck with calibrate sensor message but will not calibrate | Changed sensor, still same Took logs and service dump to create TechLine case | SWO-204247 | [5 empty] | Heather Canavan on 0141 4521324 | Monitor stuck with calibrate sensor message but will not calibrate | Changed sensor, still same Took logs and service dump to create TechLine case
No | Visit | Complaint | RCHN | 报错1031 | 报错1031，需要借件 903-364备件一个 |  | SWO-203946 | Defect | HW | Process Module | Optical Unit | 1031 | 1031 | 1031, need to borrow 903-364 spare parts one
No | Visit | Complaint | RCHN | 定标失败 1059 | 定标失败 1059，借944-258（12398）一张， |  | SWO-203944 | Defect | HW | Process Module | Optical Unit | 1059 | Bid failure 1059 | Bid failure 1059, borrow 944-258 (12398),
No | Visit | Complaint | RCHN | 定标中止 1999 | 定标中止 1999，借944-216(12758)一张。（无货） | 更换944-216后修复 by jesson | SWO-203557 | Cleaned | HW | Process Module | Process wheel | 1999 | Bid Suspension 1999 | The bid was suspended 1999, with a loan of 944-216 (12758). (No goods) | Repair by jesson after replacement 944-216
No | Visit | Complaint | RDEU | PH Empfindlichkeitsfehler | PH Elektrode gebrochen | PH Elektrode (E777 XB-02) geprüft und ausgetauscht, Funktionskontrolle ok | SWO-204393 | Assisted Replacement | Electrode | Electrode | pH |  | PH Empfindlichkeitsfehler | PH electrode broken | PH electrode (E777 XB-02) tested and replaced, function control ok
No | 3 | Y | RAME | 2 pCO2 not passing level 1 and level 3 QC. |  | Examined data logs. Examined pCO2 and Cl electrodes. | D2C043AE4A18BFDB | Assisted Replacement | Electrode | Membrane | pCO2 |  | 2 pCO2 not passing level 1 and level 3 QC. |  | Examined data logs. Examined pCO2 and Cl electrodes.
No | Visit | Complaint | RLTD | mori - 01865 221596 Lisa Tyler called back 23/06/2017, new electrode has not solved the issue. 01865 221596 | Failing pco2 and po2 QCs. | This only started after a gas2 and cal solution change. Customer has replaced electrode. Tried new membranes and gas, both from other lot numbers. Replaced ref membrane. Chambers cleaned and inspected on previous visit, so swapped chambers BG for el $met, also swapped over gas valve assembly. Replaced solutions. Replaced black tubing. After all this still the QC reading is almost identical to original error values. | SWO-204300 | Revisit required | HW | Gas module | Gas valve | 1 & 2 | mori - 01865 221596 Lisa Tyler called back 23/06/2017, new electrode has not solved the issue. 01865 221596 | Failing pco2 and po2 QCs. | This only started after a gas2 and cal solution change. Customer has replaced electrode. Tried new membranes and gas, both from other lot numbers. Replaced ref membrane. Chambers cleaned and inspected on previous visit, so swapped chambers BG for el $met, also swapped over gas valve assembly. Replaced solutions. Replaced black tubing. After all this still the QC reading is almost identical to original error values.
No | Visit | Complaint | RDEU | Gerät meldet unter Status und System Meldungen: Neon Lamp Voltage too high | OXI Modul defekt. | OXI Modul zur Reparatur Überbrückung eingesetzt. | SWO-205186 | Defect | HW | Oxi module | Oxi module, Replaced |  | Device reports under Status and System Messages: Neon Lamp Voltage too high | OXY module Puncture. | OXI module used for repair bridging.
No | 2 | Y | RAME | Z The ABL827 would not initialize. |  | Tried restarting the unit and it failed inlet calibration. Ran the inlet calibration from the service screen and failed. | D2C0289A4A0B0269 | Defect | HW | Inlet module | Friction detector unit |  | Z The ABL827 would not initialize. |  | Tried restarting the unit and it failed inlet calibration. Ran the inlet calibration from the service screen and failed.
No | 3 | Y | RFRA | Blank | Blocage de l'AQT90, la porte des tubes ne se ferme pas bien. |  | D2C054D04A42D0C5 | Cleaned | HW | Cover and Chassis | Inlet Door |  | Blank | Blocking the AQT90, the tube door does not close well.
No | Visit | Complaint | RDEU | 27.06.2017 Sr. Elke 02345092085 meldet, das sie immer wieder das LP tauschen soll. | Datenspeicher ausgewertet | Funktion im Serviceprogramm überprüft, Fließwege gespült, Einlassdichtung aus Kundenbestand ersetzt. Funktion z. Zt. OK. Dem Anwender empfohlen das LP zu erneuern, wenn das Problem nochmals auftritt. | SWO-204597 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | 27.06.2017 Sr. Elke 02345092085 announces that she will exchange the LP again and again. | Data stores evaluated | Function checked in the service program, flow paths flushed, inlet seal replaced from customer stock. Function currently OK. The user is advised to renew the LP if the problem occurs again.
No | Visit | Complaint | RDEU | 27.06.2017, Sr. Christie, Tel: 017655597898(Handy privat-Telefonanlage defekt): Kunden möchten Dialysat messen und brauchen niedrigere Ca- Werte | Kunde mißt Blut mit angereichertem Dialysat und erhält zum Teil Ca Werte unter 0,4 ( derzeitigem Meßbereich) wünscht Meßbereichserweiterung | Meßbereich Ca von 0,4 auf 0,1 abgesenkt | SWO-204628 | Setup | SW | Measuring Range | Dialysis mode (range) |  | 27.06.2017, Sr. Christie, Tel: 017655597898(Mobile private telephone system defective): Customers want to measure dialysis and need lower Ca- values | Customer measures blood with enriched dialysis and receives ca values below 0.4 (current measuring range) wishes to expand the measuring range | Measuring range Ca lowered from 0.4 to 0.1
No | Visit | Complaint | RIND | Sensor Cassette not initializing. | Parameter failed. SC reinstallation not OK. | Cl Parameter failed. At the time of re-installation V $C fail msg coming. Found SC failed. SC replaced. QC run found all with in range. Sampling perform and found values not reliable as per clinical condition of Patient. Ask to conduct a Pre-analytical training. Dr given time for next day. Analyser serviced and calib rated properly. Handed over in a fine working condition. | SWO-205740 | [5 empty] | Sensor Cassette not initializing. | Parameter failed. SC reinstallation not OK. | Cl Parameter failed. At the time of re-installation V $C fail msg coming. Found SC failed. SC replaced. QC run found all with in range. Sampling perform and found values not reliable as per clinical condition of Patient. Ask to conduct a Pre-analytical training. Dr given time for next day. Analyser serviced and calib rated properly. Handed over in a fine working condition.
No | Visit | Complaint | RCHN | 泵轮故障 | 泵轮故障，需要借件902-971备件一个 | 更换902-971后修复(by tim) | SWO-204341 | [5 empty] | Pump wheel failure | Pump wheel failure, need to borrow 902-971 spare parts one | Repair after replacement 902-971 (by tim)
No | Visit | Complaint | RLTD | mori - 01865 221596 Lisa Tyler called back 23/06/2017, new electrode has not solved the issue. 01865 221596 | From viewing the data logs, it can be seen that the QC's began to fail on the 19th June. On the 19th June Cal2, Gas2 and pCO2 membranes were replaced. These have all since been replaced again with new lot number. It was also noted that glucose has been drifting. The AutoCheck module temperature was checked in the service program and it was reading 25.86 degrees. On speaking to the customer the night before, it was reading 27.5 degrees. The AutoCheck module was removed and inspected. The fan appeared to be working however there was a slight concern that it was not fully free to rotate, (slight resistance), this was unlikely to cause a problem but was replaced to rule this out. After replacing the fan, the temperature fell to 24.58 degrees. The measuring modules were removed and inspected. The measuring chambers and valves were removed in both the pH/BG module and the El/Met module. The pH and pCO2 chambers were dirty and were cleaned. The pH sealing was potentially not sitting correctly on the electrode and was replaced. The lactate measuring chamber was potentially in the early stages of the chamber separating from the base, therefore the measuring chamber was replaced. The pO2 and pCO2 electrodes were thoroughly cleaned. The gas module was tested in the service program and was regulating well. When the valves were opened 100% and the gas bottles were removed, the transducer actual voltage fell below that of the transducer offset voltage on both gas 1 and 2 suggesting that the valves may be leaking. Glucose was failing calibrations with low sensitivity and the membrane was replaced. | New gas regulators have been ordered. Radiometer to return to complete the repair once the spares become available. | SWO-204737 | Revisit required | HW | Gas module | Gas valve | 1 & 2 | mori - 01865 221596 Lisa Tyler called back 23/06/2017, new electrode has not solved the issue. 01865 221596 | From viewing the data logs, it can be seen that the QC's began to fail on the 19th June. On the 19th June Cal2, Gas2 and pCO2 membranes were replaced. These have all since been replaced again with new lot number. It was also noted that glucose has been drifting. The AutoCheck module temperature was checked in the service program and it was reading 25.86 degrees. On speaking to the customer the night before, it was reading 27.5 degrees. The AutoCheck module was removed and inspected. The fan appeared to be working however there was a slight concern that it was not fully free to rotate, (slight resistance), this was unlikely to cause a problem but was replaced to rule this out. After replacing the fan, the temperature fell to 24.58 degrees. The measuring modules were removed and inspected. The measuring chambers and valves were removed in both the pH/BG module and the El/Met module. The pH and pCO2 chambers were dirty and were cleaned. The pH sealing was potentially not sitting correctly on the electrode and was replaced. The lactate measuring chamber was potentially in the early stages of the chamber separating from the base, therefore the measuring chamber was replaced. The pO2 and pCO2 electrodes were thoroughly cleaned. The gas module was tested in the service program and was regulating well. When the valves were opened 100% and the gas bottles were removed, the transducer actual voltage fell below that of the transducer offset voltage on both gas 1 and 2 suggesting that the valves may be leaking. Glucose was failing calibrations with low sensitivity and the membrane was replaced. | New gas regulators have been ordered. Radiometer to return to complete the repair once the spares become available.
No | 2 | Y | RAME | G user called stating that LAC and PCO2 out on qc level 3 low for LAC and high for PCO2 |  | found that the user keeps getting inhomogeneous sample when they run a qc | D2C18A5EC03E06A1 | [5 empty] | G user called stating that LAC and PCO2 out on qc level 3 low for LAC and high for PCO2 |  | found that the user keeps getting inhomogeneous sample when they run a qc
No | 3 | Y | RAME | Z Customer reports that gas 1 has needed replacement 3x within the past week. |  | 0636, Gas 1 low and 0326, gas 1 empty, errors are present. Gas 1 has been replaced repeatedly as the customer indicated. | D2C18DFEC01B2776 | Defect | HW | Gas module | Gas valve | 636 | Z Customer reports that gas 1 has needed replacement 3x within the past week. |  | 0636, Gas 1 low and 0326, gas 1 empty, errors are present. Gas 1 has been replaced repeatedly as the customer indicated.
No | 13 | Y | RPNZ | Blank | pO2 calibration and QC failures - solution pack replacement fixed issue. |  | D2C275A636328D47 | Assisted Replacement | SP | SP | pO2 |  | Blank | pO2 calibration and QC failures - solution pack replacement fixed issue.
No | Visit | Complaint | RBNL | Afgebroken bloedmonster | Er zijn problemen met de capillaire metingen bij de ABL90. Tijdens de meting (meestal capillaire metingen), de inletgasket blijft vast zitten en daardoor de meting wordt afgebroken. | Inletmodule heb vervangen. | SWO-205216 | Defect | HW | Auto Inlet module | Auto Inlet module | Cap. | Broken Blood sample | There are problems with the capillary measurements at the ABL90. During the measurement (usually capillary measurements), the inletgasket remains stuck and therefore the measurement is broken down. | Have replaced the inletmodule.
Yes | 2 | Y | RPAC | Blank | pCO2 QC above control range |  | D2C29BAF3663D125 | [5 empty] | Blank | pCO2 QC above control range
No | 5 | Y | RFRA | Blank | Calibration LAC en erreur |  | D2C29B2D361915E7 | Defect | HW | Basic Unit | PCB, Sensor interface | Glu/Lac | Blank | LAC calibration in error
No | Visit | Complaint | RCHN | LAC定标失败 | LAC定标失败 | 更换测试卡后修复 | SWO-204997 | [5 empty] | LAC calibration failure | LAC calibration failure | Repair after replacing test card
No | Visit | Complaint | RCHN | 数字信号板故障 | 数字信号板故障，需要借件903-334备件一个 | 更换903-334后未修复，要重新借903-334和902-972各一个 by jesson | SWO-204392 | [5 empty] | Number signal board failure | Digital signal board failure, need to borrow 903-334 spare parts one | Replace 903-334 after not repaired, to borrow 903-334 and 902-972 each by jesson
No | 3 | Y | RAME | F Oxi error, error code 292, 581,688 on blood samples |  | checked data logs for errors, checked diagnostics, checked oxi pinch valve, | D2C310A236018F53 | [5 empty] | F Oxi error, error code 292, 581,688 on blood samples |  | checked data logs for errors, checked diagnostics, checked oxi pinch valve,
No | Visit | Complaint | RCHN | 进样模块损坏，进样针的弹簧断成两截无法恢复 | 进样模块损坏，进样针的弹簧断成两截无法恢复，借903-139一个。 | 更换903-139备件后修复(By tim) | SWO-204036 | Defect | HW | Inlet module | Inlet module |  | The sample module is damaged and the spring of the feeding needle is broken in two and cannot be recovered | The sample module is damaged, the spring of the feeding needle is broken in two and cannot be recovered, borrowing 903-139 one. | Repair after replacement of 903-139 spare parts (By tim)
No | 4 | Y | RAME | F OXI SPECTRUM MISMATCH ON CONTROL MATERIAL (INTERMITTENT) |  | INTERMITTENT PROBLEM PROBABLY SPECTROMETER OR HEMOLYZER | D2C28CEE36583C97 | Defect | HW | Oxi module | Hemolyzer |  | F OXI SPECTRUM MISMATCH ON CONTROL MATERIAL (INTERMITTENT) |  | INTERMITTENT PROBLEM PROBABLY SPECTROMETER OR HEMOLYZER
No | Visit | Complaint | RLTD | mori - 01865 221596 Lisa Tyler called back 23/06/2017, new electrode has not solved the issue. 01865 221596 | Arrived with the parts Tony had ordered following his visit yesterday. | Fitted new flow regulators for gas 1 & gas 2 as well as a new pressure transducer chassis. The analysers was left to start up and then the QCs run, for all levels the pCO2 & pO2 were back to the levels prior to the problem starting. The firmware for the touchscreen was also updated to solve the issue with the touchscreen becoming unresponsive. | SWO-205104 | Defect | HW | Gas module | Gas valve | 1 & 2 | mori - 01865 221596 Lisa Tyler called back 23/06/2017, new electrode has not solved the issue. 01865 221596 | Arrived with the parts Tony had ordered following his visit yesterday. | Fitted new flow regulators for gas 1 & gas 2 as well as a new pressure transducer chassis. The analysers was left to start up and then the QCs run, for all levels the pCO2 & pO2 were back to the levels prior to the problem starting. The firmware for the touchscreen was also updated to solve the issue with the touchscreen becoming unresponsive.
Yes | Visit | Complaint | RDEU | 29.06.2017, Herr Müller, tel.: 0401818879644: AQC Fehler bei pCO2 29.06.2017, herr Müller, Tel: 0401818879644: Weiterhin AQC Fehler bei pCO2 | Fehlermeldung nach Kal. PCO2 Empfindlichkeit zu niedrig. (48,4%) | Dem Kunden erklärt das es Probleme gibt. SC getauscht. Kunde bestellt neu. | SWO-205538 | Re classified RMED | Phone support |  |  |  | 29.06.2017, Mr. Müller, tel.: 0401818879644: AQC error with pCO2 29.06.2017, mr Müller, Tel: 0401818879644: Continue AQC error with pCO2 | Error message according to Kal. PCO2 sensitivity too low. (48,4%) | Explains to the customer that there are problems. SC. Customer reorders.
No | Visit | Complaint | RDEU | Sample Mixer gebrochen | Sample Mixer durch Fremdeinwirkung gebrochen | Sample Mixer Unit ersetzt, Gerät kalibriert und mit QC überprüft Defekt: R0946N005 | SWO-205662 | Defect | HW | Basic Unit | Sampler Mixer |  | Sample Mixer broken | Sample mixer broken by external influence | Sample Mixer Unit replaced, device calibrated and checked with QC Defect: R0946N005
No | Visit | Complaint | RDEU | Probleme mit Radiance. | Probleme mit Radiance Anbindung | Mit Herrn Samuel und Herrn Kozelj Rücksprache gehalten und den Analysator richtig an Radiance angebunden. ABL nun mit Radiance verbunden. | SWO-205636 | Setup | SW | His/lis | Radiance |  | Problems with Radiance. | Problems with Radiance connection | Consultation with Mr Samuel and Mr Kozelj and the analyzer properly connected to Radiance. ABL is now connected to Radiance.
No | Visit | Complaint | RDEU | Gerätestatus über LIVE Connect auf gelb | QC Level B und C fehlerhaft LP wurde gestern gewechselt | Pfl. Simon und Pfl. Christoph nicht im Dienst Beim Eintreffen Kal und alle QC Level fehlerhaft bei Crea $Urea => LP Wechsel (von Lot YN-01 auf YN-01) von gestern brachte keine Besserung Reinstallierung SC ServiceDump und DataLogs erstellt => Daten und Info an Ingo Später nochmal Rückruf auf Station durch RenRo | SWO-205505 | Assisted Replacement | SC | SC | Crea/Urea |  | Device status via LIVE Connect on yellow | QC Level B and C faulty LP was changed yesterday | Pfl. Simon and Pfl. Christoph not on duty Upon arrival Kal and all QC levels faulty at Crea $Urea => LP change (from Lot YN-01 to YN-01) from yesterday brought no improvement Reinstall SC ServiceDump and DataLogs created => Data and info to Ingo Later callback to station by RenRo
No | Visit | Complaint | RCHN | 电解质测量不准确，系统质检时会出现定标类型C/A | 电解质测量不准确，系统质检时会出现定标类型C/A，借902-972和903-334各一个 | 更换902-972和903-334后修复 by jesson | SWO-204779 | [5 empty] | Electrolyte measurement is not accurate, system quality inspection will appear calibration type C/A | Electrolyte measurement is not accurate, the system quality inspection will appear calibration type C/A, borrow 902-972 and 903-334 each one | Repair by jesson after replacing 902-972 and 903-334
No | Visit | Complaint | RCHN | 数字信号板故障 | 加借902-972和903-334各一个 | 更换902-972后修复 by jesson | SWO-204999 | [5 empty] | Number signal board failure | Borrow one each for 902-972 and 903-334 | Repair by jesson after replacement 902-972
No | Visit | Complaint | RCHN | 电池无法充电，无法使用 | 电池无法充电，无法使用，需要借件431-018C备件一个 | 更换备用电池431-018C后修复 by jesson | SWO-204898 | [5 empty] | Battery cannot be charged and cannot be used | Battery can not be charged, can not be used, need to borrow 431-018C spare one | Repair by jesson after replacing backup battery 431-018C
Yes | 5 | Y | RESP | dicen que han cambiado 2 SP y les sigue pidiendo cambio |  | Chequeo completo de equipo y comprobación de registro de datos. Sustitución de SP y SC Calibraciones y pasado qc | D2C42138AC531834 | [5 empty] | they say they've changed 2 SP and keep asking them for change |  | Complete computer check and data logging check. Replacing SP and SC Calibrations and past qc
No | Visit | Complaint | RCHN | 无法开机 | 无法开机，借910-285 | 更换910-285后修复 by jesson | SWO-204475 | Defect | HW | Power Supply | Power Supply |  | Cannot power on | Unable to power on, borrow 910-285 | Repair by after replacement 910-285 jesson
No | Visit | Complaint | RCHN | 试剂包液体漏液 | 试剂包液体漏液，需要借件903-199备件一个 | 液体传感器错误，更换试剂包，冲洗管路后，暂时解决。未更换配件 | SWO-205368 | Revisit required | HW | Basic Unit | Cassette retainer |  | Reagent pack liquid leakage | Reagent package liquid leakage, need to borrow 903-199 spare parts one | Liquid sensor error, replace the reagent package, rinse the line, temporarily solve. Unreplaced accessories
No | 3 | Y | RESP | Fallo en el transporte de liquidos. |  | Limpieza del circuito neumatico. Cambio de inlet gasket. Realizacion del test de servicio. Cal, qc. | D2C40878AC499E0B | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | Liquid transport failure. |  | Cleaning the tire circuit. Change of inlet gasket. Performing the service test. Cal, qc.
No | Visit | Complaint | RKK | CAL pH感度ｴﾗｰ。pH電極交換。pCO2感度不安定（100%超え頻発）の為、pCO2電極交換。（手持ち） | CAL pH感度ｴﾗｰ。pCO2感度不安定。（100%超え頻発）pCO2電極交換。（手持ち） | pH電極交換。pCO2電極交換。（手持ち） | SWO-205499 | Assisted Replacement | Electrode | Electrode | pH |  | CAL pH sensitivity error. pH electrode replacement. PCO2 electrode replacement due to pCO2 sensitivity instability (more than 100% frequent). (Hand held) | CAL pH sensitivity error. PCO2 sensitivity instability. (more than 100% frequent) pCO2 electrode replacement. (Hand held) | pH electrode replacement. PCO2 electrode replacement. (Hand held)
Yes | 5 | Y | RFRA | Blank | Erreur rinçage |  | D2C7E3B50E296C28 | [5 empty] | Blank | Rinse error
No | 7 | Y | RAME | F Neon lamp voltage too low erros on analyzer. Will not calibrate OXI. |  | Changed the oxi module PCB, power supply, oxi module completed. | D2C4B38822498B2F | Defect | HW | Oxi module | PCB, Oxi module |  | F Neon lamp voltage too low erros on analyzer. Will not calibrate OXI. |  | Changed the Oxygen module PCB, power supply, oxygen module completed.
No | Visit | Complaint | RKK | キャル時、Hb値のドリフトエラーが多発 | 修理完了 | モジュール返却し代替えモジュール引き上げ コンプリートテスト・ｔHbCalOK | SWO-205423 | Defect | HW | Oxi module | Hemolyzer |  | Frequent drift errors of Hb values at the time of | It's done. | Module return and replacement module pull Complete Test tHbCalOK
No | Visit | Complaint | RIND | Touch Hanging again and again. | USB Key to install. | USB stick ready but found BIOS setting not to be update as per requirement. Inform to Manager for further investigation and updates. Touch calibration done. Analyzer calibrated and found parameters well with in range. External Key board and Mouse connected and handed over in fine working condition. | SWO-205745 | Defect | HW | DMS_PC | Touch screen | Firmware | Touch Hanging again and again. | USB Key to install. | USB stick ready but found BIOS setting not to be update as per requirement. Inform to Manager for further investigation and updates. Touch calibration done. Analyzer calibrated and found parameters well with in range. External Key board and Mouse connected and handed over in fine working condition.
No | Visit | Complaint | RDEU | Probleme mit Radiance | ABL hat keine Verbindung mit Radiance | Mit Herrn Samuel und Herrn Kozelj Rücksprache gehalten und den Analysator richtig an Radiance angebunden. ABL nun mit Radiance verbunden. | SWO-205637 | Setup | SW | Roma.ini | His/lis |  | Problems with Radiance | ABL has no connection to Radiance | Consultation with Mr Samuel and Mr Kozelj and the analyzer properly connected to Radiance. ABL is now connected to Radiance.
No | Visit | Complaint | RIND | cartridges not scanning. | cartridge not scanning properply | checked and found scratches on bar code scanner . | SWO-205409 | Defect | HW | Storage Module | Barcode reader, Storage wheel |  | cartridges not scanning. | cartridge not scanning properply | checked and found scratches on bar code scanner .
Yes | Visit | Complaint | RCHN | 报错1031 | 报错1031，借903-364一个 | 更换903-364后修复 by jesson | SWO-205511 | Re classified RMED | Stock recovery |  |  | 1031 | 1031 | 1031, loan 903-364 one | Repair by jesson after replacing 903-364
No | Visit | Complaint | RCHN | 报错1031 | 报错1031，借903-364一个 | 更换903-364后修复 by jesson | SWO-205498 | Defect | HW | Process Module | Optical Unit | 1031 | 1031 | 1031, loan 903-364 one | Repair by jesson after replacing 903-364
No | Visit | Complaint | RKK | FLEX Q ポジション3 点灯せず | FLEX Q ポジション3 点灯せず | サンプラーディテクターPCB交換 CAL・QC測定データ確認 | SWO-206265 | Defect | HW | FlexQ Module | PCB, FLEXQ |  | FLEX Q Position 3 Without Lights | FLEX Q Position 3 Without Lights | Sampler Detector PCB Replacement Cal/QC measurement data confirmation
No | Visit | Complaint | RCHN | 报错2168 | 报错2168，借备件903-030一个 | 更换903-030后修复 by jesson | SWO-205448 | Defect | HW | Storage Module | PCB, Cartridge Lock | 2168 | 2168 error | 2168, borrowed spare parts 903-030 one | Repair by jesson after replacement 903-030
No | Visit | Complaint | RDEU | Status Fehler bei Kalibration pO2 | System überprüft => LP und Sensokassette defekt Material aus UKE geholt | LP und Sc erneuert, Kalibration und QC durchgeführt | SWO-206142 | Cleaned | Fluidic pathway | Fluidic pathway | pO2 | Decontamination | Status Error at Calibration pO2 | System checked => LP and senso cassette defective Material fetched from UKE | LP and Sc renewed, calibration and QC performed
No | 8 | Y | RAME | 8 Glu not present on the parameter bar or printout. | Changed the complaint visit tab from N to Y. | Troubleshot and found that the customer has a 946-013 sensor cassette installed on the analyzer, which does not have Glu. The customer is unsure how they got this part number because the only one they use is 946-005. | D2C7A5410E1AFE30 | User error | SC | SC | Wrong SC |  | 8 Glu not present on the parameter bar or printout. | Changed the complaint visit tab from N to Y. | Troubleshot and found that the customer has a 946-013 sensor cassette installed on the analyzer, which does not have Glu. The customer is unsure how they got this part number because the only one they use is 946-005.
No | Visit | Complaint | RDEU | 03.07.2017 10:10 Uhr Frau Wirth Tel.: 02863209926 Fehlermeldung: pCO2 Kalibrierdrift, pCO2-QC-Fehler 03.07.2017, 13.05Uhr, Frau Dr. Wirth, Tel: 0286320990: weiterhin Kalibrationsfehler und QC Fehler bei pCO2 | pCO2 Fehler QC alle Level | Gerät überprüft, Sensorkassette erneuert | SWO-206257 | Assisted Replacement | SC | SC | pCO2 |  | 07/03/2017 10:10 AM Ms Wirth Tel:02863209926 Error message: pCO2 calibration drift, pCO2 QC error 03.07.2017, 13.05h, Dr. Wirth, Tel: 0286320990: continued calibration error and QC error in pCO2 | pCO2 Fehler QC alle Level | Device checked, sensor cartridge renewed
No | Visit | Complaint | RDEU | 03.07.2017, 09:41 Uhr, Fr. Lenz, Tel. 0221890715190: Pat. Probe wird verworfen | Einlass bleibt bei Kapillarmessung in der Pushed Position hängen Zusätzlich: 1405 inkonsistente Datenbank | Einlass im Servicemenü in die Kapillarposition gebracht, eingedrückt - bleibt hängen. Bei weiteren Betätigungen rutscht der Einlass regelmäßig zurück, wirkt aber schwergängig. Einlass abgebaut, Nadel überprüft, war geringfügig (!) gebogen => nachgebogen und relevante Teile der Mechanik leicht geölt. 1405: SP entfernt, neu initialisiert, Analysator neu gestartet, Fehlermeldung beseitigt. | SWO-206068 | Adjustment | HW | Auto Inlet module | Auto Inlet module |  | 07/03/2017, 09:41 AM, Ms. Lenz, tel. 0221890715190: Pat. Sample is discarded | Inlet gets stuck in the pushed position when capillary measurement is measured Additional: 1405 inconsistent database | Inlet in the service menu brought into the capillary position, pressed in - gets stuck. With further actuation, the inlet regularly slips back, but appears difficult to operate. Inlet dismantled, needle checked, was slightly (!) bent => bent and relevant parts of the mechanics slightly oiled. 1405: SP removed, reinitialized, analyzer restarted, error message fixed.
No | Visit | Complaint | RDEU | Einlassdichtung erneuern | Einlassdichtung erneuern | Einlassdichtung erneuert | SWO-206113 | Defect | HW | Auto Inlet module | Inlet gasket |  | Renew ingess | Renew ingess | Inlet seal renewed
No | Visit | Complaint | RCHN | CO2，O2定标不过 | CO2，O2定标不过，借备件902-972,902-982各一个 | 更换902-972，接口板 | SWO-205889 | [5 empty] | CO2, O2 calibration is not | CO2, O2 calibration, spare parts 902-972,902-982 each | Replace 902-972, interface board
No | 33 | Y | RAME | O Failed the profiency test. |  | Checked the data logs. | D2C782E50E3AC952 | [5 empty] | O Failed the profiency test. |  | Checked the data logs.
No | 3 | Y | RPAC | Blank | zero's dont compare well to 7XX analyser |  | D2C76ADE0E3184EB | [5 empty] | Blank | zero's dont compare well to 7XX analyser
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 测试卡参比电极错误 | 更换测试卡后修复 | SWO-207789 | Assisted Replacement | SC | SC | Reference |  | Test card, reagent package failure | Test card reference electrode error | Repair after replacing test card
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 测试卡漏液 | 更换测试卡后修复 | SWO-207790 | Re classified RMED | PER |  |  |  | Test card, reagent package failure | Test card leakage | Repair after replacing test card
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | LAC，GLU质控错误 | 更换测试卡后修复 | SWO-207792 | Re classified RMED | PER |  |  |  | Test card, reagent package failure | LAC, GLU QUALITY CONTROL ERROR | Repair after replacing test card
Yes | Visit | Complaint | RCHN | 测试卡参比电极错误 1361 | 测试卡参比电极错误 1361 | 更换测试卡正常。 | SWO-206203 | [5 empty] | Test card reference electrode error 1361 | Test card reference electrode error 1361 | The replacement test card is OK.
No | Visit | Complaint | RCHN | Lac定标不过 | Lac定标不过 | 更换测试卡后修复 | SWO-205914 | Assisted Replacement | SC | SC | Lac |  | Lac calibration is not | Lac calibration is not | Repair after replacing test card
Yes | Visit | Complaint | RCHN | 报错1031 | 报错1031，借903-364一个 | 之前借的903-364有质量问题，现重新借903-364一个，此单先关单处理。(By Ran) The first borrowed OU may can't pass factory reference calibration. need another one. | SWO-205721 | Re classified RMED | HW | Process Module | Optical Unit | 1031 | 1031 | 1031, loan 903-364 one | The previous loan of 903-364 has quality problems, now re-borrow 903-364 one, this single first to close the single processing. (By Ran) The first borrowed ou may can't pass factory reference reference. need another one.
No | Visit | Complaint | RCHN | 报错1025 | 报错1025，借903-364一个 | 更换903-364备件后修复(By Ran) | SWO-205710 | Defect | HW | Process Module | Optical Unit | 1025 | 1025 error | 1025, loan 903-364 one | Repair after replacement of 903-364 spare parts (By Ran)
No | Visit | Complaint | RCHN | 干燥前温度超范围 | 干燥前温度超范围，报错1068，借903-210 | 更换备件903-210后，故障解决 | SWO-205880 | Revisit required | HW | Process Module | Cable, Cup dryer | 1068 | Overrange temperature before drying | Temperature above the range before drying, reported error 1068, borrowed 903-210 | Fault solving after replacement of spare parts 903-210
No | Visit | Complaint | RKK | 修理返却後ドリフト発生しランプ調整してもなおらず。再度引き上げ |  |  | SWO-206012 | Revisit required | HW | Oxi module | Oxi module |  | Drift occurs after repair return and the lamp is not adjusted. raise d'etre again
No | Visit | Complaint | RLTD | Analyser has a system message saying options will expire soon. Chris Bromley | Check analyser options, options had been installed several unused options were expiring in two days, however analyser was purchased with Crea and Crea options were only demo. | Made several phone calls to find out what had happened on installation and why there were no Crea options installed. Customer had definitely ordered Crea option but option had never been sent from RMED even though RLTD had ordered it, JA ordered Crea option to be sent ASAP. | SWO-206460 | Setup | SW | RMED SW | Demo Key exp | LEMAIL | Analyser has a system message saying options will expire soon. Chris Bromley | Check analyser options, options had been installed several unused options were expiring in two days, however analyser was purchased with Crea and Crea options were only demo. | Made several phone calls to find out what had happened on installation and why there were no Crea options installed. Customer had definitely ordered Crea option but option had never been sent from RMED even though RLTD had ordered it, JA ordered Crea option to be sent ASAP.
No | 5 | Y | RAME | F INTERMITTENT OXI SPECTRUM ERRORS CONTINUED |  | TRIED OXI MODULE PCB AND LAMP UNIT | D2C8273B0E33A924 | Revisit required | HW | Oxi module | Oxi module |  | F INTERMITTENT OXI SPECTRUM ERRORS SIZE |  | TRIED OXYGEN MODULE PCB AND LAMP UNIT
No | 4 | Y | RFRA | Blank | Erreur Autocheck |  | D2C7B4DE0E5AD1AB | Defect | HW | Autocheck module | Spindle |  | Blank | Autocheck error
No | 3 | Y | RFRA | Blank | Ampoule CQ non cassées par autocheck en automatique. |  | D2C7BF3E0E5AE5DC | Defect | HW | Autocheck module | Autocheck module | Breaker arm | Blank | CQ bulb not broken by autocheck automatically.
No | Visit | Complaint | RKK | SET SAFE 22 インレットドア開閉動作不良 | SET SAFE 22 インレットドア開閉動作不良 | インレットドア開閉レール清掃。サービスプログラムにて開閉動作確認。システムチェック実施。 | SWO-206244 | Cleaned | HW | Cover and Chassis | Inlet Door |  | SET SAFE 22 Inlet door opening and closing operation failure | SET SAFE 22 Inlet door opening and closing operation failure | Inlet door opening and closing rail cleaning. Check the opening and closing operation in the service program. System check was carried out.
No | 4 | Y | RFRA | Blank | plus de test disponnible |  | D2C81D8E0E01BA99 | Assisted Replacement | Consumables | SP | Solution Pack | Ejection | Blank | more available test
No | Visit | Complaint | RIND | AQT goes in maintenance and not allowing for any new test. | AQT goes in maintenance and not allowing for any new test. | Other activity data log checked and found Cleaning limit exceeds to 220. Need to do Blank cartridge cleaning . TO sending for Blank Cartridge and cleaning solution. | SWO-206164 | Revisit required | Consumables | Replacement | Blank cartridge Cleaning |  | AQT goes in maintenance and not allowing for any new test. | AQT goes in maintenance and not allowing for any new test. | Other activity data log checked and found Cleaning limit exceeds to 220. Need to do Blank cartridge cleaning . TO sending for Blank Cartridge and cleaning solution.
No | Visit | Complaint | RDEU | 21.06.2017 13:50Uhr Herr Hoxbergen 01708259944: Der Kunde möchte das Gerät an das Netz anschließen und dazu die IP-Adresse eingeben. | ASTM.INI korrekte Version für Conworx einspielen | astm.ini eingespielt und auf CF Karte gesichert, Applikation neu gestartet und Patmessungen nachgesendet, Hr. Korte(Conworx) bestätigt korrekte Übermittlung | SWO-206260 | Setup | SW | His/lis | His/lis |  | 21.06.2017 13:50h Mr. Hoxbergen 01708259944: The customer wants to connect the device to the network and enter the IP address. | Astm. Import INI correct version for Conworx | astm.ini recorded and saved on CF card, application restarted and pat measurements sent, Mr. Korte(Conworx) confirms correct transmission
No | Visit | Complaint | RDEU | 03.07.2017 18:00 Fr. Herd 030440307030: Touch ohne Funktion | Touchpanel lässt sich nicht kalibrieren, Mousezeiger zieht immer zur oberen Kante | Kalibrationen des Touchpanels gescheitert, Display ausgebaut und ersetzt, Neustart, Kal und QC | SWO-206466 | Defect | HW | DMS_PC | Touch screen |  | 07/03/2017 18:00 Fr. Herd 030440307030: Touch without function | Touch panel cannot be calibrated, mouse pointer always pulls to the top edge | Calibrations of the touch panel failed, display removed and replaced, restart, Kal and QC
No | Visit | Complaint | RDEU | 03.07.2017, 11:09 Uhr, Herr Studt, Tel. 0401818879728: Pumpe rattert laut | System überprüft | Kein Fehler feststellbar , Kalibration und QC durchgeführt | SWO-206307 | No fault found | HW | Basic Unit | Pump |  | 07/03/2017, 11:09 AM, Mr. Studt, Tel. 0401818879728: Pump rattles loudly | System checked | No error detectable , calibration and QC performed
No | Visit | Complaint | RSCH | Blank | Impossible de lire la puce SP (sporadique) garantie | Nettoyage du capteur d'enregistrement de la puce e.o Graissage pour le verrouillage du SP e.o Calibration e.o contrôle de qualité e.o | SWO-206469 | Cleaned | HW | Basic Unit | Solution Pack ID PCB |  | Blank | Can't read the Guaranteed SP Chip (sporadic) | Cleaning the e.o chip recording sensor Grease for locking the SP e.o Calibration e.o quality control e.o
No | Visit | Complaint | RBNL | Het is ons opgevallen dat de tijd op de C1CC ABLflex+ regelmatig een uur afwijkt. Na het herstellen van de tijd zie we dat hij in de loop der tijd, zomaar weer terug gaat naar de oude tijd. Een uur vroeger. | Zomer- wintertijds instelling stond niet aangevinkt. | Zomer-wintertijds instelling aangevinkt en bevestigd in windows. | SWO-206470 | Setup | SW | Roma.ini | Time setting |  | We noticed that the time on the C1CC ABLflex + regularly deviates an hour. After recovering from time, we see that in the course of time, it just goes back to the old days. An hour earlier. | Summer-Wintertime setting was unchecked. | Summer-wintertime setting checked and confirmed in Windows.
No | Visit | Complaint | RBNL | Beste Rheza, Zo af en toe de knoppen Analyzer Status en Data Overzicht boven in het beeldscherm van de ABL 90+ het niet. Alle andere knoppen wel. Het zijn alleen deze 2 knoppen die het dan niet doen en na een paar minuten werkt alles ineens weer. Enig idee?? Het is nu al 3 keer voorgekomen. Lastig maar nog niet onwerkbaar. Kan ik er zelf iets aan doen? Met vriendelijke groet, logo-handtekening-Outlook-2016-klein (2)Marja de Lely ¦ Gespecialiseerd Analiste POCT ¦ KCL Flevoziekenhuis E MdLely@flevoziekenhuis.nl ¦ T 036-868 8978 ¦ full time aanwezig Hospitaalweg 1 1315 RA Almere │Postbus 3005 1300 EG Almere T 036-868 8888 │www.flevoziekenhuis.nl | Beste Rheza, Zo af en toe de knoppen Analyzer Status en Data Overzicht boven in het beeldscherm van de ABL 90+ het niet. Alle andere knoppen wel. Het zijn alleen deze 2 knoppen die het dan niet doen en na een paar minuten werkt alles ineens weer. Enig idee?? Het is nu al 3 keer voorgekomen. Lastig maar nog niet onwerkbaar. Kan ik er zelf iets aan doen? Met vriendelijke groet, | Het scherm heb vervangen vervolgens heb gecontroleerd, OK. | SWO-206478 | Defect | HW | DMS_PC | Display |  | Dear Rheza, So occasionally the buttons Analyzer Status and Data Overview at the top of the display of the ABL 90 + it does not. All the other buttons though. It's only these 2 buttons that don't do it and after a few minutes everything will work all of a sudden. Any idea??? It has already been featured 3 times now. Tricky but not yet unworkable. Can I do something about it myself? Sincerely Logo-Signature-Outlook-2016-Small (2) Marja de Lely ¦ Specialized analyste POCT ¦ KCL Flevoziekenhuis E MdLely@flevoziekenhuis.nl ¦ T 036-868 8978 ¦ Full Time present Hospitaalweg 1 1315 RA Almere │ P.O. Box 3005 1300 EG Almere T 036-868 8888 │ www.flevoziekenhuis.nl | Dear Rheza, So occasionally the buttons Analyzer Status and Data Overview at the top of the display of the ABL 90 + it does not. All the other buttons though. It's only these 2 buttons that don't do it and after a few minutes everything will work all of a sudden. Any idea??? It has already been featured 3 times now. Tricky but not yet unworkable. Can I do something about it myself? Sincerely | The screen got replaced then checked, OK.
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | PCO2定标失败，结果偏高 | 更换测试卡 | SWO-206426 | [5 empty] | Test card, reagent package failure | PCO2 calibration fails, result is high | Replace the test card
No | Visit | Complaint | RHUN | pO2 QC hiba | - | zéró áramok beállítása, 2 pontos kalibráció, QC mérés | SWO-206384 | [5 empty] | pO2 QC hiba | - | Zero currents, 2-point calibration, QC measurement
No | Visit | Complaint | RBNL | Courant de fuite détecté | Pendant la nuit toutes les calibrations étaient refusés à cause d'un problème de courant de fuite. Remplacement de la vanne rinçage (EL/Met) parce qu'elle était remplacé la semaine dernière par une vanne option (module fluidique) Déconnecté toutes les vannes afin de déterminer la cause : toujours le même problème. Vérification toutes les modules : ils sont tous propre sans fuite de liquide. Finalement l'appareil c'est mis en mode prêt sans erreur. La valeur de 'Isolation' en service reste sur 350 (max 500) La cause reste donc inconnu. Si le problème persiste on doit le reprendre pour réparation à la firme. | Cause inconnu. | SWO-206530 | Revisit required | HW | Basic Unit | Valve, leak current | Leak Current | Leak current detected | During the night all calibrations were refused due to a leak current problem. Replacement of the rinse valve (EL/Met) because it was replaced last week by an option valve (fluidmodule) Disconnected all the valves to determine the cause: always the same problem. Check all modules: they are all clean without leaking liquid. Finally the device is put in error-free ready mode. The value of 'Isolation' in service remains on 350 (max 500) The cause remains unknown. If the problem persists, it must be taken back for repair at the firm. | Unknown cause.
No | Visit | Complaint | RDEU | 03.07.2017 10:30Uhr Herr Wegner 030450675277: Temperaturfehler im AutoQC-Modul | AC-Temperaturfehler | Fehlersuche; Gerät aus; AC-Modul ausgebaut; opt. Prüfung = Flachbandkabelauffällig geknickt; Kabel ersetzt; AC-Modul gereinigt; Gerät ein; Kal.2 + Reinigungslsg. gew.; Neustart; Kal. + QC i.O. | SWO-206507 | Defect | HW | Autocheck module | FPC Cable |  | 07/03/2017 10:30AM Mr Wegner 030450675277: Temperature error in the AutoQC module | AC-Temperaturfehler | Troubleshooting; device off; AC module removed; Opt. Testing = ribbon cable conspicuously bent; Cable replaced; AC module cleaned; device; Cal.2 + Cleaning w.; Restart; Kal. + QC in the first place
No | Visit | Complaint | RLTD | Analyser has a system message saying options will expire soon. Chris Bromley | Return visit to install relieved option key. | Install Crea option. | SWO-206459 | Setup | SW | RMED SW | Demo Key exp | LEMAIL | Analyser has a system message saying options will expire soon. Chris Bromley | Return visit to install relieved option key. | Install Crea option.
No | Visit | Complaint | RIND | AQT goes in maintenance and not allowing for any new test. | Cleaning to do. | Blank cartridge received from Nagpur. But Barcode found corrupt. Asked new barcodes by mail but not useful. need to get new Blank Cartridge and cleaning solution. TO Sent and asked one cartridge from Delhi. | SWO-207157 | Revisit required | Replacement | Replacement | Blank cartridge Cleaning |  | AQT goes in maintenance and not allowing for any new test. | Cleaning to do. | Blank cartridge received from Nagpur. But Barcode found corrupt. Asked new barcodes by mail but not useful. need to get new Blank Cartridge and cleaning solution. TO Sent and asked one cartridge from Delhi.
No | Visit | Complaint | RIND | cartridge barcode fail to scan . | Check the barcode reader. | Clean the dust & reset the beam position still. intermittent error. need to replace the barcode. | SWO-207674 | Cleaned | HW | Storage Module | Barcode reader, Storage wheel |  | cartridge barcode fail to scan . | Check the barcode reader. | Clean the dust & reset the beam position still. intermittent error. need to replace the barcode.
No | 31 | Y | RPAC | Blank | Leaked detected |  | D2CA31ABFA288C67 | Assisted Replacement | SP | SP | Leakage |  | Blank | Leaked detected
No | Visit | Complaint | RDEU | Aufforderung zu LP Tausch | Aufforderung zu LP Tausch | Kunden LP nochmals installiert, Kal, QC | SWO-207281 | Revisit required | SP | SP | SP |  | Prompt for LP Exchange | Prompt for LP Exchange | Customer LP reinstalled, Kal, QC
No | Visit | Complaint | RDEU | Aufforderung zu LP Tausch | Aufforderung zu LP Tausch | Serviceprogramme durchgeführt, Gleiches LP erneut gewechselt, Analysator komplett aus- und wieder eingeschaltet, Kal, Fehler Flüssigkeitstransport gescheitert, Abbruch, da kein LP mehr vorrätig | SWO-207282 | Assisted Replacement | SP | SP | SP | Wetability | Prompt for LP Exchange | Prompt for LP Exchange | Service programs carried out, same LP changed again, analyzer completely switched off and back on, Kal, fault fluid transport failed, cancellation, since no LP is in stock anymore
No | 2 | Y | RPOL | Blank | Wymiana uszczelki wlotu oraz pakietu roztworów. Kalibracja tHb. Analizator sprawny technicznie. |  | D2CA0C36FA29B77F | Defect | HW | Auto Inlet module | Inlet gasket |  | Blank | Replacing the inlet gasket and the solution package. THb Calibration. Technically efficient analyzer.
No | Visit | Complaint | RDEU | 09.06.2017, Sr. Marion, Tel.: 0211904361241: mögliche Verstopfung der Fließwege 03.07.2017 09:50Uhr Frau Janeck 021190430: Regelmäßig erscheint die MeldungFließwege verstopft 03.07.2017, 09:56 Uhr, Fr. Janeck, Tel. 021190430: Kd. informiert, dass der Fehler möglicherweise mit einer Proteinentfernung zusammenhängt. | Schlauchwege verstopft. | Schlauchwege erneuert. Frau Janik hatte vorab angekündigt, dass Betriebsmittel vorhanden wären. Kundin hat kein LP mehr vorhanden. LP in der Zentrale geholt und installiert. Gerät kalibriert. | SWO-206627 | [5 empty] | 09.06.2017, Sr. Marion, Tel.: 0211904361241: possible blockage of the flow paths 07/03/2017 09:50PM Ms Janeck 021190430: The message appears clogged regularly 07/03/2017, 09:56, Ms. Janeck, tel. 021190430: Kd. informs that the error may be related to protein removal. | hoseways clogged. | Hose ways renewed. Mrs Janik had announced in advance that resources would be available. Customer no longer has an LP. LP at the head office and installed. device calibrated.
No | 3 | Y | RESP | Fallo de impresión |  | Ajuste del rodillo de impresión, limpieza de conectores y limpieza de cabezal | D2CA5D64FA517B17 | [5 empty] | Printing failure |  | Printing roller adjustment, connector cleaning and head cleaning
No | Visit | Complaint | RCHN | 定标失败，灵敏度错误 | 定标失败，灵敏度错误，借903-365,902-778 | 选择送修 | SWO-206656 | [5 empty] | Bid failed, sensitivity error | Bid failure, sensitivity error, borrowing 903-365, 902-778 | Select to send for repair
Yes | Visit | Complaint | RCHN | 报错1031 | 之前借的903-364有质量问题，现重新借903-364一个 | 更换903-364后修复(By Ran) | SWO-206005 | Re classified RMED | HW | Process Module | Optical Unit | 1031 | 1031 | The previous loan of 903-364 has quality problems, now re-borrow903-364 one | Repair after replacement 903-364 (By Ran)
No | 16 | Y | RPOL | Blank | Wymiana zaworu Waste, membran zaworowych w module pH/BG, wê¿yka pompowego pH/BG oraz uszczelki wlotu. Testy serwisowe - ok. Probki pacjentów - ok. Analizator sprawny technicznie. |  | D2CA26C3FA51863D | Defect | HW | pH/BG Module | Valve, air |  | Blank | Replacement of the Waste valve, valve membranes in the pH/BG module, pH/BG pump and inlet gasket. Service Tests-OK. Patient Test-OK. Technically efficient analyzer.
No | 7 | Y | RAME | 9 LACTATE UNSTABLE DURING CALIBRATIONS |  | FOUND ORING MISSING ON LACTATE, PO2 AND GLUCOSE. | D2C9FFB4FA3F3308 | Defect | Electrode | Electrode | Sealing ring | Sealing ring, pO2 | 9 LACTATE UNSTABLE DURING CALIBRATIONS |  | FOUND ORING MISSING ON LACTATE, PO2 AND GLUCOSE.
No | 7 | Y | RAME | 1 Ph unstable errors on blood samples. Occurring about 1 every week. |  | No errors showing up on calibrations or qc. Only blood samples about 1 per 5-7 days. Sample will be rerun and errors will be gone. Happening in all kinds of samples not a certain sample type or operator. This is a continuation of previous service performed for same problem. | D2CA3018FA1B5BFB | Defect | HW | pH/BG Module | PCB, ph/BG |  | 1 Ph unstable errors on blood samples. Occurring about 1 every week. |  | No errors showing up on calibrations or qc. Only blood samples about 1 per 5-7 days. Sample will be rerun and errors will be gone. Happening in all kinds of samples not a certain sample type or operator. This is a continuation of previous service performed for same problem.
No | 1 | Y | RPOL | Blank | Zmiana ustawieñ analizatora oraz upgrade softwearu w celu rozwi¹zania problemu z blokuj¹cym siê ekranem. Testy kontrolne. Analizator sprawny. |  | D2CA30ACFA499EE3 | Defect | HW | DMS_PC | Touch screen | Firmware | Blank | Change the analyzer settings and the Softwearu upgrade to resolve the screen lock problem. Control tests. Functional Analyzer.
Yes | Visit | Complaint | RDEU | 29.06.2017 Frau Beier 023234995291 meldet das die CRP-Kontolle außer Bereich ist. 29.06.2017 11:30Uhr Frau Beier 023234995291: Alle Maßnahmen wurden durchgeführt, aber das Gerät meldet weiterhin QC-Fehler | Siehe Jahreswartung | Fehlerbehebung bei Jahreswartung . | SWO-206933 | [5 empty] | 29.06.2017 Ms Beier 023234995291 reports that the CRP-Conftolle is out of range. 06/29/2017 11:30 AM Ms Beier 023234995291: All measures have been taken, but the device continues to report QC errors | See Annual Maintenance | Troubleshooting Annual Maintenance .
No | Visit | Complaint | RKK | 圧力テスト吸引エラー | 圧力テスト吸引エラー | ポンプローター交換 インレットガスケット交換 インレット分解・調整 各部動作確認 内部QCエラーの為、SC交換 CAL・QC測定データ確認 | SWO-208593 | Defect | HW | Basic Unit | Pump Rotor |  | Pressure test suction error | Pressure test suction error | Pump rotor replacement Inlet gasket replacement Inlet decomposition and adjustment Check the operation of each part SC replacement due to internal QC error Cal/QC measurement data confirmation
No | Visit | Complaint | RDEU | Gerät Konfiguration für stetigen Ping erforderlich. Kunde hat Verbindungsprobleme mit dem Radiance Server | Geräte Konfig für stetigen Ping und Refrenz Werte erforderlich. | Geräte konfiguriert. Ref Werte nach Kundenangaben konfiguriert. | SWO-207295 | Setup | SW | His/lis | His/lis |  | Device configuration required for continuous ping. Customer has connection problems with Radiance Server | Devices Config required for continuous ping and refrenz values. | devices configured. Ref values configured according to customer specifications.
Yes | 4 | Y | RPAC | Blank | Qc adaptor for ABL90FlexPlus |  | D2CB4E3B704F8DF2 | Re classified RMED |  |  |  |  | Blank | Qc adaptor for ABL90FlexPlus
Yes | Visit | Complaint | RBNL | Afgebroken bloedmonster | Op verzoek van RMED, in verband met de Techline case RBNL 784190857, Servicedump file heb gekopieerd. | Servicedump file heb gekopieerd. | SWO-207191 | Re classified RMED | to collect dat for TL |  |  |  | Broken Blood sample | At the request of RMED, in connection with the Techline case RBNL 784190857, have copied Service dump file. | Have copied service dump file.
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | Lac定标不过 | 更换测试卡后修复 | SWO-207371 | [5 empty] | Test card, reagent package failure | Lac calibration is not | Repair after replacing test card
No | Visit | Complaint | RCHN | 条码扫描失败 | 条码扫描失败，借902-992一个 | 更换902-992后修复(By Ran) | SWO-206373 | [5 empty] | Barcode scanning failed | Barcode scanning failed, borrowing 902-992 one | Repair after replacement 902-992 (By Ran)
No | Visit | Complaint | RCHN | 接口板腐蚀 | 接口板腐蚀，借902-972 | 更换902-972后修复(By Ran) | SWO-206681 | [5 empty] | Interface plate corrosion | Interface plate corrosion, loan 902-972 | Repair after replacement 902-972 (By Ran)
No | Visit | Complaint | RCHN | 不均匀定标2溶液 | 不均匀定标2溶液 | 更换试剂包后修复 | SWO-206892 | Cleaned | Fluidic pathway | Inlet module | Inlet gasket | 1083 | Uneven calibration 2 solution | Uneven calibration 2 solution | Repair after replacement of reagent kit
Yes | Visit | Complaint | RCHN | 不均匀定标2溶液 | Lac，cl定标不过 | 更换测试卡后修复 | SWO-206893 | Re classified RMED |  |  |  |  | Uneven calibration 2 solution | Lac, cl-de-standardbutd | Repair after replacing test card
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 测试卡参比电极错误 | 更换测试卡后修复 | SWO-207831 | Assisted Replacement | SC | SC | Reference |  | Test card, reagent package failure | Test card reference electrode error | Repair after replacing test card
No | Visit | Complaint | RCHN | 试剂包更换检测失败 | 试剂包更换检测失败 | 更换试剂包后修复 | SWO-207336 | Assisted Replacement | SC | SC | Wetability |  | Reagent pack replacement test failed | Reagent pack replacement test failed | Repair after replacement of reagent kit
No | Visit | Complaint | RCHN | 试管架没有滚轮 | 试管架没有滚轮，借886-731/2个 Tube insert was broken. | 更换配件886-731(By Ran) | SWO-206277 | Defect | HW | Aspiration Module | Tube holder |  | There are no rollers in the test tube holder | No rollers in the test tube rack, 886-731/2 Tube insert was broken. | Replacement accessory 886-731 (By Ran)
No | Visit | Complaint | RCHN | 干燥前温度超范围 | 干燥前温度超范围，现借636-561 | 更换636-561一个正常。(By Ran) | SWO-206622 | Defect | HW | Process Module | Cable, Cup dryer | 1068 | Overrange temperature before drying | Over-range temperature before drying, 636-561 now | Replace 636-561 one normal. (By Ran)
No | Visit | Complaint | RCHN | 报错1031 | 报错1031，借903-364一个 | 更换903-364后修复(By Ran) | SWO-206621 | Defect | HW | Process Module | Optical Unit | 1031 | 1031 | 1031, loan 903-364 one | Repair after replacement 903-364 (By Ran)
No | Visit | Complaint | RCHN | 试管架没有滚轮 | 试管架没有滚轮，借886-731/个 | 更换配件886-731(By Ran) | SWO-206270 | Defect | HW | Aspiration Module | Tube holder |  | There are no rollers in the test tube holder | No rollers in the test tube rack, borrowed 886-731/each | Replacement accessory 886-731 (By Ran)
No | Visit | Complaint | RCHN | 报错1029，无其他报错提示 | 报错1029，借903-364 |  | SWO-206921 | Defect | HW | Process Module | Optical Unit | 1029 | 1029 error, no other error prompt | 1029, loan 903-364
No | Visit | Complaint | RBNL | Courant de fuite détecté | Toujours le même problème, la valeur 'Isolation' en service est à 1250. Dernier essai avec by-pass du Inlet preheater, mais le problème persiste. | Reprise de l'appareil pour diagnostiquer à la société | SWO-207228 | Revisit required | HW | Basic Unit | Valve, leak current | Leak Current | Leak current detected | Always the same problem, the value 'Isolation' in service is 1250. Last try with by-pass of the Inlet preheater, but the problem persists. | Takeover of the device to diagnose to society
Yes | 6 | Y | RAME | P Fluid flow issues |  | Customer complained about a pool of blood built up inside the measuring chamber and leaking out as well. pH and reference electrode were filled with fluid and blood. | D2CB8B2F704E08D8 | [5 empty] | P Fluid flow issues |  | Customer complained about a pool of blood built up inside the measuring chamber and leaking out as well. pH and reference electrode were filled with fluid and blood.
No | 10 | Y | RPNZ | Blank | Calibration and QC issues Na, Ca. |  | D2CCCEDDE622F078 | Assisted Replacement | Electrode | Membrane | Reference |  | Blank | Calibration and QC issues Na, Ca.
No | 8 | Y | RAME | Z Customer changing Gas1 every 7-10 days | Changed the complaint visit tab from N to Y. | checked replacement logs, Gas1 changed on 5 $29, 6 $5, 6 $12, 6 $19, 6 $26, 7 $3 | D2CBD7CC704F8F77 | Defect | HW | Gas module | Gas valve |  | Z Customer changing Gas1 every 7-10 days | Changed the complaint visit tab from N to Y. | checked replacement logs, Gas1 changed on 5 $29, 6 $5, 6 $12, 6 $19, 6 $26, 7 $3
No | 2 | Y | RAME | F calibrations fail thb halogen lamp voltage too high |  | lamp intensity was 3800. | D2CB7C7570436379 | Adjustment | HW | Oxi module | Lamp unit | Halogen | F calibrations fail thb halogen lamp voltage too high |  | lamp intensity was 3800.
No | Visit | Complaint | RIND | AQT goes in maintenance and not allowing for any new test. | Black Cartridge cleaning. | Blank cartridge and cleaning solution received from Delhi. Blank cartridge installed and cleaning done. Cleaning solution and blank cartridge taken back. Analyzer Serviced and calibrated properly. Handed over in a fine working condition. | SWO-207176 | User error | Replacement | Replacement | Blank cartridge Cleaning |  | AQT goes in maintenance and not allowing for any new test. | Black Cartridge cleaning. | Blank cartridge and cleaning solution received from Delhi. Blank cartridge installed and cleaning done. Cleaning solution and blank cartridge taken back. Analyzer Serviced and calibrated properly. Handed over in a fine working condition.
No | Visit | Complaint | RIND | trop-t calibration failed. | Check the system. check the logs. check the activity log. | Calibration aborted. check the incubation wheel. check the cartridge. Recalibrate the trop-t. calibration fail due to background cup fail. clean the process wheel & run cleaning process. | SWO-207681 | Cleaned | HW | Process Module | Process wheel |  | trop-t calibration failed. | Check the system. check the logs. check the activity log. | Calibration aborted. check the incubation wheel. check the cartridge. Recalibrate the trop-t. calibration fail due to background cup fail. clean the process wheel & run cleaning process.
No | 10 | Y | RPAC | Blank | SC failed on Lactate Sensitivity |  | D2CC7774E65D5057 | Assisted Replacement | SC | SC | Lac |  | Blank | SC failed on Lactate Sensitivity
No | 10 | Y | RPAC | Blank | Liquid Sensor errors |  | D2CC7451E6435054 | Defect | HW | Basic Unit | LS Sensor |  | Blank | Liquid Sensor errors
No | 2 | Y | RFRA | Blank |  |  | D2CC8CEBE6507E12 | Defect | HW | Inlet module | Inlet module |  | Blank
No | Visit | Complaint | RCHN | 多参数定标不过，K,NA,CA,CL,LAC | 多参数定标不过，借备件903-334一个 | 更换903-334备件后修复(By Ran) | SWO-206919 | [5 empty] | Multi-parameter calibration, K, NA, CA, CL, LAC | Multi-parameter calibration, however, borrowed spare parts 903-334 one | Repair after replacement of 903-334 spare parts (By Ran)
Yes | 33 | Y | RAME | 3 po2 fails internal qc level 3 low |  | instrument ran system cycle at 4 am failed po2. changed solution pack | D2CCA11CE635D167 | [5 empty] | 3 po2 fails internal qc level 3 low |  | instrument ran system cycle at 4 am failed po2. changed solution pack
No | 2 | Y | RAME | P Customer called with leakage around the solution pack. |  | Troubleshooting found possible waste port and door assembly loose. | D2CCA39AE6555F31 | [5 empty] | P Customer called with leakage around the solution pack. |  | Troubleshooting found possible waste port and door assembly loose.
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | Lac定标不过 | 更换测试卡后修复 | SWO-207830 | Assisted Replacement | SC | SC | Lac |  | Test card, reagent package failure | Lac calibration is not | Repair after replacing test card
No | 6 | Y | RAME | F OXI SPECTRUM MISMATCH CONTINUED |  | REMOVED LAMP UNIT AND OXI PCB FOR STOCK RETURN AND TRIED NEW HEMOLYZER WITH NO HELP | D2CC8641E63F2AAA | Defect | HW | Oxi module | Oxi module, Replaced |  | F OXI SPECTRUM MISMATCH CONTINUED |  | REMOVED LAMP UNIT AND OXI PCB FOR STOCK RETURN AND TRIED NEW HEMOLYZER WITH NO HELP
No | 2 | Y | RAME | 2 Customer stating pCO2 is not passing qc levels 1, 2, and 4. |  | Arrived on site and noted that QC was within package insert but the customer was failing on there peer comparison | D2CCC126E631E020 | Setup | SW | DMS_PC | pCO2 | 6.17 | 2 Customer stating pCO2 is not passing qc levels 1, 2, and 4. |  | Arrived on site and noted that QC was within package insert but the customer was failing on there peer comparison
No | Visit | Complaint | RDEU | Immer wieder Kal Driften und Parameter instabil bei Parametern aus dem E-Lyte Modul | Fehlersuche, immer wieder Kal Drift und Statusfehler, QC Lavel3 instabil bei Lyte und PH | Ventile BG und ELyte Modul ausgebaut kontrolliert und gereinigt, Oxi Modul Ventil hackt zeitweise -Ventil ausgebaut und gereinigt, Fluidiq Modul ausgebaut - Ventile Rinse Air und Waste Flussigkeitsschaden ( Rost) - Ventile ersetzt, Flüssigkeits Transport Fehler behoben, Kal und QC durchgeführt | SWO-207651 | Defect | HW | El/Met Module | Valve, air | Liquid leak | Again and again Kal drifting and parameters unstable with parameters from the E-Lyte module | Troubleshooting, again and again Kal Drift and status error, QC Lavel3 unstable at Lyte and PH | Valves BG and ELyte module removed controlled and cleaned, Oxi module valve hacks intermittent valve removed and cleaned, Fluidiq module removed - valves Rinse Air and Waste flow damage (rust) - Valves replaced, fluid transport error fixed, Kal and QC performed
No | Visit | Complaint | RDEU | Remote Support funktioniert nicht | Demo Key beinhaltet keine Remote Support | Key installiert. | SWO-207835 | Setup | SW | RMED SW | Remote support |  | Remote support is not working | Demo Key does not include Remote Support | Key installed.
No | 22 | Y | RESP | Error 5784 Safe 22 |  | Comprobar posiciones de seguridad, desatascar cartucho de test en la rueda de almacén. Revisión del equipo, limpiando el polvo de interior, filtros de aire y muestras, lectores de código de barras, circuito de paso muestras, engrase de mecanismos, borrado de datos no útiles, prueba de blancos y controles Actualizar el software a la versión V.8.9 ES | D2CC944EE618CC0A | Defect | Consumables | Test Kit | Cartridge | Broken | Error 5784 Safe 22 |  | Check safety positions, unblock test cartridge on the storage wheel. Equipment review, cleaning indoor dust, air filters and samples, barcode readers, pass-through circuit samples, greasing mechanisms, erasure of inuseful data, white test and controls Update the software to v.8.9 ES version
Yes | Visit | Complaint | RIND | trop-t calibration failed. | Check the system.chekc the logs. | Recalibrate the system & calibration pass. | SWO-207684 | Re classified RMED | Courtesy |  |  |  | trop-t calibration failed. | Check the system.chekc the logs. | Recalibrate the system & calibration pass.
No | Visit | Complaint | RCHN | 开机开不了，没电源 | 开机开不了，没电源，借903-663电源模块 | 更换备用机备用件903-663后修复(By Ran) | SWO-207364 | [5 empty] | It's not turned on, there's no power. | Can't turn on, no power, borrow 903-663 power supply | Repair after replacing spare 903-663 (By Ran)
No | Visit | Complaint | RCHN | 血氧模块硬件错误 0581 | 血氧模块硬件错误 0581 ，借902-980一个。 | 更换902-980后修复(By Ran) | SWO-207476 | Defect | HW | Oxi module | Hemolyzer | Spectrum mismatch | Blood oxygen module hardware error 0581 | Blood oxygen module hardware error 0581 , borrow 902-980 one. | Replacement 902-980 repair (By Ran)
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 试剂包错误 | 更换试剂包后修复 | SWO-207828 | Assisted Replacement | SP | SP | SP |  | Test card, reagent package failure | Reagent pack error | Repair after replacement of reagent kit
Yes | Visit | Complaint | RCHN | 光学单元工厂调准失败 | 光学单元工厂调准失败，借903-209,903-198 | 机器未修复，加借备件903-236一个，此单先关单(By Ran) Have not finished this case, need one more spare parts 903-236. | SWO-207007 | Re classified RMED | HW | Process Module | Optical Unit | 1031 | Optical unit factory alignment failure | Optical unit factory registration failure, loan 903-209,903-198 | The machine has not been repaired, plus spare parts 903-236 one, this single first clearance (By Ran) Have not finished this case, need one more spare parts 903-236.
No | Visit | Complaint | RCHN | CPU单元接口松动 | CPU单元接口松动，借备件903-381一个 |  | SWO-207723 | [5 empty] | CPU unit interface loose | CPU unit interface loose, spare part 903-381 one
No | Visit | Complaint | RKK | スペクトロメーター不良にて、感度エラー | 感度レンジオーバー | スペクトロメーター交換実施 | SWO-207726 | Defect | HW | Oxi module | Spectrometer |  | Sensitivity error due to a faulty spectrometer | Sensitivity range over | Spectrometer replacement
No | Visit | Complaint | RKK | QCエラー（Na） PCB交換で改善 | NaのQC範囲外 | PCB交換 | SWO-207728 | Defect | HW | El/Met Module | PCB, El/Met | Na | QC Error (Na) IMPROVED WITH PCB REPLACEMENT | Na QC out of range | PCB switch
No | 5 | Y | RAME | F CUSTOMER HAVING OXI FAILURES DUE TO POSSIBLE NEON INTENSITY TOO LOW ERRORS. INTERMITTENTLY. |  | INTERMITTENT FAILURE OF NEAN HAD BEEN OCCURRING SINCE MID JUNE WHEN CUSTOMER CONTACTED TECHNICAL SUPPORT. THEY HAD BEEN RESETING ANALYZER AS TECH SUPPORT HAD INSTRUCTED THEM. | D2CFB10ED25A9C5B | Defect | HW | Oxi module | Lamp unit |  | F CUSTOMER HAVING OXI FAILURES DUE TO POSSIBLE NEON INTENSITY TOO LOW ERRORS. INTERMITTENTLY. |  | INTERMITTENT FAILURE OF NEAN HAD BEEN OCCURRING SINCE MID JUNE WHEN CUSTOMER CONTACTED TECHNICAL SUPPORT. THEY HAD BEEN RESETING ANALYZER AS TECH SUPPORT HAD INSTRUCTED THEM.
No | Visit | Complaint | RHUN | a készülék elég sokszor kér felszívótű tartó és oldat cserét | 06.16-tól kezdődően 1031 hibakód ~ 40 alkalommal szerepel a naplóban | felszívótű tartó cseréje, járulékos folyamatok elvégzése, QC mérés | SWO-207938 | Defect | HW | Inlet module | Inlet gasket |  | The appliance requests a replacement needle holder and solution | From 06.16, 1031 error code is logged in ~ 40 times in the log | Replacing the suction needle bracket, carrying out ancillary processes, QC measurement
No | 4 | Y | RFRA | Blank | Erreur 0581, |  | D2D094A6484DEF6A | Defect | HW | Oxi module | Hemolyzer |  | Blank | Error 0581,
Yes | 4 | Y | RFRA | Blank | Erreur de spectre sur mesure patient |  | D2D02D2A480F4B2A | Re classified RMED |  |  |  |  | Blank | Patient custom spectrum error
Yes | Visit | Complaint | RDEU | 10.07.2017, Herr Weiterer, Tel: 05121905600: Intervention erforderlich-LP erneuern. | System mit entferntem Lösungspack aus- und eingeschaltet | Lösungspack erneuert | SWO-208105 | Re classified RMED | Phone support |  |  |  | 10.07.2017, Mr. Others, Tel: 05121905600: Intervention Required-LP Renew. | System with removed solution pack switched off and switched on | Solution pack renewed
No | 4 | Y | RAME | F Customer reported tHB calibration will not pass. |  | Data logs show co-ox errors which include 581. Unit will not pass tHB calibration. Ran oximetry tests with pass results. | D2D08A4B483936D3 | [5 empty] | F Customer reported tHB calibration will not pass. |  | Data logs show co-ox errors which include 581. Unit will not pass tHB calibration. Ran oximetry tests with pass results.
No | Visit | Complaint | RCHN | 泵头断裂 | 泵头断裂，借备件902-971一个 |  | SWO-207815 | [5 empty] | The pump head is broken. | Pump head broken, spare part 902-971 one
Yes | 33 | Y | RAME | z fails initilization |  | took to westlake for technical support | D2D17AC8BE2DD172 | [5 empty] | z fails initilization |  | took to westlake for technical support
No | 34 | Y | RAME | 2 PCO2 will not pass level 1 of external QC. |  | Checked data logs and verified the problem. | D2D06169480EC18A | [5 empty] | 2 PCO2 will not pass level 1 of external QC. |  | Checked data logs and verified the problem.
No | Visit | Complaint | RCHN | Lac，pO2定标不过 | Lac定标不过 | 更换测试卡后修复 | SWO-207834 | [5 empty] | Lac, pO2 calibration is not | Lac calibration is not | Repair after replacing test card
No | Visit | Complaint | RCHN | 触摸屏故障 | 触摸屏故障，借552-039 | 更换552-039修订 by andrew | SWO-207374 | [5 empty] | Touch screen failure | Touch screen fault, borrow 552-039 | Replacement 552-039 revision by andrew
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | Lac，GLU，定标不过 | 连续更换4张测试卡后修复 | SWO-207826 | Assisted Replacement | SC | SC | Lac/Glu |  | Test card, reagent package failure | Lac, GLU, bid but | Repair after 4 consecutive test cards
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | NA,GLU,LAC定标失败 | 更换试剂包 | SWO-208304 | Assisted Replacement | SP | SP | Glu/Lac |  | Test card, reagent package failure | NA, GLU, LAC calibration failed | Replace the reagent kit
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | LAC质控超范围 | 更换测试卡后修复 | SWO-210335 | Assisted Replacement | SC | SC | Lac |  | Test card, reagent package failure | LAC QUALITY CONTROL OVER RANGE | Repair after replacing test card
Yes | Visit | Complaint | RCHN | 光学单元工厂调准失败 | 加借备件903-236一个 | 更换903-236后修复 by andrew | SWO-207701 | Re classified RMED | HW | Process Module | Optical Unit | 1031 | Optical unit factory alignment failure | Borrow spare parts 903-236 one | Repair by andrew after replacement 903-236
No | Visit | Complaint | RCHN | 样本量不足 | 样本量不足，借958-201 |  | SWO-207848 | Defect | HW | Aspiration Module | Cavro pump |  | Insufficient sample size | Insufficient sample size, 958-201
No | Visit | Complaint | RCHN | 机件损坏 | 机件故障，借636-541 |  | SWO-207842 | Defect | HW | Cover and Chassis | Cable, front | 5784 | Damage to the machine | Machine failure, borrow 636-541
No | Visit | Complaint | RIND | Intervantion required. | Intervantion required. PM Service. | Found hypocloride detected. Cleaning done. Try to calibrated but oxymetry values not comming. cheked for pump cal, LS adjust and Complete Oxymetry Test. Anlaser calibrated propelry and found all parameters well with in range. PM Service done. Calibration done. Handed over in a fine working condition. | SWO-211094 | User error | Fluidic pathway | Fluidic pathway | Hypocloride |  | Intervantion required. | Intervantion required. PM Service. | Found hypocloride detected. Cleaning done. Try to calibrated but oxymetry values not comming. cheked for pump cal, LS adjust and Complete Oxymetry Test. Anlaser calibrated propelry and found all parameters well with in range. PM Service done. Calibration done. Handed over in a fine working condition.
No | Visit | Complaint | RDEU | 10.07.2017, 08:30 Uhr, Fr. Achatz (H. Zierfuss), Tel. 06963015479: Netzteil defekt | Analysator zeigt beim Einschalten einen leicht grauen Bildschirm, aber keine Informationen auf dem Bildschirm. Display abgeklemmt, keine Verbesserung. Neues Display angesteckt, keine Änderung, Rückbau. Neues Netzteil eingesbaut, keine Veränderung. | Vmtl. CPU oder Festplatte defekt. Ersatzteil per Nacht Express bestellt. Anwender haben Analysator von Station 95-1 auf Station gebracht. Analysator weiterhin ohne Funktion. Service wenn Ersatzteile geliefert. | SWO-208109 | Revisit required | HW | DMS_PC | DMS_PC |  | 10.07.2017, 08:30, Ms. Achatz (H. Zierfuss), Tel. 06963015479: Power supply defective | Analyzer displays a slightly gray screen when powering on, but no information on the screen. Display disconnected, no improvement. New display plugged in, no change, dismantling. New power supply installed, no change. | Vmtl. CPU or hard drive defective. Spare part ordered by Night Express. Users have brought analyzer from station 95-1 to station. Analyzer still unfunctioned. Service when spare parts are delivered.
No | Visit | Complaint | RIND | Oxymetry values not coming. | Oxymetry values not OK. | Protean removal done but not OK. Oxymetry drained, OFF and removed. Flused properly and refixed by service. Analyser swithed on. Protean removal 3-4 times done. Cleaning done. Pump cal done. LS ajset done. Oxy compleate test done. Calibration done. sampling done for 4-5 samples and found all value well with in range. Analyser Serviced and calibrated properly. Handed over in a fine working condition. | SWO-211092 | Cleaned | Fluidic pathway | Oxi module | Oxi module, Cleaned |  | Oxymetry values not coming. | Oxymetry values not OK. | Protean removal done but not OK. Oxymetry drained, OFF and removed. Flused properly and refixed by service. Analyser swithed on. Protean removal 3-4 times done. Cleaning done. Pump cal done. LS ajset done. Oxy compleate test done. Calibration done. sampling done for 4-5 samples and found all value well with in range. Analyser Serviced and calibrated properly. Handed over in a fine working condition.
No | 3 | Y | RPOL | Blank | Na kalibracjach pojawia siê dryf od 30 do 50. Wymiana elektrody, klapki oraz membrany. Testy kontrolne |  | D2D30430345F489F | Defect | HW | El/Met Module | Electrode connector |  | Blank | A drift from 30 to 50 is displayed on the calibrations. Replacing electrodes, flip flops and diaphragms. Control tests
No | 4 | Y | RPOL | Blank | Aparat nie przechodzi³ Complete Oximetry Test. Próba regulacji lampy nie przynios³a efektu. Wymieniono modu³ OXI. Diagnostyka podzespo³ów. Kalibracja tHb od serwisu. Analizator sprawny. |  | D2D050AB484D2E99 | Defect | HW | Oxi module | Oxi module, Replaced |  | Blank | The camera does not pass the Complete Oximetry Test. The lamp adjustment has not resulted in an effect. Listed OXI module. Subsystem diagnostics. THb Calibration from service. Functional Analyzer.
No | Visit | Complaint | RDEU | Optimierungskalibration, obwohl SC nicht erneuert | Überprüfung des Gerätes nach Ausfall, Rücktausch von Labor in Dialyse | Gerät bei eintreffen grün & messbereit, Probemessungen durch Laborpersonal ohne Beanstandungen, Austausch Deckel. Schlauchmaterials, thb Kal durchgeführt, 5 kapilläre Patientenprobenmessung aus EDTA Blut durchgeführt: i.O. | SWO-208389 | No fault found | No fault found | No fault found | No fault found |  | Optimization calibration, although SC is not renewed | Checking the device after failure, replacing laboratory for dialysis | Device when arrived green & ready for measurement, sample measurements by laboratory personnel without complaints, replacement lid. Hose material, thb Kal performed, 5 capillary patient sample measurement performed from EDTA blood: i.O.
No | 2 | Y | RAME | Z Unable to pump solutions |  | Unable to pump solutions error. They also had issues with this SC, which they brougt from a different analyzer. I found the clip was set the wrong way causing the flow error. | D2D18C1FBE090AF0 | User error | HW | Inlet module | Inlet Clip | Turned | Z Unable to pump solutions |  | Unable to pump solutions error. They also had issues with this SC, which they brougt from a different analyzer. I found the clip was set the wrong way causing the flow error.
Yes | Visit | Complaint | RDEU | 10.07.2017, 08:21 Uhr, H. Wagner, Tel. 065194783743: Einlassdichtung springt immer ab | Einlassmodul an meine private Adresse für vor 8 Uhr, für den 10.07.2017 bestellt. E-Teil wurde an die GS 15 Bad Neunahr Ahrweiler geliefert. | In die GS gefahren und das E- Teil abgeholt. | SWO-208514 | Re classified RMED | Transport doc | Auto Inlet module | Auto Inlet module |  | 10.07.2017, 08:21, H. Wagner, Tel. 065194783743: Inlet seal always jumps off | Entrance module to my private address for before 8 a.m., ordered for 10.07.2017. E-part was delivered to GS 15 Bad Neunahr Ahrweiler. | Driven into the GS and picked up the E- part.
No | Visit | Complaint | RDEU | 10.07.2017, 08:21 Uhr, H. Wagner, Tel. 065194783743: Einlassdichtung springt immer ab | Einlassmodul ersetzt, im Service Menü alle Programme getestet. I.O. Startphase pO2 besteht die Kal. Nicht Emp. 36%. Sensorkassette ersetzt. | Startphase QC und Kalibration | SWO-208516 | Defect | HW | Auto Inlet module | Auto Inlet module |  | 10.07.2017, 08:21, H. Wagner, Tel. 065194783743: Inlet seal always jumps off | Inlet module replaced, all programs tested in the service menu. I.O. Start phase pO2 consists of the Kal. Not Emp. 36%. Sensor cartridge. | Start phase QC and calibration
No | Visit | Complaint | RDEU | Monitor dunkel | Fehlersuche --> Bildschirm überprüft --> Bildschirm im Tausch erneuert --> ohne Erfolg --> Bildschirm wieder zurück getauscht in Urzustand | Bildschirm funktioniert einwandfrei --> vermutlich CF-Karte oder CPU-Unit defekt --> keine CPU-Unit im Technikerbestand --> Analysator beim Kunden abgezogen um diesen in der GS zu reparieren!!!!! | SWO-208633 | Revisit required | HW | DMS_PC | CPU |  | Monitor dunkel | Troubleshooting --> Screen checked --> Screen replaced --> without success --> Screen returned to its original state | Screen works fine --> probably CF card or CPU unit defective --> no CPU unit in the technician's inventory --> Analyzer withdrawn from the customer to repair it in the GS!!!!!
No | Visit | Complaint | RKK | 温度エラーで測定不可 | 引上げ検証で再現せず、 | ウエットセクションPCB交換 | SWO-208385 | Defect | HW | Basic Unit | PCB, Wet section Control |  | Temperature error not measurable | Without reproducing in the raise verification, | Wet Section PCB Replacement
Yes | 6 | Y | RCZE | Blank |  |  | D2D1A5B2BE340C46 | [5 empty] | Blank
No | Visit | Complaint | RDEU | pO2 QC Verletzung Level B. | Hr. Wagner berichtet, das der ABL ausgestiegen ist. Mit Anwendern zusammen in der Datenbank nachgeschaut seit wann der QC Fehler besteht, seit dem 07.07 an dem Tag wurde der LS Pack ersetzt. Mit den Anwendern den LS Pack ersetzt. | Startphase QC und Kalibration. | SWO-208515 | Assisted Replacement | SP | SP | pO2 |  | pO2 QC Verletzung Level B. | Mr. Wagner reports that the ABL has left. With users looking together in the database since when the QC error exists, since 07.07 on the day the LS Pack has been replaced. Replace the LS Pack with the users. | Start phase QC and calibration.
Yes | 7 | Y | RCZE | Blank |  |  | D2D1A520BE3E0C42 | Re classified RMED | SWU |  |  |  | Blank
Yes | 1 | Y | RPAC | Blank | Calibration failure on Glu |  | D2D1EA3FBE18066B | [5 empty] | Blank | Calibration failure on Glu
No | Visit | Complaint | RDEU | Intervention erforderlich und Einstellungen für externen Monitor vornehmen. | LP Fehlerhaft | LP erneuert und im servicemenü alle fehlermeldungen beseitigt. Einstellungen für den Anschluss von einen Beamer vorgenommen. | SWO-208492 | Assisted Replacement | SP | SP | SP |  | intervention required and settings for external monitor. | LP Faulty | LP is renewed and all error messages are removed in the service menu. settings for connecting a beamer.
Yes | 33 | Y | RAME | Blank |  |  | D2D17A7ABE36D16F | [5 empty] | Blank
No | Visit | Complaint | RCHN | 泵头断裂 | 泵头断裂，借备件902-971一个，无货工程师用备用机备件修 |  | SWO-208195 | [5 empty] | The pump head is broken. | Pump head broken, borrowed spare parts 902-971 one, no cargo engineer with spare machine spare parts repair
No | 34 | Y | RAME | Z Customer indicated that the PCO2 will not pass on level one external QC. |  | Verified once on site. Level 1 QC consistently fails high around 71-73 mmHg. | D2D1BA5BBE0DE91C | [5 empty] | Z Customer indicated that the PCO2 will not pass on level one external QC. |  | Verified once on site. Level 1 QC consistently fails high around 71-73 mmHg.
No | Visit | Complaint | RCHN | 报错5766和45487 | 报错5766和45487，需要借件903-198一个 |  | SWO-208180 | Defect | HW | Analyzer Control Module | PCB, Main Board |  | 5766 and 45487 | 5766 and 45487, need to borrow 903-198 one
No | Visit | Complaint | RCHN | BNP温度传感器损坏 1235 | BNP温度传感器损坏 1235 | 更换905-908一个正常。 by andrew | SWO-208301 | Defect | HW | Process Module | Sensor, BNP | 1235 | BNP temperature sensor damage 1235 | BNP temperature sensor damage 1235 | Replace 905-908 one normal. by andrew
No | Visit | Complaint | RKK | OXI QC外れ。OXI Spectrummismatch。ﾍﾓﾗｲｻﾞｰ交換。pH感度ﾊﾞﾗﾂｷの為、pH/BG、AU PCBｹｰﾌﾞﾙ再接続にて改善。 | OXI QC外れ。異常値。 | ﾍﾓﾗｲｻﾞｰ交換にて改善。 | SWO-210298 | Defect | HW | Oxi module | Hemolyzer | Spectrum mismatch | OXI QC out. OXI Spectrummismatch. Hemolyzer replacement. Improved by pH/BG and AU PCB cable reconnect due to pH sensitivity variation. | OXI QC out. The outlier. | Improved by hemolizer replacement.
No | Visit | Complaint | RKK | 装置がホールドモードになっているのに、DMS側の画面が測定可能の画面のままになっている | 機器状況確認 | アナライザーコントロールPCB交換作業一式 各部動作チェックOK CALチェックOK、QCチェックOK | SWO-210778 | Defect | HW | DMS_PC | Touch screen | off/on | The device is in hold mode, but the screen on the DMS side remains measurable. | Equipment status check | Analyzer Control PCB Replacement Work Set Each part operation check OK CAL check OK, QC check OK
Yes | 9 | Y | RAME | I Numerous temp errors. |  | Numerous temp errors in activity log - Met II module mostly, but numerous others in last two weeks. | D2D208A1BE054F68 | [5 empty] | I Numerous temp errors. |  | Numerous temp errors in activity log - Met II module mostly, but numerous others in last two weeks.
No | 10 | Y | RAME | F Call continuation for lamp voltage too low errors. |  | Waiting for complete oxi module to ship to customer. | D2D17C6CBE207B96 | Defect | HW | Oxi module | Oxi module, Replaced |  | F Call continuation for lamp voltage too low errors. |  | Waiting for complete oxi module to ship to customer.
No | 4 | Y | RPOL | Blank | Wymiana ig³y, diagnostyka podzespo³ów. Kalibracje. Analizator sprawny. |  | D2D21BE7BE2986A0 | Defect | HW | Inlet module | Probe |  | Blank | Replacing the chips, diagnosing the sub-assembly. Calibrations. Functional Analyzer.
No | Visit | Complaint | RDEU | 06.07.2017 14:20Uhr Sr.Silvia 024622045091: Der Kunde hat den Ext. Scanner getestet und festgestell, dass er nicht funktioniert. | Sexterner Scanner funktioniert nicht. | Eintrag in Roma ergänzt. Danach war Scanner funktionstüchtig 2 Punkt kal durchgeführt. QCs für alle Level durchgeführt. --> i.O. ABL wieder messbereit. | SWO-208539 | Setup | SW | Roma.ini | Barcode reader, Setup |  | 06.07.2017 14:20h Sr.Silvia 024622045091: The customer has the ext. scanner tested and determined that it was not working. | External scanner does not work. | entry in Roma. After that, scanner was functional 2 point kal performed. QCs for all levels. --> i.O. ABL ready to be measured again.
Yes | 16 | Y | RESP | El cliente solicita eliminar los caracteres de inicio y fin en la lectura del código de barras en el lector de tubos. |  | Actualizar el software a la versión 8.9.ES. | D2D1AA53BE0A1CD2 | Re classified RMED | SWU |  |  |  | The client requests to remove the start and end characters in the barcode reading in the pipe reader. |  | Update the software to the 8.9.ES version.
No | 33 | Y | RAME | W Customer reported unit shut down and re-booted for no apparent reason. |  | Checked data logs for errors. They only showed shutdown and restart of analyzer. No errors logged. | D2D30FFD34045B2B | Defect | SW | SW | Reboot |  | W Customer reported unit shut down and re-booted for no apparent reason. |  | Checked data logs for errors. They only showed shutdown and restart of analyzer. No errors logged.
No | 2 | Y | RFRA | Blank |  |  | D2D33D6C3437428F | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | Blank
Yes | Visit | Complaint | RDEU | Monitor dunkel | Rückgabe des reparierten Kundengerätes --> erneuter Service fällig, wenn bestellte CPU-Unit geliefert wurde | Repariertes Kundengerät zurück gebracht, System und Funktion i.O. | SWO-208712 | Re classified RMED | HW | DMS_PC | CPU |  | Monitor dunkel | Return of the repaired customer device --> re-service due when ordered CPU unit was delivered | Repaired customer device brought back, system and function i.O.
No | 7 | Y | RAME | Z Analyzer will not print. |  | The printer will not print. It will not even advance the paper. | D2D2DE1134393CA4 | [5 empty] | Z Analyzer will not print. |  | The printer will not print. It will not even advance the paper.
No | Visit | Complaint | RCHN | 光学敏感度超范围 | 光学敏感度超范围，借903-364 | 更换903-364后修复 by andrew | SWO-208583 | Defect | HW | Process Module | Optical Unit | 1029 | Optical sensitivity over range | Optical sensitivity over range, borrowed 903-364 | Repair by andrew after replacing 903-364
No | 64 | Y | RESP | Aguja doblada. |  | Comprobar aguja y mecanismos de posición, cambiar la aguja, calibrar y probar con controles. | D2D2C1C0345D1CD7 | Defect | HW | Inlet module | Probe bend |  | Needle bent. |  | Check needle and position mechanisms, change the needle, calibrate and test with controls.
No | 29 | Y | RAME | Z Inconsistent Results |  | Customer reported that the results on one sample were inconsistent when the sample was run twice within 5 minutes. The THB,PC02 and Glucose were significantly different. | D2E3007D321207BD | Unknown | Performance | Performance | Performance | Techline ? | Z Inconsistent Results |  | Customer reported that the results on one sample were inconsistent when the sample was run twice within 5 minutes. The THB,PC02 and Glucose were significantly different.
No | 7 | Y | RAME | Z Customer changing Gas2 14-20 days, Update software |  | checked data logs for errors and replacement logs, | D2D35DD634468FB2 | Defect | HW | Gas module | Gas valve |  | Z Customer changing Gas2 14-20 days, Update software |  | checked data logs for errors and replacement logs,
No | 3 | Y | RPAC | Blank | Inlet calibration error |  | D2D30854341DC030 | Defect | HW | Inlet module | Friction detector unit |  | Blank | Inlet calibration error
No | 4 | Y | RPOL | Blank | Uszkodzony zawór OXI. Wymiana na u¿ywany. Testy kontrolne. Analizator sprawny. |  | D2D303673438489A | Defect | HW | Oxi module | Pinch valve tube |  | Blank | A faulty OXI valve. Replacement on use. Control tests. Functional Analyzer.
Yes | 14 | Y | RFRA | Blank | blocage du dispositif : erreur 4801 |  | D2D2F927341BA84D | [5 empty] | Blank | Device blocking: error 4801
Yes | 10 | Y | RFRA | Blank | investigation suite à problème de mesure de troponine (faux positif ) |  | D2D2F223342EA84A | [5 empty] | Blank | investigation following troponin measurement problem (false positive)
Yes | Visit | Complaint | RIND | AQT goes in maintenance and not allowing for any new test. | Blank cartridge cleaning. | Blank cartridge installed. Cleaning done by cleaning solution. Analyser Serviced and calibrated properly. handed over in a fine working condition. | SWO-211098 | Re classified RMED | Logistical upSWO |  |  |  | AQT goes in maintenance and not allowing for any new test. | Blank cartridge cleaning. | Blank cartridge installed. Cleaning done by cleaning solution. Analyser Serviced and calibrated properly. handed over in a fine working condition.
No | Visit | Complaint | RDEU | 11.07.2017, 08:43 Uhr, Fr. Fraas, Tel. 06062797708: Temperatur außer Bereich vor Trocknungsschritt | Analysator zeigt zwei Fehlermeldungen. Bei Patientenmessungen, ohne Ergebnis mit Fragezeichen und Meldung Temperatur außer Bereich vor Trocknungsschritt . Die zweite Meldung erscheint sporadisch unter Aktivitäten ersichtlich; 1020 Inkubationstemperatur außer Bereich . | Tausch der Incubator PCB. Ausbau der alten und Einbau der Neuen. Ausbau der Cup Dryer Unit und Tausch der Cup Dryer PCB. Beim Cup wash temerature Sensor hat sich ein Kabel gelöst, erneuert. | SWO-210249 | Defect | HW | Process Module | PCB, Cup dryer | 1068 | 11.07.2017, 08:43, Fr. Fraas, Tel. 06062797708: Temperature out of range before drying step | Analyzer displays two error messages. For patient measurements, without result with question mark and message temperature out of range before drying step. The second message appears sporadically under Activities; 1020 Incubation temperature out of range . | Replacement of the Incubator PCB. Extension of the old and installation of the new ones. Expansion of the Cup Dryer Unit and exchange of the Cup Dryer PCB. With the Cup wash temerature sensor, a cable has loosened, renewed.
No | Visit | Complaint | RLTD | Lactate locked out issue | Lactate locked out. | Followed troubleshooting advice from RMED. Reloaded system software. Lactate parameter still locked out on AQURE. Passed to IT to view AQURE device destined messages. Sensor cassette on board due to be changed due to sensitivity issue on calibrations and QCs. 3.3 MR1 software on board. | SWO-210708 | Setup | SW | RMED SW | Parameter setup |  | Lactate locked out issue | Lactate locked out. | Followed troubleshooting advice from RMED. Reloaded system software. Lactate parameter still locked out on AQURE. Passed to IT to view AQURE device destined messages. Sensor cassette on board due to be changed due to sensitivity issue on calibrations and QCs. 3.3 MR1 software on board.
Yes | 4 | Y | RAME | Z User called to get analyzer checked out before proficiency survey. |  | user wanted analyzer checked out before proficiency survey due to failure on the last one they did. | D2D45431AA2CBA26 | Re classified RMED |  |  |  |  | Z User called to get analyzer checked out before proficiency survey. |  | user wanted analyzer checked out before proficiency survey due to failure on the last one they did.
No | 3 | Y | RFRA | Blank |  |  | D2D41E6DAA1B9E7B | Defect | HW | Basic Unit | PCB, Sensor interface | LEH Mail | Blank
No | Visit | Complaint | RLTD | Lactate locked out issue | Lactate locked out | Followed troubleshooting advice from RMED. Reloaded system software. Lactate parameter still locked out on AQURE. Passed to IT to view AQURE device destined messages. 3.3 MR1 software on board. | SWO-210709 | Setup | SW | RMED SW | Parameter setup |  | Lactate locked out issue | Lactate locked out | Followed troubleshooting advice from RMED. Reloaded system software. Lactate parameter still locked out on AQURE. Passed to IT to view AQURE device destined messages. 3.3 MR1 software on board.
No | Visit | Complaint | RDEU | Monitor dunkel | Fehlersuche --> CPU-Unit defekt --> CPU-Unit erneuern | CPU-Unit erneuert, System und Funktion i.O. | SWO-210463 | Defect | HW | DMS_PC | CPU |  | Monitor dunkel | Troubleshooting --> CPU unit defective --> Renew CPU unit | CPU unit renewed, system and function in the end
No | Visit | Complaint | RBNL | Inletmodule heb vervangen | De klant ziet dat de inletgasket soms vast blijft zitten, dit gebeurt vooral nadat een capillaire meting is uitgevoerd. | Inletmodule heb vervangen. | SWO-210707 | Defect | HW | Auto Inlet module | Auto Inlet module | LEH Mail, module return (Requested 24/7-2017) | Inletmodule have replaced | The customer sees that the inletgasket sometimes remains stuck, this happens especially after a capillary measurement has been performed. | Have replaced the inletmodule.
No | 4 | Y | RPNZ | Blank | Solution pack replacement. |  | D2D41E6FAA2BF085 | [5 empty] | Blank | Solution pack replacement.
No | 4 | Y | RPNZ | Blank | Solution pack replacement due to pH QC Level 4 below range. |  | D2D41E6FAA2BF085 | [5 empty] | Blank | Solution pack replacement due to pH QC Level 4 below range.
No | 11 | Y | RAME | R Solution Pack issues |  | every month when scheduled replacement, SP leaks or is difficult to stabilize (other ABL80 has no issues. | D2D44D27AA4C41FF | [5 empty] | R Solution Pack issues |  | every month when scheduled replacement, SP leaks or is difficult to stabilize (other ABL80 has no issues.
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 参比电极错误 | 更换测试卡后修复 | SWO-210318 | Assisted Replacement | SC | SC | Reference |  | Test card, reagent package failure | Reference electrode error | Repair after replacing test card
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | GLU定标错误 | 更换测试卡后修复 | SWO-211747 | Assisted Replacement | SC | SC | Glu | 1310 | Test card, reagent package failure | GLU calibration error | Repair after replacing test card
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | LAC敏感度漂移错误，借903-135,903-165,903-183 | 仪器没问题，怀疑电源问题，未使用借件，现已修复 by andrew | SWO-208227 | Defect | HW | Basic Unit | PCB, Sensor interface |  | Test card, reagent package failure | LAC Sensitivity Drift Error, borrowed 903-135, 903-165, 903-183 | The instrument is OK, suspected of power problems, not using the debit, has been repaired by andrew
No | Visit | Complaint | RCHN | 报错1031 | 报错1031，借903-364一个，944-394（12920）1张缺货 | 更换光学单元后修复 by andrew | SWO-208302 | Defect | HW | Process Module | Optical Unit | 1031 | 1031 | 1031, loan 903-364 one, 944-394 (12920) 1 out of stock | Repair by andrew after replacing optical unit
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 光学单元超出响应范围，需要借件944-258批次号12955、944-212 批次号：12758、 944-216 批次号：12758 944-214 批次号：21841（没货） | 重新定标后修复 by andrew | SWO-208581 | Adjustment | HW | Process Module | Optical Unit |  | Test card, reagent package failure | Optical unit out of the response range, need to borrow 944-258 batch number 12955, 944-212 batch number: 12758, 944-216 batch number: 12758 944-214 Batch Number: 21841 (out of stock) | Fix by andrew after re-calibration
No | Visit | Complaint | RCHN | 光学强度超出范围 | 光学强度超出范围,需要借件903-364（缺货） |  | SWO-210553 | Defect | HW | Process Module | Optical Unit |  | Optical strength out of range | Optical strength out of range, need to borrow 903-364 (out of stock)
No | Visit | Complaint | RDEU | 13.07.2017 08:40Uhr Herr Kohl 003441860814: Das Gerät meldet Kal-fehler | Kla werden abgebrochen, Gerät steht in Intervention erforderlich | Fluidicmodul zerlegt, Ventile geprüft, Ventilmembran von V7 ersetzt, Air Anschluß im Fluidicmodul verstopft -> gereinigt unstabiles Flüssigkeitsverhalten im pH $BG Modul, Ventile in den Messmodulen gereinigt, Ventilblöcke der Module getauscht, Airventilmembran ersetzt, im pH $BG Modul Rinse Ventil ersetzt Dichtring an pH Elektrode sitzt nicht korrekt-> ersetzt, Dichtring an Ref -Elektrode doppelt -> entfernt Fehler bleibt bestehen, weitere Untersuchungen ->pH $BG Messkammer gerissen-> Ersatzteil bestellt, Info an Kunden | SWO-210703 | Revisit required | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | 07/13/2017 08:40Am Mr Kohl 003441860814: The device reports Kal error | Kla is aborted, device is required in intervention | Fluidic module disassembled, valves tested, valve membrane replaced by V7, air connection clogged in fluidic module -> cleaned unstable fluid behavior in the pH $BG module, valves cleaned in the measuring modules, valve blocks of the modules replaced, air valve membrane replaced, replaced in the pH $BG module Rinse valve Sealing ring on pH electrode sits not correctly-> replaced, sealing ring on ref electrode double -> removed Fault remains, further investigations ->pH $BG measuring chamber torn-> spare part ordered, info to customers
No | 11 | Y | RPNZ | Blank | Frequent calibraiton drift on electrolytes and Na QC unstable with high bias. |  | D2D46A9BAA3C4938 | Defect | HW | pH/BG Module | Valve |  | Blank | Frequent calibraiton drift on electrolytes and Na QC unstable with high bias.
No | Visit | Complaint | RDEU | 12.07.2017; Vanessa Braunstein; Tel.: 063120381704; Au nicht mit DMS verbunden | Nach einer telefonische Abfrage, hat man mir berichtet dass der Monitor Probleme bereitet. | Vorort habe ich festgestellt dass die Phase vom Bildschirmschoner zu normalen Bild nach dem ersten berühren der Bildschirm etwas verzögert ist. Dies ist kein Fehler den ich behebn kann. | SWO-210755 | Defect | HW | Basic Unit | PCB, Analyzer control |  | 12.07.2017; Vanessa Braunstein; Tel.: 063120381704; Au not connected to DMS | After a telephone query, I was told that the monitor is causing problems. | Suburb I noticed that the phase from the screen saver to normal image after the first touch of the screen is somewhat delayed. This is not a mistake I can fix.
No | Visit | Complaint | RDEU | 12.07.2017, Pfl. Holger, Tel: 03645952821: Kalibrationsfehler bei Laktat | Empfindlichkeitsfehler Lac , dadurch auch QC-Fehler Lösungspack wurde 2 x gewechselt, Sensorkassette wurde 3x gewechselt, wegen Lactatfehler (durch den Kunden) | Fehlersuche, Kontrolle Sensorkontakte, Reinitialisierung Sensorkassette, Funktionskontrolle | SWO-210835 | Assisted Replacement | SC | SC | Lac |  | 12.07.2017, Pfl. Holger, Tel: 03645952821: Calibration error in lactate | Sensitivity error Lac , thus also QC error Solution pack was changed 2 times, sensor cartridge was changed 3times, due to lactate error (by the customer) | Troubleshooting, control of sensor contacts, reinitialization sensor cassette, function control
No | Visit | Complaint | RCHN | 温度不稳定，屏幕不显示 | 温度不稳定，,需要借件902-972备件一个 | 更换902-972备件后修复 by andrew | SWO-210572 | [5 empty] | Temperature is unstable, screen does not display | Temperature is unstable, requires a loan of 902-972 spare parts | Repair by andrew after replacement of 902-972 spare parts
No | Visit | Complaint | RIND | Na and K value not acceptable. analyser swithed off. | Parameter reading not OK *** coming. | Zeroing calibration done. Calibration done. analyser Serviced and calibrated properly. Sapling done on both analyser and found identical readings. Inform to Dr Sharma and Dr Pathak. They ask us to monitor analyser for next 30 days on both the analyser they will run the sample and cross check if all well they will continew with us. | SWO-211099 | [5 empty] | Na and K value not acceptable. analyser swithed off. | Parameter reading not OK *** coming. | Zeroing calibration done. Calibration done. analyser Serviced and calibrated properly. Sapling done on both analyser and found identical readings. Inform to Dr Sharma and Dr Pathak. They ask us to monitor analyser for next 30 days on both the analyser they will run the sample and cross check if all well they will continew with us.
No | Visit | Complaint | RCHN | 扫描器灯不亮 | 扫描器灯不亮，需要借件903-929备件一个 | 更换903-929后修复 by andrew | SWO-210363 | Defect | HW | Storage Module | PCB, Cup transport Front |  | Scanner light is not on | Scanner light is not on, need to borrow 903-929 spare parts one | Repair by andrew after replacement 903-929
No | Visit | Complaint | RCHN | 定标不走终止 | 定标步骤终止,需要借件944-212批次号：13011各2个（缺货），13073（2张） | 重新定标正常。 by andrew | SWO-208694 | Adjustment | SW | RMED SW | off/on |  | Bid does not go to end | The calibration step is terminated and the loan 944-212 batch number: 13011 2 each (out of stock), 13073 (2 sheets) | Recalibration is normal. by andrew
No | Visit | Complaint | RCHN | 光学单元超出范围 | 光学单元超出范围，需要借件903-364备件一个 | 更换光单后修复 by andrew | SWO-208631 | Defect | HW | Process Module | Optical Unit | Temp | Optical unit out of range | Optical unit out of range, need to borrow 903-364 spare parts one | Fix by andrew after replacing a light sheet
No | Visit | Complaint | RCHN | 报1025错误 | 报1025错误，借903-236一个，903-364因缺货暂时不借 | 更换903-236一个正常。 by andrew | SWO-210103 | Defect | HW | Process Module | Optical Unit | 1025 | Reported 1025 error | Reported 1025 error, borrow903-236 one, 903-364 temporarily not borrowed due to out-of-stock | Replace the 903-236 one normal. by andrew
No | Visit | Complaint | RCHN | 光学敏感度超范围 | 光学敏感度超范围，借903-364，944-258（批号13003）/2张 |  | SWO-210808 | Defect | HW | Process Module | Optical Unit | 1029 | Optical sensitivity over range | Optical sensitivity over range, 903-364,944-258 (lot number 13003)/2 sheets
No | Visit | Complaint | RDEU | 13.07.2017 08:40Uhr Herr Kohl 003441860814: Das Gerät meldet Kal-fehler | Siehe SWO's | ABL 800 zur Rep. mitgenommen Rep-Ersatzgerät ( 754R01235N006 ) aufgebaut --> Konfig eingespielt --> QC lässt sich nicht starten --> AQC-Mechanik-Fehler--> AQC-Modul aus Kundengerät in Rep-Ersatzgerät eingebaut --> Modulupgrade--> Neustart | SWO-210988 | Revisit required | Unknown | Unknown | Unknown |  | 07/13/2017 08:40Am Mr Kohl 003441860814: The device reports Kal error | See SWO's | ABL 800 to Rep. Taken Rep Replacement Device ( 754R01235N006 ) built --> Config recorded --> QC cannot be started --> AQC-Mechanics-Error--> AQC module from customer device built into rep replacement device --> Module upgrade--> Restart
No | Visit | Complaint | RCHN | O2模块硬件无法检测 | 血氧模块硬件错误，借902-833 | 拆备用机更换902-833一个正常。 by andrew | SWO-210532 | Defect | HW | Oxi module | Oxi module, Replaced |  | O2 module hardware cannot be detected | Blood oxygen module hardware error, borrow 902-833 | Dismantling the spare machine to replace 902-833 one normal. by andrew
Yes | Visit | Complaint | RHUN | a windows elindul, de az analizátor sw már nem. Disk space full hiba üzenet jön helyette. | a C:\....\temp könyvtárba irányuló írási kísérlet nem sikerült, az analizátor szoftver ezért nem indul | lemez terület felszabadítása után az analizátor sw elindul. WIN-XPE és analizátor szoftver újra telepítése, beállítások visszaállítása SW frissítése v6.17-re. | SWO-211045 | [5 empty] | Windows starts but the Analyzer SW is no longer. Disk Space Full error message comes instead. | attempt to write to C: \ .... \ temp directory failed, therefore parser software does not start | Disk space after unlocking the analyzer SW starts. WIN-XPE and Analyzer software re-install, restore settings SW upgrade to V 6.17.
No | Visit | Complaint | RDEU | 10.07.2017, 08:30 Uhr, Fr. Achatz (H. Zierfuss), Tel. 06963015479: Netzteil defekt | Analysator ohne Funktion | CPU getauscht, Display wieder sichtbar. Neue Mac AD: 00-04-5F-9C-BC-70. Kal und QCs i.O Gerät wieder auf Station 95-3 gebracht. | SWO-211101 | Defect | HW | DMS_PC | CPU |  | 10.07.2017, 08:30, Ms. Achatz (H. Zierfuss), Tel. 06963015479: Power supply defective | Analyzer without function | CPU exchanged, display visible again. New Mac AD: 00-04-5F-9C-BC-70. Kal and QC's i.O device brought back to station 95-3.
Yes | 6 | Y | RAME | P Customer reporting error 374, inhomogeneous sample at pH/BG module. |  | Confirmed reported issue with customer. Discussed some of the probable reason and some of the areas that may be causing the problem. | D2D87E370C3C540B | Re classified RMED | Phone support |  |  |  | P Customer reporting error 374, inhomogeneous sample at pH/BG module. |  | Confirmed reported issue with customer. Discussed some of the probable reason and some of the areas that may be causing the problem.
No | 2 | Y | RAME | w software crashed roma.ini false |  | re installed software upgrade to 6.17 database is corrupt had to wipe and re install | D2D5354120310273 | Defect | SW | Roma.ini | Roma.ini |  | w software crashed roma.ini false |  | re installed software upgrade to 6.17 database is corrupt had to wipe and re install
No | 5 | Y | RAME | G Customer reported by email Glucose and Lactate failing calibrations. |  | Checked data logs for errors. Found large drift values for both parameters. Visual check good. No leakage. Pump calibration factors on low side. | D2D59C4F20232759 | Defect | HW | Inlet module | Probe bend |  | G Customer reported by email Glucose and Lactate failing calibrations. |  | Checked data logs for errors. Found large drift values for both parameters. Visual check good. No leakage. Pump calibration factors on low side.
No | 3 | Y | RCZE | Blank |  |  | D2D53BE7202441AB | Defect | HW | Storage Module | PCB, Cartridge Lock |  | Blank
No | Visit | Complaint | RCHN | HCP值不能确定，报错1223 | HCP值不能确定，报错1223，HCP值测量出错，报错1267，借903-329 | 更换903-329后修复 by andrew | SWO-210963 | Defect | HW | Aspiration Module | Needle Wash Unit | 1223 | HCP value can not be determined, error 1223 | HCP value can not be determined, error 1223, HCP value measurement error, error 1267, borrow 903-329 | Repair by andrew after replacement 903-329
No | Visit | Complaint | RDEU | 14.07.2017 14:30Uhr Herr Kurscheid 02241187361: Das Gerät hat aus dem laufenden Betrieb den Fehler Vacuumfehler gemeldet | Einlassnadel ist nicht korrekt montiert. | Einlassnadel richtig angebracht. Im Servicemenü alle Fehler beseitigt. Nach Initialisierung Fehler in der Sensorkassette. Kassette erneuert. | SWO-211531 | User error | HW | Inlet module | Probe |  | 14.07.2017 14:30Pm Mr Kurscheid 02241187361: The device reported the error vacuum error from the ongoing operation | Inlet pin is not mounted correctly. | Inlet needle correctly attached. Fixed all errors in the service menu. After initialization error in the sensor cartridge. Cassette renewed.
No | Visit | Complaint | RIND | Lac parameter is not active. | Checked SC data and found that sensitivity of lac parameter was higher that required. | I had run liquid heparin in SC and then kept in fridge for 30 minutes. Then reinstalled SC but still problem is there.Need to replace SC. | SWO-212851 | [5 empty] | Lac parameter is not active. | Checked SC data and found that sensitivity of lac parameter was higher that required. | I had run liquid heparin in SC and then kept in fridge for 30 minutes. Then reinstalled SC but still problem is there. Need to replace SC.
No | Visit | Complaint | RCHN | lac不稳定 | lac不稳定,需要借件902-972备件一个 |  | SWO-211217 | [5 empty] | lac unstable | lac unstable, need to borrow 902-972 spare parts one
No | Visit | Complaint | RCHN | 接口板腐蚀 | 接口板腐蚀，需要借件902-972备件一个 |  | SWO-211215 | [5 empty] | Interface plate corrosion | Interface plate corrosion, need to borrow 902-972 spare parts one
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | Lac定标不过 | 更换测试卡后修复 | SWO-211324 | [5 empty] | Test card, reagent package failure | Lac calibration is not | Repair after replacing test card
No | Visit | Complaint | RCHN | 光学强度超范围 | 光学强度超范围，借903-364 |  | SWO-211344 | Defect | HW | Process Module | Optical Unit | 1024 | Optical strength over range | Optical strength over range, borrowed 903-364
No | 9 | Y | RAME | 2 PCO2 fails QC |  | Checked data logs for any previous issues. PCO2 on level2 QC is on the high range of packet insert. Cleaned PH and PCO2 electrode connectors and chambers. | D2D92EC0823F6A4A | Defect | Electrode | Electrode | Sealing ring | Sealing ring, pH | 2 PCO2 fails QC |  | Checked data logs for any previous issues. PCO2 on level2 QC is on the high range of packet insert. Cleaned PH and PCO2 electrode connectors and chambers.
No | 3 | Y | RAME | 9 Lactate results on samples low then other 3 analyzers |  | checked electrode, electrode membrane and chamber | D2D8E5A782338FC7 | Assisted Replacement | Electrode | Electrode | Lac |  | 9 Lactate results on samples low then other 3 analyzers |  | checked electrode, electrode membrane and chamber
No | Visit | Complaint | RLTD | Original case (CAS-233765-Z6D2) was incorrect serial number (090R0812N0004) . This serial number which is showing as being in cardiac theatre is physically in ICU. | Analyser is not accepting logon barcodes. On checking the operators list, there appears to be only a couple of operators. Able to log on with radiometer passcode. | On arrival i checked the serial number of analyser, i seen that it did not match what was recorded in CRM. I asked customer to log on to AQURE and show me which analyser they had tried to push operators onto. This also did not match with the physical anaylyser. I then checked all ABL90 located at the Infirmary and found that the serial number belonged to cardiac theatres. I pushed operators over to this analyser and the continued to reboot. After warm up the operators were still not showing in the list. A full windows and application software re-install was carried out. After re-loading data and setup problem had still not been resolved. After further investigation, it was found that the analyser in the ward did not match the department allocated on AQURE. Reset to correct department on AQURE and then rep usher operators. Checked on analyser and all operators now appeared as they should. Tested logon with nurse I'd and this now works correctly. Informed customer of issue. No further action required. | SWO-211862 | Setup | SW | His/lis | Aqure |  | Original case (CAS-233765-Z6D2) was incorrect serial number (090R0812N0004) . This serial number which is showing as being in cardiac theatre is physically in ICU. | Analyser is not accepting logon barcodes. On checking the operators list, there appears to be only a couple of operators. Able to log on with radiometer passcode. | On arrival i checked the serial number of analyser, i seen that it did not match what was recorded in CRM. I asked customer to log on to AQURE and show me which analyser they had tried to push operators onto. This also did not match with the physical anaylyser. I then checked all ABL90 located at the Infirmary and found that the serial number belonged to cardiac theatres. I pushed operators over to this analyser and the continued to reboot. After warm up the operators were still not showing in the list. A full windows and application software re-install was carried out. After re-loading data and setup problem had still not been resolved. After further investigation, it was found that the analyser in the ward did not match the department allocated on AQURE. Reset to correct department on AQURE and then rep usher operators. Checked on analyser and all operators now appeared as they should. Tested logon with nurse I'd and this now works correctly. Informed customer of issue. No further action required.
Yes | Visit | Complaint | RDEU | Bei der Auslieferung wurde ein fast abgelaufenes LP mit geliefert | Kunde benötigt ein neues LP | LP ausgeliefert | SWO-212480 | Re classified RMED | Logistical upSWO |  |  |  | At the time of delivery, an almost expired LP was delivered with | Customer needs a new LP | LP delivered
No | 5 | Y | RAME | 2 pCO2 failing calibrations and QC |  | customer replaced solution pack and sensor cassette last friday due to PCO2 failing ; they ran cal and QC Monday and pCO2 is failing external QC running high outside the range; | D2DB79BB6E589561 | Assisted Replacement | SP | SP | Leakage | Techline RAME *4075 | 2 pCO2 failing calibrations and QC |  | customer replaced solution pack and sensor cassette last friday due to PCO2 failing ; they ran cal and QC Monday and pCO2 is failing external QC running high outside the range;
No | Visit | Complaint | RLTD | Unit blocked unable to run samples Lynne 07909002861 | Unit blocked unable to pump solution . | Blood clot found in cassette module at the inlet end. Flushed with protein remover and water clot removed. Flushed all the system with water ok. Run service test all passed ok. | SWO-211854 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway | Clot | Unit blocked unable to run samples Lynne 07909002861 | Unit blocked unable to pump solution . | Blood clot found in cassette module at the inlet end. Flushed with protein remover and water clot removed. Flushed all the system with water ok. Run service test all passed ok.
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | LAC超限度，分析仪测卡故障 | 更换946-005后修复 | SWO-212526 | Revisit required | SC | SC | Lac |  | Test card, reagent package failure | LAC over-limit, analyzer card failure | Repair after replacement 946-005
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | pO2定标不过 | 更换944-157后修复 | SWO-212527 | Assisted Replacement | SC | SC | pO2 |  | Test card, reagent package failure | pO2 calibration but | Repair after replacement 944-157
Yes | 4 | Y | RPAC | Blank | Travelled to site and replaced item. |  | D2DA5402F82A1461 | [5 empty] | Blank | Travelled to site and replaced item.
No | Visit | Complaint | RCHN | 报错 1031，1024 | 报错 1031，1024，借903-364一个。 | 更换903-364后修复 by jesson | SWO-211377 | Defect | HW | Process Module | Optical Unit | 1031 | Error 1031,1024 | Error 1031,1024, borrowed 903-364 one. | Repair by jesson after replacing 903-364
No | Visit | Complaint | RCHN | 固定样本的齿轮断了 | 固定样本的齿轮断了，借备件903-584一个 |  | SWO-211794 | Defect | HW | Aspiration Module | Tube holder |  | The gear for the fixed sample is broken. | The gear for the fixed sample is broken, with spare parts 903-584 one
No | 8 | Y | RPAC | Blank | Assymetry errors intermittently occurring |  | D2DA83B9F83A02BE | Defect | HW | El/Met Module | Pump Tubing | Assymetri | Blank | Assymetry errors intermittently occurring
No | Visit | Complaint | RDEU | 13.07.2017 08:40Uhr Herr Kohl 003441860814: Das Gerät meldet Kal-fehler | Rücklieferung des Gerätes nach Rep. In GS | Gerät zurückgeliefert, AC Modul zurückgetauscht, pO2 Fehler bei QC Messung -> pO2 Elektrode neu membraniert | SWO-211756 | Defect | HW | Autocheck module | Autocheck module |  | 07/13/2017 08:40Am Mr Kohl 003441860814: The device reports Kal error | Return of the device to Rep. In GS | Device returned, AC module replaced, pO2 error in QC measurement -> pO2 electrode re-membraneed
No | 12 | Y | RPNZ | Blank | Follow up from SWO: BAEL5HID8 calibrations drifting on electrolytes, Na QC unstable. |  | D2DD3FD4E40AB10C | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | Blank | Follow up from SWO: BAEL5HID8 calibrations drifting on electrolytes, Na QC unstable.
No | 6 | Y | RAME | Z Customer reports that QC ampoules are not aspirated from the AutoCheck carousel. |  | Ampoules are scanned and correctly identified. Ampoules are also cracked but solution is not aspirated. | D2DA7990F843EACE | Adjustment | HW | Autocheck module | Carroussel |  | Z Customer reports that QC ampoules are not aspirated from the AutoCheck carousel. |  | Ampoules are scanned and correctly identified. Ampoules are also cracked but solution is not aspirated.
No | 6 | Y | RFRA | Blank | CQ Radiometer et non Radiometer hors gamme |  | D2DAF1B4F844B9C2 | Adjustment | Consumables | Test Kit | TnT |  | Blank | CQ Radiometer et non Radiometer hors gamme
No | Visit | Complaint | RDEU | Rotate Storage Wheel 24 | Analysator steht im Fehler Set Save 22 Rotate Storage Wheel 24 | Storage Wheel demontiert - Testkassettenabbruch entfernt - Schwarze Testkassettenhalterung fixiert (aus Verankerung einseitig gelöst) - Erneuerung Lager Storage Wheel - Allgemeine Überprüfung und Reinigung - Erneuerung Pumpenkopf 2 (Geräuchentwicklung) - Funktionsüberprüfung etc. ohne Beanstandung ! | SWO-211945 | Defect | HW | Storage Module | Cassette guide | 5784 | Rotate Storage Wheel 24 | Analyzer is in error Set Save 22 Rotate Storage Wheel 24 | Storage Wheel disassembled - Test cartridge abort removed - Black test cassette holder fixed (detached from anchorage on one side) - Renewal Bearing Storage Wheel - General inspection and cleaning - Renewal pump head 2 (smoke development) - Functional check etc. without complaint !
No | Visit | Complaint | RDEU | Ableitströme im LP erkannt | System überprüft => Sensorkassette und Bügelventil Schlauch defekt , Ventil schwergänig | Sensorkassette und Schlauch erneuert, Ventil gangbar gemacht,Kalibration und QC durchgeführt | SWO-212128 | Defect | HW | Oxi module | Pinch valve tube |  | Discharge currents detected in LP | System checked => Sensor cassette and handle valve hose defective, valve heavy-duty | Sensor cassette and hose renewed, valve made workable,calibration and QC performed
No | 9 | Y | RAME | H error 521 PO2 inhomogeneous errors on qc |  | Issue with PO2 giving erratic results and error 521 on external qc only | D2DBA60B6E00F84B | Revisit required | Fluidic pathway | Fluidic pathway | Wetability | pO2 | H error 521 PO2 inhomogeneous errors on qc |  | Issue with PO2 giving erratic results and error 521 on external qc only
No | Visit | Complaint | RDEU | Lactat-Fehler nach mehrfachem SC Wechsel bei Kal. Empf. mit .... | Aktivitätendatei analysiert. Nach SC Wechsel Lot.1067 Fehlermeldung in Kalibrierung bei Lactat vorhanden. | Kontakte für Sensorkassette überprüft. Sensorkassette aus GS17 -bestand ( auch Lot. 1067 ) erneuert. Kein Erfolg, Lactat nach mehreren Kal. Fehlerhaft, kein Empf. Wert. Wetsektion aus GS17 -Gerät testweise ausgetauscht, kein Erfolg , Lactat Kal. -Fehler Keine Empfindlichkeit. Inzwischen wurden vom Kunden und mir 4 SC mit Lot. 1067 getauscht, alle mit Lac. und gelegentlich Glu. Kal. Fehler. Wetsektion wieder zurückgetauscht und Gebrauchte SC aus GS17 -Gerät ( allerdings mit pCO2 Fehler aber guten Glu. und Lac Werten ) in Kunden Gerät eingesetzt. Alle Kal. Werte bis auf pCO2 o.k. . Mehrere Kal. durchgeführt. Neue SC aus Kundenbestand mit Lot. 1076 eingesetzt. Jetzt wieder alle Kal. Werte vorhanden und QC im Kontrollbereich , Glu. und Lac . Kal. allerdings noch nicht ganz auf Sollwert. Einlaufphase über Nacht muß abgewartet werden , Kunde wird sich bei weiteren Problemen melden. | SWO-212459 | Assisted Replacement | SC | SC | Lac | Sensor Interface PCB ? | Lactate error after multiple SC change at Kal. Empf. with.... | Activity file analyzed. After SC change lot.1067 error message in calibration present at lactate. | Checked contacts for sensor cartridge. Sensor cassette made of GS17 -existing ( also lot. 1067 ) renewed. No success, lactate after several kal. Faulty, no empf. sth means a lot to sb. Wetsection from GS17 device tested, no success, Lactate Kal. -Error No sensitivity. Meanwhile, the customer and me were given 4 SC with lot. 1067 exchanged, all with Lac. and occasionally Glu. Kal. Error. Wetsection exchanged again and used SC from GS17 device (but with pCO2 error but good glu. and lac values) used in customer device. All Kal. values except pCO2 ok. . Several Kal. Performed. New SC from customer base with lot. 1076. Now again all Kal. values present and QC in the control area, Glu. and Lac . Kal. not quite on target value yet. Run-in phase overnight must be waited for, customer will contact in case of further problems.
No | Visit | Complaint | RCHN | 气流超范围 | 气流超范围，借备件903-111一个 | 更换903-111后修复 | SWO-211729 | [5 empty] | Airflow over range | Airflow over range, spare parts 903-111 one | Repair after replacement 903-111
Yes | Visit | Complaint | RCHN | 报错1999 | 报错1999，借903-113 | 机器未修复，加借备件903-034,848-037各一个 by jesson | SWO-211764 | Re classified RMED | Double entry |  |  | 1999 | Error 1999 | Error 1999, loan 903-113 | Machine not repaired, plus spare parts 903-034, 848-037 each by jesson
No | Visit | Complaint | RCHN | 报错1999 | 报错1999，借备件903-198，636-541一个 | 机器未修复，加借备件903-034,848-037各一个，此单先关单处理 by jesson | SWO-211819 | Revisit required | HW | Aspiration Module | PCB, Needle | 1999 | Error 1999 | Error 1999, borrowed spare parts 903-198, 636-541 one | The machine has not been repaired, with spare parts 903-034, 848-037 each, this single first processing by jesson
Yes | Visit | Complaint | RCHN | 报错1999 | 机器未修复，加借备件903-034,848-037各一个，848-037广办自提 |  | SWO-212032 | Re classified RMED | Double entry |  |  | 1999 | Error 1999 | The machine has not been repaired, with spare parts 903-034, 848-037 each, 848-037
Yes | Visit | Complaint | RCHN | 报错1999 | 机器未修复，加借备件903-034,848-037各一个，848-037广办自提 |  | SWO-212033 | Re classified RMED | Double entry |  |  | 1999 | Error 1999 | The machine has not been repaired, with spare parts 903-034, 848-037 each, 848-037
No | Visit | Complaint | RDEU | Demokey wird am 21.4.2017 ablaufen | Entgültige Keys installieren | Keys installiert | SWO-212324 | Setup | SW | RMED SW | Demo Key exp |  | Demokey will expire on April 21, 2017 | Installing Final Keys | Keys installed
Yes | 21 | Y | RAME | P Intervention required, failing 2pt Cals, |  | checked data logs for errors, checked valves and liquid sensors, | D2DBDD4C6E2F8FE1 | [5 empty] | P Intervention required, failing 2pt Cals, |  | checked data logs for errors, checked valves and liquid sensors,
No | 29 | Y | RAME | V Sample errors and barcode not reading |  | Upon arrival, unit was all green and operating properly. Could not replicate barcode error. Customer said she was trying to swipe her badge and no red light was emitting at all. | D2E306CD325C07C6 | Cleaned | HW | Autocheck module | Barcode reader, Autocheck Module |  | V Sample errors and barcode not reading |  | Upon arrival, unit was all green and operating properly. Could not replicate barcode error. Customer said she was trying to swipe her badge and no red light was emitting at all.
No | Visit | Complaint | RDEU | 18.07.2017; Hr. Roth; Tel.: 04421891526; Fehler Dunkelstrommessung | Fehler 1035, Dunkelstrommessung gescheitert | Gerät überprüft, Gerät ausgesaugt, optische Einheit und Referenzflächen gesäubert, optische Einheit und Processwheel ausgetauscht, 2 Parameter kalibriert, QC o.k. | SWO-212315 | Defect | HW | Process Module | Optical Unit | 1035 | 18.07.2017; Mr. Roth; Tel.: 04421891526; Error Dark current measurement | Error 1035, Dark current measurement failed | Device checked, device sucked out, optical unit and reference surfaces cleaned, optical unit and processwheel replaced, 2 parameters calibrated, QC ok.
Yes | 18 | Y | RPAC | Blank | Multiple parameter failure post sensor cassette replacement. |  | D2DD2307E454BB6D | [5 empty] | Blank | Multiple parameter failure post sensor cassette replacement.
Yes | 8 | Y | RAME | Z When they replace a SC the Glucose gets activated |  | They do not report Glucose, but when a new SC is installed it gets activated. | D2DCF8D6E4232E0C | [5 empty] | Z When they replace a SC the Glucose gets activated |  | They do not report Glucose, but when a new SC is installed it gets activated.
No | 4 | Y | RAME | P FLUID FLOW |  | Client advised of leak detection issues not able to get analyzer to complete start up sequence. | D2DCDD71E431820A | Assisted Replacement | SC | SC | Leakage |  | P FLUID FLOW |  | Client advised of leak detection issues not able to get analyzer to complete start up sequence.
No | Visit | Complaint | RDEU | 17.07.2017 10:45 H. Baldini 08517205301: SC-Probleme Lac 18.07.2017; 13:45 Uhr; Hr. Baldini; Tel.: 08517205301; weiterhin Lac Kal Fehler 19.07.2017, 10:20 Uhr, H. Baldini, Tel. 085172054302:Weiter Probleme mit Lac. Kal wird abgebrochen, Keine QC möglich | Kalibrationen Lac fehlerhaft. | Sensor Interface Board und Ventilschlauch getauscht. Lac wieder aktiviert. SC (GS Bestand) und LP (Kundenbestand) erneuert. SC (aus Kundenbestand) wegen Kal Fehler Glu erneuert. Kalibrationen und QC's vorgenommen. | SWO-212749 | Defect | HW | Basic Unit | PCB, Sensor interface |  | 07/17/2017 10:45 H. Baldini 08517205301: SC Problems Lac 18.07.2017; 1:45 p.m.; Mr. Baldini; Tel.: 08517205301; Continue Lac Kal Error 19.07.2017, 10:20 a.m., H. Baldini, tel. 085172054302:Further problems with Lac. Kal is canceled, No QC possible | Calibrationlac faulty. | Sensor interface board and valve hose. Lac reactivated. SC (GS Stock) and LP (Customer Stock). SC (from customer base) renewed due to Kal Error Glu. calibrations and QC's.
No | 5 | Y | RPAC | Blank | Consumable failure - leaking solution pack |  | D2DE13C55A303638 | [5 empty] | Blank | Consumable failure - leaking solution pack
No | 8 | Y | RAME | Z Not reading solution pack x3 |  | Installed 3 solution packs could not reproduced the problem. Also customer request to check analyzer due to pO2 4 $5 unacceptable results on API results. | D2DCCC3FE41A27E2 | [5 empty] | Z Not reading solution pack x3 |  | Installed 3 solution packs could not reproduced the problem. Also customer request to check analyzer due to pO2 4 $5 unacceptable results on API results.
No | Visit | Complaint | RCHN | 质控超范围 | 质控超范围，借903-183 | 工程师序列号报错 by jesson | SWO-212431 | Defect | HW | Basic Unit | PCB, Sensor interface | 1397 : Replace SP ????? | Quality control over range | Quality control over range, borrowed 903-183 | Engineer serial number error by jesson
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 仪器报错卡盘 | 测试卡断裂，更换测试卡后修复 | SWO-212415 | Defect | Consumables | Test Kit | Cartridge | Broken | Test card, reagent package failure | Instrument error card | The test card breaks and is repaired after replacing the test card
No | 8 | Y | RPAC | Blank | Flex Q mixer not working, mixer unit on backorder, fitted new Flex Q tray First flex Q R1345N003 was out of box failure |  | D2DCC52BE4454306 | Revisit required | HW | FlexQ Module | FlexQ tray |  | Blank | Flex Q mixer not working, mixer unit on backorder, fitted new Flex Q tray First flex Q R1345N003 was out of box failure
No | Visit | Complaint | RKK | Lac Cal Error | Lac Cal Error | Lacの応答不良のエラーがメンブランを交換しても改善されず。電極及びコネクター交換実施。 | SWO-214837 | Assisted Replacement | Electrode | Electrode | Lac | 379 | Lac Cal Error | Lac Cal Error | Lac's poor response error does not improve when replacing the membrane. Electrode and connector exchange was implemented.
No | Visit | Complaint | RDEU | 17.07.2017 13:30 H. Diller 02345093116: Kalfehler Na, Ka, Lac 19.07.2017, Herr Breitmoser, Tel: 02345092016: weiterhin Kalibrationsfehler bei Na | Sporadisch Elektrolyt und Laktat Driften | Funktion überprüft, Messkammer gereinigt, O-Ring der pH Elektrode (klebt in der Messkammer ) und O-Ring an der Glucoseelektrode ersetzt. | SWO-212653 | Cleaned | Fluidic pathway | El/Met Module | Sealing ring | Sealing Ring left in Meas ch | 17.07.2017 13:30 H. Diller 02345093116: Cal error Na, Ka, Lac 19.07.2017, Mr. Breitmoser, Tel: 02345092016: continued calibration error at Na | Sporadic Electrolyte and Lactate Drifting | Function checked, measuring chamber cleaned, O-ring of the pH electrode (sticks in the measuring chamber) and O-ring replaced at the glucose electrode.
No | Visit | Complaint | RIND | Neon intensity too low. | neon intesity too low. neon position changed. | required new neon lamp. | SWO-212788 | Defect | HW | Oxi module | Lamp unit | Neon lamp | Neon intensity too low. | neon intesity too low. neon position changed. | required new neon lamp.
Yes | 12 | Y | RFRA | Blank |  |  | D2DCE951E44B782E | [5 empty] | Blank
No | 12 | Y | RAME | 3 pO2 |  | pO2 failing level 1 QC (S7730) | D2DE69F15A4E4E23 | Revisit required | HW | Inlet module | Connector gasket | pO2 | 3 pO2 |  | pO2 failing level 1 QC (S7730)
No | 12 | Y | RAME | H ERROR 521 on qc samples |  | unit is giveing inhomgeneous sample errors on qc internal qc and patient samples passing | D2DE28395A076027 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway | pO2 | H ERROR 521 on qc samples |  | unit is giveing inhomgeneous sample errors on qc internal qc and patient samples passing
No | Visit | Complaint | RDEU | 20.07.2017 13:15Uhr Frau Gerke 05622/9977758 Einlass gesperrt 20.07.2017, Sr. Ina, Tel: 056229977758: man. Spülung nicht möglich | Gerät verstopft | Fließwege, Einlassnadel und Schläuche verstopft. Nachdem am 15.07.2017 um 5:45 eine Messung abgebrochen wurde, zeigte das Gerät den Fehler 1190: Einlass geöffnet. Danach stand das Gerät eine Woche, da die Stationen wegen eines Brandes geschlossen wurden. Sehr wahrscheinlich wurde die Messung fehlerhaft durchgeführt und der Einlass schloss nicht mehr richtig, sodass das Blut nicht wieder herausgespült werden konnte. Gerät überprüft. Qualitätskontrollwerte schon im Bereich, aber pCo2 Empf. noch zu niedrig. Kunde lässt die vorhandene Kassette im Gerät und wird sich am Montag melden, ob die Werte in Ordnung sind. | SWO-213095 | Defect | HW | Auto Inlet module | Inlet Position Cal. |  | 20.07.2017 13:15pm Mrs Gerke 05622/9977758 Entrance blocked 20.07.2017, Sr. Ina, Tel: 056229977758: man. Flushing not possible | Device clogged | flow paths, inlet needle and hoses. After a measurement was canceled at 5:45 a.m. on Jul 15, 2017, the device showed error 1190: Inlet open. After that, the device stood for a week as the stations were closed due to a fire. Most likely, the measurement was performed incorrectly and the inlet did not close properly, so that the blood could not be flushed out again. device checked. Quality control values already in the range, but pCo2 Empf. still too low. Customer leaves the existing cassette in the device and will report on Monday whether the values are okay.
No | Visit | Complaint | RDEU | Patienten Barcode wird nicht erkannt. | Barcode Konfiguration | Gerät konfiguriert und getestet Barcode wird erkannt. Verbindung zum Radiance ist OK. Für die Entschlüsselung der Patientendaten durch die Darstellung vom Namen und Geburtsdatum ist eine Rücksprache mit Herrn Depta erforderlich. | SWO-214556 | [5 empty] | Patient barcode is not recognized. | Barcode configuration | Device configured and tested barcode is detected. Connection to the radiance is OK. A consultation with Mr. Depta is required to decrypt the patient data by presenting the name and date of birth.
No | Visit | Complaint | RCHN | 无法阅读测试卡智能芯片 | 无法阅读测试卡智能芯片 ，借902-972一个。 | 更换902-972备件后修复by jesson | SWO-212161 | [5 empty] | Can't read test card smart chip | Can't read test card smart chip , borrowed 902-972 one. | Repair by jesson after replacement of 902-972 spare parts
No | 20 | Y | RAME | 2 pco2 out high on level 1 qc. Reading 174 s/b between 141-169. |  | Investigation was to inspect Sc,SP for damage or leaking and none was found. | D2DE1C095A041183 | [5 empty] | 2 pco2 out high on level 1 qc. Reading 174 s/b between 141-169. |  | Investigation was to inspect Sc,SP for damage or leaking and none was found.
No | 2 | Y | RAME | 3 R PO2 out on linearity. |  | Tried replacing sensor cassette. That didnt work. Tried Solution pack. | D2DE37BE5A3EC4AB | [5 empty] | 3 R PO2 out on linearity. |  | Tried replacing sensor cassette. That didnt work. Tried Solution pack.
No | Visit | Complaint | RCHN | 浑浊度过高 | 浑浊度过高，借902-980 | 更换902-980后修复 by jesson | SWO-212492 | Defect | HW | Oxi module | Hemolyzer |  | Turbidity over high | Turbidity over high, loan 902-980 | Repair by jesson after replacement 902-980
No | Visit | Complaint | RCHN | lac定标状态超出限度范围 | lac定标状态超出限度范围，需要借件903-183备件一个 | 更换903-183(903-581)修复by jesson | SWO-212433 | Defect | HW | Basic Unit | PCB, Sensor interface |  | lac calibration state out of limit | Lac calibration status is out of limit and requires a spare part of 903-183 | Replace 903-183 (903-581) Repair by jesson
No | Visit | Complaint | RCHN | 光学强度超范围 ，定标失败 1031，1025 | 光学强度超范围 ，定标失败 1031，1025，借903-364，903-236各一个。 | 更换903-364清洁处理盘灰尘，定标仍然出现1024.1025. 更换903-236后修复 by jesson | SWO-212505 | Defect | HW | Process Module | Optical Unit | 1031 | Optical strength over range, calibration failure 1031,1025 | Optical strength over range, calibration failure 1031, 1025, borrow 903-364, 903-236 each. | Replace the 903-364 cleaning processing disc dust, the calibration still appears 1024.1025. Repair by jesson after replacement 903-236
No | Visit | Complaint | RCHN | 报1999错误 | 报1999错误,需要借件867-631一个 | 867-631 损坏，使用备用机867-631配件修复 by jesson | SWO-212512 | Defect | HW | Aspiration Module | Timing belt | 1999 | Reported 1999 error | reported 1999 error, need to borrow 867-631 one | 867-631 Damage, repair by jesson with spare machine 867-631 accessories
No | Visit | Complaint | RCHN | 重新定标 | 重新定标,需要借件944-394（批次号：12450） | 升级至8.9以后，发现PCT定标结果丢失。重新定标后修复 by jesson | SWO-212405 | User error | SW | Anlyzer SW | No copy of DB |  | Re-calibration | Re-bid, borrowing 944-394 (batch number: 12450) | After upgrading to 8.9, it was found that the PCT calibration results were missing. Fix by jesson after re-calibration
No | Visit | Complaint | RCHN | 定标中止 3557 | 定标中止 3557，借944-258（12301）一张。 | 清洁内部灰尘 。重新定标后修复 by jesson | SWO-212419 | Cleaned | HW | Incubator Module | PCB, Incubator | 3557 | Bid Abort 3557 | The bid was suspended 3557, with a loan of 944-258 (12301). | Clean the internal dust. Fix by jesson after re-calibration
Yes | 1 | Y | RFRA | Blank | Difficulté à démarrer - erreurs wifi |  | D2DE28E35A2A8F5C | Re classified RMED | Double entry | See D2DE28E35A2A8F5C |  |  | Blank | Difficulty getting started - wifi errors
Yes | 1 | Y | RFRA | Blank | Difficulté à démarrer - erreurs wifi |  | D2DE28E35A2A8F5C | [5 empty] | Blank | Difficulty getting started - wifi errors
Yes | Visit | Complaint | RHUN | a windows elindul, de az analizátor sw már nem. Disk space full hiba üzenet jön helyette. | a készülék továbbra is sokszor újraindul, mostanra már az OXI modul inaktív marad az újraindulások után | csere készülék beüzemelése, eredeti készülék elszállítása Radiometer MO. Kft telephelyére szerviz és tesztelés céljából | SWO-213034 | [5 empty] | Windows starts but the Analyzer SW is no longer. Disk Space Full error message comes instead. | The device continues to restart many times, now the OXY module remains inactive after resumption | Installation of the replacement apparatus, removal of the original equipment to Radiometer MO. Kft. For service and testing purposes
No | Visit | Complaint | RDEU | 21.07.2017 08:30Uhr Herr Ostertag 076223951247: Das Gerät zeigt nach dem Einschalten: Bitte Boot-CD einsetzen | Gerät bootet nicht. Im BIOS wird HDD zwar erkannt, aber Parameter nicht eingelesen. | Netzteil erneuert. Identischer Fehler. Ersatzteile bestellt. Kunde (Dr. Meine) wünscht Samstagseinsatz. | SWO-213163 | Revisit required | HW | DMS_PC | HDD |  | 21.07.2017 08:30Am Mr Ostertag 076223951247: The device shows after switching on: Please insert boot CD | Device does not boot. HdD is detected in the BIOS, but parameters are not read in. | power supply is renewed. Identical error. spare parts ordered. Customer (Dr. Meine) wishes Saturday use.
Yes | 15 | Y | RFRA | Blank |  |  | D2DE04B75A3D7841 | [5 empty] | Blank
No | Visit | Complaint | RHUN | oldat csere néhány napja volt, most felhasználói beavatkozást kér | pressure teszt vákum érték általában > 0,75 (0,25-0,75) | hiba tesztelése során kiderül, hogy a felszívó tű tartó okozza a hibát, alaktrészek cseréje, kalibráció, QC mérés | SWO-213223 | Defect | HW | Inlet module | Inlet gasket |  | Replacement solution was a few days ago, it now asks for user intervention | Pressure test vacuum value usually > 0.75 (0.25-0.75) | Fault testing reveals that the absorption needle holder is causing the defect, Replacement of the shape, calibration, QC measurement
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 参比电极错误，1361 | 更换测试卡后修复 | SWO-213232 | [5 empty] | Test card, reagent package failure | Reference electrode error, 1361 | Repair after replacing test card
Yes | Visit | Complaint | RDEU | 21.07.2017 08:30Uhr Herr Ostertag 076223951247: Das Gerät zeigt nach dem Einschalten: Bitte Boot-CD einsetzen | Gerät bootet nicht | E-Teil nicht verfügbar nach Rücksprache mit Lieferdienst, Anlieferung erfolgt Montagmorgen Herr Dr. Kammerer informiert und Service für Montag abgesprochen | SWO-213244 | Re classified RMED | For Comm doc |  |  |  | 21.07.2017 08:30Am Mr Ostertag 076223951247: The device shows after switching on: Please insert boot CD | Device does not boot | E-part not available after consultation with delivery service, delivery takes place On Monday morning Dr. Kammerer informs and service agreed for Monday
No | Visit | Complaint | RIND | The calibration and the patient sample measurement aborted and showing the message optical intensity out of range. | The calibration and the patient sample measurement aborted and showing the message optical intensity out of range. 1025:Calibration adjustment analyte outliers failed Error Code:1031,1025 | Has to change the optical unit and process wheel. part number: 903-364-Optical Unit 903-236:process wheel | SWO-213271 | Defect | HW | Process Module | Optical Unit | 1031 | The calibration and the patient sample measurement aborted and showing the message optical intensity out of range. | The calibration and the patient sample measurement aborted and showing the message optical intensity out of range. 1025:Calibration adjustment analyte outliers failed Error Code:1031,1025 | Has to change the optical unit and process wheel. part number: 903-364-Optical Unit 903-236:process wheel
No | Visit | Complaint | RIND | continous failure of cassette. | sensor board was faulty. | arranged new analyzer standby at hospital installed cassette successfully. require sensor board. | SWO-214384 | [5 empty] | continous failure of cassette. | sensor board was faulty. | arranged new analyzer standby at hospital installed cassette successfully. require sensor board.
No | Visit | Complaint | RSAF | Analyzer keeps on giving pressure test failed and pump calibration error Tshepo gave telephonic assistance and myself but could not get analyzer corrected. | Entered service mode and done trouble shooting, Fault found to be blocked inlet probe. | Unblocked probe and done complete Fluid transport test and all passed. | SWO-214924 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway | Clot | Analyzer keeps on giving pressure test failed and pump calibration error Tshepo gave telephonic assistance and myself but could not get analyzer corrected. | Entered service mode and done trouble shooting, Fault found to be blocked inlet probe. | Unblocked probe and done complete Fluid transport test and all passed.
No | Visit | Complaint | RDEU | 19.07.2017 10:00Uhr Frau Jannick 0211/904361209 Pat. Barcodes werden nicht erkannt | Barcode Konfig muss angepasst werden. | Gerät konfiguriert, Barcode Laser für Interleaf 2-5 any Length konfiguriert. | SWO-215464 | [5 empty] | 07/19/2017 10:00AM Ms Jannick 0211/904361209 Pat. Barcodes are not recognized | Barcode Config must be adjusted. | Device configured, barcode laser configured for Interleaf 2-5 any Length.
No | Visit | Complaint | RCHN | 屏幕黑屏 | 屏幕黑屏，借备件552-039一个 |  | SWO-214452 | [5 empty] | Screen black screen | Screen black screen, borrowed spare parts 552-039 one
No | 6 | Y | RAME | Z Customer unable to get sensor cassette to initialize |  | Tried to reinitialize sensor cassette. Unable to pass. | D2E20CB7BC34268E | [5 empty] | Z Customer unable to get sensor cassette to initialize |  | Tried to reinitialize sensor cassette. Unable to pass.
No | Visit | Complaint | RCHN | 剂样针错误 | 剂样针错误，借903-158 |  | SWO-214440 | Defect | HW | Basic Unit | Solution Pack ID PCB |  | Dossing needle error | Dose needle error, borrow 903-158
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | O2，CL定标不过 | 更换测试卡后修复 | SWO-214419 | Assisted Replacement | SC | SC | pO2 |  | Test card, reagent package failure | O2, CL calibration but | Repair after replacing test card
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 传感器加温器重复定标 | 更换试剂包后修复 | SWO-214897 | Assisted Replacement | SC | SC | Thermistor | 1325 | Test card, reagent package failure | Sensor booster repeat calibration | Repair after replacement of reagent kit
No | 1 | Y | RFRA | Blank | Erreur de calibration |  | D2E1FC5EBC01D0D8 | [5 empty] | Blank | Calibration error
Yes | Visit | Complaint | RHUN | a windows elindul, de az analizátor sw már nem. Disk space full hiba üzenet jön helyette. | Cserekészüléknél is hasonló hibajelenségek. | Analizátor szoftver megsérült. Újratelepítés. Kalibrálás, QC mérés. | SWO-214477 | [5 empty] | Windows starts but the Analyzer SW is no longer. Disk Space Full error message comes instead. | Similar symptoms in replacement plants. | Analyzer software is damaged. Reinstallation. Calibration, QC measurement.
No | Visit | Complaint | RDAN | Jan melder fejl 579. | Afprøvet. | Neonlamp udskiftet, kalibreret og slutafprøvet. | SWO-214725 | Defect | HW | Oxi module | Lamp unit | Neon lamp | Jan reports error 579. | Tested. | Neon lamp replaced, calibrated and final tested.
No | 9 | Y | RPNZ | Blank | RCPA electrolytes high and Na QC unstable (high), also calibration drift. |  | D2E22DFEBC3C80BB | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | Blank | RCPA electrolytes high and Na QC unstable (high), also calibration drift.
No | 7 | Y | RAME | F spectrometer missing trigger |  | all oxi fail | D2E220CEBC579E8B | Revisit required | HW | Oxi module | PCB, Oxi module |  | F spectrometer missing trigger |  | All Oxy fail
No | 3 | Y | RAME | F oximetry errors |  | intermittent missing triiger and low voltge | D2E221E9BC389E8E | Defect | HW | Oxi module | PCB, Oxi module |  | F oximetry errors |  | intermittent missing triiger and low voltge
No | Visit | Complaint | RDEU | 21.07.2017 08:30Uhr Herr Ostertag 076223951247: Das Gerät zeigt nach dem Einschalten: Bitte Boot-CD einsetzen | Gerät bootet nicht | Test-Netzteil wieder ausgebaut, CPU ersetzt, keine Veränderung, Festplatte eingebaut, BIOS angepasst und Windows 7 aufgespielt, Initialisierungsprozess von Windows bleibt hängen ....mehrfach versucht, Backup Gerät (ABL80 CO-OX) aus Lörrach abgeholt und aufgestellt, Kunden eingewiesen | SWO-214780 | Revisit required | HW | DMS_PC | DMS_PC |  | 21.07.2017 08:30Am Mr Ostertag 076223951247: The device shows after switching on: Please insert boot CD | Device does not boot | Test power supply removed, CPU replaced, no modification, hard drive built-in, BIOS adapted and Windows 7 played up, initialization process of Windows remains hanging .... tried several times to pick up backup device (ABL80 CO-OX) from Lörrach and set it up, instructed customers
No | Visit | Complaint | RDEU | 25.07.2017, 08:08 Uhr, H. Kotz, Tel.: 071164893256: Leckage erkannt | Fehleranalyse durchgeführt. Gerät fordert ständig LP und Einlassdichtung wechseln. | Einlasssonde und Dichtungen aus Kundenbestand erneuert. Kein Erfolg. Oxitubing erneuert, Schlauch ist vor dem Jahreswartungs - Intervall durchgeschlagen - Undicht. Gerät gespült und Einlaufverfahren durchgeführt . Alles wieder o.k. | SWO-216158 | Defect | HW | Oxi module | Pinch valve tube |  | 25.07.2017, 08:08, H. Kotz, Tel.: 071164893256: Leakage detected | error analysis. Device constantly requires LP and inlet seal to change. | Inlet probe and seals from customer base renewed. No success. Oxitubing renewed, hose is leaked before the annual maintenance - interval . Device rinsed and run-in procedure performed. All again OK.
No | Visit | Complaint | RLTD | Leak | Checked the various parts of the sample path and found that there was still a clot in the inlet probe. | Flushed the probe and removed the clot. In the service programs ran the fluid transport checks which all passed. Ran the clot detect $remove program which passed. Left the service programs and pressed test again, waited until all the fluid transport checks had completed. | SWO-215022 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway | Clot | Leak | Checked the various parts of the sample path and found that there was still a clot in the inlet probe. | Flushed the probe and removed the clot. In the service programs ran the fluid transport checks which all passed. Ran the clot detect $remove program which passed. Left the service programs and pressed test again, waited until all the fluid transport checks had completed.
No | Visit | Complaint | RLTD | Involving two ABL 90s in ICU one analyser gave a Haemoglobin level of 62 the other 84 but the laboratory result was 74 Bahar Razzaghi 01689 863699 | Customer reported significant difference in tHb results between the 2 ABL90 analysers in the unit, (serial nos. 092R0097N007 and 092R0097N008) for one patient on 21/07/2017. Checked both of the ABL90 analysers calibration and quality control data logs - no errors or out of range results on the 21/07 or on the preceding and following days. Obtained fresh blood sample from unit and ran the same sample 5 times on each analyser, results were within 2 g/L (between 96 & 98 g/L) for all samples on both analysers. This is well within Radiometer specification of +/- 4 g/L at 100 g/L. 2nd sample obtained and run 3 times on each analyser, again all results within 2 g/L. Also upgraded the software on both analysers from v3.3 to v3.3 MR1. | Analysers checked and both found to be working correctly. | SWO-215131 | No fault found | HW | Oxi module | tHb |  | Involving two ABL 90s in ICU one analyser gave a Haemoglobin level of 62 the other 84 but the laboratory result was 74 Bahar Razzaghi 01689 863699 | Customer reported significant difference in tHb results between the 2 ABL90 analysers in the unit, (serial nos. 092R0097N007 and 092R0097N008) for one patient on 21/07/2017. Checked both of the ABL90 analysers calibration and quality control data logs - no errors or out of range results on the 21/07 or on the preceding and following days. Obtained fresh blood sample from unit and ran the same sample 5 times on each analyser, results were within 2 g/L (between 96 & 98 g/L) for all samples on both analysers. This is well within Radiometer specification of +/- 4 g/L at 100 g/L. 2nd sample obtained and run 3 times on each analyser, again all results within 2 g/L. Also upgraded the software on both analysers from v3.3 to v3.3 MR1. | Analysers checked and both found to be working correctly.
No | Visit | Complaint | RCHN | 泵头断裂 | 泵头断裂，借备件902-971一个 | 更换902-971后修复 by tim | SWO-214502 | [5 empty] | The pump head is broken. | Pump head broken, spare part 902-971 one | Repair by tim after replacement 902-971
No | 14 | Y | RAME | F None of the oximetry parameters will work. |  | The halogen lamp channel fails in the hemolyzer service test. | D2E3640532300E09 | [5 empty] | F None of the oximetry parameters will work. |  | The halogen lamp channel fails in the hemolyzer service test.
No | 21 | Y | RAME | Z OXIMETRY FAILURE ALL OXI READINGS LOCKING OUT. |  | UNIT WOULD END SYSTEM CYCLES WITH AN ERROR CODE OF 1218 AND LOCK OUT ALL OXIMETRY READINGS | D2E36AAB325F5B7A | [5 empty] | Z OXIMETRY FAILURE ALL OXI READINGS LOCKING OUT. |  | UNIT WOULD END SYSTEM CYCLES WITH AN ERROR CODE OF 1218 AND LOCK OUT ALL OXIMETRY READINGS
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | GLU,LAC定标不过 | 更换测试卡后修复 | SWO-214842 | Assisted Replacement | SC | SC | Glu/Lac |  | Test card, reagent package failure | GLU, LAC calibration but | Repair after replacing test card
No | Visit | Complaint | RESP | MULTIPLES ERRORES DE FLUIDICA. | SE DESMONTA MODULO INLET Y SE DETECTA PIEZA ROTA DEL MISMO. | SUSTITUCION MODULO INLET. | SWO-215186 | Defect | HW | Auto Inlet module | Auto Inlet module |  | MULTIPLE FLUIDIC ERRORS. | MODULO INLET IS DISMONTATED AND ROTATED PART OF IT IS DETECTED. | SUBSTITUTION MODULE INLET.
Yes | Visit | Complaint | RESP | MULTIPLES ERRORES DE FLUIDICA. | MODULO INLET ROTO | SE HACE OTRO SWO POR OLVIDO DE PIEZAS. | SWO-215192 | Re classified RMED | Double entry |  |  |  | MULTIPLE FLUIDIC ERRORS. | BROKEN INLET MODULE | ANOTHER SWO IS MADE BY FORGET OF PARTS.
No | Visit | Complaint | RCHN | 定标卡报错 | 定标卡报错，借944-258（批号12413） |  | SWO-214853 | Cleaned | HW | Process Module | Process wheel | 1025 | Bid card error | Bid card error, borrow 944-258 (lot number 12413)
No | Visit | Complaint | RKK | キャリブレーション不能 | キャリブレーション不能 | 機器引き上げ | SWO-215217 | [5 empty] | Not calibrated | Not calibrated | Equipment pulling up
No | Visit | Complaint | RKK | pH電極、電極ｺﾈｸﾀ交換し感度ﾊﾞﾗﾂｷ改善。ｲﾝﾚｯﾄLSﾁｭｰﾌﾞ、ｲﾝﾚｯﾄﾆｯﾌﾟﾙ交換。測定ﾓｼﾞｭｰﾙ、ﾌﾙｲﾄﾞの各ﾊﾞﾙﾌﾞ分解清掃注油、ﾌﾗｯﾄﾒﾝﾌﾞﾗﾝ、ﾘﾝｽﾊﾞﾙﾌﾞﾒﾝﾌﾞﾗﾝ交換。 | pH感度ﾊﾞﾗﾂｷ。ﾄﾞﾘﾌﾄ大きい。 測定時ｲﾝﾚｯﾄでの空気導入なしによりｻﾝﾌﾟﾙ移動ｴﾗｰ。 | pH電極、電極ｺﾈｸﾀ交換し感度ﾊﾞﾗﾂｷ改善。ｲﾝﾚｯﾄLSﾁｭｰﾌﾞ、ｲﾝﾚｯﾄﾆｯﾌﾟﾙ交換。測定ﾓｼﾞｭｰﾙ、ﾌﾙｲﾄﾞの各ﾊﾞﾙﾌﾞ分解清掃注油、ﾌﾗｯﾄﾒﾝﾌﾞﾗﾝ、ﾘﾝｽﾊﾞﾙﾌﾞﾒﾝﾌﾞﾗﾝ交換 | SWO-215228 | Defect | HW | Inlet module | Tubes, Inlet |  | Improved sensitivity variation by replacing pH electrodes and electrode connectors. Inlet LS tube, inlet nipple replacement. Measuring module, fluid valve disassembly cleaning lubrication, flat membrane, rinse valve membrane replacement. | pH sensitivity variation. The drift is great. More sample movement error without air introduction in the inlet at the time of measurement. | Improved sensitivity variation by replacing pH electrodes and electrode connectors. Inlet LS tube, inlet nipple replacement. Measuring module, fluid valve decomposition cleaning lubrication, flat membrane, rinse valve membrane replacement
No | Visit | Complaint | RBNL | Problème oximètre | De temps en temps une erreur apparait : #0256 (neon lamp voltage too high) Une remise en route ne résouds pas le problème. | Remplacement de la carte pilote du module oximètre, chargement du logiciel (de la carte), remis en route, calibration, cal tHb et contrôle : OK | SWO-215010 | Defect | HW | Oxi module | PCB, Oxi module |  | Oximeter problem | Occasionally an error appears: #0256 (neon lamp too high voltage) A restarted does not solve the problem. | Replacing the pilot card of the oximeter module, loading the software (of the map), restarted, calibration, cal tHb and control: OK
No | Visit | Complaint | RIND | Touch Screen Problem. | Touch Screen Problem. | Replaced the display unit, from Lotus instrument, calibrated found display is working satisfactorily. | SWO-216557 | Defect | HW | DMS_PC | Touch screen | Firmware should have been SWU | Touch Screen Problem. | Touch Screen Problem. | Replaced the display unit, from Lotus instrument, calibrated found display is working satisfactorily.
No | Visit | Complaint | RSCH | 0509: Oberer ph/BG Flüssigkeitsfühler hat keine Spüllösung von oben erhalten. | 0509: Oberer ph/BG Flüssigkeitsfühler hat keine Spüllösung von oben erhalten. | Allgemeine Kontrolle ok. Ventile im Servicemenü überprüft ok. Rinse Ventil ph $BG Modul hat keine funktion und fehler. Messkammer ers. Ok. Rinseventil ersetzt ok. System i.o. Kal. + QC i.o. | SWO-215092 | Defect | HW | pH/BG Module | Valve, rinse |  | 0509: Upper ph/BG liquid sensor has not received a rinsing solution from above. | 0509: Upper ph/BG liquid sensor has not received a rinsing solution from above. | General control ok. Valves checked in the service menu ok. Rinse valve ph $BG module has no function and error. measuring chamber hes. Ok. Rinse valve replaces ok. System i.o. Kal. + QC i.o.
No | 7 | Y | RAME | f all oxi value fail |  | spectrometer missing trigger, neon number, ect | D2E33BFB325982E6 | Defect | HW | Oxi module | PCB, Oxi module |  | F All Oxy value fail |  | spectrometer missing trigger, neon number, ect
No | 12 | Y | RAME | Erro 521 on po2 |  | Start techline case on error521 po2 | D2E46CA8A81EB79E | [5 empty] | Erro 521 on po2 |  | Start techline case on error521 PO2
No | 4 | Y | RAME | V user called stating that the door to the SC will not open. |  | found that if you tell analyzer to open SC door it will not open. | D2E49801A8380D9C | Adjustment | HW | Basic Unit | SC Door |  | V user called stating that the door to the SC will not open. |  | found that if you tell analyzer to open SC door it will not open.
Yes | 7 | Y | RFRA | Blank | RAS |  | D2E4D971A80274F2 | Re classified RMED | SWU |  |  |  | Blank | RAS
No | Visit | Complaint | RDEU | 25.07.2017 Sr Steffi 06519473930 meldet, dass sich der Bildschirm nicht bedienen lässst. | Touch Screen defekt | Touch Screen ausgetauscht. | SWO-216170 | Defect | HW | DMS_PC | Touch screen |  | 07/25/2017 Sr Steffi 06519473930 reports that the screen is not being able to operate. | Touch screen defective | Touch screen.
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | Lac定标不过 | 更换测试卡后修复 | SWO-215718 | [5 empty] | Test card, reagent package failure | Lac calibration is not | Repair after replacing test card
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 定标错误 | 更换测试卡后修复 | SWO-215729 | Assisted Replacement | SC | SC | SC |  | Test card, reagent package failure | Calibration error | Repair after replacing test card
No | Visit | Complaint | RKK | リンスエラー・サンプル移動エラー：エアライン洗浄、タンパク除去、ガスケット交換実施。LS、バルブ動作問題なし。流路内に血液の汚れがあり一時的なものと考えられる。 | リンスエラー・サンプル移動エラー：エアライン洗浄、タンパク除去、ガスケット交換実施。LS、バルブ動作問題なし。流路内に血液の汚れがあり一時的なものと考えられる。 | リンスエラー・サンプル移動エラー：エアライン洗浄、タンパク除去、ガスケット交換実施。LS、バルブ動作問題なし。流路内に血液の汚れがあり一時的なものと考えられる。 | SWO-216091 | Cleaned | HW | Fluidic pathway | Fluidic pathway |  | Rinsing error sample movement error: Airline cleaning, protein removal, gasket replacement. LS, no valve operation problem. There is a dirt of blood in the flow path is considered to be temporary. | Rinsing error sample movement error: Airline cleaning, protein removal, gasket replacement. LS, no valve operation problem. There is a dirt of blood in the flow path is considered to be temporary. | Rinsing error sample movement error: Airline cleaning, protein removal, gasket replacement. LS, no valve operation problem. There is a dirt of blood in the flow path is considered to be temporary.
Yes | 19 | Y | RAME | 8 Glu is giving unstable results |  | Glu is giving unstable results on Patient and QC results | D2E46582A80DBB11 | [5 empty] | 8 Glu is giving unstable results |  | Glu is giving unstable results on Patient and QC results
No | Visit | Complaint | RDEU | Gerät meldet unter Status und System Meldungen: Neon Lamp Voltage too high | OXI Modul defekt. | OXI Modul erneuert und nachjustiert. | SWO-215758 | Defect | HW | Oxi module | Oxi module, Replaced | Neon lamp | Device reports under Status and System Messages: Neon Lamp Voltage too high | OXY module Puncture. | OXI module renewed and readjusted.
No | Visit | Complaint | RDEU | 25.07.2017 11:30 Fr. Benz : Interventionsmeldung; Gasdruck | Gas Unit defekt. Gas 1 defekt, Druck Transducer an der Gas Unit defekt. | Messkammer gereinigt, Gas Wege freigespült, Air Membranventile in den Messmodulen demontiert und gereinigt. Drucktranducer Einheit demontiert und gereinigt, Gaswege können nicht mehr komplett gereinigt werden. Gas 1 Flasche lässt sich nicht mehr gerade in das Gas1 Ventil Gewinde Schrauben. Nach der Reinigung der Gas Unit hat sich die Messung des Betriebsdruck in den Gasflaschen verbessert aber immer nicht nicht optimal. Fehler 0637: Gasbehälter geringer Druck bleibt bestehen. Kal 1 und 2 sind OK. THB Kal durchgeführt. Ersatzteile werden bestellt. | SWO-215757 | Revisit required | HW | Gas module | Gas module |  | 25.07.2017 11:30 Fr. Benz : Intervention notification; Gas | Gas Unit defective. Gas 1 defective, pressure transducer defective at the gas unit. | Measuring chamber cleaned, gas paths cleared, air diaphragm valves in the measuring modules dismantled and cleaned. Pressure tranducer unit dismantled and cleaned, gas paths can no longer be completely cleaned. Gas 1 bottle can no longer be straight into the Gas1 valve thread screws. After cleaning the gas unit, the measurement of the operating pressure in the gas cylinders has improved but is not always optimal. Error 0637: Low pressure gas tank remains. Kal 1 and 2 are OK. THB Kal. Spare parts are ordered.
No | Visit | Complaint | RDEU | 21.07.2017 08:30Uhr Herr Ostertag 076223951247: Das Gerät zeigt nach dem Einschalten: Bitte Boot-CD einsetzen | Gerät bootet nicht | Windows 7 aufgespielt, AU Software aufgespielt, Konfiguration angepasst, Anbindung hergestellt, Gerät einkalibriert und Testmessung und QC durchgeführt, i.O. | SWO-215505 | Defect | SW | Windows | Windows |  | 21.07.2017 08:30Am Mr Ostertag 076223951247: The device shows after switching on: Please insert boot CD | Device does not boot | Windows 7 played up, AU software played up, configuration adapted, connection established, device calibrated and test measurement and QC performed, i.o.
No | 3 | Y | RAME | Z General error |  | General error caused by a faulty sensor. Sensor 84 lot 84-66543 out of warranty and no contract. | D2E4B58DA855282C | [5 empty] | Z General error |  | General error caused by a faulty sensor. Sensor 84 lot 84-66543 out of warranty and no contract.
No | 3 | Y | RAME | z Not reading correctly | Changed the complaint visit tab from N to Y and changed the health hazard tab from 'blank' to N. | changed the front panel where power button is housed and issues were fixed | D2E43C4AA8284313 | [5 empty] | z Not reading correctly | Changed the complaint visit tab from N to Y and changed the health hazard tab from 'blank' to N. | changed the front panel where power button is housed and issues were fixed
No | 5 | Y | RAME | Z - NO CONNECTION TO RADIANCE. |  | FOUND STATIC IP ADDRESS ASSIGNED IN NETWORK PROPERTIES. | D2E60B7A1E359AB0 | Setup | SW | His/lis | His/lis |  | Z - NO CONNECTION TO RADIANCE. |  | FOUND STATIC IP ADDRESS ASSIGNED IN NETWORK PROPERTIES.
No | 5 | Y | RAME | G user called stating that COhb and SO2 will not pass internal qc B |  | found that it was failing internal qc B low for SO2 and COHb | D2E5E2CE1E001C3F | Adjustment | SW | Oxi module | tHb Cal |  | G user called stating that COhb and SO2 will not pass internal qc B |  | found that it was failing internal qc B low for SO2 and COHb
Yes | 1 | Y | RPOL | Blank | Bledy OXI na kalibracji. |  | D2E6E0BF9424D252 | [5 empty] | Blank | Bledy OXI on kalibracji.
No | Visit | Complaint | RDEU | 27.07.2017 9:30 Fr. Becke 0221467771346: Interventionsmeldung | Fehlersuche durchgeführt. | Einlassdichtung erneuert. | SWO-216181 | Defect | HW | Auto Inlet module | Inlet gasket |  | 27/07/2017 9:30 Ms. Becke 0221467771346: Intervention notification | Troubleshooting. | Inlet seal renewed.
No | 6 | Y | RFRA | Blank | Probleme de stabilité de l'analyseur |  | D2E6C81B945CD9F4 | Unknown | Unknown | Unknown | Unknown |  | Blank | Analyzer stability problem
Yes | 14 | Y | RAME | P Unit leaking at sensor. |  | Customer put on 2 solution packs. One pack has recurring VC check failure after reset. | D2E5E3931E0591BF | [5 empty] | P Unit leaking at sensor. |  | Customer put on 2 solution packs. One pack has recurring VC check failure after reset.
No | 7 | Y | RAME | H Air Detection Disabled. |  | Air detection was being disabled after installation of new sensor cassette. Suspect is malfunctioning sensor cassette. | D2E6527E1E580429 | [5 empty] | H Air Detection Disabled. |  | Air detection was being disabled after installation of new sensor cassette. Suspect is malfunctioning sensor cassette.
No | Visit | Complaint | RCHN | LAC定标值偏高 | LAC定标值偏高，借备件903-135一个 | 由于用户反映机器测量Lac的结果不准确，使用检测工具903-135检测机器后，发现机器正常。继续寻找影响结果的其他原因。 by jesson | SWO-215367 | Unknown | Unknown | Unknown | Unknown |  | High LAC calibration value | LAC calibration value is high, borrowed spare parts 903-135 one | As the user reported that the results of the machine measurement Lac were not accurate, the machine was found to be ok after the machine was detected using the detection tool 903-135. Continue to look for other reasons that affect the results. by jesson
No | Visit | Complaint | RCHN | 光单故障 | 光单故障，借903-364 |  | SWO-215654 | Defect | HW | Process Module | Optical Unit |  | Light single fault | Light single fault, borrow 903-364
No | Visit | Complaint | RDEU | 24.07.2017, 11:05 Uhr, H. Penczok, Tel. 030450575776: Temp. Fehler im Lyt/Met. Modul | ABL aktuell ohne Fehler, Speicher durchgesehen, viele Tempfehler im El/Met Modul in den letzten 2 Wochen, Regulierung im Service überprüft | ABL abgeschalten und Modul ausgebaut, Regleplatine und Fühler überprüft - Wärmeleitpaste erneuert, alles zusammengebaut, Einlaßdichtung stark verschmutzt - ersetzt, Neustart, mehrere Kal, QC, CFM und Station darauf hingewiesen das der Fehler durchaus wiederkommen kann und dann Teile gebraucht werden | SWO-216202 | Revisit required | HW | El/Met Module | Temp |  | 24.07.2017, 11:05, H. Penczok, tel. 030450575776: Temp. Error in Lyt/Met module | ABL currently without errors, memory reviewed, many temp errors in the El/Met module in the last 2 weeks, regulation in the service checked | ABL switch off and module removed, regle board and probe checked - thermal paste renewed, everything assembled, inlet seal heavily soiled - replaced, restart, several Kal, QC, CFM and station pointed out that the error may well come back and then parts are needed
No | 11 | Y | RPNZ | Blank | Urgent visit due to error 0509: pH/BG upper liquid sensor did not receive rinse from above. Isuru unavailable so I drove from Whakatane to Rotorua airport to get to Wellington next flight. |  | D2E61B9F1E5A7112 | Defect | HW | El/Met Module | Valve, rinse |  | Blank | Urgent visit due to error 0509: pH/BG upper liquid sensor did not receive rinse from above. Isuru unavailable so I drove from Whakatane to Rotorua airport to get to Wellington next flight.
No | Visit | Complaint | RDEU | Es tritt ein Spül- und LS-Fehler auf | Fehlersuche Kalibrierung gescheitert | Ventilmembrane an den AIR-Valves in den Meßmodulen erneuert Glu-Elektrode membraniert 2-Pkt.-Kalibrierungen durchgeführt | SWO-215998 | Defect | HW | pH/BG Module | Valve, air |  | A flushing and LS error occurs | Troubleshooting Calibration failed | Valve membrane on the AIR valves in the measuring modules renewed Glu-electrode membraneed 2-pt. calibrations performed
No | Visit | Complaint | RDEU | Gerät friert ein, bei mehrmaligem Öffnen des Einlasses. | Bei Eintreffen kein Fehler feststellbar/reproduzierbar. | 40-poliges Flachbandkabel zwischen Inlet und Mainboard erneuert. Ca. 20 malignes öffnen und schließen im Servicemenü und bei der Probeneingabe ohne Beanstandung. | SWO-216052 | Defect | HW | Cover and Chassis | Cable, front |  | device freezes when the inlet is opened several times. | No error can be detected/reproducible on arrival. | 40-pin ribbon cable between inlet and motherboard renewed. Approximately 20 malignant sands open and close in the service menu and when entering samples without complaint.
No | 5 | Y | RAME | G Customer stating unit giving multiple parameter failure. |  | Arrived on site and checked logs, unit giving cal back log errors. | D2E7846A9443169D | Assisted Replacement | SC | SC | FT |  | G Customer stating unit giving multiple parameter failure. |  | Arrived on site and checked logs, unit giving cal back log errors.
No | Visit | Complaint | RDEU | 27.07.2017 13:00 Trompka 003084370625: Drucker druckt nicht vollständig | Andruckrolle aus dem Deckel gerissen - Messingring fehlt | Drucker gewechselt - Test i.O. | SWO-216402 | Defect | HW | Printer Module | Printer head |  | 07/27/2017 13:00 Trompka 003084370625: Printer does not print completely | Pressure roller ripped from the lid - brass ring is missing | Printer changed - Test i.O.
No | 3 | Y | RPOL | Blank | P³ytka uzupe³niona z wys³an¹ do Kuby na Wa³cz. Ró¿ne b³edy OXI. |  | D2E6DA9A9434D247 | Defect | HW | Oxi module | PCB, Wet section Control |  | Blank | The disc is completely sent to the Cuba on the part. Various errors OXI.
No | Visit | Complaint | RCHN | LAC定标失败，初始化失败1117 | LAC定标失败，初始化失败1117 ，借903-334一个。 | 更换903-334后修复 by jesson | SWO-215055 | [5 empty] | LAC calibration failed, initialization failure 1117 | LAC calibration failed, initialization failure 1117 , borrow edith of 903-334. | Repair by jesson after replacing 903-334
No | 1 | Y | RFRA | Blank | CQ1: PH hors normes |  | D2E7502F941B51F2 | [5 empty] | Blank | CQ1: Outstanding PH
No | Visit | Complaint | RCHN | 液路故障，报错3412 | 液路故障，报错3412，借备件903-329一个 |  | SWO-216173 | Defect | HW | Aspiration Module | Needle Wash Unit | 3412 | Liquid road fault, error 3412 | Liquid road fault, error 3412, spare parts 903-329 one
No | Visit | Complaint | RCHN | lac超出范围，定标失败 | 干燥前温度超出范围，定标失败，需要借件944-258,和944-212（批次号：13121、12988） | 重新校准温度。重新定标后修复。 by jesson | SWO-215372 | Adjustment | HW | Process Module | Cup dryer | 1068 | lac out of range, bid failed | The pre-drying temperature is out of range, the calibration fails, the borrowing s944-258, and 944-212 (batch number: 13121, 12988) | Recalibrate the temperature. Repair after re-calibration. by jesson
No | Visit | Complaint | RBNL | Oxi-mètre (valeur de tHb et sO2 en dehors des tolerances pour CQ) | Les valeurs toutes les paramètres oximétrie sont en dehors des tolérances pour CQ. Intensité de la lampe : OK | Calibration de l'Hb en service, vérification des CQ (3 niveaux) : OK ! Si le problème revient je viendrais remplacer l'hémolyseur. | SWO-216290 | Adjustment | HW | Oxi module | tHb |  | Oxi-meter (tHb and sO2 value outside tolerances for CQ) | The values all the parameters oximetry are outside the tolerances for CQ. Lamp Intensity: OK | Hb calibration in service, check of CQ (3 levels): OK! If the problem comes back I'll come and replace the hemolyser.
No | Visit | Complaint | RDEU | Anwender müssen in letzter Zeit sehr häufig membranieren. | Aus Kopenhagen empfohlene Maßnahme 'Dekontamination durchführen', lt. Dateiarchiv immer noch auftretende Kal.fehler bei Glu und Lac | Dekontamination gemäß Herstellervorgaben durchgeführt, mit Kunden (Herr Dünnebeil u. Frau Kupferschmidt-Bode) besprochen, dass bei weiterhin aufkommenden Kal.fehlern direkte Rückmeldung an VBM oder KDB erfolgt, Kunde beobachtet Verhalten des Gerätes | SWO-216436 | Cleaned | HW | El/Met Module | Fluidic pathway |  | Users have had to membrane very frequently lately. | Recommended action from Copenhagen 'perform decontamination', according to file archive still occurring cal.errors in Glu and Lac | Decontamination carried out according to manufacturer's specifications, discussed with customers (Mr. Dünnebeil and Ms. Kupferschmidt-Bode) that in case of further occurring Kal.errors direct feedback to VBM or KDB takes place, customer observes behavior of the device
No | Visit | Complaint | RDEU | 27.07.2017; Sr. Anett; Tel.: 03381791614; CRP QC Messfehler 28.07.2017 10:30 Sr. Nicole; Tel.: 03381791612; weiterhin CRP QC Messfehler | LQC zu hoch in Level 1 u. 2 | Reinigung durchgeführt, Nadel und Wash Unit überprüft - i.O., LQC Level 2 i.O. SD gezogen | SWO-216661 | Cleaned | HW | Aspiration Module | Needle Wash Unit |  | 27.07.2017; Sr. Anett; Tel.: 03381791614; CRP QC Messfehler 28.07.2017 10:30 Sr. Nicole; Tel.: 03381791612; weiterhin CRP QC Messfehler | LQC too high in Level 1 &i.g. 2 | Cleaning performed, needle and wash unit checked - i.O., LQC Level 2 i.O. SD drawn
No | Visit | Complaint | RIND | Multiple parameters failed in calibration | Checked and found sensor board corroded. | Replaced sensor board from ABL80 in CP's stock, 2 point calibration performed successfully. | SWO-216569 | [5 empty] | Multiple parameters failed in calibration | Checked and found sensor board corroded. | Replaced sensor board from ABL80 in CP's stock, 2 point calibration performed successfully.
No | Visit | Complaint | RCHN | 测试卡下层温度超范围 | 测试卡下层温度超范围，借备件903-644一个 |  | SWO-216550 | Defect | HW | Basic Unit | PCB, Wet section Control | Temp | Test card lower temperature over range | Test card lower temperature over range, spare parts 903-644 one
No | 18 | Y | RFRA | Blank | Lenteur de connexion |  | D2EAE092F63F3E73 | Setup | SW | Roma.ini | His/lis |  | Blank | Slow connection
Yes | 19 | Y | RPOL | Blank | Zdalne i telefoniczne usuniêcie awarii polegaj¹cej na wymianie uszczelki po wymianie pakietu. Z biura wys³ana uszczelka. Kalibracja, pomiary QC |  | D2EACDC0F613E58E | Re classified RMED | Phone support |  |  |  | Blank | Remote and dial-up failure to replace the seal after replacing the package. A seal is sent from the office. Calibration, QC measurements
No | 15 | Y | RPNZ | Blank | Sensor cassette replacement due to pO2 QC below range. |  | D2EAAAA6F610E0D2 | Assisted Replacement | SC | SC | pO2 | QC. SP also replaced | Blank | Sensor cassette replacement due to pO2 QC below range.
Yes | Visit | Complaint | RLTD | Install | Install | Set up 3 machines. Will revisit site Wed 2nd to set up machines and apply IT settings to 3 machines. | SWO-217009 | Re classified RMED | Installation |  |  |  | Install | Install | Set up 3 machines. Will revisit site Wed 2nd to set up machines and apply IT settings to 3 machines.
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | LAC定标失败 | 更换测试卡后修复 | SWO-216673 | [5 empty] | Test card, reagent package failure | LAC calibration failure | Repair after replacing test card
No | 18 | Y | RAME | S Failed VK |  | Customer reported that the slope was low on the MetHb when running a VK kit. | D2EFA000CE4B40D6 | [5 empty] | S Failed VK |  | Customer reported that the slope was low on the MetHb when running a VK kit.
No | Visit | Complaint | RCHN | 压力测试真空错误 | 压力测试真空错误 | 更换夹紧阀后修复(By Ran) | SWO-216824 | Defect | HW | Oxi module | Pinch valve tube | 1063 | Pressure test vacuum error | Pressure test vacuum error | Repair after replacing clamping valve (By Ran)
No | Visit | Complaint | RCHN | 光学单元超范围 | 光学单元超范围 |  | SWO-216635 | Defect | HW | Process Module | Optical Unit | 1031 | Optical unit over range | Optical unit over range
No | Visit | Complaint | RKK | 装置が予期せぬシャットダウンを起こし正常起動しない。立ち上がるも黒い画面で英文のメッセージが出る為、装置背面の電源を切り、再度立ち上げるも同様の現象が起きてしまう。 |  |  | SWO-217107 | Defect | HW | Basic Unit | Power Supply |  | The device causes an unexpected shutdown and does not start normally. Because the message of English comes out on a black screen though it stands up, the power on the back of the device is turned off, and a similar phenomenon occurs even if it starts it up again.
No | Visit | Complaint | RKK | Auto check cover failure オートチェック・カバーが開いてそのまま閉まり、閉まっているにも関わらず閉まり続ける。 | Auto check cover failure オートチェック・カバーが開いてそのまま閉まり、閉まっているにも関わらず閉まり続ける。 | 代替モジュール設置致しました。 オートチェック動作及びQC測定確認・・・良好 | SWO-221269 | Defect | HW | Auto Check Module | Auto Check Module | Cover | Auto check cover failure The autocheck cover opens and closes as it is, and it continues to close even though it is closed. | Auto check cover failure The autocheck cover opens and closes as it is, and it continues to close even though it is closed. | We have installed an alternative module. Auto check operation and QC measurement confirmation...
No | 7 | Y | RAME | F OXI SPECTRUM MISMATCH ON QC |  | NONE NECESSARY | D2EAB67AF660959A | Defect | HW | Oxi module | Oxi module, Replaced | Spectrum mismatch | F OXI SPECTRUM MISMATCH ON QC |  | NONE NECESSARY
No | 13 | Y | RPNZ | Blank | QC S7735 aborting when new LOT number installed. |  | D2EB19F4F62FE0DC | Assisted Replacement | QC | QC | QC |  | Blank | QC S7735 aborting when new LOT number installed.
No | 3 | Y | RCZE | Blank |  |  | D2EB387FF606B6AA | Assisted Replacement | Electrode | Membrane | Cl |  | Blank
No | 12 | Y | RAME | F oxy spectrum mismatch on linearity |  | VK kit ran on July 7th, oxy parameters had results of 581 for level 3 only; customer ran another VK on July 28th, within the 21 day lock out will replace VK kit | D2EAC0E5F6557FF6 | Defect | HW | Oxi module | Hemolyzer | Spectrum mismatch | F oxy spectrum mismatch on linearity |  | VK kit ran on July 7th, oxy parameters had results of 581 for level 3 only; customer ran another VK on July 28th, within the 21 day lock out will replace VK kit
No | 4 | Y | RFRA | Blank | FUITE JOINT INTRO |  | D2EAFA0BF628A821 | Defect | HW | Inlet module | Inlet gasket |  | Blank | LEAK JOINT INTRO
No | Visit | Complaint | RDEU | Bildschirm reagiert falsch | 1- vom Bildschirmschoner zu normalen Bild dauert sehr lange 2- Bildschirm reagiert entweder verzögert oder gar nicht | Neustart hilft manchmal. 1-SI -17.004 eingestellt. Regedit 2- Keine Mitteilung über Firmware Installation in Win 7.!!! Firmware nicht installiert. | SWO-217054 | Defect | SW | Screen Saver | Screen Saver |  | Screen reacts incorrectly | 1- from screen saver to normal image takes a long time 2- Screen either reacts delayed or not at all | Restart sometimes helps. 1-SI -17.004. Regedit 2- No notification about firmware installation in Win 7.!!! Firmware not installed.
No | Visit | Complaint | RDEU | Touchscreen defekt | Bildschirm gereinigt. | Bildschirmfunktion OK. | SWO-217243 | Defect | HW | Analyzer Control Module | Touch screen |  | Touchscreen defective | Screen cleaned. | Bildschirmfunktion OK.
No | Visit | Complaint | RHUN | Vacuum error | Szivárgás hiba. | Nyomás teszt vákum hiba. Bementi tűtartó csere. Kalibrálás, QC mérés. | SWO-217407 | Replaced | HW | Inlet module | probe | vacuum | Vacuum error | Leakage failure. | Pressure test vacuum defect. Input needle holder replacement. Calibration, QC measurement.
No | Visit | Complaint | RLTD | Annette on 0300 3001761 before 10:30 or after 11:30 | The Analyser is trying to turn on, but just keeps rebooting. Suspected battery failure. | Disconnected battery and powered on Analyser. The Analyser is now turning on. Inspected battery and found that one of the cells has leaked. Screwed battery back in place, but left disconnected. Replacement battery required. Analyser has passed calibrations and is ready for use. | SWO-217283 | Defect | HW | Battery module | Battery | Cell leaked | Annette on 0300 3001761 before 10:30 or after 11:30 | The Analyser is trying to turn on, but just keeps rebooting. Suspected battery failure. | Disconnected battery and powered on Analyser. The Analyser is now turning on. Inspected battery and found that one of the cells has leaked. Screwed battery back in place, but left disconnected. Replacement battery required. Analyser has passed calibrations and is ready for use.
No | Visit | Complaint | RDEU | 18.07.2017; Fr. Dreyer; Tel.: 0213152953144; Oxy Hardware Fehler; der erste Anruf erfolgte durch Fr. Krause-Reinbrodt vom Labor Stein. 18.07.2017; Fr. Dreyer; Tel.: 0213152953144; Nach dem Neustart ist die Kal ok, QC weiterhin mit Fehlern 18.07.2017 14:00Uhr Sr. Alush 0213152953134: Weiterhin QC-Fehler bei tHb 31.07.2017 9:15 Fr. Krause-Reinbrodt 01702245229: Weiterhin QC-B Fehler Oximetrie SO2 01.08.2017, 10:25 Uhr ==> Anrufer Fr. Krause-Reinbrodt, Tel.: 01702245229 ==> QC-Messfehler in Lösung B für sO2 besteht trotz SP-Wechsel weiter | sO2 ist in der QC B oft oberhalb der Messgrenze. Der Kunde hat schon mehrfach das Lösungspack ausgetauscht und die tHb Kalibrierung durchgeführt. | Oximetrie Ventil gereinigt, Hemolyzer mit Hypochlorid gereinigt und gespült, Kalibriert, tHb Kalibrierung und QC's durchgeführt. QC B jetzt wieder Zielwertig. | SWO-217567 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway | Oxi - sO QC B | 18.07.2017; Ms. Dreyer; Tel.: 0213152953144; Oxy Hardware Error; the first call was made by Fr. Krause-Reinbrodt from the Stein laboratory. 18.07.2017; Ms. Dreyer; Tel.: 0213152953144; After reboot, the Cal ok, QC continues to be affected by errors 07/18/2017 14:00Pm Sr. Alush 0213152953134: Continued QC error at tHb 07/31/2017 9:15 Fr. Krause-Reinbrodt 01702245229: Continue QC-B Error Oximetry SO2 01.08.2017, 10:25 a.m. ==> Caller Fr. Krause-Reinbrodt, tel.: 01702245229 ==> QC measurement error in solution B for sO2 persists despite SP change | sO2 is often above the measurement limit in QC B. The customer has already replaced the solution pack several times and performed the tHb calibration. | Oximetry valve cleaned, hemolyzer cleaned and rinsed with hypochloride, calibrated, tHb calibration and QC's performed. QC B now again target-worthy.
No | 4 | Y | RFRA | Blank | 1391: VOLET INTRO BLOQUE |  | D2EC36B76C5D8499 | Adjustment | HW | Auto Inlet module | Inlet gasket | Or connector gasket | Blank | 1391: INTRO SHUTTER BLOCKS
Yes | 3 | Y | RPAC | Blank | ABL90 Flex Plus installation |  | D2EE9B5B585F5C25 | Re classified RMED | Installation |  |  |  | Blank | ABL90 Flex Plus installation
Yes | 3 | Y | RPAC | Blank | Consumable failure on installation |  | D2EE965E584A0A4D | [5 empty] | Blank | Consumable failure on installation
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | LAC定标不过 | 更换测试卡后修复 | SWO-217526 | [5 empty] | Test card, reagent package failure | LAC calibration is not | Repair after replacing test card
No | 18 | Y | RAME | S Failed VK |  | Returned too site to run linearity. Found all levels ok with the exception of PC02 running high on level 1 only. | D2EFA174CE4740DA | [5 empty] | S Failed VK |  | Returned too site to run linearity. Found all levels ok with the exception of PC02 running high on level 1 only.
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | PCO2质控失败，0589,1316 | 更换测试卡后修复 | SWO-217946 | [5 empty] | Test card, reagent package failure | PCO2 QC Failure, 0589,1316 | Repair after replacing test card
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 参比电极错误，1361 | 更换测试卡后修复 | SWO-217783 | Assisted Replacement | SC | SC | Reference | 1361 | Test card, reagent package failure | Reference electrode error, 1361 | Repair after replacing test card
No | Visit | Complaint | RCHN | 定标失败报错5769 | 定标失败报错5769，借944-394定标卡批号12726一张 |  | SWO-217125 | Defect | HW | Incubator Module | Punch Up | 5769 | Bid failure error 5769 | Bid failure error 5769, borrow 944-394 calibration card lot number 12726 one
No | 6 | Y | RAME | P Random sample errors. |  | Random sample errors. Checked all tubing and flow. | D2EC8EE66C2A1E89 | Defect | Electrode | Electrode | Sealing ring | Sealing ring, pCO2 | P Random sample errors. |  | Random sample errors. Checked all tubing and flow.
No | 19 | Y | RAME | 9 Customer states that the LAC is out after six month maintenance performed. |  | Found that the P02 , Glu and Lac having calibration drift errors on analyzer. Ran zero current calibration in service. | D2EC43826C5D3EAF | Adjustment | HW | Met2 Module | Zero current |  | 9 Customer states that the LAC is out after six month maintenance performed. |  | Found that the P02 , Glu and Lac having calibration drift errors on analyzer. Ran zero current calibration in service.
No | 4 | Y | RFRA | Blank |  |  | D2EC51E66C1910AC | Defect | HW | Cover and Chassis | Cartridge door |  | Blank
No | Visit | Complaint | RDEU | 31.07.2017, Frau Funke, QC Fehler bei NT PRO BNP Level 1 31.07.2017, 12:04 Uhr ==> Anrufer Fr. Funke, Tel.: 05313489900 ==> die vorgeschlagenen Maßnahmen haben nicht zum Erfolg geführt, der Fehler0211 besteht weiterhin | NTproBNP außerhalb der Grenzen im Multicheck L1 | Analysator geprüft, Multicheck gelöscht, neue Multicheck wird von A. Wischerhoff geliefert | SWO-217380 | Assisted Replacement | Consumables | LQC Kit | LQC | LQC Replaced | 31.07.2017, Ms Funke, QC Error at NT PRO BNP Level 1 31.07.2017, 12:04 o'clock ==> Caller S.A. Fr. Funke, tel.: 05313489900 ==> the proposed measures have not been successful, the error0211 still exists | NTproBNP outside the limits in Multicheck L1 | Analyzer checked, Multicheck deleted, new multicheck supplied by A. Wischerhoff
No | Visit | Complaint | RDEU | 31.07.2017, 11:41 Uhr, Fr. Usbeck, Tel. 071164893016: Messerg. bei tHb und Gluc. werden angezweifelt | tHb und Glucose waren kurzfristig unstimmig mit den Messungen im Labor. Kalibrierung , QC und Patientendaten analysiert. | Kal. und QC unauffällig. Zum Zeitpunkt der Unstimmigkeiten der Glucose wurde kurz zuvor die Sensorkassette erneuert. Anwender verwenden zur Zeit nicht immer die sPico Spritzen. Habe Anwender beobachtet wie die Sarstett Spritze ohne vorheriges Mischen gemessen wurde. Die zu niedrigen tHb Werte können durch zu langes liegen der Spritze bis zur Messung (ohne Durchmischung) fehlerhaft sein. 8 Blutmessungen mit der gleichen Spritze durchgeführt. Messwerte nahezu identisch, keine bemerkenswerten Abweichungen festzustellen. Stellvertretenden Stationsleiter Herr Weissmann unterrichtet über die getroffenen Erkenntnisse und gebeten bei weiteren Vorfällen dieser Art sofort eine QC Messung durchzuführen um zu sehen ob sich der Fehler in der QC widerspiegelt. Des Weiteren sollen wieder ausschließlich die sPico Spritzen verwendet werden. | SWO-217993 | User error | Pre analytical | Mixing | Glu & tHb | Sarstett syringes | 31.07.2017, 11:41 Uhr, Fr. Usbeck, Tel. 071164893016: Messerg. bei tHb und Gluc. werden angezweifelt | tHb and glucose were inconsistent with the measurements in the laboratory at short notice. calibration , QC and patient data. | Kal. and QC inconspicuous. At the time of the glucose inconsistencies, the sensor cartridge was recently renewed. Users do not always use the sPico syringes at the moment. Have observed users how the Sarstett syringe was measured without prior mixing. The tHb values that are too low can be faulty due to too long the syringe is up to the measurement (without mixing). 8 blood measurements performed with the same syringe. measurements are almost identical, no notable deviations can be observed. Deputy station manager Mr Weissmann briefed on the findings and asked for further incidents of this nature to be taken immediately by qc to see if the error was reflected in the QC. Furthermore, only the sPico syringes should be used again.
No | 5 | Y | RAME | E instrument will not turn off |  | instrument will not turn off or reset. not even unplugged | D2ED6B52E2153FC2 | Defect | HW | Battery module | Battery | Not turn off | E instrument will not turn off |  | instrument will not turn off or reset. not even unplugged
No | Visit | Complaint | RSCH | Blank | Instabilité calibration malgré le changement des réactifs 1XSP et 2XSC (garantie) | Nettoyage du système fluidique e.o Remplacé le kit de maintenance e.o Test complet avec le programme de service e.o Calibration e.o CQ e.o | SWO-217836 | Defect | HW | Basic Unit | PM kit | & cleaned fluidic pathway | Blank | Instability calibration despite the change of 1XSP and 2XSC reagents (guarantee) | Cleaning the fluid system e.o Replaced the e.o maintenance kit Full test with the e.o service program Calibration e.o CQ e.o
No | 10 | Y | RAME | F The oximetry fails calibrations and QC. |  | Found the oximeter not working properly. | D2ED3F7BE23DB775 | [5 empty] | F The oximetry fails calibrations and QC. |  | Found the oximeter not working properly.
No | Visit | Complaint | RCHN | PO2质控偏高 | PO2质控偏高，需要借件902-972和903-334各一个 | 更换902-972和903-334后修复 （By Ran） | SWO-216839 | [5 empty] | PO2 QUALITY CONTROL IS HIGH | PO2 quality control is high, need to borrow 902-972 and 903-334 each one | Replacement 902-972 and 903-334 post-repair (By Ran)
No | Visit | Complaint | RCHN | 做样本时屏幕不能点击 | 做样本时屏幕不能点击，借903-334一个 |  | SWO-217633 | [5 empty] | The screen cannot be clicked when making a sample | Do sample when the screen can not click, borrow 903-334 a
No | Visit | Complaint | RCHN | 接口板腐蚀 | 接口板腐蚀，借902-972一个 | 更换902-972后修复(By Ran) | SWO-216733 | [5 empty] | Interface plate corrosion | Interface plate corrosion, borrow 902-972 one | Repair after replacement 902-972 (By Ran)
Yes | Visit | Complaint | RCHN | 前面板坏了 | 前面板坏了，借886-571一个 | 更换886-571一个正常。(By Ran) Short of spare parts. Opened a new SWO. | SWO-205502 | Re classified RMED | OOB |  |  |  | The front panel is broken. | The front panel is broken, borrow 886-571 one. | 更换886-571一个正常。 (By Ran) Short of spare parts. Opened a new SWO.
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | PCO2质控值高于统计范围0589 | 更换测试卡正常。 | SWO-217963 | Assisted Replacement | SC | SC | pCO2 | QC high - 589 | Test card, reagent package failure | PCO2 QC value is higher than statistical range 0589 | The replacement test card is OK.
No | Visit | Complaint | RESP | Coágulo interno. | coagulo detectado | Se coloco correctamente el sistema de entrada de muestras, se calibro y se comprobó el correcto funcionamiento del equipo mediante QC. | SWO-218458 | Cleaned | Fluidic pathway | Fluidic pathway | inlet probe | clot - 1169 | Internal clot. | clot detected | The sample entry system was correctly placed, calibro and the correct operation of the equipment was checked by QC.
No | Visit | Complaint | RCHN | 报3415，hct板和排线板腐蚀 | 报3415，hct板和排线板腐蚀需要借件636-561和903-329 |  | SWO-217668 | Defect | HW | Aspiration Module | Needle Wash Unit | 3415 | Reported 3415, hct plate and line plate corrosion | Reported 3415, hct plate and line plate corrosion needs to borrow 636-561 and 903-329
No | Visit | Complaint | RCHN | 定标错误 | 定标错误，借944-212（12862） | 重新定标后修复(By Ran) Recal, OK | SWO-216689 | Defect | HW | Process Module | Optical Unit | 1024 | Calibration error | Bid error, borrow 944-212 (12862) | Re-calibration repair (By Ran) Recal, OK
No | Visit | Complaint | RCHN | 光学单元超范围 | 光学单元超范围，借903-364 |  | SWO-217541 | Defect | HW | Process Module | Optical Unit |  | Optical unit over range | Optical unit over range, 903-364
No | Visit | Complaint | RCHN | HCT值无法测试 | HCT值无法测试，借备件903-329一个 |  | SWO-217655 | Defect | HW | Aspiration Module | Needle Wash Unit | 1258 | HCT value cannot be tested | HCT value cannot be tested, spare part 903-329 one
No | 9 | Y | RAME | G pH, Ca, and K constanly failing 2-point calibration. |  | Examined data logs, measuring chamber, and reference electrode. | D2ED4FE9E2460DE2 | Assisted Replacement | Electrode | Electrode | Reference |  | G pH, Ca, and K constanly failing 2-point calibration. |  | Examined data logs, measuring chamber, and reference electrode.
No | Visit | Complaint | RHUN | poros | Autocheck nem olvas | Poros a lézer leolvasó, tisztítás után megfelelően működik. Kalibrálás, QC mérés. | SWO-217805 | Cleaned | HW | Autocheck module | Barcode reader, Autocheck Module |  | Pores | AutoCheck not read | Dusty the laser scanner, after cleaning it works properly. Calibration, QC measurement.
Yes | Visit | Complaint | RIND | pH measurement Unstable. | pH measurement Unstable for all the samples. | Cleaned the electrode with protein removal and electrode cleaning solution but problem was remaining same, need to replace with new pH electrode. | SWO-217628 | Re classified RMED | Electrode | Electrode | pH | Next, Same Day visit | pH measurement Unstable. | pH measurement Unstable for all the samples. | Cleaned the electrode with protein removal and electrode cleaning solution but problem was remaining same, need to replace with new pH electrode.
No | Visit | Complaint | RIND | pH measurement Unstable. | pH measurment Unstable. | Cleaned the electrode with protein removal and electrode cleaning solution, but problem was remaining same, replaced the Standby pH electrode calibrated found results are OK, Old pH Electrode No. XB-05 performed the samples results are ok for pH measurement, found instrument is working satisfactorily. | SWO-217785 | Assisted Replacement | Electrode | Electrode | pH |  | pH measurement Unstable. | pH measurment Unstable. | Cleaned the electrode with protein removal and electrode cleaning solution, but problem was remaining same, replaced the Standby pH electrode calibrated found results are OK, Old pH Electrode No. XB-05 performed the samples results are ok for pH measurement, found instrument is working satisfactorily.
No | 6 | Y | RAME | V VALVES (3) IN THE FLUIDIC MODULE SHOWED SIGNS OF CRYSTALIZATION HAVING FORMED AROUND THE BODY OF THE VALES. ANALYZER INTERMITTENTLY FAILING RINSE ERROR. |  | FOUND THREE VALVES WITH CORROSION AND FORMATION OF RESIDUE CRYSTALIZING AROUND THE BODY OF THE VALVES. | D2ED57DCE224DF41 | Defect | HW | Basic Unit | Fluidic Module |  | V VALVES (3) IN THE FLUIDIC MODULE SHOWED SIGNS OF CRYSTALIZATION HAVING FORMED AROUND THE BODY OF THE VALES. ANALYZER INTERMITTENTLY FAILING RINSE ERROR. |  | FOUND THREE VALVES WITH CORROSION AND FORMATION OF RESIDUE CRYSTALIZING AROUND THE BODY OF THE VALVES.
No | 3 | Y | RAME | 9 LACTATE TAKES DAYS USUALLY FOUR TO FIVE TO PASS LEVEL THREE |  | ALL SYSTEMS CHECKED OUT | D2ED87F8E239830E | Assisted Replacement | Electrode | Electrode | Lac |  | 9 LACTATE TAKES DAYS USUALLY FOUR TO FIVE TO PASS LEVEL THREE |  | ALL SYSTEMS CHECKED OUT
Yes | 10 | Y | RFRA | Blank | Annualtion mesure echantillon |  | D2ED50F0E2292AA8 | [5 empty] | Blank | Annualtion sample measurement
No | Visit | Complaint | RLTD | Andy Luker on 0207 3518662 ex 2767 or bleep 7253 Derek Proudfoot 02073528121 Bleep 7253 | On arrival battery showing as 0% charged even though connected to mains supply. Replaced faulty battery unit serial no. P0073N008 with new unit. Verified analyser operation on restart - all ok. | Replaced battery unit | SWO-218489 | Defect | HW | Battery module | Battery | 0% | Andy Luker on 0207 3518662 ex 2767 or bleep 7253 Derek Proudfoot 02073528121 Bleep 7253 | On arrival battery showing as 0% charged even though connected to mains supply. Replaced faulty battery unit serial no. P0073N008 with new unit. Verified analyser operation on restart - all ok. | Replaced battery unit
No | Visit | Complaint | RLTD | 07887 428 417 - Hannah, 01865 221 595 Rosie steel. | Bilirubin not being reported on blood samples. Checked the report layout and ctBil was on the layout. Checked that the OUH Standard layout was the default for both syringe and capillary samples, which it was. Checked ctBil is being reported on QCs and it is. Checked all the other parameters and found the FHbF was not enabled. | Enabled FHbF but was unable to run a blood sample to check to see if ctBil was reported. The unit is going to report back to clinical measurements when they have run a sample to say whether ctBil is reported. | SWO-218154 | Setup | SW | Report layout | ctBil | HbF enabled | 07887 428 417 - Hannah, 01865 221 595 Rosie steel. | Bilirubin not being reported on blood samples. Checked the report layout and ctBil was on the layout. Checked that the OUH Standard layout was the default for both syringe and capillary samples, which it was. Checked ctBil is being reported on QCs and it is. Checked all the other parameters and found the FHbF was not enabled. | Enabled FHbF but was unable to run a blood sample to check to see if ctBil was reported. The unit is going to report back to clinical measurements when they have run a sample to say whether ctBil is reported.
No | 7 | Y | RAME | z sensor door will not shut durring sensor change |  | checked pins went into service and opened and shut the door checked allignment with bottom screw | D2EE68A05820EF93 | Adjustment | HW | Basic Unit | SC Door |  | z sensor door will not shut durring sensor change |  | checked pins went into service and opened and shut the door checked allignment with bottom screw
Yes | 15 | Y | RAME | 3 - PO2 ERRORS IN LOGS AND SOFTWARE UPGRADE. |  | REVIEWED LOGS AND FOUND PO2 ERRORS BEGAN WITH SENSOR CASSETTE REPLACEMENT ON 7 $25. PO2 ERRORS CONTINUED THROUGH MULTIPLE SOLUTION PACK REPLACEMENTS AND ANOTHER SENSOR CASSETTE REPLACEMENT. 5 BLOOD SAMPLES RAN ON 8 $1 FACILITATED QC STABILITY AND NO FAILURES FOR 48 HOURS. | D2EEA35B58449B89 | Re classified RMED | SW upgrade |  |  |  | 3 - PO2 ERRORS IN LOGS AND SOFTWARE UPGRADE. |  | REVIEWED LOGS AND FOUND PO2 ERRORS BEGAN WITH SENSOR CASSETTE REPLACEMENT ON 7 $25. PO2 ERRORS CONTINUED THROUGH MULTIPLE SOLUTION PACK REPLACEMENTS AND ANOTHER SENSOR CASSETTE REPLACEMENT. 5 BLOOD SAMPLES RAN ON 8 $1 FACILITATED QC STABILITY AND NO FAILURES FOR 48 HOURS.
No | Visit | Complaint | RDEU | 03.08.2017, 09:45 Uhr ==> Anrufer Schw. Ute und Schw. Daniela, Tel.: 0221467771750 ==> Benutzerintervention erforderlich | Vakuumfehler | Fehlersuche durchgeführt. Schlauch zwischen der sensorkassette und dem Hemolyzer hat ein Loch. Schlauch aus JW Kit erneuert. | SWO-218345 | Defect | HW | Oxi module | Pinch valve tube | vacuum | 03.08.2017, 09:45 ==> Caller Sweto Ute and Schw. Daniela, Tel.: 0221467771750 ==> User intervention required | Vakuumfehler | Troubleshooting. Hose between the sensor cassette and the Hemolyzer has a hole. Hose from JW Kit renewed.
No | 8 | Y | RAME | Barcode scanner not working. |  |  | D2EE9F0D5830A516 | Defect | HW | Barcode scanner | Barcode Scanner | Internal | Barcode scanner not working.
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | Lac定标不过 | 更换测试卡后修复 | SWO-219090 | Assisted Replacement | SC | SC | Lac |  | Test card, reagent package failure | Lac calibration is not | Repair after replacing test card
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | Lac定标不过 | 更换测试卡后修复 | SWO-219092 | Re classified RMED | Double entry & OOB |  |  |  | Test card, reagent package failure | Lac calibration is not | Repair after replacing test card
Yes | Visit | Complaint | RCHN | 报错1024,1025,1031错误 | 报错1024,1025,1031错误，借903-364和944-394（12726）各一个 | 更换903-364和944-394后修复(By Ran) Changed the optical unit. Recal, OK. | SWO-216866 | Re classified RMED | HW | Process Module | Optical Unit | 1031 | Error 1024,1025,1031 error | Error 1024,1025,1031 errors, borrowing 903-364 and 944-394 (12726) each | 更换903-364和944-394后修复(By Ran) Changed the optical unit. Recal, OK.
No | Visit | Complaint | RIND | Instrument is showing Oximetry Hardware error. | Checked instrument and done complete oximetry test. Instrument is showing Neon Lamp intensity is high error.Also done Thb calibration procedure but calibration is failed. | Need to replace oximetry Unit. | SWO-218416 | Revisit required | HW | Oxi module | Lamp unit | Neon lamp | Instrument is showing Oximetry Hardware error. | Checked instrument and done complete oximetry test. Instrument is showing Neon Lamp intensity is high error. Also done Thb calibration procedure but calibration is failed. | Need to replace oximetry Unit.
No | 7 | Y | RAME | 2 pCO2 is failing out of peer 2 range for level 1 qc on WDC. it is high outside of peer 2 range. |  | found that all levels of qc are running high on wdc but level 1 is out. | D2EF134D580EA9D9 | Setup | SW | QC-limits | QC-Limits |  | 2 pCO2 is failing out of peer 2 range for level 1 qc on WDC. it is high outside of peer 2 range. |  | found that all levels of qc are running high on wdc but level 1 is out.
No | Visit | Complaint | RKK | ウィルスチェック | ウィルスチェック | ウィルスチェック | SWO-219903 | Setup | SW | Virus Check | Virus Check |  | Virus Check | Virus Check | Virus Check
No | 4 | Y | RAME | Z DMS not connected to AU error |  | checked network connection, checked software | D2EF128658548B5A | Revisit required | HW | Basic Unit | PCB, Analyzer control |  | Z DMS not connected to AU error |  | checked network connection, checked software
No | Visit | Complaint | RDEU | 25.07.2017 11:30 Fr. Benz : Interventionsmeldung; Gasdruck | Gas Unit verunreinigt, Gas1 Flasche lässt sich Schwer einschrauben. | Gas Unit erneuert, Gas 1 und Gas 2 Ventil erneuert. | SWO-218401 | Defect | HW | Gas module | Valve Gas | Does not correspond to Error 509, which is pH/BG upper liquid sensor | 25.07.2017 11:30 Fr. Benz : Intervention notification; Gas | Gas Unit contaminated, Gas1 bottle is difficult to screw in. | Gas Unit renewed, Gas 1 and Gas 2 valve renewed.
No | Visit | Complaint | RDEU | Oximetry Hardware fehler | 258 Neon Lampe intensity too low. 316 Neon nummber too low | Oxi Compl. test gescheitert. Neustart bringt keine Veränderung. Oxi- Module getauscht. | SWO-218659 | Defect | HW | Oxi module | Oxi module, Replaced | Neon lamp | Oximetry Hardware fehler | 258 Neon Lampe intensity too low. 316 Neon nummber too low | Oxi Compl. test failed. Restart brings no change. Oxi modules.
No | Visit | Complaint | RDEU | 03.08.2017, Herr Penzock, Tel: 030450675530: Neon number to low; das Gerät ist erst am Montag ausgeliefert worden.. | Kalarchiv durchgesehen- in den letzten Tagen mehrfach Neon Number to low Fehler | OxiCompleteTest - Neon Position extrem knapp an unterer Grenze, Oximodul ersetzt, Software aktualisiert, justiert und kalibriert, Kal, tHb Kal, QC | SWO-218578 | Defect | HW | Oxi module | Oxi module, Replaced | Neon lamp, Low New analyzer | 03.08.2017, Mr Penzock, Tel: 030450675530: Neon number to low; the device was not delivered until Monday. | Kalarchiv reviewed - in the last few days several Neon Number to low error | OxiCompleteTest - Neon position extremely tight at the lower limit, Oximodul replaced, software updated, adjusted and calibrated, Kal, tHb Kal, QC
Yes | Visit | Complaint | RKK | ウィルスチェック | ウィルスチェック | ウィルスチェック | SWO-219905 | Re classified RMED | Virus check |  |  |  | Virus Check | Virus Check | Virus Check
No | Visit | Complaint | RDEU | 03.08.2017 Herr Frerk 05219436623 meldet, das er immer wieder aufgefordert wird das LP zu Wechseln. | Leckage erkannt | Ventil überholt, Ventilschlauch, Anschlussdichtung und inletgasket ersetzt, SP erneuert, Gerät kalibriert und mit QCs überprüft, OK | SWO-218951 | Defect | HW | Basic Unit | PM kit | Leak detected | 03.08.2017 Mr. Frerk 05219436623 reports that he is repeatedly asked to change the LP. | Leakage detected | Valve overhauled, valve hose, connection seal and inletgasket replaced, SP renewed, device calibrated and checked with QCs, OK
No | Visit | Complaint | RDEU | 02.08.2017, Kundin, tel.: 035867621528: AQC Fehler bei ctHb und Kalibrationsfehler bei pO2 04.08.2017; Fr. Klein; Tel.: 035867621528; Oxy QC Messfehler Level B und Level C | Oxi completete Test durchgeführt --> i.O | Div QC und Pat-Etgebnisse mit Oxispectrum mismatch--> ABL ist nach aus $ ein wieder fehlerfrei --> Oxi complete Test i.O Fehler tritt seit einigen Tagen gehäuft auf Anwender verwendet Probennehmer von Kabe und Radiometer Ersatzteile werden bestellt --> Labor ist informiert | SWO-218800 | Revisit required | HW | Oxi module | parts ordered | 581 | 02.08.2017, customer, tel.: 035867621528: AQC error with ctHb and calibration error at pO2 04.08.2017; Ms. Klein; Tel.: 035867621528; Oxy QC Measurement Error Level B and Level C | Oxi completete test performed --> i.O | Div QC and Pat-Results with Oxispectrum mismatch--> ABL is again error-free after a series of Error has been accumulating for several days User uses samplers of kabe and radiometer Spare parts are ordered --> Laboratory is informed
No | Visit | Complaint | RLTD | Please call Sandra on 02392 286 000 ext: 4326 | Leak detected and oxi spectrum mismatch errors point towards a hole in the pinch valve tube. Removed and inspected the tube and there was a hole in it. | Replaced the pinch valve tube with a new one. | SWO-218773 | Defect | HW | Oxi module | Pinch valve tube | Leak detected | Please call Sandra on 02392 286 000 ext: 4326 | Leak detected and oxi spectrum mismatch errors point towards a hole in the pinch valve tube. Removed and inspected the tube and there was a hole in it. | Replaced the pinch valve tube with a new one.
No | Visit | Complaint | RLTD | The battery status is showing 0%. | Checked service programs. The display is showing 0% for the battery and the charging light is flashing. Battery status is showing it is charging, cycle count is 4 and percent remaining is 0%. Removed the power lead and observed the status change to discharging. When the power was restored the charging state stayed as discharging. | Turned off the Analyser, removed and reset the battery. Replaced the battery and tested again in service programs. The battery is now showing 100%, cycle count is 0 and is charging. Turned off the mains power, display now shows discharging and battery level is dropping. Turned on the mains power, charging state shows charging, charging light is flashing and stops when it is at 100%. | SWO-219356 | Adjustment | HW | Battery module | Battery | 0% | The battery status is showing 0%. | Checked service programs. The display is showing 0% for the battery and the charging light is flashing. Battery status is showing it is charging, cycle count is 4 and percent remaining is 0%. Removed the power lead and observed the status change to discharging. When the power was restored the charging state stayed as discharging. | Turned off the Analyser, removed and reset the battery. Replaced the battery and tested again in service programs. The battery is now showing 100%, cycle count is 0 and is charging. Turned off the mains power, display now shows discharging and battery level is dropping. Turned on the mains power, charging state shows charging, charging light is flashing and stops when it is at 100%.
Yes | 2 | Y | RPAC | Blank | Repeated failed calibrations - In homogeneous cal 3 solution |  | D2F033A3CE314D73 | [5 empty] | Blank | Repeated failed calibrations - In homogeneous cal 3 solution
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | LAC定标不过 | 更换测试卡后修复 | SWO-218541 | [5 empty] | Test card, reagent package failure | LAC calibration is not | Repair after replacing test card
No | Visit | Complaint | RCHN | 光学单元超范围 | 光学单元超范围，借903-364,903-236 | 更换903-364后修复，903-236没有使用(By Ran) | SWO-218021 | Defect | HW | Process Module | Optical Unit |  | Optical unit over range | Optical unit over range, 903-364, 903-236 | Replacement 903-364 repair, 903-236 not used (By Ran)
No | Visit | Complaint | RDEU | 03.08.2017, Herr Kohl, Tel: 03441860814: Intervention erforderlich 03.08.2017, Herr Kohl, Tel.: 03441860814: weiterhin Intervention erforderlich | Intervention erforderlich Beim Kalibrationsvorgang baut sich Druck im pH/BG Modul auf und führt zum Abbruch der Kalibration | Fehlersuche und diverse Tests mit Ventilen, Verschlauchung, anderem Inletmodul und Rollenkränzen Austausch der Rollenkränze der Abfall- und Lösungsstrecke Airventil ersetzt Verschlauchung der Messmodule erneuert Beim Kalibrationsvorgang baut sich Druck im pH $BG Modul auf und führt zum Abbruch der Kalibration Gas Unit am Transducer verschmutzt Gas1 wird der Füllstand gemessen und gefördert Gas2 wird der Füllstand gemessen und gefördert Rechte Stahlkapillare der Gas Unit ist dicht Rechten Ausgang am Dualvalvehouse des Airventils im pH $ BG Modul abgezogen - Kalibration wird vollständig durchgeführt Austausch der Gas Unit nötig Herr Roßbach tauscht diese am Montag Herr Kohl ist über weiteres vorgehen informiert Gerät zur Zeit nicht Messbereit | SWO-218856 | Revisit required | HW | Basic Unit | Gas module |  | 03.08.2017, Mr Kohl, Tel: 03441860814: Intervention required 03.08.2017, Mr Kohl, Tel.: 03441860814: continued intervention required | Intervention needed During the calibration process, pressure builds up in the pH/BG module and leads to the termination of the calibration | Troubleshooting and various tests with valves, hoses, other inlet modules and roller cranes Replacement of the roller wreaths of the waste and solution line Air valve replaced Hoseing of the measuring modules renewed During the calibration process, pressure builds up in the pH $BG module and leads to the termination of the calibration Gas Unit polluted at the Transducer Gas1 the level is measured and conveyed Gas2 the level is measured and conveyed Right steel capillary of the gas unit is dense Right output at the dual valve house of the air valve is subtracted in the pH BG module - calibration is carried out completely Replacement of the gas unit necessary Mr. Roßbach will exchange them on Monday Mr Kohl is informed of further action Device currently not ready for measurement
No | 7 | Y | RAME | Z File checkSum error, instrument will not boot |  | checked all connection, checked computer board and hard drive | D2F0556FCE558B62 | Revisit required | HW | DMS_PC | HDD |  | Z File checkSum error, instrument will not boot |  | checked all connection, checked computer board and hard drive
No | Visit | Complaint | RDEU | 04.08.2017, 11:15 Uhr ==> Anrufer Fr. Dill, Tel.: 071314946415 ==> pO2-Nullpunktfehler und häufig schwarzer Bildschirm (Softwareabsturz) | PO2 Nullwertfehler bei Kal. und spontaner Ausfall des Monitors , lässt sich spontan nicht mehr bedienen. | Monitor testweise getauscht von Demogerät GS17. PO2 Elektrode erneuert Nullwert zu tief trotz Nullpunktkallibrierung im Servicemenue. Mehrere Kal. durchgeführt. pO2 Nullwert jetzt stabil bei 1,1. | SWO-219111 | Defect | HW | DMS_PC | Touch screen | Blank | 04.08.2017, 11:15 a.m. ==> Caller Ms. Dill, tel.: 071314946415 ==> pO2 zero point error and frequent black screen (software crash) | PO2 Null value error at Kal. and spontaneous failure of the monitor, can no longer be operated spontaneously. | Monitor tested by demo device GS17. PO2 electrode renews zero value too deep despite zero-point cacanning in the service menu. Several Kal. Performed. pO2 zero value now stable at 1.1.
No | Visit | Complaint | RDEU | 02.08.2017, 09:20 Uhr ==>Tel.: 03533603351 ==> Fehler 0361: Neon lamp voltage too low, Oxi-Hardwarefehler 02.08.2017, 12.00Uhr, Sr. Michaels: 03533603351: weiterhin Oximetrie-Hardware Fehler 316 | Immer wiederkehrend Fehler 579, 581 und 316 | Neon Position Wert bei 0,102 sehr niedrig (Tausch der Oxymodul PCB brachte keine Besserung), Oxymodul ersetzt, AU Software eingespielt, Kal, Service tHb Kal, Kal, Anwender tHb Kal, Probemessung, QC ausgebautes defektes Oxymodul 902-833R2127N009 | SWO-218866 | Defect | HW | Oxi module | Oxi module, Replaced | Neon Lamp | 02.08.2017, 09:20 ==>Tel.: 03533603351 ==> Error 0361: Neon lamp voltage too low, Oxi hardware error 02.08.2017, 12.00 pm, Sr. Michaels: 03533603351: continue Oximetry Hardware Error 316 | Recurring Errors 579, 581 and 316 | Neon position value at 0.102 very low (exchange of oxymodul PCB brought no improvement), Oxymodul replaced, AU software plugged in, Kal, service tHb Kal, Kal, user tHb Kal, sample measurement, QC removed defective oxymodule 902-833R2127N009
No | 1 | Y | RPOL | Blank | Wymieniono membranê LAC, kalibracja, testy analizatora. Aparat sprawny. |  | D2EF9DEBCE514FFE | Assisted Replacement | Electrode | Membrane | Lac |  | Blank | The LAC membrane, calibration, analyzer tests are listed. The camera is functional.
No | Visit | Complaint | RCHN | 传感器定标失败 | 传感器定标失败，借备件903-199,903-644各一个 | 更换903-199一个正常。(By Ran) | SWO-218676 | Defect | HW | Basic Unit | Sensor Cassette retainer |  | Sensor calibration failed | Sensor calibration failed, with spare parts 903-199,903-644 each | Replace 903-199 one normal. (By Ran)
No | Visit | Complaint | RCHN | 测试卡卡口断裂 | 测试卡卡口断裂 | 更换测试卡正常。 | SWO-223262 | Defect | Consumables | Test Kit | Cartridge | Broken | Test ingress card mouth break | Test ingress card mouth break | The replacement test card is OK.
No | Visit | Complaint | RCHN | 气泡液体超出范围 | 气流温度超出范围，加借加一个903-209 | cup dryer pcb 和inlet pcb 上所有排线拔插(By Ran) | SWO-218733 | Defect | HW | Process Module | PCB, Cup dryer |  | Bubble liquid out of range | Airflow temperature out of range, plus a 903-209 | All line plugs on cup dryer pcb and inlet pcb (By Ran)
No | Visit | Complaint | RCHN | 仪器做样本时报错5698 | 仪器做样本时报错5698 | 更换测试卡后修复 | SWO-219147 | Defect | Consumables | Test Kit | Cartridge | 5698 | Instrument to do sample times wrong 5698 | Instrument to do sample times wrong 5698 | Repair after replacing test card
No | Visit | Complaint | RCHN | 气流温度超范围 | 气流温度超范围 借903-210一个。 | 更换903-210一个正常。(By Ran) | SWO-218969 | Defect | HW | Process Module | PCB, Cup dryer | 1068 | Airflow temperature over range | Airflow temperature over range Borrow 903-210 one. | Replace the 903-210 one normal. (By Ran)
No | Visit | Complaint | RDEU | Die Aushilfe meldet das der Drucker ohne Funktion ist. | Drucker soll nicht gehen. | Analysator ausgeschaltet und aufgeschraubt. Alle Kabelanschlüsse für den Drucker auf festen Sitz geprüft. Drucker vom Staub befreit. Analysator wieder hochgefahren und Testdruck durchgeführt alles OK. | SWO-219251 | Adjustment | HW | Printer Module | cables reinstalled | 2nd hand - UKE | The helper reports that the printer is without function. | Printer should not go. | Analyzer switched off and screwed on. Checked all cable connections for the printer for fixed fit. printer cleared of dust. Analyzer restarted and test pressure performed all OK.
No | 3 | Y | RAME | R Battery pack will not charge. Charge always shows 0%. Found no type of battery error displayed on the analyzer. |  | Checked logs for any battery error but found none. | D2F3954D3004F5E4 | Defect | HW | Battery module | Battery | 0% | R Battery pack will not charge. Charge always shows 0%. Found no type of battery error displayed on the analyzer. |  | Checked logs for any battery error but found none.
No | Visit | Complaint | RDEU | 02.08.2017, Kundin, tel.: 035867621528: AQC Fehler bei ctHb und Kalibrationsfehler bei pO2 04.08.2017; Fr. Klein; Tel.: 035867621528; Oxy QC Messfehler Level B und Level C | Siehe SWO | ABL steht beim Eintreffen im Halt --> Lösungen können nicht gepumpt werden Hemolyzer und div. Schläuche erneuert--> Neustart--> Oxi complete Test i.O --> tHb-Kal Testmessung | SWO-219398 | Defect | HW | Oxi module | Hemolyzer | 581 | 02.08.2017, customer, tel.: 035867621528: AQC error with ctHb and calibration error at pO2 04.08.2017; Ms. Klein; Tel.: 035867621528; Oxy QC Measurement Error Level B and Level C | See SWO | ABL is in stop-by-stop on arrival --> Solutions cannot be pumped Hemolyzer and various hoses renewed--> Restart--> Oxi complete Test i.O --> tHb-Kal Test measurement
Yes | 3 | Y | RAME | Z screen went black PER CUSTOMER |  | COULD NOT VERIFY ISSUE WHEN ONSITE. ANALYZER TURNED ON AND WAS ABLE TO PERFORM SYSTEM CYCLE NORMALLY. | D2F387C4304C1632 | [5 empty] | Z screen went black PER CUSTOMER |  | COULD NOT VERIFY ISSUE WHEN ONSITE. ANALYZER TURNED ON AND WAS ABLE TO PERFORM SYSTEM CYCLE NORMALLY.
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | LAC定标失败 | 更换测试卡后修复 | SWO-219225 | [5 empty] | Test card, reagent package failure | LAC calibration failure | Repair after replacing test card
No | Visit | Complaint | RDEU | 03.08.2017, Herr Kohl, Tel: 03441860814: Intervention erforderlich 03.08.2017, Herr Kohl, Tel.: 03441860814: weiterhin Intervention erforderlich | Kalibrierungen werden abgebrochen Gas Unit defekt | Austausch Gas Unit aus GS Bestand Wechsel der Gasflaschen Gasnachfüllung i.O. Schlauchnachfüllung, LS-Justage und PumpenKal i.O. Mehrmals Kal => immer wieder Abbruch an gleicher Stelle wie durch JorPr berichtet Vermutlich Gas1 oder Gas2 Ventil defekt | SWO-219370 | Defect | HW | pH/BG Module | Valve Gas |  | 03.08.2017, Mr Kohl, Tel: 03441860814: Intervention required 03.08.2017, Mr Kohl, Tel.: 03441860814: continued intervention required | Calibrations are canceled Gas Unit defective | Exchange gas unit from GS stock Changing the gas cylinders Gas refilling as a.o. Hose refilling, LS-adjustment and pumpsKal i.O. Several times Kal => repeatedly cancel in the same place as reported by JorPr Probably Gas1 or Gas2 valve defective
Yes | Visit | Complaint | RSCH | Reparatur | Demo key's Elektrolyte und Metabolyte abgelaufen | Key's aus Archiv installiert i.o Nach mehreren Kal. K, Na, Ca instabil Lac bei QC instabil Membranenwartung Ca, K, Na und Lac durchgeführt i.o Kal. und QC i.o Blutmessung i.o Daten werden ins Labor gesendet i.o | SWO-219328 | Re classified RMED | SW | Demo-Key | Demo-Key |  | Repair | Demo key's electrolytes and metabolyte expired | Key's from archive installed i.o After several Kal. K, Na, Ca unstable Lac unstable at QC Membrane maintenance Ca, K, Na and Lac carried out i.o Kal. and QC i.o Blood measurement i.o Data is sent to the laboratory i.o
No | Visit | Complaint | RIND | Rinse error. | Check the activity log & found inlet flow impeded. | Clean the rinse flow path,clean the gasket,found that waste tubing was pinched,after that restart the instrument ,run the protein removal & cleaning few times,run a patient sample,got the desired result,handover the instrument to the customer. | SWO-221161 | Defect | HW | Basic Unit | Waste Tube | Pinched | Rinse error. | Check the activity log & found inlet flow impeded. | Clean the rinse flow path,clean the gasket,found that waste tubing was pinched,after that restart the instrument ,run the protein removal & cleaning few times,run a patient sample,got the desired result,handover the instrument to the customer.
No | 8 | Y | RAME | 2 user called stating that pCO2 out on qc according to peer2 range on level 1 qc S7335 |  | Found that the new electrode would not calibrate keeps drifting on status and sensitivity. | D2F40E963060B846 | Assisted Replacement | Electrode | Electrode | pCO2 |  | 2 user called stating that pCO2 out on qc according to peer2 range on level 1 qc S7335 |  | Found that the new electrode would not calibrate keeps drifting on status and sensitivity.
No | 4 | Y | RAME | Z At main screen page is displayed up 1 inch up from normal position. |  | Display is calibrated and operates normally. Instrument functions normally and display is in correct position for other menu pages. | D2F3E0E4300AAB1E | Adjustment | HW | DMS_PC | Touch screen | Touch | Z At main screen page is displayed up 1 inch up from normal position. |  | Display is calibrated and operates normally. Instrument functions normally and display is in correct position for other menu pages.
Yes | Visit | Complaint | RLTD | Lisa Tyler Tel 01865 221596 | The permanent option keys had not been installed when the analyser was installed and the demo keys have now expired. | Installed the permanent option keys and ran calibrations & QCs to get the analyser back to ready. | SWO-219652 | Re classified RMED | SW | Demo-Key | Demo-Key | Kontakt RLTD | Lisa Tyler Tel 01865 221596 | The permanent option keys had not been installed when the analyser was installed and the demo keys have now expired. | Installed the permanent option keys and ran calibrations & QCs to get the analyser back to ready.
No | 6 | Y | RAME | Z DMS not connected to AU error |  | checked network connection, checked software | D2F4147430008B6D | Defect | HW | DMS_PC | PCB, Main Board |  | Z DMS not connected to AU error |  | checked network connection, checked software
No | 15 | Y | RAME | 0 Customer reporting difference of 1gdl between ABL820 and ABL80COOX during sample comparison testing. |  | Reviewed datalogs, QC is recovering well within specifications. Observed tHB pA @ nominal 486. | D2F3EF18303AC470 | Adjustment | HW | Oxi module | Lamp unit |  | 0 Customer reporting difference of 1gdl between ABL820 and ABL80COOX during sample comparison testing. |  | Reviewed datalogs, QC is recovering well within specifications. Observed tHB pA @ nominal 486.
No | Visit | Complaint | RCHN | 打印机无法打印结果 | 打印机无法打印结果，借910-286一个 |  | SWO-219086 | Defect | HW | Analyzer Control Module | Printer module |  | The printer could not print the results | The printer cannot print the results, borrow 910-286 one
No | 4 | Y | RAME | P Pump calibration error detected. |  | The analyzer keeps going into intervention required presenting message 1178 A pump calibration error was detected. | D2F4F027A606719B | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway | 1178 | P Pump calibration error detected. |  | The analyzer keeps going into intervention required presenting message 1178 A pump calibration error was detected.
No | 4 | Y | RAME | P Unable to pump solutions error 1178 |  | The analyzer keeps going in to intervention required presenting error 1178. The analyzer kept also asking for a flush, replace the sensor cassette and, solution pack. | D2F4ECBBA6107195 | Replaced | HW | Inlet module | Probe bend | 1178 | P Unable to pump solutions error 1178 |  | The analyzer keeps going in to intervention required presenting error 1178. The analyzer kept also asking for a flush, replace the sensor cassette and, solution pack.
No | Visit | Complaint | RDEU | 07.08.2017, 10:40 Uhr, H. Fritsch, Tel. 03623350230: Touch Screen reagiert nicht 08.08.2017, 10:02 Uhr, Fr. Neidardt, Tel. 016098735800: Touch Screen reagiert nicht | Fehlersuche | Touch screen erneuert | SWO-219700 | Defect | HW | DMS_PC | Touch screen |  | 07.08.2017, 10:40, H. Fritsch, tel. 03623350230: Touch screen not responding 08.08.2017, 10:02, Fr. Neidardt, tel. 016098735800: Touch Screen not responding | Troubleshooting | Touch screen renewed
No | 8 | Y | RAME | 0 Instrument external QC is out of control. |  | COHb & sO2 are the only reported parameters and both failed external QC for all levels. Instrument SP had been replaced 8-3 and user experienced some startup issues but instrument came up and operated normally. Inspection of data logs shows all system cycle results to be normal. | D2F4E61BA60B97D9 | [5 empty] | 0 Instrument external QC is out of control. |  | COHb & sO2 are the only reported parameters and both failed external QC for all levels. Instrument SP had been replaced 8-3 and user experienced some startup issues but instrument came up and operated normally. Inspection of data logs shows all system cycle results to be normal.
No | Visit | Complaint | RCHN | 验证失败 | 验证失败，借903-334一个 | 更换903-334后修复 by Andrew | SWO-219223 | [5 empty] | Validation failed | Failed to validate, borrow 903-334 one | Repair by Andrew after replacing 903-334
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | K定标失败 | 更换测试卡后修复 | SWO-221007 | [5 empty] | Test card, reagent package failure | K bid failure | Repair after replacing test card
No | 14 | Y | RAME | F All oximetry failing on System Cycles. |  | Upon arrival, analysis was disabled due to all oximetry failing System Cycles. Ran Service Test and the hemolyzer and lamp portions failed. | D3050726A4204A34 | [5 empty] | F All oximetry failing on System Cycles. |  | Upon arrival, analysis was disabled due to all oximetry failing System Cycles. Ran Service Test and the hemolyzer and lamp portions failed.
No | 15 | Y | RAME | P LS errors and leaking |  | the analyzer was leaking and would not pass the LS cal | D2F4A75BA633B6BF | [5 empty] | P LS errors and leaking |  | the analyzer was leaking and would not pass the LS cal
No | Visit | Complaint | RSCH | Rep | Analysator immer mit möglicher Verstopfung. Und / oder Luft in der Probe | Schlauch Werkseitig bei Anschluss Ventilblock abgeknickt. Schlauch und Pumpe ausgetauscht. Funktionskontrolle i.o | SWO-219569 | [5 empty] | Rep | Analyzer always with possible constipation. And / or air in the sample | Hose factory side bent at connection valve block. hose and pump. Functional control i.o
No | Visit | Complaint | RIND | sensor cassette failed. | checked and found that the sensor cassette not calibrating properply | done zeroing, reset the SP and also replaced the sensor cassette, now working satisfactorily. | SWO-221149 | [5 empty] | sensor cassette failed. | checked and found that the sensor cassette not calibrating properply | done zeroing, reset the SP and also replaced the sensor cassette, now working satisfactorily.
No | Visit | Complaint | RIND | Na and K value not acceptable. analyser swithed off. | Electrolyte values doubtful . | Analyser Serviced and calibrated properly. Handed over in a fine working condition. Analyser shifting to Pathlab. Training given to staff. Dr's Meeting done. Analyser working fine. Dr ask to give 3 month time to satisfy then relies final order. | SWO-221154 | [5 empty] | Na and K value not acceptable. analyser swithed off. | Electrolyte values doubtful . | Analyser Serviced and calibrated properly. Handed over in a fine working condition. Analyser shifting to Pathlab. Training given to staff. Dr's Meeting done. Analyser working fine. Dr ask to give 3 month time to satisfy then relies final order.
No | Visit | Complaint | RCHN | 定标步骤终止 | 定标步骤终止，借BNP定标卡944-258（12979）一张缺货 | 工程告知，机器已正常，想增加BNP这个项目，现无货的话，就不增加了，故此单关单 by Andrew | SWO-219494 | Defect | HW | Incubator Module | Punch Up | 5140 | Bid step termination | End of bid step, with BNP bid card 944-258 (12979) one out of stock | The project informed that the machine has been normal, want to increase BNP this project, now no goods, it will not increase, so this single single by Andrew
No | Visit | Complaint | RDEU | 13.07.2017 08:40Uhr Herr Kohl 003441860814: Das Gerät meldet Kal-fehler | Siehe SWO | Gas 1 und 2 zur Fehlersuche inst. --> Fehler bleibt bestehen ABL in halt Rep-Ersatzgerät wird vorbereitet Dialyse ist informiert | SWO-219753 | Revisit required | HW | Basic Unit | Gas module |  | 07/13/2017 08:40Am Mr Kohl 003441860814: The device reports Kal error | See SWO | Gas 1 and 2 for troubleshooting inst. --> Error remains ABL in halt Preparing Rep Replacement Device Dialysis is informed
No | Visit | Complaint | RIND | Sample aborted. | Sample aborted & query in all the parameter for few samples. | Found the blood clots in the gasket,clean the gasket ,check all the electrode chamber,nothing found,run the protein removal & cleaning ,run the 2 point calibration,run few sample,got the desired result ,handover the instrument to | SWO-221166 | Cleaned | HW | Fluidic pathway | Fluidic pathway |  | Sample aborted. | Sample aborted & query in all the parameter for few samples. | Found the blood clots in the gasket,clean the gasket ,check all the electrode chamber,nothing found,run the protein removal & cleaning ,run the 2 point calibration,run few sample,got the desired result ,handover the instrument to
No | 9 | Y | RAME | Z File checkSum error, instrument will not boot |  | checked all connection, checked computer board and hard drive | D2F5630DA6628B7F | Defect | HW | DMS_PC | DMS_PC |  | Z File checkSum error, instrument will not boot |  | checked all connection, checked computer board and hard drive
No | Visit | Complaint | RDEU | 08.08.2017, 08:39 Uhr, H. Meiners, Tel. 017618104595: Verstopfung im Abfallweg | Verstopfung im Abfallweg. | Abfallschläuche komplett blutig - Abfallschläuche ausgetauscht und Abfallwege gespült. Das Gerät hat einen Oximetrie Hardware Fehler - Gerät neu gestartet, Kalibriert und QC's durchgeführt. | SWO-219794 | Cleaned | HW | Fluidic pathway | Fluidic pathway |  | 08.08.2017, 08:39, H. Meiners, Tel. 017618104595: Constipation in the waste path | Constipation in the waste path. | Waste hoses completely bloody - waste hoses replaced and waste paths flushed. The device has restarted an oximetry hardware error device, calibrated and QC performed it.
No | 21 | Y | RPNZ | Blank | ABL/DMS PC disconnected from AU |  | D2F53FF4A643AE4F | Adjustment | HW | Basic Unit | PCB, Analyzer control | Reboot | Blank | ABL/DMS PC disconnected from AU
Yes | Visit | Complaint | RLTD | Hanna | ctBil was unavailable because the OXI30 demo key had expired (the permanent key had not been loaded when the analyser was installed) | Loaded the permanent OXI30 key on to the analyser and then ran all levels of QCs to clear the red light on ctBil & replaced Gas 2 as it was low. While in the unit I checked the sensor connection pins on the HDU analyser and they all looked ok, I downloaded the QC log so that the pH failures could be looked at. | SWO-219651 | Re classified RMED | SW | Demo-Key | Demo-Key |  | Hanna | ctBil was unavailable because the OXI30 demo key had expired (the permanent key had not been loaded when the analyser was installed) | Loaded the permanent OXI30 key on to the analyser and then ran all levels of QCs to clear the red light on ctBil & replaced Gas 2 as it was low. While in the unit I checked the sensor connection pins on the HDU analyser and they all looked ok, I downloaded the QC log so that the pH failures could be looked at.
No | Visit | Complaint | RDEU | 27.07.2017; Sr. Anett; Tel.: 03381791614; CRP QC Messfehler 28.07.2017 10:30 Sr. Nicole; Tel.: 03381791612; weiterhin CRP QC Messfehler 31.07.2017: Rücksprache mit KerZi - neues QC wird bestellt, telefonische Info an Kunden 03.08.2017, 11:36 Uhr ==> Anrufer Pfl. Sebastian, Tel.: 03381791611 ==> LQC-Messfehler CRP 04.08.2017, 12:04 Uhr ==> Anrufer Schw. Anett, Tel.: 03381791614 ==> LQC-Fehler CRP besteht trotz neuer Charge weiterhin | Nadel neu kalibriert Process Wheel gereinigt inkl. der Ref. Flächen neue Charge CRP kalibriert, anschließende QC wieder zu hoch | Blank Messungen durchgeführt - Raw Count bei allen 5 Messungen >1000 opt. Einheit wird bestellt | SWO-219878 | Revisit required | HW | Process Module | Optical Unit | LQC | 27.07.2017; Sr. Anett; Tel.: 03381791614; CRP QC Measurement Error 07/28/2017 10:30 Sr. Nicole; Tel.: 03381791612; further CRP QC measurement error 31.07.2017: Consultation with KerZi - new QC is ordered, telephone information to customers 03.08.2017, 11:36 a.m. ==> Caller Pfl. Sebastian, Tel.: 0338179111 ==> LQC measurement error CRP 04.08.2017, 12:04 o'clock ==> Caller Schw. Anett, tel.: 03381791614 ==> LQC error CRP persists despite new batch | Needle recalibrated Process Wheel cleaned incl. the ref. Surfaces new batch CRP calibrated, subsequent QC too high again | Blank measurements performed - Raw Count on all 5 measurements >1000 Opt. Unit is ordered
No | 4 | Y | RAME | Z Calibration error. gas low out of range. In warranty | Changed the complaint visit tab from N to Y. | Was able to duplicate the fault. Replaced COBIM module ran calibration and all passed | D2F515B4A60025A3 | [5 empty] | Z Calibration error. gas low out of range. In warranty | Changed the complaint visit tab from N to Y. | Was able to duplicate the fault. Replaced COBIM module ran calibration and all passed
No | Visit | Complaint | RHUN | PC meghibásodott |  | PC csere. | SWO-220194 | Defect | HW | DMS_PC | DMS_PC |  | PC Defective |  | PC replacement.
No | Visit | Complaint | RCHN | 传感器电路错误 | 传感器电路错误，报错1306，借备,902-972,903-334各一个 | 更换902-972后修复，903-334未使用 by Andrew | SWO-219078 | [5 empty] | Sensor circuit error | Sensor circuit error, error 1306, loan, 902-972, 903-334 each | Replacement 902-972 repair, 903-334 unused by Andrew
No | 14 | Y | RAME | 3 pO2 failing QC |  | Visited yesterday for oximetry errors but, pO2 was failing on qc levels 1 and 2. Revisit to continue troubleshooting. | D305097CA4214A39 | [5 empty] | 3 pO2 failing QC |  | Visited yesterday for oximetry errors but, pO2 was failing on qc levels 1 and 2. Revisit to continue troubleshooting.
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | pO2，GLU定标不过 报错0377 | 更换测试卡后修复 | SWO-220256 | Assisted Replacement | SC | SC | Glu & pO2 | 0377 | Test card, reagent package failure | pO2, GLU calibration but error 0377 | Repair after replacing test card
No | Visit | Complaint | RESP | salta continuamente la pantalla de recambios y ya han cambiado tanto sonda como gasket como reactivos | se sustituye gasket nueva, sonda visiblemente doblada y se ejecutan programas de servicio | sustitución de gasket, sonda y ejecución de programas de servicio | SWO-220102 | Replaced | HW | Inlet module | Probe bend |  | continuously skips the spare parts screen and has already changed both probe and gasket and reagents | new gasket is replaced, visibly bent probe and service programs are executed | gasket substitution, probe and service program execution
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 定标液传输失败 | 更换试剂包后修复 | SWO-219969 | Assisted Replacement | SP | SP | FT |  | Test card, reagent package failure | Calibration fluid transmission failed | Repair after replacement of reagent kit
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 不均匀的定标3溶液 1084 | 更换试剂包后修复 | SWO-221284 | Re classified RMED | same day UpSWO | SWO-219969 |  |  | Test card, reagent package failure | Uneven calibration 3 solution 1084 | Repair after replacement of reagent kit
No | Visit | Complaint | RCHN | 无法读取测试卡芯片数据 | 无法读取测试卡芯片数据，借903-183（现已被903-581替代）一个 | 更换903-581后修复 by andrew | SWO-219653 | Defect | HW | Basic Unit | PCB, Sensor interface |  | Failed to read test card chip data | Failed to read test card chip data, borrow 903-183 (now replaced by 903-581) | Repair by andrew after replacement 903-581
No | Visit | Complaint | RCHN | 光单检测不准确，报错1031 | 光单检测不准确，报错1031，借备件903-364一个 | 更换903-364后修复 by Andrew | SWO-219523 | Defect | HW | Process Module | Optical Unit | 1031 | Light single detection is not accurate, reported error 1031 | Inaccurate light single detection, error 1031, borrowed spare parts 903-364 one | Repair by Andrew after replacing 903-364
No | Visit | Complaint | RCHN | 软件系统不匹配，报错0002 | 软件系统不匹配，借备件903-198一个 | 更换903-198后修复 by Andrew | SWO-219560 | Revisit required | HW | Analyzer Control Module | PCB, Main Board | 0002 | Software system mismatch, error 0002 | Software system mismatch, spare parts 903-198 one | Repair by Andrew after replacement 903-198
No | Visit | Complaint | RDEU | 03.08.2017, Herr Kohl, Tel: 03441860814: Intervention erforderlich 03.08.2017, Herr Kohl, Tel.: 03441860814: weiterhin Intervention erforderlich | Abbruch der Kalibration | Austausch des pH $BG Moduls Modul wird zur Reparatur mitgenommen Elektroden aus Originalmodul umgesetzt Fünf Kalibrationen ohne Fehler durchgeführt Ausstehende QC Messungen erfolgreich durchgeführt | SWO-220055 | Defect | HW | Basic Unit | Gas module |  | 03.08.2017, Mr Kohl, Tel: 03441860814: Intervention required 03.08.2017, Mr Kohl, Tel.: 03441860814: continued intervention required | Cancellation of calibration | Replacing the pH $BG module Module is taken with you for repair Electrodes implemented from original module Five calibrations performed without errors Pending QC measurements successfully performed
No | Visit | Complaint | RDEU | 09.08.2017 09:15Uhr Herr Zierfuss 015776400501: Das Gerät meldet seit 23:00 einen Spülfehler und bricht die Kal ab. Das Gerät ist nicht zu benutzen. | Spülfehler | Defekte Membrane des Airventils auf BG Modul. Ventil ausgebaut und Membrane ersetzt. Initialisierungsphase und Test des Ventils. Kalibrationen durchgeführt. QC ohne Beanstandung | SWO-220198 | Defect | HW | pH/BG Module | Valve Air |  | 09.08.2017 09:15Am Mr. Zierfuss 015776400501: The device has been reporting a flushing error since 23:00 and cancels the calf. The device is not to be used. | Spülfehler | Defective diaphragm of the air valve on BG module. valve removed and membrane replaced. Initialization phase and valve test. calibrations. QC without complaint
No | Visit | Complaint | RDAN | Hanna Andresen har ringet på vagttelefonen at ABL er blank(ingen parameterbar). Der er fejlmddelse 0734 Ikke forbudnet til AU . | AU kortet virker ikke. Inletproben er fundet bukket. | Jeg har skiftet AU kort og strømforsyning. Jeg har skiftet inletprobe. Der er lagt software 6.17 på ABLen. Alla data er genindlæst. Der er kørt 2 punkts kal. Der er kørt tHb kal. Der er kørt QCer. Alt er fundet OK | SWO-220349 | Defect | HW | Basic Unit | PCB, Analyzer control |  | Hanna Andresen has called on the Watchphone that ABL is blank (no parametric). There is error 0734 not forbidden to AU. | The AU card does not work. The inletproben is found bent. | I changed the AU card and power supply. I changed the inletprobe. Software 6.17 has been added to the ABLen. Alla data is reloaded. Two-point Kal has been run. A tHb Kal has been run. QCer has been run. Everything is found OK
No | Visit | Complaint | RDEU | Häufig verbogene Einlßsonde bei ausschließlicher Nutzung FlexQ | Reinigungsarbeiten, Fehlersuche | Justage FlexQ, Probemessungen mit safePICO ok | SWO-220208 | Adjustment | HW | FlexQ Module | FlexQ Module |  | Frequently bent single-use probe with exclusive use FlexQ | Cleaning work, troubleshooting | Adjustment FlexQ, sample measurements with safePICO ok
Yes | 4 | Y | RCZE | Blank | urgují chybu na ls oxi error, opakovanì øešeno pøes telefon ( + Marek pøes dovolenou ), i pøes informaci, že je vpoøádku, požadují opravu |  | D2F62BED1C3D8DBD | Re classified RMED | HW | Oxi module | Oxi module | Phone support | Blank | Urgues an error on LS Oxi error, repeated by the phone (+ Marek via leave), even if the message is in order, require repair
No | 3 | Y | RCZE | Blank | nefunkèní tlaèítko start, doporuèen restart, po provedení OK, ale je stále zpomalená reakce,stanièní požaduje kontrolu displaye, + èasto kalibruje - v nastavení úporava |  | D2F62D501C618DC0 | Defect | SW | 6.16 Win 7 | 6.16 Win 7 |  | Blank | Not function start key, restart recommended, after performing OK, but is still slowed reaction, station requires control of display, + often calibrates-in the settings of the
No | Visit | Complaint | RKK | インレットバーコードを読まない事がある。 | インレットバーコードを読まない事がある | バーコードリーダー交換を実施しました。 | SWO-223731 | Defect | HW | Analyzer Control Module | Barcode reader, display |  | You may not read the inlet bar code. | Sometimes I don't read inlet barcodes. | Bar code reader exchange was carried out.
No | Visit | Complaint | RDEU | 28.07.2017 12:00 Fr. Kirstein 0355543922: QC-B SO2-Fehler | Hämolysator Austausch | Hämolysator ausgebaut und erneuert, Kal und QC | SWO-220797 | Defect | HW | Oxi module | Hemolyzer | QC B - sO2 | 28.07.2017 12:00 Fr. Kirstein 0355543922: QC-B SO2-Fehler | Hemolyser Exchange | Hemolyser removed and renewed, Kal and QC
Yes | Visit | Complaint | RDEU | 28.07.2017 12:00 Fr. Kirstein 0355543922: QC-B SO2-Fehler | Datenspeicher ausgelesen, sporadisch Fehler bei QC2 Oximetrie | Fließwege gereinigt, Ventil geprüft, Ventilschlauch erneuert. Fehler trat weiterhin auf. Hämolysator aus der GS geholt | SWO-220800 | Re classified RMED | same day UpSWO | see SWO-220797 |  |  | 28.07.2017 12:00 Fr. Kirstein 0355543922: QC-B SO2-Fehler | Data memory read out, sporadic error in QC2 oximetry | Flow paths cleaned, valve tested, valve hose renewed. Errors continued to occur. Hemolyser taken out of the GS
No | Visit | Complaint | RSAF | tHb cal reran but FMetHb and Fo2Hb and tBil out of range | Checked | Back flush done and hemolyzer was flushed. Replaced faulty hemolyzer unit and replaced parts as per yearly service. | SWO-221583 | Defect | HW | Oxi module | Hemolyzer |  | tHb cal reran but FMetHb and Fo2Hb and tBil out of range | Checked | Back flush done and hemolyzer was flushed. Replaced faulty hemolyzer unit and replaced parts as per yearly service.
No | Visit | Complaint | RLTD | The Analyser is failing LV B & C QC's | Checked data logs. |  | SWO-220682 | Defect | HW | Oxi module | Hemolyzer | QC B/C | The Analyser is failing LV B & C QC's | Checked data logs.
No | Visit | Complaint | RESP | Varias pérdidas de líquido del Csol. En casos anteriores se le ha cambiado Csol.No mensajes de error. Equipo funciona normalmente. | Se ven restos de líquido rosa. Limpieza con hipoclorito y agua de todos los circuitos neumáticos. Tras tests y hacerle pasar líquidos se sigue viendo restos por la parte del desperdicio. | Se decide llevar equipo a taller Madrid. Desmontaje y limpieza de todos los restos. Realización kit tubos y cambio de puerta Csol (por si acaso no hiciera presión suficiente). Puesta en marcha, tests y ciclos de calidad. Se devuelve a cliente. Queda en observación | SWO-220564 | [5 empty] | Various fluid losses from Csol. In previous cases, you have changed Csol.No error messages. Equipment works normally. | You can see traces of pink liquid. Cleaning with hypochlorite and water of all pneumatic circuits. After testing and having liquids pass through, you still see traces from the waste. | It is decided to bring equipment to Madrid workshop. Disassembly and cleaning of all debris. Realization kit tubes and change of Csol door (just in case it did not make enough pressure). Commissioning, tests and quality cycles. It is returned to the client. It's under observation
No | Visit | Complaint | RCHN | Lac定标不过 | Lac定标不过，借903-135 |  | SWO-220313 | Assisted Replacement | SC | SC | Lac | Sensor simulator ? | Lac calibration is not | Lac bid is not, borrow 903-135
No | Visit | Complaint | RCHN | 血氧硬件问题 | 血氧硬件问题，借903-644,903-197,902-980各一个 | 更换备件902-980和903-197备件后修复 by Andrew | SWO-219996 | Defect | HW | Oxi module | Spectrometer | & Hemolyzer + Wetsection PCB | Blood oxygen hardware problems | Blood oxygen hardware problems, borrow 903-644, 903-197, 902-980 each | Repair by Andrew after replacement of spare parts 902-980 and 903-197 spare parts
No | Visit | Complaint | RCHN | 系统软件不匹配 | 系统软件不匹配,需要借件903-198备件一个 | 工程师表示已经修好了，不需要借件了，取消借件 by Andrew | SWO-220275 | Defect | HW | Analyzer Control Module | PCB, Main Board | 0002 | System software does not match | System software does not match, need to borrow 903-198 spare parts one | The engineer says it's fixed, no loanised, cancels borrowed by Andrew
No | 1 | Y | RFRA | Blank | ARRET ALEATOIRE |  | D2F8E27E084A9A45 | [5 empty] | Blank | ARRET ALEATOIRE
No | Visit | Complaint | RDAN | Jan klager over drift på Lac på +/-0,3-4 mmol på alt for mange kalibreringer. Han har prøvet andet membran lot og elektrode samt 0-punktsjustering uden resultat. | Afprøvet og undersøgt. Loggen viser alt for mange kal med drift på Lac. | Skiftet luftventilen og pumpeslanger samt kørt 0-punktsjustering, slutafprøvet. | SWO-220456 | Defect | HW | El/Met Module | Air Valve |  | Jan complains of operation on LAC at +/-0.3-4 mmol on too many calibrations. He has tried other membrane lot and electrode as well as 0-point adjustment without result. | Tested and investigated. The log shows too many Kal with drift on Lac. | Switched the air valve and pump hoses and driven 0-point adjustment, final tested.
No | 1 | Y | RFRA | Blank | Pb de mesure de l'oxy |  | D2F73DBB923DE52C | Defect | HW | Oxi module | PCB, Oxi module |  | Blank | Pb measuring oxy
No | 5 | Y | RAME | V Waste Line Popping Off. |  | Customer reported that the clear waste line was coming off at the clear Y junction during a rinse. Customer attempted to flush towards waste bottle but said there was a lot of pressure. Suspect a clot. | D2FEF98D561E1682 | Adjustment | HW | Autocheck module | Waste Tube | Pinched | V Waste Line Popping Off. |  | Customer reported that the clear waste line was coming off at the clear Y junction during a rinse. Customer attempted to flush towards waste bottle but said there was a lot of pressure. Suspect a clot.
No | Visit | Complaint | RIND | Touch Screen error. | Touch Screen error. | Found Touch screen problem, rebooted the instrument calibrated the display unit, found cal 1 solution was empty stock was nil. | SWO-221015 | Defect | HW | DMS_PC | Touch screen |  | Touch Screen error. | Touch Screen error. | Found Touch screen problem, rebooted the instrument calibrated the display unit, found cal 1 solution was empty stock was nil.
No | 5 | Y | RAME | Z Battery light is flickering |  | Battery light is flickering when Analyzer is turned off | D2F78E5592497EC5 | [5 empty] | Z Battery light is flickering |  | Battery light is flickering when Analyzer is turned off
No | 10 | Y | RPAC | Blank | Na Sensitivity error |  | D2FDD525E042BCAA | Assisted Replacement | SC | SC | Na |  | Blank | Na Sensitivity error
No | 14 | Y | RAME | L Customer stated that she is having problems with patient identification coming up and patient results coming over there LIS |  | Customer stated that sometimes it does come over and other times its not everything. Customer requested visit to verifty analyzer. | D2F8B31808560E88 | Setup | SW | Patient layout | Patient layout |  | L Customer stated that she is having problems with patient identification coming up and patient results coming over there LIS |  | Customer stated that sometimes it does come over and other times its not everything. Customer requested visit to verifty analyzer.
No | 8 | Y | RAME | Z - MISC |  | Greg advised that he has yellow parameters showing on the screen. verified calibrations passed, QCs passed, parameters still yellow | D2FBB1FBF41AFE1B | Setup | SW | ABL SW | QC setup | External QC | Z - MISC |  | Greg advised that he has yellow parameters showing on the screen. verified calibrations passed, QCs passed, parameters still yellow
No | 3 | Y | RAME | Z The printer is printing to lightly. |  | Checked the data logs and print out. | D2F895CC081EEE70 | [5 empty] | Z The printer is printing to lightly. |  | Checked the data logs and print out.
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 全参数不过，lac ph pco2 po2 na ca cl hct 定标不过 | 更换测试卡后修复 | SWO-221964 | [5 empty] | Test card, reagent package failure | Full parameter, however, lac ph pco2 po2 na ca cl hct calibration but | Repair after replacing test card
No | Visit | Complaint | RIND | continous failure of cassette. | Sensor board defective | Replaced sensor board, set zero performed successfully, SC installed successfully. | SWO-221146 | [5 empty] | continous failure of cassette. | Sensor board defective | Replaced sensor board, set zero performed successfully, SC installed successfully.
No | Visit | Complaint | RLTD | Rachel Leach 02087251667 | Error code 0500002A - Sensor Placement detection error during calibration. Error code 06100001 - SpO2 measurement Open LED connection. Suggested customer check sensor cable, check red LED stays on if cable gently flexed. If it goes off return sensor to RLTD for testing. Alternatively open and close calibration chamber door to see if error occurs. | Check sensor lead for faults, return to RLTD if faulty. | SWO-221113 | [5 empty] | Rachel Leach 02087251667 | Error code 0500002A - Sensor Placement detection error during calibration. Error code 06100001 - SpO2 measurement Open LED connection. Suggested customer check sensor cable, check red LED stays on if cable gently flexed. If it goes off return sensor to RLTD for testing. Alternatively open and close calibration chamber door to see if error occurs. | Check sensor lead for faults, return to RLTD if faulty.
No | Visit | Complaint | RLTD | Mark Spector Tel 07939 069137 | Checked OXI module in service programs. Complete OXI test has passed with no errors. Checked valve operation and it is showing both the EL/MET valves have a lot of errors. Solution pump tubing was not pushed down far enough to attach to the fluidic module nipple. | Removed and replaced EL $MET valves, but the each time the valves activated it is flagging an error. Restarted the Analyser, but the errors are still going up. The EL $MET PCB may need to be replaced. Reseated solution pump tubing. Ran QC and found QC material running out of the OXI module, showing the pinch valve s not closing properly. Even testing showed it as OK the pinch valve has been replaced and the Analyser is now passing QC. | SWO-221488 | No fault found | HW | Oxi module | Oxi module |  | Mark Spector Tel 07939 069137 | Checked OXI module in service programs. Complete OXI test has passed with no errors. Checked valve operation and it is showing both the EL/MET valves have a lot of errors. Solution pump tubing was not pushed down far enough to attach to the fluidic module nipple. | Removed and replaced EL $MET valves, but the each time the valves activated it is flagging an error. Restarted the Analyser, but the errors are still going up. The EL $MET PCB may need to be replaced. Reseated solution pump tubing. Ran QC and found QC material running out of the OXI module, showing the pinch valve s not closing properly. Even testing showed it as OK the pinch valve has been replaced and the Analyser is now passing QC.
No | 3 | Y | RPNZ | Blank | Abnormal high glucose results reported. |  | D2F8D2710850B139 | Cleaned | HW | Fluidic pathway | Fluidic pathway |  | Blank | Abnormal high glucose results reported.
No | 21 | Y | RAME | F When running VK kit L-3 gives result errors for all OXI values. |  | L-3 VK material generates OXI spectrum mismatch error. | D2F88A580807561C | Defect | HW | Oxi module | Hemolyzer |  | F When running VK kit L-3 gives result errors for all OXI values. |  | L-3 VK material generates OXI spectrum mismatch error.
No | Visit | Complaint | RDEU | 11.09.2017 12:50Uhr Frau Vent 02037330: Seit gestern meldet das Gerät bei allen Messungen inhomogene Probe. Die Werte sind plausibel. Keine weitere Fehlermeldung am Gerät. | Inhomogene Probe im BG Modul bzw. ungenügende Probe im BG Modul, Fehlersuche | Einlassnadel verbogen - Einlassnadel ersetzt, Messkammern BG Modul gereinigt -PH Elektrode untere Dichtung in Messkammer hängen geblieben - - Dichtung ersetzt, Einlassschlauch Inlet gedehnt - Schlauch ersetzt, Testmessungen und QC durchgeführt | SWO-221204 | Defect | HW | Inlet module | Probe bend |  | 11.09.2017 12:50pm Mrs. Vent 02037330: Since yesterday, the device has been reporting inhomogeneous samples for all measurements. The values are plausible. No more error messages on the device. | Inhomogeneous sample in the BG module or insufficient sample in the BG module, troubleshooting | Inlet needle bent - inlet needle replaced, measuring chambers BG module cleaned -PH electrode lower seal stuck in measuring chamber - - seal replaced, inlet hose stretched - hose replaced, test measurements and QC performed
No | 11 | Y | RAME | Z Communication Errors. |  | Confirmed the reported issue. | D2F8619D082B19A9 | [5 empty] | Z Communication Errors. |  | Confirmed the reported issue.
No | Visit | Complaint | RIND | iCa Values very low. not as per clinical condition and for every patient it showing low only. | iCa Values very low. not as per clinical condition and for every patient it showing low only. | Found analyser working well but reference ranges of iCa was as per Searum sample from 8 to 11 mg $dL. Ask Dr to confirm the reference as per there convenience and clinical purpose what they sugest. Dr ask to change as per global slandered. Reference changed as per Blood gas App. Sampling done and found well. Analyser Serviced and calibrated properly. Handed over in a fine working condition. | SWO-221202 | [5 empty] | iCa Values very low. not as per clinical condition and for every patient it showing low only. | iCa Values very low. not as per clinical condition and for every patient it showing low only. | Found analyser working well but reference ranges of iCa was as per Searum sample from 8 to 11 mg $dL. Ask Dr to confirm the reference as per there convenience and clinical purpose what they sugest. Dr ask to change as per global slandered. Reference changed as per Blood gas App. Sampling done and found well. Analyser Serviced and calibrated properly. Handed over in a fine working condition.
No | Visit | Complaint | RCHN | 测试卡卡口断裂 | 测试卡卡口断裂 | 更换测试卡正常。 | SWO-223263 | Defect | Consumables | Test Kit | Cartridge | Broken | Test ingress card mouth break | Test ingress card mouth break | The replacement test card is OK.
No | Visit | Complaint | RCHN | 电极定标反复错误 | 电极定标反复错误，借945-377一个 | 更换945-377后修复 by Andrew | SWO-220893 | [5 empty] | Repeated errors in electrode calibration | Repeated errors in electrode calibration, with a 945-377 | Repair by Andrew after replacing 945-377
No | Visit | Complaint | RCHN | 报2167错误 | 报2167错误，借636-539一个 | 更换636-539后修复 by Andrew | SWO-220970 | Cleaned | HW | Incubator Module | Sensor, Cup | 2167 | Reported 2167 errors | Reported 2167 error, borrow636-539 one | Repair by Andrew after replacement 636-539
No | Visit | Complaint | RCHN | 报1031错误 | 报1031错误，借903-364一个 | 更换903-364后修复 by Andrew | SWO-221150 | Defect | HW | Process Module | Optical Unit | 1031 | Reported 1031 error | Reported 1031 error, borrow 903-364 one | Repair by Andrew after replacing 903-364
No | 5 | Y | RFRA | Blank | PAS DE CONNECTION AQURE ERREUR CAL PCO2 |  | D2FC62916A00FF33 | Setup | SW | His/lis | AQURE |  | Blank | NO CONNECTION AQURE ERROR CAL PCO2
No | 10 | Y | RAME | C MetHb and COHb proficiency test out of limits. Unit also does not print. | Changed the complaint visit tab from N to Y. | Printer connections OK. Got copy of proficiency results to review. | D2FC8A906A491CFE | Defect | HW | Basic Unit | Printer module |  | C MetHb and COHb proficiency test out of limits. Unit also does not print. | Changed the complaint visit tab from N to Y. | Printer connections OK. Got copy of proficiency results to review.
No | 7 | Y | RPAC | Blank | Multiple parameter failures |  | D2FC71A26A3EDF7C | Assisted Replacement | SC | SC | Multi |  | Blank | Multiple parameter failures
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | LAC定标不过 | 更换测试卡后修复 | SWO-221958 | [5 empty] | Test card, reagent package failure | LAC calibration is not | Repair after replacing test card
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | PCO2定标失败，0210 | 更换测试卡 | SWO-224978 | [5 empty] | Test card, reagent package failure | PCO2 calibration failure, 0210 | Replace the test card
No | Visit | Complaint | RESP | Error 1063 Pressure test vacuum error | Chequeo de neumatica y sustitución de tubo de la válvula, limpieza y desobstrucción de circuito. | Chequeo de neumatica y sustitución de tubo de la válvula, limpieza y desobstrucción de circuito. | SWO-221521 | Defect | HW | Oxi module | Pinch valve tube | 1063 | Error 1063 Pressure test vacuum error | Tire check and valve tube replacement, circuit cleaning and deobstruction. | Tire check and valve tube replacement, circuit cleaning and deobstruction.
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 参比电极错误，PH，PCO2，LAC，NA，CA定标失败，1361，1315，1319，1316，1321 | 连续更换2张测试卡后修复 | SWO-221503 | Assisted Replacement | SC | SC | FT | 1310 | Test card, reagent package failure | Reference electrode error, PH, PCO2, LAC, NA, CA calibration failure, 1361, 1315, 1319, 1316, 1321 | Repair after 2 consecutive test cards
No | Visit | Complaint | RCHN | 测试杯传送故障 2168 | 测试杯传送故障 2168 ，借580-421一个。 | 更换580-421备件后修复完成 by jesson | SWO-221471 | Defect | HW | Storage Module | Stepper motor | 2168 | Test Cup Transfer Fault 2168 | Test Cup Transfer Fault 2168 , borrow 580-421 one. | Repair completed by jesson after replacement of 580-421 spare parts
No | Visit | Complaint | RCHN | 光学单元报错 | 光学单元报错 需要借件903-364备件一个 |  | SWO-221318 | Defect | HW | Process Module | Optical Unit |  | Optical unit error | Optical unit error needs to borrow 903-364 spare parts one
No | 6 | Y | RAME | 2 PCO2 is failing calibration intermittently and failing level 1 and 2 QC. Also was getting intermittent cal failures with PH. |  | Looked in data logs and found PCO2 failing a few times but even with a good PCO2 calibration level 1 and level 2 was below package insert range. Both PH electrode and PCO2 electrode are under warranty. 945-614XK-01, 945-612YM-01. | D2FC5DB16A06EB9E | Assisted Replacement | Electrode | Electrode | pCO2 |  | 2 PCO2 is failing calibration intermittently and failing level 1 and 2 QC. Also was getting intermittent cal failures with PH. |  | Looked in data logs and found PCO2 failing a few times but even with a good PCO2 calibration level 1 and level 2 was below package insert range. Both PH electrode and PCO2 electrode are under warranty. 945-614XK-01, 945-612YM-01.
No | 7 | Y | RAME | V EL/MET module error not receiving rinse from above. |  | Rinse and air valves are operating slightly slower causing small segments of rinse solution causing error. | D2FC5AE06A1C00E0 | Defect | HW | El/Met Module | Valve |  | V EL/MET module error not receiving rinse from above. |  | Rinse and air valves are operating slightly slower causing small segments of rinse solution causing error.
No | 22 | Y | RPNZ | Blank | Membrane replacement due to Crea QC failure. |  | D2FC46596A30B914 | Assisted Replacement | Electrode | Membrane | Crea |  | Blank | Membrane replacement due to Crea QC failure.
No | Visit | Complaint | RSCH | Rep Garantie | pCO2 Empfindlichkeit ohne Einwirkung unter 85% Kunde hat mehrmals membraniert ohne Erfolg | pCO2 Elektrode ausgetauscht. Empf mit neuer Elektrode 36.9% 1. neue Elektrode ebenfalls defekt pH Mainbord ausgetauscht. Kal knapp i.o 2. pCo2 Elektrode YN03 Kal i.o pH Mainbord wieder ausgetauscht QC i $o | SWO-221593 | Defect | HW | pH/BG Module | PCB, ph/BG |  | Rep Guarantee | pCO2 Sensitivity without exposure below 85% Customer has membraned several times without success | pCO2 electrode. Empf with new electrode 36.9% 1. new electrode also defective pH Mainbord. Kal knapp i.o 2. pCo2 electrode YN03 Kal i.o pH Mainbord replaced again QC i $o
No | Visit | Complaint | RDAN | Skærmen er ude af kalibrering. | NA | Jeg har kalibreret skærmen. Alt er OK igen. | SWO-223573 | Adjustment | HW | DMS_PC | Touch screen |  | The screen is out of calibration. | In | I have calibrated the screen. Everything is OK again.
No | 6 | Y | RAME | I Customer called over the weekend with PCO2 errors and Temp errors. |  | Found PCO2 errors on QC and temp errors (355 - Temp error PH $BG module) (678 - Heater error). FOund crystals in PH measuring chamber. | D2FC62C26A5C5C41 | Defect | HW | pH/BG Module | PCB, Thermostat |  | I Customer called over the weekend with PCO2 errors and Temp errors. |  | Found PCO2 errors on QC and temp errors (355 - Temp error PH $BG module) (678 - Heater error). FOund crystals in PH measuring chamber.
No | 7 | Y | RFRA | Blank | ERR 1794 PIPETTE 38 |  | D2FC44AC6A56F170 | Defect | HW | Aspiration Module | Bend needle |  | Blank | ERR 1794 PIPETTE 38
No | Visit | Complaint | RDEU | Kunde beachtete schon länger, dass thb QC Werte sich in allen 3 Leveln an der oberen Grenzenbefinden | tHb QC Werte in Level B & C permanent N oberer Kontrollgrenze mit verschiedenen LP Lot thb Kal brachte keine Verbesserung | Austausch der Hemolyzer Unit ( 902-980R0747N003 verbaut), Befestigungen am Unterteil des Oxymoduls zerbrochen -> ersetzt, kpl. Verschlauchung erneuert, thb Kal durchgeführt, QC Messung: thb Werte 0,1 mmol $l über ZW, Kunde beobachtet Geräteverhalten und meldet sich ggf. wieder | SWO-221894 | Defect | HW | Oxi module | Hemolyzer |  | Customer has long noticed that thb QC values are at the upper limit in all 3 levels | tHb QC values in level B & C permanent N upper control limit with different LP lot thb Kal brought no improvement | Replacement of the Hemolyzer Unit ( 902-980R0747N003 installed), fastenings to the lower part of the oxymodule broken -> replaced, cpl. Hose ingestion renewed, thb Kal performed, QC measurement: thb values 0.1 mmol $l via ZW, customer observes device behavior and reports again if necessary
No | Visit | Complaint | RLTD | Annette on 0300 3001761 before 10:30 or after 11:30 | Checked data logs. | Replaced faulty battery (P0063N022) with new battery (P0138N087). Analyser is ready for use. | SWO-222068 | Defect | HW | Battery module | Battery |  | Annette on 0300 3001761 before 10:30 or after 11:30 | Checked data logs. | Replaced faulty battery (P0063N022) with new battery (P0138N087). Analyser is ready for use.
No | Visit | Complaint | RDEU | 15.08.2017 12:50Uhr Sr.Andrea 022146771750: Wegen einem Temperaturfehler wurde das Gerät neu gebootet. Nach dem Booten meldet das Gerät | Fehlersuche per Service Menü | Manuelle Spülung der Blutwege, Hemolyzer frei bekommen, Sensor ersetzt da verstopft, Neustart Kalibration und QC Kontrolle Kein technischer defekt sonder Verstopfung | SWO-221947 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway | Clot | 15.08.2017 12:50h Sr.Andrea 022146771750: Due to a temperature error the device was rebooted. After booting, the device reports | Troubleshooting via service menu | Manual flushing of bloodpathies, get hemolyzerfree, sensor replaces there clogged, restart calibration and QC control No technical defect but constipation
No | 4 | Y | RAME | Z Analyzer will not boot up has blank screen. |  | Customer tried to reboot analyzer several times but will not boot up. Check analyzer and ordered a power supply. | D2FD8304E01E78DD | Defect | HW | Basic Unit | Power Supply |  | Z Analyzer will not boot up has blank screen. |  | Customer tried to reboot analyzer several times but will not boot up. Check analyzer and ordered a power supply.
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | LAC定标失败 | 更换测试卡后修复 | SWO-221847 | [5 empty] | Test card, reagent package failure | LAC calibration failure | Repair after replacing test card
No | Visit | Complaint | RCHN | GLU，Lac定标不过 | GLU，Lac定标不过，借903-183一个（被903-581替代） | 信号板探针调整位置后修复 by jesson | SWO-221400 | Adjustment | HW | Basic Unit | PCB, Sensor interface |  | GLU, Lac calibration is not | GLU, Lac calibration, borrow edgy 903-183 (replaced by 903-581) | Repair by jesson after signal board probe adjusts position
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | PCO2,CL,GLU,LAC质控失控 | 更换测试卡正常。 | SWO-221953 | [5 empty] | Test card, reagent package failure | PCO2, CL, GLU, LAC control out of control | The replacement test card is OK.
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | O2质控不过，0090 | 更换测试卡后修复 | SWO-221778 | Assisted Replacement | SP | SP | pO2 |  | Test card, reagent package failure | O2 quality control, 0090 | Repair after replacing test card
No | Visit | Complaint | RCHN | 报5684 | 报5684，5670错误 | 自行更换580-098、636-541和重新定标后修复 by jesson | SWO-221817 | Defect | HW | Process Module | Stepper motor | 5684 | Reported 5684 | 5684,5670 errors | Repair by jesson after self-replacement 580-098, 636-541 and re-calibration
Yes | 7 | Y | RAME | W Remove demo keys. |  | Analyzer was working properly upon arrival. There was a message stating that the demo keys would be expiring shortly. | D30043A1CC022555 | Re classified RMED | Demo key will expire |  |  |  | W Remove demo keys. |  | Analyzer was working properly upon arrival. There was a message stating that the demo keys would be expiring shortly.
No | Visit | Complaint | RDEU | 15.08.2017 11:40Uhr Pickler 03080505352: Die Cal-Cartrigde wurde eingesetzt, aber das Gerät blockiert und die Cartrigde lässt sich nicht mehr entfernen. | - Einlaßfenster ausgehakt - Cal-Cartridge klemmt | Fehlersuche; Gerät aus; Einlaßfenster korrekt positioniert; Federmechanismus justiert; Cal-Cartridge aus Storage Wheel entfernt = ein Nippel zur Führung d. Kassette fehlt; Anwender hat keine neue Kal.-Kassette vor Ort; Anwender bestellt neu; Gerät ein; Thermostatisierung abgewartet; QC i.O. | SWO-221885 | Defect | Consumables | Test Kit | Cartridge | Broken | 08/15/2017 11:40 Am Pickler 03080505352: The Cal-Cartrigde was used, but the device was blocked and the cartrigde can no longer be removed. | - Inlet window checked out - Cal cartridge clamps | Troubleshooting; device off; Inlet window correctly positioned; spring mechanism adjusted; Cal cartridge removed from Storage Wheel = a nipple for the guide of the cassette is missing; User does not have a new Cal. cassette on site; User orders new; device; thermostatization awaited; QC in the first
Yes | 3 | Y | RAME | E batteries overheated while in storage | Changed the complaint visit tab from N to Y. | Needs new PS because the battery connector is broken | D2FE1D06E009A209 | [5 empty] | E batteries overheated while in storage | Changed the complaint visit tab from N to Y. | Needs new PS because the battery connector is broken
Yes | 3 | Y | RAME | E batteries overheated while in storage |  | Needs new PS because the battery connector is broken | D2FE1811E003A205 | [5 empty] | E batteries overheated while in storage |  | Needs new PS because the battery connector is broken
Yes | 3 | Y | RAME | E Battery connector broken |  | Battery connector broken will need a PS | D2FF575C562820FE | [5 empty] | E Battery connector broken |  | Battery connector broken will need a PS
Yes | 3 | Y | RAME | E Battery connector broken due to heat |  | Battery connector broken will need new PS | D2FF5722563620FC | [5 empty] | E Battery connector broken due to heat |  | Battery connector broken will need new PS
Yes | 3 | Y | RAME | E Battery connector broken Due to heat |  | Battery connector broken will need PS | D2FF56E0563020FB | [5 empty] | E Battery connector broken Due to heat |  | Battery connector broken will need PS
Yes | 3 | Y | RAME | E batteries overheated while in storage | Changed the complaint visit tab from N to Y. | Needs new PS because the battery connector is broken | D2FE208AE05AA20D | [5 empty] | E batteries overheated while in storage | Changed the complaint visit tab from N to Y. | Needs new PS because the battery connector is broken
No | Visit | Complaint | RBNL | Ca, Gluc, Lact erreur | Les contrôles externe n'étaient pas correct pour le Ca++ Depuis 4 jours les calibrations pour Gluc et Lact donnaient des valeurs d'état trop élevées ! Remplacement de la cassette des capteurs n'a pas résoud le problème. | Remplacement de la carte électronique (Sensor Interface) Calibration et contrôle : NOK, les valeurs d'état et sensibilité pour Glu et Lac sont assez mauvais. Remplacement de la carte électronique originale, le problème vient donc de la nouvelle cassette de capteurs , elle devra donc remplacé prochainement. Contrôle externe Radiometer (autocheck 5+) : OK pour le Ca (voir note). | SWO-222301 | Assisted Replacement | SC | SC | Multi | LEH Mail | Ca, Gluc, Lact erreur | The external controls were not correct for the Ca For 4 days the calibrations for Gluc and Lact gave too high state values! Replacing the sensor cassette did not solve the problem. | Replacing the e-card (Sensor Interface) Calibration and control: NOK, the state and sensitivity values for Glu and Lac are pretty bad. Replacing the original electronic card, the problem comes from the new cassette of sensors, so it will have to replace soon. External control Radiometer (autocheck 5): OK for ca (see note).
No | Visit | Complaint | RBNL | pO2 instable | Depuis quelque temps la valeur de la pO2 n'est plus affichée pour les résultats sanguin. Le problème était 2 fois lié à une erreur de calibration causé par une cassette senseur (solution non homogène pendant calibration), après remplacement de la cassette l'erreur était solutionnée. Les qraphiques des cq étaient tout le temps très stable ! | Remplacement préventif d'une carte électronique (Sensor Interface). | SWO-222300 | Defect | HW | Basic Unit | PCB, Sensor interface |  | unstable pO2 | For some time the value of pO2 is no longer displayed for blood results. The problem was twice related to a calibration error caused by a sensor cassette (non-homogeneous solution during calibration), after replacement of the cassette the error was resolved. The qraphics of the cq were always very stable! | Preventive replacement of an electronic card (Sensor Interface).
No | Visit | Complaint | RDEU | QC-Fehler sO2 | QC-Fehler sO2 | Fehlersuche; Gerät aus; Hemolyzer gew.; Gerät ein; div. Einstellungen vorgenommen; Startphase abgewartet; Kal. = Glu-Fehler; SC gew.; weitere Glu-Fehler; Gerät neu gestartet; Kal. + QC i.O. | SWO-222226 | Defect | HW | Oxi module | Hemolyzer |  | QC-Fehler sO2 | QC-Fehler sO2 | Troubleshooting; device off; Hemolyzer. device; div. Settings made; start-up phase awaited; Kal. = Glu error; SC. further Glu errors; Device restarted; Kal. + QC in the first place
No | Visit | Complaint | RDEU | 15.08.2017, 09:42 Uhr, Fr. Dill, Tel. 07131492115: tBil Werte bei Pat. Messungen falsch | Analysatordaten und Konfiguration überprüft. | Konfiguration fehlerhaft , da HbF Korrektion deaktiviert war. Korrektion auf alle Anteile eingestellt. Vorsorglich Oximetrie - Fließwege über Serviceprogramm mit Proteinentferner und Wasser gespült. Oxitubing erneuert. Einlaufverfahren, tHb Kal. und QC durchgeführt. | SWO-222431 | Setup | SW | Oxi module | HbF corr. |  | 15.08.2017, 09:42, Ms. Dill, tel. 07131492115: tBil values at Pat. Measurements incorrect | Analyzer data and configuration checked. | configuration is incorrect because HbF correction was disabled. correction to all shares. Precautionary oximetry - flow paths rinsed via service program with protein remover and water. Oxitubing renewed. inlet method, tHb Kal. and QC.
No | Visit | Complaint | RDEU | GLU/Lac Fehler trotz Wechsel der SC. Nach Rüchsprache mit Rmed soll die Sensor Interface PCB getauscht werden. | Sensor Interface PCB defekt. | Interface PCB getauscht. | SWO-222181 | Defect | HW | Basic Unit | PCB, Sensor interface |  | GLU/Lac error despite changing the SC. The sensor interface PCB is to be replaced after the speech with Rmed. | Sensor Interface PCB defekt. | Interface PCB getauscht.
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | O2,LAC定标不过 | 更换测试卡后修复 | SWO-223296 | Assisted Replacement | SC | SC | Lac |  | Test card, reagent package failure | O2, LAC calibration is not | Repair after replacing test card
No | Visit | Complaint | RCHN | 报错1035 | 报错1035，借备件903-364一个 | 更换903-364一个正常。by jesson | SWO-221405 | Defect | HW | Process Module | Optical Unit | 1035 | 1035 error | Error 1035, borrowed spare parts 903-364 one | Replace 903-364 one normal. by jesson
No | Visit | Complaint | RCHN | 报错1031 | 报错1031，需要借件903-364一个 | 更换903-364后修复 by jesson | SWO-221664 | Defect | HW | Process Module | Optical Unit | 1031 | 1031 | 1031, need to borrow 903-364 one | Repair by jesson after replacing 903-364
No | Visit | Complaint | RCHN | 定标错误 | 定标终止 ，借944-258（12979）一个 | 清除内部灰尘。重新定标后修复 by jesson | SWO-221723 | Cleaned | HW | Incubator Module | Sensor | 2877 | Calibration error | Bid termination, borrowing 944-258 (12979) one | Remove internal dust. Fix by jesson after re-calibration
No | 25 | Y | RCZE | Blank |  |  | D2FEAE2056444DFF | Defect | Electrode | Electrode | Sealing ring | Sealing ring, pH | Blank
No | Visit | Complaint | RIND | ABL800 is not working | Checked and found fuse on analyzer control PCB were corroded | Cleaned fuse contacts and remounted fuse. Restarted ABL800 successfully. 2 point calibration performed successfully. | SWO-222315 | Cleaned | HW | Power supply | Fuse contacts | corrosion | ABL800 is not working | Checked and found fuse on analyzer control PCB were corroded | Cleaned fuse contacts and remounted fuse. Restarted ABL800 successfully. 2 point calibration performed successfully.
No | Visit | Complaint | RDAN | Jan Pedersen melder at R's ABL driver alt for tit på GLU. Han har prøvet med nyt lot membraner og anden elektrode uden resultat. | Loggen giver Jan ret. | Luftventil og pumpeslange udskiftet, slutafprøvet. | SWO-222090 | Defect | HW | El/Met Module | Valve, air |  | Jan Pedersen reports that R's ABL runs too often on GLU. He has tried with new lot membranes and other electrode without result. | The log giver Jan ret. | Air valve and pump hose replaced, final tested.
Yes | 4 | Y | RCZE | Blank |  |  | D2FEB591564B4E01 | Re classified RMED | UPS failure |  |  |  | Blank
No | Visit | Complaint | RDEU | Sehr häufig Glukose Kalibrationsfehler Fehler:0379 Ansprachefehler, QC im Bereich | Überprüfung (Garantie) Elektrode wurde bereits mehrfach membraniert | QC Verlauf in Ordnung. Kalibrationsfehler Glu, Ansprachefehler. Glu Elektrode ausgetauscht (alt:YK04, neu:XT03) Kalibrationen mehrfach durchgeführt. Weiterhin Ansprachefehler. Rücktausch der Elektrode. Austausch des Elektrodendeckels. Kalibrationen mehrfach durchgeführt. Kal und QC ohne Beanstandung | SWO-222173 | Defect | HW | El/Met Module | Electrode connector |  | Very common glucose calibration errors Error:0379 Address error, QC in the area | Review (guarantee) Electrode has already been membraneed several times | QC course fine. Calibration error Glu, speech error. Glu electrode replaced (old:YK04, new:XT03) calibrations performed several times. Continued speech errors. Replacement of the electrode. Replacement of the electrode cover. calibrations performed several times. Kal and QC without complaint
No | Visit | Complaint | RDEU | Kalium und Glucose haben hohe driften seit der Auslieferung | Kalibrierfehler Status und drift Fehler bei Kalium, Glucose in Ordnung | Messkammern überprüft, Ref Elektrode nicht bis anschlag durchgedrückt. Elektrodenflüssigkeit ausgelaufen. Kammern gesäubertund neu membraniert. Pumpenkalibrierung PH $Blutgasmodul zu niedrig, pumpe Schlauch gewechselt. LS Justage in Ordnung | SWO-222588 | User error | Electrode | Membrane | Reference |  | Potassium and glucose have drifted high since delivery | Calibration error status and drift error in potassium, glucose fine | Measuring chambers checked, ref electrode not pressed through to stop. electrode fluid leaked. chambers cleaned and re-membraneed. Pump calibration PH $Blutgasmodul too low, pump hose changed. LS Adjustment in order
No | Visit | Complaint | RKK | 圧力テスト吸引エラー | 圧力テスト吸引エラー | ホルダ付インレットガスケットの破損確認→交換実施。 顧客要望にてSC交換。 キャル、QC確認・・・良好 | SWO-222861 | Defect | HW | Inlet module | Inlet gasket | Request Data (LEH) | Pressure test suction error | Pressure test suction error | Check for damage to inlet gasket with holder → replacement. SC exchange at customer request. Cal, QC confirm... good.
No | 23 | Y | RAME | 9 reported variations on blood samples on LAC while running sample comparison. |  | Reviewed logs and found some sensor casette cal errors during the comparison run. | D3014ED4422F210A | Assisted Replacement | SC | SC | Lac |  | 9 reported variations on blood samples on LAC while running sample comparison. |  | Reviewed logs and found some sensor casette cal errors during the comparison run.
Yes | Visit | Complaint | RLTD | Training & live connection connections X 3 | Training & live connection connections X 3 | Provided AQURE express user training to lab staff, ABL90 troubleshooting training and set 3 X ABL90s for connection to live connect. | SWO-222813 | Re classified RMED | Training |  |  |  | Training & live connection connections X 3 | Training & live connection connections X 3 | Provided AQURE express user training to lab staff, ABL90 troubleshooting training and set 3 X ABL90s for connection to live connect.
No | Visit | Complaint | RCHN | 模块有故障 | 模块有故障，借902-478一个 | 配件需重新借件902-778一个 by jesson | SWO-222515 | [5 empty] | There is a problem with the module | The module is faulty, with a 902-478 | Accessories need to be re-borrowed 902-778 a by jesson
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 响应错误 1310 | 更换测试卡后修复 | SWO-225616 | Assisted Replacement | SC | SC | Glu/Lac | 1310 | Test card, reagent package failure | Response Error 1310 | Repair after replacing test card
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | PO2质控失败 | 更换测试卡正常。 | SWO-222576 | Assisted Replacement | SC | SC | pO2 |  | Test card, reagent package failure | PO2 QC FAILURE | The replacement test card is OK.
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | O2定标不过 | 更换测试卡后修复 | SWO-223297 | Assisted Replacement | SC | SC | pO2 |  | Test card, reagent package failure | O2 calibration but | Repair after replacing test card
Yes | 1 | Y | RPAC | Blank | TCm5 Gas used up in x2 days due to Sensor 84 membrane issue. |  | D3002C90CC52E6E0 | [5 empty] | Blank | TCm5 Gas used up in x2 days due to Sensor 84 membrane issue.
No | Visit | Complaint | RCHN | 报错5297 | 报错5297，借备件580-098一个 | 更换580-098后修复 by jesson | SWO-221813 | Defect | HW | Incubator Module | Stepper motor | 5297 | 5297 error | Error 5297, borrowed spare parts 580-098 one | Repair by jesson after replacement 580-098
No | Visit | Complaint | RLTD | Mark Spector Tel 07939 069137 | Checked QC data logs and found a lot of samples being detected prematurely during aspiration. Viewed service programs and tested valves. EL/MET PCB needs to replaced due to the valves showing errors even though they are functioning correctly. | Replaced EL $MET PCB and programmed. Replaced both valves and also found the c clip on the air valve is loose. This is the suspected root cause of the QC error. There are no failures showing for the valves. Analyser has passed calibrations and QC's. | SWO-222836 | Defect | HW | El/Met Module | Valve, air |  | Mark Spector Tel 07939 069137 | Checked QC data logs and found a lot of samples being detected prematurely during aspiration. Viewed service programs and tested valves. EL/MET PCB needs to replaced due to the valves showing errors even though they are functioning correctly. | Replaced EL $MET PCB and programmed. Replaced both valves and also found the c clip on the air valve is loose. This is the suspected root cause of the QC error. There are no failures showing for the valves. Analyser has passed calibrations and QC's.
No | 9 | Y | RAME | P LEAK ON RIGHT SIDE OF INSTRUMENT |  | FOUND ORING ON LACTATE ELECTRODE MISSING. | D3005315CC096E56 | Defect | Electrode | Electrode | Sealing ring | Sealing ring, Lac | P LEAK ON RIGHT SIDE OF INSTRUMENT |  | FOUND ORING ON LACTATE ELECTRODE MISSING.
No | 9 | Y | RAME | 2 QC 1 & 2 RECOVERING BELOW ACCEPTABLE RANGES FOR PCO2. DRIFT ON Ca+. |  | QC 1 & 2 FOR PCO2 HAVE BEEN RECOVERING TOWARDS MEAN VALUES COMPARED TO JUL7 WHEN FIRST REPORTED. L1 NOW READING ABOUT 67+ $- 0.5 mmHg (62.1-72.1) WHILE L2 IS RECOVERING ABOUT 39.5 + $- 0.5 mmHg (37.2-43.2). DID FIND BUBBLE TRAPPED BETWEEN PCO2 ELECTRODE & MEMBRANE AND BROUGHT TO GIA'S ATTENTION FOR BETTER TECHNIQUE WHEN MEMBRANING BG. REPLACED Ref & Ca ELECTRODES FOR TROUBLESHOOTING. IF Ca CALIBRATION CONTINUES TO STABILIZE, THEN GIA WILL INFORM RYAN TO REPLACE BOTH ELECTRODES SINCE THEY ARE NEW AND UNDER WARRANTY. | D300337FCC1098EC | Assisted Replacement | Electrode | Electrode | pCO2 |  | 2 QC 1 & 2 RECOVERING BELOW ACCEPTABLE RANGES FOR PCO2. DRIFT ON Ca+. |  | QC 1 & 2 FOR PCO2 HAVE BEEN RECOVERING TOWARDS MEAN VALUES COMPARED TO JUL7 WHEN FIRST REPORTED. L1 NOW READING ABOUT 67+ $- 0.5 mmHg (62.1-72.1) WHILE L2 IS RECOVERING ABOUT 39.5 + $- 0.5 mmHg (37.2-43.2). DID FIND BUBBLE TRAPPED BETWEEN PCO2 ELECTRODE & MEMBRANE AND BROUGHT TO GIA'S ATTENTION FOR BETTER TECHNIQUE WHEN MEMBRANING BG. REPLACED Ref & Ca ELECTRODES FOR TROUBLESHOOTING. IF Ca CALIBRATION CONTINUES TO STABILIZE, THEN GIA WILL INFORM RYAN TO REPLACE BOTH ELECTRODES SINCE THEY ARE NEW AND UNDER WARRANTY.
No | Visit | Complaint | RDEU | 27.07.2017; Sr. Anett; Tel.: 03381791614; CRP QC Messfehler 28.07.2017 10:30 Sr. Nicole; Tel.: 03381791612; weiterhin CRP QC Messfehler 31.07.2017: Rücksprache mit KerZi - neues QC wird bestellt, telefonische Info an Kunden 03.08.2017, 11:36 Uhr ==> Anrufer Pfl. Sebastian, Tel.: 03381791611 ==> LQC-Messfehler CRP 04.08.2017, 12:04 Uhr ==> Anrufer Schw. Anett, Tel.: 03381791614 ==> LQC-Fehler CRP besteht trotz neuer Charge weiterhin | CRP-Messfehler bei QK | Datensicherung, Optische Einheit ersetzt, Prozeßrad gesäubert, Startphase abgewartet, FactoryReferenz kalibriert, CRP(Lot13125) kalibriert, QK weiterhin zu hoch in beiden Leveln, wg. defektem Neuteil (optische Einheit) Gerät nach Rücksprache mit lt. OÄ zu Werkstattreparatur abgezogen ! | SWO-222697 | Defect | HW | Process Module | Optical Unit | LQC | 27.07.2017; Sr. Anett; Tel.: 03381791614; CRP QC Measurement Error 07/28/2017 10:30 Sr. Nicole; Tel.: 03381791612; further CRP QC measurement error 31.07.2017: Consultation with KerZi - new QC is ordered, telephone information to customers 03.08.2017, 11:36 a.m. ==> Caller Pfl. Sebastian, Tel.: 0338179111 ==> LQC measurement error CRP 04.08.2017, 12:04 o'clock ==> Caller Schw. Anett, tel.: 03381791614 ==> LQC error CRP persists despite new batch | CRP-Messfehler bei QK | Data backup, optical unit replaced, process wheel cleaned, start phase awaited, FactoryReference calibrated, CRP(Lot13125) calibrated, QK still too high in both levels, wg. defective new part (optical unit) device deducted after consultation with lt. OÄ for workshop repair !
No | Visit | Complaint | RLTD | Failing Glu calibration across three sensor cassettes. | Failing Glu calibration across two sensor cassettes. | Inspected and cleaned contact pins. No issues found with contacts. | SWO-223070 | Assisted Replacement | SC | SC | Glu |  | Failing Glu calibration across three sensor cassettes. | Failing Glu calibration across two sensor cassettes. | Inspected and cleaned contact pins. No issues found with contacts.
No | Visit | Complaint | RDEU | 16.08.2017, Heb. Marie, Tel.: 06021324150: Einlass schließt nicht mehr; Sie ist zum Frühdienst gekommen und der Einlass stand offen, mehr Infos hat Sie nicht | Bei Ankunft: Intervention Erforderlich >Einlassdichtungshalter tauschen. | Intervention durchgeführt, Startphase 50min. Kal und QCs i.O. | SWO-223321 | Adjustment | HW | Auto Inlet module | Inlet Position Cal. |  | 16.08.2017, Heb. Marie, Tel.: 06021324150: Entrance no longer closes; She came to the early service and the entrance was open, you don't have more information | On arrival: Intervention Required >Exchange inlet seal holder. | Intervention carried out, start phase 50min. Kal and QCs i.O.
No | Visit | Complaint | RLTD | Lactate is intermittently failing calibrations. | Checked data logs and service programs. | Cleaned sensor interface PCB contact pins. Flushed Analyser, cleaned gasket and probe. The Analyser is still failing on Lactate. Swapped sensor cassette (LOT 47-21) with new ABL90 in SAU (LOT 50-24). All calibrations and QC's are now passing. SAU Analyser has low sensitivity on Lactate, but we are able to leave it over the weekend to see if it recovers. New sensor interface PCB will be fitted 21 $08 $17. | SWO-223317 | Assisted Replacement | SC | SC | Lac |  | Lactate is intermittently failing calibrations. | Checked data logs and service programs. | Cleaned sensor interface PCB contact pins. Flushed Analyser, cleaned gasket and probe. The Analyser is still failing on Lactate. Swapped sensor cassette (LOT 47-21) with new ABL90 in SAU (LOT 50-24). All calibrations and QC's are now passing. SAU Analyser has low sensitivity on Lactate, but we are able to leave it over the weekend to see if it recovers. New sensor interface PCB will be fitted 21 $08 $17.
No | 3 | Y | RPOL | Blank | Wymiana kasetki sensorowej nt LOT R0098N028. Instalacja nowej kasety sensorowej. Kalibracja oraz kontrola jakoœci prawid³owa. Analizator sprawny. |  | D3016B4F4212F460 | Assisted Replacement | SC | SC | pH & Elek |  | Blank | Replacing the sensor cassette NT LOT R0098N028. Installing a new sensor cassette. Calibration and quality control. Functional Analyzer.
No | 6 | Y | RAME | H Error message 1218, possible air in sample. | Changed the health hazard tab from N to Y. | Checked the data logs. | D301783E4262DEFD | [5 empty] | H Error message 1218, possible air in sample. | Changed the health hazard tab from N to Y. | Checked the data logs.
No | 5 | Y | RAME | Z ANALYZER FAILED STARTUP / INTERNAL SYSTEM CHECK |  | FOUND PARAMETERS ALL LOCKED OUT. AFTER RE-ENABLING, SYSTEM CHECK FAILED AGAIN. | D30197E04242A65B | [5 empty] | Z ANALYZER FAILED STARTUP / INTERNAL SYSTEM CHECK |  | FOUND PARAMETERS ALL LOCKED OUT. AFTER RE-ENABLING, SYSTEM CHECK FAILED AGAIN.
No | Visit | Complaint | RDEU | Lösungsaustritt am Abfallstutzen der Sensorkassette | Abfallschlauch abgeknickt | Gerät gesäubert, Lösungspackklappe gewechselt, Schlauchknick beseitigt, Funktionskontrolle | SWO-222961 | [5 empty] | Solution outlet at the waste nozzle of the sensor cassette | Waste hose bent | Device cleaned, solution packing flap changed, hose buckle removed, function control
No | Visit | Complaint | RCHN | 模块有故障 | 模块有故障，借902-478一个,现需要加借902-778一个 |  | SWO-222920 | [5 empty] | There is a problem with the module | The module has a fault, borrow 902-478 one, now need to borrow 902-778 one
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | pO2质控高于范围 | 更换试剂包后修复 | SWO-223257 | Assisted Replacement | SP | SP | pO2 |  | Test card, reagent package failure | pO2 QC is higher than range | Repair after replacement of reagent kit
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | pO2质控超范围 | 更换试剂包后修复 | SWO-223292 | Assisted Replacement | SP | SP | pO2 |  | Test card, reagent package failure | pO2 QC over range | Repair after replacement of reagent kit
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 不均匀的冲洗溶液 1318,1319,1324 | 更换试剂包后修复 | SWO-222980 | Assisted Replacement | SP | SP | FT | 1310 | Test card, reagent package failure | Uneven flushing solution 1318,1319,1324 | Repair after replacement of reagent kit
No | Visit | Complaint | RESP | Fuga detectada | Chequeo neumático. Se observa aguja doblada | Sustitución aguja de entrada. Limpieza circuito neumático. Tests de neumática. Arranque . Tras más de un día parado se presentan problemas en pH (fallo de impedancia), Lac y Glu. Sustitución SC. inició ,cal y QC | SWO-223166 | Defect | HW | Auto Inlet module | Probe bend | 1184 | Leak detected | Pneumatic check. Bent needle is observed | Input needle replacement. Pneumatic circuit cleaning. Pneumatic tests. Start. After more than a day of standing there are problems in pH (impedance failure), Lac and Glu. SC substitution. started ,cal and QC
No | Visit | Complaint | RCHN | 测试卡无法扫描 | 测试卡无法扫描，借备件910-304一个 | 使用备用机的备件910-304后修复 by jesson | SWO-223173 | Defect | HW | Analyzer Control Module | Barcode reader, display | BR001 | Test card can't be scanned | Test card can not be scanned, borrowed spare parts 910-304 one | Repair by jesson after using spare parts 910-304 of the spare machine
No | Visit | Complaint | RCHN | 机器频繁出现黑屏 | 机器频繁出现黑屏，借备件910-343,910-329各一个 | 更换910-343和910-329备件后修复 by jesson | SWO-221968 | Defect | HW | Analyzer Control Module | DMS_PC |  | Frequent black screens appear on the machine | Machines frequently appear black screen, borrowed spare parts 910-343, 910-329 each one | Repair by jesson after replacing 910-343 and 910-329 spare parts
No | Visit | Complaint | RESP | Error de lavado | Chequeo de circuitos neumáticos y válvulas | Sustitución de válvulas modulo ph $bg y limpieza de el $met Sustitucion tubos de bomba | SWO-223042 | Defect | HW | pH/BG Module | Valve |  | Washing error | Checking pneumatic circuits and valves | Replacing valves module ph $bg and cleaning the $met Replacement pump tubes
No | 23 | Y | RPNZ | Blank | pCO2 QC low bias again |  | D3017631423767DF | Defect | HW | Inlet module | Inlet nipple kit |  | Blank | pCO2 QC low bias again
No | Visit | Complaint | RDEU | Kalium und Glucose haben hohe driften seit der Auslieferung |  |  | SWO-223021 | Assisted Replacement | Electrode | Electrode | K+ | LEH Mail | Potassium and glucose have drifted high since delivery
No | Visit | Complaint | RCHN | 检测到泄露，压力测试真空错误 | 检测到泄露，压力测试真空错误 | 冲洗分析仪后修复 by jesson | SWO-223182 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway | 1342 | Leaks detected, pressure test vacuum errors detected | Leaks detected, pressure test vacuum errors detected | Fix by jesson after flushing analyzer
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | pO2质控超范围 0589 | 更换测试卡后修复 | SWO-223293 | Assisted Replacement | SC | SC | pO2 | 589 | Test card, reagent package failure | pO2 QC over range 0589 | Repair after replacing test card
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | CO2定标失败，压力测试流量错误 | 更换测试卡后修复 | SWO-223200 | Assisted Replacement | SC | SC | pCO2 |  | Test card, reagent package failure | CO2 calibration failure, stress test flow error | Repair after replacing test card
No | Visit | Complaint | RCHN | 血氧模块硬件错误 | 血氧模块硬件错误，借备件902-980一个 |  | SWO-223181 | Defect | HW | Oxi module | Hemolyzer | 1034 | Blood oxygen module hardware error | Blood oxygen module hardware error, borrowed spare parts 902-980 one
No | Visit | Complaint | RCHN | 定标终止 | 定标终止，需要借件944-214 批次号：13002 | 清洁内部灰尘，重新定标后修复 by jesson | SWO-222912 | Cleaned | HW | Incubator Module | Sensor | 2877 | Bid termination | Bid termination, need to borrow 944-214 batch number: 13002 | Clean the internal dust and repair by jesson after re-calibration
No | Visit | Complaint | RCHN | CPU单元损坏 | cpu单元损坏，借备件903-381,933-306各一个 | 需重新加借备件，故重新开单by tim | SWO-223227 | [5 empty] | CPU unit damaged | Cpu unit damage, spare parts 903-381,933-306 each | Spare parts need to be re-borrowed, so re-billing by tim
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 测试卡参比电极错误 1361 | 更换测试卡后修复 | SWO-223294 | Assisted Replacement | SC | SC | Reference |  | Test card, reagent package failure | Test card reference electrode error 1361 | Repair after replacing test card
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | CO2定标不过0210，LAC定标不过，1310 | 连续更换2张测试卡后修复 | SWO-223232 | Assisted Replacement | SC | SC | pCO2 |  | Test card, reagent package failure | CO2 calibration but 0210, LAC calibration but, 1310 | Repair after 2 consecutive test cards
No | Visit | Complaint | RCHN | 报错5784,4726 | 报错5784,4726，借备件903-198一个 | 更换903-198后修复 by jesson | SWO-223212 | Defect | HW | Analyzer Control Module | PCB, Main Board | 5784 | 57844726 | 5784,4726, spare parts 903-198 one | Repair by jesson after replacement 903-198
No | Visit | Complaint | RDEU | 21.08./08:15 Uhr, Tel.: 006921967326, Frau Fürstmann: Kundin meldet Probenaspiration gescheitert - ABL steht in messbereit | Ventil hängt sporadisch | Ventil gewechselt und in Betrieb genommen,alle Tests bestanden | SWO-223910 | Defect | HW | Oxi module | Pinch valve tube |  | 21.08./08:15 o'clock, Tel.: 006921967326, Mrs. Fürstmann: Customer reports sample aspiration failed - ABL is ready for measurement | Valve hangs sporadically | Valve changed and put into operation,all tests passed
No | 5 | Y | RFRA | Blank | PERTE CONNECTION AQURE |  | D305624AA435484E | Setup | SW | His/lis | AQURE |  | Blank | PERTE CONNECTION AQURE
No | Visit | Complaint | RDEU | 18.08.2017 12:30 Sr. Lisa 04088908652: Touch-Screen dejustiert auf der angegebenen Telefonnummer reagiert niemand, ralse, korrekt ~601 | Touchscreen lässt sich nicht richtig bedienen. | Neuen Touchscreen eingebaut. | SWO-223350 | Defect | HW | DMS_PC | Touch screen |  | 08/18/2017 12:30 Sr. Lisa 04088908652: Touch-Screen readjusted no one responds to the specified phone number, ralse, correctly . . . . . . . . . . . . . . . . . . . . . . . . | Touchscreen is not operated properly. | New touchscreen built-in.
No | 2 | Y | RAME | P Analyzer stuck in loop. |  | Analyzer stuck in loop. Randomly failing Pump, SP, or LS tests. Swapped SP, SC, inlet with limited success. Backflush and repeated cals with no success. | D305B02AA41478F2 | Defect | HW | Basic Unit | PM kit | FT | P Analyzer stuck in loop. |  | Analyzer stuck in loop. Randomly failing Pump, SP, or LS tests. Swapped SP, SC, inlet with limited success. Backflush and repeated cals with no success.
No | Visit | Complaint | RBNL | regelmatig error code 0964: Lekstroom gedetecteerd. Oorzaak is het reagenspack lot nr. YN-56 | Regelmatig error code 0964: Lekstroom gedetecteerd. Oorzaak is het reagenspack lot nr. YN-56 | Defecte RP lot nr YN-56 heb vervangen. | SWO-223668 | Assisted Replacement | SP | SP | Leak current | 964 | Regular error code 0964: leakage current detected. Cause is the reagent pack lot Nr. YN-56 | Regular error code 0964: leakage current detected. Cause is the reagent pack lot Nr. YN-56 | Defective RP lot nr YN-56 have replaced.
Yes | 7 | Y | RCZE | Blank |  |  | D305A595A407C190 | Re classified RMED | SWO/FAN |  |  |  | Blank
No | 4 | Y | RPAC | Blank | Sensor cassette and Solution pack incompatible after SW upgrade |  | D3059444A4315A10 | Defect | SW | Incompatible consumables | after SW upgrade | 1525&1526 | Blank | Sensor cassette and Solution pack incompatible after SW upgrade
No | Visit | Complaint | RSAF | cant get inlet module removed to replace the Solution pack. | Analyzer in the intervention required field replacement of solution pack required, entered service mode and done all necessary tests and got a vacuum error. Analyzer not working since 19/08/2017 | Done back flush and cleaned badly contaminated inlet gasket, done pressure test again and all passed. | SWO-223947 | Defect | SW | SP replacement | FPGA | Solved 3.2MR1 | cant get inlet module removed to replace the Solution pack. | Analyzer in the intervention required field replacement of solution pack required, entered service mode and done all necessary tests and got a vacuum error. Analyzer not working since 19/08/2017 | Done back flush and cleaned badly contaminated inlet gasket, done pressure test again and all passed.
No | Visit | Complaint | RCHN | CPU单元损坏 | cpu单元损坏，借备件903-381,933-306各一个,现加借905-914一个 |  | SWO-223289 | [5 empty] | CPU unit damaged | Cpu unit damage, one spare part 903-381,933-306 each, now 905-914 one
No | 9 | Y | RAME | P Possible fluid blockage during patient sample. |  | They were running patient samples in the Cath lab and it gave a possible fluid blockage. I did an annual on the analyzer on 8-15-17. | D305301CA44BA41C | [5 empty] | P Possible fluid blockage during patient sample. |  | They were running patient samples in the Cath lab and it gave a possible fluid blockage. I did an annual on the analyzer on 8-15-17.
No | Visit | Complaint | RCHN | 定标失败 | 定标失败，借945-377一个。 |  | SWO-223283 | [5 empty] | Bid failed | The bid failed, borrowing 945-377 one.
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | Lac定标不过 | 连续更换三张测试卡后修复 | SWO-223295 | Assisted Replacement | SC | SC | Lac |  | Test card, reagent package failure | Lac calibration is not | Repair after replacing three test cards in a row
No | 3 | Y | RAME | V Customer states unit not running AutoCheck modules. |  | Upon arrival, found Inlet caught in Sng Probe to oftware Loop not allow extend. | D3050DB6A4105A2B | Adjustment | HW | Inlet module | Inlet module |  | V Customer states unit not running AutoCheck modules. |  | Upon arrival, found Inlet caught in Sng Probe to oftware Loop not allow extend.
No | 17 | Y | RAME | F Customer reports various oximetry errors |  | Neon lamp too low, neon lamp intensity too low and hemolyzer phase regulation errors are present. | D305494BA4223478 | Defect | HW | Oxi module | Hemolyzer |  | F Customer reports various oximetry errors |  | Neon lamp too low, neon lamp intensity too low and hemolyzer phase regulation errors are present.
No | Visit | Complaint | RIND | frequently instrument going to Hold mode and it showing inlet flap is open. | NA | upgrade the application software and checked the instrument for same error and checked the status in service instrument working fine instrument under observation. | SWO-225458 | Adjustment | SW | SWU | SWU |  | frequently instrument going to Hold mode and it showing inlet flap is open. | In | upgrade the application software and checked the instrument for same error and checked the status in service instrument working fine instrument under observation.
No | Visit | Complaint | RDEU | 21.08.2017, Sr. Alexandra, Tel.: 020885084343: Halterung im Einlass abgebrochen; 2 Proben sin schon im AQT verschwunden | Halterung für die Proben ist abgebrochen, | neue Tube holder eingebaut, | SWO-223625 | Defect | HW | Aspiration Module | Tube holder |  | 21.08.2017, Sr. Alexandra, Tel.: 020885084343: Bracket in the entrance aborted; 2 samples already disappeared in AQT | Bracket for the samples is aborted, | new tube holder built-in,
No | Visit | Complaint | RSCH | Rep | Immer Fehler 1401 Einlassdichtungshalter Bricht Messungen ab. Fehler passiert nur Kapillär | Kein Fehler festgestellt. Testmessungen alle i.o | SWO-223814 | Defect | HW | Auto Inlet module | Auto Inlet module | Capilary 1401 | Rep | Always error 1401 inlet seal holder Cancels measurements. Mistake only happens to Kapillär | No error detected. Test measurements all i.o
No | 18 | Y | RAME | PUnit sent in due to fluid flow issue. Possible fluid blockage | Changed the complaint vist tab from N to Y. | There is a leak in the SP housing area. Suspect waste line clogged. | D306C84C1A570EB8 | [5 empty] | PUnit sent in due to fluid flow issue. Possible fluid blockage | Changed the complaint vist tab from N to Y. | There is a leak in the SP housing area. Suspect waste line clogged.
No | 10 | Y | RAME | P Intermittent rinse error, aborting Cal1 |  | checked data logs for errors, checked flow of solutions, checked pump tubing and pump rotors, checked valves and liquid sensors | D306ACFE1A521C27 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | P Intermittent rinse error, aborting Cal1 |  | checked data logs for errors, checked flow of solutions, checked pump tubing and pump rotors, checked valves and liquid sensors
No | 10 | Y | RAME | Z Unit is using gas 2 tanks at a rate of one a month. |  | Checked data logs. Cal 1 gas was recently changed within a month's time. | D306A06E1A62164B | Defect | HW | Gas module | Tubings, Gas |  | Z Unit is using gas 2 tanks at a rate of one a month. |  | Checked data logs. Cal 1 gas was recently changed within a month's time.
No | 7 | Y | RAME | G Customer reported calibration will not pass for Crea. |  | Checked data logs. Found high Crea check failing. Parameters are passing for Crea except for the high Crea check at -9. Cleaning solution and membranes already replaced multiple times. | D306BE241A4F6533 | Assisted Replacement | Consumables | Solution bottle | Cleaning, Crea |  | G Customer reported calibration will not pass for Crea. |  | Checked data logs. Found high Crea check failing. Parameters are passing for Crea except for the high Crea check at -9. Cleaning solution and membranes already replaced multiple times.
No | 3 | Y | RAME | V Inlet calibration errors. |  | Unable to calibrate inlet in user or service modes. | D3067CE31A2D1643 | Defect | HW | Inlet module | Friction detector unit |  | V Inlet calibration errors. |  | Unable to calibrate inlet in user or service modes.
No | 8 | Y | RAME | Z Will not power on. |  | Was able to recreate the customers complaint. | D3068EA91A22B8F5 | [5 empty] | Z Will not power on. |  | Was able to recreate the customers complaint.
No | Visit | Complaint | RHUN | Tű elgörbült, kiegyenesítés | Mintafelszívási hiba. | Bemeneti tű elgörbült, kiegyenesítése után a mintafelszívás megfelelő. | SWO-224289 | Adjustment | HW | Inlet module | Probe bend |  | Needle Bent, straightening | Mintafelszívási hiba. | Input needle is bent and straightened after the sample is absorbed properly.
No | 10 | Y | RPAC | Blank | 1361: Internal Reference electrode error |  | D307E4219009AB30 | Assisted Replacement | SC | SC | Reference | 1361 | Blank | 1361: Internal Reference electrode error
Yes | 10 | Y | RPAC | Blank | Install loan instrument, take away customer instrument |  | D307CFC9905218FB | Re classified RMED | same day UpSWO |  |  |  | Blank | Install loan instrument, take away customer instrument
No | Visit | Complaint | RDEU | 22.08.2017, 09:18 Uhr, Fr. Neidhardt, Tel. 016098735800: Intervention erforderlich | Intervension erforderlich Lösungen konnten nicht gepumpt werden LP wurde durch den Kunden erneuert | Fehlersuche Überprüfung Rollenkranz, Einlaß gewartet LP erneut initialisiert Kal.lösung 3 inhomogen, bzw. wird nicht richtig gepumpt LP müßte whr. erneuert werden, aber kein LP mehr vorhanden ( ist bestellt), Kunde wird nach Erhalt LP erneuern | SWO-224352 | Assisted Replacement | SP | SP | FT | Inhom Cal3 | 22.08.2017, 09:18, Ms. Neidhardt, tel. 016098735800: Intervention required | Intervension required Solutions could not be pumped LP has been renewed by the customer | Troubleshooting Review roller wreath, admission serviced LP reinitialized Cal.solution 3 inhomogeneous, or is not pumped properly LP would have to be whr. be renewed, but no LP is available (is ordered), Customer will renew LP after receipt
No | Visit | Complaint | RCHN | 报浑浊度过高，OXI光谱不匹配 | 报浑浊度过高，OXI光谱不匹配，借902-980一个 | 更换902-980后修复by tim | SWO-223706 | Defect | HW | Oxi module | Hemolyzer | 292 | Reported cloudy through high, OXI spectrum mismatch | Reported turbidity over high, OXI spectrum mismatch, borrow 902-980 a | Repair by tim after replacement 902-980
No | Visit | Complaint | RESP | Sp no detectado durante la instalacion | Chequeo de solution packID Pcb y sustitucion de ésta | Sustitucion de Solution pack pcb | SWO-224370 | Defect | HW | Basic Unit | PCB Solution Pack ID | 1152 & 1134 | Sp not detected during installation | Checking solution packID Pcb and substituting it | Sustitucion de Solution pack pcb
No | Visit | Complaint | RESP | No transmite a infinity Error en calibracion de bombas | Chequeo de conexiones y datos de configuracion Chequeo de neumatica e inicio completo y calibracion | Se habla con informatica y de activa punto de red Incio y calibracion | SWO-224305 | Setup | SW | His/lis | His/lis |  | Doesn't transmit to infinity Pump calibration error | Checking connections and configuration data Tire check and full start and calibration | Talks about informatics and active network point Application and calibration
No | Visit | Complaint | RDAN | Skærm fryser eller går i sort | Når der trykkes på skærmen sker der ikke noget. | Der er lagt ny firmware i skærmen. Dette skulle afhlælpe fejlen. Alt er OK nu. | SWO-224385 | Defect | HW | DMS_PC | Touch screen |  | Screen freezes or goes in black | When the screen is pressed, nothing happens. | New firmware has been added to the screen. This should be the fault. Everything is OK now.
No | 6 | Y | RAME | F Customer reported oximetry errors. |  | System messages showing error code 0258: Neon Intensity Too Low and 0316: Neon Number Too Low. Ran complete oximetry tests with failures on lamp test. | D307D8C49001552E | Defect | HW | Oxi module | Lamp unit |  | F Customer reported oximetry errors. |  | System messages showing error code 0258: Neon Intensity Too Low and 0316: Neon Number Too Low. Ran complete oximetry tests with failures on lamp test.
No | Visit | Complaint | RDEU | 23.08.2017, 10:15 Uhr, H. Brevis, Tel. 02037333204: 2 Pkt. Kal wird abgebrochen | Pumpenschlauch im pH / BG Modul defekt. Pumpenasymetrie Fehler | Pumpenschlauch im pH $ BG Modul erneuert. Gerät kalibriert und QC durchgeführt. | SWO-224509 | Defect | HW | pH/BG Module | Pump tubes |  | 23.08.2017, 10:15 a.m., H. Brevis, tel. 02037333204: 2 pts. Kal is canceled | Pump hose in pH / BG module defective. Pump aymetry error | Pump hose in the pH BG module renewed. calibrated and QC performed.
Yes | 5 | Y | RFRA | Blank | Err démarrage application Radiometer/ |  | D308D9CB061F6A3A | [5 empty] | Blank | Err start-up application Radiometer/
No | 5 | Y | RPAC | Blank | Software Update causes consumable issue |  | D308F15E06521893 | Defect | SW | Incompatible consumables | after SW upgrade | 1525&1526 | Blank | Software Update causes consumable issue
No | Visit | Complaint | RCHN | 触摸屏失灵 | 触摸屏失灵，借552-039一个 | 更换552-039后修复by tim | SWO-223930 | [5 empty] | Touch screen failure | Touch screen failed, borrow 552-039 a | Repair by tim after replacement 552-039
No | Visit | Complaint | RCHN | 测量仓内针损坏 | 测量仓内针损坏，借备件903-581一个 | 更换903-581后修复by tim | SWO-224612 | Defect | HW | Basic Unit | PCB, Sensor interface | Contact broken | Measuring needle damage in the chamber | Measure the needle damage in the warehouse, borrow spare parts 903-581 one | Repair by tim after replacement 903-581
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 液体传输失败，1243 | 更换试剂包后修复 | SWO-224797 | Assisted Replacement | SP | SP | FT |  | Test card, reagent package failure | Liquid transfer failure, 1243 | Repair after replacement of reagent kit
No | Visit | Complaint | RCHN | 报错1024 | 报错1024，借PCT定标卡944-394（12726）一张 | 重新定标后修复by tim | SWO-224174 | Assisted Replacement | Consumables | Test Kit | Cartridge | PCT | 1024 error | 1024, borrowpct PCT calibration card 944-394 (12726) one | Fix by tim after re-calibration
No | Visit | Complaint | RCHN | 报错3412 | 报错3412，需要借件944-394（批次号：12726） | 重新定标后修复by tim | SWO-224175 | Defect | HW | Aspiration Module | Needle Wash Unit | 3412 | Error 3412 | 3412, need to borrow 944-394 (batch number: 12726) | Fix by tim after re-calibration
No | Visit | Complaint | RIND | Printer head found broken . | printer head found broken . | Has to replace the printer head. Part Number:910-318 | SWO-225528 | Defect | HW | Printer Module | Printer module |  | Printer head found broken . | printer head found broken . | Has to replace the printer head. Part Number:910-318
Yes | 4 | Y | RCZE | Blank |  |  | D308CF0306438ED4 | Re classified RMED | FAN |  |  |  | Blank
No | Visit | Complaint | RBNL | Techline case wegens Health Hazard | Wegens vals verhoogde uitslagen, Health Hazard Case heb aangemaakt. | Techline Case en Servicedump gemaakt. | SWO-225538 | Unknown | False results | Health hazard | Data exported | RBNL 788752285 | Techline case due to Health Hazard | Due to false increased results, Health Hazard Case have been created. | Techline Case and Service dump created.
No | Visit | Complaint | RKK | 装置底部から廃液もれ発生 | 状態確認 | 排液マニホールド交換 | SWO-225592 | [5 empty] | Waste liquid from the bottom of the device | Status Check | Drainage manifold replacement
No | Visit | Complaint | RCHN | 报错5669,5670 | 报错5669,5670，借备件910-343一个 | 更换910-343后修复by tim | SWO-224588 | Defect | HW | Analyzer Control Module | CPU unit |  | 56695670 | 5669,5670, spare parts 910-343 one | Repair by tim after replacement 910-343
No | Visit | Complaint | RCHN | 报错1024 | 报错1024，借备件903-364一个 | 更换903-364一个正常。by tim | SWO-224685 | Defect | HW | Process Module | Optical Unit | 1024 | 1024 error | Error 1024, borrowed spare parts 903-364 one | Replace 903-364 one normal. by tim
No | Visit | Complaint | RCHN | 样本夹损坏 | 样本夹损坏，借备件903-584一个 | 更换903-584后修复by tim | SWO-224560 | Defect | HW | Aspiration Module | Tube holder |  | Damaged sample clip | Sample clip damage, spare parts 903-584 one | Repair by tim after replacement 903-584
No | Visit | Complaint | RESP | Fuga | Comprobación de registros de datos del equipo. | Comprobación de tubos, de la parte neumática, comprobación de válvulas, comprobación de cámaras de medida, comprobación de electrodos, se observa que la membrana del Ca+ está desmontada por la parte inferior, se sustituye la membrana del Ca+ y verificamos el funcionamiento con QC. | SWO-225344 | User error | replacement | replacement | replacement |  | flight | Checking computer data records. | Checking pipes, pneumatic part, valve check, measuring chamber check, electrode check, seeing that the Ca+ membrane is disassembled at the bottom, replacing the Ca+ membrane and checking the qc.
No | Visit | Complaint | RDEU | 03.08.2017, Herr Kohl, Tel: 03441860814: Intervention erforderlich 03.08.2017, Herr Kohl, Tel.: 03441860814: weiterhin Intervention erforderlich | Kalibrationen werden abgebrochen | Leihgerät abgeholt | SWO-225333 | Revisit required | HW | HW | HW |  | 03.08.2017, Mr Kohl, Tel: 03441860814: Intervention required 03.08.2017, Mr Kohl, Tel.: 03441860814: continued intervention required | Calibrations are canceled | Rental device picked up
No | Visit | Complaint | RDEU | 25.08.2017 10:30 H. Neitzel 02151324460: Kal-Fehler Elythe/pH 25.08.2017 13:30 H. Neitzel 02151324460: pH bleibt bei der Kal fehlerhaft | Messkammer und pH Elektrode verunreinigt. | pH Elektrode und Messkammer gereinigt. Gerät kalibriert und QC durchgeführt. | SWO-225444 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | 08/25/2017 10:30 H. Neitzel 02151324460: Kal Error Elythe/pH 25.08.2017 13:30 H. Neitzel 02151324460: pH remains faulty with the Kal | measuring chamber and pH electrode. | pH electrode and measuring chamber. calibrated and QC performed.
No | Visit | Complaint | RCHN | 信号板故障 | 信号板故障，借备件903-334一个 | 更换903-334备件后修复by tim | SWO-224570 | [5 empty] | Signal plate failure | Signal board failure, spare parts 903-334 one | Repair by tim after replacement of 903-334 spare parts
No | Visit | Complaint | RCHN | 报错1029,1031,3410 | 报错1029,1031,3410，借备件903-364,903-236,944-258（13152）一张 | 更换903-364,903-236,944-258（13152）后重新定标修复by tim | SWO-224594 | Defect | HW | Process Module | Optical Unit | 1031 | 102910313410 | 1029,1031,3410, spare parts 903-364,903-236,944-258 (13152) | Replace 903-364, 903-236, 944-258 (13152) after re-calibration repair by tim
No | Visit | Complaint | RCHN | 报错1024，定标底物外漏 | 报错1024，定标底物外漏，借TNI944-212定标卡（13079）一张 | 更换944-212后修复by tim | SWO-224693 | Assisted Replacement | Consumables | Test Kit | Cartridge | TnI | Error 1024, bid substrate leakage | Report error 1024, bid substrate leakage, borrow TNI944-212 calibration card (13079) one | Repair by tim after replacement 944-212
No | Visit | Complaint | RCHN | 血氧模块硬件错误 | 血氧模块硬件错误，借备件902-833一个 | 更换902-833备件后修复by tim | SWO-225155 | Defect | HW | Oxi module | Oxi module, Replaced | 258 | Blood oxygen module hardware error | Blood oxygen module hardware error, borrowed spare parts 902-833 one | Repair by tim after replacement of 902-833 spare parts
No | Visit | Complaint | RCHN | 每次更换试剂包或者测试卡时有些参数质控都不过 | 每次更换试剂包或者测试卡时有些参数质控都不过，借903-135和903-183各一个 | 更换903-135和903-183(903-581)备件后修复by tim | SWO-225128 | Defect | HW | Basic Unit | PCB, Sensor interface | 0089 | Some parameters are not enough to be replaced with a reagent kit or test card | Each time the reagent kit or test card is replaced some parameters are not how much, borrow 903-135 and 903-183 each one | Repair by tim after replacement of 903-135 and 903-183 (903-581) spare parts
No | Visit | Complaint | RCHN | 报错1031 | 报错1031，借备件903-364一个 | 更换903-364备件后修复by tim | SWO-225098 | Defect | HW | Process Module | Optical Unit | 1031 | 1031 | Error 1031, borrowed spare parts 903-364 one | Repair by tim after replacement of 903-364 spare parts
No | Visit | Complaint | RCHN | 血氧检测失败 | 血氧检测失败，借902-527一个 | 已使用备用件修复(By Ran) | SWO-225773 | Defect | HW | Oxi module | Pinch valve | 581 | Failed blood oxygen test | Blood oxygen test failed, borrow edi-902-527 one | Repaired with spare parts (By Ran)
No | 7 | Y | RAME | 9 Lac having issue on analyzer with qc internal level 3 running high. |  | Lac having issue on analyzer with qc internal level 3 running high. | D30E6E18DE34BF85 | Assisted Replacement | SC | SC | Lac | QC C High | 9 Lac having issue on analyzer with qc internal level 3 running high. |  | Lac having issue on analyzer with qc internal level 3 running high.
Yes | 5 | Y | RPAC | Blank | Software Update to V3.3 MR1 |  | D30E0863DE1D78FA | Re classified RMED | SWU |  |  |  | Blank | Software Update to V3.3 MR1
No | 2 | Y | RPAC | Blank | Inlet jamming |  | D30DF712DE2D1898 | Adjustment | HW | Auto Inlet module | Inlet drive gear | Service tip issued | Blank | Inlet jamming
Yes | 2 | Y | RPAC | Blank |  |  | D30E0529DE6071EE | Re classified RMED | Courtesy |  |  |  | Blank
No | 1 | Y | RPOL | Blank | Kalibracja inletu, odbia³czanie analizatora, kalibracje, testy kontrolne. Analizator sprawny. |  | D30E15D5DE559782 | Adjustment | HW | Auto Inlet module | Probe cal. |  | Blank | Calibration of the inlet, analyser, calibrations, control tests. Functional Analyzer.
No | 1 | Y | RPOL | Blank | Kalibracja inletu, odbia³czanie analizatora, kalibracje, testy kontrolne. Analizator sprawny. |  | D30E1697DE549784 | Adjustment | HW | Auto Inlet module | Probe cal. |  | Blank | Calibration of the inlet, analyser, calibrations, control tests. Functional Analyzer.
No | 3 | Y | RCZE | Blank |  |  | D30F9289545E8F0A | [5 empty] | Blank
No | Visit | Complaint | RESP | Botle Ilumination | Comprobación de registros de datos del equipo. Comprobación de sensores de botes de líquidos. | Sustitución de sensores de Botes de líquidos. Comprobación del correcto funcionamiento del equipo con QC | SWO-225848 | Defect | HW | Basic Unit | PCB, Bottle illumination | Used analyzer | Botle Ilumination | Checking computer data records. Checking liquid boat sensors. | Replacing liquid boat sensors. Checking the correct operation of the equipment with QC
Yes | 7 | Y | RCZE | Blank |  |  | D30DDE59DE193CA7 | Re classified RMED | FAN |  |  |  | Blank
No | Visit | Complaint | RKK | pCO2キャルエラー感度が全くとれない。 |  |  | SWO-226847 | Unknown | Electrode | Electrode | pCO2 | LEH mail | pCO2 cal error sensitivity can not take at all.
No | Visit | Complaint | RKK | サンプル移動エラー：ｐＨ／ＢＧ側リンスバルブ、エアバルブ交換実施。 | サンプル移動エラー：ｐＨ／ＢＧ側リンスバルブ、エアバルブ交換実施。 | サンプル移動エラー：ｐＨ／ＢＧ側リンスバルブ、エアバルブ交換実施。 | SWO-225815 | Defect | HW | pH/BG Module | Valve |  | Sample transfer error: pH/BG side rinse valve, air valve replacement. | Sample transfer error: pH/BG side rinse valve, air valve replacement. | Sample transfer error: pH/BG side rinse valve, air valve replacement.
Yes | 4 | Y | RCZE | Blank |  |  | D30DE65DDE383CA9 | Re classified RMED | FAN |  |  |  | Blank
Yes | 5 | Y | RCZE | Blank |  |  | D30DDA1FDE258F04 | Re classified RMED | FAN |  |  |  | Blank
Yes | 4 | Y | RCZE | Blank |  |  | D30DD3EDDE3C8EFA | Re classified RMED | FAN |  |  |  | Blank
Yes | 3 | Y | RCZE | Blank |  |  | D30DD61CDE5B8EFC | Re classified RMED | FAN |  |  |  | Blank
Yes | 3 | Y | RCZE | Blank |  |  | D30DD6DEDE048EFE | Re classified RMED | FAN |  |  |  | Blank
No | 7 | Y | RAME | P leaking from inlet after maintenance. |  | Inspected all tubings that were replaced and all looked good. | D30DCE3ADE0EB84A | Defect | HW | Inlet module | Nipple kit |  | P leaking from inlet after maintenance. |  | Inspected all tubings that were replaced and all looked good.
Yes | 3 | Y | RCZE | Blank |  |  | D30DD7EEDE148F00 | Re classified RMED | FAN |  |  |  | Blank
Yes | 2 | Y | RCZE | Blank |  |  | D30DD96FDE2F8F02 | Re classified RMED | FAN |  |  |  | Blank
No | 6 | Y | RCZE | Blank | po opravì instalace na ARO, z pùvodního pøeneseny elektrody a spotøební materiál, |  | D30EF88054083D39 | Revisit required | HW | Basic Unit | PCB, Analyzer control |  | Blank | After the installation has been repaired to ARO, the electrodes and consumables
No | Visit | Complaint | RLTD | Mori on 01865 221596 | On arrival the analyser was in intervention required and failing most of the solution pack tests. I put in my test cassette and ran 2 solution pack tests which both passed and I also ran a solution pack test using another expired cassette which also passed. Checking through the datalogs it looks like a clotted sample was introduced into the analyser as a sample was run 25/08 23:20 & the fluid transport errors started straight after that. It looks like the flushing done at the time probably cleared most of the clot and anything else remaining has worked its way out of the analyser over the weekend. | Fitted a new sensor cassette and that started up ok. | SWO-226415 | No fault found | No fault found | No fault found | Clot |  | Mori on 01865 221596 | On arrival the analyser was in intervention required and failing most of the solution pack tests. I put in my test cassette and ran 2 solution pack tests which both passed and I also ran a solution pack test using another expired cassette which also passed. Checking through the datalogs it looks like a clotted sample was introduced into the analyser as a sample was run 25/08 23:20 & the fluid transport errors started straight after that. It looks like the flushing done at the time probably cleared most of the clot and anything else remaining has worked its way out of the analyser over the weekend. | Fitted a new sensor cassette and that started up ok.
No | 2 | Y | RPOL | Blank | Awaria systemu, analizator nie uruchamia siê , zabrano do serwisu. |  | D30F5E6C54170251 | Revisit required | Analyzer removed | Analyzer removed | Analyzer removed |  | Blank | System failure, analyzer does not start, taken to service.
No | 16 | Y | RPNZ | Blank | Sensor cassette replacement due to 1361: internal reference electrode error. |  | D30FB2225459E321 | Assisted Replacement | SC | SC | Reference | LEH PER kontrol : Én SC PER modtaget i Marts 2017, Error Code 1361 SP : 3 stk Leak Current (Dec-Feb-Apr)..1 stk i July pO2 QC | Blank | Sensor cassette replacement due to 1361: internal reference electrode error.
No | Visit | Complaint | RDEU | 28.08.2017, 11:11 Uhr, OA Dr. Bockenholt, Tel. 0221890712598: Einlassnadel fährt nicht automatisch zurück | Einlassnadel hakt an der Einlassdichtung. Öfters Probenabbrüche. | Einlassnadel ist leicht krum. Einlassnadel (Kundenbestand) und einlassdichtung erneuert. Einlassanschlussdichtung ist beschädigt. Anschlussdichtung aus Kundenbestand erneuert. Mehrere Funktionskontrollen mit tHb Kallösung waren alle OK. | SWO-226165 | Replaced | HW | Auto Inlet module | Probe bend | Gasket | 28.08.2017, 11:11 a.m., OA Dr. Bockenholt, Tel. 0221890712598: Inlet pin does not return automatically | Inlet pin hooks on the inlet seal. Frequently sample interruptions. | Inlet needle is slightly crooked. Inlet pin (customer stock) and inlet seal renewed. Inlet connection seal is damaged. Connection seal from customer stock renewed. Several functional controls with tHb calsolution were all OK.
No | Visit | Complaint | RSCH | Blank | Bruit important de la carte mère (Main Board CPU) Latence du système (application lente) Sensibilité du lactate très bas malgré le changement de SC | Remplacé le Main Board CPU e.o Test du canal Lactate avec le simulateur ODIN sensor MK2 e.o Remplacé le kit de maintenance e.o Calibration e.o CQ e.o | SWO-226282 | Defect | HW | Basic Unit | PCB, Main Board |  | Blank | Major Motherboard Noise (Main Board CPU) System latency (slow application) Lactate sensitivity very low despite SC change | Replaced the Main Board CPU e.o Lactate canal test with the ODIN sensor MK2 e.o simulator Replaced the e.o maintenance kit Calibration e.o CQ e.o
No | Visit | Complaint | RDEU | QC-Fehler bei Level 1 und Level 2 (QC-Werte außer Bereich) | Fehlersuche --> Dateiarchive (Kal., QC und Aktivitäten) überprüft --> Problem tritt seit 23.8.17 öfters auf --> SC-Wechsel erfolgte am 23.8.17 und LP-Wechsel am 31.7.17 SC (240890) und LP (239741) --> SC (238256) aus GS19-Bestand eingebaut -->nicht in Ordnung (Kal.-Werte pH 40,4 (40-65); pCO2 55,8 (30-75); pO2 0,008 (0,002-0,052) --> SC nochmals erneuern!!! --> Schlauchwege überprüft --> Schlauch eingequetscht --> Schlauchwege erneuern --> Zeroing Anweisung durchgeführt | SC (241498) aus Kundenbestand eingebaut, SP (237313) aus GS19-Bestand eingebaut und (239741) ausgebaut, Schlauchwege erneuert, Zeroing Anweisung durchgeführt, System und Funktion soweit in Ordnung, bis auf QC (sporadisch) --> Signal Data Board und Sensor Board bestellt!!! | SWO-226361 | [5 empty] | QC error at Level 1 and Level 2 (QC values out range) | Troubleshooting --> File archives (Cal., QC and activities) checked --> Problem occurs more often since 23.8.17 --> SC change took place on 23.8.17 and LP change on 31.7.17 SC (240890) and LP (239741) --> SC (238256) built from GS19 stock -->not okay (cal.-values pH 40.4 (40-65); pCO2 55.8 (30-75); pO2 0.008 (0.002-0.052) --> Check SC again!!! --> Hose paths checked --> Hose squeezed --> Renew hose paths --> Zeroing instruction performed | SC (241498) built from customer base, SP (237313) from GS19 stock installed and (239741) removed, hoseways renewed, zeroing instruction carried out, system and function as far as in order, except for QC (sporadically) --> Signal Data Board and Sensor Board!!!
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | K,HCT定标失败气泡探测定标失败 | 更换测试卡 | SWO-226141 | [5 empty] | Test card, reagent package failure | K, HCT calibration failed bubble detection calibration failure | Replace the test card
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | LAC定标失败 | 更换测试卡 | SWO-226202 | [5 empty] | Test card, reagent package failure | LAC calibration failure | Replace the test card
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | pO2定标不过 | 更换试剂包后修复 | SWO-226271 | Assisted Replacement | SP | SP | pO2 |  | Test card, reagent package failure | pO2 calibration but | Repair after replacement of reagent kit
No | Visit | Complaint | RESP | INSTALACION DE EQUIPO REPARADO EN MADRID. | INSTALAR | INSTALACION Y DESINSTALACION DE EQUIPO PRESTADO, PRUEBAS DE FUNCIONAMIENTO Y CONTROLES DE CALIDAD. | SWO-226398 | Assisted Replacement | SC | SC | Lac | Mail | INSTALLATION OF REPAIRED EQUIPMENT IN MADRID. | Install | INSTALLATION AND DISINSTALLATION OF PROVIDED EQUIPMENT, OPERATION TESTS AND QUALITY CONTROLS.
Yes | 2 | Y | RFRA | Blank |  |  | D30F6909544B2E02 | Re classified RMED | SWU |  |  |  | Blank
No | Visit | Complaint | RLTD | When reporting a Sodium problem customer mentioned they had cleaned up a leak around the H electrode. | Inspect pH chamber, still dries blood around electrode jacket and under sealing ring and tip of electrode. Also blood behind chamber turret, and discovered bad bend in inlet probe. | Remove electrode block, remove pH $Bg measuring chamber and thoroughly clean pH turret inside and out, no damage to chamber just blood ingress around outside. Clean pH electrode and jacket and fit new pH sealing ring to tip of electrode. Straighten bent inlet probe. | SWO-226434 | Cleaned | HW | pH/BG Module | Measuring chamber |  | When reporting a Sodium problem customer mentioned they had cleaned up a leak around the H electrode. | Inspect pH chamber, still dries blood around electrode jacket and under sealing ring and tip of electrode. Also blood behind chamber turret, and discovered bad bend in inlet probe. | Remove electrode block, remove pH $Bg measuring chamber and thoroughly clean pH turret inside and out, no damage to chamber just blood ingress around outside. Clean pH electrode and jacket and fit new pH sealing ring to tip of electrode. Straighten bent inlet probe.
No | 11 | Y | RAME | 9 Customer reported Lactate will not calibrate or pass QC. |  | Checked data logs for errors. Lac drifting and has not passed QC. | D30F3EC55402CE74 | Defect | HW | El/Met Module | Pump tubes |  | 9 Customer reported Lactate will not calibrate or pass QC. |  | Checked data logs for errors. Lac drifting and has not passed QC.
No | Visit | Complaint | RDEU | 28.08.2017 17:45 Sr. Andrea 021028512912 pH-Elektrode weiterhin undicht. | Flüssigkeit kommt aus dem Gerät, Dichtungsring an der Lac.elektrode muss ersetzt werden, Dichtung an der Ca.membran fehlt, | Div.Dichtungsringe ersetzt, Elektrode neu bezogen, Fließwege und Messkammer gereinigt, Luftweg gereinigt, Gerät kalibriert und QC durchgeführt, | SWO-226375 | Defect | Electrode | Electrode | Ca+ | Electrode gasket | 28.08.2017 17:45 Sr. Andrea 021028512912 pH electrode continues to leak. | Liquid comes from the device, sealing ring on the lac.electrode must be replaced, seal on the ca.membrane is missing, | Div.Seal rings replaced, electrode re-covered, flow ways and measuring chamber cleaned, airway cleaned, device calibrated and QC performed,
No | 6 | Y | RCZE | Blank | osazení testovacími elektorodami ze Zlína, u Ca instalována nová membrány, zbytek z OIPOo, hodnoty ISE z El/met modulu bez statusu a se špatnou citùlivostí, chyba odezvy |  | D30EFB2454003D3C | Defect | HW | Basic Unit | PCB, Analyzer control |  | Blank | Installation of test electors from Zlín, U Ca installed new membrane, rest from OIPOo, ISE values from El/Met module without status and with poor sensitivity, error response
Yes | 3 | Y | RFRA | Blank | Pas d'impression |  | D30F384F540763F5 | Re classified RMED | OOB |  |  |  | Blank | No printing
No | 8 | Y | RAME | Z Calibration light blinks even after it is charged. |  | Calibration light blinks even after it is charged. | D30F301354620664 | [5 empty] | Z Calibration light blinks even after it is charged. |  | Calibration light blinks even after it is charged.
No | Visit | Complaint | RCHN | 提示测试卡未连接 | 提示测试卡未连接，借902-972一个。 | 更换902-972一个正常。(By Ran) | SWO-225689 | [5 empty] | Prompt the test card is not connected | Prompt the test card is not connected and borrows 902-972 one. | Replace 902-972 one normal. (By Ran)
No | Visit | Complaint | RCHN | 仪器报异常的应用程序，系统关机 | 仪器报异常的应用程序，系统关机，借933-364一个。 | 工程师报错了机器。关闭此派工单。(By Ran) Wrong SN. Case closed. | SWO-226472 | [5 empty] | Instrument reported abnormal application, system shutdown | Instrument reported abnormal application, system shutdown, borrowed 933-364 one. | The engineer reported the wrong machine. Close this ticket. (By Ran) Wrong SN. Case closed.
No | Visit | Complaint | RCHN | 仪器报异常的应用程序，系统要关机 | 仪器报异常的应用程序，系统要关机，借933-364一个。 |  | SWO-226478 | [5 empty] | Instrument reported abnormal application, system to shut down | Instrument reported abnormal application, the system to shut down, borrow 933-364 one.
No | Visit | Complaint | RDEU | Rote Flüssigkeit tritt im Beteich des LP aus | LP Verriegelung erneuert | LP Verriegelung erneuert, Abfallstutzen an Analysator überprüft ok, LP mehrfach herausgenommen und eingesetzt, mehrfach Kal. in IQC Kein Flüssigkeitsaustritt mehr feststellbar | SWO-226819 | [5 empty] | Red liquid leaks in the bepond of the LP | LP locking renewed | LP locking renewed, waste nozzle checked to analyzer ok, LP removed several times and inserted, multiple cal. in IQC No more liquid leakage detectable
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 试剂包废液口堵塞 | 更换试剂包正常。 | SWO-226964 | [5 empty] | Test card, reagent package failure | Reagent pack waste mouth blockage | The replacement kit is OK.
No | Visit | Complaint | RCHN | 打印机打印不清楚 | 打印机打印不清楚，借910-286一个。 | 更换910-286一个正常。(By Ran) | SWO-226197 | Defect | HW | Basic Unit | Printer module |  | The printer does not print clearly | The printer is not clearly printed, borrowing 910-286 one. | Replace 910-286 one normal. (By Ran)
Yes | 2 | Y | RPAC | Blank | TCm5 Gas used up in x2 days even when in standby & Turned off. Possible gas leak from connector valve. |  | D310827FCA065C4D | [5 empty] | Blank | TCm5 Gas used up in x2 days even when in standby & Turned off. Possible gas leak from connector valve.
No | Visit | Complaint | RCHN | PCT定标失败 2877 | PCT定标失败 2877 ，借944-394（12921） 一张。 |  | SWO-226574 | Assisted Replacement | Consumables | Test Kit | Cartridge, PCT | 2877 | PCT calibration failure 2877 | PCT calibration failure 2877 , borrow 944-394 (12921).
No | Visit | Complaint | RCHN | 安全位置报错 5784 | 安全位置报错 5784 ，借903-198一个。 | 更换903-198后修复(By Ran) | SWO-226170 | Defect | HW | Analyzer Control Module | PCB, Main Board | 5784 | Safe Location Error 5784 | Safe Location Error 5784 , borrowing 903-198. | Post-replacement 903-198 repair (By Ran)
No | Visit | Complaint | RCHN | AU检测失败 2912 | AU检测失败 2912，借903-198，636-564，636-541各一个。 |  | SWO-226459 | Defect | HW | Analyzer Control Module | PCB, Main Board | 2912 | AU Detection Failure 2912 | AU test failure 2912, borrowing 903-198, 636-564, 636-541 each.
No | Visit | Complaint | RCHN | 氧电极故障 | 氧电极故障，借945-613一个 | 更换945-613后修复(By Ran) | SWO-226482 | Assisted Replacement | Electrode | Electrode | pO2 |  | Oxygen electrode failure | Oxygen electrode failure, borrow 945-613 one | Repair after replacement 945-613 (By Ran)
No | Visit | Complaint | RDAN | Hanne + Anette melder om at inlet drypper og pH + salte er ustabile. | Undersøgt | pH $blodgas målekammer renset, pH elektrode renset indvendigt og forsynet med ny pakring, pumpeslanger udskiftet, slutafprøvet. | SWO-227447 | Revisit required | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | Hanne + Anette reports that the inlet drips and pH + salty are unstable. | Examined | pH $blodgas Measuring Chamber purified, pH electrode cleaned inside and equipped with new packing ring, pump hoses replaced, final tested.
No | 4 | Y | RAME | F Customer reported oximetry errors. |  | Checked data logs for errors. Found error codes 258 and 316. Ran complete oximetry test with failures on lamp unit for neon intensity and position. | D31057FCCA5DD87F | Defect | HW | Oxi module | Lamp unit |  | F Customer reported oximetry errors. |  | Checked data logs for errors. Found error codes 258 and 316. Ran complete oximetry test with failures on lamp unit for neon intensity and position.
No | Visit | Complaint | RDAN | Lene Jacobsen fortæller, at der er stor variation i Cl. Ebbe har mistanke til Cl elektrode. | Afprøvet og undersøgt. Logen viser, at Cl QC svinger meget. | 30-8-17: Cl elektrode nr 07-06 isat og kalibreret. | SWO-226705 | Assisted Replacement | Electrode | Electrode | Cl | QC / CAPA | Lene Jacobsen says there is a lot of variation in Cl. Ebbe suspect Cl electrode. | Tested and investigated. The gallery shows that the Cl QC fluctuates a lot. | 30-8-17: Cl electrode nr 07-06 inserted and calibrated.
No | 6 | Y | RAME | Y Customer is sending unit in for evaluation due to repeated calibration failures. | Changed the complaint visit tab from N to Y. | Able to replicate error. | D31040F6CA60BE85 | [5 empty] | Y Customer is sending unit in for evaluation due to repeated calibration failures. | Changed the complaint visit tab from N to Y. | Able to replicate error.
Yes | 4 | Y | RPOL | Blank | sprawdzenie analizatora, wymiana uszczelki, kalibracja. Analizator sprawny technicznie. |  | D311BAAF404D6892 | [5 empty] | Blank | Analyzer check, gasket replacement, calibration. Technically efficient analyzer.
Yes | 11 | Y | RAME | 3 Continuous p02 errors. |  | Arrived on stie and verified unit down for p02 errors, constantly failing external calibrations for p02, all three levels fail at a value of 130-132. Ran several qc samples no change, ran side by side with another analyzer, the other analyzer rand everything fine, while the bad analyzer continued to run p02 values at 130-132. | D311FFD9404A8A62 | [5 empty] | 3 Continuous p02 errors. |  | Arrived on stie and verified unit down for p02 errors, constantly failing external calibrations for p02, all three levels fail at a value of 130-132. Ran several qc samples no change, ran side by side with another analyzer, the other analyzer rand everything fine, while the bad analyzer continued to run p02 values at 130-132.
No | 6 | Y | RAME | 0 Failed CAP survey on ctHb twice. |  | Bo contacted me and, advised me that this analyzer has failed on total hemoglobin twice on the CAP proficiency testing. | D311EB144041EFEC | Unknown | Performance | Oxi module | tHb | CAP Survey | 0 Failed CAP survey on ctHb twice. |  | Bo contacted me and, advised me that this analyzer has failed on total hemoglobin twice on the CAP proficiency testing.
No | 2 | Y | RAME | P Clot in the system |  | While doing training they introduce a clot $They tried backflush, but liquid is not moving freely. | D311C30B4043066C | Defect | HW | Basic Unit | Sensor Bay Unit | Clot | P Clot in the system |  | While doing training they introduce a clot $They tried backflush, but liquid is not moving freely.
No | 4 | Y | RAME | Z screen blank. |  | Checked for power and there was power coming into instrument. Power on and off still no help. | D3119FB840456E10 | Defect | HW | Basic Unit | CPU unit |  | Z screen blank. |  | Checked for power and there was power coming into instrument. Power on and off still no help.
No | Visit | Complaint | RDEU | QC-Fehler bei Level 1 und Level 2 (QC-Werte außer Bereich) | Fehlersuche --> Platinen (Sensorboard und Signal Data Board) überprüft --> Sensorboard und Signal Data Board tauschen, da bei der Überprüfung mit der Zeroing Prozedur dies auffällig war | Tausch des Sensorboard und des Signal Data Boards --> brachten keine Besserung --> Abzug des Kundengerätes ((308212) und aufstellen eines Reparaturüberbrückungsgerätes (306469) --> Reparaturüberbrückungsgerät i.O. (QC und Kal. i.O.) | SWO-227218 | [5 empty] | QC error at Level 1 and Level 2 (QC values out range) | Troubleshooting --> Check boards (sensor board and signal data board) --> Replace sensor board and signal data board, as this was noticeable during the verification with the zeroing procedure | Replacement of the sensor board and the Signal Data Board --> brought no improvement --> Deduction of the customer device ((308212) and set up a repair bridging device (306469) --> Repair bridging device in the way (QC and Kal. i.O.)
Yes | 6 | Y | RAME | P Possible fluid pathway blockage, reports parameter PO2 is red and did not finish system cycle |  | Getting error code 1206, possible fluid pathway blockage. customer, turn unit off and on two times and still having error, then had them replace solution pack. Still having the same problem. | D311CC40402A6BF3 | [5 empty] | P Possible fluid pathway blockage, reports parameter PO2 is red and did not finish system cycle |  | Getting error code 1206, possible fluid pathway blockage. customer, turn unit off and on two times and still having error, then had them replace solution pack. Still having the same problem.
No | Visit | Complaint | RKOR | Interface loading error | . | Data full error. Sample, activity log delete. Interface check => OK | SWO-227268 | [5 empty] | Interface loading error | . | Data full error. Sample, activity log delete. Interface check => OK
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | Na定标不过 | 更换测试卡后修复 | SWO-226992 | [5 empty] | Test card, reagent package failure | Na calibration is no more than | Repair after replacing test card
No | 2 | Y | RPOL | Blank | sprawdzenie analizatora, wymiana wê¿yków, odka¿enie analizatora, kalibracja tHb. Analizator sprawny technicznie. |  | D31F5DEE520EE614 | [5 empty] | Blank | Analyzer, replacement of the device, separation of the analyser, calibration of the tHb. Technically efficient analyzer.
No | Visit | Complaint | RCHN | GLU,Lac定标不过 | GLU,Lac定标不过，借903-183一个（被903-581替代） | 更换903-581后修复(By Ran) | SWO-226115 | Defect | HW | Basic Unit | PCB, Sensor interface |  | GLU, Lac calibration is not | GLU, Lac calibration, borrow edgy 903-183 (replaced by 903-581) | Repair after replacement 903-581 (By Ran)
No | Visit | Complaint | RCHN | 定标失败 1024，1035 | 定标失败 1024，1035 ，借903-364一个。 | 更换903-364后修复(By Ran) | SWO-226449 | Defect | HW | Process Module | Optical Unit | 1024 | Bid failure 1024,1035 | Bid failure 1024,1035 , borrow edith of 903-364. | Repair after replacement 903-364 (By Ran)
No | Visit | Complaint | RCHN | 报错1024 | 报错1024，借903-364一个 | 更换903-364后修复(By Ran) | SWO-225697 | Defect | HW | Process Module | Optical Unit | 1024 | 1024 error | 1024, loan 903-364 one | Repair after replacement 903-364 (By Ran)
No | Visit | Complaint | RIND | Oximetry parameters are showing problem. | Checked instrument and found that there was oximetry PCB problem. | Replaced Oxi PCB and given oximetry calibration for 2 times. Then performed 1 point and 2 point calibration. Unit is working ok. | SWO-227228 | Defect | HW | Oxi module | PCB, Oxi module |  | Oximetry parameters are showing problem. | Checked instrument and found that there was oximetry PCB problem. | Replaced Oxi PCB and given oximetry calibration for 2 times. Then performed 1 point and 2 point calibration. Unit is working ok.
No | Visit | Complaint | RIND | Printout not working | Checked and found printer head is defective. | Replaced new printer head from stock, 2 point calibration performed successfully. | SWO-227871 | Defect | HW | Printer Module | Printer head |  | Printout not working | Checked and found printer head is defective. | Replaced new printer head from stock, 2 point calibration performed successfully.
Yes | 4 | Y | RAME | V user called stating that they keep getting gas error 689 |  | gas off set voltage error on pressure transducer. | D311E8C1401CF25C | [5 empty] | V user called stating that they keep getting gas error 689 |  | gas off set voltage error on pressure transducer.
No | Visit | Complaint | RIND | Touch Hanging again and again. | USB Stick dose not work. | Try with with new USB same problem occures. Complete software upgrade done. Flound touch working well. Analyser Serviced and calibrated properly. Handed over in a fine working condition. | SWO-228143 | Defect | HW | DMS_PC | Touch screen | Firmware update | Touch Hanging again and again. | USB Stick dose not work. | Try with with new USB same problem occures. Complete software upgrade done. Flound touch working well. Analyser Serviced and calibrated properly. Handed over in a fine working condition.
No | 9 | Y | RAME | 2 pCO2 calibrations is drifting on one pt ph/bg cal |  | found that the analyzer keeps getting random drift on one pt ph $BG cal. | D312407E4034F5CE | Defect | HW | pH/BG Module | Valve, air |  | 2 pCO2 calibrations is drifting on one pt ph/bg cal |  | found that the analyzer keeps getting random drift on one pt ph $BG cal.
No | Visit | Complaint | RLTD | Neil Juland on 02476 968510 | Valve errors and issues with electrode seals. | Replaced valve and electrodes | SWO-227251 | Defect | HW | El/Met Module | Valve, air |  | Neil Juland on 02476 968510 | Valve errors and issues with electrode seals. | Replaced valve and electrodes
Yes | 12 | Y | RAME | 3 pCO2 QC results are running high. |  | pCO2 QC results are inside 2SD range but has been trending higher. The Ag $Cl reference is showing some deterioration. | D311B21C403D14CE | Re classified RMED | No RMED spec violation |  |  |  | 3 pCO2 QC results are running high. |  | pCO2 QC results are inside 2SD range but has been trending higher. The Ag $Cl reference is showing some deterioration.
No | Visit | Complaint | RDEU | LQC Level 2 und 3 werden bei TnT sehr häufig zu hoch gemessen | LOT Nummern überprüft: 944-515 LOT 09163, 944-514 LOT 06162 | Schlauchsystem gereinigt Handhabung mit Frau Uecker besprochen Lagerung überprüft-ok | SWO-227159 | Defect | Consumables | Test Kit | Cartridge, Lot | LQC | LQC Levels 2 and 3 are very often over-measured at TnT | LOT numbers checked: 944-515 LOT 09163, 944-514 LOT 06162 | Hose system cleaned Handling discussed with Mrs. Uecker Storage checked-ok
No | Visit | Complaint | RDEU | 30.08.2017, Herr Kamp, Tel: 06946086123: Probenabbrüche mit 1067 und 3553 | Gerät zeigt Fehler: 1067 und 3553. Blankmessung mit 5 Tests, Fehler tritt sporadisch auf, bei 3 Messung bricht das Gerät alle 3 übrigen Messungen ab. Gerät ausgebaut, Incubatorwheel, Prozesswheel, Storagewheel, ausgebaut alles gereinigt, Punch Out und andere Gewinde gefettet., Service Tests alle bestanden, mehrmals durchgeführt. 5 Blankmessungen noch mal gestartet, alle i.O, kein Probenabruch. Kunden informiert. | Fehler taucht nach paar Stunden wieder auf. Nach Rücksprache mit Michael Brand, E- Teile bestellt. Müssen aus CPH geschickt werden. Laut Kunde ist ein Leihgerät zwingend notwendig. | SWO-228286 | Revisit required | HW | Unknown | Unknown | 1067 | 30.08.2017, Mr. Kamp, Tel: 06946086123: Sample cancellations with 1067 and 3553 | Device shows errors: 1067 and 3553. Blank measurement with 5 tests, error occurs sporadically, with 3 measurement the device cancels all 3 remaining measurements. Device removed, Incubatorwheel, process wheel, storagewheel, removed everything cleaned, punch out and other threads greased, service tests all passed, performed several times. 5 blank measurements started again, all i.O, no sample break. informed customers. | Error reappears after a few hours. After consultation with Michael Brand, E- Parts ordered. Must be sent from CPH. According to the customer, a rental device is absolutely necessary.
No | Visit | Complaint | RDEU | 23.08.2017, 09:43 Uhr, Fr. Nowack, Tel. 02118604030: Kal Fehler bei pCO2 und Lac 23.08.2017, 10.55. Uhr, frau Nowack, tel: 02118604030: weiterhin Kal Fehler bei Lac. 28.0/8.2017, 9.20Uhr, Frau Nowak: Tel.: 02118604030: auch mit neuen Betriebsmittel-ständig Kalibrationsfehler | Lac Probleme. | Gerät durchgespült, Flieswege OK. Sensor Kassette und LP erneuert. Nach Kal und QC bleiben die Kal Probleme mit Lac weiterhin bestehen. | SWO-227610 | Assisted Replacement | SC | SC | Lac |  | 23.08.2017, 09:43, Fr. Nowack, Tel. 02118604030: Kal error in pCO2 and Lac 23.08.2017, 10.55. a.m., mrs. Nowack, tel: 02118604030: continue Kal error at Lac. 28.0/8.2017, 9.20a,mrs. Nowak: Tel.: 02118604030: also with new equipment-permanent calibration error | Problem lake. | Device rinsed, flow paths OK. Sensor cassette and LP renewed. According to Kal and QC, the Kal problems with Lac remain.
No | Visit | Complaint | RKK | Lac ステータスエラーでカセット交換したが、pHの安定が悪いので訪問 | 確認 | 訪問時には安定していたがキャル確認、キャルパック交換、QC確認 | SWO-227810 | Assisted Replacement | SC | SC | Lac |  | I replaced the cassette with lac status error, but i visited because the pH was not stable. | Check | It was stable at the time of the visit, but the car check, the cal pack replacement, and the QC confirmation
Yes | 3 | Y | RPAC | Blank | Pre-delivery checks |  | D3132E76B63C4D1F | Re classified RMED | OOB |  |  |  | Blank | Pre-delivery checks
No | Visit | Complaint | RCHN | 报错 4548 | 报错 4548，借903-028一个。 |  | SWO-227309 | Defect | HW | Cover and Chassis | PCB, Cabinet Top | 4548 | Error 4548 | 4548, borrow edith 903-028.
Yes | 68 | Y | RPAC | Blank | Analyser failing all pump calibrations. |  | D3133885B603C6E9 | [5 empty] | Blank | Analyser failing all pump calibrations.
No | Visit | Complaint | RCHN | 血氧模块检测硬件错误 | 血氧模块检测硬件错误，借902-833,902-458各一个 |  | SWO-227308 | Defect | HW | Oxi module | Oxi module, Replaced | 258 | Blood oxygen module detects hardware errors | Blood oxygen module detection hardware error, borrow 902-833, 902-458 each one
No | 11 | Y | RAME | Z Gas 2 indicating low after 34 days on analyzer. |  | No apparent leaks. Gas program in service programs checks OK. | D3131800B64B412C | Defect | HW | Gas module | Gas valve |  | Z Gas 2 indicating low after 34 days on analyzer. |  | No apparent leaks. Gas program in service programs checks OK.
No | Visit | Complaint | RDAN | Hanne klager over, at ABL'en stadigvæk drypper fra inlettet. | Afprøvet. Laver ikke fejlen nu men har helt tydeligt dryppet. | Probearm, studs f. inletpakning og inletpakning udskiftet, slutafprøvet. | SWO-227558 | Defect | HW | Inlet module | Inlet gasket |  | Hanne complains that the Abbey is still dripping from the inrelieved. | Tested. Do not make the mistake now but have quite clearly sprinkled. | Probearm, studs f. Inlet packaging and inlet pack replaced, final tested.
No | 2 | Y | RAME | z Out of Box failure from site: Repair completed by QA in Westlake. |  | DMS not functioning: Instrument won't initiate, screen goes blank, touch display won't work. | D317E41B8E3F051E | Defect | HW | DMS_PC | CPU unit |  | z Out of Box failure from site: Repair completed by QA in Westlake. |  | DMS not functioning: Instrument won't initiate, screen goes blank, touch display won't work.
No | Visit | Complaint | RDEU | 30.08.2017, Herr Kamp, Tel: 06946086123: Probenabbrüche mit 1067 und 3553 | Leihgerät aufstellen. | Leihgerät abgeholt, in CPU aufgestellt, Startphase, QCs eingescant, TNT und D-Dimer Level 2 Durchgeführt, alles i.O Gerät von CPU abgeholt, nach Reparatur wird das Gerät zurückgebracht. | SWO-228285 | Revisit required | Loaner installed | Loaner installed | Loaner installed | 1067 | 30.08.2017, Mr. Kamp, Tel: 06946086123: Sample cancellations with 1067 and 3553 | borrower. | Rental device picked up, set up in CPU, start phase, QCs scanned, TNT and D-Dimer Level 2 Performed, all i.O device is picked up from CPU, after repair the device is returned.
No | Visit | Complaint | RLTD | Lactate wont pass calibrations after replacing SC two days ago. Please call MRS DEBBIE CICKLES - 07867 663 201 | Site visit to investigate lactate with low sensitivity after two new sensor cassettes. | Arrived on site and attended analyser. Checked logs and found lactate was consistently low on sensitivity for lactate. Status value was good. Entered service programs and checked sensor values. Lactate looked low at 0.4 Cleaned sensor connector pins with alcohol, allowed to air dry and then closed the sensor bay again. Retested in service programs and sensor value for lactate now up towards 500. Exited service programs and allowed analyser to run 10 minute startup. After this I ran calibrations. Calibrations showed a small increase in sensitivity but only up to 18 from 10. Checked service programs again and lac reading were back down to the values they were at before. Opened sensor bay again to take another look and saw no bent pins or obvious dirt inside. Closed the bay and checked readings again. They were now up at 450. Checked the values being read on the second analyser on site and these were 0.38 to 0.4. Opened sensor bay and closed again before rechecked. Values read at 400 to 450. Suspect possible sensor cassette issue or interface pcb. Asked customer to install new sensor cassette from new lot whilst awaiting part to arrive. Ordered new interface pcb and will return to fit this when it arrives if required. | SWO-228063 | Assisted Replacement | SC | SC | Lac | Possibly later swo Sensor if PCB | Lactate wont pass calibrations after replacing SC two days ago. Please call MRS DEBBIE CICKLES - 07867 663 201 | Site visit to investigate lactate with low sensitivity after two new sensor cassettes. | Arrived on site and attended analyser. Checked logs and found lactate was consistently low on sensitivity for lactate. Status value was good. Entered service programs and checked sensor values. Lactate looked low at 0.4 Cleaned sensor connector pins with alcohol, allowed to air dry and then closed the sensor bay again. Retested in service programs and sensor value for lactate now up towards 500. Exited service programs and allowed analyser to run 10 minute startup. After this I ran calibrations. Calibrations showed a small increase in sensitivity but only up to 18 from 10. Checked service programs again and lac reading were back down to the values they were at before. Opened sensor bay again to take another look and saw no bent pins or obvious dirt inside. Closed the bay and checked readings again. They were now up at 450. Checked the values being read on the second analyser on site and these were 0.38 to 0.4. Opened sensor bay and closed again before rechecked. Values read at 400 to 450. Suspect possible sensor cassette issue or interface pcb. Asked customer to install new sensor cassette from new lot whilst awaiting part to arrive. Ordered new interface pcb and will return to fit this when it arrives if required.
No | Visit | Complaint | RDEU | 04.09.2017; Fr. Spätling; Tel.: 09123180121; Luftblasen erkannt 04.09.2017 10:15 Fr. Spätling 09123180121 : nach SC-Wechsel weiterhin Luftblasen im System; LP läuft zudem weiterhin aus. 04.09.2017 10:36 Fr. Spätling 09123180121 Fehler bleibt bestehen. Ein Vorortservice wird gewünscht. | Fehlersuche --> am Display des ABL80 ist zu sehen Unzureichende Zyklen im Lösungspack vorhanden --> LP erneuern --> Luftblasenerkennung schlägt wieder an --> Schläuche überprüfen --> oberer dicker Schlauch (welcher vom Manifold zum Frontteil geht) ist geknickt --> Schläuche erneuern | LP () aus Kundenbestand eingebaut, Schlauchsystem erneuert, System und Funktion i.O. | SWO-227996 | [5 empty] | 04.09.2017; Ms. Spätling; Tel.: 09123180121; Air bubbles detected 04.09.2017 10:15 Fr. Spätling 09123180121 : after SC change, air bubbles in the system continue; LP is also still running out. 09/04/2017 10:36 Fr. Spätling 09123180121 Error remains. A suburban service is desired. | Troubleshooting --> on the display of the ABL80 is visible Insufficient cycles present in the solution pack --> LP renew --> Air bubble detection strikes again --> Check hoses --> upper thick hose (which goes from manifold to front part) is bent --> hoses to renew | LP () built-in from customer base, hose system renewed, system and function in the end
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | PCO2,CL定标失败 | 更换测试卡 | SWO-227760 | [5 empty] | Test card, reagent package failure | PCO2, CL calibration failed | Replace the test card
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 测试卡参比电极错误 1361 | 更换测试卡正常。 | SWO-227887 | [5 empty] | Test card, reagent package failure | Test card reference electrode error 1361 | The replacement test card is OK.
No | 3 | Y | RPOL | Blank | Okazjonalne blokowanie siê ekranu dotykowego podcas odbia³czania. Diagnoza problemu, wymiana ekranu wraz z panelem dotykowym nie rozwi¹za³a problemu. Ponowna zmiana ekranu na ekran pocz¹tkowy. Konieczna dalsza diagnoza w celu wyeliminowania usterki. Analizator sprawny pomiarowo. |  | D31BFC8EF03D204C | Revisit required | HW | DMS_PC | Touch screen | Email | Blank | Occasional blocking of the touch screen of the subcas of the reconnection. Diagnosing the problem, replacing the screen with the touchpad does not resolve the problem. Change the screen to the initial screen again. Further diagnosis is needed to eliminate the defect. Analyser measuring.
No | Visit | Complaint | RIND | Calibration error | Checked the data logs found cl Calibration sensitivity out of range replaced new cl membrane & checked then also same problem Calibration sensitivity out of range crosschecked the cl electrode found working | Need to replace cl electrode | SWO-229778 | Assisted Replacement | Electrode | Electrode | Cl |  | Calibration error | Checked the data logs found cl Calibration sensitivity out of range replaced new cl membrane & checked then also same problem Calibration sensitivity out of range crosschecked the cl electrode found working | Need to replace cl electrode
No | Visit | Complaint | RDEU | 30.08.2017, Herr Kamp, Tel: 06946086123: Probenabbrüche mit 1067 und 3553 | Gerät ausgebaut, Cup Transport zeigt keine Fehler. Incubatorweehl und storageweehl neu kalibriert. Systemreinigung 4 mal durchgeführt. 5 Blankmessungen Tests alle OK. QC D-Dimer wieder Probenabruch. Rücksprache mit Michael Brand, E-Teile bestellt. | E-Teile bestellt. | SWO-232813 | Revisit required | Unknown | Unknown | off/on | 1067 | 30.08.2017, Mr. Kamp, Tel: 06946086123: Sample cancellations with 1067 and 3553 | Device removed, Cup transport shows no errors. Incubatorweehl and storageweehl recalibrated. System cleaning performed 4 times. 5 blank measurements tests all OK. QC D-Dimer again sample breakage. Consultation with Michael Brand, E-Parts ordered. | E-parts.
No | Visit | Complaint | RDEU | Calibrate 42 | Fehlersuche, Cup aus dem Inkubator Rad entfernt. | Startphase, QC. | SWO-227927 | Defect | HW | Incubator Module | Punch Up |  | Calibrate 42 | Troubleshooting, Cup removed from the incubator wheel. | Startphase, QC.
No | Visit | Complaint | RHUN | Tű újra elgörbült, megfelelő kezelés átbeszélése. | Tű újra elgörbült. | Megfeflő kezelés átbeszélése, tű cseréje újra. | SWO-228356 | Replaced | HW | Inlet module | Probe bend |  | Needle re-bent, discussing appropriate treatment. | The needle is bent again. | Relaunches, needle replacement.
No | 5 | Y | RFRA | Blank | MESURES ANNULEES / PAS D'ASPIRATION ECHANTILLON |  | D3184CD88E09D007 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | Blank | NEMIES / NOT ASPIRATION ECHANTILLON
No | Visit | Complaint | RCHN | 提示测试卡未连接 | 提示测试卡未连接，借924-738一个 |  | SWO-228209 | [5 empty] | Prompt the test card is not connected | Prompt test card is not connected, borrow 924-738 one
No | 11 | Y | RAME | 0 Thb low on blood samples |  | Customer compared to another unit and the THB is running lower. | D318172D8E44FF5A | [5 empty] | 0 Thb low on blood samples |  | Customer compared to another unit and the THB is running lower.
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | pco2定标失败和pco2灵敏度超出范围 | 更换测试卡后修复 | SWO-228361 | Assisted Replacement | SC | SC | pCO2 | Sensitivity | Test card, reagent package failure | pco2 calibration failure and pco2 sensitivity out of range | Repair after replacing test card
No | 17 | Y | RPOL | Blank | Aparat zalany, wymiana kruæców pod uszczelk¹ wlotu, uszczelki wlotu oraz pogiêtej ig³y. Testy funkcjonalne. Analizator sprawny. |  | D317ED808E37CCA1 | Defect | HW | Inlet module | Nipple kit |  | Blank | The engine floated, replacing the dealers under the gasket, inlet seals and the bent needle. Functional Tests. Functional Analyzer.
No | 14 | Y | RAME | Z Autocheck not running QC. |  | QC housing tilts up but the breaker does not move. In service everyhting works fine. | D31801D18E60670E | Defect | HW | Auto Inlet module | FPC Cable |  | Z Autocheck not running QC. |  | QC housing tilts up but the breaker does not move. In service everyhting works fine.
No | 4 | Y | RPNZ | Blank | Calcium results |  | D3229BDB3E008797 | Assisted Replacement | Electrode | Electrode | Ca |  | Blank | Calcium results
No | 4 | Y | RAME | 4 COHB ON CALIBRATION VERIFICATION HIGH OUT |  | RAN VK-R7 TO VERIFY | D31814248E3EE88A | Defect | HW | Oxi module | Lamp unit |  | 4 COHB ON CALIBRATION VERIFICATION HIGH OUT |  | RAN VK-R7 TO VERIFY
No | Visit | Complaint | RDEU | 04.09.2017, Frau Erfurt, tel: 08315301793: Kalibrationsfehler bei pCO2 und Cl 04.09.2017, Frau Erfurt, Tel: 08315301793: weiterhin Kalibrationfehler bei pCO2; Frau Erfurt hat bereits 2x neu membraniert. | Kalibrationsfehler bei pCO2 und Cl | pCO2 Elektrode überprüft, Elektrode defekt, pCO2 Elektrode erneuert und neu membraniert, Kalibrationen und QC's vorgenommen. | SWO-228538 | Assisted Replacement | Electrode | Electrode | pCO2 |  | 04.09.2017, Mrs. Erfurt, tel: 08315301793: Calibration error in pCO2 and Cl 04.09.2017, Mrs. Erfurt, Tel: 08315301793: continued calibration errors in pCO2; Ms. Erfurt has already re-logged 2times. | Calibration error for pCO2 and Cl | pCO2 electrode checked, electrode defective, pCO2 electrode renewed and re-membraneed, calibrations and QC's made.
No | 5 | Y | RCZE | Blank | nelze vložit data pøes klávesnici ani display, obrazovka neustále pøepíná do nápovìdy nebo do dalších oken, kurzor zamrzá a problikává na jednom místì, nelze ovládat, problém i se vstupem do servisu pøes klávesnici |  | D3181A778E0ED283 | Defect | HW | DMS_PC | DMS_PC |  | Blank | Cannot insert data via keyboard or display, the screen constantly switches to help or other windows, the cursor freezes and flashes in one place, can not be controlled, problem even with access to service via keyboard
No | 7 | Y | RAME | I Heat Failure Error. |  | Verified Heat Failure Error. | D3183DB08E499925 | [5 empty] | I Heat Failure Error. |  | Verified Heat Failure Error.
No | 5 | Y | RAME | z unit will not power on |  | checked power supply checked battery checked all fuses checked wall power replaced cpu - reloaded software | D3193BDB045C8048 | Defect | HW | DMS_PC | CPU unit |  | z unit will not power on |  | checked power supply checked battery checked all fuses checked wall power replaced cpu - reloaded software
Yes | 21 | Y | RAME | Z Came in to perform procedures for techline cases 783271241 and 788617906 suggested by Sendx. Customer not complaining. Just came to to do what Sendx suggest. |  | Sendx suggested running the cleaning procedure for the possible fluid blockage issue we are seeing. I ran the cleaning but also replaced the manifold unit. I also did the SW upgrade for the corrupt SC file issue after reboot. The 3.16 SW appears to address this. | D31941D804326FCD | [5 empty] | Z Came in to perform procedures for techline cases 783271241 and 788617906 suggested by Sendx. Customer not complaining. Just came to to do what Sendx suggest. |  | Sendx suggested running the cleaning procedure for the possible fluid blockage issue we are seeing. I ran the cleaning but also replaced the manifold unit. I also did the SW upgrade for the corrupt SC file issue after reboot. The 3.16 SW appears to address this.
No | 2 | Y | RCZE | Blank | reklamace vadný pack, po reinstalaci dalšího opìt teèe |  | D31A3D157A39669B | [5 empty] | Blank | Defective pack complaints, after reinstalling the next
No | Visit | Complaint | RDEU | Nach Eingabe der Netzwerkeingaben/ -Änderungen werden diese nach dem Bestätigen nicht gespeichert. | Einstellungen geprüft | Die Netzwerkeinstellungen werden nach Bestätigung der Änderung nicht gespeichtert--> SW V3.16 inst.--> Neustart --> Einstellungen erneut eingegeben --> Bestätigen der Änderungen --> ABL benötigt ein paar Sekunden dafür startet den ABL aber nicht neu --> Aufruf der Konfiguration ( Netzwerkeinstellungen ) -->alle Eingaben wurden nicht gespeichert --> Netzwerkeinstellungen stehen wieder auf DHCP --> keine Anbindung an Radiance $ Netzwerk möglich Auf dem ABL isst Windows 7 inst. | SWO-228867 | [5 empty] | After entering the network inputs/changes, these are not saved after confirmation. | Settings checked | The network settings are not speculated after confirmation of the change--> SW V3.16 inst.--> Restart --> Settings are re-entered --> Confirming the changes --> ABL takes a few seconds to restart the ABL but not again --> Calling the Configuration ( Network Settings ) -->all inputs have not been saved --> Network settings are back on DHCP --> no connection to Radiance - Network possible On the ABL, Windows 7 eats inst.
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 质控C不过 | 更换试剂包正常。 | SWO-228701 | [5 empty] | Test card, reagent package failure | Quality control C but | The replacement kit is OK.
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | PCO2定标失败 | 更换测试卡 | SWO-228719 | Assisted Replacement | SC | SC | pCO2 | Cal | Test card, reagent package failure | PCO2 calibration failure | Replace the test card
No | Visit | Complaint | RCHN | 光学单元超范围 | 光学单元超范围，借903-364一个 | 更换903-364后修复 by jesson | SWO-227855 | Defect | HW | Process Module | Optical Unit | 1024 | Optical unit over range | Optical unit over range, borrowed 903-364 one | Repair by jesson after replacing 903-364
No | Visit | Complaint | RCHN | 报错1024 | 报错1024，借903-364,944-212（批号13147），944-216（批号13072），944-214（批号13002缺货），944-258（批号13224）各一个 | 更换903-364和重新定标后修复 by jesson | SWO-228330 | Defect | HW | Process Module | Optical Unit | 1024 | 1024 error | 1024, borrow 903-364,944-212 (lot number 13147), 944-216 (lot number 13072), 944-214 (lot number 13002 out of stock), 944-258 (lot number 13224) each one | Fix by jesson after replacement 903-364 and re-calibration
No | 7 | Y | RAME | 1 pH and Ca sensitivity error |  | checked data logs for errors and replacement, checked valves, checked reference membrane and electrode | D319CE55044368EA | Assisted Replacement | Electrode | Membrane | Reference |  | 1 pH and Ca sensitivity error |  | checked data logs for errors and replacement, checked valves, checked reference membrane and electrode
No | Visit | Complaint | RKK | 初期化が通らず、トラブルシュート画面になってしまう。 | サービス画面では特に異常がなかった。 | 流路洗浄及び、ｲﾝﾚｯﾄﾓｼﾞｭｰﾙ周りの調整を実施。 | SWO-234677 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway | & clean inlet | The initialization does not pass, and it becomes a trouble shoot screen. | There was no particular abnormality in the service screen. | Flow path cleaning and adjustment around the inlet module.
No | Visit | Complaint | RSCH | Probleme mit PH 1070: Sensor Ansprachefehler | 1070: Sensor Ansprachefehler (PH) | Allgemeine Kontrolle ok. Kunde bereits 2 Sensorkassetten ers., gleicher Fehler. Kontktstift gebogen, Gerät nach Thalwil. Gerät zerlegt, alle Kontaktstifte ausgebaut und gereinigt. 1 Stift ersetzt (krumm). Alle Werte ausser Glu i.o. (Reaktionfehler). Kontakstifte ausgebaut, Glu und Lac Kontaktstifte ers. Lösungspack und Sensorkassette nochmals installiert ok. Kal. + QC + Thb i.o. *** Gerät retour an Kunde, Inbetriebnahme i.o. *** | SWO-229356 | Defect | HW | Basic Unit | SC Contacts | MRM + HP offline | Problems with PH 1070: Sensor address error | 1070: Sensor Ansprachefehler (PH) | General control ok. Customer already ers., same error. Kontktstift bent, device to Thalwil. device, removes and cleans all contact pins. 1 pin replaced (crooked). All values in other than Glu i.o. (reaction error). Contact ingesis, Glu and Lac contact pins. Solution pack and sensor cartridge reinstalled ok. Kal. + QC + Thb i.o. Device return to customer, commissioning i.o. ***
No | Visit | Complaint | RDEU | 07.09.2017; Hr. Schnitzler; Tel.: 0815823682; pCO2 Kal Fehler | pCO2 Kal Fehler | Sensor Kassette (von Klinik Füssen) erneuert, Kalibrationen und QC's vorgenommen, Glu und Lac besteht Kal und QC's nicht, Mehrere Kalibrationen und QC's vorgenommen, BGA ok. | SWO-229376 | Assisted Replacement | SC | SC | pCO2 | Cal | 07.09.2017; Mr. Schnitzler; Tel.: 0815823682; pCO2 Kal Error | pCO2 Kal Error | Sensor cassette (renewed by clinic feet), calibrations and QC's made, Glu and Lac does not consist of Kal and QC's, Several calibrations and QC's made, BGA ok.
No | Visit | Complaint | RSCH | Rep | Probenaspiration scheitert 1235 FTC 1189 | Einlassdichtungshalter ausgetauscht (klemmt) Softwareupgrade V3.3MR1 ausgeführt Präanalytik mit Frau Ganz kurz besprochen Funktionskontrolle i.o | SWO-229156 | User error | Pre analytical | Sample handling | Sample handling |  | Rep | Probenaspiration scheitert 1235 FTC 1189 | Inlet seal holder replaced (clamped) Software upgrade V3.3MR1 running Pre-analysis with Mrs. Very briefly discussed Functional control i.o
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | Lac定标不过 0375 pO2低于质控范围的质控值 0090 | 连续更换2张测试卡和1个试剂包后修复 | SWO-229437 | Assisted Replacement | SC | SC | Lac | Cal | Test card, reagent package failure | Lac calibration is no more than 0375 pO2 below THE QC VALUE OF THE QC RANGE 0090 | Repair after 2 test cards and 1 reagent kit in a row
No | Visit | Complaint | RCHN | 气流温度超范围 | 气流温度超范围，借903-210一个 | 更换903-210后修复by jesson | SWO-228261 | Defect | Consumables | SP | Cup Stack | 1067 | Airflow temperature over range | Airflow temperature over range, borrow 903-210 one | Repair by jesson after replacement 903-210
No | 10 | Y | RAME | 2 User called stating that analyzer keeps drifting on calibrations for status |  | user keeps randomly getting error for drift on pco2 for status | D31AB4A17A1D4561 | Assisted Replacement | Electrode | Electrode | pCO2 |  | 2 User called stating that analyzer keeps drifting on calibrations for status |  | user keeps randomly getting error for drift on pco2 for status
No | Visit | Complaint | RCHN | PH电极膜玻璃破损 | PH电极膜玻璃破损，借945-614一个 | 更换945-614后修复 by jesson | SWO-228582 | Assisted Replacement | Electrode | Electrode | pH | Broken | DAMAGED PH ELECTRODE MEMBRANE GLASS | PH electrode membrane glass broken, borrowed 945-614 one | Repair by jesson after replacing 945-614
No | 19 | Y | RAME | V Oxi Liquid Sensor |  | numerous error 537 | D31ADFF87A16D6FF | Defect | HW | Oxi module | LS | 537 | V Oxygen Liquid Sensor |  | numerous error 537
No | 4 | Y | RAME | F Customer reported by email that analyzer was getting oximetry errors on all levels of QC. |  | Checked data logs for errors. Found 581 errors on all levels of QC. Ran complete oximetry tests with pass results. tHB drifting intermittently on calibrations. | D31A9FEB7A34E7C8 | Adjustment | HW | Oxi module | Lamp unit | Halogen | F Customer reported by email that analyzer was getting oximetry errors on all levels of QC. |  | Checked data logs for errors. Found 581 errors on all levels of QC. Ran complete oximetry tests with pass results. tHB drifting intermittently on calibrations.
No | 7 | Y | RAME | 2 CUSTOMER REQUESTED SERVICE AFTER PCO2 AND CA++ SHOWED OUTSIDE OF PEER GROUP REPORT RECEIVED FROM RADIOMETER |  | FOUND SIGNS OF COMTAMINATION IN MEASUREMENT CHAMBERS. HEAVY SIGNS OF BLOOD IN THE CA++ WELL AND ON THE BODY OF THE ELECTRODE | D31DA9AF66301307 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | 2 CUSTOMER REQUESTED SERVICE AFTER PCO2 AND CA++ SHOWED OUTSIDE OF PEER GROUP REPORT RECEIVED FROM RADIOMETER |  | FOUND SIGNS OF COMTAMINATION IN MEASUREMENT CHAMBERS. HEAVY SIGNS OF BLOOD IN THE CA++ WELL AND ON THE BODY OF THE ELECTRODE
No | Visit | Complaint | RDEU | 06.09.2017; Hr. Kolbe; Tel.: 082140065694; Neon lamp intensity too low, Neon number too low. | 0258: Neon lamp intensity too low 0316: Neon number too low | Oxi PCB erneuert - Fehler weiterhin vorhanden. Neue Oxi PCB wieder ausgebaut. Oxi Modul ausgetauscht. Modulsoftware angeglichen. Lampenabgleich durchgeführt. Oxi Complete Test ok. Service tHb Kalibration durchgeführt - ok. tHb Kalibration (Fcuv = 1,0032), Kalibrationen und QC's vorgenommen. | SWO-229209 | Defect | HW | Oxi module | Oxi module, Replaced | Neon Lamp | 06.09.2017; Hr. Kolbe; Tel.: 082140065694; Neon lamp intensity too low, Neon number too low. | 0258: Neon lamp intensity too low 0316: Neon number too low | Oxi PCB renewed - error still present. New Oxi PCBs expanded again. Oxi module replaced. module software. lamp leveling. Oxi Complete Test ok. Service tHb calibration carried out - ok. tHb calibration (Fcuv = 1.0032), calibrations and QC's.
No | Visit | Complaint | RDEU | 23.08.2017, 09:43 Uhr, Fr. Nowack, Tel. 02118604030: Kal Fehler bei pCO2 und Lac 23.08.2017, 10.55. Uhr, frau Nowack, tel: 02118604030: weiterhin Kal Fehler bei Lac. 28.0/8.2017, 9.20Uhr, Frau Nowak: Tel.: 02118604030: auch mit neuen Betriebsmittel-ständig Kalibrationsfehler | Sensor Interface PCB defekt. | Gerät durchgespült, Au Board und Sensor Interface PCB erneuert. Fehler bleibt bestehen, Jw Kit verbaut, OXI Ventil demontiert und gereinigt, Kundeneigene Ersatzteile: AU Board und Sensor Interface PCB wieder zurück getauscht. Einlasssonde erneuert, Gerät in der RDEU Zentrale gebracht. Modifizierte PCB wird über die TS Werkstatt bestellt und eingebaut. Leihgerät aufbereitet und aufgestellt. Beim Kunden eine Einweisung durchgeführt. | SWO-229673 | Revisit required | HW | Basic Unit | PCB, Sensor interface |  | 23.08.2017, 09:43, Fr. Nowack, Tel. 02118604030: Kal error in pCO2 and Lac 23.08.2017, 10.55. a.m., mrs. Nowack, tel: 02118604030: continue Kal error at Lac. 28.0/8.2017, 9.20a,mrs. Nowak: Tel.: 02118604030: also with new equipment-permanent calibration error | Sensor Interface PCB defekt. | Device flushed through, Au Board and Sensor Interface PCB renewed. Errors remain, Jw Kit installed, OXI valve dismantled and cleaned, customer-owned spare parts: AU board and sensor interface PCB replaced again. Inlet probe renewed, device brought to the RDEU headquarters. Modified PCBs are ordered and installed via the TS workshop. Rental device is prepared and set up. A briefing was carried out at the customer's site.
No | Visit | Complaint | RBNL | Afgebroken bloedmonsters | Beste, Bij deze stuur ik je de gegevens die wij uit de ABL90 hebben getrokken. De afgelopen twee maanden hebben wij gezien dat de ABL KICrood 1700 alalyses heeft verricht waarvan 104 met error 0331 (geen monster) = 6% fout Op dezelfde afdeling heeft de ABL KICgroen 1800 analyses verricht, waarvan 9 fout = 0.5% Tijdens het vergelijk van de twee ABLen op de KIC, van de afgelopen twee dagen, met hetzelfde bloedmonster, heeft de ABL kic rood 4 keer een error 0331 gegeven en de KIC groen geen fout. Twee verpleegkundigen die error 0331 hadden heb ik gesproken. Zij gaven beide aan dat het bloedmonster werd aangeboden, maar dat de ABL geen bloed opzoog, terwijl de naald in de spuit zat. Het beeldscherm bleef “blank”. Na het verwijderen van de spuit ging de ABL spoelen en werd als fout 0331 gegeven. Zelfde verschijnsel zoals ik in de mail van 5-9-17 heb aangegeven. De gebruikers geven in het algemeen aan dat zij mogelijk te snel handelen / het monster te snel aanbieden. De optie dat het heparine kussentje de opzuigopening afsluit sluit jij uit als optie, omdat wij alleen pico spuiten gebruiken en dat bij deze spuit een dergelijke fout niet kan voorkomen. Wij hebben afgesproken dat je maandagochtend de ABL90 KICrood opnieuw gaat afstellen, alle slangen vervangt en een nieuwe monsterunit plaatst. Ik zal zorgen dat de ABLrood op het lab staat om deze te repareren. Je kunt naar het algemene lab / tegenover mijn kamer, op Y1.119 komen of even telefonisch contact met mij opnemen als je bij de fontein in het UMCG staat, dan kom ik je bij de fontein ophalen. met vriendelijke groet, Jeroen J.J. van Leeuwen Staf Medewerker / Point Of Care Coördinator Laboratoriumgeneeskunde Universitair Medisch Centrum Groningen (UMCG) Hanzeplein 1 9713 GZ Groningen Kamer Y1.119 Postadres: J.J.J. van Leeuwen, code EA 20 Postbus 30001 9700 RB Groningen | Inletmodule $Inletconnector heb vervangen. Wetsection met alcohol daarna demiwater heb doorgespoten, hierna diverse metingen heb uitgevoerd, OK. Opmerking: ABL90 geeft QC responsfout voor de glucose en daardoor is nog niet vrijgegeven. Klant heb geadviseerd om indien na twee uur nog steeds responsfout aanwezig is, de SC te vervangen. | SWO-229880 | Revisit required | HW | Auto Inlet module | PCB, Auto Inlet | Mail to luc | Broken down blood samples | Best With this I send you the data that we have pulled from the ABL90. Over the past two months we have seen the ABL KICrood 1700 alalyses of which 104 with error 0331 (no sample) = 6% error In the same department, the ABL KICgroen has carried out 1800 analyses, of which 9 error = 0.5% During the comparison of the two ABLen on the KIC, of the last two days, with the same blood sample, the ABL KIC Red 4 times an error 0331 given and the KIC Green no error. Two nurses who had error 0331 I talked. They both indicated that the blood sample was offered, but that the ABL did not suckler blood, while the needle was in the syringe. The display remained "blank". After removing the syringe, the ABL coils went and was given as error 0331. Same phenomenon as I indicated in the Mail of 5-9-17. Users generally indicate that they may be acting too fast/offering the sample too fast. The option that the heparin cushion closes the suction opening excludes you as an option, because we only use Pico syringes and that this syringe cannot prevent such an error. We agreed that you will reset the ABL90 KICrood on Monday morning, replace all hoses and place a new monster unit. I will make sure that the ABLrood is on the lab to repair it. You can go to the general lab/in front of my room, on Y 1.119 or contact me by phone if you are at the fountain in the UMCG, then I come to pick you up at the fountain. Sincerely Vijay Kumar Verma Staff member/Point Of Care Coordinator Laboratory Medicine University Medical Center Groningen (UMCG) Hanzeplein 1 9713 GZ Groningen Room Y 1.119 Postal address: J.J.J. van Leeuwen, code EA 20 P.O. Box 30001 9700 RB Groningen | Inletmodule $Inletconnector have replaced. Wetsection with alcohol then demi water sprayed, after which several measurements have been made, OK. Note: ABL90 gives QC response error for glucose and therefore has not been released yet. Customer have advised to replace the SC if after two hours of response error is still present.
No | 3 | Y | RPOL | Blank | sprawdzenie analizatora, wymiana kasety, kalibracje parametrów, kontrola jakoœci skalibrowanych parametrów. Analizator sprawny technicznie |  | D31F5B245224E4B9 | Assisted Replacement | SC | SC | pH | Cal | Blank | Checking of the analyzer, replacing the cassette, calibrating the parameters, checking the quality of the calibrated parameters. Technically efficient analyzer
No | 4 | Y | RPOL | Blank | sprawdzenie analizatora, sprawdzenie po³¹czeñ do modu³u. Wymiana kasety sensorowej. Kalibracja parametrów oraz zrobienie kontrolek jakoœci. |  | D31BF820F0174888 | Assisted Replacement | SC | SC | pH | Cal | Blank | Inspect the analyzer, check the connection to the module. Replacing the sensor cassette. Calibrating parameters and making quality controls.
Yes | 3 | Y | RPAC | Blank | Qualicheck adaptor |  | D3237B4BB4136A27 | Re classified RMED | Not upSWO |  |  |  | Blank | Qualicheck adaptor
No | Visit | Complaint | RCHN | 泵轮老化 | 泵轮老化，需要借件902-971一个 | 更换902-971一个正常。by jesson | SWO-229157 | [5 empty] | Pump wheel aging | Pump wheel aging, need to borrow 902-971 one | Replace the 902-971 one normal. by jesson
No | Visit | Complaint | RCHN | 样本夹损坏 | 样本夹故障，借备件903-584 4个 | 更换903-584 4个正常。by jesson | SWO-228680 | Defect | HW | Aspiration Module | Tube holder |  | Damaged sample clip | Sample clip failure, spare parts 903-584 4 | Replace 903-584 4 normal. by jesson
No | Visit | Complaint | RCHN | Lac敏感度偏低 | Lac敏感度偏低 | 更换电极膜后修复 | SWO-229584 | Assisted Replacement | Electrode | Membrane | Lac | Sensitivity | Lac sensitivity is low | Lac sensitivity is low | Repair after replacement of electrode film
No | 12 | Y | RAME | 9 Customer called with LAC unstable errors |  | Checked logs to verify performance and found several patient results with Lac unstable. Lac sensitivity drop from 590s to 340s. | D31BCA42F045C6C9 | Cleaned | Fluidic pathway | Fluidic pathway | Measuring chamber |  | 9 Customer called with LAC unstable errors |  | Checked logs to verify performance and found several patient results with Lac unstable. Lac sensitivity drop from 590s to 340s.
Yes | 7 | Y | RAME | Blank |  |  | D31BD078F0504D3A | Re classified RMED |  |  |  |  | Blank
Yes | 7 | Y | RAME | Opened by mistake. |  |  | D31BD078F0504D3A | Re classified RMED |  |  |  |  | Opened by mistake.
No | 2 | Y | RCZE | Blank |  |  | D31BAE45F060BC70 | Defect | HW | Oxi module | Lamp unit | Neon Lamp | Blank
No | Visit | Complaint | RDEU | 06.09.2017; Fr. Scherb; Tel.: 03905481913; QC Barcodes können nicht gescannt werden 06.09.2017 Frau Scherb 03905481913 QC Barcodes können nicht eingelesen werden. 06.09.2017; 12:10 Uhr; Fr. Scherb; Tel.: QC Barcodes können nicht gescannt werden | Fehlersuche, Bestandsaufnahme QK | auf zuständige Laborantin des Hauses gewartet, Einlesen des QK-Beipackzettel zeigen lassen und die Fehler erklärt, fehlende QK's eingelesen | SWO-230058 | No fault found | No fault found | No fault found | No fault found |  | 06.09.2017; Ms. Scherb; Tel.: 03905481913; QC barcodes cannot be scanned 06.09.2017 Ms Scherb 03905481913 QC Barcodes cannot be read in. 06.09.2017; 12:10 p.m.; Ms. Scherb; Tel.: QC barcodes cannot be scanned | Troubleshooting, stocktaking QC | waited for the responsible laboratory technician of the house, have the QK package leaflet shown and the errors explained, missing QC read it
No | Visit | Complaint | RCHN | 无法阅读测试卡智能芯片 | 无法阅读测试卡智能芯片，借902-972,903-334各一个 |  | SWO-229819 | [5 empty] | Can't read test card smart chip | Unable to read the test card smart chip, borrow 902-972, 903-334 each one
No | Visit | Complaint | RCHN | THB定标失败 | THB定标失败，借902-980一个 | 更换902-980后修复 by jesson | SWO-228976 | Defect | HW | Oxi module | Hemolyzer | 378 | THB calibration failure | THB calibration failure, borrow 902-980 one | Repair by jesson after replacement 902-980
No | Visit | Complaint | RCHN | TNI定标不出结果 | TNI定标不出结果，借备件903-364一个 | 更换903-364后修复 by jesson | SWO-229484 | Defect | HW | Process Module | Optical Unit | 1031 | TNI calibration results | TNI bid does not produce results, borrowed spare parts 903-364 one | Repair by jesson after replacing 903-364
No | Visit | Complaint | RDEU | Sporadisch Oximetriefehler 581 | Sporadisch Oximetriefehler 581 | Gerät gespült, div. Probemessungen, OK | SWO-230293 | Cleaned | Fluidic pathway | Oxi module | Fluidic pathway |  | Sporadisch Oximetriefehler 581 | Sporadisch Oximetriefehler 581 | Device rinsed, various sample measurements, OK
No | Visit | Complaint | RIND | SC not installing. | SC not installaing. | SC found burn on back and at same time Analyser Sensor Board also got burnt. Both Scs got damage. Need to replace Sensor board. Demo Analyser Installed. TO sendind for Sensor board. | SWO-230244 | [5 empty] | SC not installing. | SC not installaing. | SC found burn on back and at same time Analyser Sensor Board also got burnt. Both Scs got damage. Need to replace Sensor board. Demo Analyser Installed. TO sendind for Sensor board.
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | Na,PH定标不过 LAC灵敏度超范围 | 连续更换2张测试卡和1个试剂包后修复 | SWO-229945 | Assisted Replacement | SC | SC | Lac | Response | Test card, reagent package failure | Na, PH calibration is no more than LAC sensitivity over range | Repair after 2 test cards and 1 reagent kit in a row
No | Visit | Complaint | RCHN | 样本浑浊度过高 | 样本浑浊度过高，借902-980,903-197各一个 |  | SWO-229901 | Defect | HW | Oxi module | Hemolyzer | 0292 | Sample turbidity over high | Sampleturly, borrow 902-980, 903-197 each
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | ca na k ph cl定标失败 | 更换测试卡后修复 | SWO-230328 | Assisted Replacement | SC | SC | Reference |  | Test card, reagent package failure | ca na k ph cl定标失败 | Repair after replacing test card
No | Visit | Complaint | RCHN | 液体传输错误 | 液体传输错误，借902-918一个 | 备件未使用，液体传输问题已解决，现机器出现定标失败，故需重新借件，此单先关单 by Andrew | SWO-229840 | No fault found | No fault found | No fault found | No fault found | Liquid sensor?? | Liquid transfer error | Liquid transfer error, borrow 902-918 one | Spare parts are not used, liquid transfer problems have been resolved, the machine now has a calibration failure, so need to re-borrow, this single first close by Andrew
No | Visit | Complaint | RCHN | 光学强度超范围 | 光学强度超范围，借903-364一个 |  | SWO-229935 | Defect | HW | Process Module | Optical Unit | 1024 | Optical strength over range | Optical strength over range, borrowed 903-364 a
No | 1 | Y | RFRA | Blank | Er 9612 Patients data base pleine |  | D31FC485524D8082 | [5 empty] | Blank | Er 9612 Patients full data base
No | Visit | Complaint | RDAN | Skærm fryser | NA | Hussein har fået sw til at løse problemet | SWO-230090 | Defect | HW | DMS_PC | Touch screen | Firmware update | Screen freezes | In | Hussein has got SW to solve the problem
No | 5 | Y | RAME | W Customer reported analyzer will not start-up. Software errors. |  | Start-up reporting error code status 9. Error in .ini file. Software at version 6.16 needs upgraded to 6.17 | D31FF067521632FC | Defect | SW | SWU | Roma.ini |  | W Customer reported analyzer will not start-up. Software errors. |  | Start-up reporting error code status 9. Error in .ini file. Software at version 6.16 needs upgraded to 6.17
No | Visit | Complaint | RDEU | Kunde meldet einen einen immer wiederkehrenden Pumpen Asymetrie Fehler. | Air Membranventil im pH/BG Modul verunreinigt. | Air Membranventil im pH $ BG Modul gereinigt und deren Flat Membran erneuert, Gerät kalibriert und QC durchgeführt. | SWO-230304 | Defect | HW | pH/BG Module | Valve, air |  | Customer reports a recurring pump asymetry error. | Air diaphragm valve contaminated in pH/BG module. | Air diaphragm valve cleaned in pH BG module and renewed its flat membrane, device calibrated and QC performed.
No | 6 | Y | RAME | F Customer called with 315 (Neon number too high) |  | Checked all logs to verify error code. Error 315 several times over the last couple of days. | D31FC51652338ACB | Defect | HW | Oxi module | Oxi module, Replaced | Neon Lamp | F Customer called with 315 (Neon number too high) |  | Checked all logs to verify error code. Error 315 several times over the last couple of days.
No | Visit | Complaint | RDEU | 11.09.2017, Herr Gerling, Tel: 051218941850: diverse LS Fehler; Herr Gerling berichtet, dass er zuerst die verbogene Nadel getauscht hat, dann einen Oxi LS Fehler hatte. Er hat versucht mit einer Spritze den Lufteinlass zu spülen. | LS Fehler | LS Sensoren und Schlauchsystem geprüft und gereinigt, LS Justage, Funktionskontrolle ok | SWO-230432 | Cleaned | Fluidic pathway | Fluidic pathway | LS, Oxi module |  | 11.09.2017, Mr. Gerling, Tel: 051218941850: various LS errors; Mr Gerling says he first swapped the bent needle, then had an Oxi LS error. He tried to flush the air inlet with a syringe. | LS Error | LS sensors and hose system tested and cleaned, LS adjustment, function control ok
No | Visit | Complaint | RKK | インレットより液漏れ | インレットより液漏れ | インレットプローブの曲がりが原因。 インレットプローブ交換。 | SWO-230322 | Defect | HW | Inlet module | Probe bend |  | Leakage from inlet | Leakage from inlet | Caused by the bend of the inlet probe. Inlet probe replacement.
No | Visit | Complaint | RDAN | Martin melder, at den drypper fra inlet og erutæt ved pH. | Afprøvet. | Inlet probe, probeholder, pakning, pakring, slange og spild pumpeslanger udskiftet, slutafprøvet. | SWO-230248 | Defect | HW | Inlet module | Inlet gasket | Same hosp | Martin reports, that it drips from the inlet and erutæt at pH. | Tested. | Inlet probe, Procontainer, packing, packing ring, hose and waste pump hoses replaced, final tested.
No | Visit | Complaint | RDAN | Martin melder, at ABL'en drypper fra inlet og er utæt ved pH elektroden samt har pixelfejl. | Afprøvet/set. | Skærm, inletprobe, pakning, probeholder, inletslange, pakring og spild pumpeslanger udskiftet, slutafprøvet. | SWO-230238 | Defect | HW | Inlet module | Inlet gasket | Same hosp | Martin reports that the ABL is dripping from the inlet and is outtied by the pH electrode and has pixelfejl. | Tested/seen. | Display, inletprobe, packing, Procontainer, inlethose, packing and waste pump hoses replaced, final tested.
No | Visit | Complaint | RDAN | Martin melder, at den drypper fra inlet og er utæt ved pH | Afprøvet. | Inlet probe, probeholder, inlet slange, inlet pakning, pakring, spild pumpeslanger udskiftet, parpirholder justeret, slutafprøvet. | SWO-230242 | Defect | HW | Inlet module | Inlet gasket | Same hosp | Martin reports that it drips from inlet and is leaking at pH | Tested. | Inlet probe, Procontainer, inlet hose, inlet packing, packing ring, spillage pump hoses replaced, pair-holder adjusted, final tested.
No | Visit | Complaint | RDEU | 27.07.2017; Sr. Anett; Tel.: 03381791614; CRP QC Messfehler 28.07.2017 10:30 Sr. Nicole; Tel.: 03381791612; weiterhin CRP QC Messfehler 31.07.2017: Rücksprache mit KerZi - neues QC wird bestellt, telefonische Info an Kunden 03.08.2017, 11:36 Uhr ==> Anrufer Pfl. Sebastian, Tel.: 03381791611 ==> LQC-Messfehler CRP 04.08.2017, 12:04 Uhr ==> Anrufer Schw. Anett, Tel.: 03381791614 ==> LQC-Fehler CRP besteht trotz neuer Charge weiterhin | Auslieferung und Inbetriebnahme beim Kunden nach Werkstattreparatur durch AngAn | O.g. Arbeiten, QK ok, Aqureverbindung ok | SWO-230162 | Defect | HW | Process Module | Optical Unit |  | 27.07.2017; Sr. Anett; Tel.: 03381791614; CRP QC Measurement Error 07/28/2017 10:30 Sr. Nicole; Tel.: 03381791612; further CRP QC measurement error 31.07.2017: Consultation with KerZi - new QC is ordered, telephone information to customers 03.08.2017, 11:36 a.m. ==> Caller Pfl. Sebastian, Tel.: 0338179111 ==> LQC measurement error CRP 04.08.2017, 12:04 o'clock ==> Caller Schw. Anett, tel.: 03381791614 ==> LQC error CRP persists despite new batch | Delivery and commissioning at the customer after workshop repair by AngAn | O.g. work, QK ok, aqure connection ok
No | 5 | Y | RAME | I Heater error, recalibration errors. |  | Verified error. When connected sensor 54-61745 on upper arm for performance test heater error came up. Used known good sensor 54 for testing no error occurred. | D31FD4F8522FF6C9 | [5 empty] | I Heater error, recalibration errors. |  | Verified error. When connected sensor 54-61745 on upper arm for performance test heater error came up. Used known good sensor 54 for testing no error occurred.
No | Visit | Complaint | RDEU | Datenabfrage nach dem Einscannen eines Patienten Barcodes funktioniert nicht. | Anwender verwenden eine Barcode ID, diese nicht kompatibel ist zur zum Laborsystem. | Patienten Barcode eingespannt. Barcode wird korrekt interpretiert, jedoch wird laut Hr Depta der Barcode nicht korrekt wieder gegeben, weil die Laborsoftware nur Befunde verarbeitet deren Patienten ID mit 014 beginnt. Patienten Barcode von einer anderen Station geholt und zum Test eingespannt. Die Patientendaten werden anschließend korrekt dargestellt. Die Funktion ist einwandfrei gegeben. | SWO-230770 | Setup | Ext.SW | His/lis | His/lis |  | Data retrieval after scanning a patient's barcode does not work. | Users use a barcode ID, which is not compatible with the laboratory system. | Patient barcode clamped. Barcode is interpreted correctly, but according to Hr Depta, the barcode is not returned correctly because the laboratory software only processes findings whose patient ID starts with 014. Patient barcode is fetched from another ward and clamped for testing. The patient data is then displayed correctly. The function is given correctly.
No | 35 | Y | RAME | G Numerous Replace Solution Pack messages. | Changed the complaint visit tab from N to Y. | Pack installed on 8 $23 $2017. During the time the pack was installed, numerous Leakage Current Detected error messages have occurred, but the pack was never replaced, only removed and reinstalled. Performed sensor simulator test to test system. Found no faults. | D320CBD6C816F6AA | Assisted Replacement | SP | SP | Leak current | MRM Mail, request Service Dump | G Numerous Replace Solution Pack messages. | Changed the complaint visit tab from N to Y. | Pack installed on 8 $23 $2017. During the time the pack was installed, numerous Leakage Current Detected error messages have occurred, but the pack was never replaced, only removed and reinstalled. Performed sensor simulator test to test system. Found no faults.
No | Visit | Complaint | RBNL | Lekkage geconstateerd bij de OXI-valve | Er is lekkage te zien bij de OXI-valve. In het servicemenu pressure test uitgevoerd, deze geeft bij 'Vacuum Value' = 0,79 dit duidt op een lek. Derhalve de ABL gedemonteerd en de OXI-valve tubing en de waste tubing vervangen. OXI-valve schoongemaakt. Bij demontage breekt de schroefbus van het oximetrie-compartiment af, derhalve deze ook vervangen. ABL gemonteerd, en alle testen in het servicemenu doorgelopen, deze zijn nu allemaal goed. Toestel opgestart en gekalibreerd -->oke. QC gemeten --> oke. | tubing for valve & waste tubing vervangen. | SWO-230543 | Defect | HW | Oxi module | Pinch valve tube |  | Leakage detected with the OXI-valve | Leakage is visible at the OXI-valve. In the service menu run pressure test, it indicates ' Vacuum Value ' = 0.79 This indicates a leak. Therefore disassembled the ABL and replaced the OXI-valve tubing and the waste tubing. OXI-valve cleaned. When disassembling, the screw socket of the oximetry compartment breaks down, therefore it also replaces it. ABL mounted, and all the tests in the service menu went through, these are now all good. Device booted and calibrated--> okay. QC measured--> okay. | tubing for valve & waste tubing vervangen.
No | Visit | Complaint | RBNL | Lekkage geconstateerd bij de OXI-valve | Er is lekkage te zien bij de OXI-valve. In het servicemenu pressure test uitgevoerd, deze geeft bij 'Vacuum Value' = 0,79 dit duidt op een lek. Derhalve de ABL gedemonteerd en de OXI-valve tubing en de waste tubing vervangen. OXI-valve schoongemaakt. Bij demontage breekt de schroefbus van het oximetrie-compartiment af, derhalve deze ook vervangen. Tevens is te zien dat de accu op 0% staat. Bij ontkoppelen van de netstroom blijft ABL wel functioneren. Resetten middels de jumper geeft geen resultaat. Accu vervangen, nu wordt wel weer juiste status en laadpercentage weergegeven. ABL gemonteerd, en alle testen in het servicemenu doorgelopen, deze zijn nu allemaal goed. Toestel opgestart en gekalibreerd -->oke. QC gemeten --> oke. | OXI-valve tubing en de waste tubing vervangen. Accu vervangen. | SWO-230587 | Defect | HW | Basic Unit | Battery | Service tip not followed LEH Mail to Luc | Leakage detected with the OXI-valve | Leakage is visible at the OXI-valve. In the service menu run pressure test, it indicates ' Vacuum Value ' = 0.79 This indicates a leak. Therefore disassembled the ABL and replaced the OXI-valve tubing and the waste tubing. OXI-valve cleaned. When disassembling, the screw socket of the oximetry compartment breaks down, therefore it also replaces it. It is also shown that the battery is at 0%. ABL will continue to function when disconnecting the mains power. Resetting through the jumper does not result. Replace battery, now it displays correct status and charge percentage. ABL mounted, and all the tests in the service menu went through, these are now all good. Device booted and calibrated--> okay. QC measured--> okay. | OXI-valve tubing en de waste tubing vervangen. Accu vervangen.
Yes | 2 | Y | RFRA | Blank | Plus de connexion à Aqure |  | D320C802C82E42A3 | Re classified RMED | Phone support |  |  |  | Blank | More connection to Aqure
No | 4 | Y | RPOL | Blank | sprawdzenie analizatora, prostowanie igly, puszczenie testowego badania. Analizator sprawny technicznie. |  | D3212029C81ADE08 | Adjustment | HW | Auto Inlet module | Probe bend |  | Blank | The analyser, straighten the needle, releasing the test. Technically efficient analyzer.
Yes | 2 | Y | RPAC | Blank | Inlet jammed |  | D320C730C84C414C | Re classified RMED |  |  |  |  | Blank | Inlet jammed
No | 5 | Y | RPOL | Blank | Diagnoza problemu, tetsy serwisowe, analiza historii pomiarów. Konieczna wymiana pakietu odczynnikowego. |  | D3213CADC8182053 | Assisted Replacement | SP | SP | SP | Oxi, Crea, Urea, Glu | Blank | Diagnosis of the problem, service Tetsy, analysis of the history of measurements. Replacement of the reagent package is necessary.
No | Visit | Complaint | RCHN | 温度不稳定 | 温度不稳定，借备件903-334一个 |  | SWO-230335 | [5 empty] | Temperature instability | Temperature instability, spare parts 903-334 one
No | Visit | Complaint | RCHN | 多参数定标失败 | 多参数定标失败，借902-972一个。 | 更换902-972后修复 by Andrew | SWO-230476 | [5 empty] | Multi-parameter calibration failed | Multi-parameter calibration failed, borrowing 902-972 one. | Repair by Andrew after replacement 902-972
No | Visit | Complaint | RCHN | 测试卡试剂包提示验证失败 | 测试卡试剂包提示验证失败，借备件903-334一个 |  | SWO-230324 | [5 empty] | Test card kit prompt verification failed | Test card kit prompt verification failure, borrow spare parts 903-334 one
No | Visit | Complaint | RCHN | 定标不过，屏幕有断电现象 | 定标不过，屏幕有断电现象，需要借件903-365和902-778各一个 | 需要借件903-365和902-778 | SWO-230453 | [5 empty] | But the screen has a power failure | However, the screen has a power loss phenomenon, need to borrow 903-365 and 902-778 each one | Borrowed 903-365 and 902-778
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | PO2质控C超范围 0090 | 更换试剂包后修复 | SWO-230357 | [5 empty] | Test card, reagent package failure | PO2 QC C over range 0090 | Repair after replacement of reagent kit
No | Visit | Complaint | RCHN | 定标失败 | 定标失败，借备件903-364一个 | 未使用借件修复 by Andrew | SWO-230334 | Defect | HW | Process Module | Optical Unit | 1025 | Bid failed | Bid failure, borrowed spare parts 903-364 one | Not using a debit to fix by Andrew
No | Visit | Complaint | RCHN | 测试卡无法扫描 | 测试卡无法扫描，借备件910-304一个 |  | SWO-230336 | Defect | HW | Storage Module | Barcode reader, Storage wheel | BR001 | Test card can't be scanned | Test card can not be scanned, borrowed spare parts 910-304 one
No | Visit | Complaint | RCHN | TNI定标失败 | TNI定标失败,借 944-212（13147）一张。 |  | SWO-230342 | Assisted Replacement | Consumables | Test Kit | Cartridge, Tni |  | TNI calibration failed | TNI bid failed, borrowing 944-212 (13147).
No | Visit | Complaint | RBNL | Calibration refusées | Depuis quelques jours les calibrations pH , Ca, K sont refusées à cause de la sensibilité trop élevée. Actions correctives du client : décontaminations, remplacement de l'électrode de réf., membranes de réf., électrode de Ca. J'ai échangé les 2 électrodes de réf. des 2 ABL dans le service mais sans amélioration. Ensuite échangé les 2 supports noirs au dessus des électrodes (Cl et Réf), pas OK, remplacé le module pH, le problème reste toujours. Echangé les électrodes Ca et du coup le problème se déplace vers l'autre appareil ! | Remplacé la membrane de Ca sur l'ancienne électrode et remis en route (module original) : OK | SWO-230833 | Defect | HW | pH/BG Module | Electrode connector |  | Recalibration Denied | For a few days the calibrations pH , Ca, K are refused because of the sensitivity too high. Customer corrective actions: decontaminations, ref. electrode replacement, ref. membranes, Ca electrode. I exchanged the two ref. electrodes of the 2 ABLs in the service but without improvement. Then exchanged the 2 black supports above the electrodes (Cl and Ref), not OK, replaced the pH module, the problem still remains. Interchanged the Ca electrodes and suddenly the problem moves to the other device! | Replaced the Ca membrane on the old electrode and restarted (original module): OK
No | 5 | Y | RAME | F Customer reported oximetry errors intermittently. | Changed the complaint visit tab from N to Y. | Error code 316 (neon number too low) in activity log. Complete oximetry test run with no errors. Software version at 6.16 | D320F7E6C84D59A1 | Defect | HW | Oxi module | Spectrometer |  | F Customer reported oximetry errors intermittently. | Changed the complaint visit tab from N to Y. | Error code 316 (neon number too low) in activity log. Complete oximetry test run with no errors. Software version at 6.16
No | 10 | Y | RAME | 1 pH UNSTABLE INTERMITTENT ON PATIENT SAMPLES, CUSTOMER DECLINED SOFTWARE UPGRADE. |  | REVIEWED LOGS. FOUND PROBLEM INCREASED FREQUENCY AFTER 8 $28 BUT WAS PRESENT PRIOR. INSPECTED ELECTRODE (XT-01) FOUND BLOOD ON TIP AND SEALING RING WORN. | D320CF2EC803BD33 | Defect | Electrode | Electrode | Sealing ring | Sealing ring, pH | 1 pH UNSTABLE INTERMITTENT ON PATIENT SAMPLES, CUSTOMER DECLINED SOFTWARE UPGRADE. |  | REVIEWED LOGS. FOUND PROBLEM INCREASED FREQUENCY AFTER 8 $28 BUT WAS PRESENT PRIOR. INSPECTED ELECTRODE (XT-01) FOUND BLOOD ON TIP AND SEALING RING WORN.
No | Visit | Complaint | RDEU | ABL reagiert verzögert | Fehlersuche vor Ort. | Workaround für Verzögerung des Bildschirmschoners angewendet. Mehrere Kalibrierungen durchgeführt. QCs für alle Messungen durchgeführt. Mehrere Testmessung durchgeführt. ABL wieder messbereit. | SWO-230684 | Defect | SW | Windows | Windows | Slow response | ABL reacts delayed | Troubleshooting on site. | Workaround for screen saver delay. Multiple calibrations performed. QCs for all measurements. Multiple test measurements have been performed. ABL ready to be measured again.
No | 5 | Y | RPOL | Blank | Diagnoza problemu. Reinstalacja Windows 7 oraz aplikacji Radiometer 6.16. Wgranie ustawieñ, tetsy kontrolne. Analizator sprawny. |  | D320F674C8272051 | Defect | SW | Windows | Windows |  | Blank | Diagnosis of the problem. Reinstalling Windows 7 and the Radiometer 6.16 application. Settings, Audit Tetsy. Functional Analyzer.
Yes | 1 | Y | RPAC | Blank | Reference Electrode |  | D3237A93B4046A25 | Re classified RMED | OOB |  |  |  | Blank | Reference Electrode
No | Visit | Complaint | RDEU | 12.09.2017, Herr Knorr, tel: 015254847560: ABL bleibt nach Neustart auf der Windows Ebene hängen | Windows Fehlermeldung beim laden File: \Windows\system32\config\system | Windows Software neu installiert DMS Software installiert, Gerät konfiguriert und angebunden Gerät kalibriert und überprüft | SWO-230605 | Defect | SW | Windows | Windows |  | 12.09.2017, Mr. Knorr, tel: 015254847560: ABL gets stuck on the Windows level after reboot | Windows error message when loadingFile: \ Windows \ system32 \ config \ system | Windows software reinstalled DMS software installed, device configured and connected Device calibrated and checked
No | Visit | Complaint | RDAN | Jane melder, at efter utæthed ved reference elektroden er målekammeret rødt på ydersiden så hun ikke kan gøre det rent. | Afprøvet. | Målekammer udskiftet, slutafprøvet. | SWO-230607 | Defect | HW | pH/BG Module | Measuring chamber, REF | Reference | Jane reports that after leakage at the reference electrode, the measuring chamber is red on the outside so she cannot clean it. | Tested. | Measuring chamber replaced, final tested.
No | Visit | Complaint | RDAN | Martin melder, at den drypper fra inlet. | Afprøvet. | Inlet pakning, probe, probe holder, slange, spild pumpeslanger og pH pakring udskiftet, slutafprøvet. | SWO-230603 | Defect | HW | Inlet module | Inlet gasket | LEH Mail | Martin reports that it is dripping from the inlet. | Tested. | Inlet packing, probe, probe holder, hose, waste pump hoses and pH packing replaced, final tested.
No | 5 | Y | RFRA | Blank | Erreur 4794 |  | D32105EAC822F152 | Defect | HW | Aspiration Module | Tube holder |  | Blank | Error 4794
No | 12 | Y | RAME | 2 PCO2 unable to calibrate on analyzer after multiple cassettes and solution packs changed. |  | Found that the analyzer was not able to successfully calibrate due to the sensor interface board. | D3220C153E50D1F9 | Defect | HW | Basic Unit | PCB, Sensor interface | pCO2 | 2 PCO2 unable to calibrate on analyzer after multiple cassettes and solution packs changed. |  | Found that the analyzer was not able to successfully calibrate due to the sensor interface board.
No | 5 | Y | RAME | L instrument customer spent 2 months trying to get redults to cross to Sunquest. |  | worked with Todd, debbie and Juanita to get the system working, turned out to be a sunquest settning issue | D3221B9E3E152F38 | Setup | Ext.SW | His/lis | His/lis |  | L instrument customer spent 2 months trying to get redults to cross to Sunquest. |  | worked with Todd, debbie and Juanita to get the system working, turned out to be a sunquest settning issue
No | Visit | Complaint | RDEU | Sporadisch Oximetriefehler 581 | Fehler 581 und zu niedrige tHB Werte | Hemolyzer ersetzt, Schlauchventil überholt, checkout and adjustment durchgeführt, div. Blutmessungen überprüft, OK Defekt: 908-552R1936N011 902-980 ohne Seriennummer | SWO-231138 | Defect | HW | Oxi module | Hemolyzer |  | Sporadisch Oximetriefehler 581 | Error 581 and tHB values that are too low | Hemolyzer replaced, hose valve overhauled, checkout and adjustment performed, various blood measurements checked, OK Defect: 908-552R1936N011 902-980 without serial number
No | Visit | Complaint | RCHN | 前面板开裂 | 前面板开裂，借924-738一个。 | 更换924-738后修复 by Andrew | SWO-230447 | [5 empty] | Front panel cracking | The front panel is cracked, borrowing 924-738 one. | Repair by Andrew after replacement 924-738
No | Visit | Complaint | RCHN | PO2定标失败 | PO2定标失败 ，借902-972一个。 | 更换902-972备件后修复 by Andrew | SWO-230534 | [5 empty] | PO2 calibration failure | PO2 calibration failure , borrowed 902-972 one. | Repair by Andrew after replacement of 902-972 spare parts
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 温度错误 | 更换测试卡后修复 | SWO-233726 | [5 empty] | Test card, reagent package failure | Temperature error | Repair after replacing test card
No | Visit | Complaint | RIND | SC not installing. | Sensor board to replace. | Sensor board replaced. Zeroing calibration done. Analyser Serviced and calibrated properly. Handed over in a fine working condition. | SWO-231618 | [5 empty] | SC not installing. | Sensor board to replace. | Sensor board replaced. Zeroing calibration done. Analyser Serviced and calibrated properly. Handed over in a fine working condition.
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 参比电极错误，1361 | 更换测试卡后修复 | SWO-230863 | Assisted Replacement | SC | SC | Reference | 1361 | Test card, reagent package failure | Reference electrode error, 1361 | Repair after replacing test card
No | Visit | Complaint | RCHN | 启动时无法连接 | 启动时无法连接，借备件910-285一个 | 更换910-285后修复 by Andrew | SWO-230525 | Defect | HW | Basic Unit | Power Supply | 1223 | Unable to connect at startup | Unable to connect at start-up, spare parts 910-285 | Repair by Andrew after replacement 910-285
No | Visit | Complaint | RDEU | 13.09.2017 08:35Uhr Frau Hergarten 02114092367: Das Gerät meldet bei ca. jeder vierten Probe Probenfehler | Flat Membran am Air Membranventil defekt. | Air Membranventil im pH & BG Modul demontiert, gereinigt und Flat Membran erneuert, Messkammer gereinigt, pH Elektrode gereinigt, Pumpen Kal und LS Justage durchgeführt, Gerät kalibriert und QC durchgeführt. | SWO-231123 | Defect | HW | pH/BG Module | Valve, air | Membrane | 13.09.2017 08:35pm Mrs Hergarten 02114092367: The device reports sample errors for about one in four samples | Flat membrane defective on air diaphragm valve. | Air diaphragm valve in pH & BG module dismantled, cleaned and flat membrane renewed, measuring chamber cleaned, pH electrode cleaned, pumps Kal and LS adjustment performed, device calibrated and QC performed.
No | 13 | Y | RAME | F Customer reported oximetry errors while running VK-Kit level 3. |  | Checked data logs, found 581 errors when VK level 3 solution was run. | D32261B53E005B7C | Defect | HW | Oxi module | Hemolyzer |  | F Customer reported oximetry errors while running VK-Kit level 3. |  | Checked data logs, found 581 errors when VK level 3 solution was run.
No | 6 | Y | RAME | P Customer reported waste detector error in system messages. Waste bottle getting over filled. |  | Checked data logs for errors. Error code 618 showing in activity log. | D32238233E28D458 | Defect | HW | Basic Unit | Waste detector |  | P Customer reported waste detector error in system messages. Waste bottle getting over filled. |  | Checked data logs for errors. Error code 618 showing in activity log.
No | Visit | Complaint | RLTD | Leakage detected during installation in new department | Faulty hemolyser noticed during installation | Replaced hemolyser unit with new | SWO-231446 | Defect | HW | Oxi module | Hemolyzer | Leakage | Leakage detected during installation in new department | Faulty hemolyser noticed during installation | Replaced hemolyser unit with new
No | Visit | Complaint | RDEU | 13.09.2017 Sr. Bettina 02114092520 meldet, dass die Einlassnadel geknickt ist. | Einlasssonde leicht verbogen. | Einlasssonde gerichtet, Gerät kalibriert und QC durchgeführt. | SWO-231279 | Adjustment | HW | Inlet module | Probe bend |  | 13.09.2017 Sr. Bettina 02114092520 reports that the inlet needle has buckled. | Inlet probe slightly bent. | inlet probe, device calibrated and QC performed.
No | Visit | Complaint | RLTD | The Analyser is failing LV B & C QC's | Inspected data logs for fault history. Level B & C are consistently failing on co-oximetry values. | Replaced Hemolyser, ran fluid transport checks, Complete OXI test and updated Hemolyser status. Performed calibrations, tHb calibration and QC. Analyser is ready for use. | SWO-231427 | Defect | HW | Oxi module | Hemolyzer | QC B-C | The Analyser is failing LV B & C QC's | Inspected data logs for fault history. Level B & C are consistently failing on co-oximetry values. | Replaced Hemolyser, ran fluid transport checks, Complete OXI test and updated Hemolyser status. Performed calibrations, tHb calibration and QC. Analyser is ready for use.
No | Visit | Complaint | RCHN | 接口板值错误 | 接口板值错误，需要借件902-972 | 更换902-972后修复 by Andrew | SWO-230845 | [5 empty] | Interface board value error | Interface board value error, need to borrow 902-972 | Repair by Andrew after replacement 902-972
Yes | 22 | Y | RAME | P Update for Techline case RAME 783271241 for possible fluid blockage. |  | Came in to replace the hemolyzer and spectrometer to see if it solves the possible fluid blockage issue. Customer will monitor and let me know. | D323517AB44443EF | [5 empty] | P Update for Techline case RAME 783271241 for possible fluid blockage. |  | Came in to replace the hemolyzer and spectrometer to see if it solves the possible fluid blockage issue. Customer will monitor and let me know.
No | Visit | Complaint | RSCH | Blank | Mehrmals Glu. Probleme beim ersetzen von SC. |  | SWO-231602 | [5 empty] | Blank | Several times Glu. Problems replacing SC.
No | 4 | Y | RAME | Z Error 2100 CO-OX Control Board failure. Considered out of box failure |  | found error messages. Obtained service dump | D3235D95B42673F7 | [5 empty] | Z Error 2100 CO-OX Control Board failure. Considered out of box failure |  | found error messages. Obtained service dump
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | PO2定标质控低于质控范围 0090 | 连续更换2个试剂包修复 | SWO-231797 | [5 empty] | Test card, reagent package failure | PO2 calibration qc control is lower than QC range 0090 | 2 reagent kits for continuous replacement
No | Visit | Complaint | RCHN | 定标不过 | 定标不过，借备件903-364一个，903-584两个，944-212（13079）一张 | 更换903-364一个，903-584两个，944-212（13079）备件后修复 by Andrew | SWO-230790 | Defect | HW | Process Module | Optical Unit | 1031 | The calibration is not the | However, spare parts 903-364 one, 903-584 two, 944-212 (13079) one | Replace 903-364 one, 903-584 two, 944-212 (13079) spare parts after repair by Andrew
No | Visit | Complaint | RCHN | 光学强度超范围 | 光学强度超范围，借903-364一个 | 工程师要求加借2张定标，944-258（12979，13036）各一张、944-212（12862）一张 by Andrew | SWO-231150 | Revisit required | HW | Process Module | Optical Unit | 1031 | Optical strength over range | Optical strength over range, borrowed 903-364 a | Engineerrequests to borrow 2 bids, 944-258 (12979, 13036) each, 944-212 (12862) by Andrew
No | Visit | Complaint | RCHN | 报错1068，干燥强度超出范围 | 报错1068，干燥强度超出范围借903-118一个 | 更换903-118后修复 by andrew | SWO-230816 | Defect | HW | Process Module | Cup dryer | 1068 | Error 1068, drying strength out of range | Error 1068, drying strength out of range borrowed 903-118 one | Repair by andrew after replacement 903-118
No | Visit | Complaint | RCHN | CRP定标中止 5784 | CRP定标中止 5784，借 944-267（13125）一张。 | 清理处理盘异物，重新定标 | SWO-231146 | Cleaned | HW | Process Module | Process wheel | 5784 | CRP Calibration Abort 5784 | CRP calibration suspension 5784, borrow 944-267 (13125) one. | Clean up the process disc foreign body and re-label
Yes | 9 | Y | RPAC | Blank | Screen intermittently going black |  | D3232F2CB4267B08 | [5 empty] | Blank | Screen intermittently going black
Yes | 7 | Y | RPAC | Blank |  |  | D3236F29B45DDFBE | Re classified RMED |  |  |  |  | Blank
No | 13 | Y | RAME | 1 pH electrode well filling with fluid, pH, ca, Na out on QC |  | checked data logs for errors, checked electrodes and electrode membranes, checked electrode chambers | D323DB03B4106940 | User error | Electrode | Electrode | Sealing ring | Sealing ring, pH | 1 pH electrode well filling with fluid, pH, ca, Na out on QC |  | checked data logs for errors, checked electrodes and electrode membranes, checked electrode chambers
No | 5 | Y | RPOL | Blank | Diagnoza problemu, zmiana ustawieñ d¿wiêku w Windowsie. Testy kontrolne. Analizator sprawny. |  | D3236E10B42F2056 | Setup | SW | SWU | Sound | Why not install 6.17 | Blank | Diagnose the problem, change the sound settings in Windows. Control tests. Functional Analyzer.
Yes | 1 | Y | RPAC | Blank | pCO2 Electrode no good |  | D328A2638C00BFF1 | [5 empty] | Blank | pCO2 Electrode no good
No | Visit | Complaint | RDEU | LQC Level 2 und 3 werden bei TnT sehr häufig zu hoch gemessen | Erneut ist Level 3 der LQC (Technopath LOT 06173) oberhalb der Grenzwerte | Cup Wash Unit demontiert und gereinigt -anschließend LS-Fehler (FTC103) Wetsection PCB aus Leihgerät eingebaut | SWO-232193 | Cleaned | HW | Process Module | Cup wash unit | LQC Level 3, high | LQC Levels 2 and 3 are very often over-measured at TnT | Once again Level 3 of the LQC (Technopath LOT 06173) is above the limit values | Cup Wash Unit dismantled and cleaned -following LS error (FTC103) Wetsection PCB built-in from rental device
No | 2 | Y | RAME | 9 Lactate will not pass calibration. |  | Lactate will not pass sensitivity. It just shows ... with error code 1310 response error and 0378 and calibration error. | D324F0F02A46BD3C | Defect | HW | Basic Unit | PCB, Sensor interface | Lac | 9 Lactate will not pass calibration. |  | Lactate will not pass sensitivity. It just shows ... with error code 1310 response error and 0378 and calibration error.
No | Visit | Complaint | RBNL | ABL gaat in vervangingsscherm | ABL90 springt soms in het vervangingsscherm, meestal na een meting uit het buisje. ABL90 heb doorgespoten en voorzien van nieuwe gasket $connector en OXI valvetubing. | Reiniging en vervanging tubings. | SWO-232582 | Defect | HW | Auto Inlet module | Inlet gasket |  | ABL gaat in vervangingsscherm | ABL90 sometimes jumps in the replacement screen, usually after a measurement from the tube. ABL90 have repainted and equipped with new gasket /connector and OXI valvetubing. | Cleaning and replacement tubings.
No | Visit | Complaint | RKK | 初期化ｐHエラーが多いので訪問 | 確認 | シュミレーターで確認して機械側には問題無、カセットの再初期化でOK | SWO-232177 | Assisted Replacement | SC | SC | pH | Re install SC | There are a lot of initialization pH errors, so visit | Check | There is no problem on the machine side by confirming it with the simulator, and it is OK by reinitializing the cassette.
No | 2 | Y | RPAC | Blank | Inlet error, sensor door unable to close |  | D324B4352A244FD3 | Adjustment | HW | Auto Inlet module | Cable disconnected |  | Blank | Inlet error, sensor door unable to close
No | Visit | Complaint | RDEU | 15.09.2017, Herr Rehmann, Tel: 03455577433: Luftblasen Erkennung gescheitert | Interner Lösungsschlauch vom Anschlussstutzen gelöst | Fehlersuche, Korrektur Lösungsschlauch, Funktionskontrolle | SWO-231784 | [5 empty] | 15.09.2017, Mr. Rehmann, Tel: 03455577433: Air bubble detection failed | Internal solution hose detached from the connection nozzle | Troubleshooting, correction solution hose, function control
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | PO2定标不过 0210 | 更换试剂包后修复 | SWO-231798 | [5 empty] | Test card, reagent package failure | PO2 calibration is no more than 0210 | Repair after replacement of reagent kit
Yes | 2 | Y | RPAC | Blank | Replaced |  | D324B2EC2A22099B | [5 empty] | Blank | Replaced
No | Visit | Complaint | RCHN | 光学强度超范围 | 工程师要求加借2张定标，944-258（12979，13036缺货）各一张、944-212（12862）缺货一张 | 更换903-364和重新定标后修复 by Andrew | SWO-231189 | Defect | HW | Process Module | Optical Unit | 1031 | Optical strength over range | Engineers requested 2 additional bids, 944-258 (12979,13036 out of stock) one each, 944-212 (12862) out of stock | Repair by Andrew after replacement 903-364 and re-calibration
No | Visit | Complaint | RCHN | 报错5766 | 报错5766，借备件903-198一个 | 更换903-198后修复 by Andrew | SWO-230931 | Defect | HW | Analyzer Control Module | PCB, Main Board | 5766 | 5766 error | 5766, borrowed spare parts 903-198 one | Repair by Andrew after replacement 903-198
No | Visit | Complaint | RIND | Lac is out of range and need to do a routine visit. | Checked and found that the Lac membrane need to be change | changed the membrane, now working satisfactorily. | SWO-232294 | Assisted Replacement | Electrode | Membrane | Lac | LEH Mail | Lac is out of range and need to do a routine visit. | Checked and found that the Lac membrane need to be change | changed the membrane, now working satisfactorily.
No | 6 | Y | RAME | Z Ca sensitivity high outside upper limit. Inlet module broken. Spectrum mismatch errors on level 3 VK kit. |  | Ca issue mysteriously went away. Unable to resolve spectrum mismatch errors in VK Kit level 3. | D324B8492A1F23F1 | Unknown | Unknown | Unknown | Unknown |  | Z Ca sensitivity high outside upper limit. Inlet module broken. Spectrum mismatch errors on level 3 VK kit. |  | Ca issue mysteriously went away. Unable to resolve spectrum mismatch errors in VK Kit level 3.
No | 8 | Y | RPAC | Blank | Analyser will not oas the Pump calibration or Liquid Sensor calibration. |  | D3264254A03B86EA | Revisit required | HW | Basic Unit | Cassette retainer |  | Blank | Analyser will not oas the Pump calibration or Liquid Sensor calibration.
No | Visit | Complaint | RSAF | Sol pack properly activated needle straight but no inlet gasket in stock. Analyzer fails startup | Inlet gasket dirty | Cleaned inlet gasket with holder start up ok | SWO-233282 | Cleaned | HW | Inlet module | Inlet gasket |  | Sol pack properly activated needle straight but no inlet gasket in stock. Analyzer fails startup | Inlet gasket dirty | Cleaned inlet gasket with holder start up ok
No | Visit | Complaint | RCHN | 电路开路 | 电路开路，借备件902-972一个 | 更换902-972后修复 by Andrew | SWO-231638 | [5 empty] | Circuit open | Circuit open, spare parts 902-972 one | Repair by Andrew after replacement 902-972
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 1361，参比电极错误 | 更换测试卡后修复 | SWO-232201 | Assisted Replacement | SC | SC | Reference | 1361 | Test card, reagent package failure | 1361, Reference Electrode Error | Repair after replacing test card
No | Visit | Complaint | RIND | Printer head lock broken. | Printer head lock broken. | Replace the printer head lock from other faulty printer head,check the print out & it was ok | SWO-234608 | Defect | HW | Printer Module | Printer head |  | Printer head lock broken. | Printer head lock broken. | Replace the printer head lock from other faulty printer head,check the print out & it was ok
No | Visit | Complaint | RHUN | OXI hardver hibák, FLEXQ hibák | OXI hardver hibák: OXI paraméter mérése nem lehetséges FLEXQ hibák: Mintafelszívás sikertelen | OXI PCB csereéje, AU install, OXI tesztek, Felszívótű kiegyenesítése, Kalibrálás, tHb kal, QC mérés, Páciens minta teszt FLEXQ modullal | SWO-232196 | Defect | HW | Oxi module | PCB, Oxi module |  | OXY hardware failures, FLEXQ errors | OXY hardware failures: OXY parameter measurement not possible FLEXQ defects: Sample absorption failed | OXY PCB Replacement, AU Install, OXY tests, suction needle straightening, calibration, tHb cal, QC measurement, patient sample test with FLEXQ module
Yes | Visit | Complaint | RLTD | SAYING THE LACTATE IS OUT. INTERVENTION NEEDED ON THE SCREEN THE WHOLE WEEKEND. HAVE BEEN ADVISED NOT TO TOUCH IT AND TO RING RADIOMETER Key Scorrier No 01225 824 964 | Customer reporting that lactate is out and has been over the weekend. | Asked customer and they informed me that all other parameters are working just lactate not come in after sensor cassette change. Advised customer that the sensor cassette need to be changed. Customer told me that they had no other cassettes and that their new ones were on order. Customer said that they had another analyser that they could use for lactate until the new cassettes arrived. | SWO-232267 | Re classified RMED | Phone support |  |  |  | SAYING THE LACTATE IS OUT. INTERVENTION NEEDED ON THE SCREEN THE WHOLE WEEKEND. HAVE BEEN ADVISED NOT TO TOUCH IT AND TO RING RADIOMETER Key Scorrier No 01225 824 964 | Customer reporting that lactate is out and has been over the weekend. | Asked customer and they informed me that all other parameters are working just lactate not come in after sensor cassette change. Advised customer that the sensor cassette need to be changed. Customer told me that they had no other cassettes and that their new ones were on order. Customer said that they had another analyser that they could use for lactate until the new cassettes arrived.
No | Visit | Complaint | RCHN | 无法阅读测试卡智能芯片 | 无法阅读测试卡智能芯片，借902-972和903-334各一个 |  | SWO-232245 | [5 empty] | Can't read test card smart chip | Unable to read the test card smart chip, borrow 902-972 and 903-334 each one
No | Visit | Complaint | RDEU | 18.09.2017, Frau Pfilippsen, Tel: 021596795132: Lösungen können nicht gepumpt werden 19.09.2017, Frau Philippsen, Tel: 021596795132: weiterhin Lösungen können nicht gepumpt werden. | Kal-3-Lösung kann nicht gefördert werden. Fehler 1084: Inhomogene Kal3-Lösung | Gerät durchgespült, über das Servicemenü die Fluidic-Transport-Programme durchgeführt. Wiederkehrender Fehler bei der Kal-3-Lösung. Auch im Visual Check kann die Kal-3-Lösung nicht beobachtet werden. Erneuerung des Lösungspacks aus Kundenbestand (alt und neu jeweils Lot YR-54). Während der Startphase keine Fehlermeldungen. | SWO-232770 | Assisted Replacement | SP | SP | FT | Cal 3 | 18.09.2017, Mrs Pfilippsen, Tel: 021596795132: Solutions cannot be pumped 19.09.2017, Mrs. Philippsen, Tel: 021596795132: solutions cannot be pumped. | Kal-3 solution cannot be promoted. Error 1084: Inhomogeneous Kal3 Solution | device, the Fluidic transport programs are carried out via the service menu. Recurring error with the Kal-3 solution. Also in Visual Check, the Kal-3 solution cannot be observed. Renewal of the solution pack from customer base (old and new lot YR-54). No error messages during the startup phase.
No | Visit | Complaint | RDEU | 18.09.2017 Frau Giacomazza 0213152953152 meldet, dass die Blutprobe nicht angesaugt wird. | Proben werden nicht mehr angesaugt. | Einlassnadel verbogen - erneuert. Fehlerbehebung durchgeführt. Gerät steht in der Startphase (50min). | SWO-232585 | Replaced | HW | Inlet module | Probe bend |  | 18.09.2017 Ms Giacomazza 0213152953152 reports that the blood sample is not sucked in. | Samples are no longer sucked in. | Inlet needle bent - renewed. Troubleshooting. Device is in the start-up phase (50min).
No | 8 | Y | RPAC | Blank | Analyser has faulty Inlet Liquid Senosr. |  | D32862518C22391F | Defect | HW | Basic Unit | Cassette retainer |  | Blank | Analyser has faulty Inlet Liquid Senosr.
No | Visit | Complaint | RLTD | Sue O'Brien Wheeler Tel 07954576051 | Checked data logs and service programs. | Installed an old sensor cassette and flushed with hypochlorite and water. Straightened slightly bent probe. Cleaned p02 sensor. Replaced OXI tubing and ensured connection gasket was clear. Swapped solution pack YN-20 with Biochem Analyser solution pack YR-02 and is now passing QC. Performed calibrations and QC. Analyser is ready for use. | SWO-232775 | Assisted Replacement | SP | SP | SP | pO2 QC | Sue O'Brien Wheeler Tel 07954576051 | Checked data logs and service programs. | Installed an old sensor cassette and flushed with hypochlorite and water. Straightened slightly bent probe. Cleaned p02 sensor. Replaced OXI tubing and ensured connection gasket was clear. Swapped solution pack YN-20 with Biochem Analyser solution pack YR-02 and is now passing QC. Performed calibrations and QC. Analyser is ready for use.
No | 4 | Y | RAME | V Customer reported unit making an unusual noise. |  | Checked pump housings and fans. Found computer fan making noise intermittently. | D32893278C5CBAB0 | [5 empty] | V Customer reported unit making an unusual noise. |  | Checked pump housings and fans. Found computer fan making noise intermittently.
No | Visit | Complaint | RDEU | Analysator steht mit roter Ampel und Meldung CS Temp. instabil. | SC Temp. instabil Kal. Fehler Keine IQC Messungen | ABL Datei-Archiv überprüft, ABL mit Zeroing Kassette überprüft - alles Ok, SC80 aus Kundenbestand getauscht, nach Init.Phase der neuen SC80 und fehlerhafter IQC Level 2 (pCO2) LP80 getauscht, ABL nach LP80 Tausch und mehreren Kal. wieder einsatzbereit, IQC Ok, System- und Funktionskontrolle in Ordnung | SWO-232666 | [5 empty] | Analyzer stands with red traffic light and message CS Temp. Unstable. | SC Temp. Unstable Kal. Error No IQC measurements | ABL file archive checked, ABL checked with zeroing cassette - all Ok, SC80 swapped from customer inventory, exchanged for init.phase of the new SC80 and faulty IQC Level 2 (pCO2) LP80, ABL to LP80 exchange and several Cal. ready for use again, IQC Ok, system and function control in order
No | Visit | Complaint | RCHN | 泵轮不转 | 泵轮不转，借903-334一个 |  | SWO-232405 | [5 empty] | Pump wheel does not turn | Pump wheel does not turn, borrow 903-334 one
No | Visit | Complaint | RIND | Unable to switch on the instrument. | Unable to switch the instrument | Replace the power cable,change the plug point but still the problem persist,need to replace the power supply for further checking. | SWO-232511 | [5 empty] | Unable to switch on the instrument. | Unable to switch the instrument | Replace the power cable,change the plug point but still the problem persist,need to replace the power supply for further checking.
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 定标漂移 0377 | 更换测试卡后修复 | SWO-232757 | Revisit required | SC | SC | SC | Unknown | Test card, reagent package failure | Calibration Drift 0377 | Repair after replacing test card
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 定标灵敏度超范围 | 更换试剂包后修复 | SWO-232758 | Assisted Replacement | SP | SP | SP | Unknown See SWO-232757 | Test card, reagent package failure | Calibration sensitivity over range | Repair after replacement of reagent kit
No | Visit | Complaint | RCHN | 报错1024 | 报错1024，借903-364一个 |  | SWO-232431 | Defect | HW | Process Module | Optical Unit | 1024 | 1024 error | 1024, loan 903-364 one
No | Visit | Complaint | RCHN | 定标终止 | 定标终止，借944-394（13041）一张 | 工程师来电告知仪器序列号报错，此单关闭处理(By Ran) Wrong SN. Case closed. | SWO-232354 | Assisted Replacement | Consumables | Test Kit | Cartridge, PCT | 1065 | Bid termination | Bid termination, loan 944-394 (13041) one | The engineer called to inform the instrument of an error in the serial number, and this single closed processing (By Ran) Wrong SN. Case closed.
No | Visit | Complaint | RCHN | PCT定标终止 | PCT定标终止，借944-394（批号13041）一个 |  | SWO-232392 | Assisted Replacement | Consumables | Test Kit | Cartridge, PCT | 1056 | PCT calibration termination | PCT calibration terminated, borrowing 944-394 (lot number 13041) one
Yes | Visit | Complaint | RCHN | MOY定标失败，卡盘 | MOY定标失败，卡盘，借944-214（12841）一张。 | 机器序列号报错了。该Case取消。(By Ran)Wrong SN. Case Closed. | SWO-232432 | Re classified RMED | Wrong ID |  |  |  | MOY calibration failure, card disk | MOY bid failure, card plate, borrow 944-214 (12841) one. | The machine serial number was reported incorrectly. The Case is canceled. (By Ran) Wrong SN. Case Closed.
No | Visit | Complaint | RCHN | 处理包弹不出，报错8957 | 处理包弹不出，报错8957，需要借件903-036备件一个 |  | SWO-232300 | Defect | HW | Cover and Chassis | PCB, replenisment PCB | 8957 | The handling bag could not be ejected, the error 8957 | Handling package can not be ejected, reported error 8957, need to borrow 903-036 spare parts one
No | 7 | Y | RAME | Z Fails leak test. |  | Gasket missing from K electrode. | D32876C38C616799 | Defect | Electrode | Membrane | K+ | Mempack | Z Fails leak test. |  | Gasket missing from K electrode.
No | Visit | Complaint | RDEU | Anwender haben immer wieder Probleme mit der Bedienung des Bildschirms. das ist schon länger und sind dadurch genervt. | Einstellungen für den Bildschirmschoner nach RDEU Vorgabe überprüft>>I.O. ABL Neugestartet Das Display funktioniert. | Herr Kehl ist mit der Funktion des Displays unzufrieden und möchte eine dauerhafte Lösung. Neues Display für den 19.10 bestellt. | SWO-232525 | Revisit required | HW | DMS_PC | Touch screen |  | Users have always had problems with the operation of the screen. this has been a long time and is therefore annoyed. | Checked settings for the screen saver according to RDEU specification>>I.O. ABL Restarted The display works. | Mr Kehl is dissatisfied with the function of the display and wants a permanent solution. New display ordered for 19.10.
No | Visit | Complaint | RDEU | 30.08.2017, Herr Kamp, Tel: 06946086123: Probenabbrüche mit 1067 und 3553 | 10 Blankmessungen durchgeführt, alle i.O, QCs: MCC1, MCC2 D-Dimer Level 1 und 2 je 2 mal durchgeführt. Probenabrüche tauchen nicht mehr auf. MCC1 Lot 09161 wird zu Hoch gemessen, nach Rücksprache mit Frau Löwenkamp; bekanntes Problem mit Lot. | Gerät wieder in CPU gebracht, Startphase, QCs alle i.O, Leihgerät abgeholt. Kunden informiert. | SWO-232812 | Defect | Unknown | Unknown | off/on | 1067 | 30.08.2017, Mr. Kamp, Tel: 06946086123: Sample cancellations with 1067 and 3553 | 10 blank measurements performed, all i.O, QCs: MCC1, MCC2 D-Dimer Level 1 and 2 performed 2 times each. Sample breaks no longer appear. MCC1 Lot 09161 is measured too high, after consultation with Ms. Löwenkamp; known problem with Lot. | Device brought back into CPU, start phase, QCs all i.O, loaner picked up. informed customers.
No | Visit | Complaint | RDEU | 18..09.2017 Herr Theel (Medizintechnik) 0581836051 meldet eine inkonsistente Datenbank. | System überprüft => Gerät muß neu gestartet werden | Gerät neugestartet | SWO-232870 | Adjustment | SW | off/on | Data base inconsistent | 1405 | 18..09.2017 Mr. Theel (medical technology) 0581836051 reports an inconsistent database. | System checked => Device must be restarted | Device restarted
Yes | 4 | Y | RFRA | Blank | Erreur oxi sur QC auto et calibration |  | D329D623023DCEC2 | Re classified RMED | Phone support |  |  | see D32C2E59EE32C612 | Blank | Error oxi on QC auto and calibration
No | Visit | Complaint | RDEU | 18.09.2017, Frau Pfilippsen, Tel: 021596795132: Lösungen können nicht gepumpt werden 19.09.2017, Frau Philippsen, Tel: 021596795132: weiterhin Lösungen können nicht gepumpt werden. | Lac Empfindlichkeit 2946pA $mM. Der Kunde hat keine gültige Sensorkassette mehr vor Ort. 3-4 Sensorkassetten welche der Kunde vor Ort hat sind abgelaufen. | Sensorkassette organisiert und am Gerät eingesetzt, Startphase durchgeführt, nach der Startphase ist pO2, Na, CI, Glu und Lac gelb - erneut Kalibriert und QC's durchgeführt. | SWO-233078 | Assisted Replacement | SC | SC | SC | Lac | 18.09.2017, Mrs Pfilippsen, Tel: 021596795132: Solutions cannot be pumped 19.09.2017, Mrs. Philippsen, Tel: 021596795132: solutions cannot be pumped. | Lac Sensitivity 2946pA /mM. The customer no longer has a valid sensor cartridge on site. 3-4 Sensor cartridges which the customer has on site have expired. | Sensor cassette organized and inserted on the device, start phase performed, after the start phase is pO2, Na, CI, Glu and Lac yellow - recalibrated and QC's performed.
No | Visit | Complaint | RSCH | Blank | Erreur QC PO2 malgré changement 2X SC et SP | Contrôle le canal de mesure de la PO2 avec le simulateur système e.o Nettoyage du système fluidique e.o Remplacement du Pack de solution 3X e.o Calibration e.o contrôle de qualité e.o | SWO-233066 | Assisted Replacement | SP | SP | SP | pO2 QC | Blank | ERROR QC PO2 despite change 2X SC and SP | Controls the PO2's measurement channel with the e.o system simulator Cleaning the fluid system e.o Replacing the 3X e.o Solution Pack Calibration e.o quality control e.o
No | Visit | Complaint | RBNL | Hoi Siemen, Zou je de batterij van de KIC groen willen vervangen? Deze is defect (0% stroom in de batterij) met vriendelijke groet, Jeroen J.J. van Leeuwen Staf Medewerker / Point Of Care Coördinator Laboratoriumgeneeskunde Universitair Medisch Centrum Groningen (UMCG) Hanzeplein 1 9713 GZ Groningen Kamer Y1.131 | Status geeft 0% aan maar instrument blijft wel normaal functioneren als het netsnoer uitgetrokken wordt. In servicemenu precentage remaining 0. Resetten van battery unit verhelpt probleem niet. | Nieuwe batterij unit gemonteerd nu alles weer OK, percentage 84% en stijgend. Gekalibreerd en QC gemeten. | SWO-232888 | Defect | HW | Battery module | Battery | 0% | Hi Siemen, Would you like to replace the battery of the KIC green? This is defective (0% power in the battery) Sincerely Vijay Kumar Verma Staff member/Point Of Care Coordinator Laboratory Medicine University Medical Center Groningen (UMCG) Hanzeplein 1 9713 GZ Groningen Room Y 1.131 | Status indicates 0% but instrument will continue to function normally when the power cord is pulled out. In Service Menu Precentage remaining 0. Resetting battery unit does not resolve problem. | New battery unit mounted now everything OK again, percentage 84% and rising. Calibrated and QC measured.
No | 18 | Y | RPNZ | Blank | Fluid transport (Clot detected). Analyzer in Intervention Required mode |  | D32A206202560754 | Revisit required | HW | Oxi module | Hemolyzer |  | Blank | Fluid transport (Clot detected). Analyzer in Intervention Required mode
Yes | 1 | Y | RFRA | Blank |  |  | D329D36D02515FF2 | Re classified RMED | Analyzer installation |  |  |  | Blank
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | Lac敏感度偏低 | 更换测试卡后修复 | SWO-232715 | [5 empty] | Test card, reagent package failure | Lac sensitivity is low | Repair after replacing test card
No | Visit | Complaint | RESP | no inicializa ni O2 ni pCO2 | no inicializa ni O2 ni pCO2 | Se limpio internamente el equipo, se comprobo el cero electronico del O2, se realizaron varias inicializaciones del SCI y continuan fallando los parametros de pCO2 y O2. El equipo necesita un cambio de SCI, del cual el cliente no dispone de mas. Se quedo el equipo trabajando sin los parametros de O2 ni pCO2 hasta que el cliente sustituya este SCI como defectuoso. | SWO-233574 | [5 empty] | it doesn't initialize Either O2 or PCO2 | it doesn't initialize Either O2 or PCO2 | The equipment was cleaned internally, the Electronic zero of the O2 was compiled, several initializations of the SCI were performed and the pCO2 and O2 parameters continue to fail. The equipment needs a change of SCI, of which the customer does not have more. The equipment was stayed working without the O2 or pCO2 parameters until the customer replaces this SCI as defective.
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | Lac,GLU定标不过 0210 | 更换测试卡后修复 | SWO-232751 | Assisted Replacement | SC | SC | Glu/Lac | 210 | Test card, reagent package failure | Lac, GLU calibration is no more than 0210 | Repair after replacing test card
No | Visit | Complaint | RESP | MODULO INLET ROTO. | MODULO INLET ROTO | SUSTITUCION Y AJUSTE DE MODULO INLET COMPLETO POR ROTURA DEL MISMO. CALIBRADO Y VERIFICADO CON Q.C. | SWO-233073 | Defect | HW | Auto Inlet module | Auto Inlet module |  | BROKEN INLET MODULE. | BROKEN INLET MODULE | REPLACEMENT AND ADJUSTMENT OF COMPLETE INLET MODULE FOR BREAKING THE SAME. CALIBRATED AND VERIFIED WITH Q.C.
No | 1 | Y | RPAC | Blank | TCM5 has been damaged, the Calibration Chamber door has been broken off. The Serial Port connector has been pushed in at the rear. |  | D32C1965EE1536CC | [5 empty] | Blank | TCM5 has been damaged, the Calibration Chamber door has been broken off. The Serial Port connector has been pushed in at the rear.
No | Visit | Complaint | RCHN | 无法读取试剂包 | 无法读取测试卡，借910-304一个 | 更换910-304后修复(by tim) | SWO-232371 | Defect | HW | Storage Module | Barcode reader, Storage wheel |  | Unable to read the package | Failed to read test card, borrow 910-304 one | Repair after replacement 910-304 (by tim)
No | Visit | Complaint | RCHN | CKMB定标失败 | CKMB定标失败，借944-216（13052）一张 （缺货） | 重新定标后修复（by tim 操作失误，卡过期） | SWO-232366 | Assisted Replacement | Consumables | Test Kit | Cartridge, CKMB | Email - err code | CKMB calibration failed | CKMB bid failure, borrow 944-216 (13052) one (out of stock) | Fix after recalibration (by tim operation error, card expired)
No | Visit | Complaint | RDEU | Anwender haben immer wieder Probleme mit der Bedienung des Bildschirms. das ist schon länger und sind dadurch genervt. | Monitor ersetzt. | Startphase | SWO-233083 | Defect | HW | DMS_PC | Touch screen | W firmware | Users have always had problems with the operation of the screen. this has been a long time and is therefore annoyed. | Monitor replaced. | Start
No | Visit | Complaint | RDEU | 12.09.2017, Herr Knorr, tel: 015254847560: ABL bleibt nach Neustart auf der Windows Ebene hängen 19.09.2017; Hr. Knorr; Tel.: 015254847560; der ABL bleibt erneut beim Booten hängen | erneut Windows Fehlermeldung beim laden File: $Windows $system32 $config $system vermutlich Festplatte defekt | Festplatte erneuert Windows Software neu installiert DMS Software installiert, Gerät konfiguriert und angebunden Gerät kalibriert und überprüft | SWO-232756 | Defect | HW | DMS_PC | HDD |  | 12.09.2017, Mr. Knorr, tel: 015254847560: ABL gets stuck on the Windows level after reboot 19.09.2017; Mr. Knorr; Tel.: 015254847560; the ABL gets stuck again when booting | Windows error message when loading again File: /Windows /system32 /config /system probably hard drive defective | Hard drive renewed Windows software reinstalled DMS software installed, device configured and connected Device calibrated and checked
Yes | Visit | Complaint | RLTD | Barcode reader replacement | Barcode reader replacement | Replaced barcode reader for tubes. Tested. Passed. | SWO-232908 | Re classified RMED | HW | Aspiration Module | Barcode reader, Inlet |  | Barcode reader replacement | Barcode reader replacement | Replaced barcode reader for tubes. Tested. Passed.
No | Visit | Complaint | RKK | 以前よりヘモグロビンの値がおかしいとユーザーから指摘があったそうです。 血液測定では混濁度が高すぎるｴﾗｰ、コンプリートテストにて、スペクトルメータノイズが外れていた。 | ｵｷｼｴﾗｰ | ﾍﾓﾗｲｻﾞｰ、ｽﾍﾟｸﾄﾙﾒｰﾀ交換しても、ｵｷｼｺﾝﾌﾟﾘｰﾄﾃｽﾄのｽﾍﾟｸﾄﾙﾒｰﾀノイズがパスしなかった。 後日、代替機対応となった。 | SWO-235346 | Defect | HW | Oxi module | Spectrometer | Mayby later visit? | It seems to have been pointed out by the user that the value of hemoglobin is more interesting than before. In the blood measurement, the error of the turbidity was too high, and the spectral meter noise was off in the complete test. | Oxy error | Even when the hemolizer and spectrum meter were replaced, the spectrum meter noise of the oxycomplete test did not pass. Later, it became an alternative machine correspondence.
No | Visit | Complaint | RDEU | 20,09,2017 Sr Christa 0213152954216 meldet, das der Analysator das LP immer wieder gewechselt haben möchte. | Das Gerät vorgefunden mit den Fehlern: Lösungen können nicht gepumpt werden, Flüssigkeitstransport Fehler, Kalibrierfehler Flow Selektor und Chipdaten des Lösungspack können nicht gelesen werden. | LP nicht richtig aktiviert - nur einseitig heruntergedrückt, LP komplett aktiviert und im Servicemenü überprüft - LP Check wird nicht bestanden. LP erneut initialisiert aber ohne Erfolg. LP erneuert. Transportfehler behoben. Startphase läuft. | SWO-234167 | Assisted Replacement | SP | SP | FT | Incorrect activation | 20,09,2017 Sr Christa 0213152954216 reports that the analyzer wants to change the LP over and over again. | The device found with the errors: solutions cannot be pumped, liquid transport errors, calibration error flow selector and chip data of the solution pack cannot be read. | LP not activated correctly - only one-sided lyre, LP completely activated and checked in the service menu - LP Check is not passed. LP reinitializes but to no avail. LP renewed. Transport error fixed. Start-up phase is underway.
No | 18 | Y | RPNZ | Blank | Failing Oximetry QC and analyzer in intervention required mode |  | D32B5DE97858CD7D | Defect | HW | Oxi module | Hemolyzer | QC | Blank | Failing Oximetry QC and analyzer in intervention required mode
No | Visit | Complaint | RDEU | Empf. Kalibrationsfehler bei PH zu niedrig | System überprüft => PH Elektrode defekt | PH Elektrode erneuert | SWO-233317 | Assisted Replacement | Electrode | Electrode | pH |  | Empf. Calibration error at PH too low | System checked => PH electrode defective | PH electrode renewed
No | Visit | Complaint | RCHN | 接口板腐蚀 | 接口板腐蚀，借902-972一个 | 更换902-972后修复by tim | SWO-232827 | [5 empty] | Interface plate corrosion | Interface plate corrosion, borrow 902-972 one | Repair by tim after replacement 902-972
No | Visit | Complaint | RDEU | Nach Eingabe der Netzwerkeingaben/ -Änderungen werden diese nach dem Bestätigen nicht gespeichert. | . | Speicherkarte getauscht | SWO-234121 | [5 empty] | After entering the network inputs/changes, these are not saved after confirmation. | . | Memory card exchanged
No | Visit | Complaint | RDEU | Nach Eingabe der Netzwerkeingaben/ -Änderungen werden diese nach dem Bestätigen nicht gespeichert. | Siehe SWO | Speicherkarte mit Win XP inst. ( 933-438 ) SP und SC inst. --> Neustart --> Netzwerkdaten eingegeben und Verbindung zu Radiance getestet --> keine Verbindung möglich Station ist informiert | SWO-233551 | [5 empty] | After entering the network inputs/changes, these are not saved after confirmation. | See SWO | Memory card with Win XP inst. ( 933-438 ) SP and SC inst. --> Restart --> Network data entered and connection to Radiance tested --> no connection possible Station is informed
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 报错1249 液体传输错误 | 更换测试卡后修复 | SWO-233206 | Assisted Replacement | SC | SC | FT | 1249 | Test card, reagent package failure | Report error 1249 liquid transfer error | Repair after replacing test card
No | Visit | Complaint | RCHN | 定标失败，卡盘 | 定标失败，卡盘 | 重新定标正常。 | SWO-233255 | No fault found | No fault found | No fault found | No fault found | Email - err code | Bid failed, card disk | Bid failed, card disk | Recalibration is normal.
No | Visit | Complaint | RIND | Patient result printout not clearly visible | Checked the machine patient result printout not clearly visible found printer head was defective | Need to replace printer head | SWO-233292 | Defect | HW | Printer Module | Printer head |  | Patient result printout not clearly visible | Checked the machine patient result printout not clearly visible found printer head was defective | Need to replace printer head
No | Visit | Complaint | RDEU | Pumpenasymetrie-Fehler | Pumpenasymetrie-Fehler im pH $BG-Modul pCO2-Flush-Fehler | pH $BG-Modul Pumpenschlauch ersetzt Probelauf | SWO-233879 | Defect | HW | pH/BG Module | Pump tubes |  | Pump asymetry error | Pump asymetry error in pH /BG module pCO2 flush error | pH $BG module pump hose replaced Rehearsal
No | Visit | Complaint | RDEU | Auftrag per Fax-Barcode Scanner sporadisch ohne Funktion/Benutzer logt sich nicht mehr aus. | Kunde hat ZBM die Abmeldezeit bei Benutzer war auf 4 Minuten eingestellt, diese habe ich auf 1,5 Minuten runtergesetzt. | Scanner für den 21.09 für vor 8:00 Uhr bestellt. | SWO-233354 | Revisit required | HW | Basic Unit | Barcode reader, Basic Unit |  | Order by fax barcode scanner sporadically without function/user no longer logs out. | Customer has ZBM the unsubscribe time for user was set to 4 minutes, this I have lowered to 1.5 minutes. | Scanner for the 21.09 ordered for before 8:00 a.m.
Yes | Visit | Complaint | RSCH | Neudefekt | QC Level 3 pCO2 sprunghaft $zu tief | Allgemeine Kontrolle i.o Messkammer ersetzt i.o Kal. und QC i.o | SWO-233580 | Re classified RMED | OOB |  |  |  | Neudefekt | QC Level 3 pCO2 jumps /zu low | General control i.o Measuring chamber replaces i.o Kal. and QC i.o
No | 4 | Y | RFRA | Blank | problèmes oxymètre récurrents depuis le 14 septembre sur calibration 2 points et cq : messages code n°1010,1011,1012 |  | D32C2E59EE32C612 | Defect | HW | Oxi module | Spectrometer | 1010 | Blank | recurrent oxymeter problems since September 14 on calibration 2 points and cq: code messages no.1010,1011,1012
No | 2 | Y | RFRA | Blank | Fermeture de la bay unit impossible |  | D32C1E48EE2882A2 | Adjustment | HW | Basic Unit | SC door |  | Blank | Closure of the bay unit impossible
No | Visit | Complaint | RDEU | Kunde hat präanalytische Probleme mit den MBU Messungen. Geräteeinweisung erforderlich. |  | Mit Frau Rusch die Problematik besprochen, Einlasssonde gerichtet, Einlasshalterung erneuert, Einweisung durchgeführt. | SWO-234315 | Defect | HW | Inlet module | Inlet gasket |  | Customer has preanalytical problems with The MBU measurements. Device instruction required. |  | Discussed the problem with Ms. Rusch, inlet probe directed, inlet holder renewed, briefing carried out.
No | Visit | Complaint | RLTD | Inlet position failure | Inlet appeared to be obstructed but no obstruction found | Removed inlet module, set position in misc service program with no module in place, fitted inlet module back in place and calibrated position ok | SWO-233784 | Defect | SW | Auto Inlet module | Inlet open |  | Inlet position failure | Inlet appeared to be obstructed but no obstruction found | Removed inlet module, set position in misc service program with no module in place, fitted inlet module back in place and calibrated position ok
Yes | 3 | Y | RPAC | Blank | S box |  | D32C1A1DEE0A2AC8 | Re classified RMED | Logistical upSWO | S box delivery |  |  | Blank | S box
Yes | Visit | Complaint | RDEU | Sporadisch Oximetriefehler 581 | Sporadisch Oximetriefehler 581 | Daten aufgenommen, nach AT der SC ist kein Fehler mehr aufgetreten, SC reklamiert | SWO-234223 | Re classified RMED | PER report |  |  |  | Sporadisch Oximetriefehler 581 | Sporadisch Oximetriefehler 581 | Data recorded, after AT the SC no more error occurred, SC complains
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 试剂包液体传输失败 | 更换试剂包后修复 | SWO-234274 | Assisted Replacement | SP | SP | FT |  | Test card, reagent package failure | Reagent pack liquid transfer failed | Repair after replacement of reagent kit
No | Visit | Complaint | RCHN | 显示器条码器不亮 | 显示器条码器不亮，需要借件902-924一个 | 工程师告知未使用备件，重启后仪器恢复正常by tim | SWO-233269 | Defect | HW | Analyzer Control Module | Barcode reader, display |  | Display barcode is not lit | The monitor barcode is not lit, requires a loan 902-924 | Engineer tells no spare parts, instrument returns to normal after restart
No | Visit | Complaint | RCHN | 处理包卡盘 | 处理包卡盘 | 更换试剂包后修复 | SWO-233985 | Defect | Consumables | Solution Pack | Solution Pack | 5684 | Handling the package chuck | Handling the package chuck | Repair after replacement of reagent kit
No | Visit | Complaint | RCHN | pco2质控偏低 | pco2质控偏低，需要加件945-612 | 更换945-612，902-446后修复by tim | SWO-233357 | Assisted Replacement | Electrode | Electrode | pCO2 |  | low quality control of pco2 | pco2 quality control is low, need to add 945-612 | Replace 945-612, 902-446 after repair by tim
No | 5 | Y | RCZE | Blank | QC dlhodobo ide na hornom limite, pre vsetky hladiny, respektive nad |  | D32C3C72EE553E38 | Assisted Replacement | Electrode | Electrode | Na+ |  | Blank | QC for the long term goes to the upper limit, for all levels, respectively over
Yes | 21 | Y | RPAC | Blank | Faulty solution pack giving various liquid transport errors. |  | D32D6AA7640C7703 | [5 empty] | Blank | Faulty solution pack giving various liquid transport errors.
Yes | 21 | Y | RPAC | Blank | Annual Service and software |  | D32D57C9643EC636 | [5 empty] | Blank | Annual Service and software
No | 4 | Y | RFRA | Blank | Erreur oximétrie sur patients et contrôles |  | D32D5A5B64223939 | Defect | HW | Basic Unit | PM kit |  | Blank | Patient-oximetry error and controls
No | Visit | Complaint | RLTD | Dave Gardner 0790 359 1896 | Checked condition of contact pins for Sensor Cassette - no obvious fault, but cleaned with alcohol as precaution. Checked alignment of pins to Sensor Cassette Bay, pins offset to the right. Adjusted pin position to central. Restarted analyser - Lactate still faulty. Further visit required on Monday 25th September with Sensor Simulator and spare parts to further troubleshoot. | Adjusted alignment of sensor cassette pins to sensor bay. | SWO-234636 | Revisit required | HW | Basic Unit | SC Contacts | Lac | Dave Gardner 0790 359 1896 | Checked condition of contact pins for Sensor Cassette - no obvious fault, but cleaned with alcohol as precaution. Checked alignment of pins to Sensor Cassette Bay, pins offset to the right. Adjusted pin position to central. Restarted analyser - Lactate still faulty. Further visit required on Monday 25th September with Sensor Simulator and spare parts to further troubleshoot. | Adjusted alignment of sensor cassette pins to sensor bay.
Yes | 1 | Y | RPAC | Blank | Upon Installation the sensor cassette would not calibrate and gave errors 0378, 0377 and 1310 Lacate Sensor Response Error |  | D32D2AD864637EBC | [5 empty] | Blank | Upon Installation the sensor cassette would not calibrate and gave errors 0378, 0377 and 1310 Lacate Sensor Response Error
No | Visit | Complaint | RAME | Customer reports leakage of fluid from under solution pack. | Found leakage at solution pack cavity. Possible seating issue with solution pack. | Cleaned entire analyzer and table top. Replaced hinge assembly and waste port. Ran multiple rinse and system cycles. Checked for additional leaks. | SWO-235564 | [5 empty] | Customer reports leakage of fluid from under solution pack. | Found leakage at solution pack cavity. Possible seating issue with solution pack. | Cleaned entire analyzer and table top. Replaced hinge assembly and waste port. Ran multiple rinse and system cycles. Checked for additional leaks.
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | Lac,K,HCT定标不过 | 连续更换2张测试卡后修复 | SWO-234715 | [5 empty] | Test card, reagent package failure | Lac, K, HCT calibration but | Repair after 2 consecutive test cards
No | Visit | Complaint | RIND | Temp-F error. | Temp-F error. | Found performed the SC installation but error remaining same, performed the Zeroing, values are above the range need to replace Sensor connector board PCB. | SWO-234514 | [5 empty] | Temp-F error. | Temp-F error. | Found performed the SC installation but error remaining same, performed the Zeroing, values are above the range need to replace Sensor connector board PCB.
No | Visit | Complaint | RCHN | 接口板腐蚀 | 接口板腐蚀，借902-972一个 |  | SWO-234303 | [5 empty] | Interface plate corrosion | Interface plate corrosion, borrow 902-972 one
No | Visit | Complaint | RCHN | 试剂包传感器错误 | 试剂包传感器错误，借902-975一个 | 使用库存机修复(By Ran)Changed 902-975, Ok. | SWO-233724 | [5 empty] | Reagent pack sensor error | Reagent pack sensor error, borrow 902-975 one | Use Stock Machine Repair (By Ran) Change 902-975, Ok.
No | Visit | Complaint | RIND | Unable to switch on the instrument. | Instrument unable to switch on. | Replace the power supply with new one ,after power on machine became ready.Hanover the instrument to the customer. | SWO-234619 | [5 empty] | Unable to switch on the instrument. | Instrument unable to switch on. | Replace the power supply with new one ,after power on machine became ready. Hanover the instrument to the customer.
No | Visit | Complaint | RCHN | 黑屏 | 黑屏，借备件903-365一个 |  | SWO-234452 | [5 empty] | Black screen | Black screen, borrowed spare parts 903-365 one
No | Visit | Complaint | RCHN | 离线状态 0734，1223 | 离线状态 0734，1223 ，借910-331，903-165各一个，现加借636-568,910-342,903-644各一个 |  | SWO-234344 | Defect | HW | DMS_PC | CPU unit | 734 DMS not connected to AU | Offline 0734,1223 | Offline 0734,1223 , borrow 910-331,903-165 each, now 636-568, 910-342, 903-644 each
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | PO2质控失败，0090 | 更换试剂包后修复 | SWO-234443 | Assisted Replacement | SP | SP | pO2 QC | 90 | Test card, reagent package failure | PO2 QC Failure, 0090 | Repair after replacement of reagent kit
No | Visit | Complaint | RKK | 起動時にシステムエラーとなり、エラー表示が解除されない現象 | ケース詳細の通り | 引上げ後再現性なし、TST秋山さんに確認してもらうも再現性はなし。立ち上げ時のシステムチェックで何かが引っかかっていたと推測に至る。 | SWO-237911 | [5 empty] | System error occurs at startup and the error is not displayed. | As the case details | There is no reproducibility after the hike, and there is no reproducibility though it is confirmed by Mr. Akiyama. It leads to the guess that something was caught in the system check at the time of the launch.
No | Visit | Complaint | RCHN | 报错5766 | 报错5766，借903-198一个 |  | SWO-234264 | Defect | HW | Analyzer Control Module | PCB, Main Board | 5766 | 5766 error | 5766, loan 903-198 one
No | Visit | Complaint | RESP | PO2 | Comprobación de registros de datos del equipo. | Comprobación del estado de la cámara de medida, sustitución de la membrana del Cl, se sustituye la cámara de medida por sospecha de posible defecto, e sustituye el Inlet por estar defectuoso o doblado, de calibra el equipo y verificamos el correcto funcionamiento con QC. | SWO-234537 | Defect | HW | pH/BG Module | Measuring chamber, pO2 | pO2 QC | PO2 | Checking computer data records. | Checking the condition condition of the measuring chamber, replacing the Cl membrane, replacing the measuring chamber on suspicion of possible defect, and replacing the Inlet with being defective or bent, calibrates the equipment and verifies the correct operation with QC.
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | LAC定标不过 | 更换新机器 | SWO-234631 | [5 empty] | Test card, reagent package failure | LAC calibration is not | Replace a new machine
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | pO2质控不过 0090 | 更换试剂包后修复 | SWO-235223 | Assisted Replacement | SP | SP | pO2 QC | 90 | Test card, reagent package failure | pO2 QC but 0090 | Repair after replacement of reagent kit
No | Visit | Complaint | RCHN | pO2定标不过 | NA,GLU,LAC定标不过，0375,0377,0378 | 更换测试卡后修复 | SWO-234693 | Assisted Replacement | SC | SC | pO2 | Na, Glu, Lac | pO2 calibration but | NA, GLU, LAC calibration,0375,0377,0378 | Repair after replacing test card
No | Visit | Complaint | RDEU | Seit einiger Zeit gibt es Oxi Störungen am ABL90 die QC und Kal Störungen auslösen. Bitte Prüfen. FSE Einsatz | Kal,QCs und tHb bringen dieselbe Fehler Meldungen: Oxi Hardware Fehler, Fortlaufende Spülung unterbrochen. Service Tests durchgeführt, alles i.O. Oxi complete Test, alles i.O, Fliessweg geprüft, gespült, alles i.O. Fehler bestehend. Fehler traten auf seit dem Zeitpunkt wo LP getauscht wurde. Kein Ersatz LP vor Ort. Den Kunden gebeten neunen LP bei erhalt zu installieren. Falls die Probleme weiter bestehen sollten, wird sich der Kunde melden. | Kunden informiert. | SWO-235331 | Unknown | HW | Oxi module | Oxi module |  | For some time there have been Oxi faults on the ABL90 which trigger QC and Kal faults. Please check. FSE Use | Kal,QCs and tHb bring the same error messages: Oxi hardware errors, continuous flushing interrupted. Service tests carried out, all i.O. Oxi complete test, all i.O, flow path tested, rinsed, all i.O. error existing. Errors occurred since the time LP was swapped. No replacement LP on site. Asked the customer to install nine LP upon receipt. If the problems persist, the customer will get in touch. | informed customers.
No | 3 | Y | RPOL | Blank | Regulacja mechanizmu drzwi kasety sensorowej. Testy serwisowe - ok. Analizator sprawny technicznie. |  | D3313108C61423C1 | Adjustment | HW | Basic Unit | Sensor cassette Door |  | Blank | Adjustment of the sensor cassette door mechanism. Service Tests-OK. Technically efficient analyzer.
No | Visit | Complaint | RLTD | Dave Gardner 0790 359 1896 | Calibration still failing on lactate with sensitivity out of limits. Two SC already in stalled. | Installed new sensor interface PCB and after calibration still having same error for lactate . SP replaced and after calibration lactate is failing. Traced fault to 3rd SC being faulty. Refitted original interface PCB and new SC calibration and QC ok. | SWO-235080 | Assisted Replacement | SC | SC | SC | Lac | Dave Gardner 0790 359 1896 | Calibration still failing on lactate with sensitivity out of limits. Two SC already in stalled. | Installed new sensor interface PCB and after calibration still having same error for lactate . SP replaced and after calibration lactate is failing. Traced fault to 3rd SC being faulty. Refitted original interface PCB and new SC calibration and QC ok.
No | Visit | Complaint | RKK | 患者結果の送信が遅い | 同時にチケットプリンタへ出力している為、遅くなっていました。 | チケットに送らないように設定し完了 | SWO-237172 | [5 empty] | Slow transmission of patient results | Because i output it to the ticket printer at the same time, it was late. | Set not to send to ticket and complete
No | Visit | Complaint | RIND | Sodium results have variation. | Checked SC and cleaned it. Reinstalled Sc and checked with Quality control. | Qc values are out of range for Sodium.Need to change SC. | SWO-237325 | [5 empty] | Sodium results have variation. | Checked SC and cleaned it. Reinstalled Sc and checked with Quality control. | Qc values are out of range for Sodium.Need to change SC.
No | Visit | Complaint | RCHN | 机器无法开机 | 机器无法开机，需要借件910-246一个 |  | SWO-234849 | [5 empty] | The machine does not power on | The machine cannot be switched on, need to borrow 910-246 one
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | GLU灵敏度超范围 | 更换测试卡后修复 | SWO-235163 | Assisted Replacement | SC | SC | Glu | Sensitivity | Test card, reagent package failure | GLU sensitivity over range | Repair after replacing test card
No | Visit | Complaint | RCHN | 血氧单元故障 | 血氧单元故障,需要借件902-980和903-197各一个 |  | SWO-234881 | Defect | HW | Oxi module | Spectrometer | & hemolyzer | Blood oxygen unit failure | Blood oxygen unit failure, need to borrow 902-980 and 903-197 each one
No | Visit | Complaint | RLTD | Total software crash. Needs help to reboot/fix. Mrs Hanna McPeak - 01865 221 596 / 07887 428 417 | The analyser software was not running on arrival so I rebooted the analyser and when the analyser software started to load it failed due to a status code 9 error which is a missing Roma.ini file. When I looked at the file the Roma.ini was there but it contained no data. I re-installed an old version of the Roma.ini file and then there was a problem with the astm.ini file. I copied this file from another analyser and after that I encountered errors due to problems with the astm6xxini & hl7.ini files. These were copied from another analyser but the software still wouldn't start up. | Installed the latest version of software and this re-installed the corrupted files and the analyser started up ok. There was an inlet calibration error when the analyser had started up and I found that the inlet probe was bent. A new probe and a new inlet gasket (this was cracked) were fitted and this resolved the inlet calibration error. | SWO-235078 | Defect | SW | RMED SW | RMED SW | HDD defective ? | Total software crash. Needs help to reboot/fix. Mrs Hanna McPeak - 01865 221 596 / 07887 428 417 | The analyser software was not running on arrival so I rebooted the analyser and when the analyser software started to load it failed due to a status code 9 error which is a missing Roma.ini file. When I looked at the file the Roma.ini was there but it contained no data. I re-installed an old version of the Roma.ini file and then there was a problem with the astm.ini file. I copied this file from another analyser and after that I encountered errors due to problems with the astm6xxini & hl7.ini files. These were copied from another analyser but the software still wouldn't start up. | Installed the latest version of software and this re-installed the corrupted files and the analyser started up ok. There was an inlet calibration error when the analyser had started up and I found that the inlet probe was bent. A new probe and a new inlet gasket (this was cracked) were fitted and this resolved the inlet calibration error.
No | Visit | Complaint | RCHN | 血氧无法测氧 | 血氧无法测氧，借902-527一个 |  | SWO-234811 | Defect | HW | Oxi module | Pinch valve |  | Blood oxygen can't be oxygenated. | Blood oxygen can not be measured, borrow 902-527 one
No | Visit | Complaint | RSCH | Stampante difettosa | Stampante difettosa. | Sostituita stampante e controllata. | SWO-235130 | Defect | HW | Printer Module | Printer head |  | Faulty printer | Faulty printer. | Printer and controlled replaced.
No | Visit | Complaint | RDAN | Inlet fejl | NA | Jeg har skiftet probeholder for at se som dette ville ændre på kalibreringsdate fra Inlet. Men dette har ikke ændret noget. Alt ser fint ud. | SWO-236784 | Defect | HW | Inlet module | Inlet gasket | probe holder replaced?? | Inlet fejl | In | I changed the Procontainer to see as this would change the calibration date from Inlet. But this has not changed anything. Everything looks fine.
No | Visit | Complaint | RAME | Customer reported by text, analyzer battery not showing any charge. Reading 0% | Checked battery condition. Showing 0% charge in main screen and diagnostics. Power down will not reset. | Replaced battery module. Test operation. | SWO-235789 | Defect | HW | Battery module | Battery |  | Customer reported by text, analyzer battery not showing any charge. Reading 0% | Checked battery condition. Showing 0% charge in main screen and diagnostics. Power down will not reset. | Replaced battery module. Test operation.
No | Visit | Complaint | RDEU | 26.09.2017 12:50Uhr 0401818854442: Nach dem Tausch des LPs zeigt der Bildschirm eine Sanduhr und Intervention erforderlich | System überprüft => LP nicht richtig installiert | LP erneut installiert ,Kalibration und QC durchgeführt | SWO-235556 | User error | Installation | SP Installation | SP Installation | LEH | 26.09.2017 12:50h 0401818854442: After swapping the PL, the screen shows an hourglass and intervention required | System checked => LP not installed correctly | LP reinstalled ,calibration and QC performed
No | 7 | Y | RFRA | Blank | pO2 sur les CQ externe mauvais |  | D332652E3C4CB772 | Defect | HW | Basic Unit | PM kit |  | Blank | pO2 on bad external CQ
No | Visit | Complaint | RSCH | Ago di aspirazione bloccato. | Ago di aspirazione bloccato. | Ago di aspirazione piegato di conseguenza sostituito, installata nuova versione Software 3.3 MR1. Sostituita SP e SC apparecchio ricalibrato. | SWO-235692 | Replaced | HW | Auto Inlet module | Probe bend |  | Shalction needle blocked. | Shalction needle blocked. | Suction needle bent accordingly replaced, installed new version Software 3.3 MR1. Replaced SP and SC recalibrated appliance.
Yes | 4 | Y | RFRA | Blank |  |  | D332AC873C5D4467 | Re classified RMED | FAN 915-371 |  |  |  | Blank
Yes | 1 | Y | RFRA | Blank |  |  | D3328F6C3C5AB778 | Re classified RMED | FAN 915-371 |  |  |  | Blank
No | Visit | Complaint | RCHN | 试剂包芯片无法读取 | 试剂包芯片无法读取，需要借件903-158备件一个 |  | SWO-235168 | Defect | HW | Basic Unit | SP ID PCB |  | Reagent package chip cannot be read | Reagent package chip can not be read, need to borrow 903-158 spare parts one
No | Visit | Complaint | RCHN | 进样针皮带断了 | 进样针皮带断了，借867-631一个。 | 更换867-631备件后修复(By Ran) | SWO-235379 | Defect | HW | Aspiration Module | Trimming belt |  | The in-the-needle belt is broken. | The intake needle belt is broken, borrowing 867-631. | Repair after replacement of 867-631 spare parts (By Ran)
No | Visit | Complaint | RESP | NO LEE CODIGOS DE BARRAS | MALFUNCION DEL LECTOR | SUSTITUCION DE LECTOR DE CODIGO DE BARRAS Y PRUEBAS DE FUNCIONAMIENTO. APROVECHANDO LA VISITA SE DA FORMACION A NUEVOS USUARIOS DEL LABORATORIO TRAS LAS VACACIONES. CALIBRADO Y VERIFICADO CON Q.C. | SWO-235623 | Defect | HW | Basic Unit | Barcode reader, Basic Unit |  | DO NOT READ BAR CODES | READER MALFUNCTION | REPLACEMENT OF BARCODE READER AND OPERATION TESTS. TAKING ADVANTAGE OF THE VISIT IS PROVIDED TO NEW LABORATORY USERS AFTER THE HOLIDAYS. CALIBRATED AND VERIFIED WITH Q.C.
No | Visit | Complaint | RAME | Liquid in pH chamber | Checked data logs for errors, checked valves and liquid sensors | Removed and modules and cleaned valves, removed and replaced pH electrode, | SWO-237947 | Cleaned | Fluidic pathway | pH/BG Module | Measuring chamber |  | Liquid in pH chamber | Checked data logs for errors, checked valves and liquid sensors | Removed and modules and cleaned valves, removed and replaced pH electrode,
No | Visit | Complaint | RAME | Tried different lots of membranes and verified QC range | Lactate out on S7755 QC. Tried different lot of membranes-no change | Replaced the lactate electrode. Values slightly better, customer will monitor | SWO-237881 | Assisted Replacement | Electrode | Electrode | Lac |  | Tried different lots of membranes and verified QC range | Lactate out on S7755 QC. Tried different lot of membranes-no change | Replaced the lactate electrode. Values slightly better, customer will monitor
No | Visit | Complaint | RAME | PO2 is on the higher range on WDC report. | PO2 is high on WDC report. This has been high since the install of the analyzer in December 2016. | Cleaned electrode and chamber. Performed zero channel in service diagnostic. Still the same issue. Remembraned PO2 with a different lot#. Still the same issue. All tests are within manufacturers specifications. | SWO-235722 | No fault found | Performance | Ext Specs compromised, pCO2 | Ext Specs compromised, pCO2 |  | PO2 is on the higher range on WDC report. | PO2 is high on WDC report. This has been high since the install of the analyzer in December 2016. | Cleaned electrode and chamber. Performed zero channel in service diagnostic. Still the same issue. Remembraned PO2 with a different lot#. Still the same issue. All tests are within manufacturers specifications.
No | Visit | Complaint | RAME | PO2 is on higher end of WDC report. | Checked data logs and PO2 is high 151 in comparison to peer group.(packet inserts 135-155). Barometric pressure is 763 mmHg and temperatures is 22 degrees C. Lot # of PO2 electrode YM-02. | Cleaned PO2 electrode chamber. Performed zero check on electrode channels. Replaced another PO2 gas cylinder with a different lot number. Remembraned PO2 with a different lot#. Level 1 QC still reads 151. Would open a Techline case for further investigations. All tests are within manufacturer's specifications. | SWO-235721 | No fault found | Performance | Ext Specs compromised, pCO2 | Ext Specs compromised, pCO2 |  | PO2 is on higher end of WDC report. | Checked data logs and PO2 is high 151 in comparison to peer group. (packet inserts 135-155). Barometric pressure is 763 mmHg and temperatures is 22 degrees C. Lot # of PO2 electrode YM-02. | Cleaned PO2 electrode chamber. Performed zero check on electrode channels. Replaced another PO2 gas cylinder with a different lot number. Remembraned PO2 with a different lot#. Level 1 QC still reads 151. Would open a Techline case for further investigations. All tests are within manufacturer's specifications.
No | Visit | Complaint | RDEU | 26.09.2017 14:10Uhr Herr Kempf 06021324272: Das Gerät wollte einen vorzeitigen Wechsel des LPs. Nach dem Wechsel bleibt das Gerät im Fehlersuchverfahren hängen. | Bei Ankunft: Intervention erforderlich, Pumpen Kal Fehler, Flüssigkeitstransport Fehler, Inhomogene Lösungen, LS Fehler, Drucküberprüfung Fehler, usw. LP nicht richtig erkannt. LP läuft in 2 Tagen ab. | Gerät Rückwärts gespült (Gerinnsel beseitigt), LP Sensoren gereinigt, LP neu installiert, nach Initialisierungsphase pCO2, Glu und Lac Kal Fehler aufgrund das die fortlaufende Spülung unterbrochen wurde. Mit Anwender besprochen neunen LP und neue SK zu installieren. Beides installiert. Alles i.O. Kunden Informiert. | SWO-236251 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | 26.09.2017 14:10Mr Kempf 06021324272: The device wanted an early change of the LPs. After the change, the device gets stuck in the error search procedure. | On arrival: intervention required, pumping kal errors, fluid transport errors, inhomogeneous solutions, LS errors, pressure verification errors, etc. LP not recognized correctly. LP expires in 2 days. | Device reflushed (clot removed), LP sensors cleaned, LP reinstalled, after initialization phase pCO2, Glu and Lac Kal error due to which the continuous flushing was interrupted. Discussed with users nine LP and new SK to install. Both installed. Everything i.O. Customers Informed.
No | 6 | Y | RFRA | Blank | MESURES ANNULEES $ PAS D'ASPIRATION ECHANTILLON |  | D33407F8B25658A7 | Defect | HW | Auto Inlet module | Inlet gasket |  | Blank | RESORED MESURES / NO ASPIRATION ECHANTILLON
No | Visit | Complaint | RESP | No calibra el Cloro (parámetro desactivado) Fallan los controles de calidad. | No calibra el CL (parámetro desactivado) Fallan los controles de calidad internos. | Se realizó kit de tubos anual, se desmontó y limpió manifold, se ajustó cero del equipo, se sustituyó casete de electrodos defectuoso. se calibró equipo y se comprobó correcto funcionamiento mediante ciclos de calidad. | SWO-236296 | [5 empty] | Do not calibrate Chlorine (parameter off) Quality controls fail. | Do not calibrate the CL (parameter off) Internal quality controls fail. | Annual pipe kit was made, disassembled and cleaned manifold, zero equipment adjusted, defective electrode cassette replaced. Equipment was graded and proper operation was proven through quality cycles.
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | Lac定标不过 | 更换测试卡后修复 | SWO-236422 | [5 empty] | Test card, reagent package failure | Lac calibration is not | Repair after replacing test card
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 测试卡参比电极错误 1361 | 更换测试卡后修复 | SWO-236606 | Assisted Replacement | SC | SC | Reference | 1361 | Test card, reagent package failure | Test card reference electrode error 1361 | Repair after replacing test card
No | Visit | Complaint | RAME | AutoCheck not running QC, Liquid in reference chamber | Checked data logs for errors, checked valves and liquid sensors, checked AutoCheck module | Removed modules and cleaned valves, replaced ripped air valve membrane, removed AutoCheck module and replaced FPC cable. | SWO-237950 | Defect | HW | pH/BG Module | Valve, air | Membrane | AutoCheck not running QC, Liquid in reference chamber | Checked data logs for errors, checked valves and liquid sensors, checked AutoCheck module | Removed modules and cleaned valves, replaced ripped air valve membrane, removed AutoCheck module and replaced FPC cable.
No | 10 | Y | RPNZ | Blank | Oxi failure (Phaseregulation error on Hemolyzer) |  | D33394BEB24EB5F5 | Defect | HW | Oxi module | Hemolyzer | phase regulation | Blank | Oxi failure (Phaseregulation error on Hemolyzer)
No | Visit | Complaint | RAME | PCO2 QC results are outside of range | PCO2 Ag $Cl coating is diminished and QC result have trended up out of control | Replaced CO2 electrode. Calibrated and ran QC,instrument operates normally | SWO-235699 | Assisted Replacement | Electrode | Electrode | pCO2 | AgCl Coating diminished | PCO2 QC results are outside of range | PCO2 Ag /Cl coating is diminished and QC result have trended up out of control | Replaced CO2 electrode. Calibrated and ran QC,instrument operates normally
No | Visit | Complaint | RLTD | Mory Tel 01865 221596 | It was reported that the barcode scanner would read ID badges but would not scan the bottles or membrane boxes. I checked the lens of the barcode scanner and you could see that there had been a leak over the scanner in the past. | Cleaned the barcode scanner lens to clear the signs of a leak. After that removed all the bottles and gas cylinders and rescanned them in successfully, I was also able to use the scanner to log on using the barcode on my badge. | SWO-236201 | Cleaned | HW | Basic Unit | Barcode reader, Basic Unit |  | Mory Tel 01865 221596 | It was reported that the barcode scanner would read ID badges but would not scan the bottles or membrane boxes. I checked the lens of the barcode scanner and you could see that there had been a leak over the scanner in the past. | Cleaned the barcode scanner lens to clear the signs of a leak. After that removed all the bottles and gas cylinders and rescanned them in successfully, I was also able to use the scanner to log on using the barcode on my badge.
No | 5 | Y | RFRA | Blank | inlet se tord régulièrement |  | D333D434B21EA05C | Defect | HW | Inlet module | Probe bend | Motor replaced | Blank | inlet twists regularly
No | Visit | Complaint | RBNL | Na vervangen van een testcassette blokkeert het storagewheel. | Na vervangen van een testcassette (TnI) blokkeert het storagewheel. Onderzoek laat zien dat de nieuw geplaatste cassette de blokkade veroorzaakt. Het blijkit dat de lipjes van de cassette teveel naar binnen staan, wat zorgt voor de blokkade. Cassette verwijdert, en van hetzelfde lotnummer een nieuwe cassette geplaatst (controle vooraf laat zien dat deze lipjes goed zijn), AQT werkt nu weer normaal. | Testcassette verwijdert. | SWO-236066 | Defect | Consumables | Test Kit | Cartridge |  | After replacing a test cassette, the Storagewheel is blocked. | After replacing a test cassette (TnI) it blocks the Storagewheel. Research shows that the newly placed cassette is causing the blockage. The Happy kit that the tabs of the cassette are too inwards, which ensures the blockade. Cassette, and from the same lot number a new cassette is inserted (check in advance shows that these tabs are good), AQT now works normally again. | Test cassette.
No | Visit | Complaint | RHUN | DEMO közben felszívótű tartót kell cserélni | - | felszívótű csere, kalibráció, QC mérés | SWO-237411 | Defect | HW | Inlet module | Inlet gasket |  | You need to replace the suction needle holder during DEMO | - | Suction needle replacement, calibration, QC Measurement
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | Lac定标不过 | 更换测试卡后修复 | SWO-236447 | [5 empty] | Test card, reagent package failure | Lac calibration is not | Repair after replacing test card
No | Visit | Complaint | RESP | O cliente queixa-se que o equipamento está constantemente em estabilização e não calibra. Foi substituída a membrana e o problema mantem-se. Foi testado o eléctrodo no TCMTosca e também não calibra e foi testado outro eléctrodo neste equipamento e funciona corretamente, deduz-se então que o problema é do eléctrodo. Vou fazer o pedido de um novo em garantia. | Não calibra | Substituição do eléctrodo (Foi retirado o SN92-67797 e instalado o SN92-78309). Instalação de membrana. Calibração e verificação do correcto funcionamento. | SWO-236635 | [5 empty] | The customer complains that the equipment is constantly stabilizing and does not calibrate. The membrane has been replaced and the problem is maintained. The electrode has been tested in the TCMTosca and it also does not calibrate and another electrode has been tested in this equipment and works correctly, then deduces that the problem is the electrode. I'm going to order a new one in warranty. | No Calibra | Replacing the electrode (SN92-67797 was removed and the SN92-78309 was installed). Membrane installation. Calibration and verification of the correct operation.
No | Visit | Complaint | RESP | O cliente queixa-se que o equipamento está constantemente em estabilização e não calibra. Foi substituída a membrana e o problema mantem-se. Foi testado o eléctrodo no TCMTosca e também não calibra e foi testado outro eléctrodo neste equipamento e funciona corretamente, deduz-se então que o problema é do eléctrodo. Vou fazer o pedido de um novo em garantia. | Não calibra | Substituição do eléctrodo (Foi retirado o SN92-67800 e instalado o SN92-78187). Instalação de membrana. Calibração e verificação do correcto funcionamento. | SWO-236637 | [5 empty] | The customer complains that the equipment is constantly stabilizing and does not calibrate. The membrane has been replaced and the problem is maintained. The electrode has been tested in the TCMTosca and it also does not calibrate and another electrode has been tested in this equipment and works correctly, then deduces that the problem is the electrode. I'm going to order a new one in warranty. | No Calibra | Replacing the electrode (SN92-67800 was removed and the SN92-78187 was installed). Membrane installation. Calibration and verification of the correct operation.
No | Visit | Complaint | RCHN | BNP温度传感器损坏，显示不存在 | BNP温度传感器损坏，显示不存在，需要借件905-908备件一个 | 更换905-908一个正常。(By Ran) | SWO-236000 | Defect | HW | Process Module | BNP Temp Sensor |  | BNP temperature sensor damaged, shows no | BNP temperature sensor damage, display does not exist, need to borrow 905-908 spare one | Replace 905-908 one normal. (By Ran)
No | Visit | Complaint | RCHN | 板子坏了 2168 | 板子坏了 2168，借903-138一个。 | 更换903-138一个正常。(By Ran) | SWO-235851 | Defect | HW |  | PCB, Incubator | 2168 | The board is broken 2168. | The board broke 2168, borrowing 903-138. | Replace 903-138 one normal. (By Ran)
No | 1 | Y | RFRA | Blank |  |  | D3351BB32859432A | [5 empty] | Blank
No | Visit | Complaint | RDEU | Für tHb wird weder in der Kalibrierung 1&2 noch in der QC ein Wert ermittelt. In den Patientenmessungen ebenfalls keine Werte Ermittlung ---- . Oximetrie Hardwarefehler und Oxi Water Spectrum intensity unsatisfactory. | Für tHb wird weder in der Kalibrierung 1&2 noch in der QC ein Wert ermittelt. In den Patientenmessungen ebenfalls keine Werte Ermittlung ---- . Oximetrie Hardwarefehler und Oxi Water Spectrum intensity unsatisfactory. | Die Lampenspannung im Hemolyzer steht auf dem maximal Wert 4095. Lampenspannung abgeglichen. Kalibriert und QC's durchgeführt. | SWO-236726 | Adjustment | HW | Oxi module | Lamp unit |  | For tHb, neither calibration 1&2 nor QC determines a value. Also no values ---- . in the patient measurements Oximetry hardware failure and Oxi Water Spectrum intensity unsatisfactory. | For tHb, neither calibration 1&2 nor QC determines a value. Also no values ---- . in the patient measurements Oximetry hardware failure and Oxi Water Spectrum intensity unsatisfactory. | The lamp voltage in the Hemolyzer is adjusted to the maximum value 4095. Lamp voltage is matched. Calibrated and QC's performed.
No | Visit | Complaint | RKK | 測定中止される | ﾊﾟｯｸを交換してからｴﾗｰが出始めたが、都度再起動して測定可能となっていた。 | 溶液パック‐プロセスﾎｲｰﾙ間に廃液カップが詰まっていた。 溶液パック交換にて良好。 | SWO-237941 | Defect | Consumables | Solution Pack | Solution Pack | Stack | The measurement is stopped. | After replacing the pack, an error began to appear, but it was possible to measure it by restarting each time. | Solution pack - the waste water cup was jammed between the process wheels. Good in solution pack exchange.
No | Visit | Complaint | RHUN | SC csere után GLU nem áll be, elég sűrűn 1066-os hiba jön: A hőmérséklet a szenzorkazetta szubsztrátumán kívül esik | szerviz tesztek elvégzése | SC csere, kalibrációk, QC mérés | SWO-237394 | Assisted Replacement | SC | SC | Thermistor |  | After SC replacement GLU does not stand, quite densely 1066 error is coming: temperature is outside the sensor cartridge substrate | Performing service Tests | SC Exchange, calibrations, QC measurement
No | Visit | Complaint | RAME | Customer advised that nothing on display, just blue background, attempted reboot, issue not resolved. | OSPD found analyzer frozen in windows, unable to load ROMA, unable to interact with keyboard. Suspect corrupted CF card | Replaced CF card, reloaded windows, reloaded Roma. All tested ok. FSE to wait for site to finish configuration in ramp up to go-live before returning to complete configuration. | SWO-237627 | Defect | HW | DMS_PC | CF-Card |  | Customer advised that nothing on display, just blue background, attempted reboot, issue not resolved. | OSPD found analyzer frozen in windows, unable to load ROMA, unable to interact with keyboard. Suspect corrupted CF card | Replaced CF card, reloaded windows, reloaded Roma. All tested ok. FSE to wait for site to finish configuration in ramp up to go-live before returning to complete configuration.
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | po2质控低于范围 | 更换试剂包后修复 | SWO-237165 | Assisted Replacement | SP | SP | pO2 QC | 1373 | Test card, reagent package failure | po2 QC below range | Repair after replacement of reagent kit
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 定标错误和电极错误 | 更换测试卡后修复 | SWO-237158 | Assisted Replacement | SC | SC | Unknown |  | Test card, reagent package failure | Calibration errors and electrode errors | Repair after replacing test card
No | Visit | Complaint | RCHN | 血氧浑浊度高，OXI光谱不匹配 0292，0581，0291 | 血氧浑浊度高，OXI光谱不匹配 0292，0581，0291，借903-197，902-980各一个。 | 备件今天才收到，这台机器拿回代理商公司仓库维修。拿了一台新机器在医院用。 by jesson | SWO-235853 | Defect | HW | Oxi module | Hemolyzer |  | High blood oxygen turbidity, OXI spectral mismatch 0292,0581,0291 | The oxygen turbidity is high, and the OXI spectra do not match 0292,0581,0291, borrowing 903-197,980 each. | The spare parts were only received today, and the machine was taken back to the dealer's warehouse for repair. Take a new machine for use in the hospital. by jesson
No | 4 | Y | RPAC | Blank | Gas leaking |  | D33623AE9E05F558 | [5 empty] | Blank | Gas leaking
No | Visit | Complaint | RCHN | 定标失败 | 定标失败，借944-214（批号13099）缺货 | 重新定标后修复完成(By Ran) Don't have the Cal cartridge lot 13099. Case closed. | SWO-235995 | Assisted Replacement | Consumables | Test Kit | Cartridge, Myo | 1061 | Bid failed | Failed to bid, out of stock by 944-214 (lot number 13099) | 重新定标后修复完成(By Ran) Don't have the Cal cartridge lot 13099. Case closed.
No | Visit | Complaint | RSAF | Lab manager called to inform us of error - Oximetry hardware error not possible to measure and the patient results are printing with Red question marks on the results and he has tried everything, | Found analyzer in replacement, Entered Service Program and done complete Oxi Test - all passed. Exited Service program and restarted from replacement screen, Checked Activity schedule and found calibration errors on pH and electrolytes on Status and Sense with erratic drifts since 14H00 on the 26 $09 $2017 after membrane replacement was performed. Activity log also had pH $BG asymmetry pump errors. Removed membrane housing cover and saw that the was fluid leaking from the pH $BG module's drip tray as well. | Replaced pH electrode sealing ring and cleaned all the chambers, done pump calibrations all ok, Liquid sensor adjustment passed, done 4 x two point calibrations and had to replace Reference membrane due to calibrations not wanting to stabilize, New Reference membrane installed and calibrated 4x and all ok, all Qc's ok. Error was not what customer had reported. | SWO-237632 | Defect | Electrode | Electrode | Sealing ring | Sealing ring, pH | Lab manager called to inform us of error - Oximetry hardware error not possible to measure and the patient results are printing with Red question marks on the results and he has tried everything, | Found analyzer in replacement, Entered Service Program and done complete Oxi Test - all passed. Exited Service program and restarted from replacement screen, Checked Activity schedule and found calibration errors on pH and electrolytes on Status and Sense with erratic drifts since 14H00 on the 26 /09 /2017 after membrane replacement was performed. Activity log also had pH /BG asymmetry pump errors. Removed membrane housing cover and saw that the was fluid leaking from the pH /BG module's drip tray as well. | Replaced pH electrode sealing ring and cleaned all the chambers, done pump calibrations all ok, Liquid sensor adjustment passed, done 4 x two point calibrations and had to replace Reference membrane due to calibrations not wanting to stabilize, New Reference membrane installed and calibrated 4x and all ok, all Qc's ok. Error was not what customer had reported.
No | Visit | Complaint | RAME | Error 579 - Oximetry Hardware problem Error 316 - Neon number too low | Reviewed data logs and performed service diagnostics. Observed error code 258 - neon lamp intensity too low, 316 - neon number too low. Hemolyzer failed due to phase regulation error. | Replace lamp unit to resolve the neon lamp failure, removed and reseated fiber optic cable between hemolyzer and spectrometer to resolve the hemolyzer failure. Reran Oxi Complete test several times, no further failures. | SWO-237091 | Defect | HW | Oxi module | Lamp unit | Neon Lamp | Error 579 - Oximetry Hardware problem Error 316 - Neon number too low | Reviewed data logs and performed service diagnostics. Observed error code 258 - neon lamp intensity too low, 316 - neon number too low. Hemolyzer failed due to phase regulation error. | Replace lamp unit to resolve the neon lamp failure, removed and reseated fiber optic cable between hemolyzer and spectrometer to resolve the hemolyzer failure. Reran Oxi Complete test several times, no further failures.
No | Visit | Complaint | RDEU | Level 3 pH QC Verletzung. | Messkammer und pH Elektrode gereinigt, ABL ein Kalibriert. Weiterhin QC Verletzung Level 3. pH Elektrode ersetzt. Fehler bleibt bestehen. Kalibrationslösung 1 ersetzt LOT UT04 | Startphase QC und Kalibration | SWO-237544 | Assisted Replacement | Consumables | Solution bottle | Cal1 | Lot UT04 | Level 3 pH QC Verletzung. | Measuring chamber and pH electrode cleaned, ABL a calibrated. Further QC violation level 3. pH electrode replaced. Errors remain. Calibration solution 1 replaces LOT UT04 | Start phase QC and calibration
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | HCT,LAC定标失败 | 更换测试卡正常。 | SWO-237741 | [5 empty] | Test card, reagent package failure | HCT, LAC calibration failed | The replacement test card is OK.
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | LAC,GLU定标失败 | 更换测试卡 | SWO-240794 | Assisted Replacement | SC | SC | Glu/Lac |  | Test card, reagent package failure | LAC, GLU bid failed | Replace the test card
No | Visit | Complaint | RKK | インレットから液漏れ | インレットから液漏れ | インレットガスケット交換、排液回路洗浄 | SWO-237928 | Defect | HW | Inlet module | Inlet gasket |  | Liquid leak from inlet | Liquid leak from inlet | Inlet gasket replacement, drainage circuit cleaning
No | Visit | Complaint | RDEU | Keine Funktion | Gerät hat keine Funktion | Sicherungen Netzteil überprüft i.O., Netzteil ersetzt, Kal und QC durchgeführt | SWO-237782 | Defect | HW | DMS_PC | Power Supply |  | No function | Device has no function | Fuses power supply checked i.O., power supply replaced, Kal and QC performed
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | Lac定标不过 | 更换测试卡后修复 | SWO-241412 | Assisted Replacement | SC | SC | Lac |  | Test card, reagent package failure | Lac calibration is not | Repair after replacing test card
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | Lac定标不过 0210 | 更换测试卡后修复 | SWO-240780 | Assisted Replacement | SC | SC | Lac |  | Test card, reagent package failure | Lac calibration is no more than 0210 | Repair after replacing test card
No | Visit | Complaint | RLTD | Leak in analyser detected 0963 Martin Berry 01324 566151 | Leak detected reported out of hours, customer already replaced SC and SP | Arrived on site and customer had alter au replaced inlet probe and gasket at 07:45 Pressed test again to leave intervention screen and analyser failed leakage current again, new SP fitted Left replacement probe and gasket | SWO-238081 | Assisted Replacement | SP | SP | Leak current | 963, could be leak from tube | Leak in analyser detected 0963 Martin Berry 01324 566151 | Leak detected reported out of hours, customer already replaced SC and SP | Arrived on site and customer had alter au replaced inlet probe and gasket at 07:45 Pressed test again to leave intervention screen and analyser failed leakage current again, new SP fitted Left replacement probe and gasket
No | Visit | Complaint | RAME | Multiple SP's having the same issue | Error code 0954 leak detected. Went thru all fluid movement service program | Replaced the inlet gasket with holder. Will talk about not cleaning the inlet gasket with customer | SWO-238406 | Assisted Replacement | SP | SP | Leak current | 964 | Multiple SP's having the same issue | Error code 0954 leak detected. Went thru all fluid movement service program | Replaced the inlet gasket with holder. Will talk about not cleaning the inlet gasket with customer
Yes | Visit | Complaint | RAME | Multiple SP's having the same issue | Customer has changed the SP numerous times. Always happens right after midnight during the maintenance program | Went thru all the fluid flow programs in service. Replaced the inlet gasket. Will talk about not cleaning the inlet gasket any more | SWO-238398 | Re classified RMED | Double entry | SP | Leak current | 964 | Multiple SP's having the same issue | Customer has changed the SP numerous times. Always happens right after midnight during the maintenance program | Went thru all the fluid flow programs in service. Replaced the inlet gasket. Will talk about not cleaning the inlet gasket any more
No | Visit | Complaint | RDEU | 02.10.2017, 10:27 Uhr, Fr. Rohlsing, Tel. 023631082704: Leak current detected | Das Gerät erkennt Ableitströme am LP, | Fließwege und LP kontrolliert, alles o.k., SC erneuert und Gerät kalibriert, | SWO-238288 | Assisted Replacement | SC | SC | Leak current | 964 | 02.10.2017, 10:27 Uhr, Fr. Rohlsing, Tel. 023631082704: Leak current detected | The device detects discharge currents on the LP, | Flow paths and LP checked, everything OK, SC renewed and device calibrated,
No | Visit | Complaint | RAME | User called stating analyzer keeps failing calibration for Lac getting sensor response error. | Found that the analyzer keeps getting lac sensor response error 1310 on calibrations. | Replaced SC and SP to try and correct the issue. It did not work. | SWO-239447 | Defect | HW | Basic Unit | PCB, Sensor interface | 1310 | User called stating analyzer keeps failing calibration for Lac getting sensor response error. | Found that the analyzer keeps getting lac sensor response error 1310 on calibrations. | Replaced SC and SP to try and correct the issue. It did not work.
Yes | 2 | Y | RPOL | Blank | Okresowy przegl¹d techniczny wg procedury. Kompletna wymiana owê¿owania oraz uszczelki wlotu. Oczyszczenie obudowy oraz wnêtrza analizatora. Kalibracja tHb, tetsy kontrolne. Analizator sprawny. |  | D33A0328004782B2 | [5 empty] | Blank | Periodic technical Overview by procedure. Complete replacement of the intake and seal of the inlet. Cleaning the casing and the analyzer's inside. THb calibration, Control Tetsy. Functional Analyzer.
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | ca定标不过 | 更换测试卡后修复 | SWO-238013 | [5 empty] | Test card, reagent package failure | ca calibration but | Repair after replacing test card
No | Visit | Complaint | RDEU | 02.10.2017, 08:22 Uhr, Sr. Alexandra, Tel. 02114092277: Touch Screen reagiert nicht mehr | touch screen reagiert nicht, | neue Bildschirmsoftware eingespielt, Gerät kalibriert, Touch screen getestet, | SWO-238269 | Defect | SW | DMS_PC | Touch screen | Firmware update | 02.10.2017, 08:22, Sr. Alexandra, Tel. 02114092277: Touch Screen stops responding | touch screen does not respond, | new screen software, device calibrated, touch screen tested,
Yes | 6 | Y | RFRA | Blank |  |  | D33A1D67001F95AA | Re classified RMED |  |  |  |  | Blank
No | Visit | Complaint | RBNL | CDS waarden zijn te hoog | Regelmatig error Stolsel in SC terwijl geen stolsel in de SC aanwezig. Bij nader onderzoek blijkt dat de Clot Detection waarden te hoog of aan de hoge kant zijn. De waarden zijn als volgt: - Inlet value: 1,17 (0,80 - 1,22) - Sensor cassette value: 1,22 (0,80 - 1,20) - OXI module value: 0,96 (0,80 - 1,20) De SC Bay Unit heb gedemonteerd, alle sensoren heb schoongemaakt, nu geen foutmelding maar de waarden zijn nog steeds aan de hoge kant. Een nieuwe SC Bay Unit is besteld. | SC Bay unit sensoren heb schoongemaakt. | SWO-254886 | Cleaned | HW | Basic Unit | Sensor Bay Unit |  | CDS values are too high | Regular error clot in SC While no clot is present in the SC. Further investigation reveals that the Clot Detection values are too high or high-end. The values are as follows: -Inlet Value: 1.17 (0.80-1.22) -Sensor Cassette value: 1.22 (0.80-1.20) -OXI Module Value: 0.96 (0.80-1.20) The SC Bay Unit got disassembled, all the sensors got cleaned, now no error message but the values are still on the high side. A new SC Bay Unit has been ordered. | SC Bay unit sensors have been cleaned.
No | Visit | Complaint | RAME | Customer states unit having Temperature errors | Customer states unit going into recal mode on intermittent basis | Upon arrival, found unit having intermittent Temperature errors causing issue. moved unit to warmer stable location. Unit Temp error gone. Ran unit calibrations and QCS. Unit functioning to Specifications. NOTE: Trained customer on appropriate use of unit. | SWO-238871 | User error | Installation | Analyzer | Temp vs Location |  | Customer states unit having Temperature errors | Customer states unit going into recal mode on intermittent basis | Upon arrival, found unit having intermittent Temperature errors causing issue. moved unit to warmer stable location. Unit Temp error gone. Ran unit calibrations and QCS. Unit functioning to Specifications. NOTE: Trained customer on appropriate use of unit.
No | Visit | Complaint | RAME | The sO2 works for a short time the deactivates. Also there are numerous Possible Fluid Blockage errors. | Analyzer has had numerous and various errors. The sO2 is currently deactivated and QC Levels 2 and 3 need to be performed. | Performed tHb calibration and ran cleaning pack. Ran external QC. Analyzer is working right now. Will have customer monitor. | SWO-238865 | [5 empty] | The sO2 works for a short time the deactivates. Also there are numerous Possible Fluid Blockage errors. | Analyzer has had numerous and various errors. The sO2 is currently deactivated and QC Levels 2 and 3 need to be performed. | Performed tHb calibration and ran cleaning pack. Ran external QC. Analyzer is working right now. Will have customer monitor.
No | 1 | Y | RFRA | Blank | Bouteille de gaz bloquée dans la chambre |  | D33B70EC76167C0D | [5 empty] | Blank | Gas bottle stuck in room
Yes | 2 | Y | RPAC | Blank | Operators manual |  | D33B0CEA7609A374 | Re classified RMED |  |  |  |  | Blank | Operators manual
Yes | 2 | Y | RPAC | Blank | Operator Manual |  | D33B0DD9763CA378 | Re classified RMED |  |  |  |  | Blank | Operator Manual
No | Visit | Complaint | RIND | Oxymetry values not coming. | Oxymetry Values Not coming. | Try to do Complete Oxymetry Test and found lam unit failed as neon lamp intensity and Potential not in range. Try to shutdown and restart analyzer after one hour lap and again try to do complete oxy test but same problem reported. TO sending for oxi control PCB. | SWO-238821 | Revisit required | HW | Oxi module | PCB, Oxi module |  | Oxymetry values not coming. | Oxymetry Values Not coming. | Try to do Complete Oxymetry Test and found lam unit failed as neon lamp intensity and Potential not in range. Try to shutdown and restart analyzer after one hour lap and again try to do complete oxy test but same problem reported. TO sending for oxi control PCB.
Yes | 1 | Y | RPAC | Blank | OPerator Manual |  | D33B0E887614A37B | Re classified RMED |  |  |  |  | Blank | OPerator Manual
No | 2 | Y | RFRA | Blank | aiguille tordue |  | D33B23AF7659D1BD | Defect | HW | Aspiration Module | Bend needle |  | Blank | twisted needle
No | Visit | Complaint | RDAN | D-Dimer målt til 4000, 400 og 500 på samme prøve | NA | Der er kørt 5 prøver på vand. Der er taget et Service Dump. Alt er OK. | SWO-238710 | Cleaned | HW | Process Module | Process wheel | Techline | D-Dimer measured at 4000, 400 and 500 on the same sample | In | 5 samples have been taken on water. A Service Dump is being taken. Everything is OK.
No | Visit | Complaint | RAME | Customer advised that nothing on display, just blue background, attempted reboot, issue not resolved. | Completion of Network config and analyzer setup now that new supplies have been installed. | Configured Network Adapter, LIS Server, Time $Formatting, parameters and ranges. | SWO-239597 | Defect | SW | RMED SW | Frozen - His/lis |  | Customer advised that nothing on display, just blue background, attempted reboot, issue not resolved. | Completion of Network config and analyzer setup now that new supplies have been installed. | Configured Network Adapter, LIS Server, Time $Formatting, parameters and ranges.
No | Visit | Complaint | RAME | Client advised of error 964: Leak current in relation to solution pack detected. | OSPD found analyzer stuck on user intervention after customer troubleshooting. | Ran service tests, no errors found. Startup run and leak error returned with 5 seconds left on timer. Client to replace solution pack when new inventory arrives and continue testing. | SWO-239357 | Assisted Replacement | SP | SP | Leak current | 964 | Client advised of error 964: Leak current in relation to solution pack detected. | OSPD found analyzer stuck on user intervention after customer troubleshooting. | Ran service tests, no errors found. Startup run and leak error returned with 5 seconds left on timer. Client to replace solution pack when new inventory arrives and continue testing.
No | Visit | Complaint | RAME | Possible leak in system | Inspected system and found small hole in pinch tubing. Sensor cassette cracked and leaking.solution pack replaced for poor pump cal results | Replaced solution pack and sensor cassette. Installed Annual kit and tested system in service. | SWO-238940 | Defect | HW | Oxi module | Pinch valve tube |  | Possible leak in system | Inspected system and found small hole in pinch tubing. Sensor cassette cracked and leaking.solution pack replaced for poor pump cal results | Replaced solution pack and sensor cassette. Installed Annual kit and tested system in service.
No | Visit | Complaint | RLTD | Tim 07546906874 | Customer unable to log on during trouble shooting. | Couldn't find any issues, customer can now log on. Customer to monitor. | SWO-239481 | No fault found | No fault found | No fault found | Log on |  | Tim 07546906874 | Customer unable to log on during trouble shooting. | Couldn't find any issues, customer can now log on. Customer to monitor.
No | 4 | Y | RPOL | Blank | Reinstalacja Windows. Aktualizacja oprogramowania do wersji 3.3 MR1. Analizator sprawny technicznie. |  | D33CBCE6EC1D7B65 | Defect | SW | RMED SW | RMED SW |  | Blank | Reinstalling Windows. Software update to version 3.3 MR1. Technically efficient analyzer.
No | 5 | Y | RPAC | Blank | Liquid transport error |  | D33CA44BEC2E0E7F | Assisted Replacement | SP | SP | FT |  | Blank | Liquid transport error
Yes | 5 | Y | RFRA | Blank |  |  | D33C6299EC1EF813 | Re classified RMED | TL |  |  |  | Blank
Yes | 1 | Y | RFRA | Blank |  |  | D33C7E71EC3EB911 | Re classified RMED | FAN |  |  |  | Blank
Yes | 5 | Y | RFRA | Blank | Bay unit non fixé |  | D33C5AFBEC58F80F | Re classified RMED | TL |  |  |  | Blank | Bay unit unfixed
No | Visit | Complaint | RAME | Intervention required. Inlet gasket required changing due to schedule. | Intervention required. Inlet gasket needed adjustment and changing. | Adjusted inlet gasket holder and followed video guidance. No super users available for phone support. Visit required. | SWO-239371 | User error | HW | Auto Inlet module | Inlet gasket |  | Intervention required. Inlet gasket required changing due to schedule. | Intervention required. Inlet gasket needed adjustment and changing. | Adjusted inlet gasket holder and followed video guidance. No super users available for phone support. Visit required.
Yes | Visit | Complaint | RAME | The sO2 works for a short time the deactivates. Also there are numerous Possible Fluid Blockage errors. | The analyzer has had numerous oximetry issues, including sO2 out of range, unable to measure Hb, and fluid blockage errors with error code 2106. | Replaced the hemolyzer unit and the spectrometer unit. Performed tHb calibration and ran QC. | SWO-239429 | Re classified RMED | Double entry | SWO-238865 |  |  | The sO2 works for a short time the deactivates. Also there are numerous Possible Fluid Blockage errors. | The analyzer has had numerous oximetry issues, including sO2 out of range, unable to measure Hb, and fluid blockage errors with error code 2106. | Replaced the hemolyzer unit and the spectrometer unit. Performed tHb calibration and ran QC.
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | cl定标失败 | 更换测试卡后修复 | SWO-248007 | [5 empty] | Test card, reagent package failure | cl bid failed | Repair after replacing test card
No | Visit | Complaint | RAME | Single parameter failure Co2 | PCO2 unstable after any SP $SC change (compared to other analyzer | Ran Zero current calibration. Ran system cycle. | SWO-239494 | [5 empty] | Single parameter failure Co2 | PCO2 unstable after any SP /SC change (compared to other analyzer | Ran Zero current calibration. Ran system cycle.
Yes | 1 | Y | RPAC | Blank | Installation |  | D33C92ACEC0F393E | Re classified RMED | Setup / Installation |  |  |  | Blank | Installation
No | Visit | Complaint | RAME | Fluid leakage from syringe port. | Fluid leakage from syringe port. Checked all waste tubing. Found waste line under inlet disconnected. Also found tubing fitting loosely on probe. | Reseat all waste tubing. Replaced inlet probe tubing. | SWO-239435 | Defect | HW | Inlet module | Tubings |  | Fluid leakage from syringe port. | Fluid leakage from syringe port. Checked all waste tubing. Found waste line under inlet disconnected. Also found tubing fitting loosely on probe. | Reseat all waste tubing. Replaced inlet probe tubing.
Yes | 7 | Y | RPOL | Blank | sprawdzenie analizatora, wymiana uszczelki wlotu, kalibracja parametrów , kontrola jakoœci. Aparat sprawny technicznie |  | D3446205B0379A87 | [5 empty] | Blank | Analyzer check, inlet gasket replacement, parameter calibration, quality control. Technically efficient apparatus
No | Visit | Complaint | RDEU | 02.10.2017 10:15Uhr Herr Söllner 095150315242: Das Gerät meldet 1342 | Anruf am Mittwoch 04.10.17, 14:31 Uhr, von Medizintechnik Hr. Robert Söllner (i.A. OA), mit der Bitte das Gerät zu überprüfen --> Termin für Donnerstag 05.10.17 vereinbart --> Fehlersuche: Dateiarchive (Aktivitäten, QC, Kal., Betriebsmittel), Einlass und Schlauchwege überprüft --> alles i.O., die 2 vorletzten Tage (03.10. und 04.10. , je um 00:03 Uhr, Ableitströme erkannt --> LP (Lot YR-51)) --> Jahreswartung in Rücksprache mit Hr. Söllner vorgezogen, siehe SWO-239734 der Jahreswartung | Analysator manuell mit Spritze gespült, Systemüberprüfung im Service (Adjust Manifold, Pumpenkal., LS-Justage, Clot detect $remove, Pressure Test), System und Funktion i.O. | SWO-239732 | Defect | HW | Basic Unit | PM kit |  | 02.10.2017 10:15Am Mr Söllner 095150315242: The device reports 1342 | Call on Wednesday 04.10.17, 14:31, from medical technology Mr. Robert Söllner (i.A. OA), with the request to check the device --> Date for Thursday 05.10.17 --> Troubleshooting: File archives (activities, QC, cal., equipment, equipment), inlet and hoseways checked --> all the penultimate days (03.10. and 04.10. , each at 00:03 o'clock, discharge currents detected --> LP (Lot YR-51)) --> Annual maintenance in consultation with Mr. Söllner preferred, see SWO-239734 of the annual maintenance | Analyzer manually rinsed with syringe, system check in service (Adjust Manifold, Pump Scale., LS-Adjustment, Clot detect $remove, Pressure Test), system and function i.O.
No | Visit | Complaint | RBNL | ABL geeft errormelding 1559, dit betekend dat de accu stuk is. | ABL geeft systeem status geel ,met de melding error 1559 (accu error). Daarom de accu vervangen. Foutmelding is verdwenen en statuscode is nu weer normaal. | Accu vervangen, onder garantie. | SWO-240282 | Defect | HW | Battery module | Battery | 1559 | ABL gives error message 1559, this means that the battery is part. | ABL Displays system status yellow, with the message Error 1559 (ACCUERROR). Therefore replace the battery. Error message is gone and status code is now back to normal. | Replace the battery, under warranty.
No | 2 | Y | RPOL | Blank | Naprawa modu³u inlet. Analizator sprawny |  | D33DB0FE623E5001 | Defect | HW | Auto Inlet module | Auto Inlet module |  | Blank | Repair of the inlet module. Functional analyzer
Yes | 3 | Y | RFRA | Blank | Intervention nécessaire |  | D33D95A8620BF05F | Re classified RMED | Phone support |  |  |  | Blank | Intervention needed
Yes | 1 | Y | RFRA | Blank |  |  | D33D93466229F05D | Re classified RMED | TL |  |  |  | Blank
No | Visit | Complaint | RAME | ALERE reports: We are not receiving the device ID or firmware version in the device. | Troubleshot and found that the RALS setup in the analyzer was not correct | Corrected the RALS setup, and this resolved the problem | SWO-240147 | [5 empty] | ALERE reports: We are not receiving the device ID or firmware version in the device. | Troubleshot and found that the RALS setup in the analyzer was not correct | Corrected the RALS setup, and this resolved the problem
No | Visit | Complaint | RAME | Linking pink liquid. The FSE has already been on site once for this issue. | Analyzer leaking solution from sensor cassette. Also found fluid underneath the solution pack. | Removed the wet section to find waste tube partially kinked right behind the waste port where it connects to the sensor cassette. Straightened waste tube, flushed air through to make sure it was a clean path. | SWO-240090 | [5 empty] | Linking pink liquid. The FSE has already been on site once for this issue. | Analyzer leaking solution from sensor cassette. Also found fluid underneath the solution pack. | Removed the wet section to find waste tube partially kinked right behind the waste port where it connects to the sensor cassette. Straightened waste tube, flushed air through to make sure it was a clean path.
No | Visit | Complaint | RESP | El analizador no conecta con Aqure, no recibe datos ni envía. | Volvió a coger la IP antigua del equipo (10.35.17.87) y no la nueva que es 10.35.16.7 | Se cambia IP y hace commit | SWO-239921 | Setup | SW | His/lis | His/lis |  | The scanner does not connect to Aqure, does not receive data, or send data. | He picked up the old IP of the computer (10.35.17.87) again and not the new one which is 10.35.16.7 | IP changes and commits
No | Visit | Complaint | RAME | Customer called to report the display screen will turn black intermittently during use. | Unit will not complete startup. Locks up with status code 9. | Removed and reinstalled Kira application software. Tried to reload Roma.db but unit would not accept. Reinstalled QC and setup. | SWO-240042 | Defect | SW | RMED SW | RMED SW |  | Customer called to report the display screen will turn black intermittently during use. | Unit will not complete startup. Locks up with status code 9. | Removed and reinstalled Kira application software. Tried to reload Roma.db but unit would not accept. Reinstalled QC and setup.
No | Visit | Complaint | RAME | GLU Unstable on patient samples only | Checked all logs to verify errors. Found calibrations and QC passing but several unstable errors on patient samples. Found Glu drifting and low sensitivity. Found clot in PH and Ref. Measuring chambers. | Cleaned the measuring chambers, PH seal ring and replaced the Ref. Membrane. Calibrations and blood sample with unstable errors. Replaced the Glu electrode and ran several bloods to confirm the no unstable errors. Calibrations have less drift and sensitivity is still increasing. I will follow up with customer in the morning to confirm performance. | SWO-239507 | Assisted Replacement | Electrode | Electrode | Glu |  | GLU Unstable on patient samples only | Checked all logs to verify errors. Found calibrations and QC passing but several unstable errors on patient samples. Found Glu drifting and low sensitivity. Found clot in PH and Ref. Measuring chambers. | Cleaned the measuring chambers, PH seal ring and replaced the Ref. Membrane. Calibrations and blood sample with unstable errors. Replaced the Glu electrode and ran several bloods to confirm the no unstable errors. Calibrations have less drift and sensitivity is still increasing. I will follow up with customer in the morning to confirm performance.
No | Visit | Complaint | RDEU | Gerät meldet Kal drift Fehler. Gasflaschen werden als leer darstellt obwohl Sie erneuert worden sind. | Flat Membran defekt. Undichtigkeiten an den Membranventilen im pH $ BG und EL $Met Modul. | Gerät kalibriert, Membraniventile in den Messmodulen demontiert und gereinigt, Flat Membran an den Air Membranventilen im pH $ BG Modul und EL $Met Modul erneuert. Gerät kalibriert. Gas Flaschen erneuert initialisiert und Bestände OK. QC durchgeführt und OK. | SWO-240263 | Defect | HW | pH/BG Module | Valve, air | Membrane | Device reports Kal drift error. Gas cylinders are presented as empty even though they have been renewed. | Flat membrane defective. Leaks on the diaphragm valves in pH BG and EL /Met module. | Device calibrated, diaphragm valves in the measuring modules disassembled and cleaned, flat membrane on the air diaphragm valves in the pH BG module and EL $Met module renewed. device calibrated. Gas bottles renewed initialized and stocks OK. QC performed and OK.
No | Visit | Complaint | RDAN | Når Pia scanner sit kort for at logge ind på ABLen, så giver scanneren et bib, men Pia bliver ikke logget ind. Lidt efter giver scanneren 2 bib fra sig. | NA | Skanner skiftet uden det løste problemmet. Stikforbindelser tjekket | SWO-240316 | Defect | HW | Basic Unit | Barcode reader, Basic Unit |  | When Pia scans his card to log in to the ABLen, the scanner gives a bib, but Pia is not logged in. | In | Scan the switch without the resolved prob lem. Plug Connections checked
No | Visit | Complaint | RSCH | Neudefekt | Sporadisch Fehler 0597 während Kal. Flüssigkeit während Gas Kalibrierung in pH $BG-Modul erkannt | Allgemeine Kontrolle i.o Ventil-Block komplett unter pH und EL-Modul getauscht i.o O-Ring pH ersetzt i.o Nach zwei Tagen erneut Fehler Messkammer pH $BG ausgebaut und ersetzt i.o Oberer Flüssigkeitsfühler pH $BG ausgebaut und ersetzt i.o Kal. und QC i.o | SWO-239745 | Defect | HW | pH/BG Module | Valve | OOB | Neudefekt | Sporadic error 0597 during Kal. Liquid detected in pH /BG module during gas calibration | General control i.o Valve block completely under pH and EL module replaced i.o O-ring pH replaces i.o Again after two days of errors Measuring chamber pH $BG removed and replaces i.o Upper liquid sensor pH $BG removed and replaces i.o Kal. and QC i.o
Yes | 7 | Y | RFRA | Blank | erreur 2877 |  | D33DE0BD62582686 | [5 empty] | Blank | error 2877
No | 6 | Y | RFRA | Blank | 4794: PIPETTE 38 |  | D33D9E506216F86B | Defect | HW | Aspiration Module | Needle Sensor | 4794 | Blank | 4794: PIPETTE 38
No | Visit | Complaint | RLTD | Ranvir Tel 01865 221596 | Ran the sensor simulator test and it showed and error on the Cl channel. | Replaced the sensor interface PCB and re-ran the sensor simulator test. | SWO-240301 | Defect | HW | Basic Unit | PCB, Sensor interface | Cl | Ranvir Tel 01865 221596 | Ran the sensor simulator test and it showed and error on the Cl channel. | Replaced the sensor interface PCB and re-ran the sensor simulator test.
No | Visit | Complaint | RDEU | Hb | Fehlersuche --> Dateiarchive (Aktivitäten, QC, Kal.) überprüft --> pCO2-QC-Fehler in allen 3 Leveln (oberhalb Statistikbereich), seit 29.9.17 (letzter Wechsel der SC) --> SC erneuern | SC (R1132) aus Kundenbestand eingebaut und (R1107) ausgebaut, System und Funktion i.O. | SWO-240253 | Assisted Replacement | SC | SC | pCO2 |  | Hb | Troubleshooting --> File archives (Activities, QC, Cal.) checked --> pCO2-QC errors in all 3 levels (above statistics area), since 29.9.17 (last change of SC) --> Renew SC | SC (R1132) built from customer base and (R1107) removed, system and function in the
No | Visit | Complaint | RAME | Corrupt Win 7 | Corrupt Windows 7, will need to reload win 7, calibrated touch screen and ran the commit command. Reloaded the setup. | Reloaded Win 7 and 3.3 MR1 application SW, ran cal's and QC's, abl 90 running at manufactures spec. Also left new inlet gaskets with holder for each unit at the hospital and agreed for go live week. | SWO-240110 | Defect | SW | Windows | Windows |  | Corrupt Win 7 | Corrupt Windows 7, will need to reload win 7, calibrated touch screen and ran the commit command. Reloaded the setup. | Reloaded Win 7 and 3.3 MR1 application SW, ran cal's and QC's, abl 90 running at manufactures spec. Also left new inlet gaskets with holder for each unit at the hospital and agreed for go live week.
No | Visit | Complaint | RSCH | Blank | Ableitströme erkannt. | Gerät und sämtliche Flüssigkeitswege auf Dichtheit überprüft. Im SC-Fach Pins teils verkrustet, teils feucht. Fach kompl. gereinigt und getrocknet. Gleiche SC neu installiert, alles i.o. Kal., QC und System i.o. Funktionskontrolle i.o. | SWO-240450 | Adjustment | HW | Basic Unit | Sensor Bay Unit | Contact pin bend | Blank | discharge currents detected. | device and all fluid pathways checked for leak. In the SC compartment pins partly encrusted, partly moist. Fach kompl. cleaned and dried. Same SC reinstalled, all i.o. Kal., QC and System i.o. Functional control i.o.
No | Visit | Complaint | RLTD | Fault observed on arrival in Biochem Lab via AQURE, with Stephanie Sandy POC. | Found Oxi pinch valve tube to be leaking slightly from pin hole in the tube where the valve pinch arm operates to close it. Fitted new Oxi inch tube. Verified operation on restart - all ok. Also upgraded software from v3.3 to v3.3 MR1. | Replaced worn Oxi pinch valve tube. | SWO-240470 | Defect | HW | Oxi module | Pinc valve tube |  | Fault observed on arrival in Biochem Lab via AQURE, with Stephanie Sandy POC. | Found Oxi pinch valve tube to be leaking slightly from pin hole in the tube where the valve pinch arm operates to close it. Fitted new Oxi inch tube. Verified operation on restart - all ok. Also upgraded software from v3.3 to v3.3 MR1. | Replaced worn Oxi pinch valve tube.
No | Visit | Complaint | RLTD | Analyser has been playing up for last 2 days, mentioned when on site by Stephanie Sandy POC. | Checked data logs, two errors are causing entry into troubleshooting: 1) Error 1401 - Inlet Gasket Holder error 2) Error 1086 Inhomogeneous QC2 solution. Checked inlet mechanism and tightened screws for the clutch mechanism, probe etc all ok. Checked solution pack and pushed down hard to ensure it was fully activated. Carry out extended service flush of sample pathway. Restarted analyser - all ok. Also carried out software upgrade from v3.3 to v3.3 MR1. | Tightened inlet clutch mechanism. Flushed sample pathway. | SWO-240469 | Adjustment | HW | Auto Inlet module | Inlet drive gear | 1401 | Analyser has been playing up for last 2 days, mentioned when on site by Stephanie Sandy POC. | Checked data logs, two errors are causing entry into troubleshooting: 1) Error 1401 - Inlet Gasket Holder error 2) Error 1086 Inhomogeneous QC2 solution. Checked inlet mechanism and tightened screws for the clutch mechanism, probe etc all ok. Checked solution pack and pushed down hard to ensure it was fully activated. Carry out extended service flush of sample pathway. Restarted analyser - all ok. Also carried out software upgrade from v3.3 to v3.3 MR1. | Tightened inlet clutch mechanism. Flushed sample pathway.
No | 2 | Y | RFRA | Blank | Surconsommation Pack de solution |  | D33F1332D831D3E0 | [5 empty] | Blank | Overconsumption Solution Pack
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | GLU，Lac定标不过 | 更换测试卡后修复 | SWO-241410 | Assisted Replacement | SC | SC | Glu |  | Test card, reagent package failure | GLU, Lac calibration is not | Repair after replacing test card
No | Visit | Complaint | RAME | Bottle transport error, aborting calibrations | Checked data logs for errors, checked flow of solutions, checked valves and liquid sensors | Removed modules and cleaned valves, replaced ripped air valve membrane, replaced pH seal ring | SWO-240475 | Defect | HW | pH/BG Module | Valve, air |  | Bottle transport error, aborting calibrations | Checked data logs for errors, checked flow of solutions, checked valves and liquid sensors | Removed modules and cleaned valves, replaced ripped air valve membrane, replaced pH seal ring
Yes | 6 | Y | RCZE | Blank |  |  | D33F0C77D84FAAD1 | Re classified RMED | Fan 915-318 |  |  |  | Blank
No | Visit | Complaint | RAME | Customer reports tHb failure during calibrations. 0579 oxy hardware impossible to measure errors are present. | The lamp unit results are observed to fail during the complete oximetry test. The halogen intensity value is above the upper threshold of 3900. | Performed the lamp unit adjustment procedure, lowering the halogen intensity value to 3200. Successfully performed the complete oximetry test. Performed all other necessary calibrations to bring he analyzer to a functional state. Analyzer is within acceptable operating specifications. | SWO-240510 | Adjustment | HW | Oxi module | Lamp unit |  | Customer reports tHb failure during calibrations. 0579 oxy hardware impossible to measure errors are present. | The lamp unit results are observed to fail during the complete oximetry test. The halogen intensity value is above the upper threshold of 3900. | Performed the lamp unit adjustment procedure, lowering the halogen intensity value to 3200. Successfully performed the complete oximetry test. Performed all other necessary calibrations to bring he analyzer to a functional state. Analyzer is within acceptable operating specifications.
No | Visit | Complaint | RIND | Touch Screen Problem. | Touch Screen problem. | Found problem with Touch screen, tried calibrating but problem was remaining same, need to replace with new Touch Screen waiting for part. | SWO-240750 | Defect | HW | Basic Unit | Touch screen | LEH Mail | Touch Screen Problem. | Touch Screen problem. | Found problem with Touch screen, tried calibrating but problem was remaining same, need to replace with new Touch Screen waiting for part.
No | Visit | Complaint | RCHN | glu、lac定标失败 | glu、lac定标失败，需要借件903-183一个 |  | SWO-240716 | Defect | HW | Basic Unit | PCB, Sensor interface |  | glu, lac calibration failed | glu, lac calibration failure, need to borrow 903-183 one
No | Visit | Complaint | RCHN | 测试卡锁不住 | 测试卡锁不住，需要借件903-030一个 |  | SWO-240718 | Defect | HW | ?? | PCB, Cartridge Lock |  | The test card won't lock. | The test card is not locked and needs to borrow 903-030 a
Yes | 2 | Y | RFRA | Blank | Impossible de pomper solutions |  | D342C57D3A62E97C | [5 empty] | Blank | Unable to pump solutions
No | 3 | Y | RFRA | Blank | Divers problèmes : bouchage, erreur pO2, plantage logiciel |  | D343137D3A437673 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | Blank | Various problems: blocking, pO2 error, software crash
No | Visit | Complaint | RESP | INICIALIZACION. | REVISION DE TODO EL SISTEMA | REALIZAR REVISION DE TODO EL SISTEMA PARA ANALIZAR POSIBLES CAUSAS, NO DETECTANDO NINGUN PROBLEMA Y NO TENIENDO REACTIVOS PARA SU VERIFICACION. TODOS LOS TEST PASAN, TANTO ELECTRONICOS COMO NEUMATICOS. | SWO-241192 | [5 empty] | INICIALIZACION. | SYSTEM-WIDE REVIEW | REVIEW THE ENTIRE SYSTEM TO ANALYSIS POSSIBLE CAUSES, NOT DETECTING ANY PROBLEMS AND NOT HAVING REACTIVE FOR VERIFICATION. ALL TESTS PASS, BOTH ELECTRONIC AND NEUMATIC.
No | Visit | Complaint | RESP | SPECTRUM MISTMATCH | ERRORES DE MUESTRAS EN OXI | SE EFECTUAN PRUEBAS SE DESMONTA DODO EL SISTEMA Y SE SUSTITUYE HEMOLIZADOR POR MAL FUNCIONAMIENTO. AJUSTES DE THB Y DE CANALES DE PARAMETROS. CALIBRADO Y VERIFICADO CON Q.C. | SWO-241191 | [5 empty] | SPECTRUM MISTMATCH | OXY SAMPLE ERRORS | TESTS ARE EFFECTED THE SYSTEM IS DISMOUNTED AND HEMOLIZADOR IS SUBSTITUTEFORD BY THE WORLD OPERATION. THB AND PARAMETER CHANNEL SETTINGS. CALIBRATED AND VERIFIED WITH Q.C.
No | Visit | Complaint | RCHN | 光学强度超范围 | 光学强度超范围，借903-364一个 |  | SWO-240883 | Defect | HW | Process Module | Optical Unit | 1031 | Optical strength over range | Optical strength over range, borrowed 903-364 a
No | Visit | Complaint | RAME | Customer reported oximetry error messages, Neon Lamp and Neon Number Too Low. | Analyzer failing with error code 316 and 258. Oximetry test failing with low neon intensity and low neon number. | Replaced Lamp Unit. Adjusted halogen lamp intensity. Ran tHB calibration with pass results. Restart calibrations. | SWO-241290 | Defect | HW | Oxi module | Lamp unit | 316 | Customer reported oximetry error messages, Neon Lamp and Neon Number Too Low. | Analyzer failing with error code 316 and 258. Oximetry test failing with low neon intensity and low neon number. | Replaced Lamp Unit. Adjusted halogen lamp intensity. Ran tHB calibration with pass results. Restart calibrations.
Yes | 7 | Y | RPOL | Blank | pomimo wymiany uszczelki aparat nie umo¿liwia³ oznaczanie parametrów. Sprawdzenie analizatrora, wymiana wê¿yków, kalibracja parametrów. Kontrola jakoœci. Wykonanie testowego badania. Aparat sprawny technicznie. |  | D34462ECB00E9A89 | [5 empty] | Blank | Despite replacing the gasket, the camera does not allow parameter determination. Checking the analytor, replacing the parts, calibrating the parameters. Quality control. Perform a test study. The camera is technically efficient.
No | Visit | Complaint | RLTD | Please call Mark Harper on 01895 828 757 | Fluid transportation message, won't start up properly | Solution transport errors leak detected and inhomogeneous solution . Different SP and SC fitted still getting same errors. Wet section removed and all tubing removed and checked for leaks ok. Refitted all parts and passing service test ok. Test sample but though lubricant system. Cal and QC run passing ok. Setup installed ok and software ok. | SWO-241907 | Cleaned | Fluidic pathway | Fluidic pathway | Waste Blood |  | Please call Mark Harper on 01895 828 757 | Fluid transportation message, won't start up properly | Solution transport errors leak detected and inhomogeneous solution . Different SP and SC fitted still getting same errors. Wet section removed and all tubing removed and checked for leaks ok. Refitted all parts and passing service test ok. Test sample but though lubricant system. Cal and QC run passing ok. Setup installed ok and software ok.
No | Visit | Complaint | RAME | The analyzer keeps telling the user to replace the solution pack even though they just did. | Reviewed the data logs. QC failures started occurring after the sensor cassette was replaced on 10 $5. The pO2 is reading low on S9030 and high on S9050. | Swapped sensor cassettes with 090R0900N0031 as there weren't any more available. Left a spare inlet gasket with holder. I am going to order more supplies. | SWO-241978 | Assisted Replacement | SC | SC | pO2 QC | Wetability | The analyzer keeps telling the user to replace the solution pack even though they just did. | Reviewed the data logs. QC failures started occurring after the sensor cassette was replaced on 10 /5. The pO2 is reading low on S9030 and high on S9050. | Swapped sensor cassettes with 090R0900N0031 as there weren't any more available. Left a spare inlet gasket with holder. I am going to order more supplies.
Yes | 2 | Y | RPAC | Blank | Installation and training of ABL90 Flex Plus |  | D356CE539A4D027E | Re classified RMED | Installation |  |  |  | Blank | Installation and training of ABL90 Flex Plus
No | Visit | Complaint | RAME | Thb cal solution | Checked data logs. | Changed the setup for thb calls to 0 dats. | SWO-241885 | [5 empty] | Thb cal solution | Checked data logs. | Changed the setup for thb calls to 0 dats.
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | K定标失败 | 更换测试卡后修复 | SWO-242025 | [5 empty] | Test card, reagent package failure | K bid failure | Repair after replacing test card
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | LAC定标失败，0210；温度下层超范围，1066 | 连续更换2张测试卡后修复 | SWO-241657 | Assisted Replacement | SC | SC | Temp | 1066 | Test card, reagent package failure | LAC calibration failure, 0210; temperature sub-range, 1066 | Repair after 2 consecutive test cards
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | LAC,GLU定标失败，0378 | 更换测试卡后修复 | SWO-242233 | Assisted Replacement | SC | SC | Glu/Lac | 1373 | Test card, reagent package failure | LAC, GLU calibration failure, 0378 | Repair after replacing test card
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | Lac，PCO2定标不过 | 更换测试卡后修复 | SWO-241499 | Assisted Replacement | SC | SC | Lac/pCO2 |  | Test card, reagent package failure | Lac, PCO2 calibration is not | Repair after replacing test card
No | Visit | Complaint | RCHN | 光学敏感性超出范围，报错1029 | 光学敏感性超出范围，报错1029，需要借件903-364备件一个 | 更换903-364备件后修复 by jesson | SWO-241524 | Defect | HW | Process Module | Optical Unit | 1029 | Optical sensitivity out of range, error 1029 | Optical sensitivity is out of range, error 1029, need to borrow 903-364 spare parts one | Repair by jesson after replacement of 903-364 spare parts
No | Visit | Complaint | RIND | Oximetry hardware problem. Possible to measure | Checked the data logs found Oximetry hardware problem. Possible to measure for all the samples switch off the machine for 1 hr & on the machine then also same problem performed complete oximetry test found neon voltage was getting failed found oxi controller pcb is defective | Need to replace oxi controller pcb | SWO-241780 | Revisit required | HW | Oxi module | PCB, Oxi module | Neon Lamp | Oximetry hardware problem. Possible to measure | Checked the data logs found Oximetry hardware problem. Possible to measure for all the samples switch off the machine for 1 hr & on the machine then also same problem performed complete oximetry test found neon voltage was getting failed found oxi controller pcb is defective | Need to replace oxygen controller PCB
No | Visit | Complaint | RDEU | pCO2 scheitert häufig in QC LEVEL3 | Membran und Elektrode wurde offenbar schon ersetzt, Membran überprüft, Gasregelung getestet | Messkammer gereinigt, Kontaktkappe gebürstet, Gasventile überprüft, Kal und QC | SWO-242290 | Cleaned | Fluidic pathway | pH/BG Module | Measuring chamber |  | pCO2 often fails in QC LEVEL3 | Membrane and electrode apparently has already been replaced, membrane checked, gas control tested | Measuring chamber cleaned, contact cap brushed, gas valves checked, Kal and QC
No | Visit | Complaint | RKK | 検体測定にてPH応答不良頻発 | 検体測定にてPH応答不良頻発 | PH電極及び測定チャンバー洗浄。PH電極シーリングリング交換。エアーライン洗浄。 | SWO-242067 | Cleaned | Fluidic pathway | pH/BG Module | Measuring chamber | 476 | Frequent ph response failure in sample measurement | Frequent ph response failure in sample measurement | PH electrode and measurement chamber cleaning. PH electrode sealing ring replacement. Air line cleaning.
No | Visit | Complaint | RCHN | 扫码器无法扫描 | 扫码器无法扫描，需要借件902-822后修复 |  | SWO-241431 | Defect | HW | Basic Unit | Barcode reader, Basic Unit |  | The scanner cannot be scanned | The scanner cannot be scanned and needs to be repaired after borrowing 902-822
No | Visit | Complaint | RDEU | ABL reagiert verzögert | Display reagiert nur teilweise. | Display Firmware neu installiert Danach war die Bedienung wieder vollständig hergestellt. Mehrere Kalibrierungen durchgeführt. QCs für alle Level durchgeführt. ABL wieder messbereit. | SWO-242107 | Defect | SW | DMS_PC | Touch screen | Firmware update | ABL reacts delayed | Display reacts only partially. | Display firmware reinstalled After that, the operation was fully restored. Multiple calibrations performed. QCs for all levels. ABL ready to be measured again.
Yes | 6 | Y | RFRA | Blank | Pas de lecture code barre ecran |  | D343FF55B04E3617 | [5 empty] | Blank | No screen barcode reading
Yes | 33 | Y | RPAC | Blank | Solution pack chip would not read |  | D347EA7012055E64 | [5 empty] | Blank | Solution pack chip would not read
No | Visit | Complaint | RAME | Customer reported multiple consumable failures over the course of one week. | Data Logs needed for investigation. | Retrieved data logs. Upload data into spreadsheets for analysis. | SWO-242639 | Assisted Replacement | SC | SC | Unknown |  | Customer reported multiple consumable failures over the course of one week. | Data Logs needed for investigation. | Retrieved data logs. Upload data into spreadsheets for analysis.
No | Visit | Complaint | RBNL | pO2 trop bas (cq niveau C) | La pO2 est trop bas pour cq niveau 3 depuis quelques semaines. Le client a déjà changé la cassette et pack de solution au moins 2 fois. J'ai constaté que la sensibilité de la pO2 (calibration) est devenu plus haute (96%) après avoir changé la cassette la première fois. Nettoyage du sensor pO2 (dans le compartiment 'bay unit') n'a rien apporté. Remplacement de la cassette mais le problème reste (sensibilité 96%) | Remplacement du pack de solution (YT-07) : la pO2 mesure plus haute (20mmHg) | SWO-242759 | Assisted Replacement | SP | SP | pO2 - QC C low |  | pO2 too low (cq level C) | The pO2 is too low for cq level 3 for a few weeks. The customer has already changed the cassette and solution pack at least 2 times. I found that the sensitivity of pO2 (calibration) has become higher (96%) after changing the tape the first time. Cleaning the sensor pO2 (in the bay unit compartment) did not bring anything. Replacement of the cassette but the problem remains (sensitivity 96%) | Solution Pack Replacement (YT-07): pO2 measures higher (20mmHg)
No | 1 | Y | RCZE | Blank | pøi mìøení výsledku s chybou 1070 u Ca - chyba odpovìdi senzoru, chyba u více instalovaných kazet, všechny kalibrace a interní kontroly OK |  | D3459B7E262D81D6 | Adjustment | HW | Basic Unit | Sensor bay Unit (Pin alignment) | Ca | Blank | When measuring the result with error 1070 for CA-sensor response error, error for multiple installed cartridges, all calibrations and internal checks OK
Yes | 1 | Y | RFRA | Blank | FAN 915-371 |  | D345538026191646 | Re classified RMED | FAN |  |  |  | Blank | FAN 915-371
No | Visit | Complaint | RSAF | oxi hadware error on calibrations intermittently | inermittent oxi neon intensity error, cable connections on oxi module and WSC PCB ok, hemolyzer ok. | Replaced WSC PCB not compatible- part number changed. New PCB to be ordered and replaced. | SWO-243310 | Revisit required | HW | Basic Unit | PCB, Wet Sec Control |  | oxi hadware error on calibrations intermittently | Inermittent Oxygen Neon intensity error, cable connections on Oxygen module and WSC PCB OK, hemolyzer OK. | Replaced WSC PCB not compatible- part number changed. New PCB to be ordered and replaced.
No | 3 | Y | RPOL | Blank | Analizator zabrany w celu usuniêcia usterki. |  | D3467E769C4DB3EA | Defect | HW | Basic Unit | Pump Rotor |  | Blank | The analyser taken to remove the defect.
No | Visit | Complaint | RIND | Display collar improper like fluorescence. | Display problem. | Display improper collar. At the time of restart display comes like fluorescence collared. Cross check with other display. working fine. But unable to boot. CF card found faulty in bios setting. Need to replace Display unit and CF card. TO Sending. | SWO-250376 | [5 empty] | Display collar improper like fluorescence. | Display problem. | Display improper collar. At the time of restart display comes like fluorescence collared. Cross check with other display. working fine. But unable to boot. CF card found faulty in bios setting. Need to replace Display unit and CF card. TO Sending.
No | 1 | Y | RPOL | Blank | sprawdzenie analizatora, zerowanie kanalow, kalibracja, czyszczenie serwisowe analizatora, kalibracja tHb, wykonanie próbnego badania. Analizator sprawny technicznie. Pobrano logi do analizy. |  | D3457E1A2625C39C | [5 empty] | Blank | Analyzer check, reset channels, calibration, service cleaning of the analyser, calibration of the tHb, testing test. Technically efficient analyzer. Downloaded logs for analysis.
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 温度错误，1066 | 更换新机器后修复 | SWO-242299 | Assisted Replacement | SC | SC | Temp |  | Test card, reagent package failure | Temperature error, 1066 | Repair after replacing a new machine
No | 1 | Y | RCZE | Blank | pøístroj na Vinièní v Brnì, chyba storage wheel, po vložení vzorku se zasekne |  | D3459C77264881D7 | Defect | Consumables | Test Kit | Cartridge, Broken |  | Blank | The device at the wine Bar in Brno, the storage wheel error is stuck
No | Visit | Complaint | RDEU | Reagenzpack wird nicht erkannt | Reagenzpack pcb defekt | Reagenzpack pcb ausgetauscht. | SWO-242149 | Defect | HW | Storage Module | PCB, Solution Pack |  | Reagent pack is not detected | Reagenzpack pcb defekt | Reagent pack pcb.
No | 3 | Y | RPOL | Blank | Wymiana karty CF. Reinstalacja Windows i aktualizacja oprogramowania do wersji 3.3 MR1. Instalacja zewnêtrznej karty sieciowe USB TP-LINK UE200 S $N 216C244000336 UN $1.0. Konfiguracja analizatora Testy - ok. Analizator sprawny technicznie. |  | D346D7829C3DE2DD | Defect | HW | DMS_PC | CF Card |  | Blank | Replacing the CF card. Reinstalling Windows and upgrading the software to version 3.3 MR1. Installing an external USB network adapter TP-LINK UE200 S /N 216C244000336 UN /1.0. Analyzer configuration Tests-OK. Technically efficient analyzer.
No | Visit | Complaint | RDEU | 12.10.2017 08:05Uhr Essebier 03455572733: Gerät bleibt im Fehlersuchverfahren | Vakuumfehler Pumpen-Kal Fehler | Überprüfung Dateiarchiv => Fehler tritt seit mehreren Tagen immer wieder auf In Vergangenheit immer wieder fehlgeschlagener Wechsel der Einlassdichtung und anschließend LP Tausch Überprüfung Einlassdichtung => schwarze Dichtung sitzt falsch in Hülse => Dichtung fehlerhaft => Austausch Einlassdichtung aus Kundenbestand Während Startphase mit Hr. Essebier offene Fragen, Überprüfung der Einlassdichtung und Handlingsprobleme besprochen => weiterhin ist die kurze Haltbarkeit der LPs aufgefallen (4 von 5 eingesetzten LPs mit Haltbarkeit weit unter 30 Tagen) => Info an RaiBe Startphase läuft ohne Fehler durch | SWO-242945 | Defect | HW | Inlet module | Inlet gasket |  | 12.10.2017 08:05H Essebier 03455572733: Device remains in the error search procedure | Vacuum error Pump-Kal Error | Checking File Archive => Error has been recurring for several days In the past repeatedly failed change of the inlet seal and then LP exchange Inspection inlet seal => black seal sits wrong in sleeve => Seal faulty => Replacement inlet seal from customer stock During the start phase discussed with Mr. Essebier open questions, inspection of the inlet seal and handling problems => the short durability of the LPs has been noticed (4 out of 5 LPs used with shelf life well under 30 days) => Info to RaiBe Start-up phase runs without error
No | Visit | Complaint | RLTD | pO2 failing calibration with status errors frequently. | Investigated slow start-up for Glucose & Lactate sensors - data showed start-up for last SC was around 1 hour - so ok. Performed Sensor simulator test to verify electronic channels - all channels passed. Frequent pO2 status errors on calibration were confirmed from the data logs. Decontaminated analyser sample pathway for 20 minutes with hypochlorite solution. Inlet gasket soaked in hypochlorite also. Analyser then dismantled and all tube connectors cleaned externally with hypochlorite solution. Re-assembled fitting new waste tube, Oxi pinch valve tube and inlet connector gasket. Fitted new Sensor Cassette and restarted analyser. Verified analyser operation - all ok. Also checked the pO2 performance of the other ABL90s on site. All analyses will need to be decontaminated, this will be arranged at future visits soon. | Decontaminated analyser sample pathway & replaced Sensor Cassette. | SWO-243101 | Cleaned | Fluidic pathway | Fluidic pathway | Decontamination |  | pO2 failing calibration with status errors frequently. | Investigated slow start-up for Glucose & Lactate sensors - data showed start-up for last SC was around 1 hour - so ok. Performed Sensor simulator test to verify electronic channels - all channels passed. Frequent pO2 status errors on calibration were confirmed from the data logs. Decontaminated analyser sample pathway for 20 minutes with hypochlorite solution. Inlet gasket soaked in hypochlorite also. Analyser then dismantled and all tube connectors cleaned externally with hypochlorite solution. Re-assembled fitting new waste tube, Oxi pinch valve tube and inlet connector gasket. Fitted new Sensor Cassette and restarted analyser. Verified analyser operation - all ok. Also checked the pO2 performance of the other ABL90s on site. All analyses will need to be decontaminated, this will be arranged at future visits soon. | Decontaminated analyser sample pathway & replaced Sensor Cassette.
No | Visit | Complaint | RSAF | oxi hadware error on calibrations intermittently | WSC PCB possibly faulty | Replaced WSC PCB and Software upgrade done. Startup test ok. | SWO-243312 | Defect | HW | Basic Unit | PCB, Wet Sec Control |  | oxi hadware error on calibrations intermittently | WSC PCB possibly faulty | Replaced WSC PCB and Software upgrade done. Startup test ok.
No | Visit | Complaint | RDEU | nach Wechsel d. Sensor Interface PCB laut FAN (915-371), Gerät mit Crea-Fehler nach Kal. | Rücktausch d. Sensor Interface PCB | Gerät aus; PCB zurückgetauscht; Gerät ein; Startphase abgewartet; Servicedump gezogen; div. Einstellungen im Serviceprogr. vorgenommen; Kal. + QC i.O. | SWO-242891 | Revisit required | HW | Basic Unit | PCB, Sensor interface | Changed back to old | after changing the Sensor Interface PCB according to FAN (915-371), device with Crea error to Kal. | Rücktausch d. Sensor Interface PCB | device off; PCB sexchanged back; device; start-up phase awaited; Service dump dragged; div. Settings in the Serviceprogr. carried out; Kal. + QC in the first place
No | Visit | Complaint | RDEU | 10.10.2017 16:15 Fr. Hülseeger 02519762820: Oxi-Fehler Patient 11.10.2017, 14.10 Uhr, Herr Noltensmeier, Tel: 02519765605: weiterhin werden ctHb und SO2 nicht gemessen. | Gerät überprüft und gereinigt | Hemolyzer erneuert, Parameter in der Software wieder frei geschaltet, Gerät neu gestartet, tHb Kal, Kal und QCs durchgeführt | SWO-242848 | [5 empty] | 10/10/2017 16:15 Ms. Hülseeger 02519762820: Oxi Error Patient 11.10.2017, 14.10, Mr. Noltensmeier, Tel: 02519765605: ctHb and SO2 are still not measured. | Device checked and cleaned | Hemolyzer renewed, parameters in the software unlocked again, device restarted, tHb Kal, Kal and QCs performed
No | Visit | Complaint | RCHN | 处理盘故障 | 处理盘故障 | 更换库存机的636-550,580-087后修复 by jesson | SWO-243273 | Defect | HW | Process Module | Stepper motor, Processwheel | 5784 | Dealing with disk failure | Dealing with disk failure | Repair by jesson after replacing the stock machine 636-550,580-087
No | Visit | Complaint | RCHN | 报1024错误 | 报1024错误，借903-364一个，944-212（13147）、944-216（13052）、944-214（13002）、944-258（13034）各一个（定标卡全部缺货） | 更换903-364一个备件和重新定标后修复 by jesson | SWO-242175 | Defect | HW | Process Module | Optical Unit | 1024 | Reported 1024 error | Reported 1024 errors, borrowing 903-364 one, 944-212 (13147), 944-216 (13052), 944-214 (13002), 944-258 (13034) each (all cards are out of stock) | Replace 903-364 a spare part and repair by jesson after recalibration
No | Visit | Complaint | RCHN | 测试卡卡口坏了 | 测试卡卡口坏了 | 更换测试卡后修复 | SWO-243542 | Assisted Replacement | Consumables | Test Kit | Cartridge unknown |  | The test card mouth is broken. | The test card mouth is broken. | Repair after replacing test card
No | Visit | Complaint | RKK | 納品依頼pO2ゼロ点が高い。メンブラン交換しても改善せず。 | lot違うメンブランに交換してもゼロ点高め。pO2電極初期不良疑い。 | pO2電極交換実施。 | SWO-242861 | Assisted Replacement | Electrode | Electrode | pO2 |  | Delivery request pO2 zero point is high. It does not improve even if the membrane exchange. | It raises zero point even if it exchanges it with a lot different Membranes. PO2 electrode initial defect suspected. | PO2 electrode replacement was implemented.
No | Visit | Complaint | RDAN | Når Pia scanner sit kort for at logge ind på ABLen, så giver scanneren et bib, men Pia bliver ikke logget ind. Lidt efter giver scanneren 2 bib fra sig. | Skiftet div. | Kabel til skanner skiftet CPU skiftet Skærm skiftet Hele computer enheden skiftet W7 + ABL software geninlæst - derefter ok | SWO-243324 | Defect | HW | Basic Unit | Cable, Barcode Scanner |  | When Pia scans his card to log in to the ABLen, the scanner gives a bib, but Pia is not logged in. | Shift Div. | Scanner Switch Cable CPU Change Screen change The entire computer unit switched W7 + ABL software re-read-then OK
No | Visit | Complaint | RSCH | Reparatur | Vermehrt Fehler 0882 Mischerträger Schlupf erkannt Wurde schon mehrmals kalibriert | Allgemeine Kontrolle i.o Komplettes FlexQ ersetzt i.o FlexQ kalibriert i.o System i.o | SWO-243484 | Defect | HW | FlexQ Module | FlexQ Module |  | Repair | Increases error 0882 Mixer carrier slip detected Has been calibrated several times | General control i.o Complete FlexQ replaces i.o FlexQ calibrates i.o System i.o
No | Visit | Complaint | RDEU | Nach einem kontrolliertem Neustart steht der AQT mit Fehler, auch ein weiterer aufgeforderter Neustart bringt keine Besserung | Reinigungsarbeiten, Fehlersuche, Fehlercode 3264 (Environment) | Im Service SafePositions kontrolliert ok, Datum und Uhrzeit kontrolliert und korrigiert sowie gespeichert, Neustart ok mit anschließender Wartung | SWO-242975 | Setup | SW | RMED SW | Time setting | 3264 Email (LEH) | After a controlled restart, the AQT is errored, even another prompted restart does not improve | Cleaning, troubleshooting, error code 3264 (Environment) | In the safepositions service checks ok, checks and corrects date and time as well as saved, restart ok with subsequent maintenance
No | 19 | Y | RFRA | Blank | Lenteur dans la remonté de demo |  | D347AA4112492916 | Defect | SW | RMED SW | RMED SW | Slow | Blank | Slowness in the demo upwelling
No | Visit | Complaint | RAME | CAS team member called on behalf of customer to report that the demographics are not coming from the LIS/HIS middle ware. | Demographics do not display on Patient ID screen after accession # is imputed. Analyzer says request timed out. | Turn on trace log in Roma.ini file and reboot analyzer. Got a trace file from analyzer and it doesn't show a query response from Sunquest. Also, got a trace file from Sunquest and it shows same thing. Customer will work with Sunquest. I will leave trace log enabled for now. | SWO-243652 | No fault found | His/Lis | His/Lis | His/Lis | sunquest | CAS team member called on behalf of customer to report that the demographics are not coming from the LIS/HIS middle ware. | Demographics do not display on Patient ID screen after accession # is imputed. Analyzer says request timed out. | Turn on trace log in Roma.ini file and reboot analyzer. Got a trace file from analyzer and it doesn't show a query response from Sunquest. Also, got a trace file from Sunquest and it shows same thing. Customer will work with Sunquest. I will leave trace log enabled for now.
Yes | Visit | Complaint | RDEU | 23.08.2017, 09:43 Uhr, Fr. Nowack, Tel. 02118604030: Kal Fehler bei pCO2 und Lac 23.08.2017, 10.55. Uhr, frau Nowack, tel: 02118604030: weiterhin Kal Fehler bei Lac. 28.0/8.2017, 9.20Uhr, Frau Nowak: Tel.: 02118604030: auch mit neuen Betriebsmittel-ständig Kalibrationsfehler | Geräte Auslieferung erforderlich. | Sensor Kassette installiert und LP installiert und Gerät ausgeliefert. Leihgerät wieder mitgenommen. | SWO-243679 | Re classified RMED | Logistical upSWO |  |  |  | 23.08.2017, 09:43, Fr. Nowack, Tel. 02118604030: Kal error in pCO2 and Lac 23.08.2017, 10.55. a.m., mrs. Nowack, tel: 02118604030: continue Kal error at Lac. 28.0/8.2017, 9.20a,mrs. Nowak: Tel.: 02118604030: also with new equipment-permanent calibration error | Devices delivery required. | Sensor cassette installed and LP installed and device delivered. borrower.
No | Visit | Complaint | RSCH | Lac Kalibration fehler. | Parametro non correttamente funzionante. | Sostituita sensor interface PCB. | SWO-244306 | Defect | HW | Basic Unit | PCB, Sensor interface | 378 | Lac calibration error. | Parameter is not functioning properly. | Replaced PCB sensor interface.
No | Visit | Complaint | RIND | Sc temp not stable. | Checked Sc and flushed it. Then reinstalled SC.calibration was successful. Then checked with samples and again problem repeat. | Replaced SC and reinstalled it. Then also same problem happened. Need to check again after replace sensor connector board. | SWO-243746 | [5 empty] | Sc temp not stable. | Checked Sc and flushed it. Then reinstalled SC.calibration was successful. Then checked with samples and again problem repeat. | Replaced SC and reinstalled it. Then also same problem happened. Need to check again after replace sensor connector board.
No | Visit | Complaint | RCHN | 自动关机，再开机测试卡显示失效 | 自动关机，再开机测试卡显示失效，借431-020一个 | 更换431-020后修复 by jesson | SWO-242722 | [5 empty] | Automatic shutdown, and then power on test card display is invalid | Automatic shutdown, and then power on test card display failure, borrow 431-020 one | Repair by jesson after replacement 431-020
No | Visit | Complaint | RESP | Error de fuga | Error de fuga | Limpieza interna de tubos Sustitucion de inlet gasket y sp Cal , qc y muestras | SWO-243791 | Assisted Replacement | SP | SP | Leak current |  | Leak error | Leak error | Internal pipe cleaning Substitution of inlet gasket and sp Cal , qc and samples
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | Lac定标不过，质控也不过 | 连续更换3张测试卡后和1个试剂包修复 | SWO-243881 | Assisted Replacement | SC | SC | Lac |  | Test card, reagent package failure | Lac calibration, but quality control is not | After 3 consecutive test cards and 1 reagent kit repair
No | Visit | Complaint | RESP | INLET NO FUNCIONA | ERROR 1391 | INLET FUERA DE POSISCION POR DETECTAR UN OBSTACULO QUE NO EXISTE, POR SERVICIO SE INTENTA AJUSTAR, SIENDO NECESARIO SUSTITUIR LA SONDA Y VORVER A POSICIONAR MECANICAMENTE EL MODULO INLET. CALIBRADO Y VERIFICADO. SE EFECTUAN PRUEBAS DE CALIBRACIONES TRAS INICIALIZACION Y PRUEBAS DE MUESTRAS Y Q.C. | SWO-243779 | Adjustment | HW | Auto Inlet module | Inlet Position Cal. |  | INLET NO WORKS | ERROR 1391 | INLET OUT OF POSISCION FOR DETECTING AN OBSTACLE THAT DOES NOT EXIST, BY SERVICE IS TRYING TO ADJUST, BEING NECESSARY TO SUBSTITUTE THE SOUND AND VORVER TO POSITION THE MODULE INLET. CALIBRATED AND VERIFIED. CALIBRATION TESTS ARE EFFECTED AFTER INITIALIZATION AND SAMPLE TESTS AND Q.C.
No | Visit | Complaint | RCHN | 系统软件不匹配 | 系统软件不匹配，需要借件910-343 | 更换910-343后修复by jesson | SWO-242790 | Defect | HW | Analyzer Control Module | CPU unit | 0002 | System software does not match | System software does not match, need to borrow 910-343 | Repair by jesson after replacement 910-343
Yes | Visit | Complaint | RCHN | TNT定标失败 | 由于升级软件后，显示不正常，重新安装软件后，需要重新定标TNT，借944-268（13188）和944-220（12412）缺货各一张 | 由于升级软件后，显示不正常，重新安装软件后，需要重新定标 更换944-268（13188）后修复 by jesson | SWO-242704 | Re classified RMED | FAN |  |  |  | TNT calibration failure | Due to the upgrade of the software, the display is not normal, reinstall the software, need to re-label TNT, borrow 944-268 (13188) and 944-220 (12412) out of stock each one | Due to the incorrect display after upgrading the software, the recalibration of the software requires recalibration Fix by jesson after replacement 944-268 (13188)
No | Visit | Complaint | RCHN | 光学强度超范围 | 光学强度超范围，借903-364一个 | 工程师称因收到的光单有问题，故需加借备件，故此单关单 by jesson | SWO-243404 | Revisit required | HW | Process Module | Optical Unit | 1031 | Optical strength over range | Optical strength over range, borrowed 903-364 a | The engineer said that due to the problem with the light sheet received, it was necessary to borrow spare parts, so the single pass by jesson
No | Visit | Complaint | RAME | Leaking from waste line | Checked data logs for errors, checked flow of solutions, checked waste system | Replaced waste line from waste connector to wast bottle | SWO-243691 | Defect | HW | Basic Unit | Waste Tube |  | Leaking from waste line | Checked data logs for errors, checked flow of solutions, checked waste system | Replaced waste line from waste connector to wast bottle
No | Visit | Complaint | RAME | TREND FOR PCO2 FOUR LAST SEVERAL MONTHS HAS BEEN MOVING HIGH FOR LEVELS 1 & 2 OF QC COMPARED TO PEER GROUP. NOW OUTSIDE PEER RANGE. | TESTED RESULTS FROM ANOTHER PCO2 ELECTRODE (NOT NEW) AND FOUND SLIGHT IMPROVEMENT OF RESULTS, HOWEVER DID NOT CORRECT CLOSE TO MEAN OF PEER RESULTS. ORDERED NEW ELECTRODE AND INLET HEATER. | NEW ELECTRODE INITIALLY CAME IN AT PEER MEAN, BUT DRIFTED BACK UP. REPLACED INLET PREHEATER. LET SETTLE FOR SEVERAL HOURS WHILE ATTENDING SERVICE CALL AT ANOTHER LOCATION. RETURNED TO FIND QC VALUES IN A MUCH MORE ACCEPTABLE RANGE. | SWO-242003 | Defect | HW | Inlet module | Preheater |  | TREND FOR PCO2 FOUR LAST SEVERAL MONTHS HAS BEEN MOVING HIGH FOR LEVELS 1 & 2 OF QC COMPARED TO PEER GROUP. NOW OUTSIDE PEER RANGE. | TESTED RESULTS FROM ANOTHER PCO2 ELECTRODE (NOT NEW) AND FOUND SLIGHT IMPROVEMENT OF RESULTS, HOWEVER DID NOT CORRECT CLOSE TO MEAN OF PEER RESULTS. ORDERED NEW ELECTRODE AND INLET HEATER. | NEW ELECTRODE INITIALLY CAME IN AT PEER MEAN, BUT DRIFTED BACK UP. REPLACED INLET PREHEATER. LET SETTLE FOR SEVERAL HOURS WHILE ATTENDING SERVICE CALL AT ANOTHER LOCATION. RETURNED TO FIND QC VALUES IN A MUCH MORE ACCEPTABLE RANGE.
No | 3 | Y | RCZE | Blank |  |  | D34B6166744A2975 | Assisted Replacement | Electrode | Electrode | Glu |  | Blank
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 定标失败 | 更换测试卡 | SWO-247790 | [5 empty] | Test card, reagent package failure | Bid failed | Replace the test card
No | Visit | Complaint | RCHN | 进压口断裂 | 进压口断裂，借903-584一个 | 更换903-584后修复 by jesson | SWO-243911 | Defect | HW | Cover and Chassis | Tube holder |  | Press inlet break | The inlet is broken, borrowed 903-584 one | Repair by jesson after replacing 903-584
No | Visit | Complaint | RSAF | Instrument Failing to aspirate | Instrument Failing to aspirate samples | Replaced with a new waste tubing line, replaced with a new probe due blockage and also repklaced a sesnsor cassette. | SWO-244119 | Replaced | HW | Auto Inlet module | Probe |  | Instrument Failing to aspirate | Instrument Failing to aspirate samples | Replaced with a new waste tubing line, replaced with a new probe due blockage and also repklaced a sesnsor cassette.
No | Visit | Complaint | RCHN | 检测到汞定标不过 | 检测到汞定标不过，借903-165（被903-644替代）一个 |  | SWO-243969 | Defect | HW | Basic Unit | PCB, Wet Sec Control | 1178 | Mercury calibration detected but | Mercury was detected, however, by borrowing 903-165 (replaced by 903-644) one
No | Visit | Complaint | RCHN | 光学强度超范围 | 光学强度超范围，借903-364一个，需加借903-364一个 |  | SWO-243952 | Defect | HW | Process Module | Optical Unit | 1031 | Optical strength over range | Optical strength over range, borrow 903-364 one, need to borrow 903-364 one
No | Visit | Complaint | RAME | Leak detect error | Leak detect errors # 964, looked at diagnostics and all were passing, flushed unit and reinstalled SP but error came back, will order parts and return 10 $17 $2017 am | Flush and retest, | SWO-247734 | Assisted Replacement | SP | SP | Leak current | 964 | Leak detect error | Leak detect errors # 964, looked at diagnostics and all were passing, flushed unit and reinstalled SP but error came back, will order parts and return 10 /17 /2017 am | Flush and retest,
No | 6 | Y | RFRA | Blank | Fermeture sensor bay avec sensor simuilator impossible |  | D34BAD1A74184F20 | Adjustment | HW | Basic Unit | Sensor Bay Unit |  | Blank | Sensor bay closure with sensor simuilator impossible
No | 3 | Y | RCZE | Blank | teèe pod RP, viz pøedchozí servis, kde byl použit manifold s roèního kitu |  | D34B7DEF745384DC | [5 empty] | Blank | Teèe pod RP, viz pøedchozí Servicing, KDE Byl Použit manifold S roèního kitu
No | Visit | Complaint | RKK | インレット付近から液漏れ | インレット付近から液漏れ | 内部のチューブが踏みつぶされていた おそらく製造段階で踏まれていた様子 早い段階で交換が必要 | SWO-247775 | [5 empty] | Leaking liquid from near inlet | Leaking liquid from near inlet | The tube inside was trampled. Probably stepped on the manufacturing stage Need to be replaced early
No | Visit | Complaint | RCHN | 样本报290和581错误，溶血器不过 | 样本报290和581错误，溶血器不过，借902-980一个 | 更换902-980后修复by tim | SWO-243399 | Defect | HW | Oxi module | Hemolyzer | 290-581 | sample this newspaper 290 and 581 errors, hemolytic stoics but | Sample this newspaper 290 and 581 errors, hemolytic, however, borrow 902-980 a | Repair by tim after replacement 902-980
No | Visit | Complaint | RCHN | 汞定标不过 | 汞定标不过，借903-205一个 | 更换备用机上的903-205后修复by tim | SWO-243951 | Defect | HW | Basic Unit | Waste connector |  | Mercury calibration is no more than | Mercury calibration, however, borrows 903-205 a | Repair by tim after replacing the 903-205 on the standby
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | GLU定标失败 1310 | 更换测试卡正常。 | SWO-244232 | [5 empty] | Test card, reagent package failure | GLU calibration failure 1310 | The replacement test card is OK.
No | Visit | Complaint | RESP | NO INICIALIZA. | PRUEBAS CON TEST DE SERVICIO | TRAS CAMBIO DE DIFERENTES REACTIVOS SIN EXITO Y COMPROBACION DE NEUMATICA SE DETECTA VALVULA OXI LA CUAL EJERCE POCA PRESION Y MOVIMIENTO, SE DESMONTA ESTA SE LIMPIA Y SE EFECTUAN PRUEBAS SIENDO ESTAS CORRECTAS. SE DEJAN INLET PARA VARIOS SERVICIOS. CALIBRADO Y VERIFICADO CON Q.C. | SWO-249448 | Cleaned | HW | Oxi module | Pinch valve |  | IT DOES NOT INITIALIZE. | TESTS WITH SERVICE TEST | AFTER CHANGE OF DIFFERENT REACTIVES WITHOUT EXIT AND NEUMATICA CHECK, OXI VALVE IS DETECTED WHICH SOME PRESSURE AND MOVEMENT, THIS IS CLEANED AND TESTS ARE EFFECTED BEING THESE CORRECT. THEY ARE LEFT INLET FOR VARIOUS SERVICES. CALIBRATED AND VERIFIED WITH Q.C.
No | 11 | Y | RPNZ | Blank | Advice line call re AQURE connection failing as results were not going through to the LIS |  | D34B5150744E121A | User error | HW | His/lis | Cable connection |  | Blank | Advice line call re AQURE connection failing as results were not going through to the LIS
No | Visit | Complaint | RAME | Customer reported leakage from syringe port and waste tubing not staying connected. | Examine all waste tubing. Found clots in tubing leading into waste bottle. | Replaced clear tubing leading into waste bottle from extra tubing stock. Also replaced inlet tubing. Glucose unstable on level 1 QC run. Check and cleaned electrode. Restart calibrations. Replaced Glucose and pO2 membranes. Restart calibrations. | SWO-247630 | Defect | HW | Basic Unit | Waste Tube | Clot | Customer reported leakage from syringe port and waste tubing not staying connected. | Examine all waste tubing. Found clots in tubing leading into waste bottle. | Replaced clear tubing leading into waste bottle from extra tubing stock. Also replaced inlet tubing. Glucose unstable on level 1 QC run. Check and cleaned electrode. Restart calibrations. Replaced Glucose and pO2 membranes. Restart calibrations.
No | 8 | Y | RCZE | Blank | rozhonová mìnila membránu na pCO2 kde nevycházela citlivost ( byla nízká ), i po opakované výmìnì stále stejné |  | D34B6C0F74357C6A | Revisit required | Electrode | Electrode | pCO2 |  | Blank | The membrane was changed to pCO2 where there was no sensitivity (low), even after repeated replacement
No | Visit | Complaint | RSCH | Reparatur | Halogen intensity too high | Halogen justiert i.o Kal. und QC i.o Sytem i.o | SWO-249347 | Adjustment | HW | Oxi module | Lamp unit | Halogen | Repair | Halogen intensity too high | Halogen justiert i.o Kal. und QC i.o Sytem i.o
No | Visit | Complaint | RDEU | Herr Buchholz MTA (Tel. 0152 0872 0951) hat Troponin Kal gestartet. Diese wurde mit Hintergrundausreißer abgebrochen. Weitere Kal Troponin / neues Kit ist nicht vorhanden. Er bittet um Überprüfung des Gerätes und um weiteres Material. | während der Kalibrierung ist der Fehler: 5684: Rotate Process Wheel 40 aufgetreten eine Pat. Messungen wurde ebenfalls abgebrochen mit 1218: keine ausreichende Probe | visueller Check Incubator Wheel und Process Wheel $ im Service Mode beides getestet $ Process Wheel gereinigt Blank Messungen durchgeführt Liquid Path Needle - > über P4 zum Abfall überprüft $ Filter gereinigt SD gezogen - Durchfluß i.O. neue Charge TnT kalibriert Test Kit Lot: 13424 mit dem abgebrochenem Lot: 13423 1:1 getauscht | SWO-247601 | Defect | Consumables | Test Kit | Cartridge, Broken | 5684 | Mr Buchholz MTA (tel. 0152 0872 0951) started Troponin Kal. This was aborted with background outlier. More Kal Troponin/new kit is not available. He asks for inspection of the device and other material. | during calibration, error: 5684: Rotate Process Wheel 40 occurred a Pat. Measurements were also aborted with 1218: insufficient sample | Visual Check Incubator Wheel and Process Wheel - tested both in service mode - Process Wheel cleaned Blank measurements performed Liquid Path Needle - > checked via P4 for waste . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . SD drawn - flow i.O. new batch TnT calibrated Test Kit Lot: 13424 with the aborted lot: 13423 1:1 exchanged
No | Visit | Complaint | RDEU | 29.09.2017; Fr. Haug; Tel.: 0832396390; CRP QC Messfehler Level 2 | CRP LQC oberhalb Bereich. | Servicedump gezogen - Fließwegwerte ok (1179 - 1265). Vorsorglich P1 und P2 erneuert und AQT mit Hypochlorid gereinigt. 5 Blankmessungen durchgeführt. Servicedump gezogen und ausgewertet. Fließwege ok (1135 - 1178). Photocounts und Messergebnisse der Blankmessungen ok. Systemüberprüfung und LQC's durchgeführt. | SWO-247634 | Defect | HW | Process Module | Pump1 & 2 |  | 29.09.2017; Fr. Haug; Tel.: 0832396390; CRP QC Messfehler Level 2 | CRP LQC above range. | Service dump pulled - flow path values ok (1179 - 1265). As a precaution P1 and P2 renewed and AQT cleaned with hypochloride. 5 blank measurements performed. Service dump dragged and evaluated. Flow paths ok (1135 - 1178). Photocounts and measurement results of blank measurements ok. system verification and LQC's.
Yes | Visit | Complaint | RDEU | nach Wechsel d. Sensor Interface PCB laut FAN (915-371), Gerät mit Crea-Fehler nach Kal. | erneuter Tausch der Sensor Interface PCB | Gerät aus; Platine gew. (alt: 903-213R2003N0508); Gerät ein; Startphase abgewartet; Kal. i.O.; QC = Urea-Fehler; LP gew.; Initialisierung; Kal. + QC i.O. | SWO-247979 | Re classified RMED | FAN |  |  |  | after changing the Sensor Interface PCB according to FAN (915-371), device with Crea error to Kal. | re-exchange of the Sensor Interface PCB | device off; board. (old: 903-213R2003N0508); device; start-up phase awaited; Kal. i.O.; QC = Urea error; LP. Initialization; Kal. + QC in the first place
No | Visit | Complaint | RDEU | Kalibrationsfehler Crea | cCrea instabil | Gerät überprüft, Sensorkassette ausgetauscht und reklamiert, manuelle QC durchgeführt und Servicedump gezogen | SWO-247921 | Assisted Replacement | SC | SC | Crea |  | Kalibrationsfehler Creates | cCrea unstable | Device checked, sensor cartridge replaced and complained, manual QC performed and service dump pulled
No | Visit | Complaint | RCHN | 试剂包锁松不开 | 试剂包锁松不开，借580-400一个 | 更换580-400 | SWO-247851 | Defect | HW | Basic Unit | Stepper motor, Actuator |  | Reagent pack lock loose | Reagent pack lock loose, borrow 580-400 one | Replace 580-400
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | GLU,LAC,PCO2定标失败 | 更换三张测试卡正常。 | SWO-247900 | Assisted Replacement | SC | SC | Glu |  | Test card, reagent package failure | GLU, LAC, PCO2 bid failed | Replacing the three test cards is OK.
No | Visit | Complaint | RLTD | Customer TCM5 monitors are cracking near sensor chamber. | Attended site to swap out monitors for repair and install loan units. Removed customer TCM5s serial nos. R0021N003 (9), R0025N003 (2), R0025N007 (4), R0026N006 (11), R0026N009 (6) all removed due to cracks in the casing. Also removed customer TCM5 serial no. R0026N007 (10) for upgrade. Radiometer loan unit R0043N001 giving monitor error 05000010 - Heating error, removed from site for investigation. Radiometer loan unit R0057N015 giving monitor error 05000029 - Gas flow error during calibration, removed from site for investigation. Supplied 8 loan TCM5 monitors serial nos. R0039N001, R0040N001, R0040N009, R0041N002, R0041N005, R0041N009, R0054N007, R0054N010. Customer set-up copied to loan monitors, batteries, gas bottles and sensors swapped from removed units into the loan units. | Removed 8 TCM5 monitors for repair, installed 8 loan TCM5s. | SWO-249413 | [5 empty] | Customer TCM5 monitors are cracking near sensor chamber. | Attended site to swap out monitors for repair and install loan units. Removed customer TCM5s serial nos. R0021N003 (9), R0025N003 (2), R0025N007 (4), R0026N006 (11), R0026N009 (6) all removed due to cracks in the casing. Also removed customer TCM5 serial no. R0026N007 (10) for upgrade. Radiometer loan unit R0043N001 giving monitor error 05000010 - Heating error, removed from site for investigation. Radiometer loan unit R0057N015 giving monitor error 05000029 - Gas flow error during calibration, removed from site for investigation. Supplied 8 loan TCM5 monitors serial nos. R0039N001, R0040N001, R0040N009, R0041N002, R0041N005, R0041N009, R0054N007, R0054N010. Customer set-up copied to loan monitors, batteries, gas bottles and sensors swapped from removed units into the loan units. | Removed 8 TCM5 monitors for repair, installed 8 loan TCM5s.
No | 27 | Y | RCZE | Blank |  |  | D34D2E30EA45298B | Defect | HW | pH/BG Module | Pump tubes |  | Blank
No | Visit | Complaint | RIND | Oxymetry not calibrating. | Oxymetry not calibrating. intervention required is the message. | Try to Service by replacing one by one parts and reconnect the oxy PCB. and calibrate. No repalcement required. analyser serviced and Calibrated properly. Handed over in a fine working condition. | SWO-250433 | No fault found | HW | Oxi module | Oxi module |  | Oxymetry not calibrating. | Oxymetry not calibrating. intervention required is the message. | Try to Service by replacing one by one parts and reconnect the oxy PCB. and calibrate. No repalcement required. analyser serviced and Calibrated properly. Handed over in a fine working condition.
Yes | Visit | Complaint | RAME | TREND FOR PCO2 FOUR LAST SEVERAL MONTHS HAS BEEN MOVING HIGH FOR LEVELS 1 & 2 OF QC COMPARED TO PEER GROUP. NOW OUTSIDE PEER RANGE. | COURTESY VISIT FOLLOW UP WITH CUSTOMER WHILE IN TULSA. INTRODUCTION OF NEW SLAES REP TO HILLCREST PERSONNEL. | NONE REQUIRED. SPOKE WITH RAY NORTH, SUPERVISOR. INTRODUCED HIM TO NEW SALES REP CASSIDY. | SWO-249583 | Re classified RMED | Courtesy |  |  |  | TREND FOR PCO2 FOUR LAST SEVERAL MONTHS HAS BEEN MOVING HIGH FOR LEVELS 1 & 2 OF QC COMPARED TO PEER GROUP. NOW OUTSIDE PEER RANGE. | COURTESY VISIT FOLLOW UP WITH CUSTOMER WHILE IN TULSA. INTRODUCTION OF NEW SLAES REP TO HILLCREST PERSONNEL. | NONE REQUIRED. SPOKE WITH RAY NORTH, SUPERVISOR. INTRODUCED HIM TO NEW SALES REP CASSIDY.
No | Visit | Complaint | RDAN | Lotte Petersen melder at der bliver skiftet gas 1 ofte på ABLen. Lotte har lagt 2 gasbeholder som der er skiftet i dag den 16/10 2017. | Ventil tjek Gas tjek | Ingen fejl fundet - alle flasker med fejl havde samme lot nr. | SWO-249344 | Assisted Replacement | Consumables | Gas Cylinder | Gas Cylinder |  | Lotte Petersen reports that gas 1 is often switched on the ABLen. Lotte has put 2 gas tank as changed today on 16/10 2017. | Valve Check Gas Check | No errors found-all bottles with defects had the same lot nr.
No | Visit | Complaint | RAME | Getting inlet liquid sensor error, | Looked in data logs found inlet liquid sensor error and pump calibration error for ph $BG, EL $Met,and oxi. | Replaced all pump tubing and gray tubing from maintenance kit. Cleaned measuring chamber and noticed leak on bottom of measuring chamber. Tried it clean measuring chamber and liquid sensor. Still failing replaced inlet liquid sensor. | SWO-247765 | Revisit required | HW | Inlet module | LS |  | Getting inlet liquid sensor error, | Looked in data logs found inlet liquid sensor error and pump calibration error for ph /BG, EL /Met,and oxi. | Replaced all pump tubing and gray tubing from maintenance kit. Cleaned measuring chamber and noticed leak on bottom of measuring chamber. Tried it clean measuring chamber and liquid sensor. Still failing replaced inlet liquid sensor.
No | Visit | Complaint | RAME | Inconsistent software versions. Pls contact tech support. | Confirmed issue when onsite. | Proceed to install system software v 6.17 and reboot analyzer. Error message is removed. | SWO-249579 | Defect | SW | RMED SW | RMED SW |  | Inconsistent software versions. Pls contact tech support. | Confirmed issue when onsite. | Proceed to install system software v 6.17 and reboot analyzer. Error message is removed.
No | Visit | Complaint | RIND | Oxymetry values not coming. | Oxy Metry not OK. | Try to replce Oxy PCB and Helogen lam Iintencity to set. But Calibration Drift not OK. Try different combinations but the sam e problem reported. TO Sending for Hemolyser unit, | SWO-250438 | Revisit required | HW | Oxi module | Hemolyzer |  | Oxymetry values not coming. | Oxy Metry not OK. | Try to replce Oxy PCB and Helogen lam Iintencity to set. But Calibration Drift not OK. Try different combinations but the sam e problem reported. TO Sending for Hemolyser unit,
No | Visit | Complaint | RDAN | Blank Data | NA | Beslag til døren monteret korekt Skanner til test kassetter renset | SWO-249341 | Defect | HW | Cover and Chassis | Cartridge door |  | Blank Data | In | Bracket for door mounted Corekt Scanner for test cassettes cleaned
No | Visit | Complaint | RAME | Po2 out on External QC | pO2 out on multiple levels of QC. Possible silicone contamination from syringes. | Flushed fluid path multiple times with warm water. Restart calibrations. | SWO-250277 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway | Silicone? | Po2 out on External QC | pO2 out on multiple levels of QC. Possible silicone contamination from syringes. | Flushed fluid path multiple times with warm water. Restart calibrations.
No | Visit | Complaint | RAME | Printer disconnected | Printer continues to disconnect | Installed new printer | SWO-250269 | Defect | HW | Printer Module | Printer |  | Printer disconnected | Printer continues to disconnect | Installed new printer
No | Visit | Complaint | RAME | Blank screen | Blank screen. | Replaced display | SWO-250164 | Defect | HW | DMS_PC | Touch screen |  | Blank screen | Blank screen. | Replaced display
No | Visit | Complaint | RLTD | The inlet gasket was not returning to the correct position after a sample. | When a sample was being run the inlet would not close and the sample was lost. Found that the inlet probe had a bend in it which prevented the inlet gasket returning to the correct position after running the sample. | Replaced the inlet probe. | SWO-250030 | Replaced | HW | Auto Inlet module | Probe bend |  | The inlet gasket was not returning to the correct position after a sample. | When a sample was being run the inlet would not close and the sample was lost. Found that the inlet probe had a bend in it which prevented the inlet gasket returning to the correct position after running the sample. | Replaced the inlet probe.
Yes | 7 | Y | RPAC | Blank | wrong SWO |  | D34E178D604CF95A | Re classified RMED |  |  |  |  | Blank | wrong SWO
No | Visit | Complaint | RAME | Analyzer detected leak. | Analyzer down. Customer installed new Solution Pack. Leak detected. | Re-installed Solution Pack. Re-installed an old Sensor cassette of the customers. Completed all service tests. Installed other Solution pack. Completed all service tests. Field service engineer will put in for credit of Solution pack and Sensor cassette. | SWO-250138 | Assisted Replacement | SP | SP | Leak detected |  | Analyzer detected leak. | Analyzer down. Customer installed new Solution Pack. Leak detected. | Re-installed Solution Pack. Re-installed an old Sensor cassette of the customers. Completed all service tests. Installed other Solution pack. Completed all service tests. Field service engineer will put in for credit of Solution pack and Sensor cassette.
No | Visit | Complaint | RCHN | 无法开机 | 无法开机，加借903-381一个 | 更换903-381,905-914后修复by tim | SWO-247839 | [5 empty] | Cannot power on | Cannot power on, plus 903-381 one | Repair by tim after replacing 903-381, 905-914
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | QC2溶液不均匀 1086 | 更换试剂包正常。 | SWO-251040 | [5 empty] | Test card, reagent package failure | QC2 solution uneven 1086 | The replacement kit is OK.
Yes | 5 | Y | RPAC | Blank | TCM5 constantly failing to calibrate. Customer followed Radiometer instructions to remembrane numerous times without success. Possible Sensor 92 issue. |  | D34E126D602483CA | [5 empty] | Blank | TCM5 constantly failing to calibrate. Customer followed Radiometer instructions to remembrane numerous times without success. Possible Sensor 92 issue.
Yes | 5 | Y | RPOL | Blank | sprawdzenie analizatora, wymiana kruæca od strzykawki, kalibracja czujników, pomp. Analizator sprawny technicznie. |  | D34E4D43603EEE2D | [5 empty] | Blank | Analyzer, replacing the syringe, calibrating the sensors, pumps. Technically efficient analyzer.
No | Visit | Complaint | RAME | OXI hardware error 0579 | OXI hardware error. Lamp unit fails complete OXI test. | Replaced Lamp unit. Performed complete alignment. | SWO-249602 | Defect | HW | Oxi module | Oxi module, Replaced | 579 | OXI hardware error 0579 | OXY hardware error. Lamp unit fails complete OXI test. | Replaced Lamp unit. Performed complete alignment.
No | Visit | Complaint | RLTD | Ranvir 01865 221596 | The line from the waste pump tubes y-piece to the waste bottle was blocked. | The blockage was where the waste line connected to the waste bottle nipple, I disconnected the waste line and was able to flush the blockage out. I then flushed water through several times to make sure the line was completely clear. | SWO-250047 | Cleaned | Fluidic pathway | Fluidic pathway | Y-Piece |  | Ranvir 01865 221596 | The line from the waste pump tubes y-piece to the waste bottle was blocked. | The blockage was where the waste line connected to the waste bottle nipple, I disconnected the waste line and was able to flush the blockage out. I then flushed water through several times to make sure the line was completely clear.
No | 8 | Y | RCZE | Blank | nelze zkalibrovat pCO2, citlivost kolem 50% i po výmìnì membrány |  | D34F1974D62BDE02 | Assisted Replacement | Electrode | Electrode | pCO2 |  | Blank | Cannot calibrate pCO2, sensitivity around 50% even after membrane replacement
Yes | 1 | Y | RCZE | Blank |  |  | D34F1F20D635DE06 | Re classified RMED | Setup / Installation |  |  |  | Blank
No | Visit | Complaint | RSAF | Customer called to inform us that the analyzer is in Automatic recovery in process, Switches off and reboots, undefined look up string and process error | at arrival at analyzer it was in working condition and traffic light green, looked up in activity log for errors - Automatic recovery successful, process unit 24 and undefined lookup string 50.15.1008.0.-1.-1 | Entered service mode and ran safe positions and also done a system check which passed, Ran 2 x LQC's Multi checks L2 - ok 2 x LQC's Multi checks L1 - ok | SWO-250405 | No fault found | No fault found | No fault found | No fault found |  | Customer called to inform us that the analyzer is in Automatic recovery in process, Switches off and reboots, undefined look up string and process error | at arrival at analyzer it was in working condition and traffic light green, looked up in activity log for errors - Automatic recovery successful, process unit 24 and undefined lookup string 50.15.1008.0.-1.-1 | Entered service mode and ran safe positions and also done a system check which passed, Ran 2 x LQC's Multi checks L2 - ok 2 x LQC's Multi checks L1 - ok
Yes | 1 | Y | RPOL | Blank | FAN sensor board replacement |  | D34F37ADD63D1CBA | Re classified RMED | FAN |  |  |  | Blank | FAN sensor board replacement
No | Visit | Complaint | RCHN | 文件目录错误 | 文件目录错误，借903-381一个。 | 更换903-381后修复by tim | SWO-249799 | [5 empty] | File directory error | File directory error, borrow 903-381 one. | Repair by tim after replacement 903-381
Yes | Visit | Complaint | RDEU | 19.10.2017 14:15Uhr Sr. Gisela 061157714590: Kal Fehler bei Elektrolyten |  |  | SWO-251703 | Re classified RMED | Request |  |  |  | 19.10.2017 14:15pm Sr. Gisela 061157714590: Kal fault in electrolytes
No | Visit | Complaint | RAME | Fluid build into ph chanber. | Fluid is building up in the pH well of the measuring chamber and leaking from the area inlet gasket. As a result, calibrations and QC's are failing, rendering multiple parameters red. | Cleared a a clot in the capillary side of the waste line and waste fitting connection. Verified that leak is no longer present and fluid is not observed in the pH well. Performed the necessary calibrations and QC's to bring all parameters into calibration range. | SWO-250867 | Cleaned | Fluidic pathway | pH/BG Module | Measuring chamber, Clot |  | Fluid build into ph chanber. | Fluid is building up in the pH well of the measuring chamber and leaking from the area inlet gasket. As a result, calibrations and QC's are failing, rendering multiple parameters red. | Cleared a a clot in the capillary side of the waste line and waste fitting connection. Verified that leak is no longer present and fluid is not observed in the pH well. Performed the necessary calibrations and QC's to bring all parameters into calibration range.
No | Visit | Complaint | RKK | QCでｐCO2にエラー | pCO2がレベル２と３で範囲外。レベル3を再測定でｐHも範囲外になる | pCO2メンブランの気泡除去、PH電極先端クロット除去、Cal実施後、QCOK | SWO-251080 | User error | Electrode | Membrane | pCO2 |  | Error in pCO2 in QC | pCO2 is out of range at levels 2 and 3. Level 3 re-measure sph out of range | Bubble removal of pCO2 membranes, PH electrode tip clot removal, Cal after implementation, QCOK
No | Visit | Complaint | RDAN | Blank | NA | Renset pH målekammer Skiftet fodpakning | SWO-250690 | Cleaned | Fluidic pathway | pH/BG Module | Sealing ring | Sealing Ring left in Meas ch | Blank | In | Purified pH Measuring Chamber Changed foot gasket
No | Visit | Complaint | RAME | Instrument is leaking from the front | Inspected waste tubing for proper install. Inspected inlet gasket, inlet tubing for damage. | Found inlet tubing with a small hole and also flushed out waste line. Used customers inlet tubing for replacement. | SWO-250878 | Defect | HW | Inlet module | Tube, Inlet |  | Instrument is leaking from the front | Inspected waste tubing for proper install. Inspected inlet gasket, inlet tubing for damage. | Found inlet tubing with a small hole and also flushed out waste line. Used customers inlet tubing for replacement.
No | Visit | Complaint | RKK | pCO2のQCが、全レベル下限値ギリギリか外れる。 | 機器状況確認 | BGモジュール引き上げ、代替品、取り付け | SWO-255303 | Defect | HW | pH/BG Module | pH/BG Module | 588 | QC of pCO 2, the total level lower limit or deviate. | Equipment status check | BG module pulling, replacement, installation
No | Visit | Complaint | RAME | Needs help setting network to Aqure express | Assistance with connecting an analyzer to the AQURE server has been requested. Connection through a term server has been unsuccessful. | Bypassed he term server and connected the analyzer directly to the hospital's network. Configured the analyzer's Ethernet card with the appropriate IP address and network settings provided by the customer. Verified network connection by pinging the AQURE server and the Radiance connection established status. | SWO-250974 | Setup | His/Lis | His/Lis | His/Lis |  | Needs help setting network to Aqure express | Assistance with connecting an analyzer to the AQURE server has been requested. Connection through a term server has been unsuccessful. | Bypassed he term server and connected the analyzer directly to the hospital's network. Configured the analyzer's Ethernet card with the appropriate IP address and network settings provided by the customer. Verified network connection by pinging the AQURE server and the Radiance connection established status.
No | Visit | Complaint | RLTD | Software reinstallation - R94N2 | Software reinstallation - R94N2 | Reinstalled software. Clean install. Loaded calibration parameter for CKMB. Passed. Loaded 3 levels QC for this parameter. Technopath to discuss with radiometer RMED on next course of action. | SWO-250791 | Defect | SW | RMED SW | RMED SW |  | Software reinstallation - R94N2 | Software reinstallation - R94N2 | Reinstalled software. Clean install. Loaded calibration parameter for CKMB. Passed. Loaded 3 levels QC for this parameter. Technopath to discuss with radiometer RMED on next course of action.
No | Visit | Complaint | RLTD | Software reinstallation - R110N3 | Software reinstallation - R110N3 | Reinstalled software. Clean install. Loaded calibration parameter for CKMB. Passed. Loaded 3 levels QC for this parameter. Technopath to discuss with radiometer RMED on next course of action. | SWO-250792 | Defect | SW | RMED SW | RMED SW |  | Software reinstallation - R110N3 | Software reinstallation - R110N3 | Reinstalled software. Clean install. Loaded calibration parameter for CKMB. Passed. Loaded 3 levels QC for this parameter. Technopath to discuss with radiometer RMED on next course of action.
No | Visit | Complaint | RSAF | Customer called to inform us that the analyzer is in Automatic recovery in process, Switches off and reboots, undefined look up string and process error | Customer complain that the SP only lasted 3 days. Inspected used SP and found the cups stuck in the waste well and cups behind the SP in the SP bay on the instrument | Remove cups in the SP bay on the instrument and replaced SP. | SWO-253134 | Defect | Consumables | Solution Pack | Cup Stack |  | Customer called to inform us that the analyzer is in Automatic recovery in process, Switches off and reboots, undefined look up string and process error | Customer complain that the SP only lasted 3 days. Inspected used SP and found the cups stuck in the waste well and cups behind the SP in the SP bay on the instrument | Remove cups in the SP bay on the instrument and replaced SP.
No | Visit | Complaint | RDEU | 20.10.2017 Sr. Jelly 02037332314 meldet, das der LP nicht mehr einrastet. | Einlassdichtung mit Halter defekt. | LP über Service Menü auswerfen lassen und wieder eingesetzt. LP konventionell auswerfen lassen und wieder eingesetzt. Einlassdichtung mit Halter durch Kundeneigenen ersetzt. Gerät initialisiert. Kal und QC sind OK. | SWO-251704 | Defect | HW | Inlet module | Inlet gasket |  | 20.10.2017 Sr. Jelly 02037332314 reports that the LP no longer snaps into place. | Inlet seal with holder defective. | Let LP eject via service menu and re-inserted. Let LP be ejected conventionally and re-inserted. Inlet seal with holder replaced by customer's own. device initialized. Kal and QC are OK.
No | Visit | Complaint | RAME | Lactate problems after replacing the sensor cassette | Lactate takes longer than normal to pass calibration. Internal QC is failing level 1-2 high on Lactate | Replaced with a Pcb R2004N0206. Brought system to ready $Qc came in without errors | SWO-251479 | Defect | HW | Basic Unit | PCB, Sensor interface |  | Lactate problems after replacing the sensor cassette | Lactate takes longer than normal to pass calibration. Internal QC is failing level 1-2 high on Lactate | Replaced with a Pcb R2004N0206. Brought system to ready $Qc came in without errors
Yes | Visit | Complaint | RBNL | Techline case wegens Health Hazard | Van: Hamer, Henrike Verzonden: dinsdag 10 oktober 2017 9:49 Aan: 'Bouwhuis, Wim WIMBO'; Treskes, Marco CC: Groot, Hans de; Gharib, Reza REZGH Onderwerp: vraag over bloedgas; verkeerde waarde HCO3 zonder foutcode Beste Wim, We hebben eerder een casus bij jullie gemeld waarin een aantal parameters waren gerapporteerd zonder “?”, terwijl die achteraf toch niet bleken te kloppen (zie ook onderstaande correspondentie). Gister hadden we weer een vergelijkbaar geval. In dit monster was alleen een HCO3 aangevraagd, en deze werd gerapporteerd zonder “?”. Echter, de oximetrie waarden en bilirubine gaven de foutcode 0593 (onvoldoende (bloed)monster) en de K en Na gaven de foutcode 1070 (sensor resonsfout). Bij het overmeten (na 3 minuten) werd er echter een HCO3 waarde van 32.4 gemeten, deze was sterk afwijkend van de eerdere waarde van 20.6. Zie hiervoor bijgevoegde prints. Kortom: hebben jullie er een verklaring voor waarom in dit geval de HCO3 niet met een “?” was gerapporteerd? Alvast bedankt voor je reactie, Met vriendelijke groet, Dr. H.M. (Henrike) Hamer Klinisch chemicus i.o. Klinische chemie M 06 24109219 T 020 510 8897 h.m.hamer@olvg.nl | Techline $Health Hazard case heb aangemaakt, datalogs en SD file heb gekopieerd en opgestuurd naar RMED. | SWO-252491 | Re classified RMED | Techline |  |  |  | Techline case due to Health Hazard | From: Hammer, Henrike Sent: Tuesday 10 October 2017 9:49 To: ' Bouwhuis, Wim WIMBO '; Treskes, Marco CC: Great, Hans de; Gharib, Reza REZGH Subject: Question about blood gas; Wrong value HCO3 without error code Dear Wim, We have previously reported a case with you in which a number of parameters had been reported without "?", while they were not found to be correct afterwards (see also correspondence below). We had a similar case again. In This sample, only a HCO3 was requested, and it was reported without "?". However, the oximetry values and bilirubin gave the error code 0593 (Insufficient (blood) sample) and the K and after gave the error code 1070 (sensor Resonserror). When Overmeasuring (after 3 minutes), however, a HCO3 value of 32.4 was measured, which was significantly different from the previous value of 20.6. See attached prints. In short: Do you have an explanation for why in this case the HCO3 was not reported with a "?"? Thanks in advance for your comment, Sincerely Dr. H.M. (Henrike) Hamer Clinical Chemist I.O. Clinical Chemistry M 06 24109219 T 020 510 8897 h.m.hamer@olvg.nl | Techline $Health Hazard case, copied data logs and SD file and sent it to RMED.
No | 5 | Y | RPAC | Blank | Faulty Solution Pack |  | D354460DAE4D6FFC | Assisted Replacement | SP | SP | SP |  | Blank | Faulty Solution Pack
No | Visit | Complaint | RCHN | PH,CO2定标不过 | PH,CO2定标不过，借903-334一个 | 更换903-334一个正常。by tim | SWO-250533 | [5 empty] | PH, CO2 calibration but | PH, CO2 calibration but, borrowing 903-334 a | Replace 903-334 one normal. by tim
No | Visit | Complaint | RCHN | 报错4794 | 报错4794，借备件636-549一个 | 更换636-549一个正常。by tim | SWO-250451 | Defect | HW | Aspiration Module | Cable, Aspiration Module | 4794 | 4794 error | 4794, borrowed spare parts 636-549 one | Replace 636-549 one normal. by tim
No | Visit | Complaint | RCHN | 没有气体 | 没有气体，借902-505,902-833各一个 |  | SWO-251020 | Defect | HW | Gas Cylinder | Flow regulator | 513 | No gas. | No gas, borrow 902-505,902-833 each
No | Visit | Complaint | RDEU | Oxi LS Fehler EL LS Fehler | Oxi LS Fehler EL LS Fehler | Fließwege überprüft Luftansaugstutzen gereinigt Demontageind Wartung Ventile BG und EL Modul LS Justage, Pumpenkal. | SWO-251329 | Cleaned | Fluidic pathway | El/Met Module | Air vent |  | Oxi LS Error EL LS Error | Oxi LS Error EL LS Error | Flow paths checked Air intake nozzle cleaned Demontageind Maintenance Valves BG and EL Module LS Adjustment, Pump Scale.
No | Visit | Complaint | RCHN | 报错0656 | 报错0656，借备件902-832,902-833各一个 | 更换902-832一个正常。by tim | SWO-250599 | Defect | HW | Oxi module | LS |  | Error 0656 | Error 0656, one spare part 902-832,902-833 each | Replace 902-832 one normal. by tim
No | Visit | Complaint | RSCH | Neudefekt | pCO2 Level 4 zu tief | Allgemeine Kontrolle i.o Membranenwartung pCO2 durchgeführt, ohne Erfolg Messkammer pH $BG-Modul ersetzt i.o Alle QC i.o System i.o | SWO-251420 | Defect | HW | pH/BG Module | Measuring chamber, pCO2 | pCO2 | Neudefekt | pCO2 Level 4 too deep | General control i.o Membrane maintenance pCO2 carried out, without success Measuring chamber pH $BG module replaces i.o All QC i.o System i.o
Yes | Visit | Complaint | RSCH | Neudefekt | Empfindlichkeit pH und Ca zu hoch | Allgemeine Kontrolle i.o O-Ring pH ersetzt i.o Membranenwartung Ca durchgeführt i.o Fehler weiterhin Neue pH und Ca Elektrode i.o Kal. und QC i.o System i.o | SWO-251428 | Re classified RMED | Double entry |  |  |  | Neudefekt | Sensitivity pH and Ca too high | General control i.o O-ring pH replaces i.o Membrane maintenance Ca carried out i.o Errors continue New pH and Ca electrode i.o Kal. and QC i.o System i.o
No | Visit | Complaint | RSCH | Neudefekt | pCO2 Level 4 zu tief | Allgemeine Kontrolle i.o Membranenwartung pCO2 durchgeführt, ohne Erfolg Messkammer pH $BG-Modul ersetzt i.o Alle QC i.o System i.o | SWO-251419 | Defect | HW | pH/BG Module | Measuring chamber, pCO2 | pCO2 | Neudefekt | pCO2 Level 4 too deep | General control i.o Membrane maintenance pCO2 carried out, without success Measuring chamber pH $BG module replaces i.o All QC i.o System i.o
No | Visit | Complaint | RSCH | Neudefekt | Sporadisch ctHb drift | Allgemeine Kontrolle i.o Lampeneinheit ersetzt i.o Lampe justiert i.o tHb Kal. durchgeführt i.o ctHb drift stabil Kal. und QC i.o System i.o | SWO-251344 | Defect | HW | Oxi module | Lamp unit |  | Neudefekt | Sporadisch ctHb drift | General control i.o Lamp unit replaces i.o Lamp adjusted i.o tHb Kal. carried out i.o ctHb drift stable Kal. and QC i.o System i.o
No | Visit | Complaint | RCHN | 无法开机 | 无法开机，借910-285，910-342各一个。 | 更换910-285，910-342各一个正常。 by Andrew | SWO-251639 | Defect | HW | Basic Unit | Power Supply |  | Cannot power on | Cannot be switched on, borrow 910-285, 910-342 each. | Replace 910-285, 910-342 each normal. by Andrew
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 不均匀的Q3溶液 | 连续更换2个试剂包后修复 | SWO-251732 | Revisit required | HW | Oxi module | Hemolyzer |  | Test card, reagent package failure | Uneven Q3 solution | Repair after 2 reagent kits in a row
No | Visit | Complaint | RCHN | 光学强度超范围 1031 | 光学强度超范围 1031 ，用备用光学单元903-364R8198N005 | 更换903-364后修复 by Andrew | SWO-251675 | Defect | HW | Process Module | Optical Unit | 1031 | Optical strength over range 1031 | Optical strength over range 1031 , with spare optical unit 903-364R8198N005 | Repair by Andrew after replacing 903-364
No | Visit | Complaint | RCHN | 开机提示错误 2158 | 开机提示错误 2158，借903-127一个。 | 更换903-127一个正常 by Andrew | SWO-251640 | Defect | HW | Storage Module | PCB, Storage Module | 2158 | Boot prompt error 2158 | Boot prompt error 2158, borrow 903-127 one. | Replace 903-127 a normal by Andrew
No | Visit | Complaint | RDEU | 20.10.2017 Herr Spannhoff 04921981713 meldet, dass die Installation von SP & SC gescheitert ist. | Diverse Flüssigkeitsfehler trotz Erneuerung von Lösungspack, Kassette, Einlaßdichtung, Durchspülung ohne Erfolg | Gerät überprüft, Schlauchsystem erneuert, Meßkassette und Hemolyser gespült, | SWO-252388 | Revisit required | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | 20.10.2017 Mr. Spannhoff 04921981713 reports that the installation of SP & SC has failed. | Various liquid errors despite renewal of solution pack, cassette, inlet seal, flushing without success | device checked, hose system renewed, measuring cassette and hemolyser flushed,
No | Visit | Complaint | RAME | Unable to connect to the LIS | No send button available. Analyzer not communicating. Customer was panning to use AQURE but now plans to use serial connections with DI. | Completed changes in Roma.ini file to allow for serial connection. Completed on all units. | SWO-252279 | Setup | His/Lis | His/Lis | His/Lis |  | Unable to connect to the LIS | No send button available. Analyzer not communicating. Customer was panning to use AQURE but now plans to use serial connections with DI. | Completed changes in Roma.ini file to allow for serial connection. Completed on all units.
No | Visit | Complaint | RLTD | Fault observed in logs during site visit. | Flushed analyser sample pathway for 30 minutes with fresh hypochlorite solution to disinfect. Swabbed all tube connectors with hypochlorite solution. Reassembled and fitted new Sensor Cassette. Verified operation on restart - all ok. | Disinfected analyser. | SWO-252036 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway | Disinfection | Fault observed in logs during site visit. | Flushed analyser sample pathway for 30 minutes with fresh hypochlorite solution to disinfect. Swabbed all tube connectors with hypochlorite solution. Reassembled and fitted new Sensor Cassette. Verified operation on restart - all ok. | Disinfected analyser.
No | Visit | Complaint | RDEU | Sporadisch Probenabbrüche | Sporadisch Probenabbrüche 1253 Probenaspiration gescheitert. | Einlassnadel überprüft - in Ordnung. Oxi Complete Test - ok. Präventiv Einlassnadel, Pinch Valve, Abfallschlauch, Ventilschlauch und Einlassdichtungshalter erneuert. Ventil im Service überprüft. Schaltfunktion in Ordnung. Kalibrationen und QC's vorgenommen. Testmessung in Ordnung. Kunde wird Gerät beobachten. | SWO-252007 | Defect | HW | PM Kit | PM Kit |  | Sporadisch Probenabbrüche | Sporadic sample cancellations 1253 sample aspiration failed. | Inlet needle checked - fine. Oxi Complete Review - ok. Preventive inlet needle, pinch valve, waste hose, valve hose and inlet seal holder renewed. valve in the service. Switching function in order. calibrations and QC's. Test measurement in order. Customer will watch device.
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | LAC,GLU定标不过 0210 0378 | 连续更换2张测试卡修复 | SWO-252934 | [5 empty] | Test card, reagent package failure | LAC, GLU calibration is no more than 0210 0378 | 2 consecutive test card repair
No | Visit | Complaint | RCHN | LAC定标失败 0375，0378 | LAC定标失败 0375，0378 ，借903-581一个。 | 更换903-581一个正常。 by Andrew | SWO-251637 | Defect | HW | DMS_PC | PCB, Sensor interface |  | LAC calibration failure 0375,0378 | LAC calibration failure 0375,0378 , borrow edith of 903-581. | Replace the 903-581 one normal. by Andrew
No | Visit | Complaint | RCHN | LAC灵敏度超范围 0378 | LAC灵敏度超范围 0378，借903- 581一个。 | 未使用备件，调整探针后修复 by Andrew | SWO-251682 | Defect | HW | DMS_PC | PCB, Sensor interface |  | LAC Sensitivity Over range 0378 | LAC sensitivity over range 0378, loan 903- 581 one. | Spare parts not used, adjust probe to repair by Andrew
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | LAC,GLU响应错误 1310 | 连续更换3张测试卡后修复 | SWO-255236 | [5 empty] | Test card, reagent package failure | LAC, GLU Response Error 1310 | Repair after 3 consecutive test cards
No | Visit | Complaint | RCHN | 需要升级扫描枪 | 需要升级扫描枪，需要借件905-560备件一个 |  | SWO-251941 | Setup | SW | SW | External Barcode Reader |  | Need to upgrade scan gun | Need to upgrade scanning gun, need to borrow 905-560 spare parts one
No | Visit | Complaint | RCHN | 干燥前温度错误 | 干燥前温度错误，借903-210，636-561各一个。 | 更换903-210R2000N0955 | SWO-251638 | Defect | HW | Process Module | Cup dryer | 1068 | Temperature error before drying | The temperature before drying is wrong, borrow 903-210, 636-561 each. | Replace 903-210R2000N0955
No | Visit | Complaint | RCHN | 频繁报样本量不足 1218 | 频繁报样本量不足 1218，借903-329一个。 | 更换903-329后修复 by Andrew | SWO-251684 | Revisit required | HW | Analyzer Control Module | Needle Wash Unit | 1218 | Frequently reported sample size less than 1218 | The frequently reported sample size was less than 1218, borrowing 903-329 one. | Fix by Andrew after replacing 903-329
No | Visit | Complaint | RCHN | 进样位置无法固定样本 | 进样位置无法固定样本，借903-584一个。 | 更换903-584一个正常。 by Andrew | SWO-251642 | Defect | HW | Cover and Chassis | Tube holder |  | Sample cannot be fixed at the sample position | The sample position cannot be fixed, borrowing 903-584 one. | Replace the 903-584 one normal. by Andrew
No | Visit | Complaint | RESP | Equipamento não detecta abertura da entrada para fazer a gasimetria | Verificação dos registros | Remoção da cassete de electrodos e inspecção e limpeza do opto-sensor da entrada, verificação do correcto funcionamento com a realização de duas gasimetrias | SWO-255176 | [5 empty] | Equipment does not detect entry opening to make Gasimetry | Checking Records | Removal of the electro-cassette and inspection and cleaning of the opto-sensor of the inlet, verification of the correct operation with the realization of two gasimetries
No | Visit | Complaint | RSCH | Rep | Rep Elektolyte trotz mehrmals SC und LP tausch mit Fehler (Ableitströme etc) | Neue Sensorkassette undicht und ausgelaufen. Komplettes Sensor Interface PCB ausgetauscht Alles justiert Softwarupgrade auf V3.3MR1 durchgeführt und angepasste Roma Datei installiert (Für Crea vorbereitet) SC initialisiert. Na nicht kalibrierbar (Status+Empf) Neuen Sensor installiert tHb Kal ausgeführt (Überfällig seit 06.2017) Funktionskontrolle i.o | SWO-252532 | Assisted Replacement | SC | SC | Leaking |  | Rep | Rep Elektolyte despite several SC and LP exchanges with errors (discharge currents etc) | New sensor cartridge leaks and leaks. Complete Sensor Interface PCB replaced Everything adjusted Softwarupgrade performed on V3.3MR1 and customized Roma file installed (prepared for Crea) SC initialized. Na not calibrated (Status+Empf) New sensor installed tHb Kal executed (overdue since 06.2017) Functional control i.o
No | 2 | Y | RPAC | Blank | Inlet probe bent during use for demonstration |  | D355DE862438ACDF | Replaced | HW | Auto Inlet module | Probe bend |  | Blank | Inlet probe bent during use for demonstration
No | Visit | Complaint | RCHN | 样本抽不上去 | 样本抽不上去，借902-971一个。 | 更换902-971后修复 by Andrew | SWO-251683 | [5 empty] | The sample won't go up. | Samples can not be drawn up, borrow 902-971 one. | Repair by Andrew after replacement 902-971
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 液体传输失败 1245 | 更换试剂包后修复 | SWO-253694 | Assisted Replacement | SP | SP | FT | 1245 | Test card, reagent package failure | Liquid transfer failure 1245 | Repair after replacement of reagent kit
No | Visit | Complaint | RCHN | 报错0290,0291,0586 | 报错0290,0291,0586，借备件902-980,903-197各一个 |  | SWO-252369 | Defect | HW | Oxi module | Hemolyzer |  | Error 0290,0291,0586 | Error 0290,0291,0586, spare parts 902-980,903-197 each
No | 6 | Y | RPAC | Blank | Keeps calibrating after membrane change |  | D356065F2410A2C7 | [5 empty] | Blank | Keeps calibrating after membrane change
No | Visit | Complaint | RCHN | 报5784错误 | 报5784错误，借636-550一个 |  | SWO-252371 | Defect | HW | Process Module | Cable, Process Module | 5784 | Reported 5784 errors | Reported 5784 error, borrow 636-550 one
No | Visit | Complaint | RKK | 検体測定にてサンプルエラー頻発 | 検体測定にてサンプルエラー頻発 | インレットプローブ曲がりが原因。インレットプローブ交換。エアーライン洗浄。Ref,PO2,PCO2メンブラン交換。 | SWO-255028 | Defect | HW | Inlet module | Probe bend |  | Frequent sample errors in sample measurement | Frequent sample errors in sample measurement | Inlet probe due to bending. Inlet probe replacement. Air line cleaning. Ref, PO2, PCO2 membrane replacement.
No | Visit | Complaint | RIND | Thb/so2 ? Mark For all samples | Thb $so2 giving error for all samples | checked and found that clott sample was given so cleaned the gasket and probles done pump calibartion and liquid sensor calibration after thta tested sample instrument working fine | SWO-253941 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | Thb/so2 ? Mark For all samples | Thb /so2 giving error for all samples | checked and found that clott sample was given so cleaned the gasket and probles done pump calibartion and liquid sensor calibration after thta tested sample instrument working fine
No | Visit | Complaint | RAME | pH failing on external QC level 3, after unit rebooting from software upgrade process. | Examined data logs and sensor cassette. | Replaced sensor cassette, with customer stock. Replaced solution pack for troubleshooting purposes. | SWO-253630 | Assisted Replacement | SC | SC | pH |  | pH failing on external QC level 3, after unit rebooting from software upgrade process. | Examined data logs and sensor cassette. | Replaced sensor cassette, with customer stock. Replaced solution pack for troubleshooting purposes.
No | Visit | Complaint | RAME | MetHb and COHb blood samples that have low values are being reported as low out of range. | During blood comparisons, user noted the most MerHb and many COHb results have three down arrows and no results printed. | Found Reportable Range set to 1 - 91 for MetHb and 1 - 92 for COHb. Changed low range setting to zero and the analyzer now reports results that are less than 1%. | SWO-253549 | Setup | Reportable range | Reportable range | Oxi module |  | MetHb and COHb blood samples that have low values are being reported as low out of range. | During blood comparisons, user noted the most MerHb and many COHb results have three down arrows and no results printed. | Found Reportable Range set to 1 - 91 for MetHb and 1 - 92 for COHb. Changed low range setting to zero and the analyzer now reports results that are less than 1%.
No | Visit | Complaint | RLTD | Please call Sue O'Brien Wheeler on 07954 576 051 | Checked mechanical operation of door - able to open and close manually. Dismantled analyser and checked door sensor - ok. Checked all cable connections to the wet section control pcb, cable for analyser ID pcb not clamped in place, fixed clamp to hold cable, (possibly unrelated to door). Re-assembled analyser and tested door in service program - now opening and closing correctly. Checked in consumable replacement that door opens and closes - all ok. Fault rectified. | Secured loose cable at wet section control pcb. Operated mechanism manually to ensure it was in a known position. | SWO-253376 | Adjustment | HW | Basic Unit | Cable, Wet sec PCB |  | Please call Sue O'Brien Wheeler on 07954 576 051 | Checked mechanical operation of door - able to open and close manually. Dismantled analyser and checked door sensor - ok. Checked all cable connections to the wet section control pcb, cable for analyser ID pcb not clamped in place, fixed clamp to hold cable, (possibly unrelated to door). Re-assembled analyser and tested door in service program - now opening and closing correctly. Checked in consumable replacement that door opens and closes - all ok. Fault rectified. | Secured loose cable at wet section control pcb. Operated mechanism manually to ensure it was in a known position.
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | PH,CO2，NA,CA,CL,K,HCT定标不过 | 更换测试卡后修复 | SWO-254269 | [5 empty] | Test card, reagent package failure | PH,CO2，NA,CA,CL,K,HCT定标不过 | Repair after replacing test card
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | LAC定标不过 | 更换测试卡后修复 | SWO-253688 | [5 empty] | Test card, reagent package failure | LAC calibration is not | Repair after replacing test card
No | Visit | Complaint | RCHN | 报错1024 定标失败 | 报错1024 定标失败，需要借件944-214（13265）和944-394（批次号随机发） | 重新定标后修复 by Andrew | SWO-252500 | Adjustment | Consumables | Test Kit | Cartridge, Myo | 1024 | Error 1024 bid failure | Misquoting 1024 bid failure, need to borrow 944-214 (13265) and 944-394 (batch number random send) | Fix by Andrew after re-calibration
No | Visit | Complaint | RCHN | 报错5784 | 报错5784 | 更换备用机件636-550后修复 by Andrew | SWO-253637 | Defect | HW | Process Module | Sensor, Process Unit | 5784 | 5784 | 5784 | Repair by Andrew after replacing spare parts 636-550
No | Visit | Complaint | RCHN | 报错4726,4584 | 报错4726,4584，借备件903-198,903-028各一个 | 更换903-198,903-028后修复 by Andrew | SWO-252344 | Defect | HW | Analyzer Control Module | PCB, Main Board |  | 47264584 | Error 4726,4584, spare parts 903-198,903-028 each | Fix by Andrew after replacing 903-198,903-028
No | Visit | Complaint | RIND | Service Visit | Analyzer is not printing out when give the print command. | After observing, the latch of the printer had broken and need to change the printer head. Change the printer head. Now analyzer is in well working condition and handed over to staff. | SWO-254929 | Defect | HW | Printer Module | Printer Module |  | Service Visit | Analyzer is not printing out when give the print command. | After observing, the latch of the printer had broken and need to change the printer head. Change the printer head. Now analyzer is in well working condition and handed over to staff.
No | Visit | Complaint | RDAN | Dorte melder, at ABL stadigvæk har ca. 15 til 20 gange neon lamp too... pr. måned. | Da EBJ har skiftet neon lamp to gange og oxi pcb en gang bliver det et nyt modul da der er garanti på ABL'en. | Nyt oxi modul monteret, opjusteret, service tHb kalibreret og slutafprøvet. | SWO-253328 | Defect | HW | Oxi module | Oxi module, Replaced | Neon Lamp | Dorte reports that ABL still has about 15 to 20 times the neon lamp too... Per month. | Since the EBJ has switched the neon lamp twice and oxi PCB once it becomes a new module as there is a warranty on the ABL. | New Oxi module mounted, upwards adjusted, service tHb calibrated and final tested.
No | Visit | Complaint | RDEU | Oxi LS Fehler EL LS Fehler | Leckage Fehler Hotline-Meldung | Fehlersuche Pumpenkalibrierung, LS Justage pH Elekt. und Meßkammer gereinigt da Messung instabil Einlaßdichtung erneuert | SWO-253291 | Defect | HW | Inlet module | Inlet gasket |  | Oxi LS Error EL LS Error | Leakage error Hotline message | Troubleshooting Pump calibration, LS adjustment pH Elekt. and measuring chamber cleaned as measurement unstable Inlet seal renewed
No | Visit | Complaint | RAME | Unit failing el met pump cal | Unit had a raise in el met pump motor. | Cleaned and Re seated pump motor ran calibrations. | SWO-253450 | Adjustment | HW | El/Met Module | Pump Rotor | Revisit ?? | Unit failing el met pump cal | Unit had a raise in el met pump motor. | Cleaned and Re seated pump motor ran calibrations.
No | Visit | Complaint | RDEU | 24.10.2017 15:00 Andreas Rönne 04152803410: Am 23.10.2017 um 6:51Uhr wurde ein Patient gemessen. Das TnI Ergebnis war 600ng/L. Die Blutprobe, die ins Labor geschickt wurde, zeigte den Wert 13ng/L. | Überprüfung des Analysator wegen unstimmiger Werte. | Zur GS gefahren und vorsichtshalber das Leihgerät abgeholt. Analysator überprüft. Festgestellt wurde eine Massive Verschmutzung der Cup Wash Unit. Nach Reinigung der Unit eine Reinigung durchgeführt. Dann die QC für alle Parameter durchgeführt. Danach Blankmessungen zur Überprüfung der Optischen Einheit. Danach drei Propemessungen mit dem Lapor verglichen. Alles ok. | SWO-253271 | Cleaned | HW | Process Module | Cup wash unit |  | 24.10.2017 15:00 Andreas Rönne 04152803410: On 23.10.2017 at 6:51o a patient was measured. The TnI result was 600ng/L. The blood sample sent to the laboratory showed a value of 13ng/L. | Check the analyzer for inconsistent values. | Driven to the GS and as a precaution the rental device was picked up. Analyzer checked. A massive pollution of the Cup Wash Unit was detected. After cleaning the unit, a cleaning is carried out. Then the QC is performed for all parameters. Then blank measurements for checking the optical unit. Then compared three prope measurements with the lapor. All ok.
No | Visit | Complaint | RDEU | Leckage erkannt (Benutzerintervention erforderlich); Kunde Hr. Söllner hat bereits alles durchgeführt, was der Analysator aufgefordert hat (manuelle Spülung, Einlassdichtung, LP, SC; Einlassnadel ist gerade) --> Herr Robert Söllner (Medizintechnik), bittet um einen Service-Einsatz | Fehlersuche --> Dateiarchive (Aktivitäten, QC, Kal., Betriebsmittel), Schlauchwege und Analysator im Service überprüft --> Leckage erkannt (Drucküberprüfung Vakuumfehler) --> sämtliche Schlauchwege erneuern --> siehe SWO- der Jahreswartung | Fehler konnte durch die Jahreswartung behoben werden!!! | SWO-253773 | Defect | HW | Basic Unit | PM Kit |  | Leakage detected (user intervention required); Customer Mr. Söllner has already carried out everything that the analyzer has requested (manual flushing, inlet seal, LP, SC; Inlet needle is straight) --> Mr. Robert Söllner (medical technology), asks for a service | Troubleshooting --> File archives (activities, QC, cal., equipment), hoseways and analyzer checked in the service --> Leakage detected (pressure check vacuum error) --> renew all hoseways --> see SWO- of annual maintenance | The annual maintenance could fix errors!!!
No | 1 | Y | RCZE | Blank | 1391 - vstup neni ve spravne poloze, ani po restarte - ihla ostava v polohe cca pre striekacku |  | D35841D31045CA3C | Adjustment | HW | Auto Inlet module | Inlet Position Cal. |  | Blank | 1391-The input does not carry the semi-finished, even after the Restarte-the needle Ostava in the position approx. for the spray
No | Visit | Complaint | RCHN | 定标时机器为空 | 定标时机器为空.需要借件962-187两个 | 更换962-187后修复 by Andrew | SWO-253045 | [5 empty] | The machine is empty at the time of calibration | The machine is empty at the time of calibration. | Repair by Andrew after replacement 962-187
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | pco2定标失败，0210 | 更换测试卡后修复 | SWO-254524 | Assisted Replacement | SC | SC | pCO2 | 210 | Test card, reagent package failure | pco2 calibration failure, 0210 | Repair after replacing test card
No | Visit | Complaint | RCHN | lac定标通过时间太长 | lac定标通过时间太长,需要借件903-581、903-158、903-135各一个 | 更换903-581、903-158备件后修复 by Andrew | SWO-253110 | Defect | HW | Basic Unit | PCB, Sensor interface |  | lac calibration pass time is too long | Lac calibration time is too long, need to borrow 903-581, 903-158, 903-135 each one | Repair by Andrew after replacing 903-581, 903-158 spare parts
No | Visit | Complaint | RAME | FCOHb out intermittently on level 2 QC. | Found all levels of QC intermittently failing and passing on rerun. Opened lid to measuring chamber and found small leak coming from PH chamber. | Replaced o-ring for PH electrode. | SWO-254152 | Defect | Electrode | Electrode | Sealing ring | Sealing ring, pH | FCOHb out intermittently on level 2 QC. | Found all levels of QC intermittently failing and passing on rerun. Opened lid to measuring chamber and found small leak coming from PH chamber. | Replaced o-ring for PH electrode.
No | Visit | Complaint | RIND | Software Error. | Software Error. | found problem with Software, tried the upgrading windows software but it was showing error, replaced the HDD upgraded the analyzer software done the setup found instrument is working satisfactorily. Note: Need standalone 1KVA UPS Urgently. | SWO-253824 | Defect | HW | DMS_PC | HDD | UPS | Software Error. | Software Error. | found problem with Software, tried the upgrading windows software but it was showing error, replaced the HDD upgraded the analyzer software done the setup found instrument is working satisfactorily. Note: Need standalone 1KVA UPS Urgently.
No | Visit | Complaint | RSCH | Reparatur | pH instabil bei installation | O-Ring pH 2x gewechselt i.o Kal. und QC i.o System i.o | SWO-254629 | Defect | Electrode | Electrode | Sealing ring | Sealing ring, pH | Repair | pH unstable at installation | O-Ring pH 2x gewechselt i.o Kal. und QC i.o System i.o
No | Visit | Complaint | RDEU | 26.10.2017, 09:13 Uhr, Frau Kaiser-Mahr, Tel. 0952128326: Halterung für Probenröhrchen abgebrochen | Fehlersuche --> Einlass überprüft --> Halterung für Probenröhrchen defekt | Probenhalter erneuert, Testmessung durchgeführt, System und Funktion i.O. | SWO-253920 | Defect | HW | Cover and Chassis | Tube holder |  | 26.10.2017, 09:13, Mrs. Kaiser-Mahr, tel. 0952128326: Bracket for sample tubes aborted | Troubleshooting --> Inlet checked --> Bracket for sample tubes defective | Sample holder renewed, test measurement performed, system and function in the order
Yes | Visit | Complaint | RDEU | Leckage erkannt (Benutzerintervention erforderlich); Kunde Hr. Söllner hat bereits alles durchgeführt, was der Analysator aufgefordert hat (manuelle Spülung, Einlassdichtung, LP, SC; Einlassnadel ist gerade) --> Herr Robert Söllner (Medizintechnik), bittet um einen Service-Einsatz | Siehe SWO-253773 (eigentliche Maßnahmen, sind in SWO-253773 vorgenommen worden) --> Zeitenänderung!!! --> die angegebene Arbeitszeit, Ankunft- und Abfahrtszeit passen nicht wie im SWO-253773 angegeben --> diese Zeiten werden mit diesem hier abgeändert | Zeitenänderung: 1-0100 = 1,3h --> (vorher 0,2h) Ankunft: 09:40 Uhr, 26.10.17 Abfahrt: 11:00 Uhr, 26.10.17 --> (vorher 10:00Uhr) | SWO-254474 | Re classified RMED | To document action in other SWO |  |  |  | Leakage detected (user intervention required); Customer Mr. Söllner has already carried out everything that the analyzer has requested (manual flushing, inlet seal, LP, SC; Inlet needle is straight) --> Mr. Robert Söllner (medical technology), asks for a service | See SWO-253773 (actual measures taken in SWO-253773) --> Time change!!! --> the specified working time, arrival and departure time do not match as indicated in SWO-253773 --> these times are changed with this here | Time change: 1-0100 = 1.3h --> (previously 0.2h) Arrival: 09:40, 26.10.17 Departure: 11:00 a.m., 26.10.17 --> (was made at 10:00 a.m.)
No | Visit | Complaint | RBNL | De laatste dagen heeft deze analyzer problemen met opzuigen bloedmonsters. Alleen bij spuitmetingen, capillairen, QC en calibraties gaan wel goed. | Veel metingen: Afgebroken. Al een nieuwe inletgasket en naald geplaatst door gebruiker. Ook al doorgespoeld. Safetipcap spuiten met dop gaan soms fout, zonder dop niet. Capillaren goed. | Complete inlethouder gewisseld met instrument dat ernaast staat (lab2). Zien of fout mee verplaatst. | SWO-255055 | Unknown | HW | Auto Inlet module | Auto Inlet module | VTC/User error? | The last few days this analyzer has problems with sucking up blood samples. Only for spraying measurements, capillaries, QC and calibrations will be good. | Many measurements: aborted. Already a new inletgasket and needle placed by user. Even though it has been washed. Safetipcap spraying with cap sometimes go wrong, without cap. Capillaren well. | Complete Holder exchanged with instrument that stands next to it (LAB2). See if error moved with it.
No | Visit | Complaint | RBNL | CDS waarden zijn te hoog | Regelmatig error Stolsel in SC terwijl geen stolsel in de SC aanwezig. Bij nader onderzoek blijkt dat de Clot Detection waarden te hoog of aan de hoge kant zijn. De waarden zijn als volgt: - Inlet value: 1,17 (0,80 - 1,22) - Sensor cassette value: 1,22 (0,80 - 1,20) - OXI module value: 0,96 (0,80 - 1,20) De SC Bay Unit heb vervangen vervolgens heb gecontroleerd, nu zijn de waarden als volgt: - Inlet value: 0,90 (0,80 - 1,22) - Sensor cassette value: 1,02 (0,80 - 1,20) - OXI module value: 0,96 (0,80 - 1,20) ABL90 heb kalibratie $QC's laten doen, OK. | Defecte SC Bay Unit heb vervangen. | SWO-254888 | Defect | HW | Basic Unit | Sensor Bay Unit |  | CDS values are too high | Regular error clot in SC While no clot is present in the SC. Further investigation reveals that the Clot Detection values are too high or high-end. The values are as follows: -Inlet Value: 1.17 (0.80-1.22) -Sensor Cassette value: 1.22 (0.80-1.20) -OXI Module Value: 0.96 (0.80-1.20) The SC Bay Unit have replaced then checked, now the values are as follows: -Inlet Value: 0.90 (0.80-1.22) -Sensor Cassette value: 1.02 (0.80-1.20) -OXI Module Value: 0.96 (0.80-1.20) ABL90 have calibration /QC's Let it do, OK. | Defective SC Bay Unit.
No | Visit | Complaint | RLTD | Lynn Shaw on 01582 497991 | Clot detected errors and cracked mixing chamber cover. | Sensor cassette value stating high. Disassembled analyser and individually flushed component parts. No issues found, checked all fluid paths. Will ask customer to monitor. Replaced cracked sample mixer tray. | SWO-254812 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | Lynn Shaw on 01582 497991 | Clot detected errors and cracked mixing chamber cover. | Sensor cassette value stating high. Disassembled analyser and individually flushed component parts. No issues found, checked all fluid paths. Will ask customer to monitor. Replaced cracked sample mixer tray.
No | Visit | Complaint | RAME | pO2 parameter flagged yellow, observed by customer | pCO2 failure observed by customer | New Inlet connector gasket installed. Analyzer flushed by DM water. All service tests performed and passed. | SWO-254934 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | pO2 parameter flagged yellow, observed by customer | pCO2 failure observed by customer | New Inlet connector gasket installed. Analyzer flushed by DM water. All service tests performed and passed.
No | Visit | Complaint | RCHN | PCO2定标失败 | PCO2定标失败，借902-972和903-334各一个 |  | SWO-254425 | [5 empty] | PCO2 calibration failure | PCO2 bid failure, borrow 902-972 and 903-334 each one
No | Visit | Complaint | RCHN | 间歇性定标失败，需要更换903-581 | 间歇性定标失败，需要更换903-581，借903-581一个。 | 更换903-581后修复 by jesson | SWO-253730 | Defect | HW | Basic Unit | PCB, Sensor interface | 378 | Intermittent calibration failed and 903-581 needs to be replaced | Intermittent calibration fails and needs to be replaced with 903-581, borrowing 903-581. | Repair by jesson after replacement 903-581
No | Visit | Complaint | RCHN | 报错5684 | 报错5684，需要借件903-236和636-533两个 |  | SWO-254422 | Revisit required | HW | Process Module | Sensor, Process wheel | 5684 | 5684 error | 5684, need to borrow 903-236 and 636-533 two
No | Visit | Complaint | RCHN | 定标卡无法扫描 | 定标卡无法扫描,借备件902-924一个 | 更换备用902-924后修复 by jesson | SWO-253698 | Defect | HW | Analyzer Control Module | Barcode reader, display | BR002 | Bid card cannot be scanned | The bid card cannot be scanned, one spare part 902-924 | Repair by jesson after replacing spare 902-924
No | Visit | Complaint | RCHN | 触摸屏故障 | 触摸屏故障，借636-571,903-036,903-584 $2个 |  | SWO-254284 | Defect | HW | Analyzer Control Module | Touch screen |  | Touch screen failure | Touchscreen fault, borrow 636-571,903-036,903-584 /2
No | Visit | Complaint | RLTD | Please call on Hannah McPeak on 01865 221 596 01865 741 166 ask switch bleep 1624 They have tried to clean it but can't get to all of the blood and the tube keeps popping off and it is causing further issues. | Checked the waste line where it connects to the waste pump tubes and the connection was still secure. | Removed the fluidic manifold and cleaned up the remaining blood from underneath it. | SWO-254485 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | Please call on Hannah McPeak on 01865 221 596 01865 741 166 ask switch bleep 1624 They have tried to clean it but can't get to all of the blood and the tube keeps popping off and it is causing further issues. | Checked the waste line where it connects to the waste pump tubes and the connection was still secure. | Removed the fluidic manifold and cleaned up the remaining blood from underneath it.
No | Visit | Complaint | RDAN | Revne i målekammer | NA | Målekammer skiftet | SWO-254614 | Defect | HW | pH/BG Module | Measuring chamber, pH | Cracked | Crack in the measuring chamber | In | Measuring chamber Change
No | Visit | Complaint | RAME | Upon arrival onsite, customer reported inhomogeneous errors during QC runs. | Checked all tubing for leakage. Checked data logs for errors. Intermittently failing with inhomogeneous errors when any level QC is run. | Replaced lower inlet and module tubings with extra kit supplies. Start calibrations and QC testing. Additional failure on second QC run. Replaced Inlet tube restart calibrations. | SWO-254922 | Defect | HW | Inlet module | Tubings |  | Upon arrival onsite, customer reported inhomogeneous errors during QC runs. | Checked all tubing for leakage. Checked data logs for errors. Intermittently failing with inhomogeneous errors when any level QC is run. | Replaced lower inlet and module tubings with extra kit supplies. Start calibrations and QC testing. Additional failure on second QC run. Replaced Inlet tube restart calibrations.
No | Visit | Complaint | RAME | Customer states calibration errors and transport errors. | failed calibrations, fluid transport and pump errors. | Replaced the flat membrane in the PH $BG dual valve housing. Replaced the pump tubing on both modules. LS adjust no pump calibration. Replaced he PH seal ring. Calibrations and QC passed. | SWO-254911 | Defect | HW | pH/BG Module | Valve, air | Membrane | Customer states calibration errors and transport errors. | failed calibrations, fluid transport and pump errors. | Replaced the flat membrane in the PH $BG dual valve housing. Replaced the pump tubing on both modules. LS adjust no pump calibration. Replaced he PH seal ring. Calibrations and QC passed.
No | Visit | Complaint | RSCH | Neudefekt | Mischerträger schlupf erkannt | Mischer träger ersetzt i.o | SWO-254628 | Defect | HW | FlexQ Module | FlexQ - Mixer |  | Neudefekt | Mixer carrier slip detected | Mixer carrier replaces i.o
No | Visit | Complaint | RESP | No transmite datos. | Se ha borrado la dirección IP. | Colocar la IP en la configuración de Windows, guardarla en memoria permanente y comprobar. | SWO-254503 | Setup | His/Lis | His/Lis | His/Lis |  | It does not transmit data. | The IP address has been deleted. | Place the IP in Windows settings, save it to permanent memory, and check.
Yes | 4 | Y | RFRA | Blank | PC ne démare pas |  | D35990EC860D5F26 | [5 empty] | Blank | PC doesn't start
No | Visit | Complaint | RBNL | Na plaatsen van een buisje worden lang niet alle barcodes gelezen. | Met het halfjaarlijks onderhoud hier ook al aandacht aan besteed. | Nieuwe Barcode scanner gemonteerd. Afgeregeld. Diverse malen ID correct ingelezen. | SWO-255051 | Defect | HW | Aspiration Module | Barcode reader, Inlet |  | After inserting a tube, not all barcodes are read for a long time. | With the half-yearly maintenance here also already paid attention to. | New Barcode scanner mounted. Adjusted. Several times ID correctly read.
Yes | 20 | Y | RPNZ | Blank | Solution pack replacement due to 964: leak current in relation to solution pack detected. |  | D35D0512E82C2736 | [5 empty] | Blank | Solution pack replacement due to 964: leak current in relation to solution pack detected.
Yes | Visit | Complaint | RCHN | 光学单元超出范围，报错1031 | 光学单元超出范围，报错1031，需要借件903-364和944-394（13419） | 工程师反馈光单有问题，报错3268。需重新再借一个903-364 by Andrew | SWO-253127 | Defect | HW | Process Module | Optical Unit | 1031 | Optical unit out of range, error 1031 | Optical unit out of range, error 1031, need to borrow 903-364 and 944-394 (13419) | The engineer reported a problem with the light sheet, which misstated 3268. Need to borrow another 903-364 by Andrew
No | Visit | Complaint | RKK | 溶液をポンプできません | 溶液をポンプできません | 回路洗浄実施 Ｃａｌ ＯＣ 実施 | SWO-255949 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway | 1343 | Unable to pump solution | Unable to pump solution | Circuit cleaning Cal OC implemented
No | Visit | Complaint | RCHN | 报5784，定标失败 | 报5784，定标失败，借944-394（13418）一张（缺货 ） | 重新定标后修复 by Andrew | SWO-255032 | Assisted Replacement | Consumables | Test Kit | Cartridge, PCT | 5784 | Reported 5784, bid failed | Reported 5784, bid failure, borrow 944-394 (13418) one (out of stock) | Fix by Andrew after re-calibration
Yes | Visit | Complaint | RCHN | 光学单元超出范围，报错1031 | 光学单元超出范围，报错1031，借的光单有质量问题，报错3268，重新再借一个903-364 | 更换903-364，重新定标后修复 by Andrew | SWO-255034 | Defect | HW | Process Module | Optical Unit | 1031 | Optical unit out of range, error 1031 | Optical unit out of range, reported error 1031, borrowed light sheet has quality problems, reported error 3268, re-borrowed a 903-364 | Replace 903-364, fix by Andrew after recalibration
No | Visit | Complaint | RAME | Customer called with unable to pump solution errors | Blockage reported from previous troubleshooting. Start diagnostic flow testing all passed. | Flush analyzer to check flow. Flow checks ok. Restart initialization of newly replaced consumables. | SWO-255824 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway | 1169 | Customer called with unable to pump solution errors | Blockage reported from previous troubleshooting. Start diagnostic flow testing all passed. | Flush analyzer to check flow. Flow checks ok. Restart initialization of newly replaced consumables.
No | Visit | Complaint | RDEU | 30.10.2017 11:30Uhr Frau Szymczak 091285043200 (091285048499): Das Gerät nimmt den LP nicht an. | ABL bootet nicht | ABL stand ausgeschaltet und mit LP90 entnommen da, Netzschalter stand auf Ein, drücken des Reset-Knopfes startet den ABL, LP90 soll installiert werden, nach Thermostatisierung beginnt die Startphase - 1 Std., Startphase abgewartet, QC Ok, Taub Kal. durchgeführt, System- und Funktionskontrolle in Ordnung | SWO-256008 | User error | SW | Basic Unit | off/on |  | 10/30/2017 11:30 AM Ms. Szymczak 091285043200 (091285048499): The device does not accept the LP. | ABL does not boot | ABL was switched off and with LP90 removed there, power switch stood on, pressing the reset button starts the ABL, LP90 is to be installed, after thermostatization the start phase begins - 1 hr. , start phase waited, QC Ok, deaf Kal. performed, system and function control in order
No | Visit | Complaint | RBNL | Druk test fout | ABL90 FLEX PLUS geeft error code 1062: Druk test fout aan. Druktest is te hoog. Bij nader onderzoek blijkt dat de pinchvalve vast zit. | De pinchvalve heb vervangen vervolgens heb gecontroleerd,OK. | SWO-255719 | Defect | HW | Oxi module | Pinch valve | 1062 | Druk test fout | ABL90 FLEX PLUS gives error code 1062: Press test error. Pressure test is too high. On closer examination it appears that the pinchvalve is stuck. | The pinchvalve have replaced then checked, OK.
No | 4 | Y | RPAC | Blank | Glu calibration error |  | D35D463EE80566A0 | Defect | HW | Basic Unit | PCB, Sensor interface |  | Blank | Glu calibration error
Yes | 2 | Y | RPAC | Blank | New install training |  | D369E303840A781B | Re classified RMED | Training/Installation |  |  |  | Blank | New install training
No | Visit | Complaint | RLTD | Ranvir on 01865 221 596 | Leakage from analyser and waste tube popping off y piece. | Replaced blocked waste tube, flushed waste bottle nipple and waste bottle. Remove and cleaned waste pump rollers, tubing and covers. Cleaned a large amount of blood from around waste pump. Ran pump and LS calibrations and leakage check, also ran a 1 point calibration, all with no signs of leakage. | SWO-255731 | Cleaned | Fluidic pathway | Fluidic pathway | Clot, Waste Line |  | Ranvir on 01865 221 596 | Leakage from analyser and waste tube popping off y piece. | Replaced blocked waste tube, flushed waste bottle nipple and waste bottle. Remove and cleaned waste pump rollers, tubing and covers. Cleaned a large amount of blood from around waste pump. Ran pump and LS calibrations and leakage check, also ran a 1 point calibration, all with no signs of leakage.
No | Visit | Complaint | RAME | Oxi liquid sensor error | Oxi Liquid Sensor error in system messages | Flush air line through V7. Ran LS adjust with pass results. | SWO-255906 | Cleaned | Fluidic pathway | Fluidic pathway | Clot, Manifold |  | Oxygen Liquid Sensor Error | Oxygen Liquid Sensor Error in System messages | Flush air line through V7. Ran LS adjust with pass results.
No | Visit | Complaint | RDEU | Monitor Display ohne Funktion | Monitor Display ohne Funktion | Funktionstest KeinMängel feststellbar, Display fkt. ohne Beanstandung | SWO-255594 | Revisit required | HW | DMS_PC | Touch screen | Firmware not possible | Monitor display without function | Monitor display without function | Functional No defects detectable, display fkt. without complaint
No | Visit | Complaint | RDEU | Sehr häufig Na, Ca und Cl Kalibrationsdriften | Überprüfung | Referenzelektrode optisch ohne Beanstandung. Pumpenschläuche der Module stark verschlissen, erneuert. Elektrodendeckel der Referenzelektrode erneuert. LS Justage (up ELMet 17) und Pumpen Kal durchgeführt. Kalibrationen mehrfach durchgeführt. QC ohne Beanstandung Kalibrationen werden beobachtet und bei weiteren Fehlern gemeldet. Drifttolleranzen wurden am 01.11.17 via Remote angepasst (V6.17) | SWO-255705 | Defect | HW | pH/BG Module | Electrode connector |  | Very common Na, Ca and Cl calibration drifts | Review | Reference electrode optically without complaint. Pump hoses of the modules are heavily worn, renewed. electrode cover of the reference electrode. LS Adjustment (up ELMet 17) and pumps Kal performed. calibrations performed several times. QC without complaint Calibrations are observed and reported in case of further errors. Drift toll shards were adjusted on 01.11.17 via remote (V6.17)
No | Visit | Complaint | RSCH | Rep OXI | 0251 und 0579, 0238 Diverse Oxifehler Lampenwert 4095 Na bei QC ausserhalb Bereich | OXI Halogenlampe neu justiert auf 3400 Mehrere Kal stabil QC Level 1 und 2 wiederholt. Kal 1 Lösung erneuert trotz 11% Füllstand leer. Na wieder i.o Funktionskontrolle i.o | SWO-255575 | Adjustment | HW | Oxi module | Lamp unit | Halogen | Rep OXYGEN | 0251 and 0579, 0238 Various Oxi errors Lamp value 4095 Na at QC outside area | OXI halogen lamp re-adjusted to 3400 Multiple Kal stable QC Level 1 and 2. Kal 1 solution renewed despite 11% level empty. Well again i.o Functional control i.o
No | Visit | Complaint | RSCH | Rep Na QC | Na bei CQ ausser Bereich | Kal 1 Lösung ausgetauscht. 12% aber trotzdem leer. CREA membraniert Kalibrierung wiederholt i.o QC wiederholt i.o Funktionskontrolle i.o | SWO-255578 | Assisted Replacement | Consumables | Solution Bottle | Cal 1 bottle |  | Rep Na QC | Na at CQ out of range | Kal 1 solution replaced. 12% but still empty. CREA membraneed Calibration repeated i.o QC repeats i.o Functional control i.o
No | Visit | Complaint | RDEU | Halogen Lamp intensity too high | Überprüfung der Oximetrie | Lamp Unit neu justiert, tHb Kalibration, Funktionskontrolle. | SWO-255397 | Adjustment | HW | Oxi module | Lamp unit | Halogen | Halogen Lamp intensity too high | Examination of oximetry | Lamp Unit re-adjusted, tHb calibration, function control.
No | Visit | Complaint | RBNL | temperature du sèchage cupule | De temps en temps une erreur apparait pendant la mesure : #1068 température hors gamme avant le sèchage. | la température dans le local était très froid (entre 10 et 15°C), la température de 70°C ne pouvait donc pas etre atteint. Chauffage de u local jusqu'à 20°C et fait le test en service du sèchage : OK Passage de 2 controles et 3 échantillons sans faute. | SWO-255852 | User error | HW | Temperature vs Enviroment | Temperature vs Enviroment |  | temperature of the cupule dryer | Occasionally an error appears during the measurement: #1068 out-of-range temperature before the dryer. | the temperature in the room was very cold (between 10 and 15 degrees Celsius), so the temperature of 70oC could not be reached. Heating up to 20oC and doing the in-service dry-dry test: OK Passing 2 controls and 3 samples without fault.
No | Visit | Complaint | RBNL | Diverse Foutcodes | AQT90 staat in FOUT en geeft diverse foutcodes aan zoals error 4794: PIPETTE_38 etc. Bij nader onderzoek blijkt dat de naald helemaal gebogen is en de optical sensor 636-549 defect te zijn. | Zowel de naald als de sensor heb vervangen vervolgens zonder fout opgestart, Systeemcheck $systeemreiniging heb uitgevoerd, OK. | SWO-255700 | Defect | HW | Aspiration Module | Cable, Aspiration Module | 4794 | Several Foutcodes | AQT90 is in ERROR and indicates various error codes such as Error 4794: PIPETTE_38 etc. Further examination shows that the needle is completely curved and the optical sensor 636-549 is defective. | Both the needle and the sensor have replaced then without error booted, System check $systeemreiniging have run, OK.
No | Visit | Complaint | RLTD | Battery error detected 1559 | Battery cell failed, confirmed in service program | Replaced battery module with new Old battery unit 910-311P0138N009 under warranty | SWO-256091 | Defect | HW | Battery module | Battery | 1559 | Battery error detected 1559 | Battery cell failed, confirmed in service program | Replaced battery module with new Old battery unit 910-311P0138N009 under warranty
No | Visit | Complaint | RESP | no baja la puerta de entrada | No baja la puerta de entrada | Se observo puerta y aguja de entrada defectuosas. se procedio a sustituirlas. se calibro equipo y se comprobo correcto funcionamiento mediante QCs | SWO-257148 | Defect | HW | Auto Inlet module | Connector gasket | and Probe | doesn't lower the front door | Doesn't lower the front door | Faulty door and entry needle were observed. were replaced. equipment is cadescribed and properly functioning by QCs
Yes | Visit | Complaint | RESP | no baja puerta de entrada | no baja puerta de entrada | Se duplica intervención para descontar pieza 834-662. | SWO-257156 | Re classified RMED | HW | Auto Inlet module | Connector gasket | and Probe | doesn't go down the front door | doesn't go down the front door | Duplicate severing intervention to decount part 834-662.
Yes | 6 | Y | RPOL | Blank | Wymiana membrany referencyjnej oraz Ca2+, kalibracja, testy. Analizator sprawny. |  | D35E5D9B5E4EBEBB | [5 empty] | Blank | Reference membrane replacement and CA2 +, calibration, testing. Functional Analyzer.
No | 5 | Y | RCZE | Blank |  |  | D35F86D1D41AE50A | Assisted Replacement | Electrode | Electrode | Reference |  | Blank
No | 4 | Y | RCZE | Blank |  |  | D35F84FFD462E506 | Defect | HW | DMS_PC | CPU |  | Blank
No | Visit | Complaint | RIND | General Visit | Clot disengaged in between inlet probe. | Flush the inlet probe with the Nacol, and dip the tubings which is conneted to inlet probe to pH $BG module. Reconnect to same as before, give the protein removal and after that give the cleaning command to the analyzer. Do the 2 point calibration, analyzer handed over to staff in well working condition. | SWO-256193 | Cleaned | HW | Inlet module | Clot, Inlet |  | General Visit | Clot disengaged in between inlet probe. | Flush the inlet probe with the Nacol, and dip the tubings which is conneted to inlet probe to pH $BG module. Reconnect to same as before, give the protein removal and after that give the cleaning command to the analyzer. Do the 2 point calibration, analyzer handed over to staff in well working condition.
No | Visit | Complaint | RAME | Instrument shut down after being unplugged during testing | Unit not holing power or not switching to battery properly when unplugged | Replaced Power supply and Battery, showing 100% charge at this time. | SWO-256682 | Defect | HW | Battery module | Battery |  | Instrument shut down after being unplugged during testing | Unit not holing power or not switching to battery properly when unplugged | Replaced Power supply and Battery, showing 100% charge at this time.
Yes | 5 | Y | RPAC | Blank | Sensor cassette failure (1361) - Internal reference electrode error in sensor cassette |  | D35FAAE2D450A7FD | [5 empty] | Blank | Sensor cassette failure (1361) - Internal reference electrode error in sensor cassette
No | Visit | Complaint | RCHN | 显示器模块故障 | 显示器模块故障，借备件910-342,910-298各一个 | 更换910-342,910-298后修复(By Ran) | SWO-255341 | Defect | HW | DMS_PC | Display |  | Display module failure | Display module failure, spare parts 910-342, 910-298 each one | Replacement 910-342, 910-298 post-repair (By Ran)
Yes | Visit | Complaint | RCHN | 光学单元超出范围，报错1031 | 定标不过，借944-394（批号13419）一个 | 重新定标后修复 by Andrew | SWO-255383 | Defect | HW | Process Module | Optical Unit | 1031 | Optical unit out of range, error 1031 | However, borrow 944-394 (lot number 13419) a | Fix by Andrew after re-calibration
No | 4 | Y | RCZE | Blank |  |  | D35FAF66D40EE511 | [5 empty] | Blank
No | Visit | Complaint | RKK | QCにてCaの値が上限に外れる。（レベル２、４） pH電極がすぐに汚れる。 | 機器状況確認 | ELモジュール部アンプチェック（電極シュミレーター使用） BGモジュール部ベースプレート上側ハウジング取り付け穴不良の為、代替ベースプレートと交換。BG側測定室交換。各部調整動作チェックOK。 記録データー患者測定結果記録レイアウト変更。 CALチェックOK。QCチェックOK | SWO-256861 | Defect | HW | pH/BG Module | Sealing ring | Sealing ring, pH | The value of Ca is removed from the upper limit in QC. (Level 2, 4) The pH electrode gets dirty immediately. | Equipment status check | EL module amplifier check (using electrode simulator) Bg module part base plate upper housing mounting hole failure, replaced with an alternate base plate. BG side measuring chamber replacement. Each part adjustment operation check OK. Recording data - Patient measurement results record layout change. CAL check OK. QC Check OK
Yes | Visit | Complaint | RAME | PO2 is on the higher range on WDC report. | PO2 2SD out of WDC ranges since go live of instrument. | Gather new data logs and service dump for further investigations. | SWO-256752 | Re classified RMED | Courtesy | WDC Setup |  |  | PO2 is on the higher range on WDC report. | PO2 2SD out of WDC ranges since go live of instrument. | Gather new data logs and service dump for further investigations.
No | Visit | Complaint | RAME | PO2 is on higher end of WDC report. | PO2 above 2SD on WDC reports since go live of instrument. | Replaced rubber sleeve on gas unit.and air valve membrane on PH &BG module. Exported data logs and service dump for further investigations on CQC case. | SWO-256689 | Defect | HW | pH/BG Module | Valve, air | Membrane | PO2 is on higher end of WDC report. | PO2 above 2SD on WDC reports since go live of instrument. | Replaced rubber sleeve on gas unit.and air valve membrane on PH &BG module. Exported data logs and service dump for further investigations on CQC case.
No | Visit | Complaint | RKK | pHリング外れ。正規リングで再発のため、やむなくダミー電極リングに交換。 | チャンバー内にリング付着。電極から外れた。 | 再発のためやむなくダミー電極Oリングに交換 | SWO-256541 | Defect | HW | pH/BG Module | Sealing ring | Sealing ring, pH | pH ring out. Because of the recurrence in the regular ring, it is forced to replace it with a dummy electrode ring. | Ring adhering into the chamber. It came off the electrode. | Replacement for dummy electrode O-ring due to recurrence
No | Visit | Complaint | RAME | 0643 ABL/DMS PC Disconnected from AU. | Inspection of the AU PCB revealed a possible poor connection of the fuses F1,f2, f3, &f4. | Replaced fuses F1, f2, f3, & f4. | SWO-256799 | Revisit required | HW | Basic Unit | PCB, Analyzer control | 0355 | 0643 ABL/DMS PC Disconnected from AU. | Inspection of the AU PCB revealed a possible poor connection of the fuses F1,f2, f3, &f4. | Replaced fuses F1, f2, f3, & f4.
Yes | 5 | Y | RFRA | Blank | Ecran noir |  | D35F8B83D40A66B0 | [5 empty] | Blank | Black screen
No | Visit | Complaint | RDEU | 13.09.2017 Sr. Bettina 02114092520 meldet, dass die Einlassnadel geknickt ist. | Einlasssonde defekt. | Einlasssonde erneuert, Gerät mehrfach kalibriert. | SWO-257390 | Defect | HW | Inlet module | Probe |  | 13.09.2017 Sr. Bettina 02114092520 reports that the inlet needle has buckled. | Inlet probe defective. | Inlet probe renewed, device calibrated several times.
No | Visit | Complaint | RAME | Spectrum errors | Spectrum errors $failed total oxi test | Replaced hemolyzer- performed THB calls | SWO-257423 | Defect | HW | Oxi module | Hemolyzer |  | Spectrum errors | Spectrum errors /failed Total Oxi test | Replaced hemolyzer- performed THB calls
No | 3 | Y | RPAC | Blank | SP 964: Leak current error in SP |  | D362281DC04D9F05 | Assisted Replacement | SP | SP | Leak current |  | Blank | SP 964: Leak current error in SP
No | Visit | Complaint | RCHN | GLU,LAC定标失败 | GLU,LAC定标失败，借903-581一个 | 更换903-581后修复(By Ran) | SWO-256537 | Defect | HW | Basic Unit | PCB, Sensor interface | Glu/Lac | GLU, LAC calibration failed | GLU, LAC calibration failed, borrowed 903-581 one | Repair after replacement 903-581 (By Ran)
No | Visit | Complaint | RCHN | 测试卡，试剂包上机无反应 | 测试卡，试剂包上机无反应，借备件903-183（被903-581替代）一个 | 更换903-581后修复(By Ran) | SWO-256394 | Defect | HW | Basic Unit | PCB, Sensor interface |  | Test card, reagent package on the machine without reaction | Test card, reagent package on the machine no response, borrowed spare parts 903-183 (replaced by 903-581) one | Repair after replacement 903-581 (By Ran)
No | Visit | Complaint | RESP | Error 05000028 TC Core (OMSG10): Gas valve error | Desinstalación del equipo para su reparación en taller | Desinstalación del equipo para su reparación en taller. | SWO-257638 | [5 empty] | Error 05000028 TC Core (OMSG10): Gas valve error | Uninstalling equipment for repair in workshop | Uninstalling the equipment for repair in the workshop.
No | Visit | Complaint | RCHN | 定标卡无法扫描 | 定标卡无法扫描,借备件902-924一个 | 更换902-924后修复(By Ran) | SWO-256015 | Defect | HW | Analyzer Control Module | Barcode reader, display | BR002 | Bid card cannot be scanned | The bid card cannot be scanned, one spare part 902-924 | Post-replacement 902-924 (By Ran)
No | Visit | Complaint | RCHN | 试管夹轮子脱落 | 试管夹轮子脱落,需要借件903-584备件一个 | 更换903-584后修复(By Ran) | SWO-256517 | Defect | HW | Cover and Chassis | Tube holder |  | Test tube clip wheels fall off | Test tube clip wheel shaved off, need to borrow 903-584 spare parts one | Repair after replacement 903-584 (By Ran)
No | Visit | Complaint | RIND | Calibration Aborted | Found Calibration Aborted And Rinse error due to Air Valve failed of PH $BG module. Air valve was failing after the connection of pressure transducer. | Check and found that pressure transducer is defective. Change pressure transducer. | SWO-256913 | Defect | HW | Gas module | Pressure Transducer |  | Calibration Aborted | Found Calibration Aborted And Rinse error due to Air Valve failed of PH /BG module. Air valve was failing after the connection of pressure transducer. | Check and found that pressure transducer is defective. Change pressure transducer.
No | Visit | Complaint | RDEU | Empfindlichkeit der pO2 Elektrode ist außerhalb der Grenzwerte, bei 4,0 mmHg. | Trotz mehrfacher Kalibration bleibt der Empfindlichkeitswert der pO2 Elektrode stabil außer Grenzen. Im Haus unterwegs zu weiteren Service. | Erneuerung der pO2 Elektrode. Nach der ersten Kalibration ist die Empfindlichkeit deutlich über Grenzwert. Alte Elektrode Reklamation, Garantie. | SWO-257255 | Assisted Replacement | Electrode | Electrode | pO2 |  | Sensitivity of the pO2 electrode is outside the limit values, at 4.0 mmHg. | Despite multiple calibration, the sensitivity value of the pO2 electrode remains stable beyond limits. In the house on the way to further service. | Renewal of the pO2 electrode. After the first calibration, the sensitivity is well above limit. Old electrode complaint, warranty.
No | Visit | Complaint | RDEU | Gerät erkennt kein LP | Kontakt der Solution Pack ID PCB abgebrochen | Solution Pack ID PCB ersetzt | SWO-257709 | Defect | HW | Basic Unit | PCB, Solution Pack Id |  | Device does not recognize an LP | Contact of Solution Pack ID PCB aborted | Solution Pack ID PCB ersetzt
Yes | Visit | Complaint | RAME | Uninstall Analyzer as per sales request. It will be replaced with ABL90 Flex Plus. | Unisntall $ Decontaminate analyzer and prepare for shipping. | Uninstalled and decontaminated analyzer. Boxed and returned on RGA#Y17-26283E. | SWO-262886 | Re classified RMED | Un install |  |  |  | Uninstall Analyzer as per sales request. It will be replaced with ABL90 Flex Plus. | Unisntall / Decontaminate analyzer and prepare for shipping. | Uninstalled and decontaminated analyzer. Boxed and returned on RGA#Y17-26283E.
No | Visit | Complaint | RAME | Po2 failed after installing a new sensor | Customer installed a new SC, because it was due. The po2 would not calibrate $Unable to calculate Po2. I performed the Po2 check and it failed the energy value for Po2 | Replaced the interface Pcb after installing the new one the SC data chip would not be read $ Ordered another Sensor interface Pcb same problem $I also tried several SC $ I have a wet section PCB ordered to install on Monday $Installed wet section PCB now sensor bay doesn't close $ | SWO-257972 | Revisit required | HW | Basic Unit | PCB, Wet Sec Control |  | Po2 failed after installing a new sensor | Customer installed a new SC, because it was due. The po2 would not calibrate /Unable to calculate Po2. I performed the Po2 check and it failed the energy value for Po2 | Replaced the interface Pcb after installing the new one the SC data chip would not be read $ Ordered another Sensor interface Pcb same problem $I also tried several SC $ I have a wet section PCB ordered to install on Monday $Installed wet section PCB now sensor bay doesn't c lose $
No | Visit | Complaint | RLTD | Lynn Shaw 01582 497991 | Flow selector error, unable to calibrate manifold in service. | Checked IR led with mobile phone camera, no light. Disassembled analyser and found cable not fitted correctly, tested cable and refitted. Led now emitting IR light and manifold calibration ok. | SWO-258392 | Adjustment | HW | Basic Unit | Cable, Flow selector detector |  | Lynn Shaw 01582 497991 | Flow selector error, unable to calibrate manifold in service. | Checked IR led with mobile phone camera, no light. Disassembled analyser and found cable not fitted correctly, tested cable and refitted. Led now emitting IR light and manifold calibration ok.
No | Visit | Complaint | RSAF | neon intensity too low, after following replacement procedure of the SERVICE MANUAL lastly power supply to be replaced as no stock available when replacing other items. | Queries on all Qc's and patient result intermittingly on oximetry values, | Replaced Spectrometer and done the complete oxi test and value on neon intensity remains low 0.24 ? pA | SWO-258962 | Revisit | HW | Oxi module | Oxi module |  | neon intensity too low, after following replacement procedure of the SERVICE MANUAL lastly power supply to be replaced as no stock available when replacing other items. | Queries on all Qc's and patient result intermittingly on oximetry values, | Replaced Spectrometer and done the complete oxi test and value on neon intensity remains low 0.24 ? pA
No | Visit | Complaint | RCHN | LAC,GLU定标不过 | LAC,GLU定标不过，借备件903-581一个 |  | SWO-257459 | Defect | HW | Basic Unit | PCB, Sensor interface | Glu/Lac | LAC, GLU calibration but | LAC, GLU calibration, borrowed spare parts 903-581 one
No | Visit | Complaint | RCHN | LAC定标不过 | LAC定标不过，借备件903-581一个 | 更换903-581后修复(By Ran) | SWO-256524 | Defect | HW | Basic Unit | PCB, Sensor interface | Lac | LAC calibration is not | LAC calibration is not the same, borrowed spare parts 903-581 one | Repair after replacement 903-581 (By Ran)
No | Visit | Complaint | RCHN | LAC定标不过 | LAC定标不过,借备件903-581一个 | 更换903-581后修复 | SWO-256527 | Defect | HW | Basic Unit | PCB, Sensor interface | Lac | LAC calibration is not | LAC calibration is not the same, borrowed spare parts 903-581 one | Repair after replacement 903-581
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | CO2质控低于正常范围 1397 | 更换试剂包后修复 | SWO-258230 | Assisted Replacement | SP | SP | pCO2 |  | Test card, reagent package failure | CO2 QC BELOW NORMAL RANGE 1397 | Repair after replacement of reagent kit
No | Visit | Complaint | RCHN | 报错5257 | 报错5257，借636-550一个 | 更换636-550备件后修复(By Ran) | SWO-256538 | Defect | HW | Process Module | Process Unit sensor | 5257 | 5257 error | 5257, loan 636-550 one | Repair after replacement of 636-550 spare parts (By Ran)
No | Visit | Complaint | RCHN | 管夹故障 | 管夹故障，借886-731（被903-584替换） $2个 | 更换903-584 2个后修复by tim | SWO-256865 | Defect | HW | Cover and Chassis | Tube holder |  | Tube clip failure | Tube clip failure, borrow 886-731 (replaced by 903-584) /2 | Replace 903-584 2 after repair by tim
No | Visit | Complaint | RKK | サンプル移動エラーとの事でBML検査センターより連絡あり。訪問し確認、ガスケット洗浄、排液ライン洗浄、タンパク除去実施。 | サンプル移動エラーとの事でBML検査センターより連絡あり。訪問し確認、ガスケット洗浄、排液ライン洗浄、タンパク除去実施。 | サンプル移動エラーとの事でBML検査センターより連絡あり。訪問し確認、ガスケット洗浄、排液ライン洗浄、タンパク除去実施。 | SWO-261465 | Cleaned | Fluidic pathway | Fluidic pathway | Clot, Inlet |  | There is a communication from the BML inspection center by the thing with the sample movement error. Visited and confirmed, gasket cleaning, drainage line cleaning, protein removal. | There is a communication from the BML inspection center by the thing with the sample movement error. Visited and confirmed, gasket cleaning, drainage line cleaning, protein removal. | There is a communication from the BML inspection center by the thing with the sample movement error. Visited and confirmed, gasket cleaning, drainage line cleaning, protein removal.
No | Visit | Complaint | RDEU | Monitor Display ohne Funktion 03.11.2017 10:30 Fr. Milke 036458543436: Monitor weiterhin ohne Funktion | Display läßt sich schlecht bedienen, hängt | Überprüfung Hardware check Monitor sollte mit Firmware aufgerüstet werden Service Tip17.008 entsprechende SW wird besorgt und in ABL eingespielt Geräteneustart - Problem bleibt bestehen | SWO-257748 | Defect | HW | DMS_PC | Touch screen |  | Monitor display without function 03/11/2017 10:30 Fr. Milke 036458543436: Monitor still without function | Display is difficult to operate, hangs | Review Hardware check Monitor should be upgraded with firmware Service Tip17.008 corresponding SW is taken care of and recorded in ABL Device restart - problem remains
No | Visit | Complaint | RAME | Level 3 oximetry questioning on linearity. | Oxi questioning while running linearity. Made sure customer was not in lock out period. Validated errors. | Replaced oxi module complete. Replaced all the membranes for customer as they were due to expire soon, and analyzer kept questioning during calibrations. | SWO-257435 | Defect | HW | Oxi module | Oxi module, Replaced |  | Level 3 oximetry questioning on linearity. | Oxi questioning while running linearity. Made sure customer was not in lock out period. Validated errors. | Replaced oxi module complete. Replaced all the membranes for customer as they were due to expire soon, and analyzer kept questioning during calibrations.
No | Visit | Complaint | RDEU | 03.11.2017; Fr. Unterländer; Tel.: 063524053972; Neon lamp intensity too low | Oximodul defekt nein Lamp 46 % intensit | Oximetriemodul gewechselt , Oxi Complete Test bestanden , Pump Kal. Durchgeführt. | SWO-257826 | Defect | HW | Oxi module | Lamp Unit | Neon Lamp | 03.11.2017; Ms. Unterländer; Tel.: 063524053972; Neon lamp intensity too low | Oximodul defective no Lamp 46% intensit | Oximetry module changed , Oxi Complete Test passed , Pump Kal. Performed.
Yes | Visit | Complaint | RSCH | Neudefekt | OXI Lampenwert nimmt extrem schnell ab pO2 QC ausserhalb | OXI Halogen Lampe ausgetauscht und neunjustiert auf 3400 (Blackstock, kann drinn bleiben) pH Dichtung ausgetauscht da original Dichtung undicht ung in der Messkammer geklebt. Funktionskontrolle i.o | SWO-257728 | Re classified RMED | OOB |  |  |  | Neudefekt | OXI lamp value decreases extremely fast pO2 QC outside | OXI halogen lamp replaced and nine-adjusted to 3400 (Blackstock, can stay in) pH seal replaced as original seal leaked in the measuring chamber. Functional control i.o
No | Visit | Complaint | RAME | Clinical trial analyzer down due to occasional sample indoor failure to close, therefore aborting all samples onboard or in progress. | Could not replicate error. Open $close inlet door multiple times successfully. Also finished a QC without issue. Did hear squeaky noise when door open $close. | To avoid error recurring, affecting trial study results, replaced inlet drive belt, pn 835-493, from pm kit (drive belt alone was unavailable in stock). Also transferred parts from front panel Vincent Sigmund sent. Cleaned and lubricated sample inlet doors. Squeaky noise eliminated. | SWO-257859 | Adjustment | HW | Cover and Chassis | Cartridge door | squeking noise | Clinical trial analyzer down due to occasional sample indoor failure to close, therefore aborting all samples onboard or in progress. | Could not replicate error. Open /close inlet door multiple times successfully. Also finished a QC without issue. Did hear squeaky noise when door open /close. | To avoid error recurring, affecting trial study results, replaced inlet drive belt, pn 835-493, from pm kit (drive belt alone was unavailable in stock). Also transferred parts from front panel Vincent Sigmund sent. Cleaned and lubricated sample inlet doors. Squeaky noise eliminated.
No | Visit | Complaint | RDEU | 02.11.2017, 11:56 Uhr, Fr. Fraas, Tel. 06062797708: Inkubationstemp. zu hoch 02.11.2017 14:10 Fr. Fraas 06062797708: Inkubationstemp. weiterhin zu hoch | Überprüfung Temperatur des Incubator Wheels. Es liegt bei über 44 Grad Celsius, die Temperatur der Heizplatte bei über 51 Grad Celsius. Vmtl. Regelung der Incubator PCB defekt. | Ausbau des alten Incubator PCP und Einbau des neuen Incubator PCB. Funktion ist in Ordnung. Temperatur nicht mehr außerhalb der Grenzwerte. Reklamation und Gewährleistung, da bereits das gleiche Ersatzteil bereits dieses Jahr erneuert. | SWO-257890 | Defect | HW | Analyzer Control Module | PCB, Incubator | 3264 | 02.11.2017, 11:56, Fr. Fraas, tel. 06062797708: incubation temp. too high 02.11.2017 14:10 Fr. Fraas 06062797708: Incubation Temp. continue to be too high | Check temperature of the Incubator Wheel. It is over 44 degrees Celsius, the temperature of the heating plate is over 51 degrees Celsius. Vmtl. Control of the Incubator PCB defective. | Expansion of the old Incubator PCP and installation of the new Incubator PCB. Function is fine. temperature no longer outside the limits. Complaint and warranty, as the same spare part has already been renewed this year.
No | Visit | Complaint | RDEU | AQT läßt sich nicht bedienen, SW hat sich aufgehangen | AQT läßt sich nicht bedienen, Aufnahme Kassetten nicht möglich | Fehlersuche im Service Programm safe position durchlaufen lassen, anschl. o.k., Aufnahme wieder möglich PCT wurde nicht kalibriert, AQT hat Kalibrierkas. verworfen,erneute Kalibrierung nicht möglich, da Kal.kassette nicht erkannt wird Geräteneustart Systemreinigung und QC Messung durchgeführt | SWO-257637 | No fault found | No fault found | No fault found | No fault found |  | AQT cannot be operated, SW has hung up | AQT cannot be operated, Recording cassettes not possible | Troubleshooting in the service program safe position, followed by ok, recording possible again PCT has not been calibrated, AQT has calibration kas. discarded,recalibration is not possible because Kal.cassette is not recognized Device restart System cleaning and QC measurement performed
No | Visit | Complaint | RDEU | Der neue AQT macht bei Messungen sehr laute Geräusche! | Vermutlich Software Problem. Fehler nicht festgestellt. Datensicherung wird an den Radiometer Support kommuniziert. | Gerät macht sporadisch einen Neustart. Es wird automatisch auch ein verzehrter Ausdruck ausgedruckt. Der Touch Screen reagiert in diesen Momenten verzögert. Das Einscannen von Barcodes funktioniert in diesen Momenten, wenn der Fehler auftritt auch nicht. Bezüglich des Messverhaltens und deren Qualität gibt es keine Beanstandung. Service Dump und Datensicherung durchgeführt. | SWO-257724 | No fault found | HW | HW | Noise during measuring |  | The new AQT makes very loud noises when measuring! | Probably software problem. Error not detected. Data backup is communicated to the Radiometer Support. | Device makes a sporadic restart. A consumed printout is also printed automatically. The touch screen reacts delayed in these moments. Scanning barcodes works in these moments, even if the error does not occur. There is no complaint regarding the measurement behaviour and its quality. Service dump and data backup performed.
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | lac定标不过 | 更换测试卡后修复 | SWO-258170 | [5 empty] | Test card, reagent package failure | lac calibration is not | Repair after replacing test card
Yes | Visit | Complaint | RCHN | 光学单元超范围 | 光学单元超范围，借903-364,905-960,910-324各一个. Error code 1031. Can't read the barcode of Tni Cal cartridge. | 更换903-364,905-960,910-324后修复(By Ran) | SWO-257058 | Defect | HW | Process Module | Optical Unit | 1031 | Optical unit over range | Optical unit over range, borrow 903-364, 905-960, 910-324 each. Error code 1031. Can't read the barcode of Tni Cal cartridge. | Replacement 903-364, 905-960, 910-324 post-repair (By Ran)
No | Visit | Complaint | RCHN | 测试卡无法扫描，报错3410 | 测试卡无法扫描and报错3410，借备件910-304 903-329各一. Can't read the barcode, and error code 3410. | 更换备件910-304及903-329后修复(By Ran) | SWO-257995 | Defect | HW | Storage Module | Barcode reader, Storage wheel | BR001 | Test card can't be scanned, error 3410 | The test card was unable to scan and report error 3410, borrowing spare parts 910-304 903-329 each one. Can't read the barcode, and error code 3410. | Replacement of spare parts 910-304 and 903-329 after repair (By Ran)
Yes | 12 | Y | RPAC | Blank | Glu and Lact parameters take more than 4 hours to be ready |  | D3660F18221277E0 | [5 empty] | Blank | Glu and Lact parameters take more than 4 hours to be ready
Yes | 13 | Y | RPAC | Blank | QC Level C low on p02 |  | D36716449863756D | [5 empty] | Blank | QC Level C low on p02
No | Visit | Complaint | RAME | Battery error code 1559 | Battery Mode Status: Battery cell fault | 1) Replaced battery and tested. 2) Enabled DST. Service Tip 721029809. | SWO-258704 | Defect | HW | Battery module | Battery | 1559 | Battery error code 1559 | Battery Mode Status: Battery cell fault | 1) Replaced battery and tested. 2) Enabled DST. Service Tip 721029809.
No | Visit | Complaint | RSAF | oxi hadware error on calibrations intermittently | Hemolyzer, spectrometer and WSC PCB already replaced, possible power supply fault. | Unit taken for repair and loan installed 092R0129N0012 | SWO-259163 | Revisit | HW | Oxi module | Oxi module |  | oxi hadware error on calibrations intermittently | Hemolyzer, spectrometer and WSC PCB already replaced, possible power supply fault. | Unit taken for repair and loan installed 092R0129N0012
No | Visit | Complaint | RCHN | 试剂包卡不住 | 试剂包卡不住，借903-363一个 |  | SWO-258278 | [5 empty] | The reagent bag won't work. | Reagent bag can't be stuck, borrow 903-363 one
No | Visit | Complaint | RCHN | 开机一直不停重启 | 开机一直不停重启，需要借件910-329一个 | 更换备用件910-329后修复 by jesson | SWO-258189 | Defect | HW | DMS_PC | CF Card |  | Power-on keeps restarting | Power on has been non-stop restart, need to borrow 910-329 one | Repair by jesson after replacing spare parts 910-329
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | O2定标不过 0378 | 更换测试卡后修复 | SWO-259439 | Assisted Replacement | SC | SC | pO2 | 378 | Test card, reagent package failure | O2 calibration is no more than 0378 | Repair after replacing test card
No | Visit | Complaint | RCHN | glu和lac定标不过 | glu和lac定标不过，需要借件903-581 |  | SWO-258201 | Defect | HW | Basic Unit | PCB, Sensor interface | Glu/Lac | glu and lac calibration however | glu and lac calibration, however, need to borrow 903-581
No | Visit | Complaint | RLTD | Mori Tel 01865 221596 | On arrival the waste line from the fluidic module was disconnected but the one from the y-piece to waste was still connected, you could see that blood had leaked out around the waste pump. | Flushed both of the waste lines and there was no sign of a blockage in either one, also visually inspected the waste lines and no blood clots could be seen. Reconnected the waste tubes and checked the connections and they both seemed to be secure. Cleaned up the blood from around the waste pump roller and ran 10 rinse cycles with no problems. I will order a new waste connector just in case there is a problem with this, | SWO-258635 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway, Clot |  | Mori Tel 01865 221596 | On arrival the waste line from the fluidic module was disconnected but the one from the y-piece to waste was still connected, you could see that blood had leaked out around the waste pump. | Flushed both of the waste lines and there was no sign of a blockage in either one, also visually inspected the waste lines and no blood clots could be seen. Reconnected the waste tubes and checked the connections and they both seemed to be secure. Cleaned up the blood from around the waste pump roller and ran 10 rinse cycles with no problems. I will order a new waste connector just in case there is a problem with this,
No | Visit | Complaint | RAME | Customer reported Oximetry errors. Neon Number and Intensity too low (codes 0316 and 0258) | Unit failing oximetry tests for the Lamp unit. Neon Intensity and Neon Position. | Replaced lamp unit. Adjusted halogen lamp Intensity. Ran complete oximetry tests and tHB calibration with pass results. restart calibrations. | SWO-258699 | Defect | HW | Oxi module | Lamp unit |  | Customer reported Oximetry errors. Neon Number and Intensity too low (codes 0316 and 0258) | Unit failing oximetry tests for the Lamp unit. Neon Intensity and Neon Position. | Replaced lamp unit. Adjusted halogen lamp Intensity. Ran complete oximetry tests and tHB calibration with pass results. restart calibrations.
Yes | Visit | Complaint | RDEU | Monitor Display ohne Funktion 03.11.2017 10:30 Fr. Milke 036458543436: Monitor weiterhin ohne Funktion | Monitor mit fehlfunktion | Upgrade Display mit Firmware nicht möglich,da keine funktionierende Anmeldung im BIOS; Austausch Display Unit Funktionstest i.O. Einlaufverfahren, da ABL abgeschaltet war | SWO-258288 | Re classified RMED |  |  |  | LEH email "what going on" | Monitor display without function 03/11/2017 10:30 Fr. Milke 036458543436: Monitor still without function | Monitor with malfunction | Upgrade display with firmware not possible,because no working login in the BIOS; Replacement Display Unit Functional test i.O. Inlet procedure because ABL was switched off
No | Visit | Complaint | RDEU | Schwester Heike meldet die Fehlermeldung 887 | Datenbank geprüft | Problem mit Station besprochen. Flexmodus ist gem. Handbuch beim ABL827 nicht verfügbar | SWO-258486 | User error | SW | Setup | Adaptive Meas Mode |  | Sister Heike reports the error message 887 | Database checked | Issue discussed with station. Flex mode is not available according to manual on The ABL827
No | Visit | Complaint | RAME | Autocheck door malfuntioning | AutoCheck motor seems too weak to pull open lid. Checked lid motor resistance. | Exchanged AutoCheck module for fault isolation. Run AU engine update. Restart calibrations. | SWO-258789 | Defect | HW | Auto Check Module | Auto Check Module | door not opening | Autocheck door malfuntioning | AutoCheck motor seems too weak to pull open lid. Checked lid motor resistance. | Exchanged AutoCheck module for fault isolation. Run AU engine update. Restart calibrations.
Yes | Visit | Complaint | RDEU | Fehler FlexQ / wird nicht verwendet | Fehler FlexQ | Analysator ausgeschaltet. FlexQ Modul deaktiviert. SafePico werden nicht verwendet. Abdeckung wird bestellt und ausgetauscht. Neustart Initialisierungsphase QC ohne Beanstandung | SWO-258523 | Re classified RMED | HW | FlexQ Module | Cover |  | Error FlexQ / is not used | Error FlexQ | Analyzer switched off. FlexQ module deactivated. SafePico are not used. Cover is ordered and replaced. Restart Initialization phase QC without complaint
No | Visit | Complaint | RSCH | pH instabil | pH instabil | Zwei O-Ringe Medizintechnik abgegeben | SWO-258588 | Defect | Electrode | Electrode | Sealing ring | Sealing ring, pH | pH unstable | pH unstable | Two O-rings medical technology delivered
No | Visit | Complaint | RSCH | pH instabil | Sporadisch pH instabil | Allgemeine Kontrolle i.o O-Ring pH ersetzt i.o Kal. und QC i.o Testmessung i.o | SWO-258589 | Defect | Electrode | Electrode | Sealing ring | Sealing ring, pH | pH unstable | Sporadic pH unstable | General control i.o O-ring pH replaces i.o Kal. and QC i.o Test measurement i.o
No | Visit | Complaint | RAME | Intermittent error 0315 neon number too high. | Performed a complete OXI test with no errors. Replaced the OXI module PCB had same problem. Replaced the lamp unit had same problem. | Replaced OXI module PC board for diagnostic reasons. | SWO-258866 | Unknown | HW | Oxi module | Lamp unit | 315 | Intermittent error 0315 neon number too high. | Performed a complete OXI test with no errors. Replaced the OXI module PCB had same problem. Replaced the lamp unit had same problem. | Replaced OXI Module PC Board for diagnostic reasons.
No | Visit | Complaint | RDEU | 06.11.2017, 09:11 Uhr, Fr. Eilers-Gladrow, Tel. 05315952045: Gerät reagiwert auf keine Eingaben mehr | Sicherung auf der AU PCB defekt | Computereinheit demontiert und geprüft, Power Supply ausgetauscht, Sicherung auf der AU PCB ersetzt, Analysator in die Werkstatt abgezogen | SWO-259631 | Defect | HW | Basic Unit | PCB, Analyzer control |  | 06.11.2017, 09:11, Ms. Eilers-Gladrow, tel. 05315952045: Device reagiwert on no more inputs | Fuse on the AU PCB defective | Computer unit dismantled and tested, Power Supply replaced, fuse replaced on AU PCB, analyzer subtracted into workshop
No | 2 | Y | RCZE | Blank | pøi inicializaci chyby 5784:SETSAFE_22 a 5689:ROTATESTORAGEWHEEL_24, pøístroj v Ifcoru na Vinièní 2.patro |  | D365F8A722146477 | Adjustment | HW | Storage Module | Storage Module |  | Blank | When initializing the error 5784: SETSAFE_22 and 5689: ROTATESTORAGEWHEEL_24, the device in Ifcoru on the Vinière 2nd floor
No | Visit | Complaint | RIND | Process wheel error. | Process wheel error. | Check the analyzer found process wheel was not moving, found test cup was fell inside the process wheel, removed the process wheel, removed the test cup, restarted analyzer found instrument is working satisfactorily. | SWO-261378 | Defect | Consumables | Solution Pack | Stack | LEH | Process wheel error. | Process wheel error. | Check the analyzer found process wheel was not moving, found test cup was fell inside the process wheel, removed the process wheel, removed the test cup, restarted analyzer found instrument is working satisfactorily.
No | Visit | Complaint | RAME | Oxi Hardware Error | Error codes 1016, 1019, 1029, 375, 376, 1034. Oxi hardwar problems | Replaced the hemolyzer. Tried rebooting, running service programs and replacing the spectrometer but none of them worked. | SWO-258857 | Revisit required | HW | Oxi module | Oxi module |  | Oxi Hardware Error | Error codes 1016, 1019, 1029, 375, 376, 1034. Oxygen Haridwar Problems | Replaced the hemolyzer. Tried rebooting, running service programs and replacing the spectrometer but none of them worked.
No | Visit | Complaint | RDEU | 02.11.2017; Fr. Jisiew; Tel.: 01777970834; Intervention erforderlich, Lösungen können nicht gepumpt werden 06.11.2017, 10:32 Uhr, Fr. Yildiz, Tel. 0307201100: QC Fehler | Fehlersuche, häufig QC Fehler bei Oxiparametern, Aktivitätenspeicher durchsucht - Spüllösungsfehler - wird oft nicht erkannt, Dichtungen und Oxischlauch überprüft, | Einlaßdichtung korrekt angebracht, Schlauch Messkammer zu Oxi erneuert - rutschte immer runter, System mehrfach gespült, Oxitest und Systemtest durchgeführt, Neustart, Kal, tHb Kal, QC | SWO-258986 | Defect | HW | Basic Unit | Tubings | 1169 | 02.11.2017; Ms. Jisiew; Tel.: 01777970834; Intervention required, solutions cannot be pumped 06.11.2017, 10:32, Ms. Yildiz, tel. 0307201100: QC Error | Troubleshooting, common QC errors in oxi parameters, activity memory searched - flushing solution error - is often not detected, seals and oxi hose checked, | Inlet seal correctly attached, hose measuring chamber renewed to Oxi - always slipped down, system repeatedly flushed, Oxitest and system test performed, restart, Kal, tHb Kal, QC
No | 7 | Y | RPAC | Blank | Faulty Solution Pack |  | D3675EB09827028C | Assisted Replacement | SP | SP | pO2 | B Low | Blank | Faulty Solution Pack
No | Visit | Complaint | RCHN | 定标终止 | 定标终止，借944-258（批号13280）一个 |  | SWO-258858 | Assisted Replacement | Consumables | Test Kit | Cartridge, NT-PROBNP |  | Bid termination | Bid termination, borrowing 944-258 (lot number 13280) one
No | Visit | Complaint | RIND | Service Visit | In printer head latch is not working properly. | Open the CPU section, remove the screw of printer section, analyze it and found the pin of the printer head not in its position, after open full module of printer head observe that right side gear section also found broken. Need to change the printer head. | SWO-259146 | Defect | HW | Printer Module | Printer head | Revisit ?? | Service Visit | In printer head latch is not working properly. | Open the CPU section, remove the screw of printer section, analyze it and found the pin of the printer head not in its position, after open full module of printer head observe that right side gear section also found broken. Need to change the printer head.
No | Visit | Complaint | RKK | QCにてCaの値が上限に外れる。（レベル２、４） pH電極がすぐに汚れる。 | 機器状況確認 | オキシモジュール部ピンチバルブ分解、清掃、取り付け金具調整(QCにて1度だけOXIエラー発生の為） 電源 $UPSライン電源線取り回し変更、BG,El,ダミー,オキシモジュールアースチェックOK Elモジュール代替品と入れ替え | SWO-261463 |  | Unknown |  |  | LEH Mail | The value of Ca is removed from the upper limit in QC. (Level 2, 4) The pH electrode gets dirty immediately. | Equipment status check | Oxy module part pinch valve decomposition, cleaning, mounting bracket adjustment (for oxi error generation only once in QC) Power supply $UPS line power line change, BG, El, dummy, oxy module ground check OK Replacing el module replacements
Yes | 9 | Y | RCZE | Blank |  |  | D36762AB9827A3B8 | [5 empty] | Blank
No | Visit | Complaint | RAME | New oxi module just installed is now giving Neon lamp and voltage errors. Not possible to measure. All oxi values questioned. | I had just installed new oxi module complete last week. A day later it appeared to be giving Neon voltage and lamp errors. | Tried replacing spectrometer on the new oxi module. This had no affect. Replaced the whole oxi module again. Problem went away. Performed all necessary calibrations. | SWO-260141 | Defect | HW | Oxi module | Oxi module, Replaced | Second replacment | New oxi module just installed is now giving Neon lamp and voltage errors. Not possible to measure. All oxi values questioned. | I had just installed new oxi module complete last week. A day later it appeared to be giving Neon voltage and lamp errors. | Tried replacing spectrometer on the new oxi module. This had no affect. Replaced the whole oxi module again. Problem went away. Performed all necessary calibrations.
No | Visit | Complaint | RAME | PH ELECTRODE FAILURE | PH ELECTRODE FAILURE TO CALIBRATE. | REPLACED PH ELECTRODE WITH NEW ELECTRODE. FOUND NO OTHER SIGNS OF FAILURES ON THE ANALYZER. | SWO-259340 | Assisted Replacement | Electrode | Electrode | pH |  | PH ELECTRODE FAILURE | PH ELECTRODE FAILURE TO CALIBRATE. | REPLACED PH ELECTRODE WITH NEW ELECTRODE. FOUND NO OTHER SIGNS OF FAILURES ON THE ANALYZER.
No | Visit | Complaint | RAME | Oxi HW error 0579 & 316 Neon number too low. | Inspection of the Oxi system revealed the initial error was bean lamp voltage too low, replaced Ofilia lamp unit & Oxi PCB. Now getting neon position error. Ordering additional parts. | Installed a lamp unit & Oxi PCB. Suspect replacement Oxi lamp needs to be replaced again. | SWO-259365 | Defect | HW | Oxi module | Lamp unit | 0316 | Oxi HW error 0579 & 316 Neon number too low. | Inspection of the Oxi system revealed the initial error was bean lamp voltage too low, replaced Ofilia lamp unit & Oxi PCB. Now getting neon position error. Ordering additional parts. | Installed a lamp unit & oxygen PCB. Suspect replacement oxygen lamp needs to be replaced again.
Yes | Visit | Complaint | RAME | Installation procedure revealed level 3 VK-R7 lot 490 is failing Oxi parameters and voting out of range on SO2 & negative -49.5 FO2Hb. | Installation procedure revealed level 3 VK-R7 lot 490 is failing Oxi parameters and voting out of range on SO2 & negative -49.5 FO2Hb. | Ordering a hemolyzer. | SWO-259350 | Re classified RMED | doc in OOB | Oxi module | Hemolyzer |  | Installation procedure revealed level 3 VK-R7 lot 490 is failing Oxi parameters and voting out of range on SO2 & negative -49.5 FO2Hb. | Installation procedure revealed level 3 VK-R7 lot 490 is failing Oxi parameters and voting out of range on SO2 & negative -49.5 FO2Hb. | Ordering a hemolyzer.
No | Visit | Complaint | RAME | The battery has been plugged in the wall outlet 24/7 but the charge displays shows 0%. | Checked and, confirmed the battery option is not charging. | Replaced a new battery. Checked and, confirmed the battery is now charging properly. The battery had a charge of 67% and, after 35 minutes the charge went up to 100%. | SWO-260078 | Defect | HW | Battery module | Battery |  | The battery has been plugged in the wall outlet 24/7 but the charge displays shows 0%. | Checked and, confirmed the battery option is not charging. | Replaced a new battery. Checked and, confirmed the battery is now charging properly. The battery had a charge of 67% and, after 35 minutes the charge went up to 100%.
No | Visit | Complaint | RDEU | 08.11.2017; Hr. Buschhausen; Tel.: 01718112403; ABL kalibriert ständig seit der Aufstellung | sporadisch Fehler 1324 und um 0:00 Uhr Fehler 1350 | LPs ersetzt, Funktion z.Zt. OK. | SWO-260784 | Assisted Replacement | SP | SP | Fluid transport |  | 08.11.2017; Mr. Buschhausen; Tel.: 01718112403; ABL continuously calibrated since installation | sporadical error 1324 and at 0:00 o'clock error 1350 | LPs replaced, function currently OK.
No | Visit | Complaint | RSAF | Analyzer failing calibrations & Qc's on pH and electrolytes due to sensor impedance error & Sensor thermistor error/re calibration failed (replace Sensor Cassette) | looked for most frequent error codes and trouble shooting done in accordance, which was Sensor Thermistor re-calibration failed, Parameter response fault, Sensor impedance error, | Inspected Sensor Bay complete and removed white crystals from the contact probe array, Ran Sensor Simulator Test and all passed - but keeps failing initialization of S $C, Removed and inspected Sensor Interface PCB and most of the contact pins on the Sensor probe contact array where all Black, Replaced Sensor Interface PCB and passed the Thermistor test first time, But still failed to initialize the S $C, installed a new S $C and passed all Qc's and Calibrations. | SWO-260177 | Defect | SC | SC | Leakage | Sensor IF replaced, crystals | Analyzer failing calibrations & Qc's on pH and electrolytes due to sensor impedance error & Sensor thermistor error/re calibration failed (replace Sensor Cassette) | looked for most frequent error codes and trouble shooting done in accordance, which was Sensor Thermistor re-calibration failed, Parameter response fault, Sensor impedance error, | Inspected Sensor Bay complete and removed white crystals from the contact probe array, Ran Sensor Simulator Test and all passed - but keeps failing initialization of S $C, Removed and inspected Sensor Interface PCB and most of the contact pins on the Sensor probe contact array where all Black, Replaced Sensor Interface PCB and passed the Thermistor test first time, But still failed to initialize the S $C, installed a new S $C and passed all Qc's and Calibrations.
Yes | 24 | Y | RPNZ | Blank | Sensor cassette replacement due to 1361: internal reference electrode error in sensor cassette. |  | D36DC31BE663A510 | [5 empty] | Blank | Sensor cassette replacement due to 1361: internal reference electrode error in sensor cassette.
No | Visit | Complaint | RCHN | 接口板腐蚀 | 接口板腐蚀，借902-972一个。 | 更换902-972一个正常。 by Andrew | SWO-258988 | [5 empty] | Interface plate corrosion | Interface plate corrosion, borrow 902-972 one. | Replace 902-972 one normal. by Andrew
No | Visit | Complaint | RBNL | Afgekeurde controles | Klant heeft aangegeven dat de controles regelmatig buiten de grenswaarden liggen. Hierbij komt het verkeerslicht niet op geel te staan. Bij controle is gebleken dat de verbindingsconnector tussen sensorcassette en de analyzer gecorrodeerd is, deze moet vervangen worden, bestelling wordt doorgegeven aan kantoor. | Bij controle van 3 levels blijkt dat level 1 een te lage pH en te hoge pCO2 rapporteert. Afwachten of dit gecorrigeerd wordt met een nieuwe connector. Instelling aangepast zodat bij afgekeurde parameter voor qc deze een geel licht geeft bij de status. | SWO-259843 | [5 empty] | Rejected Checks | Customer has indicated that the checks are regularly outside the limit values. The traffic light does not appear to be yellow. When checked, the connection connector between the sensor cassette and the analyzer has been corroded, it has to be replaced, order is passed to the office. | When checking 3 levels it appears that level 1 reports too low a pH and too high pCO2. To see if this is corrected with a new connector. Adjusted so that in the rejected parameter for QC, it gives a yellow light to the status.
No | Visit | Complaint | RKK | キャルが取れない | ケース詳細の通り | キャリブレーションチャンバーの液汚れ、除去。メンブラン交換。 | SWO-269379 | [5 empty] | I can't get the cal. | As the case details | Liquid stains and removal of calibration chamber. Membrane replacement.
No | Visit | Complaint | RCHN | 报错1068 | 报错1068，借903-210一个 | 更换903-210后修复 by Andrew | SWO-259438 | Defect | HW | Process Module | PCB, Cup dryer | 1068 | 1068 error | 1068, borrow edith 903-210 | Fix by Andrew after replacing 903-210
No | 28 | Y | RCZE | Blank |  |  | D3684A4B0E4CE531 | Defect | HW | pH/BG Module | Pump tubes |  | Blank
Yes | Visit | Complaint | RKK | QCにてCaの値が上限に外れる。（レベル２、４） pH電極がすぐに汚れる。 | 機器状況確認 | モジュール復旧作業 CALチェックOK。QCチェックOK | SWO-261469 | Re classified RMED | Courtesy |  |  |  | The value of Ca is removed from the upper limit in QC. (Level 2, 4) The pH electrode gets dirty immediately. | Equipment status check | Module recovery work CAL check OK. QC Check OK
No | Visit | Complaint | RDEU | 08.11.2017, 09:24 Uhr, Fr. Eschmann, Tel. 020885083200: Halterung für Probenröhrchen abgebrochen | Probenhalterung abgebrochen, | Neue Halterung eingesetzt, | SWO-259701 | Defect | HW | Cover and Chassis | Tube holder |  | 08.11.2017, 09:24, Ms. Eschmann, Tel. 020885083200: Bracket for sample tubes aborted | Sample holder aborted, | New bracket inserted,
No | Visit | Complaint | RAME | Unable to log in to analizer | Customer not able to connect new instrument to LIS. After review with LIS, instrument not set up in external system. | Customer requested LIS to set up new instruments. | SWO-261031 | Setup | His/Lis | His/Lis | His/Lis |  | Unable to log in to analizer | Customer not able to connect new instrument to LIS. After review with LIS, instrument not set up in external system. | Customer requested LIS to set up new instruments.
Yes | 6 | Y | RPAC | Blank | Error (378) Calibration sensitivity out of range on pH |  | D369927484274788 | [5 empty] | Blank | Error (378) Calibration sensitivity out of range on pH
No | 2 | Y | RPOL | Blank | Aktualizacja orpogramownia do wersji 3.3. Kalibracja tHb. Analizator sprawny technicznie. |  | D369C7DD843C142C | Defect | SW | RMED SW | Restart auto |  | Blank | Update the Orpogramownia to version 3.3. THb Calibration. Technically efficient analyzer.
No | Visit | Complaint | RCHN | 接口板腐蚀 | 接口板腐蚀，借902-972一个 |  | SWO-260221 | [5 empty] | Interface plate corrosion | Interface plate corrosion, borrow 902-972 one
No | Visit | Complaint | RKK | 血液漏れ。 | 血液漏れのため、修理代替器と入れ替え | 入れ替え実施 | SWO-261404 | [5 empty] | blood leak. | Replacement for repair replacement due to blood leak | Replacement implementation
No | Visit | Complaint | RCHN | 测试卡扫描器不能扫描，测试卡槽拖落 | 测试卡扫描器不能扫描，测试卡槽拖落，借910-304（北办发货） | 更换910-304一个正常。 by jesson | SWO-259595 | Defect | HW | Storage Module | Barcode reader, Storage wheel | BR001 | Test card scanner cannot be scanned, test card slot is dragged down | Test card scanner can not be scanned, test card slot drag, borrow 910-304 (North to ship) | Replace 910-304 one normal. by jesson
No | Visit | Complaint | RAME | Customer reported intermittent leakage from inlet port. | Checked fluid flow from inlet gasket. | Exchanged multiple inlet components and tubing for fault isolation. Replaced inlet nipple kit and pump rotor type B. | SWO-260053 | Defect | HW | Inlet module | Nipple kit |  | Customer reported intermittent leakage from inlet port. | Checked fluid flow from inlet gasket. | Exchanged multiple inlet components and tubing for fault isolation. Replaced inlet nipple kit and pump rotor type B.
No | Visit | Complaint | RIND | Service Visit | Clotting appeared in pH chamber, leakage symptoms also appeared from pH chamber. | Open the chamber and remove the electrode from the chamber, remove the cap and O-ring from the electrode and dip in to NaOCl, after 5 min take away from solution and wash with the water and wipe out with gauze. Now clean the pH chamber. Remove the inlet gas kit, inlet probe tubing which is connect to probe and pH $BG module. All these parts flush with the NaOCl solution and wash with the water, wipe out with gauze and reconnect all the parts same as before. Now restart the analyzer, wait for auto calibration. Do protein removal and cleaning command to the analyzer and then give 2 point calibration. Now tell the user for sample, user do the sample, no query message appearing in pH during result. Measurement unstable problem solved. Analyzer handed over to staff in well working condition. | SWO-261213 | Cleaned | Fluidic pathway | pH/BG Module | Measuring chamber | Clot | Service Visit | Clotting appeared in pH chamber, leakage symptoms also appeared from pH chamber. | Open the chamber and remove the electrode from the chamber, remove the cap and O-ring from the electrode and dip in to NaOCl, after 5 min take away from solution and wash with the water and wipe out with gauze. Now clean the pH chamber. Remove the inlet gas kit, inlet probe tubing which is connect to probe and pH $BG module. All these parts flush with the NaOCl solution and wash with the water, wipe out with gauze and reconnect all the parts same as before. Now restart the analyzer, wait for auto calibration. Do protein removal and cleaning command to the analyzer and then give 2 point calibration. Now tell the user for sample, user do the sample, no query message appearing in pH during result. Measurement unstable problem solved. Analyzer handed over to staff in well working condition.
No | Visit | Complaint | RIND | Ph and Ca++ ? for all calibration | Ph and ca++ sensitivy vallues too low | Observed that Clott was there in ph and ca++ chamber cleaned the chamber and protein removal tested then calibrated to raedy mode | SWO-260775 | Cleaned | Fluidic pathway | Fluidic pathway | Measuring chamber | Clot | Ph and Ca++ ? for all calibration | Ph and ca++ sensitivy vallues too low | Observed that Clott was there in ph and ca++ chamber cleaned the chamber and protein removal tested then calibrated to raedy mode
No | Visit | Complaint | RAME | Printer wrinkles left side of paper when printing. | Cleaned printer still had the same problem. | Replaced printer head. | SWO-261062 | Defect | HW | Printer Module | Printer head | Paper wrinkles | Printer wrinkles left side of paper when printing. | Cleaned printer still had the same problem. | Replaced printer head.
No | Visit | Complaint | RAME | Oxi HW error 0579 & 316 Neon number too low. | Installation procedure revealed level 3 VK-R7 lot 490 is failing Oxi parameters and voting out of range on SO2 & negative -49.5 FO2Hb. Also getting neon voltage too low prior to Oxi PCB & lamp unit replacement.after Oxi PCB $lamp got a neon position error, replaced second lamp unit and optic fiber cable. | Replaced Oxi PCB, hemolyzer, lamp unit twice, & optic fiber cable. | SWO-260079 | Defect | HW | Oxi module | Lamp unit | 0316 | Oxi HW error 0579 & 316 Neon number too low. | Installation procedure revealed level 3 VK-R7 lot 490 is failing Oxi parameters and voting out of range on SO2 & negative -49.5 FO2Hb. Also getting neon voltage too low prior to Oxi PCB & lamp unit replacement.after Oxi PCB /lamp got a neon position error, replaced second lamp unit and optic fiber cable. | Replaced Oxi PCB, hemolyzer, lamp unit twice, & optic fiber cable.
Yes | Visit | Complaint | RAME | Installation procedure revealed level 3 VK-R7 lot 490 is failing Oxi parameters and voting out of range on SO2 & negative -49.5 FO2Hb. | Installation procedure revealed level 3 VK-R7 lot 490 is failing Oxi parameters and voting out of range on SO2 & negative -49.5 FO2Hb. | Replaced hemolyzer & ran another VK-R7 kit. | SWO-260063 | Re classified RMED | doc in OOB | Oxi module | Hemolyzer | 0581 | Installation procedure revealed level 3 VK-R7 lot 490 is failing Oxi parameters and voting out of range on SO2 & negative -49.5 FO2Hb. | Installation procedure revealed level 3 VK-R7 lot 490 is failing Oxi parameters and voting out of range on SO2 & negative -49.5 FO2Hb. | Replaced hemolyzer & ran another VK-R7 kit.
No | Visit | Complaint | RDEU | 08.11.2017; Sr. Anne-Marie; Tel.: 06044615906; CRP QC Messfehler, Messwerte zu hoch | LQC Fehler CRP Level2 Lot 12696 (Ergebnis zu hoch, grenzwertig) | Analysator geöffnet. Reinigungsmaßnahmen durchgeführt. Reinigung P1-P2. Nadelwashblock und Nadel gereinigt. Nadel kalibriert. Reinigung der CupwashUnit und des Processwheel Initialisierungsphase, Aufwärmphase LQC im Level 1&2 wiederholt. QC nahe dem Zielwert | SWO-260510 | Cleaned | HW | Aspiration Module | Needle Wash Unit |  | 08.11.2017; Sr. Anne-Marie; Tel.: 06044615906; CRP QC measurement error, readings too high | LQC Error CRP Level2 Lot 12696 (Result too high, borderline) | Analyzer open. cleaning measures. Cleaning P1-P2. Needle wash block and needle cleaned. Needle calibrated. Cleaning the CupwashUnit and the Processwheel Initialization phase, warm-up phase LQC in Level 1&2. QC close to target
No | Visit | Complaint | RAME | Intermittently will not print. | Intermittently will not print. | Replaced printer. | SWO-261259 | Defect | HW | Printer Module | Printer Module |  | Intermittently will not print. | Intermittently will not print. | Replaced printer.
Yes | Visit | Complaint | RAME | pO2 QC failures | Inspected instrument. Reviewed logs. Found vacuum test failures. | Customer needs to replace sensor cassette. Frank to follow-up Monday. | SWO-261193 | Re classified RMED | Remote support |  |  |  | pO2 QC failures | Inspected instrument. Reviewed logs. Found vacuum test failures. | Customer needs to replace sensor cassette. Frank to follow-up Monday.
No | Visit | Complaint | RLTD | Whilst on site David Ruaux has reported an issue with the battery not holding charge. No contact on site to inform whilst he was there so no details as to how long this has been happening. Contact is Kim Padfield, RF, please contact when back from holiday. | Battery icon present, but showing 0%. Battery status in service also showing 0 charge, but all 6 cells also showing as above 4V. Turned off mains power and analyser ran on battery for 30 minutes to discharge cells. Then turned analyser off and removed battery unit and reset charger pcb by removing strap on pcb for 10 seconds. Refitted battery and tested operation. Battery display now showing correct percentage of charge in the battery. Tested battery charging ok, tested analyser still able to operate on battery, also ok. Battery Unit now working correctly. Verified analyser operation - all ok. | Reset battery charger pcb. | SWO-261071 | Defect | SW | RMED SW | Battery module - Firmware |  | Whilst on site David Ruaux has reported an issue with the battery not holding charge. No contact on site to inform whilst he was there so no details as to how long this has been happening. Contact is Kim Padfield, RF, please contact when back from holiday. | Battery icon present, but showing 0%. Battery status in service also showing 0 charge, but all 6 cells also showing as above 4V. Turned off mains power and analyser ran on battery for 30 minutes to discharge cells. Then turned analyser off and removed battery unit and reset charger pcb by removing strap on pcb for 10 seconds. Refitted battery and tested operation. Battery display now showing correct percentage of charge in the battery. Tested battery charging ok, tested analyser still able to operate on battery, also ok. Battery Unit now working correctly. Verified analyser operation - all ok. | Reset battery charger pcb.
No | 4 | Y | RCZE | Blank | 2106 - mozne ucpani behem cyklu systemu |  | D36AF347FA4A4F8E | [5 empty] | Blank | 2106-Mozne Ucpani Behem system cycle
No | Visit | Complaint | RESP | No inicializa SP | no inicializa SP | Se detectó error de calibración de bomba y aguja doblada. se colocó aguja correctamente y se calibró bomba. Tras esto, se inicializa SP y se observa fallo en varios parámetros. se procedió a sustituir SC. se calibró equipo y se compruebó correcto funcionamiento mediante QC,s. Equipo operativo. | SWO-261177 | Adjustment | HW | Auto Inlet module | Inlet Position Cal. |  | Sp does not initialize | doesn't initialize SP | Pump and bent needle calibration error was detected. needle was placed correctly and bomb was calibrated. After this, SP is initialized and several parameters fail. SC was replaced. equipment was graded and proper operation was verified by QC,s. Operational Equipment.
No | Visit | Complaint | RCHN | 测试卡扫描器不能扫描，测试卡槽拖落 | 测试卡扫描器不能扫描，测试卡槽拖落，借886-625三个（广办发货） | 更换3个886-625后修复 by jesson | SWO-259565 | Defect | HW | Storage Module | Cassette guide | Pizza guide | Test card scanner cannot be scanned, test card slot is dragged down | Test card scanner can not be scanned, test card slot drag, borrow 886-625 three (wide delivery) | Repair by jesson after replacing 3 886-625
No | Visit | Complaint | RCHN | 试管夹断裂 | 试管夹断裂，借903-584一个 | 更换903-584后修复 by jesson | SWO-260122 | Defect | HW | Cover and Chassis | Tube holder |  | Test tube clip break | Test tube clip broken, borrow 903-584 one | Repair by jesson after replacing 903-584
No | Visit | Complaint | RCHN | 试管夹坏了 | 试管夹故障，借903-584一个 | 更换903-584后修复 by jesson | SWO-260203 | Defect | HW | Cover and Chassis | Tube holder |  | The test tube clip is broken. | Test tube clip failure, borrow 903-584 one | Repair by jesson after replacing 903-584
No | Visit | Complaint | RLTD | Please call on Hannah McPeak on 01865 221 596 01865 741 166 ask switch bleep 1624 They have tried to clean it but can't get to all of the blood and the tube keeps popping off and it is causing further issues. |  |  | SWO-261415 | Cleaned | Fluidic pathway | Fluidic pathway | Clot, Fluidic Pathway |  | Please call on Hannah McPeak on 01865 221 596 01865 741 166 ask switch bleep 1624 They have tried to clean it but can't get to all of the blood and the tube keeps popping off and it is causing further issues.
Yes | Visit | Complaint | RAME | Installation procedure revealed level 3 VK-R7 lot 490 is failing Oxi parameters and voting out of range on SO2 & negative -49.5 FO2Hb. OOBF- new installation. VK-R7 lev 3 Oxi failure. | Installation procedure revealed level 3 VK-R7 lot 490 is failing Oxi parameters and voting out of range on SO2 & negative -49.5 FO2Hb. | Replaced Hemolyzer & Spectrometer units. | SWO-260685 | Re classified RMED | doc in OOB | Oxi module | Hemolyzer | 0581 | Installation procedure revealed level 3 VK-R7 lot 490 is failing Oxi parameters and voting out of range on SO2 & negative -49.5 FO2Hb. OOBF- new installation. VK-R7 lev 3 Oxi failure. | Installation procedure revealed level 3 VK-R7 lot 490 is failing Oxi parameters and voting out of range on SO2 & negative -49.5 FO2Hb. | Replaced Hemolyzer & Spectrometer units.
No | Visit | Complaint | RAME | Lactate calibration and QC out of control multiple SC &SP changes have not made any difference | Lactate calibration errors began on 10-27. The errors were intermittent but multiple SC & SP changes have occurred to rectify until 11-9 consumable changes did not help. | Changed sensor interface PCB and SC. Instrument calibrated and ran QC instrument, operating normally . | SWO-261272 | Defect | HW | Basic Unit | PCB, Sensor interface |  | Lactate calibration and QC out of control multiple SC &SP changes have not made any difference | Lactate calibration errors began on 10-27. The errors were intermittent but multiple SC & SP changes have occurred to rectify until 11-9 consumable changes did not help. | Changed sensor interface PCB and SC. Instrument calibrated and ran QC instrument, operating normally .
No | 4 | Y | RPAC | Blank | Inlet positioning error |  | D36CBC12704D54AA | Adjustment | HW | Auto Inlet module | Inlet Position Cal. |  | Blank | Inlet positioning error
No | Visit | Complaint | RCHN | 舱门间接性的无法关上 | 舱门间接性的无法关上，需要借件924-609 | 924-609 断裂，更换924-609备件后修复 by jesson | SWO-260890 | Defect | HW | Basic Unit | Sensor Door |  | The indirectness of the hatch cannot be closed. | The door is not closed indirectly and requires a loan of 924-609 | 924-609 Break, repair by jesson after replacing 924-609 spare parts
No | Visit | Complaint | RCHN | 报错2167 | 报错2167，需要借件886-621 | 由于此故障偶发性出现，且固定电机的两个螺纹已经滑丝，不能正常的将电机固定，先使用备用机备件修复。需要日后继续换观察。 by jesson | SWO-261309 | Defect | HW | Storage Module | Bracket, Punch Out Motor | 2167 | 2167 error | 2167, 886-621 borrowed | Due to the occasional occurrence of this fault, and the two threads of the fixed motor have slipped wire, the motor can not be properly fixed, first using spare parts to repair. It is necessary to continue to change observation at a later date. by jesson
No | Visit | Complaint | RIND | Touch Screen Problem. | Touch Screen problem, | Check the analyzer upgraded firmware for touch screen. touch screen was working, will be under Observation. | SWO-261357 | Defect | SW | DMS_PC | Touch screen | Firmware upgrade | Touch Screen Problem. | Touch Screen problem, | Check the analyzer upgraded firmware for touch screen. touch screen was working, will be under Observation.
No | Visit | Complaint | RIND | Rotatesotragewheel. | Rotatesotragewheel. | Found Storage wheel error, Cleaned the Analyzer with vacuum cleaner, removed the cartridge broken piece, restarted analyzer and performed the system clean, performed the test found results are OK. found instrument is working satisfactorily. | SWO-261356 | Defect | Consumables | Test Kit | Cartridge, Broken | - | Rotatesotragewheel. | Rotatesotragewheel. | Found Storage wheel error, Cleaned the Analyzer with vacuum cleaner, removed the cartridge broken piece, restarted analyzer and performed the system clean, performed the test found results are OK. found instrument is working satisfactorily.
No | Visit | Complaint | RDEU | ABL meldet Intervention erforderlich. LP/ SC und ED wurden mehrfach erneuert. | Über entstehende Kosten hingewiesen. Freigabe durch OÄ Fr. Hochrein. Analysatorüberprüfung. Quetschventilschlauch undicht. | Ventilschlauch erneuert. Funktionskontrolle iO. | SWO-261409 | Defect | HW | Oxi module | Pinch valve tube |  | ABL reports intervention required. LP/ SC and ED have been renewed several times. | Informed about any costs incurred. Release by OÄ Fr. Hochrein. Analyzer check. Squeezing valve hose leaks. | Valve hose renewed. Functional control iO.
No | Visit | Complaint | RSAF | Telephonic assistance requested from Ishmael to try and correct the calibration and Qc issues on pCo2 and Chloride, assisted Ishmael but could not correct the issue. | Done trouble shooting according to the manual and could not find any issues on analyzer which indicated error to Sensor Cassette | Had to replace sensor cassette as the analyzer showed no cause for reported error, Customer had no spare Cassette and I installed the one which I had with me a previous call but never used. | SWO-261594 | Assisted Replacement | SC | SC | pCO2 and Cl |  | Telephonic assistance requested from Ishmael to try and correct the calibration and Qc issues on pCo2 and Chloride, assisted Ishmael but could not correct the issue. | Done trouble shooting according to the manual and could not find any issues on analyzer which indicated error to Sensor Cassette | Had to replace sensor cassette as the analyzer showed no cause for reported error, Customer had no spare Cassette and I installed the one which I had with me a previous call but never used.
No | Visit | Complaint | RCHN | 测试卡漏液导致机器无法启动 | 测试卡漏液导致机器无法启动，需要借件903-183一个 |  | SWO-261379 | Defect | SC | SC | Leakage | Sensor IF replaced, crystals | Test card leak prevents machine from starting | Test card leakage causes the machine to fail to start, requires a loan 903-183
No | Visit | Complaint | RDEU | Hotline- Pumpenkal Fehler | Gerät überprüft und gereinigt, Gerät manuell gespült, Lösungspack erneuert, | Einlassdichtung und Einlasssonde erneuert, Gerät neu gestartet, Kaltstartphase abgewartet, Kal und QCs durchgeführt | SWO-261751 | Defect | HW | Inlet module | Inlet gasket |  | Hotline pump scale error | device checked and cleaned, device manually rinsed, solution pack renewed, | Inlet seal and inlet probe renewed, device restarted, cold start phase awaited, Kal and QCs performed
No | Visit | Complaint | RDEU | 08.11.2017 10:45 Fr. Martens 04817852600: Barcodescanner ohne Funktion | Interner Scanner defekt. | Neuen Scanner eingebaut. Nach Test ist dieser auch ohne Funktion. Danke. Kunde darüber informiert. | SWO-261706 | Defect | HW | Basic Unit | Barcode reader, Basic Unit |  | 08.11.2017 10:45 Fr. Martens 04817852600: Barcode scanner without function | Internal scanner defective. | New scanner installed. After testing, it is also without function. Thank you. customer informed.
No | 4 | Y | RPOL | Blank | Czyszczenie czujnika selektora przep³ywów, aktualizacja oprogramowania 3.3MR1. Analizator sprawny |  | D36EB6B85C52BAF5 | Assisted Replacement | SP | SP | Flow Selector |  | Blank | Cleaning the overflowing selector sensor, software update 3.3 MR1. Functional analyzer
No | Visit | Complaint | RLTD | Cameifha Oyabyo on 0203 2991895 | Inlet was stuck open in syringe position and analyser was in Inlet gasket replacement screen. Stepped through inlet gasket replacement without replacing gasket - inlet closed and analyser restarted. After restart all ok except for Lactate giving cal response error 1310. Checked data logs to find it had been failing like this before inlet was stuck open. Customer to replace Sensor Cassette themselves. | Close inlet, by completing inlet gasket replacement without replacing. Recommended Sensor Casette replaced to resolve Lactate error. | SWO-263770 | Adjustment | HW | Auto Inlet module | Inlet Position Cal. |  | Cameifha Oyabyo on 0203 2991895 | Inlet was stuck open in syringe position and analyser was in Inlet gasket replacement screen. Stepped through inlet gasket replacement without replacing gasket - inlet closed and analyser restarted. After restart all ok except for Lactate giving cal response error 1310. Checked data logs to find it had been failing like this before inlet was stuck open. Customer to replace Sensor Cassette themselves. | Close inlet, by completing inlet gasket replacement without replacing. Recommended Sensor Casette replaced to resolve Lactate error.
No | Visit | Complaint | RCHN | 主机屏幕花屏 | 主机屏幕花屏 | 更换备用机后修复 | SWO-261614 | [5 empty] | Host screen screen | Host screen screen | Repair after replacing the spare machine
Yes | Visit | Complaint | RCHN | 主机屏幕花屏 | 主机屏幕花屏 |  | SWO-261617 | Re classified RMED | Double entry |  |  |  | Host screen screen | Host screen screen
No | Visit | Complaint | RCHN | 血氧单元硬件故障，报错代码：1030 | 血氧单元硬件故障，报错代码：1030，需要902-980一个 |  | SWO-261510 | Defect | HW | Oxi module | Hemolyzer | 1030 | Blood oxygen unit hardware failure, error code: 1030 | Blood oxygen unit hardware failure, error code: 1030, need 902-980 one
No | Visit | Complaint | RCHN | 机器无法启动，po2无法正常定标，报错代码：0581和1361 | 机器无法启动，po2无法正常定标，报错代码：0581和1361，需要借件902-980和903-183各一个 | 更换902-980及备用的备件903-183后修复by tim | SWO-261319 | Defect | HW | Oxi module | Hemolyzer |  | The machine cannot start, po2 cannot be properly calibrated, error codes: 0581 and 1361 | The machine cannot start, po2 can not be properly calibrated, error code: 0581 and 1361, need to borrow 902-980 and 903-183 each one | Repair by tim after replacement of 902-980 and spare parts 903-183
No | Visit | Complaint | RIND | Service Visit | Need to change the printer head. | Shut down the analyzer. Remove the safety back covers, remove the screw which is interconnect, take out from the chamber and remove the printer head. Connect the new printer head and reconnect all the parts same as before. Switch on the analyzer and wait for auto 2 point calibration after initializing and start up process. Calibration is successful and then connect the paper in to the printer, tested it successfully. Analyzer handed over to staff in well working condition. | SWO-261923 | Defect | HW | Printer Module | Printer head | 0 | Service Visit | Need to change the printer head. | Shut down the analyzer. Remove the safety back covers, remove the screw which is interconnect, take out from the chamber and remove the printer head. Connect the new printer head and reconnect all the parts same as before. Switch on the analyzer and wait for auto 2 point calibration after initializing and start up process. Calibration is successful and then connect the paper in to the printer, tested it successfully. Analyzer handed over to staff in well working condition.
No | Visit | Complaint | RLTD | Intermittent touch screen freezes Ian Malin | Touch screen freezing | Install Citron touch screen firmware 4.00.39 Clear patient profile log. | SWO-263629 | Defect | HW | Touch Screen | Firmware Update | 0 | Intermittent touch screen freezes Ian Malin | Touch screen freezing | Install Citron touch screen firmware 4.00.39 Clear patient profile log.
No | Visit | Complaint | RKK | サンプル移動エラー。ｐH測定室にあったクロットが原因。測定室の清掃で復旧。ガスケット交換、エアライン洗浄、タンパク除去実施。 | サンプル移動エラー。ｐH測定室にあったクロットが原因。 | 測定室の清掃で復旧。ガスケット交換、エアライン洗浄、タンパク除去実施。 | SWO-261684 | Cleaned | Fluidic pathway | pH/BG Module | Clot, Measuring Chamber | 0 | Sample move error. The clot in the pH measurement room is the cause. The measurement room was cleaned to restore it. Gasket replacement, airline cleaning, and protein removal. | Sample move error. The clot in the pH measurement room is the cause. | The measurement room was cleaned to restore it. Gasket replacement, airline cleaning, and protein removal.
No | Visit | Complaint | RDAN | Karussel er svær at sætte på plads | NA | Motor til karrusel renset | SWO-261671 | Cleaned | HW | Auto Check Module | Carousel engine | 0 | Carousel is hard to put in place | In | Motor for Carousel Cleaned
No | Visit | Complaint | RAME | Analyzer aborting QC runs because of AutoCheck mechanical error. | Ran AutoCheck in service mode. Carousel not always moving to zero position when house is up. | Replaced FPC cable in AutoCheck assembly. Also replaced tubing nipple kit for small crack detected. | SWO-262128 | Defect | HW | Autocheck Module | FPC cable | 0 | Analyzer aborting QC runs because of AutoCheck mechanical error. | Ran AutoCheck in service mode. Carousel not always moving to zero position when house is up. | Replaced FPC cable in AutoCheck assembly. Also replaced tubing nipple kit for small crack detected.
No | Visit | Complaint | RIND | Lactate Calibration error | found problem with lactate membrane. | replaced the lactate membrane and given demo for membrane changing to users. did calibration and samples found instrument working satisfactorily. | SWO-262603 | Other | Electrode | Membrane | Lac | Training | Lactate Calibration error | found problem with lactate membrane. | replaced the lactate membrane and given demo for membrane changing to users. did calibration and samples found instrument working satisfactorily.
No | Visit | Complaint | RIND | Calibration Error | found problem with reference, glucose and lactate membranes. found problem with pH $BG pump error also | replaced the pump tubing for pH $BG module and did calibration for pump.replaced the reference, glucose and lactate membranes and did 2 point calibration and samples found instrument working satisfactorily. | SWO-262645 | Defect | HW | pH/BG Module | Pump tubes | half year service | Calibration Error | found problem with reference, glucose and lactate membranes. found problem with pH /BG pump error also | replaced the pump tubing for pH $BG module and did calibration for pump.replaced the reference, glucose and lactate membranes and did 2 point calibration and samples found instrument working satisfactorily.
No | Visit | Complaint | RAME | Analyzer will not boot up. | Found BIOS would not start up. | Replaced the CPU unit. | SWO-262733 | Defect | HW | Basic Unit | DMS_PC | CPU to return for investigation | Analyzer will not boot up. | Found BIOS would not start up. | Replaced the CPU unit.
No | Visit | Complaint | RLTD | pO2 failing calibration with status errors frequently. | Follow up to previous visit on 12 $10 when analyser was decontaminated. Analysed data logs since visit and concluded that the pO2 status errors were still occurring. Visited all analysers on site to gather data. | Issue not resolved, now with RMED. | SWO-265169 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway/Decon | pO2 status | pO2 failing calibration with status errors frequently. | Follow up to previous visit on 12 /10 when analyser was decontaminated. Analysed data logs since visit and concluded that the pO2 status errors were still occurring. Visited all analysers on site to gather data. | Issue not resolved, now with RMED.
No | Visit | Complaint | RDEU | 08.11.2017 10:45 Fr. Martens 04817852600: Barcodescanner ohne Funktion | Neuer Scanner von gestern defekt | Scanner aus neuen Gerät eingebaut , Kalibration und QC durchgeführt | SWO-262573 | Defect | HW | Basic Unit | Barcode reader, Basic Unit |  | 08.11.2017 10:45 Fr. Martens 04817852600: Barcode scanner without function | New scanner from yesterday defective | Scanner built from new device, calibration and QC performed
No | Visit | Complaint | RAME | Internal reference electrode error. | Internal reference electrode error. Bad Sensor cassette. | Replaced Ensor cassette. Completed all service tests. Customer requested automatic patient results printing, added this function. Sensor had 82 tests left on it Credit done onsite with replacement cassette with 100 tests. | SWO-262982 | Assisted Replacement | SC | SC | Reference |  | Internal reference electrode error. | Internal reference electrode error. Bad Sensor cassette. | Replaced Ensor cassette. Completed all service tests. Customer requested automatic patient results printing, added this function. Sensor had 82 tests left on it Credit done onsite with replacement cassette with 100 tests.
No | Visit | Complaint | RKK | SPドアラッチ破損の為部品交換。SP交換。QC確認。 | SPドアラッチ破損。 | SPドアラッチ破損の為部品交換。SP交換。QC確認。 | SWO-263472 | [5 empty] | Parts replacement for SP door latch breakage. SP replacement. QC confirmation. | SP door latch breakage. | Parts replacement for SP door latch breakage. SP replacement. QC confirmation.
No | Visit | Complaint | RCHN | 打印机故障 | 打印机故障，借备件903-001一个 | 更换新打印机模块 | SWO-262255 | [5 empty] | Printer failure | Printer failure, spare parts 903-001 one | Replace the new printer module
No | Visit | Complaint | RAME | All oxi outside measuring range on sample runs | Hemolyzer hardware status = Cavitation absorb failed. | Replaced hemolyzer. Cavitation absorb passed. | SWO-262999 | [5 empty] | ALL oxygen outside measuring range on sample runs | Hemolyzer hardware status = Cavitation absorb failed. | Replaced hemolyzer. Cavitation absorb passed.
No | Visit | Complaint | RCHN | 报错3412 | 报错3412 | 更换备用件903-112后修复by tim | SWO-263197 | Defect | HW | Process Module | PCB, Wet section | 3412 | Error 3412 | Error 3412 | Repair by tim after replacing spare parts 903-112
No | Visit | Complaint | RIND | Oxy parameters not coming. | Oxy parameters no calibrated. | Found drift in parameters. Protean removal done. Cleaning Done. Drain and swithoff for some time. Re-calibrated. Analyser Serviced and calibrated properly. Handed over in a fine working condition. | SWO-265001 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic Pathway | 0 | Oxy parameters not coming. | Oxy parameters no calibrated. | Found drift in parameters. Protean removal done. Cleaning Done. Drain and swithoff for some time. Re-calibrated. Analyser Serviced and calibrated properly. Handed over in a fine working condition.
No | Visit | Complaint | RAME | Customer reported inlet nipple broke off when changing inlet gasket. | Plastic nipple broken under inlet gasket. | Replaced nipple kit and waste tubing connected to it. Restart calibrations. | SWO-262966 | Defect | HW | Inlet module | Nipple kit | 0 | Customer reported inlet nipple broke off when changing inlet gasket. | Plastic nipple broken under inlet gasket. | Replaced nipple kit and waste tubing connected to it. Restart calibrations.
No | Visit | Complaint | RDEU | Gerät bricht die Kal ab. Einlasssonde ist verbogen. | Einlasssonde defekt, Dichtungsring an der Lac Elektrode fehlt. Gas Unit oder Druckaufnehmer defekt. | Einlasssonde erneuert, Membranventile demontiert und gereinigt. Defekte Flat Membran am pH $BG Modul - Air Membranventil erneuert, Fehlender unterer Dichtungsring an der Lac Elektrode ersetzt. Drucktransducer an der Gas Unit zerlegt und gereinigt. Ersatzteil wird bestellt | SWO-263983 | Defect | HW | Inlet module | bent probe | 0 | device breaks off the cal. Inlet probe is bent. | Inlet probe defective, sealing ring on the Lac electrode is missing. gas unit or pressure transducer defective. | Inlet probe renewed, diaphragm valves dismantled and cleaned. Defective flat membrane on pH $BG module - Air diaphragm valve renewed, missing lower sealing ring replaced at the Lac electrode. Pressure transducer at the gas unit disassembled and cleaned. Spare part is ordered
No | Visit | Complaint | RAME | El/Met upper liquid sensor error. | Opened the compartment lid and found module tubing not installed properly. | Replaced the module tubing. | SWO-263066 | Defect | HW | El/Met Module | Tubings | 0 | El/Met upper liquid sensor error. | Opened the compartment lid and found module tubing not installed properly. | Replaced the module tubing.
No | Visit | Complaint | RAME | THb cal not within range | Arrived onsite and found that the inlet gasket would not calibrate. When replacing, found small leak. | Replaced inlet gasket. | SWO-263091 | Defect | HW | Inlet module | Inlet gasket | 0 | THb cal not within range | Arrived onsite and found that the inlet gasket would not calibrate. When replacing, found small leak. | Replaced inlet gasket.
No | Visit | Complaint | RDEU | 06.11.2017, 09:11 Uhr, Fr. Eilers-Gladrow, Tel. 05315952045: Gerät reagiwert auf keine Eingaben mehr | das Gerät bleibt im Netz ein stehen, die Sicherung F2 auf der AU PCB defekt | Gerät in die Werkstatt abgezogen und an die Zentrale Krefeld versendet. | SWO-262752 | Defect | HW | Basic Unit | AU Module | F2 fuse | 06.11.2017, 09:11, Ms. Eilers-Gladrow, tel. 05315952045: Device reagiwert on no more inputs | the device stops in the mains, the fuse F2 on the AU PCB is defective | device is removed to the workshop and sent to the headquarters in Krefeld.
No | Visit | Complaint | RAME | Customer stuck at user intervention screen with leak detected error | User intervention leak detected | Replaced leak detect sensor and motor assembly. | SWO-263823 | Defect | HW | Basic Unit | LS |  | Customer stuck at user intervention screen with leak detected error | User intervention leak detected | Replaced leak detect sensor and motor assembly.
No | Visit | Complaint | RAME | Flow Issues | Customer has two SP that have chip could not be read errors. No more consumables on site. | Replaced two SP for troubleshooting. Checked the SP Id PCB and checked for any flow issues. SC failing PCO2 with sensitivity errors due to no fluids run overnight. Performed several calibrations and internal QC to hydrate the SC but could not obtain a sensitivity over 50%. Replaced the SC and tested with passing results. | SWO-263752 | Assisted Replacement | SP | SP | Chip |  | Flow Issues | Customer has two SP that have chip could not be read errors. No more consumables on site. | Replaced two SP for troubleshooting. Checked the SP Id PCB and checked for any flow issues. SC failing PCO2 with sensitivity errors due to no fluids run overnight. Performed several calibrations and internal QC to hydrate the SC but could not obtain a sensitivity over 50%. Replaced the SC and tested with passing results.
No | Visit | Complaint | RAME | Po2 failed after installing a new sensor | Analyzer removed | Analyzer removed and replaced by SN 090R0952N024 | SWO-263653 | Revisit required | Revisit required | Revisit required | Revisit required |  | Po2 failed after installing a new sensor | Analyzer removed | Analyzer removed and replaced by SN 090R0952N024
No | Visit | Complaint | RSCH | Blank | Consommation importante de solution SP et calibration fréquente selon M. Bardy | Contrôle du système fluidique e.o Vérification de la configuration système ( nombre d’analyse + contrôle cycle ) e.o Calibration e.o CQ e.o | SWO-263559 | User error | Calibration schedule | Calibration schedule | Calibration schedule |  | Blank | High consumption of SP solution and frequent calibration according to Bardy | Fluid system control e.o Checking the system configuration (number of analysis - cycle control) e.o Calibration e.o CQ e.o
No | Visit | Complaint | RSCH | Reparatur pO2 | Fehler 1002: pO2 Dunkeldaten ausser Bereich | Ersatzgerät vor Ort installiert i.o IT Radiometer: Ersatzgerät im Aqure eingetragen i.o Test-Messung durchgeführt i.o IT USZ: Test-Messung überprüft i.o System i.o Kunden-Gerät wird in Thalwil repariert i.o Sensor Interface PCB ersetzt i.o | SWO-263667 | Defect | HW | Basic Unit | pO2 dark data | Loaner installed | Repair pO2 | Error 1002: pO2 Dark data off range | Replacement device installed on site i.o IT Radiometer: Replacement device entered in Aqure i.o Test measurement performed i.o IT USZ: Test measurement checked i.o System i.o Customer device is repaired in Thalwil i.o Sensor Interface PCB replaces i.o
No | Visit | Complaint | RDEU | 13.11.2017, Herr Fürst, Tel: 065194783789: ständige Interventionsmeldungen | Kapillar Messung wird fast immer abgebrochen. Einlassdichtung hängt> Modul ersetzt. | Startphase QC und Kalibration | SWO-263436 | Defect | HW | Auto Inlet module | Cap - Spring | Barnevogns fjeder | 13.11.2017, Mr Fürst, Tel: 065194783789: constant intervention reports | Capillary measurement is almost always aborted. Inlet seal hangs> Module replaced. | Start phase QC and calibration
No | Visit | Complaint | RDEU | Nach Software Update moniert der Kunde das bei Kapillarenmessungen das erwartete pO2 zu niedrig ist. | pO2 Patientenmessungen zu niedrig | Daten gesichert für Rückmeldung Technischer Probleme | SWO-263402 | [5 empty] | After software update, the customer complains that the expected pO2 is too low for capillary measurements. | pO2 patient measurements too low | Data backed up for feedback of technical problems
No | Visit | Complaint | RFRA | Paramètre pco2 et hct bloqués. | QC pco2 et hct hors normes. | Passage qc hct ok. Passage qc pco2 ok Paramètres gds et hct ok. Activation Windows suite message erreur. | SWO-263672 | [5 empty] | Blocked pco2 and hct setting. | QC pco2 and hct out of norms. | Passage qc hct ok. Passage qc pco2 ok Paramètres gds et hct ok. Activation Windows suite message erreur.
No | Visit | Complaint | RDEU | Gerät bricht die Kal ab. Einlasssonde ist verbogen. | Kein Fehler mehr vorhanden. | Gerät ohne Fehler vorgefunden. Feuchtigkeit in der Gas Uni ist vermutlich diffundiert. Einbau einer neuen Gas Unit ist nicht mehr erforderlich. | SWO-263985 | Revisit | HW | Inlet module | bent probe | Revisit | device breaks off the cal. Inlet probe is bent. | No more errors. | Device found without errors. Moisture in the gas uni is probably diffused. Installation of a new gas unit is no longer required.
No | Visit | Complaint | RAME | Ca not passing on levels 2 and 3. | Arrived onsite and could not find any problems with the fluid path. Replaced the Ca electrode and ran calibration. QC all passed. Will send electrode in under warranty replacement. | Replaced Ca electrode. | SWO-263170 | Assisted Replacement | Electrode | Electrode | Ca | 0 | Ca not passing on levels 2 and 3. | Arrived onsite and could not find any problems with the fluid path. Replaced the Ca electrode and ran calibration. QC all passed. Will send electrode in under warranty replacement. | Replaced Ca electrode.
No | Visit | Complaint | RDEU | Sehr häufig Na, Ca und Cl Kalibrationsdriften | Austausch pH $BG Modul Stark schwankendes QC Ergebnis des pH und Lyte | Altes Modul (902-823R2697N005) ausgebaut. Neues Modul eingebaut. Initialisierungsphase AU Software neu installiert. Initialisierungsphase LS Justage und Pumpen Kal. Kalibrationen mehrfach durchgeführt. QC ohne Beanstandung, Verlauf wird in den nächsten Tagen kontrolliert. Garantieaustausch | SWO-263389 | Revisit required | Electrode | Electrode | Reference | 0 | Very common Na, Ca and Cl calibration drifts | Replacement of pH /BG module Highly fluctuating QC result of pH and lyte | Old module (902-823R2697N005). New module built-in. Initialization phase AU software reinstalled. Initialization phase LS Adjustment and Pumps Kal. calibrations performed several times. QC without complaint, course will be checked in the next few days. Warranty exchange
Yes | Visit | Complaint | RDEU | Fehler FlexQ / wird nicht verwendet | Ausbau FlexQ, defekt und nicht verwendet Mischerträgerschlupf erkannt | FlexQ Modul (R1350N005) ausgebaut und gereinigt. Abdeckung ersetzt. Rücksendung des Moduls in Zentrale (Garantie) | SWO-263388 | Re classified RMED | HW | FlexQ Module | FlexQ Module | 0 | Error FlexQ / is not used | Removal FlexQ, defective and not used Mixer carrier slip detected | FlexQ module (R1350N005) removed and cleaned. Cover replaced. Return of the module to headquarters (guarantee)
No | 2 | Y | RPOL | Blank | Wymiana pompy œciekowej wraz z ³¹cznikiem. Czyszczenie aparatu. Analizator sprawny. |  | D371504E483DBAFD | [5 empty] | Blank | Replacing the liquid pump and the connector. Cleaning the camera. Functional Analyzer.
No | Visit | Complaint | RDEU | 16.11.2017, Sr. Ulrike, Tel: 015222815149: Probenaspiration gescheitert | Probenaspiration scheitert öfters | Einlassdichtung erneuert, Anschlussdichtung, Abfallschlauch sowie oxi Schlauch erneuert, pinch Ventil gereinigt, systemüberprüfung durchgeführt i. O. , Kalibrierung und Qualitätskontrolle i. O. Mehrere Testmessungen durchgeführt i. O. | SWO-264078 | Defect | HW | PM Kit | PM Kit |  | 16.11.2017, Sr. Ulrike, Tel: 015222815149: Rehearsal aspiration failed | Sample aspiration often fails | Inlet seal renewed, connection seal, waste hose as well as oxi hose renewed, pinch valve cleaned, system inspection carried out in the following, calibration and quality control as the standard Several test measurements are carried out in the first place
No | Visit | Complaint | RLTD | Annette 03003001761 | Intervention required failing pressure test. Unit switched off. | Switched unit on started up ok. Flushed system and run service programs passed ok. Suspect clot causing error. Checked battery status ok. Run calibration and THb cal as due passed ok. QC passed ok. | SWO-264203 | Cleaned | Fluidic pathway | Fluidic pathway | Clot, Fluidic Pathway |  | Annette 03003001761 | Intervention required failing pressure test. Unit switched off. | Switched unit on started up ok. Flushed system and run service programs passed ok. Suspect clot causing error. Checked battery status ok. Run calibration and THb cal as due passed ok. QC passed ok.
Yes | Visit | Complaint | RLTD | Asking to be flushed all the time |  |  | SWO-264304 | Re classified RMED |  |  |  |  | Asking to be flushed all the time
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | K定标不过 | 更换测试卡后修复 | SWO-264625 | [5 empty] | Test card, reagent package failure | K calibration but | Repair after replacing test card
No | Visit | Complaint | RBNL | Afgekeurde controles | Verbindingsconnector tussen Cassette en analyzer ziet er smerig uit. pH te laag en pCO2 te hoog voor QC level1 Vervanging van Sensor board, nul-punt calibratie in diagnostiek programma. Na calibratie een controle uitgevoerd , maar waarden liggen nog steeds slecht. Nieuwe Sensor Cassette geïnstalleerd maar qc-waarden blijven slecht. | Uiteindelijk een nieuwe Reagenspack geinstalleerd, nu liggen alle controles mooi binnen de grenzen ! Bestanden gedownload om vervangen consumables te claimen (SC lotnr 241913 en SP 241688) | SWO-264290 | [5 empty] | Rejected Checks | Connection connector between Cassette and Analyzer looks nasty. pH too low and pCO2 too high for QC Level1 Replacement of Sensor board, Zero-point calibration in diagnostics program. After calibration a check performed, but values are still bad. New Sensor Cassette installed but QC values remain bad. | Finally installed a new reagent pack, Now all controls are beautiful within the boundaries! Files downloaded to claim replaced consumables (SC Lot no 241913 and SP 241688)
No | Visit | Complaint | RAME | Leaking from the waste | Fluid leaking out the bottom | The waste was leaking onto the counter. I found the waste lug not tight and think that may be the cause. I replaced the case due to the stain on it. | SWO-264935 | [5 empty] | Leaking from the waste | Fluid leaking out the bottom | The waste was leaking onto the counter. I found the waste lug not tight and think that may be the cause. I replaced the case due to the stain on it.
No | Visit | Complaint | RIND | sensor connector not responding | sensor casette connector was not sensing Sc | checked the analyser connector was not sensing the sensor so cleaned the connector then tried but still its failed to calibrate then placed other with newconnector then it got calibrated | SWO-264202 | [5 empty] | sensor connector not responding | sensor casette connector was not sensing Sc | checked the analyser connector was not sensing the sensor so cleaned the connector then tried but still its failed to calibrate then placed other with newconnector then it got calibrated
No | Visit | Complaint | RCHN | 光学强度超范围 | 光学强度超范围 | 更换903-364后修复by tim | SWO-264627 | Defect | HW | Process Module | Optical Unit | 1031 | Optical strength over range | Optical strength over range | Repair by tim after replacing 903-364
No | Visit | Complaint | RAME | Customer reported oxy liquid sensor. | Oxy liquid sensor error after LS adjust. Noticed level 3 QC had no scheduled time entered for automatic runs. Notified user and added scheduled time. | Flush air line. Restart calibrations. | SWO-264417 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic manifold | 0 | Customer reported oxy liquid sensor. | Oxy liquid sensor error after LS adjust. Noticed level 3 QC had no scheduled time entered for automatic runs. Notified user and added scheduled time. | Flush air line. Restart calibrations.
No | 3 | Y | RCZE | Blank |  |  | D3729288BE55D4D7 | Unknown | Unknown | Unknown | Cl- |  | Blank
No | Visit | Complaint | RBNL | AQT90 staat in het FOUT met de melding, 1999: UNDEFINED_LOOKUP_STRING22 66 In het service menu komen alle onderdelen in SAFE POSITIE zonder foutmelding. Diverse onderdelen los heb gecontroleerd, geen afwijkingen heb geconstateerd. Incubatorwheel en bijbehorende sensoren heb gereinigd, deksel heb omgewisseld. Vervolgens Systeemcheck, systeemreiniging en een bloedmetingen heb uitgevoerd, OK. | AQT90 staat in het FOUT met de melding, 1999: UNDEFINED_LOOKUP_STRING22 66 In het service menu komen alle onderdelen in SAFE POSITIE zonder foutmelding. Diverse onderdelen los heb gecontroleerd, geen afwijkingen heb geconstateerd. Incubatorwheel en bijbehorende sensoren heb gereinigd, deksel heb omgewisseld. Vervolgens Systeemcheck, systeemreiniging en een bloedmetingen heb uitgevoerd, OK. | Sensoren heb gereinigd. | SWO-264277 | Cleaned | HW | Incubator module | Sensors |  | AQT90 is in ERROR with the message, 1999: UNDEFINED_LOOKUP_STRING22 66 In the service menu all parts are in SAFE POSITION without error message. Several parts have been checked, no abnormalities have been observed. Incubatorwheel and associated sensors have been cleaned, swapped lid. Then system check, system cleaning and a blood measurements performed, OK. | AQT90 is in ERROR with the message, 1999: UNDEFINED_LOOKUP_STRING22 66 In the service menu all parts are in SAFE POSITION without error message. Several parts have been checked, no abnormalities have been observed. Incubatorwheel and associated sensors have been cleaned, swapped lid. Then system check, system cleaning and a blood measurements performed, OK. | Sensors have been cleaned.
No | Visit | Complaint | RDEU | 15.11.2017 10:14Uhr Sr. Angela 03033609486: Die Eingabe im Patienten-ID-Feld sind nicht möglich | Tastatur hakt | SW update durchgef. (siehe folgende Cases; Kal. + QC i.O. | SWO-264933 | Defect | SW | RMED SW | RMED SW |  | Nov 15, 2017 10:14am Sr Angela 03033609486: It is not possible to enter the Patient ID field | Keyboard hakt | SW update. (see the following cases; Kal. + QC in the first place
No | Visit | Complaint | RAME | The analyzer will not transmit patient demographic to the LIS | The analyzer will not transmit patient demographic to the lab computer. Need to modify the ROMA.ini file. | The Connection1 is missing from the ROMA.ini file. Modified the ROMA.ini file to include Connection1 and added command line ETXBTXClassic=False to the file. Analyzer still will not transmit demographics. IT is checking with Sunquest. Waiting for information back from Sunquest. | SWO-265392 | Defect | SW | roma.ini | roma.ini |  | The analyzer will not transmit patient demographic to the LIS | The analyzer will not transmit patient demographic to the lab computer. Need to modify the ROMA.ini file. | The Connection1 is missing from the ROMA.ini file. Modified the ROMA.ini file to include Connection1 and added command line ETXBTXClassic=False to the file. Analyzer still will not transmit demographics. IT is checking with Sunquest. Waiting for information back from Sunquest.
No | Visit | Complaint | RDEU | 17.11.2017 12:00Uhr Herr Sperling 02591231224: Wird eine Probe in den Mischer eingelegt, dann startet der Mischer nicht | Mischermotor läuft nicht | Gerät überprüft, Verbindungskabel kontrolliert, Neustart durchgeführt | SWO-265422 | Defect | HW | Basic Unit | Steppermotor, Mixer |  | 17.11.2017 12:00o Mr. Sperling 02591231224: If a sample is inserted into the mixer, the mixer does not start | Mixer motor does not run | Device checked, connection cable checked, restart performed
No | Visit | Complaint | RLTD | Caroline Brown 01225 824964 | On looking at the data logs, the analyser was reporting - 1601 pressure test flow errors, 1063 pressure test vacuum errors, 1081 $ 1083 $ 1084 $ 1085 $ 1086 $ 1087 Inhomogeneous rinse $ cal 2 $ cal 3 QC1 $ QC2 $QC3 errors, and 1178 pump calibration error. On entering service, the fluid tests were ran in sequence and the pressure test failed on all tests and no further tests were ran. The inlet handle was removed and the inlet connector gasket was inspected to ensure that it was seated correctly. It was noted that the inlet probe was bent and this was straightened. The analyser was restarted and it failed once again. On removing the inlet gasket, a clot was seen to be hanging from the end of the probe. A back flush of the analyser was performed and the fluid tests ran and passed. Analyser restarted and passed calibrations and QC’s with no further errors. | Performing a back flush of the analyser and straightening the inlet probe resolved the issues with the analyser. | SWO-264984 | Cleaned | HW | Auto Inlet module | Clot, probe |  | Caroline Brown 01225 824964 | On looking at the data logs, the analyser was reporting - 1601 pressure test flow errors, 1063 pressure test vacuum errors, 1081 / 1083 / 1084 / 1085 / 1086 / 1087 Inhomogeneous rinse / cal 2 / cal 3 QC1 / QC2 /QC3 errors, and 1178 pump calibration error. On entering service, the fluid tests were ran in sequence and the pressure test failed on all tests and no further tests were ran. The inlet handle was removed and the inlet connector gasket was inspected to ensure that it was seated correctly. It was noted that the inlet probe was bent and this was straightened. The analyser was restarted and it failed once again. On removing the inlet gasket, a clot was seen to be hanging from the end of the probe. A back flush of the analyser was performed and the fluid tests ran and passed. Analyser restarted and passed calibrations and QC’s with no further errors. | Performing a back flush of the analyser and straightening the inlet probe resolved the issues with the analyser.
No | Visit | Complaint | RAME | The oximetry parameters will work for a week or two after installing a sensor cassette and solution pack, and then all the oxi parameters are out of range. | All oximeter parameters are red. Performed complete oximeter test and found the hemolyzer failing the lamp intensity and cavitation tests. | Replaced the hemolyzer unit. Performed tHb calibration and system cycle and ran QC. Analyzer checks okay. | SWO-264546 | [5 empty] | The oximetry parameters will work for a week or two after installing a sensor cassette and solution pack, and then all the oxi parameters are out of range. | All oximeter parameters are red. Performed complete oximeter test and found the hemolyzer failing the lamp intensity and cavitation tests. | Replaced the hemolyzer unit. Performed tHb calibration and system cycle and ran QC. Analyzer checks okay.
No | Visit | Complaint | RCHN | 进样针定标错误 | 进样针定标错误，借902-885一个 | 更换902-885后修复by tim | SWO-263213 | Defect | HW | Inlet module | Friction detection | 0 | Sample pin calibration error | Sample pin calibration error, borrow 902-885 one | Repair by tim after replacement 902-885
No | 9 | Y | RPAC | Blank | reference electrode failure |  | D373F1B13453A86E | Assisted Replacement | Electrode | Electrode | Reference | 0 | Blank | reference electrode failure
No | Visit | Complaint | RKK | NaとCaに機種間差が出る為引き上げ、１号機とNa３～５高く出る | 引き上げ | 引き上げ修理 | SWO-264709 | Unknown | Unknown | Unknown | Unknown | LEH mail | Raise dare to come out between the models, the first unit and Na3- 5 high | Raising | Raise repair
No | Visit | Complaint | RIND | Na and Ca status out of limit. | Replaced reference membrane, still error persists. Cross checked reference electrode in a working ABL800 and found defective( Lot#40-13) | Replaced new reference electrode from stock, 2 point calibration performed successfully. | SWO-265368 | Assisted Replacement | Electrode | Electrode | Reference | lot 4013 | Na and Ca status out of limit. | Replaced reference membrane, still error persists. Cross checked reference electrode in a working ABL800 and found defective( Lot#40-13) | Replaced new reference electrode from stock, 2 point calibration performed successfully.
No | Visit | Complaint | RDEU | Sehr häufig Na, Ca und Cl Kalibrationsdriften | QC im Bereich der Lyte und pH weiter schwankend Na RiliBÄK Fehler oder komplett aus QC Bereich Garantieaustausch | Referenzelektrode und Na Elektrode ausgetauscht. Mehrfach Kalibrationen durchgeführt. Weitere Kalibrationen und Überwachung via Remote Support. QC ohne Beanstandung. QC in allen 3 Leveln zielwertig | SWO-265003 | Assisted Replacement | Electrode | Electrode | Reference | LEH mail ?? | Very common Na, Ca and Cl calibration drifts | QC in the field of lyte and pH continues to fluctuate Na RiliBÄK error or completely out of QC range Warranty exchange | Reference electrode and Na electrode. Multiple calibrations performed. Further calibrations and monitoring via remote support. QC without complaint. QC in all 3 levels targeted
No | Visit | Complaint | RAME | Having correlation issues with the MetHb and COHB. | Customer reported still having correlation issues with the MetHb and COHB. | When I repaired and tested the analyzer before I found the results on 5 blood samples to be no further off than 0.2 on either MetHb or COHB. Customer reported that the results were as much as 0.5 apart. Upon arriving ran complete oxi test and found the halogen lamp results a little high (3455). Went ahead and replaced hemolyzer and re-adjusted lamp. Adjusted to 3300 but when I ran the service thb I got errors 250 and 579 Oxi hardware issue. Replaced lamp unit and optical cable at this point. After replacing the lamp and optical fiber I performed the lamp adjustment ( much easier with new parts) and set to 3300.Ran service thb cal, user thb cal and found all results ok. | SWO-265059 | Defect | HW | Oxi module | Lamp unit | 0 | Having correlation issues with the MetHb and COHB. | Customer reported still having correlation issues with the MetHb and COHB. | When I repaired and tested the analyzer before I found the results on 5 blood samples to be no further off than 0.2 on either MetHb or COHB. Customer reported that the results were as much as 0.5 apart. Upon arriving ran complete oxi test and found the halogen lamp results a little high (3455). Went ahead and replaced hemolyzer and re-adjusted lamp. Adjusted to 3300 but when I ran the service thb I got errors 250 and 579 Oxi hardware issue. Replaced lamp unit and optical cable at this point. After replacing the lamp and optical fiber I performed the lamp adjustment ( much easier with new parts) and set to 3300.Ran service thb cal, user thb cal and found all results ok.
Yes | Visit | Complaint | RAME | When on site to run method comparison with NICU 835 I found errors 791 had been occurring for a couple of days. | Upon returning to site I found that the analyzer had worked fine overnight. All of the oxi parameters were resulted. | Parts replaced on 9 $5 must have corrected the issue, however it took a few blood samples before the results did not have error 791.Continued with method comparison between analyzers. | SWO-265107 | Re classified RMED | 0 | 0 | 0 | CRM consequence | When on site to run method comparison with NICU 835 I found errors 791 had been occurring for a couple of days. | Upon returning to site I found that the analyzer had worked fine overnight. All of the oxi parameters were resulted. | Parts replaced on 9 $5 must have corrected the issue, however it took a few blood samples before the results did not have error 791.Continued with method comparison between analyzers.
No | Visit | Complaint | RSCH | Rep Processwheel | 5684: Rotateprocesswheel_40 Bricht Kalibrierungen ab |  | SWO-265056 | Defect | SW | Process Module | Process wheel | 5684 | Rep Processwheel | 5684: Rotateprocesswheel_40 Cancels calibrations
No | Visit | Complaint | RDEU | PO2 Fehler bei Kal 1006 Case schon über Hotline dokumentiert | PO2 Fehler 1006 | Gerät zur Rep mitgenommen, Leihgerät 090R0349N005 zur Rep Überbrückung aufgestellt - konfiguriert und an Netzwerk angebunden, Kal und QC durchgeführt, Allgemeine Geschäftsbedingungen für Leihgeräte durchgesprochen und wird akzeptiert, | SWO-265434 | Revisit required | Install loaner | Basic Unit | pO2 | 1006 | PO2 error in Kal 1006 case already documented via hotline | PO2 Error 1006 | Device taken to rep, loaner 090R0349N005 set up for rep bridging - configured and connected to network, Kal and QC performed, general terms and conditions for loaners discussed and accepted,
No | Visit | Complaint | RAME | Part 905-847 tubing nipples broken from analyzer. | Examined inlet module. Found broken waste nipple connectors. | Replaced nipple connectors. Verified operation of ABL 825. | SWO-265363 | Defect | HW | Inlet module | Nipple kit | 0 | Part 905-847 tubing nipples broken from analyzer. | Examined inlet module. Found broken waste nipple connectors. | Replaced nipple connectors. Verified operation of ABL 825.
No | Visit | Complaint | RAME | Analyzer has oximetry hardware error messages. | Oximetry Hardware errors 0250, 0251, 0579. | Replaced hemolyzer unit. Programmed oximeter controller board, aligned Hb Lamp, performed complete oximeter alignment, performed calibrations and ran QC. Analyzer checks okay. | SWO-265394 | Defect | HW | Oxi module | Hemolyzer | 0 | Analyzer has oximetry hardware error messages. | Oximetry Hardware errors 0250, 0251, 0579. | Replaced hemolyzer unit. Programmed oximeter controller board, aligned Hb Lamp, performed complete oximeter alignment, performed calibrations and ran QC. Analyzer checks okay.
No | Visit | Complaint | RCHN | 自动关机 | 自动关机，借910-285一个 | 更换910-285后修复tim | SWO-265385 | Defect | HW | Basic Unit | Power Supply |  | Auto-shutdown | Automatic shutdown, borrowing 910-285 one | Repair tim after replacement 910-285
No | Visit | Complaint | RCHN | 报错1031 | 报错1031，需要借件903-364 | 更换903-364后修复(By Ran) | SWO-264003 | Defect | HW | Process Module | Optical Unit | 1031 | 1031 | 1031, need to borrow 903-364 | Repair after replacement 903-364 (By Ran)
No | Visit | Complaint | RCHN | 进样门故障 | 进样门故障，借903-028一个 | 更换903-028后修复by tim | SWO-265365 | Defect | HW | Cover and Chassis | PCB, Cabinet top | 4548 | Sample door failure | Sample door fault, borrow 903-028 one | Repair by tim after replacement 903-028
No | Visit | Complaint | RDEU | Lactat ist nicht messbereit, verbrauchsmittel wurden bereits erneuert | Lactat nicht messbereit | Fehleranalyse, sensorkasette aus kondi gerät eingesetzt und gerät kalibriert, weiterhin Lactat Fehler, Abbruch zwecks sectio op | SWO-266067 | Cleaned | HW | Basic Unit | SC Contact |  | Lactate is not ready for measurement, consumables have already been renewed | Lactate not ready for measurement | Error analysis, sensor casette from kondi device used and gets calibrated, continue lactate error, cancellation for sectio op
No | Visit | Complaint | RDEU | Gerät geht zeitweise ohne Grund in Intervention | Zeitweise inhomogene Lösungen | Gerät überprüft und gereinigt, Schlauchsytem erneuert, Ventil kontrolliert, tHb Kal durchgeführt | SWO-265694 | Defect | HW | Basic Unit | PM Kit | Cleaned | Device goes into intervention at times for no reason | Intermittently inhomogeneous solutions | Device checked and cleaned, hose system renewed, valve inspected, tHb Kal performed
Yes | 7 | Y | RPAC | Blank | Analyser encountering parameter issues on calibration and QC. |  | D377AD47960022C5 | [5 empty] | Blank | Analyser encountering parameter issues on calibration and QC.
Yes | 7 | Y | RPAC | Blank | Sensor Cassette leaked |  | D377A966960B2F12 | [5 empty] | Blank | Sensor Cassette leaked
Yes | Visit | Complaint | RDEU | Lösungspack lässt sich nicht entriegeln | Lösungspack lässt sich nicht heraus nehmen | Da der Fehler bei der Auslieferung auftrat, Gerät zur Reparatur mit in Werkstatt genommen. | SWO-267804 | Re classified RMED | OOB |  |  |  | Solution pack cannot be unlocked | Solution pack can't be taken out | Since the error occurred during delivery, the device was taken to the workshop for repair.
No | Visit | Complaint | RLTD | Contact Sergio Guimaraes 01689 865318 | Analyser failing pressure valve drawn test with too high a value as well as the pump calibration. Tested Oxi pinch valve, opening and closing ok. Performed manual flush - sample pathway clear. Removed analyser wet section and inspected Oxi pinch valve & tube - no visible fault. Cleaned & lubricated pinch valve and replaced pinch valve tube. After re-assembly tested in service. - pressure test and pump calibration now passing. Also reset battery unit as showing 0% charge when fully charged. Refitted older solution pack and restarted analyser and verified operation - all ok including battery charge % now correct. | Replaced Oxi pinch valve tube. | SWO-266031 | Defect | HW | Oxi module | Pinch valve tube |  | Contact Sergio Guimaraes 01689 865318 | Analyser failing pressure valve drawn test with too high a value as well as the pump calibration. Tested Oxi pinch valve, opening and closing ok. Performed manual flush - sample pathway clear. Removed analyser wet section and inspected Oxi pinch valve & tube - no visible fault. Cleaned & lubricated pinch valve and replaced pinch valve tube. After re-assembly tested in service. - pressure test and pump calibration now passing. Also reset battery unit as showing 0% charge when fully charged. Refitted older solution pack and restarted analyser and verified operation - all ok including battery charge % now correct. | Replaced Oxygen Pinch valve tube.
No | Visit | Complaint | RDEU | Analysator startet nicht | Analysator startet nicht Bleibt bei Neustart im Start up hängen ! (Meldung von WE) | Neustart - Analysator startet ohne Probleme (Messbereit in 55 Minuten) - nach Ablauf erneut Neustart durchgeführt (Analysator hängt im Start Up) - Windows Installation (wird nicht abgeschlossen - Blauer Bildschirm!) - CF Karte erneuert (Windows Installation wird ebenfalls nicht abgeschlossen) - Bestellung CPU erfolgt ! Aus Medizintechnik ABL 90Flex Plus (092R0015N001) zur Reparatur-Überbrückung aufgestellt | SWO-265983 | Revisit required | HW | DMS_PC | CPU |  | Analyzer does not start | Analyzer does not start Gets stuck in the start-up when you restart ! (Message from WE) | Restart - Analyzer starts without any problems (ready to be measured in 55 minutes) - Restart after expiration (Analyzer hangs in Start Up) - Windows installation (not completed - Blue screen!) - CF card renewed (Windows installation will also not be renewed completed) - Order CPU is done ! From medical technology ABL 90Flex Plus (092R0015N001) for repair bridging
No | Visit | Complaint | RCHN | 检测不到试剂包 | 检测不到试剂包，借903-158一个 | 自行更换903-158后修复 | SWO-265637 | Defect | HW | Basic Unit | PCB, Solution Pack PCB Id |  | Reagent pack not detected | No reagent pack detected, borrow 903-158 one | Repair after replacing 903-158
No | Visit | Complaint | RCHN | 间接性定标失败 | 间接性定标失败，需要借件903-581一个 |  | SWO-265528 | Defect | HW | Basic Unit | PCB, Sensor interface |  | Indirect calibration failure | Indirect calibration failure, need to borrow 903-581 one
Yes | 10 | Y | RPAC | Blank | FAN915-318 |  | D37CFB286E2F6F9A | Re classified RMED | FAN |  |  |  | Blank | FAN915-318
No | Visit | Complaint | RLTD | Please call Rosie on 01865 221 596 | Intermittent 0791 errors unable to measure parameter for Co-Ox parameters, this indicates inhomogeneous or insufficient sample. Complete Oxi test run and all parameters passed. Checked the pinch valve to make sure it was working and there were no problems with this as the analyser was due its annual service I decided to do this. | Annual service performed and after this it was giving Co-Ox results on patient samples again. | SWO-266119 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | Please call Rosie on 01865 221 596 | Intermittent 0791 errors unable to measure parameter for Co-Ox parameters, this indicates inhomogeneous or insufficient sample. Complete Oxi test run and all parameters passed. Checked the pinch valve to make sure it was working and there were no problems with this as the analyser was due its annual service I decided to do this. | Annual service performed and after this it was giving Co-Ox results on patient samples again.
No | Visit | Complaint | RAME | Customer reported No leading air segment in inlet's liquid sensor within timeframe. errors during patient samples. | Checked data logs for errors. Error code 0328 (No leading air segment inlets liquid sensor within timeframe) intermittently during patient samples. | Checked fluid flow. Replaced pH seal ring and air valve membrane in pH $BG module Restart calibrations and ran several blood samples with no errors. Continued to have issue. Replaced additional air membranes on other modules and Inlet gasket. Restart testing with blood samples. | SWO-266181 | Defect | HW | El/Met Module | Valve, air | 328 | Customer reported No leading air segment in inlet's liquid sensor within timeframe. errors during patient samples. | Checked data logs for errors. Error code 0328 (No leading air segment inlets liquid sensor within timeframe) intermittently during patient samples. | Checked fluid flow. Replaced pH seal ring and air valve membrane in pH $BG module Restart calibrations and ran several blood samples with no errors. Continued to have issue. Replaced additional air membranes on other modules and Inlet gasket. Restart testing with blood samples.
Yes | Visit | Complaint | RDEU | Geräte Umstellung an den neuen Standort im Neubau | Standortwechsel erforderlich. | Gerät sollte auf der Station P12 abgeholt und anschließen am neuen Standort Inbetrieb genommen. Gerät auf der Station P12 gesucht, wurde von der Station vorab transportiert. Am neuen Standort Gerät in Betrieb genommen. | SWO-266422 | Re classified RMED | Moving analyzer to new Loc. |  |  |  | Devices Conversion to the new location in the new building | Relocation required. | Device should be picked up at station P12 and connected to the new location put into operation. device on station P12 was transported from the station in advance. Put into operation at the new location device.
No | Visit | Complaint | RAME | Gas pressure offset out of range, error 689. | Customer getting error 689 on analyzer. Ongoing oxi errors since installation. | Replaced power supply and AU board due to intermittent errors. | SWO-266205 | Defect | HW | Basic Unit | PCB, Analyzer control | 689 | Gas pressure offset out of range, error 689. | Customer getting error 689 on analyzer. Ongoing oxi errors since installation. | Replaced power supply and AU board due to intermittent errors.
No | Visit | Complaint | RKK | タッチパネル反応しない。ファームウェアアップグレードも効果なし。ディスプレイ交換。 | タッチパネル反応しない。ファームウェアアップグレードも効果なし。 | ディスプレイ交換。 | SWO-265625 | Defect | HW | DMS_PC | Touch screen | Firmware not possible | Touch panel does not react. Firmware upgrade swelled. Display replacement. | Touch panel does not react. Firmware upgrade swelled. | Display replacement.
No | Visit | Complaint | RDEU | Probleme mit Bildschirmbedienung. | Fehlersuche vor Ort. | Herr Höllriegl hat mir mitgeteilt dass der Bildschirm des Öfteren nicht richtig bedienbar ist. Erst nach Neustart ist der Fehler weg, kommt aber nach einiger Zeit wieder. Ein Bildschirmsoftwareupdate brachte keine Besserung. Neuen Bildschirm bestellt. | SWO-265771 | Revisit required | HW | DMS_PC | Touch screen | Firmware not possible | Problems with screen operation. | Troubleshooting on site. | Mr Höllriegl told me that the screen is often not properly operated. Only after reboot the error is gone, but comes back after some time. A screen software update did not improve. New screen ordered.
No | Visit | Complaint | RAME | Ca not passing on levels 2 and 3. | Ionized calcium running on the low end of WDC Peer 2 Mean. Peer 2 SD range of Level 2 is .51-.59. Level 3 Peer 2 SD range is .34-.43 | Investigated data logs and examined electrodes and measuring chamber. Replaced measuring chamber for investigative reasons | SWO-266179 | Defect | HW | El/Met Module | Measuring Chamber |  | Ca not passing on levels 2 and 3. | Ionized calcium running on the low end of WDC Peer 2 Mean. Peer 2 SD range of Level 2 is .51-.59. Level 3 Peer 2 SD range is .34-.43 | Investigated data logs and examined electrodes and measuring chamber. Replaced measuring chamber for investigative reasons
No | Visit | Complaint | RDEU | zeitw. werden die Proben abgebrochen, inhomogene Probe, zu wenig Probe, | Einl.nadel ist verbogen, | Einl.nadel ersetzt, Gerät kalibriert, QC durchgeführt, | SWO-265795 | Defect | HW | Inlet module | Probe bend |  | zeitw. samples are aborted, inhomogeneous sample, too little sample, | Einl.nadel is bent, | One-needle replaced, device calibrated, QC performed,
No | Visit | Complaint | RAME | Intermittent error 0315 neon number too high. | Isolated problem to a defective spectrometer. | Replaced the spectrometer unit. | SWO-266163 | Defect | HW | Oxi module | Spectrometer |  | Intermittent error 0315 neon number too high. | Isolated problem to a defective spectrometer. | Replaced the spectrometer unit.
Yes | 12 | Y | RPAC | Blank | Software Upgrade FAN915-365R4 V8.9.613.0 |  | D37776B79625E27D | Re classified RMED | SWU |  |  |  | Blank | Software Upgrade FAN915-365R4 V8.9.613.0
No | Visit | Complaint | RDEU | 15.11.2017, 15:36 Uhr ==> Anrufer Fr. Irlinger, Tel.: 08654606534 ==> Halterung für Probenröhrchen am Inlet wheel defekt | Probenhalter gebrochen | Probenhalter erneuert Systemreinigung -und überprüfung durchgeführt | SWO-265841 | Defect | HW | Cover and Chassis | Tube holder |  | 15.11.2017, 15:36 ==> Caller Fr. Irlinger, Tel.: 08654606534 ==> Bracket for sample tubes on the inlet wheel defective | Sample holder broken | Sample holder renewed System cleaning and verification performed
No | Visit | Complaint | RDEU | 17.11.2017, 11:20Uhr, Anrufer Frau Hernick Tel.: 0211904361209 - 1083 Durckerstatus offline | Printer Unit defekt. | Drucker meldet das kein Papier eingelegt ist. Drucker gereinigt, weiterhin keine Funktion. Printer Unit erneuert. | SWO-266420 | Defect | HW | DMS_PC | Printer Module |  | 17.11.2017, 11:20a.m., caller Mrs. Hernick Tel.: 0211904361209 - 1083 Durcker status offline | Printer Unit defective. | Printer reports that no paper is loaded. printer cleaned, still no function. Printer Unit.
No | Visit | Complaint | RDEU | Lactat ist nicht messbereit, verbrauchsmittel wurden bereits erneuert | Messing Kontakte korrodiert | Messingkontakte gereinigt und auf Gangbarkeit kontrolliert , Kalibration und QC durchgeführt | SWO-267385 | Cleaned | HW | Basic Unit | SC Contact |  | Lactate is not ready for measurement, consumables have already been renewed | Brass contacts corroded | Brass contacts cleaned and checked for gaitability, calibration and QC performed
Yes | Visit | Complaint | RDEU | Lactat ist nicht messbereit, verbrauchsmittel wurden bereits erneuert | System überprüft => Messingkontakte korrediert | Messingkontakte gereinigt und auf Gangbarkeit kontrolliert , Kalibration und QC durchgeführt | SWO-267294 | Re classified RMED | Double entry |  |  |  | Lactate is not ready for measurement, consumables have already been renewed | System checked => Brass contacts corred | Brass contacts cleaned and checked for gaitability, calibration and QC performed
No | Visit | Complaint | RLTD | Glucose faking calibration and QC over a period of months | Analyser intermittently failing calibration and QC on Glucose | Checked alignment of sensor interface PCB with sensor bay, ok Replaced sensor interface PCB with new | SWO-266758 | Defect | HW | Basic Unit | PCB, Sensor interface |  | Glucose faking calibration and QC over a period of months | Analyser intermittently failing calibration and QC on Glucose | Checked alignment of sensor interface PCB with sensor bay, ok Replaced sensor interface PCB with new
No | Visit | Complaint | RSCH | Blank | Impossible de pomper les solutions malgré le changement de 2 SP | Nettoyage du système fluidique e.o Vérification de la pompe péristaltique e.o Calibration e.o contrôle de qualité e.o | SWO-266693 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | Blank | Unable to pump solutions despite 2 SP change | Cleaning the fluid system e.o Checking the peristaltic pump e.o Calibration e.o quality control e.o
No | Visit | Complaint | RAME | Sensor cassette due to be replaced Lactate didn't calibrated | SC due to be replaced sensitivity less than 10. | Replaced the Sensor interface Pcb following service tip with a Pcb with a greater SN than R2003N0800 $There was no significant change $Opening a techline case | SWO-267127 | Revisit required | Unknown | Unknown | Lac |  | Sensor cassette due to be replaced Lactate didn't calibrated | SC due to be replaced sensitivity less than 10. | Replaced the Sensor interface Pcb following service tip with a Pcb with a greater SN than R2003N0800 $There was no significant change $Opening a techline case
No | Visit | Complaint | RLTD | Please call Stephanie Royalington on 01225 825 010 | The extended log-on button had been pressed which was preventing people logging in. This had been rectified by rebooting the analyser. | Changed the log-on to barcode only to prevent this happening again. | SWO-266915 | Setup | SW | User Access | User Access |  | Please call Stephanie Royalington on 01225 825 010 | The extended log-on button had been pressed which was preventing people logging in. This had been rectified by rebooting the analyser. | Changed the log-on to barcode only to prevent this happening again.
Yes | 2 | Y | RPAC | Blank | Lactate failure on two DIFFERENT sensor cassettes |  | D3791A740C446020 | [5 empty] | Blank | Lactate failure on two DIFFERENT sensor cassettes
No | 4 | Y | RPOL | Blank | Wymiana kasety, testy serwisowe analizatora, kalibracje. Aparat sprawny |  | D3794A070C46D367 | Assisted Replacement | SC | SC | Unknown |  | Blank | Cartridge replacement, analyser service tests, calibrations. Efficient camera
Yes | 5 | Y | RPAC | Blank | Solution Pack leaked |  | D378BB060C101929 | [5 empty] | Blank | Solution Pack leaked
No | Visit | Complaint | RDEU | AQC Probleme | CREA außerhalb der QC Grenzen | Gerät überprüft, Sensorkassette erneuert | SWO-266715 | Assisted Replacement | SC | SC | Crea |  | AQC Problems | CREA outside the QC limits | Device checked, sensor cartridge renewed
No | Visit | Complaint | RDEU | Analysator startet nicht | Erneuerung CPU | Erneuerung CPU - Software Installation - SC und SP Installation - Einstellungen und Funktionsüberprüfung etc. - keine Zeitangabe und Zeitbalken während Aktivitäten ! Gespeicherte Daten (alle Daten) werden nicht eingelesen (Wiederherrstellung gescheitert) - Analysator mitgenommen ! Neue MAC Adr. : 00-04-5F-9D-DA-32 | SWO-266878 | Defect | HW | DMS_PC | CPU |  | Analyzer does not start | Renewal CPU | Renewal CPU - Software Installation - SC and SP Installation - Settings and Function Verification etc. - No time and time bar during activities ! Saved data (all data) are not read in (re-mastery failed) - Analyzer is taken with you ! New MAC Adr. : 00-04-5F-9D-DA-32
No | Visit | Complaint | RIND | SC temp is not stable. | Checked Sc and cleaned with distilled water. Then kept in fridge for 1 hour. | Done SC zeroing procedure and completed it successfully. Then reinstalled Sc and done calibrations. Checked with samples. Unit is working fine. | SWO-269042 | [5 empty] | SC temp is not stable. | Checked Sc and cleaned with distilled water. Then kept in fridge for 1 hour. | Done SC zeroing procedure and completed it successfully. Then reinstalled Sc and done calibrations. Checked with samples. Unit is working fine.
Yes | Visit | Complaint | RCHN | 报错1024 | 报错1024，借903-364,903-329各一个 | 更换903-364及903-329后修复(By Ran) | SWO-266296 | Defect | HW | Process Module | Optical Unit | 1024 | 1024 error | 1024, lend ingress 903-364, 903-329 each | Replacement 903-364 and 903-329 post-repair (By Ran)
No | Visit | Complaint | RCHN | 皮带断裂 | 皮带断裂，借867-631一个 | 更换867-631后修复(By Ran) | SWO-265659 | Defect | HW | Aspiration Module | Timing belt | LEH TL | Belt break | The belt is broken, borrowed 867-631 one. | Repair after replacement 867-631 (By Ran)
No | Visit | Complaint | RDEU | Elyte außerhalb der AC Grenzen | Na und Cl Werte außerhalb der Grenzen | Gerät geprüft, Elektroden neu membraniert, Ref. Elektrode geprüft, Kal 1 $2 (Lot YB01 $XT05) ausgetauscht, Funktionskontrolle ok | SWO-267235 | Assisted Replacement | Electrode | Membrane | Na/Cl |  | Elyte outside the AC borders | Na and Cl values outside the limits | Device tested, electrodes re-membraneed, ref. Electrode tested, Kal 1 x 2 (Lot YB01 $XT05) replaced, function control ok
Yes | Visit | Complaint | RDEU | Lösungspack lässt sich nicht entriegeln | Gerät zurückgeleitest | Gerät wieder beim Kunden aufgestellt. Gerät überprüft. | SWO-267803 | Re classified RMED | OOB |  |  |  | Solution pack cannot be unlocked | Device retracted test | device is set up again at the customer's device checked.
No | Visit | Complaint | RBNL | Accustoring, error 1559. | ABL geeft error 1559, accustoring. Derhalve de accu vervangen (onder garantie). Melding is nu weg. ABL opgestart en meetklaar achtergelaten. | Accu vervangen. | SWO-267563 | Defect | HW | Battery module | Battery |  | Accustoring, error 1559. | ABL gives error 1559, battery failure. Therefore replace the battery (under warranty). Notification is now gone. ABL booted and measurement ready left behind. | Replace the battery.
Yes | Visit | Complaint | RDEU | Analysator startet nicht | Kein Zeitbalken und Zeitangabe | In der Medizintechnik aus Radiometer Bestandsanalysator Wetsection Control Board demontiert und zu Testzwecken in Analysator montiert (keine Beanstandung) - Rückmontage und Neuteil-Bestellung ! | SWO-267352 | Re classified RMED | Repeated visit |  |  |  | Analyzer does not start | No time bar and time indication | In medical technology from Radiometer inventory analyzer Wetsection Control Board dismantled and assembled for testing purposes in analyzer (no complaint) - reassembly and new part order !
No | Visit | Complaint | RSCH | Rep. Mögliche Verstopfung | Mögliche Verstopfung der Fliesswege Gerät undicht Lösung läuft unter Gerät aus | Kundengerät durch Leihgerät ersetzt i.o Konfiguration geladen i.o Leihgerät System i.o Kundengerät wird in Thalwil repariert | SWO-267519 | [5 empty] | Rep. Possible constipation | Possible blockage of the flow paths Device leaks Solution expires under device | Customer device replaced by rental device i.o Configuration loaded i.o Rental device System i.o Customer device is repaired in Thalwil
Yes | Visit | Complaint | RAME | Oxi Hardware error | Neon lamp Voltage to low | Needs oxi controls pcb replaced | SWO-267770 | Re classified RMED | Phone |  |  |  | Oxi Hardware error | Neon lamp Voltage to low | Needs Oxygen Controls PCB replaced
No | Visit | Complaint | RDEU | 27.07.2017; Sr. Anett; Tel.: 03381791614; CRP QC Messfehler 28.07.2017 10:30 Sr. Nicole; Tel.: 03381791612; weiterhin CRP QC Messfehler 31.07.2017: Rücksprache mit KerZi - neues QC wird bestellt, telefonische Info an Kunden 03.08.2017, 11:36 Uhr ==> Anrufer Pfl. Sebastian, Tel.: 03381791611 ==> LQC-Messfehler CRP 04.08.2017, 12:04 Uhr ==> Anrufer Schw. Anett, Tel.: 03381791614 ==> LQC-Fehler CRP besteht trotz neuer Charge weiterhin | Opt. einheit ersetzen (da Leihteil) | Opt. Einheit ersetzt, Startphase, Factoryreference kalibriert, Blankmessungen ok, QK bei niedriger CRP-Kontolle zu hoch und bei hoher CRP-Kontrolle zu niedrig, Gerät stillgelegt und zur Werkstattreparatur abgeholt | SWO-267517 | Defect | HW | Process Module | Optical Unit | QC | 27.07.2017; Sr. Anett; Tel.: 03381791614; CRP QC Measurement Error 07/28/2017 10:30 Sr. Nicole; Tel.: 03381791612; further CRP QC measurement error 31.07.2017: Consultation with KerZi - new QC is ordered, telephone information to customers 03.08.2017, 11:36 a.m. ==> Caller Pfl. Sebastian, Tel.: 0338179111 ==> LQC measurement error CRP 04.08.2017, 12:04 o'clock ==> Caller Schw. Anett, tel.: 03381791614 ==> LQC error CRP persists despite new batch | Opt. replace unit (because loan part) | Opt. Unit replaced, start-up phase, factory reference calibrated, blank measurements ok, QK too high at low CRP contolls and too low at high CRP control, device shut down and picked up for workshop repair
No | 13 | Y | RPAC | Blank | Unable to pump solutions |  | D37B53BAF83168F0 | Defect | HW | Basic Unit | Cassette retainer |  | Blank | Unable to pump solutions
Yes | Visit | Complaint | RAME | po2 Validation Issue | Collected data files and submitted CQC case. | n $a | SWO-267934 | Re classified RMED | PER |  |  |  | po2 Validation Issue | Collected data files and submitted CQC case. | no $a
No | Visit | Complaint | RDEU | 17.11.2017 12:00Uhr Herr Sperling 02591231224: Wird eine Probe in den Mischer eingelegt, dann startet der Mischer nicht | Mischermotor läuft nicht | Gerät überprüft, Mischer Unit (R1221N001) erneuert, Kal und QC o.k. | SWO-268588 | Defect | HW | Basic Unit | Mixer Unit |  | 17.11.2017 12:00o Mr. Sperling 02591231224: If a sample is inserted into the mixer, the mixer does not start | Mixer motor does not run | Device checked, mixer unit (R1221N001) renewed, Kal and QC ok
No | 4 | Y | RPAC | Blank | Magnet for mixing |  | D37B3783F85E095B | Defect | HW | Mixer Unit | Magnet |  | Blank | Magnet for mixing
Yes | 8 | Y | RPAC | Blank | Replace failed SC |  | D37B4543F81A095F | [5 empty] | Blank | Replace failed SC
No | Visit | Complaint | RLTD | Jackie on 01604 545058 | Sensor cassette door won't close after cassette change. | Sensor on cassette bay doors misaligned, adjusted sensor. | SWO-268633 | Adjustment | HW | Basic Unit | SC door |  | Jackie on 01604 545058 | Sensor cassette door won't close after cassette change. | Sensor on cassette bay doors misaligned, adjusted sensor.
No | Visit | Complaint | RDEU | Analysator startet nicht | Analysator im Normalbetrieb, aber ohne Zeitbalken und Zeitangaben | Erneuerung Wet Section Control PCB - Software Installation - Startphase - Konfigurationseinstellungen etc. Analysator auf Station 30-4 in Betrieb genommen | SWO-268238 | Defect | HW | DMS_PC | PCB, Wet section Control | LEH Mail | Analyzer does not start | Analyzer in normal operation, but without time bars and time data | Renewal Wet Section Control PCB - Software Installation - Start Phase - Configuration Settings, etc. Analyzer put into operation at station 30-4
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | O2,GLU,LAC定标不过 | 连续更换2张测试卡后修复 | SWO-269309 | Assisted Replacement | SC | SC | pO2/Glu/Lac |  | Test card, reagent package failure | O2, GLU, LAC calibration but | Repair after 2 consecutive test cards
No | Visit | Complaint | RCHN | 试剂包由于没有芯片断裂 | 试剂包由于没有芯片断裂，需要借件903-158和910-329各一个 | 更换903-158和910-329后修复(By Ran) | SWO-267221 | Defect | HW | Basic Unit | PCB, Solution Pack PCB Id |  | The reagent pack didn't break because of the chip. | Reagent package slots with no chip break, requires a loan of 903-158 and 910-329 each | Replacement 903-158 and 910-329 post-repair (By Ran)
No | Visit | Complaint | RCHN | 显示屏扫码器故障 | 显示屏扫码器故障，借902-924一个 | 更换902-924后修复(By Ran) | SWO-267942 | Defect | HW | Cover and Chassis | Barcode reader, display |  | Display scanner failure | Display scanner failure, borrow 902-924 one | Post-replacement 902-924 (By Ran)
Yes | 10 | Y | RPAC | Blank | inlet nipple broken |  | D37B4A6AF82883B7 | [5 empty] | Blank | inlet nipple broken
No | 5 | Y | RPAC | Blank | Ongoing Spectrometers A to D errors, when the analyser had been turned off. |  | D37B51F2F83E3316 | Defect | HW | Oxi module | Oxi module, Replaced |  | Blank | Ongoing Spectrometers A to D errors, when the analyser had been turned off.
No | Visit | Complaint | RESP | PCO2 | Comprobación de registros de datos del analizador | Se sustituye el electrodo de PCO2 y su membrana, se calibra el equipo en varias ocasiones, y verificamos el correcto funcionamiento con QC. | SWO-268160 | Assisted Replacement | Electrode | Membrane | pCO2 |  | PCO2 | Checking analyzer data records | The PCO2 electrode and its membrane are replaced, the equipment is calibrated several times, and we verify the proper operation with QC.
No | Visit | Complaint | RCHN | 扫码器无法扫描 | 扫码器无法扫描，需要借件902-822和902-833各一个 | 更换902-822和902-833备件后修复(By Ran) | SWO-267223 | Defect | HW | Basic Unit | Barcode reader, Basic Unit |  | The scanner cannot be scanned | The scanner cannot be scanned and requires a loan of 902-822 and 902-833 each | Repair after replacement of 902-822 and 902-833 spare parts (By Ran)
No | Visit | Complaint | RDAN | QC ligger udenfor grænserne på elektrolytter(Na og Ca) | Ca ligger for lavt på QC S7735 og Ca og Na ligger for højt på QC S7765. Der er skiftet ref. membran. men dette har ikke løst problemet. | Jeg har skiftet ref. elektroden. Den nye ref. elektrode har lot 1804, den gamle har lot 7701. Der er kørt 2 punkts kal. Der er kørt QCer Alt er fundet OK | SWO-268324 | Assisted Replacement | Electrode | Electrode | Reference |  | QC is located outside the boundaries of the electrolytes (Na and Ca) | CA is too low on QC S7735 and Ca and Na is too high on QC S7765. The ref. membrane has been changed. But this has not solved the problem. | I changed the ref. electrode. The new ref. Electrode has lot 1804, the old has lot 7701. Two-point Kal has been run. QCer has been run Everything is found OK
No | Visit | Complaint | RIND | Service Visit | Continuous 2 same sample run by customer, the result of electrolytes comes different. | Open the cover of measuring module, check the chamber of Na, K and Ref, no clot found. Close the cover of measuring module and record no replacement, then restart the analyzer, after restarting open the activity archive data in data log section and click on filter and select replacement. No replacement record of reference membrane found from the date of installation. Tell the user and pathologist that need to change the reference membrane. | SWO-268917 | User error | Electrode | Membrane | Reference |  | Service Visit | Continuous 2 same sample run by customer, the result of electrolytes comes different. | Open the cover of measuring module, check the chamber of Na, K and Ref, no clot found. Close the cover of measuring module and record no replacement, then restart the analyzer, after restarting open the activity archive data in data log section and click on filter and select replacement. No replacement record of reference membrane found from the date of installation. Tell the user and pathologist that need to change the reference membrane.
Yes | Visit | Complaint | RIND | Service Visit | Need to change the reference membrane. | Open the cover of measuring modules, remove the reference electrode from the measuring chamber of pH $BG module, remove the electrode jacket by pulling it off and rinse with the tap water and shake gently for removing the drop of tap water from the electrode. Now open the seal of new jacket in membrane box and press the electrode in to the jacket and place it same as before in to the chamber and close the electrode connector. Now record the replacement of reference membrane and close the measuring module cover, then press restart, after restarting the analyzer wait for auto 2 point calibration, successfully done. Now run the same sample two times value comes same, problem has sorted out and analyzer handed over to staff in well working condition. | SWO-268981 | Re classified RMED | Training |  |  |  | Service Visit | Need to change the reference membrane. | Open the cover of measuring modules, remove the reference electrode from the measuring chamber of pH $BG module, remove the electrode jacket by pulling it off and rinse with the tap water and shake gently for removing the drop of tap water from the electrode. Now open the seal of new jacket in membrane box and press the electrode in to the jacket and place it same as before in to the chamber and close the electrode connector. Now record the replacement of reference membrane and close the measuring module cover, then press restart, after restarting the analyzer wait for auto 2 point calibration, successfully done. Now run the same sample two times value comes same, problem has sorted out and analyzer handed over to staff in well working condition.
No | Visit | Complaint | RAME | Client advise Crea our of Calibration even after changing membrane and electrodes | Crea not passing calibration. | Cleaned connection gate contacts, verified no cracks in measuring chamber, cleaned the measuring chamber, verified valve performance logs, ran multiple 2 point caps, analyzer came into spec. | SWO-268491 | Cleaned | HW | MetII Module | Electrode connector | Crea | Client advise Crea our of Calibration even after changing membrane and electrodes | Crea not passing calibration. | Cleaned connection gate contacts, verified no cracks in measuring chamber, cleaned the measuring chamber, verified valve performance logs, ran multiple 2 point caps, analyzer came into spec.
No | Visit | Complaint | RLTD | Contact Dave Gardner 0203 299 9000 x32264 | Customer is experiencing that each tie they need to replace a Sensor Cassette, the new one fails to work on Glucose & Lactate. Sometimes have to try 3 new Sensor Cassettes to get one that will work. Replaced Sensor Interface pcb and once re-assembled fitted a new Sensor Cassette. Also blue light of sample mixer flashes intermittently when nothing is present in the mixer. Removed mixer, cleaned sample mixer tray and refitted cable at sample mixer as slightly askew in the socket. Restarted analyser and verified operation - | Replaced Sensor Interface pcb. | SWO-268911 | Defect | HW | Basic Unit | PCB, Sensor interface |  | Contact Dave Gardner 0203 299 9000 x32264 | Customer is experiencing that each tie they need to replace a Sensor Cassette, the new one fails to work on Glucose & Lactate. Sometimes have to try 3 new Sensor Cassettes to get one that will work. Replaced Sensor Interface pcb and once re-assembled fitted a new Sensor Cassette. Also blue light of sample mixer flashes intermittently when nothing is present in the mixer. Removed mixer, cleaned sample mixer tray and refitted cable at sample mixer as slightly askew in the socket. Restarted analyser and verified operation - | Replaced Sensor Interface pcb.
Yes | 0 | Y | RPAC | Blank | UPS needed for ABL90 |  | D37C76546E4FA45B | Re classified RMED | OOB |  |  |  | Blank | UPS needed for ABL90
Yes | 1 | Y | RPAC | Blank | UPS |  | D37C77536E1BA45D | Re classified RMED | OOB |  |  |  | Blank | UPS
No | Visit | Complaint | RCHN | 定标错误 | 定标错误 | 更换945-737一个正常。 | SWO-269454 | [5 empty] | Calibration error | Calibration error | Replace the 945-737 one normal.
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | SO2质控失败，0090 | 更换试剂包后修复 | SWO-268663 | Assisted Replacement | SP | SP | pO2 |  | Test card, reagent package failure | SO2 QC Failure, 0090 | Repair after replacement of reagent kit
No | Visit | Complaint | RCHN | 定标不过 | 定标不过，借903-581一个 | 更换903-581后修复(By Ran) | SWO-267273 | Defect | HW | Basic Unit | PCB, Sensor interface |  | The calibration is not the | But, borrow 903-581 a | Repair after replacement 903-581 (By Ran)
Yes | Visit | Complaint | RCHN | 电线版线路故障 | 电线版线路故障，需要借件902-925和903-581各一个 |  | SWO-268554 | Re classified RMED | Double entry | Basic Unit | PCB, Sensor interface |  | Wire line failure | Wire line failure, need to borrow 902-925 and 903-581 each one
No | Visit | Complaint | RCHN | LAC定标不过 | LAC定标不过，借903-581一个 | 更换903-581后修复(By Ran) | SWO-267990 | Defect | HW | Basic Unit | PCB, Sensor interface |  | LAC calibration is not | LAC calibration, however, borrow 903-581 a | Repair after replacement 903-581 (By Ran)
No | Visit | Complaint | RCHN | 报错3998 | 报错3998，借888-654,944-394（13419）缺货 | 更换888-654后修复(By Ran) | SWO-267334 | Defect | HW | Incubator module | Wheel Lock | 3998 | Error 3998 | 3998, 888-654,944-394 (13419) out of stock | Repair after replacement 888-654 (By Ran)
No | Visit | Complaint | RCHN | 进样门无法关闭 | 进样门无法关闭，借903-198一个 | 更换903-198备件后修复(By Ran) | SWO-267275 | Defect | HW | DMS_PC | PCB, Main Board |  | The sample door cannot be closed | The sample door cannot be closed, borrow 903-198 one | Repair after replacement of 903-198 spare parts (By Ran)
No | Visit | Complaint | RCHN | 报错1020 | 报错1020，借备件903-138一个 | 备件未使用，插拔排线后修复Spare parts were not used.(By Ran) | SWO-267192 | Defect | HW | Incubator module | PCB, Incubator | 1020 | 1020 | Error 1020, borrowed spare parts 903-138 one | Spare parts are not in use and repair spare parts parts were not used after wiring. (By Ran)
No | Visit | Complaint | RCHN | 报错5784 | 报错5784，借944-214（批号13410）一个 | 重新定标正常。 Recal ,Then OK. | SWO-267991 | Defect | Consumables | Test Kit | Cartridge, Myo | Broken | 5784 | 5784, borrow 944-214 (lot number 13410) one | Recalibration is normal. Recal, Then OK.
No | Visit | Complaint | RCHN | 质控不在耙值范围内 | 质控不在耙值范围内，借903-364一个 | 更换903-364后修复 | SWO-268028 | Defect | HW | Process Module | Optical Unit |  | Quality control is not within the range of values | The control is not in the range of the value, borrow 903-364 one | Repair after replacement 903-364
No | Visit | Complaint | RCHN | Pushout无法固定 | Pushout无法固定，借886-664一个 |  | SWO-268638 | Defect | HW | Incubator module | Incubator Lit |  | Pushout can't be fixed | Pushout can't be fixed, borrow 886-664 a
No | Visit | Complaint | RSCH | Rep ausdruck | Ausdruck mit Langem Zeilenabstand. Tritt nur bei Arterieller Messung die Fehlerhaft ist auf. | Andere Konfiguration installiert und Arterielle Testmessung durchgeführt. Fehler immer noch vorhanden. Fehler kann reproduziert werden. Messung Nr. 96 | SWO-268846 | Setup | Printout | Printout | Patient Layout |  | Rep print | Expression with long line spacing. If only with arterial measurement the fault yuncturn is occurs. | Other configuration installed and arterial test measurement performed. Errors still present. Errors can be reproduced. Measurement No. 96
No | Visit | Complaint | RKK | PCB一部不良 | PCB一部不良 | PCB交換 | SWO-268769 | Defect | HW | Storage Module | PCB, Cup Transport |  | A bad PCB | A bad PCB | PCB switch
No | 4 | Y | RPOL | Blank | Usuniêcie awarii w module Inlet. Czyszczenie selektora przep³ywów. Analizator sprawny |  | D38BA71EF6516E8E | Adjustment | HW | Auto Inlet module | Inlet Position Cal. |  | Blank | Fault removal in the Inlet module. Cleaning the overflow selector. Functional analyzer
No | Visit | Complaint | RCHN | 提示温度不稳定1302 | 提示温度不稳定1302 | 更换902-972一个正常。(By Ran) | SWO-269207 | [5 empty] | Tip Temperature Instability 1302 | Tip Temperature Instability 1302 | Replace 902-972 one normal. (By Ran)
Yes | Visit | Complaint | RCHN | 血氧质控失败 | 血氧质控失败，借903-197和902-980各一个 | 更换903-197一个正常。(By Ran) | SWO-268652 | Re classified RMED | RCHN Dealer |  |  |  | Failed in blood oxygen control | Blood oxygen control failure, borrow 903-197 and 902-980 each one | Replace 903-197 one normal. (By Ran)
No | Visit | Complaint | RCHN | 干燥前温度超范围 1068 | 干燥前温度超范围 1068，借903-210一个。 |  | SWO-269204 | Defect | HW | Process Module | PCB, Cup Dryer | 1068 | Over-range temperature before drying 1068 | The pre-drying temperature exceeds the range of 1068, borrowing 903-210 one.
No | Visit | Complaint | RCHN | 做样本时打问号 | 做样本时打问号，借903-329一个。 | 更换903-329一个正常。 | SWO-269091 | Defect | HW | Process Module | Needle Wash Unit |  | Mark the question mark when you take the sample. | Ask a question mark when you take the sample and borrow 903-329. | Replace the 903-329 one normal.
No | Visit | Complaint | RBNL | Printer stuk | T.g.v. foutieve manipulatie is de binnenkant (rolgedeelte) afgebroken. | Omdat het onderdeel niet te herstellen viel werd de volledige printer vervangen : test OK | SWO-269836 | Defect | HW | Printer Module | Printer Module |  | Printer piece | T.G.V. faulty manipulation, the inside (roller part) is broken down. | Because the component failed to repair the entire printer was replaced: Test OK
No | Visit | Complaint | RAME | The analyzer will not transmit patient demographic to the LIS | Analyzer transmits but Sunquest system does not send demographics to analyzer. | Modified the ini file to include ETXETBClassic=False code. Performed transmission test, and the analyzer still does not receive demographics from the LIS. The analyzer still will not communicate with the LIS. | SWO-269969 | Revisit required | His/Lis | His/Lis | His/Lis |  | The analyzer will not transmit patient demographic to the LIS | Analyzer transmits but Sunquest system does not send demographics to analyzer. | Modified the ini file to include ETXETBClassic=False code. Performed transmission test, and the analyzer still does not receive demographics from the LIS. The analyzer still will not communicate with the LIS.
No | Visit | Complaint | RLTD | Contact Joyce Tel 01689 864839 | On arrival, screen stuck in inlet gasket replacement and unable to progress, so rebooted analyser. On restart, checked activity log and a lot of 1391 errors present (Inlet is not in the correct position). Analyser restarted without errors. Removed and inspected inlet probe - probe has been bent and straightened again at some point, leaving a small kink. Tested movement of inlet gasket over probe and is ok. Tested inlet within service program, working correctly. Analyser operation restarted - all ok. Performed due tHb calibration - ok. Recommended to purchase spare inlet probe part no. 924-455 price £99.70 plus VAT. | Rebooted analyser software. Further straightened probe. | SWO-269826 | Adjustment | SW | Auto Inlet module | Inlet Position Cal. |  | Contact Joyce Tel 01689 864839 | On arrival, screen stuck in inlet gasket replacement and unable to progress, so rebooted analyser. On restart, checked activity log and a lot of 1391 errors present (Inlet is not in the correct position). Analyser restarted without errors. Removed and inspected inlet probe - probe has been bent and straightened again at some point, leaving a small kink. Tested movement of inlet gasket over probe and is ok. Tested inlet within service program, working correctly. Analyser operation restarted - all ok. Performed due tHb calibration - ok. Recommended to purchase spare inlet probe part no. 924-455 price £99.70 plus VAT. | Rebooted analyser software. Further straightened probe.
No | Visit | Complaint | RLTD | Always failing Glucose across 3 different sensor cassettes. | 1310 response error. | Pin alignment looks good, sensor cassettes from different lots. Will have to return and replace sensor interface PCB. | SWO-269786 | Revisit required | HW | Basic Unit | PCB, Sensor interface | 1310 | Always failing Glucose across 3 different sensor cassettes. | 1310 response error. | Pin alignment looks good, sensor cassettes from different lots. Will have to return and replace sensor interface PCB.
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | CL定标不过 | 更换测试卡后修复 | SWO-273201 | [5 empty] | Test card, reagent package failure | CL calibration is not | Repair after replacing test card
No | Visit | Complaint | RCHN | 胆红素测试不准 | 胆红素测试不准，借902-980一个 |  | SWO-269381 | Defect | HW | Oxi module | Hemolyzer | ctBill | The bilirubin test is not allowed. | Bilic erythema test is not allowed, borrow 902-980 one
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | PCO2,CL定标失败 | 更换测试卡后修复 | SWO-269514 | Assisted Replacement | SC | SC | pCO2 |  | Test card, reagent package failure | PCO2, CL calibration failed | Repair after replacing test card
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | GLU,LAC定标失败 | 更换测试卡后修复 | SWO-269522 | Assisted Replacement | SC | SC | Glu/Lac |  | Test card, reagent package failure | GLU, LAC calibration failed | Repair after replacing test card
No | Visit | Complaint | RESP | 1221 temperatura del sistema fuera de rango. | 1221 temperatura del sistema fuera de rango. Sci defectuoso | Se reviso el equipo y se detecto el sci en mal estado. Se sustituyo sci y se comprobó correcto funcionamiento del equipo mediante QC. | SWO-273982 | Assisted Replacement | SC | SC | Temp |  | 1221 out-of-range system temperature. | 1221 out-of-range system temperature. Faulty Sci | The equipment was checked and the sci was unhealthy. Sci was replaced and the equipment was tested for proper operation by QC.
Yes | 13 | Y | RPAC | Blank | FAN915-318R5 |  | D381A00F462F0B1A | Re classified RMED | FAN |  |  |  | Blank | FAN915-318R5
Yes | Visit | Complaint | RAME | Auto check mechanical error | Tried to run QC and auto check raises up but did not run the QC, code 0742, mechanical error | Replaced the auto check flex cable Alignment in service | SWO-269999 | [5 empty] | Auto check mechanical error | Tried to run QC and auto check raises up but did not run the QC, code 0742, mechanical error | Replaced the auto check flex cable Alignment in service
Yes | Visit | Complaint | RLTD | Creatinine failing low on QC for all levels excluding high Crea check | The QC logs were examined and Crea has been failing low on all levels of QC, (not including high Crea check). The Crea membranes were replaced on the 20th and 24th November. A QC measurement was run and the fluid transport was viewed, no fluid transport errors were witnessed. As error 955 was listed in the messages for the failed QC measurements, the Cal 2 bottle was replaced. Prior to running the calibration, the Crea measuring chambers were inspected and the tips of the Crea electrodes were cleaned. On running QC, Crea once again failed with low measurements. | New Crea electrodes ordered and are being sent FAO Rosie Steele, Clinical Measurements, John Radcliffe Hospital | SWO-269833 | [5 empty] | Creatinine failing low on QC for all levels excluding high Crea check | The QC logs were examined and Crea has been failing low on all levels of QC, (not including high Crea check). The Crea membranes were replaced on the 20th and 24th November. A QC measurement was run and the fluid transport was viewed, no fluid transport errors were witnessed. As error 955 was listed in the messages for the failed QC measurements, the Cal 2 bottle was replaced. Prior to running the calibration, the Crea measuring chambers were inspected and the tips of the Crea electrodes were cleaned. On running QC, Crea once again failed with low measurements. | New Crea electrodes ordered and are being sent FAO Rosie Steele, Clinical Measurements, John Radcliffe Hospital
No | Visit | Complaint | RAME | Customer states that the barcode reader is no longer scanning labels. Customer opted not to phone troubleshoot as analyzer is due for 6 month pump tubing replacement and VK-R7 and would like this completed by Radiometer | Barcode scanner is observed to have no beam visible. | Upon reboot of analyzer, barcode scanner again functions normally. Also completed the pump tubing replacement as a courtesy to the customer. VK-R7 is refrigerated and can not be completed. Explained that a training visit can be scheduled to instruct on pump tubing replacement and running of the VK-R7. Analyzer is within acceptable operating specifications. | SWO-269911 | Defect | SW | Windows | SW Barcode Reader | Reboot | Customer states that the barcode reader is no longer scanning labels. Customer opted not to phone troubleshoot as analyzer is due for 6 month pump tubing replacement and VK-R7 and would like this completed by Radiometer | Barcode scanner is observed to have no beam visible. | Upon reboot of analyzer, barcode scanner again functions normally. Also completed the pump tubing replacement as a courtesy to the customer. VK-R7 is refrigerated and can not be completed. Explained that a training visit can be scheduled to instruct on pump tubing replacement and running of the VK-R7. Analyzer is within acceptable operating specifications.
No | Visit | Complaint | RDEU | Probleme mit Bildschirmbedienung. | Bildschirm Tausch. | Neuen Bildschirm verbaut. Bildschirm Firmware aktualisiert. Bildschirm kalibriert. Bildschirmschonereinstellungen überprüft. Mehrere Kalibrierungen durchgeführt. QCs für alle Level durchgeführt. ABL wieder messbereit. | SWO-269897 | Defect | HW | DMS_PC | Touch screen | Find Revisit | Problems with screen operation. | Screen swap. | New screen installed. Screen firmware updated. screen calibrated. Screen saver settings checked. Multiple calibrations performed. QCs for all levels. ABL ready to be measured again.
No | Visit | Complaint | RDEU | 27.11.2017 Herr Theel 0581836051 meldet einen Vaccum Fehler 27.11.2017, 14.45 Uhr, Herr Theel, Tel: 0581836051: Drucküberprüfung Vakuumfehler | System überprüft => schwarze Dichtung in der Einlaßdichtung nicht richtig positioniert => defekt | Einlassdichtung erneuert, Kalibration und QC durchgeführt | SWO-270407 | Defect | HW | Inlet module | Inlet gasket |  | 27.11.2017 Mr Theel 0581836051 reports a Vaccum error 27.11.2017, 14.45, Mr. Theel, Tel: 0581836051: Pressure check vacuum error | System checked => black seal in the inlet seal not positioned correctly => defective | Inlet seal renewed, calibration and QC performed
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 试剂包验证失败 | 连续更换2个试剂包和1张测试卡后修复 | SWO-276600 | [5 empty] | Test card, reagent package failure | Reagent pack validation failed | Repair after 2 reagent kits and 1 test card in a row
No | Visit | Complaint | RCHN | 无法读取试剂包智能芯片 | 无法读取试剂包智能芯片，借903-158一个 | 更换903-158后修复 by jesson | SWO-269466 | Defect | HW | Basic Unit | PCB, Solution Pack IF |  | Unable to read the reagent package smart chip | Unable to read reagent package smart chip, borrow 903-158 one | Repair by jesson after replacement 903-158
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 温度超范围，借903-581一个 | 更换903-581备件后修复 by jesson | SWO-269374 | Defect | HW | Basic Unit | PCB, Sensor interface | Temp | Test card, reagent package failure | Temperature over range, borrow 903-581 one | Repair by jesson after replacement of 903-581 spare parts
No | Visit | Complaint | RCHN | 报1024错误 | 报1024错误 | 更换备用903-364后修复 by Jesson | SWO-271012 | Defect | HW | Process Module | Optical Unit | 1024 | Reported 1024 error | Reported 1024 error | Repair by Jesson after replacing spare 903-364
No | Visit | Complaint | RCHN | 报错8130 | 报错8130，借备件903-329一个 | 更换903-329后修复 by jesson | SWO-269500 | Defect | HW | Process Module | Needle Wash Unit | 8130 | 8130 | 8130, spare parts 903-329 one | Repair by jesson after replacing 903-329
No | Visit | Complaint | RIND | analyzer showing frequently error neon number too high. | checked the oxi module and conformed oximetry having neon number too error. | performed complete oximetry test . oximetry test completed successfully with out any error. but after three hours again same error repeated neon number too high . we need to replace oximetry module for this case. | SWO-271801 | Revisit required | HW | Oxi module | Oxi module | Neon Lamp | analyzer showing frequently error neon number too high. | checked the oxi module and conformed oximetry having neon number too error. | performed complete oximetry test . oximetry test completed successfully with out any error. but after three hours again same error repeated neon number too high . we need to replace oximetry module for this case.
No | Visit | Complaint | RKK | Gas2の減りが著しく早い(約１か月) | ガス２の減りが早い | フローレギュレーターガス２交換 レギュレーター接続チューブ交換 Ｃａｌ２ ＱＣ ＯＫ | SWO-270167 | Defect | HW | Gas module | Flow regulator | Gas2 | Gas2 decreases significantly faster (about a month) | Gas 2 is decreasing quickly. | Flow regulator gas 2 replacement Regulator connection tube replacement Cal2 QC OK
No | Visit | Complaint | RIND | ? mark error for all samples | sample in ph $bg upper liquid sensor prematurely during aspiration this error was occured | so chekced the gasket and cleaned the gas ket and sample inlet . then performed the pump calibration and liquid sensor calibration after that performed the 2point calibartion , observed the analyser for an hour was working fine | SWO-271423 | Cleaned | Fluidic pathway | pH/BG Module | Measuring chamber |  | ? mark error for all samples | sample in ph /bg upper liquid sensor prematurely during aspiration this error was occured | so chekced the gasket and cleaned the gas ket and sample inlet . then performed the pump calibration and liquid sensor calibration after that performed the 2point calibartion , observed the analyser for an hour was working fine
No | 13 | Y | RPAC | Blank | Radiance connectivity & sample introduce button missing |  | D381DE2A46516FAD | Defect | SW | RMED SW | Expired SW Version |  | Blank | Radiance connectivity & sample introduce button missing
No | Visit | Complaint | RAME | Shown in the Activity Log 1341 - when User-intervention required has been entered due to this reason. 1342 - Shown in the Activity Log when User-intervention required has been entered due to this reason. | Fails oxi parameters SO2 | Replaced hemolyzer unit. Run the calibration and QC. | SWO-270808 | Defect | HW | Oxi module | Hemolyzer |  | Shown in the Activity Log 1341 - when User-intervention required has been entered due to this reason. 1342 - Shown in the Activity Log when User-intervention required has been entered due to this reason. | SO2 Oxy Parameters | Replaced hemolyzer unit. Run the calibration and QC.
No | Visit | Complaint | RAME | Error messages indicate to replace Inlet gasket. Customer replaced inlet gasket and still same issue. SC an SP was also replaced. | Back flushes is very difficult to perform. Replaced tubing and still backflushes is difficult. Debris is stuck in left side of Sensor bay unit and it is difficult to clear. | Ordered sensor bay unit. Part#903-199 | SWO-270816 | Revisit required | HW | Basic Unit | Sensor Bay Unit |  | Error messages indicate to replace Inlet gasket. Customer replaced inlet gasket and still same issue. SC an SP was also replaced. | Back flushes is very difficult to perform. Replaced tubing and still backflushes is difficult. Debris is stuck in left side of Sensor bay unit and it is difficult to clear. | Ordered sensor bay unit. Part#903-199
No | Visit | Complaint | RDEU | 28.11.2017 14:40Uhr Her Porthen 04081914572: Kal-Fehler bei Elektrolyten 29.11.2017, 10:26 Uhr, H. Neumann: Tel. 04081914993: Fehler nach allen besprochenen und durchgeführten Maßnahmen wieder aufgetreten. | System überprüft => Sensorinterface PCB defekt Funktionierende SC aus UKE installiert ohne Erfolg | Sensorinterface PCB bestellt | SWO-271203 | Revisit required | HW | Basic Unit | PCB, Sensor interface |  | Nov 28, 2017 2:40pm Her Porthen 04081914572: Kal fault in electrolytes 29.11.2017, 10:26 a.m., H. Neumann: Tel. 04081914993: Errors reappeared after all measures discussed and implemented. | System checked => Sensor interface PCB defective Functioning SC from UKE installed without success | Sensor interface PCB ordered
No | Visit | Complaint | RDEU | 22.11., Herr Deiml, Tel.: 023729082257: OXI Probleme bei Pat. Messung | Gerät überprüft und gereinigt | Hemolyzer erneuert, Gerät neu gestartet, Startphase abgewartet, tHb Kal, Kal und QCs durchgeführt | SWO-271095 | Defect | HW | Oxi module | Hemolyzer |  | 22.11., Mr. Deiml, Tel.: 023729082257: OXI Problems at Pat. Measurement | Device checked and cleaned | Hemolyzer renewed, device restarted, start-up phase awaited, tHb Kal, Kal and QCs performed
No | Visit | Complaint | RAME | Possible Blockage During Sample Analysis (Primary Liquid Sensor Status) | Checked data logs and verified the problem. Took the wet section out and found that the top tube on the hemolyzer was pulled tight over the top of the guard on the pump. | Trimmed the guard on top of the pump to allow the tube not too be pulled so tight. Ran the cleaning pack. | SWO-271427 | [5 empty] | Possible Blockage During Sample Analysis (Primary Liquid Sensor Status) | Checked data logs and verified the problem. Took the wet section out and found that the top tube on the hemolyzer was pulled tight over the top of the guard on the pump. | Trimmed the guard on top of the pump to allow the tube not too be pulled so tight. Ran the cleaning pack.
No | Visit | Complaint | RCHN | lac定标失败 | lac定标失败，需要借件903-581 | 更换903-581后修复 by jesson | SWO-270253 | Defect | HW | Basic Unit | PCB, Sensor interface |  | lac calibration failed | Lac bid failed, need to borrow 903-581 | Repair by jesson after replacement 903-581
No | Visit | Complaint | RCHN | 无法开机 | 无法开机 | 自行更换备用机910-299后修复 by jesson | SWO-271028 | Defect | HW | DMS_PC | CPU |  | Cannot power on | Cannot power on | Repair by jesson after replacing the spare 910-299
No | Visit | Complaint | RCHN | 报软件有问题 | 报软件有问题，借910-343一个 | 更换910-343备件后修复by tim | SWO-270255 | Defect | HW | DMS_PC | CPU |  | There is a problem with the report software | Reported software problems, borrow 910-343 one | Repair by tim after replacement of 910-343 spare parts
No | Visit | Complaint | RDAN | OXI fejl | OXI test kal. data tjekket | Lampe justeret Service tHb kal. QC | SWO-272453 | Adjustment | HW | Oxi module | Lamp unit |  | OXY Headl | OXI test kal. data tjekket | Lampe justeret Service tHb kal. QC
No | Visit | Complaint | RCHN | 条码扫码器无法扫描 | 条码扫码器无法扫描，需要借件902-822备件一个 | 更换902-822后修复 by jesson | SWO-270187 | Defect | HW | Basic Unit | Barcode reader, Basic Unit |  | Barcode scanner cannot be scanned | Barcode scanner cannot be scanned and requires a loan of 902-822 spare parts | Repair by jesson after replacement of 902-822
No | 4 | Y | RCZE | Blank | selhávání vzorkù, zejména kapilárních, chyba pumpy |  | D3850D38A852565C | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | Blank | Failure of samples, especially capillary, pump error
No | Visit | Complaint | RDEU | pO2 Empfindlichkeitsfehler | pO2 Empfindlichkeits fehler | Elektroden Membranierung und Bürsten ist bereits vom Anwender durchgeführt. Elektrode ausgetauschen. | SWO-271284 | Assisted Replacement | Electrode | Electrode | pO2 |  | pO2 Empfindlichkeitsfehler | pO2 Sensitivity error | Electrode membraneing and brushing is already carried out by the user. electrode.
No | Visit | Complaint | RDEU | pH/BG Pumpenkalibrationsfehler | Gerät überprüft und gereinigt, Messkammern gereinigt, | Pumpenschlauch erneuert, Kals und QCs durchgeführt | SWO-271137 | Defect | HW | pH/BG Module | Pump tubes |  | pH/BG Pumpenkalibrationsfehler | device checked and cleaned, measuring chambers cleaned, | Pump hose renewed, kals and QCs performed
No | Visit | Complaint | RAME | El/Met lower LS error. | Inspection of the El $Met module revealed the El $Met lower LS is failing service LS test. There is a leak at the lower El $Met module grey tubing 841-775. Getting lamp halogen intensity too high. | Adjusted lamp intensity. Replaced lower El $Met LS and all electrode grey module tubing from a PM kit. | SWO-270844 | Defect | HW | El/Met Module | Tubes |  | El/Met lower LS error. | Inspection of the El /Met module revealed the El /Met lower LS is failing service LS test. There is a leak at the lower El /Met module grey tubing 841-775. Getting lamp halogen intensity too high. | Adjusted lamp intensity. Replaced lower El $Met LS and all electrode grey module tubing from a PM kit.
Yes | 5 | Y | RPAC | Blank | Solution Pack could not be read by analyser. |  | D3831226BC32CCB1 | [5 empty] | Blank | Solution Pack could not be read by analyser.
No | Visit | Complaint | RSAF | Customer failing to load cartridges on the analyzer | Follow up | Checked the activity log to verify that the were no errors. Inserted test cartridges successfully | SWO-276050 | Revisit required | HW | Storage Module | Barcode reader, Storage wheel |  | Customer failing to load cartridges on the analyzer | Follow up | Checked the activity log to verify that the were no errors. Inserted test cartridges successfully
No | Visit | Complaint | RAME | Error messages indicate to replace Inlet gasket. Customer replaced inlet gasket and still same issue. SC an SP was also replaced. | Replaced sensor cassette bay unit that was not available on previous visit. | Replaced sensor cassette bay unit. Adjusted unit in diagnostic. Run thb calibration. | SWO-272145 | Defect | HW | Basic Unit | Sensor Bay Unit |  | Error messages indicate to replace Inlet gasket. Customer replaced inlet gasket and still same issue. SC an SP was also replaced. | Replaced sensor cassette bay unit that was not available on previous visit. | Replaced sensor cassette bay unit. Adjusted unit in diagnostic. Run thb calibration.
Yes | Visit | Complaint | RSAF | Pressure Vacuum error | Pressure Vacuum error | Inspected the pinch valve and the pinch tubing valve to check if the are still functional. Replaced with a new inlet gasket. | SWO-276069 | [5 empty] | Pressure Vacuum error | Pressure Vacuum error | Inspected the pinch valve and the pinch tubing valve to check if the are still functional. Replaced with a new inlet gasket.
No | Visit | Complaint | RBNL | Accu defect | Accu is defect.. Reset van de accu lost het probl;eem niet op. In het serviceprogramma worden alle cellen met het juiste voltage weergegeven. Zodoende accu vervangen. Serienummer defecte accu: P0134N007 Opmerking: accu valt nog onder garantie. | Accu vervangen. | SWO-271800 | Defect | HW | Battery module | Battery |  | Accu defect | Battery is defective.. Battery Reset does not solve the probl; Eem. The service program displays all cells with the correct voltage. Replace Battery. Serial number Defective battery: P0134N007 Note: battery is still under warranty. | Replace the battery.
No | Visit | Complaint | RDEU | Fehler: Einlassdichtunghalter | Fehler: Einlassdichtunghalter | Einlasssonde erneuert, Einlassdichtunghalter erneuert, Jahreswartung durchgeführt, ABL manuell gespült. Kalibrationen und QC's vorgenommen. BGA ok. | SWO-272137 | Defect | HW | Auto Inlet module | Inlet gasket |  | Error: Inlet seal holder | Error: Inlet seal holder | Inlet probe renewed, inlet seal holder renewed, annual maintenance carried out, ABL manually rinsed. calibrations and QC's. BGA ok.
Yes | 4 | Y | RPAC | Blank | Error 0964. leak current detected. |  | D3844F6C3201B868 | [5 empty] | Blank | Error 0964. leak current detected.
No | Visit | Complaint | RSAF | Hold error | Inlet error Unable to pump solution. QC 3 error Inlet gasket replaced by technologist | Remove inlet gasket with needle. Readjust wheel position. Refit gasket. Refit inlet gasket. Perform calibration and QC. Show them how to replace the inlet needle and gasket . | SWO-272382 | Adjustment | HW | Auto Inlet module | Inlet Position Cal. |  | Hold error | Inlet error Unable to pump solution. QC 3 error Inlet gasket replaced by technologist | Remove inlet gasket with needle. Readjust wheel position. Refit gasket. Refit inlet gasket. Perform calibration and QC. Show them how to replace the inlet needle and gasket .
No | Visit | Complaint | RLTD | Always failing Glucose across 3 different sensor cassettes. Update Failing Sodium and Glucose please contact Rhiannon Tel 01604 544080 | 1310 response error on glucose channel. | After receiving latest technical information from senior service engineer replaced sensor interface PCB and replaced sensor cassette. Now ok. | SWO-273701 | Defect | HW | Basic Unit | PCB, Sensor interface |  | Always failing Glucose across 3 different sensor cassettes. Update Failing Sodium and Glucose please contact Rhiannon Tel 01604 544080 | 1310 response error on glucose channel. | After receiving latest technical information from senior service engineer replaced sensor interface PCB and replaced sensor cassette. Now ok.
No | Visit | Complaint | RLTD | Mark Spencer Tel 01604 545542 | Screen doing random things, even with keyboard fitted. | Reloaded software and the issue remained. Will return tomorrow with new screen | SWO-272827 | Revisit required | HW | DMS_PC | Touch screen |  | Mark Spencer Tel 01604 545542 | Screen doing random things, even with keyboard fitted. | Reloaded software and the issue remained. Will return tomorrow with new screen
Yes | 2 | Y | RPOL | Blank | Wymiana Sensor Bay Unit. Analizator sprawny technicznie. |  | D39B5B5A7E433981 | [5 empty] | Blank | Replacing the Sensor Bay Unit. Technically efficient analyzer.
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | PCO2,PO2质控失败 | 更换测试卡后修复 | SWO-276598 | [5 empty] | Test card, reagent package failure | PCO2, PO2 QC Failure | Repair after replacing test card
No | Visit | Complaint | RCHN | Backlog error | backlog error | valve tubing has been renewed | SWO-271678 | Defect | HW | Oxi module | Pinch valve tube |  | Backlog error | backlog error | valve tubing has been renewed
No | Visit | Complaint | RCHN | GLU响应错误 | GLU响应错误，借903-581一个 | 更换903-581后修复 by jesson | SWO-270866 | Defect | HW | Basic Unit | PCB, Sensor interface |  | GLU response error | GLU response error, borrow 903-581 one | Repair by jesson after replacement 903-581
No | Visit | Complaint | RCHN | 定标数据丢失 | 定标数据丢失，借944-214（13002）缺货、944-258（13238）、944-220（12568）各一张 | 由于定标数据丢重新定标 后修复。 更换944-258（13238）、944-220（12568）各一张后修复 by jesson | SWO-271045 | Assisted Replacement | Consumables | Test Kit | Cartridge, NT-PROBNP |  | Target data loss | Bid data loss, out of stock 944-214 (13002), 944-258 (13238), 944-220 (12568) each | Fix because the calibration data is dropped and re-calibrated. Replace 944-258 (13238), 944-220 (12568) each after repairing by jesson
No | Visit | Complaint | RCHN | 报1267错误 | 报1267错误，借903-329一个 | 更换903-329备件后修复 by jesson | SWO-270906 | Defect | HW | Process Module | Needle Wash Unit | 1267 | Reported 1267 errors | Reported 1267 error, borrowed 903-329 one | Repair by jesson after replacement of 903-329 spare parts
Yes | Visit | Complaint | RLTD | Peter Tandy emailed to report the fault. | Since the service was performed on the analyser there have been repeated Hemolyzer temp errors and it has also been failing QC due to spectrum mismatch errors | New Hemolyzer fitted to the Oxi module. Lamp unit adjustment performed and complete Oxi test run and passed. | SWO-271818 | [5 empty] | Peter Tandy emailed to report the fault. | Since the service was performed on the analyser there have been repeated Hemolyzer temp errors and it has also been failing QC due to spectrum mismatch errors | New Hemolyzer fitted to the Oxi module. Lamp unit adjustment performed and complete Oxi test run and passed.
Yes | Visit | Complaint | RLTD | Creatinine failing low on QC for all levels excluding high Crea check | Investigating further to see if the cause of the Crea QC failure can be found. The dummy electrodes in the Crea module were removed and the chambers cleaned but it still failed QC. The Crea electrodes were removed and inspected and it was found that there is tarnishing on the contacts at the top of the electrodes, the current advice from Copenhagen is that if this tarnishing is present to replace the electrodes with new ones. | New electrodes were ordered for this analyser on Monday but they have not been brought up to it yet, they should be being fitted tomorrow. | SWO-271939 | [5 empty] | Creatinine failing low on QC for all levels excluding high Crea check | Investigating further to see if the cause of the Crea QC failure can be found. The dummy electrodes in the Crea module were removed and the chambers cleaned but it still failed QC. The Crea electrodes were removed and inspected and it was found that there is tarnishing on the contacts at the top of the electrodes, the current advice from Copenhagen is that if this tarnishing is present to replace the electrodes with new ones. | New electrodes were ordered for this analyser on Monday but they have not been brought up to it yet, they should be being fitted tomorrow.
Yes | Visit | Complaint | RDAN | Studs under inletpakning er knækket. Pernille ønsker tekniker til at skifte den knækket studs. | Spildstuds til kapillærrør er knækket | Jeg har vejledt i hvordan man skifter spildstuds. Spildstuds er skiftet. Alt er OK | SWO-271938 | [5 empty] | Studs under the inlets are broken. Pernille wants technician to change the broken studs. | Waste studs for capillary tubes are broken | I have been guided in how to change the waste studs. Waste studs are switched. Everything is OK
No | Visit | Complaint | RDEU | Flüssigkeit tritt am Inlet aus | Tubing Nipple am Einlassschlauch angebrochen | Tubing Nipple Kit eingebaut. | SWO-272446 | Defect | HW | Inlet module | Nipple kit |  | Liquid leaks at the inlet | Tubing Nipple am Einlassschlauch angebrochen | Tubing Nipple Kit eingebaut.
No | Visit | Complaint | RCHN | 报错0258 | 报错0258，借902-833一个 | 更换902-833一个 by jesson | SWO-270864 | Defect | HW | Oxi module | Oxi module, Replaced | Neon Low | Error 0258 | 0258, loan 902-833 one | Replace 902-833 a by jesson
No | Visit | Complaint | RIND | Undefined Lookup String. 35.16.1.0-1-1 | Undefined Lookup String. 35.16.1.0-1-1 | Found error, check the Service programs and performed the safe position it was OK, check the all service programs, but problem was remaining same. need to replace Incubator PCB and check. | SWO-272785 | Revisit required | Unknown | Unknown | Unknown | JGF Communicate | Undefined Lookup String. 35.16.1.0-1-1 | Undefined Lookup String. 35.16.1.0-1-1 | Found error, check the Service programs and performed the safe position it was OK, check the all service programs, but problem was remaining same. need to replace Incubator PCB and check.
Yes | Visit | Complaint | RDEU | Oxi Spectrum Mismatch Nacharbeit zu SWO 271191 | Nacharbeit zu SWO 271191 | Oxi Spectrum Mismatch, Schlauchwege ok, Pinch Valve ok, Hemolyzer Unit defekt->erneuert, Startphase tHb Kal. und Kal. jetzt wieder ok ohne Fehler System-und Funktionskontrolle i.O. | SWO-272650 | [5 empty] | Oxi Spectrum Mismatch Rework on SWO 271191 | Rework on SWO 271191 | Oxi Spectrum Mismatch, Hose Paths ok, Pinch Valve ok, Hemolyzer Unit defective->renewed, Start phase tHb Kal. and Kal. now ok again without errors System and function control in the
No | Visit | Complaint | RDEU | 28.11.2017 14:40Uhr Her Porthen 04081914572: Kal-Fehler bei Elektrolyten 29.11.2017, 10:26 Uhr, H. Neumann: Tel. 04081914993: Fehler nach allen besprochenen und durchgeführten Maßnahmen wieder aufgetreten. | Trotz Tausch Sc und LP immer wieder Fehlermeldung | Sensor Bord getauscht. Danach gleiche Fehlermeldung. Verbrauchsmaterial erneut getauscht. Und Fehlersuche durchgeführt. Analysator mehrfach gespült. Danach erneut mehrere Kalibrationen durchgeführt. Analysator wieder einsatzbereit. | SWO-272570 | Defect | HW | Basic Unit | PCB, Sensor interface |  | Nov 28, 2017 2:40pm Her Porthen 04081914572: Kal fault in electrolytes 29.11.2017, 10:26 a.m., H. Neumann: Tel. 04081914993: Errors reappeared after all measures discussed and implemented. | Despite swapping Sc and LP again and again error message | Sensor board exchanged. Then the same error message. consumables are exchanged again. And troubleshooting performed. Analyzer rinsed several times. Then several calibrations were carried out again. Analyzer ready for use again.
No | Visit | Complaint | RDEU | 30.11.2017 16:30Uhr Frau Trote 059311523970: Der Drucker arbeitet nicht | Fehler : 1297 Drucker offline | Gerät demontiert , Drucker und Verkabelung überprüft. Drucker auf andere USB Schnittstelle gesteckt und Analysator neu gestartet. Alles ohne Erfolg! Software neu aufgespielt , - alles ohne Erfolg. Gerät wieder zusammengebaut und Startphase abgewartet. Service abgebrochen. - Kunde besteht noch heute auf ein voll funktionsfähiges Gerät! ABL 90 ( Neugerät )aus dem OP geholt und Drucker umgebaut,- Fehler bleibt. PC Board defekt. OP Gerät für den Kreissaal fertig gemacht. CF-Karte ,SC und LP getauscht. | SWO-272768 | Revisit | HW | DMS_PC | Printer | 1297 | Nov 30, 2017 4:30 pm Ms Trote 059311523970: The printer is not working | Error : 1297 printer offline | Device dismantled, printer and wiring checked. Printer plugged into other USB interface and analyzer restarted. All without success! Software replayed , - all without success. reassembled and start-up phase awaited. Service aborted. - Customer insists on a fully functional device today! ABL 90 (new device )taken out of the operating room and printer converted,- error remains. PC board defective. SURGICAL device for the district hall finished. CF card ,SC and LP exchanged.
Yes | 8 | Y | RPAC | Blank | Various Solution Pack issues with the analyser. |  | D3857453A802F882 | [5 empty] | Blank | Various Solution Pack issues with the analyser.
No | Visit | Complaint | RIND | Checked with Zeroing Cassete and found PO2 values are in abnormal range. | Contionous problem in Po2, checked with zeroing cassete and found po2 values showing abnormal range. | Signal data board need to be change. | SWO-272541 | [5 empty] | Checked with Zeroing Cassete and found PO2 values are in abnormal range. | Contionous problem in Po2, checked with zeroing cassete and found po2 values showing abnormal range. | Signal data board need to be change.
No | Visit | Complaint | RCHN | 打印机不工作 | 打印机不工作 | 自行更换903-001后修复 by jesson | SWO-272415 | [5 empty] | The printer does not work | The printer does not work | Fix by jesson after replacing 903-001
No | Visit | Complaint | RCHN | LAC响应错误 | LAC响应错误，借903-581一个 | 更换903-581后修复 by jesson | SWO-271732 | Defect | HW | Basic Unit | PCB, Sensor interface |  | LAC response error | LAC response error, borrow 903-581 one | Repair by jesson after replacement 903-581
No | Visit | Complaint | RCHN | 试管夹破损 | 试管夹破损，需要借件903-584两个 | 更换903-584两个后修复 | SWO-271797 | Defect | HW | Cover and Chassis | Tube holder |  | The test tube clip is broken | The test tube clip is broken and requires a loan of 903-584 two | Replace 903-584 two post-repair
Yes | Visit | Complaint | RCHN | 定标不过 | 定标不过，借903-364,903-236,944-212（13499），944-216（13434（缺货）），944-214（13492），944-258（13313缺货），944-220（13435） | 更换备用机配件 903-364，903-236.重新定标后修复 by Jesson | SWO-270915 | [5 empty] | The calibration is not the | However, 903-364,903-236,944-212 (13499), 944-216 (13434 (out of stock),944-214 (13492), 944-258 (13313 out of stock), 944-220 (13435) | Replace spare parts 903-364,903-236.fix by Jesson after re-calibration
Yes | Visit | Complaint | RCHN | 处理盘卡盘 | 处理盘卡盘，借944-212（13607,13551）各一张 | 重新定标后修复 by jesson | SWO-271661 | Cleaned | HW | Storage Module | Barcode reader, Storage wheel |  | Handling disks | Process the disk, borrow 944-212 (13607, 13551) each one | Fix by jesson after re-calibration
No | Visit | Complaint | RCHN | 报错1999 | 报错1999，借636-541一个 | 更换636-541后修复 by jesson | SWO-270889 | Defect | HW | Analyzer Control Module | Cable, Pipette | 1999 | Error 1999 | Misquoted 1999, loan 636-541 one | Repair by jesson after replacement 636-541
Yes | Visit | Complaint | RDEU | AC Ampullen werden nicht geöffnet | AC Ampullen werden nicht geöffnet | Gerät demontiert, AC Modul geprüft und instand gesetzt, Funktionskontrolle ok | SWO-274107 | [5 empty] | AC ampoules will not be opened | AC ampoules will not be opened | Device dismantled, AC module tested and repaired, function control ok
No | Visit | Complaint | RFRA | Le CV% est trop important en comparaison avec les 2 autres analyseurs . Le contrôle AC5+ sur Na+ est parfois en dessous ou au dessus des bornes fournisseur | Verification historique AC5+ | Remplacement électrode de référence Remembranage référence Modification fréquence de calibration 2 pts ( 8 heures passé à 2 Heures ) durant la nuit | SWO-272892 | Assisted Replacement | Electrode | Electrode | Reference |  | The CV% is too important in comparison with the other 2 analyzers. The AC5 control on Na is sometimes below or above the supplier terminals | AC5 Historical Verification | Reference electrode replacement Reference reembranage Change calibration frequency 2 pts (8 hours spent at 2 Hours) during the night
No | Visit | Complaint | RFRA | Autocheck planifié non excuté du fait erreur | Test de l'autocheck ( Marteau cassé ) Lors du démontage de l'autocheck problème fixation nappe Autocheck PCB $Laser Test complet du module autocheck OK Alignement Laser OK Alignement site d'introduction OK | Remplacement du marteau Remplacement de Autocheck module PCB Install AUEngine Relance analyseur Calibration +passage AC5+ | SWO-272954 | Defect | HW | Auto Inlet module | Breaker arm |  | Planned autocheck not executed by error | Autocheck test (Broken hammer) When dismantling the autocheck problem fixing Autocheck PCB /Laser Full test of the AUtocheck OK module Laser Alignment OK Alignment introducing site OK | Replacing the hammer Replacement of Autocheck MODULE PCB Install Auengine Relaunch analyzer Calibration -AC5 pass
No | Visit | Complaint | RAME | Inlet leaking fluid onto autocheck lid. | Random dripping occurs from below the inlet gasket. | All lines appreciated to be clear, valves functioning, and electrodes are sealing ok. Replaced inlet waste nipples and inlet gasket. After numerous wet section activities, no sign of leaks are detected. | SWO-272660 | Defect | HW | Inlet module | Nipple kit |  | Inlet leaking fluid onto autocheck lid. | Random dripping occurs from below the inlet gasket. | All lines appreciated to be clear, valves functioning, and electrodes are sealing ok. Replaced inlet waste nipples and inlet gasket. After numerous wet section activities, no sign of leaks are detected.
No | Visit | Complaint | RIND | Creatinine sensitivity error. | Cal 1 & Cleaning solution got expired. | Replaced the CAL1 & cleaning solution,run the 2 point calibration few times,but still the error still persist,replace the crea A & CreaB membrane & also the Cal2 solution,then run the calibration few times & it got passed successfully. | SWO-272847 | User error | Consumables | Crea Solution exp | Crea Solution exp |  | Creatinine sensitivity error. | Cal 1 & Cleaning solution got expired. | Replaced the CAL1 & cleaning solution,run the 2 point calibration few times,but still the error still persist,replace the crea A & CreaB membrane & also the Cal2 solution,then run the calibration few times & it got passed successfully.
No | Visit | Complaint | RESP | Error 3414 | Se detecta needle torcida | Sustitución de ,needle y pruebas de ajustes y funcionamiento. Calibración de parámetros y q.c | SWO-272588 | Defect | HW | Aspiration Module | Sensor, Needle | 3414 | Error 3414 | If it detects twisted needle | Substitution of ,needle and tests of adjustments and operation. Parameter calibration and q.c
No | Visit | Complaint | RIND | Undefined Lookup String. 35.16.1.0-1-1 | Undefined Lookup String. 35.16.1.0-1-1 | Performed the Cup transportation but it was failed, replaced following PCBs, Incubator PCB, storage wheel PCB, Solution Pack PCB, but problem was remaining same, need to replace with Main Board and check waiting for spare. | SWO-272763 | Revisit required | Unknown | Unknown | Unknown |  | Undefined Lookup String. 35.16.1.0-1-1 | Undefined Lookup String. 35.16.1.0-1-1 | Performed the Cup transportation but it was failed, replaced following PCBs, Incubator PCB, storage wheel PCB, Solution Pack PCB, but problem was remaining same, need to replace with Main Board and check waiting for spare.
Yes | Visit | Complaint | RSAF | Move the analyzer , Tnt giving false readings | Nove the analyzer to a new location, TnT reading high | Changed the units for tnt, moved the analyzer to a new location | SWO-277470 | Re classified RMED | Re location |  |  |  | Move the analyzer , Tnt giving false readings | Nove the analyzer to a new location, TnT reading high | Changed the units for tnt, moved the analyzer to a new location
No | Visit | Complaint | RLTD | Mark Spencer Tel 01604 545542 | Screen issues | Replaced and calibrated touch screen. | SWO-273021 | Defect | HW | DMS_PC | Touch screen |  | Mark Spencer Tel 01604 545542 | Screen issues | Replaced and calibrated touch screen.
No | Visit | Complaint | RCHN | 放试剂包的位置漏液 | 放试剂包的位置漏液，借902-990、903-169、835-492各一个 | 更换835-492后修复 by jesson | SWO-272332 | [5 empty] | Position leakage of reagent pack | Place the reagent pack location leak, borrow 902-990, 903-169, 835-492 each one | Repair by jesson after replacing 835-492
No | Visit | Complaint | RCHN | 蓝屏 | 蓝屏，借910-342一个 | 更换910-342后修复 by jesson | SWO-272366 | Defect | HW | DMS_PC | DMS_PC |  | Blue screen | Blue screen, borrow 910-342 one | Repair by jesson after replacing 910-342
Yes | Visit | Complaint | RCHN | 黑屏 | 黑屏，借备件636-571一个 | 更换636-571后修复 by jesson | SWO-272321 | [5 empty] | Black screen | Black screen, borrowed spare parts 636-571 one | Repair by jesson after replacement 636-571
No | Visit | Complaint | RFRA | Signalement une mauvaise corrélation entre les 2 ABL825 Flex sur la Bilirubine | Verification historique AC5+ Bili Test lampe | Déprotéinisation Calibration 1 pts Cal ThB en mode service et utilisateur Passage des 4 niveaux de AC5+ | SWO-272963 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | Reporting a poor correlation between the 2 ABL825 Flex on the Bilirubin | Historical Verification AC5 Bili Lamp test | Deproteinization 1 pts calibration Cal ThB in service and user mode Passing the 4 levels of AC5
No | Visit | Complaint | RAME | pH & Ca failing out high sensitivity calibrations. | Inspection of the system revealed Ref channel appears to be erratic causing pH sensitivity of 103.5 & Ca 110.1 | Replaced El $Met PCB & disconnected failed Flex-Q module & ordered a replacement. | SWO-272882 | Revisit required | HW | El/Met Module | PCB, El/Met |  | pH & Ca failing out high sensitivity calibrations. | Inspection of the system revealed Ref channel appears to be erratic causing pH sensitivity of 103.5 & Ca 110.1 | Replaced El $Met PCB & disconnected failed Flex-Q module & ordered a replacement.
Yes | Visit | Complaint | RCHN | PO2超出响应范围 | PO2超出响应范围，借903-334一个 | 更换903-334后修复 by jesson | SWO-272949 | [5 empty] | PO2 out of range | PO2 out of range of response, borrow 903-334 a | Repair by jesson after replacing 903-334
No | Visit | Complaint | RCHN | 流量选择器错误 | 流量选择器错误，借903-185和580-417各一个 |  | SWO-273057 | Defect | HW | Basic Unit | Flow Selector Detector |  | Traffic Selector Error | Flow selector error, borrow 903-185 and 580-417 each one
No | Visit | Complaint | RCHN | 孵化器PCB损坏 | 孵化器PCB损坏，借903-138一个 |  | SWO-273046 | Defect | HW | Incubator module | PCB, Incubator |  | Incubator PCB damage | Incubator PCB damage, borrow 903-138 one
No | Visit | Complaint | RIND | Calibration expired. | demo key got expired | Need to update the option key. | SWO-276105 | Setup | SW | SW | Demo Key Exp | LEH mail | Calibration expired. | demo key got expired | Need to update the option key.
No | Visit | Complaint | RDEU | WE-Leckage erkannt | Leckage erkannt LP am 01.12. gewechselt SC durch MT Hr. Gräfe heute früh gewechselt Startphase nie abgeschlossen => immer wieder Interventionsmeldung Uhrzeit verstellt | Beim Eintreffen Gerät in Intervention mit Aufforderung zum Wechsel der Einlassdichtung => Dichtung wurde schon mehrfach gewechselt Überprüfung Fließwege über Service => Flüssigkeit unterhalb des Quetschventils => Quetschventilschlauch auf Höhe vom Quetschventilbügel perforiert => Austausch Schlauch und Einlassanschlussdichtung => Einlassdichtung in Kunden-SBox Alle Fluid Transport Programme anschließend ohne Fehler Startphase mit Reinstallation LP i.O. Durchführung aller QC Level Uhrzeit angepasst | SWO-273349 | Defect | HW | Oxi module | Pinch valve tube |  | WE leak age detected | Leakage detected LP changed on 01.12. SC changed by MT Mr. Gräfe this morning Start-up phase never completed => repeatedly intervention notification Time adjusted | When device arrives in intervention with request to change the inlet seal => seal has already been changed several times Check flow paths via service => liquid below the crushing valve => crushing valve hose perforated at level of the squeezing valve bracket => Exchange hose and inlet connection seal => inlet seal in customer SBox All Fluid Transport programs then without errors Start-up phase with reinstallation LP i.O. Execution of all QC levels Time adjusted
No | Visit | Complaint | RDEU | 30.11.2017, Herr Richter, tel: 03455573194: Probenaspiration gescheitert | LS auf Sensorbay blind | Fehlersuche, Reinigung Flüssigkeitssystem und Sensoren, Bestellung Ersatzteil | SWO-273558 | Revisit required | HW | Basic Unit | Cassette retainer |  | 30.11.2017, Mr. Richter, tel: 03455573194: Rehearsal aspiration failed | LS blind on Sensorbay | Troubleshooting, cleaning liquid system and sensors, ordering spare part
No | Visit | Complaint | RDEU | 30.11.2017, Herr Richter, tel: 03455573194: Probenaspiration gescheitert | LS auf Sensorbay blind | Austausch Sensorbay, Funktionskontrolle | SWO-273559 | Defect | HW | Basic Unit | Cassette retainer |  | 30.11.2017, Mr. Richter, tel: 03455573194: Rehearsal aspiration failed | LS blind on Sensorbay | Sensorbay replacement, function control
No | Visit | Complaint | RDEU | 04.12.2017, Frau Fink, Tel: 03078822878: Intervention erforderlich- Pumpenkal Fehler 04.12.2017, Sr. Christiane, Tel: 03078822855: weiterhin Intervention erforderlich - Pumpenkal Fehler | Lösungen können nicht gepumpt werden | Fehlersuche; Fließwege gespült; vorhandenes LP erneut eingesetzt -> Fehler; div. Serviceprogr. angewandt; LP gew.; Initialisierung abgewartet; Kal. + QC i.O. | SWO-273937 | Assisted Replacement | SP | SP | FT |  | 04.12.2017, Mrs. Fink, Tel: 03078822878: Intervention Required- Pumpscale Error 04.12.2017, Sr. Christiane, Tel: 03078822855: continued intervention required - Pumpscale Error | Solutions cannot be pumped | Troubleshooting; Flow paths rinsed; existing LP re-inserted -> error; div. Serviceprogr. applied; LP. initialization awaited; Kal. + QC in the first place
No | Visit | Complaint | RLTD | Blank |  |  | SWO-273416 | Defect | HW | Auto Inlet module | Probe bend |  | Blank
No | Visit | Complaint | RDEU | Kunde meldet das sich der Einlass nicht schließt. Analysator Meldung: Entfernung von Hindernissen aus dem Einlassbereich. Kunde ist mit der Bedienung überfordert. Kunden darauf hingewiesen das der Einsatz am Wochenende mit Kosten verbunden ist. Kunde klärt das mit leitenden Oberarzt. Kunde wünscht für Montag einen Service Einsatz. | Einlass schließt nicht richtig aufgrund verbogener Einlasssonde | Einlasssonde erneuert, Systemüberprüfung i. O. , fließwege überprüft i. O. , sensorkassette erneuert, Kalibrierung und Qualitätskontrolle nach telefonischer Rücksprache i. O. | SWO-273271 | Defect | HW | Auto Inlet module | Probe bend |  | customer reports that the entrance does not close. Analyzer Message: Removal of obstacles from the inlet area. Customer is overwhelmed with the operation. Customers pointed out that the use on weekends is associated with costs. Customer clarifies this with senior physician. Customer wants a service for Monday. | Inlet does not close properly due to bent inlet probe | Inlet probe renewed, system check in accordance with , flow paths checked i.o., sensor cassette renewed, calibration and quality control after telephone consultation as a standard
No | Visit | Complaint | RDEU | 13.11.2017, Frau Dzinic, Tel: 01776266679: AQC Fehler bei pO2 Level C seit Wochen | Gerät überprüft, Kal und QC sind seit Wochen i.O, | Mit Frau Dzinic festgestellt das in KIS System die Grenzbereiche falsch eingetragen sind. | SWO-273737 | User error | Setup | QC Limits | QC Limits |  | 13.11.2017, Mrs Dzinic, Tel: 01776266679: AQC error at pO2 Level C for weeks | device has been checked, Kal and QC have been | With Mrs Dzinic, it was established that in KIS System the border areas are incorrectly entered.
No | Visit | Complaint | RLTD | Melinda Kormany 07787 598407 | The analyser was blocked o |  | SWO-273397 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | Melinda Kormany 07787 598407 | The analyser was blocked o
No | Visit | Complaint | RLTD | Always failing Glucose across 3 different sensor cassettes. Update Failing Sodium and Glucose please contact Rhiannon Tel 01604 544080 | 1310 response error on glucose. | Replaced sensor interface PCB, however issues remained. Will have to return with main control board and new Interface PCB. | SWO-273700 | Defect | HW | Basic Unit | PCB, Sensor interface | LEH mail | Always failing Glucose across 3 different sensor cassettes. Update Failing Sodium and Glucose please contact Rhiannon Tel 01604 544080 | 1310 response error on glucose. | Replaced sensor interface PCB, however issues remained. Will have to return with main control board and new Interface PCB.
No | Visit | Complaint | RESP | Error en la cal de glucosa con casete diferntes | Error en la cal de glucosa con valores cercanos a 0 | Sustitucion de la placa sensor interface board Cambio de la casete Cal, qc y muestra | SWO-273435 | Defect | HW | Basic Unit | PCB, Sensor interface |  | Glucose lime error with cassette diferntes | Glucose lime error with values close to 0 | Replacement of the sensor interface board Changing the cassette Cal, qc and sample
No | Visit | Complaint | RCHN | 间歇性定标失败 | 间歇性定标失败，借903-581一个 |  | SWO-273181 | Defect | HW | Basic Unit | PCB, Sensor interface |  | Intermittent calibration failure | Intermittent calibration failure, borrowing 903-581 one
No | Visit | Complaint | RCHN | 定标质控不过 | 定标质控不过，借903-581一个 |  | SWO-273218 | Defect | HW | Basic Unit | PCB, Sensor interface |  | The calibration is not a good quality control. | The calibration is not a quality control, borrow 903-581 one
No | Visit | Complaint | RESP | No calibra parametros de o2,cl y na | No calibra parametros de o2,cl y na. Sci defectuoso. | Se sustituto sci y se comprobo correcto funcionamiento del equipo mediante qc. | SWO-273574 | Assisted Replacement | SC | SC | pO2, Cl, Na |  | No calibra parameters de O2, cl y na | Does not calibrate parameters of o2,cl and na. Sci defective. | The equipment is replaced and properly functioned by qc.
No | Visit | Complaint | RESP | Fuga detectada | Fuga detectada, sol pack defectuoso. | Se sustituyo solution pack y se comprobo correcto funcionamiento del equipo mediante qc. | SWO-273573 | Assisted Replacement | SP | SP | Leakage |  | Leak detected | Leak detected, sun pack defective. | Solution pack is replaced and the equipment is compatible with the proper operation by qc.
No | Visit | Complaint | RCHN | 报错1067 | 报错1067，借944-212（13315），944-212（13352）各一个 |  | SWO-273285 | Unknown | Consumables | Test Kit | Cartridge, Tni | 1067 | 1067 error | Error 1067, loan 944-212 (13315), 944-212 (13352) each
Yes | Visit | Complaint | RCHN | 报错1024 | 新机器装机前检测，报错1024，借903-364一个 | 更换903-364后修复 by Andrew | SWO-272974 | [5 empty] | 1024 error | Pre-installation detection of new machine, error 1024, loan 903-364 one | Repair by Andrew after replacing 903-364
No | Visit | Complaint | RCHN | 测试卡断裂，报错5784 | 测试卡断裂，报错5784,需要借件903-584两个 |  | SWO-273120 | Defect | HW | Cover and Chassis | Tube holder | 5784 | Test card break, 5784 error | Test card break, error 5784, need to borrow 903-584 two
Yes | 11 | Y | RPAC | Blank | Software not loading |  | D38952D30A43591D | [5 empty] | Blank | Software not loading
No | Visit | Complaint | RKK | AUと接続できないエラーで再起動後、改善する | 装置ログチェック、エラー記録確認 | 再起動後、正常稼働。AQURE経由でクイックアクセス使用時にAUengineを停止した可能性がある。様子を見る事に | SWO-274532 | Setup | SW | His/Lis | His/Lis |  | Improve after reboot with an error that cannot be connected to AU | Device log check, error record check | After reboot, it is up and running normally. AUengine may have been stopped when using Quick Access via AQURE. I'm going to see how it goes.
Yes | Visit | Complaint | RAME | Customer reported Carboxyhemoglobin running low in range for all levels of QC. | Checked data logs for errors. Research QC logs and check QC ranges. Ran complete oximetry tests with pass results. Checked lamp Intensity. | Ran tHB is service mode with pass results. Restart calibrations. | SWO-273694 | [5 empty] | Customer reported Carboxyhemoglobin running low in range for all levels of QC. | Checked data logs for errors. Research QC logs and check QC ranges. Ran complete oximetry tests with pass results. Checked lamp Intensity. | Ran tHB is service mode with pass results. Restart calibrations.
No | Visit | Complaint | RIND | Calibration expired. | Need to install the option key. | Option key installed successfully,run the 2 point calibration few times,run few patient sample & got the satisfactory result. | SWO-276118 | Setup | SW | SW | Demo Key Exp | LEH mail | Calibration expired. | Need to install the option key. | Option key installed successfully,run the 2 point calibration few times,run few patient sample & got the satisfactory result.
No | Visit | Complaint | RAME | pH & Ca failing out high sensitivity calibrations. | Inspection of the system revealed Ref channel appears to be erratic causing pH sensitivity of 103.5 & Ca 110.1 | Replaced AU PCB, Replaced pH measuring chamber, power supply, same issue. Ordering additional parts. Removing AU & power supply back to original 2 parts leaving measuring chamber in place. | SWO-273728 | Revisit required | Unknown | Unknown | Unknown |  | pH & Ca failing out high sensitivity calibrations. | Inspection of the system revealed Ref channel appears to be erratic causing pH sensitivity of 103.5 & Ca 110.1 | Replaced AU PCB, Replaced pH measuring chamber, power supply, same issue. Ordering additional parts. Removing AU & power supply back to original 2 parts leaving measuring chamber in place.
No | Visit | Complaint | RSCH | Blank | DMS non connecté au AUSW(2jrs après l’installation) | Remplacé le Power Supply (carte d’alimentation) e.o Remplacé l’ Analyzer Control PCB e.o Chargé WRES7 complet e.o Chargé le software système 3.3MR1 e.o | SWO-273516 | Defect | HW | Basic Unit | PCB, Analyzer control |  | Blank | DMS not connected to AUSW (2jrs after installation) | Replaced the Power Supply e.o Replaced the Analyzer Control PCB e.o Loaded WRES7 full e.o Loaded software system 3.3MR1 e.o
No | Visit | Complaint | RFRA | Erreur 1999 | consultation des fichiers | Contrôle et déplacement de la clé usb de sauvegarde. Collecte de données pour une éventuelle techline si le problème persiste après la mise à jour. Essais | SWO-273529 | Unknown | Unknown | Unknown | Unknown, data collected | 1999 | Error 1999 | file view | Control and move the backup USB key. Collecting data for a possible techline if the problem persists after the update. Tests
No | 4 | Y | RPAC | Blank | Clot deteced in OXI module |  | D391FCD244189E0A | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | Blank | Clot deteced in OXI module
Yes | 5 | Y | RPOL | Blank | Sprawdzenie aparatu, aktualizacja oprogramowania 3.3MR1. Czyszczenie selektora przep³ywów. Kalibracja hemoglibiny . Kalibracja oraz kontrola jakoœci prawid³owa. Analizator sprawny technicznie. |  | D38A632980426E81 | Re classified RMED | Planned |  |  |  | Blank | Camera check, software update 3.3 MR1. Cleaning the overflow selector. Calibration of Hemoglibine. Calibration and quality control. Technically efficient analyzer.
No | Visit | Complaint | RLTD | Tim Gray on 07546 906 874 | Analyser frozen in mid boot up with data upgrade failed, status update 2. Unable to determine cause. | Reloaded software and manually moved setup, analyser has been of for over 24 hours, sensor cassette maybe damaged. Loaded set from original ed transfer. | SWO-274229 | Defect | SW | Windows | Windows | Status Code 2 | Tim Gray on 07546 906 874 | Analyser frozen in mid boot up with data upgrade failed, status update 2. Unable to determine cause. | Reloaded software and manually moved setup, analyser has been of for over 24 hours, sensor cassette maybe damaged. Loaded set from original ed transfer.
No | Visit | Complaint | RLTD | Please call Sam Wakefield - 01604 634 700 | Failing QCs. | No one answering on number provided, arrived on site re-ran calibration and QCs, now ok. | SWO-274228 | Adjustment | SC | SC | pO2 |  | Please call Sam Wakefield - 01604 634 700 | Failing QCs. | No one answering on number provided, arrived on site re-ran calibration and QCs, now ok.
Yes | 6 | Y | RPOL | Blank | Zmiana drukarki z brother na hp LaserJet Pro M501. Nie mo¿na zainstalowaæ drukarki. Konieczne dalsze wizyty. |  | D38D3B5C6C4BB473 | Re classified RMED | DEMO Setup |  |  |  | Blank | Changing the printer from brother to HP LaserJet Pro M501. You cannot install the printer. Further visits are necessary.
No | Visit | Complaint | RKK | 安定化中が消えない。チャンバー内多少の汚れの為、清掃。メンブラン交換後復帰。大気測定問題なし。 | 安定化中が消えない。 | チャンバー内多少の汚れの為、清掃。メンブラン交換後復帰。大気測定問題なし。 | SWO-274050 | [5 empty] | During stabilization does not disappear. Cleaning for some dirt in the chamber. Returned after the menbran exchange. No atmospheric measurement problems. | During stabilization does not disappear. | Cleaning for some dirt in the chamber. Returned after the menbran exchange. No atmospheric measurement problems.
Yes | Visit | Complaint | RCHN | 报错1024，1025，1031 | 报错1024，1025，1031，903-364,944-212（13315）缺货一张 | 更换903-364后修复 by Andrew | SWO-273051 | [5 empty] | Error 1024,1025,1031 | 1024,1025,1031,903-364,944-212 (13315) out of stock | Repair by Andrew after replacing 903-364
No | 13 | Y | RPAC | Blank | High Crea check failed |  | D38A828B8022E87D | Assisted Replacement | Electrode | Electrode | Crea | CAPA | Blank | High Crea check failed
No | Visit | Complaint | RAME | Blue screen and won't boot up | Analyzer would not boot. Had a The blue screen of death. | The customer and I rebooted the analyzer with no change. I disconnected the AU and it came up with the start up screen but no farther. I replaced the CPU and the analyzer came to life. | SWO-275043 | Defect | HW | DMS_PC | CPU |  | Blue screen and won't boot up | Analyzer would not boot. Had a The blue screen of death. | The customer and I rebooted the analyzer with no change. I disconnected the AU and it came up with the start up screen but no farther. I replaced the CPU and the analyzer came to life.
No | Visit | Complaint | RLTD | Hannah McPeak Tel 01865 741166 Bleep 1624 | The touchscreen was working on arrival but at times there did seem to be a delay when pressing the buttons. | Cleaned and dried the touchscreen to make sure that there wasn't any dirt on the screen that could be effecting its operation. Updated the touchscreen firmware which should also resolve the issue where the touchscreen becomes completely unresponsive and requires a reboot to get it working again. After cleaning the screen and updating the firmware the touchscreen seems to be working normally. | SWO-274421 | Defect | HW | DMS_PC | Touch screen | Firmware | Hannah McPeak Tel 01865 741166 Bleep 1624 | The touchscreen was working on arrival but at times there did seem to be a delay when pressing the buttons. | Cleaned and dried the touchscreen to make sure that there wasn't any dirt on the screen that could be effecting its operation. Updated the touchscreen firmware which should also resolve the issue where the touchscreen becomes completely unresponsive and requires a reboot to get it working again. After cleaning the screen and updating the firmware the touchscreen seems to be working normally.
No | Visit | Complaint | RAME | Rinse Errors | Unit aborting calibrations due to 0597 errors. Liquid in pH $BG module during gas calibration. | Replaced flat valve membrane in pH $BG module. Also replaced missing seal ring on glucose electrode and electrode module pump tubes. | SWO-274458 | Defect | Electrode | Electrode | Sealing ring | Sealing ring, Glu | Rinse Errors | Unit aborting calibrations due to 0597 errors. Liquid in pH /BG module during gas calibration. | Replaced flat valve membrane in pH $BG module. Also replaced missing seal ring on glucose electrode and electrode module pump tubes.
No | Visit | Complaint | RAME | analyzer not recognizing QC on carousel. | Verified customer's stated complaint. Upon further investigation it was discovered that the customer had set level 2 and level 3 to run manually. | Set both control levels to run automatically and corrected the problem. | SWO-274406 | Setup | SW | QC Schedule | QC Schedule |  | analyzer not recognizing QC on carousel. | Verified customer's stated complaint. Upon further investigation it was discovered that the customer had set level 2 and level 3 to run manually. | Set both control levels to run automatically and corrected the problem.
No | Visit | Complaint | RAME | 0735: Ampoule temperature error and Autocheck will not run QC. | Found temperature for autocheck in service mode @ 55 degrees. | Replaced carousel $zero detect pcb. | SWO-274348 | Defect | HW | Auto Inlet module | PCB, AC Dist |  | 0735: Ampoule temperature error and Autocheck will not run QC. | Found temperature for autocheck in service mode @ 55 degrees. | Replaced carousel $zero detect pcb.
No | Visit | Complaint | RIND | Thb/so2 values ? | Thb $so2 ? | Decontamination process was done and pump calibartion was given and cleaned the gaskets and calibrated to raedy mode and tested fews amples was working fine | SWO-275698 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | Thb/so2 values ? | Thb /so2 ? | Decontamination process was done and pump calibartion was given and cleaned the gaskets and calibrated to raedy mode and tested fews amples was working fine
No | Visit | Complaint | RDAN | I forbindelse med afslutningsforretning klager Jan, Jane og Marianne over div. sjove fejl. | Afprøvet. Målemoduls ventiler har tørre $brændte ankre. | Målemoduls ventiler renset og smurt, slutafprøvet. | SWO-274217 | Defect | HW | pH/BG Module | Valve | Cleaned and re-greased | In connection with closing business, Jan, Jane and Marianne complain about Div. funny mistakes. | Tested. Measuring module valves have dry /br-lit anchors. | Measuring modul valves cleaned and lubricated, final tested.
No | 2 | Y | RPAC | Blank | Halogen too high |  | D38A6A06802F6A93 | Adjustment | HW | Oxi module | Lamp unit | Halogen | Blank | Halogen too high
No | Visit | Complaint | RIND | Undefined Lookup String. 35.16.1.0-1-1 | Undefined Lookup string. | As suggested by RMED replaced the SP and restarted instrument found instrument is working satisfactorily. | SWO-276458 | Assisted Replacement | Consumables | SP | SP | FSE misjudged | Undefined Lookup String. 35.16.1.0-1-1 | Undefined Lookup string. | As suggested by RMED replaced the SP and restarted instrument found instrument is working satisfactorily.
No | Visit | Complaint | RFRA | Problème optique Calibration tni en erreur ,les blancs sont trop hauts | Vérification des calibrations et tests système Réalisation du background Check -> fail | Remplacement du bloc optique $ $ calibration factory $ $ essai du background Check -> ok $ $ calibration crp lot13012 $ $ calibration Bhcg lot13491 $ $ calibration Tni lot13551 $ $ passage de contrôle $ $ essais bons | SWO-274352 | Defect | HW | Process Module | Optical Unit | 1024 | Optical problem Calibration tni in error, the whites are too high | Checking calibrations and system tests Making the background check -fail | Replacement of the optical block $$ calibration factory $$ test of the background check - ok $$ calibration crp lot13012 $$ calibration Bhcg lot13491 $ calibration Tni lot13551 $ passage control $$ trials
No | Visit | Complaint | RSCH | Parameter Kalibrieren | 05.12.2017 DANKO Repariertes Kundengerät retour Nach Neustart vor Ort Kalibratoren nicht mehr vorhanden | 05.12.2017 JOSLO Vor Ort TnT kalibriert, danach Neustart durchgeführt i.o Neuer Kalibrierten TnT Parameter vorhanden i.o Alte Parameter nach Neustart wieder vorhanden i.o D-dimer, NT-proBNP, CRP und TnT sind aktiviert i.o System i.o Kunde bekommt aus Kulanz NT-proBNP Test Kit (ohne Cal Cartridge) Da vorhandene Test andere Lot 13290 Kundentestkassetten retour 7x Retour nach Thalwil TnT Test Kit + 2 Cal Cartridge 13631 D-dimer Test Kit + 2 Cal Cartridge 13566 3x Cal Cartridge NT-proBNP | SWO-274799 | User error | User error | User error | Calibration data erased |  | Parameters Calibration | 12/05/2017 THANKS Repaired customer device back Calibrators no longer available after restart on site | 12/05/2017 JOSLO On-site TnT calibrated, then rebooted i.o New calibrated TnT parameters available i.o Old parameters re-existing after restart i.o D-dimer, NT-proBNP, CRP and TnT are activated i.o System i.o Customer gets out of goodwill NT-proBNP Test Kit (without Cal Cartridge) Since existing test other lot 13290 Customer test cassettes return 7x Return to Thalwil TnT Test Kit + 2 Cal Cartridge 13631 D-dimer Test Kit + 2 Cal Cartridge 13566 3x Cal Cartridge NT-proBNP
No | Visit | Complaint | RLTD | Please call - see below contact detailsAnnette Neill Specialist Medical Scientist POCT Coordinator/Laboratory Accreditation Pathology Department St John’s Hospital St John’s Square Limerick Ireland Email: annette.neill@stjohnshospital.ie Website: http://www.stjohnshospital.ie Phone: 00 353 61 462 119 Set up and static IP address applied | Not connecting to Aqure. | Analyser was moved to another site & IP settings were changed. Input correct settings & tested connection. | SWO-274729 | Setup | SW | His/Lis | His/Lis |  | Please call - see below contact detailsAnnette Neill Specialist Medical Scientist POCT Coordinator/Laboratory Accreditation Pathology Department St John’s Hospital St John’s Square Limerick Ireland Email: annette.neill@stjohnshospital.ie Website: http://www.stjohnshospital.ie Phone: 00 353 61 462 119 Set up and static IP address applied | Not connecting to Aqure. | Analyser was moved to another site & IP settings were changed. Input correct settings & tested connection.
No | Visit | Complaint | RLTD | Please call Julie Knutson on 01642 850 850 and ask for Cardiac Cath Lab, Lab 4 | battery error | Powered analyser down and replaced battery. On analyser power on the error had been resolved. | SWO-274863 | Defect | HW | Battery module | Battery |  | Please call Julie Knutson on 01642 850 850 and ask for Cardiac Cath Lab, Lab 4 | battery error | Powered analyser down and replaced battery. On analyser power on the error had been resolved.
No | Visit | Complaint | RDEU | Flüssigkeitstransportfehler. | Intervention erforderlich, Lösungspack wechseln | Fehlersuche: mehrfach alle relevanten Serviceprogramme absolviert, => Vacuumtest Fehler, immer wiederkehrend trotz Spülung und einwandfrei durchgängigen sowie dichten Fließwegen => SC ersetzt, Fehlermeldung sofort weg. Neustart, SW Upgrade (s. weitere SWOs), Kal und QC Messungen ok. | SWO-274786 | Assisted Replacement | SC | SC | FT |  | Flüssigkeitstransportfehler. | Intervention required, solution pack change | Troubleshooting: completed several times all relevant service programs, => Vacuumtest error, repeatedly repeated despite flushing and perfectly continuous and dense flow paths => SC replaced, error message immediately removed. Restart, SW Upgrade (see more SWOs), Kal and QC measurements ok.
No | Visit | Complaint | RAME | pO2 issues | pO2 is high outside qc range for qc level C. | After inspecting the analyzer, I found that the analyzer was also presenting pO2 inhomogeneous error from the internal qc level C and the external qc level 3. The auto detection was not working properly due to the pO2 error. The analyzer software recommendation was to 'Replace the solution pack' which Jodi already did on two occasions as per the analyzer software 'recommended actions'. I flushed the analyzer. Ran the service fluid transport checks all the pressure test values failing outside of range. I ran the solution pack check, passing ok. I checked the qc data log, found that the qc failures started occurring after the sensor cassette was replaced on 11 $27 $17. Determined that the pO2.errors were being caused by the sensor cassette, replaced a new sensor cassette. Submitting the credit form for this defective cassette which still had 192 remaining tests. | SWO-275251 | Assisted Replacement | SC | SC | pO2 |  | pO2 issues | pO2 is high outside qc range for qc level C. | After inspecting the analyzer, I found that the analyzer was also presenting pO2 inhomogeneous error from the internal qc level C and the external qc level 3. The auto detection was not working properly due to the pO2 error. The analyzer software recommendation was to 'Replace the solution pack' which Jodi already did on two occasions as per the analyzer software 'recommended actions'. I flushed the analyzer. Ran the service fluid transport checks all the pressure test values failing outside of range. I ran the solution pack check, passing ok. I checked the qc data log, found that the qc failures started occurring after the sensor cassette was replaced on 11 $27 $17. Determined that the pO2.errors were being caused by the sensor cassette, replaced a new sensor cassette. Submitting the credit form for this defective cassette which still had 192 remaining tests.
No | Visit | Complaint | RLTD | Tim Gray on 07546 906 874 | No set up due to clean install. | Had to clean install software last visit, and no one was available to provide setup information. Setup analyser as per local instructions. | SWO-275443 | Setup | SW | User Setup | User Setup |  | Tim Gray on 07546 906 874 | No set up due to clean install. | Had to clean install software last visit, and no one was available to provide setup information. Setup analyser as per local instructions.
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 不均匀的QC2溶液 1541 | 更换试剂包后修复 | SWO-276017 | [5 empty] | Test card, reagent package failure | Uneven QC2 solution 1541 | Repair after replacement of reagent kit
Yes | Visit | Complaint | RCHN | 报错1031 | 报错1031，借903-364一个 | 更换903-364后修复 by Andrew | SWO-273862 | [5 empty] | 1031 | 1031, loan 903-364 one | Repair by Andrew after replacing 903-364
No | Visit | Complaint | RCHN | PH质控结果偏高 | PH质控结果偏高，借905-792一个 | 更换905-792后修复 by Andrew | SWO-274613 | Defect | HW | Basic Unit | Fluidic manifold |  | PH QUALITY CONTROL RESULTS ARE HIGH | PH QC AND CONTROL RESULTS ARE HIGH, BORROWING 905-792 ONE | Repair by Andrew after replacing 905-792
Yes | Visit | Complaint | RDEU | Kunde hat Probleme mit Glu und Lac. | Probleme mit Glu $ Lac Membranen | Glu und Lac Elektroden gereinigt und membraniert. | SWO-274939 | [5 empty] | Customer has problems with Glu and Lac. | Problems with Glu-Lac Membranes | Glu and lac electrodes cleaned and membraneed.
No | 5 | Y | RCZE | Blank |  |  | D38BCB07F64AF95B | Defect | HW | FlexQ Module | Barcode reader, FlexQ | LEH Mail | Blank
Yes | Visit | Complaint | RDAN | ABL har kørt med gas 1 lot Y10H25 og med dette lot har ABL meldt at gas 1 var tom før tid. | Alle de gang hvor ABLen har meldt om at gas 1 er tom, har det været lot Y10H25 som har været brugt. | Der er nu send en kasse gas 1 hvor lot er Y31H33. Jeg afventer nu om der igen kommer fejl. | SWO-279652 | Re classified RMED | Logistical upSWO |  |  |  | ABL has been running with gas 1 lot Y10H25 and with this lot, ABL has reported that gas 1 was empty before time. | All the time the ABLen has reported that gas 1 is empty, it has been lot Y10H25 which has been used. | There is now send a box gas 1 where lot is Y31H33. I am now waiting for mistakes to be made again.
No | Visit | Complaint | RDAN | I forbindelse med afslutningsforretning klager Jan, Jane og Marianne over div. sjove fejl. | Afprøvet. | Målemoduls ventiler renset og smurt, slutafprøvet. | SWO-274960 | Defect | HW | pH/BG Module | Valve |  | In connection with closing business, Jan, Jane and Marianne complain about Div. funny mistakes. | Tested. | Measuring modul valves cleaned and lubricated, final tested.
No | 8 | Y | RPAC | Blank | IP addresses install $ Firewall issue |  | D38BEA32F610E4C9 | Setup | SW | His/Lis | His/Lis |  | Blank | IP addresses install / Firewall issue
No | Visit | Complaint | RAME | pH & Ca failing out high sensitivity calibrations. | Inspection of the system revealed Ref channel appears to be erratic causing pH sensitivity of 103.5 & Ca 110.1. The original pH electrode is lot YT-01. | Replaced electrode module V6 & V7 valves, all tubing from a pm kit & the pH electrode to resolve the issue. System is operational. | SWO-274506 | Defect | HW | pH/BG Module | Valve |  | pH & Ca failing out high sensitivity calibrations. | Inspection of the system revealed Ref channel appears to be erratic causing pH sensitivity of 103.5 & Ca 110.1. The original pH electrode is lot YT-01. | Replaced electrode module V6 & V7 valves, all tubing from a pm kit & the pH electrode to resolve the issue. System is operational.
No | Visit | Complaint | RAME | Oxi HW error 0579 & 316 Neon number too low. | CAS Kim S. is requesting oxi system be verified due to COHb is running high on lev 2 QC at 3.6 per EP evaluation. | Re-verified oxi module. No changes made. | SWO-275178 | No fault found | No fault found | No fault found | No fault found |  | Oxi HW error 0579 & 316 Neon number too low. | CAS Kim S. is requesting oxi system be verified due to COHb is running high on lev 2 QC at 3.6 per EP evaluation. | Re-verified oxy module. No changes made.
No | Visit | Complaint | RSAF | Customer failing to load cartridges on the analyzer | Customer failing to load cartridges on the analyzer | Opened the analyzer and inspected for any obstructions on the storage wheel, cleaned the barcode scanner. Ran Safe positions to verify the functionality of the analyzer. | SWO-274597 | Cleaned | HW | Storage Module | Barcode reader, Storage wheel |  | Customer failing to load cartridges on the analyzer | Customer failing to load cartridges on the analyzer | Opened the analyzer and inspected for any obstructions on the storage wheel, cleaned the barcode scanner. Ran Safe positions to verify the functionality of the analyzer.
No | Visit | Complaint | RDEU | 05.12.2017, Frau Schick, Tel: 07631883171: Fehler Incubatorwheelmove 31- No cup under puncher after punch - | Probenmessung abgebrochen, Incubatorwheel Move-Error | Incubatorwheel überprüft, diverse Testprogramme durchlaufen, keine Einschränkung feststellbar, Testmessungen ohne Beanstandung | SWO-274726 | User error | Consumables | Test Kit | Cartridge reinstalled |  | 05.12.2017, Frau Schick, Tel: 07631883171: Fehler Incubatorwheelmove 31- No cup under puncher after punch - | Sample measurement aborted, Incubatorwheel Move-Error | Incubatorwheel checked, various test programs go through, no limitation can be detected, test measurements without complaint
Yes | Visit | Complaint | RDEU | 06.12.2017, Frau Jackmat, Tel: 0832396390: Intervention erforderlich-Lösungen können nicht gepumpt werden. 06.12.2017 09:15Uhr Frau Jackmaci 0832396390: Lösungen können nicht gepumpt werden. | Intervention erforderlich-Lösungen können nicht gepumpt werden. | ABL manuell gespült, Einlassnadel überprüft, Einlassdichtung erneuert, Kalibrationen und QC's vorgenommen, BGA ok. | SWO-276475 | [5 empty] | 06.12.2017, Mrs Jackmat, Tel: 0832396390: Intervention required Solutions cannot be pumped. 06.12.2017 09:15Pm Mrs Jackmaci 0832396390: Solutions cannot be pumped. | Intervention required solutions cannot be pumped. | ABL manually rinsed, inlet needle checked, inlet seal renewed, calibrations and QC's done, BGA ok.
No | Visit | Complaint | RLTD | Steven on 01604 545611 | Disk boot failure | Fault finding revealed corrupt flash card, replaced flash card now ok. | SWO-275585 | Defect | HW | DMS_PC | SD Card |  | Steven on 01604 545611 | Disk boot failure | Fault finding revealed corrupt flash card, replaced flash card now ok.
Yes | 6 | Y | RPOL | Blank | Reinstalacja systemu Windows, zmiana ustawieñ. Pod³¹czenie drukarki oraz testy funkcjonalnoœci. Drukarka nie drukuje polskich liter. Kalibracja tHb. Deinstalacja drukarki brother. Konieczne dalsze dzia³ania naprawcze. Analizator sprawny technicznie i pomiarowo. |  | D38D3C486C43B476 | Re classified RMED | DEMO Setup |  |  |  | Blank | Reinstallation of Windows, change settings. Printer connection and functional tests. The printer does not print Polish letters. THb Calibration. Uninstalling the Brother printer. Further corrective action necessary. The analyser is technically efficient and measuring.
No | Visit | Complaint | RCHN | 散热风扇声音过大 | 散热风扇声音过大，借636-520一个 | 更换636-520后修复 by Andrew | SWO-274701 | [5 empty] | Cooling fan sounds too loud | Cooling fan sound is too large, borrow 636-520 one | Repair by Andrew after replacement 636-520
No | Visit | Complaint | RCHN | 无法打印 | 无法打印，借903-001一个 | 更换903-001后修复 by Andrew | SWO-274719 | [5 empty] | Can't print | Cannot print, borrow 903-001 one | Fix by Andrew after replacing 903-001
No | Visit | Complaint | RCHN | 测试卡漏液导致接口板腐蚀 | 测试卡漏液导致接口板腐蚀，需要借件902-972 | 更换902-972后修复 by Andrew | SWO-274542 | [5 empty] | Test card leak causes interface plate corrosion | Test card leakage causes interface plate corrosion, needs to borrow 902-972 | Repair by Andrew after replacement 902-972
No | Visit | Complaint | RAME | Customer notified Radiometer sales & service about a possible patient maltreatment based on blood gas results | Blood gas was ran on the patient in question on 12 $4 $17 @ 8:37 am. Patient was transferred from Outpatient to ER based on results of HCO3(P)=12.6 on this unit. ER Dr. requested a rerun of the blood gas at 10:24 am and results for HCO3(P) were 12.0 which the patient was treated on and released with a prescription based on the results of the blood gas. Customer was unable to run the samples on the other analyzer (315248) due to not having consumables. After consumables were installed patient was retested on 315248 and pH & HCO3(P) were different. Customer notified Radiometer. Upon arrival I researched the logs and found that external QC's were out for pH & pCO2. Analyzer was not set for QC Lockout. Performed zero check and all passed but sensor connector board showed signs of corrosion. | Replaced sensor connector pcb due to corrosion. Setup analyzer to Disable Analysis if QC has errors with QC. Input reference, critical & reportable ranges. Ran blood sample on both analyzers and pH on this analyzer was 7.19 and 7.43 on other analyzer. Replaced sensor cassette, ran 1 QC which then failed pO2. Replaced solution pack, reran 3 levels of QC, all passed. Reran blood sample and the pH on this analyzer was 7.41 and 7.43 on the other analyzer. Customer agreed those were acceptable results. | SWO-275145 | [5 empty] | Customer notified Radiometer sales & service about a possible patient maltreatment based on blood gas results | Blood gas was ran on the patient in question on 12 /4 /17 @ 8:37 am. Patient was transferred from Outpatient to ER based on results of HCO3(P)=12.6 on this unit. ER Dr. requested a rerun of the blood gas at 10:24 am and results for HCO3(P) were 12.0 which the patient was treated on and released with a prescription based on the results of the blood gas. Customer was unable to run the samples on the other analyzer (315248) due to not having consumables. After consumables were installed patient was retested on 315248 and pH & HCO3(P) were different. Customer notified Radiometer. Upon arrival I researched the logs and found that external QC's were out for pH & pCO2. Analyzer was not set for QC Lockout. Performed zero check and all passed but sensor connector board showed signs of corrosion. | Replaced sensor connector pcb due to corrosion. Setup analyzer to Disable Analysis if QC has errors with QC. Input reference, critical & reportable ranges. Ran blood sample on both analyzers and pH on this analyzer was 7.19 and 7.43 on other analyzer. Replaced sensor cassette, ran 1 QC which then failed pO2. Replaced solution pack, reran 3 levels of QC, all passed. Reran blood sample and the pH on this analyzer was 7.41 and 7.43 on the other analyzer. Customer agreed those were acceptable results.
No | Visit | Complaint | RCHN | 无法开机 | 无法开机，借910-285和910-342各一个 | 更换910-285后修复 by Andrew | SWO-274625 | Defect | HW | Basic Unit | Power Supply |  | Cannot power on | Cannot be switched on, borrow 910-285 and 910-342 each | Repair by Andrew after replacement 910-285
No | Visit | Complaint | RCHN | 无法做样本 | 无法做样本，借902-980和902-918各一个 | 更换902-980和902-918后修复 by Andrew | SWO-274529 | Defect | HW | Oxi module | Hemolyzer |  | Can't make a sample | Cannot make a sample, borrow 902-980 and 902-918 each | Fix by Andrew after replacing 902-980 and 902-918
No | Visit | Complaint | RCHN | 测试卡温度加热 | 测试卡温度加热，借903-581一个 | 更换备用机上的903-581备件后修复 by Andrew | SWO-275275 | Defect | HW | Basic Unit | PCB, Sensor interface | Temp | Test card temperature heating | Test card temperature heating, borrow 903-581 one | Repair by Andrew after replacing 903-581 spare parts on a spare machine
No | Visit | Complaint | RESP | 0964 corriente de fuga en paquete de soluciones. | 0964 corriente de fuga en paquete de soluciones. |  | SWO-275393 | Assisted Replacement | SP | SP | Leak current | 0964 | 0964 leakage current in solution package. | 0964 leakage current in solution package.
No | Visit | Complaint | RCHN | 报错1035 | 报错1035，借903-364一个 | 更换903-364后修复 by Andrew | SWO-274723 | Defect | HW | Process Module | Optical Unit | 1035 | 1035 error | 1035, loan 903-364 one | Repair by Andrew after replacing 903-364
No | Visit | Complaint | RCHN | AU板故障，显示屏花屏，检测不到进样针 | AU板故障，显示屏花屏，检测不到进样针，加借903-198一个 | 更换903-198,910-298,636-549后修复 by Andrew | SWO-275258 | Defect | HW | Analyzer Control Module | PCB, Main Board |  | AU board fault, display screen, can not detect the sample needle | AU board fault, display screen, can not detect the sample needle, plus 903-198 one | Repair by Andrew after replacing 903-198, 910-298, 636-549
No | Visit | Complaint | RCHN | 试剂包弹片掉了 | 试剂包弹片掉了，借903-036一个 | 更换903-036后修复 by Andrew | SWO-273876 | Defect | HW | Storage Module | PCB, Replesnish Pack |  | The reagent pack shrapnel's gone. | Reagent pack shrapnel dropped, borrow 903-036 one | Repair by Andrew after replacement 903-036
No | Visit | Complaint | RIND | Calibration error. | Check the system. Check the calibration logs. | run calibration multiple time. flush run cassette & Reinstalled. calibration are in ranges & qc is pass. | SWO-276611 | [5 empty] | Calibration error. | Check the system. Check the calibration logs. | run calibration multiple time. flush run cassette & Reinstalled. calibration are in ranges & qc is pass.
No | Visit | Complaint | RIND | QC Ampoules not being scanned in Autocheck module | Checked autocheck module and found scan engine is defective. | Laser scan engine (910-304) requires replacement. | SWO-275711 | Revisit required | HW | Auto Check Module | Barcode reader, Autocheck Module |  | QC Ampoules not being scanned in Autocheck module | Checked autocheck module and found scan engine is defective. | Laser scan engine (910-304) requires replacement.
Yes | Visit | Complaint | RDEU | heller Bildschirm | helles Display | Aufgef. Defekte CPU ersetzt, Kal, QC | SWO-275428 | Defect | HW | DMS_PC | CPU |  | bright screen | bright display | Aufgef. Defective CPU replaced, Kal, QC
No | Visit | Complaint | RAME | Blood leaking from ph module trough, ph gaskets replaced with no help | Found po2 gasket missing causing leak | Replaced gasket and tested for leaks | SWO-275761 | Defect | Electrode | Electrode | Sealing ring | Sealing ring, pO2 | Blood leaking from ph module trough, ph gaskets replaced with no help | Found po2 gasket missing causing leak | Replaced gasket and tested for leaks
No | 3 | Y | RCZE | Blank |  |  | D391D6CF442DEA61 | Assisted Replacement | Electrode | Electrode | pCO2 | Sensitivity | Blank
Yes | Visit | Complaint | RFRA | Après débouchage, erreur pompage réactifs | Sonde tordue, joint connexion n'était pas bien positionné. Plus problème avec le pack neuf. | Nettoyage circuit fluide, remise en place joint connexion, chgt sonde. Contrôle en mode service ok, plus formation. | SWO-276121 | [5 empty] | After unblocking, reactive pumping error | Probe twisted, joint connection was not well positioned. No more problem with the new pack. | Fluid circuit cleaning, re-establishment joint connection, chgt probe. Control in service mode ok, plus training.
No | Visit | Complaint | RDEU | Lösungsp. soll immer wieder getauscht werden, Fließfehler pO2 fehlerhaft bei QC -Messungen, | pO2 QC-Fehler, LP soll immer wieder getauscht werden, Fließstrecken fehler, | LP ersetzt, Fließwege gespült, Gerät neu gestartet, Flowsensor gereinigt, Gerät kalibriert, | SWO-276156 | Assisted Replacement | SP | SP | pO2 |  | Solutionsp. should be exchanged again and again, flow error pO2 faulty in QC measurements, | pO2 QC error, LP should be exchanged again and again, flow lines error, | LP replaces, flow paths flushed, device restarted, flow sensor cleaned, device calibrated,
No | Visit | Complaint | RKK | レベルCにて、pO2の値が選択値(4.0～29.5)が36.3にと高めに表示される | 機器状況確認 | 排液 $インレットライン洗浄。プラスチックポンプローター確認。ピンチバルブ動作確認。pO2用オプトセンサーレンズ清掃。各部動作チェックOK。CALチェックOK。QCチェックOK（3レベル） | SWO-276594 | Cleaned | Fluidic pathway | Fluidic pathway | pO2 |  | At level C, the value of pO2 is displayed higher than the selected value (4.0 to 29.5) to 36.3 | Equipment status check | Drainage $Inlet line cleaning. Plastic pump rotor confirmed. Pinch valve operation confirmed. Optosensor lens cleaning for pO2. Each part operation check OK. CAL check OK. QC check OK (3 levels)
No | 6 | Y | RPAC | Blank | Failing Na on QC |  | D3931A64BA1D4AA6 | Assisted Replacement | SC | SC | Na - QC |  | Blank | Failing Na on QC
No | Visit | Complaint | RIND | CPU sound form ABL80 | Check the system. check the sound. check the all fans. | the sound is generated from cpu section & speaker. check the all fans & restart the cpu but continuous sound generated from cpu section & speaker. reset the ram & reset the speaker. the sound stop but the red strip on display.after resting the cable also.machine under observation. | SWO-276533 | [5 empty] | CPU sound form ABL80 | Check the system. check the sound. check the all fans. | the sound is generated from cpu section & speaker. check the all fans & restart the cpu but continuous sound generated from cpu section & speaker. reset the ram & reset the speaker. the sound stop but the red strip on display.after resting the cable also.machine under observation.
No | Visit | Complaint | RAME | Sending data that is not necessary. | Unnecessary calculated parameters being sent to Telco and LIS | Edit XML files in ABL folder and test results by sending patient samples to LIS | SWO-278376 | [5 empty] | Sending data that is not necessary. | Unnecessary calculated parameters being sent to Telco and LIS | Edit XML files in ABL folder and test results by sending patient samples to LIS
No | Visit | Complaint | RCHN | 开机报错，有异响声 | 开机报错，有异响声,需要借件903-111 | 更换备件 | SWO-275929 | [5 empty] | Turn on the error, there's a noise. | Boot error, there is a noise, need to borrow 903-111 | Replacement of spare parts
No | Visit | Complaint | RCHN | GLU定标失败 | GLU定标失败，连续更换2张测试卡都无法恢复，故借903-581一个 | 使用备用机备件903-581后修复by Andrew | SWO-275266 | Defect | HW | Basic Unit | PCB, Sensor interface | Glu | GLU calibration failed | GLU calibration failure, continuous replacement of 2 test cards can not be recovered, so borrow 903-581 one | Repair by Andrew after using spare parts 903-581
Yes | Visit | Complaint | RCHN | 定标卡扫描器无法扫描 | 定标卡扫描器无法扫描，借902-924一个 | 更换902-924后修复 by Andrew | SWO-274676 | [5 empty] | The calibration card scanner cannot be scanned | The calibration card scanner cannot be scanned, borrowing 902-924 one | Repair by Andrew after replacing 902-924
No | Visit | Complaint | RCHN | 报错5784 | 报错5784，借13147（944-212）一张 缺货 | 更换13147（944-212）后修复 by Andrew | SWO-275277 | Adjustment | HW | Unknown module | Steppermotor, unknown | 5784 | 5784 | 5784, 13147 (944-212) out of stock | Fix by Andrew after replacing 13147 (944-212)
No | Visit | Complaint | RCHN | 报错1024 | 报错1024，需要借件903-364、944-212（13464）（缺货）、944-214（13265）、944-216（13072）（缺货） | 更换903-364、944-214（13265）、备件后修复 by Andrew | SWO-274598 | Defect | HW | Process Module | Optical Unit | 1024 | 1024 error | 1024, need to borrow 903-364, 944-212 (13464) (out of stock), 944-214 (13265), 944-216 (13072) (out of stock) | Replace 903-364, 944-214 (13265), repair by Andrew after spare parts
Yes | Visit | Complaint | RCHN | 光学单元故障 | 光学单元故障，需要借件903-364 | 硬件问题，调整后，现已排除故障 by Andrew A brand new optical unit, there are not screw lines. it is only difficult to install at the first time to tight screws. no issues, no problems there are. | SWO-275897 | [5 empty] | Optical unit failure | Optical unit failure, need to borrow 903-364 | 硬件问题，调整后，现已排除故障 by Andrew A brand new optical unit, there are not screw lines. it is only difficult to install at the first time to tight screws. no issues, no problems there are.
No | Visit | Complaint | RCHN | PH测量不稳定 | PH测量不稳定，借945-614一个 | 更换945-614后修复 by Andrew | SWO-275300 | Assisted Replacement | Electrode | Electrode | pH | Unstable | PH measurement instability | PH measurement unstable, borrow 945-614 one | Repair by Andrew after replacing 945-614
No | Visit | Complaint | RAME | Leakage at elmetII module. | Leakage at Met2 module on Crea B electrode chambers. | Replaced electrode connector spring on Crea B electrode. | SWO-275866 | Defect | HW | MetII Module | Electrode connector |  | Leakage at elmetII module. | Leakage at Met2 module on Crea B electrode chambers. | Replaced electrode connector spring on Crea B electrode.
No | Visit | Complaint | RAME | analyzer not recognizing QC on carousel. | Ampoules not breaking when scheduled to run auto QC. | Replaced FPC cable in autocheck module. Adjusted auto check laser. | SWO-275863 | Defect | HW | Auto Check Module | FPC cable | Ampoule breaker ? | analyzer not recognizing QC on carousel. | Ampoules not breaking when scheduled to run auto QC. | Replaced FPC cable in autocheck module. Adjusted auto check laser.
No | Visit | Complaint | RIND | Ca Calibration Error | Ca Membrane Need To Be Replace | Replaced Ca Membrane , Machine Passed 2 Point Calibration, Machine Working Fine | SWO-276080 | Assisted Replacement | Electrode | Membrane | Ca |  | Ca Calibration Error | Ca Membrane Need To Be Replace | Replaced Ca Membrane , Machine Passed 2 Point Calibration, Machine Working Fine
No | Visit | Complaint | RIND | Oxymetry values not coming. | Hemolyser to replace. | Hemolyser replaced. and analyser Calibrated. Analyser Serviced and calibrated properly. Handed over in a fine working condition. | SWO-276487 | Defect | HW | Oxi module | Hemolyzer |  | Oxymetry values not coming. | Hemolyser to replace. | Hemolyser replaced. and analyser Calibrated. Analyser Serviced and calibrated properly. Handed over in a fine working condition.
No | 2 | Y | RCZE | Blank | neaktivní èteèka pod autocheckem, nepøenáší do LISu - pøes pøipojený pøevadìè, |  | D392338744317470 | Setup | SW | His/Lis | His/Lis | LEH Mail, Barcode or His/Lis?? | Blank | Inactive timer Under the AutoCheck, not transferred to the Press, or connected to the receiver,
Yes | Visit | Complaint | RAME | pH & Ca failing out high sensitivity calibrations. | No investigation needed. | Delivering replacement Ca & Ref membranes to Earl Bonner that were used up during a previous call. Also had pre go-live meeting. | SWO-275842 | Re classified RMED | Courtesy |  |  |  | pH & Ca failing out high sensitivity calibrations. | No investigation needed. | Delivering replacement Ca & Ref membranes to Earl Bonner that were used up during a previous call. Also had pre go-live meeting.
No | Visit | Complaint | RIND | Display collar improper like fluorescence. | Display unit to change. | Display Unit replaced. Working fine. Analyser Serviced and calibrated properly. Handed over in a fine working condition. | SWO-276488 | [5 empty] | Display collar improper like fluorescence. | Display unit to change. | Display Unit replaced. Working fine. Analyser Serviced and calibrated properly. Handed over in a fine working condition.
No | Visit | Complaint | RCHN | 报错1029 | 报错1029，借903-364,903-236,944-216（13434缺货），944-258（13238），944-212（13499缺货）各一个720 | 更换903-364后修复 by Andrew | SWO-275244 | Defect | HW | Process Module | Optical Unit | 1029 | 1029 error | 1029, loan 903-364,903-236, 944-216 (13434 out of stock), 944-258 (13238), 944-212 (13499 out of stock) each 720 | Repair by Andrew after replacing 903-364
No | Visit | Complaint | RCHN | 打印机无法打印 | 打印机无法打印，借903-001一个 |  | SWO-276512 | [5 empty] | The printer can't print | Printer can't print, borrow 903-001 one
No | Visit | Complaint | RCHN | 触摸屏失灵 | 触摸屏失灵，借552-036一个 | 更换552-036后修复 by Andrew | SWO-276493 | [5 empty] | Touch screen failure | Touch screen failed, borrow 552-036 a | Repair by Andrew after replacement 552-036
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | pO2,LAC定标不过 0378 | 更换测试卡后修复 | SWO-276618 | [5 empty] | Test card, reagent package failure | pO2, LAC calibration is no more than 0378 | Repair after replacing test card
No | Visit | Complaint | RCHN | 处理包芯片信息错误 | 处理包芯片信息错误 | 更换试剂包后修复 | SWO-276586 | Assisted Replacement | Consumables | SP | Chip read |  | Processing packet chip information errors | Processing packet chip information errors | Repair after replacement of reagent kit
No | Visit | Complaint | RCHN | 报错2167 | 报错2167，借903-138一个 |  | SWO-276519 | Defect | HW | Incubator module | PCB, Incubator | 2167 | 2167 error | 2167, loan 903-138 one
Yes | Visit | Complaint | RDEU | Vakuum Fehler | Sichtkontrolle | Rep zus. Mit Jahreswartung | SWO-277175 | [5 empty] | Vacuum error | Visual inspection | Rep to. With annual maintenance
Yes | Visit | Complaint | RDEU | 30.11.2017 16:30Uhr Frau Trote 059311523970: Der Drucker arbeitet nicht | Auslieferung nach Rep. | Gerät nach Reparatur ausgeliefert. Startphase , kal und QC abgewartet. | SWO-277832 | Re classified RMED | Sameday, return after repair | See SWO-277831 |  |  | Nov 30, 2017 4:30 pm Ms Trote 059311523970: The printer is not working | Extradition to Rep. | Device delivered after repair. start phase, kal and QC waited.
No | Visit | Complaint | RDEU | 30.11.2017 16:30Uhr Frau Trote 059311523970: Der Drucker arbeitet nicht | Drucker druckt nicht - Lac und GLu kal Fehler | Gerät demontiert, CPU Unit ein und wieder ausgebaut. - Fehler war erst weg , kam aber dann wieder. Druckerkabel aus anderem Gerät in der GS ausgebaut und in das defekte Gerät eingebaut. Drucker Funktion wieder ok. Diverse kal und QC durchgeführt, - SC erneuert. | SWO-277831 | Defect | HW | DMS_PC | Cable, Printer | Customer claim revisit | Nov 30, 2017 4:30 pm Ms Trote 059311523970: The printer is not working | Printer does not print - Lac and GLu kal errors | device, CPU Unit on and removed. - Mistake was gone first, but then came back. Printer cable removed from other device in the GS and installed in the defective device. Printer function ok again. Various kal and QC performed, - SC renewed.
No | 5 | Y | RPAC | Blank | Magnet |  | D391DF7A44639A0F | Defect | HW | Mixer Unit | Mixer Unit | Magnet | Blank | Magnet
Yes | 8 | Y | RPAC | Blank | Solution pack failures |  | D3920EC644455CB1 | [5 empty] | Blank | Solution pack failures
Yes | 9 | Y | RPAC | Blank | Replaced failed SC |  | D391DB57443A9A0A | [5 empty] | Blank | Replaced failed SC
No | Visit | Complaint | RLTD | Miriam Martin on 0131 242 2544 She also said that the ABL8xx is not working either but this is no longer showing in CRM so I was unable to log a separate case. The ABL8xx is showing 2 point calibration failed and not accepted and all PH boxes are showing in red. | Leakage current in SP detected, error 0964 Customer has already replaced SP several times | Checked analyser in service programs and found that leakage current was coming from SC, replaced SC with new and leakage currents returned to normal levels | SWO-276942 | Assisted Replacement | SP | SP | Leak current | 0964 | Miriam Martin on 0131 242 2544 She also said that the ABL8xx is not working either but this is no longer showing in CRM so I was unable to log a separate case. The ABL8xx is showing 2 point calibration failed and not accepted and all PH boxes are showing in red. | Leakage current in SP detected, error 0964 Customer has already replaced SP several times | Checked analyser in service programs and found that leakage current was coming from SC, replaced SC with new and leakage currents returned to normal levels
No | 2 | Y | RPOL | Blank | Wymiana uszczelki wlotu. |  | D391C4FD4453D267 | Defect | HW | Auto Inlet module | Inlet gasket |  | Blank | Replacing the inlet gasket.
No | 4 | Y | RPOL | Blank | Wymiana wykrzywionej ig³y, pomiary kontorne w normie, analziator sprawny technicznie. |  | D391DE9A442D073B | Replace | HW | Auto Inlet module | Probe bend |  | Blank | Replacement of curved needle, standard and technically efficient analysers.
No | Visit | Complaint | RFRA | Erreur de mesures PH et PO2 | Retour atelier Écran fendu | Retour atelier | SWO-277068 | [5 empty] | Error of PH and PO2 measurements | Return workshop Split screen | Return workshop
No | Visit | Complaint | RSCH | Blank | Impossible d’initialiser le pack de solutions et le sensor de mesure Demande de garantie 2X SP et 2X SC | Changement de la date intempestive e.o Nettoyage de la section fluidique e.o Réinstallation du software système e.o Configuration du « Date & Time » e.o Cycle de qualité e.o | SWO-277563 | [5 empty] | Blank | Can't start the solution pack and sensor measurement 2X SP and 2X SC warranty request | Change of untimely date e.o Cleaning the fluidsection e.o Reinstallation of the e.o system software Setting up the "Date and Time" e.o E.o quality cycle
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | PO2质控偏高 0589 | 更换测试卡正常。 | SWO-276765 | [5 empty] | Test card, reagent package failure | PO2 QC HIGH 0589 | The replacement test card is OK.
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 压力测试真空错误 1063 | 更换测试卡 | SWO-276749 | [5 empty] | Test card, reagent package failure | Pressure test vacuum error 1063 | Replace the test card
No | Visit | Complaint | RCHN | 一直报泄漏 | 一直报泄漏，借903-129一个 |  | SWO-276663 | Defect | HW | Basic Unit | Pinch valve |  | Keep reporting leaks. | have been reported leak, borrow 903-129 a
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | pco2定标失败，0378；LAC质控失败，0210 | 连续更换2张测试卡后修复 | SWO-277441 | [5 empty] | Test card, reagent package failure | pco2 calibration failure, 0378; LAC QC failure, 0210 | Repair after 2 consecutive test cards
No | Visit | Complaint | RCHN | 报错1029 | 报错1029，借903-364一个 |  | SWO-276596 | Defect | HW | Process Module | Optical Unit | 1029 | 1029 error | 1029, loan 903-364 one
No | Visit | Complaint | RCHN | 采血针弯曲 | 采血针弯曲，借903-584一个 |  | SWO-276654 | Defect | HW | Cover and Chassis | Tube holder |  | Blood needle bending | Blood needle bend, borrow 903-584 one
Yes | Visit | Complaint | RIND | Calibration error. | Check the system. check the logs. | Check the system. check the logs.flush cassette multiple time. calibrate the system. Calibration error multiple time. replace the sp9 & calibrate the system. calibration are in ranges. Machine under observation. | SWO-280109 | [5 empty] | Calibration error. | Check the system. check the logs. | Check the system. check the logs.flush cassette multiple time. calibrate the system. Calibration error multiple time. replace the sp9 & calibrate the system. calibration are in ranges. Machine under observation.
Yes | 16 | Y | RPNZ | Blank | Replacement pCO2 membranes (also for DCC BAIT382RV) |  | D391BD6A44262D65 | [5 empty] | Blank | Replacement pCO2 membranes (also for DCC BAIT382RV)
No | 5 | Y | RCZE | Blank | chyby v qc u oxi, trvá i po èištìní a odstranìní proteinù, Beòovská požaduje øešení, mají prý audit |  | D392388244137475 | Revisit required | HW | Oxi module | Oxi module |  | Blank | Errors in the QC, it persists after cleaning and removal of proteins, the Beòovská requires solutions,
No | Visit | Complaint | RAME | Halogen Lamp Errors | Troubleshot and found the halogen lamp was out of adjustment | Adjusted halogen lamp, and this resolved the problem | SWO-276988 | Adjustment | HW | Oxi module | Lamp unit | Halogen | Halogen Lamp Errors | Troubleshot and found the halogen lamp was out of adjustment | Adjusted halogen lamp, and this resolved the problem
No | Visit | Complaint | RAME | Received Neon no too high error 315 | Neon number too high - checked oxi system possible errors lamp hemolyzer spectrometer or pcb | Replaced module due to multiple errors on a new unit | SWO-277362 | Defect | HW | Oxi module | Oxi module, Replaced |  | Received Neon no too high error 315 | Neon number too high - checked oxi system possible errors lamp hemolyzer spectrometer or pcb | Replaced module due to multiple errors on a new unit
No | 3 | Y | RPOL | Blank | Wyczyszczono uk³ad optyczny. Testy analizatora. Kontrola QC, aparat sprawny. |  | D3920EA844147E00 | Cleaned | HW | Process Module | Process wheel |  | Blank | The optical disc is cleared. Analyzer tests. QC control, the camera is functional.
No | Visit | Complaint | RDEU | 27.07.2017; Sr. Anett; Tel.: 03381791614; CRP QC Messfehler 28.07.2017 10:30 Sr. Nicole; Tel.: 03381791612; weiterhin CRP QC Messfehler 31.07.2017: Rücksprache mit KerZi - neues QC wird bestellt, telefonische Info an Kunden 03.08.2017, 11:36 Uhr ==> Anrufer Pfl. Sebastian, Tel.: 03381791611 ==> LQC-Messfehler CRP 04.08.2017, 12:04 Uhr ==> Anrufer Schw. Anett, Tel.: 03381791614 ==> LQC-Fehler CRP besteht trotz neuer Charge weiterhin | Lieferung Leihgerät AQT90 (838R0099N004) wg. weiterhin bestehenden CRP-Fehlern !! | Lieferung Leihgerät, Konfiguration, Netzwerkanbindung, Aqureverbindung, QK ok | SWO-276770 | Revisit required | Unknown | Unknown | Unknown, installed Loaner |  | 27.07.2017; Sr. Anett; Tel.: 03381791614; CRP QC Measurement Error 07/28/2017 10:30 Sr. Nicole; Tel.: 03381791612; further CRP QC measurement error 31.07.2017: Consultation with KerZi - new QC is ordered, telephone information to customers 03.08.2017, 11:36 a.m. ==> Caller Pfl. Sebastian, Tel.: 0338179111 ==> LQC measurement error CRP 04.08.2017, 12:04 o'clock ==> Caller Schw. Anett, tel.: 03381791614 ==> LQC error CRP persists despite new batch | Delivery of rental device AQT90 (838R0099N004) wg. existing CRP errors !! | Delivery of rental device, configuration, network connection, aqure connection, QK ok
Yes | Visit | Complaint | RFRA | Erreur de vide | Vérification: en service( fuite du vide), circuit fluide. Chgt joint d'introduction et connexion. Intervertir les packs des 2 ABL90 | Chgt du pack de solution | SWO-280427 | [5 empty] | Empty error | Check: in service (leakfrom from the void), fluid circuit. Chgt introductory joint and connection. Switch the packs of the 2 ABL90s | Chgt of the solution pack
No | Visit | Complaint | RDEU | 11.12.2017, 12:57, Fr. Stelzer, Tel. 02088813188: Benutzerintervention erforderlich | inhomogene Lösungen, Intervent. erforderlich, LP können nicht installiert werden, YW 55 und YW56 ohne Pumpenschlauch, | neues LP mit anderer Lotnr. angeschlossen, Gerät kalibriert, neue SC angeschlossen, | SWO-277507 | Assisted Replacement | SP | SP | Pump Tubes | Tubes missing | 12/11/2017, 12:57 PM, Ms. Stelzer, Tel. 02088813188: User intervention required | inhomogeneous solutions, intervening. required, LP cannot be installed, YW 55 and YW56 without pump hose, | new LP with different Lotnr. connected, device calibrated, new SC connected,
No | 3 | Y | RPAC | Blank | Lac failed on calibration |  | D3951509302C02BB | Assisted Replacement | SC | SC | Lac |  | Blank | Lac failed on calibration
No | Visit | Complaint | RLTD | pO2 failing calibration with status errors frequently. | Carried out thorough decontamination of sample pathway within the analyser. Complete sample pathway flushed, including scrubbing the line, with hypochlorite solution for 30 minutes. Wet section then removed and inlet probe, gaskets and tubes soaked in hypochlorite, all metal connection pipes also cleaned with hypochlorite solution. All areas cleaned with hypochlorite flushed through with sterile water. Analyser re-assembled and new Sensor CAssette and Solution pack fitted. Verified start-up of analyser - all ok. | Decontaminated analyser sample pathway. | SWO-277896 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway | Decontamination | pO2 failing calibration with status errors frequently. | Carried out thorough decontamination of sample pathway within the analyser. Complete sample pathway flushed, including scrubbing the line, with hypochlorite solution for 30 minutes. Wet section then removed and inlet probe, gaskets and tubes soaked in hypochlorite, all metal connection pipes also cleaned with hypochlorite solution. All areas cleaned with hypochlorite flushed through with sterile water. Analyser re-assembled and new Sensor CAssette and Solution pack fitted. Verified start-up of analyser - all ok. | Decontaminated analyser sample pathway.
No | Visit | Complaint | RLTD | Stephanie Sandy 07863 330762 | Carried out thorough decontamination of sample pathway within the analyser. Complete sample pathway flushed, including scrubbing the line, with hypochlorite solution for 30 minutes. Wet section then removed and inlet probe, gaskets and tubes soaked in hypochlorite, all metal connection pipes also cleaned with hypochlorite solution. All ares cleaned with hypochlorite flushed through with sterile water. Analyser re-assembled and new Sensor Cassette and Solution pack fitted. Verified start-up of analyser - all ok. | Analyser sample pathway decontaminated. | SWO-277895 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway | Decontamination | Stephanie Sandy 07863 330762 | Carried out thorough decontamination of sample pathway within the analyser. Complete sample pathway flushed, including scrubbing the line, with hypochlorite solution for 30 minutes. Wet section then removed and inlet probe, gaskets and tubes soaked in hypochlorite, all metal connection pipes also cleaned with hypochlorite solution. All ares cleaned with hypochlorite flushed through with sterile water. Analyser re-assembled and new Sensor Cassette and Solution pack fitted. Verified start-up of analyser - all ok. | Analyser sample pathway decontaminated.
No | Visit | Complaint | RLTD | Fault observed in logs during site visit. | Carried out thorough decontamination of sample pathway within the analyser. Complete sample pathway flushed, including scrubbing the line, with hypochlorite solution for 30 minutes. Wet section then removed and inlet probe, gaskets and tubes soaked in hypochlorite, all metal connection pipes also cleaned with hypochlorite solution. All areas cleaned with hypochlorite flushed through with sterile water. Analyser re-assembled and new Sensor CAssette and Solution pack fitted. Verified start-up of analyser - all ok. | Analyser sample pathway decontaminated | SWO-277897 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway | Decontamination | Fault observed in logs during site visit. | Carried out thorough decontamination of sample pathway within the analyser. Complete sample pathway flushed, including scrubbing the line, with hypochlorite solution for 30 minutes. Wet section then removed and inlet probe, gaskets and tubes soaked in hypochlorite, all metal connection pipes also cleaned with hypochlorite solution. All areas cleaned with hypochlorite flushed through with sterile water. Analyser re-assembled and new Sensor CAssette and Solution pack fitted. Verified start-up of analyser - all ok. | Analyser sample pathway decontaminated
No | 3 | Y | RPAC | Blank | Lac failed in new SC |  | D395149D301002B7 | Assisted Replacement | SC | SC | Lac |  | Blank | Lac failed in new SC
No | 6 | Y | RPAC | Blank | Failing SP 964:Leak current in relation to solution pack detected |  | D3931C02BA624AAA | Assisted Replacement | SP | SP | Leak current |  | Blank | Failing SP 964:Leak current in relation to solution pack detected
No | Visit | Complaint | RSAF | oxi hadware error on calibrations intermittently | analyzer repaired at RSAF workshop | Returned analyzer and start up tests ok. | SWO-278989 | Defect | HW | Oxi module | Oxi module | Neon | oxi hadware error on calibrations intermittently | analyzer repaired at RSAF workshop | Returned analyzer and start up tests ok.
No | Visit | Complaint | RDEU | 12.12.2017 Herr Soest 020172384285 meldet, dass der Analysator immer wieder zum zum Spülen auffordert. | Fehler 1352 Verdacht auf Gerinnsel im Oxymetriemodul . | Fehlersuche. Oxi-Module-Status im Service-Menü: 1,26 [Soll: 0,80 - 1,20]. Mehrfaches Spülen bringt kaum Besserung (zuletzt 1,19: grenzwertig ok). Fließwege überprüft. Durchführung einer Jahreswartung. Anschließend liegt der entsprechende Wert bei 0,96 (i.O.) | SWO-277724 | Defect | HW | Basic Unit | PM Kit |  | 12.12.2017 Mr. Soest 020172384285 reports that the analyzer repeatedly asks for rinsing. | Error 1352 Suspected clot in the oxymetry module . | Troubleshooting. Oxi module status in the service menu: 1.26 [should: 0.80 - 1.20]. Multiple rinsing brings hardly any improvement (last 1.19: borderline ok). Flow paths checked. Execution of an annual maintenance. The corresponding value is then 0.96 (usually)
Yes | Visit | Complaint | RDEU | 12.12.2017 11:00 Sr. Debby 03070604145: Gerät ist weihin nach LP-Wechsel undicht | Undichtigkeit unter LP | LP ersetzt ( Kundenmaterial), Lösung kommt am Übergang Abfallstutzen zum Adapter LP ( Abfall) heraus beanstandetes LP Lot 248701 | SWO-277920 | [5 empty] | 12.12.2017 11:00 Sr. Debby 03070604145: Device is weihin after LP change leak | Leakage under LP | LP replaces (customer material), solution comes at the transition waste nozzle to adapter LP (waste) challenged LP Lot 248701
No | Visit | Complaint | RIND | pO2 not calibrating. | pO2 not calibrating. | Tried by flush and reinstall. Reset and reinstall but problem repeating. need to replace SC. Asked new. | SWO-279314 | [5 empty] | pO2 not calibrating. | pO2 not calibrating. | Tried by flush and reinstall. Reset and reinstall but problem repeating. need to replace SC. Asked new.
No | 3 | Y | RPAC | Blank | Gas |  | D3935B62BA01B99F | [5 empty] | Blank | Gas
Yes | Visit | Complaint | RCHN | 进样门故障 | 进样门故障，借903-028,903-117,903-198各一个 | 更换903-117后修复by tim | SWO-276525 | [5 empty] | Sample door failure | Sample door fault, borrow 903-028, 903-117, 903-198 each | Repair by tim after replacement 903-117
Yes | Visit | Complaint | RLTD | Hannah Tel 01865 221596 | The analyser was in standby because of a full waste container. Replaced the waste container and allowed the analyser to start-up and run its pending QCs. | The QCs were still failing so the annual service was performed however this didn't rectify the fault. I found that the Crea electrodes on the analyser were not the ones fitted on 14 $11, I cleaned the contacts on these electrodes and re ran Cals & QCs but Crea was still failing level 1 QC, new electrodes have been ordered. | SWO-277786 | [5 empty] | Hannah Tel 01865 221596 | The analyser was in standby because of a full waste container. Replaced the waste container and allowed the analyser to start-up and run its pending QCs. | The QCs were still failing so the annual service was performed however this didn't rectify the fault. I found that the Crea electrodes on the analyser were not the ones fitted on 14 $11, I cleaned the contacts on these electrodes and re ran Cals & QCs but Crea was still failing level 1 QC, new electrodes have been ordered.
No | Visit | Complaint | RLTD | Replace PCB to see if it helps with the start button issue | A colleague had mentioned that they had changed the inlet flap PCB on an analyser that was having problems with the start button not working and since then the problem has not come back. | Replaced the inlet flap PCB to see if it resolves the issue with the start button not working. | SWO-277785 | Defect | HW | Inlet module | PCB, Flap |  | Replace PCB to see if it helps with the start button issue | A colleague had mentioned that they had changed the inlet flap PCB on an analyser that was having problems with the start button not working and since then the problem has not come back. | Replaced the inlet flap PCB to see if it resolves the issue with the start button not working.
Yes | Visit | Complaint | RAME | Customer called with the analyzer will not run QC. The autocheck will not recognize the QC carousel. | Checked the autocheck in service to find the QC carousel is not being recognized. The sensor for the carousel dropped and it will not recognize the carousel when installed. | Replaced the autocheck module for fault isolation. The sensor for the carousel is faulty and will need replaced. | SWO-277932 | [5 empty] | Customer called with the analyzer will not run QC. The autocheck will not recognize the QC carousel. | Checked the autocheck in service to find the QC carousel is not being recognized. The sensor for the carousel dropped and it will not recognize the carousel when installed. | Replaced the autocheck module for fault isolation. The sensor for the carousel is faulty and will need replaced.
No | Visit | Complaint | RESP | ERROR GENERAL PC | REVISION DIFERENTES MODULOS DE PC | SE SUSTITUYEN LOS DIFERENTES ELEMENTOS QUE SE DESCRIBEN POR ESTAR DEFECTUOSOS. AJUSTE Y RECONFIGURACION DE TODO EL SISTEMA. CALIBRADO Y VERIFICADO CON Q.C. | SWO-281540 | Defect | HW | DMS_PC | CPU |  | ERROR GENERAL PC | REVIEW DIFFERENT PC MODULE | THE DIFFERENT ELEMENTS DESCRIBED BY BEING DEFECTIVE ARE SUBSTITUTED. ADJUSTMENT AND RELOCATION OF THE ENTIRE SYSTEM. CALIBRATED AND VERIFIED WITH Q.C.
No | Visit | Complaint | RAME | Leak detected, unable to clear issue | Client advised that they are unable to bypass leak detection and ref chamber is full of blood | Cleaned ref measuring chamber inside and out. Leak detect d again. Verified membranes, potassium membrane missing white o-ring. Swapped membrane for new with o-ring. Leak detected issue resolved. | SWO-278580 | Defect | Electrode | Membrane | Sealing ring | Sealing ring, K+ | Leak detected, unable to clear issue | Client advised that they are unable to bypass leak detection and ref chamber is full of blood | Cleaned ref measuring chamber inside and out. Leak detect d again. Verified membranes, potassium membrane missing white o-ring. Swapped membrane for new with o-ring. Leak detected issue resolved.
No | Visit | Complaint | RAME | Printer not working | No Print | 12 $6 $17 Customer reported no print, Confirmed printer attempts to print but the paper will not feed. 12 $12 $17 Replaced printer unit. | SWO-279934 | Defect | HW | Printer Module | Printer Module |  | Printer not working | No Print | 12 $6 $17 Customer reported no print, Confirmed printer attempts to print but the paper will not feed. 12 $12 $17 Replaced printer unit.
No | Visit | Complaint | RAME | failed vk kit | Hemolyzer not showing data, | Installed hemolyzer after calverification shutout period to test level 3 . Adjusted lamp ran thb cal , ran linearity pass | SWO-277923 | Defect | HW | Oxi module | Hemolyzer |  | failed vk kit | Hemolyzer not showing data, | Installed hemolyzer after calverification shutout period to test level 3 . Adjusted lamp ran thb cal , ran linearity pass
No | Visit | Complaint | RAME | Reference and pH electrode cambers clotted | Checked data logs, checked electrodes and electrode chambers, | Removed pH and reference electrode and cleaned clotting from electrode chamber and electrode housing replaced pH seal ring, replaced reference membrane, replaced all pump tubing | SWO-277212 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | Reference and pH electrode cambers clotted | Checked data logs, checked electrodes and electrode chambers, | Removed pH and reference electrode and cleaned clotting from electrode chamber and electrode housing replaced pH seal ring, replaced reference membrane, replaced all pump tubing
No | Visit | Complaint | RAME | Cream out of cal, once stable 2,3 cals later sensitivity 1 is above 15 limit | Crea sensitivity 15.9, | Replaced electrodes, cal in 3 cals later out at. 15.2, replaced membrane, lot 772 good | SWO-277927 | Assisted Replacement | Electrode | Membrane | Crea |  | Cream out of cal, once stable 2,3 cals later sensitivity 1 is above 15 limit | Crea sensitivity 15.9, | Replaced electrodes, cal in 3 cals later out at. 15.2, replaced membrane, lot 772 good
No | Visit | Complaint | RSCH | Spülfehler | Spülfehler | Ventil V6 weisst 26 Fehler auf. Ventil V6 ausgetauscht Funktionskontrolle i.o | SWO-277409 | Defect | HW | Fluidic manifold | Valve | V6 | Spülfehler | Spülfehler | Valve V6 whites out 26 faults. Valve V6 replaced Functional control i.o
No | Visit | Complaint | RDEU | Daten werden nicht übertragen | Daten werden nicht übertragen | IP Daten überprüft - sind korrekt Gerät neu gestartet - Daten werden wieder übertragen | SWO-277690 | Defect | SW | RMED SW | Off/On |  | Data is not transferred | Data is not transferred | IP data checked - are correct Device restarted - data is transferred again
No | Visit | Complaint | RDEU | Seit 2 Tagen werden Proben abgebrochen Nach Wechsel von Einlass- und Anschlussdichtung wird Leckage angezeigt | Patientenmessungen werden abgebrochen => Kal und QC i.O. Leckage erkannt => Drucküberprüfung scheitert Wechsel Einlass- und Anschlussdichtung ohne Besserung | Überprüfung Quetschventil => Ventil hängt und schaltet nicht sauber Quetschventil ausgebaut und gereinigt Geräteneustart i.O. Durchführung aller Fluid Progamme i.O. Testmessung Spritze und Kapillare mit tHb Kalibrator ohne Abbruch | SWO-278268 | Revisit required | HW | Oxi module | Pinch valve |  | Samples have been cancelled for 2 days Leakage is displayed after changing the inlet and connection seal | Patient measurements are canceled => Kal and QC i.O. Leakage detected => Pressure check fails Change inlet and connection seal without improvement | Check crushing valve => Valve hangs and does not switch cleanly Squeezing valve removed and cleaned Device restart i.O. Execution of all Fluid Progamme i.O. Test measurement syringe and capillary with tHb calibrator without demolition
No | Visit | Complaint | RSCH | Rilevata dispersione di corrente connessa al SP | Rilevata dispersione nel SP. Apparecchio si blocca. | Sostituito Solution Pack ed eseguita manutenzione annuale. Rimesso in funzione, apparecchio testato in ordine. Sostituito emolizzatore e SC. | SWO-278562 | Assisted Replacement | SP | SP | Leak current |  | Detected current dispersion connected to the SP | Detected dispersion in the SP. Device hangs. | Replaced Solution Pack and performed annual maintenance. Put back into operation, device tested in order. Replaced molizer and SC.
No | Visit | Complaint | RDEU | Patientenbarcode kann nicht eingelesen werden | Fehlersuche vor Ort | Patientenbarcode kann nicht eingelesen werden. Roma angepasst. Danach könnte Barcode wieder eingelesen werden. Mehrere Kalibrierungen durchgeführt. QCs für alle Level durchgeführt. ABL wieder messbereit. | SWO-278763 | Setup | SW | RMED SW | roma.ini | Barcode | Patient bar code cannot be read in | Troubleshooting on site | Patient bar code cannot be read in. Roma. After that, barcode could be read again. Multiple calibrations performed. QCs for all levels. ABL ready to be measured again.
No | Visit | Complaint | RLTD | Phillip Harding Payne - Tel. 01206 742493 | Failing fluid transport, pressure value. Also failing to aspirate samples. | Customer had already replaced pluton pack and inlet gasket. Replaced connection gasket and waiting for analyser to start up. Analyser failing to aspirate when sample is presented, asked user not to be so forceful when introducing sample, users a | SWO-279093 | Defect | HW | Basic Unit | PM Kit |  | Phillip Harding Payne - Tel. 01206 742493 | Failing fluid transport, pressure value. Also failing to aspirate samples. | Customer had already replaced pluton pack and inlet gasket. Replaced connection gasket and waiting for analyser to start up. Analyser failing to aspirate when sample is presented, asked user not to be so forceful when introducing sample, users a
No | Visit | Complaint | RIND | pO2 not calibrating. | SC to repalce. | SC and SP replaced. Analyser Serviced and calibrated. Handed over in a fine working condition. | SWO-279316 | [5 empty] | pO2 not calibrating. | SC to repalce. | SC and SP replaced. Analyser Serviced and calibrated. Handed over in a fine working condition.
Yes | Visit | Complaint | RCHN | 报错5784 | 报错5784 | 更换PCT（13041）卡后修复 | SWO-278041 | [5 empty] | 5784 | 5784 | Repair after replacement of PCT (13041) card
No | Visit | Complaint | RIND | Touch Screen Problem. | Touch Screen Problem. | Replaced the Display unit restarted instrument found instrument is working its under observation for one week. | SWO-278714 | Defect | HW | DMS_PC | Touch screen |  | Touch Screen Problem. | Touch Screen Problem. | Replaced the Display unit restarted instrument found instrument is working its under observation for one week.
No | Visit | Complaint | RAME | Oxi Parameters Failing on Level 3 of VK Kit | Reviewed all results from linearity and found only the level 3 oxi results to be off. Thiese results were also accompanied by error 581. Ran complete oxi test and found all passed. | Replaced hemolyzer first and tried level 3 from VK kit. Found same issues. Replaced spectrometer but had out of box failure. Could not measure oxi, put original spectrometer back in. Performed lamp adjust, thb cals. Ran level 3 VK again and found the same result, errors 581. Analyzer has also entered the lockout period so customer must wait the 14 days before running linearity again. Will speak with Jim Gibbons in tech support about this issue. Customer may have an option for level 3 of the VK kit. I will discuss with customer how to proceed once the lockout period is over. | SWO-282818 | Revisit required | HW | Oxi module | Oxi module | VRK Test | Oxi Parameters Failing on Level 3 of VK Kit | Reviewed all results from linearity and found only the level 3 oxi results to be off. Thiese results were also accompanied by error 581. Ran complete oxi test and found all passed. | Replaced hemolyzer first and tried level 3 from VK kit. Found same issues. Replaced spectrometer but had out of box failure. Could not measure oxi, put original spectrometer back in. Performed lamp adjust, thb cals. Ran level 3 VK again and found the same result, errors 581. Analyzer has also entered the lockout period so customer must wait the 14 days before running linearity again. Will speak with Jim Gibbons in tech support about this issue. Customer may have an option for level 3 of the VK kit. I will discuss with customer how to proceed once the lockout period is over.
No | Visit | Complaint | RAME | Leak current detected on analyzer. | Arrived onsite and found error 963 leak current detected. Found leak at grey tubing between the BG and MET module. Nipple broken at measuring chamber. | Replaced measuring chamber and liquid sensor. | SWO-278556 | Defect | HW | pH/BG Module | Measuring chamber | New analyzer | Leak current detected on analyzer. | Arrived onsite and found error 963 leak current detected. Found leak at grey tubing between the BG and MET module. Nipple broken at measuring chamber. | Replaced measuring chamber and liquid sensor.
No | Visit | Complaint | RAME | QC identification issues | Can't select external QC | Deleted duplicate lots in solutions file. | SWO-279483 | Setup | SW | RMED SW | QC Schedule |  | QC identification issues | Can't select external QC | Deleted duplicate lots in solutions file.
No | Visit | Complaint | RDEU | 14.12.2017 Pfl. Johannes 0213152953200 meldet einen Leckagefehler | Leckage erkannt, Flüssigkeit tritt am Schlauch vor dem Oximodul aus, Schlauch ist def. | Ventilschlauch erneuert, Gerät kalibriert, | SWO-279209 | Defect | HW | Oxi module | Pinch valve tube |  | 14.12.2017 Pfl. Johannes 0213152953200 reports a leakage error | Leakage detected, liquid leaks on the hose in front of the oximodul, hose is def. | Valve hose renewed, device calibrated,
No | Visit | Complaint | RDEU | 13.12.2017, Herr Knobloch, Tel: 05115309961: keine Anwender Anmeldung möglich; es wurde noch nicht damit gearbeitet 13.12.2017; 12.20 Uhr; Hr. Knobloch; Tel.: 05115309961; Fehlermeldung Zeitüberschreitung bei der Pat. Profilabfrage | Keine Datenabfrage | Einstellungen geprüft, ASTM.ini V1.6 neu installiert, Funktionskontrolle ok | SWO-279873 | Setup | SW | RMED SW | ASTM.INI |  | 13.12.2017, Mr. Knobloch, Tel: 05115309961: no user registration possible; it has not yet been worked on 13.12.2017; 12:20 p.m.; Mr. Knobloch; Tel.: 05115309961; Error message Timed out at the Pat. Profile query | No data query | Settings checked, ASTM.ini V1.6 reinstalled, function control ok
Yes | Visit | Complaint | RBNL | Er is een metalen deel van de aandrukrol verdwenen. Hoe?? Er zal een nieuwe printer geïnstalleerd moeten worden. Bestellen 910-286 Opdrachtnummer aangevraagd. | Kreeg gisteren een mailbericht dat de printer weer defect is en tijdelijk door medewerkers is gerepareerd. Messing busje zit verkeerdom op het asje waardoor het neit geborgd wordt. | Busjeandersom op as geplaatst. Print gemaakt. ok. | SWO-279165 | [5 empty] | A metal part of the pressure roller has disappeared. How?? A new printer will need to be installed. Ordering 910-286 Order number requested. | Received a mail message yesterday that the printer is defective and temporarily repaired by employees. Brass canister is traffic on the loin which makes the Neit secured. | Busjevice versa on as placed. Print made. Ok.
No | Visit | Complaint | RLTD | Lynn Tel 01582 497991 | Inlet gasket sticking to probe. | Checked probe, all seems ok. Advised customer if issue persists they should change inlet gasket and holder. | SWO-282781 | Adjustment | HW | Auto Inlet module | Inlet gasket | Capillary | Lynn Tel 01582 497991 | Inlet gasket sticking to probe. | Checked probe, all seems ok. Advised customer if issue persists they should change inlet gasket and holder.
No | Visit | Complaint | RLTD | Inlet not raising | Inlet not raising | Couldn't see any issues however inlet would not raise, even with inlet removed. Stripped entire inlet mechanism and reassembled, now ok. Customer to monitor. | SWO-279094 | Adjustment | HW | Auto Inlet module | Auto Inlet module |  | Inlet not raising | Inlet not raising | Couldn't see any issues however inlet would not raise, even with inlet removed. Stripped entire inlet mechanism and reassembled, now ok. Customer to monitor.
No | 3 | Y | RPAC | Blank | Calibration failure on Glu |  | D3960DE2A62AB9A6 | Assisted Replacement | SC | SC | Glu |  | Blank | Calibration failure on Glu
No | Visit | Complaint | RDEU | 14.12.2017 9:15 Fr. Püttman 020644417351: Einlaß-Fehler | Einl.dichtung ist etwas schwergängig, | Einl.system ausgebaut, alles gereinigt und gängig gemacht, Probemessungen durchgeführt, | SWO-279210 | Defect | HW | Auto Inlet module | Gasket movement |  | 14.12.2017 9:15 Ms. Püttman 020644417351: Admission error | Single-seal is somewhat difficult, | single system removed, everything cleaned and made common, sample measurements carried out,
No | Visit | Complaint | RDEU | 12.12.2017, Klaus, Tel: 02265995914: keine Probenmessungen möglich, Fehlermeldunge nicht bekannt 13.12.2017; Sr. Katrin; Tel.: 02265995914; Proben werden abgebrochen | Fehler 0331 (keine Probe bei Ansaugung erkannt) mehrfach $ sehr oft, 1271 sporadisch und 1190 gestern zweimal | Fehlersuche: Gestern 2x Fehler 1190 (Einlass geöffnet), vermutlich aufgrund leicht verbogener Einlasssonde - Sonde gerichtet. Regelmäßig in den letzten Wochen Fehler 0331 (keine Probe erkannt) - laut Anwender wurden Pico50 Spritzen mit aufgesetzter Kappe angesetzt, dadurch war keine Aspiration möglich! Mehrfach Test-Messungen mit Clinitubes und mit Pico50 durchgeführt, Funktion einwandfrei. QC Messung, ok | SWO-279247 | Defect | HW | Auto Inlet module | Hall Element |  | 12.12.2017, Klaus, Tel: 0226595914: no sample measurements possible, error messages not known 13.12.2017; Sr. Katrin; Tel.: 02265995914; Samples are aborted | Error 0331 (no sample detected at suction) multiple times , 1271 sporadically and 1190 yesterday twice | Troubleshooting: Yesterday 2x error 1190 (inlet opened), probably due to slightly bent inlet probe - probe directed. Regularly in the last weeks error 0331 (no sample detected) - according to the user, Pico50 syringes were applied with the cap on, so no aspiration was possible! Multiple test measurements with Clinitubes and performed with Pico50, function flawless. QC measurement, ok
No | Visit | Complaint | RFRA | Erreurs CQ COOX | Test COOX Contrôle circuits fluides | Échange joint et sonde aspirations Nettoyage circuits fluides Calibration tHB | SWO-279301 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | COOX CQ Errors | COOX test Fluid circuit control | Joint exchange and probe aspirations Cleaning fluid circuits THB calibration
Yes | Visit | Complaint | RCHN | 仪器无法启动 | 仪器无法启动，借903-381，903-334各一个。 | 更换903-381，903-334后修复 bu tim | SWO-278192 | [5 empty] | The instrument does not start | The instrument cannot be started, borrowing one each at 903-381,903-334. | Repair bu tim after replacing 903-381,903-334
No | Visit | Complaint | RCHN | 定标无法通过 | 定标无法通过，借903-111一个 | 更换903-111后修复 | SWO-278721 | [5 empty] | The bid failed to pass | The bid cannot be passed, borrow 903-111 one | Repair after replacement 903-111
No | Visit | Complaint | RCHN | CA,K响应错误 | CA,K响应错误，借903-581一个 |  | SWO-278676 | Defect | HW | Basic Unit | PCB, Sensor interface |  | CA, K response error | CA, K response error, borrow 903-581 one
No | Visit | Complaint | RCHN | GLU,LAC定标不过 | GLU,LAC定标不过，借903-581一个 |  | SWO-278765 | Defect | HW | Basic Unit | PCB, Sensor interface |  | GLU, LAC calibration but | GLU, LAC calibration, borrow 903-581 a
No | Visit | Complaint | RESP | Error 05000028 TC Core (OMSG10): Gas valve error | Se detecta que el equipo no funciona correctamente. | Sustiuida las siguientes piezas: Tarjeta SD. Placa base, Regulador del GAS, Cámara de cablibración. Actualización de software a la versión 1.08. Comprobado buen funcionamiento del equipo. | SWO-279242 | [5 empty] | Error 05000028 TC Core (OMSG10): Gas valve error | The computer is found to be not working properly. | Replaced the following parts: SD card. Base plate, GAS regulator, Cablibration chamber. Software update to version 1.08. Proven good operation of the equipment.
Yes | Visit | Complaint | RCHN | 报错1068，1031，1029，1024，1025 | 报错1068，1031，1029，1024，1025，借903-364、903-210一个，944-212（13643、13632）各一张 定标卡有货 | 更换903-364，903-210各一个，重新定标正常。by tim | SWO-277460 | [5 empty] | Error 1068,1031,1029,1024,1025 | Error1068,1031,1029,1024,1025, borrow 903-364, 903-210 one, 944-212 (13643, 13632) each with a bid card available | Replace the 903-364, 903-210 each one, re-calibration normal. by tim
Yes | Visit | Complaint | RCHN | 报错3412 | 报错3412，借903-329一个 | 更换903-329后修复by tim | SWO-278059 | [5 empty] | Error 3412 | 3412, loan 903-329 one | Repair by tim after replacement 903-329
No | Visit | Complaint | RCHN | 报错1031 | 报错1031，借903-364一个 | 更换903-364后修复by tim | SWO-278038 | Defect | HW | Process Module | Optical Unit | 1031 | 1031 | 1031, loan 903-364 one | Repair by tim after replacing 903-364
Yes | Visit | Complaint | RKK | 突然リンスボトル交換画面になり、リンスボトルの再読み込みでリスタート可能になるが、1日に何回か同じようになってしまう。 | リンスボトル交換画面になる。 | リンス用ボトルイルミネーションPCB交換。動作チェック。 | SWO-284146 | [5 empty] | It suddenly becomes a rinse bottle replacement screen, and it becomes possible to restart by reloading the rinse bottle, but it becomes the same several times a day. | It becomes a rinse bottle replacement screen. | Bottle illumination PCB replacement for rinsing. Operation check.
No | Visit | Complaint | RIND | QC Ampoules not being scanned in Autocheck module | laser scan engine(910-304) to be replaced. | Replaced laser scan engine, Aligning laser position was not possible as displayed value remains fixed at 192 . | SWO-280387 | Defect | HW | Autocheck Module | Barcode reader, Autocheck Module |  | QC Ampoules not being scanned in Autocheck module | laser scan engine(910-304) to be replaced. | Replaced laser scan engine, Aligning laser position was not possible as displayed value remains fixed at 192 .
No | 15 | Y | RPNZ | Blank | Oximetry (0581- OXI Spectrum mismatch) and rinse errors |  | D395D0A4A632AB52 | Cleaned | Fluidic pathway | Oxi module | Fluidic pathway |  | Blank | Oximetry (0581- OXI Spectrum mismatch) and rinse errors
No | Visit | Complaint | RLTD | Ian | Rana QC, house comes up, ampoule is selected and broken, carousel then rotates and does not seem to find its position and house goes own and lid closes. | Remove and thoroughly clean autocheck module including main and laser scanning unit mirrors. Inspect all, PCBs and sensors no damage found, autocheck 14 pole cable in good condition. After calibrating red QC unstable on Glu, recalibrate analyser twice before red QC ok on Glu, had to then top up carousel and pO2 took 2 x 2 point calibrations and several QCs to come in on red level. | SWO-279899 | Cleaned | HW | Autocheck Module | Autocheck Module |  | Ian | Rana QC, house comes up, ampoule is selected and broken, carousel then rotates and does not seem to find its position and house goes own and lid closes. | Remove and thoroughly clean autocheck module including main and laser scanning unit mirrors. Inspect all, PCBs and sensors no damage found, autocheck 14 pole cable in good condition. After calibrating red QC unstable on Glu, recalibrate analyser twice before red QC ok on Glu, had to then top up carousel and pO2 took 2 x 2 point calibrations and several QCs to come in on red level.
No | Visit | Complaint | RIND | Oximetry hardware problem,neon lamp intensity low. | Neon lamp need to be replaced | Replace the the neon lamp,perform the complete oximetry test,it got passed successfully,perform 2 point cal & got the desired result | SWO-280080 | Defect | HW | Oxi module | Lamp unit | Neon Lamp, Low | Oximetry hardware problem,neon lamp intensity low. | Neon lamp need to be replaced | Replace the the neon lamp,perform the complete oximetry test,it got passed successfully,perform 2 point cal & got the desired result
No | Visit | Complaint | RDAN | Marianne og Jan har ringet ang. denne. Man har forsøgsvis skiftet samtlige pumpeslanger samt elektroder og membraner. Det har stået på i nogen tid, specielt pCO2, hvor der skiftes membraner meget ofte. Det er aftalt at Ebbe kigger forbi torsdag den 14. december. | Afprøvet og undersøgt. | Udskiftet pH $blodgasmodulet, slutafprøvet. | SWO-279255 | Defect | HW | pH/BG Module | pH/BG Module |  | Marianne and Jan have called Ang. This. All pump hoses, electrodes and membranes have been changed in experimental manner. It has been going on for some time, especially pCO2, where the membranes are changed very often. It is agreed that Ebbe is looking past Thursday, 14. December. | Tested and investigated. | Replaced pH $blodgasmodulet, final tested.
No | Visit | Complaint | RAME | Error code 1189 over 30 times in 1.5 months and 3 solution packs | Inspected logs. Found many 1189, 1084, 1085, and 1083 errors across multiple solution packs and sensor cassettes. Inlet gasket dirty. | Replaced all pumps $ tubings and hemolyzer. | SWO-280060 | Defect | HW | Basic Unit | PM Kit | 1189 | Error code 1189 over 30 times in 1.5 months and 3 solution packs | Inspected logs. Found many 1189, 1084, 1085, and 1083 errors across multiple solution packs and sensor cassettes. Inlet gasket dirty. | Replaced all pumps $ tubings and hemolyzer.
No | Visit | Complaint | RDEU | Seit 2 Tagen werden Proben abgebrochen Nach Wechsel von Einlass- und Anschlussdichtung wird Leckage angezeigt | Erneute Leckage Drucküberprüfung scheitert Ventil-Reinigung vor 2 Tagen ohne langanhaltende Besserung Gerät seit gestern Abend 17:5 Uhr ausgeschaltet | Quetschventil hängt erneut Austausch Quetschventil Geräteneustart und Fluid Programme i.O. Startphase i.O. Fällige Kal und QC Level A und C durchgeführt | SWO-279627 | Defect | HW | Oxi module | Pinch Valve |  | Samples have been cancelled for 2 days Leakage is displayed after changing the inlet and connection seal | Renewed leakage Pressure check fails Valve cleaning 2 days ago without long-term improvement Device switched off since last night 5:5 p.m. | Squeeze valve hangs again Replacement of squeezing valve Device restart and fluid programs i.O. Start-up phase i.O. Due Kal and QC Levels A and C performed
No | Visit | Complaint | RDEU | 13.12.2017, Sr. Alin, tel: 035145817110: Gerät hat sich ausgeschaltet und läßt sich nicht mehr neu starten | Netzteil defekt | Netzteil erneuert--> Neustart SC verfällt am 15.12.2017 --> SC erneuert Station überwacht die Einlaufphase | SWO-279739 | Defect | HW | Power Supply | Power Supply |  | 13.12.2017, Sr. Alin, tel: 035145817110: Device has turned off and cannot be restarted | Power supply defective | Power supply renewed--> Restart SC expires on 15.12.2017 --> SC renewed Station monitors the run-in phase
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | Na质控不过 | 连续1张更换测试卡和1个试剂包后修复 | SWO-279582 | [5 empty] | Test card, reagent package failure | Na qc but not | 1 consecutive replacement test card and 1 reagent kit after repair
Yes | Visit | Complaint | RESP | Error 05000028 TC Core (OMSG10): Gas valve error | Registrar pieza no puesta en parte anterior | Registrar pieza no puesta en parte anterior | SWO-279775 | Re classified RMED | Double entry, see previous day |  |  |  | Error 05000028 TC Core (OMSG10): Gas valve error | Record unplaced part not previous part | Record unplaced part not previous part
No | Visit | Complaint | RCHN | 定标数据丢失 | 定标数据丢失，借944-258（13338）一张 | 重新定标后修复by tim | SWO-278092 | User error | Consumables | Test Kit | Cartridge data missing |  | Target data loss | Bid data loss, borrow 944-258 (13338) one | Fix by tim after re-calibration
No | Visit | Complaint | RCHN | 干燥前温度超范围 | 干燥前温度超范围，借903-210一个 | 更换903-210后修复by tim | SWO-279713 | Defect | HW | Process Module | PCB, Cup Dryer |  | Overrange temperature before drying | Over-range temperature before drying, borrow 903-210 one | Repair by tim after replacement 903-210
No | Visit | Complaint | RCHN | 报错1024错误 | 报错1024错误，借903-364、944-212（13499）缺货、944-214（13673）、944-216（13335）缺货 各一张 |  | SWO-279628 | Defect | HW | Process Module | Optical Unit | 1024 | 1024 errors | 1024 errors, 903-364, 944-212 (13499) out of stock, 944-214 (13673), 944-216 (13335) out of stock each
No | Visit | Complaint | RIND | Oximetry values is not coming in patient report. | Found inlet gasket and inter connecting tubes are getting old and taking air during aspiration of blood sample. | Change inlet gasket, all inter connecting tubes solution pump tube and waste pump tube. Run LS calibration and pump calibration. Now machine is working satisfactorily. | SWO-280196 | Defect | HW | Basic Unit | Pump tubes | Oxi Values | Oximetry values is not coming in patient report. | Found inlet gasket and inter connecting tubes are getting old and taking air during aspiration of blood sample. | Change inlet gasket, all inter connecting tubes solution pump tube and waste pump tube. Run LS calibration and pump calibration. Now machine is working satisfactorily.
No | Visit | Complaint | RIND | Print error. | Check the print head condition. & found roller is not working. | Remove the printer head & found the gear wheel was broken from the the base plate,replace the printer head with the new one,check the print out & it was ok. | SWO-280089 | Defect | HW | Printer Module | Printer Module |  | Print error. | Check the print head condition. & found roller is not working. | Remove the printer head & found the gear wheel was broken from the the base plate,replace the printer head with the new one,check the print out & it was ok.
Yes | 3 | Y | RPAC | Blank | installation of MET II Module |  | D39707F71C32F950 | Re classified RMED | Installation/Setup |  |  |  | Blank | installation of MET II Module
No | 2 | Y | RPOL | Blank | Regulacja lampy w module OXI, testy, kalibracja tHb od serwisu. Aparat sprawny. |  | D396EFC41C18E68C | Adjustment | HW | Oxi module | Lamp unit |  | Blank | Lamp adjustment in OXI module, testing, calibration of tHb from service. The camera is functional.
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 压力测试真空故障 | 更换试剂包后修复 | SWO-280099 | Assisted Replacement | SP | SP | Vacuum |  | Test card, reagent package failure | Pressure test vacuum failure | Repair after replacement of reagent kit
No | Visit | Complaint | RCHN | pCO2敏感数值偏高 | pCO2敏感数值偏高，借902-972,903-334各一个 | ，更换耗材后修复by tim | SWO-280199 | [5 empty] | pCO2 sensitivity value is high | pCO2 sensitive value is high, borrow 902-972, 903-334 each | Repair by tim after replacing consumables
No | Visit | Complaint | RCHN | 条码扫描器扫不到测试卡 | 条码扫描器扫不到测试卡，借910-304一个 | 更换910-304后修复by tim910 | SWO-280093 | Defect | HW | Storage Module | Barcode reader, Storage wheel |  | Barcode scanner can't sweep test card | Barcode scanner can't sweep the test card, borrow 910-304 one | Repair by tim910 after replacement 910-304
Yes | Visit | Complaint | RCHN | 报错1031 | 报错1031，借903-364,944-212（13518缺货），944-216（13552），944-214（13583）各一个 | 更换903-364,944-212（13518），944-216（13552），944-214（13583）后修复by tim | SWO-279655 | [5 empty] | 1031 | 1031, loan 903-364,944-212 (13518 out of stock), 944-216 (13552), 944-214 (13583) each | Repair by tim after replacement 903-364, 944-212 (13518), 944-216 (13552), 944-214 (13583)
No | Visit | Complaint | RSCH | Drucker funktioniert nicht mehr, erst nach Neustart wieder. | Drucker funktioniert nicht mehr, erst nach Neustart wieder. | Allgemeine Kontrolle ok. Windows 7 und DMS neu installieren. Windows Einstellungen i.o. Kundenkonfig geladen. Drucker und USB Key funktionieren im Moment. Kunde wird sich melden. | SWO-280505 | Defect | SW | Windows | RMED SW | Printer | Printer no longer works until restarted. | Printer no longer works until restarted. | General control ok. Reinstall Windows 7 and DMS. Windows Settings i.o. Customer confectionery loaded. Printer and USB Key are working at the moment. Customer will get in touch.
No | Visit | Complaint | RAME | API failures | PO2 is out of range on the API proficiency testing. Analyzer has failed the last two tests. | Checked QC log and everything is measuring very near the mean except sO2 measures between mean and 1 SD low but still well within the QC range. Ran linearity kit and all results are near the mean. Requesting to be present during the next API survey. | SWO-280795 | No fault found | Performance | Performance | Proficiency Test | pO2 | API failures | PO2 is out of range on the API proficiency testing. Analyzer has failed the last two tests. | Checked QC log and everything is measuring very near the mean except sO2 measures between mean and 1 SD low but still well within the QC range. Ran linearity kit and all results are near the mean. Requesting to be present during the next API survey.
No | Visit | Complaint | RAME | Analyzer Will Not Power On | Analyzer will not power up. The analyzer was placed in the trauma room in which the ambient temperature was 88F and, the power supply may have overheated causing the analyzer to shutdown. | Checked and, replaced a new power supply. Replaced a new sensor cassette due to the cassette not having fluid flow during the down time. However, Na presented 'Sensor response errors' during calibrations. Replaced the sensor cassette again. Submitted credit forms for both faulty sensor cassettes. Showed the customer the ambient temperature specifications in the operator's manual (15C to 32C) or ( 59F to 89.60F). Advised the customer to try to keep the analyzer outside of the trauma room to prevent exposure to high temperatures. | SWO-280978 | Defect | HW | Basic Unit | Power Supply (temp) | Temp | Analyzer Will Not Power On | Analyzer will not power up. The analyzer was placed in the trauma room in which the ambient temperature was 88F and, the power supply may have overheated causing the analyzer to shutdown. | Checked and, replaced a new power supply. Replaced a new sensor cassette due to the cassette not having fluid flow during the down time. However, Na presented 'Sensor response errors' during calibrations. Replaced the sensor cassette again. Submitted credit forms for both faulty sensor cassettes. Showed the customer the ambient temperature specifications in the operator's manual (15C to 32C) or ( 59F to 89.60F). Advised the customer to try to keep the analyzer outside of the trauma room to prevent exposure to high temperatures.
No | Visit | Complaint | RAME | The customer states the analyzer showed temperatures errors and, it the. Shutdown. Won't power up anymore. | Analyzer is no powering up. | Upon arrival, found the analyzer shutdown. The customer explained the analyzer was probably in use at the trauma room on which the ambient temperature was at 88F. I explained and, showed the customer in the operators manual the ambient temperatures that the analyzer needs to be working in (15C to 32C) or (59F to 89.6F). The analyzer may have overheated due to the high ambient temperature in the room. When checking the analyzer, found the power supply with a slight burning smell. I immediately powered down the analyzer and, determined that the power supply is at fault and, needs to be replaced. Ordered a new power supply. | SWO-280791 | Revisit required | HW | Basic Unit | Power Supply | Temp | The customer states the analyzer showed temperatures errors and, it the. Shutdown. Won't power up anymore. | Analyzer is no powering up. | Upon arrival, found the analyzer shutdown. The customer explained the analyzer was probably in use at the trauma room on which the ambient temperature was at 88F. I explained and, showed the customer in the operators manual the ambient temperatures that the analyzer needs to be working in (15C to 32C) or (59F to 89.6F). The analyzer may have overheated due to the high ambient temperature in the room. When checking the analyzer, found the power supply with a slight burning smell. I immediately powered down the analyzer and, determined that the power supply is at fault and, needs to be replaced. Ordered a new power supply.
No | Visit | Complaint | RBNL | Metingen worden afgebroken, Inlet blijft open | Regelmatig afgebroken bloedmetingen. Na de metingen blijft Inlet ook vaak open staan. Bij aankomst zijn alle parameters groen en geen foutmelding. Dit toestel heeft al eerder problemen gehad met de afgebroken metingen en openstaan van de Inlet na de meting. | Alle Inlet gerelateerde onderdelen heb vervangen vervolgens diverse metingen heb uitgevoerd zonder probleem. | SWO-280842 | Defect | HW | Auto Inlet module | Inlet Position Cal. |  | Measurements are aborted, Inlet remains open | Regularly broken down blood measurements. After the measurements, Inlet is also often open. On arrival all parameters are green and no error message. This unit has previously had problems with the aborted measurements and open the Inlet after the measurement. | All Inlet related parts have replaced then several measurements have performed without problem.
Yes | 3 | Y | RPAC | Blank | Lactate calibration error. Lactate calibration sensitivity out of range. Measured QC value higher than statistical range. |  | D39AD6A77E5FC8FD | [5 empty] | Blank | Lactate calibration error. Lactate calibration sensitivity out of range. Measured QC value higher than statistical range.
Yes | Visit | Complaint | RSCH | Gemäss Absprache Hr. Dr. Bucher und U. Stolz wird Gerät ausgetauscht. | Gemäss Absprache Hr. Dr. Bucher und U. Stolz wird Gerät ausgetauscht. | ABL 80 Co-ox SN 302068 mit ABL 80 Co-ox SN 305905. Zählerstand Austauschgerät in Betrieb genommen. Man. QA Zyklus i.o. | SWO-280713 | Re classified RMED | Replaced Unit |  |  |  | According to the agreement, Mr. Dr. Bucher and U. Stolz will exchange equipment. | According to the agreement, Mr. Dr. Bucher and U. Stolz will exchange equipment. | ABL 80 Co-ox SN 302068 with ABL 80 Co-ox SN 305905. Meter reading replaced device. somebody. QA cycle i.o.
No | Visit | Complaint | RCHN | lac和glu定标不过 | lac和glu定标不过，需要借件903-581一个 |  | SWO-280278 | Defect | HW | Basic Unit | PCB, Sensor interface |  | lac and glu calibration however | lac and glu calibration however, need to borrow 903-581 one
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | GLU,Lac定标不过 | 更换测试卡后修复 | SWO-283020 | [5 empty] | Test card, reagent package failure | GLU, Lac calibration is not | Repair after replacing test card
No | Visit | Complaint | RCHN | 做樣本時機器報錯: 1031 Optical intensity out of range | 客戶操作時出現1031 Optical intensity out of range | 1.更換903-364 optic unit 2.執行factory calibration 3.Assay calibration 4.Run QC ok | SWO-283805 | Defect | HW | Process Module | Optical Unit | 1031 | Machine error when taking samples: 1031 Optical intensity out of range | 1031 Optical Intensity out of range during customer operation | 1.更換903-364 optic unit 2.執行factory calibration 3.Assay calibration 4.Run QC ok
No | Visit | Complaint | RCHN | 报错5960错误 | 报错5960错误，借944-258（13224）一个 |  | SWO-280488 | Assisted Replacement | Consumables | Test Kit | NT-ProBNP | 5960 | 5960 error | 5960 error, borrow 944-258 (13224) one
No | Visit | Complaint | RCHN | 报错5784错误 | 报错5784错误，借925-016一个 |  | SWO-280475 | Defect | HW | Cover and Chassis | Cartridge Lid | 5784 | 5784 errors | 5784 error, borrow 925-016 one
No | 16 | Y | RPNZ | Blank | Reference electrode replaced due to electrolyte drift |  | D39BEF5CF4439738 | Assisted Replacement | Electrode | Electrode | Reference | Drift | Blank | Reference electrode replaced due to electrolyte drift
No | 16 | Y | RPNZ | Blank | Calibration failure (Na, Lac, Glu failures) |  | D39B1F147E3695A3 | Defect | HW | El/Met Module | Measuring chamber | Glu/Lac/Na | Blank | Calibration failure (Na, Lac, Glu failures)
Yes | Visit | Complaint | RIND | Calibration Error In Lactate, Glucose & Sodium | Check the logs. check the calibration. | Replace the membrane & calibrate the system. calibration are in ranges. | SWO-283641 | [5 empty] | Calibration Error In Lactate, Glucose & Sodium | Check the logs. check the calibration. | Replace the membrane & calibrate the system. calibration are in ranges.
Yes | Visit | Complaint | RKK | 測定時、Lac応答不良頻発。 | 測定時、Lac応答不良頻発。 | Lac電極交換実施。測定、Cal.QCOK。 | SWO-280675 | [5 empty] | Lac response failure frequently at the time of measurement. | Lac response failure frequently at the time of measurement. | Lac electrode replacement was implemented. Measurement, Cal.QCOK.
Yes | Visit | Complaint | RAME | Customer run LCR and instrument failed high Crea checks. | Failed high Crea Checks on LCR. Customer cleaned CreaA and Crea B electrode contacts and connectors. Replaced a different lot# of rinse solution and still the same issue. Crea A electrode Lot#YI-03. Crea B electrode lot# YI-02 | Replaced electrode connectors on Annual PM. Remembraned Crea electrodes. Lot#R0765. | SWO-280924 | [5 empty] | Customer run LCR and instrument failed high Crea checks. | Failed high Crea Checks on LCR. Customer cleaned CreaA and Crea B electrode contacts and connectors. Replaced a different lot# of rinse solution and still the same issue. Crea A electrode Lot#YI-03. Crea B electrode lot# YI-02 | Replaced electrode connectors on Annual PM. Remembraned Crea electrodes. Lot#R0765.
No | Visit | Complaint | RAME | Missing pH,Base,Bicarb on results | Found pH parameter logging unstable errors during patient testing intermittently. Checked measuring chamber electrode wells. Found some blood residue in pH well. Also noticed inlet tube on probe not fully seated on inlet nipple. | Cleaned measuring chamber wells. Checked electrode and replaced seal ring on original electrode. Customer had replaced electrode with a different one for troubleshooting. Replaced inlet tube. Restart calibrations. | SWO-280955 | Defect | Electrode | Electrode | Sealing ring | Sealing ring, pH | Missing pH,Base,Bicarb on results | Found pH parameter logging unstable errors during patient testing intermittently. Checked measuring chamber electrode wells. Found some blood residue in pH well. Also noticed inlet tube on probe not fully seated on inlet nipple. | Cleaned measuring chamber wells. Checked electrode and replaced seal ring on original electrode. Customer had replaced electrode with a different one for troubleshooting. Replaced inlet tube. Restart calibrations.
No | 6 | Y | RCZE | Blank | po pøedchozí dohodì s Beòovskou provedena kontrola a výmìna hadièky ( asymetrie pumpy ), |  | D39BCC86F4022155 | Defect | HW | pH/BG Module | Pump tubes |  | Blank | After the previous agreement with Beòov, the control and replacement of the hose (asymmetry of the pump),
No | Visit | Complaint | RAME | Customer reports that analyzer leaked overnight and appears to be under control but would like anachecked out. | Analyzer leak has been resolved but dry. Loodnis observed in and around the rear of the inlet gasket and on the inlet probe. | Flushed the waste system to ensure that the clot is no longer present in the system. Analyzer is within acceptable operating specifications. | SWO-280879 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | Customer reports that analyzer leaked overnight and appears to be under control but would like anachecked out. | Analyzer leak has been resolved but dry. Loodnis observed in and around the rear of the inlet gasket and on the inlet probe. | Flushed the waste system to ensure that the clot is no longer present in the system. Analyzer is within acceptable operating specifications.
No | Visit | Complaint | RDAN | Per fra medicoteknik klager over larm fra ABL'en. | Afprøvet. Kan umiddelbart ikke høre noget usædvaneligt. Når man styrer de to store blæsere kan det lige anes. | A $C blæser udskiftet, slutafprøvet. | SWO-280507 | Defect | HW | DMS_PC | FAN |  | Per from Medtech complains about the noise of the abbey. | Tested. Can not hear anything year. When you control the two big fans, it can just be. | A $C Blower replaced, final tested.
Yes | Visit | Complaint | RDEU | Hotline - PH Kal Fehler | pH Kal Fehler. | Der Fehler wurde im laufe des Vormittags schon behoben. Fehlerhafte QC durchgeführt. Analysator Status grün. | SWO-281072 | Re classified RMED |  |  |  |  | Hotline - PH Kal Error | pH Kal error. | The bug has already been fixed in the morning. Incorrect QC performed. Analyzer status green.
No | Visit | Complaint | RAME | pH Cal error | pH Cal error. Confirmed when onsite. | Sealing o-ring on end of pH electrode missing. Replaced for customer. | SWO-280883 | Defect | Electrode | Electrode | Sealing ring | Sealing ring, pH | pH Cal error | pH Cal error. Confirmed when onsite. | Sealing o-ring on end of pH electrode missing. Replaced for customer.
Yes | Visit | Complaint | RDEU | KAL Fehler pCO2 und tHb | KAL Fehler pCO2 (Status 0 und Empf. 160%) und tHb (kein Wert) | Gas und Ventilüberprüfungen - Elektrode und Kontakte gereinigt - Lampeneinstellung Hemolysator - mehrfache Kalibrationen durchgeführt und Funktionsüberprüfung etc. | SWO-280715 | Re classified RMED | OOB |  |  |  | KAL Fehler pCO2 und tHb | KAL Error pCO2 (Status 0 and Mpf. 160%) and tHb (no value) | Gas and valve checks - electrode and contacts cleaned - lamp setting Hemolysator - multiple calibrations performed and function verification etc.
No | Visit | Complaint | RLTD | Lucy Tel 0208 9095126 | 31 blank cups available. No TNT test cups on board. | Advised customer the blank cups are for the cleaning cycle only. Showed customer on how to install the TNT test cups and they have installed 48 test. Explain the difference on the consumable screen between Blank cups and TNT test. | SWO-280681 | User error | Consumables | Test Kit | Cartridge, Blank |  | Lucy Tel 0208 9095126 | 31 blank cups available. No TNT test cups on board. | Advised customer the blank cups are for the cleaning cycle only. Showed customer on how to install the TNT test cups and they have installed 48 test. Explain the difference on the consumable screen between Blank cups and TNT test.
No | Visit | Complaint | RAME | Pump Calibration failure in auxiliary service mode happens. | Analyzer keep asking to replace the solution pack Pump Calibration failed in auxiliary service mode. | New inlet gasket and connector installed Performed all service tests in auxiliary page. Flushed the analyzer with DM water and air. The sensor cassette replaced as was left with no action on analyzer for 24 hours. Analyzer passed all the tests and it is ready for use. | SWO-281607 | Defect | HW | Inlet module | Inlet gasket | & connector gasket | Pump Calibration failure in auxiliary service mode happens. | Analyzer keep asking to replace the solution pack Pump Calibration failed in auxiliary service mode. | New inlet gasket and connector installed Performed all service tests in auxiliary page. Flushed the analyzer with DM water and air. The sensor cassette replaced as was left with no action on analyzer for 24 hours. Analyzer passed all the tests and it is ready for use.
No | Visit | Complaint | RAME | The analyzer is not communicating with Sunquest. | The analyzer is connected via a serial adapter to the wall jack. | Tested the serial port with my laptop. It is communicating successfully. | SWO-281809 | Setup | His/Lis | His/Lis | His/Lis | Serial port | The analyzer is not communicating with Sunquest. | The analyzer is connected via a serial adapter to the wall jack. | Tested the serial port with my laptop. It is communicating successfully.
No | Visit | Complaint | RLTD | Chris Davidson contact the office. The ABL90 FLEX PLUS on ICU has an error code 1343 and is askinig for an engineer to be contacted. | Reviewed activity log to assess issues. Current issue of unable to pump solutions found to be due to Inhomogeneous QC2 & QC3 solution errors. These were found to have started after replacement of the solution pack on 18 $12 $17 at 08:32. Therefore removed Solution Pack and installed another new one. Tested analyser operation - | Replaced faulty solution pack. | SWO-281220 | Assisted Replacement | SP | SP | FT | 1343 | Chris Davidson contact the office. The ABL90 FLEX PLUS on ICU has an error code 1343 and is askinig for an engineer to be contacted. | Reviewed activity log to assess issues. Current issue of unable to pump solutions found to be due to Inhomogeneous QC2 & QC3 solution errors. These were found to have started after replacement of the solution pack on 18 /12 /17 at 08:32. Therefore removed Solution Pack and installed another new one. Tested analyser operation - | Replaced faulty solution pack.
No | Visit | Complaint | RDEU | 18.12.2017, Frau Neiswirth, Tel: 0089857915605: Kalibrationsfehler bei Lac, Glu, Na und K | Fehlersuche vor Ort | Am Vortag mit Frau Neiswirth telefoniert und nochmals die Sensorkassette und den LP tauschen lassen. Bei Eintreffen war nur noch die Empf. von Lac nicht in Ordnung bei der Kal. Deswegen nochmals die Sensorkassette und Lösungspack getauscht. Nach Startphase war der Fehler noch vorhanden (Reaktionsfehler bei Lac) Sensorinterface PCB bestellt. Sobald diese da ist wird sie eingebaut. Frau Neiswirth darüber informiert. QCs für alle Level durchgeführt. ABL wieder messbereit. | SWO-281262 | Revisit required | HW | Basic Unit | PCB, Sensor interface |  | 18.12.2017, Mrs Neiswirth, Tel: 0089857915605: Calibration error in Lac, Glu, Na and K | Troubleshooting on site | The day before, Ms. Neiswirth phoned and had the sensor cassette and the LP exchanged again. When we arrived, only the empf was left. of Lac is not okay with the Cal. Therefore, the sensor cassette and solution pack are replaced again. After the start phase the error was still present (reaction error at Lac) sensor interface PCB. As soon as it is there, it is installed. Mrs Neiswirth has been informed. QCs for all levels. ABL ready to be measured again.
Yes | 5 | Y | RPAC | Blank | Unable to access analyser |  | D39C0A77F4037AE6 | Re classified RMED | Remote support |  |  |  | Blank | Unable to access analyser
No | Visit | Complaint | RDEU | 14.12.2017 9:15 Fr. Püttman 020644417351: Einlaß-Fehler 19.12.2017; 08:15 Uhr; Fr. Püttmann; Tel.: 0206444417351; wieder Probleme mit dem Einlass | Zeitweise Fehler 1199 Einlass geöffnet bei Probenmessung mit Kapillaren Einlassdichtung rutscht nicht in Ausgangsposition | Einlass im Servicemenü überprüft, kein Fehler festgestellt, Einlassdichtung mit Halter aus Kundenbestand ersetzt, Führung des Einlassmoduls gefettet, Anwender beobachtet und meldet sich bei weiteren Problemen, Kal und QC durchgeführt | SWO-281876 | Defect | HW | Auto Inlet module | Auto Inlet module | Capillary | 14.12.2017 9:15 Ms. Püttman 020644417351: Admission error 19.12.2017; 8:15 a.m.; Ms. Püttmann; Tel.: 0206444417351; again problems with the inlet | Intermittent error 1199 Inlet opened when sample measurement with capillaries Inlet seal does not slip into starting position | Inlet checked in the service menu, no error detected, inlet seal replaced with holder from customer stock, guide of the inlet module greased, user observed and reports in case of further problems, Kal and QC performed
No | Visit | Complaint | RLTD | Dan Rickets 07787 598407 | The main QC failure was O2 high on level 3, it has been like this since the solution pack was installed so the solution pack will need replacing. Also found that the inlet probe was bent. | No obvious reason was found for the fluid transport problems when replacing the solution pack so it was decided to fit a service kit and also to soak the sample path with protein remover. The reason for this was that the problems seemed to start after a recent blockage in the analyser so I wanted to ensure that the fluid path was clear. The fluid path was thoroughly flushed with water to remove all of the protein remover. The analyser was allowed to start up but the O2 was still failing level 3 QC so a new solution pack was fitted but the O2 still failed level 3 QC so a new sensor cassette was fitted and left to start-up. I contacted the customer the next day and the new cassette had started up ok. | SWO-281622 | Assisted Replacement | SC | SC | pO2 | QC 3(C?) high | Dan Rickets 07787 598407 | The main QC failure was O2 high on level 3, it has been like this since the solution pack was installed so the solution pack will need replacing. Also found that the inlet probe was bent. | No obvious reason was found for the fluid transport problems when replacing the solution pack so it was decided to fit a service kit and also to soak the sample path with protein remover. The reason for this was that the problems seemed to start after a recent blockage in the analyser so I wanted to ensure that the fluid path was clear. The fluid path was thoroughly flushed with water to remove all of the protein remover. The analyser was allowed to start up but the O2 was still failing level 3 QC so a new solution pack was fitted but the O2 still failed level 3 QC so a new sensor cassette was fitted and left to start-up. I contacted the customer the next day and the new cassette had started up ok.
No | Visit | Complaint | RCHN | 报错5689 | 报错5689，借944-394（13533）一个 |  | SWO-281039 | Defect | Consumables | Test Kit | Cartridge | 5689 | 5689 error | 5689, loan 944-394 (13533) one
No | Visit | Complaint | RLTD | Email received from Tara Goddard | Tested monitor together with 2 sensor 92 electrodes. With Sensor 92 sn. 92-78304 The monitor starts calibrating but after a short period gives a Sensor error 05100008, thermistor error, followed by monitor error 05000010, heating error. With Sensor 92 sn. 92-78568 connected the monitor will not start gives error FAILED selfTestTCCORE and error code 02040000. Both sensors tested on loan TCM5 monitor serial no. 500R0008N002 and on returned monitor serial no. 500R0063N004 and errors remain with sensors. Monitor tested without sensor connected and is working correctly. Both faulty sensors removed from site and warranty replacement sensors ordered from Radiometer Limited office for urgent delivery. | Replace 2 x faulty Sensor 92. | SWO-281593 | [5 empty] | Email received from Tara Goddard | Tested monitor together with 2 sensor 92 electrodes. With Sensor 92 sn. 92-78304 The monitor starts calibrating but after a short period gives a Sensor error 05100008, thermistor error, followed by monitor error 05000010, heating error. With Sensor 92 sn. 92-78568 connected the monitor will not start gives error FAILED selfTestTCCORE and error code 02040000. Both sensors tested on loan TCM5 monitor serial no. 500R0008N002 and on returned monitor serial no. 500R0063N004 and errors remain with sensors. Monitor tested without sensor connected and is working correctly. Both faulty sensors removed from site and warranty replacement sensors ordered from Radiometer Limited office for urgent delivery. | Replace 2 x faulty Sensor 92.
No | Visit | Complaint | RAME | Customer called to report getting EC 0258 and 0316 after the unit was recently shut down and restarted. | Complete oxi test revealed neon failure | Replaced lamp unit and recalibrate system. | SWO-281676 | Defect | HW | Oxi module | Lamp unit | Neon | Customer called to report getting EC 0258 and 0316 after the unit was recently shut down and restarted. | Complete Oxi test revealed neon failure | Replaced lamp unit and recalibrate system.
Yes | 7 | Y | RPAC | Blank | Replace the pH tip Seal BLACK |  | D39E66CAE018291B | Re classified RMED | FAN |  |  |  | Blank | Replace the pH tip Seal BLACK
Yes | 7 | Y | RPAC | Blank | Bi Annual Service |  | D39D41596A061CF1 | Re classified RMED | Service |  |  |  | Blank | Bi Annual Service
No | Visit | Complaint | RDEU | QC wird nicht gemessen | QC wird nicht gemessen Keine Probe bei Ansaugung erkannt | Überprüfung Analysator (Nadel verbogen) - Erneuerung Einlassnadel und Reinigung Einlassbereich - Erneuerung Einlassdichtung (aus Kundenbestand) - Funktionsüberprüfung etc. ohne Beanstandung ! Bemerkung: Aqure-Verbindung aktiviert | SWO-281633 | Defect | HW | Inlet module | Probe bend |  | QC is not measured | QC is not measured No sample detected during suction | Inspection analyzer (needle bent) - renewal inlet needle and cleaning inlet area - renewal inlet seal (from customer stock) - function check etc. without complaint ! Note: Aqure connection enabled
Yes | Visit | Complaint | RDEU | Gerät saugt kapiläre BGAs nicht an | Gerät überprüft | Einlasssonde erneuert, kapiläre BGAs durchgeführt | SWO-281820 | [5 empty] | Device does not suck in kapilary BGAs | Device checked | Entrance probe renewed, capilary BGAs carried out
Yes | Visit | Complaint | RDEU | QC wird nicht gemessen | QC wird nicht gemessen (Abbruch- keine Meldung) Ampulle wird geknackt und Einheit fährt ohne Messung herunter ! | Auto Check demontiert - Erneuerung FPC Kabel - Rückmontage und Funktionsüberprüfung (keine Fehlerbeseitigung) - Heute bereits 3 Qualitätskontrollen bestanden ! Neuteil-Bestellung erfolgt | SWO-281635 | [5 empty] | QC is not measured | QC is not measured (cancellation- no message) Ampoule is cracked and unit shuts down without measurement ! | Auto Check dismantled - Renewal FPC cable - Reassembly and functional verification (no troubleshooting) - Today already passed 3 quality controls ! New part order is made
No | Visit | Complaint | RAME | The customer states the analyzer showed temperatures errors and, it the. Shutdown. Won't power up anymore. | Analyzer is not powering up. Revisit to continue working on this analyzer. | Revisit to setup the wireless connection and, make sure the instrument connects with AQURE. Connection is successful., the customer was able to connect to the analyzer thru AQURE. In windows, I setup the time to show in military time. Last, I reloaded the current qc the customer is using. Note: The battery pack needs to be replaced, ordering a new battery. The customer is now using the analyzer as a bench top analyzer and, it will not be used on the rolling stand until the new battery is replaced. | SWO-282372 | Setup | SW | His/Lis | His/Lis | CPU replaced | The customer states the analyzer showed temperatures errors and, it the. Shutdown. Won't power up anymore. | Analyzer is not powering up. Revisit to continue working on this analyzer. | Revisit to setup the wireless connection and, make sure the instrument connects with AQURE. Connection is successful., the customer was able to connect to the analyzer thru AQURE. In windows, I setup the time to show in military time. Last, I reloaded the current qc the customer is using. Note: The battery pack needs to be replaced, ordering a new battery. The customer is now using the analyzer as a bench top analyzer and, it will not be used on the rolling stand until the new battery is replaced.
No | Visit | Complaint | RAME | The customer states the analyzer showed temperatures errors and, it the. Shutdown. Won't power up anymore. | Analyzer will not power 'ON'. Revisit to replace the power supply. | This analyzer was exposed to high ambient temperature which caused the power supply to be damaged and, the analyzer software deleted from the analyzer. Replaced the power supply. When reinstalling the analyzer software the analyzer just kept rebooting itself. Had to replace a new CPU unit and, the flash card as it was found to be corrupted. Installed the analyzer software. Due the the software re-installation the sensor cassette and, the solution pack needed to be replaced. The new solution pack part #944-157 lot YW-53 was causing error 1087 in homogeneous QC3 solution. The analyzer kept asking to change a new solution pack and, replace the inlet gasket. Replaced a new pack again. Back-flush the analyzer. Replaced the inlet gasket, the inlet connector gasket and, the inlet probe. Performed the service fluid transport checks, now passing ok. Ran the ctHb calibration. Reinstall the Asus wireless adapter but, I do not know what is the wireless network the analyzer needs to be connected to. I will need to speak with the IT person so they can provide the wireless network information. Also found the battery pack not working properly and, causing the analyzer to reboot. Ordered a new battery option. | SWO-281800 | Defect | HW | Power Supply | Power Supply | & CPU + SW | The customer states the analyzer showed temperatures errors and, it the. Shutdown. Won't power up anymore. | Analyzer will not power 'ON'. Revisit to replace the power supply. | This analyzer was exposed to high ambient temperature which caused the power supply to be damaged and, the analyzer software deleted from the analyzer. Replaced the power supply. When reinstalling the analyzer software the analyzer just kept rebooting itself. Had to replace a new CPU unit and, the flash card as it was found to be corrupted. Installed the analyzer software. Due the the software re-installation the sensor cassette and, the solution pack needed to be replaced. The new solution pack part #944-157 lot YW-53 was causing error 1087 in homogeneous QC3 solution. The analyzer kept asking to change a new solution pack and, replace the inlet gasket. Replaced a new pack again. Back-flush the analyzer. Replaced the inlet gasket, the inlet connector gasket and, the inlet probe. Performed the service fluid transport checks, now passing ok. Ran the ctHb calibration. Reinstall the Asus wireless adapter but, I do not know what is the wireless network the analyzer needs to be connected to. I will need to speak with the IT person so they can provide the wireless network information. Also found the battery pack not working properly and, causing the analyzer to reboot. Ordered a new battery option.
No | Visit | Complaint | RLTD | Not printing when paper is changed they keep having to turn off and reboot please call Tony Malone 02072885428 | Not printing when paper changed and unit has to be rebooted. Then works ok. | As this fault is intermittent the printer has been replaced. Test replacing paper several times and each time the printer has worked ok. | SWO-282233 | Defect | HW | Printer Module | Printer Module |  | Not printing when paper is changed they keep having to turn off and reboot please call Tony Malone 02072885428 | Not printing when paper changed and unit has to be rebooted. Then works ok. | As this fault is intermittent the printer has been replaced. Test replacing paper several times and each time the printer has worked ok.
No | Visit | Complaint | RLTD | Please call Tony Malone on 0207 288 5428 | Found analyser to be working ok | Customer had replaced inlet gasket and this has resolved the issue. | SWO-282234 | No fault found | No fault found | Auto Inlet module | Probe | Customer resolved issue | Please call Tony Malone on 0207 288 5428 | Found analyser to be working ok | Customer had replaced inlet gasket and this has resolved the issue.
No | Visit | Complaint | RDEU | 18.12.2017, Frau Neiswirth, Tel: 0089857915605: Kalibrationsfehler bei Lac, Glu, Na und K | Neue Sensor Interface PCB eingebaut | Neue Sensor Interface PCB eingebaut Sensorkassette und Lösungspack aus Kundenbestand erneuert. Mehrere Kalibrierungen durchgeführt --> Lac wieder i.O. Bei KAL QCs für alle Level durchgeführt--> i.O. ABL wieder messbereit. | SWO-282275 | Defect | HW | Basic Unit | PCB, Sensor interface |  | 18.12.2017, Mrs Neiswirth, Tel: 0089857915605: Calibration error in Lac, Glu, Na and K | New Sensor Interface PCB built-in | New Sensor Interface PCB built-in Sensor cassette and solution pack from customer base renewed. Several calibrations performed --> Lac again i.O. At KAL QCs performed for all levels--> i.O. ABL ready to be measured again.
No | 3 | Y | RPOL | Blank | Uszkodzony pakkiet, wymieniono. |  | D3A4F01E2E2B2424 | Assisted Replacement | SP | SP | pO2 |  | Blank | Faulty packing, listed.
No | 5 | Y | RPAC | Blank | repeat sensor issues |  | D39D731F6A2D763E | Defect | HW | Basic Unit | PCB, Sensor interface |  | Blank | repeat sensor issues
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | ph,pco2,po2质控失败 | 更换测试卡后修复 | SWO-282449 | [5 empty] | Test card, reagent package failure | ph, pco2, po2 QC failure | Repair after replacing test card
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 试剂包漏液 | 更换试剂包后修复 | SWO-284229 | [5 empty] | Test card, reagent package failure | Reagent bag leak | Repair after replacement of reagent kit
No | Visit | Complaint | RCHN | hct响应超出范围 | hct响应超出范围，需要借件903-329一个 | 清洁HCT模块后修复 Cleaned the HCT, then OK.(By Ran) | SWO-281782 | Cleaned | HW | Aspiration Module | Needle Wash Unit |  | hct response out of range | hct response out of range, need to borrow 903-329 one | 清洁HCT模块后修复 Cleaned the HCT, then OK. (By Ran)
No | Visit | Complaint | RCHN | 干燥前温度超范围 | 干燥前温度超范围，借903-118一个 | 更换903-118后修复(By Ran) | SWO-281784 | Defect | HW | Process Module | PCB, Cup Dryer |  | Overrange temperature before drying | Over-range temperature before drying, borrow 903-118 | Repair after replacement 903-118 (By Ran)
No | Visit | Complaint | RDEU | 20.12.2017, Frau Laskowski, tel: 021618922929: kein vorderes Luftsegment im ELS | Proben werden abgebrochen mit dem Fehler 0328 Kein vorderes Luftsegment im Einlass-Flüssigkeitsfühler im Zeitrahmen erkannt . | Flüssigkeiten treten immer mal wieder im Kapillar Einlass und am Schlauch zwischen den Messkammern aus. Einlassnippel und den Luftweg gereinigt. LS justage durchgeführt. Pumpen Kalibrierung durchgeführt. 4 Testmessungen durchgeführt - keine Fehler. Es läuft weiterhin wenige Mengen Flüssigkeiten oberhalb der Schlauches zwischen den Messkammern aus. Messkammer undicht - Flüssigkeiten sind bereits auf die Platine gelaufen und haben Korrosionen verursacht - Ableitströme erkannt. Neues El $Met Modul aus der Zentrale organisiert und eingebaut. Startphase, Kalibrierung und QC's durchgeführt. | SWO-282466 | Defect | HW | El/Met Module | El/Met Module | Leak | 20.12.2017, Mrs. Laskowski, tel: 021618922929: no front air segment in the ELS | Samples are aborted with error 0328 No front air segment detected in the inlet liquid sensor in the time frame. | Liquids repeatedly leak in the capillary inlet and on the hose between the measuring chambers. inlet nipples and the airway. LS adjustment. Pump calibration performed. 4 test measurements performed - no errors. There are still a few amounts of liquids running out above the hose between the measuring chambers. Measuring chamber leaking - liquids have already run onto the board and caused corrosion - leaking currents detected. New El $Met module from the headquarters organized and installed. start-up phase, calibration and QC's.
No | Visit | Complaint | RIND | Leaking in Gaskets | Gasket leaking | Removed the gasket and kept hypo solution for one hour and cleaned all blood fibrins and cleaned the needle after that done pump calibration calibrated the analyser to raedy mode observation was done for while | SWO-283339 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | Leaking in Gaskets | Gasket leaking | Removed the gasket and kept hypo solution for one hour and cleaned all blood fibrins and cleaned the needle after that done pump calibration calibrated the analyser to raedy mode observation was done for while
No | Visit | Complaint | RDAN | Marianne og Jan har ringet ang. denne. Man har forsøgsvis skiftet samtlige pumpeslanger samt elektroder og membraner. Det har stået på i nogen tid, specielt pCO2, hvor der skiftes membraner meget ofte. Det er aftalt at Ebbe kigger forbi torsdag den 14. december. | Afprøvet. Jan siger, at fejl 478 er forsvundet efter skift af pH $blodgas modul men PCO2 vipper i status +følsomhed hver 8-10 kal. | Udskiftet ventil $flow enheder i gasmodulet, slutafprøvet. | SWO-282101 | Defect | HW | pH/BG Module | Valve, air |  | Marianne and Jan have called Ang. This. All pump hoses, electrodes and membranes have been changed in experimental manner. It has been going on for some time, especially pCO2, where the membranes are changed very often. It is agreed that Ebbe is looking past Thursday, 14. December. | Tested. Jan says that error 478 has disappeared after changing the pH /blodgas module but PCO2 lashes in status + sensitivity every 8-10 kal. | Replaced valve $flow units in the gas module, final tested.
No | Visit | Complaint | RAME | PH drifting and instable | PH calibration erreur | Remplacé électrode PH | SWO-282342 | Assisted Replacement | Electrode | Electrode | pH |  | PH drifting and instable | PH calibration error | Replaced electrode PH
Yes | Visit | Complaint | RDEU | QC wird nicht gemessen | Auto Check defekt QC wird nicht gemessen | Analysator heruntergefahren - Test Auto Check Modul montiert - komplette Software Installation notwendig - Meldung Inkonsistente Software bleibt - Platinen aus OXi Modulen quer getauscht - Funktionsüberprüfung etc. ohne Beanstandung Bemerkung: Auto Check Modul bereits bestellt ! | SWO-282167 | [5 empty] | QC is not measured | Auto Check defective QC is not measured | Analyzer shut down - Test Auto Check module mounted - complete software installation necessary - message Inconsistent software remains - boards of OXi modules cross-swapped - function check etc. without complaint Note: Auto Check module already ordered !
Yes | Visit | Complaint | RDEU | 14.11.2017, Herr Müller(MT), Tel: 05861838400: nach jedem Wechsel der P/A Kassette friert der Bildschirm ein. 28.11.2017; Hr. Müller; Tel.: 05861838400; Hr. Müller wartet noch auf einen Technikereinsatz. | System überprüft => ReagenzPack PCB defekt und zerstörte Caps in LP Fach | ReagenzPack PCB erneuert und zerstörte Caps entfernt, QC durchgeführt | SWO-282054 | [5 empty] | 14.11.2017, Mr. Müller(MT), Tel: 05861838400: after each change of the P/A cassette, the screen freezes. 28.11.2017; Mr. Müller; Tel.: 05861838400; Mr. Müller is still waiting for a technician. | System checked => ReagentPack PCB defective and destroyed caps in LP compartment | ReagentPack PCB renewed and destroyed caps removed, QC performed
No | Visit | Complaint | RFRA | Lors du dosage sang totale erreur aléatoire 1068 | cf maintenance préventive | calibration cup dryer temp | SWO-282195 | Adjustment | HW | Process Module | PCB, Cup Dryer | 1068 | During total blood dosing random error 1068 | cf preventative maintenance | calibration cup dryer temp
No | Visit | Complaint | RDEU | 30.11.2017 14:55Uhr Herr Kamh 01724295700: Das Gerät soll über die serielle Schnittstelle angebunden werden. | Anbindung an Praxis EDV | Konfiguration eingestellt und geprüft | SWO-282053 | Setup | SW | Analyzer Control Module | His/Lis |  | 30.11.2017 14:55Mr. Kamh 01724295700: The device is to be connected via the serial interface. | Connection to practice COMPUTER | Configuration set and checked
No | Visit | Complaint | RSAF | New lot cal adj failed | No fault found, all ok and process wheel clean | Reran TnT Cal adj ok. | SWO-282566 | Assisted Replacement | Consumables | Test Kit | Cartridge, TNT Cal |  | New lot cal adj failed | No fault found, all ok and process wheel clean | Reran TnT Cal adj ok.
No | Visit | Complaint | RLTD | Possible bacterial contamination spared from other ABL90s. | Carried out thorough decontamination of sample pathway within the analyser. Complete sample pathway flushed, including scrubbing the line, with hypochlorite solution for 30 minutes. Wet section then removed and inlet probe, gaskets and tubes soaked in hypochlorite, all metal connection pipes also cleaned with hypochlorite solution. All areas cleaned with hypochlorite flushed through with water. Analyser re-assembled and new Sensor Cassette and Solution pack fitted. Verified start-up of analyser - all ok. | Decontaminated sample pathway. | SWO-282797 | Cleaned | Fluidic pathway | Fluidic pathway | Decontamination |  | Possible bacterial contamination spared from other ABL90s. | Carried out thorough decontamination of sample pathway within the analyser. Complete sample pathway flushed, including scrubbing the line, with hypochlorite solution for 30 minutes. Wet section then removed and inlet probe, gaskets and tubes soaked in hypochlorite, all metal connection pipes also cleaned with hypochlorite solution. All areas cleaned with hypochlorite flushed through with water. Analyser re-assembled and new Sensor Cassette and Solution pack fitted. Verified start-up of analyser - all ok. | Decontaminated sample pathway.
No | Visit | Complaint | RLTD | Decontamination required due to possible bacterial contamination from other ABL90s. | Carried out thorough decontamination of sample pathway within the analyser. Complete sample pathway flushed, including scrubbing the line, with hypochlorite solution for 30 minutes. Wet section then removed and inlet probe, gaskets and tubes soaked in hypochlorite, all metal connection pipes also cleaned with hypochlorite solution. All areas cleaned with hypochlorite flushed through with water. Analyser re-assembled and new Sensor Cassette and Solution pack fitted. Verified start-up of analyser - all ok. | Decontaminated sample pathway. | SWO-282798 | Cleaned | Fluidic pathway | Fluidic pathway | Decontamination |  | Decontamination required due to possible bacterial contamination from other ABL90s. | Carried out thorough decontamination of sample pathway within the analyser. Complete sample pathway flushed, including scrubbing the line, with hypochlorite solution for 30 minutes. Wet section then removed and inlet probe, gaskets and tubes soaked in hypochlorite, all metal connection pipes also cleaned with hypochlorite solution. All areas cleaned with hypochlorite flushed through with water. Analyser re-assembled and new Sensor Cassette and Solution pack fitted. Verified start-up of analyser - all ok. | Decontaminated sample pathway.
No | Visit | Complaint | RBNL | uitslagen van een vorig monster onder gescand monsternummer | Het volgende klacht is opgetreden; Wat er gebeurd is: 2 collega's willen een capillair bepalen op de ABL90 Volgens mijn collega's stond het apparaat meetklaar. De ene scant het monsternummer terwijl de andere de capillair voorbereid om aan te bieden. Mengen, dopjes eraf... Op het moment dat het monster aangeboden wordt, print de ABL ineens uitslagen uit onder het gescande monsternummer, tijd: 08.28 uur Die uitslagen worden meteen naar glims gestuurd en staan dus dan bij het ordernummer gerapporteerd. 7 minuten later print de ABL de uitslagen uit van het capillaire monster onder hetzelfde monsternummer, tijd:08.35 uur. De uitslagen van 08.28 uur worden op verzoek van de 2 analisten verwijderd door de GA analist en die van 08.35 uur handmatig ingevoerd. De vraag is wat er mis is gegaan, hoe dan ook mag zoiets natuurlijk nooit gebeuren Zoals telefonisch afgesproken bijgevoegd de gegevens van het incident. De print outs van de ABL. De gegevens van het kindje waar het capillaire monster van was met de gerectificeerde uitslagen. Reza belde net al dat hij langs komt om meteen gegevens uit de ABL te halen voor onderzoek. | Health Hazard Case $Techline Case naar RMED gemeld. ServiseDump file en datalogs heb gekopieerd voor het onderzoek door RMED. | SWO-284584 | Defect | SW | Patient Mixup | Patient Mixup | Techline raised | Results of a previous sample under scanned sample number | The following complaint has occurred; What happened: 2 colleagues want to determine a capillary on the ABL90 According to my colleagues, the device was ready for measurement. One scans the sample number while the other is preparing the capillary to offer. Mixing, caps off... At the moment the sample is offered, the ABL will suddenly print results under the scanned sample number, time: 08.28 hours These results are immediately sent to Glims and are then reported with the order number. 7 minutes later, the ABL prints the results of the capillary sample under the same sample number, time: 08.35 hours. The results of 08.28 hours will be removed by the GA analyst at the request of the 2 analysts and entered manually from 08.35 hours. The question is what has gone wrong, anyway such a thing should of course never happen As agreed by telephone, the details of the incident are attached. The print outs of the ABL. The data of the child where the capillary sample was with the rectified results. Reza just called him to come by to get data from the ABL for research immediately. | Health Hazard Case $Techline Case to RMED reported. ServiseDump file and data logs have been copied for research by RMED.
Yes | Visit | Complaint | RCHN | 报温度过低 | 报温度过低，借902-972一个 |  | SWO-282543 | [5 empty] | The temperature is too low. | The temperature is too low, borrow 902-972 one
Yes | Visit | Complaint | RCHN | 进样架故障 | 进样架故障，借903-584一个 | 更换903-584后修复(By Ran) | SWO-281122 | [5 empty] | Sample frame failure | Sample frame failure, borrow 903-584 one | Repair after replacement 903-584 (By Ran)
Yes | Visit | Complaint | RCHN | 报错5684 | 报错5684 | 更换试剂包后修复 | SWO-282511 | [5 empty] | 5684 error | 5684 error | Repair after replacement of reagent kit
No | Visit | Complaint | RCHN | BNP定标不过，干燥前温度超范围 | BNP定标不过，干燥前温度超范围，借903-210一个(高工确认取消借件） | 调整Cup dryer Pcb后修复(By Ran) | SWO-281161 | Defect | HW | Process Module | PCB, Cup Dryer |  | BNP calibration is not the case, the temperature exceeds the range before drying | BNP calibration however, before drying temperature over range, borrow 903-210 one (high work confirmation cancellation of the loan) | Post-Fix (By Ran) Adjustcup erpcb
No | Visit | Complaint | RCHN | 干燥前温度超范围 | 干燥前温度超范围，借903-210一个 | 调整Cup dryer Pcb后修复(By Ran) | SWO-281154 | Defect | HW | Process Module | PCB, Cup Dryer |  | Overrange temperature before drying | Over-range temperature before drying, borrow 903-210 one | Post-Fix (By Ran) Adjustcup erpcb
No | Visit | Complaint | RIND | QC Ampoules not being scanned in Autocheck module | Autocheck barcode scanner not working | Replaced barcode scanner from stock. Performed laser alignment successfully. 2 point calibration performed successfully. | SWO-283054 | Defect | HW | Autocheck Module | Barcode reader, Autocheck Module |  | QC Ampoules not being scanned in Autocheck module | Autocheck barcode scanner not working | Replaced barcode scanner from stock. Performed laser alignment successfully. 2 point calibration performed successfully.
Yes | Visit | Complaint | RDAN | Hanne ringer og fortæller, at hun har problemer med Na og Ca målingerne på ITA 1 apparatet i forhold til ITA 2 apparatet. ABL1 måler umiddelbart højere end ABL1. Hanne har forsøgt at bytte elektrode med ABL2 og kan se, at hun ikke kan flytte problemet. | Afprøvet. | Årsservice udført, service tHb kalibreret, målekamre udskiftet, slutafprøvet. | SWO-282107 | [5 empty] | Hanne rings and tells me that she has problems with Na and Ca measurements on the ITA 1 apparatus compared to the ITA 2 apparatus. ABL1 measures immediately higher than ABL1. Hanne has tried to swap the electrode with ABL2 and can see that she can not move the problem. | Tested. | Annual service performed, service tHb calibrated, measuring chambers replaced, final tested.
No | 4 | Y | RPOL | Blank | Mimo czyszczenia optical unit oraz process wheel, wystêpowa³y powtarzaj¹ce siê b³êdy QC na troponinie. Wymieniono Optical Unit ora Process Wheel. Kalibracja, testy analizatora. Aparat sprawny. |  | D39EADE3E059DA41 | Defect | HW | Process Module | Optical Unit |  | Blank | Although the optical unit and process wheel were cleaned, the QC errors were repeated on the troponin. Listed Optical Unit or Process Wheel. Calibration, analyzer tests. The camera is functional.
No | Visit | Complaint | RAME | Analyzer posted an error message to call radiometer | Pop up box stated analyzer could not continue call radiometer after changing a SC and SP | Reloaded win 7 and application SW, and setup and changed SC due to length of time inactive. All systems are back running normally | SWO-283554 | Defect | SW | Windows | Win Pop Up |  | Analyzer posted an error message to call radiometer | Pop up box stated analyzer could not continue call radiometer after changing a SC and SP | Reloaded win 7 and application SW, and setup and changed SC due to length of time inactive. All systems are back running normally
No | Visit | Complaint | RIND | No reportable parameter except PO2. | Check the calibration log & found no reportable parameter except PO2 | Check the SC80 and found micro clot in the cassette,flush the cassette with the distilled water & heparin,activate all the parameter & the calibration passed successfully. | SWO-283625 | [5 empty] | No reportable parameter except PO2. | Check the calibration log & found no reportable parameter except PO2 | Check the SC80 and found micro clot in the cassette,flush the cassette with the distilled water & heparin,activate all the parameter & the calibration passed successfully.
No | Visit | Complaint | RCHN | 无法开机 | 无法开机，借910-342一个 | 更换910-342后修复 | SWO-281814 | Defect | HW | DMS_PC | DMS_PC |  | Cannot power on | Cannot power on, borrow 910-342 one | Repair after replacement 910-342
Yes | Visit | Complaint | RCHN | 报错1031 | 报错1031，借903-364一个 | Changed the optical unit, recal, then OK(By Ran) | SWO-282439 | [5 empty] | 1031 | 1031, loan 903-364 one | Changed the optical unit, recal, then OK(By Ran)
No | Visit | Complaint | RLTD | Please call Johanna on 02381 208 368 | Checked the results on the monitor and the data from its last use was on the monitor and it could be downloaded to a USB drive. When the monitor was connected to Viso Viso found the monitor but it showed nothing to download. The download setting were checked on the monitor and these all appeared to be correct. | A system dump was taken from the monitor so that it could be sent for analysis to see if it could pin-point the problem. A system dump was also take from the other TCM5 which has an intermittent error code occurring. | SWO-283353 | [5 empty] | Please call Johanna on 02381 208 368 | Checked the results on the monitor and the data from its last use was on the monitor and it could be downloaded to a USB drive. When the monitor was connected to Viso Viso found the monitor but it showed nothing to download. The download setting were checked on the monitor and these all appeared to be correct. | A system dump was taken from the monitor so that it could be sent for analysis to see if it could pin-point the problem. A system dump was also take from the other TCM5 which has an intermittent error code occurring.
No | Visit | Complaint | RAME | Analyzer is unable to run auto QC | Troubleshot and found that the autocheck house would not operate. I also found a damaged wire on the ampoule breaker motor. | Replaced the FPC cable for the autocheck housing, and replaced the ampoule breaker motor, but the autocheck housing issue was not resolved. | SWO-284556 | Defect | HW | Autocheck Module | Stepper Motor, Breaker arm |  | Analyzer is unable to run auto QC | Troubleshot and found that the autocheck house would not operate. I also found a damaged wire on the ampoule breaker motor. | Replaced the FPC cable for the autocheck housing, and replaced the ampoule breaker motor, but the autocheck housing issue was not resolved.
No | Visit | Complaint | RFRA | Erreur 5670 et 4036 | Manipulations infructueuses par téléphone avec technicienne | Présence de cupules derrière le pac de solution $ $ une fois les cupules retirées remise en route du dispositif : erreur 5670 disparut mais lors du passage d'un contrôle qualité MCC 1 : erreur 4036 . Démontage de l'inlet wheel et de son capteur pour repositionnement ok $ $ redémarrage du dispositif ok | SWO-283514 | Revisit required | HW | Aspiration Module | Inlet Wheel | 5670, 4036 | Error 5670 and 4036 | Unsuccessful phone manipulations with technician | Presence of cups behind the $$ solution pac once the cups removed restarted the device: error 5670 disappeared but during the passage of a quality control MCC 1: error 4036 . Unwinding the inlet wheel and its sensor for repositioning ok $$ reboot of the device ok
No | Visit | Complaint | RCHN | Lac定标不过 | Lac定标不过，借903-581（缺货） | 使用公司备用机的903-581修复 by Andrew | SWO-284209 | Defect | HW | Basic Unit | PCB, Sensor interface |  | Lac calibration is not | Lac bid is not, borrow 903-581 (out of stock) | 903-581 Repair by Andrew using the company's spare machine
Yes | Visit | Complaint | RIND | Calibration error. | Check the system. | check the logs. Replace the new sc9. Calibrate the system multiple time. | SWO-283656 | [5 empty] | Calibration error. | Check the system. | check the logs. Replace the new sc9. Calibrate the system multiple time.
No | Visit | Complaint | RBNL | Error 581 | Error code 581, OXI spectrum mismatch. De FO2Hb en de FMetHb en de Bili QC's zijn afwijkend. | Flush procedure en daarna complete oximetrie test vervolgens kalibratie $QC's heb bepaald, OK. | SWO-283731 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | Error 581 | Error code 581, OXI spectrum mismatch. From FO2Hb to FMetHb en de Bili QC's zijn afwijkend. | Flush procedure and then complete oximetry test then calibration $QC's have determined, OK.
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | LAC,CA定标失败，0375、0378 | 更换测试卡后修复 | SWO-284152 | [5 empty] | Test card, reagent package failure | LAC, CA calibration failure, 0375, 0378 | Repair after replacing test card
No | Visit | Complaint | RCHN | 光学单元超出范围，报错：1029 | 光学单元超出范围，报错：1029,需要借件903-364 | 更换903-364后修复 by Andrew | SWO-283618 | Defect | HW | Process Module | Optical Unit | 1029 | Optical unit out of range, error: 1029 | Optical unit out of range, error: 1029, need to borrow 903-364 | Repair by Andrew after replacing 903-364
Yes | Visit | Complaint | RCHN | 报温度过低 | 报温度过低，需要借件903-334一个 | 更换903-334后修复 by Andrew | SWO-283683 | [5 empty] | The temperature is too low. | The temperature is too low, need to borrow 903-334 one | Repair by Andrew after replacing 903-334
No | Visit | Complaint | RIND | machine suddenly switched off customer tried to switch on but its not switch on. | Checked the machine and found waste leaking from solution pack and found solution pack waste port having clot. | cleaned the solution pack waste port and checked the machine due to waste leakage power supply unit short circuit .we need to replace power supply unit for this issue. | SWO-284169 | [5 empty] | machine suddenly switched off customer tried to switch on but its not switch on. | Checked the machine and found waste leaking from solution pack and found solution pack waste port having clot. | cleaned the solution pack waste port and checked the machine due to waste leakage power supply unit short circuit .we need to replace power supply unit for this issue.
No | Visit | Complaint | RKK | 装置がランニング中にランダムに発生。 | SCのロットが変わっても発生 | センサボード、シグナルデータボード交換。 | SWO-285705 | [5 empty] | The device occurs randomly while running. | Occurs even if the sc lot changes | Sensor board, signal data board exchange.
Yes | Visit | Complaint | RCHN | tni定标失败，光学单元超出范围 | tni定标失败，光学单元超出范围，需要借件903-364和903-236一个 | 更换903-364和903-236备件后修复 by Andrew | SWO-283700 | [5 empty] | tni calibration failed, optical unit out of range | tni calibration failure, optical unit out of range, need to borrow 903-364 and 903-236 one | Repair by Andrew after replacing 903-364 and 903-236 spare parts
No | Visit | Complaint | RCHN | 新机装机钱检测，报错1024 | 报错1024，借944-212（13264缺货） | 重新定标后修复 | SWO-283953 | Assisted Replacement | Consumables | Test Kit | Cartridge, TNi | 1024 | New machine installed money detection, reported error 1024 | 1024, 944-212 (13264 out of stock) | Repair after re-calibration
Yes | Visit | Complaint | RESP | Pierde líquido por la entradsa de capilares. | Pierde líquidos por la entrada de capilares y no pasa el test de fugas. | Colocar tubo suelto en la salida al desperdicio, limpiar electrodo de pH, calibrar y probar con controles. | SWO-284058 | [5 empty] | Lose fluid through the entradsa of capillaries. | Loses fluids through capillary ingress and does not pass the leak test. | Place loose tube in the outlet to waste, clean pH electrode, calibrate and test with controls.
No | Visit | Complaint | RKK | ガスの減りが早い | ガスの減りが早 | フローレギュレータ交換 | SWO-284796 | Defect | HW | Gas module | Flow regulator |  | The gas is running low quickly. | The decrease in gas is fast. | Flow regulator replacement
No | Visit | Complaint | RIND | Display is not working or showing anything. | Found display is not showing anything. | Reconnect all display cable & CPU. But still problem is same. CPU required to change. | SWO-284948 | Revisit required | HW | DMS_PC | DMS_PC |  | Display is not working or showing anything. | Found display is not showing anything. | Reconnect all display cable & CPU. But still problem is same. CPU required to change.
No | Visit | Complaint | RKK | setsafe22 | 角田さん指示のもと、カートリッジスライド時の動きを確認。 2＆3辺りでカートリッジの動きが渋くなる。 ストレージホイールベースのバリを除去して改善した。 様子見。 | 角田さん指示のもと、カートリッジスライド時の動きを確認。 2＆3辺りでカートリッジの動きが渋くなる。 ストレージホイールベースのバリを除去して改善した。 様子見。 | SWO-283980 | Defect | Consumables | Test Kit | Cartrdige | Broken | setsafe22 | Under Tsunoda's instructions, we checked the movement at the time of cartridge slide. The movement of the cartridge becomes reluctant around 2&3. Improved by removing storage wheelbase burrs. Wait. | Under Tsunoda's instructions, we checked the movement at the time of cartridge slide. The movement of the cartridge becomes reluctant around 2&3. Improved by removing storage wheelbase burrs. Wait.
No | Visit | Complaint | RLTD | Jackie on 01225 824964 | Found that the tube between the sensor bay and the oxi module was leaking. | The annual service was due so this was done as there is a replacement tube in the service kit. | SWO-284561 | Defect | HW | Oxi module | Pinch valve tube |  | Jackie on 01225 824964 | Found that the tube between the sensor bay and the oxi module was leaking. | The annual service was due so this was done as there is a replacement tube in the service kit.
Yes | Visit | Complaint | RAME | The parameter keep dropping out intermittently. Everyday the parameters have to be deactivated manually. | Fails the system cycle most of the time. Re-activating the parameter may or may not work. | Replaced the spectrometer and hemolyzer and the analyzer finally passed a system cycle. These same parts were replaced in October. | SWO-284605 | [5 empty] | The parameter keep dropping out intermittently. Everyday the parameters have to be deactivated manually. | Fails the system cycle most of the time. Re-activating the parameter may or may not work. | Replaced the spectrometer and hemolyzer and the analyzer finally passed a system cycle. These same parts were replaced in October.
No | Visit | Complaint | RIND | Service Visit | Barrier between male and female connector of solution pack and analyzer respectively. | Sanitize the solution pack smart chip as well as female connector of analyzer. Check the molded hinge door assembly, tighten the screw of that. Connect the solution pack, now its connected. Take the permission for the preventive maintenance. | SWO-285595 | [5 empty] | Service Visit | Barrier between male and female connector of solution pack and analyzer respectively. | Sanitize the solution pack smart chip as well as female connector of analyzer. Check the molded hinge door assembly, tighten the screw of that. Connect the solution pack, now its connected. Take the permission for the preventive maintenance.
No | Visit | Complaint | RESP | NO CALIBRA | NO CALIBRA | SE DETECTA CONECTOR DE CASETE EN MAL ESTADO, SE SUSTITUYE ESTE Y SE EFECTUAN PRUEBAS DE FUNCIONAMIENTO. CALIBRADO Y VERIFICADO | SWO-284510 | [5 empty] | DON'T CALIBRATE | DON'T CALIBRATE | CASE CONNECTOR IS DETECTED IN THE WORLD STATE, THIS IS SUBSTITUTE AND OPERATION TESTS ARE EFFECTED. CALIBRATED AND VERIFIED
Yes | Visit | Complaint | RCHN | 前面板损坏，条码扫码器无法扫描 | 前面板损坏，条码扫码器无法扫描，需要借件903-224、902-924、910-304各一个 | 更换903-224、902-924、910-304后修复 by Andrew | SWO-283842 | [5 empty] | Damaged front panel, barcode scanner unable to scan | Damaged front panel, bar code sweeper can not be scanned, need to borrow 903-224, 902-924, 910-304 each one | Repair by Andrew after replacing 903-224, 902-924, 910-304
No | Visit | Complaint | RIND | Calibration error. | Routine visit | routine visit. | SWO-285781 | [5 empty] | Calibration error. | Routine visit | routine visit.
No | Visit | Complaint | RDEU | WE - Neonlamp intensity too high | Oxi-Complete-Test durchgeführt Fehlermeldung 258 und 316 | Neon-Int. und Neon-Position fehlerhaft --> Ersatzteil wird bestellt | SWO-284283 | Defect | HW | Oxi module | Lamp unit | Neon | WE - Neonlamp intensity too high | Oxi Complete Test Performed Error message 258 and 316 | Neon Int. and neon position faulty --> Spare part is ordered
No | Visit | Complaint | RKK | チケットプリンター（CBM820）が使用できなくなる時があり再起動が必要になる。 | 訪問当日は問題なくここ最近は症状なしと。 | Loｇの収集 | SWO-284353 | No fault found | SW | Printer external | Printer external |  | The ticket printer (CBM820) may become unavailable and a reboot will be required. | There are no problems on the day of the visit, and there are no symptoms recently. | Collecting Log
No | Visit | Complaint | RDEU | QC wird nicht gemessen | Auto Check defekt (Karussell fährt hoch und ohne QC Messung wieder herunter) | Erneuerung Auto Check Modul - Software Installation - Funktionsüberprüfung etc. | SWO-284418 | Defect | HW | Autocheck Module | Autocheck Module |  | QC is not measured | Auto Check defective (carousel goes up and down again without QC measurement) | Renewal Auto Check Module - Software Installation - Functional Check etc.
No | Visit | Complaint | RDEU | 27.12.2017; Fr. Mengesdorf; Tel.: 062150335589; Oxy Hardwarefehler | 251 Halogen Lampe Intensität zu hoch 579 Oximetrie Hardware Problem | Einstellungen Oximetrie - Funktionsüberprüfung etc. | SWO-284416 | Adjustment | HW | Oxi module | Lamp unit | Halogen | 27.12.2017; Ms. Mengesdorf; Tel.: 062150335589; Oxy Hardware Error | 251 Halogen lamp intensity too high 579 Oximetry Hardware Problem | Settings Oximetry - Function check, etc.
Yes | Visit | Complaint | RAME | Analyzer has Temperature Error messages | Sensor Cassette overheating. | User turned off analyzer for two days. User turned the analyzer back on and replaced the sensor cassette. After four hours the pCO2, Glu and Lac had very low sensitivities. Performed temperature calibration and replaced the sensor cassette. Analyzer now passes calibrations and QCs and the temperature is stable. | SWO-285127 | [5 empty] | Analyzer has Temperature Error messages | Sensor Cassette overheating. | User turned off analyzer for two days. User turned the analyzer back on and replaced the sensor cassette. After four hours the pCO2, Glu and Lac had very low sensitivities. Performed temperature calibration and replaced the sensor cassette. Analyzer now passes calibrations and QCs and the temperature is stable.
Yes | Visit | Complaint | RAME | Battery showing a 0% charge | Battery showing 0 charge but analyzer did not shut down while plugged in and lasted a while after unplugged $ battery had power but showed no power | Swapped battery with ER unit that is not mobile until new battery comes in. | SWO-285288 | [5 empty] | Battery showing a 0% charge | Battery showing 0 charge but analyzer did not shut down while plugged in and lasted a while after unplugged / battery had power but showed no power | Swapped battery with ER unit that is not mobile until new battery comes in.
No | Visit | Complaint | RAME | Battery Error | Battery error on analyzer | Disconnect the battery until new battery comes in. | SWO-285866 | Revisit required | HW | Battery module | Battery |  | Battery Error | Battery error on analyzer | Disconnect the battery until new battery comes in.
No | Visit | Complaint | RDEU | Lösungspack muss erneuert werden, Anwender hat dies bereits durchgeführt jedoch kommt Aufforderung zur Benutzeraktion wieder | Ständige Benutzeraktion gefordert auch nach LP Wechsel | Fehleranalyse, Lösungspack und Sensorkassette erneuert, Einlassdichtung erneuert, Einlasssonde gerichtet, Service Complete Check i. O., Kalibrierung incl. tHb und Qualitätskontrolle i. O. | SWO-284939 | Assisted Replacement | SP | SP | FT |  | Solution pack needs to be renewed, user has already done this but prompt for user action comes back | Constant user action required even after LP change | Fault analysis, solution pack and sensor cartridge renewed, inlet seal renewed, inlet probe directed, Service Complete Check in the direction of the Calibration incl. tHb and quality control as a.
No | Visit | Complaint | RFRA | erreur vaccum value | erreur vaccum value | chagement tuyau , joint de connexion, cassette => erreur toujours présente, changement joint d'introduction => OK | SWO-285114 | Defect | HW | Auto Inlet module | Inlet gasket |  | erreur vaccum value | erreur vaccum value | hose, connection seal, cassette - error always present, change intle in , OK
No | Visit | Complaint | RAME | Keeps disconnecting from the network | Analyzer keeps disconnecting from the network. Upon arrival, I found the analyzer connected to the network and, radiance. Cynthia states she had the IT department checking the network connection. | I checked the communication settings and, they are correct. I disconnected and, reconnected the network cable and, the analyzer reconnected. However, after waiting a few minutes the connection was lost and, it was slow to reconnect. Cynthia and, I moved the analyzer to one of the cath labs and, compared the connection settings against another analyzer and, they were all the same settings. While in the room, we also connected the analyzer to the network and, waited for 10 minutes with seeing any connection loss. We brought the analyzer back to Cynthia's office and, after connecting it to the network, the analyzer got disconnected again in less than 3 minutes. Cynthia tried another network cable but this didn't help solve the problem. Recommended Cynthia to contact the IT department and, ask the to check the network connection again. | SWO-284666 | [5 empty] | Keeps disconnecting from the network | Analyzer keeps disconnecting from the network. Upon arrival, I found the analyzer connected to the network and, radiance. Cynthia states she had the IT department checking the network connection. | I checked the communication settings and, they are correct. I disconnected and, reconnected the network cable and, the analyzer reconnected. However, after waiting a few minutes the connection was lost and, it was slow to reconnect. Cynthia and, I moved the analyzer to one of the cath labs and, compared the connection settings against another analyzer and, they were all the same settings. While in the room, we also connected the analyzer to the network and, waited for 10 minutes with seeing any connection loss. We brought the analyzer back to Cynthia's office and, after connecting it to the network, the analyzer got disconnected again in less than 3 minutes. Cynthia tried another network cable but this didn't help solve the problem. Recommended Cynthia to contact the IT department and, ask the to check the network connection again.
No | Visit | Complaint | RCHN | 报错0581和0299 | 报错0581和0299.需要借件902-980一个 | 更换902-980解决 | SWO-284815 | Defect | HW | Oxi module | Hemolyzer |  | Error0581 and 0299 | 0581 and 0299. Need to borrow 902-980 one | Replacement 902-980 resolution
No | Visit | Complaint | RCHN | 温度过高 | 温度过高，需要借件903-581和903-199各一个 | 更换903-581和903-199备件后修复 by Andrew 903-581 was used for solving Glu parameter failed at calibration. | SWO-283945 | Defect | HW | Basic Unit | Cassette retainer |  | Too high temperature | The temperature is too high, and one each of the 903-581 and 903-199 loans is required. | 更换903-581和903-199备件后修复 by Andrew 903-581 was used for solving Glu parameter failed at calibration.
No | Visit | Complaint | RCHN | TNI定标失败，1024 | TNI定标失败，1024，借903-364,944-212（13643）各一个 | 更换903-364后重新定标后修复 by Andrew | SWO-284257 | Defect | HW | Process Module | Optical Unit | 1024 | TNI bid failure, 1024 | TNI bid failure, 1024, borrowing 903-364, 944-212 (13643) each | Repair by Andrew after re-calibration after replacement of 903-364
No | Visit | Complaint | RCHN | 报错3414 | 报错3414，借903-329，905-960各一个 | 更换903-329,905-960后修复 by Andrew | SWO-283964 | Defect | HW | Aspiration Module | Needle Wash Unit | 3414 | Error 3414 | Error 3414, lending 903-329, 905-960 each | Repair by Andrew after replacing 903-329, 905-960
Yes | Visit | Complaint | RCHN | 报错1035 | 报错1035，借903-364一个 | 更换903-364后修复 by Andrew | SWO-284199 | [5 empty] | 1035 error | 1035, loan 903-364 one | Repair by Andrew after replacing 903-364
No | Visit | Complaint | RCHN | 温度传感器与WASH UNIT连接线断裂 | 温度传感器与WASH UNIT连接线断裂，借905-908一个 | 更换905-908备件后修复 by Andrew | SWO-283997 | Defect | HW | Process Module | Sensor, BNP |  | The temperature sensor breaks with the WASH UNIT cable | The temperature sensor is disconnected from the WASH UNIT cable, with a 905-908 | Repair by Andrew after replacement of 905-908 spare parts
No | Visit | Complaint | RCHN | TNI定标不过 | TNI定标不过，借903-364,944-212（13315缺货），944-258（13313缺货），（13355），944-394（13419缺货），944-220（13566,13354）各一个 | 光学单元和定标卡都未使用。客户要求，暂时不做TNI了。先将库存的测试卡用完再说。by Andrew Please refer to SWO-284709 mentioned. | SWO-284228 | Defect | HW | Process Module | Optical Unit |  | TNI calibration is not the | TNI bid, however, borrows 903-364, 944-212 (13315 out of stock), 944-258 (13313 out of stock), (13355), 944-394 (13419 out of stock), 944-220 (13566, 13354) each one | Neither the optical unit nor the calibration card is used. Customer request, do not do TNI for the time being. Use up the stock test card first. by Andrew Please refer to SWO-284709 mentioned.
No | Visit | Complaint | RCHN | 报错1024 | 报错1024，借903-364,636-541各一个 | 更换903-364,636-541后修复 by Andrew | SWO-284151 | Defect | HW | Process Module | Optical Unit | 1024 | 1024 error | 1024, borrow 903-364, 636-541 each | Fix by Andrew after replacing 903-364, 636-541
No | Visit | Complaint | RCHN | 测试卡舱门损坏 | 测试卡舱门损坏,需要借件925-016一个 | 更换925-016后修复 by Andrew | SWO-283916 | Defect | HW | Cover and Chassis | Cartridge Lid |  | Test card door damage | Test card door damage, need to borrow 925-016 one | Fix by Andrew after replacing 925-016
No | Visit | Complaint | RESP | ERRORES BOMBAS | FUGAS | SE DETECTA FUGA EN ELECTRODO FALSO DE CLORO PROVOCANDO ERRORES EN MUESTRAS Y CALIBRACIONES DE BOMBAS. SE DOTA DE JUNTA RORICA DE GOMA Y SE EFECTUAN CALIBRACIONES DE BOMBAS Y PRUEBAS CON MUESTRAS. CALIBRADO Y VERIFICADO CON Q.C. | SWO-285087 | Defect | HW | pH/BG Module | Pump tubes |  | BOMB ERRORS | Leaks | FAILURE IS DETECTED IN FALSE CLORO ELECTRODE CAUSING ERRORS IN SAMPLES AND PUMP CALIBRATIONS. GOMA RORIC JOINT IS AVAILABLE AND CALIBRATIONS OF PUMPS AND TESTS WITH SAMPLES ARE EFFECTED. CALIBRATED AND VERIFIED WITH Q.C.
Yes | Visit | Complaint | RIND | Print not coming. | Print not coming. | Printer head found broken. Need to replace printer head. Serviced printer head replaced mean while. Analyser Serviced and calibrated properly. Handed over in a fine working condition. | SWO-285555 | [5 empty] | Print not coming. | Print not coming. | Printer head found broken. Need to replace printer head. Serviced printer head replaced mean while. Analyser Serviced and calibrated properly. Handed over in a fine working condition.
Yes | Visit | Complaint | RLTD | Ranvir on 01865 221596 | The Crea electrodes were replaced on 22 $12 but Crea is still failing level 1 & 3 QCs. | Replaced the measuring chamber and swapped the electrode connectors with ones from the dummy electrodes on the MET2 module. Ran 2 point calibrations until they passed and ran level 1 QC but the Crea was still below the lower limit. Changed the Crea membranes and the Cal 1 & Cal 2 solutions. 2 point calibrations ran until passed and then the failed QCs were rerun and this time they passed. | SWO-284879 | [5 empty] | Ranvir on 01865 221596 | The Crea electrodes were replaced on 22 /12 but Crea is still failing level 1 & 3 QCs. | Replaced the measuring chamber and swapped the electrode connectors with ones from the dummy electrodes on the MET2 module. Ran 2 point calibrations until they passed and ran level 1 QC but the Crea was still below the lower limit. Changed the Crea membranes and the Cal 1 & Cal 2 solutions. 2 point calibrations ran until passed and then the failed QCs were rerun and this time they passed.
No | Visit | Complaint | RDEU | WE - Neonlamp intensity too high | Siehe SWO | Oxi-Module-PCB erneuert --> Oxi-complete-Test -->Fehler bleibt bestehen Ersatzteil wird bestellt --> Station ist informiertQC-Ampullen werden nicht erkannt --> beide Umlenkspiegel des AQC-Moduls gereinigt | SWO-284972 | Revisit required | HW | Oxi module | Lamp unit | Neon | WE - Neonlamp intensity too high | See SWO | Oxi-Module-PCB renewed --> Oxi-complete-Test -->Error remains Spare part is ordered --> Station is informedQC ampoules are not detected --> both deflection mirrors of the AQC module cleaned
No | Visit | Complaint | RAME | Analyzer is unable to run auto QC | Troubleshot to a faulty housing tilt motor. | Replaced housing tilt motor, and this resolved the problem. | SWO-285257 | Defect | HW | Autocheck Module | Stepper Motor, Tilt |  | Analyzer is unable to run auto QC | Troubleshot to a faulty housing tilt motor. | Replaced housing tilt motor, and this resolved the problem.
No | Visit | Complaint | RFRA | Lors du dosage sang totale erreur aléatoire 1068 | Erreur cup dryer PCB | Changement de la cup dryer PCB $ $ Nettoyage avec cleaning solution $ $ passage d'un test CRP OK | SWO-285111 | Defect | HW | Process Module | PCB, Cup Dryer | 1068 | During total blood dosing random error 1068 | Erreur cup dryer PCB | Change of cup dryer PCB $$ Cleaning with cleaning solution $$ passing a CRP OK test
No | Visit | Complaint | RFRA | Erreur 5670 et 4036 | Appel téléphonique : de nouveau erreurs 5670 et 4036. Manipulations par téléphone infructueuses | Changement de l'inlet wheel et des photo-interrupteurs de position de la roue ( DIV ) ok $ $ inspection du système de verrouillage de la roue ok $ $ redémarrage du dispositif $ $ tests de l'inlet wheel ok | SWO-285008 | Adjustment | HW | Aspiration Module | Inlet Wheel | 5670, 4036 | Error 5670 and 4036 | Phone call: again errors 5670 and 4036. Unsuccessful phone manipulations | Change of the wheel inlet and photo-switches of wheel position (DIV) ok $$ inspection of the wheel locking system ok $$ reboot of the device $$ wheel test ok
Yes | Visit | Complaint | RAME | Unable to pump solution | Unable to pump solution | Flushed system within warm water several times. Replaced sensor cassette and several solution packs. | SWO-285843 | [5 empty] | Unable to pump solution | Unable to pump solution | Flushed system within warm water several times. Replaced sensor cassette and several solution packs.
No | Visit | Complaint | RAME | Customer reported inhomogeneous rinse error during during precision QC testing. | Inhomogeneous rinse error and vacuum level high during pressure test. | Replaced solution pack. Error still present. Adjusted inlet gasket and error cleared. Replaced inlet gasket. Replaced sensor cassette due to errors created by previous cassette not being hydrated while analyzer sat. | SWO-285594 | Defect | HW | Inlet module | Inlet gasket |  | Customer reported inhomogeneous rinse error during during precision QC testing. | Inhomogeneous rinse error and vacuum level high during pressure test. | Replaced solution pack. Error still present. Adjusted inlet gasket and error cleared. Replaced inlet gasket. Replaced sensor cassette due to errors created by previous cassette not being hydrated while analyzer sat.
No | Visit | Complaint | RCHN | 连接数据串口烧坏 | 连接数据串口烧坏，借903-381一个 | 更换903-381后修复 by Andrew | SWO-284857 | [5 empty] | Connection data serial sinters burned out | Connection data serial snout burned, borrow 903-381 one | Fix by Andrew after replacing 903-381
No | Visit | Complaint | RCHN | 仪器把手故障 | 仪器把手故障，借备件886-606一个 | 更换886-606一个正常。by Andrew | SWO-284737 | Defect | HW | Cover and Chassis | Handle |  | Instrument handle failure | Instrument handle failure, spare parts 886-606 one | Replace 886-606 one normal. by Andrew
No | Visit | Complaint | RCHN | 报错：1342 | 报错：1342。需要借件903-129一个 | 更换903-129一个正常。by Andrew | SWO-284712 | Defect | HW | Oxi module | Pinch valve |  | Error: 1342 | Error: 1342. Need to borrow 903-129 one | Replace the 903-129 one normal. by Andrew
Yes | Visit | Complaint | RLTD | 11 repaired monitors ready for return to customer. | Returned 11 repaired TCM5 monitors serial nos: 500R0021N003, 500R0023N007, 500R0025N003, 500R0025N005, 500R0025N007, 500R0026N002, 500R0026N006, 500R0026N007, 500R0026N008, 500R0026N009 & 500R0029N010. Fitted batteries, gas bottles and sensors from loan monitors - set date & time and restored customer set-up. Verified that every monitor calibrated without errors. Deleted all patient data on loan monitors and removed from site serial nos. 500R0032N004, 500R0035N004, 500R0055N014 (RLTD) 500R0039N001, 500R0040N001, 500R0040N009, 500R0041N002, 500R0041N005, 500R0041N009, 500R0054N007, 500R0054N010 (RBSL) | Returned repaired monitors. Removed loan monitors. | SWO-285443 | [5 empty] | 11 repaired monitors ready for return to customer. | Returned 11 repaired TCM5 monitors serial nos: 500R0021N003, 500R0023N007, 500R0025N003, 500R0025N005, 500R0025N007, 500R0026N002, 500R0026N006, 500R0026N007, 500R0026N008, 500R0026N009 & 500R0029N010. Fitted batteries, gas bottles and sensors from loan monitors - set date & time and restored customer set-up. Verified that every monitor calibrated without errors. Deleted all patient data on loan monitors and removed from site serial nos. 500R0032N004, 500R0035N004, 500R0055N014 (RLTD) 500R0039N001, 500R0040N001, 500R0040N009, 500R0041N002, 500R0041N005, 500R0041N009, 500R0054N007, 500R0054N010 (RBSL) | Returned repaired monitors. Removed loan monitors.
No | Visit | Complaint | RCHN | 定标错误 | 定标错误 | 更换903-364后修复 by Andrew | SWO-285299 | Defect | HW | Process Module | Optical Unit |  | Calibration error | Calibration error | Repair by Andrew after replacing 903-364
No | Visit | Complaint | RIND | Oximetry parameters are failed in calibration. Neon lamp intensity is low. | Checked fluidic path and found there was clot. Also given complete oximetry test but it was failed. | Opened oximetry unit and reconnected all cable.Then given complete oximetry test and it was passed. Then performed 2 point calibration.Instrument is working fine. | SWO-285745 | Cleaned | Fluidic pathway | Oxi module | Oxi module, Cleaned |  | Oximetry parameters are failed in calibration. Neon lamp intensity is low. | Checked fluidic path and found there was clot. Also given complete oximetry test but it was failed. | Opened oximetry unit and reconnected all cable. Then given complete oximetry test and it was passed. Then performed 2 point calibration. Instrument is working fine.
No | Visit | Complaint | RDEU | WE - Neonlamp intensity too low Fehlermeldung 258/316 | Siehe SWO | Oxi-Modul erneuert --> SW-Update-->Halogenlampe justiert--> Oxi-Complete-Test i.O Setvice-tHb-Kal. und tHb-Kal. durchgeführt Kal und QC i.O | SWO-285390 | Defect | HW | Oxi module | Lamp unit | Neon | WE - Neonlamp intensity too low Error message 258/316 | See SWO | Oxi module renewed --> SW-Update-->Halogen lamp adjusted--> Oxi-Complete-Test i.O Setvice-tHb-Kal. and tHb-Kal. Performed Kal and QC i.O
No | Visit | Complaint | RAME | Oxi error 258,316 | Oxi calibration erreur full oxi test néon lampe hors limi | Remplacé lampe unit | SWO-285596 | Defect | HW | Oxi module | Lamp unit | Neon | Oxi error 258,316 | Oxi calibration error full oxi neon lamp test off limi | Replaced lamp unit
No | Visit | Complaint | RFRA | Pb cal inlet | Contrôle du module inlet | Changement du capteur de position. Essais | SWO-285454 | Defect | HW | Inlet module | Friction detection |  | Pb cal inlet | Control of the inlet module | Change of position sensor. Tests
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 检测到漏电流 | 更换试剂包正常。 | SWO-285821 | [5 empty] | Test card, reagent package failure | Drain current detected | The replacement kit is OK.
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | GLU,LAC定标不过 0378 | 连续更换6张测试卡后修复 | SWO-285900 | [5 empty] | Test card, reagent package failure | GLU, LAC calibration is no more than 0378 | Repair after 6 consecutive test cards
Yes | Visit | Complaint | RHUN | Gas1 kifogyott, nincs másik | Gas1 kifogyott, és nincs tartalék palack | Gas1 kiszállítása és telepítése, 2 pontos kalibrációk, esedékes tHb kalibráció elvégzése, QC mérés | SWO-285820 | Re classified RMED | Logistical upSWO |  |  |  | GAS1 expended, no other | GAS1 expended and no spare bottle | GAS1 delivery and installation, 2-point calibrations, completion of the tHb calibration, QC Measurement
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 传感器电路错误 | 更换测试卡后修复 | SWO-285883 | [5 empty] | Test card, reagent package failure | Sensor circuit error | Repair after replacing test card
Yes | Visit | Complaint | RIND | Lactate parameter not working. | checked the Lactate 2point calibration results and found Lactate sensitivity error. | replaced Lactate membrane with new one and calibrated the machine now Lactate parameter working and analyzer also working fine. | SWO-288789 | Re classified RMED | Mail from Prakatheesh : 9/1-2018 | Regular replacement of membrane. | Lac | Sensitivity | Lactate parameter not working. | checked the Lactate 2point calibration results and found Lactate sensitivity error. | replaced Lactate membrane with new one and calibrated the machine now Lactate parameter working and analyzer also working fine.
No | 24 | Y | RPOL | Blank | Sprawdzono dzia³anie aparatu, testy analizatora. Aparat sprawny. |  | D3AD7A5E684A7964 | No fault found | No fault found | No fault found | No fault found |  | Blank | The camera was checked, the analyzer tests. The camera is functional.
No | Visit | Complaint | RAME | Pump Cal error | Error 1342: leak detected Error 1351: Clot suspected in SC Error 1352: Clot suspected in OXI module | Found fluid by pinch valuve tubing and replaced. Replaced all tubing in anlayzer as well as inlet gasket. Provided cusotmer with ABL 90 TS guide for training new employees. | SWO-286214 | Defect | HW | Oxi module | Pinch valve tube | 1342 | Pump Cal error | Error 1342: leak detected Error 1351: Clot suspected in SC Error 1352: Clot suspected in OXI module | Found fluid by pinch valuve tubing and replaced. Replaced all tubing in anlayzer as well as inlet gasket. Provided cusotmer with ABL 90 TS guide for training new employees.
Yes | 6 | Y | RPAC | Blank | Faulty sensor cassette. Error code 0589 pCO2 QC failure. |  | D3ADABC168058414 | [5 empty] | Blank | Faulty sensor cassette. Error code 0589 pCO2 QC failure.
No | Visit | Complaint | RCHN | 屏幕蓝屏，cthb质控不过 | 屏幕蓝屏，cthb质控不过，需要借件902-980和910-342各一个 |  | SWO-285896 | Defect | HW | Oxi module | Hemolyzer |  | Screen blue screen, cthb qc but not | Screen blue screen, cthb QC, need to borrow 902-980 and 910-342 each one
No | Visit | Complaint | RCHN | 检测到泄漏 1342 | 检测到泄漏 1342,借903-129一个。 |  | SWO-285996 | Defect | HW | Oxi module | Pinch Valve | 1342 | Leak detected 1342 | Leak 1342 was detected, borrowing 903-129 one.
No | Visit | Complaint | RESP | Equipo entra en parada por que no puede bombear liquidos | No puede bombear liquidos | Limpieza del circuito neumático Chequeo por servicio Cal, qc | SWO-286082 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway | Clot | Equipment goes into stop because it can't pump liquids | Can't pump liquids | Cleaning the pneumatic circuit Check for service Cal, qc
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 试剂包上机失败 1343 | 更换试剂包后修复 | SWO-286356 | [5 empty] | Test card, reagent package failure | Reagent Package Failed to Board 1343 | Repair after replacement of reagent kit
No | Visit | Complaint | RCHN | 报错5784 | 报错5784 | 取出仪器内断裂的测试卡后修复 by jesson | SWO-285890 | Defect | Consumables | Test Kit | Cartridge | Broken | 5784 | 5784 | Repair by jesson after removing the faulty test card in the instrument
No | 12 | Y | RPAC | Blank | Rinse Errors all teh time. |  | D3ADA2BD6858A4D2 | Defect | HW | pH/BG Module | Valve | Membrane | Blank | Rinse Errors all teh time.
No | Visit | Complaint | RIND | Qc problem for ph and po2 | qc problem for ph and p02 | checked the analyser cleaned the p and po2 membrane after that cleaned the gaskets and fill routine was done then calibartion was done after that tested all level qc its passed for all qc | SWO-289025 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | Qc problem for ph and po2 | qc problem for ph and p02 | checked the analyser cleaned the p and po2 membrane after that cleaned the gaskets and fill routine was done then calibartion was done after that tested all level qc its passed for all qc
No | 5 | Y | RCZE | Blank | zpomalená reakce dotykového displaye |  | D3AD95CC68432CB9 | Defect | SW | RMED SW | Slow Response | LEH Mail, caused by Touch Screen FW | Blank | Slow reaction of touch display
No | 4 | Y | RCZE | Blank | zpomalená odezva displaye |  | D3AD9799681D2CBC | Defect | SW | RMED SW | Slow Response | LEH Mail, caused by Touch Screen FW | Blank | Slow-motion response display
No | Visit | Complaint | RDEU | 02.01./Frau Mengesdorf, Tel.: 062150335589: OXI Probleme - Neon Lamp intensity too low | Analysator zeigt unter Fehlermeldung folgende Meldungen: Neon intensität zu niedrig und eine weitere Meldung bezogen auf die Neon Lampe. Überprüfung der Lampenwerte ergibt im Service unter Neon Intensität eine zu niedrige Intensität und die Neon Position ist außer Grenzen, negativ gemessen. Keine neue Neon Lampe auf Lager. | Tausch des Oximetrie Moduls. Nach Neustart wird Inkonsitente Software angezeigt. Einspielung der aktuellen Software in das Modul auf die aktuelle Software Version. Modul einkalibiert, Werte in Ordnung, Qualitätskontrolle auch. Garantiereparatur. | SWO-286190 | Defect | HW | Oxi module | Lamp unit | Neon | 02.01./Mrs. Mengesdorf, Tel.: 062150335589: OXI Problems - Neon Lamp intensity too low | Analyzer displays the following messages under Error message: Neon intensity too low and another message related to the neon lamp. Checking the lamp values results in too low intensity in the service under neon intensity and the neon position is out of bounds, measured negatively. No new neon lamp in stock. | Replacement of the oximetry module. After rebooting, inconsites are displayed. Recording of the current software in the module to the current software version. module calibrated, values in order, quality control too. Warranty repair.
No | 2 | Y | RCZE | Blank |  |  | D3AD87F0684C293D | Defect | HW | Aspiration Module | Bend needle |  | Blank
No | Visit | Complaint | RAME | Leak current detected, error 964. | Checked data logs and verified the problem. | Replaced the inlet gasket and the inlet connector gasket from a pm kit. | SWO-288340 | Assisted Replacement | SP | SP | Leak current |  | Leak current detected, error 964. | Checked data logs and verified the problem. | Replaced the inlet gasket and the inlet connector gasket from a pm kit.
No | Visit | Complaint | RDEU | Schwarzer Display bzw. stromlos 02 .01 12:25 Frau lüchow | System startet nicht | Fehleranalyse, Netzteil Check i. O. , Display Check i. O. , defekte CPU Einheit, neu bestellt , Station informiert | SWO-286518 | Revisit required | HW | DMS_PC | CPU |  | Black display or powerless 02 .01 12:25 Mrs lüchow | System does not start | Error analysis, power supply check i. O. , Display Check i. O. , defective CPU unit, newly ordered , station informed
No | Visit | Complaint | RDEU | 02.01.2018 Frau Daniels 02418035092 meldet, dass vermulich Luft während der Messung gezogen wird. | LS Pack getauscht, tHb Kalibration durchgeführt, Hemolysator ersetzt, tHb Kal. | Startphase QC und Kalibration | SWO-288211 | Defect | HW | Oxi module | Hemolyzer |  | 02.01.2018 Ms. Daniels 02418035092 reports that muddy air is drawn during the measurement. | LS Pack swapped, tHb calibration performed, Hemolysator replaced, tHb Kal. | Start phase QC and calibration
No | Visit | Complaint | RAME | Unable to pump solution. | Traced problem to a defective solution pack. | Replace solution pack. | SWO-287947 | Assisted Replacement | SP | SP | FT |  | Unable to pump solution. | Traced problem to a defective solution pack. | Replace solution pack.
No | Visit | Complaint | RKOR | There was happend sample error often | Cleaning inlet gasket | Cleaning inlet gasket | SWO-289427 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | There was happend sample error often | Cleaning inlet gasket | Cleaning inlet gasket
No | Visit | Complaint | RCHN | 感应板数值不稳定 | 感应板数值不稳定，需要借件902-972一个 | 更换902-972后修复 by jesson | SWO-285920 | Defect | HW | Basic Unit | PCB, Sensor connector assembly |  | Sensing board value instability | The sensing plate value is not stable, and a loan 902-972 is required | Repair by jesson after replacement 902-972
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 试剂包漏夜 | 更换试剂包后修复 | SWO-288946 | Assisted Replacement | SP | SP | Leak current |  | Test card, reagent package failure | Reagent bag leak night | Repair after replacement of reagent kit
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 压力测试真空错误，1063 | 更换试剂包后修复 | SWO-288025 | [5 empty] | Test card, reagent package failure | Pressure test vacuum error, 1063 | Repair after replacement of reagent kit
Yes | Visit | Complaint | RCHN | 卡盘 | 卡盘 | 更换测试卡后修复 | SWO-286387 | [5 empty] | Chuck | Chuck | Repair after replacing test card
No | Visit | Complaint | RCHN | 报错1024 | 报错1024，借903-364一个 |  | SWO-286390 | Defect | HW | Process Module | Optical Unit | 1024 | 1024 error | 1024, loan 903-364 one
No | Visit | Complaint | RCHN | 报错5784 | 报错5784，借903-215一个。 |  | SWO-286481 | Defect | HW | Incubator module | Incubator wheel | 5784 | 5784 | 5784, borrowed 903-215 one.
No | 6 | Y | RCZE | Blank | problémy s PC unit viz poslední návštìva. |  | D3AEEC36DE577E78 | Defect | SW | Windows | Win 7 reinstall |  | Blank | Problems with the PC unit see last visit.
No | Visit | Complaint | RIND | Ph calibration error. | Found Ph sensitivity is down. A Blood clot is stucked in chamber of ph. | Clean ph measuring chamber, run protein removal & 2 point calibration. | SWO-289330 | Cleaned | Fluidic pathway | Fluidic pathway | Measuring chamber |  | Ph calibration error. | Found Ph sensitivity is down. A Blood clot is stucked in chamber of ph. | Clean ph measuring chamber, run protein removal & 2 point calibration.
No | Visit | Complaint | RIND | Service Visit | When place the sample tube analyzer gives the message remove sample tube. | Shutdown the analyzer, open front panel as well as back panel, remove the dust from the IR sensor which is sense the sample tube. Now connect all the parts same as before and close the panel. Start the analyzer after start up its come in ready position, insert the sample tube by user, successfully done. Analyzer handed over to staff in well working condition. Take the counter also. | SWO-289085 | Cleaned | HW | Aspiration Module | Sensor, Vacutainer |  | Service Visit | When place the sample tube analyzer gives the message remove sample tube. | Shutdown the analyzer, open front panel as well as back panel, remove the dust from the IR sensor which is sense the sample tube. Now connect all the parts same as before and close the panel. Start the analyzer after start up its come in ready position, insert the sample tube by user, successfully done. Analyzer handed over to staff in well working condition. Take the counter also.
No | Visit | Complaint | RAME | Analyzer Will Not Boot To Main Screen | On power up the CPU does not load windows or application software. Screen locks up and is inoperable within 10 seconds of powering on. | Ordering CPU and CF card. | SWO-288359 | Revisit required | HW | DMS_PC | CPU |  | Analyzer Will Not Boot To Main Screen | On power up the CPU does not load windows or application software. Screen locks up and is inoperable within 10 seconds of powering on. | Ordering CPU and CF card.
No | Visit | Complaint | RDEU | Schwarzer Display bzw. stromlos 02 .01 12:25 Frau lüchow | System Überprüfung CPU Einheit defekt | Neue CPU - Unit installiert, Kalibrierung und Qualitätskontrolle i. O. | SWO-288315 | Defect | HW | DMS_PC | CPU |  | Black display or powerless 02 .01 12:25 Mrs lüchow | System verification CPU unit defective | New CPU unit installed, calibration and quality control in the
No | Visit | Complaint | RAME | po2 Glu Discrepancies | PO2 questioned on patient sample but GLU showed a value with no question marks. There was an error code 1387: GLU not usable on bottom of printout. Doctor saw error message on bottom of result and didn't report it GLU. | Needed to go on site to collect files for CQC case. | SWO-288721 | Defect | SW | RMED SW | GLU depend pO2 | 1387 (Techline?) | po2 Glu Discrepancies | PO2 questioned on patient sample but GLU showed a value with no question marks. There was an error code 1387: GLU not usable on bottom of printout. Doctor saw error message on bottom of result and didn't report it GLU. | Needed to go on site to collect files for CQC case.
No | Visit | Complaint | RAME | Error 2016 Possible fluid pathway blockage | Possible Fluid Blockage. Found analysis disabled upon arriving. Opened analyzer and found small tubing on manifold blown off as well as the waste tubing lines clotted. The tubing is in need of replacement. | Replaced all tubings and flushed hemolyzer with hypochlorite solution. Analyzer is operating fine upon restarting. Left customer a SC as they had 0 in stock. | SWO-288352 | [5 empty] | Error 2016 Possible fluid pathway blockage | Possible Fluid Blockage. Found analysis disabled upon arriving. Opened analyzer and found small tubing on manifold blown off as well as the waste tubing lines clotted. The tubing is in need of replacement. | Replaced all tubings and flushed hemolyzer with hypochlorite solution. Analyzer is operating fine upon restarting. Left customer a SC as they had 0 in stock.
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | lac定标不过，和试剂包漏电流 | 连续更换了一个测试卡和一个试剂包后修复 | SWO-288167 | [5 empty] | Test card, reagent package failure | lac calibration but, and reagent packet leakage current | A test card and a reagent kit have been replaced continuously after repair
No | Visit | Complaint | RIND | Service Visit | Continuously Na values coming in lower side. | Remove the sensor cassette from the analyzer, and flush with the distilled water after that sensor cassette flush with the Na heparin, and keep the sensor cassette in 5 to 15 degree Celsius and after that re-initiate the sensor cassette. User do the sample, Now user satisfied with the results. Analyzer handed over to staff in well working condition. | SWO-289175 | [5 empty] | Service Visit | Continuously Na values coming in lower side. | Remove the sensor cassette from the analyzer, and flush with the distilled water after that sensor cassette flush with the Na heparin, and keep the sensor cassette in 5 to 15 degree Celsius and after that re-initiate the sensor cassette. User do the sample, Now user satisfied with the results. Analyzer handed over to staff in well working condition.
No | Visit | Complaint | RDEU | 03.01.2018 14:30 Schwester Johanna Tel.03646197910 LP läuft ständig aus | Gerät läuft sporadisch aus pO2 mit QC Fehler Level 1 | Gerät gereinigt Überprüfung der LP Verschlussklappe - ok Austausch des Abfallstutzens Abfallanschluss des LP - ok Sensorkassette neu initialisiert Kal- und QC Kein Austritt von Flüssigkeit nach Abschluss aller Maßnahmen Mit Schwester Jana den Reklamationsprozess besprochen Schwester Jana wünscht eine Kontaktaufnahme durch Herrn Zedschak zum besprechen diverser Themen | SWO-288279 | [5 empty] | 03.01.2018 14:30 Sister Johanna Tel.03646197910 LP is running out all the time | Device runs out sporadically pO2 with QC Error Level 1 | Device cleaned Check of the LP shutter - ok Replacement of the waste nozzle Waste connection of the LP - ok Sensor cartridge reinitialized Kal- and QC No leakage of liquid after completion of all measures Discussed the complaint process with Sister Jana Sister Jana wishes mr. Zedschak to contact me to discuss various topics
No | Visit | Complaint | RCHN | 无法抽吸溶液，无法读取或写入试剂包芯片信息 | 无法抽吸溶液，无法读取或写入试剂包芯片信息，借903-158一个 | 更换903-158备件后修复 by jesson | SWO-286388 | Defect | HW | Basic Unit | PCB, Solution Pack ID |  | Unable to pump solution, unable to read or write reagent package chip information | Unable to pump solution, unable to read or write reagent package chip information, borrow 903-158 a | Repair by jesson after replacement of 903-158 spare parts
No | Visit | Complaint | RCHN | 机器无法启动 | 机器无法启动,需要借件903-129一个 |  | SWO-288024 | Defect | HW | Oxi module | Pinch Valve |  | The machine does not start | The machine does not start and requires a loan of 903-129
Yes | Visit | Complaint | RCHN | 报错5684 | 报错5684，需要借件 944-394（13708） | 重新定标后修复 by jesson | SWO-285925 | [5 empty] | 5684 error | 5684, need to borrow 944-394 (13708) | Fix by jesson after re-calibration
No | Visit | Complaint | RCHN | 报错1024 | 报错1024，需要借件944-214（13673） |  | SWO-288038 | Defect | HW | Process Module | Optical Unit | 1024 | 1024 error | 1024, need to borrow 944-214 (13673)
No | Visit | Complaint | RCHN | 仪器报1024，更换了一个光单依然报1024 | 仪器定标报1024, | 更换光单解决问题。 | SWO-288294 | Defect | HW | Process Module | Optical Unit | 1024 | Instrument reported 1024, replaced a light sheet still reported 1024 | Instrument calibration 1024, | Replacing the light sheet solves the problem.
Yes | 11 | Y | RPAC | Blank | FAN915-360R2 Software upgrade to v6.17 |  | D3B001A55400D3AA | [5 empty] | Blank | FAN915-360R2 Software upgrade to v6.17
No | 4 | Y | RCZE | Blank |  |  | D3B03148544F9D82 | Defect | HW | DMS_PC | CPU |  | Blank
No | Visit | Complaint | RCHN | 重启错误 | 重启错误，需要借件902-832和903-652各一个 |  | SWO-288217 | Defect | HW | Aspiration Module | Bend needle |  | Restart error | Restart error, need to borrow 902-832 and 903-652 each one
No | 2 | Y | RPOL | Blank | Wymiana kasety sensorowej. Prze³adowanie danych kalibracji i zdarzeñ z przepe³nionej bazy danych. Aparat sprawny. |  | D3AFF87D544B771B | [5 empty] | Blank | Replacing the sensor cassette. To reload the calibration and event data from the full database. The camera is functional.
No | Visit | Complaint | RAME | Leak current detected, error 964. | Checked the data logs and verified the problem. Error 964 refers to a internal leak in the solution pack. This leak can cause different error messages. | Replaced the solution pack with a new pack. | SWO-289183 | Assisted Replacement | SP | SP | Leak current | 964 | Leak current detected, error 964. | Checked the data logs and verified the problem. Error 964 refers to a internal leak in the solution pack. This leak can cause different error messages. | Replaced the solution pack with a new pack.
No | Visit | Complaint | RDEU | 04.01.2017 Frau Stelzer 02088813188 meldet einen QC sO2 Fehler. | sO2 und HHb bei QC2 ausserhalb der Grenzen, 2XLösungspacks def. Pumpenschlauch fehlt, | Reinigung durchgeführz, tHb-Kal. durchgeführt, QC durchgeführt, Cred.Claim Form ausgefüllt, | SWO-289091 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | 04.01.2017 Ms. Stelzer 02088813188 reports a QC sO2 error. | sO2 and HHb at QC2 outside the limits, 2X solution packs def. Pump hose is missing, | cleaning, tHb-Kal. performed, QC performed, Cred.Claim form completed,
Yes | Visit | Complaint | RFRA | Erreur 1177 | Test en mode service du manifold | Démontage et nettoyage du capteur. Remise en route | SWO-289277 | [5 empty] | Error 1177 | Testing in service mode of the manifold | Disassembly and cleaning of the sensor. Restarted
No | Visit | Complaint | RBNL | Defecte batterij module | Nieuwe batterij module heb lot nr. P0146N039 geïnstalleerd, na opstarten meteen de foutmelding code 1559: Accufout gedetecteerd. | Batterij module geeft error code 1559. Een andere batterij heb geïnstalleerd. | SWO-289462 | Defect | HW | Battery module | Battery |  | Faulty battery module | New battery module have lot nr. P0146N039 installed, after booting immediately the error message code 1559: Battery error detected. | Battery module displays error code 1559. Have a different battery installed.
No | Visit | Complaint | RAME | Analyzer intermittently make a loud noise | Troubleshot to a faulty computer unit cooling fan | Replaced the computer unit cooling fan. | SWO-289245 | [5 empty] | Analyzer intermittently make a loud noise | Troubleshot to a faulty computer unit cooling fan | Replaced the computer unit cooling fan.
No | Visit | Complaint | RCHN | 打印机无法打印 | 打印机无法打印，需要借件903-001和902-985各一个 | 更换903-001和902-985后修复 by jesson | SWO-288220 | [5 empty] | The printer can't print | The printer cannot print, requires a loan of 903-001 and 902-985 each | Repair by jesson after replacing 903-001 and 902-985
No | Visit | Complaint | RAME | Patient samples has to run twice before a successful derived values are obtained. | When samples are run the Oximetry derived parameters shows N $D. Samples would be successful when ran the 2nd time. Sometimes there is no parameters reported for oxi values. Customer run thb calibration and still the same issue. | Perform cleaning procedures and zero cassette function in diagnostics. Replaced hemolyzer and perform thb calibration. Verified all setup parameters. | SWO-288651 | [5 empty] | Patient samples has to run twice before a successful derived values are obtained. | When samples are run the Oximetry derived parameters shows N /D. Samples would be successful when ran the 2nd time. Sometimes there is no parameters reported for oxi values. Customer run thb calibration and still the same issue. | Perform cleaning procedures and zero cassette function in diagnostics. Replaced hemolyzer and perform thb calibration. Verified all setup parameters.
No | Visit | Complaint | RAME | Oximetry Derived values fails when samples are run the first time. | Oximetry Derived values fails when samples are run the first time with error N $D. It passes when ran the second time. The customer ran THB calibrations and still is the same issue. | Performed zero cassette and cleaning procedures. Verified setup parameters. Checked hemolyzer. | SWO-288679 | [5 empty] | Oximetry Derived values fails when samples are run the first time. | Oximetry Derived values fails when samples are run the first time with error N /D. It passes when ran the second time. The customer ran THB calibrations and still is the same issue. | Performed zero cassette and cleaning procedures. Verified setup parameters. Checked hemolyzer.
No | Visit | Complaint | RDEU | 03.01.2018 08:30 Frau Hoft Tel. 015228505103 Patientenwerte bei pO2 sind unplausibel 05.01.2018 11:20 Frau Lenz 03518385693 wünscht einen Vorortservice. | pO2 Ergebnisse zu niedrig gegenüber Patientenzustand, Messungen in Datenbank alle fehlerfrei | mit Anwender über Präanalytische Fehler gesprochen, verwendet 100µl Kapillaren ohne Adapter, Anwender wird 125µl Kapillaren mit Adaptern verwenden, beobachtet Geräteverhalten und meldet sich bei Testmessungen wurden Luftblasen erkannt ->kein Praxisbetrieb | SWO-289041 | [5 empty] | 03.01.2018 08:30 Mrs Hoft Tel. 015228505103 Patient values at pO2 are implausible 05.01.2018 11:20 Mrs Lenz 03518385693 would like a suburban service. | pO2 results too low compared to patient status, measurements in database all error-free | Talked to users about preanalytical errors, uses 100 l capillaries without adapter, user will use 125 l capillaries with adapters, observes device behavior and reports in test measurements, air bubbles were detected ->no practical operation
No | Visit | Complaint | RIND | Printer is not working. | Check & found printer gears are dismental or not in proper position. | Repair the printer arrange all gears. now printer is working. | SWO-289326 | Defect | HW | Printer Module | Printer Head |  | Printer is not working. | Check & found printer gears are dismental or not in proper position. | Repair the printer arrange all gears. now printer is working.
No | Visit | Complaint | RAME | Nipple under inlet broken off | Inlet nipple broken | Replaced inlet nipple; replaced inlet gasket, leaking out the inlet occurred, cleaning failed analyzer went into 18 minute initialization and cleaning; | SWO-289385 | Defect | HW | Inlet module | Nipple kit |  | Nipple under inlet broken off | Inlet nipple broken | Replaced inlet nipple; replaced inlet gasket, leaking out the inlet occurred, cleaning failed analyzer went into 18 minute initialization and cleaning;
No | Visit | Complaint | RIND | Touch screen not working properly and Capillary sample demo. | NA | Cleaned the touch screen, given demo for capillary sampling found instrument working fine. | SWO-294038 | Cleaned | HW | DMS_PC | Touch screen |  | Touch screen not working properly and Capillary sample demo. | In | Cleaned the touch screen, given demo for capillary sampling found instrument working fine.
No | Visit | Complaint | RIND | gasket leaking | gasket leaking | its observed gasket was blocked cleaned with hypochlorine solution and needle was slightly bent made correct after that probe initialization was done and calibrated to raedy mode | SWO-289294 | Adjustment | HW | Inlet module | Probe bend |  | gasket leaking | gasket leaking | its observed gasket was blocked cleaned with hypochlorine solution and needle was slightly bent made correct after that probe initialization was done and calibrated to raedy mode
No | Visit | Complaint | RAME | Oxi Parameters Failing on Level 3 of VK Kit | Reviewed all results from linearity and found only the level 3 oxi results to be off. | Adding the spectrometer which was not reported on the last service visit 12 $12 $17 SWO-282818 | SWO-289084 | Defect | HW | Oxi module | Range + |  | Oxi Parameters Failing on Level 3 of VK Kit | Reviewed all results from linearity and found only the level 3 oxi results to be off. | Adding the spectrometer which was not reported on the last service visit 12 $12 $17 SWO-282818
No | Visit | Complaint | RIND | Leakage from Ph bg module | Checked measuring chamber and found that there was leakage from Ref electrode chamber. | Replaced measuring chamber and checked. There was no leakage.Unit is working fine. | SWO-289082 | Defect | HW | pH/BG Module | Measuring chamber |  | Leakage from Ph bg module | Checked measuring chamber and found that there was leakage from Ref electrode chamber. | Replaced measuring chamber and checked. There was no leakage. Unit is working fine.
No | Visit | Complaint | RIND | Calibration sensitivity out of range | Checked the data logs found pco2 calibration sensitivity out of range for new membrane. Cross checked new electrode found working fine | Need to replace pco2 electrode | SWO-289040 | Assisted Replacement | Electrode | Electrode | pCO2 | Mayby revisit required | Calibration sensitivity out of range | Checked the data logs found pco2 calibration sensitivity out of range for new membrane. Cross checked new electrode found working fine | Need to replace pco2 electrode
No | Visit | Complaint | RFRA | Aiguille tordue | Aucune | Remplacement de l'aiguille et reset du compteur. | SWO-289278 | Defect | HW | Aspiration Module | Bend needle |  | Twisted needle | No | Replacing the needle and reset the meter.
No | Visit | Complaint | RKOR | tHb calibration failed. | tHb calibration 이 계속 Fail됨, 확인 해보니 SC 입구를 여닫는 Moter 부분의 마모가 되어 제대로 open 되지 않음. | 1.Spectrometer 교체. 2.여닫는 모터 부분의 모듈 교체.(오래된 모듄이었으나 해당 파트가 없어 대체함. 3.Calibration 과 QC Pass 모두 확인 되었으나, SC입구가 열리지 않음. 4.추가적인 교체 작업 필요. | SWO-289404 | Defect | HW | Oxi module | Spectrometer |  | tHb calibration failed. | tHb calibration continues to fail, The Moter part that closes the SC entrance is worn and does not open properly. | 1.Spectrometer replacement. 2. Replacing the module of the closing motor part. (It was old, but the part is not there, so it is replaced. 3.Both Calibration and QC Pass have been confirmed, but the SC entrance is not open. 4.Additional replacement work is required.
No | 54 | Y | RPNZ | Blank | Solution pack replacement due 1177: flow selector error detected $ 978: flow selector calibration error. |  | D3B80F811863E757 | Assisted Replacement | SP | SP | Flow Selector | 1177 | Blank | Solution pack replacement due 1177: flow selector error detected / 978: flow selector calibration error.
No | Visit | Complaint | RSAF | Intervention required and no details from caller(Sister incharge: Sr Pillay). | No fault found on arrival. Fault occured previous night due to blocked inlet gasket. Lab staff cleaned inlet gasket and test again done | 2pt cal and qc ok | SWO-294027 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | Intervention required and no details from caller(Sister incharge: Sr Pillay). | No fault found on arrival. Fault occured previous night due to blocked inlet gasket. Lab staff cleaned inlet gasket and test again done | 2pt cal and qc ok
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | LAC定标失败，0089 | 更换测试卡后修复 | SWO-289447 | Assisted Replacement | SC | SC | Lac | 0089 | Test card, reagent package failure | LAC calibration failure, 0089 | Repair after replacing test card
No | Visit | Complaint | RCHN | 报错1235 | 报错1235，借905-908一个 | 更换备用905-908后修复 by jesson | SWO-289313 | Defect | HW | Process Module | Sensor, BNP | 1235 | 1235 error | 1235, loan 905-908 one | Repair by jesson after replacing spare 905-908
No | Visit | Complaint | RKOR | Wetsection control PCB issue | Wetsection PCB 의 defective가 확인 됨 | PCB 교체 | SWO-293722 | Defect | HW | Basic Unit | PCB, Wet section Control |  | Wetsection control PCB issue | Defective of Wetsection PCB confirmed | PCB replacement
No | Visit | Complaint | RAME | Leak current detected, error 964. | Checked the data logs and verified the problem. | After studying the data logs I found that the trouble started after a sensor cassette had been changed. The customer did not have another sensor cassette so one was borrowed from a nearby hospital. After replacing the sensor cassette I changed the time to just before midnight and let the Analyzer do 3 system test. There were no failures on the system test. Also replaced the pinch tube as a precaution. | SWO-290020 | Assisted Replacement | SC | SC | Leak current |  | Leak current detected, error 964. | Checked the data logs and verified the problem. | After studying the data logs I found that the trouble started after a sensor cassette had been changed. The customer did not have another sensor cassette so one was borrowed from a nearby hospital. After replacing the sensor cassette I changed the time to just before midnight and let the Analyzer do 3 system test. There were no failures on the system test. Also replaced the pinch tube as a precaution.
No | Visit | Complaint | RAME | Analyzer Will Not Boot To Main Screen | CPU not working. | Replaced the CPU and FC. Replaced SC and SP - ran calibrations. | SWO-289981 | Defect | HW | DMS_PC | CPU |  | Analyzer Will Not Boot To Main Screen | CPU not working. | Replaced the CPU and FC. Replaced SC and SP - ran calibrations.
No | Visit | Complaint | RDAN | ABL90 hænger i udskiftning af væske pakke. Uanset hvad vi prøver går den ikke videre i billedet når inlet er løftet op. Det hjælper heller ikke at fjerne den manuelt da den ikke registrerer den er væk (væskepakken) | NA | Fejlen gik væk efter en time, da Dorte åbnede inlet | SWO-290424 | Defect | SW | RMED SW | SW Paused |  | ABL90 hangs in the replacement of liquid package. Whatever we try, it does not go further in the picture when inlet is lifted up. It also does not help to remove it manually as it does not detect it is gone (the liquid package) | In | The error went away after an hour when Dorte opened the inlet
No | Visit | Complaint | RFRA | Erreur 1079 Erreur impédance du capteur. | Test service : impédance K+ = 100 | Vérification contact des aiguilles de connexion K7 ok Test lecture impédance. K+ < 50. Ok Ça libration 2 points ok Passage Cq ok | SWO-291666 | Assisted Replacement | SC | SC | K+ |  | Error 1079 Sensor impedance error. | Test service: impedance K - 100 | Checking contact of K7 connection needles ok Test reading impedance. K-50. Ok It libration 2 points ok Passage Cq ok
Yes | Visit | Complaint | RDEU | 13.11.2017, 08:45 Uhr, H. Milzarek, Tel. 022029380: Bildschirm schwarz 13.11.2017, 11:00Uhr, H. Milzarek, Tel. 022029380: Bildschirm ist wieder schwarz (Lüftergeräusch zu hören) 08.01.2018 8:00 Sr. Anna 040790202392: Bildschirm erneut schwarz | System überprüft => Netzteil defekt , aus Rücknahme Gerät Netzteil ausgebaut | Netzteil aus Rücknahme Gerät eingebaut, Kalibration und QC durchgeführt | SWO-289586 | [5 empty] | 13.11.2017, 08:45, H. Milzarek, Tel. 022029380: Screen black 13.11.2017, 11:00 a.m., H. Milzarek, tel. 022029380: screen is black again (fan noise can be heard) 08.01.2018 8:00 Sr. Anna 040790202392: Screen again black | System checked => Power supply defective , removed from return device power supply | Power supply from return device installed, calibration and QC performed
No | Visit | Complaint | RIND | sc failed to calibrate | Sc failed to calibrate, | its observed that Sc was not fixed properly and cSC connector was slightly cleaned and refixed it and calibrated to raedy mode | SWO-292640 | [5 empty] | sc failed to calibrate | Sc failed to calibrate, | its observed that Sc was not fixed properly and cSC connector was slightly cleaned and refixed it and calibrated to raedy mode
Yes | Visit | Complaint | RAME | Keeps disconnecting from the network | The customer contacted me stating the analyzers still disconnecting from the network. | Cynthia has connected the analyzer in different cath lab rooms where the current analyzers were not presenting any connection problem but, this analyzer keeps disconnecting no matter what network connection is attached to. Upon arrival, I found the analyzer connected. Cynthia states the problem began after the new software was installed. I tried copying the settings from another analyzer and, reloading the settings to this analyzer. However, I noticed that the users were not transferring over. I then replaced a new flash card. I upgraded the software to v3.16 and, reloaded all settings. During the time I stayed by the analyzer, I did not see the instrument disconnecting from the network. The analyzer is holding a steady network connection. Cynthia mentioned to me that at times the analyzer presents a message stating the 'Solution pack is not connected'. Replaced the solution pack door hinge. Tried to reproduce the problem and, no errors were presented with the new door hinge. Note: Provided to Cynthia the saved data on a flash drive and, also gave her the flash card I removed out of the analyzer. | SWO-289903 | [5 empty] | Keeps disconnecting from the network | The customer contacted me stating the analyzers still disconnecting from the network. | Cynthia has connected the analyzer in different cath lab rooms where the current analyzers were not presenting any connection problem but, this analyzer keeps disconnecting no matter what network connection is attached to. Upon arrival, I found the analyzer connected. Cynthia states the problem began after the new software was installed. I tried copying the settings from another analyzer and, reloading the settings to this analyzer. However, I noticed that the users were not transferring over. I then replaced a new flash card. I upgraded the software to v3.16 and, reloaded all settings. During the time I stayed by the analyzer, I did not see the instrument disconnecting from the network. The analyzer is holding a steady network connection. Cynthia mentioned to me that at times the analyzer presents a message stating the 'Solution pack is not connected'. Replaced the solution pack door hinge. Tried to reproduce the problem and, no errors were presented with the new door hinge. Note: Provided to Cynthia the saved data on a flash drive and, also gave her the flash card I removed out of the analyzer.
No | Visit | Complaint | RFRA | Erreur de mesures PH et PO2 | Retour atelier | Mise en route | SWO-289844 | [5 empty] | Error of PH and PO2 measurements | Return workshop | Getting started
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | LAC定标不过 | 更换测试卡后修复 | SWO-291083 | [5 empty] | Test card, reagent package failure | LAC calibration is not | Repair after replacing test card
No | Visit | Complaint | RCHN | 测试卡下层温度超范围，试剂包智能芯片无法读取 | 测试卡下层温度超范围，试剂包智能芯片无法读取，借备件903-158，903-581各一个 |  | SWO-289531 | Defect | HW | Basic Unit | PCB, Sensor interface |  | The test card undertemperature temperature is over range, the reagent package smart chip can not be read | Test card undertemperature temperature super range, reagent package smart chip can not be read, borrowed spare parts 903-158, 903-581 each one
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | PO2定标失败 0210 | 更换测试卡正常。 | SWO-293976 | [5 empty] | Test card, reagent package failure | PO2 calibration failure 0210 | The replacement test card is OK.
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 流量选择器错误 0978 | 更换试剂包修复 | SWO-293790 | [5 empty] | Test card, reagent package failure | Traffic Selector Error 0978 | Replacement reagent kit repair
No | Visit | Complaint | RCHN | 报错1066.1065 | 报错1066.1065，借903-644一个 |  | SWO-289510 | Defect | HW | Basic Unit | PCB, Wet section Control | 1065 | Error 1066.1065 | 1066.1065, loan 903-644 a
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 流量选择器错误 0978 | 更换试剂包后修复 | SWO-293778 | [5 empty] | Test card, reagent package failure | Traffic Selector Error 0978 | Repair after replacement of reagent kit
No | Visit | Complaint | RCHN | 报错1068 | 报错1068,需要借件903-210一个 |  | SWO-289565 | Defect | HW | Process Module | PCB, Cup dryer | 1068 | 1068 error | 1068, need to borrow 903-210 one
Yes | Visit | Complaint | RCHN | 干燥前温度超范围，报错1068 | 干燥前温度超范围，报错1068，借903-210一个 |  | SWO-289544 |  |  |  |  | 1068 | Temperature over range before drying, reported error 1068 | The pre-drying temperature exceeded the range, reported error 1068, borrowed 903-210 one
Yes | Visit | Complaint | RCHN | 报错0569和0510 | 报错0569和0510，需要借件902-823和902-825各一个 | 自行更换834-222后修复by tim | SWO-289562 | [5 empty] | Error0569 and 0510 | 0569 and 0510, need to borrow 902-823 and 902-825 each one | Repair by tim after replacing 834-222 on your own
No | Visit | Complaint | RIND | Received a complaint from Ruby Hall Clinic that the print of the patient reports was found to be blurred with the parameter values not clearly visible. | Need to check for any obstructions in the printer head section. | Printer head was found to be OK, but as requested by Dr. Munshi, replaced the printer head with the one from the backup analyzer, after replacing the same the prints are found to be OK. | SWO-293665 | Defect | HW | Printer Module | Printer Head |  | Received a complaint from Ruby Hall Clinic that the print of the patient reports was found to be blurred with the parameter values not clearly visible. | Need to check for any obstructions in the printer head section. | Printer head was found to be OK, but as requested by Dr. Munshi, replaced the printer head with the one from the backup analyzer, after replacing the same the prints are found to be OK.
No | Visit | Complaint | RAME | 529,532,537 liquid sensor errors | Checked data logs for errors, checked valves and liquid sensors | Removed modules and cleaned valves, performed pump cal and liquid sensor adjust | SWO-289987 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | 529,532,537 liquid sensor errors | Checked data logs for errors, checked valves and liquid sensors | Removed modules and cleaned valves, performed pump cal and liquid sensor adjust
No | Visit | Complaint | RIND | Patients results are not reliable for Doctors & Lab staff. | Check & found that all membranes are old. showing error in calibration. | Change all membrane, run protein removal and 2 point calibration. | SWO-292994 | Assisted Replacement | Consumables | Replacemnet schedule | Replacemnet schedule |  | Patients results are not reliable for Doctors & Lab staff. | Check & found that all membranes are old. showing error in calibration. | Change all membrane, run protein removal and 2 point calibration.
No | Visit | Complaint | RDEU | 27.12.2017; Fr. Mengesdorf; Tel.: 062150335589; Oxy Hardwarefehler 05.01.2017 Fr. Mengesdorf 062150335589 OXI driften | 0258 Neon Lamp Intensity too low (bei Oximetrie-Test) tHb-Driften zu hoch bei Kalibration | Einstellungen Halogenlampe (Oximetrie) - Oximetrie Test durchgeführt (Neonlamp Intensity zu niedrig 43%) - Funktionsüberprüfung (Kalibration und QC ohne Beanstandung) - Neuteil-Bestellung bereits erfolgt ! | SWO-289642 | Revisit required | HW | Oxi module | Lamp unit | Neon | 27.12.2017; Ms. Mengesdorf; Tel.: 062150335589; Oxy Hardware Error 01/05/2017 Fr. Mengesdorf 062150335589 OXI drift | 0258 Neon Lamp Intensity too low (for oximetry test) tHb drifts too high at calibration | Settings Halogen lamp (oximetry) - Oximetry test performed (Neonlamp Intensity too low 43%) - Functional check (calibration and QC without complaint) - New part order already made !
No | Visit | Complaint | RDEU | 09.01.2018, 09:30 Uhr, H. Essebier, Tel. 03455572733: Gerät kommt nicht mehr aus Endlosschleife heraus. Bitte Auftr. in der MT holen. | Intervention erforderlich Flow Selector Fehler Druckfehler Vakuum | Beim Eintreffen Gerät messbereit Überprüfung Dateiarchiv => Überprüfung Flow Selector => Fenster verschmutzt => Reinigung Überprüfung Quetschventil => Funktion i.O. Alle Fluid Programme i.O. Startphase i.O. tHb-Kal fällig => tHb Kal durchgeführt => Status 285,55 pA $ Empf. 100,0% | SWO-290662 | Cleaned | HW | Basic Unit | Flow selector detector |  | 09.01.2018, 09:30, H. Essebier, tel. 03455572733: Device no longer comes out of endless loop. Please on the way. in the MT. | Intervention needed Flow Selector Error Printing error vacuum | Ready to measure device on arrival Checking File Archive => Checking Flow Selector => Window Dirty => Cleaning Checking squeezing valve => Function in the o. All Fluid Programs i.O. Start-up phase i.O. tHb-Kal due => tHb Kal performed => Status 285.55 pA . 100,0%
No | Visit | Complaint | RAME | OXI errors. 0581 OXI spectrum mismatch. | 0581 Spectrum mismatch error on Cal. Will not do tHb Cal foe same reason. Tried complete OXI test in service. Cal passed but tHb Cal failed with 0581 error. Hemolyzer is suspect. | Replaced Hemolyzer and performed complete alignment. | SWO-290714 | Defect | HW | Oxi module | Hemolyzer |  | OXYGEN errors. 0581 OXYGEN Spectrum mismatch. | 0581 Spectrum mismatch error on Cal. Will not do tHb Cal foe same reason. Tried complete OXI test in service. Cal passed but tHb Cal failed with 0581 error. Hemolyzer is suspect. | Replaced Hemolyzer and performed complete alignment.
No | Visit | Complaint | RDEU | 09.01.2018, 10:21 Uhr, H. Holzhausen, Tel. 023312016165: Oxy Kal Fehler | Oxicomplete Test wird nicht bestandén | Hemolyzer Unit gewechselt in Betrieb genommen SC und LP gewechselt nach mehrer Fehler im LP und nicht bestehen der Kalibrierung | SWO-293962 | Defect | HW | Oxi module | Hemolyzer |  | 09.01.2018, 10:21, H. Holzhausen, Tel. 023312016165: Oxy Kal Error | Oxicomplete test will not be passed | Hemolyzer Unit changed into operation SC and LP changed after several errors in the LP and do not pass the calibration
No | Visit | Complaint | RAME | Printer will not print | Printer not printing | Replaced printer still no working, tried all three USB ports on the top of the CPU,,, none worked, rebooted each time I tried a new port. Went into control panel and compared setting with analyzer at Brea all the setting in control panel are correct, Next step is to replace the CPU--Customer is requesting a replacement analyzer. Will contact Chuck Rall and Sales | SWO-290131 | Revisit required | HW | Printer Module | Printer Module |  | Printer will not print | Printer not printing | Replaced printer still no working, tried all three USB ports on the top of the CPU,,, none worked, rebooted each time I tried a new port. Went into control panel and compared setting with analyzer at Brea all the setting in control panel are correct, Next step is to replace the CPU--Customer is requesting a replacement analyzer. Will contact Chuck Rall and Sales
No | Visit | Complaint | RAME | Glucose will not pass level 1 on VK kit. | Glu not passing linearity | Replaced solutions ran Linearity, Emailed Pauline VK-R5 lot 494 needs to be removed from my stock, used and not in the CRM | SWO-290130 | Assisted Replacement | SC | SC | Glu | VK-R5 Lot 494 | Glucose will not pass level 1 on VK kit. | Glu not passing linearity | Replaced solutions ran Linearity, Emailed Pauline VK-R5 lot 494 needs to be removed from my stock, used and not in the CRM
No | Visit | Complaint | RSCH | Blank | Bruit excessif de l’introducteur (mouvement impossible) Instabilité de mesure | Remplacé le kit de maintenance e.o Ajustage du site d’introduction e.o Divers tests dans le programme de service e.o Calibration e.o CQ e.o | SWO-290524 | Adjustment | HW | Auto Inlet module | Auto Inlet module |  | Blank | Excessive noise from the introductor (impossible movement) Measurement instability | Replaced the e.o maintenance kit Adjustment of the introductory site e.o Various tests in the e.o service program Calibration e.o CQ e.o
No | Visit | Complaint | RSCH | PH und pcO2 QA Zyklus fehler. | PH und pcO2 QA Zyklus fehler. | SC ersetzt und QA Zyklus gestartet. I.O. | SWO-291473 | [5 empty] | PH and pcO2 QA cycle error. | PH and pcO2 QA cycle error. | SC replaced and QA cycle started. Ok.
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 温度-F，NA,GLU,LAC定标失败 | 连续更换2张测试卡后修复 | SWO-290305 | [5 empty] | Test card, reagent package failure | Temperature-F, NA, GLU, LAC calibration failure | Repair after 2 consecutive test cards
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | O2定标不过 | 更换测试卡后修复 | SWO-290166 | [5 empty] | Test card, reagent package failure | O2 calibration but | Repair after replacing test card
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 温度-F,失败-FC | 连续更换两张测试卡后修复 | SWO-291887 | [5 empty] | Test card, reagent package failure | Temperature-F, Failure-FC | Repair after replacing two test cards in a row
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 失败-SC 温度-F | 连续更换2张测试卡和试剂包后修复 | SWO-291910 | [5 empty] | Test card, reagent package failure | Failure-SC Temperature-F | Repair after 2 consecutive test cards and reagent kits
No | Visit | Complaint | RIND | Patient Data full. | Patient Data Full. | Patient Data removed. Training given for available staff to removing data and save in there PC. Analyzer Serviced and calibrated properly. Handed over in a fine working condition. | SWO-292503 | [5 empty] | Patient Data full. | Patient Data Full. | Patient Data removed. Training given for available staff to removing data and save in there PC. Analyzer Serviced and calibrated properly. Handed over in a fine working condition.
No | Visit | Complaint | RESP | Cliente informa que el equipo siempre da error en parámetros de pCO2 cuando instala el reactivo mientras el mismo en el equipo de al lado funciona correctamente. | Fallo en pCO2 cuando se camba el reactivo, atrasos en Cal, Fallo de respuesta electrodos etc.. | Se procede a realizar un kit de mantenimiento por lo que se sustituyen todos los tubos, también se sustituye Inlet y válvula de gancho. Se realizan test por servicio sin detectar ningún problema. El cliente no dispone de material por lo que la próxima vez comprobara si el equipo ha mejorado. | SWO-290581 | Defect | HW | PM Kit | PM Kit |  | Client reports that the computer always fails pCO2 parameters when it installs the reagent while the reagent on the computer next door is working properly. | Failure in pCO2 when reagent is being transferred, delays in Cal, Response failure electrodes etc. | A maintenance kit is carried out so that all tubes are replaced, Inlet and hook valve are also replaced. They are performed service-by-service tests without detecting any problems. The customer does not have any material so next time he will check if the equipment has improved.
No | Visit | Complaint | RCHN | 报错1024 | 报错1024，借903-364一个 |  | SWO-290159 | Defect | HW | Process Module | Optical Unit | 1024 | 1024 error | 1024, loan 903-364 one
No | Visit | Complaint | RCHN | BNT测试卡断裂 | BNT测试卡断裂 | 更换测试卡后修复 | SWO-290217 | Defect | Consumables | Test Kit | Cartridge | Broken | BNT test card break | BNT test card break | Repair after replacing test card
Yes | Visit | Complaint | RCHN | 报错5784 | 报错5784，需要借件903-138和636-539各一个 | 更换903-138和636-539后修复by tim | SWO-289507 | [5 empty] | 5784 | 5784, need to borrow 903-138 and 636-539 each | Repair by tim after replacing 903-138 and 636-539
No | Visit | Complaint | RFRA | Reboot de l'ABL9 après écran figé. | Echange carte CF | Pasd'amélioration | SWO-290710 | [5 empty] | ABL9 reboot after frozen screen. | Exchange card CF | Pasd'improvement
No | Visit | Complaint | RLTD | Showing error code: 05000029 | Monitor returned to site following test at Radiometer Ltd. workshop. Sensor connected, remembraned and calibrated - Sensor went into stabilising for 40 minutes. Tested re-calibration now ok. Sensor connected to myself and 10-15 minutes of study done. Verified study present in records and unable to verify that it could be downloaded to Visi download software. Customer to check with IT about network settings for this monitor. | Returned to site and verified operation. | SWO-290409 | [5 empty] | Showing error code: 05000029 | Monitor returned to site following test at Radiometer Ltd. workshop. Sensor connected, remembraned and calibrated - Sensor went into stabilising for 40 minutes. Tested re-calibration now ok. Sensor connected to myself and 10-15 minutes of study done. Verified study present in records and unable to verify that it could be downloaded to Visi download software. Customer to check with IT about network settings for this monitor. | Returned to site and verified operation.
No | Visit | Complaint | RLTD | Showing error 05000029 | Monitor returned to site following test at Radiometer Ltd. workshop. Sensor connected, remembraned and calibrated - ok. Sensor connected to myself and 10-15 minutes of study done. Verified study present in records and also verified that it could be downloaded to Visi download software. | Monitor returned to site and performance verified. | SWO-290377 | [5 empty] | Showing error 05000029 | Monitor returned to site following test at Radiometer Ltd. workshop. Sensor connected, remembraned and calibrated - ok. Sensor connected to myself and 10-15 minutes of study done. Verified study present in records and also verified that it could be downloaded to Visi download software. | Monitor returned to site and performance verified.
Yes | 8 | Y | RCZE | Blank |  |  | D3B79FA5185A9876 | Re classified RMED |  |  |  |  | Blank
No | Visit | Complaint | RDEU | Auto Check Lüfter defekt | Gerät macht Geräusche Auto Check Lüfter defekt Lac Kalibrier- und QC Fehler vorhanden | Proteinentfernung durchgeführt Auto Check Modul ausgebaut, Lüfter erneuert, Modul eingebaut unteren Dichtring an Lac-Elektrode sowie Membran erneuert Gerät kalibriert und überprüft | SWO-290543 | Defect | HW | Auto Check Module | Fan |  | Auto Check fan defective | Device makes noises Auto Check fan defective Lac calibration and QC errors present | Protein removal performed Auto Check module removed, fan renewed, module built-in lower sealing ring on lac electrode as well as membrane renewed Device calibrated and checked
No | Visit | Complaint | RDEU | 27.12.2017; Fr. Mengesdorf; Tel.: 062150335589; Oxy Hardwarefehler 05.01.2017 Fr. Mengesdorf 062150335589 OXI driften | tHb Drift-Fehler bei Kalibration Oximetrie-Test-Fehler (Neon Lamp Intensity zu tief) | Austausch Oximetrie Modul - Oximetrie-Test ohne Beanstandung - Einstellung Oximetrie - tHb KAL - Funktionsüberprüfung | SWO-290382 | Defect | HW | Oxi module | Oxi module, Replaced |  | 27.12.2017; Ms. Mengesdorf; Tel.: 062150335589; Oxy Hardware Error 01/05/2017 Fr. Mengesdorf 062150335589 OXI drift | tHb Drift Error in Calibration Oximetry Test Error (Neon Lamp Intensity Too Deep) | Replacement Oximetry Module - Oximetry Test Without Complaint - Setting Oximetry - tHb KAL - Function Check
No | 4 | Y | RPOL | Blank | sprawdzenie aparatu, czyszczenie, kalibracja tHb. Aparat sprawny technicznie. Zalecenie stosowania wy³apywaczy skrzepów. |  | D3B7CB71183BFF13 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway | Clot | Blank | Camera check, cleaning, tHb calibration. The camera is technically efficient. Recommendation for the use of blood clots.
No | Visit | Complaint | RAME | Leak Detected | Error 1184: Leak Detected | Replaced analyzer tubing and PM. Tested. Note: From: Iwankovitsch, James JRIWA Sent: Wednesday, January 10, 2018 3:13 PM To: 'Huntsinger, Tanya'; Keller, Bonnie BKELL Cc: 'Reed, Michelle'; Robert Riccardi (rriccardi@radiometeramerica.com) Subject: RE: Reading Hospital update with Leak Detected ABL 90 SP and SC XL Hi Michelle and Tanya, So you have had the same Leak detected error on 2 of the 3 analyzers and replacing the tubing resolved it both times. I am reaching out to your Radiometer Clinical Applications Specialist with the details below. If you get a Leak Detected error on ABL 2; do not flush, troubleshoot, etc. Place a service call and I will come out to work on the analyzer. I will also gather Data Logs and Service Dump and create an escalation. Thanks, Jim Details of repair for ABL 1 and ABL3 What doesn’t work is Back flushing, replacing inlet probe, inlet gasket, SP and SC. What does work is conducting a Planned Maintenance (replacing tubing) on the analyzer. ABL 90 Solution Pack – SP90 XL 980 activities, 944-457 and SC90 1200 (30 days), 946-060. The Green Top Cap sodium heparin tube goes into a BD 1ml syringe tuberculin slip tip with the Nova Stat Profile Critical Care Xpress Syringe Clot catcher. Not using Safe PICO syringe. No Fluids (pleural fluid) only whole blood. | SWO-291665 | Defect | HW | PM Kit | PM Kit | Leak | Leak Detected | Error 1184: Leak Detected | Replaced analyzer tubing and PM. Tested. Note: From: Iwankovitsch, James JRIWA Sent: Wednesday, January 10, 2018 3:13 PM To: 'Huntsinger, Tanya'; Keller, Bonnie BKELL Cc: 'Reed, Michelle'; Robert Riccardi (rriccardi@radiometeramerica.com) Subject: RE: Reading Hospital update with Leak Detected ABL 90 SP and SC XL Hi Michelle and Tanya, So you have had the same Leak detected error on 2 of the 3 analyzers and replacing the tubing resolved it both times. I am reaching out to your Radiometer Clinical Applications Specialist with the details below. If you get a Leak Detected error on ABL 2; do not flush, troubleshoot, etc. Place a service call and I will come out to work on the analyzer. I will also gather Data Logs and Service Dump and create an escalation. Thanks, Jim Details of repair for ABL 1 and ABL3 What doesn’t work is Back flushing, replacing inlet probe, inlet gasket, SP and SC. What does work is conducting a Planned Maintenance (replacing tubing) on the analyzer. ABL 90 Solution Pack – SP90 XL 980 activities, 944-457 and SC90 1200 (30 days), 946-060. The Green Top Cap sodium heparin tube goes into a BD 1ml syringe tuberculin slip tip with the Nova Stat Profile Critical Care Xpress Syringe Clot catcher. Not using Safe PICO syringe. No Fluids (pleural fluid) only whole blood.
No | 5 | Y | RPOL | Blank | sprawdzenie analizatora, prostowanie igly, czyszczenie inietu. Kalibracja parametrów. Aparat sprawny technicznie. Instrukta¿ personelu |  | D3B7C972183CFF0E | Adjustment | HW | Auto Inlet module | Probe bend |  | Blank | Analyzer, straighten needles, clean the injection. Parameter calibration. The camera is technically efficient. Staff instruction
No | Visit | Complaint | RFRA | Le site d'introduction reste bloqué en position seringue lors du passage de BG | Contrôle de l'inlet, erreur de positionnement du joint d'intro. | Inlet module cassé. Un module neuf sera envoyé par RADIOMETER. | SWO-291553 | Revisit required | HW | Auto Inlet module | Auto Inlet module | Broken | The introductory site remains stuck in the syringe position during the passage of BG | Inlet control, intro joint positioning error. | Inlet module broken. A new module will be sent by RADIOMETER.
No | Visit | Complaint | RCHN | 电源损坏 | 电源损坏，借903-663一个 | 更换公司备用机903-663修复by tim | SWO-290311 | Defect | HW | Basic Unit | Power Supply |  | Power damage | Power damage, borrow 903-663 one | Replacement company spare 903-663 repair by tim
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 液体传输错误，1285 | 更换试剂包后修复 | SWO-291867 | Assisted Replacement | SP | SP | FT | 1285 | Test card, reagent package failure | Liquid transfer error, 1285 | Repair after replacement of reagent kit
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | LAC定标不过 0375 | 更换测试卡后修复 | SWO-295865 | [5 empty] | Test card, reagent package failure | LAC calibration is no more than 0375 | Repair after replacing test card
No | Visit | Complaint | RCHN | 报错1024 | 报错1024，借903-364一个 | 更换903-364备件后修复by tim | SWO-290250 | Defect | HW | Process Module | Optical Unit | 1024 | 1024 error | 1024, loan 903-364 one | Repair by tim after replacement of 903-364 spare parts
No | 17 | Y | RPNZ | Blank | Gas module (Gas 1 keep running low) |  | D3B7BA0B1847FFF8 | Defect | HW | Gas module | Flow regulator | Left | Blank | Gas module (Gas 1 keep running low)
No | Visit | Complaint | RIND | Service Visit. | In calibration message pO2 shows zero current error and pCO2, Na, Glu and lac shows sensitivity out of range. | Keep on hold the analyzer, in the pO2 chamber scratch the inlet cover pin which is connected to the pO2 electrode. Record no replacement done, restart the analyzer wait for auto two point calibration done. Need to change the membrane of glucose, lactate, pCO2, and sodium. | SWO-293557 | Revisit required | Fluidic pathway | Fluidic pathway | Fluidic pathway | Repeated visit | Service Visit. | In calibration message pO2 shows zero current error and pCO2, Na, Glu and lac shows sensitivity out of range. | Keep on hold the analyzer, in the pO2 chamber scratch the inlet cover pin which is connected to the pO2 electrode. Record no replacement done, restart the analyzer wait for auto two point calibration done. Need to change the membrane of glucose, lactate, pCO2, and sodium.
Yes | 31 | Y | RCZE | Blank | obèas selhání vzorku |  | D3B7B7D1182C7AE4 | Re classified RMED | SWU |  |  |  | Blank | Time of sample failure
No | Visit | Complaint | RAME | QC AND CAL ERRORS ON PCO2 AND PO2 | CUSTOMER EXPERIENCED ERRORS IN 2PNT CALS ON PCO2 AND PO2. WHILE TRYING TO RECOVER FROM THIS QC’S RAN ON LEVELS 1 AND 3 THAT SHOWED ERRORS FOR THE QC. ONCE THEY RECOVERED THE DRIFT FROM 2PNT CALS (AFTER REPLACING THE REF MEMBRANE) THEY WERE UNAWARE THE QC’S HAD FAILED. | I RAN SEVERAL 2PNT CALS AS I OBSERVED ADDITIONAL DRIFT IN THE RESULTS. AFTER THE 2PNT CALS PASSED THE QC’S PASSED WITH NO PROBLEM. DURING THIS TIME THE ANALYZER SHOWED THE CAL2 GAS BOTTLE WAS LOW. REPLACED THE BOTTLE WITH CUSTOMER STOCK AND RESOLVED. | SWO-291907 | Assisted Replacement | Consumables | Gas Cylinder | Gas Cylinder | pO2/pCO2 | QC AND CAL ERRORS ON PCO2 AND PO2 | CUSTOMER EXPERIENCED ERRORS IN 2PNT CALS ON PCO2 AND PO2. WHILE TRYING TO RECOVER FROM THIS QC’S RAN ON LEVELS 1 AND 3 THAT SHOWED ERRORS FOR THE QC. ONCE THEY RECOVERED THE DRIFT FROM 2PNT CALS (AFTER REPLACING THE REF MEMBRANE) THEY WERE UNAWARE THE QC’S HAD FAILED. | I RAN SEVERAL 2PNT CALS AS I OBSERVED ADDITIONAL DRIFT IN THE RESULTS. AFTER THE 2PNT CALS PASSED THE QC’S PASSED WITH NO PROBLEM. DURING THIS TIME THE ANALYZER SHOWED THE CAL2 GAS BOTTLE WAS LOW. REPLACED THE BOTTLE WITH CUSTOMER STOCK AND RESOLVED.
No | Visit | Complaint | RIND | thb/so2 values problem | thb $so2 values problem | switched off the analyser for some times , then cleaned the gasket and probe after that pump calibration was done after that checked oxymetry values then calibrated the analyser to raedy mode, tested the samples all are ok | SWO-293409 | Cleaned | HW | Oxi module | Oxi module, Cleaned |  | thb/so2 values problem | thb /so2 values problem | switched off the analyser for some times , then cleaned the gasket and probe after that pump calibration was done after that checked oxymetry values then calibrated the analyser to raedy mode, tested the samples all are ok
No | 2 | Y | RCZE | Blank |  |  | D3B7A0A5185CB396 | Assisted Replacement | Electrode | Electrode | pCO2 |  | Blank
No | Visit | Complaint | RDEU | pH Kalibrierfehler, Empfindlichkeit | pH Kalibrierfehler, Empfindlichkeit | pH Elektrode überprüft und gereinigt, pH Dichtringe erneuert, Kal. 1 Lösung erneuert, Reinigungslösung erneuert, Mehrere Kalibrationen vorgenommen, Fehler weiterhin vorhanden, Elektrode defekt, pH Elektrode erneuert, Kalibrationen und QC's vorgenommen, BGA ok. | SWO-291615 | Assisted Replacement | Electrode | Electrode | pH | Sensitivity | pH calibration error, sensitivity | pH calibration error, sensitivity | pH electrode checked and cleaned, pH sealing rings renewed, Kal. 1 solution renewed, cleaning solution renewed, several calibrations made, errors still present, electrode defective, pH electrode renewed, calibrations and QC's made, BGA ok.
No | Visit | Complaint | RIND | Machine is not working. | Found machine is not in ready mode, Showing msg DMS not Connected to analyzing unit. | Check the analyzer, clean fuse on analyzing controller pcb & restart the analyzer. | SWO-291067 | Cleaned | HW | Basic Unit | PCB, Analyzer control | Fuse | Machine is not working. | Found machine is not in ready mode, Showing msg DMS not Connected to analyzing unit. | Check the analyzer, clean fuse on analyzing controller pcb & restart the analyzer.
No | Visit | Complaint | RAME | The analyzer is not migrating patient info. | Incorrect comport assigned for customers connection to HealthLAN. Still having issues with system pulling demographics after scanning in the sample barcode. This is an issue with HealthLAN. Ron from HealthLAN is troubleshooting and will contact customer. | Assigned correct port number 51778. | SWO-292662 | Setup | His/Lis | His/Lis | His/Lis |  | The analyzer is not migrating patient info. | Incorrect comport assigned for customers connection to HealthLAN. Still having issues with system pulling demographics after scanning in the sample barcode. This is an issue with HealthLAN. Ron from HealthLAN is troubleshooting and will contact customer. | Assigned correct port number 51778.
No | Visit | Complaint | RDEU | 10.01.17, Frau Hoyer 015127648357, Probleme mit QC | pO2-Dunkeldaten außer Bereich | pO2-Sensor im Service überprüft. zpO2DarNo -Wert lag grenzwertig hoch. Nach Reinigung des Sensor dann bei 4,3 anstatt vorher bei 4,9 ( Grenzwert bei 5,00) Gerät kalibriert und tHb-Kal durchgeführt, danach mehrere Qualitätskontrollen und Patientenmessungen durchgeführt, die alle fehlerfrei waren. | SWO-292658 | Cleaned | HW | Basic Unit | pO2 Sensor | zpO2DarNo | 10.01.17, Frau Hoyer 015127648357, mit QC issues | pO2 dark data out of range | pO2 sensor in the service. zpO2DarNo value was borderline high. After cleaning the sensor then at 4.3 instead of before at 4.9 (limit value at 5.00) calibrated and tHb-Kal performed, then performed several quality controls and patient measurements, all of which were error-free.
Yes | 10 | Y | RPAC | Blank | Failed SC |  | D3BA37B30446B9E6 | [5 empty] | Blank | Failed SC
No | Visit | Complaint | RFRA | Erreur 1177 | Aucune | Remplacement de la carte électronique du sélecteur de flux. Remplacement du pack de solution. Remise en route et essais. | SWO-293396 | Revisit required | HW | Basic Unit | PCB, Flow Selector | 1177 | Error 1177 | No | Replacing the electronic card of the flow selector. Replacing the solution pack. Restarted and tested.
No | Visit | Complaint | RAME | Keeps disconnecting from the network | Analyzer keeps disconnecting from the network. | Since my last visit the analyzer was able to hold a steady network connection but, today Cynthia called reporting the analyzer is disconnecting again. Revisit and, inspected the analyzer communications settings and, all settings are correct. I confirmed with the customer that they do not use a Rals connection. Finally, I replaced the CPU unit. I brought the analyzer to the point of care office and, connected it to the network connection available for this instrument. The analyzer connected but, during the time I ran system cycles and, finished running qc the analyzer disconnected again from the network. Cynthia and, I brought the analyzer to the Cath lab Rm 5 and, the analyzer did not present any change, keeps disconnecting. I have collected all data logs and, the database files. I will be escalating this problem with our technical support team. I have informed my service manager Don Masters about the situation with this instrument. Last, I recommended Cynthia to have the IT group check the network connections for this analyzer in the hospital. | SWO-292744 | [5 empty] | Keeps disconnecting from the network | Analyzer keeps disconnecting from the network. | Since my last visit the analyzer was able to hold a steady network connection but, today Cynthia called reporting the analyzer is disconnecting again. Revisit and, inspected the analyzer communications settings and, all settings are correct. I confirmed with the customer that they do not use a Rals connection. Finally, I replaced the CPU unit. I brought the analyzer to the point of care office and, connected it to the network connection available for this instrument. The analyzer connected but, during the time I ran system cycles and, finished running qc the analyzer disconnected again from the network. Cynthia and, I brought the analyzer to the Cath lab Rm 5 and, the analyzer did not present any change, keeps disconnecting. I have collected all data logs and, the database files. I will be escalating this problem with our technical support team. I have informed my service manager Don Masters about the situation with this instrument. Last, I recommended Cynthia to have the IT group check the network connections for this analyzer in the hospital.
No | Visit | Complaint | RFRA | Erreur de mesures PH et PO2 | Mesures échantillons | Fonctionnement ok Poursuite essais mesures | SWO-292600 | [5 empty] | Error of PH and PO2 measurements | Sample measurements | How to operate ok Continuation tests measures
No | Visit | Complaint | RCHN | pO2质控偏高 | pO2质控偏高 | 测量质控正常by tim | SWO-291889 | [5 empty] | pO2 high quality control | pO2 high quality control | Measuring quality control normal by tim
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | GLU定标不过 0378 | 更换测试卡后修复 | SWO-291883 | [5 empty] | Test card, reagent package failure | GLU calibration is no more than 0378 | Repair after replacing test card
No | Visit | Complaint | RCHN | 报错1024 | 报错1024，之前的光单换上之后出现了定标错误，需要加件944-267（12765）缺货 | 重新定标后修复by tim | SWO-291075 | Defect | HW | Process Module | Optical Unit | 1024 | 1024 error | Error 1024, the previous light sheet after the replacement of a calibration error, need to add 944-267 (12765) out of stock | Fix by tim after re-calibration
No | Visit | Complaint | RKK | ｐCO2感度範囲外(↓) | pCO2感度エラー | pCO2メンブラン交換 交換期限超過Ref $Glu $Lac交換 pHシーリングリング交換(FAN) | SWO-292029 | Assisted Replacement | Electrode | Membrane | pCO2 | Sensitivity | pCO2 sensitivity outside (↓) | pCO2 sensitivity error | pCO2 Membrane Replacement Overdue Ref $Glu $Lac Exchange pH Sealing Ring Replacement (FAN)
No | Visit | Complaint | RAME | Customer reports several errors that detail the fluidic system. Rinse, liquid sensor and conditioning errors. | Checked the fluidics of the analyzer and logs to find several errors. Errors to include 751, (rinse error) 602 ( inlet not filled) 501 ( Rinse not aspirated to PH $BG) 648(calibration failed) 652 (OXI liquid sensor error)779 (cleaning failed). Restarted the analyzer with no errors. Multiple calibrations and cleaning to receive intermittent errors stated above. | Found Air Valve intermittently sticking in the fluidic module. Replaced valve membrane and inspected valve housing. Restarted analyzer and ran several calibrations, cleaning and rinse programs to ensure proper operation. | SWO-292704 | Defect | HW | pH/BG Module | Valve, air | Membrane | Customer reports several errors that detail the fluidic system. Rinse, liquid sensor and conditioning errors. | Checked the fluidics of the analyzer and logs to find several errors. Errors to include 751, (rinse error) 602 ( inlet not filled) 501 ( Rinse not aspirated to PH /BG) 648(calibration failed) 652 (OXI liquid sensor error)779 (cleaning failed). Restarted the analyzer with no errors. Multiple calibrations and cleaning to receive intermittent errors stated above. | Found Air Valve intermittently sticking in the fluidic module. Replaced valve membrane and inspected valve housing. Restarted analyzer and ran several calibrations, cleaning and rinse programs to ensure proper operation.
No | Visit | Complaint | RIND | Service Visit | Calibration aborted again and again. | Check the analyzer, in calibration message no cal 2 present, observe the reagent all are above 25%, keep the paper on sensor of the cal 2 reagent and remove the cal 2 open the rubber, observe that the piercing of tube is not available. pierce the old cal 1 tube in to the cal 2 and connected. Now give the command of two point calibration, successfully done. analyzer handed over to staff in well working condition. | SWO-293561 | User error | Consumables | Solution Bottle | Membrane piercing |  | Service Visit | Calibration aborted again and again. | Check the analyzer, in calibration message no cal 2 present, observe the reagent all are above 25%, keep the paper on sensor of the cal 2 reagent and remove the cal 2 open the rubber, observe that the piercing of tube is not available. pierce the old cal 1 tube in to the cal 2 and connected. Now give the command of two point calibration, successfully done. analyzer handed over to staff in well working condition.
No | Visit | Complaint | RDAN | Marianne og Jan har ringet ang. denne. Man har forsøgsvis skiftet samtlige pumpeslanger samt elektroder og membraner. Det har stået på i nogen tid, specielt pCO2, hvor der skiftes membraner meget ofte. Det er aftalt at Ebbe kigger forbi torsdag den 14. december. | - | Udskiftet ventiler på MET og MET 2 modulerne, smurt ventilerne på pH $Blodgas modulet, slutafprøvet. | SWO-293113 | Defect | HW | MetII Module | Valve |  | Marianne and Jan have called Ang. This. All pump hoses, electrodes and membranes have been changed in experimental manner. It has been going on for some time, especially pCO2, where the membranes are changed very often. It is agreed that Ebbe is looking past Thursday, 14. December. | - | Replaced valves on the MET and MET 2 modules, lubricated the valves on the pH $Blodgas module, final tested.
No | Visit | Complaint | RAME | PH status and sensitivity errors. | Status and sensitivity errors. Problem trace to a failed pH electrode. Electrode in warranty. | Replace pH electrode. | SWO-292673 | Assisted Replacement | Electrode | Electrode | pH |  | PH status and sensitivity errors. | Status and sensitivity errors. Problem trace to a failed pH electrode. Electrode in warranty. | Replace pH electrode.
No | Visit | Complaint | RIND | Test cartridge could not be inserted | It was found that Plastic Cover for storage door moving cable has fallen off as it's screws were missing | Used screws from left and right panels of front cover and fixed the plastic cover on storage door moving cable. | SWO-292857 | Defect | HW | Cover and Chassis | Wire Guide |  | Test cartridge could not be inserted | It was found that Plastic Cover for storage door moving cable has fallen off as it's screws were missing | Used screws from left and right panels of front cover and fixed the plastic cover on storage door moving cable.
No | Visit | Complaint | RAME | Wireless connectionis sporadically dropping. | Susana stated on her email that sometimes she has found the wireless adapter is turn 'off' thus loosing the wireless connection for the instrument. | I checked the USB power settings and, found they were setup to turn off to save power. I went in to the windows control panel and, disabled the USB power off feature to prevent the wireless connection dongle from shutting down. This will keep the analyzer from disconnecting from the wireless network. | SWO-293576 | Setup | His/Lis | His/Lis | His/Lis |  | Wireless connectionis sporadically dropping. | Susana stated on her email that sometimes she has found the wireless adapter is turn 'off' thus loosing the wireless connection for the instrument. | I checked the USB power settings and, found they were setup to turn off to save power. I went in to the windows control panel and, disabled the USB power off feature to prevent the wireless connection dongle from shutting down. This will keep the analyzer from disconnecting from the wireless network.
No | Visit | Complaint | RAME | The wireless connection is sporadically lost. | Susana states she has found the dongle turned off on this analyzer thus the wireless connection was lost. She has removed and, reseated the dongle and, the connection was reestablished. | I checked the USB connections where the dongle connects and, went in to windows to change the 'Power saving' option that would turn off the power to the USB connector where the dongle is connected. Now the power to the dongle will not be off and, so the wireless connection will not be lost. | SWO-293578 | Revisit required | His/Lis | His/Lis | His/Lis |  | The wireless connection is sporadically lost. | Susana states she has found the dongle turned off on this analyzer thus the wireless connection was lost. She has removed and, reseated the dongle and, the connection was reestablished. | I checked the USB connections where the dongle connects and, went in to windows to change the 'Power saving' option that would turn off the power to the USB connector where the dongle is connected. Now the power to the dongle will not be off and, so the wireless connection will not be lost.
Yes | Visit | Complaint | RDEU | 12.01.18, Herr Braun 036212205371. Probleme mit QC | Herr Braun meldet einen QC Fehler bei Level 2 Er hat in den letzten Tagen schon mehrmalss den LP erneuert, gespült und die Einlassdichtung erneuert. Auch die Sensorkassette wurde nochmals neu reinstalliert. Keine Besserung. | Herr Braun wird nun noch eine komplett neue Sensorkassette installieren. Falls das Problem weiterhin besteht wird er sich wieder in der Hotline melden. | SWO-293317 | Re classified RMED | Phone support |  |  |  | 12.01.18, Herr Braun 036212205371. Probleme mit QC | Mr Braun reports a QC error at Level 2 He has renewed the LP several times in the last few days, flushed it and renewed the inlet seal. The sensor cartridge was also reinstalled. No improvement. | Mr Braun will now install a completely new sensor cartridge. If the problem persists, he will contact the hotline again.
No | Visit | Complaint | RAME | PO2 failing on external QCs. | PO2 QC Outside of range all levels. | Perform warm water back flush. Replace solution pack. Replace sensor cassette. | SWO-293529 | Assisted Replacement | SC | SC | pO2 | Ext QC | PO2 failing on external QCs. | PO2 QC Outside of range all levels. | Perform warm water back flush. Replace solution pack. Replace sensor cassette.
No | Visit | Complaint | RLTD | Analyser ok now, but keeps giving issues when solution pack is changed, Dave Gardner 0203 2992264 | Checked inlet module - gasket and probe ok and movement over probe is smooth. Checked inlet drive unit - ensured grub screws for motor clutch were tight, all ok. Checked flow selector window - residue of dust from solution packs on the window, cleaned flow selector window. Re-assembled and tested inlet module in service mode - movement to all positions is correct. Checked fluid transport in service - all ok. Restarted analyser and verified operation - all ok. | Cleaned flow selector window. Fully checked inlet module - all ok. | SWO-293386 | Cleaned | HW | Basic Unit | Flow selector detector |  | Analyser ok now, but keeps giving issues when solution pack is changed, Dave Gardner 0203 2992264 | Checked inlet module - gasket and probe ok and movement over probe is smooth. Checked inlet drive unit - ensured grub screws for motor clutch were tight, all ok. Checked flow selector window - residue of dust from solution packs on the window, cleaned flow selector window. Re-assembled and tested inlet module in service mode - movement to all positions is correct. Checked fluid transport in service - all ok. Restarted analyser and verified operation - all ok. | Cleaned flow selector window. Fully checked inlet module - all ok.
No | Visit | Complaint | RFRA | Erreur 1177 et 978 | Aucune | Mise en place de l'ABL90 du site de La fierté St Bernard en dépannage. Retour de lABL90 plus du client a l'atelier pour réparation. | SWO-293548 | Defect | HW | Basic Unit | PCB, Flow Selector | 1177 | Error 1177 and 978 | No | The ABL90 of the St Bernard Pride site is set up in troubleshooting. Return of the lABL90 plus the customer to the workshop for repair.
No | Visit | Complaint | RAME | Keeps disconnecting from the network | The analyzer keeps disconnecting from the network. | Upon investigation and, speaking with our technical support specialist supervisor Frank Rowley, I revisit this analyzer today to check the network adapter power management settings.Frank added me to uncheck the settings that 'allows the computer to turn off the adapter to save power'. This was shutting down the network connection and, causing the analyzer to be disconnected from the hospital network. Once I disabled this option, the analyzer is now keeping the network connectivity uninterrupted. At this time, opening a techline case will not be necessary as the problem has now been resolved. If the problem reoccurs we will then escalate the issue with technical support and, submit a techline case. | SWO-293577 | [5 empty] | Keeps disconnecting from the network | The analyzer keeps disconnecting from the network. | Upon investigation and, speaking with our technical support specialist supervisor Frank Rowley, I revisit this analyzer today to check the network adapter power management settings. Frank added me to uncheck the settings that 'allows the computer to turn off the adapter to save power'. This was shutting down the network connection and, causing the analyzer to be disconnected from the hospital network. Once I disabled this option, the analyzer is now keeping the network connectivity uninterrupted. At this time, opening a techline case will not be necessary as the problem has now been resolved. If the problem reoccurs we will then escalate the issue with technical support and, submit a techline case.
No | Visit | Complaint | RKK | センサーカセット初期化不良との事で訪問：機械がシャットダウンしており履歴を見るとTemp-Ｄとオーバーヒートエラーも出ていた。センサーボードに腐食がありセンサーボード交換実施。久保さんにセンサーボード届けてもらう。 | センサーカセット初期化不良との事で訪問：機械がシャットダウンしており履歴を見るとTemp-Ｄとオーバーヒートエラーも出ていた。センサーボードに腐食がありセンサーボード交換実施。久保さんにセンサーボード届けてもらう。 | センサーカセット初期化不良との事で訪問：機械がシャットダウンしており履歴を見るとTemp-Ｄとオーバーヒートエラーも出ていた。センサーボードに腐食がありセンサーボード交換実施。久保さんにセンサーボード届けてもらう。 | SWO-293178 | [5 empty] | The sensor cassette was not initialized and visited by the thing: the machine was shutting down and when i looked at the history, I saw temp-D and overheating errors. There is corrosion in the sensor board and the sensor board is replaced. I have Kubo deliver the sensor board. | The sensor cassette was not initialized and visited by the thing: the machine was shutting down and when i looked at the history, I saw temp-D and overheating errors. There is corrosion in the sensor board and the sensor board is replaced. I have Kubo deliver the sensor board. | The sensor cassette was not initialized and visited by the thing: the machine was shutting down and when i looked at the history, I saw temp-D and overheating errors. There is corrosion in the sensor board and the sensor board is replaced. I have Kubo deliver the sensor board.
No | Visit | Complaint | RCHN | 温度错误，接口板腐蚀 | 温度错误，接口板腐蚀，需要借件902-972 | 使用备用机备件902-972及更换2个试剂包和1张测试卡后修复by tim | SWO-293010 | Defect | HW | Basic Unit | PCB, Sensor connector assembly |  | Temperature error, interface plate corrosion | Temperature error, interface plate corrosion, need to borrow 902-972 | Repair by tim with spare parts 902-972 and replacement of 2 reagent kits and 1 test card
No | Visit | Complaint | RIND | Service Visit. | Sensor cassette has expired install the new one. | Install the new sensor cassette in the analyzer and change the patient ranges, and sample information according to the user. Successfully done, analyzer handed over to staff in well working condition. | SWO-293579 | [5 empty] | Service Visit. | Sensor cassette has expired install the new one. | Install the new sensor cassette in the analyzer and change the patient ranges, and sample information according to the user. Successfully done, analyzer handed over to staff in well working condition.
No | Visit | Complaint | RCHN | 定标不过 | 定标不过，借903-581一个，891-051 2个 | 更换903-581后修复by tim | SWO-291151 | Defect | HW | Basic Unit | PCB, Sensor interface |  | The calibration is not the | Bid, however, borrow 903-581 one, 891-051 2 | Repair by tim after replacement 903-581
No | Visit | Complaint | RCHN | 报错3267 | 报错3267，加借944-394（13418）一张 缺货 | 重新定标后修复 by tim | SWO-291872 | Adjustment | Consumables | Test Kit | Cartridge, PCT | 3267 | Error 3267 | 3267, plus 944-394 (13418) a out of stock | Fix by tim after re-calibration
No | Visit | Complaint | RCHN | 报错1999 | 报错1999，借903-198一个 | 更换903-198后修复 by tim | SWO-291139 | Defect | HW | Analyzer Control Module | PCB, Main Board | 1999 | Error 1999 | Misquoted 1999, loan 903-198 one | Repair by tim after replacement 903-198
No | Visit | Complaint | RCHN | 报错1031 | 报错1031，借903-364一个 | 更换903-364后修复by tim | SWO-292009 | Defect | HW | Process Module | Optical Unit | 1031 | 1031 | 1031, loan 903-364 one | Repair by tim after replacing 903-364
No | Visit | Complaint | RCHN | 突然断电 | 突然断电，借944-222（13153）缺货 | 重新定标后修复by tim | SWO-292900 | User error | Consumables | Test Kit | Power Failure (no UPS) |  | Sudden power outage | Sudden power outage, 944-222 (13153) out of stock | Fix by tim after re-calibration
No | Visit | Complaint | RCHN | TNI定标不过 | TNI定标不过，借903-364一个。 | 更换备件903-364后修复by tim | SWO-291194 | Defect | HW | Process Module | Optical Unit | Unknown | TNI calibration is not the | TNI calibration, however, borrows 903-364 one. | Repair by tim after replacement of spare parts 903-364
No | Visit | Complaint | RIND | Na and K values coming high. | Na and K values coming high. | Protean removal and cleaning programs ran. Service program ran. Analyzer Serviced and calibrated properly. Handed over in a fine working condition. Asked to purchase QCs. | SWO-293538 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | Na and K values coming high. | Na and K values coming high. | Protean removal and cleaning programs ran. Service program ran. Analyzer Serviced and calibrated properly. Handed over in a fine working condition. Asked to purchase QCs.
No | Visit | Complaint | RIND | Service Visit. | Need to change the lac.Na, glu and pCO2 membrane. | Keep on hold the analyzer, Change the lac and Na membrane, record replacement of membrane, restart the analyzer and wait for auto two point calibration, drift error present in lac give the manual two point calibration again. Need to change the glu and pCO2 membrane. | SWO-293580 | Revisit required | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | Service Visit. | Need to change the lac. Na, glu and pCO2 membrane. | Keep on hold the analyzer, Change the lac and Na membrane, record replacement of membrane, restart the analyzer and wait for auto two point calibration, drift error present in lac give the manual two point calibration again. Need to change the glu and pCO2 membrane.
No | Visit | Complaint | RCHN | 屏幕不灵敏 | 屏幕不灵敏，借910-320一个 | 更换910-320后修复by tim | SWO-291916 | Defect | HW | DMS_PC | Touch screen |  | The screen is not sensitive | The screen is not sensitive, borrow 910-320 one | Repair by tim after replacement 910-320
No | Visit | Complaint | RDEU | 12.01.2018 Herr Hölleriegel 0821400165692 meldet, dass es zu Probenfehler bei der OXI-Messung kommt. | Probenproblem - Fehler #0656 und #0664 | Oxi Complete Test durchgeführt - ok Ventile aus pH $BG sowie El. $Met. Modul ausgebaut und gereinigt - wieder eingebaut Luftweg im Fluidic Modul gespült, Proteinentfernung durchgeführt LS Justage sowie Pumpenkalibrierung durchgeführt tHb kalibriert, Gerät kalibriert und überprüft | SWO-293401 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | 12.01.2018 Mr. Hölleriegel 0821400165692 reports that there are sample errors in the OXI measurement. | Sample problem - errors #0656 and #0664 | Oxi Complete Test performed - ok Valves made of pH $BG and El. $Met. Module removed and cleaned - reinstalled Airway rinsed in fluidic module, protein removal performed LS adjustment and pump calibration performed tHb calibrated, device calibrated and checked
No | Visit | Complaint | RIND | Touch screen not working | Checked and found touch screen is not working | Tried to update touch screen firmware as per service tip received but Citron Infrared Touch option is not found on control panel nor USB-key boot priority option available in Advanced BIOS Features, therefore touch screen firmware upgrade could not be performed. | SWO-293637 | Defect | HW | DMS_PC | Touch screen | could not upgrade FM | Touch screen not working | Checked and found touch screen is not working | Tried to update touch screen firmware as per service tip received but Citron Infrared Touch option is not found on control panel nor USB-key boot priority option available in Advanced BIOS Features, therefore touch screen firmware upgrade could not be performed.
No | Visit | Complaint | RAME | Po2 failing Zero calibrations. | Customer called with PO2 failing the Zero Calibrations. Zero calibration is running 6.5 to 7.5. | Replaced the PO2 membrane and performed the Zero Calibration in service with same failing results. Replaced the PO2 electrode and the calibrations passed with Zero reading 1.2. Analyzer is less than a year old and the electrode will be covered under warranty. | SWO-293435 | Assisted Replacement | Electrode | Electrode | pO2 |  | Po2 failing Zero calibrations. | Customer called with PO2 failing the Zero Calibrations. Zero calibration is running 6.5 to 7.5. | Replaced the PO2 membrane and performed the Zero Calibration in service with same failing results. Replaced the PO2 electrode and the calibrations passed with Zero reading 1.2. Analyzer is less than a year old and the electrode will be covered under warranty.
No | Visit | Complaint | RAME | Customer has had 4 oxi spectrum mismatch errors on QC over the past month. | Checked data logs, found errors. Checked pinch valve. Found it to be slightly sticking. | Lubed the pinch valve a little bit so that it will respond better. Adjusted lamp unit, ran tHb calibration. | SWO-293617 | Defect | HW | Oxi module | Pinch Valve |  | Customer has had 4 oxi spectrum mismatch errors on QC over the past month. | Checked data logs, found errors. Checked pinch valve. Found it to be slightly sticking. | Lubed the pinch valve a little bit so that it will respond better. Adjusted lamp unit, ran tHb calibration.
No | Visit | Complaint | RSCH | Blank | Flüssigkeittransportfehler Adapter zu Einlassnadel fehlt (liegt neben ABL auf dem Tisch). | Nadel überprüft, Adapter ersetzt und kalibriert. Im Servicemenu sämtliche Tests der Flüssigkeitswege und Dichtheit des Systems durchgeführt. SC neu installiert. Kal., QC und System i.o. Funktionskontrolle i.o. Kunde wurde am Telefon über Mehrkosten informiert. | SWO-293695 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | Blank | Liquid transport error Adapter to inlet pin is missing (located next to ABL on the table). | Needle checked, adapter replaced and calibrated. In the service menu all tests of the fluid pathways and tightness of the system are carried out. SC reinstalled. Kal., QC and System i.o. Functional control i.o. Customer was informed on the phone about additional costs.
No | Visit | Complaint | RSAF | Solution pack replacement recommended by analyzer, customer advised to do the manual back flush and call back if not corrected, standing in front of analyzer, no telephonic assistance called for, Sr Gloria called in the morning to say that the trouble shooting did not resolve the issue and that Sr Prabashnee said I must respond. | Found that the sample path is blocked on the hemolyzer, Done back flush and Fluid transport test but still failed, | Removed all consumables from analyzer and flushed the hemolyzer only with Hypo chloride solution and got the clot dislodged and the sample path open, reinstalled all consumables and done Fluid transport test again and all passed. | SWO-293661 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway | Hemolyzer | Solution pack replacement recommended by analyzer, customer advised to do the manual back flush and call back if not corrected, standing in front of analyzer, no telephonic assistance called for, Sr Gloria called in the morning to say that the trouble shooting did not resolve the issue and that Sr Prabashnee said I must respond. | Found that the sample path is blocked on the hemolyzer, Done back flush and Fluid transport test but still failed, | Removed all consumables from analyzer and flushed the hemolyzer only with Hypo chloride solution and got the clot dislodged and the sample path open, reinstalled all consumables and done Fluid transport test again and all passed.
Yes | Visit | Complaint | RIND | SC Not installing. | SC Not installing. | SC leakage found. SC and Sensor Board found burned. Need to replace Sensor board. | SWO-293834 | [5 empty] | SC Not installing. | SC Not installing. | SC leakage found. SC and Sensor Board found burned. Need to replace Sensor board.
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 报错0210,1310.LAC定标失败 | 更换测试卡后修复 | SWO-299074 | [5 empty] | Test card, reagent package failure | False0210,1310.LAC bid failure | Repair after replacing test card
No | Visit | Complaint | RCHN | 报错1068 | 报错1068，借903-210一个 | 维修时发现光单卡扣缺失，要加借048-319一个，此单先关单by tim | SWO-292905 | Revisit required | HW | Process Module | PCB, Cup dryer | 1068 | 1068 error | 1068, borrow edith 903-210 | Repair found that the light single snap is missing, to borrow 048-319 one, this single first close single by tim
No | Visit | Complaint | RIND | analyzer showing frequently error neon number too high. | Lamp unit faulty. | replaced Lamp unit now and performed oxi calibration and performed complete oxi test now machine working fine. | SWO-294144 | Defect | HW | Oxi module | Lamp unit | Neon | analyzer showing frequently error neon number too high. | Lamp unit faulty. | replaced Lamp unit now and performed oxi calibration and performed complete oxi test now machine working fine.
No | Visit | Complaint | RAME | 0251 Halogen Lamp Intensity too high. | Inspection of the Oxi system revealed the halogen lamp intensity is too high, adjustment needed. | Adjusted lamp unit per procedure. | SWO-293618 | Adjustment | HW | Oxi module | Lamp unit | Halogen high | 0251 Halogen Lamp Intensity too high. | Inspection of the Oxi system revealed the halogen lamp intensity is too high, adjustment needed. | Adjusted lamp unit per procedure.
No | Visit | Complaint | RCHN | 报错1068 | 维修发现光单卡扣缺失，借048-319一个 | 更换903-210及048-319后修复by tim | SWO-293632 | Defect | HW | Process Module | PCB, Cup dryer | 1068 | 1068 error | Repair found that the light single snap is missing, borrow 048-319 one | Repair by tim after replacing 903-210 and 048-319
No | Visit | Complaint | RCHN | 报错2167 | 报错2167 | 维修机器 | SWO-293836 | Defect | HW | Storage Module | Sensor, Cup | 2167 | 2167 error | 2167 error | Repairing the machine
Yes | Visit | Complaint | RKOR | P.M AQURE 서버 교체작업 | 아큐어 서버 이슈로인해 데모설치한장비를 빼고 원래 서버를 다시 넣음 | . | SWO-297267 | Re classified RMED | His/Lis Setup |  |  |  | P.m Replacing AQURE servers | Reinsert the original server without the demo installed due to the Acure server issue | .
No | Visit | Complaint | RAME | Battery Error | The analyzer is presenting a system message that one or more battery cells are failing. The system requires the removal or replacement of the battery otherwise it would lock out the analyzer. | Visited and, confirmed the problem. Checked the activity log and, found error 1348 battery low. The battery pack was not being correctly charged. Replaced a new battery pack. | SWO-294351 | Defect | HW | Battery module | Battery |  | Battery Error | The analyzer is presenting a system message that one or more battery cells are failing. The system requires the removal or replacement of the battery otherwise it would lock out the analyzer. | Visited and, confirmed the problem. Checked the activity log and, found error 1348 battery low. The battery pack was not being correctly charged. Replaced a new battery pack.
No | 5 | Y | RPOL | Blank | Diagnoza problemu, odka¿enie modu³u OXI. Wymiana zaworu w module OXI. Testy serwisowe oraz na krwi pacjenta. Analizator sprawny. |  | D3C1B6F3C84C46ED | Defect | HW | Oxi module | Pinch Valve |  | Blank | Diagnosis of the problem, the withdrawal of the OXI module. Valve replacement in OXI module. Service tests and the patient's blood. Functional Analyzer.
No | Visit | Complaint | RCHN | 打印机故障 | 打印机故障，借902-985一个 |  | SWO-293888 | [5 empty] | Printer failure | Printer failure, borrow 902-985 one
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 定标失败，质控失败 | 更换两张测试卡正常。 | SWO-294620 | [5 empty] | Test card, reagent package failure | Bid failure, quality control failure | Replacing the two test cards is ok.
No | Visit | Complaint | RCHN | 试管夹损坏 | 试管夹损坏 | 更换903-584一个正常。(By Ran) | SWO-295362 | Defect | HW | Aspiration Module | Tube holder |  | Damage to test tube clip | Damage to test tube clip | Replace the 903-584 one normal. (By Ran)
No | Visit | Complaint | RCHN | 报错2168 | 报错2168 | 更换903-030后修复(By Ran) | SWO-294453 | Defect | HW | Storage Module | PCB, Cartridge Lock | 2168 | 2168 error | 2168 error | Repair after replacement 903-030 (By Ran)
No | Visit | Complaint | RIND | Service Visit | In some sample oximetry values not came, need to change the glu membrane. | Do the protein removal 3 times in auxiliary program, and the give the command of cleaning. Now go to service program in which do the complete oximetry test in which all the three spectrometer, hemolyzer and lamp unit successfully passed, exit from the service program, and give the two point calibration to the analyzer. After that keep on hold the analyzer and replace the glucose membrane, restart the analyzer and wait for auto two point calibration which is unsuccessful drift error present, give the again two point calibration, which is successfully done for glucose. Now analyzer handed over to staff in well working condition. | SWO-295201 | Revisit required | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | Service Visit | In some sample oximetry values not came, need to change the glu membrane. | Do the protein removal 3 times in auxiliary program, and the give the command of cleaning. Now go to service program in which do the complete oximetry test in which all the three spectrometer, hemolyzer and lamp unit successfully passed, exit from the service program, and give the two point calibration to the analyzer. After that keep on hold the analyzer and replace the glucose membrane, restart the analyzer and wait for auto two point calibration which is unsuccessful drift error present, give the again two point calibration, which is successfully done for glucose. Now analyzer handed over to staff in well working condition.
No | Visit | Complaint | RAME | Customer ran a sample and allowed the inlet probe to go into the gel in a test tube which drew gel into system which let to a clot in the system. With gel and clot analyzer overflowed around inlet area. | Waste line under inlet gasket kept coming off after gel was in the system, customer cut the line shorter and then replaced with a tubing of their own. Analyzer leaked fluid. Customer found inlet probe was bent and replaced from their stock. I remoted into analyzer and found the following errors: 0357: Temp. Error: Barometer, 0662: Barometer out of range, 0689: Gas pressure offset voltage out of range, 0238: Spectrometer voltage bias too low, 0618: Waste detector error & 0963: Leak Currents detected. Instructed customer to power down analyzer and leave until I arrive. Ordered AU pcb. Upon arrival, I powered on and all the above errors were still present except 0963. | Replaced AU pcb, installed AU software, reloaded keys, performed zero current calibration, scanned in new Cal 1, Cal 2, cleaning solution and replaced Crea A & B membranes. | SWO-294254 | User error | Sample handling | Sample handling | Sample handling | Gel | Customer ran a sample and allowed the inlet probe to go into the gel in a test tube which drew gel into system which let to a clot in the system. With gel and clot analyzer overflowed around inlet area. | Waste line under inlet gasket kept coming off after gel was in the system, customer cut the line shorter and then replaced with a tubing of their own. Analyzer leaked fluid. Customer found inlet probe was bent and replaced from their stock. I remoted into analyzer and found the following errors: 0357: Temp. Error: Barometer, 0662: Barometer out of range, 0689: Gas pressure offset voltage out of range, 0238: Spectrometer voltage bias too low, 0618: Waste detector error & 0963: Leak Currents detected. Instructed customer to power down analyzer and leave until I arrive. Ordered AU pcb. Upon arrival, I powered on and all the above errors were still present except 0963. | Replaced AU pcb, installed AU software, reloaded keys, performed zero current calibration, scanned in new Cal 1, Cal 2, cleaning solution and replaced Crea A & B membranes.
No | Visit | Complaint | RDAN | ABL har kørt med gas 1 lot Y10H25 og med dette lot har ABL meldt at gas 1 var tom før tid. | Gas 1 har igen været tom for gas, før den burde. | Gas modulet er udskiftet. Der er kørt test i service og alt ser OK ud. Der er kørt 2 punkts kal. Der er kørt QCer. Alt ser ok ud nu. | SWO-293987 | Defect | HW | Gas module | Gas module |  | ABL has been running with gas 1 lot Y10H25 and with this lot, ABL has reported that gas 1 was empty before time. | Gas 1 has again been empty for gas before it should. | The Gas module has been replaced. There are run tests in service and everything looks OK. Two-point Kal has been run. QCer has been run. Everything looks OK now.
No | Visit | Complaint | RDEU | Das Anmeldefenster überdeckt den Analysator/Statusfeld | Kunde ist mit dem Zustand nicht zufrieden und möchte seit langem dass diese Zustand geändert wird. | RMED schlägt vor den Analysator $Status Icon an den unteren Taskleiste zu platzieren. Das habe ich versucht. Nachteil dabei ist dass man erst auf dem Anmeldebotten drücken muss damit das Anmeldefenster erscheint. BORWE fragt bei RMED nach. | SWO-294190 | Revisit required | SW | Windows | Win Log On | Techline raised | The login window covers the analyzer/status field | Customer is not satisfied with the condition and has long wanted this condition to be changed. | RMED suggests placing the analyzer $Status icon at the bottom taskbar. That's what I tried to do. The disadvantage here is that you first have to press on the login vat so that the login window appears. BORWE asks RMED.
No | Visit | Complaint | RFRA | Pb d'intensité de la lampe homogène | Contrôle du réglage de la lampe | Changement de l'unité lampe. Realignement de la lampe.Cal tHb. Essais | SWO-294226 | Defect | HW | Oxi module | Lamp unit |  | Pb of homogeneous lamp intensity | Controlling the lamp setting | Change of the lamp unit. Realignment of the lamp. Cal tHb. Tests
No | Visit | Complaint | RSCH | Rep. CPU | Gerät startet nicht mehr pH instabil | Vor Ort Feinsicherungen getestet i.o Powersupply ersetzt, Fehler weiterhin In Thalwil Ersatz CPU geholt Vor Ort eingebaut i.o O-Ring pH ersetzt i.o Kal. und QC i.o System i.o Neue CPU wird bei nächster Wartung eingebaut | SWO-294340 | Defect | HW | DMS_PC | CPU |  | Rep. CPU | Device no longer starts pH unstable | On-site fine fuses tested i.o Powersupply replaced, errors continue In Thalwil replacement CPU fetched Built on site i.o O-ring pH replaces i.o Kal. and QC i.o System i.o New CPU will be installed at next maintenance
No | Visit | Complaint | RESP | Multicheck nível 1, 2 e 3 com valores altos na TnT | QC multicheck alto nos níveis 1, 2 e 3 | Verificação dos registros, calibração de um novo lote de TnT e controlos mantêm-se altos, brancos com valores à volta de 600 contagens nos 5 testes, calibração de TnI com QC Multicheck OK. Passagem de QC com novo lote OK | SWO-294638 | Assisted Replacement | Consumables | LQC Kit | LQC Kit, TnT |  | Multicheck Level 1, 2 and 3 with high values on TnT | High Multicheck QC at levels 1, 2 and 3 | Verification of records, calibration of a new batch of TnT and controls remain high, white with values around 600 counts in the 5 tests, calibration of TnI with QC Multicheck OK. QC Pass with new batch OK
Yes | Visit | Complaint | RDEU | 12.01.2018, Sr. Bianca, Tel: 03518566312: LQC CRP außer Bereich, letzte Woche auch schon. | Datenbank geprüft | QC nach Jahreswartung weiterhin außer Bereich ( CRP Level1 Lot13623 ) Station ist informiert Daten gesichert --> Daten an TS gesendet Test mit Multicheck i.O TnT und CKMB i.O | SWO-294105 | [5 empty] | 12.01.2018, Sr. Bianca, Tel: 03518566312: LQC CRP out of range, last week already. | Database checked | QC after annual maintenance continues out of range ( CRP Level1 Lot13623 ) Station is informed Data backed up --> Data sent to TS Test with Multicheck i.O TnT and CKMB i.O
No | 21 | Y | RPOL | Blank | wymiana silnika pompy. Kalibracja parametrów. QC. Aparat sprawny technicznie. Aktualizacja oprogramowania do 3.3 MR1 |  | D3BF71A6DC30FF22 | Defect | HW | Basic Unit | Steppermotor, main |  | Blank | replacing the pump motor. Parameter calibration. Qc. The camera is technically efficient. Software Update for 3.3 MR1
Yes | 21 | Y | RPOL | Blank | zabranie aparatu do naprawy |  | D3BF3EABDC36FF20 | Re classified RMED | Just pickup, see D3BF71A6DC30FF22 |  |  |  | Blank | Taking the camera for repair
No | Visit | Complaint | RAME | The analyzer is asking to replace the solution pack that changed yesterday but, it is stuck. Nestor cannot remove it. | The solution pack is stuck and, cannot be removed. | Nestor had replaced a new solution pack and, sensor cassette yesterday but, today the analyzer is requesting a solution pack change again. Upon inspection of the analyzer, I found that the solution pack was not properly activated. The right top part of the pack looked raised up and so, when Nestor tried to remove it then it would not come up. I used the ejection lever in the back to remove the solution pack and, reactivated it. I cleaned the flow selector sensor window and, reinstall the solution pack. The analyzer restarted properly with no errors. Performed all the fluid transport checks in service mode, all passed within manufacturer specifications. | SWO-295112 | User error | SP Activation | SP Activation | SP Activation |  | The analyzer is asking to replace the solution pack that changed yesterday but, it is stuck. Nestor cannot remove it. | The solution pack is stuck and, cannot be removed. | Nestor had replaced a new solution pack and, sensor cassette yesterday but, today the analyzer is requesting a solution pack change again. Upon inspection of the analyzer, I found that the solution pack was not properly activated. The right top part of the pack looked raised up and so, when Nestor tried to remove it then it would not come up. I used the ejection lever in the back to remove the solution pack and, reactivated it. I cleaned the flow selector sensor window and, reinstall the solution pack. The analyzer restarted properly with no errors. Performed all the fluid transport checks in service mode, all passed within manufacturer specifications.
No | Visit | Complaint | RAME | Leak detected | Customer called Radiometer On-Call Tech Support on 1 $14 regarding Leak detected. | Radiometer On-Call Tech Support instructed customer to replace sensor cassette. All calibrations and QC passed. Conducting Leak detected investigation with all three analyzers. | SWO-294968 | Assisted Replacement | SC | SC | Leak |  | Leak detected | Customer called Radiometer On-Call Tech Support on 1 /14 regarding Leak detected. | Radiometer On-Call Tech Support instructed customer to replace sensor cassette. All calibrations and QC passed. Conducting Leak detected investigation with all three analyzers.
No | Visit | Complaint | RAME | The wireless connection is sporadically lost. | Visit to download data log and, service dump files. | Downloaded service dump and, data log files. Checked the wireless dongle status and, the blue LED light is 'On' at this time. I also checked the WIFI-Lan status and, the signal strength looks good at 100%. The network connection status icon shows connected. I have discussed the connectivity issues that Susana is concerned about and, we will submit a CQC (Techline) case. | SWO-294395 | Setup | His/Lis | His/Lis | His/Lis | Techline raised | The wireless connection is sporadically lost. | Visit to download data log and, service dump files. | Downloaded service dump and, data log files. Checked the wireless dongle status and, the blue LED light is 'On' at this time. I also checked the WIFI-Lan status and, the signal strength looks good at 100%. The network connection status icon shows connected. I have discussed the connectivity issues that Susana is concerned about and, we will submit a CQC (Techline) case.
No | Visit | Complaint | RKOR | Sample data error | . | Sample data check. Calibration check QC check. | SWO-300559 | User error | Preanalytical error | Preanalytical error | Preanalytical error |  | Sample data error | . | Sample data check. Calibration check QC check.
No | Visit | Complaint | RAME | liquid transport error during rinse | 1343 Unable to pump solutions, 508 Liquid transport error during rinse.. tHB calibration overdue. vacuum value test is failing @ 1.00, inlet gasket holder dirty, caked with blood inside, fluid path not clear. | Replaced inlet gasket holder with customer spare, cleaned probe, replaced solution pack due to excessive buildup on selector balance, flushed fluidics, retested, issues resolved, performed tHB calibration, no further errors observed or noted. | SWO-294409 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | liquid transport error during rinse | 1343 Unable to pump solutions, 508 Liquid transport error during rinse.. tHB calibration overdue. vacuum value test is failing @ 1.00, inlet gasket holder dirty, caked with blood inside, fluid path not clear. | Replaced inlet gasket holder with customer spare, cleaned probe, replaced solution pack due to excessive buildup on selector balance, flushed fluidics, retested, issues resolved, performed tHB calibration, no further errors observed or noted.
No | Visit | Complaint | RAME | PO2 status errors during calibration. | PO2 status errors during calibration. PO2 within parameters for QC, but questioning due to calibration errors. | Swapped solution pack with known good solution pack. Completed all internal and external Quality controls along with two calibrations. Data collected for further investigation. | SWO-295041 | Assisted Replacement | SP | SP | pO2 |  | PO2 status errors during calibration. | PO2 status errors during calibration. PO2 within parameters for QC, but questioning due to calibration errors. | Swapped solution pack with known good solution pack. Completed all internal and external Quality controls along with two calibrations. Data collected for further investigation.
No | Visit | Complaint | RLTD | Lynn Tel 01582 497991 | Battery showing 0% | Reset battery as per Radiometers service tip, now ok. | SWO-295467 | Defect | SW | Battery module | Battery |  | Lynn Tel 01582 497991 | Battery showing 0% | Reset battery as per Radiometers service tip, now ok.
No | Visit | Complaint | RDEU | Oxymetrieprobleme bei Kal., QC, tHb-Kal. --> trotz Tausch des SP und durchführen einer tHb-Kal. | Fehlersuche --> Dateiarchive (Kal., QC, tHb-Kal., Aktivitäten) überprüft --> Oxymetrie Hardware Problem (kein Oxy) vorhanden --> Oxy Complete Test durchführen --> tHB-Kal. durchführen --> Fehler 1012: Kein Probenspektrum für Oxi --> Abhilfe durch Neustart des Analysators | Oxy Complete Test durchgeführt --> Test nicht OK (Fehler bei Cavitation Absorbance und Neon position, zu hohe bzw niedrige Werte), Analysator neu gestartet und nochmals Oxy Complete Test --> 2ter Oxy complete Test OK, Gerät nochmals neu gestartet und nochmals einen Oxy Complete Test durchgeführt --> Test OK, tHb-Kal. durchgeführt, System und Funktion i.O. | SWO-294635 | Defect | SW | Oxi module | off/on |  | Oxymetrieprobleme bei Kal., QC, tHb-Kal. --> trotz Tausch des SP und durchführen einer tHb-Kal. | Troubleshooting --> File archives (Cal., QC, tHb-Kal., Activities) checked --> Oxymetry Hardware problem (no oxy) present --> Oxy Complete Test --> tHB-Kal. Perform --> Error 1012: No sample spectrum for Oxi --> Remedy by restarting the analyzer | Oxy Complete Test performed --> Test not OK (error in Cavitation Absorbance and Neon position, too high or low values), Analyzer restarted and again Oxy Complete Test --> 2ter Oxy complete Test OK, device restarted and another Oxy Complete Test --> Test OK, tHb-Kal. system and function in the
No | Visit | Complaint | RBNL | Inlet module sluit niet na aanbieden capillair | Na aanbieden van een capillair sluit de inlet module niet. | Na vervangen van de inlet module capillairen gemeten: nu OK. Serienummer vervangen inlet module: R0161N0012. Deze module valt onder garantie en moet terug naar RMED. | SWO-295541 | Defect | HW | Auto Inlet module | Cap - Spring |  | Inlet module does not close after offering capillary | After offering a capillary, the inlet module does not close. | After replacing the inlet module capillaries measured: now OK. Serial number replace inlet module: R0161N0012. This module is under warranty and must be returned to RMED.
No | Visit | Complaint | RCHN | 触摸屏不灵敏 | 触摸屏不灵敏，借552-036一个。 | 更换552-036一个正常。(By Ran) | SWO-293932 | [5 empty] | The touch screen is not sensitive | The touch screen is not sensitive, borrow 552-036 one. | Replace 552-036 one normal. (By Ran)
No | Visit | Complaint | RCHN | 抽吸样本失败 | 抽吸样本失败，借903-199,903-644各一个 | 机器未修复，备件903-199,903-644未使用，现要加借903-129一个，此单先关单。(By Ran) Need more spare parts. SWO closed. | SWO-293865 | Revisit required | HW | Basic Unit | PCB, Wet section Control |  | Failed to pump samples | Failed to pump samples, borrow 903-199,903-644 each | The machine has not been repaired, spare parts 903-199,903-644 is not used, now to borrow 903-129 one, this single first clearance. (By Ran) Need more spare parts. SWO closed.
No | Visit | Complaint | RCHN | 报错1068 | 报错1068，借848-037一个 | 清洁泵里面的脏东西。备件好还。(By Ran) Cleaned the membrane pump. | SWO-293868 | Cleaned | HW | Process Module | Membrane Pump | 1068 | 1068 error | 1068, loan 848-037 one | Clean the dirt inside the pump. Spare parts are good also. (By Ran) Cleaned the pump pump.
No | Visit | Complaint | RFRA | Reboot de l'ABL9 après écran figé. | Installation ABL9 de prêt n°série=400333 | Réception ABL9 de prêt. Mise en route Installation des consommables. calibration 2 points et passage des Cq Qc4+ ok. | SWO-294896 | [5 empty] | ABL9 reboot after frozen screen. | ABL9 Loan Installation Series 400333 | ABL9 loan reception. Getting started Installing consumables. calibration 2 points and passage of the Cq Qc4 ok.
No | Visit | Complaint | RIND | Service Visit. | Need to change the glucose membrane. | Keep on hold the analyzer, replace the glucose membrane, record the replacement of glucose membrane, close the chamber and restart the analyzer, wait for auto two point calibration, which is not successfully done due to drift present in the calibration of glucose. Give the manual 2 point calibration command for glucose to the analyzer which is successfully done. Need to change the pCO2 membrane. | SWO-296973 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway | Zero adjust required | Service Visit. | Need to change the glucose membrane. | Keep on hold the analyzer, replace the glucose membrane, record the replacement of glucose membrane, close the chamber and restart the analyzer, wait for auto two point calibration, which is not successfully done due to drift present in the calibration of glucose. Give the manual 2 point calibration command for glucose to the analyzer which is successfully done. Need to change the pCO2 membrane.
No | Visit | Complaint | RIND | Oximetery values of the machine is not reliable. | Check the analyzer found complete oximetry test is failed again and again. Due to Neon lamp voltage going high. | Change oxi module pcb, run complete oximetry test after change pcb now it is pass. Run 2 point calibration. | SWO-295237 | Defect | HW | Oxi module | PCB, Oxi module |  | Oximetery values of the machine is not reliable. | Check the analyzer found complete oximetry test is failed again and again. Due to Neon lamp voltage going high. | Change oxi module pcb, run complete oximetry test after change pcb now it is pass. Run 2 point calibration.
Yes | 12 | Y | RPAC | Blank | Ref Parameters failing qc - Na Cl |  | D3BF771EDC5E461E | [5 empty] | Blank | Ref Parameters failing qc - Na Cl
No | 3 | Y | RPAC | Blank | Replace FlexQ Tray |  | D3BF5886DC50BF88 | Defect | HW | FlexQ Module | FlexQ Tray |  | Blank | Replace FlexQ Tray
No | Visit | Complaint | RDEU | 17.01.2018, 10:11 Uhr, Fr. Hägerich, Tel. 091318539119: Endlosschleife LP wechseln | Fehlersuche --> Dateiarchive (Aktivitäten, Kal., QC, Betriebsmittel), Einlass-Modul, SP --> Fehler des öfteren Inhomogene QC3-Lösung und Flüssigkeitstransport gescheitert --> Flow-Selector mit Blut verschmutzt, Einlass-Modul defekt, SP erneuern (Fehler 1200: Lösungspack leer) | Einlassmodul (R0010N003) eingebaut und anderes ausgebaut, Flow-Selector gereinigt, Analysator im Service (Adjust Manifold, Pumpenkal., LS-Justage, Clot detect $remove, Drucktest, SP Check) durchgeführt --> Service-Test OK, Lösungspack aus Kundenbestand (von INZ, Lungenfunktion ausgeliehen) (944-197, BA-50) eingebaut und (944-497, YW-50) ausgebaut, mehrere Testmessungen durchgeführt, System und Funktion i.O. | SWO-295483 | Cleaned | HW | Basic Unit | Flow Selector |  | 17.01.2018, 10:11 a.m., Fr. Hägerich, Tel. 091318539119: Change endless loop LP | Troubleshooting --> File archives (activities, cal., QC, equipment), inlet module, SP --> Error of the frequent inhomogeneous QC3 solution and fluid transport failed --> Flow-Selector with blood dirty, inlet module defective, SP renew (Error 1200: solution pack empty) | Inlet module (R0010N003) installed and other removed, flow selector cleaned, analyzer in service (Adjust Manifold, pump scale, LS-Adjustment, Clot detect $remove, print test, SP check) carried out --> Service test OK, solution pack from customer base (from INZ, lung function borrowed) (944-197, BA-50) and (944-497, YW-50) removed, several test measurements carried out, system and function in the
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 质控C超范围 | 更换试剂包后修复 | SWO-295919 | [5 empty] | Test card, reagent package failure | Quality control C over-range | Repair after replacement of reagent kit
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | glu,LAC定标失败 | 更换测试卡，试剂包正常。 | SWO-295454 | [5 empty] | Test card, reagent package failure | glu, LAC calibration failed | Replace the test card and the reagent package is OK.
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 测试卡参比电极错误 1361 | 更换测试卡正常。 | SWO-300664 | Assisted Replacement | SC | SC | Reference | 1361 | Test card, reagent package failure | Test card reference electrode error 1361 | The replacement test card is OK.
No | 2 | Y | RPAC | Blank | Battery |  | D3C08AC85203AC02 | [5 empty] | Blank | Battery
No | Visit | Complaint | RCHN | 光单故障 | 光单故障，借903-364一个 | 更换903-364一个正常。(By Ran) | SWO-295163 | Defect | HW | Process Module | Optical Unit |  | Light single fault | Light single fault, borrow 903-364 one | Replace 903-364 one normal. (By Ran)
No | Visit | Complaint | RCHN | 温度传感器断裂 | 温度传感器断裂,需要借件905-908一个 | 更换905-908后修复(By Ran) | SWO-293887 | Defect | HW | Process Module | Sensor, BnP |  | The temperature sensor is broken. | Temperature sensor is broken and requires a loan of 905-908 | Repair after replacement 905-908 (By Ran)
No | Visit | Complaint | RIND | Calibration error. | Check the system. Calibration logs. | Need to observe the abl9. | SWO-297196 | [5 empty] | Calibration error. | Check the system. Calibration logs. | Need to observe the abl9.
No | Visit | Complaint | RHUN | Oxi teszt hiba: | Oxi teszt hiba: Hemolyzer phase regulation failed | Alkatrész cseréje, Lámpa beállítás, OXI tesztek, Kalibrálás, tHb, QC mérés ** A javítás garanciális ** | SWO-295292 | Defect | HW | Oxi module | Hemolyzer |  | Oxy Test Error: | Oxy Test Error: Hemolyzer Phase regulation failed | Replacement parts, lamp adjustment, OXY tests, calibration, tHb, QC measurement ** The repair is under warranty **
No | Visit | Complaint | RAME | Fails measurement with oximetry 581 error | 581 spectrum mismatch | Replaced hemolyzer | SWO-295718 | Defect | HW | Oxi module | Hemolyzer | 581 | Fails measurement with oximetry 581 error | 581 spectrum mismatch | Replaced hemolyzer
No | 3 | Y | RCZE | Blank | po instalaci modulu Jirkou nefunkcni |  | D3C18904C827DC6C | Defect | HW | FlexQ Module | FlexQ Module |  | Blank | After the installation of the Jirkou module
No | Visit | Complaint | RKK | オキシハードウェアトラブルで測定出来ません。Halogen Lamp intensity too high。ランプ調整、tHbキャル実施。Glu,Lac,Refメンブラン交換。 | オキシハードウェアトラブルで測定出来ません。Halogen Lamp intensity too high。 | ランプ調整、tHbキャル実施。Glu,Lac,Refメンブラン交換。 | SWO-295450 | Adjustment | HW | Oxi module | Lamp unit | Halogen high | It is not possible to measure it by oxy hardware trouble. Halogen Lamp intensity too high. Lamp adjustment, tHb cal. Glu, Lac, Ref Membrane replacement. | It is not possible to measure it by oxy hardware trouble. Halogen Lamp intensity too high. | Lamp adjustment, tHb cal. Glu, Lac, Ref Membrane replacement.
No | Visit | Complaint | RLTD | FTC errors | Software issue | Reloaded software version 8.9. Machine loaded software on reboot. Loaded solution pack and reconnected system to radiometer live connect. Machine operational for use. | SWO-295462 | Defect | SW | RMED SW | RMED SW |  | FTC errors | Software issue | Reloaded software version 8.9. Machine loaded software on reboot. Loaded solution pack and reconnected system to radiometer live connect. Machine operational for use.
No | Visit | Complaint | RLTD | Front cover unit replacement | Front cover replacement | Unit broken at side clip and cartridge window cover. Replaced entire front cover. Tested passed. | SWO-295463 | Defect | HW | Cover and Chassis | Cartridge door |  | Front cover unit replacement | Front cover replacement | Unit broken at side clip and cartridge window cover. Replaced entire front cover. Tested passed.
No | Visit | Complaint | RSAF | Intermittant scan error | Check scanner Solution pack empty | Replce solution pack. Clean scanner on the outside. Check laser position checkout. Remove scanner and clean . Load slices and check. Scan all slices and everything that is on the wheel , in the service program. Load more slices , all scan . | SWO-296035 | Cleaned | HW | Storage Module | Barcode reader, Storage wheel |  | Intermittant scan error | Check scanner Solution pack empty | Replce solution pack. Clean scanner on the outside. Check laser position checkout. Remove scanner and clean . Load slices and check. Scan all slices and everything that is on the wheel , in the service program. Load more slices , all scan .
Yes | Visit | Complaint | RDEU | 12.01.2018, Sr. Bianca, Tel: 03518566312: LQC CRP außer Bereich, letzte Woche auch schon. | Siehe SWO | Geliefertes QC-Material unter Einhaltung der Kühlkette getestet --> CRP zu hoch --> Problem mit Station besprochen | SWO-296063 | [5 empty] | 12.01.2018, Sr. Bianca, Tel: 03518566312: LQC CRP out of range, last week already. | See SWO | Supplied QC material tested in compliance with the cold chain --> CRP too high --> Problem discussed with station
No | Visit | Complaint | RAME | User called stating they keeping having to replace the SP or SC on analyzer. | Found that they are getting inhomogeneous rinse cal and sample errors. | Replaced the wet section and SP ID pcb on analyzer as well as back flushed and had to do software upgrade on it. | SWO-296470 | Defect | HW | Basic Unit | PCB, Wet section Control |  | User called stating they keeping having to replace the SP or SC on analyzer. | Found that they are getting inhomogeneous rinse cal and sample errors. | Replaced the wet section and SP ID pcb on analyzer as well as back flushed and had to do software upgrade on it.
No | Visit | Complaint | RAME | Solution pack failure for in homogeneous rinse,cal3,cal2 | Service test results indicated solution pack failure | Replaced solution pack twice and finished with cassette | SWO-296555 | Assisted Replacement | SC | SC | FT |  | Solution pack failure for in homogeneous rinse,cal3,cal2 | Service test results indicated solution pack failure | Replaced solution pack twice and finished with cassette
No | Visit | Complaint | RDEU | 17.01.2018 Frau Krause Reinbot 01702245229 meldet, dass fehlerhafte QCs keine Auswirkung auf den Ampelstatus haben. | Qualitätskontrollfehler haben keinen Einfluss auf den Analysatorstatus. Managerpasswort 123456 hat keine Funktion mehr. | Konfiguration angepasst. Benutzer: Manager wurde gelöscht - wieder eingepflegt. QC Lvl 1 und Lvl 2 bei pO2 Fehlerhaft - wird zu hoch gemessen. LP erneuert. QC Lvl 2 jetzt Zielwertig aber Lvl 3 weiterhin zu hoch. | SWO-296692 | Setup | SW | User access | User access |  | 17.01.2018 Ms. Krause Reinbot 01702245229 reports that faulty QCs have no effect on traffic light status. | Quality control errors do not affect the analyzer status. Manager password 123456 no longer has any function. | configuration. User: Manager has been deleted - re-maintained. QC Lvl 1 and Lvl 2 at pO2 Faulty - is measured too high. LP renewed. QC Lvl 2 now targetd but Lvl 3 still too high.
Yes | 2 | Y | RPAC | Blank | Soultion pack failure - inhomogeneous QC |  | D3C1A80BC8412EBC | [5 empty] | Blank | Soultion pack failure - inhomogeneous QC
No | 5 | Y | RPOL | Blank | Diagnoza problemu. Wymiana wygiêtej ig³y aspiracyjnej. Regulacja ustawienia modu³u Inlet. Upgrade oprogramowania. Testy serwisowe. Analizator sprawny. |  | D3C1B7F3C84F46F0 | Defect | HW | Auto Inlet module | Probe bend |  | Blank | Diagnosis of the problem. Replacement of bent aspiration. Adjusting the Inlet module setting. Software Upgrade. Service tests. Functional Analyzer.
No | 1 | Y | RCZE | Blank | Milanove demo pre Nitru Po nastartovani Milanom doma opakovane status code 2 |  | D3C1E723C8030AAC | Defect | SW | Windows | Windows | Status Code 2 | Blank | Milanove Demo for Nitru after the start of Milan at home repeatedly status code 2
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 无法读取测试卡智能芯片 | 更换测试卡后修复 | SWO-295910 | [5 empty] | Test card, reagent package failure | Failed to read test card smart chip | Repair after replacing test card
Yes | Visit | Complaint | RIND | SC Not installing. | Sensor board to replace. | Sensor board replaced. Zeroing calibration done. Analyser Serviced and calibrated properly. Handed over in a fine working condition. | SWO-296897 | [5 empty] | SC Not installing. | Sensor board to replace. | Sensor board replaced. Zeroing calibration done. Analyser Serviced and calibrated properly. Handed over in a fine working condition.
Yes | Visit | Complaint | RIND | Lactate calibration error. | Check the calibration logs. | Re-Calibrate SC80 & Replace sc80 after continuous failure. new sc80 calibration pass. | SWO-297169 | [5 empty] | Lactate calibration error. | Check the calibration logs. | Re-Calibrate SC80 & Replace sc80 after continuous failure. new sc80 calibration pass.
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | LAC,GLU定标失败，0377,0210,0378 | 电话指导更换测试卡后未修复，工程师去现场处理，更换第二张测试卡后修复 | SWO-296531 | Assisted Replacement | SC | SC | Glu/Lac |  | Test card, reagent package failure | LAC, GLU calibration failure, 0377,0210,0378 | Phone guidance replacement test card after not repaired, engineer sits on site processing, replace the second test card after repair
No | Visit | Complaint | RKK | 測定中tCpO2とtcpCO2フリーズしたように数値が変化しないとのこと。念のため機器引上げ、データを確認するも、小数点以下の数値の変動がどのデータも見られ正常であると思われる。次回発生時は、一度大気を測定し数値変化があるか確認してもらう流れ。 | ケース詳細の通り | データ検証の結果、問題ないと判断する。TCチームからの解析と回答。 | SWO-295861 | [5 empty] | It is said that the numerical value does not change as tCpO2 and tcpCO2 freeze during measurement. Just in case, the equipment is raised and the data is checked, but the variation in the number of decimal places seems to be normal to see any data. The next time it occurs, the flow is to measure the atmosphere once and to check whether there is a numerical change. | As the case details | As a result of data validation, it is determined that there is no problem. Analysis and answers from tc team.
No | Visit | Complaint | RCHN | 报错1029 | 报错1029，借903-364一个 | 更换903-364后修复(By Ran) | SWO-295241 | Defect | HW | Process Module | Optical Unit | 1029 | 1029 error | 1029, loan 903-364 one | Repair after replacement 903-364 (By Ran)
No | Visit | Complaint | RCHN | 报错1031 | 报错1031，借903-364一个 | 更换903-364备件后修复(By Ran) | SWO-295465 | Defect | HW | Process Module | Optical Unit | 1031 | 1031 | 1031, loan 903-364 one | Repair after replacement of 903-364 spare parts (By Ran)
No | Visit | Complaint | RIND | Instrument showing intervention require. | found Gas-1 empty and inlet gasket have problem. | replaced the inlet gasket and Gas -1 and did 2-point calibration ,QC and Samples. All results found with in the range found instrument working satisfactorily. | SWO-297219 | Assisted Replacement | Consumables | Gas Cylinder | Gas Cylinder |  | Instrument showing intervention require. | found Gas-1 empty and inlet gasket have problem. | replaced the inlet gasket and Gas -1 and did 2-point calibration ,QC and Samples. All results found with in the range found instrument working satisfactorily.
No | Visit | Complaint | RDEU | Das Anmeldefenster überdeckt den Analysator/Statusfeld | Anmeldefenster überdeckt den Statusampel und ist nicht beim vorbeigehen ersichtlich. | Ananyme Zugriff ist gestatt als Option1 . Option 1 eins ist konfiguriert mit einem Statusampel. Bildschirmschoner ist eingeschaltet mit eine Minute. Einzige Nachteil ist dass, man für die Anmeldungsfenster auf dem Bildschirm zu bekommen, die Anmelde- Icon betätigen muss. Anwender sind mit dieser Lösung einverstanden. | SWO-296268 | Defect | SW | Windows | Win Log On | Techline raised | The login window covers the analyzer/status field | Login window covers the status light and is not visible when passing by. | Ananyme access is granted as option1 . Option 1 one is configured with a status light. Screen saver is turned on with one minute. The only drawback is that you have to press the login icon for the login window on the screen. Users agree with this solution.
No | Visit | Complaint | RCHN | 氖灯光强太低 | 氖灯光强太低，借902-833,902-458各一个 | 更换902-458一个正常。(By Ran) | SWO-295170 | Defect | HW | Oxi module | Lamp unit | Halogen low | The light is too low. | The light is too low, borrow 902-833,902-458 each | Replace 902-458 one normal. (By Ran)
No | Visit | Complaint | RAME | Capillary handle is loose. The pin holding the capillary handle stays in the capillary handle when removed. | Capillary handle loose. Mounting pin stays in the capillary handle when removed. | Replaced the inlet module chassis. | SWO-296207 | Defect | HW | Inlet module | Cover |  | Capillary handle is loose. The pin holding the capillary handle stays in the capillary handle when removed. | Capillary handle loose. Mounting pin stays in the capillary handle when removed. | Replaced the inlet module chassis.
No | Visit | Complaint | RAME | Correlation Issues | Tbili was higher on samples compared to other analyzers | Adj the lamp on the two new analyzers both were around 3200. Then ran thb cal in service and customer. Now the customer will run samples | SWO-296543 | Adjustment | HW | Oxi module | Lamp unit |  | Correlation Issues | Tbili was higher on samples compared to other analyzers | Adj the lamp on the two new analyzers both were around 3200. Then ran thb cal in service and customer. Now the customer will run samples
No | Visit | Complaint | RAME | The customer states the analyzer showed temperatures errors and, it the. Shutdown. Won't power up anymore. | Follow up service visit to replaced the battery module. | On my last visit from 12 $19 $17 I had found the battery module causing the analyzer to reboot and, I had to leave the battery module disconnected in order for the analyzer to work properly. We had also left this analyzer working as a bench top instrument in the ICU2 area. Today, I visited to replace the battery module. After replacement, I monitored the battery via service mode and, found it to be charging properly and, each cell had the correct voltages. The battery started with 88% charged and, it gradually moved to 100% indicating it is fully charged. Note: Respiratory department needs to keep this analyzer working as a bench top analyzer so it will stay in this area. The rolling stand belonging to this analyzer will be used for the ER ABL90 so that it is used a a mobile analyzer. The respiratory department is planning to order additional rolling stands. | SWO-297047 | Defect | HW | Battery module | Battery |  | The customer states the analyzer showed temperatures errors and, it the. Shutdown. Won't power up anymore. | Follow up service visit to replaced the battery module. | On my last visit from 12 $19 $17 I had found the battery module causing the analyzer to reboot and, I had to leave the battery module disconnected in order for the analyzer to work properly. We had also left this analyzer working as a bench top instrument in the ICU2 area. Today, I visited to replace the battery module. After replacement, I monitored the battery via service mode and, found it to be charging properly and, each cell had the correct voltages. The battery started with 88% charged and, it gradually moved to 100% indicating it is fully charged. Note: Respiratory department needs to keep this analyzer working as a bench top analyzer so it will stay in this area. The rolling stand belonging to this analyzer will be used for the ER ABL90 so that it is used a a mobile analyzer. The respiratory department is planning to order additional rolling stands.
Yes | Visit | Complaint | RDEU | Kunde hat Probleme mit Instand Ringversuch beim Glu | Kundin hatte im Vorfeld Probleme mit einer Sensor Kassette. Aktuelle alles OK. Instand Ergebnisse müssen abgewartet werden. | Gerät kalibriert, QC durchgeführt, Ringversuch gemäß SOP durchgeführt. Radiometer Service Konzept (Störungsmeldungen an die Hotline) mit Kunden besprochen. | SWO-297079 | Re classified RMED | Courtesy |  |  |  | Customer has problems with instand ring test at Glu | Customer had problems with a sensor cassette beforehand. Current everything OK. In-repair results must be waited for. | Device calibrated, QC performed, ring test performed according to SOP. Radiometer Service Concept (fault reports to the hotline) discussed with customers.
Yes | 11 | Y | RPAC | Blank | SP |  | D3C2F2BF3E3CE97C | [5 empty] | Blank | SP
No | Visit | Complaint | RIND | Received complaint that PO2 Parameter is red. | PO2 Parameter was red. PO2 Value was incorrect. | First Check Calibration Report to analyse status. after that remove SP and reinstall then remove SC and then flush with distilled water then reinstall. after that give 2 point calibration. after this process PO2 Parameter was green. and status was ok. | SWO-297046 | [5 empty] | Received complaint that PO2 Parameter is red. | PO2 Parameter was red. PO2 Value was incorrect. | First Check Calibration Report to analyse status. after that remove SP and reinstall then remove SC and then flush with distilled water then reinstall. after that give 2 point calibration. after this process PO2 Parameter was green. and status was ok.
No | Visit | Complaint | RAME | Analyzer will not connect to RALS | Troubleshot to a corrupt ASTM file | Replaced the ASTM file, and this resolved the problem | SWO-297118 | [5 empty] | Analyzer will not connect to RALS | Troubleshot to a corrupt ASTM file | Replaced the ASTM file, and this resolved the problem
No | Visit | Complaint | RCHN | 打印机故障 | 机器未修复，之前借的902-985装上后，打印几次后，又故障了，现要借903-001,902-985各一个 | 更换了903-001,902-985备件后修复(By Ran) | SWO-295883 | [5 empty] | Printer failure | The machine has not been repaired, the previous lying 902-985 installed, printed several times, and then failed, now to borrow 903-001, 902-985 each one | Repair after replacement of 903-001,902-985 spare parts (By Ran)
No | Visit | Complaint | RCHN | 抽吸样本失败 | 机器未修复，备件903-199,903-644未使用，现要加借903-129一个 | 更换903-129后修复(By Ran) | SWO-294411 | Defect | HW | Basic Unit | Pinch Valve |  | Failed to pump samples | The machine has not been repaired, spare parts 903-199,903-644 is not used, now to borrow 903-129 one | Repair after replacement 903-129 (By Ran)
No | Visit | Complaint | RCHN | lac漂移错误 | lac漂移错误，借903-581一个 | 更换903-581后修复(By Ran) | SWO-295264 | Defect | HW | Basic Unit | PCB, Sensor interface |  | lac drift error | lac drift error, borrow 903-581 one | Repair after replacement 903-581 (By Ran)
No | Visit | Complaint | RESP | 1169 no se pueden bombear soluciones. | 1169 no se pueden bombear soluciones. Atasco en el SCI. Tras el cliente intentar realizar, con ayuda telefonica por Servicio tecnico, el procedimiento de limpieza manual interno del equipo sin exito debido a una excesiva obstruccion, acudimos a revisar el problema. y se detecto que el SCI era el que ocasionaba el problema. | Se sustituyo SCI y se comprobo correcto funcionamiento del equipo. | SWO-299680 | Assisted Replacement | SC | SC | Clot |  | 1169 solutions cannot be pumped. | 1169 solutions cannot be pumped. Jam in the SCI. After the customer attempts to carry out, with telephone help by Technical Service, the procedure of internal manual cleaning of the equipment without success due to excessive obstruction, we go to review the problem. and it was detected that the SCI was the one causing the problem. | SCI is replaced and the equipment is compatible with proper operation.
No | Visit | Complaint | RCHN | TNI定标失败 | TNI定标失败，需要借件944-258（13719）缺货、903-364一个 | 更换903-364和重新定标后修复(By Ran) | SWO-295950 | Defect | HW | Process Module | Optical Unit | NT-PRO BnP | TNI calibration failed | TNI bid failed, need to borrow 944-258 (13719) out of stock, 903-364 one | Replacement 903-364 and re-calibration repair (By Ran)
No | Visit | Complaint | RCHN | 试管夹损坏 | 试管夹损坏,需要借件903-584一个 | 更换903-584后修复(By Ran) | SWO-295954 | Defect | HW | Aspiration Module | Tube holder |  | Damage to test tube clip | The test tube clip is damaged and requires a loan of 903-584 | Repair after replacement 903-584 (By Ran)
No | Visit | Complaint | RCHN | 管夹断裂 | 管夹断裂，借903-584一个 | 更换903-584后修复(By Ran) | SWO-295931 | Defect | HW | Aspiration Module | Tube holder |  | Pipe clip break | The tube clip is broken, borrowed 903-584 one | Repair after replacement 903-584 (By Ran)
No | Visit | Complaint | RIND | Na K values too high. | Na K values too high. | Protean removal done. Cleaning done. Electrode cleaning done. Chamber cleaning done. Reference membrane replaced. Analyser Serviced and calibrated properly. Handed over in a fine working condition. | SWO-297076 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | Na K values too high. | Na K values too high. | Protean removal done. Cleaning done. Electrode cleaning done. Chamber cleaning done. Reference membrane replaced. Analyser Serviced and calibrated properly. Handed over in a fine working condition.
No | Visit | Complaint | RIND | electrolyte values problem | electrolytes values problem | checked the ref electrode puss formation was found then cleaned it after that cleaned the chamber , observed few sample all are ok working fine | SWO-296869 | Cleaned | Fluidic pathway | Fluidic pathway | Measuring chamber |  | electrolyte values problem | electrolytes values problem | checked the ref electrode puss formation was found then cleaned it after that cleaned the chamber , observed few sample all are ok working fine
Yes | 3 | Y | RCZE | Blank |  |  | D3C2CD9F3E5CB3DD | Re classified RMED | FAN |  |  |  | Blank
Yes | 5 | Y | RCZE | Blank |  |  | D3C2CD043E10B3D9 | Re classified RMED | FAN |  |  |  | Blank
No | Visit | Complaint | RAME | DMS not connected to AU | 0643 ABL $DMS PC disconnected from AU error is present. Customer has rebooted but error persists. | Saved customer data and performed a complete software install of current software version, 6.18. Upon software install, DMS not connected error was resolved. Performed the necessary calibrations to get analyzer to a usable status. Analyzer is within acceptable operating specifications. | SWO-297091 | Defect | SW | RMED SW | Version 6.18 |  | DMS not connected to AU | 0643 ABL /DMS PC disconnected from AU error is present. Customer has rebooted but error persists. | Saved customer data and performed a complete software install of current software version, 6.18. Upon software install, DMS not connected error was resolved. Performed the necessary calibrations to get analyzer to a usable status. Analyzer is within acceptable operating specifications.
No | Visit | Complaint | RAME | Oxi parameters failing level 3 of VK kit | Visited 01 $16 $18 Upon arriving I realized I was one day early on the cal verification lockout. Returned 01 $19 $18 Returned too site with Oxi module and linearity kits. | On 01 $16 I tried a thb cal followed by the level 3 from the vk kit. I still receive 581 errors on the level 3. Will have to return on 1 $18 $18. Also hemolyzers that were ordered overnight did not arrive due to the weather. 01 $19 Attempted the level 3 from a new lot linearity kit and found no change. Still receiving 581 errors. Replaced oxi module and checked lamp adjustment. Adjusted lamp slightly and ran service and user THB cals, found all ok. Ran level 3 of the VK kit and no longer have 581 errors. | SWO-298911 | Defect | HW | Oxi module | Oxi module, Replaced | 581 | Oxi parameters failing level 3 of VK kit | Visited 01 /16 /18 Upon arriving I realized I was one day early on the cal verification lockout. Returned 01 /19 /18 Returned too site with Oxi module and linearity kits. | On 01 $16 I tried a thb cal followed by the level 3 from the vk kit. I still receive 581 errors on the level 3. Will have to return on 1 $18 $18. Also hemolyzers that were ordered overnight did not arrive due to the weather. 01 $19 Attempted the level 3 from a new lot linearity kit and found no change. Still receiving 581 errors. Replaced oxi module and checked lamp adjustment. Adjusted lamp slightly and ran service and user THB cals, found all ok. Ran level 3 of the VK kit and no longer have 581 errors.
Yes | 5 | Y | RCZE | Blank |  |  | D3C2CC013E00B3D6 | Re classified RMED | FAN |  |  |  | Blank
Yes | 5 | Y | RCZE | Blank |  |  | D3C2CD593E23B3DB | Re classified RMED | FAN |  |  |  | Blank
No | Visit | Complaint | RAME | Sample error 331 | Vérifier liquid sensor et pompe flush waste line | Remplacé probe et probe lock | SWO-297078 | Defect | HW | Inlet module | Probe |  | Sample error 331 | Check liquid sensor and flush waste line pump | Replaced probe and probe lock
No | Visit | Complaint | RCHN | 溶血器故障 | 溶血器故障 | 更换902-833后修复(By Ran) | SWO-297163 | Defect | HW | Oxi module | Hemolyzer |  | Hemolysis failure | Hemolysis failure | Repair after replacement 902-833 (By Ran)
No | Visit | Complaint | RAME | Issues with tbili | Tbili issues with the three analyzers | I adj the lamp and ran thb cal. The lamp was around 3200 | SWO-296544 | Adjustment | HW | Oxi module | Lamp unit |  | Issues with tbili | Tbili issues with the three analyzers | I adj the lamp and ran thb cal. The lamp was around 3200
No | Visit | Complaint | RAME | Neon # too low | Neon # too low. Perform complete OXI test. Neon position is .104, too low. | Replaced Spectrometer. Performed complete OXI alignment. Neon position is now .527. | SWO-296902 | Defect | HW | Oxi module | Spectrometer | Neon to low | Neon # too low | Neon # too low. Perform Complete OXI test. Neon position is. 104, too low. | Replaced Spectrometer. Performed complete OXI alignment. Neon position is now .527.
No | Visit | Complaint | RSAF | Date format is incorrect currently month, day, year and should be day, month, year. | incorrect date format | Changed date format in regional settings and committed to the C drive, restarted analyzer and all ok. | SWO-296720 | User error | SW | Date format | Date format |  | Date format is incorrect currently month, day, year and should be day, month, year. | incorrect date format | Changed date format in regional settings and committed to the C drive, restarted analyzer and all ok.
No | Visit | Complaint | RDEU | 19.01.2018 8:00 Sr. Jaqueline 03905482943: Reagenzienpack kann nicht entnommen werden nach Stromausfall | Nach einem Stromausfall werden keine Kalibrierungen erkannt | Gerät geprüft, Kassetten neu gescannt, Systemreinigung durchgeführt, CRP QC durchgeführt, Funktionskontrolle ok. Die Verwendung einer USV wird empfohlen. | SWO-296865 | Defect | SW | Power Supply | Power Failure (no UPS) |  | 19.01.2018 8:00 Sr. Jaqueline 03905482943: Reagent pack cannot be removed after power failure | No calibrations are detected after a power failure | Device checked, cassettes rescanned, system cleaning performed, CRP QC performed, function control ok. It is recommended to use a UPS.
No | Visit | Complaint | RCHN | 定标失败，报错2877和5694 | 定标失败，报错2877和5694，需要借件944-394（13715） |  | SWO-297172 | Defect | Consumables | Test Kit | Cartridge, Broken | 5694 | Failed to bid, error 2877 and 5694 | Failed bid, error 2877 and 5694, need to borrow 944-394 (13715)
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | Lac定标不过 0378 | 更换测试卡后修复 | SWO-299313 | Assisted Replacement | SC | SC | Lac | 0378 | Test card, reagent package failure | Lac calibration is no more than 0378 | Repair after replacing test card
No | Visit | Complaint | RCHN | glu定标超出范围，报错0378 | glu定标超出范围，报错0378，需要借件903-183一个 |  | SWO-297261 | Defect | HW | Basic Unit | PCB, Sensor interface |  | glu calibration out of range, error 0378 | glu calibration out of range, error 0378, need to borrow 903-183 one
No | Visit | Complaint | RCHN | 进样口支架损坏 | 进样口支架损坏，借2个903-584 | 更换903-584后修复(By Ran) | SWO-296730 | Defect | HW | Aspiration Module | Tube holder |  | Damage to the inlet bracket | Damaged in-in stand, borrow 2 903-584 | Repair after replacement 903-584 (By Ran)
No | Visit | Complaint | RCHN | 报1031 | 报1031，借903-364一个 | 更换903-364后修复(By Ran) | SWO-296533 | Defect | HW | Process Module | Optical Unit | 1031 | Reported 1031 | Reported 1031, loan 903-364 a | Repair after replacement 903-364 (By Ran)
No | 9 | Y | RPAC | Blank | Clot issues |  | D3C70068A0243315 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway | Clot | Blank | Clot issues
No | Visit | Complaint | RAME | Customer called with the ABL 90 disconnecting from the wireless | Analyzer is intermittently disconnecting from the wireless. | Changed the power settings and discussed possible wireless issues. We will continue to monitor with Jim and the IT department. Pulled activity logs for investigation. | SWO-297826 | Revisit required | His/Lis | His/Lis | His/Lis |  | Customer called with the ABL 90 disconnecting from the wireless | Analyzer is intermittently disconnecting from the wireless. | Changed the power settings and discussed possible wireless issues. We will continue to monitor with Jim and the IT department. Pulled activity logs for investigation.
No | Visit | Complaint | RAME | Wireless connectionis sporadically dropping. | The Asus dongle for wireless network keeps shutting down and, disconnected from the wireless network. | Checked the wireless network adapter wifi screen and, found the analyzer is not connected to the network as the ASUS dongle was turned off. Logged in to windows to recheck the USB power saving settings. Found the power saving settings 'Enabled', this time I disabled this feature again and, performed the commit command to ensure the settings stay in effect. I restarted the analyzer and, then reentered in to windows to reconfirm the USB power saving settings are not enabled again. | SWO-297867 | Revisit required | His/Lis | His/Lis | His/Lis |  | Wireless connectionis sporadically dropping. | The Asus dongle for wireless network keeps shutting down and, disconnected from the wireless network. | Checked the wireless network adapter wifi screen and, found the analyzer is not connected to the network as the ASUS dongle was turned off. Logged in to windows to recheck the USB power saving settings. Found the power saving settings 'Enabled', this time I disabled this feature again and, performed the commit command to ensure the settings stay in effect. I restarted the analyzer and, then reentered in to windows to reconfirm the USB power saving settings are not enabled again.
No | Visit | Complaint | RAME | Solution pack failure for in homogeneous rinse,cal3,cal2 | Received a call from FSE on call stating that customer complained that the pO2 is outside of range on levels 1 and 3. There was also mention of inhomogeneous QC solution errors. Once on-site, analyzer issues are confirmed. | Multiple solution packs were replaced from the customer's stock, which lead to various fluid flow errors. The sensor cassette was replaced and the analyzer then became functional, showing a fully green parameter bar. Suspected silicone contamination occurred. | SWO-298197 | Assisted Replacement | SC | SC | FT | Silicone contamination | Solution pack failure for in homogeneous rinse,cal3,cal2 | Received a call from FSE on call stating that customer complained that the pO2 is outside of range on levels 1 and 3. There was also mention of inhomogeneous QC solution errors. Once on-site, analyzer issues are confirmed. | Multiple solution packs were replaced from the customer's stock, which lead to various fluid flow errors. The sensor cassette was replaced and the analyzer then became functional, showing a fully green parameter bar. Suspected silicone contamination occurred.
Yes | 3 | Y | RPAC | Blank | Calibration errors present for pCO2. |  | D3C6B366A005B91C | [5 empty] | Blank | Calibration errors present for pCO2.
No | Visit | Complaint | RAME | Network issues. Duplicate case. | As recommended by our technical support team, I revisit to check the analyzer IP configuration. | I logged in to windows and, performed command prompts to test the IP configuration. I performed an ipconfig all, then I release ipconfig and, then renewed ipconfig. During my checks, I confirmed the IP address to this analyzer did not change. As per our technical support team recommendations, I will ask Susana to contact her IT department for help in setting up a temporary static IP address for this analyzer to check if the analyzer continues to disconnect. We will need to use the static IP address to monitor the analyzer connection to the network. | SWO-297868 | [5 empty] | Network issues. Duplicate case. | As recommended by our technical support team, I revisit to check the analyzer IP configuration. | I logged in to windows and, performed command prompts to test the IP configuration. I performed an ipconfig all, then I release ipconfig and, then renewed ipconfig. During my checks, I confirmed the IP address to this analyzer did not change. As per our technical support team recommendations, I will ask Susana to contact her IT department for help in setting up a temporary static IP address for this analyzer to check if the analyzer continues to disconnect. We will need to use the static IP address to monitor the analyzer connection to the network.
No | Visit | Complaint | RAME | Liquid sensor calibration failure | Liquid sensor failure. Mary called this in last week but due to inclement weather said it was OK to come in today. Liquid sensor failing in the service mode. All OXI passing. I found the hemolyzer leaking at the plastic nipple. Will replace hemolyzer and liquid sensor. Hemolyzer and liquid sensor still under warranty. Calibrated display touch screen. | Replaced hemolyzer and liquid sensor. Ran liquid sensor cal and passed. Ran cals, tHb cal and QC. | SWO-297533 | Defect | HW | Oxi module | Hemolyzer |  | Liquid sensor calibration failure | Liquid sensor failure. Mary called this in last week but due to inclement weather said it was OK to come in today. Liquid sensor failing in the service mode. All OXI passing. I found the hemolyzer leaking at the plastic nipple. Will replace hemolyzer and liquid sensor. Hemolyzer and liquid sensor still under warranty. Calibrated display touch screen. | Replaced hemolyzer and liquid sensor. Ran liquid sensor cal and passed. Ran cals, tHb cal and QC.
No | Visit | Complaint | RCHN | HCT模块故障 | HCT模块故障，借903-329,636-561各一个 | 更换了903-329,636-561后修复 by Andrew | SWO-297395 | Defect | HW | Aspiration Module | Needle Wash Unit |  | HCT module failure | HCT module failure, one each by 903-329, 636-561 | Repair edited by Andrew after replacing 903-329, 636-561
Yes | Visit | Complaint | RAME | Analyzer wont pass leakage test. | Found analyzer in hold upon arriving and leaking out the front of the EL $Met module. Examined electrodes and found no seal ring on the Glucose electrode. | Replaced seal ring and restarted, analyzer passes leakage test but still leaks from the EL $Met module. Replaced the K membrane, still leaks,notice rinse not flowing all the way to the bottom of the chamber. Replaced the NA membrane as this was the point the fluid was coming out of the chamber. Once replaced the leak had stopped. | SWO-299159 | [5 empty] | Analyzer wont pass leakage test. | Found analyzer in hold upon arriving and leaking out the front of the EL /Met module. Examined electrodes and found no seal ring on the Glucose electrode. | Replaced seal ring and restarted, analyzer passes leakage test but still leaks from the EL $Met module. Replaced the K membrane, still leaks,notice rinse not flowing all the way to the bottom of the chamber. Replaced the NA membrane as this was the point the fluid was coming out of the chamber. Once replaced the leak had stopped.
No | Visit | Complaint | RAME | Methb and cohb reading too high compared to other analyzers | Cohb and methb outside peer 2 group high | Replaced hemolyzer assembly,complete oxi test , service thb cal. | SWO-297816 | Defect | HW | Oxi module | Hemolyzer |  | Methb and cohb reading too high compared to other analyzers | Cohb and methb outside peer 2 group high | Replaced hemolyzer assembly, complete oxygen test, service THB Cal.
No | Visit | Complaint | RKK | Neon number too low / Neon Lamp intensity too low | Neon number too low Neon Lamp intensity too low | OXIモジュール交換 | SWO-298605 | Defect | HW | Oxi module | Oxi module, Replaced |  | Neon number too low / Neon Lamp intensity too low | Neon number too low Neon Lamp intensity too low | OXI module replacement
No | Visit | Complaint | RIND | NA Level 3 QC getting failed | Checked the data logs NA Level 3 QC getting failed found problem NA membrane | Replaced new Na membrane & performed QC again then also Na was getting failed need to run another level of QC & need to confirm the problem | SWO-300200 | Assisted Replacement | Electrode | Membrane | Na+ |  | NA Level 3 QC getting failed | Checked the data logs NA Level 3 QC getting failed found problem NA membrane | Replaced new Na membrane & performed QC again then also Na was getting failed need to run another level of QC & need to confirm the problem
No | Visit | Complaint | RDEU | Martin Suttner meldet kurzeitiges Problem, inkonsitente Software. Meldung zur Zeit nicht erichtlich | Überprüfung in Aktivitäten | Meldung: Oximodul nicht aktiv, inkonsistente Software. Zur Zeit ohne Beanstandung. AU Software nochmals installiert. Initialisierungsphase Kalibrationen durchgührt. tHb Kal QC ohne Beanstandung Fehler wird beobachtet | SWO-297620 | Defect | SW | RMED SW | Inconsistent SW | OXI | Martin Suttner reports short-term problem, inconfixed software. Notification not available at the moment | Review in activities | Message: Oximodul inactive, inconsistent software. At the moment without complaint. AU software is reinstalled. Initialization phase calibrations. tHb Kal QC without complaint Error is observed
No | Visit | Complaint | RESP | Erro de temperatura 1068 | Verificação de erro de temperatura de secagem com um valor de 63,6 graus. | Desmontagem do equipamento, substituição do material indicado, montagem, calibração do cup dryer, actualização do config, verificação do sistema com um valor de secagem de 70,1 e passagem do QC. Tudo ok. | SWO-297542 | Defect | HW | Process Module | PCB, Cup dryer | 1068 | Temperature Error 1068 | Check drying temperature error with a value of 63.6 degrees. | Disassembly of the equipment, replacement of the indicated material, assembly, calibration of the cup Dryer, config update, system check with a drying value of 70.1 and QC pass. Okay, all right.
Yes | Visit | Complaint | RDEU | AQT bricht Probenmessung ab | Datenbank geprüft Fehlermeldung : 2873;2877;3553 | Beim Eintreffen steht der AQT in messbereit--> Test im Serviceprogramm durchgeführt --> i.O div. Sensorkalibrierungen durchgeführt --> i.O 2 Testmessungen ( CRP ) --> i.O --> Testmessung mit TNT wird abgebrochen --> Testkassette TNT entfernt und neue Testkassette eingesetzt --> Testmessung ( TNT und CKMB ) i.O | SWO-297492 | [5 empty] | AQT cancels sample measurement | Database checked Error: 2873;2877;3553 | Upon arrival, the AQT is ready for measurement --> test in the service program --> i.O div. Sensor calibrations performed --> i.O 2 test measurements ( CRP ) --> i.O --> Test measurement with TNT is aborted --> Test cassette TNT is removed and new test cassette inserted --> Test measurement ( TNT and CKMB ) i.O
Yes | Visit | Complaint | RDEU | 19.01.2018; Fr. Funke; Tel.: 05313489900; Messabbrüche (Undefind lookopstring 12.-1 und Corrupt xml) | Proben werden abgebrochen | Siehe SWO der Halbjahreswartung | SWO-297548 | Re classified RMED | Preventive maintenance |  |  |  | 19.01.2018; Fr. Funke; Tel.: 05313489900; Messabbrüche (Undefind lookopstring 12.-1 und Corrupt xml) | Samples are aborted | See SWO of Half-Year Maintenance
Yes | Visit | Complaint | RLTD | Angela Sullivan Tel 01753 633464 | Following a visit from R&D it looks like the problems are being caused by silicon contamination from the blood gas syringes being used. | Radiometer syringes are going to be trailed and before doing that the analyser was flushed and a service kit fitted to remove any silicon contamination present. | SWO-298442 | Re classified RMED | Doc of RMED R&D |  |  |  | Angela Sullivan Tel 01753 633464 | Following a visit from R&D it looks like the problems are being caused by silicon contamination from the blood gas syringes being used. | Radiometer syringes are going to be trailed and before doing that the analyser was flushed and a service kit fitted to remove any silicon contamination present.
No | Visit | Complaint | RDEU | 23.01.2018, 11:04 Uhr ==> Anrufer Fr. Frahm, Tel.: 04523405211 ==> Fehler 0292: THb-Messfehler bei Patientenproben | THB Messfehler. | THB Kalibration durchgeführt kein Fehler feststellbar. Analysator zerlegt und Verschlauchung geprüft und gereinigt. Sollte der Fehler weiterhin auftreten. Sollte der Hemolyzer getauscht werden. Der Fcuv Wert liegt bei 0,968. Beim hochfahren des Analysator ist Flüssigkeit aus der SC gelaufen. Analysator wieder zerlegt und wieder Fließwege kontrolliert. Wieder hochgefahren und wieder lief Flüssigkeit aus der SC aus. Ablauf an der SC durchgespült. Danach Analysator wieder hochgefahren. Diesmal ohne Fehler. Kalibration und QC abgewartet. System OK. | SWO-298216 | [5 empty] | 23.01.2018, 11:04 ==> Caller Ms. Frahm, Tel.: 04523405211 ==> Error 0292: THb measurement error in patient samples | THB measurement error. | THB calibration performs no error detectable. Analyzer disassembled and hose tested and cleaned. If the error persists. Should the Hemolyzer be exchanged. The Fcuv value is 0.968. When the analyzer was booted up, liquid ran out of the SC. Analyzer disassembled again and controlled flow paths again. Re-started and again liquid leaked out of the SC. drain at the SC. Then the analyzer is powered up again. This time without errors. calibration and QC. System OK.
No | Visit | Complaint | RCHN | 传感器错误 | 传感器错误。需要借件903-581一个 |  | SWO-297935 | Defect | HW | Basic Unit | PCB, Sensor interface |  | Sensor error | The sensor is wrong. Need to borrow 903-581 one
No | Visit | Complaint | RCHN | 报错1999 | 报错1999，借903-117一个 |  | SWO-298247 | Defect | HW | Analyzer Control Module | PCB, Main Board | 1999 | Error 1999 | Misquoted 1999, loan 903-117 one
No | Visit | Complaint | RLTD | Email received from Kirandeep | Upgraded monitor software from v1.0.2 to latest v1.1.8 Customer verified monitor operation including download to external software - all ok. | Upgraded software to v1.1.8 | SWO-298376 | [5 empty] | Email received from Kirandeep | Upgraded monitor software from v1.0.2 to latest v1.1.8 Customer verified monitor operation including download to external software - all ok. | Upgraded software to v1.1.8
Yes | 30 | Y | RCZE | Blank |  |  | D3C864C816160CDC | Re classified RMED | FAN |  |  |  | Blank
Yes | 15 | Y | RCZE | Blank |  |  | D3C8646E16450CDA | Re classified RMED | FAN |  |  |  | Blank
No | Visit | Complaint | RIND | Service Visit | Calibration sensitivity of pCo2 is not in range. | Keep on hold the analyzer, open the chamber of pCO2 and clean it, remove the electrode from its membrane wipe out with deionized water and fix it same as before. Record no replacement in analyzer and restart the analyzer, wait for auto two point calibration successfully done. Now analyzer handed over to staff in well working condition. | SWO-300342 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | Service Visit | Calibration sensitivity of pCo2 is not in range. | Keep on hold the analyzer, open the chamber of pCO2 and clean it, remove the electrode from its membrane wipe out with deionized water and fix it same as before. Record no replacement in analyzer and restart the analyzer, wait for auto two point calibration successfully done. Now analyzer handed over to staff in well working condition.
Yes | 4 | Y | RCZE | Blank |  |  | D3C8656016040CDF | Re classified RMED | FAN |  |  |  | Blank
Yes | 2 | Y | RCZE | Blank |  |  | D3C862C016230CD6 | Re classified RMED | FAN |  |  |  | Blank
No | Visit | Complaint | RIND | Sample not aspirating. | Sample not aspirating. | Tubings flushed and cleaned, Inlet and inlet gasket cleaned. But same problem reported. Pump calibration failed for all. Pump tubing replaced from old analyser and found pump cal passed. Need to replace new pump tubing and inlet gasket. PM Needed. TO sending for PM Service. Temporarily Analyser Serviced and calibrated. Handed over in a fine working condition. | SWO-298400 | Defect | HW | Fluidic manifold | Pump tubes |  | Sample not aspirating. | Sample not aspirating. | Tubings flushed and cleaned, Inlet and inlet gasket cleaned. But same problem reported. Pump calibration failed for all. Pump tubing replaced from old analyser and found pump cal passed. Need to replace new pump tubing and inlet gasket. PM Needed. TO sending for PM Service. Temporarily Analyser Serviced and calibrated. Handed over in a fine working condition.
No | Visit | Complaint | RAME | Lactate correlation samples are higher than reference instrument at results above 4mmol | Lactate results above 4 mmol are higher than reference ABL. Zero current has not changed and overall instrument variables are normal. Prior membrane change did not improve bias. | Replaced lactate electrode & membrane. Testing of lactate | SWO-299709 | Assisted Replacement | Electrode | Electrode | Lac |  | Lactate correlation samples are higher than reference instrument at results above 4mmol | Lactate results above 4 mmol are higher than reference ABL. Zero current has not changed and overall instrument variables are normal. Prior membrane change did not improve bias. | Replaced lactate electrode & membrane. Testing of lactate
No | Visit | Complaint | RFRA | Cupules bloquées dans pack de solution | Cupules bloquées dans process Wheel => pack solution YV-57 | Déblocage des cupules pour qu’elles puissent tomber dans le pack solution Passage d’un tube cleaning solution OK $ $ passage CQ CRP level 2 trois fois d’affilié pour valider le bon passage des cupules de roue en roue OK | SWO-298739 | Defect | Consumables | Solution Pack | Cup Stack |  | Cupules stuck in solution pack | Cupules stuck in Process Wheel - YV-57 solution pack | Unlocking cups so they can fall into the solution pack Passing a tube cleaning solution OK $$ passage CQ CRP level 2 three times affiliate to validate the correct passage of wheel cups OK
No | Visit | Complaint | RAME | Unit is giving intervention required for fluid transport errors. | SP not initializing. Error 1084 inhomogenous cal3. | Replaced SP and ran 3Levels external QC. | SWO-299126 | Assisted Replacement | SP | SP | FT | 1084 | Unit is giving intervention required for fluid transport errors. | SP not initializing. Error 1084 inhomogenous cal3. | Replaced SP and ran 3Levels external QC.
No | Visit | Complaint | RDEU | 23.01.2018; Hr. Olbert; Tel.: 079618816706; Display ist ausgefallen | Schwarzer Bildschirm, Wetsection reagiert. | Monitor erneuert, Kein Erfolg. Ersatzgerät aus GS geholt und CPU Unit getauscht. Orginal - Bildschirm reagiert wieder normal. Neue CPU Unit muß bestellt und später eingebaut werden . Konfiguraton wieder hergestellt. | SWO-299920 | Revisit required | HW | Basic Unit | CPU |  | 23.01.2018; Mr. Olbert; Tel.: 079618816706; Display is down | Black screen, wetsection responds. | Monitor renewed, No success. Replacement device from GS and CPU unit exchanged. Orginal - Screen responds normally again. New CPU unit must be ordered and installed later. Configuraton.
No | Visit | Complaint | RAME | Display turns pink and purple when you adjust it. Oximetry error 1021 on QC. | The display is fine until you move it. There are a lot of oximetry error s on QC, 1021. | Removed the display and reseated all of the connections. Replaced the hemolyzer. | SWO-298629 | Defect | HW | Oxi module | Hemolyzer |  | Display turns pink and purple when you adjust it. Oximetry error 1021 on QC. | The display is fine until you move it. There are a lot of oximetry error s on QC, 1021. | Removed the display and reseated all of the connections. Replaced the hemolyzer.
No | Visit | Complaint | RKK | Tempエラー発生。センサーボードに腐食有ること確認。 | センサーボード腐食 | センサーボード交換 | SWO-298597 | Defect | HW | Basic Unit | PCB, Sensor connector assembly |  | Temp error occurred. Confirmed that there is corrosion on the sensor board. | Sensor board corrosion | Sensor board replacement
No | Visit | Complaint | RCHN | 血氧单元质控B、C无法通过 | 血氧单元质控B、C无法通过，借902-980和903-197各一个 | 更换903-197一个正常。 by Andrew | SWO-297934 | Defect | HW | Oxi module | Spectrometer |  | Blood oxygen unit quality control B, C can not pass | Blood oxygen unit quality control B, C can not pass, borrow 902-980 and 903-197 each one | Replace 903-197 one normal. by Andrew
No | Visit | Complaint | RCHN | 外部条码器不能读取条码 | 外部条码器不能读取条码，借902-924一个 | 更换备用件902-924后修复 | SWO-297898 | Defect | HW | Analyzer Control Module | Barcode reader, display |  | External barcode stors cannot read barcodes | External barcode scan cannot read barcodes, borrow 902-924 a | Repair after replacing spare parts 902-924
No | Visit | Complaint | RCHN | 定标失败 | 定标失败，需要借件944-220（13764） | Only recalibration to solve the problem. by Andrew | SWO-297924 | Adjustment | Consumables | Test Kit | Cartridge, D-Dimer | 13764 | Bid failed | Failed bid, 944-220 (13764) required | Only recalibration to solve the problem. by Andrew
No | Visit | Complaint | RCHN | 定标中止 | 定标中止，借944-258（13465）一张 |  | SWO-298673 | Adjustment | Consumables | Test Kit | Cartridge, NT-PROBNP | 13465 | Bid abort | Bid suspension, borrow 944-258 (13465) one
No | Visit | Complaint | RIND | Oximewtry measuring error. | Oximetry measuring error. | Protean removal and cleaning program done. Oxymetri complete test done and passed. Analyzer serviced and calibrated properly. Handed over in a fine working condition. | SWO-299504 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | Oximewtry measuring error. | Oximetry measuring error. | Protean removal and cleaning program done. Oxymetri complete test done and passed. Analyzer serviced and calibrated properly. Handed over in a fine working condition.
No | 28 | Y | RPNZ | Blank | Glucose membrane not sealing properly. |  | D3C94E9B8C4CABDC | Defect | Electrode | Electrode | Sealing ring | Sealing ring, Glu | Blank | Glucose membrane not sealing properly.
Yes | 7 | Y | RPNZ | Blank | Semi Annual Service |  | D3C968858C504A11 | Re classified RMED | Preventive Maintenance |  |  |  | Blank | Semi Annual Service
No | Visit | Complaint | RIND | Service Visit | There is an leakage from the inlet gas kit during rinse. | Flush the inlet gas kit, inlet probe, primary tubes with the hypochlorite solution, clean the chamber of ph and ei chamber. Now connect all the part same as before, initiate the analyzer and give two point calibration which is successfully done. Analyzer handed over to staff in well working condition. | SWO-300350 | Revisit required | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | Service Visit | There is an leakage from the inlet gas kit during rinse. | Flush the inlet gas kit, inlet probe, primary tubes with the hypochlorite solution, clean the chamber of ph and ei chamber. Now connect all the part same as before, initiate the analyzer and give two point calibration which is successfully done. Analyzer handed over to staff in well working condition.
No | Visit | Complaint | RDEU | 24.01.2018, 09:06 Uhr ==> Anrufer Herr Schraven, Tel: 015174669714 ==> Konditionier- und Spülfehler EL/MET-Modul | Sichtkontrolle | Spüllösung ersetzt, Funktionskontrolle | SWO-299010 | Assisted Replacement | Consumables | Consumables | Rinse Solution |  | 24.01.2018, 09:06 ==> Caller Mr. Schraven, Tel: 015174669714 ==> Conditioning and flushing error EL/MET module | Visual inspection | Rinse solution replaced, function control
Yes | Visit | Complaint | RDEU | 12.01.2018, Sr. Bianca, Tel: 03518566312: LQC CRP außer Bereich, letzte Woche auch schon. | Siehe SWO | CRP-Test-Kit ( 13846 ) kalibriert --> QC CRP1 (13623)-Messung durchgeführt --> Wert zu hoch 49,7 mg $L Station ist informiert Servicedump erstellt | SWO-298767 | [5 empty] | 12.01.2018, Sr. Bianca, Tel: 03518566312: LQC CRP out of range, last week already. | See SWO | CRP test kit ( 13846 ) calibrated --> QC CRP1 (13623) measurement performed --> Value to high 49.7 mg $L Station is informed Service dump created
No | Visit | Complaint | RAME | Received call about analyzer failing level 3 external QC | Found level 3 external qc to be failing | Flushed and calibrated analyzer. Reseated solution pack. | SWO-299776 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | Received call about analyzer failing level 3 external QC | Found level 3 external qc to be failing | Flushed and calibrated analyzer. Reseated solution pack.
No | Visit | Complaint | RAME | Wireless connectionis sporadically dropping. | Susana has reported to me that the analyzer keeps loosing the wireless connection and, it takes very long for it to reconnect to the network again.also the analyzer is intermittently presenting a windows message NHSTW32.exe error. | In order to correct the connection issue, I adjusted the Roma.ini file with settings to enable the reconnect ability of the analyzer if the network connection is lost. I also made adjustments so that the reconnect interval is much sooner when the connection is lost. Lastly, I made the adjustments so that the reconnection retries are more frequent. In order to prevent the window message NHSTW32.exe error, I upgraded the software from version 3.3 to version 3.3MR1. Susana will be monitoring this instrument and, She will inform me about the connectivity performance of this analyzer. She will also keep an eye to confirm the windows message does not reappear. | SWO-299836 | Setup | His/Lis | His/Lis | His/Lis |  | Wireless connectionis sporadically dropping. | Susana has reported to me that the analyzer keeps loosing the wireless connection and, it takes very long for it to reconnect to the network again.also the analyzer is intermittently presenting a windows message NHSTW32.exe error. | In order to correct the connection issue, I adjusted the Roma.ini file with settings to enable the reconnect ability of the analyzer if the network connection is lost. I also made adjustments so that the reconnect interval is much sooner when the connection is lost. Lastly, I made the adjustments so that the reconnection retries are more frequent. In order to prevent the window message NHSTW32.exe error, I upgraded the software from version 3.3 to version 3.3MR1. Susana will be monitoring this instrument and, She will inform me about the connectivity performance of this analyzer. She will also keep an eye to confirm the windows message does not reappear.
No | Visit | Complaint | RAME | Unable to edit patient ID and result layouts. | Unable to edit patient ID and result layouts. | Reverted layout setting to default and rebuilt layouts. Problem corrected. | SWO-299763 | Setup | SW | Patient layout | Patient layout |  | Unable to edit patient ID and result layouts. | Unable to edit patient ID and result layouts. | Reverted layout setting to default and rebuilt layouts. Problem corrected.
No | Visit | Complaint | RAME | Unable to edit patient ID and result layouts. | Unable to edit patient ID and result layouts. | Reverted to default and rebuilt layouts. Problem corrected. | SWO-299765 | Setup | SW | Patient layout | Patient layout |  | Unable to edit patient ID and result layouts. | Unable to edit patient ID and result layouts. | Reverted to default and rebuilt layouts. Problem corrected.
No | Visit | Complaint | RFRA | Pas d'impression des résultats de mesures | Test imprimante Vérification configuration impression Impression fonctionnelle hormis résultats cal, QC, mesures échantillons | Reset CPU | SWO-299652 | No fault found | SW | Windows | off/on |  | No printing of measurement results | Printer test Checking print configuration Functional printing except cal results, QC, sample measurements | Reset CPU
No | Visit | Complaint | RFRA | Pb de Coox pendant le cycle de qualité | Passage de la solution de deproteinisation. Contrôle des vannes. | Changement de l'hemolyseur. Cal tHb. Essais | SWO-299911 | [5 empty] | Coox pb during the quality cycle | Passage of the deproteinization solution. Control of the valves. | Changing the hemolyseur. Cal tHb. Tests
No | Visit | Complaint | RCHN | 检测屏无法操作 | 检测屏无法操作，借910-298一个 |  | SWO-299401 | Defect | HW | Basic Unit | Display |  | The detection screen cannot be operated | The detection screen is inoperable, borrowing 910-298 one
No | Visit | Complaint | RESP | Errores 1083,1084,1085,1249 | Chequeo de neumática limpieza de circuito sustitución de tubos, sustitución de sonda, gasket, conector, ejecución de programas de servicio, reinstalación de paquete de soluciones y cassette de electrodos | Sustitución de tubos sustitución de accesorios reinstalación de paquete de soluciones y cassette de electrodos limpieza de circuito neumático | SWO-299659 | Defect | HW | Basic Unit | PM Kit |  | Errors 1083,1084,1085,1249 | Pneumatic check cleaning of tube replacement circuit, probe replacement, gasket, connector, service program execution, solution package reinstallation and electrode cassette | Replacing tubes replacing accessories reinstallation of solution package and cassette of pneumatic circuit cleaning electrodes
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 定标错误 | 更换测试卡后修复 | SWO-299849 | [5 empty] | Test card, reagent package failure | Calibration error | Repair after replacing test card
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 0589 测得的质控值高于质控范围 | 更换试剂包后修复 | SWO-299981 | [5 empty] | Test card, reagent package failure | 0589 Measured QC is higher than the QC range | Repair after replacement of reagent kit
No | Visit | Complaint | RCHN | 报错2168 | 报错2168，借903-030一个 |  | SWO-299358 | Defect | HW | Storage Module | PCB, Cartridge Lock | 2168 | 2168 error | 2168, loan 903-030 one
No | Visit | Complaint | RCHN | 报错1031 | 报错1031，借903-364一个 |  | SWO-299446 | Defect | HW | Process Module | Optical Unit | 1031 | 1031 | 1031, loan 903-364 one
No | Visit | Complaint | RFRA | Reboot de l'ABL9 après écran figé. |  | Remplacement du module PC + carte identification analyseur. Remise en place carte CF client. Redémarrage OK. Remplacement du pack solution, essais calibration OK. | SWO-299557 | [5 empty] | ABL9 reboot after frozen screen. |  | Replacement of the PC module - analyzer iD card. Re-establishing CF customer card. Restart OK. Replacement of the solution pack, trial calibration OK.
No | Visit | Complaint | RIND | Service Visit | There is an leakage from the inlet gas kit during rinse. | Flush the inlet gas kit, inlet probe, primary tubes with the hypochlorite solution, clean the chamber of ph and ei chamber. Now connect all the part same as before, initiate the analyzer and give two point calibration which is successfully done. Analyzer handed over to staff in well working condition. | SWO-300355 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | Service Visit | There is an leakage from the inlet gas kit during rinse. | Flush the inlet gas kit, inlet probe, primary tubes with the hypochlorite solution, clean the chamber of ph and ei chamber. Now connect all the part same as before, initiate the analyzer and give two point calibration which is successfully done. Analyzer handed over to staff in well working condition.
No | Visit | Complaint | RAME | There was a power outage at this location and the analyzer was abruptly shutdown. Corrupted files when power was restored. | Unit not booting up correctly, uninstalling and reloading SW v6.17, after loading 6.17 all data was gone, rebuilding up analyzer security and profiles. | Installed 6.17 and reloaded all profiles and customer based data points. | SWO-299609 | Defect | SW | RMED SW | RMED SW |  | There was a power outage at this location and the analyzer was abruptly shutdown. Corrupted files when power was restored. | Unit not booting up correctly, uninstalling and reloading SW v6.17, after loading 6.17 all data was gone, rebuilding up analyzer security and profiles. | Installed 6.17 and reloaded all profiles and customer based data points.
No | Visit | Complaint | RDEU | 24.01.2018; Hr. Grimmer (Med. Technik); Tel.: 09317916105; pCO2 und Lac Kal Fehler | Fehlersuche --> Dateiarchive (Aktivitäten, Kal., QC, Betriebsmittel), Elektroden überprüft --> E788 pCO2-Elektrode defekt (Glasmembrane) --> pCO2-Elektrode wurde beim Membranieren durch den Anwender mechanisch beschädigt (Glasbruch der Membrane), vermutlich ist die Elektrode auf den Boden gefallen!!! --> Kunde auf die dadurch entstehenden Kosten hingewiesen!!! | E788 (BA07) eingebaut und (YT01) ausgebaut, Nullpunktkalibrierung durchgeführt, Kal. und QC, System und Funktion i.O. | SWO-299475 | User error | Electrode | Electrode | pCO2 |  | 24.01.2018; Mr. Grimmer (Med. Technik); Tel.: 09317916105; pCO2 and Lac Kal Errors | Troubleshooting --> File archives (activities, cal., QC, equipment), electrodes checked --> E788 pCO2 electrode defective (glass membrane) --> pCO2 electrode was mechanically damaged by the user during membraneing (glass breakage of the membrane), presumably the electrode fell to the ground!!! --> customer pointed out the resulting costs!!! | E788 (BA07) installed and (YT01) removed, zero calibration performed, Cal. and QC, system and function in the
No | Visit | Complaint | RAME | Correlation Issues | To check tbili | I switched the OXI module with the 007 analyzer to see if the high tbili moves. I ran the thb cal in service. The customer will run samples to see if any change | SWO-300391 | Revisit required | HW | Oxi module | Oxi module | ctBill | Correlation Issues | To check tbili | I switched the OXI module with the 007 analyzer to see if the high tbili moves. I ran the thb cal in service. The customer will run samples to see if any change
No | Visit | Complaint | RAME | Issues with tbili | To check the tbili | I switched the OXI modules with the 006 analyzer witch was measuring higher on blood samples but not QC. I want to see if the issue moves to the other analyzer. The customer will run samples and let us know | SWO-300390 | Revisit required | HW | Oxi module | Oxi module | ctBill | Issues with tbili | To check the tbili | I switched the OXI modules with the 006 analyzer witch was measuring higher on blood samples but not QC. I want to see if the issue moves to the other analyzer. The customer will run samples and let us know
No | Visit | Complaint | RAME | Customer unable to edit patient ID and results layouts. | Unable to edit patient ID and result layouts. | Reverted layouts to default settings and rebuilt layouts. Problem corrected. | SWO-299760 | Defect | SW | RMED SW | SW Install | Patient Layout | Customer unable to edit patient ID and results layouts. | Unable to edit patient ID and result layouts. | Reverted layouts to default settings and rebuilt layouts. Problem corrected.
No | Visit | Complaint | RAME | Door not closing | Sensor door not closing | Adjust sensor position | SWO-300597 | Adjustment | HW | Basic Unit | Sensor Cassette Door |  | Door not closing | Sensor door not closing | Adjust sensor position
No | Visit | Complaint | RDEU | Daten werden nicht gesendet/übernommen | Nach dem Gerätetausch vom ABL 80 zum ABL 90 sollen keine $unvollständige Daten über die serielle Schnittstelle zur Software der Firma ZAN gesendet werden. Im Vorfeld wurde der Firma Zan $ Herrn Müller die Schnittstellenbeschreibung zur Verfügung gestellt. Nach Anfrage aus der Med.-Technik den Standort besucht. Leider war kein Remote möglich, da das Gerät weiterhin nicht an AQURE angebunden werden konnte. Gerät Konfiguration überprüft = OK serielle Schnittstelle überprüft = OK | keine Anpassung vorgenommen, kein Sendefehler festgestellt. Der Lungenfunktionsmitarbeiterin den Datenmitschnitt gezeigt und erklärt. Herr Müller den Datenmitschnitt per E-Mail zugeschickt. Den Datensatz mit Herrn Müller telefonisch besprochen, es sind keine großen Änderungen ABL90 $ABL80 zu erkennen. aktueller Datensatz ABL90: <ENQ><ACK> 10:20:35.14 <STX>1H¦ &^&¦¦¦ABL90^MHB-Lufu-ABL90¦¦¦¦¦¦¦¦1¦20180126102159<CR><ETB>5B<CR><LF><ACK> <STX>2P¦1¦¦¦¦……^…… ¦¦¦¦¦¦¦¦¦^¦^^^¦^¦^¦¦¦¦¦¦¦¦<CR><ETB>5C<CR><LF><ACK> <STX>3O¦1¦¦Sample #^189¦^^^Kapillare¦¦¦¦¦¦¦¦¦¦¦Kapillär^¦<CR><ETB>FC<CR><LF><ACK> <STX>4R¦1¦^^^pH^M¦7.415¦¦¦N¦¦F¦¦Anonym¦20180125143329<CR><ETB>7A<CR><LF><ACK> <STX>5R¦2¦^^^pCO2^M¦53.4¦mmHg¦¦N¦¦F¦¦¦<CR><ETB>11<CR><LF><ACK> <STX>6R¦3¦^^^pO2^M¦59.0¦mmHg¦¦N¦¦F¦¦¦<CR><ETB>D2<CR><LF><ACK> <STX>7R¦4¦^^^tHb^M¦14.9¦g $dL¦¦N¦¦F¦¦¦<CR><ETB>BE<CR><LF><ACK> <STX>0R¦5¦^^^RHb^M¦10.3¦%¦¦N¦¦F¦¦¦<CR><ETB>6B<CR><LF><ACK> <STX>1R¦6¦^^^O2Hb^M¦87.6¦%¦¦N¦¦F¦¦¦<CR><ETB>AD<CR><LF><ACK> <STX>2R¦7¦^^^sO2^M¦89.5¦%¦¦N¦¦F¦¦¦<CR><ETB>79<CR><LF><ACK> <STX>3R¦8¦^^^COHb^M¦1.5¦%¦¦N¦¦F¦¦¦<CR><ETB>83<CR><LF><ACK> <STX>4R¦9¦^^^MetHb^M¦0.6¦%¦¦N¦¦F¦¦¦<CR><ETB>19<CR><LF><ACK> <STX>5R¦10¦^^^T^I¦37.0¦Cel¦¦¦¦F¦¦¦<CR><ETB>97<CR><LF><ACK> <STX>6R¦11¦^^^pH(T)^M¦7.415¦¦¦N¦¦F¦¦¦<CR><ETB>17<CR><LF><ACK> <STX>7R¦12¦^^^pCO2(T)^M¦53.4¦mmHg¦¦N¦¦F¦¦¦<CR><ETB>E9<CR><LF><ACK> <STX>0R¦13¦^^^HCO3-^C¦34.2¦mmol $L¦¦¦¦F¦¦¦<CR><ETB>90<CR><LF><ACK> <STX>1R¦14¦^^^ABE^C¦7.8¦mmol $L¦¦¦¦F¦¦¦<CR><ETB>F6<CR><LF><ACK> <STX>2R¦15¦^^^SBE^C¦9.6¦mmol $L¦¦¦¦F¦¦¦<CR><ETB>0A<CR><LF><ACK> <STX>3R¦16¦^^^SBC^C¦31.3¦mmol $L¦¦¦¦F¦¦¦<CR><ETB>32<CR><LF><ACK> <STX>4R¦17¦^^^pO2(T)^M¦59.0¦mmHg¦¦N¦¦F¦¦¦<CR><ETB>AA<CR><LF><ACK> <STX>5R¦18¦^^^p50(act)^C¦27.39¦mmHg¦¦¦¦F¦¦¦<CR><ETB>53<CR><LF><ACK> <STX>6R¦19¦^^^AaDpO2^E¦28.5¦mmHg¦¦¦¦F¦¦¦<CR><ETB>9A<CR><LF><ACK> <STX>7R¦20¦^^^tO2^C¦18.3¦Vol%¦¦¦¦F¦¦¦<CR><ETB>79<CR><LF><ACK> <STX>0L¦1¦N<CR><ETX>03<CR><LF><ACK> <EOT> | SWO-300026 | No fault found | His/Lis | His/Lis | His/Lis |  | Data is not sent/transferred | After the device exchange from the ABL 80 to the ABL 90, no /unvollständ data should be sent via the serial interface to the software of the company ZAN. In the run-up to the event, the interface description was made available to Mr. Müller. Upon request from med.technology, visit the site. Unfortunately, it was not possible to remotely because the device still could not be connected to AQURE. Device configuration checked = OK serial interface checked = OK | keine Anpassung vorgenommen, kein Sendefehler festgestellt. Der Lungenfunktionsmitarbeiterin den Datenmitschnitt gezeigt und erklärt. Herr Müller den Datenmitschnitt per E-Mail zugeschickt. Den Datensatz mit Herrn Müller telefonisch besprochen, es sind keine großen Änderungen ABL90 $ABL80 zu erkennen. aktueller Datensatz ABL90: <ENQ><ACK> 10:20:35.14 <STX>1H¦ &^&¦¦¦ ABL90^MHB-Lufu-ABL90¦¦¦¦¦¦¦¦ 1¦20180126102159<CR><ETB>5B<CR><LF><ACK> <STX>2P¦1¦¦¦¦ ……^…… ¦¦¦¦¦¦¦¦¦ ^¦^^^¦^¦^¦¦¦¦¦¦¦¦ <CR><ETB>5C<CR><LF><ACK> <STX>3O¦1¦¦ Sample #^189¦^^^Kapillare¦¦¦¦¦¦¦¦¦¦¦ Kapillär^¦ <CR><ETB>FC<CR><LF><ACK> <STX>4R¦1¦^^^pH^M¦7.415¦¦¦ N¦¦ F¦¦ Anonym¦20180125143329<CR><ETB>7A<CR><LF><ACK> <STX>5R¦2¦^^^pCO2^M¦53.4¦mmHg¦¦ N¦¦ F¦¦¦ <CR><ETB>11<CR><LF><ACK> <STX>6R¦3¦^^^pO2^M¦59.0¦mmHg¦¦ N¦¦ F¦¦¦ <CR><ETB>D2<CR><LF><ACK> <STX>7R¦4¦^^^tHb^M¦14.9¦g $dL¦¦ N¦¦ F¦¦¦ <CR><ETB>BE<CR><LF><ACK> <STX>0R¦5¦^^^RHb^M¦10.3¦%¦¦ N¦¦ F¦¦¦ <CR><ETB>6B<CR><LF><ACK> <STX>1R¦6¦^^^O2Hb^M¦87.6¦%¦¦ N¦¦ F¦¦¦ <CR><ETB>AD<CR><LF><ACK> <STX>2R¦7¦^^^sO2^M¦89.5¦%¦¦ N¦¦ F¦¦¦ <CR><ETB>79<CR><LF><ACK> <STX>3R¦8¦ ^^^COHb^M¦1.5¦%¦¦ N¦¦ F¦¦¦ <CR><ETB>83<CR><LF><ACK> <STX>4R¦9¦^^^MetHb^M¦0.6¦%¦¦ N¦¦ F¦¦¦ <CR><ETB>19<CR><LF><ACK> <STX>5R¦10¦^^^T^I¦37.0¦ Cel¦¦¦¦ F¦¦¦ <CR><ETB>97<CR><LF><ACK> <STX>6R¦11¦^^^pH(T)^M¦7.415¦¦¦ N¦¦ F¦¦¦ <CR><ETB>17<CR><LF><ACK> <STX>7R¦12¦^^^pCO2(T)^M¦53.4¦mmHg¦¦ N¦¦ F¦¦¦ <CR><ETB>E9<CR><LF><ACK> <STX>0R¦13¦^^^HCO3-^C¦34.2¦mmol $L¦¦¦¦ F¦¦¦ <CR><ETB>90<CR><LF><ACK> <STX>1R¦14¦^^^ABE^C¦7.8¦mmol $L¦¦¦¦ F¦¦¦ <CR><ETB>F6<CR><LF><ACK> <STX>2R¦15¦^^^SBE^C¦9.6¦mmol $L¦¦¦¦ F¦¦¦ <CR><ETB>0A<CR><LF><ACK> <STX>3R¦16¦^^^SBC^C¦31.3¦mmol $L¦¦¦¦ F¦¦¦ <CR><ETB>32<CR><LF><ACK> <STX>4R¦17¦^^^pO2(T)^M¦59.0¦mmHg¦¦ N¦¦ F¦¦¦ <CR><ETB>AA<CR><LF><ACK> <STX>5R¦18¦^^^p50(act)^C¦27.39¦mmHg¦¦¦¦ F¦¦¦ <CR><ETB>53<CR><LF><ACK> <STX>6R¦19¦^^^AaDpO2^E¦28.5¦mmHg¦¦¦¦ F¦¦¦ <CR><ETB>9A<CR><LF><ACK> <STX>7R¦20¦^^^tO2^C¦18.3¦ Vol%¦¦¦¦ F¦¦¦ <CR><ETB>79<CR><LF><ACK> <STX>0L¦1¦ N<CR><ETX>03<CR><LF><ACK> <EOT>
No | Visit | Complaint | RAME | Unable to edit patient ID and result layouts. | Customer is unable to add to the patient layout | Assisted Primary engineers (Matt and Chuck) with the process of returning the analyzer to factory default and rebuilding the patient layouts. Discussed with customer the importance of checking all layouts and parameters before using to eliminate any errors from the rebuild. | SWO-300262 | Setup | SW | Report layout | Report layout |  | Unable to edit patient ID and result layouts. | Customer is unable to add to the patient layout | Assisted Primary engineers (Matt and Chuck) with the process of returning the analyzer to factory default and rebuilding the patient layouts. Discussed with customer the importance of checking all layouts and parameters before using to eliminate any errors from the rebuild.
No | Visit | Complaint | RLTD | Battery showing 0% and date is in American format | Attend analyser to investigate and replace battery if required. | Entered service and ran battery test. Battery showing as present and charging but percentage charged shows 0 constantly. Cell voltages all at around 4.1-4.2 v. Powered down and fitted new battery. Restarted amd battery percentage immediately showing at 85%. Monitored this and battery charging at steady rate. Disconnected and reconnected mains to test this switched correctly. Then made changes to the Windows date and time format as this was in American format. Exited roma, removed write protection, rebooted, made changes, enabled write protection amd rebooted again. Analyser restarted five minute startup. After this passed, I ran calibration and qc to verify analyser working correctly. Calibration showed lactate out due to sensor response error. Ran another calibration and this time sensitivity went higher and still showed response error. Ran third cal and same issue. Spoke to customer and they will put new cassette on analyser if the cassette does not settle. Logs showed cassette had been on analyser for 24 hours. Sensitivity had been high but passing calibrations. | SWO-300247 | Defect | HW | Battery module | Battery |  | Battery showing 0% and date is in American format | Attend analyser to investigate and replace battery if required. | Entered service and ran battery test. Battery showing as present and charging but percentage charged shows 0 constantly. Cell voltages all at around 4.1-4.2 v. Powered down and fitted new battery. Restarted amd battery percentage immediately showing at 85%. Monitored this and battery charging at steady rate. Disconnected and reconnected mains to test this switched correctly. Then made changes to the Windows date and time format as this was in American format. Exited roma, removed write protection, rebooted, made changes, enabled write protection amd rebooted again. Analyser restarted five minute startup. After this passed, I ran calibration and qc to verify analyser working correctly. Calibration showed lactate out due to sensor response error. Ran another calibration and this time sensitivity went higher and still showed response error. Ran third cal and same issue. Spoke to customer and they will put new cassette on analyser if the cassette does not settle. Logs showed cassette had been on analyser for 24 hours. Sensitivity had been high but passing calibrations.
Yes | Visit | Complaint | RLTD | Printer stopped printing again. Had done this before Christmas and again after Christmas - sue heggerty - Telephone No 01642 28 26 76 | Printer has on a number of occasions stopped working. |  | SWO-300088 | Re classified RMED | Phone support |  |  |  | Printer stopped printing again. Had done this before Christmas and again after Christmas - sue heggerty - Telephone No 01642 28 26 76 | Printer has on a number of occasions stopped working.
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 测试卡有凝固 1351 | 更换试剂包及冲洗机器后修复 by Andrew | SWO-299936 | Assisted Replacement | SP | SP | FT | 1351 | Test card, reagent package failure | Test card has solidification 1351 | Repair by Andrew after replacing the reagent kit and flushing the machine
No | Visit | Complaint | RCHN | 试剂包安装后报错 | 试剂包安装后报错，借903-158一个 |  | SWO-299853 | Defect | HW | Basic Unit | PCB, SOLUTION PACK ID |  | Reagent pack reported error after installation | Reagent pack after installation error, borrow 903-158 one
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 流量选择器错误，报错0978 | 更换试剂包后修复 | SWO-300619 | [5 empty] | Test card, reagent package failure | Traffic selector error, error 0978 | Repair after replacement of reagent kit
No | Visit | Complaint | RCHN | 进样门开启失败 | 进样门开启失败，借886-639，0,9-028,835-493各一个 |  | SWO-299861 | Defect | HW | Cover and Chassis | Cartridge door |  | Failed to open the sample door | Failed to open the sample door, borrow 886-639, 0, 9-028, 835-493 each
No | Visit | Complaint | RCHN | PTC板生锈 | PTC板生锈，借903-036一个 |  | SWO-299869 | Defect | HW | Storage Module | PCB, Replenishment Pack |  | PTC plate rust | PTC board rust, borrow 903-036 one
Yes | Visit | Complaint | RCHN | 报错5784 | 报错5784 | 取出断卡后修复 | SWO-299959 |  |  |  |  | 5784 | 5784 | 5784 | Repair after removing the broken card
No | Visit | Complaint | RAME | Neon lamp failure | Neon lamp intensity out of range. | Replaced lamp unit per Radiometer specifications, calibrated and ran QC instrument operating normally. | SWO-300286 | Defect | HW | Oxi module | Lamp unit |  | Neon lamp failure | Neon lamp intensity out of range. | Replaced lamp unit per Radiometer specifications, calibrated and ran QC instrument operating normally.
No | Visit | Complaint | RAME | Customer called stating that they are getting halogen lamp to high on two pt call | Found 0251 Halogen lamp intensity too high and 0579 Oximetry hardware problem. Not possible to measure errors. It appears that the machine started having troubles after customer moved the physical location to it's permanent spot. My guess is the jarring of the analyzer caused a misalignment of the lamp unit. | Performed Halogen lamp intensity adjustment. Performed multiple calibrations and THB calibration. | SWO-300581 | Adjustment | HW | Oxi module | Lamp unit | Halogen | Customer called stating that they are getting halogen lamp to high on two pt call | Found 0251 Halogen lamp intensity too high and 0579 Oximetry hardware problem. Not possible to measure errors. It appears that the machine started having troubles after customer moved the physical location to it's permanent spot. My guess is the jarring of the analyzer caused a misalignment of the lamp unit. | Performed Halogen lamp intensity adjustment. Performed multiple calibrations and THB calibration.
Yes | Visit | Complaint | RCHN | 卡盘 | 卡盘，借903-378一个 | 更换903-378后修复 by Andrew | SWO-299956 | [5 empty] | Chuck | Card, borrow 903-378 one | Repair by Andrew after replacement 903-378
No | Visit | Complaint | RCHN | 黑屏 | 黑屏，借910-298一个 | 重新插拔插头后修复 by Andrew The problem was solved by reconnectting the power plug. | SWO-299974 | Defect | HW | Basic Unit | Display |  | Black screen | Black screen, borrow 910-298 one | 重新插拔插头后修复 by Andrew The problem was solved by reconnectting the power plug.
No | Visit | Complaint | RSCH | Blank | Impression impossible sporadiquement (redémarrage de l’équipement) | Réinstallation complète du système d’exploitation e.o Réinstallation du software système (application ABL90) e.o Chargement de la configuration e.o Calibration e.o contrôle de qualité e.o | SWO-300984 | Defect | SW | Windows | Windows |  | Blank | Sporadically impossible printing (restarting equipment) | Complete relocation of the e.o operating system Reinstallation of the system software (ABL90 application) e.o Loading the e.o configuration Calibration e.o quality control e.o
No | Visit | Complaint | RAME | pO2 issues might be silicon contamination | Customer has already changed the SP X2 and SC X2 | Checked fluid transport, changed parts in PM kit Replace SC for Na+ and Ca++ errors on calibration and QC | SWO-301242 | Defect | HW | Basic Unit | PM Kit |  | pO2 issues might be silicon contamination | Customer has already changed the SP X2 and SC X2 | Checked fluid transport, changed parts in PM kit Replace SC for Na+ and Ca++ errors on calibration and QC
No | Visit | Complaint | RAME | Touch screen is black and instrument is alarming | Touch screen is black and alarming. Cycling power,reseating of connections & replacement of CF card did not change symptoms. | Replaced CPU per Radiometer specifications, instrument operating normally. Calibrated and ran QC instrument operating normally. | SWO-301933 | Defect | HW | DMS_PC | CPU | Touch Screen black | Touch screen is black and instrument is alarming | Touch screen is black and alarming. Cycling power,reseating of connections & replacement of CF card did not change symptoms. | Replaced CPU per Radiometer specifications, instrument operating normally. Calibrated and ran QC instrument operating normally.
No | Visit | Complaint | RFRA | Nombreux problèmes de calibration du lactate même après 24h de conditionnement. | Pb de calibration Lac malgré remplacement fréquent de K7. | Remplacement de la carte de mesure sensor interface Pcb . Redémarrage et test service ok. | SWO-301110 | Defect | HW | Basic Unit | PCB, Sensor interface | Lac | Many problems calibrating the lactate even after 24 hours of conditioning. | Pb calibration Lake despite frequent replacement of K7. | Replacing the pcb sensor interface measurement card. Restart and test service ok.
No | Visit | Complaint | RDEU | CtBil große Messabweichung | CtBil große Messabweichung zum anderen ABL90+ - Hemolyzer wurde bereits erneuert | Gerät geprüft, oXi Complete Test im Service gestartet, Spectrometer erneuert, Oxi Initialisierung abgewartet , 1,5std. diverse kal , Thb kal und QC durchgeführt. Fehlermeldung 1031,1034 Oxymetrie Initial , Oxymetrie Hardwareproblem. - Kal erst durch erneuten Neustart ok. | SWO-301326 | Defect | HW | Oxi module | Spectrometer |  | CtBil large measurement deviation | CtBil large measurement deviation to the other ABL90+ - Hemolyzer has already been renewed | Device tested, oXi Complete Test started in service, Spectrometer renewed, Oxi initialization awaited, 1.5h. various kal , Thb kal and QC. Error message 1031,1034 Oxymetry Initial , Oxymetry Hardware Problem. - Kal ok only by restarting it again.
No | Visit | Complaint | RFRA | Erreur 1177 et 978 | Aucune | Remise en route de l'analyseur. Reprise de l'analyseur de pret. | SWO-301188 | Cleaned | HW | Basic Unit | Flow selector detector | 1177 | Error 1177 and 978 | No | Restarted the analyzer. Takeover of the pret analyzer.
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | K,CA,CL,NA无终点 | 更换测试卡正常 。 | SWO-303435 | [5 empty] | Test card, reagent package failure | K, CA, CL, NA无终点 | Replacing the test card is OK.
Yes | 1 | Y | RCZE | Blank |  |  | D3CFA04CDA0FD431 | [5 empty] | Blank
No | Visit | Complaint | RCHN | 信号板故障 | 信号板故障，借903-334一个 | 更换903-334后修复 by jesson | SWO-300464 | [5 empty] | Signal plate failure | Signal board fault, borrow 903-334 one | Repair by jesson after replacing 903-334
No | Visit | Complaint | RCHN | 报错1031 | 加借903-364一个 | 更换903-364，重新定标正常。by jesson | SWO-300485 | Defect | HW | Process Module | Optical Unit | 1031 | 1031 | Borrow 903-364 one | Replace the 903-364 and re-mark normally. by jesson
Yes | Visit | Complaint | RCHN | 卡锁组件有问题 5689 | 卡锁组件有问题 5689 | 更换903-378一个正常。by jesson | SWO-300673 | [5 empty] | Problems with the card lock assembly 5689 | Problems with the card lock assembly 5689 | Replace 903-378 one normal. by jesson
No | Visit | Complaint | RFRA | Reboot de l'ABL9 après écran figé. | Remise en route de l'ABL9 sur site. | Redémarrage, calibration 2 points ok, passage Cq ok Desinstallation de l'ABL9 de prêt. | SWO-301111 | [5 empty] | ABL9 reboot after frozen screen. | ABL9 is restarted on site. | Reboot, calibration 2 points ok, passage Cq ok Installation of the loan ABL9.
No | Visit | Complaint | RMAY | Interface TCM5 FLEX to new PSG device. | Interface data $ signal unstable. | Suspect problem with 2.5mm Jack connector. Need to replace with good quality 2.5mm Jack connector. | SWO-309469 | [5 empty] | Interface TCM5 FLEX to new PSG device. | Interface data / signal unstable. | Suspect problem with 2.5mm Jack connector. Need to replace with good quality 2.5mm Jack connector.
No | Visit | Complaint | RIND | Intervention required. | Intervantion required. | Analyser investigate and found Calibration solutions wrongly installed. reinstalled calibration solutions and training given. Analyser checked with service program. Analyser Serviced and calibrated properly. Handed over in a fine working condition. Training given for consumable replacement and sample procedure. | SWO-303629 | User error | Consumables | Replacement | Replacement |  | Intervention required. | Intervantion required. | Analyser investigate and found Calibration solutions wrongly installed. reinstalled calibration solutions and training given. Analyser checked with service program. Analyser Serviced and calibrated properly. Handed over in a fine working condition. Training given for consumable replacement and sample procedure.
No | Visit | Complaint | RIND | Service Visit. | Need to change the potassium membrane. | Variation comes in the same sample run by the user two times, there is an calibration error in the potassium. Keep on hold the analyzer, replace the membrane of potassium with the proper procedure, record replacement, restart the analyzer after initializing wait for auto two point calibration successfully done. Analyzer handed over to staff in well working condition. | SWO-300972 | Assisted Replacement | Electrode | Membrane | K+ |  | Service Visit. | Need to change the potassium membrane. | Variation comes in the same sample run by the user two times, there is an calibration error in the potassium. Keep on hold the analyzer, replace the membrane of potassium with the proper procedure, record replacement, restart the analyzer after initializing wait for auto two point calibration successfully done. Analyzer handed over to staff in well working condition.
No | Visit | Complaint | RIND | Qc values problem for all parameter | Qc values problem for all parameter | checked the analyser ref membrane was weak so changed the membrane then deconatmination process was done then pump calibation and liquid sensor calibartion done calibarted to raedy mode tested level 1 and level 2bio rad qc all values are in range , observd for few samples then | SWO-302223 | Assisted Replacement | Electrode | Membrane | Reference |  | Qc values problem for all parameter | Qc values problem for all parameter | checked the analyser ref membrane was weak so changed the membrane then deconatmination process was done then pump calibation and liquid sensor calibartion done calibarted to raedy mode tested level 1 and level 2bio rad qc all values are in range , observd for few samples then
No | Visit | Complaint | RDAN | Intensiv oplever meget nedetid på Laktat pga. afviste kalibreringer, her ser man så et drop i følsomheden. Efter et par cleaning, kommer den nogenlunde tilbage igen og kalibreringen kommer på plads. Når vi ser et lille fald i følsomheden på Laktat, så kan man typisk også se at glukose lige dykker lidt. Men de andre parametre er upåvirket. Hvordan kan det være at det kun er metabolitterne som reagerer på en evt. forurening? Eller kan der være en anden årsag? Jeg har bedt hende tjekke, at Cleaning er sat til hver 8. time, og hun må ekstra cleane hvis de oplever det opft, da vi tror, at det er patienternes IV behandling som slår igennem. | Afprøvet. | El $Met målekammer og pumpeslange udskiftet, ventiler renset og smurt, slutafprøvet. | SWO-300995 | Defect | HW | El/Met Module | Measuring chamber |  | Intensive experience a lot of downtime on lactate due to Rejected calibrations, here you see a drip in the sensitivity. After a few cleaning, it comes back roughly again and the calibration comes in place. When we see a slight decrease in the sensitivity of lactate, you can typically also see that the glucose diver is a little bit. But the other parameters are unaffected. How can it be that only the metabolites react to a possible Pollution? Or can there be another reason? I have asked her to check that Cleaning is set to every 8. Time, and she must be extra cleane if they experience it, as we believe it is the patients ' IV treatment that strikes through. | Tested. | El $Met measuring Chamber and pump hose replaced, valves cleaned and lubricated, final tested.
No | Visit | Complaint | RIND | 'Parameter not installed' message appearing for every sample in ABL800 analyzer at Aster Adhar Hospital, Kolhapur | Need to check the status of the keys installed. | Installed the keys one more time as all the keys are showing as 'expired', restarted the analyzer, run 2 samples, machine is working fine. As the QC value of Na was found to be out of range, also replaced the Na membrane. | SWO-303923 | Defect | SW | RMED SW | Demo key expired |  | Parameter not installed' message appearing for every sample in ABL800 analyzer at Aster Adhar Hospital, Kolhapur | Need to check the status of the keys installed. | Installed the keys one more time as all the keys are showing as 'expired', restarted the analyzer, run 2 samples, machine is working fine. As the QC value of Na was found to be out of range, also replaced the Na membrane.
No | Visit | Complaint | RIND | PO2, PCO2 Measurement ustable for all the sample | Checked the data logs po2, pco2 Measurement ustable for all the sample found clot in po2 & pco2 chamber | Cleand all the chamber performed 2-point calibration processed sample result was ok | SWO-303444 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway | Clot | PO2, PCO2 Measurement ustable for all the sample | Checked the data logs po2, pco2 Measurement ustable for all the sample found clot in po2 & pco2 chamber | Cleand all the chamber performed 2-point calibration processed sample result was ok
No | Visit | Complaint | RAME | 0581 on qc | Lamp out of adjustment | Alignment of lamp with complete oxi test and THB calibration. | SWO-302593 | Adjustment | HW | Oxi module | Lamp unit | Halogen | 0581 on qc | Lamp out of adjustment | Alignment of lamp with complete oxi test and THB calibration.
No | Visit | Complaint | RLTD | SAYING - RESTART ANALYSER - ERROR 5689 TRIED TO RESTART - SAME ERROR Oprah Atunrase - 0161 922 6498 | Site visit to investigate analyser error 5689 rotate storage wheel. | Arrived on site and attended analyser. Analyser stuck in error state. Customer had already tried restart multiple times but not resolved. Checked logs and found that the error code reported referred to the storage wheel rotation. Entered service and ran safe positions. These all ran fine. Ran storage wheel barcode scanning to rotate and scan all positions. Scanning aborted part way through. Ran safe positions again and then rerun the scanning. Scanning aborted again for comm22 error. Manually rotated the positions and found that when cartridge hit a particular position there was a click sound. Removed cartridges and inspected for damage or debris in path. None found. Refitted and retested. Same error occurred. Found that this happened when cartridge went past the insert cartridge position. Took storage wheel apart and inspected for blockages or damage. Cleaned all sensors and checked all cables whilst inside storage wheel. Refitted all parts and cartridges into their original positions. Powered up the analyser and it started up without error. Entered service and tested storage wheel again in scanning procedure. Multiple times and all worked correctly without issue. Suspect slight misalignment rectified by dismantle and refit. Exited service programs and waited for analyser to initialize. Ran system check which completed and passed. Ran system clean three times which all ran successfully. followed this by running three qcs of the same level to tests analyser and cup transport from the storage wheel. Qcs run failed and ran high. Opened analyser and checked various components for blockages. Cleaned process wheel, and front filter. Checked cup wash and wash block but these were both clean. Inspected needle and replaced as precaution. Ran cleaning cycles again followed by qc. Qc ran high again. Recalibrated the lot of test cartridges and then rerun the qc. Qc now running on the mean. Analyser back running correctly. | SWO-301343 | Adjustment | HW | Storage Module | Storage Wheel | 5689 | SAYING - RESTART ANALYSER - ERROR 5689 TRIED TO RESTART - SAME ERROR Oprah Atunrase - 0161 922 6498 | Site visit to investigate analyser error 5689 rotate storage wheel. | Arrived on site and attended analyser. Analyser stuck in error state. Customer had already tried restart multiple times but not resolved. Checked logs and found that the error code reported referred to the storage wheel rotation. Entered service and ran safe positions. These all ran fine. Ran storage wheel barcode scanning to rotate and scan all positions. Scanning aborted part way through. Ran safe positions again and then rerun the scanning. Scanning aborted again for comm22 error. Manually rotated the positions and found that when cartridge hit a particular position there was a click sound. Removed cartridges and inspected for damage or debris in path. None found. Refitted and retested. Same error occurred. Found that this happened when cartridge went past the insert cartridge position. Took storage wheel apart and inspected for blockages or damage. Cleaned all sensors and checked all cables whilst inside storage wheel. Refitted all parts and cartridges into their original positions. Powered up the analyser and it started up without error. Entered service and tested storage wheel again in scanning procedure. Multiple times and all worked correctly without issue. Suspect slight misalignment rectified by dismantle and refit. Exited service programs and waited for analyser to initialize. Ran system check which completed and passed. Ran system clean three times which all ran successfully. followed this by running three qcs of the same level to tests analyser and cup transport from the storage wheel. Qcs run failed and ran high. Opened analyser and checked various components for blockages. Cleaned process wheel, and front filter. Checked cup wash and wash block but these were both clean. Inspected needle and replaced as precaution. Ran cleaning cycles again followed by qc. Qc ran high again. Recalibrated the lot of test cartridges and then rerun the qc. Qc now running on the mean. Analyser back running correctly.
Yes | Visit | Complaint | RAME | API failures | The analyzer has failed the last two API studies. If the third study fails, the lab will be disallowed to run blood gasses. | Performed thorough checkout of the analyzer. Performed ODIN Sensor Simulator test. Performed tHb alignment. Ran VK kit to verify parameter results. Performed internal QC measurements. Analyzer checks okay. Had techs run the API kit. Customer will inform us of the results when received. | SWO-301719 | Re classified RMED | Jim Gibbons mail, 7/2-2018 |  |  | LEH Mail | API failures | The analyzer has failed the last two API studies. If the third study fails, the lab will be disallowed to run blood gasses. | Performed thorough checkout of the analyzer. Performed ODIN Sensor Simulator test. Performed tHb alignment. Ran VK kit to verify parameter results. Performed internal QC measurements. Analyzer checks okay. Had techs run the API kit. Customer will inform us of the results when received.
No | Visit | Complaint | RKOR | pO2 difference between installation equipment (ABL90 flex) and uninstallation equipment (GEM 3500) | Because the blood is exposed to oxygen, the oxygen content of the blood increases. | Equipment comparison evaluation | SWO-302856 | User error | Performance | Performance | Sample handling |  | pO2 difference between installation equipment (ABL90 flex) and uninstallation equipment (GEM 3500) | Because the blood is exposed to oxygen, the oxygen content of the blood increases. | Equipment comparison evaluation
No | 4 | Y | RPAC | Blank | Glu&Lac calibration error |  | D3D1503250404E61 | Assisted Replacement | SC | SC | Glu/Lac |  | Blank | Glu&Lac calibration error
No | Visit | Complaint | RAME | Analyzer is making a loud noise | Troubleshot to a bad fan on the power supply. | Placed the power supply on order | SWO-301971 | Revisit required | HW | Basic Unit | Power Supply |  | Analyzer is making a loud noise | Troubleshot to a bad fan on the power supply. | Placed the power supply on order
No | Visit | Complaint | RIND | Low Chloride value & High Anion gap in the sample result. | Check the calibration log & patient result log & found the result satisfactory but Doctor complaining about the high Anion gap & low Cl- reading in every cases. | Run the zeroing check & it was ok,run the Qualicheck QC & got the desired result.Run few samples & found the result satisfactory. | SWO-303149 | [5 empty] | Low Chloride value & High Anion gap in the sample result. | Check the calibration log & patient result log & found the result satisfactory but Doctor complaining about the high Anion gap & low Cl- reading in every cases. | Run the zeroing check & it was ok,run the Qualicheck QC & got the desired result. Run few samples & found the result satisfactory.
No | Visit | Complaint | RESP | FUGAS | FUGA | SE SUSTITUYEN VARIOS CALPACK POR CLIENTE VOLVIENDO A SALIR AL RATO EL ERROR 964 DETECTADA CORRIENTE DE FUGA RELACIONADA CON EL CALPACK, SE SUSTITUYE INLE PROBE Y ENTRADA, ASI COMO SE LE EFECTUA KIT DE MANTENIMIENTO, SE DETECTA PROBLEMA EN ELECTRODO DE REF DEL SENSOR, ASI MISMO, SE EFECTUAN TODOS LOS TEST REITERADAMENTE Y PRUEBAS DE FLUIDICAS. CALIBRADO Y VERIFICADO CON Q.C. | SWO-301724 | Assisted Replacement | SP | SP | Leak current | SC Reference as well | Leaks | flight | SEVERAL CALPACK ARE SUBSTITUTED BY CUSTOMER RETURN TO THE TIME ERROR 964 DETECTED FUGA EFFECT RELATED TO CALPACK, INLE PROBE AND INPUT IS SUBSTITUTE, AS IT IS EFFECTED MAINTENANCE KIT, ALL TEST ESMITY AND FLUIDITY TESTS ARE EFFECTED. CALIBRATED AND VERIFIED WITH Q.C.
No | Visit | Complaint | RCHN | 温度异常，机器无法开机，报错1020 | 温度异常，机器无法开机，报错1020，需要借件903-127 | 更换903-127后修复 by jesson | SWO-300693 | Defect | HW | Storage Module | PCB, Storage Module | 1020 | Abnormal temperature, machine can not be turned on, error 1020 | Abnormal temperature, machine can not be turned on, error 1020, need to borrow 903-127 | Repair by jesson after replacing 903-127
No | Visit | Complaint | RIND | Service Visit. | Need observe the membrane of the parameters. | Check the activity log of the analyzer filtered by replacement and check which membrane has replaced in last one year, reported, now check the analyzer chamber and membrane of pH and EI module by keep on hold the analyzer. Restart the analyzer and give the manual protein removal, cleaning and two point calibration. Need to change the membrane of pO2, Na, Ca, and reference membrane. Na membrane available change that membrane with proper process and wait for two point calibration which is successfully done for Na. Analyzer handed over to staff. Now need to change the pO2, calcium and reference membrane. | SWO-302016 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | Service Visit. | Need observe the membrane of the parameters. | Check the activity log of the analyzer filtered by replacement and check which membrane has replaced in last one year, reported, now check the analyzer chamber and membrane of pH and EI module by keep on hold the analyzer. Restart the analyzer and give the manual protein removal, cleaning and two point calibration. Need to change the membrane of pO2, Na, Ca, and reference membrane. Na membrane available change that membrane with proper process and wait for two point calibration which is successfully done for Na. Analyzer handed over to staff. Now need to change the pO2, calcium and reference membrane.
No | Visit | Complaint | RAME | High crea measurement is high. | High crea measurement has trended up and is going out of control range. Membrane changes do not affect the QC performance. | Sent replacement electrodes to the customer. Worked with Cameron Michael to install new electrodes over phone. Calibration and QC for crea returned to historic performance. Went to site to verify and pickup electrodes for warranty return. Instrument is operating normally. | SWO-302479 | Assisted Replacement | Electrode | Electrode | Crea |  | High crea measurement is high. | High crea measurement has trended up and is going out of control range. Membrane changes do not affect the QC performance. | Sent replacement electrodes to the customer. Worked with Cameron Michael to install new electrodes over phone. Calibration and QC for crea returned to historic performance. Went to site to verify and pickup electrodes for warranty return. Instrument is operating normally.
No | Visit | Complaint | RAME | Lactate calibration is drifting and high crea measurement is high. | Lactate calibration is drifting and high crea has trended up and is going out of control range. Membrane changes do not affect the QC performance. | Sent replacement electrodes to the customer. Worked with Cameron Michael over the phone to install new electrodes. Calibration and QC for both parameters returned to historic performance. Went to site to verify and pickup electrodes for warranty return. Instrument is operating normally. | SWO-302478 | Assisted Replacement | Electrode | Electrode | Lac | Lot no : 2 boxes of lot 3720(12/6/2017) // 1 box of lot 3764(1/29/2018) | Lactate calibration is drifting and high crea measurement is high. | Lactate calibration is drifting and high crea has trended up and is going out of control range. Membrane changes do not affect the QC performance. | Sent replacement electrodes to the customer. Worked with Cameron Michael over the phone to install new electrodes. Calibration and QC for both parameters returned to historic performance. Went to site to verify and pickup electrodes for warranty return. Instrument is operating normally.
No | Visit | Complaint | RLTD | Please call Nigel McKelvy on 061 525 741 | Blank test result errors | Site blank tests passing on 4 reps with failing 5th rep with error message - 1068. Errors on cup dryer in service mode. Cleaned all internal components of fluid path. Ran cleaning program & 1 blank test. Failed. Service dump and data logs obtained from unit. Sent to RMED for investigation. | SWO-301535 | Revisit required | Revisit required | Revisit required | Revisit required | 1068 | Please call Nigel McKelvy on 061 525 741 | Blank test result errors | Site blank tests passing on 4 reps with failing 5th rep with error message - 1068. Errors on cup dryer in service mode. Cleaned all internal components of fluid path. Ran cleaning program & 1 blank test. Failed. Service dump and data logs obtained from unit. Sent to RMED for investigation.
No | Visit | Complaint | RDEU | 30.01.18: Telefonische Fehlermeldung an HanSc -Telefonsupport durch HanSc 30.01.17: Serviceanforderung an Mailbox ca. 16.00 Uhr 1178: Fehler Pumpenkalibrierung erkannt + 1245: Flüssigkeitstransport gescheitert | Fehlersuche Pumpengeräusche deutlich hörbar Pumpenschläuche beider SP90 (LOT BC-50 und BA-54) haben Quetschstellen und schwarze Verfärbungen | Wet-Section demontiert, Rollenkranz gesäubert und gängig gemacht (Ersatz muss bestellt und eingebaut werden) Neuen SP90 installiert, Probelauf ok | SWO-302312 | Defect | HW | Basic Unit | Pump rotor |  | 01/30/18: Phone Error Message to HanSc Phone Support by HanSc 30.01.17: Service request to mailbox approx. 4 p.m. 1178: Error pump calibration detected + 1245: Liquid transport failed | Troubleshooting Pump noise clearly audible Pump hoses of both SP90 (LOT BC-50 and BA-54) have crushing points and black discoloration | Wet-Section dismantled, roller wreath cleaned and made common (replacement must be ordered and installed) New SP90 installed, test run ok
No | Visit | Complaint | RAME | PO2 and Glu failing one first sample | PO2 reportable range is at 30.1-488mmHg which is below the measured samples actual pO2 causing the results glucose linearity error. The range on comparison instrument is 12.7-592mmHg so no out of range error was generated by the instrument. | Explained setup of reportable ranges and error generation. Instrument operating normally. | SWO-302570 | User error | Reportable range | Reportable range | Reportable range (pO2) |  | PO2 and Glu failing one first sample | PO2 reportable range is at 30.1-488mmHg which is below the measured samples actual pO2 causing the results glucose linearity error. The range on comparison instrument is 12.7-592mmHg so no out of range error was generated by the instrument. | Explained setup of reportable ranges and error generation. Instrument operating normally.
Yes | Visit | Complaint | RKOR | pO2 difference between installation equipment (ABL90 flex) and uninstallation equipment (GEM 3500) | Because the blood is exposed to oxygen, the oxygen content of the blood increases. | Equipment comparison evaluation | SWO-302859 | Re classified RMED | Double entry, see previous day |  |  |  | pO2 difference between installation equipment (ABL90 flex) and uninstallation equipment (GEM 3500) | Because the blood is exposed to oxygen, the oxygen content of the blood increases. | Equipment comparison evaluation
No | Visit | Complaint | RAME | The analyzer sometimes shows a windows error message about a problem with Netop. | Susana explains to me that at times the analyzer presents a message 'NHSTW32.EXE has encountered a problem and, needs to close'. | To help resolve the problem, I had to upgrade the analyzer software from version 3.3 to 3.3MR1. We also reviewed about the possible causes for the analyzer disconnecting from the wireless network. Susana asked if the ASUS wireless adapter is moved from one analyzer to another will this change the wifi MAC address. I checked with my Senior FSE Matt Rutrough and, he confirmed that the wifi MAC address is hard coded in the actual wireless adapter and, it will not change and, it will not affect the analyzer connectivity. Susana asked about having the pending log active for the end users of the ABL90 analyzers. I showed Susana how this can be done on this analyzer. She will work on activating the pending log thru the AQURE netop. This analyzer needed a new solution pack, I tried installing a new pack but, found defective because it would not activate properly preventing it from fitting in the analyzer. The defective solution pack lot number is YR-54. I tried installing another pack but, it was expired. Lastly, I foundation third pack which I was able to install without any issues. | SWO-302630 | Defect | SW | His/Lis | Wireless | LEH kigger efter | The analyzer sometimes shows a windows error message about a problem with Netop. | Susana explains to me that at times the analyzer presents a message 'NHSTW32. EXE has encountered a problem and, needs to close'. | To help resolve the problem, I had to upgrade the analyzer software from version 3.3 to 3.3MR1. We also reviewed about the possible causes for the analyzer disconnecting from the wireless network. Susana asked if the ASUS wireless adapter is moved from one analyzer to another will this change the wifi MAC address. I checked with my Senior FSE Matt Rutrough and, he confirmed that the wifi MAC address is hard coded in the actual wireless adapter and, it will not change and, it will not affect the analyzer connectivity. Susana asked about having the pending log active for the end users of the ABL90 analyzers. I showed Susana how this can be done on this analyzer. She will work on activating the pending log thru the AQURE netop. This analyzer needed a new solution pack, I tried installing a new pack but, found defective because it would not activate properly preventing it from fitting in the analyzer. The defective solution pack lot number is YR-54. I tried installing another pack but, it was expired. Lastly, I foundation third pack which I was able to install without any issues.
No | Visit | Complaint | RAME | Customer states that over the past 48-72 hrs. they have replaced several SPs due to the unit showing EC 964 (Leak Current) | Analyzer has numerous errors. | Thoroughly checked and tested the analyzer and could find no problems. According to the data logs, no errors have been occurring since the techs started using the clot catchers except the leakage current detected errors which is a solution pack issue. Performed a data dump for engineer review. | SWO-302492 | Assisted Replacement | SP | SP | Leak current |  | Customer states that over the past 48-72 hrs. they have replaced several SPs due to the unit showing EC 964 (Leak Current) | Analyzer has numerous errors. | Thoroughly checked and tested the analyzer and could find no problems. According to the data logs, no errors have been occurring since the techs started using the clot catchers except the leakage current detected errors which is a solution pack issue. Performed a data dump for engineer review.
No | Visit | Complaint | RLTD | tHB calibration date issue | tHB calibration date issue | Compared analyser setup to ICU. Set up correct. Changed setup interval for tHB calibration to 6 months then back to 3 months. Correct timeframe for tHB now displayed on machine. Site to monitor date on next tHB calibration. | SWO-302228 | Setup | SW | RMED SW | tHb Cal schedule |  | tHB calibration date issue | tHB calibration date issue | Compared analyser setup to ICU. Set up correct. Changed setup interval for tHB calibration to 6 months then back to 3 months. Correct timeframe for tHB now displayed on machine. Site to monitor date on next tHB calibration.
No | Visit | Complaint | RFRA | Problème de saturation aléatoire | Test complet en mode service: test oximetrie $ $ test po2 $ $ test de pression $ $ manifold $ $ détecteurs de liquide $ $ pompe $ $ test à l'aide de la cassette simulateur | Installation d'un kit de maintenance annuel Après départ, appel de la technicienne, les capillaires sortent avec une erreur 581 Remplacement de l'hemoliseur $ $ Calibrations thb | SWO-302710 | Defect | HW | Oxi module | Hemolyzer |  | Random saturation problem | Full test in service mode: test oximetrie /// test in2 ///pressure test // manifold // /// /pump /pump // test using the simulator cassette | Installing an annual maintenance kit After departure, call from the technician, the capillaries come out with an error 581 Replacement of the hemolizer $ $ Thb calibrations
No | Visit | Complaint | RFRA | Le site d'introduction reste bloqué en haut, impossible de le faire redescendre depuis le menu de service. lorsque qu'on lance la calibration du site d'introduction, on entend la roue qui accroche sur quelques chose. Le site d'introduction a été retiré pour regarder la position de la roue, la gorge qui part du centre vers l'extérieur, au niveau de l'ouverture est en position 10h. | vérification site introduction position du joint sonde aspiration | remise en place site introduction essais | SWO-302879 | Adjustment | HW | Auto Inlet module | Inlet Position Cal. |  | The introductory site remains blocked at the top, impossible to bring it down from the service menu. when you launch the calibration of the introductory site, you hear the wheel hanging on something. The introductory site has been removed to look at the position of the wheel, the throat that starts from the center to the outside, at the aperture level is in the position 10am. | site verification introduction joint position suction sound | re-establishment introduction site Tests
No | Visit | Complaint | RAME | Possible blockage error. | Clot found in the tubing. Also, tubing appeared to be too long and bent. | Replaced all the tubing and pumps. Cleaned and flushed the manifold. Ran calibrations and three levels external QC. | SWO-302502 | [5 empty] | Possible blockage error. | Clot found in the tubing. Also, tubing appeared to be too long and bent. | Replaced all the tubing and pumps. Cleaned and flushed the manifold. Ran calibrations and three levels external QC.
No | Visit | Complaint | RCHN | 报错1034，1011，0521 | 报错1034，1011，0521，借902-918、902-980（上海仓库缺货）和903-165各一个 | 更换902-980后修复 | SWO-302130 | Defect | HW | Oxi module | Hemolyzer | 1034 | Error 1034,1011,0521 | Error 1034,1011,0521, borrowing 902-918, 902-980 (stock shortage in Shanghai warehouse) and 903-165 each | Repair after replacement 902-980
No | Visit | Complaint | RCHN | 报错1223错误 | 报错1223错误，借903-329一个 | 更换903-329后修复 by jesson | SWO-301437 | Defect | HW | Aspiration Module | Needle Wash Unit |  | 1223 errors | 1223 error, borrow 903-329 one | Repair by jesson after replacing 903-329
No | Visit | Complaint | RAME | Correlation Issues | After switching the OXI module with the other analyzer the issue moved. | I put the the OXI module back into 007 and replaced the module that was giving the bias. | SWO-302973 | Defect | HW | Oxi module | Oxi module, Replaced | Bias | Correlation Issues | After switching the OXI module with the other analyzer the issue moved. | I put the the OXI module back into 007 and replaced the module that was giving the bias.
Yes | Visit | Complaint | RAME | Issues with tbili | I switched the OXI module and the issue moved | I put the original OXI module back in this analyzer | SWO-302974 | Re classified RMED | Red X with R2477N006 |  |  |  | Issues with tbili | I switched the OXI module and the issue moved | I put the original OXI module back in this analyzer
No | Visit | Complaint | RDAN | Vi har nogle gange oplevet at få Fejlkode 1218: Utilstrækkelig prøvemateriale til prøvemåling selv om glasset var fyldt korrekt. Vi kunne efterfølgende sætte samme prøve på apparatet igen og få svar som vanligt. Har du nogen forklaring på det? Og hvad kan vi evt. gøre for at forhindre, det sker igen? Jan Guldberg anbefaler at kunden kører 2-3 gange cleaning, samt at vores servicetekniker udfører service tip 686803449 | Undersøgt. | Washcell renset som beskrevet i service tip 686803449, slutafprøvet. | SWO-302247 | Cleaned | HW | Aspiration Module | Needle Wash Unit | 1218 | We have sometimes experienced getting error code 1218: Insufficient sample test material even though the glass was filled correctly. We could subsequently put the same test on the device again and get answers as usual. Do you have any explanation for that? And what can we possibly do? Do to prevent it from happening again? Jan Guldberg recommends that the customer runs 2-3 times cleaning and that our service Technician performs service tip 686803449 | Examined. | Washcell cleaned as described in service tip 686803449, final tested.
No | Visit | Complaint | RKOR | Solution pack 재장착 에러가 계속 뜸 | Service mode로 확인 해보니 Manifold offset value가 out of range 임. | SP과 Manifold 닿는 부분을 거즈로 청소 해줌 | SWO-306438 | Cleaned | HW | Basic Unit | SP Contatcts |  | Solution pack re-mounting error continues to be moxibustion | I checked with Service mode and the Manifold offset value is out of range. | SP and Manifold touch areas cleaned with gauze
No | Visit | Complaint | RAME | Not connecting to Aqure | The analyzer would not connect to the wireless network | I replaced the wireless card | SWO-303259 | Setup | SW | His/Lis | Wireless |  | Not connecting to Aqure | The analyzer would not connect to the wireless network | I replaced the wireless card
No | Visit | Complaint | RAME | Customer called with pressure test failure | Customer called with pressure failures | THe analyzer was in the ready state when I arrived on site. I performed the fluidic service programs to verify performance. Checked all logs and found pressure test failures intermittently over the last coupe of weeks. I left one inlet connector gasket on site in case the failure returns. Basic training with Joel and discussed remote support to the analyzer. | SWO-303239 | No fault found | No fault found | No fault found | No fault found |  | Customer called with pressure test failure | Customer called with pressure failures | THe analyzer was in the ready state when I arrived on site. I performed the fluidic service programs to verify performance. Checked all logs and found pressure test failures intermittently over the last coupe of weeks. I left one inlet connector gasket on site in case the failure returns. Basic training with Joel and discussed remote support to the analyzer.
No | Visit | Complaint | RSCH | Rep Powersupply | Analysator schwarz | Analysator ausgetauscht und in Thalwil repariert Leihgerät installiert (ABL800) Powersupply ausgetauscht Funktionskontrolle i.o | SWO-303023 | Defect | HW | Basic Unit | Power Supply |  | Rep Powersupply | Analyzer black | Analyzer replaced and repaired in Thalwil Rental device installed (ABL800) Powersupply replaced Functional control i.o
Yes | 3 | Y | RPAC | Blank | Long Term shutdown |  | D3D398353C4703FE | Re classified RMED | Planned |  |  |  | Blank | Long Term shutdown
No | Visit | Complaint | RLTD | Stevev Ramsey or Phil Harding-Payne 01206 742493 | Failing vacuum value, customer had already changed gasket and carried out trouble shooting. | Customer had a number of faulty gaskets, replaced with know working gasket now ok. | SWO-303422 | Defect | HW | Auto Inlet module | Inlet gasket |  | Stevev Ramsey or Phil Harding-Payne 01206 742493 | Failing vacuum value, customer had already changed gasket and carried out trouble shooting. | Customer had a number of faulty gaskets, replaced with know working gasket now ok.
No | Visit | Complaint | RFRA | Erreur inlet liqui sensor et blood gas liquid sensor. Valeurs en mode service = 254. Erreur solution pack check en mode service. | Contrôle de la bay unit | Démontage et remplacement de la bay unit $Calibration menu service fluid transport ok $Remise en route $Calibration deux points et passage des CQ $Essai gaz patient ok | SWO-304563 | Defect | HW | Basic Unit | Cassette retainer |  | Erreur inlet liqui sensor et blood gas liquid sensor. Valeurs en mode service = 254. Erreur solution pack check en mode service. | Control of the bay unit | Disassembly and replacement of the bay unites $Calibration menu service fluid transport ok $Remise en route $Calibration two points and passage of the CQ $Essai patient gas ok
No | Visit | Complaint | RLTD | Battery not working new installation . | Battery error 1560 detected. | Replaced battery module and test ok. | SWO-303047 | Defect | HW | Battery module | Battery module | 1560 | Battery not working new installation . | Battery error 1560 detected. | Replaced battery module and test ok.
No | Visit | Complaint | RDEU | 31.01.2018; Fr. Rohleder; Tel.: 02611371447; Intervention erforderlich, Lösungen nicht gepumpt werden 31.01.2018 15:15 Fr. Rohleder 02611377310: weiterhin Interventionsmeldung Leckagefehler | Fehlersuche, Einlassdichtung bleibt sporadisch hängen. | Einlassmodul für den 02.02.2018 bestellt. | SWO-302963 | Revisit required | HW | Auto Inlet module | Auto Inlet module |  | 31.01.2018; Ms. Raw leather; Tel.: 02611371447; Intervention required, solutions not to be pumped 31.01.2018 15:15 Fr. Raw leather 02611377310: continue intervention notification Leakage errors | Troubleshooting, inlet seal gets stuck sporadically. | Inlet module for 02.02.2018.
No | Visit | Complaint | RAME | PO2 high on both ABL80s on VK-R5 material only | Checked data log and all calibrations and external controls are fine. | Ran Zero calibration. Ran system cycle. Ran VK-R5 kit and presented results to Respiratory | SWO-303299 | [5 empty] | PO2 high on both ABL80s on VK-R5 material only | Checked data log and all calibrations and external controls are fine. | Ran Zero calibration. Ran system cycle. Ran VK-R5 kit and presented results to Respiratory
No | Visit | Complaint | RAME | PO2 failing VK kit | PO2 high on both ABL80s on VK-R5 material only | Ran Zero calibration. Ran Linearity VK-R5. But unable to complete due to low cycles left on SP (and replacement not available today) | SWO-303303 | [5 empty] | PO2 failing VK kit | PO2 high on both ABL80s on VK-R5 material only | Ran Zero calibration. Ran Linearity VK-R5. But unable to complete due to low cycles left on SP (and replacement not available today)
No | Visit | Complaint | RESP | Cliente informa que el equipo siempre da error en parámetros de pCO2 cuando instala el reactivo mientras el mismo en el equipo de al lado funciona correctamente. | Sigue apareciendo error 1070 fallo respuesta de electrodos. | Comprobación del estado del analizador. Se revisan registros y se procede a cambiar Sensor interface PCB. Se monta el equipo y se instalan reactivos nuevos. El equipo calibra sin dar errores por lo que se pasan Qc´s. Equipo da problemas en HbF, se realiza test OXi, cal y Cal de Hb, se pasan los Qc´s y sale todo OK. Funcionamiento aparentemente correcto. | SWO-302995 | Defect | HW | Basic Unit | PCB, Sensor interface |  | Client reports that the computer always fails pCO2 parameters when it installs the reagent while the reagent on the computer next door is working properly. | You still see error 1070 electrode response failure. | Checking the status of the scanner. Logs are reviewed and PCB interface sensor is changed. The equipment is mounted and new reagents are installed. The equipment calibrates without giving errors so what Is passed Qc's. Equipment gives problems in HbF, OXi, lime and Hb Cal tests are performed, the Qc's are passed and everything OK exits. Apparently correct operation.
No | Visit | Complaint | RCHN | 光谱不匹配 0581 | 光谱不匹配 0581，借902-980一个。 | 更换902-980后修复 by jesson | SWO-301287 | Defect | HW | Oxi module | Hemolyzer |  | Spectrum Mismatch 0581 | The spectroscopy does not match 0581, borrowing 902-980 one. | Repair by jesson after replacement 902-980
No | Visit | Complaint | RCHN | 无法开机 | 无法开机，借910-285一个 | 更换910-285后修复 by jesson | SWO-301973 | Revisit required | HW | Power Supply | Power Supply |  | Cannot power on | Cannot power on, borrow 910-285 one | Repair by jesson after replacement 910-285
No | Visit | Complaint | RCHN | 测试卡扫描失败 | 测试卡扫描失败，借902-924一个。 | 更换902-924后修复 by jesson | SWO-302122 | Defect | HW | Cover and Chassis | Barcode reader, display |  | Test card scan failed | The test card scan failed, borrowing 902-924 one. | Repair by jesson after replacement 902-924
No | Visit | Complaint | RAME | Oximetry hardware error 0316. | Oximetry hardware error 0258 neon lamp intensity too low and error 0316 neon number too low. | Cleaned neon lamp and adjusted lamp unit. | SWO-303366 | Adjustment | HW | Oxi module | Lamp unit | Neon to low | Oximetry hardware error 0316. | Oximetry hardware error 0258 neon lamp intensity too low and error 0316 neon number too low. | Cleaned neon lamp and adjusted lamp unit.
No | Visit | Complaint | RLTD | Calibrations are failing with error 251 - Halogen lamp intensity too high Caller Name: MR NICHOL Telephone No 01896 82 62 45 | Halogen lamp intensity too high. | Ran complete OXI test and halogen intensity was on the high side alothough still within limits. Continued to Lamp adjustment successfully followed by a 1 point calibration. Ran another complete OXI test to double check value. After another 1 point calibration a service tHb Calibration was ran followed by user tHb cal and cal verification. | SWO-302946 | Adjustment | HW | Oxi module | Lamp unit | Halogen | Calibrations are failing with error 251 - Halogen lamp intensity too high Caller Name: MR NICHOL Telephone No 01896 82 62 45 | Halogen lamp intensity too high. | Ran complete OXI test and halogen intensity was on the high side alothough still within limits. Continued to Lamp adjustment successfully followed by a 1 point calibration. Ran another complete OXI test to double check value. After another 1 point calibration a service tHb Calibration was ran followed by user tHb cal and cal verification.
No | Visit | Complaint | RKK | CARTRIGE40エラー発生 | CARTRIGE40 | ベアリングの土台の高さを0.5mmくらい下げた バージョンＵＰ失敗 元のバージョンに戻す 溶液パック、ブランクカートリッジ使用 | SWO-303368 | Adjustment | HW | Storage Module | Storage Wheel | Hub bearing | CARTRIGE40 Error | CARTRIGE40 | The height of the bearing base was lowered by about 0.5mm. Version UP Failed Revert to original version Solution pack, blank cartridge use
No | Visit | Complaint | RBNL | Analyzer heeft een poos stil gestaan en wil nu niet opstarten na plaatsen nieuw reagenspack. Foutmelding 3410: ftc_101. | Zie zlef geen lekkage, gebruiker heeft gisteren wel iets gezien rechts achter onder instrument. Wil geen vloeistof verpompen. | Vervuilde washcup gedemonteerd en schoongemaakt. Inletfilter gedemonteerd en schoongemaakt, naald doorgespoten demiwater. Vervuilde washunit gedemonteerd en schoongemaakt. Diverse slangen doorgespoten en proceswiel schoongemaakt. Nu start systeem wel normaal op. Naald afgesteld | SWO-302939 | Cleaned | HW | Cup wash unit | Cup wash unit | 3410 | Analyzer has stood still for a while and now does not want to boot after inserting new reagent pack. Error message 3410: ftc_101. | See Zlef No leakage, user has seen something right behind under instrument yesterday. Do not want to pump liquid. | Contaminated washcup disassembled and cleaned. Inletfilter disassembled and cleaned, needle-sprayed demi water. Contaminated washunit disassembled and cleaned. Various hoses sprayed and process wheel cleaned. Now the system starts up normally. Needle adjusted
No | Visit | Complaint | RDEU | 14.12.2017, Frau Füssing, Tel: 0812888200: Display dunkel 31.12.2017 Frau Füßi Tel. 0812888200: Display dunkel,Gerät fährt herunter 31.01.2018 Fr. Balzer 0812888200: Gerät fährt immer noch beliebig runter | Gerät startet sporadisch neu | CF Karte erneuert Software komplett neu installiert Datensicherung zurückgespielt Systemreinigung und Systemüberprüfung durchführt LQC gemessen | SWO-302846 | Defect | HW | Analyzer Control Module | CF Card |  | 14.12.2017, Mrs Füssing, Tel: 0812888200: Display dark 31.12.2017 Mrs. Füßi Tel. 0812888200: Display dark, device shuts down 31.01.2018 Ms. Balzer 0812888200: Device still shuts down at will | Device restarts sporadically | CF card renewed Software completely reinstalled Data backup played back System cleaning and system verification LQC measured
No | Visit | Complaint | RDEU | Barcodes werden nur sporadisch im Inlet erkannt | Laser Engine fehlerhaft | Laser Engine ersetzt, Barcodes weiterhin nur sporadisch erkannt, Umrüstkit wird bestellt | SWO-302925 | Defect | HW | Aspiration Module | Barcode reader, Inlet |  | Barcodes are only detected sporadically in the inlet | Laser Engine faulty | Laser engine replaced, barcodes still only detected sporadically, conversion kit is ordered
No | 55 | Y | RPNZ | Blank | Sensor cassette replacement |  | D3D4B95BB22EE155 | Assisted Replacement | SC | SC | SC |  | Blank | Sensor cassette replacement
No | Visit | Complaint | RDEU | Platz lässt sich nicht entriegeln | System überprüft => Stepper Motor für Verriegelung defekt , Pumpenrad hält auf Welle nicht | Stepper Motor und Pumpenrad erneuert , Kalibration und QC durchgeführt | SWO-303686 | Defect | HW | Basic Unit | Steppermotor |  | Space cannot be unlocked | System checked => Stepper motor for locking defective, pump wheel does not hold on shaft | Stepper motor and pump wheel renewed, calibration and QC performed
No | Visit | Complaint | RAME | The analyzer is presenting error 1559 Battery error detected. | The analyzer is presenting error 1559 battery cell failure. | Checked the battery cell voltage values, found cell 4 voltage to be lo. Replaced a new battery module. | SWO-303882 | Defect | HW | Battery module | Battery module |  | The analyzer is presenting error 1559 Battery error detected. | The analyzer is presenting error 1559 battery cell failure. | Checked the battery cell voltage values, found cell 4 voltage to be lo. Replaced a new battery module.
No | Visit | Complaint | RDEU | 01.02.2018 19:00 Fr. Weißbach 02088814208: Interventionsmeldung Leckage erkannt 02.02.2018, 9.30 Uhr, Frau Kraus, Tel: 02088814181: weiterhin Intervention erforderlich-jetzt man. Spülung | Leckage erkannt, Flowsel.fehler, | Nacheinweisung durchgeführt, Flowsel. gereinigt, Fließwege durchgespült, Gerät kalibriert, | SWO-304050 | Cleaned | HW | Basic Unit | Flowselector |  | 01.02.2018 19:00 Fr. Weißbach 02088814208: Intervention notification Leakage detected 02.02.2018, 9.30 a.m., Mrs. Kraus, Tel: 02088814181: continued intervention required-now one. Flushing | Leakage detected, Flowsel.error, | Follow-up performed, Flowsel. cleaned, flow paths washed through, device calibrated,
No | Visit | Complaint | RDEU | Leckage erkannt mehrfach SP90 erneuert System bereits vom KD gespült | Fehlersuche Pumpenschlauch des SP90 gequetscht und undicht eine Rolle des Rollenkranzes schwergängig | Rollenkranz ersetzt SP90 installiert Probelauf | SWO-303560 | Defect | HW | Basic Unit | Pump rotor |  | Leakage detected multiple SP90 renewed System already flushed by KD | Troubleshooting Pump hose of the SP90 squeezed and leaked a roll of the roller wreath difficult | Roller wreath replaced SP90 installed Rehearsal
No | Visit | Complaint | RDEU | 23.01.2018; Hr. Olbert; Tel.: 079618816706; Display ist ausgefallen | Garantie CPU Modul defekt wegen black Monitor CPU Modul 910-287 aus Demogerät wieder zurückgetauscht. | Datensicherung durchgeführt. CPU Modul 910-342 eingebaut und erneuert. Software neu aufgespielt und Einlaufphase durchgeführt. | SWO-303927 | Defect | HW | DMS_PC | CPU |  | 23.01.2018; Mr. Olbert; Tel.: 079618816706; Display is down | Guarantee CPU module defective due to black monitor CPU module 910-287 from demo device. | data backup. CPU module 910-342 built-in and renewed. software replayed and run-in phase performed.
No | 3 | Y | RPOL | Blank | Problem rozwi¹zano. Aparat sprawny. Zwrócenie uwagi u¿ytkownikowi na poprawnoœæ wymiany uszczelki wlotu. |  | D3D4A7D3B2444376 | User error | Replacement | Replacement | Replacement of consumables |  | Blank | Problem resolved. The camera is functional. Attention of the user to the correct replacement of the inlet gasket.
No | Visit | Complaint | RCHN | 报错1034，1011，0521 | 报错1034，1011，0521，借一个902-980（北京仓库发出） | 更换902-980后修复 by jesson | SWO-302213 | Defect | HW | Oxi module | Hemolyzer |  | Error 1034,1011,0521 | 1034,1011,0521, borrow edgy 902-980 (issued by Beijing Warehouse) | Repair by jesson after replacement 902-980
No | Visit | Complaint | RCHN | 报错1258 | 报错1258，借903-329一个 | 更换903-329一个正常。by jesson | SWO-302731 | Cleaned | HW | Aspiration Module | Needle Wash Unit | 1258 | 1258 error | 1258, loan 903-329 one | Replace the 903-329 one normal. by jesson
No | Visit | Complaint | RCHN | 报错1068 | 报错1068，借903-210一个 | 更换903-210备件后修复 by jesson | SWO-302873 | Defect | HW | Process Module | PCB, Cup dryer | 1068 | 1068 error | 1068, borrow edith 903-210 | Repair by jesson after replacement of 903-210 spare parts
No | Visit | Complaint | RCHN | 液路部分外壳损坏 | 液路部分外壳损坏，借886-521一个 | 更换886-521后修复 by jesson | SWO-302818 | Defect | HW | Basic Unit | Front Cover, right |  | Damage to part of the liquid-road shell | Part of the liquid road shell is damaged, borrow 886-521 one | Repair by jesson after replacing 886-521
No | Visit | Complaint | RIND | Calibration sensitivity out of range | Checked dat log ph calibraiton sensitivity out of range & calibration unstable(response fault) for all the calibration replaced new reference membrane & checked then also same crosschecked new ph electrode it was working fine | Need to replace ph electrode | SWO-303637 | Revisit required | Electrode | Electrode | pH |  | Calibration sensitivity out of range | Checked dat log ph calibraiton sensitivity out of range & calibration unstable(response fault) for all the calibration replaced new reference membrane & checked then also same crosschecked new ph electrode it was working fine | Need to replace ph electrode
No | Visit | Complaint | RDEU | CO2 Empf zu niedrig 01.02.2018; Hr. Koomen; Tel.: 020644417105; weiterhin pCO2 Empf. Fehler 01.02.2018, Herr Koomen, Tel: 020644417105: weiterhin pCO2 empf. Fehler | pCO2 Kalibrationsfehler, Status und Empfehlenswert sind nicht o.k., Elektrode def. | Fließwege und Gaswege gereinigt, Meßkammer gereinigt, pCO2 Elektrode gereinigt, Elektrode ersetzt, Gerät kalibriert, QC durchgeführt, | SWO-303464 | Assisted Replacement | Electrode | Electrode | pCO2 |  | CO2 empf too low 01.02.2018; Mr. Koomen; Tel.: 020644417105; pCO2 Empf. Error 01.02.2018, Mr. Koomen, Tel: 020644417105: continue pCO2. Error | pCO2 Calibration error, status and recommended are not OK, electrode def. | Flow paths and gas paths cleaned, measuring chamber cleaned, pCO2 electrode cleaned, electrode replaced, device calibrated, QC performed,
No | Visit | Complaint | RSCH | Rep. Ca | Sporadisch Ca Drift Bereits Ca und Ref. Membraniert | Ref. Elektrode ersetzt (von Medizintechnik) Alte Lot = 6807 => Garantie | SWO-303617 | Assisted Replacement | Electrode | Electrode | Reference | Ca+ drift | Rep. Ca | Sporadic Ca Drift Already Ca and Ref. Membraneed | Ref. Electrode replaced (by medical technology) Old lot = 6807 => Warranty
No | Visit | Complaint | RSCH | Rep. Einlass | Tropft sporadisch aus Einlass Nach Einlassdichtung Wechsel 2Monate dicht Ca sporadisch Drift | Komplettes Einlass-Modul mit bestehender Einlassdichtung ersetzt i.o Leitungen Abfall gespült i.o Ref. Elektrode ersetzt i.o Alte Lot = 6807 => Garantie O-Ring pH ersetzt i.o Initialisierung i.o Kal. und QC i.o Testmessungen i.o System i.o Einlass Gerät wird 2Monate beobachtet | SWO-303616 | Defect | HW | Inlet module | Inlet module |  | Rep. Intake | Drips sporadically from inlet After inlet seal change 2months dense Ca sporadic drift | Complete inlet module with existing inlet seal replaces i.o Pipes Waste rinsed i.o Ref. Electrode replaces i.o Old lot = 6807 => Warranty O-ring pH replaces i.o Initialization i.o Kal. and QC i.o Test measurements i.o System i.o Inlet device is observed for 2months
No | Visit | Complaint | RSCH | Rep OXI Modul | Halogen Lampe ausserhalb Bereich seit Reinigung 01.02.2018 $20:00 Lampenwert 3824 | Halogen Lampe justiert (3330) Danach immer Fehler 249 Kal und QC wiederholt bis i.o Funktionskontrolle i.o | SWO-303468 | Adjustment | HW | Oxi module | Lamp unit | Halogen | Rep OXI Module | Halogen lamp outside range since cleaning 01.02.2018 Lamp value 3824 | Halogen lamp adjusted (3330) After that always error 249 Kal and QC repeated until i.o Functional control i.o
No | Visit | Complaint | RFRA | Écran noir | Test de l'écran et du cable dans les différentes positions. | Remplacement de l'écran. Essais OK | SWO-303697 | Defect | HW | Analyzer Control Module | Display |  | Black screen | Screen and cable test in different positions. | Replacing the screen. OK tests
No | Visit | Complaint | RCHN | 接口板腐蚀 | 接口板腐蚀，借902-972一个 | 更换902-972一个正常。by jesson | SWO-302667 | Defect | HW | Basic Unit | PCB, Sensor connector assembly |  | Interface plate corrosion | Interface plate corrosion, borrow 902-972 one | Replace 902-972 one normal. by jesson
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 流量选择定标错误，1177 | 更换试剂包后修复 | SWO-304154 | [5 empty] | Test card, reagent package failure | Flow selection calibration error, 1177 | Repair after replacement of reagent kit
No | Visit | Complaint | RCHN | 无法开机 | 无法开机,借910-342，910-285各一个。 | 更换备件后仍未修复。再借910-285，910-286各一个。 | SWO-303450 | Defect | HW | Power Supply | Power Supply |  | Cannot power on | Cannot be switched on, borrow 910-342, 910-285 each. | The replacement of the spare parts has not been repaired. Borrow 910-285, 910-286 each.
No | Visit | Complaint | RCHN | 测试卡卡托脱落 | 测试卡卡托脱落，借886-625 3个 | 更换886-625 3个正常。 by jesson | SWO-302857 | Defect | HW | Storage Module | Storage Wheel |  | Test Kakato fall off | Test Kakato falls off, borrow886-625 3 | Replace 886-625 3 normal. by jesson
No | Visit | Complaint | RCHN | 软件错误，定标数据丢失 | 软件错误，定标数据丢失，借944-212（13719） | 重新定标正常。by jesson | SWO-303330 | Defect | SW | Windows | Windows | LEH Mail, it is not OOB Jesson Xing | Software error, calibration data loss | Software error, calibration data loss, borrowing 944-212 (13719) | Recalibration is normal. by jesson
Yes | Visit | Complaint | RCHN | 无法开机 | 无法开机，借910-285，910-286各一个。 |  | SWO-303954 | Re classified RMED | Double entry, see previous day |  |  |  | Cannot power on | Cannot be switched on, borrow 910-285, 910-286 each.
No | Visit | Complaint | RCHN | 定标失败 1054 | 定标失败 1054，借944-212（13875）一张。 |  | SWO-303962 | Defect | HW | Process Module | Optical Unit | 1024 | Bid failure 1054 | Bid failure 1054, borrow 944-212 (13875) one.
No | 10 | Y | RCZE | Blank |  |  | D3D85D2414169943 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway | Clot | Blank
Yes | 4 | Y | RPAC | Blank | Failed calibration and QC on sensor cassette. Customer requested more Solution packs as they had none left. Solution Packs CPT. |  | D3D85F4B140E2801 | [5 empty] | Blank | Failed calibration and QC on sensor cassette. Customer requested more Solution packs as they had none left. Solution Packs CPT.
Yes | 6 | Y | RPAC | Blank | Failed sensor cassette |  | D3D86C8C14277E8F | [5 empty] | Blank | Failed sensor cassette
No | Visit | Complaint | RAME | Analyzer is making a loud noise | Troubleshot to a faulty power supply. |  | SWO-304576 | Defect | HW | Basic Unit | Power Supply |  | Analyzer is making a loud noise | Troubleshot to a faulty power supply.
No | Visit | Complaint | RCHN | 处理单元故障 | 处理单元故障，借636-550一个 |  | SWO-304032 | Defect | HW | Process Module | Sensor, J404 P116 |  | Handling unit failure | Handling unit failure, borrow 636-550 one
No | Visit | Complaint | RCHN | 分析仪报错-5784 | 分析仪报错-5784 | 更换636-550一个，这个传感器是L09的。运行季度保养 | SWO-304625 | Defect | HW | Process Module | Sensor, Process unit |  | Analyzer Error-5784 | Analyzer Error-5784 | Replace the 636-550 one, which is L09. Run quarterly maintenance
No | Visit | Complaint | RCHN | 定标扫码扫不上 | 定标扫码扫不上，借636-571,902-924各一个 |  | SWO-304019 | Defect | HW | Analyzer Control Module | Cable, Display |  | The calibration sweep code can't be swept | The mark-sweeping code is not clear, borrowing 636-571, 902-924 each
No | Visit | Complaint | RCHN | 定标失败 | 定标失败，借944-394（13715）一张。 | 更换944-394（13715）重新定标后修复 | SWO-304798 | Assisted Replacement | Consumables | Test Kit | Cartridge, PCT |  | Bid failed | The bid failed, borrowing 944-394 (13715). | Replacement 944-394 (13715) re-calibration repair
No | Visit | Complaint | RCHN | 新机器频繁报1279-测得液体流量值高于临界限 | 新机器频繁报1279-测得液体流量值高于临界限 | 调换P2泵头 | SWO-304624 | Defect | HW | Wet section | Pump, P2 | 1279 | New machine frequently reported 1279 - measured liquid flow value is higher than the limit | New machine frequently reported 1279 - measured liquid flow value is higher than the limit | Switching the P2 pump head
No | Visit | Complaint | RFRA | Plus de pco2 ni PH | Vérification des activités et calibrations Vérification en mode service | Multiple rinçage Pompage des solutions Calibrations | SWO-304533 | [5 empty] | More pco2 nor PH | Checking activities and calibrations Checking in service mode | Multiple rinse Pumping solutions Calibrations
No | Visit | Complaint | RIND | Instrument is showing sample problem in some samples. | Checked Inlet gasket and sample probe tubing. | Cleaned Inlet gasket with hypochloride solution. Then given protein removal program. Checked with samples. Unit is working fine. | SWO-306666 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | Instrument is showing sample problem in some samples. | Checked Inlet gasket and sample probe tubing. | Cleaned Inlet gasket with hypochloride solution. Then given protein removal program. Checked with samples. Unit is working fine.
No | Visit | Complaint | RDEU | QC Fehler Level 3 in Oximetrie | Überprüfung Level 2 Messung in Oximetrie zu niedrig Oxi Spectrum missmatch | Oxi Kompletttest durchgeführt, Fehler Neon Intensity 46% Lichtwerte 3100, Oximetrie geöffnet, Sicherungsschraube für Halogenlampe lose. Lampeneinheit auf 3300 justiert (sehr sprunghaft) Neustart Kalibration durchgeführt. tHb Kal QC ohne Beanstandung Meldung wird Beobachtet | SWO-304384 | Adjustment | HW | Oxi module | Lamp unit | Screw loose | QC Fehler Level 3 in Oximetrie | Review Level 2 measurement in oximetry too low Oxi Spectrum missmatch | Oxi Complete Test Performed, Error Neon Intensity 46% Light values 3100, oximetry open, safety screw for halogen lamp loose. Lamp unit adjusted to 3300 (very erratic) Restart calibration. tHb Kal QC without complaint Message is observed
No | 8 | Y | RCZE | Blank | výpadky v ovládání display a komunikaci s AU |  | D3D84A82143B79F8 | Defect | SW | Windows | Win7 |  | Blank | Outages in control display and communication with AU
No | Visit | Complaint | RAME | S02 is out of range | S02 is out of range | Customer reported that the S02 is at the bottom of the range and failing level 2 QC only. Reviewed data logs and found that the S02 has been slowly dropping on level 2 from the mean for the last month. Customer ran a protein removal but the results are the same. Ran a complete oxi test and found all levels fine, lamp value is 3300. Ran service thb cal and user thb cal, found all results passed. Ran QC level 2 and found that the result is now back in the mean. Will have customer track the results and see if it begins to fall again. | SWO-306792 | Adjustment | HW | Oxi module | tHb Cal |  | S02 is out of range | S02 is out of range | Customer reported that the S02 is at the bottom of the range and failing level 2 QC only. Reviewed data logs and found that the S02 has been slowly dropping on level 2 from the mean for the last month. Customer ran a protein removal but the results are the same. Ran a complete oxi test and found all levels fine, lamp value is 3300. Ran service thb cal and user thb cal, found all results passed. Ran QC level 2 and found that the result is now back in the mean. Will have customer track the results and see if it begins to fall again.
No | Visit | Complaint | RDEU | 05.02.18 11:00Uhr Frau Sandra Fomiatto Drucker druckt nicht | Drucker mechanisch defekt Andruckrolle dreht nicht mit | Druckermodul ausgebaut, Mechanik def, Druckermodul ersetzt, Kal und QC durchgeführt | SWO-305098 | Defect | HW | Printer Module | Printer Head |  | 05.02.18 11:00pm Mrs Sandra Fomiatto Printer does not print | Printer mechanically defective pressure roller does not rotate with | Printer module removed, mechanics def, printer module replaced, Kal and QC performed
No | Visit | Complaint | RSCH | Re. Ca | Immer wieder Ca Drift | Allgemeine Kontrolle i.o Ref. Elektrode ersetzt i.o Alte Lot = 6807 => Garantie O-Ring pH ersetzt i.o Kal. und QC i.o System i.o | SWO-304451 | Defect | Electrode | Electrode | Sealing ring, pH | pH | Re. Ca | Again and again Ca Drift | General control i.o Ref. Electrode replaces i.o Old lot = 6807 => Warranty O-ring pH replaces i.o Kal. and QC i.o System i.o
Yes | Visit | Complaint | RSCH | Rep. Ca | Immer wieder Ca Drift K Messung instabil | Allgemeine Kontrolle i.o Ref. Elektrode ersetzt i.o Alte Lot= 6807 => Garantie O-Ring pH ersetzt i.o Membranenwartung K durchgeführt i.o Kal. und QC i.o System i.o | SWO-304452 | Re classified RMED | Look left/right (SWO-304451) | WR decision in W7 | Sealing ring | pH | Rep. Ca | Again and again Ca Drift K Measurement unstable | General control i.o Ref. Electrode replaces i.o Old Lot= 6807 => Warranty O-ring pH replaces i.o Membrane maintenance K carried out i.o Kal. and QC i.o System i.o
Yes | Visit | Complaint | RSCH | Rep. Ca | Immer wieder Ca Drift | Allgemeine Kontrolle i.o Ref. Elektrode ersetzt i.o Alte Lot= 6807 => Garantie O-Ring pH ersetzt i.o Kal. und QC i.o System i.o | SWO-304449 | Re classified RMED | Look left/right (SWO-304451) | WR decision in W7 | Sealing ring | pH | Rep. Ca | Again and again Ca Drift | General control i.o Ref. Electrode replaces i.o Old Lot= 6807 => Warranty O-ring pH replaces i.o Kal. and QC i.o System i.o
Yes | Visit | Complaint | RSCH | Rep. Ca | Immer wieder Ca Drift | Allgemeine Kontrolle i.o Ref. Elektrode ersetzt i.o Alte Lot= 6809 => Garantie O-Ring pH ersetzt i.o Kal. und QC i.o System i.o | SWO-304450 | Re classified RMED | Look left/right (SWO-304451) | WR decision in W7 | Sealing ring | pH | Rep. Ca | Again and again Ca Drift | General control i.o Ref. Electrode replaces i.o Old Lot= 6809 => Warranty O-ring pH replaces i.o Kal. and QC i.o System i.o
Yes | Visit | Complaint | RSCH | Rep. pH | Immer wieder pH instabil Status und Empfindlichkeit | Bereits O-Ring pH ersetzt und Ref. membraniert Alles ohne Erfolg pH Elektrode durch Medizintechnik ersetzt i.o Elektrode von Medizintechnik Alte Lot = YT04 => Garantie Kal. und QC i.o System i.o | SWO-304453 | Re classified RMED | Look left/right (SWO-304451) | WR decision in W7 | Sealing ring | pH | Rep. pH | Time and again pH unstable Status and sensitivity | Already O-ring pH replaced and ref. membraneed All without success pH electrode replaced by medical technology i.o Electrode of medical technology Old lot = YT04 => Warranty Kal. and QC i.o System i.o
Yes | 1 | Y | RPAC | Blank | SC and SP for start up |  | D3D9AB348A602D4C | Re classified RMED | Installation/Setup |  |  |  | Blank | SC and SP for start up
No | Visit | Complaint | RDEU | 31.01.2018; Fr. Rohleder; Tel.: 02611371447; Intervention erforderlich, Lösungen nicht gepumpt werden 31.01.2018 15:15 Fr. Rohleder 02611377310: weiterhin Interventionsmeldung Leckagefehler | Einlassmodul ersetzt. | Startphase, mehrere Messungen durchgeführt I.O. | SWO-304845 | Defect | HW | Auto Inlet module | Auto Inlet module |  | 31.01.2018; Ms. Raw leather; Tel.: 02611371447; Intervention required, solutions not to be pumped 31.01.2018 15:15 Fr. Raw leather 02611377310: continue intervention notification Leakage errors | inlet module. | Start phase, several measurements performed I.O.
No | Visit | Complaint | RCHN | 进样模块自动脱落 | 进样模块自动脱落，借903-139一个 | 使用备用件903-139后修复by tim | SWO-304173 | Defect | HW | Basic Unit | Inlet module |  | The sample module comes off automatically | The sample module automatically falls off, borrowing 903-139 one | Fix by tim after using spare parts 903-139
No | Visit | Complaint | RCHN | 屏幕不亮 | 屏幕不亮，借636-571一个 | 更换636-571后by tim | SWO-304158 | Defect | HW | DMS_PC | Cable, Display |  | The screen is not lit | The screen is not bright, borrow 636-571 one | Replacement 636-571 post-by-tim
Yes | 13 | Y | RPAC | Blank | ref parameters drifting over last week |  | D3D983718A527CFB | Assisted Replacement | Electrode | Electrode | Reference |  | Blank | ref parameters drifting over last week
No | Visit | Complaint | RIND | Barcode reader not working properly. | found dust on barcode reader. | Cleaned the barcode reader and cleaned the dust in instrument but barcode not working it need to replace with new barcode. | SWO-309966 | Cleaned | HW | Basic Unit | Barcode reader, Basic Unit |  | Barcode reader not working properly. | found dust on barcode reader. | Cleaned the barcode reader and cleaned the dust in instrument but barcode not working it need to replace with new barcode.
Yes | Visit | Complaint | RSCH | Rep. Ca | Vermehrt Ca Drift | Allgemeine Kontrolle i.o O-Ring pH ersetzt i.o Ref. Elektrode ersetzt i.o Material von Medizintechnik Alte Lot= 8801 => Garantie Kal. und QC i.o | SWO-304936 | Re classified RMED | Look left/right (SWO-304451) | WR decision in W7 | Sealing ring | pH | Rep. Ca | Increases Ca Drift | General control i.o O-ring pH replaces i.o Ref. Electrode replaces i.o Material of medical technology Old Lot= 8801 => Warranty Kal. and QC i.o
Yes | Visit | Complaint | RSCH | Rep. Ca | Vermehrt Ca Drift | Allgemeine Kontrolle i.o Ref. Elektrode ersetzt i.o Material von Medizintechnik Alte Lot = 8807 => Garantie Kal. und QC i.o System i.o | SWO-304935 | Re classified RMED | Look left/right (SWO-304451) | WR decision in W7 | Sealing ring | pH | Rep. Ca | Increases Ca Drift | General control i.o Ref. Electrode replaces i.o Material of medical technology Old lot = 8807 => Warranty Kal. and QC i.o System i.o
No | Visit | Complaint | RDEU | 06.02.2018 Herr Smöra 023031068221 meldet einen Spülfehler. | Spülfehler. Flüssigkeit im Air-Ventil des EL $Met-Moduls. | Ventile der Messmodule überprüft. Beide Spülventile der Messmodule ersetzt, da defekt. Gerät einkalibriert. Funktionskontrolle: I.O. | SWO-304934 | Defect | HW | El/Met Module | Valve, air |  | 06.02.2018 Mr. Smöra 023031068221 reports a flushing error. | Flushing error. liquid in the air valve of the EL /Met module. | valves of the measuring modules. Replaced both flushing valves of the measuring modules, as defective. device calibrated. Functional control: I.O.
No | Visit | Complaint | RIND | Sample are getting aborted. | Sample are getting aborted. | Check the analzyer found pipett unit isse, removed and clened the sensors restarted but problem was reamining same. took the SD awaiting for responce. | SWO-306943 | Revisit required | Revisit required | Revisit required | RMED Contacted |  | Sample are getting aborted. | Sample are getting aborted. | Check the analzyer found pipett unit isse, removed and clened the sensors restarted but problem was reamining same. took the SD awaiting for responce.
No | Visit | Complaint | RSAF | Inlet stuck in open position gave telephonic assistance but feedback was everything is ok. | Telephonic trouble shooting never helped as feed back was given as everything is ok. | done trouble shooting in service mode and inspected inlet probe and found it bent. Installed new inlet Probe and inlet gasket. | SWO-306102 | Defect | HW | Auto Inlet module | Probe bend |  | Inlet stuck in open position gave telephonic assistance but feedback was everything is ok. | Telephonic trouble shooting never helped as feed back was given as everything is ok. | done trouble shooting in service mode and inspected inlet probe and found it bent. Installed new inlet Probe and inlet gasket.
No | Visit | Complaint | RDEU | 06.02.18/09:00 Frau Oslislo Tel.: 029413006 -> Kd meldet Intervention -> Ableitströme nach SP90-Wechsel 06.02.18/09:20 Frau Oslislo -> Kd meldet erneut Ableitstsröme | Ableitströme am LP erkannt, Kunde hat schon LP ersetzt, Fehler tritt immer wieder auf, | SC ersetzt, Gerät kalibriert, | SWO-305590 | Assisted Replacement | SC | SC | Leak current |  | 06.02.18/09:00 Ms Oslislo Tel.: 029413006 -> Kd reports intervention -> Discharge flows after SP90 change 06.02.18/09:20 Mrs. Oslislo -> Kd reports again Ableitstsröme | discharge currents on the LP detected, customer has already replaced LP, errors occur again and again, | SC, device calibrated,
Yes | 1 | Y | RCZE | Blank |  |  | D3DC160D76377DFD | Re classified RMED | Setup His/Lis |  |  |  | Blank
Yes | 1 | Y | RCZE | Blank |  |  | D3DC14EE76097DF9 | Re classified RMED | Setup His/Lis |  |  |  | Blank
No | Visit | Complaint | RAME | Error code 292 and 581 | Error code 292 and 581 | Replaced Hemolyzer assembly | SWO-306283 | [5 empty] | Error code 292 and 581 | Error code 292 and 581 | Replaced Hemolyzer assembly
No | Visit | Complaint | RDEU | pH wird rot angezeigt | pH rot in Parameterleiste Ampelstatus grün | Überprüfung Dateiarchiv Kal, QC und Aktivitäten QC und Kal ohne pH Fehler => pH wurde über Radiance gesperrt Admin-Login in Radiance und Parameter wieder entsperrt Nach Uhrzeitanpassung wird LP als verfallen angezeigt => LP zurückgesetzt und reinstalliert | SWO-305430 | [5 empty] | pH is displayed in red | pH red in parameter bar Traffic light status green | Review File Archive Kal, QC and Activities QC and Kal without pH error => pH was locked via radiance Admin login in Radiance and parameters unlocked again After time adjustment, LP is displayed as forfeited => LP resets and reinstalled
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | GLU,LAC定标失败，0378 | 连续更换5张测试卡后修复 | SWO-305354 | Assisted Replacement | SC | SC | Glu/Lac | 378 | Test card, reagent package failure | GLU, LAC calibration failure, 0378 | Repair after 5 consecutive test cards
Yes | Visit | Complaint | RCHN | 泵定标失败 | 泵定标失败 | 更换试剂包后修复 | SWO-305376 | [5 empty] | Pump calibration failed | Pump calibration failed | Repair after replacement of reagent kit
No | Visit | Complaint | RCHN | 报1029错误 | 报1029错误，借903-364和944-212（13719）、944-258（13756）各一个 |  | SWO-305224 | Defect | HW | Process Module | Optical Unit | 1029 | Reported 1029 error | 1029 errors, borrowing 903-364 and 944-212 (13719), 944-258 (13756) each
No | Visit | Complaint | RCHN | 样本终止1067 | 样本终止1067 | 清洁后修复 by tim | SWO-305812 | Cleaned |  |  |  | 1067 | Sample termination 1067 | Sample termination 1067 | Repair ingress after cleaning by tim
No | Visit | Complaint | RIND | PH and electrolytes unstable. | Check the machine found ph and electrolytes parameters is not stable. | Clean all measuring chambers, change ref membrane, run 2 point calibration but problem is same than we swap the ref electrode with another machine installed in ICU. the problem of unstable calib artion is shifted to icu. Reference electrode is found faulty. | SWO-305990 | Assisted Replacement | Electrode | Electrode | Reference |  | PH and electrolytes unstable. | Check the machine found ph and electrolytes parameters is not stable. | Clean all measuring chambers, change ref membrane, run 2 point calibration but problem is same than we swap the ref electrode with another machine installed in ICU. the problem of unstable calib artion is shifted to icu. Reference electrode is found faulty.
Yes | Visit | Complaint | RIND | Sample are getting aborted. | Sample are getting aborted. | Found repalced the Pipett PCB, restarted analyzer found issue is reamining same. awaiting for reply from RMED. | SWO-306944 | Re classified RMED |  |  |  |  | Sample are getting aborted. | Sample are getting aborted. | Found repalced the Pipett PCB, restarted analyzer found issue is reamining same. awaiting for reply from RMED.
No | Visit | Complaint | RDEU | 05.02.2018; Pfl. David; Tel.: 020916053550; Testkassetten werden nicht akzeptiert 5.2.18/14:30 Pfl. David -> keine Kal-Kassetten vorhanden für CKMB/D-Dimer | keine Kal.kassetten für D-Dimer und Ck-MB, alte Kal.werte sind gelöscht, Trop.T ist o.k., | Material zusammengestellt, Reinigung durchgeführt, Systemüberprüfung durchgeführt, D-dimer und Ck-MB neu kalibriert und aufgefüllt, | SWO-305591 | Cleaned | HW | Storage Module | Barcode reader, storage wheel |  | 05.02.2018; Pfl. David; Tel.: 020916053550; Test cartridges are not accepted 5.2.18/14:30 Pfl. David -> no Kal cassettes available for CKMB/D-Dimer | no Kal.cassettes for D-Dimer and Ck-MB, old Cal.values are deleted, Trop.T is OK, | Material assembled, cleaning carried out, system verification carried out, D-dimer and Ck-MB recalibrated and refilled,
No | Visit | Complaint | RDEU | 06.02.2019 9:15 Fr. Frühauf 051616021560: Display hat Artefakte/verpixelt | Display soll Fehler haben. | Das Display ist völlig in Ordnung. Es geht um das sporadische abschalten des Analysator. Der Raum wird sehr warm. Die Tür zum POCT Labor in der ZNA soll immer verschlossenen bleiben. Das kann im Winter und im Sommer zu sehr hohen Temperaturen in diesen Raum führen und zu einem häufigen abschalten und zu anderen Fehlermeldungen führen. Der Analysator arbeitet einwandfrei. | SWO-305341 | User error | Location vs Temp | Location vs Temp | Location vs Temp |  | 06.02.2019 9:15 Fr. early on 051616021560: Display has artifacts/pixelated | Display should have errors. | The display is perfectly fine. It's about the sporadic shutdown of the analyzer. The room becomes very warm. The door to the POCT laboratory in the ZNA should always remain closed. This can lead to very high temperatures in this room in winter and summer and can lead to a frequent shutdown and other error messages. The analyzer works flawlessly.
No | Visit | Complaint | RAME | Unable to pump solution after replacing sensor cassette and solution pack | Unable to pump solution after replacing sensor cassette and solution pack. Failing vacuum under pressure test. Pulled sensor cassette off and found liquid in the SC bay. Put in new SC and it worked. Paula is leaving and will let analyzer run overnight and will run external QC in the morning | Performed backflush. Tried solution pack. Replaced SC with customer stock. Went thru all service fluid program and all passed. Analyzer is calibrating. Already down to 40 minutes. | SWO-306291 | Assisted Replacement | SC | SC | Leakage |  | Unable to pump solution after replacing sensor cassette and solution pack | Unable to pump solution after replacing sensor cassette and solution pack. Failing vacuum under pressure test. Pulled sensor cassette off and found liquid in the SC bay. Put in new SC and it worked. Paula is leaving and will let analyzer run overnight and will run external QC in the morning | Performed backflush. Tried solution pack. Replaced SC with customer stock. Went thru all service fluid program and all passed. Analyzer is calibrating. Already down to 40 minutes.
No | Visit | Complaint | RFRA | Fuite de courant détectée | Contrôle sol pack Contrôle circuit fluides | Nettoyage circuit fluides Échange joint introduction Remplacer sol pack Remplacer sensor K7 | SWO-306278 | Assisted Replacement | SC | SC | Leak current |  | Electricity leak detected | Ground control pack Fluid circuit control | Fluid circuit cleaning Exchange joint introduction Replace ground pack Replace sensor K7
No | Visit | Complaint | RFRA | Fuite | Contrôle du pack, du bloc de vanne et du manifold de la vidange | Changement du manifold de vidange. Essais | SWO-305980 | [5 empty] | running away | Control of the pack, valve block and drain manifold | Change of the drain manifold. Tests
No | Visit | Complaint | RESP | OBSTRUCCION. | OBSTRUCCION. | SE INTENTA LIMPIEZA SIENDO IMPOSIBLE POR OBSTRUCCION EN CIRCUITO DE NEUMATICA SE SUSTITUYEN DIFERENTES PIEZA DE LA ENTRADA Y SE LE EFECTUA UN KIT DE SERVICIO. CALIBRADO Y VERIFICADO CON Q.C. | SWO-306187 | Defect | HW | PM kit | PM kit |  | OBSTRUCCION. | OBSTRUCCION. | CLEANING IS ATTEMPTED BY BEING IMPOSSIBLE BY OBSTRUCTION IN NEUMATIC Circuit DIFFERENT PART OF THE INPUT IS SUBSTITUTE AND A SERVICE KIT IS EFFECTED. CALIBRATED AND VERIFIED WITH Q.C.
No | Visit | Complaint | RCHN | 血氧定标失败 | 血氧定标失败 | 更换备用的夹紧阀（841-811）后修复by tim | SWO-305381 | Defect | HW | Oxi module | Pinch Valve |  | Blood oxygen calibration failure | Blood oxygen calibration failure | Repair by tim after replacing a spare clamping valve (841-811)
No | Visit | Complaint | RESP | FUGAS | FUGAS. | SUSTITUCION Y LIMPIEZA DE TUBOS Y PRUEBAS POR SERVIVIO TECNICO, SE DETECTA DEFECTO EN TUBO OXI-BAHIA. CALIBRADO Y VERIFICADO CON Q.C. | SWO-306188 | Defect | HW | Oxi module | Pinch valve tube |  | Leaks | Leaks. | SUBSTITUTE AND CLEANING OF TUBES AND TESTS FOR TECHNICAL SERVIVIO, DEFECT IS DETECTED IN OXI-BAHIA TUBE. CALIBRATED AND VERIFIED WITH Q.C.
No | Visit | Complaint | RCHN | 报错1031 | 报错1031，借903-364一个 | 更换903-364和905-908后修复by tim | SWO-305195 | Defect | HW | Process Module | Optical Unit | 1031 | 1031 | 1031, loan 903-364 one | Repair by tim after replacing 903-364 and 905-908
Yes | Visit | Complaint | RCHN | 报错1031 | 加借905-908一个 | 更换903-364及905-908后修复 | SWO-305308 | Re classified RMED | Double entry |  |  |  | 1031 | Borrow 905-908 one | Replacement 903-364 and 905-908 after repair
No | Visit | Complaint | RBNL | Problème capteur | Le monitor indique souvent un problème de capteur : 'capteur déconnecté du patient' Le client me confirme que c'est un capteur ancien qui a été utilisé sur un autre monitor avant. Les enfants tirent souvent à la cable pendant le test. | Test du capteur , le problème est lié au signal de saturation et pouls. Il faut donc commander un autre capteur. N° série du capteur 92-67025 défectueux. | SWO-306396 | [5 empty] | Sensor problem | The monitor often indicates a sensor problem: 'patient's disconnected sensor' The client confirms to me that it is an old sensor that has been used on another monitor before. Children often shoot at the cable during the test. | Sensor test, the problem is related to the saturation signal and pulse. So you have to order another sensor. Sensor number 92-67025 defective.
No | 2 | Y | RPOL | Blank | Testy monitora. Czujnik SN 92-78859 zabrany do reklamacji. Nowy czujnik zamówiony - zostanie dostarczony w dniu 09.02.2018. |  | D3DC1442762969EA | [5 empty] | Blank | Monitor tests. The SN 92-78859 sensor is taken to the complaint. New sensor ordered-will be delivered on 09.02.2018.
No | Visit | Complaint | RESP | Problemas de estabilización. | Sens. baja en electrodos de pCO2 y Lac. | Comprobación del estado del analizador, se comprueban registros, se verifica estado de electrodos, camara de medida y pre-amplificadores, se ve necesario el cambio de membrana de Lac, se realizan pruebas continua pCO2 con sens. baja y estado muy alto, se cambia electrodo de pCO2. Se realizan varias CAl. así como Qc´s. Funcionamiento correcto. | SWO-306376 | Assisted Replacement | Electrode | Electrode | pCO2 |  | Stabilization problems. | Sens. low in pCO2 and Lac electrodes. | Checking the condition of the analyzer, checking records, checking the status of electrodes, measuring camera and pre-amplifiers, it is necessary to change the lac membrane, continuous pCO2 tests are performed with sens. low and very high state, pCO2 electrode is changed. Several CAl. as well as Qc's. Proper Operation are performed.
Yes | 7 | Y | RCZE | Blank |  |  | D3DC5A7C76599672 | Re classified RMED | FAN |  |  |  | Blank
No | Visit | Complaint | RKK | Naのﾄﾞﾘﾌﾄが大きい。 | 比較電極交換にて落ち着いている様子。 | 比較電極交換。ユーザー備品を使用 | SWO-311470 | Assisted Replacement | Electrode | Electrode | Reference |  | Na drift is large. | It seems to be calm by the comparison electrode exchange. | Comparison electrode replacement. Use user equipment
No | Visit | Complaint | RKK | NaとCaのﾄﾞﾘﾌﾄが大きく、ﾄﾞﾘﾌﾄで外れる。 | 比較電極を交換したところ落ち着いていた。 | 比較電極交換。(ﾕｰｻﾞｰ予備品使用) | SWO-311411 | Assisted Replacement | Electrode | Electrode | Reference |  | The drift of Na and Ca is large, and it comes off by drift. | When i replaced the comparison electrodes, I was calm. | Comparison electrode replacement. (Use of user spare parts)
No | Visit | Complaint | RAME | Customer concerned there was a shift in PCO2 starting Jan 14th a. 8:58 am | 7735 lot 681 pCO2 peer 2 range is 66.2 -69.1 7745 Lot 666 pCO2 peer 2 range is 39.8 - 41.5 7755 Lot 720 pCO2 peer 2 range is 21.7 - 22.7 On Jan 14th at 6:00am Level one was running at 66.9--At 7:35 am the Ref membrane was replaced--at 8:58 level one ran at 70.0 showing a shift. Inspected the Ref electrode and found a scratch on the tip of the electrode--went through the trend on QC for pCO2 it is fluctuating between 68.1 to 70.2 for level 1. On Jan 16th QC and cals had failures, failing for sample errors 722 and 374 in homogeneous sample in ph $Bg module, Customer was also concerned about the tHb failing -- found 2 failures | Replaced the Ref electrode (lot 18-03) with Lot 18-15 customer had on hand, found the pCO2 membrane with little fluid--replaced the measuring chamber and pCO2 membrane, also found a clot in the pH chamber. Ran the gas check in service, electrode check, complete oxy test -- lamp unit is set to 3212 and ran the tHb calibration, The customer is running a 2 pt every 4 hours and a 1 pt every 1 hours and gas cal every 30 minutes; Talked to the customer and adjusted the cals to 8 hrs for 2 pt. 4 hrs for 1 cal. Level 2 and 3 are running high outside the peer 2 range after all the maintenance was performed, Discussed this with Elaine -- letting the analyzer stabilize and will call back, Called back at 5 pm and level 2 and 3 are now within the peer 2 range. | SWO-306357 | Cleaned | Fluidic pathway | Fluidic pathway | Measuring chamber |  | Customer concerned there was a shift in PCO2 starting Jan 14th a. 8:58 am | 7735 lot 681 pCO2 peer 2 range is 66.2 -69.1 7745 Lot 666 pCO2 peer 2 range is 39.8 - 41.5 7755 Lot 720 pCO2 peer 2 range is 21.7 - 22.7 On Jan 14th at 6:00am Level one was running at 66.9--At 7:35 am the Ref membrane was replaced--at 8:58 level one ran at 70.0 showing a shift. Inspected the Ref electrode and found a scratch on the tip of the electrode--went through the trend on QC for pCO2 it is fluctuating between 68.1 to 70.2 for level 1. On Jan 16th QC and cals had failures, failing for sample errors 722 and 374 in homogeneous sample in ph /Bg module, Customer was also concerned about the tHb failing -- found 2 failures | Replaced the Ref electrode (lot 18-03) with Lot 18-15 customer had on hand, found the pCO2 membrane with little fluid--replaced the measuring chamber and pCO2 membrane, also found a clot in the pH chamber. Ran the gas check in service, electrode check, complete oxy test -- lamp unit is set to 3212 and ran the tHb calibration, The customer is running a 2 pt every 4 hours and a 1 pt every 1 hours and gas cal every 30 minutes; Talked to the customer and adjusted the cals to 8 hrs for 2 pt. 4 hrs for 1 cal. Level 2 and 3 are running high outside the peer 2 range after all the maintenance was performed, Discussed this with Elaine -- letting the analyzer stabilize and will call back, Called back at 5 pm and level 2 and 3 are now within the peer 2 range.
No | Visit | Complaint | RDEU | AQC temp. Fehler | AQC Temp. Fehler 735, 736 | Gerät und AQC Modul kontrolliert, AQC Modul erneuert, Software installiert | SWO-306054 | Defect | HW | Autocheck Module | Autocheck Module |  | AQC temp. Error | AQC Temp. Error 735, 736 | Device and AQC module controlled, AQC module renewed, software installed
No | Visit | Complaint | RLTD | Rebecca Tel 0161 9226498 | On previous advice customer had found a cartridge guide out of its fixings, having refitted it the analyser was working. | Enter service, fit new cartridge guide support after finding which guide was faulty ( lug on single clip had broken off ) run safe positions, exit service, analyser ready for use. | SWO-306093 | Defect | HW | Storage Module | Cartridge guide |  | Rebecca Tel 0161 9226498 | On previous advice customer had found a cartridge guide out of its fixings, having refitted it the analyser was working. | Enter service, fit new cartridge guide support after finding which guide was faulty ( lug on single clip had broken off ) run safe positions, exit service, analyser ready for use.
No | Visit | Complaint | RDEU | 09.02.18 08:00Uhr Sr. Katharina 03303/5222870 LP wird nicht erkannt | Kal.-Fehler bei pH u. Lac | Fehlersuche; div. Serviceprogramme angewandt; Kal. = Fehler; SC gew.; Startphase abgewartet; Kal. + QC i.O. | SWO-306628 | Assisted Replacement | SC | SC | pH/Lac |  | 09.02.18 08:00H Sr. Katharina 03303/5222870 LP is not detected | Cal.-Error at pH and Lac | Troubleshooting; various service programs; Kal. = Error; SC. start-up phase awaited; Kal. + QC in the first place
No | Visit | Complaint | RDEU | Touchscreen ist nicht Richtig bedienbar | Touchscreen ist nicht Richtig bedienbar | Fehleranalyse, Touchscreen justiert, Kalibrierung und Qualitätskontrolle i. O. | SWO-306904 | Adjustment | HW | Basic Unit | Touch screen |  | Touchscreen is not properly operated | Touchscreen is not properly operated | Fault analysis, touchscreen adjusted, calibration and quality control in the
No | Visit | Complaint | RFRA | La machine n'aspire pas l'échantillon | Contrôle du site d'introduction et des différentes positions de l'inlet | Changement de la carte de détection des positions de l'inlet. Essais | SWO-306621 | Defect | HW | Auto Inlet module | PCB, Auto inlet |  | The machine does not suck up the sample | Control of the introductory site and the different positions of the inlet | Change of the inlet position detection map. Tests
No | Visit | Complaint | RFRA | Plantage de l'analyseur sur Erreur Code 2 | Suppression de la base de données impossible | Reinstallation logiciel version 3.3MR1 avec suppression base de données et configuration Restauration de la configuration Restauration de la sauvegarde du 7 $02 $2018 Relance analyseur Test de connexion OK | SWO-306557 | Defect | SW | RMED SW | RMED SW | Code 2 | Crashing the analyzer on Error Code 2 | Removing the impossible database | Software reinstallation version 3.3MR1 with database deletion and configuration Restoring the configuration Backup Restoration of $7 $02 2018 Relaunch analyzer OK connection test
No | Visit | Complaint | RDEU | 08.02.18 13:30Uhr Frau Laura Beschmann 040/790202395 Spülung erforderlich 08.02.18/15:15 Herr Jürgensen - > 040790202395 -> Intervention erforderlich | Intervention Drucküberprüfung Vakuumfehler, Gerät steht im Interventionsmodus Einlassdichtung erneuern | Fehleranalyse, Servicetests durchgeführt, Vakuumfehler vorhanden, Lufteinschlüsse im Visual Flow Check, Anschlussdichtung (Kundenmaterial) und Einlassdichtung erneuert danach Tests i. O. , Kalibrierung und Qualitätskontrolle i. O. | SWO-306439 | Defect | HW | Auto Inlet module | Inlet gasket |  | 08.02.18 13:30Pm Mrs Laura Beschmann 040/790202395 Rinse required 08.02.18/15:15 Mr Jürgensen - > 040790202395 -> Intervention required | Intervention pressure verification vacuum fault, device is in intervention mode to renew inlet seal | Fault analysis, service tests performed, vacuum faults present, air inclusions in the Visual Flow Check, connection seal (customer material) and inlet seal then renews tests according to the o. , calibration and quality control as a standard.
No | Visit | Complaint | RCHN | 流量选择器故障 | 流量选择器故障 | 更换903-185后修复(By Ran) | SWO-307719 | Defect | HW | Basic Unit | PCB, Flow Selector |  | Flow Selector Fault | Flow Selector Fault | Repair after replacement 903-185 (By Ran)
No | Visit | Complaint | RCHN | 无法读取智能芯片 | 无法读取智能芯片 | 更换903-158后修复 by jesson | SWO-313683 | Defect | HW | Basic Unit | PCB, Solution Pack PCB Id |  | Unable to read smart chip | Unable to read smart chip | Repair by jesson after replacement 903-158
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 压力测试真空错误 | 更换2个试剂包后修复 | SWO-307056 | [5 empty] | Test card, reagent package failure | Pressure test vacuum error | Repair after replacing 2 reagent kits
No | Visit | Complaint | RCHN | 光学单元定位错误 | 光学单元定位错误，借636-550一个 | 更换636-550后修复bytim | SWO-305228 | Defect | HW | Process Module | Sensor, Process Unit |  | Optical unit positioning error | Optical unit positioning error, borrow 636-550 one | Repair bytim after replacement 636-550
No | Visit | Complaint | RCHN | 报错1029 | 报错1029，借903-364，944-394（13418缺货），944-214（13673缺货），944-258（13691），944-220（13459）（13612），944-212（13871），944-216（13707）各一个 |  | SWO-306464 | Defect | HW | Process Module | Optical Unit | 1029 | 1029 error | 1029, 903-364,944-394 (13418 out of stock), 944-214 (13673 out of stock), 944-258 (13691), 944-220 (13459) (13612), 944-212 (13871), 944-216 (13707) each
Yes | Visit | Complaint | RCHN | 报错1029 | 加借944-220（13612）一个 |  | SWO-306478 | Re classified RMED | double entry |  |  |  | 1029 error | Borrow 944-220 (13612) one
No | Visit | Complaint | RAME | Autocheck module not able to run qc. QC being aborted. | No problem detected when I arrived. Customer reports autocheck module moving, but no sample detected. Typically this problem indicates 14 pole cable having a short. I also found Oxi liquid sensor error. | Replaced 14 pole cable. Checked PO2, PCO2 GLU, and LAC for O-rings because I found a spare black O-ring in machine. Cleared out V7 air valve on fluidic board and performed liquid sensor adjustment to clear the oxi liquid sensor error. Performed necessary calibrations. | SWO-307421 | Defect | HW | Autocheck Module | FPC Cable |  | Autocheck module not able to run qc. QC being aborted. | No problem detected when I arrived. Customer reports autocheck module moving, but no sample detected. Typically this problem indicates 14 pole cable having a short. I also found Oxi liquid sensor error. | Replaced 14 pole cable. Checked PO2, PCO2 GLU, and LAC for O-rings because I found a spare black O-ring in machine. Cleared out V7 air valve on fluidic board and performed liquid sensor adjustment to clear the oxi liquid sensor error. Performed necessary calibrations.
No | Visit | Complaint | RIND | Thb/so2 ? error | Thb $so2 ? for all patient results | observed the patient result datalogs after that protein removal ran for 2 to 3 times . then cleaned the gaskets , pumpcalibration was done , liquid sensor calibarton was done, then oxymetry calibartion done then calibarted the analyser for raedy mode , observed for few samples all were ok thb $so2 results were fine | SWO-306704 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | Thb/so2 ? error | Thb /so2 ? for all patient results | observed the patient result datalogs after that protein removal ran for 2 to 3 times . then cleaned the gaskets , pumpcalibration was done , liquid sensor calibarton was done, then oxymetry calibartion done then calibarted the analyser for raedy mode , observed for few samples all were ok thb $so2 results were fine
No | Visit | Complaint | RAME | Neon lamp errors | Confirmed issue when onsite. Oxi hardware error E580 on oxi panel with 2-pt cal. Performed complete oxy test and neon voltage value was 372V (acceptable range 275-325V). | Proceed to replace lamp assembly and adjusted lamp status to 3260 units. New lamp assembly’s neon voltage is 299, right on mean. Ran complete oxy test twice, both passed. Performed cals and QCs without E580 or other errors. | SWO-306352 | Defect | HW | Oxi module | Lamp unit | Neon to high | Neon lamp errors | Confirmed issue when onsite. Oxi hardware error E580 on oxi panel with 2-pt cal. Performed complete oxy test and neon voltage value was 372V (acceptable range 275-325V). | Proceed to replace lamp assembly and adjusted lamp status to 3260 units. New lamp assembly’s neon voltage is 299, right on mean. Ran complete oxy test twice, both passed. Performed cals and QCs without E580 or other errors.
No | Visit | Complaint | RDEU | 09.02.18 12:15uhr Herr Mehl 05931/1521462 Einlass tropft | Einlaß tropft | Gerät überprüft, Schlauchanschlussnippel, Abfallschläuche, Abfallpumpenschläuche, Abfallbehälter und Y-Stück erneuert, Schlauchsystem gespült, Kal und QC o.k. | SWO-308353 | Defect | HW | Basic Unit | Y-Piece |  | 09.02.18 12:15pm Mr Mehl 05931/1521462 Entrance drips | Admission drips | Device inspected, hose connection nipple, waste hoses, waste pump hoses, waste bin and Y-piece renewed, hose system flushed, Kal and QC ok.
No | Visit | Complaint | RAME | Analyzer has gone one week without running QC automatically, nor will it allow user to manually run. | Analyzer will not allow manual or automatic QC initialization. Breaker appears to be blocking normal functionality of carousel. | Replaced stepping motor for ampule opener and FPC cable 14-pole. | SWO-306802 | Defect | HW | Autocheck Module | Steppermotor | Breaker arm | Analyzer has gone one week without running QC automatically, nor will it allow user to manually run. | Analyzer will not allow manual or automatic QC initialization. Breaker appears to be blocking normal functionality of carousel. | Replaced stepping motor for ampule opener and FPC cable 14-pole.
No | Visit | Complaint | RDEU | 06.02.2018 Pfl. Ismael 09671930672 meldet, dass der AQT täglich eine Systemreinnigung aufruft. 08.02.2018 9:15 H. Zinkel 09671930244: weiterhin fordert das Gerät zu Systemreinigungen auf | Wet Section $Serviceprogrammebene Error-1279 gemessene Flüssigkeitsmenge oberhalb Warngrenze | Fehlersuche, weiterhin wird täglich mehrfach zur Systemreinigung aufgefordert, rote bzw. gelbe Ampel->d.h. LS1 Wert >1700 bzw. 1900, Pumpe P2 und P4 neu, komplette Verschlauchung zwischen Cup wash Unit und waste neu, Neustart, Systemreinigung, Filter white top neu, restlichen Schläuche neu von JW Kit, Proteinentfernung P1 $P2, LQCs, Überprüfung Flow Value jetzt bei ca.1600 , vorher 1900-2300, Testmessungen ok Remote Support aktiviert System-und Funktionskontrolle i.O. | SWO-306757 | Cleaned | HW | PM kit | PM Kit |  | 06.02.2018 Pfl. Ismael 09671930672 reports that the AQT calls a system clean-up every day. 08.02.2018 9:15 H. Zinkel 09671930244: the device continues to request system cleaning | Wet Section /Serviceprogrammebene Error-1279 measured liquid quantity above warning limit | Troubleshooting, further several times a day for system cleaning, red or yellow traffic light->i.e. LS1 value >1700 or 1900, pump P2 and P4 new, complete hose between cup wash unit and waste new, restart, system cleaning, filter white top new, remaining Hoses new from JW Kit, protein removal P1 $P2, LQCs, review flow value now at ca.1600 , previously 1900-2300, test measurements ok Remote Support Enabled System and function control in the
No | Visit | Complaint | RDEU | Barcodes werden nur sporadisch im Inlet erkannt 05.02.2018; Fr. Reither; Tel.: 077619386740; Röhrchen für QC und Systemreinigung werden nicht mehr erkannt, AQT verlangt schon wieder eine Systemreinigung. | Erfahren, dass AQT seit Aufstellung nicht richtig gelaufen ist mit unterschiedlichen Problemen. | Neue Sampler Vision Mod.Kit eingebaut. Startphase dann Test ohne Erfolg SW neu installiert....Startphase dann Test mit Saarstedt...Q C ....Cleaning Tube ohne Erfolg Gerät sollte ausgetauscht werden zur Fehlersuche | SWO-306620 | Defect | HW | Inlet module | Barcode reader, Inlet |  | Barcodes are only detected sporadically in the inlet 05.02.2018; Ms. Reither; Tel.: 077619386740; Tubes for QC and system cleaning are no longer recognized, AQT requires system cleaning again. | Learn that AQT hasn't gone well since installation with different problems. | New Sampler Vision Mod.Kit built-in. Start phase then test without success SW reinstalled.... Start phase then test with Saarstedt... Q C .... Cleaning Tube without success Device should be replaced for troubleshooting
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | LAC灵敏度超范围 0378 | 更换测试卡后修复 | SWO-313643 | [5 empty] | Test card, reagent package failure | LAC Sensitivity Over range 0378 | Repair after replacing test card
No | Visit | Complaint | RCHN | 报错3412 | 报错3412，903-112,902-977,924-591各一个 |  | SWO-306939 | Defect | HW | Wet section control | PCB, Wet section Control | 3412 | Error 3412 | Error 3412,903-112,902-977,924-591 each
No | Visit | Complaint | RCHN | 报错1031 | 更换光单后，定标不过，需要借944-220（12691），944-258（13338）各一张 944-216（13434），944-394（13418）缺货 |  | SWO-306928 | Adjustment | Consumables | Test Kit | Cartridge, Re calibration | 1031 | 1031 | After replacing the light sheet, the tender is not available, but it is required to borrow 944-220 (12691), 944-258 (13338) each 944-216 (13434), 944-394 (13418) out of stock
No | Visit | Complaint | RCHN | 报错1999.35.16 | 报错1999.35.16，借903-036一个 |  | SWO-306962 | Assisted Replacement | Consumables | Solution Pack | Solution Pack |  | False 1999.35.16 | False 1999.35.16, loan 903-036 one
Yes | Visit | Complaint | RMAY | Touch screen not responding. | Touch screen not responding due to CPU failure. | Replace with other CPU unit (DMS type CPU) | SWO-309934 | Defect | HW | DMS_PC | CPU |  | Touch screen not responding. | Touch screen not responding due to CPU failure. | Replace with other CPU unit (DMS type CPU)
No | Visit | Complaint | RLTD | Michelle 01604 545611 | Unable to pump solution error | Service programs identified solution pack errors. Calibrated pumps and liquid sensors, now passing service tests. | SWO-307148 | Assisted Replacement | SP | SP | FT |  | Michelle 01604 545611 | Unable to pump solution error | Service programs identified solution pack errors. Calibrated pumps and liquid sensors, now passing service tests.
No | Visit | Complaint | RLTD | Michelle Miller 01604 545611 12/2 Hannah 01604 545867 or 01604 525867 | Failing all fluid transport tests. | Round circle of dust on sensor cassette position window. Cleaned and carried out manual service retests. Now passing. | SWO-307147 | Cleaned | HW | Basic Unit | SC Position window |  | Michelle Miller 01604 545611 12/2 Hannah 01604 545867 or 01604 525867 | Failing all fluid transport tests. | Round circle of dust on sensor cassette position window. Cleaned and carried out manual service retests. Now passing.
No | Visit | Complaint | RLTD | Please call Stuart Kay on 01604 545 006 12/2 No memory please contact Mathew Gilmartin Tel 01604 545006 | Radiometer software failing to load due to a missing $corrupted file. | Saved customer database and settings, reloaded software version 3.3MR1. | SWO-307234 | Defect | SW | RMED SW | RMED SW | Screen blank | Please call Stuart Kay on 01604 545 006 12/2 No memory please contact Mathew Gilmartin Tel 01604 545006 | Radiometer software failing to load due to a missing /corrupted file. | Saved customer database and settings, reloaded software version 3.3MR1.
No | Visit | Complaint | RLTD | Daisy Tel 07916 732973 | Looked at analyser data logs, activity log shows quite a few aborted samples between 09 $02 $18 to 11 $02 $18. Errors 1235 Failed to aspirated sample & 1189 FTC aborted, LS change state error. Checked data logs - no calibration or quality control errors present for the period in question. Carried out manual flush of the sample pathway. Tested fluid transport in service mode - all values well inside allowable limits. Customer also questioning a difference in result for a patient sample run at 08:04 12 $02 $18 (sample no. 574) Ran test blood samples on this analyser and ABL835 in the same dept. - compared results, all ok. | Performed manual flush of the sample pathway. Performed comparative test blood samples. Analyser measuring correctly and ok for use. | SWO-307210 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | Daisy Tel 07916 732973 | Looked at analyser data logs, activity log shows quite a few aborted samples between 09 /02 /18 to 11 /02 /18. Errors 1235 Failed to aspirated sample & 1189 FTC aborted, LS change state error. Checked data logs - no calibration or quality control errors present for the period in question. Carried out manual flush of the sample pathway. Tested fluid transport in service mode - all values well inside allowable limits. Customer also questioning a difference in result for a patient sample run at 08:04 12 /02 /18 (sample no. 574) Ran test blood samples on this analyser and ABL835 in the same dept. - compared results, all ok. | Performed manual flush of the sample pathway. Performed comparative test blood samples. Analyser measuring correctly and ok for use.
No | Visit | Complaint | RCHN | 测试卡卡盘 | 测试卡卡盘 | 取出断裂的测试卡后修复(By Ran) Cartridge wing was broken. Changed a new one. OK. | SWO-307103 | Defect | Consumables | Test Kit | Cartridge, Broken |  | Testing the kaka plate | Testing the kaka plate | Repair after removing broken test card (By Ran) The Wing Wing Wing was broken. New one. Ok.
No | Visit | Complaint | RIND | Calibration error. | Check the system. check the logs. | Check the calibration logs. Check the qc logs. Need to check with new SC9 & SP9. | SWO-309871 | [5 empty] | Calibration error. | Check the system. check the logs. | Check the calibration logs. Check the qc logs. Need to check with new SC9 & SP9.
No | Visit | Complaint | RLTD | Please call Hannah on 07887 428 417 | IP address conflicts | Entered static IP address details, however this then disconnected the analyser. Customer to chase ITS to get details. | SWO-307235 | Setup | His/Lis | His/Lis | His/Lis |  | Please call Hannah on 07887 428 417 | IP address conflicts | Entered static IP address details, however this then disconnected the analyser. Customer to chase ITS to get details.
Yes | Visit | Complaint | RMAY | pH calibration sensitivity out of range | Error occur due to faulty pH electrode. | Replace with other pH electrode. | SWO-309930 | Re classified RMED | OOB | Electrode | pH |  | pH calibration sensitivity out of range | Error occur due to faulty pH electrode. | Replace with other pH electrode.
Yes | Visit | Complaint | RMAY | Halogen lamp intensity too low | Halogen lamp intensity too low. | Re-align halogen lamp accordingly. | SWO-309931 | Re classified RMED | OOB |  |  |  | Halogen lamp intensity too low | Halogen lamp intensity too low. | Re-align halogen lamp accordingly.
No | Visit | Complaint | RDEU | AutoQC erkennt keine Ampullen | AQC Ampullen werden nicht erkannt . | Servicetests durchgeführt. Karussell dreht sich nicht bei Scan-Abtastung. Scanner reagiert auch nicht. Flachbandkabel FTPC 14 Pol. und Scanner erneuert. -Kein Erfolg. AQC Ersatzmodul aus GS-Demogerät eingebaut. AQC-Modul Justage durchgeführt. Alle Funktionen wieder o.k. AQC PCB muß bestellt und Kundenmodul repariert werden. | SWO-307711 | Defect | HW | Autocheck Module | Autocheck Module |  | AutoQC does not detect ampoules | AQC ampoules are not detected. | performed service tests. Carousel does not rotate during scan scanning. Scanner also does not respond. Ribbon cable FTPC 14 pol. and scanner. -No success. AQC replacement module from GS demo device. AQC module adjustment. All functions again ok. AQC PCB must be ordered and customer module repaired.
No | Visit | Complaint | RAME | The analyzer sometimes shows a windows error message about a problem with Netop. | Analyzer intermittently disconnects from the wireless network. The dongle is also shutting 'Off' on its own. | Checked all the Asus wireless adapter settings with our Senior SIT Mike Bianconi. Replaced a new dongle. Walked the analyzer to each one of the operating rooms with Susana and, checked the wireless signal strength in each area. We noticed that the signal strength was low when we entered the operating room number 5 but, it did not disconnect from the wireless network 'MobileCart'. Susana will be monitoring the connection for this analyzer and, she will inform me if she finds or hears of any reports of the analyzer loosing connection again. | SWO-307516 | Setup | His/Lis | His/Lis | His/Lis | Wireless | The analyzer sometimes shows a windows error message about a problem with Netop. | Analyzer intermittently disconnects from the wireless network. The dongle is also shutting 'Off' on its own. | Checked all the Asus wireless adapter settings with our Senior SIT Mike Bianconi. Replaced a new dongle. Walked the analyzer to each one of the operating rooms with Susana and, checked the wireless signal strength in each area. We noticed that the signal strength was low when we entered the operating room number 5 but, it did not disconnect from the wireless network 'MobileCart'. Susana will be monitoring the connection for this analyzer and, she will inform me if she finds or hears of any reports of the analyzer loosing connection again.
No | Visit | Complaint | RKOR | Sample clot error, pO2 QC failed. | . | Service mode check. Gasket module clean. Gasket hold clean. Inlet probe clean. Liquid flow line clean. Calibration check. QC => Failed. SP replace. Cal & QC OK. | SWO-310103 | Assisted Replacement | SP | SP | FT | Clot | Sample clot error, pO2 QC failed. | . | Service mode check. Gasket module clean. Gasket hold clean. Inlet probe clean. Liquid flow line clean. Calibration check. QC => Failed. SP replace. Cal & QC OK.
No | Visit | Complaint | RDEU | 10.01.17, Frau Hoyer 015127648357, Probleme mit QC | Sporadischer Fehler1002:pO2 Dunkeldaten außer Bereich | Sensor Interface PCB erneuert. Beim Ausbau der Platine wurde im Gerät ein Kondensator gefunden, der sich von dem PCB gelöst hatte. PCB erneuert, Gerät überprüft. | SWO-307917 | Defect | HW | Basic Unit | PCB, Sensor interface |  | 10.01.17, Frau Hoyer 015127648357, mit QC issues | Sporadic Error1002:pO2 Dark Data Out Of Range | Sensor Interface PCB. When the board was removed, a capacitor was found in the device, which had detached from the PCB. PCB renewed, device checked.
No | Visit | Complaint | RDEU | Leckage Fehler Lösungspack und Meßkassette mehrfach erneuert | Fehlersuche Service-Testprogramme System gespült Nach Tausch des Ventilschlauches und Behebung des Leckage-Fehlers, trat Glu $Lak.- Status und Empf.-Fehler auf | Ventilschlauch ersetzt SC90 LOT R0199N005 ersetzt SC90 LOT R0122N005 ersetzt SP90 LOT BC-53 ersetzt Satz Wartungsteile eingebaut (siehe SWO Jahreswartung) Probelauf | SWO-308514 | Defect | HW | Oxi module | Pinch valve tube |  | Leakage error Solution pack and measuring cassette renewed several times | Troubleshooting Service test programs System flushed After replacing the valve hose and fixing the leakage error, Glu /Lak status and mpf.-error occurred | Valve hose replaced SC90 LOT R0199N005 replaced SC90 LOT R0122N005 replaced SP90 LOT BC-53 replaced Set of maintenance parts installed (see SWO annual maintenance) Rehearsal
No | Visit | Complaint | RSCH | Blank | Erreur de réponse lactate code: 1310 Erreur joint d’introduction code : 1401 (sporadique) | Recherche de panne sur le module d’introduction Remplacé le « Sensor Interface PCB » e.o Remplacé le kit de maintenance annuel (joint d’introduction) e.o Divers test de fonctionnement dans le programme de service e.o Démarrage système à froid e.o Calibration e.o contrôle de qualité e.o | SWO-308185 | Defect | HW | Basic Unit | PCB, Sensor interface |  | Blank | Lactate response error code: 1310 Joint introductory code error: 1401 (sporadic) | Breakdown search on the introductory module Replaced the Sensor Interface PCB e.o Replaced the annual maintenance kit (introductory joint) e.o Various operational tests in the e.o service program Cold system start e.o Calibration e.o quality control e.o
No | Visit | Complaint | RIND | Service Visit. | All the parameter of the sensor cassette deactivated except pO2. | Keep the sensor cassette in to refrigerator in the environment around 5 degree Celsius for 45 minutes. Keep the operating temperature of analyzer around 22 degree Celsius. Take the cassette from the refrigerator and install it into the analyzer, successfully installed. Analyzer handed over to staff in well working condition. | SWO-309254 | [5 empty] | Service Visit. | All the parameter of the sensor cassette deactivated except pO2. | Keep the sensor cassette in to refrigerator in the environment around 5 degree Celsius for 45 minutes. Keep the operating temperature of analyzer around 22 degree Celsius. Take the cassette from the refrigerator and install it into the analyzer, successfully installed. Analyzer handed over to staff in well working condition.
No | Visit | Complaint | RCHN | 报错1999 | 报错1999，借944-216（13784）一个 | 更换944-216（13784）后修复 Recal, then OK.(By Ran) | SWO-307116 | Assisted Replacement | Consumables | Test Kit | Cartridge, CKMB | 1999 | Error 1999 | Error 1999, loan 944-216 (13784) one | Fix Recal, then OK after replacing 944-216 (13784). (By Ran)
No | Visit | Complaint | RAME | Not auto running QC intermittently. | Verified customer's stated complaint. Customer is still learning the analyzer. Customer had 2 sets of lot numbers in analyzer to run automatically but only one set matched the ampoules in the carousel. The unit sat idle because it could not find the ampoules with the given lot numbers. | Showed customer and staff how to change the analyzer settings for the ampoules that are no longer in the carousel from running automatically to manual, so that the analyzer doesn't look for them when running QC's until the customer has downloaded the data and could erase the lot numbers from the analyzer. Also, ran the customers tHb calibration that was due as a courtesy. | SWO-308187 | User error | Setup AC | Setup AC | Setup AC |  | Not auto running QC intermittently. | Verified customer's stated complaint. Customer is still learning the analyzer. Customer had 2 sets of lot numbers in analyzer to run automatically but only one set matched the ampoules in the carousel. The unit sat idle because it could not find the ampoules with the given lot numbers. | Showed customer and staff how to change the analyzer settings for the ampoules that are no longer in the carousel from running automatically to manual, so that the analyzer doesn't look for them when running QC's until the customer has downloaded the data and could erase the lot numbers from the analyzer. Also, ran the customers tHb calibration that was due as a courtesy.
No | Visit | Complaint | RAME | Customer reporting leak detected error. | Confirmed reported issue. Performed visual inspection of electrodes, measuring chamber, fluidics. Reviewed data logs. CREA QC failing on level 2 only, with error or 954 below range for CREA A. | Replaced pH electrode lower sealing ring, found it stuck to bottom of chamber with dried solution around it. Replaced bent sample probe. Most likely bent while customer performed weekly maintenance related to removal of inlet gasket. Inspected CREA electrodes, suspect bubbles in membrane, removed electrode from membrane, tapped to remove bubbles, reinstalled. | SWO-309696 | Defect | Electrode | Electrode | Sealing ring, pH | pH | Customer reporting leak detected error. | Confirmed reported issue. Performed visual inspection of electrodes, measuring chamber, fluidics. Reviewed data logs. CREA QC failing on level 2 only, with error or 954 below range for CREA A. | Replaced pH electrode lower sealing ring, found it stuck to bottom of chamber with dried solution around it. Replaced bent sample probe. Most likely bent while customer performed weekly maintenance related to removal of inlet gasket. Inspected CREA electrodes, suspect bubbles in membrane, removed electrode from membrane, tapped to remove bubbles, reinstalled.
No | Visit | Complaint | RDEU | Nach Aus-Einschalten erkennt der ABL den Barcode nicht mehr. | Patientenbarcode wird nicht erkannt, andere Barcodes wohl. | Barcode-Einstellungen überprüft. Unter den Barcode-Settings hatte sich ein Zeichen(;) geschlichen. Daher funktionierte der Barcode nicht. ; entfernt und Gerät überprüft | SWO-308737 | Defect | SW | RMED SW | Roma.ini | Barcode setup lost | After switching off, the ABL no longer recognizes the barcode. | Patient barcode is not recognized, other barcodes are probably. | barcode settings checked. Among the barcode settings, there was a sign (;) sneaked in. Therefore, the barcode did not work. ; removed and device checked
Yes | 7 | Y | RPAC | Blank | Replacement of faulty Solution Pack for ABL90 Flex Plus. Errors with Quality Control A. |  | D3E3DE413A567C62 | [5 empty] | Blank | Replacement of faulty Solution Pack for ABL90 Flex Plus. Errors with Quality Control A.
No | 3 | Y | RPAC | Blank | hemolyer QCs |  | D3E3A6843A583D40 | Defect | HW | Oxi module | Hemolyzer | QC | Blank | hemolyer QCs
No | Visit | Complaint | RIND | Calibration error. | Check the logs. check the calibration logs. | Installed new set of consumable & run all 3 level qc. qc are in ranges. calibration are in ranges. machine under observation. | SWO-309882 | [5 empty] | Calibration error. | Check the logs. check the calibration logs. | Installed new set of consumable & run all 3 level qc. qc are in ranges. calibration are in ranges. machine under observation.
No | Visit | Complaint | RESP | Error en cal de bombas y en pO2. | Comprobación de registros de datos del equipo. | Se sustituyen todos los tubos de bombas después de comprobar todos los rotores de bombas, se comprueban las cámaras de medida y sus electrodos detectando la falta del O'ring inferior al electrodo de pO2, se realizan ajustes de SL y Calderón de bombas, test completo del módulo Oxi, comprobamos que el bote de limpieza no tiene activo la hexoquinasa , se añade dicho aditivo a la solución de limpieza para su adecuado tratamiento, se calibra el equipo a 2 puntos y verificamos el funcionamiento final con QC. | SWO-308516 | Defect | Electrode | Electrode | Sealing ring, pO2 | pO2 | Pump lime and pO2 error. | Checking computer data records. | All pump tubes are replaced after all pump rotors are checked, the measuring chambers are checked and their electrodes are checked by detecting the lack of the O'ring lower than the pO2 electrode, SL and Pump cauldron adjustments are made, complete test of the module Oxi, we check that the cleaning can does not have active hexokinase, this additive is added to the cleaning solution for proper treatment, the equipment is calibrated at 2 points and we verify the final operation with QC.
No | Visit | Complaint | RESP | Neón Lámpara Intensity todo los | Comprobación de registro de actividad del equipo | Sustitución de la PCB pH $BG module, Oxí module PCB, cal de la tHb, calibraciones a dos puntos, relleno del carrusel del QC, verificación final con QC. | SWO-308638 | Adjustment | HW | Oxi module | Lamp unit | Neon | Neon Lamp Intensity All | Checking the computer's activity log | Replacing the PCB pH $BG module, Oxí module PCB, tHb lime, two-point calibrations, QC carousel fill, final check with QC.
No | Visit | Complaint | RAME | pCO2 Issues | Pco2 high on level 1 qc 71.9 range 70.6 Good on other 2 Calibrations 92.1-85.7% sens 30.7 status Ph status 7.473 high Pco2 electrode looks worn looks like a 4 year old electrode Pco2 was shipped 02 $2017 Possible bad electrode | Cleaned ph electrode Remembrane ref Replaced pc02 with loaner and re membrane Ph at 7.180 Pco2 status 41.9 sens 96.5 | SWO-309433 | Assisted Replacement | Electrode | Electrode | pCO2 |  | pCO2 Issues | Pco2 high on level 1 qc 71.9 range 70.6 Good on other 2 Calibrations 92.1-85.7% sens 30.7 status Ph status 7.473 high Pco2 electrode looks worn looks like a 4 year old electrode Pco2 was shipped 02 /2017 Possible bad electrode | Cleaned ph electrode Remembrane ref Replaced pc02 with loaner and re membrane Ph at 7.180 Pco2 status 41.9 sens 96.5
No | Visit | Complaint | RIND | Oximetry values is not reflecting in patient results. | Check the machine run oximetry test and got failed, Neon Lamp is not working properly. | Change the Neon lamp and run complete oximetry test, now oximetry test is passed. Also run 2 point calibration now machine is working fine. | SWO-308348 | Defect | HW | Oxi module | Lamp unit | Neon | Oximetry values is not reflecting in patient results. | Check the machine run oximetry test and got failed, Neon Lamp is not working properly. | Change the Neon lamp and run complete oximetry test, now oximetry test is passed. Also run 2 point calibration now machine is working fine.
No | Visit | Complaint | RAME | Oxi & El/ LS error. | Inspection of the system revealed a clot in the inlet gasket and vent port at V7 in the fluidic channel. | Cleared clot, cleaned vent port at V7 on the fluidic channel. | SWO-308224 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway | Clot | Oxi & El/LS error. | Inspection of the system revealed a clot in the inlet gasket and vent port at V7 in the fluidic channel. | Cleared clot, cleaned vent port at V7 on the fluidic channel.
No | Visit | Complaint | RAME | ElMet LS error 0532 | Inspection of the system revealed the El $Met electrode module pump tube is loosely connected. | Nothing wrong with pump tube, Replaced electrode module pump tube from customer's stock to get tightest connection. | SWO-308858 | Adjustment | HW | El/Met Module | Pump tubes |  | ElMet LS error 0532 | Inspection of the system revealed the El /Met electrode module pump tube is loosely connected. | Nothing wrong with pump tube, Replaced electrode module pump tube from customer's stock to get tightest connection.
No | Visit | Complaint | RAME | CREA calibration failing due to sensitivity out of range, CREA QC Level 1 failing, low. | Observed CREA calibration sensitivity 2 & 3 at 0 with no change after numerous cycles, level 1 CREA QC below range. | Determine cause of calibration failure was missing cal 2 enzyme tablet not opened and mixed in bottle. Replaced cal 1, 2 & cleaning solution as precautionary measure. Remembrance electrodes, and tapped air bubbles out of membrane to resolve QC issues. Note - customer has run out of a precious lot of QC and the QC status will show a failed level 1. All QC are passing with the current lot numbers and this does not impact the performance of or results obtained. The customer l resolve this when they delete the old plot number. | SWO-309697 | User error | Consumables | Cal2 | Enzym Dozing |  | CREA calibration failing due to sensitivity out of range, CREA QC Level 1 failing, low. | Observed CREA calibration sensitivity 2 & 3 at 0 with no change after numerous cycles, level 1 CREA QC below range. | Determine cause of calibration failure was missing cal 2 enzyme tablet not opened and mixed in bottle. Replaced cal 1, 2 & cleaning solution as precautionary measure. Remembrance electrodes, and tapped air bubbles out of membrane to resolve QC issues. Note - customer has run out of a precious lot of QC and the QC status will show a failed level 1. All QC are passing with the current lot numbers and this does not impact the performance of or results obtained. The customer l resolve this when they delete the old plot number.
No | Visit | Complaint | RLTD | Please call Nigel McKelvy on 061 525 741 | Cup dryer PCB error | Fitted new cup dryer unit on board. Calibrated new cup dryer PCB. Ran new TnI calibration on board and 5 blank tests. All Passed. | SWO-308556 | Defect | HW | Process Module | PCB, Cup dryer |  | Please call Nigel McKelvy on 061 525 741 | Cup dryer PCB error | Fitted new cup dryer unit on board. Calibrated new cup dryer PCB. Ran new TnI calibration on board and 5 blank tests. All Passed.
No | Visit | Complaint | RDEU | 06.02.2018 Pfl. Ismael 09671930672 meldet, dass der AQT täglich eine Systemreinnigung aufruft. 08.02.2018 9:15 H. Zinkel 09671930244: weiterhin fordert das Gerät zu Systemreinigungen auf | AQT zeigt mehrmals am Tag Fehler 1278: gemessene Flüssigkeitsmenge liegt oberhalb der Warngrenze | AQT abgezogen und ein Reparaturüberbrückungsgerät aufgestellt, Überbrückungsgerät nach Kundengerät konfiguriert, IT Anbindung vorgenommen, LQC für jeden Parameter durchgeführt, Probenmessung erfolgreich, System- und Funktionskontrolle in Ordnung | SWO-308866 | Revisit required | Analyzer removed | Analyzer removed | Analyzer removed |  | 06.02.2018 Pfl. Ismael 09671930672 reports that the AQT calls a system clean-up every day. 08.02.2018 9:15 H. Zinkel 09671930244: the device continues to request system cleaning | AQT shows errors several times a day 1278: measured liquid quantity is above the warning limit | AQT withdrawn and a repair bridging device set up, bridging device configured according to customer device, IT connection made, LQC performed for each parameter, sample measurement successful, system and function control in order
Yes | 23 | Y | RPOL | Blank |  |  | D3E4FFFDB059B886 | Re classified RMED | SWU |  |  |  | Blank
Yes | 23 | Y | RPOL | Blank | sprawdzenie aparatu, wymiana wê¿yków, uszczelek. Kalibracja parametrów, pomp, czujników. Aparat sprawny techniczie. |  | D3E4FFA3B054B884 | Re classified RMED | Preventive Maintenance at SWO |  |  |  | Blank | Checking the camera, replacing the parts, seals. Calibration of parameters, pumps, sensors. The camera is an efficient techie.
No | Visit | Complaint | RAME | The Ca is out of range on level 2 QC. | The Ca measures out of range on Level 2 QC. | Analyzer started work overnight and worked well. Performed a ODIN Sensor Simulator test which passed. Performed calibration and ran QC following the simulator test and the analyzer checks okay. | SWO-309255 | No fault found | Fluidic pathway | Fluidic pathway | Ca++ | Clot | The Ca is out of range on level 2 QC. | The Ca measures out of range on Level 2 QC. | Analyzer started work overnight and worked well. Performed a ODIN Sensor Simulator test which passed. Performed calibration and ran QC following the simulator test and the analyzer checks okay.
No | Visit | Complaint | RLTD | Crack in mixing cover | Crack in mixing cover | Replaced mixing cover. | SWO-309198 | Defect | HW | Basic Unit | Mixer Cover |  | Crack in mixing cover | Crack in mixing cover | Replaced mixing cover.
No | Visit | Complaint | RLTD | Dan Rickett 07787598407 | On arrival the analyser was all in green and no errors were present. I checked the fluid transport service programs and all the pressure test values were in the middle of the ranges. Checked through the datalogs and it looks like the problems have been a combination of a clot in the analyser and faulty sensor cassettes. A clot was flushed out of the analyser yesterday by Angela and there were also pressure test flow errors in the activity log which indicates something restricting the solution flow during the pressure test. The pressure test vacuum value and the temperature errors that were being seen look to have gone since a new cassette was installed yesterday at 12:30, since then there hasn't been any fluid transport problems. The problem with the inlet being open looks like it was cause by the inlet gasket not returning to its correct position, I removed the inlet module and checked the movement of the gasket and this all seemed fine, I also removed the probe and checked it wasn't bent and also cleaned it to make sure there was nothing on its outer surface that could restrict the movement of the gasket. The reason for the analyser powering off are not 100% clear but looking through the logs it looks like on 1 occasion it was because mains power was removed, this is indicated by a battery low error and then a message saying the analyser is shutting dow due to a low battery. The other time it looks like someone may have tried to reboot the analyser because it was in intervention required and they didn't press the button again to turn it back on. I checked the battery status in the service programs and this all looked OK. | Monitor the analyser performance to see if the faults have gone. | SWO-309155 | Adjustment | HW | Auto Inlet module | Inlet Position Cal. |  | Dan Rickett 07787598407 | On arrival the analyser was all in green and no errors were present. I checked the fluid transport service programs and all the pressure test values were in the middle of the ranges. Checked through the datalogs and it looks like the problems have been a combination of a clot in the analyser and faulty sensor cassettes. A clot was flushed out of the analyser yesterday by Angela and there were also pressure test flow errors in the activity log which indicates something restricting the solution flow during the pressure test. The pressure test vacuum value and the temperature errors that were being seen look to have gone since a new cassette was installed yesterday at 12:30, since then there hasn't been any fluid transport problems. The problem with the inlet being open looks like it was cause by the inlet gasket not returning to its correct position, I removed the inlet module and checked the movement of the gasket and this all seemed fine, I also removed the probe and checked it wasn't bent and also c leaned it to make sure there was nothing on its outer surface that could restrict the movement of the gasket. The reason for the analyser powering off are not 100% clear but looking through the logs it looks like on 1 occasion it was because mains power was removed, this is indicated by a battery low error and then a message saying the analyser is shutting dow due to a low battery. The other time it looks like someone may have tried to reboot the analyser because it was in intervention required and they didn't press the button again to turn it back on. I checked the battery status in the service programs and this all looked OK. | Monitor the analyser performance to see if the faults have gone.
No | Visit | Complaint | RFRA | Ne fonctionne pas en capillaire de temps en temps Fonctionne bien en mode seringue. | Démontage, vérification système inlet RAS. | Vérification du connecteur conique Inler. Chgt sonde Tests gazo avec l'utilisatrice en mode capillaire ok | SWO-309506 | Revisit required | HW | Auto Inlet module | Cap - Spring |  | Doesn't work in hair from time to time Works well in syringe mode. | Disassembly, RAS inlet system check. | Check the Inler conical connector. Chgt probe Gazo tests with the user in hair mode ok
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | CL定标不过 | 更换测试卡后修复 | SWO-309015 | [5 empty] | Test card, reagent package failure | CL calibration is not | Repair after replacing test card
No | Visit | Complaint | RAME | Lactate electrode connector does not latch properly. | Spring missing from connector. | Replace electrode connector. | SWO-309250 | Defect | HW | El/Met Module | Electrode Connector |  | Lactate electrode connector does not latch properly. | Spring missing from connector. | Replace electrode connector.
Yes | 2 | Y | RPAC | Blank | SW Upgrade to V6.17 |  | D3E4FC97B05D43E3 | Re classified RMED |  |  |  |  | Blank | SW Upgrade to V6.17
No | Visit | Complaint | RAME | Not pulling dem with accession number | When the accession number is put in no dem comes over and they have to manually input it | I captured trace files and will look at them. I had the IT guy see what was on his side and he saw the accession number when scanned | SWO-311010 | Revisit required | His/Lis | His/Lis | His/Lis |  | Not pulling dem with accession number | When the accession number is put in no dem comes over and they have to manually input it | I captured trace files and will look at them. I had the IT guy see what was on his side and he saw the accession number when scanned
Yes | Visit | Complaint | RAME | Correlation Issues | To look at Tbili |  | SWO-309219 | Re classified RMED | Not complaint, see previous visit |  |  |  | Correlation Issues | To look at Tbili
Yes | Visit | Complaint | RAME | Issues with tbili | To check tbili | Talked with the customer and all is well with this analyzer | SWO-309223 | Re classified RMED | Not complaint, see previous visit |  |  |  | Issues with tbili | To check tbili | Talked with the customer and all is well with this analyzer
No | 16 | Y | RPAC | Blank | OXI issues |  | D3E68A9526333D50 | Defect | HW | Oxi module | Oxi module |  | Blank | OXYGEN issues
No | Visit | Complaint | RLTD | Ray Flint observed from CRM that Neonatal ABL90 still not connecting to LIVE Connect | Checked correct lines present in Roma.ini file - all ok. Certificate for LIVE Connect also ok. Analyser is connected to AQURE. Tried pinging LIVE Connect server on 195.41.216.173 - unable to connect. Tried connecting to hie dale.radiometer.dk - unable to connect. Next step verify with IT Network support that port 443 is open for this analyser. | Ask IT to check port & firewall settings. | SWO-309752 | Setup | His/Lis | His/Lis | Live Connect |  | Ray Flint observed from CRM that Neonatal ABL90 still not connecting to LIVE Connect | Checked correct lines present in Roma.ini file - all ok. Certificate for LIVE Connect also ok. Analyser is connected to AQURE. Tried pinging LIVE Connect server on 195.41.216.173 - unable to connect. Tried connecting to hie dale.radiometer.dk - unable to connect. Next step verify with IT Network support that port 443 is open for this analyser. | Ask IT to check port & firewall settings.
No | Visit | Complaint | RLTD | Please call Sue on 0208 836 5667 | Checked data logs and service programs. There have been a lot of aborted samples with failed to aspirate or no sample detected. Changed day light saving settings. | Removed inlet module and covers. Dismantled inlet arm, inspected and tested spring mechanism. All checked out ok. Reseated cables and connections. Tested and customer will monitor. | SWO-310224 | Adjustment | HW | Auto Inlet module | Auto Inlet module |  | Please call Sue on 0208 836 5667 | Checked data logs and service programs. There have been a lot of aborted samples with failed to aspirate or no sample detected. Changed day light saving settings. | Removed inlet module and covers. Dismantled inlet arm, inspected and tested spring mechanism. All checked out ok. Reseated cables and connections. Tested and customer will monitor.
No | Visit | Complaint | RSCH | Blank | Erreur d’aspiration de l’échantillon de sang en mode capillaire Erreur code 1401 : Erreur joint d’introduction | Remplacé le Sensor PCB du module d’introduction + le cable plat e.o Remplacé le module d’introduction automatique e.o Divers test de l’inlet dans le programme de service e.o Calibration e.o contrôle de qualité e.o | SWO-311680 | Revisit required | HW | Auto Inlet module | PCB, Auto inlet |  | Blank | Blood sample suction error in hair mode Code 1401 Error: Joint Introductory Error | Replaced the Sensor PCB of the introductory module - the flat cable e.o Replaced the e.o auto introduction module Various inlet testing in the e.o service program Calibration e.o quality control e.o
No | Visit | Complaint | RFRA | Ne fonctionne pas en capillaire de temps en temps Fonctionne bien en mode seringue. | Démontage, vérification: (nettoyage pinch vanne, système inlet)RAS. | Vérification du connecteur conique Inler. Chgt cassette, joint d'introduction. Tests 10 gazos en mode capillaire ok | SWO-309777 | Defect | HW | Auto Inlet module | Cap - Spring |  | Doesn't work in hair from time to time Works well in syringe mode. | Disassembly, verification: (cleaning pinch valve, inlet system)RAS. | Check the Inler conical connector. Chgt cassette, introductory seal. Tests 10 gasos in hair mode ok
No | Visit | Complaint | RLTD | Sue Hegarty Tel. 01642 282676 | Printer intermittently stops working. The printer has previously been replaced. Suspect CPU | Replaced CPU and carried out a software re-install and re-loaded all data and setup. Printer working. Handed flash card to medical engineering for disposal. Cleaned and inspected inlet gasket holder. advised customer to clean if the analyser asked for replacement. | SWO-309956 | Defect | HW | DMS_PC | CPU |  | Sue Hegarty Tel. 01642 282676 | Printer intermittently stops working. The printer has previously been replaced. Suspect CPU | Replaced CPU and carried out a software re-install and re-loaded all data and setup. Printer working. Handed flash card to medical engineering for disposal. Cleaned and inspected inlet gasket holder. advised customer to clean if the analyser asked for replacement.
Yes | Visit | Complaint | RESP | configuracion de rangos de referencia. | configuracion rangos de referencia | Se configuraron los rangos de referencia de todos los parametros del ABL y el informe a peticion del servicio. | SWO-309755 | Re classified RMED | Setup, customer request |  |  |  | configuration of reference ranges. | configuration reference ranges | The reference ranges of all ABL parameters and the report upon service request were configured.
No | Visit | Complaint | RIND | PO2 Parameter is not active. | PO2 Parameter is not active | Check Calibration Report first to find out error message. Then remove SC and reinstalled. but still PO2 was inactive. then Reset SP and reinstalled. after that Po2 was active. handed over in fine working condition. | SWO-312443 | [5 empty] | PO2 Parameter is not active. | PO2 Parameter is not active | Check Calibration Report first to find out error message. Then remove SC and reinstalled. but still PO2 was inactive. then Reset SP and reinstalled. after that Po2 was active. handed over in fine working condition.
No | Visit | Complaint | RAME | Error 0579 OXI hardware error. | Error 0376 halogen lamp intensity too high. | Adjusted halogen lamp voltage. | SWO-309923 | Adjustment | HW | Oxi module | Lamp unit | Halogen | Error 0579 OXI hardware error. | Error 0376 halogen lamp intensity too high. | Adjusted halogen lamp voltage.
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 汞定标错误 | 更换试剂包后修复 | SWO-311075 | [5 empty] | Test card, reagent package failure | Mercury calibration error | Repair after replacement of reagent kit
No | Visit | Complaint | RDEU | Sr. Angela meldet Leckage und Drucktest fehlgeschlagen am ABL90 | Intervention erforderlich, Leckage erkannt. Einlassdichtunghalter wird empfohlen auszutauschen. Nadel verbogen. | Einlassdichtungshalter und Nadel ausgetauscht. Service Tests keine Leckage mehr. Startphase i.O. Gerät ist Messbereit. Kunden informiert. | SWO-310788 | Defect | HW | Inlet module | Probe bend |  | Sr. Angela reports leak and pressure test failed on ABL90 | Intervention required, leakage detected. Inlet seal holder is recommended to replace. Needle bent. | Inlet seal holder and needle replaced. Service tests no longer leakage. Start-up phase i.O. device is ready for measurement. informed customers.
No | Visit | Complaint | RKOR | Leak detected(1184) 발생하였음. 전화를 통해 flush를 하고 service mode에서 체크하였으나, 해결되지 않음. | 서비스 모드에서 체크시 pressure감지 | External flush를 통해 Clot 제거함 | SWO-310882 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | Leak detected (1184) occurred. Flush over the phone and check in service mode, but it is not resolved. | Pressure detection when checked in service mode | Remove Clot via External Flush
No | Visit | Complaint | RAME | Unit not sending results to rals correctly. | Verified unit not sending results correctly | Changed ini file to show <> symbols when reporting lower or higher results. Unit reporting correctly. | SWO-310473 | Setup | His/Lis | His/Lis | His/Lis |  | Unit not sending results to rals correctly. | Verified unit not sending results correctly | Changed ini file to show <> symbols when reporting lower or higher results. Unit reporting correctly.
No | 2 | Y | RCZE | Blank | pO2 1002:PO2 dark data jsou mimo rozsah, provedena výmìna SC a RP, beze zmìny |  | D3E9DD6788150142 | Revisit required | Revisit required | Revisit required | pO2 |  | Blank | pO2 1002: PO2 Dark Data is out of range, the change is made to the SC and RP, unchanged
No | Visit | Complaint | RAME | Error 2302 Sensor check failed during system cycles. | During system cycles, the analyzer keeps presenting error 2302 Sensor check failed. | Reinstalled the sensor cassette and, reseated the solution pack. Flushed the hemolyzer and, replaced the sensor connector board. In diagnostics performed a complete service test, passing within specifications. Ran the ctHb calibration, passing within specifications. | SWO-310696 | Defect | HW | Basic Unit | PCB, Sensor connector assembly |  | Error 2302 Sensor check failed during system cycles. | During system cycles, the analyzer keeps presenting error 2302 Sensor check failed. | Reinstalled the sensor cassette and, reseated the solution pack. Flushed the hemolyzer and, replaced the sensor connector board. In diagnostics performed a complete service test, passing within specifications. Ran the ctHb calibration, passing within specifications.
Yes | Visit | Complaint | RIND | Calibration error. | Check the logs. check the calibration logs. | check the logs. check the calibration. | SWO-313422 | Re classified RMED | Courtesy |  |  |  | Calibration error. | Check the logs. check the calibration logs. | check the logs. check the calibration.
No | Visit | Complaint | RDEU | 19.02.2018, 08:44 Uhr, Fr. Engel, Tel. 021513210299: Intervention erforderlich; EL/MET Lower Liquid Sensor Error | Demo Key ausgelaufen, LS Fehler lower E-Lyte Modul Fehler | End Key eingespielt Luftkanal gespült, LS Sensoren kalibriert, Kal und QC durchgeführt, keine aktuelle Lot QC vorhanden , | SWO-310737 | Adjustment | HW | El/Met Module | LS Sensor |  | 19.02.2018, 08:44, Ms. Engel, tel. 021513210299: intervention required; EL/MET Lower Liquid Sensor Error | Demo Key Expired, LS Error Lower E-Lyte Module Error | End key recorded air duct flushed, LS sensors calibrated, Kal and QC performed, no current lot QC available ,
Yes | 5 | Y | RCZE | Blank |  |  | D3E9F23588430146 | Re classified RMED | PM |  |  |  | Blank
No | Visit | Complaint | RDEU | 19.02 16:45 Herr Richter Tel. 02035085716 , Nullpunkt bei pO2 bei der Kalibration zu hoch liegt bei 6,8 | Nullpunkt Fehler bei pO2 bei der Kalibration | Kunde empfohlen Elektrode zu reinigen (Bürsten ). Bürste nicht vor Ort. Kunde wischt nun Elektrode mit feuchten Tupfer aus. | SWO-310472 | Cleaned | Electrode | Electrode | pO2 |  | 19.02 16:45 Mr. Richter Tel. 02035085716 , zero point at pO2 at the calibration too high is 6.8 | Zero point error at pO2 during calibration | Customer recommended to clean electrode (brushes). Brush not on site. Customer now wipes out electrode with damp swabs.
No | Visit | Complaint | RLTD | John Meherne Tel: 07823 332 757 | Some parameters failing Cal & QC's. | Checked data logs and reviewed cal & QC results. Cleaned electrode & chamber. Checked reference chamber & electrode. Suspect faulty Calcium electrode. Electrode will be sent for MTO's to fit. Calibration & QC's passed on all parameters except Calcium. New electrode will rectify this problem. Customer will replaced electrode when it arrives. | SWO-310324 | Assisted Replacement | Electrode | Electrode | Ca++ |  | John Meherne Tel: 07823 332 757 | Some parameters failing Cal & QC's. | Checked data logs and reviewed cal & QC results. Cleaned electrode & chamber. Checked reference chamber & electrode. Suspect faulty Calcium electrode. Electrode will be sent for MTO's to fit. Calibration & QC's passed on all parameters except Calcium. New electrode will rectify this problem. Customer will replaced electrode when it arrives.
No | Visit | Complaint | RDEU | Lactat instabil | Lactat instabil Empfindlichkeitswert fällt schnell (gegen 200 pA $mM) Häufige Membranierung notwendig | Elektrode-Spitze mit Proteinentferner gereinigt etc. ohne Fehlerbeseitigung - Erneuerung Lactat Elektrode - Funktionsüberprüfung etc. | SWO-310261 | Assisted Replacement | Electrode | Electrode | Lac |  | Lactate unstable | Lactate unstable Sensitivity value drops quickly (around 200 pA /mM) Frequent membraneing necessary | Electrode tip cleaned with protein remover etc. without troubleshooting - renewal lactate electrode - function check etc.
No | Visit | Complaint | RDEU | QC MCC2 wurden mehrfach gemessen, immer wieder außerhalb Grenzbereich. Einsatz erwünscht. | In MCC2 Lot 06172 wird TnT 0,194 yg $L und CKNB mit 16,2 yg $L. Benutzer definierte Vorgabe Grenzen sind geändert. TnT Grenzen : 0,174-0,187. CKNB 16,7-16,8 | Werte wieder neu eingespannt. QC MCC2 nachgemessen. Alles i.O | SWO-310793 | User error | SW | QC Setup | QC Setup |  | QC MCC2 were measured several times, repeatedly outside the boundary area. required use. | In MCC2 Lot 06172, TnT is /L 0.194 yg /L and CKNB with 16.2 yg. User-defined default boundaries are changed. TnT Limits: 0.174-0.187. CKNB 16.7-16.8 | values re-clamped. QC MCC2. Everything i.O.
Yes | Visit | Complaint | RAME | Error 1184 Leak detected | Customer requested service Tuesday 830am Vacuum error checked fluid path Flushed Re inserted solution pack and sensor Isolated flush Replaced sensor | Replaced sensor | SWO-311390 | [5 empty] | Error 1184 Leak detected | Customer requested service Tuesday 830am Vacuum error checked fluid path Flushed Re inserted solution pack and sensor Isolated flush Replaced sensor | Replaced sensor
No | Visit | Complaint | RLTD | Mark Spector Tel 0203 4566002 or 0793 9069137 | Battery showing zero all the time on screen. | Discharge batter and removed. Reset battery module and battery now shows % 100. Battery charged to 100%. Battery icon shows when power has been turned off and battery level and when mains is plugged in. Led working correctly. | SWO-311090 | Adjustment | HW | Battery module | Battery module | Charge showing 0% | Mark Spector Tel 0203 4566002 or 0793 9069137 | Battery showing zero all the time on screen. | Discharge batter and removed. Reset battery module and battery now shows % 100. Battery charged to 100%. Battery icon shows when power has been turned off and battery level and when mains is plugged in. Led working correctly.
Yes | Visit | Complaint | RDEU | 17.01.2018, 10:11 Uhr, Fr. Hägerich, Tel. 091318539119: Endlosschleife LP wechseln | Kundengerät 092R0138N007 nach Reparatur wieder zurück gebracht | Kundengerät zurück gebracht, 5 Testmessungen durchgeführt --> alle OK, System und Funktion i.O. --> Kunde über die durchgeführte Reparatur informiert --> Kunde meldet sich wenn das genannte Problem oder ein neues Problem auftreten wieder!!! | SWO-311006 | Re classified RMED | Re-installed analyzer |  |  |  | 17.01.2018, 10:11 a.m., Fr. Hägerich, Tel. 091318539119: Change endless loop LP | Customer device 092R0138N007 brought back after repair | Customer device brought back, 5 test measurements performed --> all OK, system and function in the --> Customer informs about the repair carried out --> Customer reports again if the mentioned problem or a new problem occurs!!!
No | Visit | Complaint | RIND | Solution pack not installing. | Solution pack not installing. | Found Solution pack was expired. An-other Solution pack asked from store and installed. Analyser Serviced and calibrated properly. Handed over in a fine working condition. | SWO-312394 | [5 empty] | Solution pack not installing. | Solution pack not installing. | Found Solution pack was expired. An-other Solution pack asked from store and installed. Analyser Serviced and calibrated properly. Handed over in a fine working condition.
No | Visit | Complaint | RCHN | 报错1025 | 报错1025，借944-216（13707）一个 |  | SWO-310758 | Assisted Replacement | Consumables | Test Kit | Cartridge, CKMB | 1025 | 1025 error | 1025, loan 944-216 (13707) one
No | Visit | Complaint | RIND | Service Visit. | Analysis is disable and message shows sp is invalid due to date and time. | Observe the date and time of the analyzer and check the system log. Analyze the time of last system check, feed that time in to the analyzer after successful reset the SP. Now re-install the SP after change the time, successfully done. | SWO-313162 | [5 empty] | Service Visit. | Analysis is disable and message shows sp is invalid due to date and time. | Observe the date and time of the analyzer and check the system log. Analyze the time of last system check, feed that time in to the analyzer after successful reset the SP. Now re-install the SP after change the time, successfully done.
No | Visit | Complaint | RIND | Service Visit. | Sample is leaking from the tube which is connected from inlet section to pH $BG module. | Keep on hold the analyzer, remove the module cover as well as fluidic module cover. Take an hypochlorite solution in beaker and remove the inlet gas kit, inlet probe and tubing, and project in to the beaker. Clean the pH $BG module chamber and EI $Met module chamber. After 20 minutes wash the all projected with the tap water, and connected same as before. Now recorded no replacement in to the analyzer and restart and give the two point manual calibration to the analyzer after initializing, which is successfully done. Analyzer handed over to staff in well working condition. | SWO-313168 | Defect | HW | Basic Unit | Tube, inlet-ph/BG |  | Service Visit. | Sample is leaking from the tube which is connected from inlet section to pH /BG module. | Keep on hold the analyzer, remove the module cover as well as fluidic module cover. Take an hypochlorite solution in beaker and remove the inlet gas kit, inlet probe and tubing, and project in to the beaker. Clean the pH $BG module chamber and EI $Met module chamber. After 20 minutes wash the all projected with the tap water, and connected same as before. Now recorded no replacement in to the analyzer and restart and give the two point manual calibration to the analyzer after initializing, which is successfully done. Analyzer handed over to staff in well working condition.
No | Visit | Complaint | RAME | Spectrometer converter failed. | Examined system status. Found 4 oximetry errors: spectrometer converter failed, neon lamp intensity too low, oximetry hard are error. | Rebooted unit. Performed complete oxi test. Performed service tHb calibration. Verified operation of ABL825. | SWO-310674 | Defect | SW | Oxi module | off/on |  | Spectrometer converter failed. | Examined system status. Found 4 oximetry errors: spectrometer converter failed, neon lamp intensity too low, oximetry hard are error. | Rebooted unit. Performed complete oxi test. Performed service tHb calibration. Verified operation of ABL825.
No | Visit | Complaint | RIND | CO2 calibration error. | CO2 calibration failing repeatedly. | Found air bubble at botum of membrane. Removed and reinstalled. reference chamber found fill with clote. Cleaned properly. Analyser Serviced and calibrated properly. Handed over in a fine working condition. | SWO-312397 | Cleaned | Fluidic pathway | Fluidic pathway | Measuring chamber |  | CO2 calibration error. | CO2 calibration failing repeatedly. | Found air bubble at botum of membrane. Removed and reinstalled. reference chamber found fill with clote. Cleaned properly. Analyser Serviced and calibrated properly. Handed over in a fine working condition.
No | Visit | Complaint | RDEU | Rädchen der Probenhalterung gebrochen, | Probenhalterung gebrochen, | Probenhalterung erneuert, Probemessungen durchgeführt, | SWO-311423 | Defect | HW | Cover and Chassis | Tube holder |  | cogs of the sample holder broken, | Sample holder broken, | Sample holder renewed, sample measurements carried out,
No | Visit | Complaint | RSCH | Blank | Erreur d’aspiration de l’échantillon de sang en mode capillaire Erreur code 1401 : Erreur joint d’introduction | Remplacé le Sensor PCB du module d’introduction + le cable plat e.o Remplacé le module d’introduction automatique e.o Divers test de l’inlet dans le programme de service e.o Calibration e.o contrôle de qualité e.o | SWO-311733 | Revisit required | HW | Auto Inlet module | Auto Inlet module | Cap - Spring | Blank | Blood sample suction error in hair mode Code 1401 Error: Joint Introductory Error | Replaced the Sensor PCB of the introductory module - the flat cable e.o Replaced the e.o auto introduction module Various inlet testing in the e.o service program Calibration e.o quality control e.o
No | Visit | Complaint | RDEU | 19.02.2018 Frau Pflegel 040790202511 meldet, dass sie den ANalysator manuell gespült hat und die Verstopfung bestehen bleibt. | System überprüft => Kein Fehler feststellbar, Sensorkasette defekt weil Gerät im Intervention Modus stand. | Sensorkassette erneuert, Kalibration und QC durchgeführt | SWO-311697 | Assisted Replacement | SC | SC | SC, Clot |  | 19.02.2018 Ms Pflegel 040790202511 reports that she has manually rinsed the ANalysator and that the blockage remains. | System checked => No error detectable, sensor caette defective because device was in intervention mode. | Sensor cartridge renewed, calibration and QC performed
No | Visit | Complaint | RESP | Error al introducir una muestra | Sonda de aspiración doblada | Se repara sonda de aspiración. Se comprueba correcto funcionamiento del módulo de entrada. Equipo operativo. | SWO-312188 | Defect | HW | Auto Inlet module | Probe bend |  | Failed to enter a sample | Bent suction probe | Suction probe is repaired. The input module is checked correctly. Operational team.
No | Visit | Complaint | RSCH | FlexQ modul kann position 1 und 2 spritze nicht erkennen. | FlexQ modul erkennt position 1 und 2 nicht korrekt. | FlexQ modul in service programm neu justiert und kontrolliert I.O. | SWO-311725 | Adjustment | HW | FlexQ Module | Sample Positioning |  | FlexQ module cannot detect position 1 and 2 syringe. | FlexQ module does not correctly detect position 1 and 2. | FlexQ module in service program re-adjusted and controlled I.O.
No | Visit | Complaint | RDEU | Seit ca 2 Wochen verstärkt Kalibrierdriften der Elektrolyte | Kontaktproblem , pH-Dichtring klebte in der Messkammer fest | Fehlersuche, Renigung derKontakt, Erneuerung pH-Dichtring, Funktionskontrolle | SWO-311648 | Defect | Electrode | Electrode | Sealing ring, pH | pH | Calibration drifts of the electrolytes have been increasing for about 2 weeks | Contact problem , pH sealing ring stuck in the measuring chamber | Troubleshooting, checking the contact, renewal of pH sealing ring, function control
Yes | Visit | Complaint | RAME | Customer requested Draw Time to be added to Patient Layout. Functionality not available. | Patient layout edit function does not give user ability to add parameters. | Reverted layouts to default settings. Functionality restored. Rebuilt layouts and added parameter Draw Time requested by customer. | SWO-311861 | Re classified RMED | FAN/TU |  |  |  | Customer requested Draw Time to be added to Patient Layout. Functionality not available. | Patient layout edit function does not give user ability to add parameters. | Reverted layouts to default settings. Functionality restored. Rebuilt layouts and added parameter Draw Time requested by customer.
No | Visit | Complaint | RDEU | 06.02.2018 Pfl. Ismael 09671930672 meldet, dass der AQT täglich eine Systemreinnigung aufruft. 08.02.2018 9:15 H. Zinkel 09671930244: weiterhin fordert das Gerät zu Systemreinigungen auf | Probleme mit Flüssigkeitswarngrenze | Kundengerät wieder platziert, in Betrieb genommen, Startphase startet, abwarten bis erreichen der Inkubationstemperatur, Konfiguration überprüft, Systemreinigung durchgeführt, Systemüberprüfung durchgeführt - alles Ok, Parameter D-Dimer kalibriert, LQC für jeden Parameter durchgeführt, LQC von Parameter CRP zu hoch gemessen, neues CRP LQC Röhrchen aufgetaut und gemessen, CRP weiterhin zu hoch bei LQC, Nadel ausgebaut, gespült und Nadelgehäuse gereinigt - war jedoch keine Verschmutzung erkennbar, Systemreinigung durchgeführt, LQC für CRP erneut gemessen, LQC wieder zu hoch, CRP neu kalibriert, CRP Kal. nicht bestanden (5 Hintergrundausreißer), Systemüberprüfung gestartet - alles Ok, CRP erneut kalibriert, CRP Kal. mit einem Hintergrundausreißer bestanden, LQC für CRP gemessen und wieder zu hoch gemessen, Reparaturüberbrückungsgerät weiter stehen gelassen und Kundengerät wieder abgezogen | SWO-312105 | Revisit required | Analyzer removed | Analyzer removed | Analyzer removed |  | 06.02.2018 Pfl. Ismael 09671930672 reports that the AQT calls a system clean-up every day. 08.02.2018 9:15 H. Zinkel 09671930244: the device continues to request system cleaning | Problems with liquid warning limit | Customer device repositioned, put into operation, start-up phase starts, wait until reach the incubation temperature, configuration checked, system cleaning performed, system check performed - all Ok, parameter D-Dimer calibrated, LQC for each Parameters performed, LQC of parameter CRP too high measured, new CRP LQC tubes thawed and measured, CRP continues to be too high at LQC, needle removed, rinsed and needle housing cleaned - however, no contamination was detectable, system cleaning performed, LQC for CRP again measured, LQC too high again, CRP recalibrated, CRP Kal. failed (5 background outliers), system check started - all Ok, CRP recalibrated, CRP Kal. passed with a background outlier, LQC for CRP measured and again measured too high, repair bridging device left to stand and customer device removed again
No | Visit | Complaint | RDEU | 22.02.2018 Frau Solwieg Klimech 064312921731 meldet, dass der ABL immer wieder Störung anzeigt. | Intervention Druckfehler, Lösungen können nicht gefördert werden | Lösungspack im Sevice entfernt, Lot BC-52; Pumpenschlauch stark verlängert. Austausch des LP. Blockade im System kann nicht manuell ausgespült werden. Installation der Wartungsteile, Schläuche und Komponenten installiert. Neuen LP installiert, Initialisierungsphase 1Stunde. Initialisierungsphase Analysator seit 20.02. nicht messbereit aufgrund des Fehlers, Sensorkassette beschädigt, pCO2, Glu und Lac nicht Kalibrierbar. SC nicht vor Ort wird von Personal getauscht. | SWO-312415 | Assisted Replacement | SP | SP | Pump tubes |  | 22.02.2018 Ms Solwieg Klimech 064312921731 reports that the ABL repeatedly reports fault. | Intervention Printing errors, solutions cannot be promoted | Solution pack removed in Sevice, lot BC-52; pump hose is greatly extended. Replacement of the LP. Blockage in the system cannot be flushed out manually. installation of maintenance parts, hoses and components. New LP installed, initialization phase 1hour. Initialization phase Analyzer since 20.02. not ready to be measured due to the error, sensor cartridge damaged, pCO2, Glu and Lac not calibrated. SC not on site is exchanged by staff.
No | Visit | Complaint | RLTD | Please call Sue on 0208 836 5667 | The customer has reported that the Analyser is still aborting samples intermittently. Checked data logs and confirmed that over 10% of samples have aborted with the message no sample detected during sample aspiration. I have also witnessed a sample being prepared and processed correctly, but had aborted with the same error. | Replaced - Inlet module - 903-338 LOT R0011N003 Inlet detector PCB - 903-337 LOT R0005N002 FPC cable to wet section PCB - 636-556 Tested by running five back to back samples with no errors. | SWO-312377 | Defect | HW | Auto Inlet module | Auto Inlet module |  | Please call Sue on 0208 836 5667 | The customer has reported that the Analyser is still aborting samples intermittently. Checked data logs and confirmed that over 10% of samples have aborted with the message no sample detected during sample aspiration. I have also witnessed a sample being prepared and processed correctly, but had aborted with the same error. | Replaced - Inlet module - 903-338 LOT R0011N003 Inlet detector PCB - 903-337 LOT R0005N002 FPC cable to wet section PCB - 636-556 Tested by running five back to back samples with no errors.
No | Visit | Complaint | RDEU | 21.02.2018, 08:35 Uhr ==> Anrufer Fr. Saretzy, Tel.: 09613032925 ==> Fehler 1305: Heizkreis überprüfen, Analysator schaltet sich automatisch ab 21.02.2018, 09:00 Uhr ==> Anrufer Fr. Saretzy, Tel.: 09613032925 ==> Fehler 1305 besteht weiterhin, Analysator hat sich wieder abgeschalten | SC defekt | Neustart ohne SC, Fehlersuche, neue SC LOT 250836 defekt-> erneuert, Startphase und Intialisierung ok Kal. und QC ok, Aktivitäten gelöscht | SWO-312425 | [5 empty] | 21.02.2018, 08:35 ==> Caller Fr. Saretzy, Tel.: 09613032925 ==> Error 1305: Check heating circuit, analyzer switches off automatically 21.02.2018, 09:00 ==> Caller Fr. Saretzy, Tel.: 09613032925 ==> Error 1305 still exists, analyzer has switched off again | SC defective | Restart without SC, troubleshooting, new SC LOT 250836 defective-> renewed, start phase and intialization ok Kal. and QC ok, activities deleted
No | Visit | Complaint | RDEU | QC RiliBÄK Verletzung pH, pCO2 und pO2 | Geräterücklieferung. | Kundengerät geliefert und in Betrieb genommen. Kalibration und QC’s vorgenommen. Datenübertragung ok. | SWO-312248 | [5 empty] | QC RiliBÄK violation pH, pCO2 and pO2 | Geräterücklieferung. | customer device and put into operation. calibration and QC's. Data transfer ok.
No | Visit | Complaint | RCHN | 报错5784 | 报错5784，借636-550一个 | 更换636-550一个正常。 by Andrew | SWO-310796 | Defect | HW | Process Module | Sensor, Process Unit | 5784 | 5784 | 5784, loan 636-550 one | Replace 636-550 one normal. by Andrew
No | Visit | Complaint | RCHN | 测试卡舱门关闭报错 | 测试卡舱门关闭报错，借903-198一个 |  | SWO-312157 | Defect | HW | Analyzer Control Module | PCB, Main Board |  | Test card door closing error | Test card door closed error, borrow 903-198 one
No | Visit | Complaint | RCHN | 黑屏 | 黑屏，借910-298一个 |  | SWO-312164 | Defect | HW | Analyzer Control Module | Display |  | Black screen | Black screen, borrow 910-298 one
No | Visit | Complaint | RAME | co2 out on Calibration/QC | Contacted by customer directly. On going pCO2 1PT and 2PT calibration failing due to excessive drift 1 & 2. Customer has tried several new pCO2 membranes, 2 different lot#s, issue always return within a few hours to days. | Review data logs, issue appears to have begun approximately 2 $5 $2018. Performed applicable service diagnostics. No observed failures of Gas, measuring or fluid module valves. Inspected measuring chamber, electrode, orings, pump tubes, no issues observed. At this time suspect an intermittent $leaking rinse measuring module valve or electrode issue. Replaced both the rinse and air $gas valve, replaced pH lower sealing ring. Replaced suspect pCO2 electrode. Replaced the ref membrane. Adjusted Oximetry lamp to a nominal 553pA from 515pA. Performed all required adjustments and tests. After completing lamp adjust sO2 for L2 is now recovering low of the customer range but within the package insert and within .1 of the mean. The customer will re-evaluate their ranges as required. No further issues observed at this time. Will monitor. | SWO-312097 | Assisted Replacement | Electrode | Electrode | pCO2 |  | co2 out on Calibration/QC | Contacted by customer directly. On going pCO2 1PT and 2PT calibration failing due to excessive drift 1 & 2. Customer has tried several new pCO2 membranes, 2 different lot#s, issue always return within a few hours to days. | Review data logs, issue appears to have begun approximately 2 $5 $2018. Performed applicable service diagnostics. No observed failures of Gas, measuring or fluid module valves. Inspected measuring chamber, electrode, orings, pump tubes, no issues observed. At this time suspect an intermittent $leaking rinse measuring module valve or electrode issue. Replaced both the rinse and air $gas valve, replaced pH lower sealing ring. Replaced suspect pCO2 electrode. Replaced the ref membrane. Adjusted Oximetry lamp to a nominal 553pA from 515pA. Performed all required adjustments and tests. After completing lamp adjust sO2 for L2 is now recovering low of the customer range but within the package insert and within .1 of the mean. The customer will re-evaluate their ranges as required. No further issues observed at this time. Will monitor.
No | Visit | Complaint | RDEU | 21.02.2018, 09:06 Uhr, H. Adebar, Tel. 06213833528: Kal Fehler bei Lyten | Sporadisch auftretende Driften bei Elektrolyten und Empfindlichkeitswert-Fehler (ca. 80%) in Cl seit Aufstellung | Allgemeine Überprüfung - Reinigung Messkammern und Kontakte - Erneuerung Referenz Elektrode - Funktionsüberprüfung etc. | SWO-312297 | Assisted Replacement | Electrode | Electrode | Reference |  | 21.02.2018, 09:06, H. Adebar, Tel. 06213833528: Kal error at Lyten | Sporadic drifts in electrolytes and sensitivity error (approx. 80%) in Cl since installation | General inspection - Cleaning measuring chambers and contacts - Renewal reference electrode - function check etc.
No | Visit | Complaint | RAME | Correlation Issues | To look at Tbil issue | Performed service dump and retrieved data logs from analyzer in question(006) and from analyzer used to reference against(original 007). Created a CQC case in tech center. Will wait to here what next steps are | SWO-312647 | Revisit required | HW | Oxi module | tBil |  | Correlation Issues | To look at Tbil issue | Performed service dump and retrieved data logs from analyzer in question(006) and from analyzer used to reference against(original 007). Created a CQC case in tech center. Will wait to here what next steps are
No | Visit | Complaint | RLTD | Glynnis Tel 01324 566137 | Pin on Sensor Interface PCB bent and unable to realign | Replaced Sensor Interface PCB with new | SWO-313139 | Defect | HW | Basic Unit | PCB, Sensor interface |  | Glynnis Tel 01324 566137 | Pin on Sensor Interface PCB bent and unable to realign | Replaced Sensor Interface PCB with new
No | Visit | Complaint | RDEU | 23.02.2018, 10:15 Uhr ==> Anrufer Fr. Bernhardt, Tel.: 06213831157 ==> Oxi-Hardwarefehler, Neonkorrektur ausserhalb der Grenzen | 1034 Oxi Hardware Problem 1021 Oxymetrie Neonkorrektion außerhalb Grenzen Keine Aktivitäten-Meldungen | Oximetrie Test durchgeführt (Lamp Status Fehler : Neon Position = -1,00) - Rücktausch des am 15.2.18 erneuertem Hemolysator (R0799N004) - Erneuerung Hemolysator (R0806N001) - Erneuerung Sensor-Kassette - Fehler bleibt weiterhin bestehen (auch bei tHb KAL) - Ersatzteilbestellung erfolgt ! KAL und QC-Fehler in Oximetrie-Parametern ! | SWO-313074 | Revisit required | HW | Oxi module | Oxi module |  | 23.02.2018, 10:15 a.m. ==> Caller Fr. Bernhardt, Tel.: 06213831157 ==> Oxi hardware error, neon correction outside the limits | 1034 Oxi Hardware Problem 1021 Oxymetry Neon Correction Outside Limits No activity messages | Oximetry test performed (Lamp Status Error : Neon Position = -1.00) - Replacement of the Hemolyser renewed on 15.2.18 (R0799N004) - Renewal Hemolysator (R0806N001) - Renewal Sensor Cassette - Error remains (even with t KALHb) - Spare parts order is made ! KAL and QC errors in oximetry parameters !
No | Visit | Complaint | RAME | Will not switch from battery mode to AC | Will not switch from battery mode to AC mode | Replaced power supply. | SWO-313099 | Defect | HW | Power Supply | Power Supply |  | Will not switch from battery mode to AC | Will not switch from battery mode to AC mode | Replaced power supply.
No | Visit | Complaint | RDEU | 23.02.2018, 09:40 Uhr ==> Anrufer Schw. Angelika, Tel.: 0992280459915 ==> Bei Kalibrierung Oximetrie-Hardwarefehler (1010, 1012, 1034) | Kunden am Telefon zu Neustart geraten - danach war ABL wieder messbereit. Vorsorglich ABL Vorort überprüft. | Oxi Complete Test - ok. Ventilschlauch erneuert. Abfallschlauch auf Undichtigkeit überprüft - ok. Pinch Valve im Service überprüft - ok. ABL manuell gespült. Kalibration durchgeführt. tHb Kaibration ausgeführt (101,6%). Kal-Verifizierung ist ok. QC’s vorgenommen. | SWO-313014 | Defect | HW | Oxi module | Pinch valve tube |  | 23.02.2018, 09:40 ==> Caller Schw. Angelika, Tel.: 0992280459915 ==> For calibration oximetry hardware errors (1010, 1012, 1034) | Customers on the phone were advised to restart - after that ABL was ready to be measured again. As a precaution ABL suburb checked. | Oxi Complete Review - ok. Valve hose renewed. Waste hose checked for leakage - ok. Pinch Valve checked in service - ok. ABL manually rinsed. calibration. tHb Kaibration (101.6%). Kal verification is ok. QC's.
No | Visit | Complaint | RCHN | 无法开机 | 无法开机，借910-285一个。 |  | SWO-312758 | Defect | HW | Power Supply | Power Supply |  | Cannot power on | Cannot power on, borrow 910-285 one.
No | Visit | Complaint | RCHN | 光單報錯:Error 1031 Optical intensity out of range 更換光單後執行factory calibration ok, Run TnI cali OK. | Perform system check=>Flash lamp intensity deviation= -8.6 | Replace optic unit Perform TnI calibration pass Run QC ok | SWO-313680 | Defect | HW | Process Module | Optical Unit | 1031 | 光單報錯:Error 1031 Optical intensity out of range 更換光單後執行factory calibration ok, Run TnI cali OK. | Perform system check=>Flash lamp intensity deviation= -8.6 | Replace optic unit Perform TnI calibration pass Run QC ok
No | Visit | Complaint | RCHN | 报错1029 | 报错1029,需要借件903-364一个 |  | SWO-312762 | Defect | HW | Process Module | Optical Unit | 1029 | 1029 error | 1029, need to borrow 903-364 one
No | Visit | Complaint | RKK | 測定時アンステーブル多発 | アンステーブル | pH電極交換 | SWO-312934 | Assisted Replacement | Electrode | Electrode | pH |  | Frequent anstable at the time of measurement | Anse Table | pH electrode replacement
No | Visit | Complaint | RSCH | Rep | pO2 seit längerem immer wieder Nullstromfehler $Empfindlichkeitsfehler Mischerträger Schlupf erkannt Einlass undicht Lactat Empfindlichkeit 177 und instabil Sporadisch Barotemp Fehler (AU Board) NetOP Fehler | pO2 Elektrode ausgetauscht und neumebraniert Lactat neu membraniert FLEX Q Modul Leihweise ausgetauscht(Mischerträger wird in Thalwil ausgetauscht) Einlassdichtung ausgetauscht Funktionskontrolle i.o | SWO-313044 | Assisted Replacement | Electrode | Electrode | pO2 | Multible errors | Rep | pO2 has been repeatedly /Empfindlichkeitsfehler zero current faults for a long time Mixer carrier slip detected Entrance leaks Lactate Sensitivity 177 and unstable Sporadic Barotemp Error (AU Board) NetOP Error | pO2 electrode replaced and re-eminated Lactate re-membraneed FLEX Q Module replaced on loan (mixer carrier is replaced in Thalwil) Inlet seal replaced Functional control i.o
No | Visit | Complaint | RAME | Customer reports that QC has failed with all parameters out of range | Observed QC logs and saw that QC’s were misidentified. Level 1 was run as level 2, level 2 as level 3, etc. | Replaced 14 pole AutoCheck ribbon cable. Removed the carousel and reinstalled it to allow scanning of all QC ampoules. Verified that the physical placement of the ampoules matched what was displayed on the screen. Verified homing of the carousel in service and successfully ran all 3 levels of QC. Analyzer is within acceptable operating specifications. | SWO-312646 | Defect | HW | Autocheck Module | FPC Cable |  | Customer reports that QC has failed with all parameters out of range | Observed QC logs and saw that QC’s were misidentified. Level 1 was run as level 2, level 2 as level 3, etc. | Replaced 14 pole AutoCheck ribbon cable. Removed the carousel and reinstalled it to allow scanning of all QC ampoules. Verified that the physical placement of the ampoules matched what was displayed on the screen. Verified homing of the carousel in service and successfully ran all 3 levels of QC. Analyzer is within acceptable operating specifications.
No | Visit | Complaint | RAME | QC aspiration issues | “0739-solution’s composition not identifiable” errors are present. QC is run but one or more of the parameters are significantly higher or lower than the reportable range. | Verified that the QC solutions are properly scanned and identified. Verified that the QC ampoules in the carousel are correctly scanned and run from the AutoCheck carousel. Observed that the reference membrane had pinkish fluid in the jacket, small cracks near the top and had not been replaced since 12 $2017. Replaced the reference membrane. Analyzer is within acceptable operating specifications. | SWO-312624 | Unknown | HW | Autocheck Module | Autocheck Module | Incorrect barcode reading | QC aspiration issues | “0739-solution’s composition not identifiable” errors are present. QC is run but one or more of the parameters are significantly higher or lower than the reportable range. | Verified that the QC solutions are properly scanned and identified. Verified that the QC ampoules in the carousel are correctly scanned and run from the AutoCheck carousel. Observed that the reference membrane had pinkish fluid in the jacket, small cracks near the top and had not been replaced since 12 $2017. Replaced the reference membrane. Analyzer is within acceptable operating specifications.
No | Visit | Complaint | RDEU | Proben-Abbruch | 1067 Proben-Abbruch 4801 Pipette 425 8611 TestFlowWithoutDilution Position Needle | Analysator in GS nach Absprache Frau Löwenkamp- Funktionsüberprüfung Nadel-Sensor-Aktivitäten (ohne Beanstandung) - mehrere Messungen gestartet (ohne Beanstandung) - Service Dump erstellt | SWO-313112 | Revisit required | Revisit required | Revisit required | Revisit required | 1067 | Sample cancellation | 1067 Sample Cancellation 4801 Pipette 425 8611 TestFlowWithoutDilution Position Needle | Analyzer in GS after consultation Mrs. Löwenkamp- Function check needle sensor activities (without complaint) - multiple measurements started (without complaint) - Service Dump created
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | CO2,CL质控偏低 | 更换测试卡后修复 | SWO-313742 | [5 empty] | Test card, reagent package failure | CO2, LOW CL QUALITY CONTROL | Repair after replacing test card
No | Visit | Complaint | RCHN | 无法读取测试卡芯片信息 | 无法读取测试卡芯片信息，借903-581一个。 | 更换903-581后修复 by Andrew | SWO-312849 | Defect | HW | Basic Unit | PCB, Sensor interface | Chip | Failed to read test card chip information | The test card chip information could not be read, borrowing 903-581 one. | Repair by Andrew after replacement 903-581
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | LAC,GLU,PCO2定标灵敏度超范围 0378 | 更换测试卡后修复 | SWO-313478 | [5 empty] | Test card, reagent package failure | LAC, GLU, PCO2 calibration sensitivity over range 0378 | Repair after replacing test card
No | Visit | Complaint | RCHN | 定标失败 1054 | 加借光单一个 |  | SWO-313431 | Revisit required | Not repaired | Not repaired | Not repaired |  | Bid failure 1054 | Add a single one
No | Visit | Complaint | RCHN | 进样口支架损坏 | 进样口支架损坏，借2个903-584 | 更换2个903-584后修复 | SWO-313417 | HW | Defect | Cover and Chassis | Tube holder |  | Damage to the inlet bracket | Damaged in-in stand, borrow 2 903-584 | 2 903-584 repair
No | Visit | Complaint | RDEU | 23.02.2018, 10:15 Uhr ==> Anrufer Fr. Bernhardt, Tel.: 06213831157 ==> Oxi-Hardwarefehler, Neonkorrektur ausserhalb der Grenzen | Bestellung eingetroffen Neon Position Fehler | Rücktausch Hemolysator (R0806N001) - Erneuerung Spectrometer-Einheit - Initialisierung - Funktionsüberprüfung etc. | SWO-313906 | Defect | HW | Oxi module | Spectrometer |  | 23.02.2018, 10:15 a.m. ==> Caller Fr. Bernhardt, Tel.: 06213831157 ==> Oxi hardware error, neon correction outside the limits | Order arrived Neon Position Error | Replacement Hemolysator (R0806N001) - Renewal Spectrometer Unit - Initialization - Function Check, etc.
Yes | 1 | Y | RPAC | Blank | Software upgrade V3.3MR1 |  | D3F2CE90C2170127 | [5 empty] | Blank | Software upgrade V3.3MR1
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 气泡探测定标失败 | 连续更换一个试剂包和一个测试卡后修复 | SWO-314432 | [5 empty] | Test card, reagent package failure | Bubble probe calibration failed | Repair after continuous replacement of a reagent kit and a test card
No | Visit | Complaint | RHUN | összeméréseknél a sO2 alámér, | 3 szintű QC mérés OK, bár az alsó célértéknél a mérési eremény ~10% alámérést mutat (cél 58, határok 42-74, mért érték: 52) | analóg csatorna utánállítás, újabb QC mérés nem mutat változtást, SC csere után a QC mérések jobb eredményt adnak, összemérés reálisabb eredményt ad | SWO-313903 | [5 empty] | összeméréseknél a sO2 alámér, | 3 level QC measurement OK, although at the lower target the measured value is ~ 10% below (target 58, limits 42-74, measured values: 52) | Analogue channel adjustment, newer QC measurement does not change, after SC replacement, QC measurements give better results, and a more realistic result
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | LAC，GLU定标失败 0378 | 更换测试卡正常。 | SWO-314420 | Assisted Replacement | SC | SC | Glu/Lac | 0378 | Test card, reagent package failure | LAC, GLU calibration failure 0378 | The replacement test card is OK.
No | Visit | Complaint | RCHN | 报错1221 | 报错1221，借903-581一个 | 更换903-581后修复 by jesson | SWO-313756 | Defect | HW | Basic Unit | PCB, Sensor interface | 1221 | 1221 error | 1221, loan 903-581 one | Repair by jesson after replacement 903-581
No | Visit | Complaint | RCHN | 浑浊度过高，光谱不匹配 | 浑浊度过高，光谱不匹配。借903-197 902-980各一个 |  | SWO-313642 | Defect | HW | Oxi module | Hemolyzer |  | Turbidity is high and the spectrum does not match | The turbidity is high and the spectrum does not match. Borrow 903-197 902-980 each
No | Visit | Complaint | RCHN | 报错1999 | 报错1999，借944-220（13435）一张 |  | SWO-313792 | Cleaned | HW | Incubator Module | Bearing, Incubator wheel |  | Error 1999 | Error 1999, loan 944-220 (13435) one
No | Visit | Complaint | RCHN | 系统崩坏 | 系统崩坏，借944-212（13612缺货，13871），944-258（13756,13691），944-394（13848），944-214（13673，13800），944-220（13612），944-216（13784,13707） |  | SWO-313668 | Revisit required | Revisit required | Revisit required | Revisit required |  | System breakdown | System breakdown, loan 944-212 (13612 out of stock, 13871), 944-258 (13756, 13691), 944-394 (13848), 944-214 (13673,13800), 944-220 (13612), 944-216(13784,13707)
Yes | 6 | Y | RCZE | Blank | temp erroro hemolyzer 270, teplotní chyba 512 |  | D3F2B902C21D464E | Re classified RMED | HW | Oxi module | Temperature | To Old | Blank | Temp Erroro hemolyzer 270, temperature error 512
No | Visit | Complaint | RDEU | AutoQC erkennt keine Ampullen | Ampullen werden nicht erkannt | PCB ersetzt, Konfiguration angepasst, keine Veränderung, AutoCheckmodul ersetzt, Konfiguration angepasst, neues Modul macht ratternde Geräusche, Modul ausgebaut und überprüft, nach Reinstallation keine Auffälligkeiten mehr GARANTIEAUSTAUSCH | SWO-314245 | Defect | HW | Autocheck Module | PCBA |  | AutoQC does not detect ampoules | Ampoules are not detected | PCB replaced, configuration adapted, no change, AutoCheck module replaced, configuration adapted, new module makes rattled noises, module removed and checked, after reinstallation no more abnormalities GUARANTEE EXCHANGE
No | Visit | Complaint | RLTD | John Meherne Tel: 07823 332 757 | Ca failing cal & QC. | Replaced electrode chamber. Cleaned electrode contacts. Ran several calibrations & QC's. Calcium still failing on status. Reference electrode due later today. Site to fit electrode & advise of outcome. | SWO-313826 | Revisit required | Electrode | Electrode | Ca++ |  | John Meherne Tel: 07823 332 757 | Ca failing cal & QC. | Replaced electrode chamber. Cleaned electrode contacts. Ran several calibrations & QC's. Calcium still failing on status. Reference electrode due later today. Site to fit electrode & advise of outcome.
No | Visit | Complaint | RKOR | Software Re-installation | Software Re-installation Cal. adjust log, Activity log, Event log check Analyzer status check Consumables capacity check | Software Re-installation | SWO-317074 | Defect | SW | RMED SW | RMED SW |  | Software Re-installation | Software Re-installation Cal. adjust log, Activity log, Event log check Analyzer status check Consumables capacity check | Software Re-installation
No | Visit | Complaint | RAME | Analyzer giving 1184 & 1342 leak detected errors. | Customer had replaced solution pack, sensor cassette, inlet gasket with holder and performed backflush, errors still present. I performed visual checks and found the pinch valve tubing with a leak. | Replaced pinch valve tubing. | SWO-314908 | Defect | HW | Oxi module | Pinch valve tube |  | Analyzer giving 1184 & 1342 leak detected errors. | Customer had replaced solution pack, sensor cassette, inlet gasket with holder and performed backflush, errors still present. I performed visual checks and found the pinch valve tubing with a leak. | Replaced pinch valve tubing.
Yes | 4 | Y | RPAC | Blank | Replaced the sensor interface PCB |  | D3F686BF24117042 | Re classified RMED | Service Tip | Basic Unit | PCB, Sensor interface |  | Blank | Replaced the sensor interface PCB
No | 5 | Y | RPAC | Blank | Glu failed on calibration |  | D3F685E52405703D | Defect | HW | Basic Unit | PCB, Sensor interface |  | Blank | Glu failed on calibration
Yes | 2 | Y | RPAC | Blank | Replaced sensor interface PCB |  | D3F687CA245E7046 | Re classified RMED | Service Tip | Basic Unit | PCB, Sensor interface |  | Blank | Replaced sensor interface PCB
No | Visit | Complaint | RSAF | Stop at customer to visually inspect analyzer and found that the solution pack has no more activities left (1204) which the lab staff never mentioned as they kept telling me there are no messages on the screen. | Telephonic assistance given and not once was I told of the error on screen of 1204 No more activities left, drove to Workshop to collect sensor cassette and solution pack. | Collected Sensor cassette and solution pack from Stores and delivered to customer, | SWO-317345 | User error | User error | User error | User error |  | Stop at customer to visually inspect analyzer and found that the solution pack has no more activities left (1204) which the lab staff never mentioned as they kept telling me there are no messages on the screen. | Telephonic assistance given and not once was I told of the error on screen of 1204 No more activities left, drove to Workshop to collect sensor cassette and solution pack. | Collected Sensor cassette and solution pack from Stores and delivered to customer,
No | 1 | Y | RCZE | Blank | Po spostení ziada prepláchnutie - nie je možné vykona (upchatý analyzátor) Výmena vnútorných hadièiek servisného kitu a preèistenie Kalibrácie QC OK |  | D3F3FA4638190209 | Defect | HW | Basic Unit | PM Kit |  | Blank | After the cleaning is requested-it is not possible to perform a (stuffy analyser) replacing the inside of the service kit and cleaning the calibration QC OK
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | LAC定标失败 | 更换测试卡正常。 | SWO-317180 | [5 empty] | Test card, reagent package failure | LAC calibration failure | The replacement test card is OK.
No | Visit | Complaint | RCHN | 电极故障 | 电极故障，借902-564一个 |  | SWO-314504 | [5 empty] | Electrode failure | Electrode fault, borrow 902-564 one
No | Visit | Complaint | RESP | Error de sensibilidad y fallo de respuesta en Na | Limpieza del sistema con agua, Limpieza inlet gasket y enderezar aguja Sustitucion del casete Cal, qc | Limpieza del sistema con agua, Limpieza inlet gasket y enderezar aguja Sustitucion del casete Cal, qc | SWO-314979 | Assisted Replacement | SC | SC | Na |  | Sensitivity error and response failure in Na | Cleaning the system with water, Cleaning in gasketlet and straightening needle Case replacement Cal, qc | Cleaning the system with water, Cleaning in gasketlet and straightening needle Case replacement Cal, qc
No | Visit | Complaint | RCHN | 不均匀的冲洗溶液 | 不均匀的冲洗溶液，借903-199和903-139各一个 |  | SWO-314677 | Defect | HW | Inlet module | Inlet module |  | Uneven flushing solution | Uneven flushing solution, one for each of 903-199 and 903-139
No | Visit | Complaint | RCHN | 试剂包锁不住 | 试剂包锁不住，借580-400一个。 |  | SWO-314603 | Defect | HW | Basic Unit | Steppermotor, Actuator |  | The reagent bag won't lock. | Reagent pack can not be locked, borrow 580-400 one.
No | Visit | Complaint | RCHN | 漏液 | 漏液 | 管路堵塞，清理管路 | SWO-314881 | Cleaned | HW | Incubator Module | Bearing, Incubator wheel |  | Leak | Leak | Line blocked, clean line
No | Visit | Complaint | RCHN | 报错1029 | 报错1029 | 更换光单后修复 | SWO-314586 | HW | Defect | Process Module | Optical Unit |  | 1029 error | 1029 error | Repair after replacing a light sheet
No | 19 | Y | RPOL | Blank | B³êdy pomiarowe pacjenta pH. Wymiana elektrody pH, wymiana uszczelki w elektrodzie pH. |  | D3F47D79381B44F0 | Assisted Replacement | Electrode | Electrode | pH |  | Blank | Measurement errors of the patient pH. PH electrode replacement, gasket replacement in pH electrode.
No | Visit | Complaint | RCHN | 耗材无法扫描 | 耗材无法扫描，902-822,636-484各一个 | 更换902-822和636-484后修复 by jesson | SWO-313720 | Defect | HW | Basic Unit | Barcode Reader |  | Supplies cannot be scanned | Supplies can not be scanned, 902-822, 636-484 each | Repair by jesson after replacing 902-822 and 636-484
No | Visit | Complaint | RAME | Customer called with sample errors on patient samples only | Check inlet probe and measuring chambers. PCO2 membrane stuck in chamber. Check pH electrode seal ring. Intermittently not aspirating samples correctly. | Replaced pH electrode seal ring. Replaced PCO2 membrane after cleaning all measuring chamber wells. Replaced solution pump tubing. Monitor sample runs. Last samples started to give sample errors. Replaced seal ring on dummy electrode and rinse valve in pH $BG module. Also replaced flat membrane on air valve, Inlet tube and nipple kit to isolate intermittent failure. Ran multiple test samples. | SWO-314225 | Defect | HW | pH/BG Module | Pump tubes |  | Customer called with sample errors on patient samples only | Check inlet probe and measuring chambers. PCO2 membrane stuck in chamber. Check pH electrode seal ring. Intermittently not aspirating samples correctly. | Replaced pH electrode seal ring. Replaced PCO2 membrane after cleaning all measuring chamber wells. Replaced solution pump tubing. Monitor sample runs. Last samples started to give sample errors. Replaced seal ring on dummy electrode and rinse valve in pH $BG module. Also replaced flat membrane on air valve, Inlet tube and nipple kit to isolate intermittent failure. Ran multiple test samples.
Yes | 3 | Y | RPAC | Blank | FAN 915-360 |  | D3F411723835B99B | Re classified RMED | FAN/TU |  |  |  | Blank | FAN 915-360
No | Visit | Complaint | RDEU | Fehler 1279: Gemessene Flüssigkeitsmenge war oberhalb der Warngrenze Gerät verlangt öfter das Reinigungsverfahren | Gerät verlangt häufig Reinigung. Fehler 1279 | P2, Schlauch 3 und vier, sowie Y-Stück erneuert. Servicedump und Daten gezogen. Reinigung durchgeführt. Sollte der Fehler erneut auftreten, wird sich Herr Topel von der Medizintechnik melden. | SWO-315053 | Cleaned | HW | Wet section control | Flow restriction |  | Error 1279: Measured liquid quantity was above the warning limit Device requires the cleaning procedure more often | device often requires cleaning. Error 1279 | P2, tube 3 and four, as well as Y-piece renewed. Service dump and data dragged. Cleaning carried out. If the error recurs, Mr. Topel will report from medical technology.
Yes | 3 | Y | RPAC | Blank | Replaced sensor interface PCB |  | D3F68307240E7038 | Re classified RMED | Service Tip | Basic Unit | PCB, Sensor interface |  | Blank | Replaced sensor interface PCB
No | Visit | Complaint | RCHN | 信号板信号故障 | 信号板信号故障，借902-972一个 |  | SWO-315160 | [5 empty] | Signal plate signal failure | Signal board signal failure, borrow 902-972 one
No | Visit | Complaint | RAME | Intermittent Analog board error message | Duplicate SWO created to document time. SWO-315807 marked complete without LABOR and Travel added. | Duplicate SWO created to document time. SWO-315807 marked complete without LABOR and Travel added. | SWO-315827 | [5 empty] | Intermittent Analog board error message | Duplicate SWO created to document time. SWO-315807 marked complete without LABOR and Travel added. | Duplicate SWO created to document time. SWO-315807 marked complete without LABOR and Travel added.
No | Visit | Complaint | RAME | Intermittent Analog board error message | Intermittent Analog board error message. Message not shown in event logs. | Replaced analog pcb,analog cable, and Spine PCB. | SWO-315807 | [5 empty] | Intermittent Analog board error message | Intermittent Analog board error message. Message not shown in event logs. | Replaced analog pcb,analog cable, and Spine PCB.
No | Visit | Complaint | RCHN | 报浑浊度过高 | 报浑浊度过高，借902-980一个 |  | SWO-315197 | Defect | HW | Oxi module | Hemolyzer | turbidity to high | Reported turbidity over high | Reported turbidity over the high, borrow 902-980 a
No | Visit | Complaint | RESP | El equipo detecta fugas y no se puede usar. | El equipo presentaba fugas y no era posible utilizar el equipo | Se ha detectado que la inlet gasket venia defectuosa y no sellaba correctamente el pack de soluciones. se ha sustituido por una nueva, se han realizado los test de servicio, calibraciones y controles sin incidencias. | SWO-315586 | Defect | HW | Inlet module | Inlet gasket |  | The equipment detects leaks and cannot be used. | The equipment was leaking and it was not possible to use the equipment | The inlet gasket has been found to be defective and did not properly sellate the solution pack. replaced by a new one, service tests, calibrations and incident-free checks have been carried out.
No | Visit | Complaint | RCHN | 电源烧坏 | 电源烧坏，借910-285一个 |  | SWO-315248 | Defect | HW | Power Supply | Power Supply |  | Power burned out | The power supply burned out, borrow edith 910-285 one
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 泵定标错误，1178 | 更换测试卡后修复 | SWO-315917 | Assisted Replacement | SP | SP | Pump tubes |  | Test card, reagent package failure | Pump calibration error, 1178 | Repair after replacing test card
No | Visit | Complaint | RCHN | 系统崩坏 | 系统崩坏，重新定标后报错1999，借944-394（13418）一张。缺货 |  | SWO-315214 | Cleaned | HW | Incubator Module | Bearing, Incubator wheel |  | System breakdown | The system broke down, re-ordered and then misreported 1999, borrowing 944-394 (13418). Stock
Yes | 10 | Y | RPAC | Blank | Intermittent Neon number too low |  | D3F69C2724379661 | Re classified RMED | HW | Oxi module | Lamp unit | To old, Neon | Blank | Intermittent Neon number too low
No | Visit | Complaint | RKK | Clメンブラン交換後2週間経たずに感度低下 | 前回のClメンブラン交換後2週間経たずに感度低下 | 別ロットのメンブランと交換。経過観察。 | SWO-315887 | Assisted Replacement | Electrode | Membrane | Cl-- |  | Sensitivity decreases less than two weeks after Cl-Menbran replacement | Sensitivity decreases less than two weeks after the last Cl-Menbran replacement | Exchange it with a different lot of membranes. Follow-up.
No | Visit | Complaint | RCHN | 无法扫描试剂 | 无法扫描试剂，借902-822一个。 |  | SWO-315219 | Defect | HW | Basic Unit | Barcode Reader |  | Unable to scan reagents | Unable to scan reagents, borrow 902-822 one.
No | Visit | Complaint | RCHN | 触摸屏不灵敏 | 触摸屏不灵敏，借617-944一个。 |  | SWO-315218 | Defect | HW | DMS_PC | Touch screen | Check FW | The touch screen is not sensitive | The touch screen is not sensitive, borrowing 617-944 one.
No | 3 | Y | RPOL | Blank | Wymiana modu³u OXI, kod b³êdu 0255- Photodiode current setting too low. Kalibracja THb od serwusu, testy. Aparat sprawny. |  | D3F534EDAE2CCACC | Defect | HW | Oxi module | Oxi module |  | Blank | Wymiana modu³u OXI, kod b³êdu 0255- Photodiode current setting too low. Kalibracja THb od serwusu, testy. Aparat sprawny.
No | Visit | Complaint | RSCH | Blank | Erreur de rinçage (sporadique) lors de la calibration | Dépose du module pH-Gaz e.o Remplacé la membrane de l’électrode air $intake du module pH-Gaz e.o Calibration e.o contrôle de qualité e.o | SWO-315609 | Defect | HW | pH/BG Module | Valve, air |  | Blank | Rinse (sporadic) error during calibration | PH-Gaz e.o module drop Replaced the air electrode membrane $intake of the pH-Gaz e.o module Calibration e.o quality control e.o
No | Visit | Complaint | RAME | Calcium issues | Calcium issues | Cleaned cap assembly, replaced membranes and removed and cleaned measuring chamber assembly. Performed complete system cleaning and flush. | SWO-316231 | Cleaned | HW | Fluidic pathway | Fluidic pathway |  | Calcium issues | Calcium issues | Cleaned cap assembly, replaced membranes and removed and cleaned measuring chamber assembly. Performed complete system cleaning and flush.
No | Visit | Complaint | RAME | Have to replace the gas 1 tank every few days. | Customer replaces gas 1 every week compared to other instruments | Replaced gas regulator assembly | SWO-319377 | Defect | HW | Gas module | Valve, Gas |  | Have to replace the gas 1 tank every few days. | Customer replaces gas 1 every week compared to other instruments | Replaced gas regulator assembly
No | Visit | Complaint | RAME | All OXY with dotted lines on calibrations and qc. | All OXY values in dotted lines in qc. Ran complete OXY test and fails neon lamp to low. |  | SWO-315767 | Defect | HW | Oxi module | Lamp unit | Neon | All OXY with dotted lines on calibrations and qc. | All OXY values in dotted lines in qc. Ran complete OXY test and fails neon lamp to low.
No | 25 | Y | RPOL | Blank | Testy podzespo³ów, wymiana kasety sensorowej. Upgrade oprogramowania do wersji 3.3. Aparat sprawny. |  | D3F6A87B244CE4D7 | Assisted Replacement | SC | SC | Reference |  | Blank | Test the sub-assembly, replacing the sensor cassette. Upgrade the software to version 3.3. The camera is functional.
No | Visit | Complaint | RAME | Customer reported by email multiple sensor cassette replaced for failure to initialize glucose and lactate parameters. | Retrieve sensor cassette for research of multiple replacements when required replacement due. | Testing of sensor cassettes in different analyzer to determine root cause. | SWO-316444 | Defect | SC | SC | Glu/Lac |  | Customer reported by email multiple sensor cassette replaced for failure to initialize glucose and lactate parameters. | Retrieve sensor cassette for research of multiple replacements when required replacement due. | Testing of sensor cassettes in different analyzer to determine root cause.
No | Visit | Complaint | RDEU | 23.02.2018, 10:15 Uhr ==> Anrufer Fr. Bernhardt, Tel.: 06213831157 ==> Oxi-Hardwarefehler, Neonkorrektur ausserhalb der Grenzen | QC Fehler sO2 und FHbF (beide erhöht) in Level B und C | Oximetrie-Test (ohne Beanstandung) - Erneuerung Schlauch an Quetschventil - SP erneuert (Los BD-50 noch 194 Aktivitäten) - Funktionsüberprüfung etc. ohne Beanstandung - Service Dump und Datalogs | SWO-316082 | Defect | HW | Oxi module | Pinch valve tube | & SP | 23.02.2018, 10:15 a.m. ==> Caller Fr. Bernhardt, Tel.: 06213831157 ==> Oxi hardware error, neon correction outside the limits | QC error sO2 and FHbF (both increased) in levels B and C | Oximetry test (without complaint) - Renewal hose to squeegee valve - SP renewed (lot BD-50 still 194 activities) - functional check etc. without complaint - Service Dump and Datalogs
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | pCO2定标不过 1225 | 更换测试卡后修复 | SWO-317850 | [5 empty] | Test card, reagent package failure | pCO2 calibration is no more than 1225 | Repair after replacing test card
No | Visit | Complaint | RKOR | Sample clot error | . | Sensor cassette replace. | SWO-323522 | [5 empty] | Sample clot error | . | Sensor cassette replace.
No | Visit | Complaint | RAME | Failing blank calibrations | Failing blank calibrations | Reseated cassette, ran diagnostics and performed LS cal. Ran three levels of QC- all tests passed. | SWO-316483 | [5 empty] | Failing blank calibrations | Failing blank calibrations | Reseated cassette, ran diagnostics and performed LS cal. Ran three levels of QC- all tests passed.
No | Visit | Complaint | RFRA | Erreur SO2 sur cycle de qualité | erreur SO2, tubulure bloquée | mise en place du kit de maintenance annuelle $ $ passage du deproteneisant $ $ paramétrage du site de prélèvement (manuel service p319) | SWO-316357 | [5 empty] | SO2 error on quality cycle | SO2 error, blocked tubing | implementation of the $$ annual maintenance kit switching from the $$ deproteant setting of the sampling site (p319 service manual)
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 泄漏 | 更换测试卡后修复 | SWO-315894 | [5 empty] | Test card, reagent package failure | Leakage | Repair after replacing test card
No | Visit | Complaint | RCHN | 液体传说失败 | 液体传说失败,需要借件903-165和903-199一个 |  | SWO-315888 | Defect | HW | Basic Unit | Cassette retainer | & PCB, wetsection | Liquid Legend failed | Liquid Legends failed and needed to borrow 903-165 and 903-199 a
No | Visit | Complaint | RCHN | 血红蛋白值检测错误 | 血红蛋白值检测错误，借902-980一个 | 更换902-980后修复 by jesson | SWO-315356 | Defect | HW | Oxi module | Hemolyzer |  | Hemoglobin value detection error | Hemoglobin value detection error, borrow 902-980 one | Repair by jesson after replacement 902-980
No | Visit | Complaint | RCHN | 1031 | 1031 借903-364一个 |  | SWO-315993 | HW | Defect | Process Module | Optical Unit | 1031 | 1031 | 1031 Borrowing 903-364 one
No | 18 | Y | RPNZ | Blank |  |  | D3F6942C2427E4F1 | Unknown | Unknown | Unknown | Unknown | More Info Required | Blank
No | Visit | Complaint | RAME | PO2 zero calibration below 0.0. | PO2 zero current calibration below 0.0. PO2 electrode appears to be used up. * Per Tech Support Carlos electrode is in warranty and will be replaced under warranty. | Clean and remembrane PO2 electrode. Perform zero calibration of PO2 channel. | SWO-316446 | Adjustment | HW | pH/BG Module | Zero Current |  | PO2 zero calibration below 0.0. | PO2 zero current calibration below 0.0. PO2 electrode appears to be used up. * Per Tech Support Carlos electrode is in warranty and will be replaced under warranty. | Clean and remembrane PO2 electrode. Perform zero calibration of PO2 channel.
No | 4 | Y | RPAC | The analyser was found to have the two calibration and cleaning solutions incorrect prepared. The DosiCaps had not been correctly prepared. This was corrected. | Analyser not calibrating or performing High QC on Creatinine. Analyser not connected to RADIANCE. | The analyser was found to have the two calibration and cleaning solutions incorrect prepared. The DosiCaps had not been correctly prepared. This was corrected. | D3F6B8BA24532833 | User error | Consumables | Solution bottle | DosiCap not activated | LEH Mail | The analyser was found to have the two calibration and cleaning solutions incorrect prepared. The DosiCaps had not been correctly prepared. This was corrected. | Analyser not calibrating or performing High QC on Creatinine. Analyser not connected to RADIANCE. | The analyser was found to have the two calibration and cleaning solutions incorrect prepared. The DosiCaps had not been correctly prepared. This was corrected.
No | Visit | Complaint | RSCH | Rep Durch Kunde verursacht | Kunde hat sich beschwert da schon 2. pCO2 Elektrode in 4 Monaten. Und nach pCO2 Elektrode Tausch plötzlich Lactat Emp Instabil. | Austausch pCO2 Elektrode abgeholt und zusammen mit Frau Mijatovic Lact angeschaut. Am 21.02.18 wurde durch Personal Wil neu membraniert und bei Lactat vergessen Elektrolyt in Membrane zu füllen. Danach ist die Membrane 1 Woche ohne Probleme gelaufen. Kunde hat noch einmal membraniert und es war sofort i.o | SWO-315974 | User error | Electrode | Membrane | pCO2 | Missing inner fluid | Rep Caused by Customer | Customer has already complained since 2. pCO2 electrode in 4 months. And after pCO2 electrode exchange suddenly lactate emp unstable. | Replacement pCO2 electrode picked up and looked at together with Ms. Mijatovic Lact. On 21.02.18 was re-membraneed by Personal Wil and forgotten to fill electrolyte in membrane at lactate. After that, the membrane ran for 1 week without any problems. Customer has membraned again and it was immediately i.o
No | Visit | Complaint | RAME | PH is unstable, sensitivity and status are low and out of range. | Checked data logs. Found errors for the last two days. Had customer go through many troubleshooting steps over the last couple of days with no change in status. | Requested new electrode replacement. Electrode is still under warranty. Customer will send bad electrode back on RGA after receiving the new one. | SWO-316472 | Assisted Replacement | Electrode | Electrode | pH |  | PH is unstable, sensitivity and status are low and out of range. | Checked data logs. Found errors for the last two days. Had customer go through many troubleshooting steps over the last couple of days with no change in status. | Requested new electrode replacement. Electrode is still under warranty. Customer will send bad electrode back on RGA after receiving the new one.
No | Visit | Complaint | RKOR | Run-Time error | Reboot system Analyzer status check Event log, Activity log check LIS $HIS connection set uo change : connection 1 remove | Reboot system | SWO-317071 | No fault found | SW | Windows | off/on |  | Run-Time error | Reboot system Analyzer status check Event log, Activity log check LIS /HIS connection set uo change : connection 1 remove | Reboot system
No | Visit | Complaint | RSAF | Aborting samples | Viewed data logs and saw samples were aborting on the 28th of Feb @ 20H05 x 4 of the same patient sample number and 2 x same patient aborted in total 6 and all 6 system messages was insufficient sample. Last aborted sample was 28th feb 21H55 and the sample directly afterwards at 21H56 was fine, Service department was only called on the 1st of the March @ 15H00, | Inspected sample probe, fluid flow & inspected Solution pack to see if there are any cups that where stuck or not dropped into the waste pouch, ran 3 x LQC's which where all fine, the patient sample ran after the 28th Feb @21H56 was all ok total 7 samples and the test sample bringing it to a total of 8 x samples which where all ok, The customer only called us a day after encountering the issue. No fault found. | SWO-317375 | No fault found | No fault found | No fault found | No fault found |  | Aborting samples | Viewed data logs and saw samples were aborting on the 28th of Feb @ 20H05 x 4 of the same patient sample number and 2 x same patient aborted in total 6 and all 6 system messages was insufficient sample. Last aborted sample was 28th feb 21H55 and the sample directly afterwards at 21H56 was fine, Service department was only called on the 1st of the March @ 15H00, | Inspected sample probe, fluid flow & inspected Solution pack to see if there are any cups that where stuck or not dropped into the waste pouch, ran 3 x LQC's which where all fine, the patient sample ran after the 28th Feb @21H56 was all ok total 7 samples and the test sample bringing it to a total of 8 x samples which where all ok, The customer only called us a day after encountering the issue. No fault found.
No | Visit | Complaint | RDEU | 01.03.2018, 09:21 Uhr, Fr. Langewellpott, Tel. 0611577388: Fehler: undefined_lookup_string | Pipette 38 Fehler, Kunde hat 2 Proben Falschrum eingesteckt. Nadel Verbogen. | Gerät ausgebaut, Nadel getauscht, Nadel Justage durchgeführt. QC CRP durchgeführt, Alles i.O Mit Kunden besprochen dass der Einsatz in Rechnung gestellt wird. | SWO-316697 | HW | Defect | Aspiration Module | Needle bent |  | 01.03.2018, 09:21, Ms. Langewellpott, Tel. 0611577388: Error: undefined_lookup_string | Pipette 38 error, customer has inserted 2 samples of false rum. Needle bent. | Device removed, needle exchanged, needle adjustment performed. QC CRP performed, All i.O Discussed with customers that the use will be charged.
No | Visit | Complaint | RDEU | 02.03.2018; Hr. Schreiber; Tel.: 0343218183; weißer Monitor | Weißer Bildschirm | Gerät nochmal neugestartet => Bootsequenz wird nicht angezeigt => stattdessen weißes Bild => keine Analysatoraktivität hörbar => CF-Karten Wechsel ohne Besserung => vermutlich CPU Unit defekt Rep-Ersatzgerät (090R0143R005) aufgestellt Helios CCO Dateien eingespielt => Netzwerkdaten eingestellt LP aus Kundengerät eingesetzt => SC Chipdaten aus Kundengerät nicht lesbar => SC aus Rep-Ersatzgerät verwendet (haltbar bis 07.03.) Startphase i.O. => Rep-Ersatzgerät mit Aqure verbunden Zur Zeit noch Kal Fehler bei pCO, Glu und Lac => Info an Station SC noch einlaufen zu lassen => wenn keine Besserung eintritt später SC noch ersetzen | SWO-316839 | Revisit required | HW | Basic Unit | CPU | Replacement analyzer installed | 02.03.2018; Mr. Schreiber; Tel.: 0343218183; white monitor | White screen | Device restarted => Boot sequence is not displayed => white image => no analyzer activity audible => CF card change without improvement => probably CPU unit defective Rep Replacement Device (090R0143R005) installed Helios CCO files recorded => Network data set LP from customer device used => SC chip data from customer device not readable => SC from Rep replacement device used (sustainable until 07.03.) Start-up phase i.O. => Rep replacement device connected to Aqure Currently Kal errors with pCO, Glu and Lac => Info to allow station SC to run in => if no improvement occurs later to replace SC
No | 3 | Y | RCZE | Blank |  |  | D3F7AA689A272C01 | Defect | HW | Battery module | Battery module |  | Blank
No | Visit | Complaint | RFRA | Erreur SO2 sur cycle de qualité | maintenance | maintenance | SWO-316684 | [5 empty] | SO2 error on quality cycle | maintenance | maintenance
No | Visit | Complaint | RCHN | 报错5784 5684 | 报错5784 5684，借903-236一个 | 更换903-236后修复 by jesson | SWO-315322 | SW | Defect | Analyzer Control Module | Process Wheel |  | 5784 5684 | 5784 5684, loan 903-236 one | Repair by jesson after replacement 903-236
No | Visit | Complaint | RCHN | 质控结果检测不准 | 质控结果检测不准 | 未检测出质控不过，后续跟踪by tim | SWO-317918 | No fault found | No fault found | No fault found | No fault found | LEH Email | Quality control results are not detected | Quality control results are not detected | No qc but no detection, follow-up tracking by tim
No | Visit | Complaint | RCHN | 报错1029 | 借903-364,944-212（13632缺货），944-216（13940），944-214（13886），944-258（13918），944-394（13848） |  | SWO-316538 | HW | Defect | Process Module | Optical Unit |  | 1029 error | Loan 903-364,944-212 (13632 out of stock), 944-216 (13940), 944-214 (13886), 944-258 (13918), 944-394 (13848)
No | Visit | Complaint | RCHN | 连线断裂 | 连线断裂，需要借件636-571 | 屏幕显示不正常，更换636-571备件后修复 by jesson | SWO-315938 | HW | Defect | Analyzer Control Module | Cable, Display |  | The wire is broken. | The wire is broken and requires a loan 636-571. | Screen display is not correct, replace 636-571 spare parts after repair by jesson
No | Visit | Complaint | RDEU | The Radiance issue, was the modem needing a reboot. It was powered down and turned back on. All is now fine with this. | FlexQ Modul erkennt sporadisch die Spritzen nicht, trotz Abgleich | FlexQ Modul erneuert Abgleich durchgeführt, DMS Software neu installiert | SWO-316629 | Defect | HW | FlexQ Module | FlexQ Module |  | The Radiance issue, was the modem needing a reboot. It was powered down and turned back on. All is now fine with this. | FlexQ module sporadically does not detect the syringes, despite matching | FlexQ module renewed Matching performed, DMS software reinstalled
No | Visit | Complaint | RAME | The errors are 0258 and 0316. The neon lamp is too low. | Found neon lamp errors and neon lamp failing oxi-test. |  | SWO-316871 | Revisit required | HW | Oxi module | Lamp unit | Neon | The errors are 0258 and 0316. The neon lamp is too low. | Found neon lamp errors and neon lamp failing oxi-test.
No | Visit | Complaint | RAME | Neon lamp 580 errors | Confirm 580 errors in QC and cal results, will need lamp, and Hemolyzer Assy, new lamp may have been compromised in shipping. | Replaced Lamp and Hemolyzer assy, set up lamp adjustment to 3292, and ran OXI test complete, all passed | SWO-316893 | Defect | HW | Oxi module | Lamp unit | Neon | Neon lamp 580 errors | Confirm 580 errors in QC and cal results, will need lamp, and Hemolyzer Assy, new lamp may have been compromised in shipping. | Replaced Lamp and Hemolyzer assy, set up lamp adjustment to 3292, and ran OXI test complete, all passed
No | Visit | Complaint | RIND | Instrument not getting ON. | found 4 amp fuse are blowing at main power supply and found problem in power supply unit. | Replaced the power supply unit with working power supply unit . checked the system performances ,did two point calibration and samples found all values with in the range ,found instrument working fine. | SWO-316757 | Defect | HW | DMS_PC | Power Supply |  | Instrument not getting ON. | found 4 amp fuse are blowing at main power supply and found problem in power supply unit. | Replaced the power supply unit with working power supply unit . checked the system performances ,did two point calibration and samples found all values with in the range ,found instrument working fine.
No | Visit | Complaint | RKOR | Software Re-installation | Software Re-installation Cal. adjust log, Activity log, Event log check Analyzer status check Consumables capacity check | Software Re-installation | SWO-316595 | SW | Defect | Analyzer Control Module | CPU |  | Software Re-installation | Software Re-installation Cal. adjust log, Activity log, Event log check Analyzer status check Consumables capacity check | Software Re-installation
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 温度-A PH,CO2,CL定标不过 失败-SP | 连续更换3张测试卡及1个试剂包后修复 | SWO-317869 | [5 empty] | Test card, reagent package failure | Temperature-A PH, CO2, CL calibration but Failure-SP | Repair after replacing 3 test cards and 1 reagent kit in a row
No | Visit | Complaint | RCHN | 气压有偏差和检测到气泡，报错689和502 | 气压有偏差和检测到气泡，报错689和502，需要借件902-823、902-516、902-505、902-605 | 更换902-823、902-516、902-505、902-605备件后修复 | SWO-317010 | Defect | HW | Basic Unit | Gas module |  | Air pressure is biased and bubbles are detected, and errors 689 and 502 are reported | Air pressure is biased and bubbles are detected, errors 689 and 502 are required, and 902-823, 902-516, 902-505, 902-605 are required | Repair after replacement of 902-823, 902-516, 902-505, 902-605 spare parts
No | 25 | Y | RPNZ | Blank | Sensor cassette replacement due to 1361: internal reference electrode error in sensor cassette. |  | D417DBE1202D1988 | Assisted Replacement | SC | SC | Reference | 1361 | Blank | Sensor cassette replacement due to 1361: internal reference electrode error in sensor cassette.
Yes | Visit | Complaint | RCHN | 无法读取测试卡芯片，报错1123 | 无法读取测试卡芯片，报错1123，需要借件903-581 | 需加借备件，此单关单 by tim | SWO-317070 | Re classified RMED | HW | Basic Unit | PCB, Sensor interface | 1123 | Failed to read test card chip, error 1123 | Failed to read test card chip, error 1123, need to borrow 903-581 | Spare parts required, this single by tim
No | Visit | Complaint | RLTD | Software re-installed because the Proxy Host was not starting | On a previous visit it was noticed that the Proxy Host was not starting. | Re-installed the analyser software to get back the Proxy Host working. | SWO-317665 | Defect | SW | RMED SW | RMED SW vs PROXY |  | Software re-installed because the Proxy Host was not starting | On a previous visit it was noticed that the Proxy Host was not starting. | Re-installed the analyser software to get back the Proxy Host working.
No | Visit | Complaint | RDEU | 05.03.2018 Frau Wunsch 040790202398 meldet, dass sie immer wieder das Verbrauchsmaterial tauschen soll. | Fehlersuche am Analysator durchgeführt. | Jahreswartung durchgeführt. Und alle LS Sensoren gereinigt. | SWO-317418 | Defect | HW | Basic Unit | PM kit |  | 05.03.2018 Mrs Wunsch 040790202398 announces that she is to exchange the consumables again and again. | Troubleshooting is performed on the analyzer. | annual maintenance. And all LS sensors cleaned.
No | Visit | Complaint | RCHN | 无法读取测试卡芯片，报错1123 | 加借903-165一个 | 未使用配件，清洁机器后修复by tim | SWO-317321 | Defect | HW | Basic Unit | PCB, Wet section Control |  | Failed to read test card chip, error 1123 | Borrow 903-165 one | Unused accessories, repair by tim after cleaning the machine
No | Visit | Complaint | RCHN | 干燥前温度超范围 | 干燥前温度超范围，借903-210一个 | 更换903-210 正常by tim | SWO-317228 | Defect | HW | Process Module | PCB, Cup dryer |  | Overrange temperature before drying | Over-range temperature before drying, borrow 903-210 one | Replace 903-210 Normal by tim
No | Visit | Complaint | RIND | Intervention required. | Intervention required. Solutions Checked and found Cleaning Solution not installed properly. | Analyser turn OFF for half Hour. Service cal for Pump, LS and Oxy module done. All passed. Inform to concern for proper installation of Cleaning Solution. Protean removal and Cleaning done. Analyser Tried to calibrate and Calibrated. Analyser Serviced and calibrated properly. Handed over in a fine working condition. | SWO-318062 | User error | Consumables | Solution Bottle | Cleaning |  | Intervention required. | Intervention required. Solutions Checked and found Cleaning Solution not installed properly. | Analyser turn OFF for half Hour. Service cal for Pump, LS and Oxy module done. All passed. Inform to concern for proper installation of Cleaning Solution. Protean removal and Cleaning done. Analyser Tried to calibrate and Calibrated. Analyser Serviced and calibrated properly. Handed over in a fine working condition.
No | Visit | Complaint | RIND | Touch Screen Problem. | checked the instrument performance ,found Barcode reader not working and touch not activating properly. | Cleaned the touch screen and connected the mouse.update the firmware for touch screen. now touch screen working fine, cleaned the barcode reader but reader not working .Did calibration and sample all results found with in the range. | SWO-320834 | Defect | HW | DMS_PC | Touch screen | Firmware Update | Touch Screen Problem. | checked the instrument performance ,found Barcode reader not working and touch not activating properly. | Cleaned the touch screen and connected the mouse.update the firmware for touch screen. now touch screen working fine, cleaned the barcode reader but reader not working . Did calibration and sample all results found with in the range.
No | Visit | Complaint | RKK | キャリブレーションエラー（tHbドリフト)。ランプユニット交換。tHbキャル実施。 | キャリブレーションエラー（tHbドリフト)。 | ランプユニット交換。tHbキャル実施。 | SWO-320818 | Defect | HW | Oxi module | Lamp unit |  | Calibration error (tHb drift). Lamp unit replacement. THb cal implementation. | Calibration error (tHb drift). | Lamp unit replacement. THb cal implementation.
No | Visit | Complaint | RAME | Qc not running. | Qc not running. Upon further review the carousel would not go up. | Replaced the carousel stepping motor. | SWO-317667 | Defect | HW | Autocheck Module | Stepper Motor, Autocheck Carroussel |  | Qc not running. | Qc not running. Upon further review the carousel would not go up. | Replaced the carousel stepping motor.
No | Visit | Complaint | RKOR | Calibration, Quality Control results out of range | Perform Calibration 5 Times Calibration result check Perform Quality Control A,B,C QC result check | Perform Calibration 5 Times | SWO-320184 | Unknown | Performance | Performance | Performance (QC) |  | Calibration, Quality Control results out of range | Perform Calibration 5 Times Calibration result check Perform Quality Control A,B,C QC result check | Perform Calibration 5 Times
No | Visit | Complaint | RBNL | Fuite détecté | Le client a eu beaucoup de problèmes hier pour redémarrer l'appareil après avoir remplacer le pack de solution. Le joint d'intro a déjà été remplacé. | Remplacement du joint d'introduction par un nouveau et du coup le test 'pression' passe avec une valeur de 69. Remis en route , calibration et contrôle :OK | SWO-318097 | Defect | HW | Basic Unit | PM kit |  | Leak detected | The customer had a lot of problems yesterday to restart the device after replacing the solution pack. The intro seal has already been replaced. | Replacing the introductory seal with a new one and the 'pressure' test passes with a value of 69. Restarted, calibration and control:OK
No | Visit | Complaint | RLTD | They are unable to clear the blockage from the Analyser. Clare Bailey | Attempted to flush but the Analyser is completely blocked. Replacement Hemolyser required. | Performed intensive manual flush, but no success. Removed Hemolyser where the blockage has been identified. Attempted to unblock with hypochlorite, but still I was unable to clear the Hemolyser. | SWO-319070 | Defect | HW | Oxi module | Hemolyzer | Clot | They are unable to clear the blockage from the Analyser. Clare Bailey | Attempted to flush but the Analyser is completely blocked. Replacement Hemolyser required. | Performed intensive manual flush, but no success. Removed Hemolyser where the blockage has been identified. Attempted to unblock with hypochlorite, but still I was unable to clear the Hemolyser.
No | Visit | Complaint | RDEU | Messabbrüche und Hb Differenzen zum ABL 800 | Gerät überprüft und gereinigt | tHb Kal durchgeführt und Vergleichsmessungen durchgeführt | SWO-318017 | Adjustment | HW | Oxi module | tHb cal |  | Measurement interruptions and Hb differences to the ABL 800 | Device checked and cleaned | tHb Kal carried out and comparative measurements carried out
No | Visit | Complaint | RHUN | összeméréseknél a sO2 alámér, | szerviz csere készülékkel történő összemérés, a jelenség a csere készüléken is mutatkozik, QC mérés OK | oldat kazetta csere után a mérések reális értéket adnak | SWO-318126 | [5 empty] | összeméréseknél a sO2 alámér, | With a service replacement device, The phenomenon is reflected on the replacement device, QC Measurement OK | After a cartridge replacement, the measurements give a realistic value
No | Visit | Complaint | RFRA | Erreur de mesures PH et PO2 | Retour sur site | Retour automate après essais comparatifs | SWO-318381 | [5 empty] | Error of PH and PO2 measurements | Back on site | Automated return after comparative testing
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | LAC定标失败 | 更换测试卡正常。 | SWO-318591 | [5 empty] | Test card, reagent package failure | LAC calibration failure | The replacement test card is OK.
No | Visit | Complaint | RDEU | QC RiliBÄK Verletzung pH, pCO2 und pO2 | Level1 wird sporadisch nicht erkannt. Sporadische Verschiebung von pCO2 und pH. Sporadisch pO2 außer Bereich. | Nach Rücksprache mit Herrn Baumstark ABL90 (090R0579N006) zur Reparaturüberbrückung geliefert und in Betrieb genommen. Kalibrationen und QC’s vorgenommen. | SWO-318150 | [5 empty] | QC RiliBÄK violation pH, pCO2 and pO2 | Level1 is not detected sporadically. Sporadic displacement of pCO2 and pH. Sporadic pO2 out of range. | After consultation with Mr. Baumstark ABL90 (090R0579N006) delivered for repair bridging and put into operation. calibrations and QC's.
No | Visit | Complaint | RCHN | 定标失败 | 定标失败，借954-737，903-111各一个。 | 更换库存954-737，903-111后修复by tim | SWO-317156 | [5 empty] | Bid failed | The bid failed, borrowing 954-737, 903-111 each. | Replacement of inventory 954-737, 903-111 after repair by tim
No | Visit | Complaint | RCHN | 报错1064 | 报错1064，借903-581一个 | 更换备机903-581备件后修复by tim | SWO-317151 | Defect | HW | Basic Unit | PCB, Sensor interface | 1064 | 1064 error | 1064, loan 903-581 one | Repair by tim after replacement of spare parts 903-581
No | Visit | Complaint | RCHN | 定标失败 | 定标失败，借944-212（13696）一个 | 更换944-212（13696）备件后修复(数据丢失）by tim | SWO-317272 | Assisted Replacement | Consumables | Test Kit | Cartridge, TnI |  | Bid failed | Failed to bid, borrow 944-212 (13696) one | Repair (data loss) by tim after replacing 944-212 (13696) spare parts
No | Visit | Complaint | RLTD | Ian Tel 01865 221596 | The analyser had been failing QCs intermittently due oxi spectrum mismatch errors. | Replaced the hemolyzer and this has resolved the problem. | SWO-318418 | Defect | HW | Oxi module | Hemolyzer |  | Ian Tel 01865 221596 | The analyser had been failing QCs intermittently due oxi spectrum mismatch errors. | Replaced the hemolyzer and this has resolved the problem.
No | 6 | Y | RPOL | Blank | Aparat w czasie aspiracji ci¹gn¹³ bokiem powietrze. Wymiana wê¿yka ³¹cz¹cego inlet z komor¹ pH $BG. Testy. Aparat sprawny. |  | D3FD257E724632F1 | Defect | HW | pH/BG Module | Inlet Tube |  | Blank | The camera during the aspiration will pull the air. Replacing the inlet of the intake with the pH /BG. Tests. The camera is functional.
Yes | Visit | Complaint | RAME | Customer concerned there was a shift in PCO2 starting Jan 14th a. 8:58 am | Looked at data logs and looked over analyzer, found some dummy electrodes with blood. Also found ranges where very tight. | Cleaned dummy electrodes, assisted Chuck Rall on training 3 technicians and also having customer reestablish there ranges. | SWO-318507 | Re classified RMED | Courtesy |  |  |  | Customer concerned there was a shift in PCO2 starting Jan 14th a. 8:58 am | Looked at data logs and looked over analyzer, found some dummy electrodes with blood. Also found ranges where very tight. | Cleaned dummy electrodes, assisted Chuck Rall on training 3 technicians and also having customer reestablish there ranges.
No | Visit | Complaint | RAME | High Crea check issues. | Crea diff high out of control range on S8377. Customer remembrance crea electrodes and replaced cleaning solution. | Cleaned electrodes and electrode contacts. | SWO-319171 | Assisted Replacement | Electrode | Membrane | Crea | CAPA 625 | High Crea check issues. | Crea diff high out of control range on S8377. Customer remembrance crea electrodes and replaced cleaning solution. | Cleaned electrodes and electrode contacts.
Yes | Visit | Complaint | RFRA | Validation avant basculement sur la nouvelle plateforme |  | Validation avant basculement sur la nouvelle plateforme | SWO-318171 | Re classified RMED | SWO | Mail Eric Van Busch/ 2018-03-15: I got the confirmation that this SWO is related to a kind of check-up of the analyzer before its relocation. There wasn’t any technical issue or software upgrade. QC have been used to validate it but not reported into the SWO. |  | LEH Mail | Pre-tipping validation on new platform |  | Pre-tipping validation on new platform
No | Visit | Complaint | RAME | Oximetry Hardware Error | Error 879, no cal 1 solution during cal Error 367, ph $Bg not filled within time limit Error 251, halogen lamp intensity too high Error 579, Oximetry Hardware problem, not possible to measure. Found Cal 1 solution warning indicating solution was low, aspirating air, reagent status displayed 6%. Performed Service diagnostics, determine cause of Oximetry hardware error was halogen lamp voltage was above limit. | Replaced cal 1 solution, primed, Adjusted Oximetry lamp output to a nominal 550pA. Performed complete Oxi test, service and user tHb cal. Performed 2PT calibration, passing, No further issues observed. | SWO-317838 | Defect | HW | Oxi module | Lamp unit | Halogen | Oximetry Hardware Error | Error 879, no cal 1 solution during cal Error 367, ph /Bg not filled within time limit Error 251, halogen lamp intensity too high Error 579, Oximetry Hardware problem, not possible to measure. Found Cal 1 solution warning indicating solution was low, aspirating air, reagent status displayed 6%. Performed Service diagnostics, determine cause of Oximetry hardware error was halogen lamp voltage was above limit. | Replaced cal 1 solution, primed, Adjusted Oximetry lamp output to a nominal 550pA. Performed complete Oxi test, service and user tHb cal. Performed 2PT calibration, passing, No further issues observed.
Yes | 6 | Y | RPOL | Blank | Wyciek z aparatu. Zdiagnozowano przyczynê- spadniêty wê¿yk œciekowy. Problem rozwi¹zano, wyczyszczono aparat. Analizator sprawny. |  | D3FCB4B6722CA745 | [5 empty] | Blank | Leakage from the camera. The cause of the decreased fluid intake has been diagnosed. Problem resolved, the camera was cleared. Functional Analyzer.
No | Visit | Complaint | RDEU | D-DimerKassette lässt sich nicht einlegen. Obwohl identisches Los, wie verbrauchte Testkassette. Kalibration unter aktive Kalibrationen nicht verfügbar. | In Dateiaktivitäten wurden keine Kalibrationen gelöscht. D-Dimer musste bereits 2* kalibriert werden. | Parameter erneut (zum 3ten mal) kalibriert. Servicedump gezogen und an TSP verschickt. | SWO-318410 | User error | SW | Setup | Cal, Records deleted |  | D-DimerCassette cannot be inserted. Although identical lot, as consumed test cartridge. Calibration not available under active calibrations. | No calibrations have been deleted in file activities. D-Dimer already had to be calibrated 2*. | parameter is calibrated again (for the 3rd time). service dump and sent to TSP.
No | Visit | Complaint | RLTD | Suspected clot, can't flush analyser Angela 07787 598407 | Clot suspect unable to flush. | Removed cassette and flushed system with out cassette . System flushed ok. Tried to flush cassette and found unable to get a good flow of water. Cassette blocked. Replaced cassette and run service program all passed ok. | SWO-319055 | Assisted Replacement | SC | SC | SC, Clot |  | Suspected clot, can't flush analyser Angela 07787 598407 | Clot suspect unable to flush. | Removed cassette and flushed system with out cassette . System flushed ok. Tried to flush cassette and found unable to get a good flow of water. Cassette blocked. Replaced cassette and run service program all passed ok.
No | Visit | Complaint | RFRA | Erreur vide test de pression Circuit fluides bouché | Test circuit fluides | Remplacer cooxymetre (bouché) | SWO-320112 | Defect | HW | Oxi module | Clot |  | Empty pressure test error Clogged fluid circuit | Fluid circuit test | Replace cooxymetre (clogged)
No | Visit | Complaint | RCHN | 接口板腐蚀 | 接口板腐蚀，借902-972一个 | 更换902-972后修复by tim | SWO-317859 | Defect | HW | Basic Unit | PCB, Sensor connector assembly |  | Interface plate corrosion | Interface plate corrosion, borrow 902-972 one | Repair by tim after replacement 902-972
No | Visit | Complaint | RIND | Service Call. | V $C failed. | Reset the solution pack from the analyzer, from 5 ml syringe take out some solution from each pouch for remove the bubbles. Again install the solution pack and sensor cassette, successfully done, analyzer handed over to staff in well working condition. | SWO-320631 | [5 empty] | Service Call. | V /C failed. | Reset the solution pack from the analyzer, from 5 ml syringe take out some solution from each pouch for remove the bubbles. Again install the solution pack and sensor cassette, successfully done, analyzer handed over to staff in well working condition.
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 无法抽吸溶液，1343 | 更换试剂包后修复 | SWO-320823 | [5 empty] | Test card, reagent package failure | Unable to pump solution, 1343 | Repair after replacement of reagent kit
No | Visit | Complaint | RCHN | K,CA响应错误 | K,CA响应错误，借903-581一个 | 更换903-581后修复by tim | SWO-317861 | Defect | HW | Basic Unit | PCB, Sensor interface | Ca, K+ | K, CA response error | K, CA response error, borrow 903-581 one | Repair by tim after replacement 903-581
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 泵定标失败，不均匀的冲洗液，报错1178和1083 | 更换试剂包后修复 | SWO-318550 | Assisted Replacement | SP | SP | Pump tubes | 1178 | Test card, reagent package failure | Failed pump calibration, uneven flushing fluid, error 1178 and 1083 | Repair after replacement of reagent kit
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 泵定标失败，不均匀的冲洗液，报错1178和1083 | 更换试剂包后修复 | SWO-318551 | [5 empty] | Test card, reagent package failure | Failed pump calibration, uneven flushing fluid, error 1178 and 1083 | Repair after replacement of reagent kit
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | CA,K,NA,PH,CL,LAC定标不过 | 更换测试卡后修复 | SWO-319350 | Assisted Replacement | SC | SC | Reference |  | Test card, reagent package failure | CA, K, NA, PH, CL, LAC定标不过 | Repair after replacing test card
No | Visit | Complaint | RESP | El equipo detecta fugas y no se puede usar. | Fallo de PCO2 | El parámetro de PO2 presenta derivas se sustituye el casette de electrodos por uno nuevo, se verifica los tubos internos sin encontrar fallo. El equipo sigue sin funcionar, qued pendiente el cambio de la sensor board pcb. | SWO-320391 | Revisit required | HW | Basic Unit | PCB, Sensor interface |  | The equipment detects leaks and cannot be used. | PCO2 failure | The PO2 parameter has drifts replaces the electrode cassette with a new one, the internal tubes are checked without finding fault. The equipment is still not working, keep pending the change of the sensor board pcb.
No | Visit | Complaint | RFRA | demande intervention erreur de PO2 | Cassette avec du sang à l.interieur, | Changement de la cassette, calibration ok, passage des cq ok | SWO-328792 | [5 empty] | PO2 error intervention request | Cassette with blood inside, | Change of tape, calibration ok, passage of cq ok
No | Visit | Complaint | RDEU | 20.02.2018, 09:04 Uhr ==> Anrufer Fr. Kuhrau, Tel.: 023319742111 ==> Demo Key läuft ab | Systemkeys installiert | Systemkeys installiert | SWO-318957 | LEH | SW | Demo Key exp | Demo Key exp |  | 20.02.2018, 09:04 ==> Caller Fr. Kuhrau, Tel.: 023319742111 ==> Demo Key expires | System keys installed | System keys installed
No | Visit | Complaint | RAME | Auto check Motor is out will allow it to do automated controls | Auto check.carousel error. House will not lift up. | Tried 14 pole cable and did not fix. Stepper motor is on back order. | SWO-319123 | Revisit required | HW | Autocheck Module | Stepper Motor, Autocheck |  | Auto check Motor is out will allow it to do automated controls | Auto check.carousel error. House will not lift up. | Tried 14 pole cable and did not fix. Stepper motor is on back order.
No | Visit | Complaint | RKOR | Storage wheel is failed | Remove broken catridge | Remove broken catridge | SWO-320222 | Defect | Consumables | Test Kit | Cartridge, Broken |  | Storage wheel is failed | Remove Broken Catridge | Remove Broken Catridge
No | Visit | Complaint | RFRA | Faux positif sur mesure TnI | Relever ServiceDump Passage échantillons tests | Passages 5tubes blancs Analyse résultats Ok | SWO-321957 | Replaced | HW | Aspiration Module | Tubes |  | Fake bespoke tuning TnI | Meet ServiceDump Passing test samples | Passages 5white tubes Analysis results Ok
No | Visit | Complaint | RAME | Network Connection | Results not crossing. Roma.in file needs to be checked for serial port settings. | External=1 was not remarked out in ini file. Made change and now analyzer shoes a connection. Worked with Julie for a while and don’t see query results. | SWO-319924 | Setup | SW | Roma.ini | His/Lis |  | Network Connection | Results not crossing. Roma.in file needs to be checked for serial port settings. | External=1 was not remarked out in ini file. Made change and now analyzer shoes a connection. Worked with Julie for a while and don’t see query results.
No | Visit | Complaint | RAME | To turn on the serial port | To change the Roma.ini | I removed the barcode reader from the Roma.ini | SWO-320019 | Setup | SW | Roma.ini | Barcode Reader |  | To turn on the serial port | To change the Roma.ini | I removed the barcode reader from the Roma.ini
No | Visit | Complaint | RAME | To open the serial | To change the Roma.ini open the serial port | I changed the Roma.ini to use the serial port | SWO-320018 | Setup | SW | Roma.ini | Barcode Reader |  | To open the serial | To change the Roma.ini open the serial port | I changed the Roma.ini to use the serial port
Yes | Visit | Complaint | RAME | Rechargeable battery and mobile stand | Rechargeable battery has arrived and needs to be installed and analyzer needs to be put on mobile stand. | Installed rechargeable battery. Put analyzer on mobile chart. | SWO-319780 | Re classified RMED | Planned SWO |  |  |  | Rechargeable battery and mobile stand | Rechargeable battery has arrived and needs to be installed and analyzer needs to be put on mobile stand. | Installed rechargeable battery. Put analyzer on mobile chart.
No | 7 | Y | RPAC | Blank | Inlet not moving |  | D3FFEACD5E451C51 |  |  |  |  | LEH Update, please return | Blank | Inlet not moving
No | Visit | Complaint | RDEU | 07.03.2018; Sr. Elke; Tel.: 02017235151; Display ist schwarz | Display ist schwarz. LEDs im Mischerbereich und im Einlass nicht aktiv. Im Hintergrund Betriebsgeräusche wahrnehmbar. | Netzteil erneuert (defektes Netzteil: 910-285P0024N996). Startphase fehlerfrei durchlaufen. Funktionskontrolle: I.O. | SWO-319686 | Defect | HW | Basic Unit | Power Supply |  | 07.03.2018; Sr. Elke; Tel.: 02017235151; Display is black | Display is black. LEDs in the mixer range and in the inlet are not active. In the background operating noises can be detected. | Power supply renewed (defective power supply: 910-285P0024N996). Run through the start-up phase without errors. Functional control: I.O.
No | Visit | Complaint | RCHN | K,CA定标不过 | K,CA定标不过，借902-972一个 | 更换902-972后修复by tim | SWO-318012 | Defect | HW | Basic Unit | PCB, Sensor connector assembly |  | K, CA calibration is not | K, CA calibration, borrow 902-972 a | Repair by tim after replacement 902-972
No | Visit | Complaint | RCHN | 打印机故障 | 打印机故障，借910-286一个 | 更换910-286后修复by tim | SWO-318577 | Defect | HW | Printer Module | Printer Module |  | Printer failure | Printer failure, borrow 910-286 one | Repair by tim after replacement 910-286
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | LAC,GLU定标不过 0378 0375 | 更换测试卡后修复 | SWO-320152 | [5 empty] | Test card, reagent package failure | LAC, GLU calibration but 0378 0375 | Repair after replacing test card
Yes | Visit | Complaint | RCHN | 处理盘无法移动 5684 | 处理盘无法移动 5684，借636-545，903-031，580-098两个，903-198一个。 | 更换一个636-533后修复by tim | SWO-317268 | [5 empty] | Processing disk cannot be moved 5684 | The processing disk cannot be moved 5684, borrowing 636-545, 903-031, 580-098 two, 903-198 one. | Replace a 636-533 after repair by tim
No | Visit | Complaint | RCHN | 蓝屏 | 蓝屏，借910-343一个 | 更换910-343后修复 byi tim | SWO-318587 | Defect | HW | DMS_PC | CPU |  | Blue screen | Blue screen, borrow 910-343 one | Repair byi tim after replacing 910-343
No | Visit | Complaint | RCHN | 显示屏故障 | 显示屏故障，借910-298一个 | 更换910-298后修复by tim | SWO-318702 | Defect | HW | Analyzer Control Module | Display |  | Display failure | Display fault, borrow 910-298 one | Repair by tim after replacement 910-298
No | Visit | Complaint | RCHN | 显示屏故障 | 显示屏故障，定标卡扫描器故障，加借902-924一个 | 更换902-924后修复by tim | SWO-319347 | Defect | HW | Cover and Chassis | Barcode reader, display |  | Display failure | Display failure, calibration card scanner fault, plus 902-924 one | Repair by tim after replacement 902-924
No | Visit | Complaint | RCHN | BNP不出值 | BNP不出值，借905-908一个 | 更换905-908后故障仍存在。加借 636-561，903-210各一个。by tim | SWO-318695 | Defect | HW | Process Module | PCB, Cup dryer |  | BNP not value | BNP not value, borrow 905-908 a | The fault still exists after replacing the 905-908. Borrow 636-561, 903-210 each. by tim
Yes | 10 | Y | RPAC | Blank | Install repaired OXI module |  | D3FF3E085E3A88F9 | [5 empty] | Blank | Install Sony OXYGEN Module
No | Visit | Complaint | RFRA | Problème de tension de la lampe halogène | Contrôle de la focalisation de la lampe | Realignement de la lampe. Essais | SWO-319756 | Adjustment | HW | Oxi module | Lamp unit | Halogen | Halogen lamp voltage problem | Controlling the focus of the lamp | Realignment of the lamp. Tests
No | Visit | Complaint | RKK | pCO2ドリフトエラー | 定時のCALでPCO2のステータスが２位まで低下し、感度・ドリフトともにエラーとなっている。 | CAL中にガス１がうまく回路に流れていない状態になっているためガス用のメンブランバルブ交換。CAL等動作チェック。 | SWO-320041 | Defect | HW | Gas module | Valve, Gas |  | pCO2 Drift Error | PCO2 status is reduced to second place in the scheduled CAL, and both sensitivity and drift are errors. | Because gas 1 is not flowing well to the circuit during CAL, the membrane valve exchange for the gas. Cal operation check.
Yes | Visit | Complaint | RAME | Customer called with the ABL 90 disconnecting from the wireless | Wireless connection issues | Ongoing investigation on why and when the wireless is disconnecting on the analyzer. On site today to check logs and talk with customer about the wireless issues. Next step will be to contact hospital IT to help troubleshoot wireless issues. | SWO-320615 | Re classified RMED | Folow up visit |  |  |  | Customer called with the ABL 90 disconnecting from the wireless | Wireless connection issues | Ongoing investigation on why and when the wireless is disconnecting on the analyzer. On site today to check logs and talk with customer about the wireless issues. Next step will be to contact hospital IT to help troubleshoot wireless issues.
No | Visit | Complaint | RKOR | Liquid transport error during rinse | Re-installation Solution Pack Replacement new Solution Pack | Replacement new Solution Pack | SWO-320665 | Assisted Replacement | SP | SP | FT |  | Liquid transport error during rinse | Re-installation Solution Pack Replacement new Solution Pack | Replacement new Solution Pack
Yes | Visit | Complaint | RKOR | Liquid transport error during rinse | Re-installation Solution Pack Replacement new Solution Pack | Replacement new Solution Pack | SWO-320719 | Re classified RMED | Double entry |  |  |  | Liquid transport error during rinse | Re-installation Solution Pack Replacement new Solution Pack | Replacement new Solution Pack
No | Visit | Complaint | RAME | Customer reporting message to replace solution pack, error 1397. Customer unable to access analyzer due to wrong password entry, customer found correct password and was able to enter manager mode. | Review data logs and analyzer status. Determined problem to be failing Level 3 QC for pO2 running high out of range, 34, range is 3.2 - 28.7. Performed service diagnostics, fluidic flow failed for solution pack test, rinse, and pO2 test. | Cleaned very dirty inlet gasket holder, cleaned pO2 sensor. Retested fluid flow, all diagnostics now passing. QC still failing, L3, pO2. Replaced solution pack to resolve QC issue. | SWO-319996 | Assisted Replacement | SP | SP | QC, pO2 | 1397 | Customer reporting message to replace solution pack, error 1397. Customer unable to access analyzer due to wrong password entry, customer found correct password and was able to enter manager mode. | Review data logs and analyzer status. Determined problem to be failing Level 3 QC for pO2 running high out of range, 34, range is 3.2 - 28.7. Performed service diagnostics, fluidic flow failed for solution pack test, rinse, and pO2 test. | Cleaned very dirty inlet gasket holder, cleaned pO2 sensor. Retested fluid flow, all diagnostics now passing. QC still failing, L3, pO2. Replaced solution pack to resolve QC issue.
No | Visit | Complaint | RAME | Change inlet gasket screen | Air detected pressure vacuum error 1063 | Looked at gaskets When through analyzer troubleshooting screens Analyzer cleared error Ran cals and qc | SWO-321914 | No fault found | No fault found | No fault found | No fault found | 1063 | Change inlet gasket screen | Air detected pressure vacuum error 1063 | Looked at gaskets When through analyzer troubleshooting screens Analyzer cleared error Ran cals and qc
No | Visit | Complaint | RDEU | K zu niedrig bei Blutmessungen | Elektrolyte waren heute morgen instabil | Gerät überprüft, Meßkassette erneuert, tHb Kal durchgeführt, Kal und QC o.k. | SWO-320200 | Assisted Replacement | SC | SC | K+ |  | K too low for blood measurements | Electrolytes were unstable this morning | Device checked, measuring cassette renewed, tHb Kal performed, Kal and QC ok.
No | Visit | Complaint | RLTD | They are unable to clear the blockage from the Analyser. Clare Bailey | Hemolyser replacement required. | Replaced Hemolyser and ran complete OXI test. Performed tHb calibration and tested. | SWO-321021 | Defect | HW | Oxi module | Hemolyzer |  | They are unable to clear the blockage from the Analyser. Clare Bailey | Hemolyser replacement required. | Replaced Hemolyser and ran complete OXI test. Performed tHb calibration and tested.
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | LAC漂移错误 | 更换测试卡 | SWO-320668 | [5 empty] | Test card, reagent package failure | LAC Drift Error | Replace the test card
No | Visit | Complaint | RESP | El equipo detecta fugas y no se puede usar. | El equipo presenta fallos en la PO2 | Se sustituye la Sensor Board PCB, se realizan calibraciones, se sustituye el pack de soluciones por uno nuevo, realizadas calibraciones y controles sin incidencias | SWO-320895 | Defect | HW | Basic Unit | PCB, Sensor interface | pO2 | The equipment detects leaks and cannot be used. | The team has failures in the PO2 | The PCB Sensor Board is replaced, calibrations are performed, the solution pack is replaced with a new one, calibrations and controls without incidents
No | Visit | Complaint | RESP | INLET NO FUNCIONA | INLET NO FUNCIONA | INLET SE QUEDA A MITAD DE RECORRIDO NO PUEDIENDO FUNCIONAR, SE DESMONTA Y SE LE SUSTITUYE SONDA Y GASKET, SE AJUSTA MEDIANTE SERVICIO TECNICO Y SEE EFECTUAN PRUEBAS DE FUNCIONAMIENTO. CALIBRADO Y VERIFICADO. | SWO-320493 | Adjustment | HW | Auto Inlet module | Inlet Position Cal. |  | INLET NO WORKS | INLET NO WORKS | INLET IS IN MITAD OF REMEDY NOT WORKING, DISMONTATED AND SUBSTITUTEED BY SONDA AND GASKET, ADJUSTED THROUGH TECHNICAL SERVICE AND SEE EFFECT OPERATING TESTS. CALIBRATED AND VERIFIED.
No | Visit | Complaint | RCHN | 光学敏感性超范围1029 | 光学敏感性超范围1029 ，借903-364，944-212（13914），944-214（13886），944-258（13918）各一张。 | 更换903-364，944-212（13914），944-214（13886），944-258（13918）备件后修复by tim | SWO-318699 | Defect | HW | Process Module | Optical Unit | 1029 | Optical sensitivity over range 1029 | Optical sensitivity over range 1029 , borrow 903-364, 944-212 (13914), 944-214 (13886), 944-258 (13918) each. | Repair by tim after replacement of 903-364, 944-212 (13914), 944-214 (13886), 944-258 (13918) spare parts
Yes | Visit | Complaint | RCHN | 光学强度超范围 1031 | 光学强度超范围 1031 ，借903-364一个。 | 更换903-364一个正常。by tim | SWO-319481 | [5 empty] | Optical strength over range 1031 | Optical strength over range 1031 , borrow edith of 903-364. | Replace 903-364 one normal. by tim
Yes | Visit | Complaint | RCHN | 报错1031 | 报错1031，借903-364R8111N001一个 | 更换903-364后修复 by jesson | SWO-319401 | [5 empty] | 1031 | 1031, borrow 903-364R8111N001 one | Repair by jesson after replacing 903-364
No | Visit | Complaint | RCHN | 报错5297 | 报错5297，借903-371一个 | 更换903-371后修复by tim | SWO-318655 | Defect | HW | Incubator Module | Punch Up Unit | 5297 | 5297 error | 5297, loan 903-371 one | Repair by tim after replacement 903-371
No | Visit | Complaint | RCHN | 报错5784 | 报错5784，借903-127一个 | 更换903-127备件后修复by tim | SWO-319412 | Defect | HW | Storage Module | PCB, Storage Module | 5784 | 5784 | 5784, loan 903-127 one | Repair by tim after replacement of 903-127 spare parts
No | Visit | Complaint | RCHN | BNP不出值 | BNP不出值，借636-561，903-210各一个。 | 重新拆机后装上。故障现象消失。备件好还。by tim | SWO-319496 | Defect | HW | Process Module | PCB, Cup dryer |  | BNP not value | BNP is not worth, borrow 636-561,903-210 each. | Re-assembled and installed. The fault phenomenon disappears. Spare parts are good also. by tim
No | Visit | Complaint | RKK | プリンター印字（ー：マイナス）がかすれる | プリンター印字（ー：マイナス）がかすれる | プリンターユニット交換 903-594：Printer Assembly,ABL9 3 $9時点で部品登録無しの為、入力不可 | SWO-320159 | [5 empty] | Printer printing (-: minus) is blurred | Printer printing (-: minus) is blurred | Printer unit replacement 903-594: Printer Assembly, ABL9 3 No input because there is no part registration at $9
No | Visit | Complaint | RIND | Oximetry hardware error. | Checked with complete oximetry test and found all units failed. Then checked flow path. | Switched off analyzer and recoonceted and oxi unit. Then restarted analyzer and given complete oximetry test. Unit is working fine. | SWO-320353 | Defect | HW | Oxi module | off/on | Oxi HW error?? | Oximetry hardware error. | Checked with complete oximetry test and found all units failed. Then checked flow path. | Switched off analyzer and recoonceted and oxi unit. Then restarted analyzer and given complete oximetry test. Unit is working fine.
No | Visit | Complaint | RBNL | Autocheck | Les ampoules de l’autocheck ne sont pas tous scannés. | Nettoyage du mirroir du scanner (intern) Remis en route et contrôle : OK | SWO-320394 | Cleaned | HW | Autocheck Module | Barcode reader, Autocheck Module |  | Autocheck | Not all autocheck bulbs are scanned. | Cleaning the scanner mirror (intern) Get back on track and control: OK
No | Visit | Complaint | RDEU | Bildschirm reagiert verzögert. Alle Aktivitäten sind verlangsamt | Monitor wurde vor 2 Monaten getauscht. | CPU Einheit getauscht. Ref. Elektrode membraniert. PH gereinigt. Neustart. | SWO-320344 | Revisit required | HW | DMS_PC | CPU |  | Screen responds delayed. All activities are slowed down | Monitor was replaced 2 months ago. | CPU unit. Ref. electrode membraneed. PH cleaned. Restart.
Yes | Visit | Complaint | RAME | Customer concerned there was a shift in PCO2 starting Jan 14th a. 8:58 am | Customer is concerned about Calcium shift from .55 to .65 to .52. It is currently running at 0.52 on QC level 2 S7745 insert range is 0.43-0.63, customers range is 0.54-0.56. Talked with Chuck Rall and he authorized replacement of calcium and Reference electrodes. | Replaced reference and Calcuim electrodes with warranty electrodes. New reference electrode lot is 98-04, membrane lot R2214, new calcium electrode lot is 98-01 and membrane lot is R0748. Also did WDC training with Walter and Elaine. | SWO-319955 | Re classified RMED | Courtesy |  |  |  | Customer concerned there was a shift in PCO2 starting Jan 14th a. 8:58 am | Customer is concerned about Calcium shift from .55 to .65 to .52. It is currently running at 0.52 on QC level 2 S7745 insert range is 0.43-0.63, customers range is 0.54-0.56. Talked with Chuck Rall and he authorized replacement of calcium and Reference electrodes. | Replaced reference and Calcuim electrodes with warranty electrodes. New reference electrode lot is 98-04, membrane lot R2214, new calcium electrode lot is 98-01 and membrane lot is R0748. Also did WDC training with Walter and Elaine.
No | Visit | Complaint | RDEU | Kein Remote-Zugriff seit dem SW-Update mehr | siehe SWO-313919 vom 26.2. Fehlersuche --> Einstellungen überprüft --> Firewall für Public Network aktiv | Einstellungen (Firewall deaktiviert) geändert, System neu gestartet, System und Funktion i.O. | SWO-320171 | Setup | SW | Windows | Firewall |  | No more remote access since the SW update | see SWO-313919 of 26.2. Troubleshooting --> Settings checked --> Firewall for Public Network active | Settings (firewall disabled) changed, system restarted, system and function i.O.
No | Visit | Complaint | RAME | Auto check Motor is out will allow it to do automated controls | Auto check carousel will not move to the up position. I inspected all the cables including the 14 pole and all was good. | Replaced auto check stepping motor. | SWO-320502 | Defect | HW | Autocheck Module | Stepper Motor, Autocheck |  | Auto check Motor is out will allow it to do automated controls | Auto check carousel will not move to the up position. I inspected all the cables including the 14 pole and all was good. | Replaced auto check stepping motor.
No | Visit | Complaint | RBNL | Barcodes van de testcassette worden niet gelezen. Rode lichtstraal is wel waarneembaar. | Bij het plaatsen van een testcassette zie je wel de rode lichtstraal, maar de barcode wordt niet ingelezen. Uitlijning nagekeken, deze is goed. Barcodereader eruitgehaald, blijkt dat het venster erg vies is met een duidelijke vlek. Het venster schoongemaakt. Hierna worden alle barcodes weer zoals het hoort ingelezen. | Barcode venster schoongemaakt. | SWO-320306 | Cleaned | HW | Storage Module | Barcode reader, storage wheel |  | Barcodes of the test cassette are not read. Red light beam is perceptible. | When inserting a test cassette you can see the red beam, but the barcode is not read. Alignment checked, this is good. Barcode reader, it turns out that the window is very dirty with a clear stain. Cleaned the window. After this, all barcodes will be read as it should be. | Barcode window cleaned.
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | na lac pco2定标失败 | 更换测试卡后修复 | SWO-320787 | [5 empty] | Test card, reagent package failure | na lac pco2 calibration failed | Repair after replacing test card
No | Visit | Complaint | RCHN | 机器无法开机 | 机器无法开机，需要借件910-343 | 更换910-343一个正常。by tim | SWO-320163 | Defect | HW | DMS_PC | CPU |  | The machine does not power on | The machine does not power on, need to borrow 910-343 | Replace 910-343 one normal. by tim
No | Visit | Complaint | RCHN | 屏幕闪 | 屏幕闪，借910-298一个 | 更换910-298一个正常。by tim | SWO-319359 | Defect | HW | Analyzer Control Module | Display |  | Screen flash | Screen flash, borrow 910-298 one | Replace 910-298 one normal. by tim
No | Visit | Complaint | RAME | After the analyzer was setup to connect to the time server, the Daylight saving change did not work. | The time server connection did not update the time after the daylight saving change went in to effect this morning. | Checked and, confirmed the problem. The customer tried updating the time manually but, the time reverted back to the old time. In the windows control panel, I checked the option to make the clock automatically adjust for daylight time saving. I also enabled the Automatic daylight saving service in the windows administrator tools. | SWO-320938 | Setup | SW | Time setting | Time setting | Time Server | After the analyzer was setup to connect to the time server, the Daylight saving change did not work. | The time server connection did not update the time after the daylight saving change went in to effect this morning. | Checked and, confirmed the problem. The customer tried updating the time manually but, the time reverted back to the old time. In the windows control panel, I checked the option to make the clock automatically adjust for daylight time saving. I also enabled the Automatic daylight saving service in the windows administrator tools.
No | Visit | Complaint | RAME | After connecting the analyzer to the hospital time server, the daylight savings function didn't work. | The daylight time saving function didn't update the time for daylight time saving that went into effect this morning. | Checked the analyzer and, confirmed the problem. The customer tried updating the time manually but, the analyzer would change the time back to the old time. I went in to the windows control panel and, setup the analyzer to automatically adjust the clock for daylight saving changes. I also setup the daylight saving service in the windows administrator services tool. | SWO-320939 | Setup | SW | Time setting | Time setting | Time Server | After connecting the analyzer to the hospital time server, the daylight savings function didn't work. | The daylight time saving function didn't update the time for daylight time saving that went into effect this morning. | Checked the analyzer and, confirmed the problem. The customer tried updating the time manually but, the analyzer would change the time back to the old time. I went in to the windows control panel and, setup the analyzer to automatically adjust the clock for daylight saving changes. I also setup the daylight saving service in the windows administrator services tool.
No | Visit | Complaint | RAME | After setting up the analyzer to work with the Hospital time server, the Daylight savings feature didn't not work. The time was not updated. | The analyzer has not updated the time after the daylight saving time change. | Checked the analyzer and, confirmed the problem. The customer manually updated the time but, the analyzer reverted the time back to the old time. Thru the control panel in windows, I enabled the analyzer to automatically adjust the clock for daylight savings changes. I also enabled the Daylight saving service thru the administrative tools. | SWO-320935 | Setup | SW | Time setting | Time setting | Time Server | After setting up the analyzer to work with the Hospital time server, the Daylight savings feature didn't not work. The time was not updated. | The analyzer has not updated the time after the daylight saving time change. | Checked the analyzer and, confirmed the problem. The customer manually updated the time but, the analyzer reverted the time back to the old time. Thru the control panel in windows, I enabled the analyzer to automatically adjust the clock for daylight savings changes. I also enabled the Daylight saving service thru the administrative tools.
No | Visit | Complaint | RAME | The analyzer failed to update to the new daylight saving time after it was connected to the time server. | The analyzer daylight time change function didn't update to the new time change. | Checked the analyzer and, confirmed the problem. The customer has tried updating the time manually but, the analyzer keeps going back to the old time. In windows, I setup the analyzer so the clock is automatically adjusted for time saving changes. I also enabled the Daylight saving service in the administrator tool. | SWO-320936 | Setup | SW | Time setting | Time setting | Time Server | The analyzer failed to update to the new daylight saving time after it was connected to the time server. | The analyzer daylight time change function didn't update to the new time change. | Checked the analyzer and, confirmed the problem. The customer has tried updating the time manually but, the analyzer keeps going back to the old time. In windows, I setup the analyzer so the clock is automatically adjusted for time saving changes. I also enabled the Daylight saving service in the administrator tool.
No | Visit | Complaint | RAME | After setting up the analyzer to connect with the hospital time server, I received a call from Efraim letting me know the Daylight saving time function did not work and, it didn't update the time. | The Daylight savings time change did not update the new time after the midnight daylight saving time changed today. | I went to the control panel and, in the time and, date settings I checked the box so that the analyzer updates the clock automatically for Daylight savings changes. I also enabled the Daylight savings service. | SWO-320933 | Setup | SW | Time setting | Time setting | Time Server | After setting up the analyzer to connect with the hospital time server, I received a call from Efraim letting me know the Daylight saving time function did not work and, it didn't update the time. | The Daylight savings time change did not update the new time after the midnight daylight saving time changed today. | I went to the control panel and, in the time and, date settings I checked the box so that the analyzer updates the clock automatically for Daylight savings changes. I also enabled the Daylight savings service.
No | Visit | Complaint | RCHN | 卡盘,HCT传感器失灵,处理盘转动不灵活 | 1，卡盘，机器原存储盘黄铜卡嘴处有某些卡转过去偶尔档一下，无法彻底解决，将备用机上存储盘上部分整体给更换了； 2，HCT上的传感器失灵了（新机调试时）； 3，处理盘转动不正常，电极有漏液迹象，会偶尔报错。 | 1，将备用机上存储盘上部分整体给更换了； 2，更换HCT903-329（新机调试时）； 3，更换电极580-098 | SWO-320709 | Defect | HW | Aspiration Module | Needle Wash Unit | HCT Sensor | Card, HCT sensor failure, processing disk rotation is not flexible | 1, card disk, the machine original storage disk brass card mouth has some cards turnpast occasionally, can not completely solve, the spare machine storage disk on the whole to replace; 2, the sensor on the HCT is not working (when the new machine is debugging); 3, the processing disk rotation is not normal, the electrode has a leak, will occasionally report errors. | 1, the spare machine storage disk on the part of the whole to replace; 2, replace HCT903-329 (when the new machine is debugging); 3, replace electrode 580-098
No | Visit | Complaint | RCHN | 血氧硬件故障 | 血氧硬件故障，借902-833一个 | 灯单元、电路板重新拆了再安装后修复（by chuanwei | SWO-319389 | Defect | HW | Oxi module | Lamp unit |  | Blood oxygen hardware failure | Blood oxygen hardware failure, borrow 902-833 one | Lamp unit, board re-installed and re-installed after repair (by chuanwei
No | Visit | Complaint | RAME | The daylight saving time function didn't update to the new daylight saving time change that went into effect yesterday morning. | The daylight saving time function didn't update to the new daylight saving time change that went into effect yesterday morning. | Checked the analyzer time and, confirmed it was not updated. I setup the analyzer clock to adjust for daylight saving time changes. I also enabled the Daylight saving service. I now confirmed the analyzer has updated to the new daylight saving time change. | SWO-321651 | Setup | SW | Time setting | Time setting | Time Server | The daylight saving time function didn't update to the new daylight saving time change that went into effect yesterday morning. | The daylight saving time function didn't update to the new daylight saving time change that went into effect yesterday morning. | Checked the analyzer time and, confirmed it was not updated. I setup the analyzer clock to adjust for daylight saving time changes. I also enabled the Daylight saving service. I now confirmed the analyzer has updated to the new daylight saving time change.
No | Visit | Complaint | RAME | The analyzer daylight saving time function didn't update to the new time change. | The analyzer daylight time function didn't update the clock to the new time change. | Checked the analyzer time and, confirmed the problem. I setup the analyzer to automatically adjust the clock for daylight savings time change. I also enabled the daylight savings service. | SWO-321656 | Setup | SW | Time setting | Time setting | Time Server | The analyzer daylight saving time function didn't update to the new time change. | The analyzer daylight time function didn't update the clock to the new time change. | Checked the analyzer time and, confirmed the problem. I setup the analyzer to automatically adjust the clock for daylight savings time change. I also enabled the daylight savings service.
No | Visit | Complaint | RAME | The daylight savings function didn't work as expected. The new time was not updated. | The new daylight saving time was not updated by the analyzer. | Checked analyzer time and, confirmed the old time is still on. I setup the analyzer clock to automatically adjust for daylight saving time changes. I also enabled the Daylight saving service. I confirmed the time has now been updated to the new daylight saving time. | SWO-321657 | Setup | SW | Time setting | Time setting | Time Server | The daylight savings function didn't work as expected. The new time was not updated. | The new daylight saving time was not updated by the analyzer. | Checked analyzer time and, confirmed the old time is still on. I setup the analyzer clock to automatically adjust for daylight saving time changes. I also enabled the Daylight saving service. I confirmed the time has now been updated to the new daylight saving time.
No | Visit | Complaint | RAME | The daylight saving time change didn't not work on the analyzer. Keeps going back to the old time if updated manually. | The analyzer clock didn't not update with the daylight time savings change. | Checked the analyzer clock and, confirmed it has not been automatically updated. I setup the analyzer clock to automatically adjust for daylight saving time change. I also enabled the Automatic daylight service. I confirmed the analyzer has now updated with the new daylight saving time change. | SWO-321658 | Setup | SW | Time setting | Time setting | Time Server | The daylight saving time change didn't not work on the analyzer. Keeps going back to the old time if updated manually. | The analyzer clock didn't not update with the daylight time savings change. | Checked the analyzer clock and, confirmed it has not been automatically updated. I setup the analyzer clock to automatically adjust for daylight saving time change. I also enabled the Automatic daylight service. I confirmed the analyzer has now updated with the new daylight saving time change.
No | Visit | Complaint | RAME | After the connecting the analyzer with the time server, the daylight saving time feature didn't update to the new time. | The analyzer didn't update the new daylight saving time even after it was connected to a time server. | Checked the Ana and, confirmed the problem. I setup the analyzer to automatically adjust the clock for daylight saving time change. I also enabled the Daylight saving service. Found analyzer failing the vacuum test, replaced a new inlet gasket holder from the yearly service kit. | SWO-321652 | Setup | SW | Time setting | Time setting | Time Server | After the connecting the analyzer with the time server, the daylight saving time feature didn't update to the new time. | The analyzer didn't update the new daylight saving time even after it was connected to a time server. | Checked the Ana and, confirmed the problem. I setup the analyzer to automatically adjust the clock for daylight saving time change. I also enabled the Daylight saving service. Found analyzer failing the vacuum test, replaced a new inlet gasket holder from the yearly service kit.
No | Visit | Complaint | RAME | The analyzer daylight savings time function didn't update the new time change. | The analyzer daylight saving time function didn't update to the new time change. | Checked the analyzer time and, confirmed the problem. I now setup the to automatically adjust the clock for daylight saving time. I also enabled the Daylight saving service. | SWO-321653 | Setup | SW | Time setting | Time setting | Time Server | The analyzer daylight savings time function didn't update the new time change. | The analyzer daylight saving time function didn't update to the new time change. | Checked the analyzer time and, confirmed the problem. I now setup the to automatically adjust the clock for daylight saving time. I also enabled the Daylight saving service.
No | Visit | Complaint | RAME | The analyzer daylight saving time function didn't work as expected and, didn't change to the new daylight saving time change. | Daylight saving time function didn't update to the new daylight saving time change. | Checked the analyzer time and, confirmed the clock has not updated to the new daylight saving time change. I setup the analyzer clock to automatically adjust for daylight saving time. I also enabled the Daylight saving service. | SWO-321655 | Setup | SW | Time setting | Time setting | Time Server | The analyzer daylight saving time function didn't work as expected and, didn't change to the new daylight saving time change. | Daylight saving time function didn't update to the new daylight saving time change. | Checked the analyzer time and, confirmed the clock has not updated to the new daylight saving time change. I setup the analyzer clock to automatically adjust for daylight saving time. I also enabled the Daylight saving service.
No | Visit | Complaint | RAME | Daylight saving time function didn't update the new daylight saving time. | The daylight saving time function didn't update to the new daylight saving time change. | Checked and, confirmed the problem. I setup the analyzer to adjust the clock for daylight saving time change. I also enabled the Daylight saving time service. I now confirmed the clock has updated to the new time change. | SWO-321654 | Setup | SW | Time setting | Time setting | Time Server | Daylight saving time function didn't update the new daylight saving time. | The daylight saving time function didn't update to the new daylight saving time change. | Checked and, confirmed the problem. I setup the analyzer to adjust the clock for daylight saving time change. I also enabled the Daylight saving time service. I now confirmed the clock has updated to the new time change.
No | Visit | Complaint | RDEU | 23.02.2018, 10:15 Uhr ==> Anrufer Fr. Bernhardt, Tel.: 06213831157 ==> Oxi-Hardwarefehler, Neonkorrektur ausserhalb der Grenzen | QC-Fehler Level B und C in sO2, COHb, HbF und HHb | Oximetrie Test ohne Beanstandung - Analysator heruntergefahren und demontiert - Erneuerung Hemolysator-Einheit - tHb KAL - Funktionsüberprüfung etc. | SWO-321009 | Defect | HW | Oxi module | Hemolyzer |  | 23.02.2018, 10:15 a.m. ==> Caller Fr. Bernhardt, Tel.: 06213831157 ==> Oxi hardware error, neon correction outside the limits | QC-Fehler Level B und C in sO2, COHb, HbF und HHb | Oximetry test without complaint - Analyzer shut down and dismantled - Renewal Hemolysator unit - tHb KAL - Function check etc.
Yes | Visit | Complaint | RDEU | 02.03.2018; Hr. Schreiber; Tel.: 0343218183; weißer Monitor | Kundengerät nach Reparatur wieder aufgestellt | Repariertes Gerät nach Testlauf in VSC wieder beim Kunden aufgestellt Netzwerkverbindung überprüft Startphase i.O. Rep-Überbrückungsgerät wieder zurück in VSC | SWO-321088 | Re classified RMED | RE install Analyzer | See SWO-316839 |  |  | 02.03.2018; Mr. Schreiber; Tel.: 0343218183; white monitor | Customer device repositioned after repair | Repaired device repositioned at the customer after test run in VSC Network connection verified Start-up phase i.O. Rep bridging device back in VSC
No | 3 | Y | RPOL | Blank | Wymiana Sensor Interface PCB. Analizator sprany technicznie. |  | D4044A4736447F8B | Defect | HW | Basic Unit | PCB, Sensor interface | SC Door | Blank | Replacement Sensor Interface PCB. The analyser is technically washed.
No | Visit | Complaint | RAME | Customer states that the inlet intermittently will not not aspirate sample when presented to the sample probe. They need to remove and re-introduce the sample to aspirate. | Unable to reproduce the reported issue. | Removed and cleaned the inlet gasket holder, tested via service diagnostics, no issue observed. Demonstrated the successful result of diagnostics to customer. reminded customer to push sample completely onto the inlet gasket holder and probe to ensure activation. Will monitor. | SWO-321506 | No fault found | HW | Auto Inlet module | Auto Inlet module | HALL Element | Customer states that the inlet intermittently will not not aspirate sample when presented to the sample probe. They need to remove and re-introduce the sample to aspirate. | Unable to reproduce the reported issue. | Removed and cleaned the inlet gasket holder, tested via service diagnostics, no issue observed. Demonstrated the successful result of diagnostics to customer. reminded customer to push sample completely onto the inlet gasket holder and probe to ensure activation. Will monitor.
No | Visit | Complaint | RAME | Touch display misaligned. | Confirmed touch display is not aligned. | Ran service alignment procedure, verified operation, problem resolved. | SWO-321505 | Adjustment | HW | DMS_PC | Touch screen |  | Touch display misaligned. | Confirmed touch display is not aligned. | Ran service alignment procedure, verified operation, problem resolved.
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 参比电极错误，1361 | 更换测试卡后修复 | SWO-320805 | [5 empty] | Test card, reagent package failure | Reference electrode error, 1361 | Repair after replacing test card
No | Visit | Complaint | RCHN | 无法开机 | 无法开机，借910-342，910-285各一个。 | 更换910-342，910-285后修复 (by chuanwei) | SWO-320619 | Defect | HW | Power Supply | Power Supply |  | Cannot power on | Cannot be switched on, borrow 910-342, 910-285 each. | Replacement 910-342, 910-285 post-repair (by chuanwei)
No | Visit | Complaint | RCHN | 定标卡扫描失败 | 定标卡扫描失败，借902-924一个。 | 更换902-924后修复(by chuanwei) | SWO-320669 | Defect | HW | Cover and Chassis | Barcode reader, display |  | Bid card scan failed | The bid card scan failed, borrowing 902-924 one. | Repair after replacement 902-924 (by chuanwei)
No | 18 | Y | RCZE | Blank |  |  | D4088E7098565841 | Unknown | Unknown | Unknown | Unknown | Contact RCZE for further info. | Blank
No | Visit | Complaint | RIND | Frequent calibration and sample error. | found problem with inlet gasket. | cleaned the inlet gasket and did inlet, LS ,Pump and 2point calibration ,all results found within the range. | SWO-324037 | Cleaned | Fluidic pathway | Inlet module | Inlet gasket |  | Frequent calibration and sample error. | found problem with inlet gasket. | cleaned the inlet gasket and did inlet, LS ,Pump and 2point calibration ,all results found within the range.
No | Visit | Complaint | RIND | Thb/so2 values ?? | thb $so2 values problem | checked the analyser thb $so2 values question mark for samples so oxmetry calibartion was done and pump calibartion was done then protein removal aspirated for 3-4 times then calibarted to raedy mode analysed the sample observed for few samples | SWO-323669 | Cleaned | HW | Oxi module | Hemolyzer |  | Thb/so2 values ?? | thb /so2 values problem | checked the analyser thb $so2 values question mark for samples so oxmetry calibartion was done and pump calibartion was done then protein removal aspirated for 3-4 times then calibarted to raedy mode analysed the sample observed for few samples
No | 2 | Y | RCZE | Blank |  |  | D4088DEA9813583E | Defect | HW | DMS_PC | Touch screen | Contact RCZE for further info. | Blank
No | Visit | Complaint | RDEU | Gerät geht in Fehler; Fehler_3899: Lock 24 | Fehlersuche --> Dateiarchive (Aktivitäten, Pat.-Berichte), Motoren, Sensoren, Antriebe überprüft --> Verriegelung für Testkassetten funktioniert nicht mehr --> Spindel (Gewindestange) hat sich von der Verriegelung selbst heraus gedreht | Spindel wieder an der Verriegelung befestigt, Systemcheck im Service, Systemüberprüfung, Systemreinigungen und QCs durchgeführt, System und Funktion i.O. | SWO-321007 | Defect | HW | Storage Module | Wheel Lock | 3899 | Device goes into error; Error_3899: Lock 24 | Troubleshooting --> File archives (activities, Pat.-Reports), motors, sensors, drives checked --> Locking for test cassettes no longer works --> Spindle (thread rod) has rotated itself from the locking itself | Spindle again attached to the locking, system check in service, system check, system cleaning and QCs performed, system and function i.O.
No | Visit | Complaint | RFRA | Mesures annulées Échantillons insuffisants Tubes non maintenus | Porte tubes cassé | Remplacer porte tube Contrôle aiguille prélèvements | SWO-321320 | Defect | HW | Aspiration Module | Tube holder | LEH L3 ? | Measures cancelled Insufficient samples Unmaintained tubes | Broken tube door | Replace tube door Needle control samples
No | Visit | Complaint | RKOR | Solution pack replacement message appears repeatedly. | Analyzer Status check Activity logs check Fluid transport check Peplace Inlet gesket holder, Inlet probe Re-Install connector gasket, Solution pack | Re-Install Solution pack Peplace Inlet gesket holder, Inlet probe | SWO-324060 | Defect | HW | Inlet module | Inlet gasket |  | Solution pack replacement message appears repeatedly. | Analyzer Status check Activity logs check Fluid transport check Peplace Inlet gesket holder, Inlet probe Re-Install connector gasket, Solution pack | Re-Install Solution pack Peplace Inlet gesket holder, Inlet probe
No | Visit | Complaint | RAME | CA, Na are out on calibrations for sensitivity over 102% | Replace inlet gasket, replace SP,SC. 0964 leak current detected errors. Sensors reading below 200 mv. | Customer replaced the SC. FSE inspected all sensors and the were reading < 200mv.Flushed system.Replaced waste tubing and inspected spring on hemolyzer. Customer bought a stand and battery so I installed the battery and stand for customer. | SWO-321922 | Assisted Replacement | SC | SC | Leak Current | 0964 | CA, Na are out on calibrations for sensitivity over 102% | Replace inlet gasket, replace SP,SC. 0964 leak current detected errors. Sensors reading below 200 mv. | Customer replaced the SC. FSE inspected all sensors and the were reading < 200mv. Flushed system. Replaced waste tubing and inspected spring on hemolyzer. Customer bought a stand and battery so I installed the battery and stand for customer.
No | Visit | Complaint | RDEU | 12.03.2018, 09:23 Uhr, Fr. Schaudt, Tel. 0754253155340: ABL verlangt Spülung und LP wechseln 12.03./11:05 Uhr, Frau Schaudt, Tel.: 0754253155340: ABL wieder in Intervention erforderlich | Gerät überprüft steht bei Eintreffen. SP installiert und im Service Programm überprüft.Vacuum Test nicht bestanden | Flüss.wege gespült, waren ok. Einlassdichtung ersetzt, weil leicht oval. Jetzt ok. Vacuumfehler bleibt. SC90 eingesetzt. Startphase abgewartet. | SWO-321904 | Revisit required | Revisit required | Revisit required | Revisit required | 1063 | 12.03.2018, 09:23, Ms. Schaudt, Tel. 0754253155340: ABL requires flushing and CHANGING LP 12.03./11:05, Mrs Schaudt, Tel.: 0754253155340: ABL required to intervene again | Device is checked on arrival. SP installed and checked in the service program. Vacuum test failed | Flüss.wege rinsed, were ok. Inlet seal replaced because slightly oval. Now ok. Vacuum error remains. SC90. start-up phase.
No | 4 | Y | RPAC | Blank | OXI failed on calibration |  | D406389CAC14715C | Defect | HW | Oxi module | Oxi module |  | Blank | OXI failed on calibration
No | Visit | Complaint | RAME | Unable to scan barcodes | Observed front barcode scanner, illuminated, beeps when presented with a barcode but will not transmit code information; QC, user login, etc. | Replaced barcode scanner, did not resolve the issue. Suspect possible issue with either display interface or CPU, RS232 comm for barcode scanner. Will need to order parts. | SWO-321507 | Revisit required | HW | Basic Unit | Barcode reader |  | Unable to scan barcodes | Observed front barcode scanner, illuminated, beeps when presented with a barcode but will not transmit code information; QC, user login, etc. | Replaced barcode scanner, did not resolve the issue. Suspect possible issue with either display interface or CPU, RS232 comm for barcode scanner. Will need to order parts.
No | Visit | Complaint | RCHN | 测试卡初始化失败，接口板腐蚀 | 测试卡初始化失败，接口板腐蚀 | 干燥处理接口板，重新安装后修复 by jesson | SWO-322178 | [5 empty] | Test card initialization failed, interface board corrosion | Test card initialization failed, interface board corrosion | Drying interface board to repair by jesson after reinstallation
No | Visit | Complaint | RCHN | GLU,LAC定标不过 | GLU,LAC定标不过，借903-183一个（被903-581替代） | 更换903-581后修复(By Ran) | SWO-320902 | Defect | HW | Basic Unit | PCB, Sensor interface | Glu/Lac | GLU, LAC calibration but | GLU, LAC calibration, borrow edgy 903-183 (replaced by 903-581) | Repair after replacement 903-581 (By Ran)
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 参比电极错误，1361 | 更换测试卡后修复 | SWO-324827 | [5 empty] | Test card, reagent package failure | Reference electrode error, 1361 | Repair after replacing test card
No | Visit | Complaint | RESP | Fugas | Limpieza | Limpieza y cambio de sonda y entrada. Pruebas de funcionamiento | SWO-321826 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | Leaks | Cleaning | Cleaning and changing probe and entry. Operational tests
No | Visit | Complaint | RCHN | 报错1999 | 报错1999 | 更换处理包后修复 | SWO-321491 | Assisted Replacement | Consumables | SP | SP | 1999 | Error 1999 | Error 1999 | Repair after replacing the processing package
No | Visit | Complaint | RCHN | 定标失败 1061，5257 | ，借636-550一个。 | 重新定标后修复(By Ran) Changed 636-550.Recal , then OK. | SWO-320949 | Defect | HW | Process Module | Sensor, Process module |  | Bid failure 1061,5257 | , borrow 636-550 one. | Re-calibration repair (By Ran) Changed 636-550.Recal , then OK.
Yes | 16 | Y | RPAC | Blank | Supply of orings |  | D405972FAC0A587C | [5 empty] | Blank | Supply of orings
No | 8 | Y | RCZE | Blank | 876 flexQ positioniong warning |  | D405C27AAC068A67 | Adjustment | HW | FlexQ Module | FlexQ Module |  | Blank | 876 flexQ positioniong warning
No | 9 | Y | RCZE | Blank | neprovádí kontroly |  | D405BC11AC59D257 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | Blank | does not perform checks
No | Visit | Complaint | RKK | pHｷｬﾙｴﾗｰ。感度極端に低下。2017年11月に納品分につき電極新品交換保証。 | pHｷｬﾙｴﾗｰ。感度極端に低下。20% | 2017年11月に納品分につき電極新品交換保証。 | SWO-323378 | Assisted Replacement | Electrode | Electrode | pH | Contact RKK for return of failing elec. | pH cal error. Sensitivity decreases to extremes. In November 2017, we guaranteed a new replacement of electrodes for delivery. | pH cal error. Sensitivity decreases to extremes. 20% | In November 2017, we guaranteed a new replacement of electrodes for delivery.
No | Visit | Complaint | RIND | Oximetry values is not coming in patient reports. | Check and found that is tHb and so2 is not showing sample reports. | check all the tubing and membrane. Run protein removable program, flush waste flow path of the machine. But problem is still same than run complete oximetry test and found Lamp unit is failed. Than adjust the lamp unit. Run 2 point calibration successfully. Now machine is working fine. | SWO-324004 | Adjustment | HW | Oxi module | Lamp unit |  | Oximetry values is not coming in patient reports. | Check and found that is tHb and so2 is not showing sample reports. | check all the tubing and membrane. Run protein removable program, flush waste flow path of the machine. But problem is still same than run complete oximetry test and found Lamp unit is failed. Than adjust the lamp unit. Run 2 point calibration successfully. Now machine is working fine.
No | Visit | Complaint | RFRA | Faux positif sur mesure TnI | Vérification état analyseur | Nettoyage circuits liquides Récupérer pack de solutions pour analyse | SWO-321959 | Revisit required | HW | Wet section control | Wash Unit | LEH | Fake bespoke tuning TnI | State analyzer check | Cleaning liquid circuits Recover pack of solutions for analysis
No | Visit | Complaint | RAME | Leak detected | Leak detected also Pco2 sensitivity low | Replace hemolyzer valve tubing $ Replace the sensor cassette | SWO-323284 | Defect | HW | Oxi module | Pinch valve tube |  | Leak detected | Leak detected also Pco2 sensitivity low | Replace hemolyzer valve tubing $ Replace the sensor cassette
No | Visit | Complaint | RAME | Leak detected | Leak detected error $Pco2 sensitivity low | Replace hemolyzer valve tubing $Replace sensor cassette | SWO-322755 | Defect | HW | Oxi module | Pinch valve tube |  | Leak detected | Leak detected error /Pco2 sensitivity low | Replace hemolyzer valve tubing $Replace sensor cassette
No | Visit | Complaint | RLTD | Unable to pump solutions. | Unable to pump solutions | Customer had already changed solution pack, flushed analyser and carried out trouble shooting. I was unable to find fault with the analyser, replaced solution pack, now ok. | SWO-322617 | No fault found | No fault found | No fault found | No fault found |  | Unable to pump solutions. | Unable to pump solutions | Customer had already changed solution pack, flushed analyser and carried out trouble shooting. I was unable to find fault with the analyser, replaced solution pack, now ok.
No | Visit | Complaint | RDEU | 14.03.2018; Sr. Anette; Tel.: 0390130536; Mögliche Verstopfung der Fließwege erkannt 14.03.2018 11:30 Sr. Daniela 0390130536: Verstopfung bleibt bestehen | LS Sensor Fehler | Fehlersuche, Austausch Flüssigkeitssensor, Funktionskontrolle | SWO-322499 | [5 empty] | 14.03.2018; Sr. Anette; Tel.: 0390130536; Possible blockage of the flow paths detected 14.03.2018 11:30 Sr. Daniela 0390130536: Constipation remains | LS Sensor Error | Troubleshooting, replacement of liquid sensor, function control
No | Visit | Complaint | RCHN | 样本量不足，报错1067 | 样本量不足，报错1067，需要借件903-329 | 更换903-329备件后修复(by chuanwei) | SWO-322086 | Defect | HW | Aspiration Module | Needle Wash Unit | 1067 | Insufficient sample size, error 1067 | Insufficient sample size, error 1067, need to borrow 903-329 | Repair after replacement of 903-329 spare parts (by chuanwei)
No | 32 | Y | RCZE | Blank | nevychází QC na Lac i po opakované výmìnì membrány, provedeno èištìní Lac elektrody - beze zmìny |  | D406F5562244FA84 | Assisted Replacement | Electrode | Electrode | Lac |  | Blank | The QC to the LAC after repeated replacement of the membrane, the cleaning of the Lac electrode-without change
No | Visit | Complaint | RIND | Ca++ sensitivity low & status high. | Check the calibration log & found that the Status of Ca++ is high & sensitivity is low. | Replace the Ca++ membrane with the new one,run the 2 point calibration but problem remains same,replace the reference membrane & run the 2 point calibration but still status was high for Ca++,need to replace the reference Electrode for further check. | SWO-322837 | Assisted Replacement | Electrode | Membrane | Ca++ | Ref. electrode possibly wrongly exchanged | Ca++ sensitivity low & status high. | Check the calibration log & found that the Status of Ca++ is high & sensitivity is low. | Replace the Ca++ membrane with the new one,run the 2 point calibration but problem remains same,replace the reference membrane & run the 2 point calibration but still status was high for Ca++,need to replace the reference Electrode for further check.
No | Visit | Complaint | RAME | Inlet leaking | Inlet coule | Débloqué ABL remplacé inlet holder | SWO-322703 | Defect | HW | Inlet module | Probeholder |  | Inlet leaking | Inlet sinks | Unlocked ABL replaced inlet holder
No | Visit | Complaint | RDEU | 13.03.2018 Frau Demuth 06215034079 meldet, dass der Einlass undicht ist. 14.03.2018, 08:18 Uhr, Fr. Demuth, Tel. 062150335589: Kd. hat alle erforderlichen Maßnahmen zur Behebung des Fehlers durchgeführt - ohne Erfolg | Einlass undicht (Einlassdichtung bereits erneuert) | Überprüfung etc. - Abfallstrecke gespült - nochmalige Erneuerung Einlassdichtung - Funktionsüberprüfung etc. ohne Beanstandung | SWO-322318 | Defect | HW | Inlet module | Inlet gasket |  | 13.03.2018 Ms. Demuth 06215034079 reports that the entrance is leaking. 14.03.2018, 08:18, Ms. Demuth, tel. 062150335589: Kd. has taken all necessary measures to correct the error - without success | Inlet leaks (inlet seal already renewed) | verification, etc. - Waste section rinsed - renewed inlet seal - function check etc. without complaint
No | Visit | Complaint | RDEU | 08.03.2018; Fr. Langewellpott; Tel.: 0611577388; Meßabbrüche 12.03.2018; Fr. Langewellpott; Tel.: 0611577470; wiederholt Messabbrüche 12.03.2018 Fr. Langewellpott 0611577470 wiederholt Messabbrüche 12.03.2018; 13:50 Uhr; Fr. Langewellpott; Tel.: noch mehr Fehlermeldungen | Oberen Deckel abgenommen, Am Fluid Modul sind 2 Schläuche mit dem Deckel abgequetscht gewesen, beide Schläuche in richtige Position gerichtet. | Schlauch Nr 3, und Oxi Pumpen Schlauch neu installiert, Pumpen Kal, LS Justage etc. Funktionsprüfung i.O, QCs i.O | SWO-322393 | Defect | HW | Basic Unit | Tubes | Squeezed | 08.03.2018; Ms. Langewellpott; Tel.: 061157738; Measurement interruptions 12.03.2018; Ms. Langewellpott; Tel.: 0611577470; repeated measurement interruptions 12.03.2018 Ms. Langewellpott 0611577470 repeated measurement interruptions 12.03.2018; 1:50 p.m.; Ms. Langewellpott; Tel.: even more error messages | Upper lid removed, On the fluid module 2 hoses have been squeezed with the lid, both hoses directed in correct position. | Tube No. 3, and Oxi pumps hose newly installed, pumps Kal, LS Adjustment etc. Functional test i.O, QCs i.O
No | Visit | Complaint | RIND | Sample are getting aborted. | Cleaned the all the sensor in pipette unit related. | checked the all movement related to pipette unit. found all working fine. | SWO-324039 | Cleaned | HW | Aspiration Module | Pipette Unit |  | Sample are getting aborted. | Cleaned the all the sensor in pipette unit related. | checked the all movement related to pipette unit. found all working fine.
No | Visit | Complaint | RLTD | The Analyser is showing the error FTC_105 and is leaking at the back of the Analyser. |  | Removed wash block and cleared the blockage. Ran wet section tests and checks. | SWO-322858 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | The Analyser is showing the error FTC_105 and is leaking at the back of the Analyser. |  | Removed wash block and cleared the blockage. Ran wet section tests and checks.
No | Visit | Complaint | RKOR | Appeared Solution pack replacement message. | Re-Installation Solution Pack Analyzer Status check Activity log check Calibration, Quality Control check | Re-Installation Solution Pack | SWO-324068 | Assisted Replacement | SP | SP | RE Install |  | Appeared Solution pack replacement message. | Re-Installation Solution Pack Analyzer Status check Activity log check Calibration, Quality Control check | Re-Installation Solution Pack
No | Visit | Complaint | RAME | Battery 0% | 1) Battery 0% 2) DST needs updating | 1) Replaced battery and tested. 2) Configured RmedDstManagerService in Administrative Tools from Manual to Automatic. | SWO-323248 | Defect | HW | Battery module | Battery module | 0% | Battery 0% | 1) Battery 0% 2) DST needs updating | 1) Replaced battery and tested. 2) Configured RmedDstManagerService in Administrative Tools from Manual to Automatic.
No | Visit | Complaint | RFRA | L'analyseur se bloque en intervention nécessaire +le site reste ouvert parfois +les résultats ne sont pas cohérents | Analyse des fichiers activités Test en mode service | Installation d'un kit de maintenance annuelle $ $ vérification du site, de l'aiguille $ $ service dump et export des données $ $ essais: bons | SWO-323312 | Defect | HW | Auto Inlet module | Probe bend |  | The analyzer crashes in needed intervention The site stays open sometimes The results are not consistent | Activity file analysis Testing in service mode | Installing an annual maintenance kit $$ site check, needle $$ service dump and export data $$ trials: vouchers
No | Visit | Complaint | RAME | 581 errors | Receiving error code 581 and 292 | Replaced Hemolyzer assembly | SWO-323702 | [5 empty] | 581 errors | Receiving error code 581 and 292 | Replaced Hemolyzer assembly
No | Visit | Complaint | RIND | Results not ok or unstable, not exaptable. | Results not ok or unstable, not exaptable. | SC Flushed and Reinstalled, Calibrated properly. Trouble shooting done. Analyser Serviced and calibrated properly. Handed over in a fine working condition. | SWO-324002 | [5 empty] | Results not ok or unstable, not exaptable. | Results not ok or unstable, not exaptable. | SC Flushed and Reinstalled, Calibrated properly. Trouble shooting done. Analyser Serviced and calibrated properly. Handed over in a fine working condition.
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 检测到泵定标出错 1178 | 更换试剂包正常。 | SWO-323377 | [5 empty] | Test card, reagent package failure | Pump calibration error detected 1178 | The replacement kit is OK.
No | Visit | Complaint | RCHN | PO2定标不过，质控不过 | PO2定标不过，质控不过 借903-581一个 | 更换903-581备件后修复(by chuanwei) | SWO-322781 | Defect | HW | Basic Unit | PCB, Sensor interface |  | PO2 calibration is not, but the quality control is not | PO2 calibration but, but the quality control but borrow 903-581 one | Repair after replacement of 903-581 spare parts (by chuanwei)
No | Visit | Complaint | RCHN | LAC定标不过 | LAC定标不过 借903-581一个 | 更换903-581备件后修复(by chuanwei) | SWO-321601 | Defect | HW | Basic Unit | PCB, Sensor interface | LAC | LAC calibration is not | LAC calibration but borrow 903-581 one | Repair after replacement of 903-581 spare parts (by chuanwei)
No | Visit | Complaint | RESP | INLET SE QUEDA FUERA DE POSICION, NO ESTA EN POSICION CORRECTA. | SE DETECTA SONDA DOBLADA E INLET FUERA DE SU POSICION DE REPOSO | INLET SE QUEDA A MITAD DE RECORRIDO NO PUEDIENDO FUNCIONAR, SE DESMONTA Y SE LE SUSTITUYE SONDA Y GASKET, SE AJUSTA MEDIANTE SERVICIO TECNICO RECOLOCANDO MODULO INLET Y SE EFECTUAN PRUEBAS DE FUNCIONAMIENTO. CALIBRADO Y VERIFICADO. | SWO-323183 | Defect | HW | Auto Inlet module | Probe bend | LEH Mail | INLET IS OUT OF POSITION, NOT IN CORRECT POSITION. | DOUBLEAND SOUND DETECTED OUTSIDE YOUR REST POSITION | INLET IS HALF OF REMEDY NO BY WORKING, IT IS DISMOUNTED AND SUBSTITUTE SOUND AND GASKET, IS ADJUSTED BY TECHNICAL SERVICE REPOSITIONING INLET MODULE AND TESTS OF Operation. CALIBRATED AND VERIFIED.
No | Visit | Complaint | RCHN | 系统中毒 | 系统中毒，需要借件910-373 | 更换910-373备件后修复(by chuanwei) | SWO-321605 | Defect | HW | Analyzer Control Module | CF Card |  | Systemic poisoning | System poisoning, need to borrow 910-373 | Repair after replacement of 910-373 spare parts (by chuanwei)
No | Visit | Complaint | RCHN | 光学敏感性超范围 | 光学敏感性超范围，借903-364一个 | 更换903-364备件后修复(by chuanwei) | SWO-322108 | Defect | HW | Process Module | Optical Unit | Process wheel ? | Optical sensitivity over range | Optical sensitivity over range, borrowed 903-364 a | Repair after replacement of 903-364 spare parts (by chuanwei)
No | Visit | Complaint | RCHN | 放测试卡的门打不开 | 放测试卡的门打不开，借903-117一个 | 更换903-117后修复(by chuanwei) | SWO-321679 | Defect | HW | Analyzer Control Module | PCB, Main Board |  | The door to the test card won't open. | The door to the test card can't open, borrow 903-117 one | Repair after replacement 903-117 (by chuanwei)
No | Visit | Complaint | RESP | Error 793 Fallo en calibración de bombas | Comprobación de registros de datos del equipo. | Comprobación del motor de bomba de soluciones del módulo pH $BG, sustitución de este motor, se sustituyen los sensores de líquidos de este módulo, se realizan calibraciones a 2 puntos, ajuste de sensores, calibraciones de bombas, verificación final con QC. | SWO-322995 | Defect | HW | pH/BG Module | Pump Motor | Refurbished | Error 793 Pump calibration failure | Checking computer data records. | Pump engine check of pH module solutions $BG, replacement of this motor, liquid sensors of this module are replaced, 2-point calibrations are performed, sensor adjustment, pump calibrations, final verification with QC.
No | Visit | Complaint | RESP | Fallo cal de bombas, Fallo Oxímetro | Comprobación de registros de datos del equipo. | Comprobación del estado de funcionamiento del equipo, sustitución de sensores de líquidos del módulo de pH $BG y del módulo Oxi, ajuste de SL, cal de bombas, calibraciones a 2 puntos, sustitución del tubo de la cámara de medida superior, verificación final de funcionamiento con QC, | SWO-323039 | Defect | HW | Oxi module | Oxi module | Refurbished | Pump lime failure, Oxymeter failure | Checking computer data records. | Checking the operating status of the equipment, replacing liquid sensors of the $BG pH module and the Oxi module, SL adjustment, pump lime, 2-point calibrations, replacing the upper measuring chamber tube, final verification of operation with QC,
No | Visit | Complaint | RAME | 0581 oximetry mismatch error generated after running VK-R7 level 3 material | 0581 oximetry mismatch error. Ran several vials of VK-R7 material and found hemolyzer needs replacement | Replaced hemolyzer. Adjusted lamp unit. Ran thb calibration and complete oxi test. All passed. After that I ran VK-R7 level 3 material and problem has been resolved | SWO-323274 | Defect | HW | Oxi module | Hemolyzer | 0581 | 0581 oximetry mismatch error generated after running VK-R7 level 3 material | 0581 oximetry mismatch error. Ran several vials of VK-R7 material and found hemolyzer needs replacement | Replaced hemolyzer. Adjusted lamp unit. Ran thb calibration and complete oxi test. All passed. After that I ran VK-R7 level 3 material and problem has been resolved
No | Visit | Complaint | RDEU | 13.03.2018, 10:41 Uhr, Fr. Jannick, Tel. 021190430: Systemreinigung wird täglich angefordert | Schlauchwege verunreinigt, P2 defekt. | Schlauchwege zwischen CUP Wasch Unit und P1 erneuert, Y-Stück am Reagenzien Pack erneuert, Schlachtwege am Y Stück erneuert, HyperChlorid Reinigung durchgeführt, Blutfilter gereinigt, Service dump und Blank Messung durchgeführt. Nach Rücksprache mit TS P1 wieder zurück getauscht und P2 erneuert. | SWO-323132 | Defect | HW | Process Module | Wash unit | Revisit required?? | 13.03.2018, 10:41 a.m., Fr. Jannick, Tel. 021190430: System cleaning is requested daily | Hose ways contaminated, P2 defective. | Hoseways between CUP Washing Unit and P1 renewed, Y-piece at the reagent pack renewed, slaughter routes at the Y piece renewed, Hyperchloride cleaning carried out, blood filter cleaned, service dump and blank measurement performed. After consultation with TS P1, exchanged and renewed P2.
No | Visit | Complaint | RDEU | 16.03.2018, 15:25 Uhr, Anruferin Sr. Cordea, Tel.: 03078822545: HOTLINE - LP rastet nicht ein 16.03.2018, 15:50 Uhr, Anruferin Frau Isabell Lepp, Tel.: 03078822545: ABL fordert nun wieder zum LP Wechsel auf | Flüssigkeitstransport gescheitert | Fehlersuche; div. Serviceprogramme abgearbeitet; Startphase abgewartet; tHb-Kal. durchgef.; Kal. + QC i.O. | SWO-323879 | Assisted Replacement | SP | SP | FT |  | 16.03.2018, 15:25, caller Sr. Cordea, tel.: 03078822545: HOTLINE - LP does not snap in 16.03.2018, 15:50, caller Mrs. Isabell Lepp, tel.: 03078822545: ABL now calls for lp change again | Liquid transport failed | Troubleshooting; various service programs; start-up phase awaited; tHb-Kal. through it; Kal. + QC in the first place
No | Visit | Complaint | RAME | Printer roller dislodged and not printing | Printer roller dislodged and not printing. Went to Fedex to pick up part. | Replaced printer unit. Also went over proficiency test instructions. | SWO-323763 | Defect | HW | Printer Module | Printer Module |  | Printer roller dislodged and not printing | Printer roller dislodged and not printing. Went to Fedex to pick up part. | Replaced printer unit. Also went over proficiency test instructions.
No | Visit | Complaint | RDEU | Intervention erforderlich | Intervention erforderlich Lösungen können nicht gepumpt werden Vakuum-Fehler | Einlass, Rollenkranz (Pumpe) und Verschlauchung überprüft - Analysator gespült - Service Fluid Transport Tests etc. durchgeführt - Erneuerung Anschlussdichtung und Schlauch an Quetschventil (Vakuum-Fehler bleibt bestehen) - Erneuerung SC (Undichtigkeit) - Startphase - Funktionsüberprüfung etc. ohne Beanstandung | SWO-323658 | Assisted Replacement | SC | SC | Leak |  | Intervention needed | Intervention needed Solutions cannot be pumped Vacuum error | Inlet, roller ring (pump) and hose checked - analyzer rinsed - Service Fluid Transport tests etc. performed - Renewal connection seal and hose to squeezing valve (vacuum error remains) - Renewal SC (leakage) - Start phase - Functional check etc. without complaint
No | Visit | Complaint | RIND | Analyser is not working properly | Analysis was disable. sensor cassette was not install properly. | Remove sensor cassette first and reinstall properly. after that analyser start working properly. handed over in fine working condition. | SWO-323989 | [5 empty] | Analyser is not working properly | Analysis was disable. sensor cassette was not install properly. | Remove sensor cassette first and reinstall properly. after that analyser start working properly. handed over in fine working condition.
No | Visit | Complaint | RKK | 廃液から液漏れ。 | 廃液から液漏れ | 溶液パック交換、廃液マニフォールド・排液ポンプローター交換。 | SWO-323673 | [5 empty] | Liquid leakage from the waste liquid. | Liquid leakage from waste liquid | Solution pack replacement, waste liquid manifold and drainage pump rotor replacement.
No | Visit | Complaint | RESP | Sospecha de coagulos Fallo transporte de liquidos | Sospecha de coagulos Fallo transporte de liquidos | Limpieza del sistema Sustitucion de la aguja Limpieza inlet gasket Cal, qc | SWO-323887 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | Suspected clots Liquid transport failure | Suspected clots Liquid transport failure | Cleaning the system Needle replacement Cleaning in gasketlet Cal, qc
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 冲洗时液体传输错误，0508 | 更换试剂包后修复 | SWO-323489 | [5 empty] | Test card, reagent package failure | Liquid transfer error during flushing, 0508 | Repair after replacement of reagent kit
No | Visit | Complaint | RCHN | 报错1031 | 报错1031.需要借件903.364一个 | 更换903-364一个正常。(by chuanwei) | SWO-322891 | Defect | HW | Process Module | Optical Unit | 1031 | 1031 | 1031. A loan of 903.364 is required | Replace 903-364 one normal. (by chuanwei)
No | Visit | Complaint | RCHN | 报错 1031，1029 | 报错 1031，1029，借903-364，944-216（13695），944-214（13673）各一张。 |  | SWO-323423 | Defect | HW | Process Module | Optical Unit | 1031 | Error 1031,1029 | 1031,1029, borrow 903-364, 944-216 (13695), 944-214 (13673) each.
No | Visit | Complaint | RCHN | 报错1999.2.-1 | 报错1999.22.-1,需要借件903-030 | 更换903-030后修复(by chuanwei) | SWO-322799 | Cleaned | HW | Incubator Module | Bearing, Incubator wheel |  | Error 1999.2.-1 | False 1999.22.-1, need to borrow 903-030 | Repair after replacement 903-030 (by chuanwei)
No | Visit | Complaint | RDEU | 09.03.2018; Hr. Ahrens; Tel.: 030450651141; Hb QC Messfehler L3 | häufige Fehler bei Oximetriewerten, Reinigung und wiederholte tHb Kal ohne Erfolg, Cuv Faktor am oberen Limit | Gerät abgeschalten und Oximodul ausgebaut, Hemolysator ausgbaut und überprüft - Küvette beschädigt oder stark verschmutzt - ersetzt, Modul eingebaut, Startphase, Service tHb Kal, mehrere Kal, tHb Kal, QC | SWO-324468 | Defect | HW | Oxi module | Hemolyzer | Used analyzer | 09.03.2018; Hr. Ahrens; Tel.: 030450651141; Hb QC Messfehler L3 | common errors in oximetry values, cleaning and repeated tHb Kal without success, Cuv factor at the upper limit | Switch off the device and remove oximodul, hemolysator extended and checked - cuvette damaged or heavily soiled - replaced, module built-in, start phase, service tHb Kal, several Kal, tHb Kal, QC
Yes | Visit | Complaint | RDEU | RemoteSupport nicht möglich | Key für Remotesupport fehlt - geordert | Key installiert und Systemeinstellungen überprüft $ angepasst, Neustart, Startphase, Kal und QC | SWO-324469 | Re classified RMED | Remote Support installed |  |  |  | RemoteSupport not possible | Key for remote support missing - ordered | Key installed and system settings checked . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .
No | Visit | Complaint | RDEU | E'lyte driften im 1 Kalibr.punkt, QC fehler, | E'lyte mit ?, div. Kalibr.fehler, | Alle Messkammern mit Hypochlorid gereinigt, Dichtungsring an der Glukose und pH-Elektrode ersetzt, div.Wartungsarbeiten durchgeführt, Gerät kalibriert, QC durchgeführt, | SWO-323780 | Defect | HW | Electrode | Sealing Ring, pH |  | E'lyte driften im 1 Kalibr.punkt, QC fehler, | E'lyte mit ?, div. Kalibr.fehler, | All measuring chambers cleaned with hypochloride, sealing ring replaced on glucose and pH electrode, div.maintenance work carried out, device calibrated, QC performed,
No | Visit | Complaint | RDEU | Fehler 1279: Gemessene Flüssigkeitsmenge war oberhalb der Warngrenze Gerät verlangt öfter das Reinigungsverfahren | Weiterhin sporadische Reinigung nötig. Fehler 1279 | Anschlussstück zum Lösungspack durchgespült. Schlauch 6 erneuert. Systemüberprüfung und Reinigung durchgeführt. Servicedump gezogen. Gerät überprüft . | SWO-324241 | Defect | HW | Wet section | Needle Wash Unit | 1279 | Error 1279: Measured liquid quantity was above the warning limit Device requires the cleaning procedure more often | Further sporadic cleaning is necessary. Error 1279 | Connector to the solution pack rinsed through. Tube 6 renewed. system verification and cleaning. Service dump. Device checked .
No | Visit | Complaint | RDEU | fortlaufende Spülung der Sensorkassette unterbrochen | - Einlaßdichtung verkrustet - kein Material vor Ort - Auftrag Hr. Hennig (MT) | Einlaßdichtung gew.; Startphase abgewartet; Kal. + QC i.O. | SWO-324084 | Defect | HW | Auto Inlet module | Inlet gasket |  | continuous flushing of the sensor cassette interrupted | - Inlet seal encrusted - no material on site - Order Mr. Hennig (MT) | inlet seal. start-up phase awaited; Kal. + QC in the first place
No | Visit | Complaint | RCHN | 接口板腐蚀 | 接口板腐蚀，借902-972一个 | 更换902-972后修复(by chuanwei) | SWO-323987 | Defect | HW | Basic Unit | PCB, Sensor connector assembly |  | Interface plate corrosion | Interface plate corrosion, borrow 902-972 one | Repair after replacement 902-972 (by chuanwei)
No | Visit | Complaint | RAME | Analyzer keeps giving messages to replace the inlet gasket. | Fails pressure test vacuum error. 1063 code. In service mode it fails the pressure test . | I retubed instrument and replaced the pinch valve. It still fails pressure test. I replaced the SC and looks like that was it. Instrument taking a long time for Lactate to come in on calibrations. FSE will let instrument settle Down. | SWO-325374 | Assisted Replacement | SC | SC | Leak | 1063 | Analyzer keeps giving messages to replace the inlet gasket. | Fails pressure test vacuum error. 1063 code. In service mode it fails the pressure test . | I retubed instrument and replaced the pinch valve. It still fails pressure test. I replaced the SC and looks like that was it. Instrument taking a long time for Lactate to come in on calibrations. FSE will let instrument settle Down.
No | Visit | Complaint | RDEU | 19.03.2018, Frau Mengersdorf oder Frau Werling, Tel: 062150335589: seit einigen Tagen Pumpenkalibrationsfehler | Intervention erforderlich Lösungen können nicht gepumpt werden Einlassdichtung wurde bereits erneuert (beschädigt) | Überprüfung Analysator - Service Tests, LS-Justage und Pumpenkalibration durchgeführt - Startphase - Analysator ohne Beanstandung | SWO-324583 | Adjustment | HW | Basic Unit | Service Test | Pump | 19.03.2018, Mrs Mengersdorf or Mrs Werling, Tel: 062150335589: for several days pump calibration error | Intervention needed Solutions cannot be pumped Inlet seal has already been renewed (damaged) | Review Analyzer - Service Tests, LS-Adjustment and Pump Calibration Performed - Start Phase - Analyzer without Complaint
No | Visit | Complaint | RAME | PO2 not correlating with other ABL80 CO-OX on patient samples. Customer feels this analyzer reads lower on pO2. | Verified differences on pO2 between 80’s by running external QC. This analyzer read 130(out of range) and the adjacent 80 read 154. Found external QC lot numbers scanned in were different from what was actually being ran. | Performed zero check, ok. Scanned in new lot #’s for QC’s. | SWO-324698 | [5 empty] | PO2 not correlating with other ABL80 CO-OX on patient samples. Customer feels this analyzer reads lower on pO2. | Verified differences on pO2 between 80’s by running external QC. This analyzer read 130(out of range) and the adjacent 80 read 154. Found external QC lot numbers scanned in were different from what was actually being ran. | Performed zero check, ok. Scanned in new lot #’s for QC’s.
No | Visit | Complaint | RCHN | 报错1031 | 报错1031，需要借件903-364一个 |  | SWO-324177 | Defect | HW | Process Module | Optical Unit | 1031 | 1031 | 1031, need to borrow 903-364 one
No | 18 | Y | RPNZ | Blank | Lac Drifting in calibrations |  | D40D5F8F701F917C | Defect | HW | El/Met Module | Valve, air | Membrane | Blank | Lac Drifting in calibrations
No | Visit | Complaint | RCHN | 耗材无法扫描 | 耗材无法扫描，借902-822一个 |  | SWO-324183 | Defect | HW | Basic Unit | Barcode reader |  | Supplies cannot be scanned | Supplies can not be scanned, borrow 902-822 one
No | Visit | Complaint | RAME | Customer seeing leaking out of syringe port | Leaking from inlet. Found analyzer in hold with bent inlet probe when I arrived. | Replaced bent inlet probe. Flushed waste tubing and reinstalled inlet gasket. Calibrated analyzer, ran 3 levels of QC and 3 blood samples. Checked and found no leakage from the inlet gasket. | SWO-324776 | Defect | HW | Inlet module | Probe bend |  | Customer seeing leaking out of syringe port | Leaking from inlet. Found analyzer in hold with bent inlet probe when I arrived. | Replaced bent inlet probe. Flushed waste tubing and reinstalled inlet gasket. Calibrated analyzer, ran 3 levels of QC and 3 blood samples. Checked and found no leakage from the inlet gasket.
No | Visit | Complaint | RAME | MetHb and COHb not correlating with analyzer A and C. | Radiometer QC material correlates between analyzers. No unusual calibration failures or drift out of range for thb. Found thb and lamp value slightly higher compared to other analyzers. | Adjusted lamp and ran service thb calibration. | SWO-324752 | Adjustment | HW | Oxi module | Lamp unit |  | MetHb and COHb not correlating with analyzer A and C. | Radiometer QC material correlates between analyzers. No unusual calibration failures or drift out of range for thb. Found thb and lamp value slightly higher compared to other analyzers. | Adjusted lamp and ran service thb calibration.
No | Visit | Complaint | RKOR | Inlet does not absorb solution. | fluid line transport test. Inlet gasket with holeder, connector gasket check. Replace yearly service kit. | Replace connector gasket in yearly servyce kit. | SWO-327526 | Defect | HW | Inlet module | Connector gasket |  | Inlet does not absorb solution. | fluid line transport test. Inlet gasket with holeder, connector gasket check. Replace yearly service kit. | Replace connector gasket in yearly servyce kit.
No | Visit | Complaint | RAME | Customer called with the ABL 90 disconnecting from the wireless | Wireless dropping signal | Checked settings and wireless dongle. The analyzer will not connect to the hospital wireless. Checked power settings and wireless settings for troubleshooting. Jim Beck will contact UPMC wireless team to investigate. | SWO-325368 | Revisit required | SW | RMED SW | Wireless connection |  | Customer called with the ABL 90 disconnecting from the wireless | Wireless dropping signal | Checked settings and wireless dongle. The analyzer will not connect to the hospital wireless. Checked power settings and wireless settings for troubleshooting. Jim Beck will contact UPMC wireless team to investigate.
Yes | Visit | Complaint | RAME | I need to add the line to the ini file | I need to add a line in the ini file | I added the line in the Rome.ini to allow out of range values to show < &> | SWO-325857 | Re classified RMED | Customer request |  |  |  | I need to add the line to the ini file | I need to add a line in the ini file | I added the line in the Rome.ini to allow out of range values to show < &>
Yes | Visit | Complaint | RAME | The customer wanted the out of range turn on | A line needs to be added in the Roma.ini | I put the line in the Roma.ini | SWO-325858 | Re classified RMED | Customer request |  |  |  | The customer wanted the out of range turn on | A line needs to be added in the Roma.ini | I put the line in the Roma.ini
No | Visit | Complaint | RFRA | Erreur de position du joint d'introduction Blocage site introduction | Contrôle site introduction Contrôle circuits fluides | Repositionner site introduction Nettoyage circuit fluides | SWO-325406 | Adjustment | HW | Auto Inlet module | Inlet Position Cal. |  | Introductory seal position error Blocking site introduction | Site control introduction Fluid circuit control | Reposition introduction site Fluid circuit cleaning
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | pCO2定标不过 | 更换测试卡后修复 | SWO-326234 | [5 empty] | Test card, reagent package failure | pCO2 calibration but | Repair after replacing test card
No | Visit | Complaint | RKK | スタンバイ解除後、透析液でPHとPCO2がばらつく | PH、PCO2がばらつく | 装置引き上げ | SWO-324791 | [5 empty] | Ph and PCO2 are scattered in dialysate after standby release | PH, PCO2 are scattered | Equipment pulling up
No | Visit | Complaint | RHUN | pO2 param. piros | pO2 param. piros | Szenzorkazetta garis cseréje, Kalibrálás, Hibajelentés mentése | SWO-324998 | [5 empty] | pO2 param. Red | pO2 param. Red | Sensor module, replacing, calibration, saving error Reporting
No | Visit | Complaint | RAME | Error 292 and 581 | Checked data logs. All failed blood samples on one patient, however that patient was ran on another ABL80 and it ran successfully. | Replaced pinch valve. (Hemolyzer replaced on another SWO). | SWO-325547 | [5 empty] | Error 292 and 581 | Checked data logs. All failed blood samples on one patient, however that patient was ran on another ABL80 and it ran successfully. | Replaced pinch valve. (Hemolyzer replaced on another SWO).
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 探测到试剂包有漏电流 | 连续更换1张测试卡和1个试剂包后修复 | SWO-324832 | [5 empty] | Test card, reagent package failure | Leak current detected in reagent kit | Repair after continuous replacement of 1 test card and 1 reagent kit
No | Visit | Complaint | RCHN | BNP传感器错误提示不存在 | BNP传感器错误提示不存在，借905-908一个 | 更换905-908后修复 by Andrew | SWO-324794 | Defect | HW | Process Module | PCB, Cup dryer |  | BNP sensor error hint sits non-existent | BNP sensor error prompt does not exist, borrow 905-908 one | Repair by Andrew after replacement 905-908
No | Visit | Complaint | RCHN | 报错1035 | 更换903-364后需要重新定标，加借944-220（13459）一张。 | 更换903-364重新定标后修复 by Andrew | SWO-324981 | Defect | HW | Process Module | Optical Unit | 1035 | 1035 error | After replacing 903-364, you need to re-bid and borrow 944-220 (13459). | Repair by Andrew after replacement 903-364 recalibration
Yes | 7 | Y | RPAC | Blank | Annual service required on the analyser. Software Upgrade analyser to Version V6.18. |  | D40FB58B5C061DA8 | [5 empty] | Blank | Annual service required on the analyser. Software Upgrade analyser to Version V6.18.
No | Visit | Complaint | RCHN | 报错0636 | 报错0636，需要借件902-516 | 更换902-516备件后修复 by Andrew | SWO-324247 | Defect | HW | Gas module | Flow regulator | 0636 | Error 0636 | Error 0636, need to borrow 902-516 | Repair by Andrew after replacement of 902-516 spare parts
No | Visit | Complaint | RCHN | 触摸屏不灵敏 | 触摸屏不灵敏，借910-320一个 | 更换910-320备件后修复 by Andrew | SWO-324307 | Defect | HW | DMS_PC | Display |  | The touch screen is not sensitive | The touch screen is not sensitive, borrow 910-320 a | Repair by Andrew after replacing 910-320 spare parts
No | Visit | Complaint | RSCH | Rep. OXI | Fehler 0316 + 0258 OXI, Neon und Number too low OXI Test Neon Intensity und Neon Position Fail | Komplettes OXI-Modul ersetzt i.o Lampe justiert i.o Initialisierung i.o Kal., QC und tHb i.o System i.o *** Kunden-Modul wird in Thalwil repariert, da Spectrometer nicht lieferbar*** | SWO-325121 | Defect | HW | Oxi module | Oxi module |  | Rep. OXYGEN | Fehler 0316 + 0258 OXI, Neon und Number too low OXI Test Neon Intensity und Neon Position Fail | Complete OXI module replaces i.o Lamp adjusted i.o Initialization i.o Kal., QC and tHb i.o System i.o Customer module is repaired in Thalwil, because Spectrometer not available***
No | Visit | Complaint | RAME | PH intermittently results unstable measurements | PH and Ca++ measurements have been intermittently unstable. Reference membrane changes have not improved performance. Reference electrode AU $Cl coating is diminished | Replaced PH & reference electrodes. Calibrated, ran QC & blood samples without error. Instrument operating normally. | SWO-325926 | Assisted Replacement | Electrode | Electrode | Reference |  | PH intermittently results unstable measurements | PH and Ca++ measurements have been intermittently unstable. Reference membrane changes have not improved performance. Reference electrode AU /Cl coating is diminished | Replaced PH & reference electrodes. Calibrated, ran QC & blood samples without error. Instrument operating normally.
No | 5 | Y | RPAC | Blank | Intermittent drifts on Glu and Lac |  | D40EA8CEE60830D2 |  |  |  |  | LEH | Blank | Intermittent drifts on Glu and Lac
No | 4 | Y | RPAC | Blank | Crea taking long time to settle |  | D40EAA0BE65D30D5 |  |  |  |  | LEH | Blank | Crea taking long time to settle
No | Visit | Complaint | RAME | low pco2 QC recovery for Level 1 & 2. Results had been in range. Also observed general erratic tendencies. | Reviewed data logs, confirmed customers observation, pCO2 QC L1 & L2 have dropped from previously determined range, and exhibit erratic swings. Performed applicable service diagnostics, no errors reported or observed. PCO2 electrode had been recently replaced (SWO-312097) to resolve calibration failures due to excessive drift. Now observed unusually darkened pCO center shaft, not expected for a new electrode with date code of BC-06 (7-2019). PCO2 QC issues began shortly after this replacement occurred. | Replaced with known good used pCO2 electrode for testing. All QC are now within customers established range. Will monitor. | SWO-324792 | Assisted Replacement | Electrode | Electrode | pCO2 | Revisit required?? | low pco2 QC recovery for Level 1 & 2. Results had been in range. Also observed general erratic tendencies. | Reviewed data logs, confirmed customers observation, pCO2 QC L1 & L2 have dropped from previously determined range, and exhibit erratic swings. Performed applicable service diagnostics, no errors reported or observed. PCO2 electrode had been recently replaced (SWO-312097) to resolve calibration failures due to excessive drift. Now observed unusually darkened pCO center shaft, not expected for a new electrode with date code of BC-06 (7-2019). PCO2 QC issues began shortly after this replacement occurred. | Replaced with known good used pCO2 electrode for testing. All QC are now within customers established range. Will monitor.
No | Visit | Complaint | RLTD | Failing QC. | Reference related parameters failing QC. | Found clot in reference chamber. Cleaned out chamber. Replaced reference membrane. Ran calibration & QC's. All passed. | SWO-325435 | Cleaned | Fluidic pathway | pH/BG Module | Measuring chamber |  | Failing QC. | Reference related parameters failing QC. | Found clot in reference chamber. Cleaned out chamber. Replaced reference membrane. Ran calibration & QC's. All passed.
No | Visit | Complaint | RDEU | Software neu installieren | Windows aufgrund Fehler neu installieren ==> Gerät in Installationsphase beim Kunden Windows wurde bei Netzwerkseinstellungen beschäfigt | Windows und Analysersoftware neu installiert Konfiguration wieder hergestellt Netzwerkeinstellungen durchgeführt | SWO-325993 | Defect | SW | Windows | Windows |  | Reinstalling software | Reinstalling Windows due to errors ==> Device in installation phase at customer Windows has been used for network settings | Windows and analytics software reinstalled Configuration restored Network settings performed
No | 5 | Y | RPAC | Blank | Sensor interface PCB |  | D410EB9ED25EBF49 | Defect | HW | Basic Unit | PCB, Sensor interface |  | Blank | Sensor interface PCB
No | 6 | Y | RPAC | Blank | Sensor interface PCB |  | D410E8E6D214BF44 | Defect | HW | Basic Unit | PCB, Sensor interface |  | Blank | Sensor interface PCB
No | Visit | Complaint | RAME | Unable to scan barcodes | Follow up of issue with barcode scanner not recognizing any scanned barcode data. | Suspected hardware issue. Tried CPU, display and cable, did not resolve the issue. Suspect a software problem. Will return to reload windows and the applications. | SWO-326187 | Setup | SW | RMED SW | roma.ini | Barcode Reader | Unable to scan barcodes | Follow up of issue with barcode scanner not recognizing any scanned barcode data. | Suspected hardware issue. Tried CPU, display and cable, did not resolve the issue. Suspect a software problem. Will return to reload windows and the applications.
No | Visit | Complaint | RCHN | 测试卡初始化失败 | 测试卡初始化失败，借902-972和903-334各一个 | 更换902-972备件后修复 by Andrew | SWO-324977 | Defect | HW | Basic Unit | PCB, Sensor connector assembly |  | Test card initialization failed | Test card initialization failed, borrowing one each for 902-972 and 903-334 | Repair by Andrew after replacement of 902-972 spare parts
No | Visit | Complaint | RAME | CO-OX control board failure | Checked all board connections, | Removed wet section and replaced CO-OX control board, performed Tikhonov Matrix, performed tHb cal | SWO-326150 | [5 empty] | CO-OX control board failure | Checked all board connections, | Removed wet section and replaced CO-OX control board, performed Tikhonov Matrix, performed tHb cal
No | Visit | Complaint | RESP | Cliente informa que la bateria no se carga | Equipo no carga la bateria | Comprobación del estado del analizador. Se procede a apagar el equipo y desmontar bateria para su sustitución. Se instala y se enciende equipo. Bateria cargando, se entra en servicio y se observa equipo no detecta la nueva bateria. Por lo que se realizan pruebas así como indicaciones serevicetip correspondiente sin exito. Se vuelve a instalar bateria que no carga y se observa equipo la detecta. Se realiza Techline. | SWO-326334 | Defect | HW | Battery module | Battery module |  | Customer reports that the battery does not charge | Equipment doesn't charge the battery | Checking the status of the scanner. The equipment is switched off and the battery disassembled for replacement. The computer is installed and powered on. Battery charging, entering service and observing equipment does not detect the new battery. So tests are performed as well as corresponding serevicetip indications without success. Battery that does not charge is reinstalled and equipment is observed to detect it. Techline is performed.
No | Visit | Complaint | RCHN | 液路堵塞 | 液路堵塞 | 手动冲洗液路 by Andrew | SWO-327041 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway | Clot | Liquid road blockage | Liquid road blockage | Manual flushing of liquid road by Andrew
No | Visit | Complaint | RCHN | 机器无法开机 | 机器无法开机，需要借件910-342和910-285 | 更换910-285备件后修复 by Andrew | SWO-324878 | Defect | HW | Power Supply | Power Supply |  | The machine does not power on | The machine does not power on, 910-342 and 910-285 are required | Repair by Andrew after replacement of 910-285 spare parts
No | Visit | Complaint | RCHN | GLU定标不过 | GLU定标不过，借903-581一个 | 更换903-581备件后修复 by Andrew | SWO-324248 | Defect | HW | Basic Unit | PCB, Sensor interface | Glu | GLU calibration is not | GLU calibration, borrow 903-581 a | Repair by Andrew after replacement of 903-581 spare parts
No | Visit | Complaint | RCHN | 报错5784 | 报错5784，借944-220（13900）一个 | A lot of cups didn't been punched into reagent pack. It were cleaned completely then the error disappear. BTY, the door of reagent pack, which allow cups fall into its inside, was removed. by Andrew | SWO-324992 | Defect | Consumables | Test Kit | Cartridge | 5784 (Punch out) | 5784 | 5784, loan 944-220 (13900) one | A lot of cups didn't been punched into reagent pack. It were cleaned completely then the error disappear. BTY, the door of reagent pack, which allow cups fall into its inside, was removed. by Andrew
Yes | Visit | Complaint | RCHN | 报错1031、1024、1025 | 报错1031、1024、1025，借903-364一个 | 更换903-364备件后修复 by Andrew | SWO-324842 | [5 empty] | Errors 1031, 1024, 1025 | 1031, 1024, 1025, borrow 903-364 one | Repair by Andrew after replacement of 903-364 spare parts
No | Visit | Complaint | RIND | Po2,Pco2 parameters are yellow. | Checked analyzer status and found that gas 1 pressure was low. | Replaced Gas 1 cylinder and given 2 point calibration. Unit is working ok. | SWO-327344 | Assisted Replacement | Consumables | Gas cylinder | Gas cylinder |  | Po2,Pco2 parameters are yellow. | Checked analyzer status and found that gas 1 pressure was low. | Replaced Gas 1 cylinder and given 2 point calibration. Unit is working ok.
No | Visit | Complaint | RCHN | 血氧打问号，液路传感器出错 0537 | 血氧打问号，液路传感器出错 0537，借902-832一个。 | 借的备件未使用，用除蛋白清洗液路解决了故障。出现新的故障，PCO2定标失败，灵敏度超出正常范围，要再借945-612一个。 | SWO-324363 | Revisit required | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | Blood oxygen hit the question mark, liquid road sensor error 0537 | Blood oxygen hit the question mark, liquid road sensor error 0537, borrow 902-832 one. | The borrowed spare parts were not used and the fault was solved with the de-protein cleaning fluid route. A new fault occurs, the PCO2 calibration fails, the sensitivity is out of the normal range, to borrow another 945-612.
No | Visit | Complaint | RDEU | 20.03.2018, 11:36 Uhr, Fr. Salomon, Tel. 06213836477: Oxy Hardware Fehler. | Oximetrie Hardware-Fehler Kein Ergebnis tHb Kalibration | Oximetrie Test durchgeführt (0258 Neon Lamp Intensity too low) - Austausch Neon Lamp (aus Test-Modul) - Einstellungen Halogenlampe Hemolysator - Funktionsüberprüfung etc. ohne Beanstandung! | SWO-325725 | Defect | HW | Oxi module | Lamp unit | Neon lamp | 20.03.2018, 11:36, Ms. Salomon, tel. 06213836477: Oxy Hardware Error. | Oximetry Hardware Errors No result tHb calibration | Oximetry test performed (0258 Neon Lamp Intensity too low) - Replacement Neon Lamp (from test module) - Settings halogen lamp Hemolysator - function check etc. without complaint!
No | Visit | Complaint | RSCH | Rep. Ca | Immer wieder Ca Drift | Ref. Elektrode ersetzt i.o Ca und Ref. Membranenwartung durchgeführt i.o Kal. und QC i.o System i.o | SWO-325789 | Assisted Replacement | Electrode | Electrode | Reference |  | Rep. Ca | Again and again Ca Drift | Ref. Electrode replaces i.o Ca and Ref. Membrane maintenance carried out i.o Kal. and QC i.o System i.o
No | Visit | Complaint | RIND | PO2 Value low for all the sample | Checked the machine found PO2 Value low for all the sample when compared to another machine calibration was ok & QC done po2 was getting failed crosschecked po2 electrode & processed the sample result was correlating found po2 electrode was defective | Need to replace po2 electrode | SWO-325883 | Revisit required | Electrode | Electrode | pO2 |  | PO2 Value low for all the sample | Checked the machine found PO2 Value low for all the sample when compared to another machine calibration was ok & QC done po2 was getting failed crosschecked po2 electrode & processed the sample result was correlating found po2 electrode was defective | Need to replace po2 electrode
No | Visit | Complaint | RDEU | 20.03.2018 Sr. Corinna 051916023882 meldet einen Dunkelstromfehler. | System überprüft => Pc und optische Einheit defekt | POC und optische Einheit eingebaut => neue optische Einheit ist defekt neu bestellt | SWO-325651 | Revisit required | HW | Process Module | Optical Unit |  | 20.03.2018 Sr. Corinna 051916023882 reports a dark current error. | System checked => Pc and optical unit defective | POC and optical unit installed => new optical unit is defectively reordered
Yes | Visit | Complaint | RDEU | 20.03.2018 Sr. Corinna 051916023882 meldet einen Dunkelstromfehler. | Optische Defekt | Auf Ersatzteil in der GS Gewartet. Danach zum Kunden gefahren. Und Festgestellt das das Ersatzteil defekt ist. Neues Ersatzteil bestellt. AQT zur GS gebracht. | SWO-325653 | Re classified RMED | Double entry |  |  |  | 20.03.2018 Sr. Corinna 051916023882 reports a dark current error. | Optical defect | Waiting for spare part in the GS. Then drove to the customer. And determined that the spare part is defective. New spare part ordered. AQT to the GS.
No | Visit | Complaint | RAME | Unit is having constant problems, so onsite visit is need to retrieve data logs and service dump for analysis. | Unit is having mutiple different issues with consumables. | Retrieved data logs and service dump and sent off for evaluation. | SWO-326516 | Revisit required | Revisit required | Revisit required | Revisit required |  | Unit is having constant problems, so onsite visit is need to retrieve data logs and service dump for analysis. | Unit is having mutiple different issues with consumables. | Retrieved data logs and service dump and sent off for evaluation.
Yes | Visit | Complaint | RFRA | L'analyseur se bloque en intervention nécessaire +le site reste ouvert parfois +les résultats ne sont pas cohérents | ... | Déplacement pour installer analyseur de prêt Analyseur pas arrivé Prise de rdv pour lundi | SWO-326671 | Re classified RMED | to Document when |  |  |  | The analyzer crashes in needed intervention The site stays open sometimes The results are not consistent | ... | Moving to install loan analyzer Analyzer not arrived Rdv for Monday
No | Visit | Complaint | RDEU | 22.03.18 10:00Uhr Herr Bernerscheck 09017/578003 Leckage erkannt | Fehlersuche vor Ort | LS kage erkannt und Vakuumfehler. Lösungspack nochmals erneuert. Einlassdichtung erneuert. Ventilschlauch erneuert. Sensorkassette erneuert. Startphase abgewartet. Kal und QCs wurden während Startphase durchgeführt. ABL wieder messbereit. | SWO-327474 | Assisted Replacement | SP | SP | Leak |  | 22.03.18 10:00am Mr Bernerscheck 09017/578003 Leakage detected | Troubleshooting on site | LS kage detected and vacuum fault. solution pack renewed. Inlet seal renewed. Valve hose renewed. Sensor cartridge is renewed. start-up phase. Kal and QCs were performed during the launch phase. ABL ready to be measured again.
No | Visit | Complaint | RAME | Unable to scan barcodes | Continued investigation to resolve barcode scanner not reading barcode. | After reloading windows and application, and restoring data, determined problem with be an error in the Roma.ini file. The internal barcode setting had a text error causing the scanner not to function. Corrected the settings. Tested the scanner, issue resolved. | SWO-326188 | Defect | SW | RMED SW | roma.ini |  | Unable to scan barcodes | Continued investigation to resolve barcode scanner not reading barcode. | After reloading windows and application, and restoring data, determined problem with be an error in the Roma.ini file. The internal barcode setting had a text error causing the scanner not to function. Corrected the settings. Tested the scanner, issue resolved.
No | Visit | Complaint | RIND | Service Call. Power Supply Unit Need to be changed. | Need to change the power supply board. | Shut down the analyzer, open the back panel of the analyzer, remove the connections of it, and replace with the new one as covered in warranty, connect al the connection, check firm cable connection and close the panel, swicting on the analyzer successfully done, analyzer handed over to staff in well working condition. | SWO-326962 | Defect | HW | Basic Unit | Power Supply |  | Service Call. Power Supply Unit Need to be changed. | Need to change the power supply board. | Shut down the analyzer, open the back panel of the analyzer, remove the connections of it, and replace with the new one as covered in warranty, connect al the connection, check firm cable connection and close the panel, swicting on the analyzer successfully done, analyzer handed over to staff in well working condition.
No | Visit | Complaint | RDEU | Sporadisch Akku Meldung. Akku lt. Herrn Hartung tauschen. | Akku erneuern. | Akku erneuert. Kalibration ok. Testmessung mit an Luft hängendem Sensor ok. Kundin wird Sensor neu membranieren. | SWO-326352 | [5 empty] | Sporadic battery message. Battery according to Mr. Hartung. | Renew the battery. | Battery renewed. Calibration ok. Test measurement with air-hanging sensor ok. Customer will re-diaphragm sensor.
No | Visit | Complaint | RCHN | 光学强度超范围 1031 | 光学强度超范围 1031，借903-364，944-212*13315，944-258*13355各一张。 | 更换903-364重新定标后修复 by Andrew | SWO-325005 | Defect | HW | Process Module | Optical Unit | 1031 | Optical strength over range 1031 | Optical strength super range 1031, borrow 903-364, 944-212 x 13315, 944-258 x 13355 each. | Repair by Andrew after replacement 903-364 recalibration
No | Visit | Complaint | RCHN | 报错4548 | 报错4548，借903-028一个 | 更换903-028一个正常。by Andrew | SWO-325617 | Defect | HW | Cover and Chassis | PCB, Cabinet top | 4548 | 4548 error | 4548, loan 903-028 one | Replace 903-028 one normal. by Andrew
No | Visit | Complaint | RCHN | 血氧打问号，液路传感器出错 0537 | pco2定标失败，需要借件945-612 | 更换945-612一个正常。 by Andrew | SWO-325526 | Assisted Replacement | Electrode | Electrode | pCO2 |  | Blood oxygen hit the question mark, liquid road sensor error 0537 | pco2 calibration failed, need to borrow 945-612 | Replace 945-612 one normal. by Andrew
No | Visit | Complaint | RAME | Ca values appear low in QC levels 2 & 3 | Ca values in QC levels 2 & 3 .01 below range on analyzer. | Results values actually within insert ranges, on the mean, however analyzer ranges too narrow for results. Adjusted Ca ranges to insert ranges with customer permission. Customer will reestablish QC ranges. pCO2 became out of range during testing requiring a change in membrane. | SWO-326072 | User error | QC Ranges | QC Ranges | QC Ranges |  | Ca values appear low in QC levels 2 & 3 | Ca values in QC levels 2 & 3 .01 below range on analyzer. | Results values actually within insert ranges, on the mean, however analyzer ranges too narrow for results. Adjusted Ca ranges to insert ranges with customer permission. Customer will reestablish QC ranges. pCO2 became out of range during testing requiring a change in membrane.
No | Visit | Complaint | RDEU | 19.03.2018; Sr. Meridite; Tel.: 0211904361530; Undefinded lookup string 18.-1 | Software Fehlermeldung Corrupt XML | Aktivitäten überprüft,PISA gesichert, PISA komplett deinstalliert, AQT90 Software 8.9.613.0 installiert, Gerät konfiguriert, Parameter TNT und Ddimer kalibriert. | SWO-327586 | Defect | SW | RMED SW | RMED SW | Corrupt XML | 19.03.2018; Sr. Meridite; Tel.: 0211904361530; Undefinded lookup string 18.-1 | Software Fehlermeldung Corrupt XML | Checks activities,PISA backed up, PISA completely uninstalled, AQT90 software 8.9.613.0 installed, device configured, parameter TNT and ddimer calibrated.
No | Visit | Complaint | RKK | インレットにバーコードシールを巻いたシリンジをはめてもﾊﾞｰｺｰﾄﾞを読まない | ﾊﾞｰｺｰﾄﾞﾘｰﾀﾞ交換 | ﾊﾞｰｺｰﾄﾞﾘｰﾀﾞ交換 | SWO-327611 | Not cat |  |  | Not cat |  | Do not read the bar code even if the syringe wrapped around the bar code seal on the inlet is put on. | Bar code reader exchange | Bar code reader exchange
No | Visit | Complaint | RSCH | Rep | Verlangt Reinigung Fehler 1278 Proberörchen können nicht gelesen werden (Scanner) Proben ID wird nur manchmal gelesen(siehe Video) | Pumpe 2 und 4 ausgetauscht Konfiguration von altem Kundengerät installiert Keine Veränderung. Scannt unverändert nicht Keine Reparatur Vorort möglich Fehler 1278 behoben | SWO-326467 | Not cat |  |  | Not cat | 1278 | Rep | Requires cleaning error 1278 Sample tubes cannot be read (scanner) Sample ID is only read sometimes (see video) | Pump 2 and 4 replaced Configuration of old customer device installed No change. Does not scan unchanged No repair suburb possible Bug 1278 fixed
No | 57 | Y | RPNZ | Blank | Sensor cassette replacement due to 1310 lactate response fault. |  | D41865CE962D78E5 | Not cat | SC | SC | Lac | 1310 | Blank | Sensor cassette replacement due to 1310 lactate response fault.
No | Visit | Complaint | RAME | Customer called with the ABL 90 disconnecting from the wireless | Wireless will not connect Worked with Tom (IT) to troubleshoot the wireless connectivity issues. | Tom made the necessary changes to allow the analyzer to connect to the hospital wireless system. We verified the connectivity and verified the analyzer will reconnect on its own if the wireless signal is dropped. | SWO-327261 | Not cat | SW | RMED SW | Wireless connection |  | Customer called with the ABL 90 disconnecting from the wireless | Wireless will not connect Worked with Tom (IT) to troubleshoot the wireless connectivity issues. | Tom made the necessary changes to allow the analyzer to connect to the hospital wireless system. We verified the connectivity and verified the analyzer will reconnect on its own if the wireless signal is dropped.
No | Visit | Complaint | RAME | Turbidity to High and Spectrum mismatch errors | Confirmed issue when onsite. | Analyzer was set to proficiency mode for processing samples, causing turbidity error. Reset to normal syringe 65ul mode and ran 5 waste blood without error. | SWO-327446 | Not cat | SW | RMED SW | Sample Mode |  | Turbidity to High and Spectrum mismatch errors | Confirmed issue when onsite. | Analyzer was set to proficiency mode for processing samples, causing turbidity error. Reset to normal syringe 65ul mode and ran 5 waste blood without error.
No | Visit | Complaint | RDEU | Einstellungen überprüfen | Fehlersuche | Einstellungen in roma.ini überprüft und korrigiert | SWO-327092 | Not cat | SW | RMED SW | roma.ini | Barcode Reader | Check settings | Troubleshooting | Settings in roma.ini checked and corrected
No | Visit | Complaint | RDEU | 23.03.2018, Herr Böhm, Tel: 03447523008: AQC Fehler bei Oxiparametern alle Level. | Fehlersuche ==> QC-Messungen Oxi-Parameter mit Fehlern aufgrund von Kalibrierfehlern | QC-Messungen Lösungen A und C wiederholt | SWO-327201 | No fault found | No fault found | No fault found | No fault found |  | 23.03.2018, Mr. Böhm, Tel: 03447523008: AQC error with Oxi parameters all levels. | Troubleshooting ==> QC measurements Oxi parameters with errors due to calibration errors | QC measurements solutions A and C repeated
No | Visit | Complaint | RLTD | Tony Malone Tel 0207 2885428 Customer still having issues can you call Tony again please? | Inlet intermittent staying open. | Found inlet clutch screws to be loose. Adjust screws and tested position in service program passed ok. Blood capillary sample run through with no errors. | SWO-327563 | Adjustment | HW | Auto Inlet module | Inlet drive gear |  | Tony Malone Tel 0207 2885428 Customer still having issues can you call Tony again please? | Inlet intermittent staying open. | Found inlet clutch screws to be loose. Adjust screws and tested position in service program passed ok. Blood capillary sample run through with no errors.
No | Visit | Complaint | RFRA | Mesures PH hors limites CQ | Contrôle consommables | Remplacer sensor K7 Remplacer sol pack | SWO-327329 | Assisted Replacement | SP | SP | pH |  | PH measures off-limits CQ | Consumable stowecontrol | Replace sensor K7 Replace ground pack
No | Visit | Complaint | RCHN | 信号板故障 | 信号板故障，借903-334,902-972各一个 | 更换902-972后修复 by Andrew | SWO-324224 | Defect | HW | Basic Unit | PCB, Sensor connector assembly |  | Signal plate failure | Signal board fault, borrow 903-334, 902-972 each | Repair by Andrew after replacement 902-972
No | Visit | Complaint | RDEU | am Radiometer-Messgerät Werte um 1.5 % (Probanden auch Nicht-Raucher), ca. 50 – 70 % höhere Ergebnisse als am Messgerät von Roche im Labor | Messwerte mit verschieden Proben für MetHb im Vergleich zu einem ABL90 zu hoch (2*) | LP erneuert, Kal 2, tHb-Kal durchgeführt, keine Besserung tHb-Kal zunächst mit Fehlermeldung Verstopfung Fließwege | SWO-327160 | [5 empty] | radiometer meters values by 1.5% (subjects also non-smokers), approx. 50 - 70% higher results than on the Roche meter in the laboratory | Measurements with different samples for MetHb compared to an ABL90 too high (2*) | LP renewed, Kal 2, tHb-Kal performed, no improvement tHb-Kal initially with error message constipation flow paths
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 压力测试真空错误 1063 | 更换试剂包后修复 | SWO-326964 | [5 empty] | Test card, reagent package failure | Pressure test vacuum error 1063 | Repair after replacement of reagent kit
No | Visit | Complaint | RCHN | 探测到漏电流 964 | 探测到漏电流 964，借902-980一个。 | 备件未使用，更换测试卡后修复by Andrew | SWO-326146 | Assisted Replacement | SC | SC | Leak Current | 964 | Leak current detected 964 | The leakage current was detected at 964, borrowing 902-980. | Spare parts not used, repair by Andrew after replacement test card
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 试剂包漏电流 | 连续更换2张测试卡后修复 | SWO-326880 | [5 empty] | Test card, reagent package failure | Reagent packet leakage current | Repair after 2 consecutive test cards
No | Visit | Complaint | RCHN | 仓门无法关闭 | 仓门无法关闭，借903-198一个 | 更换903-198一个正常。 by Andrew | SWO-325615 | Defect | HW | Analyzer Control Module | PCB, Main Board |  | The warehouse door cannot be closed | The warehouse door cannot be closed, borrow 903-198 one | Replace 903-198 one normal. by Andrew
No | Visit | Complaint | RCHN | 测试卡卡盘 | 测试卡卡盘 | 取出测试卡后重新安装后修复 | SWO-327455 | Not cat |  |  | Not cat |  | Testing the kaka plate | Testing the kaka plate | Repair after removal of test card and reinstall
No | Visit | Complaint | RAME | Spectrometer converter failed.- | Examined QC data logs. | Verified QC results for service performed on 2 $19 and 2 $20. Attached image to SWO. | SWO-327590 | Defect | HW | Oxi module | off/on | Service tip issued | Spectrometer converter failed.- | Examined QC data logs. | Verified QC results for service performed on 2 $19 and 2 $20. Attached image to SWO.
No | Visit | Complaint | RIND | Ph and po2 Patient results was not proper | Ph and po2 results were not proper for patient results | checked the calibration log at that time ph and po2 sensitivity values are not in range so removed the electrdoe clott depostion was there in both ph and po2 electrode cleaned the electrode and chamber after that given cleaning and calibration all calibration was ok then tested the samples all are ok observed for few samples | SWO-327378 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway | Clot | Ph and po2 Patient results was not proper | Ph and po2 results were not proper for patient results | checked the calibration log at that time ph and po2 sensitivity values are not in range so removed the electrdoe clott depostion was there in both ph and po2 electrode cleaned the electrode and chamber after that given cleaning and calibration all calibration was ok then tested the samples all are ok observed for few samples
No | Visit | Complaint | RAME | Unable to Edit patient demo graphics | Customer unable to edit accession number. | Determine cause to be an error in the roma.ini file. Corrected the text for DoNotBlockAccNo=Yes. Tested, problem resolved. | SWO-326884 | Defect | SW | RMED SW | roma.ini |  | Unable to Edit patient demo graphics | Customer unable to edit accession number. | Determine cause to be an error in the roma.ini file. Corrected the text for DoNotBlockAccNo=Yes. Tested, problem resolved.
Yes | Visit | Complaint | RAME | Customer reporting after Health Connect went down, staff RT entered patient information in manually. Results do not shown reference ranges on results. | Confirmed customer report by review of data logs and patient results for specific date and time. The change log (audit trail) reveals numerous changes below results were eventually approved. Patient sex and DOB was not initially entered but added approximately 6 minutes and 26 minutes after results were first obtained. Although the accession number edit function was inhibited during the time the sample was run and the next morning, the RT Manager was able to edit the accession number. This was the only sample result this was possible on prior to turning the edit accession number function on . | Not corrective action taken at this time. No further impact to any other test results. All results continue to include correct reference ranges. Data logs, service dump and actual printout has been taken and a CQC case will be opened per Klaus Kjoeller for further investigation. | SWO-327649 | Re classified RMED | Pre investigation |  |  |  | Customer reporting after Health Connect went down, staff RT entered patient information in manually. Results do not shown reference ranges on results. | Confirmed customer report by review of data logs and patient results for specific date and time. The change log (audit trail) reveals numerous changes below results were eventually approved. Patient sex and DOB was not initially entered but added approximately 6 minutes and 26 minutes after results were first obtained. Although the accession number edit function was inhibited during the time the sample was run and the next morning, the RT Manager was able to edit the accession number. This was the only sample result this was possible on prior to turning the edit accession number function on . | Not corrective action taken at this time. No further impact to any other test results. All results continue to include correct reference ranges. Data logs, service dump and actual printout has been taken and a CQC case will be opened per Klaus Kjoeller for further investigation.
No | Visit | Complaint | RFRA | Erreur positions Guide trappe chargement cassé | Contrôle fonctionnement modules mécanique | Retrait trappe chargement K7 | SWO-327330 | Not cat |  |  | Not cat |  | Position error Broken loading trap guide | Control operating mechanical modules | K7 charging hatch removal
No | Visit | Complaint | RSCH | Rep | Barcodereader liest Barcodes nicht korrekt Keine Einstellungen möglich | Barcodereader und PCB alte Version eingebaut. Funktionskontrolle i.o Es wird eine Techline eröffnet. Weiteres vorgehen wird bei uns intern besprochen | SWO-327150 | Not cat |  |  | Not cat |  | Rep | Barcode reader does not read barcodes correctly No settings possible | Barcode reader and PCB old version built in. Functional control i.o A techline will be opened. Further action is discussed internally
No | Visit | Complaint | RKK | SC交換後装置メンテナンス必要のメッセージが出てSCを付け直すも改善せず。装置の再起動を行うも同様であった。サービス→TemperatureよりSubstrate TemperatureがエラーとなっておりSCの基盤の温度不良と判断。SCを交換すると改善した。 | SC交換後サーミスタ作動不良。 | SC交換。 | SWO-327516 | Assisted Replacement | SC | SC | Thermistor |  | After sc replacement the message of the equipment maintenance necessary comes out and the SC is reattached, but it does not improve. The same was true for restarting the device. Substrate Temperature from service → Temperature is an error and the temperature of the base of The SC is judged to be bad. SC was replaced and improved. | Malfunction of the thermistor after SC replacement. | SC exchange.
No | Visit | Complaint | RCHN | 不均匀的冲洗液 1324 | 不均匀的冲洗液 1324，借902-980一个 | 更换902-980后修复 by Andrew | SWO-327068 | Defect | HW | Oxi module | Hemolyzer |  | Uneven flushing fluid 1324 | Uneven flushing liquid 1324, borrowed 902-980 one | Repair by Andrew after replacement 902-980
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 无法抽吸溶液1343 | 连续更换3个试剂包修复 | SWO-327468 | [5 empty] | Test card, reagent package failure | Unable to pump solution 1343 | 3 reagent kits for continuous replacement
No | Visit | Complaint | RCHN | 报错1031 | 报错1031,借903-364一个 |  | SWO-327469 | Defect | HW | Process Module | Optical Unit | 1031 | 1031 | 1031, loan 903-364 one
No | Visit | Complaint | RCHN | 报错1035 | 报错1035，借903-364，944-212（13875）、944-220（13459）、944-258（13205缺货） | 更换903-364后重新定标修复 by Andrew | SWO-326178 | Defect | HW | Process Module | Optical Unit | 1035 | 1035 error | 1035, loan 903-364,944-212 (13875), 944-220 (13459), 944-258 (13205 out of stock) | Re-calibration repair by Andrew after replacing 903-364
No | Visit | Complaint | RKOR | Interface data communication error | Analyzer Status check Activity logs check Data communication test (ABL90 flex -> interface : OK) (Interface -> EMR : Falied) Call the interface technician Support the remote control of technician | Data communication test (ABL90 flex -> interface : OK) | SWO-330229 | No fault found | SW | Communication | Communication |  | Interface data communication error | Analyzer Status check Activity logs check Data communication test (ABL90 flex -> interface : OK) (Interface -> EMR : Falied) Call the interface technician Support the remote control of technician | Data communication test (ABL90 flex -> interface : OK)
Yes | 6 | Y | RPAC | Blank | UPS needed |  | D4161273AA13EF16 | Re classified RMED |  |  |  |  | Blank | UPS needed
No | 2 | Y | RCZE | Blank |  |  | D415E895AA594EFA | Unknown | Unknown | Unknown | Unknown |  | Blank
No | Visit | Complaint | RLTD | Ade Oremole on 0207 2885530 | Found inlet stuck out intervention | RePlaced inlet gasket as shown in video ok. Startup ok. Run service test for inlet no fault found working ok. 3 capillary put through without inlet sticking out. | SWO-328045 | Defect | HW | Auto Inlet module | Inlet gasket | Cap | Ade Oremole on 0207 2885530 | Found inlet stuck out intervention | RePlaced inlet gasket as shown in video ok. Startup ok. Run service test for inlet no fault found working ok. 3 capillary put through without inlet sticking out.
Yes | Visit | Complaint | RFRA | L'analyseur se bloque en intervention nécessaire +le site reste ouvert parfois +les résultats ne sont pas cohérents | ... Renvoi analyseur en expertise à Neuilly Plaisance | Installation d'un analyseur de prêt ABL90 flex 090R0063N010 Mise en route, installation des consommables présents sur l'ancien analyseur $ $ configuration $ $ essais et formation sur utilisation | SWO-328687 | Re classified RMED | to install Loan analyzer |  |  |  | The analyzer crashes in needed intervention The site stays open sometimes The results are not consistent | ... Expertise analyzer referral to Neuilly Plaisance | Installation of an ABL90 flex 090R0063N010 loan analyzer Start-up, installation of consumables present on the old analyzer $ configuration $ $ testing and training on use
No | Visit | Complaint | RFRA | Erreur fuite détectée | Joint d'introduction n'est pas en contact avec le pack de solution. Aiguille légèrement tordue. Mode service -->Hemolyzer bouché. Le service n'utilise pas le filtre à caillot. D'autres services viendraient utiliser l'ABL90Plus. | Débouchage du circuit fluidique, rinçage de la cassette. Nettoyage: á l'hypochlorite du circuit fluidique, connecteur sortie poubelle. Remplacement: joint d'introduction, tubulure vanne oxi. Changement: Hémolyseur. Test service: Fluid transport, Oxi $pO2 Redémarrage ABL90Plus. | SWO-328123 | Defect | HW | Oxi module | Hemolyzer | Leak | Leak error detected | Introductory seal is not in contact with the solution pack. Slightly twisted needle. Service mode ---Hemolyzer clogged. The service does not use the clot filter. Other services would use the ABL90Plus. | Unblocking the fluid circuit, rinsing the cassette. Cleaning: hypochlorite of the fluid circuit, connector out of trash. Replacement: introductory seal, oxi valve tubing. Change: Hemolyser. Test service: Fluid transport, Oxi $pO2 ABL90Plus reboot.
No | Visit | Complaint | RDEU | 23.03.2018, Frau Lipfert, tel: 07274504266: Datum läßt sich nicht einstellen. ABL jetzt ohne Funktion!!! | Garantie! Datum lässt sich nicht korrekt einstellen. - Springt in die Zukunft . SC und LP lassen sich nicht installieren. SC und LP aus VSC S $W geholt. | Konfiguration überprüft. Datum und Zeit auf Win Ebene eingestellt, gesichert und in Applikationssoftware überprüft. Datum jetzt korrekt eingestellt, Neustart durchgeführt, Datum und Zeit jetzt stabil. Daten gesichert und Gelöscht wegen falschen Datumsstempel. Neue SC und LP aus VSC S $W installiert, QC durchgeführt , alles o.k. | SWO-328294 | [5 empty] | 23.03.2018, Mrs Lipfert, tel: 07274504266: Date cannot be set. ABL now without function!!! | Guarantee! Date cannot be set correctly. - Jumps into the future . SC and LP cannot be installed. SC and LP from VSC S /W. | Configuration checked. Date and time set at win level, backed up and checked in application software. Date now set correctly, restart performed, date and time now stable. Data backed up and deleted due to incorrect date stamp. New SC and LP from VSC S $W installed, QC performed, all ok.
No | Visit | Complaint | RESP | Fallo calibración termistor. Llevan desde 16/03 con problemas intermitentes. Ya con dos SC. Temperatura ambiental adecuada y el otro equipo no está afectado. | Fallo de temperatura del casete y CL en amarillo | Se observó fallo de calibración del CL y mensajes recurrentes en los registros del equipo de fallo de temperatura en sustrato del casete. Se sustituyó tubo de conexión de la entrada al observarse deteriorado. Se limpiaron contactos del casete del electrodos. Se sustituyó casete y SP. Se realizaron todos los test de servicio sin incidencias. Se rebooteó el equipo. Tras esto el equipo queda operativo. Como se observó que en ocasiones aparece mensaje de fallo de temperatura al instalar los casetes, y que se solventa recalibrando o tras pasar tiempo de acondicionamiento del casete, se realizara seguimiento de los logs del equipo en los próximos mantenimientos para descartar problemas intermitentes del mismo. Se comprobó que las temperaturas de conservación de los consumibles en las neveras son correctas. Equipo operativo. | SWO-330716 | Assisted Replacement | SC | SC | Thermistor |  | Thermostor calibration failure. They've been with intermittent problems since 16/03. Already with two SCs. Adequate ambient temperature and other equipment is unaffected. | Cassette temperature failure and YELLOW CL | Cl calibration failure and recurring messages were observed in the cassette substrate temperature failure equipment logs. The inlet connection tube was replaced when it was observed to deteriorate. Contacts were cleaned from the electrode cassette. Cassette and SP were replaced. All service tests were performed without incident. The computer has been rebooted. After this, the computer is operational. As it was observed that temperature failure message sometimes appears when installing cassettes, and that it is solved by recalibrating or after time conditioning the cassette, the equipment logs will be tracked in the next maintenance for upcoming maintenance rule out intermittent problems. The storage temperatures of the consumables in the refrigerators were found to be correct. Operational team.
No | Visit | Complaint | RCHN | HCT无法测量 1223，1257 | HCT无法测量 1223，1257 ，借903-329一个。 |  | SWO-327633 | Defect | HW | Aspiration Module | Needle Wash Unit |  | HCT cannot measure 1223,1257 | HCT cannot measure 1223,1257 , borrow edith of 903-329.
No | Visit | Complaint | RCHN | 结果不准 | 结果不准 | 作比对，质控。验证机器准确性 by jesson | SWO-328419 | User error | Pre analytical | Pre analytical | Pre analytical |  | The result is not allowed. | The result is not allowed. | Compare, quality control. Verify machine accuracy by jesson
No | Visit | Complaint | RIND | Sample problem error | Sample problem error | checked the analyser then cleaned the gaskets and needle after that refixed the same , then pump calibration performed after that liquid sensor test performed all were ok then protein removal performed for twice , then calibrated to raedy mode analysed the sample and observed for few samples | SWO-330311 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | Sample problem error | Sample problem error | checked the analyser then cleaned the gaskets and needle after that refixed the same , then pump calibration performed after that liquid sensor test performed all were ok then protein removal performed for twice , then calibrated to raedy mode analysed the sample and observed for few samples
No | Visit | Complaint | RKOR | Storage door is not opened when click the insert cartridge button. Software loading error OPENSTORAGE_24 | Restart AQT90 FLEX Activity logs check service mode - storage door, storage rock test : OK | service mode - storage door, storage rock test : OK | SWO-330235 | Defect | HW | Cover and Chassis | Cartridge door |  | Storage door is not opened when click the insert cartridge button. Software loading error OPENSTORAGE_24 | Restart AQT90 FLEX Activity logs check service mode - storage door, storage rock test : OK | service mode - storage door, storage rock test : OK
No | Visit | Complaint | RLTD | Ann Carstairs 01851708294 Please contact Gwen regarding this issue as wants to remove solution pack | Analyser failing to reboot with blank screen Sampler mixer tray cracked | Replaced CPU with non stock spare Replaced sampler mixer tray with new | SWO-328552 | Defect | HW | DMS_PC | CPU |  | Ann Carstairs 01851708294 Please contact Gwen regarding this issue as wants to remove solution pack | Analyser failing to reboot with blank screen Sampler mixer tray cracked | Replaced CPU with non stock spare Replaced sampler mixer tray with new
No | Visit | Complaint | RDEU | 27.03.2018, 08:47 Uhr ==> Anrufer Fr. Halfmann, Tel.: 020644417212 ==> Einlass schließt nach Probenmessung nicht | Einlass schließt nicht, Mechanik rattert | Einlassnadel stark verbogen, Mechanik dadurch blockiert, Einlassnadel ersetzt, Einlass Positionen neu kalibriert, Initialisierung 50 Min, Thermostatisierung fehlgeschlagen, Neustart durchgeführt, Initialisierung 55 Min i.O.,Kal und QC durchgeführt, 45 ml Modus nur BG Parameter für MBU für die Oberärztin angelegt | SWO-328984 | Replaced | HW | Auto Inlet module | Probe bend |  | 27.03.2018, 08:47 ==> Caller Fr. Halfmann, Tel.: 020644417212 ==> Entrance does not close after sample measurement | Entrance does not close, mechanics rattle | Inlet needle strongly bent, mechanics blocked by it, inlet needle replaced, inlet positions recalibrated, initialization 50 min, thermostatization failed, restart performed, initialization 55 min i.O., Kal and QC performed, 45 ml mode only BG Parameters for MBU created for the senior doctor
Yes | 3 | Y | RPAC | Blank | Expired consumables. Customer was delivered a sensor cassette with shelf life of 6 weeks. It expired 1 week after being installed onto analyser |  | D4177C5A203B9523 | Re classified RMED | Logistical SWO |  |  |  | Blank | Expired consumables. Customer was delivered a sensor cassette with shelf life of 6 weeks. It expired 1 week after being installed onto analyser
No | Visit | Complaint | RFRA | Après comparaison l'ABL A est plus élevé sur la COHB | Vérification des historiques de contrôles | Nettoyage de la fluidique et de l'hémoliseur à l'hypochlorite $ $ relance $ $ calibration thb $ $ essais de comparaison avec ABL B (sang) $ $ toujours la différence $ $ remplacement de l'hémoliseur $ $ relance $ $ mise en route (1h) $ $ calibration thb $ $ passage des contrôles $ $ essais de comparaison avec ABL B -> | SWO-328689 | Defect | HW | Oxi module | Hemolyzer |  | After comparison the ABL A is higher on the COHB | Checking control history | Cleaning the fluidand and hemolizer to the hypochlorite $$ raise $$ calibration thb $$ test comparison with ABL B (blood) $ $ still the difference $ $ replacement of the hemolist $ $ raise $ start (1h$ $ calibration thb $ $ passage d es controls $$ testing comparison with ABL B -
Yes | Visit | Complaint | RFRA | Après comparaison l'ABL A est plus élevé sur la COHB | ... | Trajet retour | SWO-328690 | Re classified RMED | Double entry |  |  |  | After comparison the ABL A is higher on the COHB | ... | Return journey
Yes | Visit | Complaint | RFRA | Après comparaison, la Hbco est plus basse sur cet analyseur | Vérification des calibrations, des contrôles et des activités | Démontage $ $ remplacement de l'hémoliseur $ $ remplacement du tuyau de déchets et du tuyau de la vanne oxi $ $ relance $ $ mise en route (1H) $ $ calibration thb $ $ passage des contrôles $ $ essais bons $ $ passage de gaz en comparant avec ABL A -> toujours trop bas par rapport à l'ABL A -> intervention sur ABL A | SWO-328688 | Re classified RMED | See R142N12 SWOs |  |  |  | After comparison, the Hbco is lower on this analyzer | Checking calibrations, controls and activities | Disassembly $$ replacement of the hemolist $$ replacement of the waste pipe and the valve pipe oxi $$raise $$ start (1H) $ calibration thb $ $ passage of controls $ tests good $$ gas passage by comparing with ABL A - always too ba compared to ABL A - intervention on ABL A
No | Visit | Complaint | RCHN | 打印机故障 | 打印机故障，借903-001,902-985,636-523各一个 | 更换902-985备件后修复 by jesson | SWO-328292 | [5 empty] | Printer failure | Printer failure, borrow 903-001, 902-985, 636-523 each | Repair by jesson after replacement of 902-985 spare parts
No | Visit | Complaint | RIND | Service Visit. | Parameter deactivate, showing in main screen. | After observing, user run the clotted sample, flush the cassette with disttiled water and Na heparin, reinitiate the sensor cassette, successfully done, analyzer handed over to staff in well working condition. | SWO-330382 | [5 empty] | Service Visit. | Parameter deactivate, showing in main screen. | After observing, user run the clotted sample, flush the cassette with disttiled water and Na heparin, reinitiate the sensor cassette, successfully done, analyzer handed over to staff in well working condition.
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | GLU,LAC定标结果超范围 0378，1310 | 更换两张测试卡正常。 | SWO-329165 | [5 empty] | Test card, reagent package failure | GLU, LAC calibration results over range 0378,1310 | Replacing the two test cards is ok.
No | Visit | Complaint | RCHN | 系统与软件不匹配 | 系统与软件不匹配，借903-198一个 | 更换903-198一个正常。by jesson | SWO-328348 | Defect | HW | Analyzer Control Module | PCB, Main Board |  | System and software do not match | The system does not match the software, borrow 903-198 a | Replace 903-198 one normal. by jesson
No | Visit | Complaint | RCHN | hct定标失败，报错1223 | hct定标失败，报错1223，需要借件636-561 | 更换636-561后修复 by jesson | SWO-328247 | Defect | HW | Aspiration Module | Needle Wash Unit |  | hct calibration failed, error 1223 | hct calibration failure, error 1223, need to borrow 636-561 | Repair by jesson after replacement 636-561
No | Visit | Complaint | RCHN | 报错1031 | 报错1031更换903-364后，需要重新定标 944-212（13719），944-220（13612）各一张 | 更换944-212（13719），944-220（13612）后重新定标修复 by jesson | SWO-327679 | Defect | HW | Process Module | Optical Unit | 1031 | 1031 | After 1031 replacement 903-364, it is necessary to re-order 944-212 (13719), 944-220 (13612) each | Replace 944-212 (13719), 944-220 (13612) and then re-label to repair by jesson
No | Visit | Complaint | RLTD | Simon Malin 02476968510 | On arrival fitted a new AU control board, on powering back on during the reboot error window came up saying AUengine.exe could not be found. Waited for analyser to finish reboot and left it for a few minutes and it remained in the same fault condition DMS not connecting to AU. In windows explorer found AU engine folder and AU engine.exe was not in there. Carried out search for AU engine.exe and it was found, copied file and pasted it into AU engine folder. Powered off the analyser and refitted analysers original AU control board. | Located AU engine.exe, placed it back in AU engine folder, refitted analysers original AU control board, powered on and analyser rebooted fine and DMS connected to AU and analyser went through start up procedure. | SWO-329722 | Defect | SW | RMED SW | RMED SW |  | Simon Malin 02476968510 | On arrival fitted a new AU control board, on powering back on during the reboot error window came up saying AUengine.exe could not be found. Waited for analyser to finish reboot and left it for a few minutes and it remained in the same fault condition DMS not connecting to AU. In windows explorer found AU engine folder and AU engine.exe was not in there. Carried out search for AU engine.exe and it was found, copied file and pasted it into AU engine folder. Powered off the analyser and refitted analysers original AU control board. | Located AU engine.exe, placed it back in AU engine folder, refitted analysers original AU control board, powered on and analyser rebooted fine and DMS connected to AU and analyser went through start up procedure.
No | Visit | Complaint | RAME | Calibration aorted | 1 and 2 point cals aborting intermittently Error 500 rinse detected in pH $BG module when not expected | Replaced Membrane for flat membrane valve on Ph $BG module part number 834-222. | SWO-328719 | Defect | HW | pH/BG Module | Valve, air |  | Calibration aorted | 1 and 2 point cals aborting intermittently Error 500 rinse detected in pH /BG module when not expected | Replaced Membrane for flat membrane valve on Ph $BG module part number 834-222.
No | 57 | Y | RPNZ | Blank | Solution pack replacement due to |  | D4190CEF963AF8F5 | Assisted Replacement | SP | SP | SP |  | Blank | Solution pack replacement due to
No | Visit | Complaint | RKOR | Pump calibration error | detect $ remove clots Liquid transport test -> OK Solution Pack test -> Failed Replace Solution Pack -> OK Adjust liquid sensor Replace errored Solution pack -> OK | detect $ remove clots Adjust liquid sensor | SWO-330242 | Assisted Replacement | SP | SP | Pump cal |  | Pump calibration error | detect / remove clots Liquid transport test -> OK Solution Pack test -> Failed Replace Solution Pack -> OK Adjust liquid sensor Replace errored Solution pack -> OK | detect $ remove clots Adjust liquid sensor
No | Visit | Complaint | RDEU | 28.03.2018, 08:35 Uhr, Fr. Fiedler, Tel.: 061514037103: ABL verlangt ständig ein neues LP | Geräte-Überprüfung | Überprüfung Analysator - Funktionsüberprüfung etc. ohne Beanstandung ! Bem.: Lösungspack LOT BD-51 und LOT BD-52 verwendet | SWO-329168 | Assisted Replacement | SP | SP | SP |  | 28.03.2018, 08:35, Fr. Fiedler, Tel.: 061514037103: ABL constantly demands a new LP | Device verification | Review Analyzer - Function check etc. without complaint ! Bem.: Solution pack LOT BD-51 and LOT BD-52 used
No | Visit | Complaint | RAME | Oxi Spectrum error | Spectrum error on blood and Qc | Replace hemolyzer fixed the problem | SWO-329372 | Defect | HW | Oxi module | Hemolyzer |  | Oxi Spectrum error | Spectrum error on blood and Qc | Replace hemolyzer fixed the problem
No | Visit | Complaint | RLTD | Lynne Tel 01582 497991 | Battery error | Replaced battery. | SWO-329175 | Defect | HW | Battery module | Battery module |  | Lynne Tel 01582 497991 | Battery error | Replaced battery.
Yes | 6 | Y | RPAC | Blank | Failed sensor cassette |  | D4186F7896294454 | [5 empty] | Blank | Failed sensor cassette
No | Visit | Complaint | RSAF | Sr Anna called asking assistance for the analyzer keeps wanting solution pack replaced but the inl, Done telephonic trouble shooting and could not resolve the issue. | Inlet probe stuck in the closed position when wanting to replace the solution pack and then says solution pack lifetime in analyzer exceeded. | Entered Service screen, Done complete fluid transport test and had to clean all gaskets and interconnecting tubing's, reran fluid transport and all past, Changed time as it was displaying an hour in advance, committed the change to CF card. | SWO-329251 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | Sr Anna called asking assistance for the analyzer keeps wanting solution pack replaced but the inl, Done telephonic trouble shooting and could not resolve the issue. | Inlet probe stuck in the closed position when wanting to replace the solution pack and then says solution pack lifetime in analyzer exceeded. | Entered Service screen, Done complete fluid transport test and had to clean all gaskets and interconnecting tubing's, reran fluid transport and all past, Changed time as it was displaying an hour in advance, committed the change to CF card.
No | Visit | Complaint | RFRA | Erreur joint d'introduction | Vérification en mode service Déblocage site d'introduction, calibration en mode service. | Chgt: aiguille, joint d'introduction. Chgt fin de cycle: cassette et pack Solution. Dans les autres services: DAR G Chauliac, UTec, DAR St E. DAR ADV. Chgt aiguille et joint introduction. Lors du chgt des joints d'introduction, pensez à nettoyer le corps de l'aiguille pour permettre au joint de bien coulisser. | SWO-329277 | Defect | HW | Auto Inlet module | Inlet gasket |  | Introducing joint error | Checking in service mode Unlocking introductory site, calibration in service mode. | Chgt: needle, introductory joint. Chgt end of cycle: cassette and Solution pack. In other services: DAR G Chauliac, UTec, DAR St E. DAR ADV. Chgt needle and joint introduction. When chgting the introductory joints, consider cleaning the body of the needle to allow the joint to slide properly.
Yes | Visit | Complaint | RDEU | Sc und LP verfallen | Gerät überprüft | LP und SC geliefert und im Gerät eingebaut, | SWO-329844 | Re classified RMED | Logistical SWO |  |  |  | Sc and LP expire | Device checked | LP and SC and installed in the device,
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | PH,CA,K定标不过 1079 | 更换测试卡后修复 | SWO-329043 | Assisted Replacement | SC | SC | pH | 1079 | Test card, reagent package failure | PH, CA, K calibration but 1079 | Repair after replacing test card
No | Visit | Complaint | RCHN | 无法开机 | 无法开机，借910-285一个。 | 更换910-285后修复 by jesson | SWO-328266 | Defect | HW | Power Supply | Power Supply |  | Cannot power on | Cannot power on, borrow 910-285 one. | Repair by jesson after replacement 910-285
No | Visit | Complaint | RCHN | 报错 1035 | 报错 1035，借903-364一个。 | 更换903-364后修复 by jesson | SWO-328249 | Defect | HW | Process Module | Optical Unit | 1035 | Error 1035 | 1035, borrow 903-364 one. | Repair by jesson after replacing 903-364
No | Visit | Complaint | RCHN | 报错1031和1068 | 报错1031和1068，需要借件903-364和903-210、944-258（13355）、944-394（13849） | ，温度重新校准后正常。更换903-364重新定标正常。by jesson | SWO-327773 | Defect | HW | Process Module | Optical Unit | 1031 | Errors 1031 and 1068 | 1031 and 1068, need to borrow 903-364 and 903-210, 944-258 (13355), 944-394 (13849) | Normal after the temperature is recalibrated. Replace the 903-364 re-calibration normal. by jesson
No | Visit | Complaint | RAME | Oxi LS error 0656: sample detected in oxi LS prematurely during aspiration. | Inspection of the oxi system revealed the oxi pinch valve is moving hesitantly and hemolyzer is not passing service menu complete oxi test. | Replaced oxi pinch valve (tube valve), and pinch tubing as well since it was damaged during disassembly. Hemolyzer was ordered overnight and replaced. | SWO-328183 | Defect | HW | Oxi module | Valve, pinch valve |  | Oxygen LS Error 0656: Sample detected in oxygen LS prematurely during aspiration. | Inspection of the oxygen system revealed the oxygen pinch valve is moving hesitantly and Hemolyzer is not passing service menu complete oxygen test. | Replaced oxi pinch valve (tube valve), and pinch tubing as well since it was damaged during disassembly. Hemolyzer was ordered overnight and replaced.
No | Visit | Complaint | RAME | Waste detector failed. | Original waste detector failed, needs replacement. | Replaced waste detector. | SWO-328880 | Defect | HW | Basic Unit | Waste detector |  | Waste detector failed. | Original waste detector failed, needs replacement. | Replaced waste detector.
Yes | 7 | Y | RCZE | Blank | u Ca nevychazí kontroly ani po výmìnì membrány na ref a ca |  | D41982BA0C34A331 | [5 empty] | Blank | In the case of Ca, do not break the checks even after replacing the membrane
No | Visit | Complaint | RIND | Service Visit. | Analyzer shows again and again rinse error while we give the rinse command. | Temporary shutdown the analyzer, change the waste pump tube rinse tube, restart the analyzer, successfully done, analyzer handed over to staff in well working condition. | SWO-330385 | Defect | HW | Fluidic manifold | Tubes, Fluidic Manifold |  | Service Visit. | Analyzer shows again and again rinse error while we give the rinse command. | Temporary shutdown the analyzer, change the waste pump tube rinse tube, restart the analyzer, successfully done, analyzer handed over to staff in well working condition.
No | Visit | Complaint | RDEU | Anruf Hr. Seljanko Ableitstrom seit gestern früh Durch Geräteneustart wieder weg, über Nacht aber wieder da | LP wurde am 27.03. planmäßig gewechselt => bei Systemüberprüfung in der Nacht Meldung Ableitstrom => daraufhin SC und Einlassdichtung getauscht und Gerät neugestartet => Ableitstrom wieder in der Nacht erkannt Nach Neustart wieder ohne Fehler | Überprüfung auf optische Leckagen => alles i.O. Spülung Fließwege und Messung Impedanz => Leak Voltage über 1000 mV Austausch Verschlauchung, QC Rückstände an Hemolyzer und LS beseitigt => Geräteneustart => Leak Voltage weiterhin zu hoch Austausch LP aus Kundenbestand Reklamationsausdrucke erstellt und übergeben | SWO-329713 | Assisted Replacement | SP | SP | Leak Current |  | Call Mr. Seljanko Discharge current since yesterday morning Gone again by device restart, but back there overnight | LP was changed as scheduled on 27.03. => during system check in the night message dissipation current => then SC and inlet seal changed and device restarted => discharge current detected again at night After rebooting again without errors | Checking for optical leaks => all in the art Flushing flow paths and measurement impedance => Leak Voltage over 1000 mV Replacement hose, QC residues on hemolyzer and LS eliminated => Device restart => Leak Voltage still too high Replacement LP from customer base Complaint expressions created and passed
No | Visit | Complaint | RDEU | QC-Verletzung FHbF und sO2 | QC-Verletzung FHbF und sO2(in Level B und C erhöht) Siehe auch Case - Oxi-Hardware-Fehler(Spectrometer und Hemolysator bereits erneuert) | Analysator heruntergefahren - Oximetrie demontiert und Quetschventil erneuert - Rückmontage und Funktionsüberprüfung etc. | SWO-329826 | Defect | HW | Oxi module | Pinch valve |  | QC violation FHbF and sO2 | QC violation FHbF and sO2 (increased in level B and C) See also Case - Oxi hardware error (Spectrometer and Hemolysator already renewed) | Analyzer shut down - Oximetry disassembled and crushing valve renewed - reassembly and function check etc.
No | Visit | Complaint | RLTD | Please call Denise Young on 0203 299 3222 | On arrival found that new solution pack had not been activated fully. Removed pack and fully activated and refitted. Also noted Glucose sensor had been failing calibration, so replaced Sensor Cassette (failure claim printout handed to Cameisha). Verified operation on restart - all ok. Performed tHb calibration - ok. Analyser ready for use. | Correctly activated solution pack. Replaced faulty sensor cassette. Discussed $suggested refresher training for analyser replacements - Cameisha to advise Radiometer of a suitable date. | SWO-329895 | User error | SP Activation | SP Activation | SP Activation |  | Please call Denise Young on 0203 299 3222 | On arrival found that new solution pack had not been activated fully. Removed pack and fully activated and refitted. Also noted Glucose sensor had been failing calibration, so replaced Sensor Cassette (failure claim printout handed to Cameisha). Verified operation on restart - all ok. Performed tHb calibration - ok. Analyser ready for use. | Correctly activated solution pack. Replaced faulty sensor cassette. Discussed $suggested refresher training for analyser replacements - Cameisha to advise Radiometer of a suitable date.
No | 1 | Y | RPAC | Blank | Adjust grub scew |  | D419C67E0C07E65B | Defect | HW | Auto Inlet module | Inlet drive gear |  | Blank | Adjust grub scew
No | Visit | Complaint | RCHN | 接口板漏液腐蚀 | 接口板漏液腐蚀，借902-972一个 | 更换902-972备件后修复 by jesson | SWO-328324 | Defect | HW | Basic Unit | PCB, Sensor connector assembly |  | Interface plate leakage corrosion | Interface plate leakage corrosion, borrow 902-972 one | Repair by jesson after replacement of 902-972 spare parts
No | Visit | Complaint | RDEU | sO2 besteht QC Messung nicht | sO2 besteht QC Messung nicht. Fehler tritt nach nur 4 Messungen auf. | Nach Rücksprache mit Hr. Radloff ABL90 (090R0860N016) Reparaturersatzgerät geliefert. Frau Meixner führt Kalibrationen und QC’s am Ersatzgerät durch. | SWO-329878 | [5 empty] | sO2 does NOT exist QC measurement | sO2 does not exist QC measurement. Error occurs after only 4 measurements. | After consultation with Mr. Radloff ABL90 (090R0860N016) repair replacement device supplied. Ms Meixner performs calibrations and QC's on the replacement device.
No | Visit | Complaint | RCHN | LAC,GLU定标失败 | LAC,GLU定标失败 ，借903-581一个。 | 更换903-581后修复 by jesson | SWO-328944 | Defect | HW | Basic Unit | PCB, Sensor interface | Glu/Lac | LAC, GLU bid failed | LAC, GLU bid failed , borrow edith of 903-581. | Repair by jesson after replacement 903-581
No | Visit | Complaint | RCHN | 新装机开机故障+定标故障 | 1，新装机开机报5689，存储盘底部盘子卡扣断裂；2，Myo批号13886定标失败两次，买兰州荣诚定标卡两张。 | 1，存储盘底部配件稍后找机会更换，定标卡本地寻找继续定标，购买兰州荣诚两张定标卡 | SWO-329884 | Adjustment | HW | Storage Module | Storage Wheel | 5689 | New installation start-up fault and calibration fault | 1, the new installation boot reported 5689, the bottom plate of the storage disk buckle fracture; 2, Myo batch number 13886 bid failure twice, buy Lanzhou Rongcheng bid card two. | 1, storage disk bottom accessories later to find the opportunity to replace, the bid card local search to continue to bid, purchase Lanzhou Rongcheng two bid cards
No | Visit | Complaint | RCHN | 报错5257，光单不工作 | 报错5257，光单不工作，借636-550一个 | 更换636-550后修复 by jesson | SWO-328314 | Defect | HW | Process Module | Sensor, Process module | 5257 | 5257, light single does not work | 5257, light single does not work, borrow 636-550 one | Repair by jesson after replacement 636-550
No | Visit | Complaint | RBNL | Oxi-mètre | De temps en temps des résultats qc hors gamme (déjà signalé l’année dernière) | Remplacement de l’hemolyseur, cal tHb en service et contrôle sur même ampoule : OK | SWO-329837 | Defect | HW | Oxi module | Hemolyzer |  | Oxi-meter | Occasionally off-range qc results (already reported last year) | Replacement of hemolyseur, cal tHb in service and control on same bulb: OK
No | Visit | Complaint | RDEU | Immer wieder Driftfehler E‘Lyte | Immer wieder Driftfehler E‘Lyte. pH auch immer wieder hohe Driften. | Messkammern der pH, Ref, Na, K und Ca Elektrode gereinigt. Präventiv Referenzelektrode erneuert. Ref, Na, K und Ca Elektrode neu membraniert. Kontakte der pH, Ref, Na, K und Ca Elektrode, sowie Kontaktkappen gereinigt. Unteren pH Dichtring erneuert. Kalibrationen und QC’s vorgenommen. | SWO-329877 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | Immer wieder Driftfehler E‘Lyte | Again and again drift error E'Lyte. pH also high drifts again and again. | measuring chambers of pH, Ref, Na, K and Ca electrode. Preventive reference electrode renewed. Ref, Na, K and Ca electrode. Contacts of pH, Ref, Na, K and Ca electrode, as well as contact caps cleaned. Lower pH sealing ring renewed. calibrations and QC's.
No | Visit | Complaint | RDEU | 20.03.2018, 11:36 Uhr, Fr. Salomon, Tel. 06213836477: Oxy Hardware Fehler. | Erneuerung Neon Lampe | Erneuerung Neon Lampe und anschließend Halbjahreswartung durchgeführt (siehe Case) | SWO-329827 | Defect | HW | Oxi module | Lamp unit |  | 20.03.2018, 11:36, Ms. Salomon, tel. 06213836477: Oxy Hardware Error. | Renewal Neon Lamp | Renewal neon lamp and then half-year maintenance carried out (see case)
No | Visit | Complaint | RAME | Customer reports that QC has failed with all parameters out of range | QC’s are periodically run as the wrong level. All ampoules are scanned and read properly but sometimes level 1 is run as a level 2 or level 3 QC. As a result, all values are out of range. AutoCheck 14 pole ribbon cable has been replaced to address this issue on a previous visit. | Replaced the AutoCheck module and verified functionality. Analyzer is within acceptable operating specifications. | SWO-329942 | Defect | HW | Autocheck Module | Autocheck Module |  | Customer reports that QC has failed with all parameters out of range | QC’s are periodically run as the wrong level. All ampoules are scanned and read properly but sometimes level 1 is run as a level 2 or level 3 QC. As a result, all values are out of range. AutoCheck 14 pole ribbon cable has been replaced to address this issue on a previous visit. | Replaced the AutoCheck module and verified functionality. Analyzer is within acceptable operating specifications.
No | Visit | Complaint | RFRA | Erreur positions Guide trappe chargement cassé | Contrôle capot avant | Face avant à remplacer | SWO-330370 | Not cat |  |  | Not cat |  | Position error Broken loading trap guide | Front hood control | Front face to be replaced
No | Visit | Complaint | RBNL | TnT QC fout | QC waarde (enkel level 1) te laag voor TnT Test met blanco metingen toont een wederkerende afwijking op cup positie 1 van het processwheel. Ook de laatste 2 calibraties tonen dezelfde afwijking (telkens meting 1 & 6) Reiniging van het processwheel heeft niks uitgehaald, integendeel, de waarden liggen nu nog slechter. | Het processwheel zal dus moeten vervangen worden en is ondertussen besteld. | SWO-330324 | Revisit required | HW | Process Module | Process Wheel |  | TnT QC fout | QC value (only level 1) too low for TnT Test with blank measurements shows a repressing deviation at Cup position 1 of the processwheel. Also the last 2 calibrations show the same deviation (each measurement 1 & 6) Cleaning of the Processwheel has not taken off anything, on the contrary, the values are now even worse. | The Processwheel will therefore have to be replaced and ordered in the meantime.
No | Visit | Complaint | RCHN | 压力测试错误 | 压力测试错误，借903-129一个 | 更换903-129备件后修复 by jesson | SWO-329557 | Defect | HW | Oxi module | Valve, pinch valve |  | Stress test errors | Stress test error, borrow 903-129 one | Repair by jesson after replacement of 903-129 spare parts
No | Visit | Complaint | RCHN | 定标终止，仪器死机 | 定标终止，仪器死机，借944-258（14018）一张 | 重新定标后更换944-258（14018）后修复 | SWO-330232 | Assisted Replacement | Consumables | Test Kit | Cartridge, NT-PROBNP | Calibration | Calibration termination, instrument crash | Bid termination, instrument crash, borrow 944-258 (14018) one | Repair after re-calibration and replacement of 944-258 (14018)
No | Visit | Complaint | RCHN | 样本量超出范围 | 样本量超出范围，借903-209 | 重新插拔HCT管路并清洁模块后修复 by jesson | SWO-330452 | Defect | HW | Aspiration Module | Needle Wash Unit |  | Sample size out of range | Sample size out of range, borrowed 903-209 | Repair by jesson after re-plugging the HCT line and cleaning the module
No | Visit | Complaint | RDEU | Leckage erkannt | Leckage erkannt, Pumpenkalibrierung Fehler | BGA Manuell gespült, Pinch Valve erneuert, LP und SC erneuert, Kalibrationen und QC's vorgenommen, Blut Messung vorgenommen, BGA ok. | SWO-331028 | Defect | HW | Oxi module | Valve, pinch valve |  | Leakage detected | Leakage detected, pump calibration error | BGA Manually rinsed, Pinch Valve renewed, LP and SC renewed, calibrations and QC made, blood measurement made, BGA ok.
No | Visit | Complaint | RCHN | 泵定标错误 | 泵定标错误，需要借件902-939 | 更换902-939后修复 by tim | SWO-329595 | Defect | HW | Basic Unit | Pump rotor |  | Pump calibration error | Pump calibration error, need to borrow 902-939 | Repair by tim after replacing 902-939
Yes | Visit | Complaint | RCHN | 报错1024 | 报错1024，借903-364和944-212（14046）、944-258（14051）各一个 | 更换903-364和944-212（14046）、944-258（14051）和重新定标后修复 by jesson | SWO-330276 | [5 empty] | 1024 error | 1024, borrow 903-364 and 944-212 (14046), 944-258 (14051) each | Replace 903-364 and 944-212 (14046), 944-258 (14051) and repair by jesson after recalibration
No | Visit | Complaint | RCHN | 报错1999.35 | 报错1999.35 | 更换试剂包后修复 | SWO-331136 | Assisted Replacement | Consumables | SP | Chip coding | 1999 | Error 1999.35 | Error 1999.35 | Repair after replacement of reagent kit
No | Visit | Complaint | RCHN | CA，PH定标失败 0378 | CA，PH定标失败 0378，借902-828一个。 | 更换902-828备件后修复 by jesson | SWO-330286 | Revisit required | HW | pH/BG Module | PCB, ph/BG | 378 | CA, PH calibration failure 0378 | CA, PH calibration failure 0378, borrow 902-828 one. | Repair by jesson after replacement of 902-828 spare parts
No | Visit | Complaint | RAME | Jessica reported multiple errors related to fluid Flow, clot detection, and cassette temperature errors. | Client had replaced sensor cassette after high temperature reported, new sensor cassette installed, not resolving issue, unable to perform flush due to obstruction. | Disassembled analyzer, verifies all fluid paths. Paths clear. Reinstalled sensor cassette and unable to perform back flush. Replaced sensor cassette with a 2nd new cassette and analyzer is able to pass fluid properly. Startup and calibrations performed. | SWO-330844 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway | Clot | Jessica reported multiple errors related to fluid Flow, clot detection, and cassette temperature errors. | Client had replaced sensor cassette after high temperature reported, new sensor cassette installed, not resolving issue, unable to perform flush due to obstruction. | Disassembled analyzer, verifies all fluid paths. Paths clear. Reinstalled sensor cassette and unable to perform back flush. Replaced sensor cassette with a 2nd new cassette and analyzer is able to pass fluid properly. Startup and calibrations performed.
No | Visit | Complaint | RKK | 画面が真っ暗。代替機設置。引上げ確認。 | 画面が真っ暗。 | 代替機設置。設定入力。 | SWO-337251 | Revisit required | HW | DMS_PC | Display |  | The screen is pitch black. Installation of alternative machines. Raise confirmation. | The screen is pitch black. | Installation of alternative machines. Setting input.
No | Visit | Complaint | RAME | Customer getting oxi errors on patient sample error codes are 292, 581, 688 | Found that customer is getting 292, 581, and 688 intermittently on patients. | back flushed the analyzer, did complete oxi test, calibrations, and thb cal. | SWO-330775 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | Customer getting oxi errors on patient sample error codes are 292, 581, 688 | Found that customer is getting 292, 581, and 688 intermittently on patients. | back flushed the analyzer, did complete oxi test, calibrations, and thb cal.
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | NA,PH,PCO2,CL定标不过 | 更换测试卡后修复 | SWO-334669 | [5 empty] | Test card, reagent package failure | NA, PH, PCO2, CL calibration but | Repair after replacing test card
No | Visit | Complaint | RCHN | 报错4548 | 报错4548，借903-028一个 |  | SWO-330676 | Defect | HW | Cover and Chassis | PCB, Cabinet top | 4548 | 4548 error | 4548, loan 903-028 one
Yes | Visit | Complaint | RCHN | 报错1029 | 报错1029，借903-364一个 | 更换903-364后修复 by jesson | SWO-330587 | Defect | HW | Process Module | Optical Unit | 1029 | 1029 error | 1029, loan 903-364 one | Repair by jesson after replacing 903-364
No | Visit | Complaint | RESP | FLEXQ BLOQUEADO. | RUIDO AL DESPLAZARSE Y BLOQUEO TRAS EL MOVIMIENTO. | SE DETECTA EJE DE MOTOR DE TRASLACION EN MAL ESTADO, DANDO PROBLEMAS EN DICHO MOVIMIENTO. SE SUSTITUYE Y SE EFECTUAN AJUSTE DE TODO EL MODULOS ASI COMO PRUEBAS DE FUNCIONAMIENTO. | SWO-330806 | Defect | HW | FlexQ Module | Spindle |  | FLEXQ LOCKED. | NOISE WHEN DISLOCATED AND LOCKING AFTER MOVEMENT. | TRANSLATION MOTOR AXLE IS DETECTED IN POOR CONDITION, GIVING PROBLEMS IN THIS MOVEMENT. ALL MODULES ARE SUBSTITUTE AND PROVIDED AS OPERATION TESTS.
No | Visit | Complaint | RCHN | CO2质控偏低 | CO2质控偏低，借902-823,945-616各一个 |  | SWO-330634 | Defect | HW | pH/BG Module | PCB, ph/BG |  | Low CO2 quality control | Low CO2 quality control, borrowing 902-823, 945-616 each
No | Visit | Complaint | RKK | tHB CALドリフト多発 | 1日数回ドリフトエラー | ヘモライザーユニット・ピンチバルブ交換。 | SWO-330702 | Defect | HW | Oxi module | Valve, pinch valve |  | tHB CAL drift frequent | Drift error several times a day | Hemolyzer unit and pinch valve replacement.
Yes | 3 | Y | RPOL | Blank | Sprawdzenie aparatu. Aktualizacja oprogramowania 3.3MR1. Analizator sprawny technicznie. |  | D42003E85A3A83D7 | Re classified RMED | SWU |  |  |  | Blank | Check the camera. Software Update 3.3 MR1. Technically efficient analyzer.
No | Visit | Complaint | RLTD | Please call Lucy or Tracy on 01604 545 520 | Failing to aspirate sample | Failing to aspirate sample, couldn’t see any issues with fluid path, inlet and probe ok, passing all fluid transport tests, cleaned inlet and probe, tested pinch valve. reassembled, now working ok. | SWO-332391 | No fault found | HW | Auto Inlet module | Auto Inlet module |  | Please call Lucy or Tracy on 01604 545 520 | Failing to aspirate sample | Failing to aspirate sample, couldn’t see any issues with fluid path, inlet and probe ok, passing all fluid transport tests, cleaned inlet and probe, tested pinch valve. reassembled, now working ok.
No | Visit | Complaint | RAME | The ABL80 is leaking from below the inlet of the sensor cassette. | Checked the data logs. | Flushed all the tubes. | SWO-331383 | [5 empty] | The ABL80 is leaking from below the inlet of the sensor cassette. | Checked the data logs. | Flushed all the tubes.
Yes | Visit | Complaint | RIND | Calibration failed for Cl | Checked calibration data and found Cl failed | Flushed SC with NS solution, activated Cl parameter, 2 point calibration performed, still Cl failed. | SWO-331700 | [5 empty] | Calibration failed for Cl | Checked calibration data and found Cl failed | Flushed SC with NS solution, activated Cl parameter, 2 point calibration performed, still Cl failed.
No | Visit | Complaint | RCHN | k na ca cl定标失败0375、0378 | k na ca cl定标失败0375、0378,借945-603一个 |  | SWO-330967 | Assisted Replacement | Electrode | Electrode | Reference | See SWO-330286 | k na cl calibration failure 0375, 0378 | k na cl calibration failure 0375, 0378, borrow945-603 one
No | Visit | Complaint | RAME | Multiple parameter failed calibration drift and status failed. | Na,glu,lac errors | Found electrode module pump tubing with holes in black section,ph and el met chambers | SWO-331467 | Defect | HW | pH/BG Module | Pump tubes | black tubes | Multiple parameter failed calibration drift and status failed. | Na,glu,lac errors | Found electrode module pump tubing with holes in black section,ph and el met chambers
No | Visit | Complaint | RIND | Query in patient result. | Check the patient result & found the query in pH. | Check the calibration log & patient result & found that the query in the pH in the patient samples only,clean the pH electrode chamber & flush the pH $BG module path,run the protein removal & cleaning few times,after that run few patient sample & got the satisfactory result. | SWO-333349 | Cleaned | Fluidic pathway | pH/BG Module | Measuring chamber |  | Query in patient result. | Check the patient result & found the query in pH. | Check the calibration log & patient result & found that the query in the pH in the patient samples only,clean the pH electrode chamber & flush the pH $BG module path,run the protein removal & cleaning few times,after that run few patient sample & got the satisfactory result.
Yes | Visit | Complaint | RAME | Lac drift | Customer reports that their machine is not running on package insert provided mean. I called tech support to get Peer 2 mean, but this is a brand new lot and there is no data yet. | I visited the University of Utah next door and borrowed a different lot number of LAC membrane. I replaced membrane and performed all necessary calibrations. Lactate appeared to come up somewhat from what it had been running. I suggest running and seeing where peer group performs over the rest of the month. | SWO-331403 | Re classified RMED | Courtesy | peer comparison |  |  | Lac drift | Customer reports that their machine is not running on package insert provided mean. I called tech support to get Peer 2 mean, but this is a brand new lot and there is no data yet. | I visited the University of Utah next door and borrowed a different lot number of LAC membrane. I replaced membrane and performed all necessary calibrations. Lactate appeared to come up somewhat from what it had been running. I suggest running and seeing where peer group performs over the rest of the month.
No | Visit | Complaint | RBNL | TnT QC fout | zie SWO 330324 | Vervanging van Process Wheel, test met blanco stalen : ok. Calibratie van TnT + controle : OK | SWO-331310 | Defect | HW | Process Module | Process Wheel |  | TnT QC fout | zie SWO 330324 | Replacement of Process Wheel, test with blank swatches: OK. Calibration of TnT + control: OK
No | Visit | Complaint | RAME | Battery Issues | The analyzer is presenting message 1559 battery error. | Checked the power supply voltages, found ok. Checked the battery cells, found one of the cells not charging properly. replaced a new battery. Checked and, confirmed in service mode that all cells are properly charging. | SWO-332150 | Defect | HW | Battery module | Battery module |  | Battery Issues | The analyzer is presenting message 1559 battery error. | Checked the power supply voltages, found ok. Checked the battery cells, found one of the cells not charging properly. replaced a new battery. Checked and, confirmed in service mode that all cells are properly charging.
No | Visit | Complaint | RDEU | 04.04.2018, 11:26 Uhr, Fr. Gühring, Tel. 0232349892256: Endlosschleife LP erneuern | System überprüft | Einlassdichtung ersetzt und Nadel positioniert. | SWO-332193 | Defect | HW | Inlet module | Inlet gasket |  | 04.04.2018, 11:26 a.m., Ms. Gühring, tel. 0232349892256: Renew endless loop LP | System checked | inlet seal and needle positioned.
Yes | Visit | Complaint | RDEU | 04.04.2018, 08:39 Uhr, Fr. Wunsch, Tel. 040790202398: Kein Einlassdichtungshalter vorrätig | Keine Einlassdichtung | Vorhandene Einlassdichtung gereinigt und eingesetzt. Analysator gespült. Analysator geht in den normalen Startbildschirm. | SWO-331769 | Re classified RMED | Logistical SWO |  |  |  | 04.04.2018, 08:39, Fr. Wunsch, Tel. 040790202398: No inlet seal holder in stock | No inlet seal | Existing inlet seal cleaned and inserted. Analyzer rinsed. Analyzer goes to the normal home screen.
No | Visit | Complaint | RSAF | Leak current detected in relation to the solution pack | Leak current in relation to the solution pack. | Replaced with a new sensor cassette. Perfomed calibtation successfully | SWO-337445 | Assisted Replacement | SC | SC | Leak Current |  | Leak current detected in relation to the solution pack | Leak current in relation to the solution pack. | Replaced with a new sensor cassette. Perfomed calibtation successfully
No | Visit | Complaint | RCHN | 做样本量的时候机器经常死机 | 做样本量的时候机器经常死机,需要借件933-364（已被933-441替代） |  | SWO-331703 | [5 empty] | Machines often crash when you take sample stakes. | When taking sample size, the machine often crashes and needs to borrow 933-364 (has been replaced by 933-441)
No | Visit | Complaint | RCHN | 新装机之后发现定标不稳定，有腐蚀 | 新装机之后发现定标不稳定，有腐蚀，需要借件902-972 | 更换902-972一个正常 。 by jesson | SWO-331139 | Defect | HW | Basic Unit | PCB, Sensor connector assembly |  | After the new installation found that the calibration is unstable, there is corrosion | After the new installation found that the calibration is unstable, corrosion, need to borrow 902-972 | Replace 902-972 one normal. by jesson
No | Visit | Complaint | RESP | Sensor no calibra | Sustitución de membranas Chequeo de sensor | Sustitución de sensor | SWO-332084 | [5 empty] | Sensor in Calibra | Membrane replacement Sensor check | Sensor replacement
No | Visit | Complaint | RCHN | 温度超范围 | 温度超范围 | 更换903-581后修复 by Andrew | SWO-337280 | Defect | HW | Basic Unit | PCB, Sensor interface |  | Temperature over range | Temperature over range | Repair by Andrew after replacement 903-581
No | Visit | Complaint | RESP | Error 1066 Temperatura fuera de rango en sustrato de casete de electrodos | Chequeo de cables, conexiones y pcbs Chequeo de datos en menú de servicio | Sustitución de sensor interface pcb Sustitución de wet section control Pcb Reinstalación en 3 ocasiones de software | SWO-332085 | Defect | HW | Basic Unit | PCB, Sensor interface |  | Error 1066 Out-of-range temperature on electrode cassette substrate | Checking cables, connections and pcbs Checking data in service menu | Replacing sensor interface pcb Replacing wet section control Pcb 3-time software reinstallation
No | Visit | Complaint | RCHN | 报错2168 | 报错2168，借903-138一个 |  | SWO-331726 | Defect | HW | Incubator Module | PCB, Incubator | 2168 | 2168 error | 2168, loan 903-138 one
No | Visit | Complaint | RCHN | 报错5784 | 报错5784 | 更换测试卡后修复 | SWO-338043 | Defect | Consumables | Test Kit | Cartridge | 5784 | 5784 | 5784 | Repair after replacing test card
No | Visit | Complaint | RLTD | The screen is blank. It appears to be a lighter black screen with a white surrounding and no sound. The Analyser has been rebooted a few times, but no improvement has been made. Mori - 01865 221596 | The analyser is displaying a black screen with a white surrounding it has been rebooted several times and this has had no effect, suspect a faulty CPU. | A new CPU was fitted and the analyser turned on and it powered up without any problems. When the 2 point calibration was run pO2 failed with sensitivity and zero errors, the electrode was checked and there was no fluid in the membrane so a new one was fitted. | SWO-331790 | Defect | HW | DMS_PC | CPU |  | The screen is blank. It appears to be a lighter black screen with a white surrounding and no sound. The Analyser has been rebooted a few times, but no improvement has been made. Mori - 01865 221596 | The analyser is displaying a black screen with a white surrounding it has been rebooted several times and this has had no effect, suspect a faulty CPU. | A new CPU was fitted and the analyser turned on and it powered up without any problems. When the 2 point calibration was run pO2 failed with sensitivity and zero errors, the electrode was checked and there was no fluid in the membrane so a new one was fitted.
No | Visit | Complaint | RCHN | 打印机打印模糊 | 打印机打印模糊，借910-275一个 | 更换910-275一个正常。by jesson | SWO-331094 | Defect | HW | Printer Module | Printer Module |  | Printer print blur | Printer print blur, borrow 910-275 one | Replace 910-275 one normal. by jesson
No | Visit | Complaint | RDEU | Glucose zeigt seit längerem hohe Driften | Glucose werte nach membranierung und mehrmaliger Kalibrierung ok, aber immer wieder mit hohen Driften. | Glucose Elektrode gewechselt kalibriert und QC durchgeführt Garantie Reparatur! | SWO-331748 | Assisted Replacement | Electrode | Electrode | GLU |  | Glucose has been showing high drifts for a long time | Glucose values ok after membraneation and repeated calibration, but again and again with high drifts. | Glucose electrode changed calibrated and QC performed Warranty repair!
No | Visit | Complaint | RDEU | Bildschirm reagiert verzögert. Alle Aktivitäten sind verlangsamt | Beim Neustart, Dosfenster wird geöffnet und zu Ende geführt. Applikationssoftware wird nicht mehr angezeigt. AU läuft. | Softwareupdate bringt keine Veränderung. CPU- Tausch und Software installation bringt endgültig eine Verbesserung. Für die sporadische Touchscreen- Ausfälle sorgen sie für eine lokale Verbesserung der Lichtverhältnisse. | SWO-332278 | Defect | SW | RMED SW | RMED SW | slow response | Screen responds delayed. All activities are slowed down | When restarting, Dos window opens and ends. Application software is no longer displayed. AU is running. | Software update does not bring any change. CPU exchange and software installation finally brings an improvement. For the sporadic touchscreen failures, they ensure a local improvement of the lighting conditions.
No | Visit | Complaint | RAME | Autocheck runs in homogeneous samples | Found 14 pole cable shield off , interfering with operation | Remove and reattach shield | SWO-332146 | Defect | HW | Autocheck Module | FPC Cable |  | Autocheck runs in homogeneous samples | Found 14 pole cable shield off , interfering with operation | Remove and reattach shield
No | Visit | Complaint | RKOR | QC B's Oximetry values are out of range | Perform tHb calibration | Perform tHb calibration | SWO-333382 | Adjustment | HW | Oxi module | tHb cal |  | QC B's Oximetry values are out of range | Perform tHb calibration | Perform tHb calibration
No | Visit | Complaint | RAME | Battery not charging been plugged over night and stayed at 0% | Upon arrival, found Battery not charging. | Replaced Battery on unit. Checked for proper charging. | SWO-332558 | Defect | HW | Battery module | Battery module |  | Battery not charging been plugged over night and stayed at 0% | Upon arrival, found Battery not charging. | Replaced Battery on unit. Checked for proper charging.
No | Visit | Complaint | RAME | Unit is contantly failing tHb calibrations on internal calibrations, customer has replaced solution packs and performed tHb cals with no success. | Received parts in and reviews data logs and found continuous QC failure on oxi parameters. | Replaced hemolyzer. | SWO-332591 | Defect | HW | Oxi module | Hemolyzer |  | Unit is contantly failing tHb calibrations on internal calibrations, customer has replaced solution packs and performed tHb cals with no success. | Received parts in and reviews data logs and found continuous QC failure on oxi parameters. | Replaced hemolyzer.
No | Visit | Complaint | RAME | Change time to military time | Customer requested time be changed to reflect military time with no seconds. | Followed Radiometer procedure to change time stamp in Windows | SWO-339269 | Setup | SW | Time setting | Time setting | military time | Change time to military time | Customer requested time be changed to reflect military time with no seconds. | Followed Radiometer procedure to change time stamp in Windows
No | Visit | Complaint | RAME | Change Time | Customer requested clock change to military time | Followed Radiometer procedure for changing time in windows settings. | SWO-339265 | Setup | SW | Time setting | Time setting | military time | Change Time | Customer requested clock change to military time | Followed Radiometer procedure for changing time in windows settings.
No | Visit | Complaint | RFRA | Fuite de courant en liaison avec le pack de solution | Recherche erreurs dans historique activités Contrôle circuits fluides et consommables | Remplacer pack de solution Tests programme service Remplacer sensor K7 | SWO-332598 | Assisted Replacement | SC | SC | Leak Current |  | Power leak in conjunction with solution pack | Searching for errors in historical activities Fluid and consumable circuit control | Replace solution pack Service program tests Replace sensor K7
No | Visit | Complaint | RAME | During SC installation, error code 1305 is displayed, and the analyzer shuts down. | Found corrosion on the sensor board |  | SWO-332607 | [5 empty] | During SC installation, error code 1305 is displayed, and the analyzer shuts down. | Found corrosion on the sensor board
Yes | Visit | Complaint | RCHN | 报错 1031 | 报错 1031，借903-364一个 | 更换903-364一个正常。by tim | SWO-331752 | [5 empty] | Error 1031 | Error 1031, loan 903-364 one | Replace 903-364 one normal. by tim
No | Visit | Complaint | RCHN | 液路故障 3412 | 液路故障 3412 ，借903-112一个。 |  | SWO-332296 | Defect | HW | Wet section | PCB, Wet section | 3412 | Liquid Road Fault 3412 | Liquid Road Fault 3412 , borrow edith of 903-112.
Yes | 3 | Y | RCZE | Blank |  |  | D422C4C64632B562 | [5 empty] | Blank
Yes | 3 | Y | RCZE | Blank |  |  | D422C5FD4650B567 | Re classified RMED | FAN |  |  |  | Blank
No | 22 | Y | RPOL | Blank | Wymiana ig³y. Analizator sprawny technicznie. |  | D42CA823F6381AD6 | Defect | HW | Inlet module | Probe bend |  | Blank | Replacing the needle. Technically efficient analyzer.
No | Visit | Complaint | RSCH | Immer wieder fehler 0408: Lac Nullstromfehler | Immer wieder Fehler 0408: Lac Nullstromfehler | Allgemeine Kontrolle ok. Alle Ventile ausgebaut und gereinigt ok. Allle Membranen ersetzt ok Dual Valve Housing inkl. Dichtung ers EL $MET Modul. Alle Schläuche wo von Werk aus ausgetauscht ok. System i.o. Kal. + QC i.o. | SWO-332400 | Defect | HW | El/Met Module | Valve, membrane |  | Time and again error 0408: Lac zero current error | Time and again error 0408: Lac zero current error | General control ok. All valves removed and cleaned ok. Allle membranes replaces ok Dual Valve Housing incl. seal ers EL $MET module. All hoses where replaced from factory ok. System i.o. Kal. + QC i.o.
No | Visit | Complaint | RSCH | Rep Inhomogene Messungen | Viele Messungen Inhomogen | Referenz Membrane ausgetauscht (Messkammer pH und Ref voll mit Blut durch undichte Membranen Lot2213. Funktionskontrolle i.o | SWO-332466 | Revisit required | Electrode | Membrane | Reference | see SWO-336193 | Rep Inhomogeneous Measurements | Many measurements Inhomogeneous | Reference membrane replaced (measuring chamber pH and ref full with blood by leaking membranes Lot2213. Functional control i.o
No | Visit | Complaint | RSCH | Rep | Barcodereader liest Barcodes nicht korrekt auf beiden neuen AQT Keine Einstellungen möglich | Barcodereader und PCB alte Version eingebaut. Funktionskontrolle i.o Es wird eine Techline eröffnet. Weiteres vorgehen wird bei uns intern besprochen | SWO-332446 | Defect | HW | Aspiration Module | Barcode reader, Inlet | + PCB // Old Version | Rep | Barcode reader does not read barcodes correctly on both new AQT No settings possible | Barcode reader and PCB old version built in. Functional control i.o A techline will be opened. Further action is discussed internally
No | Visit | Complaint | RLTD | Taqiyah 07956903944 | Checked data logs - Oximetry quality control results made a step change after analyser was out of action due to failed sensor cassette between 30 $03 and 04 $04, but unable to establish a cause. Performed manual flush of the sample pathway to ensure pathway was clear - nothing was observed to be removed. Carried out complete oximetry test in service - passed. Checked oximetry pinch valve tube & pinch valve - both ok. Performed 2 calibrations, then carried out fresh tHb calibration. Repeated quality controls - all parameters now mid range. Analyser is ready for use. | Performed manual flush & tHb calibration. | SWO-332831 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | Taqiyah 07956903944 | Checked data logs - Oximetry quality control results made a step change after analyser was out of action due to failed sensor cassette between 30 /03 and 04 /04, but unable to establish a cause. Performed manual flush of the sample pathway to ensure pathway was clear - nothing was observed to be removed. Carried out complete oximetry test in service - passed. Checked oximetry pinch valve tube & pinch valve - both ok. Performed 2 calibrations, then carried out fresh tHb calibration. Repeated quality controls - all parameters now mid range. Analyser is ready for use. | Performed manual flush & tHb calibration.
No | Visit | Complaint | RFRA | Erreur vide test de pression | Vérification fluidique | Nettoyage circuit fluides Remplacer sonde aspirations (cintrée) et joint introduction | SWO-333105 | Defect | HW | Auto Inlet module | Probe bend |  | Empty pressure test error | Fluid check | Fluid circuit cleaning Replace aspirations probe (curved) and joint introduction
No | Visit | Complaint | RLTD | Anna 017084352621 | Checked data logs and service programs. Clot suspected in Inlet, OXI module and sensor cassette. | Removed SP and checked it is activated correctly. Removed gasket, checked and reseated. Opened the sensor bay, checked valve and is operating correctly. Checked flows in service and ran clot detection program. | SWO-333021 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | Anna 017084352621 | Checked data logs and service programs. Clot suspected in Inlet, OXI module and sensor cassette. | Removed SP and checked it is activated correctly. Removed gasket, checked and reseated. Opened the sensor bay, checked valve and is operating correctly. Checked flows in service and ran clot detection program.
No | Visit | Complaint | RFRA | Sauvegarde des données | ... | Sauvegarde et suppression des données Formation de la technicienne sur les sauvegardes | SWO-333058 | [5 empty] | Backing up data | ... | Backing up and deleting data Technician training on backups
No | Visit | Complaint | RDEU | 03.04.2018, 10:44 Uhr, Pf. Thomas, Tel. 06213831024: AutoCheck Ampullen werden nicht gelesen 06.04.2018 8:00 H. Küchelbacher 06213831024: Ampullen werden weiterhin nicht gelesen | Qualitätskontroll Ampullen werden nicht erkannt. | Ausbau des Autocheck Moduls und gründliche Reinigung des Moduls von Staub. Reinigung auch des Laserspiegels. Nach Einbau und Neustart werden alle Qualitätskontroll Ampullen erkannt. | SWO-333097 | Cleaned | HW | Autocheck Module | Mirror | dust | 03.04.2018, 10:44 a.m., Pf. Thomas, tel. 06213831024: AutoCheck ampoules are not read 06.04.2018 8:00 H. Küchelbacher 06213831024: Ampoules are still not read | Quality control ampoules are not detected. | Removal of the Autocheck module and thorough cleaning of the module of dust. Cleaning also of the laser mirror. After installation and restart, all quality control ampoules are detected.
No | Visit | Complaint | RAME | There is a blue line running down the left side of the monitor from top to bottom. | Verified blue line on the left side of monitor. Attempted to replace touch screen from parts ordered. After dismantling unit I found that the screen I was sent is the wrong type. Will need to reassemble and order a different touch screen type and return to replace. | Will order new touch screen for replacement. | SWO-333230 | Revisit required | HW | DMS_PC | Touch screen |  | There is a blue line running down the left side of the monitor from top to bottom. | Verified blue line on the left side of monitor. Attempted to replace touch screen from parts ordered. After dismantling unit I found that the screen I was sent is the wrong type. Will need to reassemble and order a different touch screen type and return to replace. | Will order new touch screen for replacement.
No | Visit | Complaint | RLTD | Harvey 01619226498 | Arrived on site and logged onto the analyser checked logs and found error messages relating to set safe and rotate storage wheel. Analyser had attempted automatic recovery and eventually succeeded. Customer had been using the analyser with the test cartridge already on board to run samples. Analyser had therefor been working. | Entered service and checked all safe positions could be met which they did. I opened the front Cover to the analyser and ran storage wheel checks more it oring for raised cartridges and or collisions from cartridge with other analyser parts. Cartridges all behaved correctly and no collisions seen or heard. Removed cover to storage wheel and inspected the whole wheel for blockages $debris in path. None found. Manually removed all cartridges to inspect for broken cartridge tabs. All cartridges intact. Took wheel hold off and manually turned wheel by hand to feel for obstructions. None found. Inspected all cartridge dividers for any loose or high sitting dividers. All in good secure position. Refitted cartridges after each inspection. Cleaned internals and sensors to remove any dust that may be on the sensors. Refitted all parts. Inspected the rest of the analyser and mechanics for signs of damage or blockage. None found. Removed solution pack and I specter pack and bay for stray test cups. None found. Refitted all and set safe positions before restarting the analyser. Analyser started correctly with only a message stating a cleaning cycle needed to be run. Ran system check followed by two cleaning cycles. All completed with no errors. Collected Ted and inserted two new test cartridges to test the installation Process. Cartridges installed without issue. Suspect the cartridge installed that caused the error was I served before analyser was ready to receive the cartridge. Ran qc to test fluid and cup transport with cartridges installed. Qc ran and passed. Performed verification service on the analyser whilst on site. Cleaned external parts of the analyser and internals to remove all dust using antistatic vacuum. Inspected the needle and cup wash units for contamination. These were both clean. Cleaned cup sensor. Cleaned incurred action wheel and process wheel. Cleaned both barcode reader's. Ran verification checks. Analyser running well. | SWO-332967 | No fault found | No fault found | No fault found | No fault found |  | Harvey 01619226498 | Arrived on site and logged onto the analyser checked logs and found error messages relating to set safe and rotate storage wheel. Analyser had attempted automatic recovery and eventually succeeded. Customer had been using the analyser with the test cartridge already on board to run samples. Analyser had therefor been working. | Entered service and checked all safe positions could be met which they did. I opened the front Cover to the analyser and ran storage wheel checks more it oring for raised cartridges and or collisions from cartridge with other analyser parts. Cartridges all behaved correctly and no collisions seen or heard. Removed cover to storage wheel and inspected the whole wheel for blockages $debris in path. None found. Manually removed all cartridges to inspect for broken cartridge tabs. All cartridges intact. Took wheel hold off and manually turned wheel by hand to feel for obstructions. None found. Inspected all cartridge dividers for any loose or high sitting dividers. All in good secure position. Refitted cartridges after each inspection. Cleaned internals and sensors to remove any dust that may be on the sensors. Refitted all parts. Inspected the rest of the analyser and mechanics for signs of damage or blockage. None found. Removed solution pack and I specter pack and bay for stray test cups. None found. Refitted all and set safe positions before restarting the analyser. Analyser started correctly with only a message stating a cleaning cycle needed to be run. Ran system check followed by two cleaning cycles. All completed with no errors. Collected Ted and inserted two new test cartridges to test the installation Process. Cartridges installed without issue. Suspect the cartridge installed that caused the error was I served before analyser was ready to receive the cartridge. Ran qc to test fluid and cup transport with cartridges installed. Qc ran and passed. Performed verification service on the analyser whilst on site. Cleaned external parts of the analyser and internals to remove all dust using antistatic vacuum. Inspected the needle and cup wash units for contamination. These were both clean. Cleaned cup sensor. Cleaned incurred action wheel and process wheel. Cleaned both barcode reader's. Ran verification checks. Analyser running well.
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | LAC灵敏度低 | 更换测试卡正常 | SWO-333247 | [5 empty] | Test card, reagent package failure | Low LAC sensitivity | Replace the test card ok
No | Visit | Complaint | RCHN | 注射器老化，密封性不好有气泡 | 注射器老化，密封性不好有气泡 ，借958-201一个。 |  | SWO-333245 | Defect | HW | Aspiration Module | Cavro Pump | Syringe | Syringe aging, poor sealing with bubbles | Syringe aging, poor sealing with bubbles , borrow edith 958-201.
No | Visit | Complaint | RIND | Oximetry values is not coming in patients reports, Showing error oxi hardware problem not possible to measure. | Oximetry values is not coming in patient reports. | Check the machine, run complete oximetry test found neon number too low and neon intensity error. Check the status of the lamp unit found neon lamp is not working. Then we replace the neon lamp from old machine in office and run complete oximetry test now its working fine all the test are passed. Run 2 point calibration and samples successfully. | SWO-333337 | Defect | HW | Oxi module | Lamp unit | Neon lamp | Oximetry values is not coming in patients reports, Showing error oxi hardware problem not possible to measure. | Oximetry values is not coming in patient reports. | Check the machine, run complete oximetry test found neon number too low and neon intensity error. Check the status of the lamp unit found neon lamp is not working. Then we replace the neon lamp from old machine in office and run complete oximetry test now its working fine all the test are passed. Run 2 point calibration and samples successfully.
Yes | Visit | Complaint | RAME | Customer call in. Customer received a SC with an expiry date of 1 month. This cassette was the same expiry date as the original date of the original cassette. Inventory problem. | A sensor cassette was relocated picked up on Saturday and delivered and installed at the account on Sunday. | The sensor cassette was installed | SWO-334632 | Re classified RMED | Logistical SWO |  |  |  | Customer call in. Customer received a SC with an expiry date of 1 month. This cassette was the same expiry date as the original date of the original cassette. Inventory problem. | A sensor cassette was relocated picked up on Saturday and delivered and installed at the account on Sunday. | The sensor cassette was installed
No | Visit | Complaint | RSAF | PO2 results reading lower than expected | Gluc and Lac still not stable Sensor simulator not showing errors on PO2 sensor | Checks done as per service tip, replaced sensor interface PCB and sensor cassette. Simulator test ok, | SWO-339563 | Defect | HW | Basic Unit | PCB, Sensor interface | Glu/Lac | PO2 results reading lower than expected | Gluc and Lac still not stable Sensor simulator not showing errors on PO2 sensor | Checks done as per service tip, replaced sensor interface PCB and sensor cassette. Simulator test ok,
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 温度错误，GLU定标失败，1065，1310 | 更换测试卡后修复 | SWO-334616 | Assisted Replacement | SC | SC | glu |  | Test card, reagent package failure | Temperature error, GLU calibration failure, 1065, 1310 | Repair after replacing test card
No | Visit | Complaint | RCHN | 报错1068 | 报错1068，借903-210一个。 | 更换903-210后修复by tim | SWO-333246 | Defect | HW | Process Module | PCB, Cup dryer | 1068 | 1068 error | 1068, borrow 903-210 one. | Repair by tim after replacement 903-210
No | Visit | Complaint | RCHN | 打印机打印不清晰 | 打印机打印不清晰，借910-275一个 |  | SWO-333392 | Defect | HW | Printer Module | Printer Module |  | Printer printing is not clear | Printer printing is not clear, borrow 910-275 one
No | Visit | Complaint | RAME | Customer noticed after change of SC, pH is failing cal continuously. | Confirmed issue when onsite that pH has failed cal continuously. Customer has tried replacing SC $SP to no avail. | Replaced needle and parts in pm kits. Also found pinch valve tubing is worn. After that, pH is passing cal now. | SWO-334983 | Defect | HW | Basic Unit | PM kit | pinch valve tube | Customer noticed after change of SC, pH is failing cal continuously. | Confirmed issue when onsite that pH has failed cal continuously. Customer has tried replacing SC /SP to no avail. | Replaced needle and parts in pm kits. Also found pinch valve tubing is worn. After that, pH is passing cal now.
No | Visit | Complaint | RLTD | Touch screen unresponsive in areas. | Touch screen unresponsive in areas. | Replaced touch screen and calibrated it. | SWO-334913 | Defect | HW | Basic Unit | Touch screen |  | Touch screen unresponsive in areas. | Touch screen unresponsive in areas. | Replaced touch screen and calibrated it.
No | Visit | Complaint | RHUN | Fogyóanyag hiba | Fogyóanyag hiba | Pumpa rotor megszorítása, Gumikupak eltávolítása a szennyestartályból, Fogyóanyagcserék, Zéró beállítás, Kalibrációk, QC mérés *** A fogyóanyanyagok megérkezéséig átmenetileg Lac nélkül üzemel a készülék **** | SWO-334900 | [5 empty] | Fogyóanyag hiba | Fogyóanyag hiba | Pump rotor restriction, removal of rubber tires from the sewage system, consumables, zero setting, calibration, QC measurement *** Until the consumables arrive, the appliance is temporarily operated without LAC ****
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 失败-FC | 更换测试卡后修复 | SWO-338049 | [5 empty] | Test card, reagent package failure | Failure-FC | Repair after replacing test card
No | Visit | Complaint | RCHN | 报错1306 | 报错1306，借902-972,903-334各一个 |  | SWO-334698 | [5 empty] | 1306 error | 1306, lend ingress 902-972, 903-334 each
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | NA,CL定保标不过 1320,1321 | 更换试剂包后修复 | SWO-338084 | [5 empty] | Test card, reagent package failure | NA, CL warranty but 1320,1321 | Repair after replacement of reagent kit
No | Visit | Complaint | RDEU | pH Elektrode muss überprüft werden. |  | pH Elektrode mitgenommen um die Elektrode auf Funktion zu überprüfen. | SWO-335537 | Assisted Replacement | Electrode | Electrode | pH |  | pH electrode must be checked. |  | pH electrode to check the electrode for function.
No | Visit | Complaint | RDEU | Spritzen Einzug defekt. Messergebnisse nicht plausibel. Probleme beim Wechseln der Lösungsflaschen. | Fehlerhafte Etikettierung von Spüllösungsflaschen im Starterkit. | Barcode Laser überprüft, gereinigt. Lösungsflaschen überprüft. Andere Lösungsflaschen geholt, tHB Kal durchgeführt, Maßeinheit angepasst. Gerät kalibriert. | SWO-335538 | Assisted Replacement | Consumables | Solution Bottle | Rinse Solution | label defect | Syringe feed defective. measurement results are not plausible. Problems changing the solution bottles. | Incorrect labelling of rinsing solution bottles in the starter kit. | Barcode laser checked, cleaned. solution bottles. Other solution bottles fetched, tHB Kal performed, unit of measure adjusted. device calibrated.
No | Visit | Complaint | RDEU | Bestellung per Fax der CFM: Einlaßdichtungseinheit tauschen - Garantie | kein Fehler mehr vorhanden, tHb Kal fällig | tHb Kal durchgeführt, Kal und QC | SWO-335716 | Defect | HW | Inlet module | Inlet gasket |  | Order by fax of the CFM: Exchange inlet seal unit - warranty | no more error, tHb Kal due | tHb Kal durchgeführt, Kal und QC
No | Visit | Complaint | RAME | Liquid sensor error | Liquid sensor error | Replaced liquid sensors, flushed instrument. Instrument operating within manufacturing specifications | SWO-336510 | Defect | HW | Basic Unit | LS Sensor |  | Liquid sensor error | Liquid sensor error | Replaced liquid sensors, flushed instrument. Instrument operating within manufacturing specifications
No | Visit | Complaint | RAME | Unit stuck in startup screen | Analyzer stopped at boot screen. CPU has failed. | Replaced CPU and moved memory card from old CPU. Calibrated and ran QC instrument operating normally. | SWO-336322 | Defect | HW | DMS_PC | CPU |  | Unit stuck in startup screen | Analyzer stopped at boot screen. CPU has failed. | Replaced CPU and moved memory card from old CPU. Calibrated and ran QC instrument operating normally.
No | Visit | Complaint | RLTD | Rob Bolton 01772523007 | Site visit to in estimate and repair analyser. | Attended analyser and since last reboot that customer had performed, the issue had not returned. Error message that customer mentioned suggests fault with printer. Powered down and replaced printer and printer cable with new before rebooting analyser. Waited for analyser to restart, ran calibratiins and qcs before testing the printer. Analyser working well. Asked customer to monitor and update me if issue returns. | SWO-335734 | Defect | HW | Printer Module | Printer Module |  | Rob Bolton 01772523007 | Site visit to in estimate and repair analyser. | Attended analyser and since last reboot that customer had performed, the issue had not returned. Error message that customer mentioned suggests fault with printer. Powered down and replaced printer and printer cable with new before rebooting analyser. Waited for analyser to restart, ran calibratiins and qcs before testing the printer. Analyser working well. Asked customer to monitor and update me if issue returns.
No | Visit | Complaint | RDEU | Gerät instabil. Drucker fällt spontan aus und diverse Messfehler. | Garantie! Gerätetausch für Ostalbklinikum Aalen St. 29 Ki Int. vorbereitet. | Aufarbeitung des Demogerätes ABL90+ R0110N004 aus VSC_SW | SWO-336336 | Revisit required | HW | Printer Module | Printer Module |  | Device unstable. Printer fails spontaneously and various measurement errors. | Guarantee! Equipment exchange for Ostalbklinikum Aalen St. 29 Ki Int. Prepared. | Refurbishment of the demo device ABL90+ R0110N004 from VSC_SW
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 锂子结果不过 | 更换测试卡后修复 | SWO-336712 | [5 empty] | Test card, reagent package failure | Lithium results are no more than | Repair after replacing test card
No | Visit | Complaint | RKK | データが勝手に何度も送られてしまう | LISに同じ検体のデータが数件送られてきている。 | CFカード交換 933-419 様子見 | SWO-335486 | [5 empty] | Data is sent many times without permission | Several data from the same specimen have been sent to LIS. | CF card exchange 933-419 Wait
No | Visit | Complaint | RESP | pO2 | Comprobación de registros de datos del analizador. | Comprobación ajuste y limpieza de pines de la PCB sensor del equipo, se realizan varias calibraciones y QC. | SWO-335897 | No fault found | No fault found | No fault found | No fault found | No fault found | pO2 | Checking analyzer data records. | Checking adjustment and cleaning of the PCB sensor of the equipment, various calibrations and QC are performed.
No | Visit | Complaint | RCHN | 多个电解质结果漂移 | 多个电解质结果漂移，借945-603一个 | 更换945-603后修复 by Andrew | SWO-334800 | Assisted Replacement | Electrode | Electrode | Reference |  | Multiple electrolyte result drift | Multiple electrolyte result drift, borrow 945-603 a | Repair by Andrew after replacing 945-603
No | Visit | Complaint | RCHN | NA,CA故障 | NA,CA故障，借903-651一个 | 更换903-651后修复 by Andrew | SWO-334785 | Defect | HW | El/Met Module | PCB, El/Met |  | NA, CA fault | NA, CA fault, borrow 903-651 one | Fix by Andrew after replacing 903-651
No | 9 | Y | RCZE | Blank |  |  | D429054C94553773 | Defect | HW | pH/BG Module | LS Sensor |  | Blank
No | Visit | Complaint | RIND | Pco2 sensitivity error in 2 point calibration. | Check the machine found pco2 membrane required to change. | change the pco2 membrane and run 2 point calibration. Now machine is working fine. | SWO-337888 | Assisted Replacement | Electrode | Membrane | pCO2 |  | Pco2 sensitivity error in 2 point calibration. | Check the machine found pco2 membrane required to change. | change the pco2 membrane and run 2 point calibration. Now machine is working fine.
No | Visit | Complaint | RLTD | Mr Nooman Tel 0161 9226498 | Remove all covers and inspect analyser internally nothing obvious observed. Remove left hand cover from over storage wheel, restart analyser and observe all processes, analyser came back to ready. Ran three test samples and all went through ok without error. Check all cabling connections between all boards as one of the errors was 2912 checkaufailed, this points to 24V supply, all cables were ok. Placed another sample on analyser to run, it had only been on the analyser for 30 seconds and I was lifting the UPS back into position and the sample aborted, secured all the UPS connections and checked mains lead in back of analyser was secure and ran further sample moving UPS around to see if I could replicate. Ran a further 2 test samples which ran without error. | Had another UPS despatched to customer as this is possibly cause of the problem. Nathan Winterbottom to order a power supply in case the problem is called in again. | SWO-335757 | Defect | HW | UPS | UPS |  | Mr Nooman Tel 0161 9226498 | Remove all covers and inspect analyser internally nothing obvious observed. Remove left hand cover from over storage wheel, restart analyser and observe all processes, analyser came back to ready. Ran three test samples and all went through ok without error. Check all cabling connections between all boards as one of the errors was 2912 checkaufailed, this points to 24V supply, all cables were ok. Placed another sample on analyser to run, it had only been on the analyser for 30 seconds and I was lifting the UPS back into position and the sample aborted, secured all the UPS connections and checked mains lead in back of analyser was secure and ran further sample moving UPS around to see if I could replicate. Ran a further 2 test samples which ran without error. | Had another UPS despatched to customer as this is possibly cause of the problem. Nathan Winterbottom to order a power supply in case the problem is called in again.
No | Visit | Complaint | RBNL | Staal afgebroken | Stalen worden soms afgebroken, analyzer staat in fout met melding #3553 Incubator wheel move , het lijkt erop dat er geen cupje geplaatst wordt in de incubator of dat de sensor vervuild is waardoor het cupje niet wordt gedetecteerd. | Reiniging van cupdetector, cuptransport controle in service met 2 cupjes : OK | SWO-335656 | Cleaned | HW | Incubator Module | cup sensor | 3553 | Steel aborted | Samples are sometimes aborted, analyzer is in error with notification #3553 incubator wheel move, it seems that no chamber is placed in the incubator or that the sensor is polluted which prevents the chamber from being detected. | Cleaning Cupdetector, Cuptransport control in service with 2 cups: OK
No | Visit | Complaint | RFRA | Blocage intermittent et impossible d'installer le pack de solution ( blocage à l'introduction du pack ) | Vérification du punch up | Vérification dans le menu service du punch up nettoyage du moteur, commande du moteur si erreur réapparaît | SWO-335623 | Cleaned | HW | Incubator Module | sensor, punch up |  | Intermittent blocking and impossible to install the solution pack (blocking at the introduction of the pack) | Checking the punch up | Check in the service menu of the engine cleaning punch up, control the engine if error reappears
No | Visit | Complaint | RIND | Service Visit. | When press the introduce sample, inlet door is not opening and annoying sound comes. | Open the front panel of the analyzer and go to service program in which the inlet door option set safe, give the command open and close gear mechanism is not working properly, tighten the screw of inlet door which is just besides of gear mechanism, now again give the command open and close successfully worked. Close the analyzer and exit from the service program, Analyzer is ready for sample measurement, Analyzer handed over to staff in well working condition. | SWO-337178 | Adjustment | HW | Cover and Chassis | inlet door |  | Service Visit. | When press the introduce sample, inlet door is not opening and annoying sound comes. | Open the front panel of the analyzer and go to service program in which the inlet door option set safe, give the command open and close gear mechanism is not working properly, tighten the screw of inlet door which is just besides of gear mechanism, now again give the comman d open and close successfully worked. Close the analyzer and exit from the service program, Analyzer is ready for sample measurement, Analyzer handed over to staff in well working condition.
No | Visit | Complaint | RDEU | 11.4.2018/16:30 Herr Felske (Tel.: 015110840511) meldet nicht erkannte SP90 | SP wird nicht erkannt, Drucküberprüfung Vakuumfehler | Fehleranalyse, SP neu installiert, Einlassdichtung gereinigt, Servicetests durchgeführt, Fließwegkontrolle, Druck Prüfung sowie Vakuum Tests i. O. , Kalibrierung und Qualitätskontrolle i. O. | SWO-336374 | Cleaned | HW | Inlet module | Inlet gasket |  | 11.4.2018/16:30 Mr. Felske (Tel.: 015110840511) reports unrecognized SP90 | SP is not detected, pressure verification vacuum error | Fault analysis, SP reinstalled, inlet seal cleaned, service tests carried out, flow path control, pressure testing as well as vacuum tests in the form of calibration and quality control as a basis
Yes | 3 | Y | RPAC | Blank | Lactate calibration failing |  | D42A9C6B0A13D91D | [5 empty] | Blank | Lactate calibration failing
No | Visit | Complaint | RDEU | Lösungspack soll häufig gewechselt werden und Proben werden gar nicht/ nicht richtig angesaugt | Fehlersuche --> lt. Kunde verlangt das Gerät in den letzten 3 Tagen, das der Lösungspack gewechselt werden soll und das Blutmessungen nicht eingezogen werden bzw. die Einlassdichtung nicht ordentlich zurückfährt --> Dateiarchive (Aktivitäten, Kal., QC, Betriebsmittel), Einlassnadel $dichtung, Schlauchwege, Quetschventil --> Nadel $Dichtung, Quetschventil, Schlauchwege OK --> Betriebsmittel (Inhomogene Kal.3-Lösung (9.4., 08:30; 11.04. 8:30)) -> Lösungspack (BE-50) | Analysator im Service (Adjust Manifold, Pumpenkal., LS-Justage, Pressuretest, SP Check, Clot detect $remove) überprüft --> Servicetest OK, 25Testmessungen durchgeführt --> alle 25 Testmessungen in Ordnung, System und Funktion i.O. --> die Gerätebeauftragte Frau Tiliouine wird in regelmäßigen Abständen angerufen $besucht um sich zu erkundigen!!! --> Dem Kunden 1 x 944-197 SP90 (BE-52) vor Ort gelassen, da dieser zur Zeit keinen auf Lager hat!!! | SWO-336225 | Cleaned | HW | Auto Inlet module | Inlet gasket |  | Solution pack should be changed frequently and samples are not/ not sucked properly | Troubleshooting --> according to the customer, the device requires in the last 3 days that the solution pack should be changed and the blood measurements are not fed or the inlet seal does not return properly --> file archives (activities, cal., QC, equipment), Inlet /dichtung, hose ways, squeegee valve --> needle /Dichtung, squeegee valve, hose pathway OK --> equipment (inhomogeneous Kal.3 solution (9.4., 08:30; 11.04. 8:30)) -> Solution Pack (BE-50) | Analyzer in service (Adjust Manifold, Pump Scale, LS-Adjustment, Pressuretest, SP Check, Clot detect $remove) checked --> Service test OK, 25 test measurements performed --> all 25 test measurements in order, system and function i.O. --> the equipment officer Ms. Tiliouine is called at regular intervals $besucht to inquire!!! --> The customer 1 x 944-197 SP90 (BE-52) left on site, as he currently has none in stock!!!
No | Visit | Complaint | RDEU | Gerät instabil. Drucker fällt spontan aus und diverse Messfehler. | Gerätetausch wegen Instabilität des neuen ABL90+ | Demogerät installiert , Konfiguriert und Netzwerkverbindung hergestellt . Überprüfung des Datentransfer durch Labor und Station - o.k. | SWO-336335 | Defect | HW | Printer Module | Printer Module | SOFVO mail | Device unstable. Printer fails spontaneously and various measurement errors. | Device replacement due to instability of the new ABL90+ | Demo device installed, configured, and network connection established. Verification of data transfer by laboratory and station - ok
No | Visit | Complaint | RESP | Se ha perdido conexión a remoto | Se comprueba red. Configuracion correcta. Esta en red ya que rime del radiance detecta cuando esta apagado y cuando conecta. Comprobacion cortafuegos. | Parece un problema de puertos. Se solicita gestión por parte de informatica. Se queda a la espera. | SWO-337024 | [5 empty] | Lost remote connection | Network is checked. Correct configuration. It is networked as it rhymes the radiance detects when it is off and when it connects. Firewall check. | Looks like a port problem. Management is requested by informatica. He's waiting.
No | Visit | Complaint | RDEU | 10.04.2018, Frau Schulz, Tel: 04101305322: Datum stimmt nicht | System überprüft => Analog Board defekt | Teil nicht vor Ort muss bestellt werden , Kunde erhält Leihgerät | SWO-336078 | [5 empty] | 10.04.2018, Mrs Schulz, Tel: 04101305322: Date is not correct | System checked => Analog board defective | Part not on site must be ordered, customer receives rental device
No | Visit | Complaint | RCHN | 报错1999 | 报错1999，借944-220（14087）一张 | 重新定标后修复 by Andrew | SWO-334735 | Assisted Replacement | Consumables | Test Kit | Catridge, D-dimer | 1999 | Error 1999 | Error 1999, loan 944-220 (14087) one | Fix by Andrew after re-calibration
No | Visit | Complaint | RCHN | D-Dimmer定标数据丢失 | D-Dimmer定标数据丢失 ，借944-220（14087）一个。 | 重新定标后修复 by Andrew | SWO-334804 | User error | Consumables | Test Kit | Cartridge, D-Dimer |  | D-Dimmer calibration data loss | D-Dimmer calibration data loss , borrowed 944-220 (14087). | Fix by Andrew after re-calibration
No | Visit | Complaint | RIND | electrolytes unstable along with pH | checked and found that the chamber is broken. | changed the chamber from other analyzer now working after changing chamber from other analyzer. need to change the measuring chamber. | SWO-337833 | Defect | HW | pH/BG Module | Measuring chamber |  | electrolytes unstable along with pH | checked and found that the chamber is broken. | changed the chamber from other analyzer now working after changing chamber from other analyzer. need to change the measuring chamber.
Yes | 6 | Y | RPAC | Blank | software upgrade to v6.18 FAN915-374 |  | D42A0D120A183BC9 | Re classified RMED | Planned SWO |  |  |  | Blank | software upgrade to v6.18 FAN915-374
No | Visit | Complaint | RIND | Sample Aborted | Check the machine found sample is aborted. And showing error of inlet liquid sensor. | Check the machine open and clean the inlet liquid sensor, Reconnect the sensor but still found inlet liquid sensor error in LS adjust program. Than we swap the inlet liquid sensor with inlet lower liquid, after the swapping of liquid sensor error start showing in inlet lower liquid. After this machine is working fine. | SWO-336698 | Defect | HW | Inlet module | LS Sensor |  | Sample Aborted | Check the machine found sample is aborted. And showing error of inlet liquid sensor. | Check the machine open and clean the inlet liquid sensor, Reconnect the sensor but still found inlet liquid sensor error in LS adjust program. Than we swap the inlet liquid sensor with inlet lower liquid, after the swapping of liquid sensor error start showing in inlet lower liquid. After this machine is working fine.
Yes | Visit | Complaint | RAME | Failed cohb on thb cal twice | Lines missing from swo 335945 | Added lines | SWO-336360 | Re classified RMED | Courtesy |  |  |  | Failed cohb on thb cal twice | Lines missing from swo 335945 | Added lines
No | Visit | Complaint | RAME | Failed cohb on thb cal twice | Customer failed cohb twice Ran different lot and complete oxi in service as well as service check of thb cal all passing | Ran a different lot of thb passing thb cal and patient sample test | SWO-335945 | No fault found | No fault found | No fault found | No fault found |  | Failed cohb on thb cal twice | Customer failed cohb twice Ran different lot and complete oxi in service as well as service check of thb cal all passing | Ran a different lot of thb passing thb cal and patient sample test
No | Visit | Complaint | RAME | Halogen lamp intensity errors. | Lamp unit was at 4095 in lamp unit adjustment. | Performed lamp unit adjust to 3300 for halogen lamp intensity. Cleared all calibration errors. Performed service and user level thb calibrations. | SWO-336515 | Adjustment | HW | Oxi module | Lamp unit | halogen | Halogen lamp intensity errors. | Lamp unit was at 4095 in lamp unit adjustment. | Performed lamp unit adjust to 3300 for halogen lamp intensity. Cleared all calibration errors. Performed service and user level thb calibrations.
No | Visit | Complaint | RSCH | Rep Inhomogene Messungen | Referenz erneut undicht und voll Blut obwohl andere Lot (R2245) verwendet. Letztesmal Lot R2213 | Messkammer ausgetauscht Funktionskontrolle i.o | SWO-336193 | Defect | HW | pH/BG Module | Measuring chamber |  | Rep Inhomogeneous Measurements | Reference leaked again and full of blood although other lot (R2245) is used. Last time Lot R2213 | Measuring chamber replaced Functional control i.o
Yes | 5 | Y | RCZE | Blank |  |  | D42A33C00A5B7707 | Re classified RMED | FAN |  |  |  | Blank
No | Visit | Complaint | RDEU | 06.02.2018 Pfl. Ismael 09671930672 meldet, dass der AQT täglich eine Systemreinnigung aufruft. 08.02.2018 9:15 H. Zinkel 09671930244: weiterhin fordert das Gerät zu Systemreinigungen auf | Rücktausch des Kundengerätes | Kundengerät zurückgebracht, AQT in Betrieb genommen, zweimal eine Systemüberprüfung durchgeführt - alles Ok, LQC für die Parameter TnI, D-Dimer, NT ProBNP und CRP durchgeführt, CRP nach LQC knapp oberhalb der Grenze, CRP mit neuer LOT kalibriert mit einem Hintergrundausreißer, Autom. Systemüberprüfung durchgeführt, CRP LQC wiederholt - wieder knapp oberhalb der Grenze gemessen, Kunden informiert, Kundengerät stehen gelassen und zurück ins TB gefahren um LQC für CRP zu besorgen, AQT wurde als nicht betriebsbereit gekennzeichnet. | SWO-336657 | Revisit required | QC Ranges | QC Ranges | QC Ranges, CRP |  | 06.02.2018 Pfl. Ismael 09671930672 reports that the AQT calls a system clean-up every day. 08.02.2018 9:15 H. Zinkel 09671930244: the device continues to request system cleaning | Backtausch of the Kundengerets | Customer device returned, AQT put into operation, twice performed a system check - all Ok, LQC for the parameters TnI, D-Dimer, NT ProBNP and CRP performed, CRP to LQC just above the limit, CRP calibrated with new LOT with a background outlier, Auto. System verification performed, CRP LQC repeated - again measured just above the limit, informed customers, customer device left standing and driven back to TB to get LQC for CRP, AQT was marked as not ready for operation.
No | Visit | Complaint | RDEU | 20.03.2018, 08:06 Uhr, Fr. Frühauf/Fr. Otto, Tel. 051616021560: Display pixelt 21.03.2018, 08:24 Uhr, Fr. Deutgen, Tel. 051916023113: Kd. fragt, wann der Techniker kommt. 11.04.2018 10:00 Frau Sonders 0519160263123 meldet, dass der Bildschirm verpixelt ist. | System überprüft => Display verpixelt sich und wird schwarz weiß und ist nicht bedienbar Da PCo2 schon erneuert wurde und Fehler bestehen blieb Display tauschen | Display wird bestellt | SWO-336223 | Revisit required | HW | Analyzer Control Module | Display | mail | 20.03.2018, 08:06, Fr. Frühauf/Fr. Otto, Tel. 051616021560: Display pixelt 21.03.2018, 08:24, Ms. Deutgen, tel. 051916023113: Kd. asks when the technician will arrive. 11.04.2018 10:00 Mrs Sonders 0519160263123 reports that the screen is pixelated. | System checks => Display pixelates and turns black and white and is not operable Since PCo2 has already been renewed and errors persisted change display | Display is ordered
No | Visit | Complaint | RLTD | Joanna Boyd 0131 2422544 | Lac response errors. | Powered analyser down and removed wet section to allow access to sensor interface PCB. This was then replaced with new and refitted to analyser. On power on the analyser went into start up. After this the sensor cassette was replaced and left to condition. Glu and Lac were failing calibrations so these were repeated after a little bit more time. | SWO-337382 | Defect | HW | Basic Unit | PCB, Sensor interface | Lac | Joanna Boyd 0131 2422544 | Lac response errors. | Powered analyser down and removed wet section to allow access to sensor interface PCB. This was then replaced with new and refitted to analyser. On power on the analyser went into start up. After this the sensor cassette was replaced and left to condition. Glu and Lac were failing calibrations so these were repeated after a little bit more time.
No | Visit | Complaint | RCRO | Leakage from waste well | Začepljena otpadna linija. | Čišćenje. | SWO-336663 | [5 empty] | Leakage from waste well | Začeplyen Otpadna Line. | Cleaning.
No | Visit | Complaint | RAME | Error code 688 and 581 | Error codes 688 and 581 | Replaced Spectrometer Assembly. Replace pinch tubing near Hemolyzer. | SWO-337311 | [5 empty] | Error code 688 and 581 | Error codes 688 and 581 | Replaced Spectrometer Assembly. Replace pinch tubing near Hemolyzer.
Yes | Visit | Complaint | RCHN | 触摸屏破了 | 触摸屏破了，借903-239一个 | 更换903-239后修复 by Andew | SWO-334851 | [5 empty] | The touch screen's broken. | The touch screen is broken, borrow edith 903-239 | Repair by Andew after replacing 903-239
Yes | Visit | Complaint | RCHN | 开机无法显示 | 开机无法显示，借903-239一个 | 更换903-239后修复 by Andrew | SWO-335956 | [5 empty] | Power on does not show | Power can not be displayed, borrow 903-239 a | Repair by Andrew after replacing 903-239
No | Visit | Complaint | RCHN | 报错1031 | 报错1031，借903-364一个 | 更换903-364后修复 by Andrew | SWO-335315 | Defect | HW | Process Module | Optical Unit | 1031 | 1031 | 1031, loan 903-364 one | Repair by Andrew after replacing 903-364
No | Visit | Complaint | RCHN | HCT无法测定 | HCT无法测定，借903-329一个 | 未使用备件，清洁管路及备件后修复 by Andrew | SWO-336071 | Cleaned | HW | Aspiration Module | Needle Wash Unit | HCT | HCT cannot be measured | HCT cannot be determined, with 903-329 one | Unused spare parts, cleaning lines and repair by Andrew after spare parts
No | Visit | Complaint | RCHN | 报错1024 | 报错1024，借903-364 944-394（13849）一张 | 更换备用机上903-364 和重新定标后修复 by Andrew | SWO-336041 | Defect | HW | Process Module | Optical Unit | 1024 | 1024 error | 1024, borrow 903-364 944-394 (13849) one | Repair by Andrew after replacing the spare 903-364 and re-calibration
No | Visit | Complaint | RCHN | 报错1070 | 报错1070 借903-210一个 | 更换903-210后修复 by Andrew | SWO-336004 | Defect | HW | Process Module | PCB, Cup dryer | 1070 | 1070 | False 1070 Borrow 903-210 one | Fix by Andrew after replacing 903-210
No | Visit | Complaint | RCHN | 显示屏扫描故障 | 显示屏扫描故障，借902-924一个 | 更换902-924后修复 by Andrew | SWO-335430 | Defect | HW | Cover and Chassis | Barcode reader, display |  | Display scanning fault | Display scanning fault, borrow 902-924 one | Repair by Andrew after replacing 902-924
No | Visit | Complaint | RCRO | Printer not printing. | Slomljen držač valjka za papir na pisaču. | Zamjena pisača. | SWO-336666 | Defect | HW | Printer Module | Printer head |  | Printer not printing. | A broken paper roll holder on the printer. | Replacing the printer.
No | Visit | Complaint | RIND | Calibration failed. | Calibartion failed for CO2, Lac, Glu and elctrolytes. | Find Sensitivity gone for all three. Membrane Asked for CO2, Lac and Glu. Reference membrane reinstalled by cleaning. and calibrated. Membranes requested. | SWO-337519 | Revisit required | Electrode | Membrane | pCO2 | no customer stock | Calibration failed. | Calibartion failed for CO2, Lac, Glu and elctrolytes. | Find Sensitivity gone for all three. Membrane Asked for CO2, Lac and Glu. Reference membrane reinstalled by cleaning. and calibrated. Membranes requested.
No | Visit | Complaint | RAME | low pco2 QC recovery for Level 1 & 2. Results had been in range. Also observed general erratic tendencies. | Follow up to SWO-324792. Customer has been using a used, know good pCO2 electrode. QC values are in line with expected values. Confirms recently replaced pCO2 electrode to resolve erratic calibration drift with the previous electrode, was the cause of erratic lower pCO2 values. | Replaced pCO2 electrode, re-membraned, performed calibration and QC. All parameters within specifications. | SWO-336567 | User expectation | Specifications | QC | electrode | pCO2 | low pco2 QC recovery for Level 1 & 2. Results had been in range. Also observed general erratic tendencies. | Follow up to SWO-324792. Customer has been using a used, know good pCO2 electrode. QC values are in line with expected values. Confirms recently replaced pCO2 electrode to resolve erratic calibration drift with the previous electrode, was the cause of erratic lower pCO2 values. | Replaced pCO2 electrode, re-membraned, performed calibration and QC. All parameters within specifications.
No | 3 | Y | RPOL | Blank | sprawdzenie aparatu, wymiana modu³u. Kalibracja pomp, czujników, parametrów. Aparat sprawny technicznie. |  | D42B6BAD8029CF42 | Defect | HW | pH/BG Module | pH/BG Module | most likely LS should be replaced | Blank | Checking the camera, replacing the module. Calibration of pumps, sensors, parameters. The camera is technically efficient.
No | Visit | Complaint | RDEU | 06.02.2018 Pfl. Ismael 09671930672 meldet, dass der AQT täglich eine Systemreinnigung aufruft. 08.02.2018 9:15 H. Zinkel 09671930244: weiterhin fordert das Gerät zu Systemreinigungen auf | CRP zu hoch bei LQC Messung | Systemüberprüfung durchgeführt - alles in Ordnung, keine Systemreinigung durchgeführt, LQC für CRP aus TB im AQT platziert und gemessen, Messzeit läuft runter, CRP im Kontrollbereich, LQC erneut gemessen als Kontrollmessung - Wert wieder oberhalb des Grenzbereichs AQT wird vom Kunden wieder abgezogen!!! | SWO-336838 | Revisit required | QC Ranges | QC Ranges | QC Ranges, CRP |  | 06.02.2018 Pfl. Ismael 09671930672 reports that the AQT calls a system clean-up every day. 08.02.2018 9:15 H. Zinkel 09671930244: the device continues to request system cleaning | CRP too high for LQC measurement | System check performed - all in order, no system cleaning performed, LQC for CRP from TB placed and measured in The AQT, measuring time runs down, CRP in the control area, LQC again measured as control measurement - value again above the limit range AQT is deducted from the customeragain!!!
No | Visit | Complaint | RDEU | 20.03.2018, 08:06 Uhr, Fr. Frühauf/Fr. Otto, Tel. 051616021560: Display pixelt 21.03.2018, 08:24 Uhr, Fr. Deutgen, Tel. 051916023113: Kd. fragt, wann der Techniker kommt. 11.04.2018 10:00 Frau Sonders 0519160263123 meldet, dass der Bildschirm verpixelt ist. | Display und Display Kabel tauschen | Display uns Display Kabel zum testen aus Demo Gerät eingebaut | SWO-336732 | Defect | HW | Analyzer Control Module | Display | mail | 20.03.2018, 08:06, Fr. Frühauf/Fr. Otto, Tel. 051616021560: Display pixelt 21.03.2018, 08:24, Ms. Deutgen, tel. 051916023113: Kd. asks when the technician will arrive. 11.04.2018 10:00 Mrs Sonders 0519160263123 reports that the screen is pixelated. | Replace display and display cable | Display us display cable built-in for testing from demo device
No | Visit | Complaint | RDEU | 12.04.2018, Herr Lehnert, Tel: 04921981654: Probenaspiration gescheitert-LS Fehler | Einsatz durch Martin Peukert durchgeführt jede zweite Messung Probenaspitration gescheitert Qwetschventil sehr schwergängig | Ventil geölt , und im Service getestet mehrere Probenmessungen durchgeführt System IO | SWO-337703 | Adjustment | HW | Oxi module | Pinch valve | oil | 12.04.2018, Mr. Lehnert, Tel: 04921981654: Sample Aspiration Failed-LS Error | Operation carried out by Martin Peukert every second measurement sample aspitration failed Qwetsch valve very difficult to use | Valve oiled , and tested in service several sample measurements performed system IO
No | Visit | Complaint | RKOR | Leak current in relation to solution pack detected | Fiuld line check replace solution pack analyzer status check | Fiuld line check replace solution pack analyzer status check | SWO-337944 | Assisted Replacement | SP | SP | Leak Current |  | Leak current in relation to solution pack detected | Fiuld line check replace solution pack analyzer status check | Fiuld line check replace solution pack analyzer status check
No | Visit | Complaint | RAME | Thb cal failure 283 Pka Mark - 863-687-1100 ext 4607 | Ran the OXI complete test in service, and the hemolyzer failed for low Cavitation Absorb. | Replaced the hemolyzer, and this resolved the problem. | SWO-337774 | Defect | HW | Oxi module | Hemolyzer |  | Thb cal failure 283 Pka Mark - 863-687-1100 ext 4607 | Ran the OXI complete test in service, and the hemolyzer failed for low Cavitation Absorb. | Replaced the hemolyzer, and this resolved the problem.
No | 6 | Y | RPAC | Blank | inlet $cassette |  | D42CB0F3F64E23AC | Defect | HW | Inlet module | probe |  | Blank | inlet /cassette
No | 3 | Y | RPOL | Blank | Przep³ukanie analizatora podchlorynem. Testy serwisowe - ok Analizator sprawn technicznie. |  | D42CB1B8F62D1ADD | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | Blank | The Hypochlorate Analyzer is being overflowed. Service Tests-OK analyser is technically efficient.
No | Visit | Complaint | RCHN | 进样口支架断裂，FTC106 | 进样口支架断裂，FTC106，借两个903-584和903-329一个 |  | SWO-337358 | Defect | HW | Cover and Chassis | Tube holder |  | Sample holder broken, FTC106 | The inlet bracket is broken, FTC106, with two 903-584 and 903-329 one
No | Visit | Complaint | RIND | Calibration failed. | Glu, Lac and CO2 membrane to replace. | Protean removal done. membrane replacement done. Analyser Serviced and calibrated properly. handed over in a fine working condition. | SWO-337531 | Assisted Replacement | Electrode | Membrane | pCO2 | no customer stock | Calibration failed. | Glu, Lac and CO2 membrane to replace. | Protean removal done. membrane replacement done. Analyser Serviced and calibrated properly. handed over in a fine working condition.
No | Visit | Complaint | RKK | サンプル測定時Oxiの液センサーの検出が早すぎました。 | サンプル移動エラー | ポンプキャリブレーション、液さんサー調整実施。QC、検体測定確認。 | SWO-338011 | Adjustment | HW | pH/BG Module | Pump cal |  | Oxi's liquid sensor was detected too quickly when measuring samples. | Sample move error | Pump calibration and liquid-san sir adjustment. QC, specimen measurement confirmation.
No | Visit | Complaint | RSCH | Rep. Ref. | Immer wieder Elektrolyten Drift | Ref. Elektrode durch MedTech ersetzt i.o Alte Lot 6807=> Garantiersatzt | SWO-337400 | Assisted Replacement | Electrode | Electrode | Reference |  | Rep. Ref. | Time and again electrolytes drift | Ref. Electrode replaced by MedTech i.o Old Lot 6807=> Warranty Set
No | Visit | Complaint | RAME | Aborting calibrations and QC samples | Checked data logs for errors, checked valves and liquid sensors, checked electrodes and electrode chambers | Removed electrodes and cleaned chambers, removed modules and cleaned valves, | SWO-337807 | Cleaned | HW | pH/BG Module | Valve, air |  | Aborting calibrations and QC samples | Checked data logs for errors, checked valves and liquid sensors, checked electrodes and electrode chambers | Removed electrodes and cleaned chambers, removed modules and cleaned valves,
No | Visit | Complaint | RCHN | 测试卡断裂 | 测试卡断裂 | 更换测试卡后修复 | SWO-337811 | Defect | Consumables | Test Kit | Cartridge, Broken |  | Test card break | Test card break | Repair after replacing test card
No | Visit | Complaint | RAME | Rinse bottle percentage is not reflecting actual remaining rinse | Confirmed issue when onsite. User is also reporting QC carousel sticky. | Reset forecast factor in user interface to 1 and then adjust actual rinse volume onboard. Will follow up with customer next week if volume is accurately reflected. Also adjust QC carousel so it is now easy to lift out of chassis. | SWO-337898 | Adjustment | HW | Autocheck Module | QC carousel |  | Rinse bottle percentage is not reflecting actual remaining rinse | Confirmed issue when onsite. User is also reporting QC carousel sticky. | Reset forecast factor in user interface to 1 and then adjust actual rinse volume onboard. Will follow up with customer next week if volume is accurately reflected. Also adjust QC carousel so it is now easy to lift out of chassis.
No | Visit | Complaint | RIND | pCO2 not calibrating. | pCO2 not calibrating. Again and again problem. | Checked replacement schedule and found pCO2 failure and membrane replacement is very frequent in last 2 months. Replacement printout taken. Electrode and membrane of pCO2 replaced and calibrated properly. Handed over in a fine working condition. | SWO-337924 | Assisted Replacement | Electrode | Electrode | pCO2 |  | pCO2 not calibrating. | pCO2 not calibrating. Again and again problem. | Checked replacement schedule and found pCO2 failure and membrane replacement is very frequent in last 2 months. Replacement printout taken. Electrode and membrane of pCO2 replaced and calibrated properly. Handed over in a fine working condition.
No | Visit | Complaint | RAME | Co2 trending high on WDC level 1, QC missed a few on level 1 and 2. trending high on qc graph | Co2 trending high on WDC level 1, QC missed a few on level 1 and 2. trending high on qc graph Checked electrodes. reference leaked on reference and pco2 electrode. | cleaned measuring chambers re membrane pco2 and ref with new lot numbers | SWO-340587 | Cleaned | Fluidic pathway | pH/BG Module | Measuring chamber | pCO2 and ref | Co2 trending high on WDC level 1, QC missed a few on level 1 and 2. trending high on qc graph | Co2 trending high on WDC level 1, QC missed a few on level 1 and 2. trending high on qc graph Checked electrodes. reference leaked on reference and pco2 electrode. | cleaned measuring chambers re membrane pco2 and ref with new lot numbers
No | Visit | Complaint | RCHN | 检测结果偏高 | 检测结果偏高，借903-236一个 |  | SWO-337964 | Defect | HW | Process Module | Process Wheel |  | The test results are high. | The test results are high, borrow 903-236 one
No | Visit | Complaint | RAME | QC high on tHb after tHb cal | 1) Error 1034 OXI HW problem | 1) Replaced hemolyzer and tested. | SWO-338538 | Defect | HW | Oxi module | Hemolyzer | 1034 | QC high on tHb after tHb cal | 1) Error 1034 OXI HW problem | 1) Replaced hemolyzer and tested.
No | Visit | Complaint | RAME | System temperature out of range | The customer has replaced two sensor cassettes today run #751-209 and, 751-161 and, both cassette presented error 1326 'Sensor thermistor recalibration failed'. The analyzer would not restart nor allow the user to press the 'Try again' tab (Grayed out). | 04 $12 $18 I had a sensor cassette on hand and, I replaced it on this analyzer for the customer. I allowed time for the sensor cassette to warm up and, the analyzer restarted successfully and, the conditioning process commenced. Calibrations for pCO2 passed but sensitivity value needed to be slightly higher as the external qc level one did not pass. All calibrations and, internal qc and, external qc levels 2 and, 3 were in. I allowed the sensor cassette to calibrate overnight so that the pCO2 sensitivity improves from 79% to 85% so that qc level one can pass. 04 $13 $18 Upon arrival, I found the calibrations for pCO2 sensitivity and, status were not passing. I replaced the sensor cassette and, the solution pack. However, the calibrations would not pass and, the temperature errors reoccurred. I ordered the sensor interface pcb. 04 $16 $18 Visited to replace the sensor interface pcb. Also replaced the sensor cassette as part of the troubleshooting process. | SWO-337456 | Defect | HW | Basic Unit | PCB, Sensor interface | temp/pCO2 | System temperature out of range | The customer has replaced two sensor cassettes today run #751-209 and, 751-161 and, both cassette presented error 1326 'Sensor thermistor recalibration failed'. The analyzer would not restart nor allow the user to press the 'Try again' tab (Grayed out). | 04 $12 $18 I had a sensor cassette on hand and, I replaced it on this analyzer for the customer. I allowed time for the sensor cassette to warm up and, the analyzer restarted successfully and, the conditioning process commenced. Calibrations for pCO2 passed but sensitivity value needed to be slightly higher as the external qc level one did not pass. All calibrations and, internal qc and, external qc levels 2 and, 3 were in. I allowed the sensor cassette to calibrate overnight so that the pCO2 sensitivity improves from 79% to 85% so that qc level one can pass. 04 $13 $18 Upon arrival, I found the calibrations for pCO2 sensitivity and, status were not passing. I replaced the sensor cassette and, the solution pack. However, the calibrations would not pass and, the temperature errors reoccurred. I ordered the sensor interface pcb. 04 $16 $18 Visited to replace the sensor interface pcb. Also replaced the sensor cassette as part of the troubleshooting process.
No | Visit | Complaint | RAME | Solution pack will not install. | Solution pack will not install. Customer says SP will not go in all the way. She tried another one and same thing. No obstructions present. | FSE found the SP not activated all the way. Smashed down on the corners and SP now installs. SC calibrations for Lactate having hard time coming in for Sens. Finally came in. | SWO-338514 | User error | SP | SP | SP Activation |  | Solution pack will not install. | Solution pack will not install. Customer says SP will not go in all the way. She tried another one and same thing. No obstructions present. | FSE found the SP not activated all the way. Smashed down on the corners and SP now installs. SC calibrations for Lactate having hard time coming in for Sens. Finally came in.
No | Visit | Complaint | RAME | The analyzer is not reconnecting to the wireless network automatically. | Checked the settings in the control panel. Found that the network was set to manual. | Changed the connection from manual to automatic. | SWO-338639 | Setup | SW | SW | Wireless connection |  | The analyzer is not reconnecting to the wireless network automatically. | Checked the settings in the control panel. Found that the network was set to manual. | Changed the connection from manual to automatic.
No | Visit | Complaint | RIND | Single Parameter error . | Single parameter error . It is found that there are some corrosion on the sensor board . | Re-installed the cassette after flushing it . Problem not solved . Replaced the sensor cassette . Problem solved . Need to change the sensor board . | SWO-340015 | [5 empty] | Single Parameter error . | Single parameter error . It is found that there are some corrosion on the sensor board . | Re-installed the cassette after flushing it . Problem not solved . Replaced the sensor cassette . Problem solved . Need to change the sensor board .
No | Visit | Complaint | RESP | Analizador muerto | Analizador muerto No se enciende la pantalla | Sustitucion de la cpu Instalacion de casete y sp | SWO-338323 | Defect | HW | Basic Unit | CPU |  | Dead analyzer | Dead analyzer The screen does not turn on | Substitution of the cpu Installing cassette and sp
No | Visit | Complaint | RCHN | 定标失败 | 定标失败，借944-212（14137）一个 | 更换903-364重新定标后修复(by chuanwei) | SWO-337881 | Defect | HW | Process Module | Optical Unit | mail | Bid failed | Failed to bid, borrow 944-212 (14137) one | Repair after replacement 903-364 recalibration (by chuanwei)
Yes | Visit | Complaint | RCHN | 定标失败 | 加借903-364一个 | 更换903-364重新定标后修复（by chuanwei） | SWO-337923 | Re classified RMED | double entry |  |  | mail | Bid failed | Borrow 903-364 one | Repair after replacement 903-364 recalibration (by chuanwei)
No | Visit | Complaint | RCHN | 报错1024 | 报错1024，借903-364，944-212（13822）各一个 | 更换903-364重新定标后修复(by chuanwei) | SWO-337865 | Defect | HW | Process Module | Optical Unit | 1024 | 1024 error | 1024, borrow 903-364, 944-212 (13822) each | Repair after replacement 903-364 recalibration (by chuanwei)
No | Visit | Complaint | RCHN | 卡盘 | 卡盘，借944-212（14137）一个 |  | SWO-338029 | Assisted Replacement | Consumables | Test Kit | Cartridge, TnI |  | Chuck | Card, borrow 944-212 (14137) one
No | 4 | Y | RPOL | Blank | Czyszczenie kruæca odpowietrzaj¹cego, konfiguracja planu QC. Aparat sprawny. |  | D4304D6D58345155 | Cleaned | HW | Oxi module | Oxi module |  | Blank | Cleaning the responding dealer, setting up the QC plan. The camera is functional.
No | Visit | Complaint | RIND | patient report not coming properly it has been some improper report while printing | checked and found that the patient as well as any other print not coming clear | cleaned the printer head as well as thermal head but problem not solved need to replace the printer head along with thermal head. | SWO-339140 | Defect | HW | Printer Module | Printer head |  | patient report not coming properly it has been some improper report while printing | checked and found that the patient as well as any other print not coming clear | cleaned the printer head as well as thermal head but problem not solved need to replace the printer head along with thermal head.
No | Visit | Complaint | RESP | Error 1273 flujo fuera de rango | Error 1273 flujo fuera de rango | Se cambian tubos, bombas, filtro, unidad de lavado copas y se hace comprobación con muestras de blanco y qcs. | SWO-338128 | Defect | HW | Process Module | Wet section | Flow | Error 1273 out-of-range flow | Error 1273 out-of-range flow | Tubes, pumps, filter, washing unit cups are changed and tested with white samples and qcs.
No | Visit | Complaint | RIND | 'CALIBRATE_13' error message was observed on AQT90 analyzer at Parkar Hospital, Ratnagiri | Need to check the incubator cup sensor for any dust deposition. | Lot of dust was observed on the incubator cup sensor, cleaned it, done the incubator wheel calibration, done the cup sensor calibration, run 2 samples, results are found to be satisfactory. Machine is working fine. | SWO-341350 | Cleaned | HW | Incubator Module | sensor, cup |  | CALIBRATE_13' error message was observed on AQT90 analyzer at Parkar Hospital, Ratnagiri | Need to check the incubator cup sensor for any dust deposition. | Lot of dust was observed on the incubator cup sensor, cleaned it, done the incubator wheel calibration, done the cup sensor calibration, run 2 samples, results are found to be satisfactory. Machine is working fine.
No | Visit | Complaint | RKOR | Sample aspiration error | . | User Training. Gasket module clean. Gasket hold clean. Liquid flow line clean. Sensor cassette clean. Solution pack clean. Service mode check. Calibration check. Control check. | SWO-338678 | User error | Sample handling | Sample handling | Sample handling |  | Sample aspiration error | . | User Training. Gasket module clean. Gasket hold clean. Liquid flow line clean. Sensor cassette clean. Solution pack clean. Service mode check. Calibration check. Control check.
No | Visit | Complaint | RDEU | 16.04.2018, Herr Tretschock, Tel: 09317912703: Einlass schließt nicht mehr | Fehlersuche --> Dateiarchive (Aktivitäten, Betriebsmittel, Kal., QC), Einlassmodul, Lösungspack, Schlauchwege überprüft --> Einlass schließt nicht komplett --> Flachbandkabel der Auto Inlet PCB ist schräg angesteckt und Zahnrad für Aufnahme des Einlassarmes nicht in Grundposition | Flachbandkabel der Auto Inlet PCB richtig angesteckt, Zahnrad des Einlasses richtig positioniert, Startphase des Analysators, zum Test den Einlass mehrfach auf und zu gefahren --> Einlass schließt wieder richtig und komplett, System und Funktion i.O. | SWO-338786 | Adjustment | HW | Auto Inlet module | PCB, Auto inlet | cable | 16.04.2018, Mr. Tretschock, Tel: 09317912703: Entrance no longer closes | Troubleshooting --> File archives (activities, equipment, cal., QC), inlet module, solution pack, hose paths checked --> Inlet does not completely close --> Ribbon cable of the Auto Inlet PCB is connected at an angle and gear for mounting the inlet arm is not in Basic position | Ribbon cable of the Auto Inlet PCB properly plugged in, gear of the inlet correctly positioned, start phase of the analyzer, for testing the inlet repeatedly on and off --> Inlet closes again correctly and completely, system and function i.O.
No | Visit | Complaint | RAME | Error 1063 vacuum error | Confirmed issue when onsite. Perform vacuum test in service program vacuum failed at 0.9, exceeded acceptable range of 0.75. Tried replacing solution pack to no avail. Error still present. | Replaced inlet gasket holder. Vacuum test now at 0.48. | SWO-339067 | Defect | HW | Auto Inlet module | Inlet gasket | 1063 | Error 1063 vacuum error | Confirmed issue when onsite. Perform vacuum test in service program vacuum failed at 0.9, exceeded acceptable range of 0.75. Tried replacing solution pack to no avail. Error still present. | Replaced inlet gasket holder. Vacuum test now at 0.48.
No | Visit | Complaint | RKOR | pCO2 Calibration error | . | Sensor cassette clean. re install, Liquid flow line clean. 2 point cal work. Sample test. | SWO-339305 | [5 empty] | pCO2 Calibration error | . | Sensor cassette clean. re install, Liquid flow line clean. 2 point cal work. Sample test.
No | Visit | Complaint | RDEU | Flüssigkeitssensor Kalibrierfehler | Auskristallisierungen am Sensor | Fehlersuche, Reinigung Flüssigkeitssensor, Erneuerung Lösungspack, Funktionskontrolle | SWO-338685 | [5 empty] | Flüssigkeitssensor Kalibrierfehler | Crystallizations on the sensor | Troubleshooting, cleaning liquid sensor, renewal solution pack, function control
Yes | Visit | Complaint | RCHN | 报错1031 | 报错1031，借903-364一个 | 更换903-364后修复(by chuanwei) | SWO-338186 | [5 empty] | 1031 | 1031, loan 903-364 one | Repair after replacement 903-364 (by chuanwei)
Yes | Visit | Complaint | RCHN | 报HCT值不能被测量 | 报HCT值不能被测量，借903-329和840-681各一个 | 更换903-329和840-681后修复(by chuanwei) | SWO-338066 | [5 empty] | Report HCT value sit | The HCT value cannot be measured, one each of 903-329 and 840-681 | Repair after replacement of 903-329 and 840-681 (by chuanwei)
Yes | Visit | Complaint | RCHN | 仪器报修 | 报错1223，仪器报修，支援代理商，HCT定标不过 | 更换HCT模块，重新进行针校准＆HCT定标(by chuanwei) | SWO-339267 | [5 empty] | Instrument repair | Report error 1223, instrument repair, support agent, HCT calibration but | Replace the HCT module, re-calibrate the needle and the HCT calibration (by chuanwei)
No | Visit | Complaint | RDAN | Har målt lavt på QC siden levering | Insertgrænser tjekket Ventil tjek, målekammer tjek, trykprøvet | Målekammer skiftet (revne) Kal. og QC | SWO-338893 | Defect | HW | pH/BG Module | Measuring chamber |  | Have measured low on QC since delivery | Insertboundaries checked Valve check, measuring chamber check, pressure tested | Measuring chamber shift (crack) Kal. and QC
No | Visit | Complaint | RSCH | Rep. Ca/Ph | Drift Elektrolyten QC pH Level 4 tief QC Na Level 4 hoch instabil QC Ca alle Level hoch instabil Kunde hat Membranen Ref. Ca etc. bereits ersetzt | Proteinentfernung durchgeführt i.o Alle Messkammern sauber i.o Einlassdichtung ersetzt i.o Einlassgummischlauch ersetzt i.o Ref. und pH Elektrode ersetzt i.o Membranenwartung Ca und Ref. durchgeführt i.o Kal. und QC i.o System i.o Alte Elektroden Lot Ref. = 8803 => Garantie pH = YP03 => Garantie | SWO-338979 | Assisted Replacement | Electrode | Electrode | Reference |  | Rep. Ca/Ph | Drift electrolytes QC pH Level 4 deep QC Na Level 4 highly unstable QC Ca all levels highly unstable Customer has membranes ref. Ca etc. already replaced | Protein removal performed i.o All measuring chambers clean i.o Inlet seal replaces i.o Inlet rubber hose replaces i.o Ref. and pH electrode replaces i.o Membrane maintenance Ca and Ref. carried out i.o Kal. and QC i.o System i.o Old Electrodes Lot Ref. = 8803 => Warranty pH = YP03 => Warranty
No | Visit | Complaint | RDEU | 17.04.2018; Hr. Hennboele | Abf.wege verstopft, | Einl.dichtung aus Kundenbestand ersetzt, Abf.wege gespült, Gerät kalibriert, Probemessung durchgeführt, | SWO-339008 | Defect | HW | Inlet module | Nipple kit |  | 17.04.2018; Hr. Hennboele | Abf.wege clogged, | Replaced from customer stock, flushed waste, device calibrated, sample measurement performed,
No | Visit | Complaint | RIND | Oxi parameter results are with ? mark. | They had issue regarding Hb and S02 results. Instrument was showing pump calibration error | completed pump calibration and checked with samples. Unit was working fine. | SWO-341135 | Revisit required | HW | Oxi module | Lamp unit |  | Oxi parameter results are with ? mark. | They had issue regarding Hb and S02 results. Instrument was showing pump calibration error | completed pump calibration and checked with samples. Unit was working fine.
No | Visit | Complaint | RIND | Leaking from inlet gasket for all the samples | Checked the machone found clot in waste tube & inlet gasket | Cleaned inlet gasket & waste tube | SWO-341168 | Cleaned | Fluidic pathway | Inlet module | Inlet gasket |  | Leaking from inlet gasket for all the samples | Checked the machone found clot in waste tube & inlet gasket | Cleaned inlet gasket & waste tube
No | Visit | Complaint | RAME | pH electrode shows clear signs of leakage. | pH electrode shows clear signs of leakage. | Cleaned blood and residue from both inside and outside of electrode sleeve and pH electrode bay. Replaced sealing ring. | SWO-339057 | Defect | Electrode | Electrode | sealing ring | sealing ring, pH | pH electrode shows clear signs of leakage. | pH electrode shows clear signs of leakage. | Cleaned blood and residue from both inside and outside of electrode sleeve and pH electrode bay. Replaced sealing ring.
No | Visit | Complaint | RIND | Leakage in Gaskets | Gasket Leakage | Chekced the analyser it was leaking in gasket near syring flap so removed the gasket and cleaned the same refixed it after that also its leaking then chekced the all tubes which will connected to gasket and checked the pump tube then inter changed the pump tubes after that also leakig was there near syring flap , cleaned the needle and given pump calibration after that also leakage was dint stopped so suspected with Gasket , placed standby gasket given rinse and leakage was stopped , then observed for few samples leakage not found so Gasket is defetctive | SWO-341057 | Defect | HW | Inlet module | Inlet gasket |  | Leakage in Gaskets | Gasket Leakage | Chekced the analyser it was leaking in gasket near syring flap so removed the gasket and cleaned the same refixed it after that also its leaking then chekced the all tubes which will connected to gasket and checked the pump tube then inter changed the pump tubes after that also leakig was there near syring flap , cleaned the needle and given pump calibration after that also leakage was dint stopped so suspected with Gasket , placed standby gasket given rinse and leakage was stopped , then observed for few samples leakage not found so Gasket is defetctive
No | Visit | Complaint | RSAF | PO2 results reading lower than expected | Simulator test ok. connenctor gasket looks warn. | Checked unit, sensor simulator test done all channels ok. Software error during upgrade, then complete software installation done.Replaced inlet connector gasket and SC . | SWO-339558 | Assisted Replacement | SC | SC | PO2 |  | PO2 results reading lower than expected | Simulator test ok. connenctor gasket looks warn. | Checked unit, sensor simulator test done all channels ok. Software error during upgrade, then complete software installation done. Replaced inlet connector gasket and SC .
No | Visit | Complaint | RLTD | Rob Bolton 01772 528430 | Site visit to investigate inlet problem. Customer had reported that they had replaced sc and sp. They had replaced inlet gasket and flushed the analyser thoroughly. They had found restriction in line which they managed to remove. Customer has replaced the inlet Probe. Logs show that the analyser has been experiencing inhomogeneous cal 3 solution and Inlet gasket errors. | Entered service and ejected solution pack. Checked pack and bay for visible issues. Cleaned flow selector window. Removed inlet module and I specter for damage. Checked gasket and probe. These are new and in good condition. Made sure probe retracts and returns to full home positioning without sticking. Refitted Inlet and ran Inlet calibration a number of times. Put Inlet to cap and strike get positions and switched between all three positions multiple times. Recalibrated. Refitted solutions pack. Ran through all fluid transport programs and found all values well within tolerances. Vacuum and pressure tests look perfect. Exited service and then ran calibrations and qcs to verify analyser working. Calibrations and qcs passed. Customer to monitor and update radiometer if issue persists. | SWO-339670 | No fault found | HW | Auto Inlet module | Auto Inlet module |  | Rob Bolton 01772 528430 | Site visit to investigate inlet problem. Customer had reported that they had replaced sc and sp. They had replaced inlet gasket and flushed the analyser thoroughly. They had found restriction in line which they managed to remove. Customer has replaced the inlet Probe. Logs show that the analyser has been experiencing inhomogeneous cal 3 solution and Inlet gasket errors. | Entered service and ejected solution pack. Checked pack and bay for visible issues. Cleaned flow selector window. Removed inlet module and I specter for damage. Checked gasket and probe. These are new and in good condition. Made sure probe retracts and returns to full home positioning without sticking. Refitted Inlet and ran Inlet calibration a number of times. Put Inlet to cap and strike get positions and switched between all three positions multiple times. Recalibrated. Refitted solutions pack. Ran through all fluid transport programs and found all values well within tolerances. Vacuum and pressure tests look perfect. Exited service and then ran calibrations and qcs to verify analyser working. Calibrations and qcs passed. Customer to monitor and update radiometer if issue persists.
No | Visit | Complaint | RDEU | 17.04.2018 13:15Uhr Teis 02611377310: Der Einlass schließ nicht bzw. fährt auch nicht mehr aus. | Einlassmodul eingebaut und Kalibriert, den Anwendern erklärt, das das Einlassmodul nicht einfach rausgenommen werden darf. | Startphase QC und Kalibration | SWO-339484 | Adjustment | HW | Auto Inlet module | Auto Inlet module |  | 17.04.2018 13:15h Teis 02611377310: The entrance does not close or does not stop. | Inlet module installed and calibrated, explains to the users that the inlet module must not be simply removed. | Start phase QC and calibration
No | Visit | Complaint | RIND | Service Visit. | Analyzer screen shows that parameter deactivated. | Keep the sensor cassette in to the 2-8 degree Celsius after flush the cassette with de-ionized water and sodium heparin for 1 hour, After 1 hour again re-initiate the SC, successfully installed, Analyzer handed over to staff in well working condition. | SWO-341055 | [5 empty] | Service Visit. | Analyzer screen shows that parameter deactivated. | Keep the sensor cassette in to the 2-8 degree Celsius after flush the cassette with de-ionized water and sodium heparin for 1 hour, After 1 hour again re-initiate the SC, successfully installed, Analyzer handed over to staff in well working condition.
No | Visit | Complaint | RIND | Ca and Cl parameter was become red. | Ca and CL parameter was failed to calibrate. | Removed sensor cassette then flush it with heparin then kept in refrigerator for 30 min. then reinstall. after that all parameter was active. handed over in fine working condition. | SWO-341329 | [5 empty] | Ca and Cl parameter was become red. | Ca and CL parameter was failed to calibrate. | Removed sensor cassette then flush it with heparin then kept in refrigerator for 30 min. then reinstall. after that all parameter was active. handed over in fine working condition.
No | Visit | Complaint | RCHN | 测试卡智能芯片无法阅读 | 测试卡智能芯片无法阅读，借902-972一个 | 更换902-972后修复(By Ran) | SWO-338642 | Defect | HW | Basic Unit | PCB, Sensor connector assembly |  | Test card smart chip can't read | Test card smart chip can not read, borrow 902-972 one | Repair after replacement 902-972 (By Ran)
Yes | Visit | Complaint | RCHN | pO2定标不过 | pco2、glu、lac定标失败 报错0378和0375 | 更换测试卡后修复 | SWO-342373 | [5 empty] | pO2 calibration but | pco2, glu, lac calibration failure reported error000s 0378 and 0375 | Repair after replacing test card
No | Visit | Complaint | RCHN | 定标失败 | 定标失败，借903-183（被903-581替代） | 更换903-581后修复(903-581 ) | SWO-338748 | Defect | HW | Basic Unit | PCB, Sensor interface |  | Bid failed | Failed to bid, borrowed 903-183 (replaced by 903-581) | Repair after replacement 903-581 (903-581) )
No | Visit | Complaint | RCHN | 报错2168 | 报错2168，需要借件903-138 | 更换903-138后修复(by chuanwei) | SWO-338788 | Defect | HW | Incubator Module | PCB, Incubator |  | 2168 error | 2168, 903-138 borrowed | Repair after replacement 903-138 (by chuanwei)
No | Visit | Complaint | RIND | IQC values found to be out of range for Na, Ca and K on ABL800 analyzer at Ruby Hall Clinic, Pune | Need to check the status of the membranes as well as any blood clots trapped in the measuring chamber. | Blood clot was observed in the El $Met Module measuring Chamber, cleaned it, done 2-point calibration , run controls, results are OK. Machine is working fine. | SWO-341360 | Cleaned | Fluidic pathway | El/Met Module | Measuring chamber |  | IQC values found to be out of range for Na, Ca and K on ABL800 analyzer at Ruby Hall Clinic, Pune | Need to check the status of the membranes as well as any blood clots trapped in the measuring chamber. | Blood clot was observed in the El $Met Module measuring Chamber, cleaned it, done 2-point calibration , run controls, results are OK. Machine is working fine.
No | Visit | Complaint | RDEU | Aqure Verbindung unterbrochen nach Technikereinsatz | Analysator hat sich nach Neustart mit Seriennummer angemeldet, Fernsupport war nicht möglich | Fehlersuche, Warten auf externe Zuarbeiten, Update Type, Funktionskontrolle | SWO-339499 | Revisit required | SW | Aqure | Aqure |  | Aqure connection interrupted after technician deployment | Analyzer logged in after reboot with serial number, remote support was not possible | Troubleshooting, Waiting for external work, Update Type, Function Control
No | Visit | Complaint | RIND | Oxi parameter results are with ? mark. | Checked complete oximetry test and found that lamp sensitivity was out of range. | Done lamp adjustment procedure and also done complete oximetry test. Then given 2 point calibration and checked with samples.Unit is working fine | SWO-341148 | Adjustment | HW | Oxi module | Lamp unit |  | Oxi parameter results are with ? mark. | Checked complete oximetry test and found that lamp sensitivity was out of range. | Done lamp adjustment procedure and also done complete oximetry test. Then given 2 point calibration and checked with samples. Unit is working fine
No | Visit | Complaint | RDEU | pH Elektrode muss überprüft werden. | pH Elektrode defekt. | pH Elektrode verhält sich instabil. pH Elektrode erneuert. | SWO-340794 | Assisted Replacement | Electrode | Electrode | pH |  | pH electrode must be checked. | pH electrode defective. | pH electrode behaves unstable. pH electrode.
No | Visit | Complaint | RDEU | pCO2 Probleme beim QC Level 3 | pCO2 Elektrode $ Messkammer verunreinigt. | Messkammer gereinigt, pCO2 Elektrode gereinigt, membraniert und dekontamination durchgeführt. | SWO-340793 | Cleaned | Fluidic pathway | pH/BG Module | Measuring chamber |  | pCO2 Probleme beim QC Level 3 | pCO2 Electrode - Measuring chamber contaminated. | Measuring chamber cleaned, pCO2 electrode cleaned, membraneed and decontamination carried out.
Yes | 10 | Y | RCZE | Blank |  |  | D4330F51443FAB86 | Re classified RMED | FAN |  |  |  | Blank
No | Visit | Complaint | RBNL | Blokkering | Analyzer staat in ‘fout’ modus omwille van foutcode #5784(setsafe_22), te wijten aan de punchup module die vastzit | Na manueel verwijderen van reagens pack kan de punchup getest worden in service, deze functioneert normaal zonder haperingen. Na opstart een 6-tal meting uitgevoerd zonder fouten. | SWO-339570 | Assisted Replacement | SP | SP | SP |  | Blocking | Analyzer is in ' error ' mode because of error code #5784 (setsafe_22), due to the punchup at module that is stuck | After manual removal of the reagent pack, the punchup at can be tested in service, it functions normally without any stuttering. After start-up a 6-meter run without errors.
No | Visit | Complaint | RDEU | Fr.Happel bemängelt die Funktion des AQT. Pat.ID und Pat.name sind über Aqure nicht zu sehen. | Der AQT wurde neu aufgestellt und an Aqure angeschlossen. Alle Funktionen sin o.k. Pat.ID einlesen und zuordnen funktioniert, Fr.Happel kann aber in Aquer keine Pat.Namen und auch die Pat.ID nicht sehen. | SveMe hat sich aufgeschaltet und ein Bild des Fehlers gemacht. Bitte eine Tech.line eröfnen. | SWO-341777 | Setup | His/Lis | His/Lis | His/Lis |  | Fr.Happel criticizes the function of the AQT. Pat.ID and Pat.name are not visible through Aqure. | The AQT was repositioned and connected to Aqure. All functions sin o.k. Pat.ID read and assign works, but Fr.Happel can't see pat.names or Pat.ID in Aquer. | SveMe turned on and made a picture of the error. Please open a Tech.line.
No | Visit | Complaint | RLTD | Rob Bolton 01772523007 | Site visit to investigate and repair as issue has returned. | Customer has rebooted analyser to resolve issue. Attended analyser and checked that auto print for patient results had not been switched off which it hadn't. Ran sample and this gave results and printed them. Took backup of analyser data. Powered down and fitted new cup and flash card. Reloaded windows and application software. Restored customer data. Made changes to dst and region settings. Ran calibrations and qcs to verify analyser running correctly. Analyser back running. Customer to monitor and inform radiometer for further investigation if issue persists. | SWO-341065 | Defect | HW | Basic Unit | CF Card |  | Rob Bolton 01772523007 | Site visit to investigate and repair as issue has returned. | Customer has rebooted analyser to resolve issue. Attended analyser and checked that auto print for patient results had not been switched off which it hadn't. Ran sample and this gave results and printed them. Took backup of analyser data. Powered down and fitted new cup and flash card. Reloaded windows and application software. Restored customer data. Made changes to dst and region settings. Ran calibrations and qcs to verify analyser running correctly. Analyser back running. Customer to monitor and inform radiometer for further investigation if issue persists.
No | Visit | Complaint | RSAF | To many queries on patient results and constant replacements required. | Observed the pre analytical procedure which the nursing staff follow. Observed that the incorrect procedure is being used for capillary gasses as no mixing wire is inserted into the capillary tubes as well as no end caps are used in conjunction with the mixing wire. | Gave the sisters in service training on the pre analytical procedures for the cap gasses and gave them a old mix magnet as they did not have in the ward, | SWO-341664 | User error | Pre analytical | Sample handling | Sample handling |  | To many queries on patient results and constant replacements required. | Observed the pre analytical procedure which the nursing staff follow. Observed that the incorrect procedure is being used for capillary gasses as no mixing wire is inserted into the capillary tubes as well as no end caps are used in conjunction with the mixing wire. | Gave the sisters in service training on the pre analytical procedures for the cap gasses and gave them a old mix magnet as they did not have in the ward,
No | Visit | Complaint | RIND | Single Parameter error . | Single parameter error . It is found that there are some corrosion on the sensor board . | Re-installed the cassette after flushing it . Problem not solved . Replaced the sensor cassette . Problem solved . Need to change the sensor board . | SWO-339997 | [5 empty] | Single Parameter error . | Single parameter error . It is found that there are some corrosion on the sensor board . | Re-installed the cassette after flushing it . Problem not solved . Replaced the sensor cassette . Problem solved . Need to change the sensor board .
No | Visit | Complaint | RCHN | 显示屏故障 | 显示屏故障，借903-239一个 | 备件未使用，线路插拔后修复(By Ran) Reconnected the cable, then OK | SWO-339349 | [5 empty] | Display failure | Display fault, borrow 903-239 one | Spare parts not used, line plug-in repair (By Ran) Reconnect the cable, then OK
No | Visit | Complaint | RCHN | 温度错误 | 温度错误，借903-644,903-199,903-183各一个 | 更换903-644,903-199后修复(By Ran) | SWO-338719 | Defect | HW | Basic Unit | PCB, Sensor interface | temp | Temperature error | Temperature error, borrow 903-644, 903-199,903-183 each | Replacement 903-644,903-199 post-repair (By Ran)
No | Visit | Complaint | RCHN | GLU,LAC没有反应 | GLU,LAC没有反应，借903-581一个 | 更换903-581后修复(By Ran) | SWO-339466 | Defect | HW | Basic Unit | PCB, Sensor interface | Glu/Lac | GLU, LAC didn't respond. | GLU, LAC did not respond, borrowed 903-581 a | Repair after replacement 903-581 (By Ran)
No | Visit | Complaint | RCHN | LAC定标失败 0378，1310 | LAC定标失败 0378，1310 ，借903-581一个。 | 更换903-581备件后修复(By Ran) | SWO-338851 | Defect | HW | Basic Unit | PCB, Sensor interface | Glu/Lac | LAC calibration failure 0378,1310 | LAC calibration failure 0378,1310 , borrow edith of 903-581. | Repair after replacement of 903-581 spare parts (By Ran)
No | Visit | Complaint | RCHN | LAC,GLU灵敏度超范围 | LAC,GLU灵敏度超范围，借903-581一个 | 更换903-581备件后修复(By Ran) | SWO-339388 | Defect | HW | Basic Unit | PCB, Sensor interface | Glu/Lac | LAC, GLU sensitivity over range | LAC, GLU sensitivity over range, borrowed 903-581 one | Repair after replacement of 903-581 spare parts (By Ran)
No | Visit | Complaint | RCHN | PCT定标不过 | PCT定标不过，借944-394（13955）一个 | 重新定标后修复(By Ran) Recal, then OK | SWO-339488 | Assisted Replacement | Consumables | Test Kit | cartridge, CRP | mail | PCT calibration is not the | The PCT bid, however, borrows 944-394 (13955) a | Re-calibration repair (By Ran) Recal, then OK
No | Visit | Complaint | RCHN | 报错4548 | 报错4548，借903-028,944-214（13673）各一个 failed when cal. | 更换903-028及944-214（13673）重新定标后修复(By Ran) | SWO-338757 | Defect | HW | Cover and Chassis | PCB, Cabinet top | 4548 | 4548 error | 4548, borrow 903-028, 944-214 (13673) each a failed time cal. | Replace 903-028 and 944-214 (13673) re-calibration repair (By Ran)
No | Visit | Complaint | RCHN | 报错1024.处理盘卡盘 | 报错1024.处理盘卡盘，需要借件903-236、903-364、944-212（13822） | 更换903-236、903-364、944-212（13822）备件后修复，和重新定标(By Ran) | SWO-338767 | Defect | HW | Process Module | Optical Unit | 1024 | 1024. Process the disk | 1024. Process disk, need to borrow 903-236, 903-364, 944-212 (13822) | Repair and re-calibration (By Ran) after replacement of 903-236, 903-364, 944-212 (13822) spare parts
No | Visit | Complaint | RCHN | 黑屏 | 黑屏，借910-298一个 | 更换910-298备件后修复(By Ran) | SWO-339432 | Defect | HW | Analyzer Control Module | Display |  | Black screen | Black screen, borrow 910-298 one | Repair after replacement of 910-298 spare parts (By Ran)
No | Visit | Complaint | RCHN | 存储盘卡盘，测试卡舱门卡门 | 存储盘卡盘，测试卡舱门卡门 | 清理灰尘后修复(By Ran) After cleaning the dust, then ok. | SWO-339961 | Cleaned | HW | Cover and Chassis | catridge door |  | Memory disk card, test card door Carmen | Memory disk card, test card door Carmen | 清理灰尘后修复(By Ran) After cleaning the dust, then ok.
No | Visit | Complaint | RESP | Error 793 Fallo en calibración de bombas | Comprobación de registros de datos del equipo. | Sustitución de la PCB del módulo de pH $BG, cal de bombas, verificación final de funcionamiento. | SWO-342115 | Defect | HW | pH/BG Module | Pump tubes | 793 | Error 793 Pump calibration failure | Checking computer data records. | PCB replacement of the pH module $BG, pump lime, final operating check.
Yes | 16 | Y | RPAC | Blank | Six Monthly service required on teh analyser. Install the METII Module, allowing analyser to measure Creatinine. |  | D4345670BA0D1C0D | Re classified RMED | semi annual service |  |  |  | Blank | Six Monthly service required on teh analyser. Install the METII Module, allowing analyser to measure Creatinine.
No | Visit | Complaint | RSCH | Immer wieder Lac unstabil während Blutmessung 476 | Fehler 0476: Lac unstabil während Blutmessung | Allgemeine Kontrolle ok. EL $MET Modul leihweise mit RSCH Modul ausgetauscht. Kal. + QC i.o. Mehrere Kapillärmessungen durchgeführt ok. *** Kunde wird in den nächsten Tagen testen *** | SWO-340106 | Defect | HW | El/Met Module | El/Met Module |  | Time and again Lac unstable during blood measurement 476 | Error 0476: Lac unstable during blood measurement | General control ok. EL $MET module on loan with RSCH module. Kal. + QC i.o. Several capillary measurements performed ok. Customer will test in the next few days ***
No | Visit | Complaint | RAME | Oxi calibration dérive | Calibration dérive halogen lampe trop haute | Remplacé lamp unit | SWO-340524 | Defect | HW | Oxi module | Lamp unit | halogen | Oxi calibration drifts | Calibration drifts halogen lamp too high | Replaced lamp unit
No | Visit | Complaint | RESP | Error 1278 flujo medido por encima del límite. | Error 1278 flujo medido por encima del límite. | Se cambia tubos, filtro, unidad de lavado copas y se hacen test de servicio, se comprueba correcto funcionamiento con muestras de blanco y qc. | SWO-340358 | Defect | HW | Process Module | Wet section | 1278 | Error 1278 flow measured above the limit. | Error 1278 flow measured above the limit. | It changes tubes, filter, washing unit cups and service tests are made, correctly functioning is checked with white and qc samples.
No | Visit | Complaint | RFRA | Problème de remonté du bloc optique | Contrôle des différents sous éléments. Contrôle du déplacement des cupules. | Nettoyage du capteur de présence de cupule dans l’incubateur. Essais sur tube de nettoyage et tube patient | SWO-340135 | Cleaned | HW | Incubator Module | sensor, cup |  | Optical block re-up problem | Control of the various sub-elements. Control of the movement of cups. | Cleaning the cup presence sensor in the incubator. Tests on cleaning tube and patient tube
No | Visit | Complaint | RFRA | une lenteur de l'interface logiciel lors du remplacement de consommable et passage de bilan | Le problème de lenteur du software . Ce problème était également présent sur les 2 analyseurs précédent BLocage retrait du pack | Inversion des branchement RJ45 et 220V ondulé entre ABL90 et AQT90 Flex Vérification du fonctionnement du punch OK Test de retrait et insertion pack de solution OK | SWO-340257 | Defect | HW | Analyzer Control Module | PC | mail | a slow software interface when replacing consumables and passing balance sheet | The problem of slow software. This problem was also present on the previous 2 analyzers BLocage withdrawal from the pack | Inversion of the RJ45 and 220V corrugated connection between ABL90 and AQT90 Flex Checking how the OK punch works Withdrawal test and insertion solution pack OK
No | Visit | Complaint | RAME | Noisy printer | Noisy printer | Replace printer | SWO-341384 | Defect | HW | Printer Module | Printer Module |  | Noisy printer | Noisy printer | Replace printer
No | Visit | Complaint | RDEU | Vacuumfehler | Einlass positioniert Einlassdichtung nicht richtig. | Mechanische Probleme. Einlassmodul erneuert. Gerät überprüft. | SWO-341034 | Defect | HW | Inlet module | Inlet module |  | Vacuumfehler | Inlet does not position inlet seal correctly. | Mechanical problems. Inlet module renewed. device checked.
No | Visit | Complaint | RLTD | John Trenfield Tel: 07900 407208 | Touch screen | Replaced touch screen, calibrated and test barcode reader | SWO-341644 | Defect | HW | DMS_PC | Touch screen |  | John Trenfield Tel: 07900 407208 | Touch screen | Replaced touch screen, calibrated and test barcode reader
No | Visit | Complaint | RLTD | Sarah 01604 54 55 42 | Passed from on call engineer | On arrival analyser failing all temperature tests, straightened ribbon cable and rebooted, now ok. Analyser failing calibrations and QC’s, maybe due to lack of sensor cassette conditioning due to inlet stuck open. Reran calibration and QC’s, now ok. | SWO-341645 | Adjustment | HW | Auto Inlet module | cable |  | Sarah 01604 54 55 42 | Passed from on call engineer | On arrival analyser failing all temperature tests, straightened ribbon cable and rebooted, now ok. Analyser failing calibrations and QC’s, maybe due to lack of sensor cassette conditioning due to inlet stuck open. Reran calibration and QC’s, now ok.
No | Visit | Complaint | RLTD | Steve Ramsley reported that the inlet gasket keeps going round in circles. Steve 01206 742493 | Customer concerns about the number of inlet gaskets being used and that new inlet gaskets are failing. | Checked over inlet and fluid transport system. No issues found. | SWO-342065 | No fault found | HW | Auto Inlet module | Inlet gasket |  | Steve Ramsley reported that the inlet gasket keeps going round in circles. Steve 01206 742493 | Customer concerns about the number of inlet gaskets being used and that new inlet gaskets are failing. | Checked over inlet and fluid transport system. No issues found.
No | Visit | Complaint | RLTD | Customer reported a very noisy fan. | Noisy fan on pc unit. | Attempted to replace fan, but the connection on the new fan is different, will have to return with a new fan. | SWO-342066 | Revisit required | HW | Basic Unit | Fan |  | Customer reported a very noisy fan. | Noisy fan on pc unit. | Attempted to replace fan, but the connection on the new fan is different, will have to return with a new fan.
No | Visit | Complaint | RDEU | Fehlerbehebung erforderlich | keine WLAN Verbindung ins Hausnetz | Überprüfung und verschiedene Test mit Konfiguration des WLAN Adapters - ohne Erfolg - Verbindung wird mehrfach kurz aufgebaut und sofort wieder getrennt XP Embedded SP3 Probleme mit WLAN Verschlüsselung WPA2-PSK AES Gerät an anderen Standort verbracht und ans LAN angeschlossen Gerät soll an seinen ursprünglichen Standort vorübergehend mit Hilfe von Powerline Adaptern angeschlossen werden Herr Lempe wünscht eine Aufrüstung auf Win7, damit die Anbindung per WLAN erfolgen kann LIS Anbindung mit TechniData eingerichtet und getestet (Patientenbericht und QC) Einheiten am ABL umgestellt (Drücke auf kPa und Lac auf mmol $L) Mitteilung technischer Probleme wird eröffnet | SWO-340959 | Setup | SW | His/Lis | Wireless connection |  | Troubleshooting required | no Wi-Fi connection to the house network | Verification and various tests with configuration of the WLAN adapter - without success - connection is established several times short and immediately disconnected again XP Embedded SP3 Wi-Fi Encryption Issues WPA2-PSK AES Device moved to another location and connected to the LAN Device should be temporarily connected to its original location using Powerline adapters Mr Lempe wants to upgrade to Win7 so that the connection can be via Wi-Fi LIS connection with TechniData set up and tested (Patient Report and QC) Units at the ABL (press kPa and Lac to mmol $L) Notification of technical problems opens
No | Visit | Complaint | RCHN | 接口板腐蚀 | 接口板腐蚀，借902-972一个 | 重新安装测试卡，清洁仪器正常。(By Ran) | SWO-340777 | Defect | HW | Basic Unit | PCB, Sensor connector assembly |  | Interface plate corrosion | Interface plate corrosion, borrow 902-972 one | Reinstall the test card to clean the instrument properly. (By Ran)
No | Visit | Complaint | RCHN | 齿轮断裂 | 断裂,借903-584 | 更换903-584后修复(By Ran) | SWO-340008 | Defect | HW | Aspiration Module | Tube holder |  | Gear break | Break, borrow 903-584 | Repair after replacement 903-584 (By Ran)
No | Visit | Complaint | RKK | カセット不良 | ｷｬﾙｴﾗｰ | カセット交換 | SWO-340754 | [5 empty] | Bad cassette | Cal Error | Cassette replacement
No | Visit | Complaint | RDEU | KAL Driften Na und Ca | KAL Driften Na und Ca | Überprüfung Analysator - Erneuerung Referenz Elektrode - Funktionsüberprüfung etc. | SWO-340947 | Assisted Replacement | Electrode | Electrode | Reference |  | KAL Drifting Na and Ca | KAL Drifting Na and Ca | Review Analyzer - Renewal Reference Electrode - Function Check etc.
No | Visit | Complaint | RDEU | KAL Driften Lactat | KAL Driften Lactat | Reinigung etc. Elektroden-Kontakte - Messventil Air im EL $MET Modul demontiert und gereinigt - Rückmontage und Funktionsüberprüfung etc. | SWO-340948 | Cleaned | HW | El/Met Module | electrode contact | Lac | KAL Drifting Lactate | KAL Drifting Lactate | Cleaning, etc. Electrode contacts - Air measuring valve in the EL $MET module disassembled and cleaned - reassembly and function verification etc.
No | Visit | Complaint | RLTD | Kauser Tel: 01619 226 498 | Site visit to investigate cartridge wheel rotate error. | Arrived on site and attended analyser. Found analyser working and in use. Checked logs and found error 5689 Rotate storage wheel error. Message for this error suggests an obstruction in storage wheel rotation. Opened front cover and entered service menus. Ran safe positions a number of times and monitored mechanical movement for signs of issue. All positioning ran correctly and passed. Entered the barcode scanning for test cartridges and ran the scan all process. Part way round scanning when group of three test cartridges reached the front of the analyser, the rotation stopped and gave error. Ran safe positions which passed. Retested in scanning process and same occurred. Removed Co er from storage wheel and retested. When wheel stopped with error I noticed that when I manually turned the wheel gently, the stuck cartridge dropped slightly. Rerun Process and error again. Manual rotation showed cartridge drop again. Swapped the stuck cartridge around with another in that group of three to test if the cartridge was causing the obstruction. No issue with cartridge as the same error occurred with a different cartridge. Removed the first two cartridges and performed the same test. Same issue occurred. No sign of anything g catching inside. Foil on bottom half of cartridge from punched out test cups was not touching base plate and lifting cartridge. Powered down analyser and removed whole storage wheel in order, making inspection at each stage. Brass cartridge lock which is at the front of storage wheel and pulls in the test cartridges looked a little low. Tightened the 4 screws that secures the locking device to the upper cover to lift it up. Refitted all parts, fitted new ups and powered up. retested the cartridges in the storage wheel multiple times. No reoccurrence of the issue seen. Exited service programs and ran two cleaning cycles followed by a system check. All passed without issue. Ran qcs to test fluid and cup transport. Qcs passed. | SWO-340979 | Adjustment | HW | Storage Module | Cartridge lock | 5689 | Kauser Tel: 01619 226 498 | Site visit to investigate cartridge wheel rotate error. | Arrived on site and attended analyser. Found analyser working and in use. Checked logs and found error 5689 Rotate storage wheel error. Message for this error suggests an obstruction in storage wheel rotation. Opened front cover and entered service menus. Ran safe positions a number of times and monitored mechanical movement for signs of issue. All positioning ran correctly and passed. Entered the barcode scanning for test cartridges and ran the scan all process. Part way round scanning when group of three test cartridges reached the front of the analyser, the rotation stopped and gave error. Ran safe positions which passed. Retested in scanning process and same occurred. Removed Co er from storage wheel and retested. When wheel stopped with error I noticed that when I manually turned the wheel gently, the stuck cartridge dropped slightly. Rerun Process and error again. Manual rotation showed cartridge drop again. Swapped the stuck cartridge around with another in that group of three to test if the cartridge was causing the obstruction. No issue with cartridge as the same error occurred with a different cartridge. Removed the first two cartridges and performed the same test. Same issue occurred. No sign of anything g catching inside. Foil on bottom half of cartridge from punched out test cups was not touching base plate and lifting cartridge. Powered down analyser and removed whole storage wheel in order, making inspection at each stage. Brass cartridge lock which is at the front of storage wheel and pulls in the test cartridges looked a little low. Tightened the 4 screws that secures the locking device to the upper cover to lift it up. Refitted all parts, fitted new ups and powered up. retested the cartridges in the storage wheel multiple times. No reoccurrence of the issue seen. Exited service programs and ran two cleaning cycles followed by a system check. All passed without issue. Ran qcs to test fluid and cup transport. Qcs passed.
No | Visit | Complaint | RIND | Service Visit. | Till 31 march to present user did 47 sample in which 6 sample has aborted. | Analyze the patient result in which 4 aborted sample message was 1067: Sample measurement aborted, 1065: Parameter measurement aborted, 1218: Insufficient sample available for measurement. These error codes comes with sample message. Open the analyzer front panel as well as back panel, remove inlet wheel and remove the Hct block and clean that, reconnect it and calibrate the needle (washblock only) in service programe. After calibration check the pumps and valves in wet section, run the sample in demo mode 3 times found ok. Need to observe the the analyzer for 15 days. | SWO-340946 | Cleaned | HW | Aspiration Module | Needle Wash Unit | 1218 | Service Visit. | Till 31 march to present user did 47 sample in which 6 sample has aborted. | Analyze the patient result in which 4 aborted sample message was 1067: Sample measurement aborted, 1065: Parameter measurement aborted, 1218: Insufficient sample available for measurement. These error codes comes with sample message. Open the analyzer front panel as well as back panel, remove inlet wheel and remove the Hct block and clean that, reconnect it and calibrate the needle (washblock only) in service programe. After calibration check the pumps and valves in wet section, run the sample in demo mode 3 times found ok. Need to observe the the analyzer for 15 days.
No | Visit | Complaint | RDEU | 20.04.2018 08:40Uhr Frau Thilmann 03834865509: Keine Bluteingabe möglich | Hauptfehler: 5658: Rotate Incubator Wheel 37 | Incubator Wheel entnommen, gereinigt; Kammer und Sensoren gereinigt Cup Sensor gereinigt Incubator Wheel und Cup Sensor kalibriert Test i.O. | SWO-341091 | Cleaned | HW | Incubator Module | PCB, Incubator | 5658 | 20.04.2018 08:40pm Ms Thilmann 03834865509: No blood entry possible | Hauptfehler: 5658: Rotate Incubator Wheel 37 | Incubator Wheel removed, cleaned; Chamber and sensors cleaned Cup sensor cleaned Incubator Wheel and Cup Sensor calibrated Test i.O.
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | LAC定标失败，0378，1310 | 更换测试卡后修复 | SWO-341366 | [5 empty] | Test card, reagent package failure | LAC calibration failure, 0378,1310 | Repair after replacing test card
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | LAC定标，响应错误；0378,1310 | 更换测试卡后修复 | SWO-341348 | Assisted Replacement | SC | SC | Lac |  | Test card, reagent package failure | LAC calibration, response error; 0378,1310 | Repair after replacing test card
No | Visit | Complaint | RCHN | 报错3412液路故障 | error code 3412液路故障and the tube inserts were broken.，借2个903-584和903-329一个 | 更换2个903-584和903-329后修复(By Ran) | SWO-340684 | Defect | HW | Aspiration Module | Needle Wash Unit | 3412 | Reported error 3412 liquid road fault | error code 3412 liquid road fault and the tube inserts have been broken. borrow 2 903-584 and 903-329 one | Replacement of 2 903-584 and 903-329 post-repair (By Ran)
No | Visit | Complaint | RKK | ｲﾝﾚｯﾄｶﾞｽｹｯﾄからの液漏れ及びﾘﾝｽｴﾗｰ Inlet not filled properly during Rinse | ケース詳細の通り | ｲﾝﾚｯﾄﾌﾟﾛｰﾌﾞ交換、ｲﾝﾚｯﾄｶﾞｽｹｯﾄ交換。共に在庫。キャル、QC確認異常なし。 | SWO-341481 | Defect | HW | Inlet module | Inlet gasket |  | Leakage and rinse errors from inlet gaskets Inlet not filled properly during Rinse | As the case details | Inlet probe replacement, inlet gasket replacement. Both in stock. Cal, QC confirmation no abnormalities.
No | Visit | Complaint | RCHN | 报错1031 | 报错1031，借903-364一个，944-212（13315）、944-394（13849）、944-258（13918）、944-220（13459）各一张 |  | SWO-341419 | Defect | HW | Process Module | Optical Unit | 1031 | 1031 | 1031, loan 903-364 one, 944-212 (13315), 944-394 (13849), 944-258 (13918), 944-220 (13459) each
No | Visit | Complaint | RCHN | 仪器卡盘，报错5784 | 仪器卡盘，报错5784 | 取出断裂的测试卡 by jesson | SWO-341482 | Defect | Consumables | Test Kit | cartridge broken | 5784 | Instrument chuck, error 5784 | Instrument chuck, error 5784 | Remove the broken test card by jesson
No | Visit | Complaint | RAME | Lac failed after replacing SC | Lac sen will not come in. Customer has replaced the SC and has the same issue on the new one | I ran the SC simulation test and all passed. I replaced the SC myself and lac still failed sen. I will return. | SWO-342213 | Revisit required | SC | SC | Lac |  | Lac failed after replacing SC | Lac sen will not come in. Customer has replaced the SC and has the same issue on the new one | I ran the SC simulation test and all passed. I replaced the SC myself and lac still failed sen. I will return.
No | Visit | Complaint | RLTD | Gary Telephone No 01604 545 611 | Barcode reader dead. | Checked all connections, no issues found. On restart barcode reader laser was working but only beeped and didn’t enter a code. Rebooted again and is now working. | SWO-342052 | Adjustment | SW | Basic Unit | off/on | Barcode Reader | Gary Telephone No 01604 545 611 | Barcode reader dead. | Checked all connections, no issues found. On restart barcode reader laser was working but only beeped and didn’t enter a code. Rebooted again and is now working.
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 冲洗时液体传输失败 0508 | 更换试剂包后修复 | SWO-342210 | [5 empty] | Test card, reagent package failure | Liquid transfer failure during flushing 0508 | Repair after replacement of reagent kit
No | Visit | Complaint | RCHN | 测试卡安装过程中出现错误，测试卡仓门一侧的一个柱子断裂了的故障表现出来了。 | 装机时箱子里发现的一个柱子就是测试卡仓门一侧断裂掉下来的，当时现场不好拆机器，就没有处理，现在表现出来了。 | 买了胶水把这个柱子给沾上了。客户的处理包也用完了，经销商首次自己订货没有订处理包，帮他们找了一个处理包安装上去，新到了一盒Tni，培训科室定标操作。 | SWO-343134 | Defect | HW | Cover and Chassis | Cartridge door |  | There was an error during the installation of the test card, and a fault on the side of the test card door broke. | A pillar found in the box at the time of installation is the test card door side broke off, when the scene is not easy to dismantle the machine, there is no processing, now show. | I bought the glue and put the post on it. The customer's processing package also used up, the dealer for the first time did not order the processing package, help them find a processing package installed, a new box of Tni, training department calibration operation.
No | Visit | Complaint | RCHN | 卡盘 | 卡盘 | 更换测试卡后修复 | SWO-342203 | Defect | Consumables | Test Kit | Cartridge, Broken | mail | Chuck | Chuck | Repair after replacing test card
No | Visit | Complaint | RCHN | 质控不过 | 质控不过 | 清洁处理盘 by jesson | SWO-343646 | Cleaned | HW | cleaned | cleaned |  | But the quality control is not | But the quality control is not | Cleaning processing disc by jesson
No | Visit | Complaint | RDEU | KAL Driften Lactat | KAL Driften bei Lactat, Calcium und Natrium | Erneuerung Referenz Elektrode etc. - Lactat Membranierung und Überprüfung - mehrfache Kalibrationen durchgeführt - Funktionsüberprüfung etc. Labor meldet sich falls Driften weiterhin auftreten ! | SWO-341944 | Assisted Replacement | Electrode | Membrane | Lac | lot 8807 | KAL Drifting Lactate | KAL Drifting in lactate, calcium and sodium | Renewal reference electrode etc. - lactate membraneing and verification - multiple calibrations performed - function verification etc. Laboratory reports if drifting continues to occur !
No | Visit | Complaint | RDEU | Gerät macht laute Geräusche. Konfiguration erforderlich. |  | Rollenkranz überprüft, Pumpenschlauch am pH BG Modul erneuert, Gerät kalibriert, Geräte Parameterprofil nach Kundenwünschen konfiguriert. | SWO-343848 | Defect | HW | pH/BG Module | pump wheel |  | Device makes loud noises. Configuration required. |  | Roller ring checked, pump hose renewed on pH BG module, device calibrated, device parameter profile configured according to customer requirements.
No | Visit | Complaint | RDEU | 20.04.2018 08:40Uhr Frau Thilmann 03834865509: Keine Bluteingabe möglich | im Service versucht das Incubator Wheel zu kalibrieren, aber ohne Erfolg Kal wird mit Error 42 abgebrochen stepweise kann verfahren werden, auch Shake fkt. - Position CW oder CCW fkt. nicht | Incubator PCB bestellt | SWO-341894 | Revisit required | HW | Incubator Module | PCB, Incubator |  | 20.04.2018 08:40pm Ms Thilmann 03834865509: No blood entry possible | in service tries to calibrate the Incubator Wheel, but to no avail Kal is aborted with Error 42 step by step can be done, also Shake fkt. - Position CW or CCW fkt. Not | Incubator PCB ordered
No | Visit | Complaint | RHUN | több oldat esetében is pumpa kalibrációs hibát jelez a készülék. Szerviz tesztben látható, hogy a pumpa kalibrációs faktor nagyon szór. OXI modul pinch valve beragadt | több oldat esetében is pumpa kalibrációs hibát jelez a készülék. Szerviz tesztben látható, hogy a pumpa kalibrációs faktor nagyon szór. OXI modul pinch valve beragadt | csere készülékre áttelepített fogyóanyagok rendben üzemelnek. Pinch valve beragadás megszűntetése, karbantartási alkatrészek cseréje, kalibráció, tHb kalibráció, QC mérés | SWO-342533 | Defect | HW | Oxi module | Pinch valve |  | For several solutions, a pump calibration fault is indicated. Service test shows that the pump calibration factor is very good. OXY Module Pinch valve stuck | For several solutions, a pump calibration fault is indicated. Service test shows that the pump calibration factor is very good. OXY Module Pinch valve stuck | Replacement equipment is properly operated. Pinch valve termination, maintenance component replacement, calibration, tHb calibration, QC measurement
No | Visit | Complaint | RKOR | Replace the Solution pack | Replace the New SP Fluid transport test -> failed Replace the another new SP | Replace the another new SP | SWO-345158 | Assisted Replacement | SP | SP | SP |  | Replace the Solution pack | Replace the New SP Fluid transport test -> failed Replace the another new SP | Replace the another new SP
Yes | Visit | Complaint | RDEU | : 13.04.2018 14:15 Fr. Anders 023312016165: Schließfach SC fährt nicht mehr zu 13.04.2018 15:15 Fr. Anders 023312016165: Fach wieder zu; kein FSE erforderlich 16.04.2018, 10.50 Uhr, Herr Holzhausen, 023312016165: Herr Holzhausen hat das SC fach nur mit mech. Hilfe zu bekommen. |  |  | SWO-343097 | Re classified RMED | phone call |  |  |  | : 13.04.2018 14:15 Fr. Anders 023312016165: Locker SC no longer goes to 13.04.2018 15:15 Fr. Anders 023312016165: Tray again; no FSE required 16.04.2018, 10.50 a.m., Mr. Holzhausen, 023312016165: Mr. Holzhausen has the SC fach only with mech. to get help.
No | 8 | Y | RPAC | Blank | Lac failed on calibration and QCs |  | D43AA1190809D26C | Assisted Replacement | SC | SC | Lac |  | Blank | Lac failed on calibration and QCs
No | Visit | Complaint | RLTD | Daphne Kelly Tel 01689 864989 (or contact 01689 863517) | The Analyser inlet was in the capillary position when I arrived. Entered service mode to investigate further. Checked SC bay, SP area, inlet probe and gasket. Inlet gasket was dirty upon inspection and has not been replaced as stated by the customer. Checked old inlet probe and found it was blocked. | Performed manual flush and replaced the inlet gasket. Straightened, unblocked old original inlet probe and handed back to the customer as a spare. | SWO-342427 | Defect | HW | Auto Inlet module | Probe | clot | Daphne Kelly Tel 01689 864989 (or contact 01689 863517) | The Analyser inlet was in the capillary position when I arrived. Entered service mode to investigate further. Checked SC bay, SP area, inlet probe and gasket. Inlet gasket was dirty upon inspection and has not been replaced as stated by the customer. Checked old inlet probe and found it was blocked. | Performed manual flush and replaced the inlet gasket. Straightened, unblocked old original inlet probe and handed back to the customer as a spare.
No | Visit | Complaint | RSCH | Rep. Nadel | Nadel krumm durch Anwender Einlasskalibrierfehler Sensor bay schliesst nicht (Schliesst und öffnet wieder) | Service-Menu lässt sich Einlass nicht Kalibrieren Gerät zerlegt i.o Flachkabel von Einlass war Fabrikseitig unter Schraube von Gehäuse eingeklemmt, kein Ersatz dabei Gerät wieder gestartet i.o Service-Menu alles i.o 50 min Initialisierung i.o System i.o Data Log und Service.dump herausgezogen Alles fotografiert *** Flachkabel Einlass muss bei nächster Wartung ersetzt werden *** Nachdem Gerät auf Abteilung gebracht wurde, Lac reaktionsfehler Mehrere Kal. durchgeführt, ohne Erfolg Sensorkassette ersetzt Danach pH und alle Elektrolyten Ansprechfehler, Impedanzfehler und Kal. Backlogfehler Im Service-Menu alles i.o Alte SC wieder installiert => gleiche Fehler Mit Gerät nach Thalwil zur Reparatur Zunge Lichtschranke war angerissen, ersetzt i.o Flachkabel Einlass-Modul ersetzt i.o System i.o 26.04.2018 retour zum Kunde | SWO-342578 | Defect | HW | Auto Inlet module | cable |  | Rep. Nadel | Needle crooked by user Inlet calibration error Sensor bay does not close (Closes and opens again) | Service menu cannot be calibrated Device disassembled i.o Flat cable from inlet was factory side under screw clamped by housing, no replacement Device restarts i.o Service menu all i.o 50 min Initialization i.o System i.o Data Log and Service.dump pulled out Everything photographed Flat cable inlet must be replaced at next maintenance *** After device has been put to department, Lac reaction error Several Kal. carried out, to no avail Sensor cassette replaced Then pH and all electrolytes response errors, impedance errors and cal. Backlog error In the service menu everything i.o Old SC reinstalled => same error With device to Thalwil for repair Tongue light barrier was torn, replaces i.o Flat cable inlet module replaces i.o System i.o 26.04.2018 return to the customer
Yes | Visit | Complaint | RCHN | 给新装机没有货,科室还有一盒PCT13708存货但是没有定标卡,找西宁昌和买了一张定标卡给定标先用上 | PCT13708有货无定标卡 | 给找定标卡 | SWO-343140 | Re classified RMED | Logistical SWO |  |  |  | To the new installation of no goods, the department also has a box of PCT13708 inventory but no bid card, find XiningChang and bought a bid card given the first use of the standard | PCT13708 Without A Bid Card | Find a calibration card.
No | Visit | Complaint | RCHN | 报错1031 | 报错1031，借903-364和944-394（13849）一张 | 更换903-364备件后发现线头断裂，需要重新借件，固重新开单by tim | SWO-341661 | Defect | HW | Process Module | Optical Unit | 1031 | 1031 | 1031, borrow 903-364 and 944-394 (13849) one | Replacing 903-364 spare parts reveals a broken wire head and needs to be re-borrowed, solid re-opening by tim
No | Visit | Complaint | RCHN | 质控结果超范围 | 质控结果超范围 | 处理盘清理，重新定标后修复 by jesson | SWO-343054 | User error | HW | PM kit | PM kit |  | Quality control results over range | Quality control results over range | Process disk clean-up, fix by jesson after recalibration
No | Visit | Complaint | RKK | Neon number too low / Neon Lamp intensity too low | Neon number too low $ Neon Lamp intensity too low | 修理完了の為、OXIモジュール返却。代替えモジュール引き上げ。Creaメンブラン、溶液類定期交換。タンパク除去実施。エアーライン洗浄。 | SWO-342198 | Defect | HW | Oxi module | Lamp unit | neon low | Neon number too low / Neon Lamp intensity too low | Neon number too low / Neon Lamp intensity too low | Oxi module returned for repair completion. Alternative module raise. Crea membrane, solution regular exchange. Protein removal was carried out. Air line cleaning.
No | Visit | Complaint | RDEU | Keine Verbindung zum AQure | Kein Fehler Lokal festzustellen. Weitere Überprüfung durch POCT IT oder Haus EDV erforderlich. | Ping vom Gerät an 10.242.25.100 gesendet und OK. Kontakt mit Radiometer POCT IT aufgenommen. RDEU POCT IT hat das AQure System überprüft und konnte keinen Fehler feststellen. Crossover -Verbindung zwischen ABL und Notebook aufgebaut und OK. | SWO-343849 | No fault found | His/Lis | His/Lis | His/Lis |  | No connection to AQure | No error to determine local. Further verification required by POCT IT or Haus EDP. | Ping from the device to 10.242.25.100 and OK. Contact edifiable to Radiometer POCT IT. RDEU POCT IT checked the AQure system and found no error. Crossover connection between ABL and notebook built and OK.
No | Visit | Complaint | RDEU | 23.04.2018 Frau Kuhrau 023319742111 meldet, dass der Monitor dunkel ist. | Gerät steht ausgeschaltet im Flur ,CPU defekt. | Neue CPU eingebaut und in betrieb genommen. Mehrfach kalibriert und Messkammern gereinigt. | SWO-343096 | Defect | HW | DMS_PC | CPU |  | 23.04.2018 Ms Kuhrau 023319742111 reports that the monitor is dark. | Device is switched off in the hallway ,CPU defective. | New CPU installed and put into operation. Multiple calibrated and measuring chambers cleaned.
No | Visit | Complaint | RAME | low pco2 QC recovery for Level 1 & 2. Results had been in range. Also observed general erratic tendencies. | Follow up to previous SWO-336567. Customer is observing consistent QC recovery for pCO2 with new pCO2 electrode but L1 is recovering at the low end of the customer's established range. Previously tested pCO2 electrode (used) was recovering close to the established mean. Customer understands a different electrode may recover differently but still within the package and group range. The concern the customer has in not wanting to experience an unexpected shift in recovery values of QC. | Replaced the pH oring with item 834-618 (as suggested by specialist). Performed calibration and QC. L1 QC changed from 65.6 (65.3-70.0) to 66.7. Customer will monitor for stability and recovery values. | SWO-342735 | User expectation | Specifications | QC | electrode | pCO2 | low pco2 QC recovery for Level 1 & 2. Results had been in range. Also observed general erratic tendencies. | Follow up to previous SWO-336567. Customer is observing consistent QC recovery for pCO2 with new pCO2 electrode but L1 is recovering at the low end of the customer's established range. Previously tested pCO2 electrode (used) was recovering close to the established mean. Customer understands a different electrode may recover differently but still within the package and group range. The concern the customer has in not wanting to experience an unexpected shift in recovery values of QC. | Replaced the pH oring with item 834-618 (as suggested by specialist). Performed calibration and QC. L1 QC changed from 65.6 (65.3-70.0) to 66.7. Customer will monitor for stability and recovery values.
Yes | Visit | Complaint | RDEU | 23.04.2018, Herr Finger, Tel: 023729082525: Demo keys laufen ab; Kalibrationsfehler bei pH |  |  | SWO-343090 | Re classified RMED | Phone support |  |  |  | 23.04.2018, Mr. Finger, Tel: 023729082525: Demo keys expire; Calibration error at pH
No | Visit | Complaint | RDEU | pH/BG upper Liquid Sensor error | pH $BG upper Liquid Sensor error | Ventile überprüft und gereinigt Defekte Airventilmembran im pH $BG Modul ersetzt LS Justage und Pumpenkal Kal- und austehende QC ausgeführt | SWO-342460 | Defect | HW | pH/BG Module | LS Sensor |  | pH/BG upper Liquid Sensor error | pH /BG upper Liquid Sensor error | Valves checked and cleaned Defective airvalve membrane replaced in pH $BG module LS Adjustment and Pump Scale Kal- and outstanding QC executed
No | Visit | Complaint | RKOR | Manual Fushing screen is displayed. | Analyzer Status check Calibration, Quality Control, Replacement, Activity logs check Cleanig Inlet gasket with holder Fiuld transport check Manual Flushing several times | Cleanig Inlet gasket with holder Fiuld transport check Manual Flushing several times | SWO-345162 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | Manual Fushing screen is displayed. | Analyzer Status check Calibration, Quality Control, Replacement, Activity logs check Cleanig Inlet gasket with holder Fiuld transport check Manual Flushing several times | Cleanig Inlet gasket with holder Fiuld transport check Manual Flushing several times
No | Visit | Complaint | RDEU | Vacuumfehler. Gerät fordert ständiges Wechseln des Einlasses und des Lösungspacks | Ventilschlauch gerissen. | Ventilschlauch erneuert. Siehe Jahreswartung. Komplette Startphase abgewartet, da Gerät vom Anwender über Nacht ausgeschaltet wurde. Kassette musste erneuert werden. Gerät überprüft. | SWO-343419 | Defect | HW | Oxi module | Pinch valve tube |  | Vacuum error. Device requires constant change of inlet and solution pack | Valve hose torn. | Valve hose renewed. See Annual Maintenance. Complete start-up phase waited, as device was switched off by the user overnight. Cassette had to be renewed. device checked.
No | Visit | Complaint | RLTD | Contact Pamela Stevenson 0203 299 9000 Ext 33222 | On arrival - analyser in intervention required due Leak detected. Found inlet probe to have slight bend, corrected so probe is straight and retested - now passing fluid tests. Allowed analyser to restart - all ok except for Glucose parameter failing calibration 0378 sensitivity and 1310 response errors. Replaced Sensor Cassette and allowed to start up - after 2 hours all parameters working except for Glucose & Lactate, both failing with 0378 & 1310 errors. Diagnosed that a new Sensor Interface pcb is required to resolve issues with Glucose & Lactate. | Corrected bend in inlet probe to resolve fluid transport issue. New Sensor Interface pcb required to resolve Glucose & Lactate issues. | SWO-344211 | Defect | HW | Auto Inlet module | Probe bend |  | Contact Pamela Stevenson 0203 299 9000 Ext 33222 | On arrival - analyser in intervention required due Leak detected. Found inlet probe to have slight bend, corrected so probe is straight and retested - now passing fluid tests. Allowed analyser to restart - all ok except for Glucose parameter failing calibration 0378 sensitivity and 1310 response errors. Replaced Sensor Cassette and allowed to start up - after 2 hours all parameters working except for Glucose & Lactate, both failing with 0378 & 1310 errors. Diagnosed that a new Sensor Interface pcb is required to resolve issues with Glucose & Lactate. | Corrected bend in inlet probe to resolve fluid transport issue. New Sensor Interface pcb required to resolve Glucose & Lactate issues.
Yes | Visit | Complaint | RFRA | Configuration rapport wdc | Configuration rapport wdc | Configuration rapport wdc $ profils d'accès. | SWO-343516 | Re classified RMED | WDC setup |  |  |  | Configuration rapport wdc | Configuration rapport wdc | Setting up report wdc $ access profiles.
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | LAC,GLU定标不过 | 更换测试卡后修复 | SWO-343056 | [5 empty] | Test card, reagent package failure | LAC, GLU calibration but | Repair after replacing test card
No | Visit | Complaint | RCHN | 报错1223和1227 | 报错1223和1227，需要借件903-329（903-672） | 校准了针后修复 by jesson | SWO-342442 | Defect | HW | Aspiration Module | Needle Wash Unit | 1223 | 1223 and 1227 | 1223 and 1227, need to borrow 903-329 (903-672) | Fix by jesson after calibration of needle
No | Visit | Complaint | RCHN | lac定标失败，报错3412 | 定标失败，报错3412 | 更换903-112备件后修复 by jesson | SWO-345104 | Defect | HW | Process Module | PCB, Wet section | 3412 | Lac bid failed, error 3412 | Failed to bid, error 3412 | Repair by jesson after replacement of 903-112 spare parts
No | Visit | Complaint | RCHN | 汽瓶漏气 | 汽瓶漏气，借902-505一个 | 更换备用机后修复 by jesson | SWO-342428 | Defect | HW | gas module | Flow regulator | 902-505 | Gas bottle leak | Cylinder leak, borrow 902-505 one | Fix by jesson after replacing the spare machine
No | Visit | Complaint | RAME | Thb calibration is drifting | Thb calibration are drifting | Replace hemolyzer $ Ran Thb cal | SWO-344254 | Defect | HW | Oxi module | Hemolyzer |  | Thb calibration is drifting | Thb calibration are drifting | Replace hemolyzer $ Ran Thb cal
No | Visit | Complaint | RAME | Customer called with THB failing calibrations and QC | Customer experiencing oximetry sO2 errors during level B internal QC runs. Solution pack replaced prior to arrival corrected issue. | Ran complete oximetry tests and tHB calibration to verify hemolyzer operation. All tests past. Restart calibrations and QC tests. Ran additional fluid flow tests to check for clots. No other issues found. | SWO-344280 | No fault found | No fault found | No fault found | No fault found | oxi | Customer called with THB failing calibrations and QC | Customer experiencing oximetry sO2 errors during level B internal QC runs. Solution pack replaced prior to arrival corrected issue. | Ran complete oximetry tests and tHB calibration to verify hemolyzer operation. All tests past. Restart calibrations and QC tests. Ran additional fluid flow tests to check for clots. No other issues found.
No | Visit | Complaint | RAME | I need to add the line to the ini file | To meet the CAS at site to check analyzers to see if they are how the customer wanted | The analyzer were asking for supplies and the customer didn't have any. We talked with Larry and answered questions he had and showed him how to adj his layout. I will return when the customer gets there supplies. | SWO-344255 | Setup | SW | Roma.ini | reportable range |  | I need to add the line to the ini file | To meet the CAS at site to check analyzers to see if they are how the customer wanted | The analyzer were asking for supplies and the customer didn't have any. We talked with Larry and answered questions he had and showed him how to adj his layout. I will return when the customer gets there supplies.
No | Visit | Complaint | RKK | 外装背面カバーに歪みあり | 納品直後より外装カバーに歪みあり。 | 外装背面カバー交換。 | SWO-343688 | Defect | HW | Basic Unit | back cover |  | Distortion on exterior back cover | There is a distortion in the exterior cover immediately after delivery. | Exterior back cover replacement.
No | 5 | Y | RPAC | Blank | Unable to load soultion pack |  | D43D89C2F424DCD8 | Assisted Replacement | SP | SP | chip |  | Blank | Unable to load soultion pack
No | Visit | Complaint | RDEU | 26.04.2018 08:20Uhr Frau Püttmann 020644417212/02064441308 Spritzenproben funktionieren, aber die Kapillareingabe funktioniert meistens nicht. | Zeitweise Probenabbrüche, Mechanik des Einl.modules klemmt, | Probemessungen durchgeführt, Einl.modul ersetzt, div. Dichtungen ersetzt, alle Fließwege mit Hypochlorid gespült, Einl.modul 908-570 R0256N003 ersetzt durch R0188N013 | SWO-343947 | Defect | HW | Auto Inlet module | Auto Inlet module |  | 26.04.2018 08:20Am Ms. Püttmann 020644417212/02064441308 Syringe samples work, but the capillary input usually does not work. | Intermittent sample breaks, mechanics of the single module clamps, | Test measurements carried out, single module replaced, various seals replaced, all flow paths rinsed with hypochloride, single module 908-570 R0256N003 replaced by R0188N013
No | Visit | Complaint | RLTD | Claire Abbott-Smith : 0131 2422584 | Analyser screen blank, customer has rebooted | Replaced SD card with new Replaced inlet gasket which was damaged and provided additional spare | SWO-344086 | Defect | HW | DMS_PC | CF Card |  | Claire Abbott-Smith : 0131 2422584 | Analyser screen blank, customer has rebooted | Replaced SD card with new Replaced inlet gasket which was damaged and provided additional spare
No | Visit | Complaint | RLTD | Contact Pamela Stevenson 0203 299 9000 Ext 33222 | Glucose & Lactate consistently slow to start up or fail with sensitivity and response errors. Replaced faulty Sensor Interface pcb. Fitted new Sensor Cassette, also supplied two Sensor Cassettes as replacements for those that failed due to faulty part. Tested operation on restart - All parameters including Glucose & Lactate ready for use within 1.5 hours. | Replaced Sensor Interface pcb. | SWO-344216 | Defect | HW | Basic Unit | PCB, Sensor interface | Glu/Lac | Contact Pamela Stevenson 0203 299 9000 Ext 33222 | Glucose & Lactate consistently slow to start up or fail with sensitivity and response errors. Replaced faulty Sensor Interface pcb. Fitted new Sensor Cassette, also supplied two Sensor Cassettes as replacements for those that failed due to faulty part. Tested operation on restart - All parameters including Glucose & Lactate ready for use within 1.5 hours. | Replaced Sensor Interface pcb.
No | Visit | Complaint | RCHN | 接口板腐蚀 | 接口板腐蚀，借902-972一个 | 更换902-972后修复 by jesson | SWO-343758 | Defect | HW | Basic Unit | PCB, Sensor connector assembly |  | Interface plate corrosion | Interface plate corrosion, borrow 902-972 one | Repair by jesson after replacement 902-972
No | Visit | Complaint | RSCH | Blank | Obstruction de l’analyseur (sporadique) | Remplacé le kit de maintenance e.o Nettoyage de la section fluidique e.o Cycle de qualité e.o | SWO-344068 | [5 empty] | Blank | Obstruction of the analyzer (sporadic) | Replaced the e.o maintenance kit Cleaning the fluidsection e.o E.o quality cycle
No | Visit | Complaint | RFRA | Pb de connexion réseau | Contrôle des configurations réseau | Changement de la prise réseau. Essais | SWO-344145 | [5 empty] | Pb network connection | Control of network configurations | Change of network socket. Tests
No | Visit | Complaint | RAME | Barcode reader inoperable | Barcode reader inoperable, doesn’t light up. | Replaced Barcode Reader and Cable. | SWO-344141 | [5 empty] | Barcode reader inoperable | Barcode reader inoperable, doesn’t light up. | Replaced Barcode Reader and Cable.
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 不均匀的QC2溶液 1541 | 更换试剂包正常。 | SWO-344386 | [5 empty] | Test card, reagent package failure | Uneven QC2 solution 1541 | The replacement kit is OK.
No | Visit | Complaint | RCHN | 报错5784 | 报错5784，借636-533一个 | 更换636-533备件后修复 by jesson | SWO-343168 | Defect | HW | Storage Module | Sensor | 5784 | 5784 | 5784, loan 636-533 one | Repair by jesson after replacement of 636-533 spare parts
No | Visit | Complaint | RCHN | 报错3415 | 报错3415，借903-112,944-214（13673）各一个 | 更换903-112,备件和944-214（13673）重新定标后修复 by jesson | SWO-343013 | Defect | HW | Process Module | PCB, Wet section | 3415 | Error 3415 | 3415, borrow edgy 903-112,944-214 (13673) each | Replace 903-112, spare parts and 944-214 (13673) after re-calibration repair by jesson
No | Visit | Complaint | RCHN | 报错1031 | 之前借了903-364备件后，发现线头断裂了重新借件905-908一个 | 更换905-908后修复 by jesson | SWO-342356 | Defect | HW | Process Module | Optical Unit | 1031 | 1031 | After borrowing 903-364 spare parts, it was found that the wire head broke and re-borrowed 905-908 one | Repair by jesson after replacement 905-908
No | Visit | Complaint | RCHN | 卡盘 | 卡盘 | 清洁传感器后修复 by jesson | SWO-343645 |  |  |  |  | mail | Chuck | Chuck | Repair by jesson after cleaning sensor
No | Visit | Complaint | RCHN | 报错3264 | 报错3264，903-198一个 | 更换903-198后修复 by jesson | SWO-342287 | Defect | HW | Analyzer Control Module | PCB, Main Board | 3264 | Error 3264 | Error 3264,903-198 | Repair by jesson after replacement 903-198
No | Visit | Complaint | RAME | Leak current detected | Leak current detected | Replaced upper Ph module valves and membrane on v6. | SWO-343664 | Defect | HW | pH/BG Module | Valve, membrane |  | Leak current detected | Leak current detected | Replaced upper Ph module valves and membrane on v6.
No | Visit | Complaint | RIND | oxymetry error | oxymetry error | observed the analyser gave complete oxymetry test calibration , then after taht switched off the analyser for 3o min then switched after that it came to raedy mode analysed the sample working fine | SWO-344599 | Cleaned | HW | Oxi module | Hemolyzer |  | oxymetry error | oxymetry error | observed the analyser gave complete oxymetry test calibration , then after taht switched off the analyser for 3o min then switched after that it came to raedy mode analysed the sample working fine
No | Visit | Complaint | RAME | Received Oximerty hardware error | Found neon lamp intensity at 24% should be 50%> | Replaced lamp. | SWO-344323 | Defect | HW | Oxi module | Lamp unit | neon low | Received Oximerty hardware error | Found neon lamp intensity at 24% should be 50%> | Replaced lamp.
No | Visit | Complaint | RAME | Aborting calibrations and QC samples | Checked data logs for errors, checked valves and liquid sensors, 597 upper liquid sensor error, | Removed modules and cleaned salts from upper liquid sensor, | SWO-344412 | Cleaned | HW | pH/BG Module | LS Sensor |  | Aborting calibrations and QC samples | Checked data logs for errors, checked valves and liquid sensors, 597 upper liquid sensor error, | Removed modules and cleaned salts from upper liquid sensor,
No | Visit | Complaint | RDEU | 20.04.2018 08:40Uhr Frau Thilmann 03834865509: Keine Bluteingabe möglich | Incubator Wheel läßt sich nicht mehr kalibrieren | Incubator PCB getauscht, dazu Gerät zerlegt und gereinigt danach konnte es wieder kalibriert werden Testmessungen durchgeführt | SWO-344229 | Defect | HW | Incubator Module | PCB, Incubator |  | 20.04.2018 08:40pm Ms Thilmann 03834865509: No blood entry possible | Incubator Wheel can no longer be calibrated | Incubator PCB exchanged, disassembled and cleaned device after that it could be calibrated again Test measurements carried out
Yes | Visit | Complaint | RAME | test | Clean new Mustang | Shine new Mustang | SWO-344702 | Re classified RMED |  |  |  |  | test | Clean new Mustang | Shine new Mustang
No | Visit | Complaint | RLTD | Issue noted by Tony Davies after fixing analyser for Temperture Errors at the weekend. | Lactate persistently failing calibration with 0378 sensitivity and 1310 response errors. Replaced faulty Sensor Interface pcb and fitted new Sensor Cassette. Allowed to start-up and verified performance - all parameters green after 1 hour 20 minutes including Lactate. Analyser ready for use. | Replaced Sensor Interface pcb. | SWO-344800 | Revisit required | HW | Basic Unit | PCB, Sensor interface | Lac | Issue noted by Tony Davies after fixing analyser for Temperture Errors at the weekend. | Lactate persistently failing calibration with 0378 sensitivity and 1310 response errors. Replaced faulty Sensor Interface pcb and fitted new Sensor Cassette. Allowed to start-up and verified performance - all parameters green after 1 hour 20 minutes including Lactate. Analyser ready for use. | Replaced Sensor Interface pcb.
No | Visit | Complaint | RAME | Frozen screen Paula Stafford 252-847-4937 | 4-26-18 Came in for display not responding. I connected a keyboard to it and it will not respond. Also the analyzer will not turn off using power button. Ordered display unit but it is on back order. Came in to see if I could adjust display. Received a call from Mark Ziegler and he has one and is sending it to me. Will return with part. 4-27-18 Went to Fedex to pick up parts. Replaced display unit and it still did not work. I opened the CPU box and found a printer unit instead do CPU unit. Went into service and still could not adjust the display. Checked device manager to see if display driver wasn't flagged and it was not. Spoke with Mark Ziegler and he is sending a CPU unit this time. | 4-26-18 Went in to service and display calibration but display will not respond. Checked all connections. 4-27-18 Tried display unit. I had ordered a CPU unit but received a printer unit. Mark Ziegler is sending the correct part. | SWO-344007 | Revisit required | HW | DMS_PC | CPU |  | Frozen screen Paula Stafford 252-847-4937 | 4-26-18 Came in for display not responding. I connected a keyboard to it and it will not respond. Also the analyzer will not turn off using power button. Ordered display unit but it is on back order. Came in to see if I could adjust display. Received a call from Mark Ziegler and he has one and is sending it to me. Will return with part. 4-27-18 Went to Fedex to pick up parts. Replaced display unit and it still did not work. I opened the CPU box and found a printer unit instead do CPU unit. Went into service and still could not adjust the display. Checked device manager to see if display driver wasn't flagged and it was not. Spoke with Mark Ziegler and he is sending a CPU unit this time. | 4-26-18 Went in to service and display calibration but display will not respond. Checked all connections. 4-27-18 Tried display unit. I had ordered a CPU unit but received a printer unit. Mark Ziegler is sending the correct part.
No | Visit | Complaint | RLTD | John Trenfield Tel 07900 407208 | Barcode reader not consistently reading. | Checked over barcode reader, no obvious issues, replaced reader, now ok. | SWO-344636 | Defect | HW | DMS_PC | Barcode reader, display |  | John Trenfield Tel 07900 407208 | Barcode reader not consistently reading. | Checked over barcode reader, no obvious issues, replaced reader, now ok.
No | 2 | Y | RPAC | Blank | SP and SC |  | D43E5A036A1F671F | Assisted Replacement | SC | SC | SC |  | Blank | SP and SC
No | Visit | Complaint | RCHN | 打印机故障 | 打印机故障，借903-001一个 | 更换903-001后修复 by jesson | SWO-343660 | [5 empty] | Printer failure | Printer failure, borrow 903-001 one | Fix by jesson after replacing 903-001
No | Visit | Complaint | RCHN | 接口板烧坏 | 接口板烧坏，借902-972一个。 | 更换902-972后修复 by jesson | SWO-343155 | Defect | HW | Basic Unit | PCB, Sensor connector assembly |  | Interface plate burned out | The interface board is burned out, borrowing 902-972 one. | Repair by jesson after replacement 902-972
No | Visit | Complaint | RCHN | 无法读取测试卡芯片 1123 | 无法读取测试卡芯片 1123 ，借903-581一个。 | 更换903-581后修复 by jesson | SWO-344515 | Defect | HW | Basic Unit | PCB, Sensor interface | chip | Failed to read test card chip 1123 | Failed to read test card chip 1123 , borrow edith of 903-581. | Repair by jesson after replacement 903-581
No | Visit | Complaint | RESP | Derivas en la calibración de la Glucosa. | Con frecuencia la Glu y Lac calibran con derivas. Se observa que los contactos de la placa sensor interface tiene los contactos negros. | Comprobar el equipo con varios SC, cambiar placa sensor interface, calibrar y probar con controles. | SWO-344550 | Defect | HW | Basic Unit | PCB, Sensor interface | Glu/Lac | You're derived from glucose calibration. | Glu and Lac often calibrate with drifts. It is observed that the contacts of the interface sensor board have black contacts. | Check the equipment with multiple SCs, change interface sensor, calibrate and test with controls.
Yes | Visit | Complaint | RCHN | 报错5694 | 报错5694 | 更换2张测试卡后修复 | SWO-344511 | [5 empty] | 5694 error | 5694 error | Repair after replacing 2 test cards
No | Visit | Complaint | RIND | Sample problem in ABL800 | Checked inlet gasket and found that there was clot inside that. | Cleaned inlet gasket with hypochloride and and reinstalled.Then given 2 point calibration and checked with samples. | SWO-344998 | Cleaned | HW | Inlet module | Inlet gasket |  | Sample problem in ABL800 | Checked inlet gasket and found that there was clot inside that. | Cleaned inlet gasket with hypochloride and and reinstalled. Then given 2 point calibration and checked with samples.
No | Visit | Complaint | RAME | Network connection error. | Unit will not connect with Aqure | Replaced CPU. Worked with R Turman over phone to check settings and MAC address. Includes maintenance labor. See SWO-344853. | SWO-344851 | Defect | HW | DMS_PC | CPU | Lan | Network connection error. | Unit will not connect with Aqure | Replaced CPU. Worked with R Turman over phone to check settings and MAC address. Includes maintenance labor. See SWO-344853.
No | Visit | Complaint | RCHN | 结果不准 | 结果不准 | 运行质控，质控结果正常。给用户样本预处理。重新做样本正常。 by jesson | SWO-345066 | [5 empty] | The result is not allowed. | The result is not allowed. | Run quality control, quality control results are normal. Pre-process the user sample. Redo the sample normally. by jesson
No | Visit | Complaint | RAME | Sensor cassette would not pass the V/C cycle after it was stalled. | Checked the data logs and verified the problem. The calibrations would not hold in the analyzer. | Replaced the analog board and the sensor connector board. Replaced the spine board. | SWO-344947 | Defect | HW | Basic Unit | PCB, Sensor connector assembly |  | Sensor cassette would not pass the V/C cycle after it was stalled. | Checked the data logs and verified the problem. The calibrations would not hold in the analyzer. | Replaced the analog board and the sensor connector board. Replaced the spine board.
No | Visit | Complaint | RCHN | 响应错误 | 响应错误 | 更换903-581后修复 by jesson | SWO-353036 | Defect | HW | Basic Unit | PCB, Sensor interface | response error | Wrong response | Wrong response | Repair by jesson after replacement 903-581
No | Visit | Complaint | RCHN | 多参数定标失败 | 多参数定标失败，借903-183（被903-581代替）一个 | 更换903-581后修复 by jesson | SWO-344500 | Defect | HW | Basic Unit | PCB, Sensor interface | multi | Multi-parameter calibration failed | Multi-parameter calibration failure, borrow 903-183 (replaced by 903-581) one | Repair by jesson after replacement 903-581
No | Visit | Complaint | RCHN | 无法读取测试卡芯片信息 | 不能读取测试卡芯片信息，借用903-581 | 更换903-581后修复 by jesson | SWO-343840 | Defect | HW | Basic Unit | PCB, Sensor interface | chip | Failed to read test card chip information | Cannot read test card chip information, borrow edith 903-581 | Repair by jesson after replacement 903-581
No | Visit | Complaint | RCHN | 报错1228、1223、1264、1267 | 报错1228、1223、1264、1267 | 更换903-329和636-561后修复 by jesson | SWO-345123 | Defect | HW | Aspiration Module | Needle Wash Unit | 1228 | Errors 1228, 1223, 1264, 1267 | Errors 1228, 1223, 1264, 1267 | Fix by jesson after replacing 903-329 and 636-561
No | Visit | Complaint | RCHN | 定标卡卡盘 | 定标卡卡盘，借944-258（13918）一张 | 更换944-258（13918）重新定标后修复 by jesson | SWO-344557 | Assisted Replacement | Consumables | Test Kit | Cartridge, NT-PROBNP |  | Calibration card disc | Bid Card Plate, borrow 944-258 (13918) | Repair by jesson after replacement 944-258 (13918) re-calibration
Yes | Visit | Complaint | RCHN | 光学单元超出范围，报错1031 | 光学单元超出范围，报错1031，需要借件903-364、944-216（14050） | 更换903-364备件和重新使用944-216（14050）定标后修复 by jesson | SWO-344490 | Defect | HW | Process Module | Optical Unit | 1031 | Optical unit out of range, error 1031 | Optical unit out of range, error 1031, need to borrow 903-364, 944-216 (14050) | Repair by jesson after replacement of 903-364 spare parts and re-use of 944-216 (14050)
Yes | Visit | Complaint | RCHN | 光学单元超出范围，报错1031 | 光学单元超出范围，报错1031，加件944-258（13691）、944-212（13875）、944-214（13800） | 重新使用件944-258（13691）、944-212（13875）、944-214（13800）定标后修复 | SWO-344545 | Defect | HW | Process Module | Optical Unit | 1031 | Optical unit out of range, error 1031 | Optical unit out of range, error 1031, plus 944-258 (13691), 944-212 (13875), 944-214 (13800) | Re-use parts 944-258 (13691), 944-212 (13875), 944-214 (13800) after calibration repair
No | Visit | Complaint | RCHN | BNP传感器损坏，报错1235 | BNP传感器损坏，报错1235，需要借件905-908、944-258（14147） | 更换905-908，重新定标后修复 by jesson | SWO-344472 | Defect | HW | Process Module | Wash unit | 1235 | BNP sensor damaged, error 1235 | BNP sensor damage, error 1235, need to borrow 905-908, 944-258 (14147) | Replace 905-908, fix by jesson after recalibration
No | Visit | Complaint | RCHN | 无法固定样本 | 无法固定样本，借903-584两个。 | 更换903-584备件后修复 by jesson | SWO-344548 | Defect | HW | Aspiration Module | Tube holder |  | Unable to pin the sample | The sample cannot be fixed, borrowing 903-584 two. | Repair by jesson after replacement of 903-584 spare parts
No | Visit | Complaint | RCHN | 试管架损坏 | 试管架损坏，借903-584一个 | 更换903-584备件后修复 by jesson | SWO-344969 | Defect | HW | Aspiration Module | Tube holder |  | Damage to the test tube rack | Damage to the test tube rack, borrow 903-584 one | Repair by jesson after replacement of 903-584 spare parts
No | Visit | Complaint | RLTD | Keith Tel 01603 286533 | Pinch valve functions intermittently. | Pinch valve is clean and appears physically sound, however failing intermittently. Replaced valve, now ok. | SWO-345472 | Defect | HW | Oxi module | pinch valve |  | Keith Tel 01603 286533 | Pinch valve functions intermittently. | Pinch valve is clean and appears physically sound, however failing intermittently. Replaced valve, now ok.
Yes | Visit | Complaint | RAME | Susana states that the users complain that the analyzer wireless connection keeps disconnecting and, patient results cannot be transmitted. | The users state that the analyzer wireless connection keeps disconnecting and, patient results cannot be transmitted. | Checked the wireless reconnect settings and, the USB connection power management settings. Set the wireless reconnection setting to enabled. Set the reconnect retries interval for every five minutes if disconnected. In the device manager, I setup the USB power management to disabled so that the USB connections are not shutdown in order to keep the wireless adapter working at all times. | SWO-345729 |  | SW | Roma.ini | Wireless connection | mail | Susana states that the users complain that the analyzer wireless connection keeps disconnecting and, patient results cannot be transmitted. | The users state that the analyzer wireless connection keeps disconnecting and, patient results cannot be transmitted. | Checked the wireless reconnect settings and, the USB connection power management settings. Set the wireless reconnection setting to enabled. Set the reconnect retries interval for every five minutes if disconnected. In the device manager, I setup the USB power management to disabled so that the USB connections are not shutdown in order to keep the wireless adapter working at all times.
Yes | Visit | Complaint | RAME | Susana states that the users are seeing this analyzer keeps loosing the wireless connectivity. | The customer is noticing that this analyzer keeps loosing the wireless connection and, they cannot transmit test results. | Checked the settings for the USB connections and, confirmed the 'Power management' settings are disabled so that the USB connections are 'not' shut down. I also enabled in Roma.ini the reconnect feature and, the reconnect interval intervals to try reconnect every five minutes if the connection is lost. | SWO-345728 |  | SW | Roma.ini | Wireless connection | mail | Susana states that the users are seeing this analyzer keeps loosing the wireless connectivity. | The customer is noticing that this analyzer keeps loosing the wireless connection and, they cannot transmit test results. | Checked the settings for the USB connections and, confirmed the 'Power management' settings are disabled so that the USB connections are 'not' shut down. I also enabled in Roma.ini the reconnect feature and, the reconnect interval intervals to try reconnect every five minutes if the connection is lost.
No | Visit | Complaint | RAME | To turn on the serial port | The port was closed for serial com | Changed the Roma.ini to open serial port, show out of range values and notify critical values | SWO-345477 | Setup | SW | Roma.ini | serial port |  | To turn on the serial port | The port was closed for serial com | Changed the Roma.ini to open serial port, show out of range values and notify critical values
No | Visit | Complaint | RAME | To open the serial | The port needed to be opened for serial com | Changed the Roma.ini to open serial port, show out of range values and notify critical values | SWO-345478 | Setup | SW | Roma.ini | serial port |  | To open the serial | The port needed to be opened for serial com | Changed the Roma.ini to open serial port, show out of range values and notify critical values
No | Visit | Complaint | RDEU | Erst nach Neustart druckt der Drucker wieder, bereits bei der Installation gab es dieses Problem | möglicherweise der USB-Chip | PC-Modul erneuert Testbetrieb über eine Woche startet | SWO-345406 | Revisit required | HW | DMS_PC | CPU |  | only after reboot the printer prints again, this problem was already in place during installation | possibly the USB chip | PC module renewed Test operation starts over one week
No | Visit | Complaint | RDEU | unplausiebele Messabweichungen bei Hb und Na zwischen ABLs | Gerät überprüft und gereinigt | Quetschventil überprüft und gereinigt, Hemolyzer testweise getauscht, Quetschventilschlauch erneuert, Gerät neu gestartet, Startphase abgewartet, tHb Kal, Kals und QCs durchgeführt | SWO-346326 | Defect | HW | Oxi module | Hemolyzer | tHb | unprudol measurement deviations in Hb and Na between ABLs | Device checked and cleaned | Squeeze valve checked and cleaned, Hemolyzer tested, hot valve hose renewed, device restarted, start phase awaited, tHb Kal, Kals and QCs performed
No | Visit | Complaint | RDEU | 10.04.2018, Frau Schulz, Tel: 04101305322: Datum stimmt nicht | Transport | Analysator getauscht. | SWO-345227 | [5 empty] | 10.04.2018, Mrs Schulz, Tel: 04101305322: Date is not correct | Transport | Analyzer.
No | Visit | Complaint | RCHN | 工作状态中样本中止 | 处理盘偶尔故障 | 检查检测，后续更换一些相关配件 | SWO-347987 | Defect | SW | Analyzer Control Module | Motor control | 5684 | Sample abort in working state | Occasional failure of the processing disk | Check the test and replace some related accessories subsequently
No | Visit | Complaint | RDEU | 30.04.2018; Fr. Werlin; Tel.: Tel.: 06215034079; Einlass tropft | Analysator in Halt Stellung bei Service. Rechter Einlassnippel abgebrochen. Qualitätskontroll Grenzverletzung bei Na+. | Tausch des Einlassnippels und Anschluss des Silikonschlauches. Einlassdichtung durch Anwender bereits erneuert. Nach Neustart ist Analysator dicht. Analysator noch kalibriert. | SWO-345497 | Defect | HW | Inlet module | Nipple kit |  | 30.04.2018; Ms. Werlin; Tel.: Tel.: 06215034079; Inlet drips | Analyzer in stop position at service. Right inlet nipple aborted. Quality control border violation at Na+. | Replacement of the inlet nipple and connection of the silicone hose. Inlet seal has already been renewed by the user. After reboot, analyzer is dense. analyzer still calibrated.
No | Visit | Complaint | RDEU | AQT bricht sporadisch Proben ab und will neu gestartet werden Fehler 3998 Lock_24 | AQT bricht sporadisch Proben ab und will neu gestartet werden Fehler 3998 Lock_24 | AQT geöffnet und ausgesaugt. Sicherungen (F101 - F104) auf Main Board erneuert. Motorgestänge des Punch Out-, Optical Unit- und Wheel Lock Motors gereinigt und neu eingefettet. Safe Positions - ok. Systemüberprüfung und Systemreinigung - ok. Je 1x LQC pro Parameter durchgeführt. | SWO-345776 | Cleaned | HW | Process Module | Optical Unit | 3998 | AQT sporadically cancels samples and wants to be restarted Error 3998 Lock_24 | AQT sporadically cancels samples and wants to be restarted Error 3998 Lock_24 | AQT opened and sucked out. Fuses (F101 - F104) renewed on main board. Motor rods of the Punch Out, Optical Unit and Wheel Lock motors cleaned and regreased. Safe Positions - ok. System check and system cleaning - ok. 1x LQC per parameter.
No | Visit | Complaint | RDEU | Fehler 1279: Gemessene Flüssigkeitsmenge war oberhalb der Warngrenze Gerät verlangt öfter das Reinigungsverfahren | Nach Auswertung des Servicedumps in Kopenhagen, soll Pumpe 1 erneuert werden. | Pumpe 1 erneuert. Systemüberprüfung und Reinigung durchgeführt. Neue CRP-Charge kalibriert. Anwender meldet sich, falls der Fehler wieder auftreten sollte. | SWO-346182 | Defect | HW | Process Module | Pump1 | 1279 | Error 1279: Measured liquid quantity was above the warning limit Device requires the cleaning procedure more often | After evaluation of the service dump in Copenhagen, pump 1 is to be renewed. | Pump 1 renewed. system verification and cleaning. New CRP batch calibrated. User logs in if the error recurs.
No | 2 | Y | RCZE | Blank |  |  | D441FA43CC5F5B42 | Cleaned | HW |  |  |  | Blank
Yes | Visit | Complaint | RAME | The battery charge percentage is showing 0% after the analyzer has been plugged-in the wall outlet the entire morning. | Battery charge percentage is showing as 0% even thought the analyzer has been plugged-in the wall outlet for hours. | Checked the battery charging in service mode. Tried resetting the battery pack but, this didn't help solve the problem. Replaced a new power supply and, after an hour the battery is showing as fully charged at 100%. | SWO-345730 | Defect | HW | Power Supply | Power Supply |  | The battery charge percentage is showing 0% after the analyzer has been plugged-in the wall outlet the entire morning. | Battery charge percentage is showing as 0% even thought the analyzer has been plugged-in the wall outlet for hours. | Checked the battery charging in service mode. Tried resetting the battery pack but, this didn't help solve the problem. Replaced a new power supply and, after an hour the battery is showing as fully charged at 100%.
No | Visit | Complaint | RAME | Thb calibration is drifting | Oximetry is totally out yellow $ I checked the Oximetry system and it all passed | Replace the waste tube and hemolyzer valve $Adjusted thermistor | SWO-346407 | Defect | HW | Oxi module | pinch valve |  | Thb calibration is drifting | Oximetry is totally out yellow / I checked the Oximetry system and it all passed | Replace the waste tube and hemolyzer valve $Adjusted thermistor
No | Visit | Complaint | RAME | Customer could not select QC from the layout menu. | Checked the data logs. Went to the syringe mode and was unable to select the ampoules QC panel. | Reloaded 3.3 MR1 software and that did not work. Tried to reload Windows 7 but it would not work. Downloaded a new copy of windows 7 and installed all the software. | SWO-346009 | Defect | SW | ABL Software | reinstall |  | Customer could not select QC from the layout menu. | Checked the data logs. Went to the syringe mode and was unable to select the ampoules QC panel. | Reloaded 3.3 MR1 software and that did not work. Tried to reload Windows 7 but it would not work. Downloaded a new copy of windows 7 and installed all the software.
No | Visit | Complaint | RCHN | 存储盘卡盘，报错5784、5689、2874 | 存储盘卡盘，报错5784、5689、2874 | 重新清理干燥机和取出卡盘之后修复by tim | SWO-346083 | Defect | Consumables | Test Kit | Cartridge, Broken | 5784 | Memory disk card, error 5784, 5689, 2874 | Memory disk card, error 5784, 5689, 2874 | Repair by tim after recleaning the dryer and removing the chuck
No | Visit | Complaint | RCHN | po2电流质偏低 | po2电流质偏低，需要借件902-833 | 更换902-833后修复by tim | SWO-345230 | Defect | HW | Oxi module | Oxi module |  | po2 current quality is low | Po2 current quality is low, need to borrow 902-833 | Repair by tim after replacement 902-833
No | Visit | Complaint | RHUN | 1559 akku hiba jelezve és 1560 akku hiba továbbra is érzékelve hibák jelezve | 1559 akku hiba jelezve és 1560 akku hiba továbbra is érzékelve hibák jelezve | akkumulátor szerviz teszt hibát jelez, akkumulátor opció megszűntetése, kalibráció, QC mérés | SWO-346276 | Defect | HW | Battery module | Battery module | 1559 | 1559 Battery error indicating and 1560 Battery Fault still detected Errors | 1559 Battery error indicating and 1560 Battery Fault still detected Errors | Service test failure, battery option termination, Calibration, QC measurement
No | Visit | Complaint | RDEU | Anwender hat schon LP, SC u. Einlaßdichtung gew., trotzdem Leckage | Leckage | Fehlersuche; Ventilschlauch u. Dichtung am Einlaß gew.; Startphase abgewartet; fällige tHb-Kal. durchgef.; Kal. + QC i.O. | SWO-346387 | Defect | HW | Oxi module | Pinch valve tube |  | User has already used LP, SC and inlet seal, still leakage | Leakage | Troubleshooting; Valve hose and seal at the inlet. start-up phase awaited; due tHb-Kal. through it; Kal. + QC in the first place
No | Visit | Complaint | RFRA | Erreur code 2 après un arrêt brutal. | Tentative sauvegarde Roma.ini, Roma.db, Roma.log Impossible de sauvegarder ou renommer Roma.db Impossible de réinstaller Windows7, après plusieurs tentatives. Carte CF semble HS. | Revenir avec une carte CF pour installer Windows7. | SWO-346557 | Defect | HW | DMS_PC | CF Card |  | Error code 2 after a sudden stop. | Attempt to save Roma.ini, Roma.db, Roma.log Can't save or rename Roma.db Can't reinstall Windows7 after several attempts. MAP CF looks HS. | Come back with a CF card to install Windows7.
No | 3 | Y | RPOL | Blank | Uszkodzona p³ytka Auto Inlet PCB. Wymina pomog³a.Analizator sprawny |  | D44492CEB862E9CB | Defect | HW | Auto Inlet module | PCB, Auto inlet |  | Blank | Uszkodzona p³ytka Auto Inlet PCB. Wymina pomog³a. Analizator sprawny
No | Visit | Complaint | RLTD | Jo Nash 02392 283232 | Checked data logs for failed patient samples - several samples from different patients and different operators failed with error 0292 turbidity too high and 0581 Oxi spectrum mismatch error. There are no calibration, quality control or other errors in the logs. Dismantled analyser and soaked hemolyser unit sample pathway with hypochlorite solution for 10 minutes to clean. Tested After restart with blood sample - turbidity error still present. Therefore replaced hemolyser unit - tested with 3 x blood samples after restart - all ok. Analyser ready for use. | Replaced faulty hemolyser unit | SWO-346316 | Defect | HW | Oxi module | Hemolyzer | 0292, 0581 | Jo Nash 02392 283232 | Checked data logs for failed patient samples - several samples from different patients and different operators failed with error 0292 turbidity too high and 0581 Oxi spectrum mismatch error. There are no calibration, quality control or other errors in the logs. Dismantled analyser and soaked hemolyser unit sample pathway with hypochlorite solution for 10 minutes to clean. Tested After restart with blood sample - turbidity error still present. Therefore replaced hemolyser unit - tested with 3 x blood samples after restart - all ok. Analyser ready for use. | Replaced faulty hemolyser unit
Yes | Visit | Complaint | RFRA | Après comparaison l'ABL A est plus élevé sur la COHB | ... | Échange du module oxi avec ABLB $ $ relance $ $ mise en route $ $ calibration thb $ $ essai de comparaison avec 4 gaz patient -> le problème a été transféré $ $ réinstallation du module oxi d'origine (avec le spectromètre d'origine) $ $ remise en route $ $ calibration thb | SWO-347781 | Re classified RMED | used for trouble shooting another analyzer |  |  |  | After comparison the ABL A is higher on the COHB | ... | Exchange of the oxi module with ABLB $$raise $$ start $$ calibration thb $$ test comparison with 4 patient gas - the problem was transferred $ $ relocation of the original oxi module (with the original spectrometer) $ restarted $$ calibration $ Thb
No | Visit | Complaint | RFRA | Après comparaison, la Hbco est plus basse sur cet analyseur | ... | Échange du module oxi avec ABLA $ $ relance $ $ mise en route $ $ calibration thb $ $ essai de comparaison avec 4 gaz patient -> le problème a été transféré $ $ réinstallation du module oxi d'origine $ $ remise en route $ $ calibration thb | SWO-347780 | Revisit required | HW | Oxi module | Spectrometer |  | After comparison, the Hbco is lower on this analyzer | ... | Exchange of the oxi module with ABLA $$raise $$ start -$ calibration thb $$ test comparison with 4 patient gas - the problem was transferred $ reinstallation of the original oxi module $$ restarted $ calibration thb
No | Visit | Complaint | RAME | The pCO2 is taking hours to start working everytime a new sensor cassette or solution pack is replaced. | No investigation done. Received report that customer has problems after replacing sensor cassette. | Replaced sensor connector board for troubleshooting purposes. | SWO-346031 | Defect | HW | Basic Unit | PCB, Sensor connector assembly |  | The pCO2 is taking hours to start working everytime a new sensor cassette or solution pack is replaced. | No investigation done. Received report that customer has problems after replacing sensor cassette. | Replaced sensor connector board for troubleshooting purposes.
No | Visit | Complaint | RKK | SC・SP接続不良 多発 | SC・SP接続不良 多発 | 装置引上げの為、代替機設置 架台ぐらつき有りの為、固定 | SWO-347322 | [5 empty] | SC/SP connection failure frequently | SC/SP connection failure frequently | Installation of an alternative machine for equipment pulling up Because there is a stand wobble, it is fixed
No | Visit | Complaint | RIND | Service Visit. | In patient report na and Hct values are outside the reference range and not calculated respectively. | Reset the SP and flush the cassette with the distilled water and again reinitiate the cassette but calibration failed again calibrate the cassette with the active parameter, need to be verified with the sample, sample not available. | SWO-347118 | [5 empty] | Service Visit. | In patient report na and Hct values are outside the reference range and not calculated respectively. | Reset the SP and flush the cassette with the distilled water and again reinitiate the cassette but calibration failed again calibrate the cassette with the active parameter, need to be verified with the sample, sample not available.
No | Visit | Complaint | RIND | Service Visit. | After sample come verification need to be done. | Again the na, ca and Hct is not giving the value, need to change the sensor cassette, Replace the sensor cassette and install it in to the analyzer, verify the user by checking the sample now all the values are coming, analyzer handed over to staff in well working condition. | SWO-347133 | [5 empty] | Service Visit. | After sample come verification need to be done. | Again the na, ca and Hct is not giving the value, need to change the sensor cassette, Replace the sensor cassette and install it in to the analyzer, verify the user by checking the sample now all the values are coming, analyzer handed over to staff in well working condition .
No | Visit | Complaint | RESP | fallo en calibración de sensor | Sustitución de sensor | Sustitución de sensor | SWO-346445 | [5 empty] | sensor calibration failure | Sensor replacement | Sensor replacement
No | Visit | Complaint | RCHN | 软件不匹配，报错0022 | 软件不匹配，报错0022，需要借件903-198和910-304 | 更换903-198和910-304后修复by tim | SWO-345721 | Defect | HW | Storage Module | Barcode reader, storage wheel | 0022 | Software mismatch, error 0022 | Software mismatch, error 0022, need to borrow 903-198 and 910-304 | Repair by tim after replacing 903-198 and 910-304
No | Visit | Complaint | RCHN | 定标卡口断裂 | 定标卡口断裂 | 更换944-214（13886）后修复by tim | SWO-346095 | Defect | Consumables | Test Kit | Cartridge, Myo |  | The calibration card opening is broken | The calibration card opening is broken | Repair by tim after replacement 944-214 (13886)
No | Visit | Complaint | RCHN | pct卡断裂 | pct卡断裂.需要借件944-394（14070） | 重新定标后修复by tim | SWO-346206 | Defect | Consumables | Test Kit | Cartridge, PCT |  | pct card break | pct card breaks . . . 944-394 (14070) | Fix by tim after re-calibration
No | Visit | Complaint | RIND | Received a complaint from Ruby hall Clinic that calibration failure was observed for K, Ca and Na after multiple calibration attempts. | Need to check the calibration values as well as status of membranes. | Found clot in the Ref measuring chamber, cleaned the chamber, cleaned the K, Na and Ca Chamber as well as membrane. Done the Calibration, after successful calibration, run internal QC, run samples, results are OK. Also cleaned the inlet casket. | SWO-347535 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway | clot, ref | Received a complaint from Ruby hall Clinic that calibration failure was observed for K, Ca and Na after multiple calibration attempts. | Need to check the calibration values as well as status of membranes. | Found clot in the Ref measuring chamber, cleaned the chamber, cleaned the K, Na and Ca Chamber as well as membrane. Done the Calibration, after successful calibration, run internal QC, run samples, results are OK. Also cleaned the inlet casket.
No | Visit | Complaint | RIND | Po2 is low for QC | po2 QC is lower side for PO2 | checked the analyser qc values was in lower range for po2 , removed the po2 elctrode puss formation was there cleaned it given calibartion then tested qc all were in range working fine | SWO-347847 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | Po2 is low for QC | po2 QC is lower side for PO2 | checked the analyser qc values was in lower range for po2 , removed the po2 elctrode puss formation was there cleaned it given calibartion then tested qc all were in range working fine
No | Visit | Complaint | RKK | Ｐｃｏ2感度エラー | Ｐｃｏ2感度エラー | Pco2電極交換 CPH：18050115C | SWO-348069 | Assisted Replacement | Electrode | Electrode | pCO2 | YR-01 | Pco2 sensitivity error | Pco2 sensitivity error | Pco2 electrode replacement CPH: 18050115C
Yes | Visit | Complaint | RAME | Susana states that the users complain that the analyzer wireless connection keeps disconnecting and, patient results cannot be transmitted. | Adjusting the wireless reconnect retries and, adjusting the patient database look up location. | Adjusted the reconnect retires to ten times if disconnected. Adjusted the analyzer so that data is only looked up in the hospital database and, not the analyzer database after patient barcode is scanned. | SWO-347904 |  |  |  | Wireless connection | mail | Susana states that the users complain that the analyzer wireless connection keeps disconnecting and, patient results cannot be transmitted. | Adjusting the wireless reconnect retries and, adjusting the patient database look up location. | Adjusted the reconnect retires to ten times if disconnected. Adjusted the analyzer so that data is only looked up in the hospital database and, not the analyzer database after patient barcode is scanned.
Yes | Visit | Complaint | RAME | Susana states that the users are seeing this analyzer keeps loosing the wireless connectivity. | Adjusting the database information look up. | Adjusted the reconnect retires to ten times if disconnected. Adjusted the analyzer so that data is only looked up in the hospital database and, not the analyzer database after patient barcode is scanned. | SWO-347907 |  |  |  | Wireless connection | mail | Susana states that the users are seeing this analyzer keeps loosing the wireless connectivity. | Adjusting the database information look up. | Adjusted the reconnect retires to ten times if disconnected. Adjusted the analyzer so that data is only looked up in the hospital database and, not the analyzer database after patient barcode is scanned.
No | Visit | Complaint | RAME | Recently there have been a number of failed to aspirated errors | There have been a number of failed to aspirated sample errors. Actual numbers are not unusual. 14 total in a 30 day period. | Double checked all flow areas and inlet parts. | SWO-347295 | No fault found | No fault found | No fault found | Sample handling |  | Recently there have been a number of failed to aspirated errors | There have been a number of failed to aspirated sample errors. Actual numbers are not unusual. 14 total in a 30 day period. | Double checked all flow areas and inlet parts.
No | Visit | Complaint | RDEU | 02.05.2018, 11:10 Uhr, Fr. Rosebrook, Tel. 05191800184: Einlass, Einlassdichtung prüfen | Einlassprüfung. Frau Rosebrook bittet um einen Service. | Alle Fliesswege im Servicemenü geprüft. Einlassdichtung gereinigt und Anschlussdichtung richtig angedrückt. Kein Fehler feststellbar. | SWO-346756 | No fault found | No fault found | No fault found | Auto Inlet |  | 02.05.2018, 11:10 a.m., Ms. Rosebrook, tel. 05191800184: inlet, check inlet seal | Entrance check. Ms Rosebrook is asking for a service. | Checked all flow paths in the service menu. Inlet seal cleaned and connection seal properly pressed. No error detectable.
No | Visit | Complaint | RAME | Sample not aspirating | Analyzer in UIR screen. Asking for new inlet gasket. Customer had issues aspirating samples while running gases. | Checked logs, replaced inlet gasket. Replaced connector gasket from customer stock. Calibrations and QC OK. | SWO-346752 | Defect | HW | Auto Inlet module | Inlet gasket |  | Sample not aspirating | Analyzer in UIR screen. Asking for new inlet gasket. Customer had issues aspirating samples while running gases. | Checked logs, replaced inlet gasket. Replaced connector gasket from customer stock. Calibrations and QC OK.
No | Visit | Complaint | RESP | Fallo SP. El equipo no consigue instalar el consumible. | El equipo no inicializa el SP. Da fallo de SP. | Se reviso el equipo internamente, se limpio y se detecto que el SP está defectuoso. Se procedió a instalar un SP nuevo y comprobación del correcto funcionamiento del equipo segun protocolo del fabricante. | SWO-348565 | [5 empty] | SP failure. The machine is unable to install the consumable. | The computer does not initialize the SP. It gives SP failure. | The equipment was checked internally, cleaned and the SP detected to be faulty. A new SP was installed and the equipment was checked for proper operation according to the manufacturer's protocol.
Yes | Visit | Complaint | RCHN | lac定标失败，报错0375和0378 | lac定标失败，报错0375和0378，需要借件903-581 | 更换903-581后修复by tim | SWO-345775 | Defect | HW | Basic Unit | PCB, Sensor interface | Lac | Lac bid failure, error 0375 and 0378 | Lac bid failure, error 0375 and 0378, need to borrow 903-581 | Repair by tim after replacement 903-581
No | Visit | Complaint | RCHN | 报错3264 | 报错3264，需要借件903-175和903-198 | 更换903-175,903-198后修复by tim | SWO-345785 | Defect | HW | Analyzer Control Module | PCB, Main Board | 3264 | Error 3264 | 3264, need to borrow 903-175 and 903-198 | Repair by tim after replacing 903-175, 903-198
No | Visit | Complaint | RCHN | HCT值不能被确定 | HCT值不能被确定，借903-329,636-561各一个 | 更换903-329 636-561后修复by tim | SWO-346211 | Defect | HW | Aspiration Module | Needle Wash Unit |  | HCT value cannot be determined | HCT value cannot be determined, borrow 903-329, 636-561 each | Repair by tim after replacing 903-329 636-561
No | Visit | Complaint | RCHN | 光学单元超范围 | 光学单元超范围，借944-212（13914），944-394（13849），944-258（13956），944-214（13886），944-216（13784）各一个 | 定标还是不通过，加借944-212（13914），944-394（13849）一张，另外还报错1235，加借905-908一个by tim | SWO-346069 | Defect | HW | Process Module | Optical Unit | 1235 | Optical unit over range | Optical unit over range, borrowed 944-212 (13914), 944-394 (13849), 944-258 (13956), 944-214 (13886), 944-216 (13784) each | Bid or not pass, add 944-212 (13914), 944-394 (13849), plus 1235, plus 905-908 one by
Yes | Visit | Complaint | RCHN | 光学单元超范围 | 加借903-364（广办自提） | 定标还是不通过，加借944-212（13914），944-394（13849）一张，另外还报错1235，加借905-908一个by tim | SWO-346070 | Re classified RMED | duplicate upSWO |  |  |  | Optical unit over range | Borrowing 903-364 (widely self-mentioned) | Bid or not pass, add 944-212 (13914), 944-394 (13849), plus 1235, plus 905-908 one by
No | Visit | Complaint | RCHN | 试管夹断裂 | 试管夹断裂 | 自行更换2个903-584后修复by tim | SWO-346789 | Defect | HW | Aspiration Module | Tube holder |  | Test tube clip break | Test tube clip break | Repair by tim after replacing 2 903-584
No | Visit | Complaint | RCHN | 测试卡断裂 2872 | 测试卡断裂 2872 | 更换测试卡正常。 | SWO-349454 | Defect | Consumables | Test Kit | Cartridge, Broken | 2872 | Test Card Break 2872 | Test Card Break 2872 | The replacement test card is OK.
No | Visit | Complaint | RFRA | Pb po2 | Vérification des historiques Relevé des informations | Extraction des données système $ $ passage des contrôles $ $ point avec le client $ $ création techline | SWO-347779 | [5 empty] | Pb po2 | Checking history Reporting | System data extraction $$ passage of controls $$ point with the customer $$ techline creation
No | Visit | Complaint | RIND | Blank Display, nothing is visible on the display. | Checked and found that the CPU unit is faulty. | Changed the CPU from other analyzer, now machine working fine. need to change the CPU unit. | SWO-347448 | Defect | HW | DMS_PC | CPU |  | Blank Display, nothing is visible on the display. | Checked and found that the CPU unit is faulty. | Changed the CPU from other analyzer, now machine working fine. need to change the CPU unit.
No | Visit | Complaint | RCHN | po2质控偏低 | po2质控偏低，需要借件902-446 | 重新调整质控后修复，未使用备件by tim | SWO-345783 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway | mail | po2 quality control is low | po2 quality control is low, need to borrow 902-446 | Re-adjusted after quality control repair, spare parts by tim
No | Visit | Complaint | RDEU | 03.05.2018 08:40Uhr Frau Waldschmitt 023729082223: Wir haben immer Probleme mit dem Gerät, ein Techniker wird hier benötigt. | S. Case | Fehlersuche: Betriebsmittel-Intervalle sind alle auf „niemals“ gesetzt, pO2 Elektrode komplett trocken und Ref. Elektrodenhülse rot (wurden vermutlich seit erst-Installation nicht gewartet), beide sowie pCO2 membraniert, Pumpenschläuche sowie Einlassdichtung und Ventilatorfilter ersetzt. Alle Betriebsmittelintervalle auf Vorgabeeinstellungen gesetzt. Ampel ist bei QC Verletzungen auf grün eingestellt => nach Absprache auf gelb korrigiert. „Korrektive Aktionen“ Einstellungen überprüft. Kalibrierungen, QC Messungen, ok. | SWO-347121 | Setup | SW | Maintenance schedule | Maintenance schedule |  | 03.05.2018 08:40Pm Mrs. Waldschmitt 023729082223: We always have problems with the device, a technician is needed here. | S. Case | Troubleshooting: Equipment intervals are all set to "never", pO2 electrode completely dry and ref. Electrode sleeve red (probably not maintained since first installation), both as well as pCO2 membraneed, pump hoses as well as inlet seal and fan filter replaced. All equipment intervals are set to default settings. Traffic light is set to green for QC injuries => corrected to yellow after consultation. "Corrective Actions" settings checked. Calibrations, QC measurements, ok.
No | Visit | Complaint | RAME | Autocheck door keeps opening. When closing, it grinds and them reopens. | Found light 2 failing on the autocheck status test. Solder joints on the open $close detector board for the autocheck door look like it had cold solder joints. | Replaced 14 pole autocheck cable. Resoldered the joints on the door open $closed detector. Replaced Lac membrane for customer as it was due. Machine also leaking from inlet when I started it up. Found bent inlet probe and nipple under inlet gasket clotted. I replaced inlet probe with customer stock and cleaned connector pieces. Restarted analyzer and performed all necessary calibrations. | SWO-347271 | Defect | HW | Autocheck Module | FPC Cable |  | Autocheck door keeps opening. When closing, it grinds and them reopens. | Found light 2 failing on the autocheck status test. Solder joints on the open /close detector board for the autocheck door look like it had cold solder joints. | Replaced 14 pole autocheck cable. Resoldered the joints on the door open $closed detector. Replaced Lac membrane for customer as it was due. Machine also leaking from inlet when I started it up. Found bent inlet probe and nipple under inlet gasket clotted. I replaced inlet probe with customer stock and cleaned connector pieces. Restarted analyzer and performed all necessary calibrations.
No | Visit | Complaint | RIND | PO2 turns yellow on the parameter bar. | Checked and found that the sensitivity of PO2 out of range due to which the PO2 turns yellow on Parameter Bar | Cleaned the chamber and done the decontamination of PO2 also change the cleaning bottle, now calibration passed and machine working satisfactorily. | SWO-347451 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | PO2 turns yellow on the parameter bar. | Checked and found that the sensitivity of PO2 out of range due to which the PO2 turns yellow on Parameter Bar | Cleaned the chamber and done the decontamination of PO2 also change the cleaning bottle, now calibration passed and machine working satisfactorily.
Yes | Visit | Complaint | RIND | Installation |  | Installed the system in appropriate location and made the setup as per customer requirement. Activated pH, pCO2, pO2, tHb, sO2, Na, K, Ca2+ and Lac options as per agreement. Performed all necessary tests and checks as per the procedure completed all checks successfully and did the Quality Control and Samples all results found within the range. Trained the technicians and Explained complete operational procedures and do's & don’ts for better performance of the system. | SWO-348010 | Re classified RMED | installation report |  |  |  | Installation |  | Installed the system in appropriate location and made the setup as per customer requirement. Activated pH, pCO2, pO2, tHb, sO2, Na, K, Ca2+ and Lac options as per agreement. Performed all necessary tests and checks as per the procedure completed all checks successfully and did the Quality Control and Samples all results found within the range. Trained the technicians and Explained complete operational procedures and do's & don’ts for better performance of the system.
Yes | Visit | Complaint | RAME | No oxi result on patient samples. Oximetry errors. | No oximetry results on patient samples. | Checked the patient data logs and, found errors 581 oxi spectrum mismatch, 688 ctHb $ceHb too low for oxi calculation and, 292 turbidity too high on patient samples results. Performed a Complete oximetry test, found an error with the neon lamp position out of range -10 (Range -0.80 to 0.80). Replaced a new hemolyzer unit and, reran the Complete oximetry test, this time now passing within specifications. Found the inlet probe bent out of shape, replaced a new probe. | SWO-347975 | Defect | HW | Oxi module | Hemolyzer | 292, 581 | No oxi result on patient samples. Oximetry errors. | No oximetry results on patient samples. | Checked the patient data logs and, found errors 581 oxi spectrum mismatch, 688 ctHb $ceHb too low for oxi calculation and, 292 turbidity too high on patient samples results. Performed a Complete oximetry test, found an error with the neon lamp position out of range -10 (Range -0.80 to 0.80). Replaced a new hemolyzer unit and, reran the Complete oximetry test, this time now passing within specifications. Found the inlet probe bent out of shape, replaced a new probe.
No | Visit | Complaint | RAME | Inlet LS failed to calibrate after midnight Calibration. | Inlet LS failed to calibrate after midnight Calibration. Error code 0529. LS value up near high end of range. | Went to FedEx to pick up part. Replace Sensor Cassette Bay Unit. Replaced parts from PM kit. | SWO-347891 | Defect | HW | Basic Unit | Sensor Bay |  | Inlet LS failed to calibrate after midnight Calibration. | Inlet LS failed to calibrate after midnight Calibration. Error code 0529. LS value up near high end of range. | Went to FedEx to pick up part. Replace Sensor Cassette Bay Unit. Replaced parts from PM kit.
No | Visit | Complaint | RAME | Wireless is intermittently dropping off | Customer requested to get looked at that afternoon because, they were using it for training on Friday. Looked in data logs, no faults found. The customer put wireless adapter on back of analyzer. Customer stated since they put it on the back rather then the top, it been working good. | I removed wireless device from back and put it on the top of the analyzer and went into the Roma.ini file and added 3 lines to make the analyzer check the wireless device more often. I also verified power saver mode was deactivated. Then did commit command. | SWO-347804 | Setup | SW | Roma.ini | Wireless connection |  | Wireless is intermittently dropping off | Customer requested to get looked at that afternoon because, they were using it for training on Friday. Looked in data logs, no faults found. The customer put wireless adapter on back of analyzer. Customer stated since they put it on the back rather then the top, it been working good. | I removed wireless device from back and put it on the top of the analyzer and went into the Roma.ini file and added 3 lines to make the analyzer check the wireless device more often. I also verified power saver mode was deactivated. Then did commit command.
No | Visit | Complaint | RAME | Inlet gasket with holder broke off in the inlet module. | Found inlet gasket broken and ABL90 in the intervention screen | Replaced inlet gasket with holder Ran fluid test in service. OK | SWO-347710 | Defect | HW | Inlet module | Inlet gasket |  | Inlet gasket with holder broke off in the inlet module. | Found inlet gasket broken and ABL90 in the intervention screen | Replaced inlet gasket with holder Ran fluid test in service. OK
No | Visit | Complaint | RKOR | Sample clot error | . | Liquid flow line clean. Service mode check. Gasket module clean. SC SP clean. Calibration, QC check. | SWO-348138 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | Sample clot error | . | Liquid flow line clean. Service mode check. Gasket module clean. SC SP clean. Calibration, QC check.
No | Visit | Complaint | RAME | On two occasions tHb varied from lab results on the same patient, using different samples. | Troubleshot to a faulty hemolyzer | Replaced the hemolyzer, and this resolved the problem | SWO-347946 | Defect | HW | Oxi module | Hemolyzer |  | On two occasions tHb varied from lab results on the same patient, using different samples. | Troubleshot to a faulty hemolyzer | Replaced the hemolyzer, and this resolved the problem
Yes | Visit | Complaint | RFRA | Après comparaison l'ABL A est plus élevé sur la COHB | ... | Réinstallation du spectromètre qui s'avérait être mieux au niveau de la COHB $ $ relance $ $ mise en route $ $ calibration thb $ $ essais: bons | SWO-347776 | Re classified RMED | used for trouble shooting another analyzer |  |  |  | After comparison the ABL A is higher on the COHB | ... | Reinstallation of the spectrometer that turned out to be better at the level of the COHB $$raise $$ start $$ calibration thb $$ tests: vouchers
No | Visit | Complaint | RFRA | Après comparaison, la Hbco et la Methb est plus basse sur cet analyseur | ... | Selon la réponse à la techline, remplacement de l'hemoliseur, du spectromètre et de la vanne oxi $ $ relance $ $ mise en route (1H) $ $ calibration thb $ $ essais avec des sang patient -> toujours pareil $ $ comparaison avec les analyseur délocalisés $ $ réinstallation de l'ancien module oxi $ $ remise en route $ $ calibration thb $ $ essais: bons | SWO-347775 | Defect | HW | Oxi module | Spectrometer |  | After comparison, the Hbco and Methb is lower on this analyzer | ... | Depending on the response to the techline, replacement of the hemoliser, spectrometer and valve oxi $$raise $$ start (1H) $$ calibration thb $ $tests with patient blood - always the same $$ comparison with the outsourced analyzer $$ relocation of the old module oxi $$ restarted $$ calibration thb $ $tests: vouchers
No | Visit | Complaint | RAME | The analyzer keeps randomly not printing. | Found that the database was corrupted. | Replaced database and reloaded software which corrected the issue. | SWO-347973 | Defect | SW | SW | reinstall | Printer | The analyzer keeps randomly not printing. | Found that the database was corrupted. | Replaced database and reloaded software which corrected the issue.
No | Visit | Complaint | RKOR | Clot issue | Clot 감지 | Flush로 Clot 제거 | SWO-349459 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway | clot | Clot issue | Clot detection | Remove Clot with Flush
Yes | Visit | Complaint | RCHN | 打印机无法打印 | 打印机无法打印，需要借件902-985 | 更换备件902-985后修复by tim | SWO-346190 | [5 empty] | The printer can't print | The printer cannot print and needs a loan 902-985 | Repair by tim after replacement of spare parts 902-985
No | Visit | Complaint | RCHN | LAC定标不过 | LAC，GLU定标不过 借903-581一个 | 更换903-581后修复by tim | SWO-347327 | Defect | HW | Basic Unit | PCB, Sensor interface | Glu/Lac | LAC calibration is not | LAC, GLU calibration but borrow 903-581 one | Repair by tim after replacement 903-581
No | Visit | Complaint | RCHN | 报错1031 | 报错1031 借903-364 一个 944-258（14051），944-394（13849 14070）944-220（13764，13785） | 更换903-364重新定标后修复by tim | SWO-346791 | Defect | HW | Process Module | Optical Unit | 1031 | 1031 | 1031 Lend 903-364 A 944-258 (14051), 944-394 (13849 14070) 944-220 (13764,13785) | Repair by tim after replacement 903-364 recalibration
No | Visit | Complaint | RIND | Service Visit. | Main Screen shows the SP memory chip is corrupt. | Remove the solution pack and sanitize the memory chip with gauss piece and also sanitize the inner memory chip connector, now reset the sp and again try to install the sp, not succeed, take all the data log from the analyzer, need to replace the SP. | SWO-347554 | [5 empty] | Service Visit. | Main Screen shows the SP memory chip is corrupt. | Remove the solution pack and sanitize the memory chip with gauss piece and also sanitize the inner memory chip connector, now reset the sp and again try to install the sp, not succeed, take all the data log from the analyzer, need to replace the SP.
No | Visit | Complaint | RSAF | Unit manager called to inform us regarding the TCM being non usable as it has a monitor error 08000200. | Sensor calibrations passes with no monitor error on screen, once sensor removed from holding latch to be applied for monitoring the monitor error message returns making the unit unusable for patient monitoring. | Unit brought into RSAF workshop for further troubleshooting, Tech line logged. | SWO-349880 | [5 empty] | Unit manager called to inform us regarding the TCM being non usable as it has a monitor error 08000200. | Sensor calibrations passes with no monitor error on screen, once sensor removed from holding latch to be applied for monitoring the monitor error message returns making the unit unusable for patient monitoring. | Unit brought into RSAF workshop for further troubleshooting, Tech line logged.
No | Visit | Complaint | RAME | Glucose unstable results excessive drift | Glucose has Excessive drift over the last two weeks, Ran a trend and it goes from -14 to 20, On QC Glucose show unstable results; Also found PCO2 failing level 4 and running on the low end for all levels; | Replaced Glu electrode under warranty, membranes and performed a zero current calibration; Replaced the PCO2 electrode; Glu is unstable for QC will let the analyzer sit over night and order electrode connecot; believe there is a bad measuring chamber in the pH module and bad electrode connector on the glu, Came on site 5.3.2018 glu still has a drift and pCO2 is better but on the low end on all QC levels replaced measuring chamber for met module electrode connector on Glu chamber; performed annual maintenance will document serparately for tracking purposes--- Glu is no longer drifting but pCO2 is failing cals and low on QC, ordering gas chassis and regulators for tomorrow; Replaced gas chassis and both regulators replaced the measuring chamber and put the original pCO2 electrode back in; | SWO-346611 | Defect | HW | gas module | Flow regulator |  | Glucose unstable results excessive drift | Glucose has Excessive drift over the last two weeks, Ran a trend and it goes from -14 to 20, On QC Glucose show unstable results; Also found PCO2 failing level 4 and running on the low end for all levels; | Replaced Glu electrode under warranty, membranes and performed a zero current calibration; Replaced the PCO2 electrode; Glu is unstable for QC will let the analyzer sit over night and order electrode connecot; believe there is a bad measuring chamber in the pH module and bad electrode connector on the glu, Came on site 5.3.2018 glu still has a drift and pCO2 is better but on the low end on all QC levels replaced measuring chamber for met module electrode connector on Glu chambe r; performed annual maintenance will document serparately for tracking purposes--- Glu is no longer drifting but pCO2 is failing cals and low on QC, ordering gas chassis and regulators for tomorrow; Replaced gas chassis and both regulators replaced the measuring chamber and put the original pCO2 electrode back in;
No | Visit | Complaint | RDEU | 04.05.2018; Hr. Henneboele; Tel.: 02093643880; Na und Ca Kal Fehler 04.05.2018; 11:40 Uhr; Hr. Henneboele; Tel.:02093644504; weiterhin Na und Ca Driftfehler | E'lytestatus nicht o.k., Na Elektrode ist def. muss noch ausgetauscht werden, | Messkammer gereinigt, Elektroden neu bezogen, Gerät kalibriert, | SWO-347646 | Revisit required | Electrode | Electrode | Na |  | 04.05.2018; Mr. Henneboele; Tel.: 02093643880; Na and Ca Kal Error 04.05.2018; 11:40 a.m.; Mr. Henneboele; Tel.:0 2093644504; continue Na and Ca drift errors | E'lytestatus not OK, Na electrode is def. has yet to be replaced, | measuring chamber cleaned, electrodes re-covered, device calibrated,
No | Visit | Complaint | RIND | Oxy liquid sensor error. | Oxy liquid sensor error. | Protean removal done. Analyser Drained and switched off for half hour. Inlet gasket cleaned and tubing replaced. Restart and calibrated. Analyser Serviced and calibrated properly. Handed over in a fine working condition. | SWO-347828 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | Oxy liquid sensor error. | Oxy liquid sensor error. | Protean removal done. Analyser Drained and switched off for half hour. Inlet gasket cleaned and tubing replaced. Restart and calibrated. Analyser Serviced and calibrated properly. Handed over in a fine working condition.
No | Visit | Complaint | RLTD | Scanner error on board | Scanner error | Unit still not picking up cartridges. New scanner unit ordered. | SWO-347599 | Revisit required | HW | Storage Module | Barcode reader, storage wheel |  | Scanner error on board | Scanner error | Unit still not picking up cartridges. New scanner unit ordered.
Yes | Visit | Complaint | RDEU | 27.07.2017; Sr. Anett; Tel.: 03381791614; CRP QC Messfehler 28.07.2017 10:30 Sr. Nicole; Tel.: 03381791612; weiterhin CRP QC Messfehler 31.07.2017: Rücksprache mit KerZi - neues QC wird bestellt, telefonische Info an Kunden 03.08.2017, 11:36 Uhr ==> Anrufer Pfl. Sebastian, Tel.: 03381791611 ==> LQC-Messfehler CRP 04.08.2017, 12:04 Uhr ==> Anrufer Schw. Anett, Tel.: 03381791614 ==> LQC-Fehler CRP besteht trotz neuer Charge weiterhin | Rücktausch Gerät | AQT90Flex zurückgetauscht, dazu Original Konfiguration und Netzwerkdaten eingespielt bzw. eingetragen, QC CRP i.O. | SWO-347764 | [5 empty] | 27.07.2017; Sr. Anett; Tel.: 03381791614; CRP QC Measurement Error 07/28/2017 10:30 Sr. Nicole; Tel.: 03381791612; further CRP QC measurement error 31.07.2017: Consultation with KerZi - new QC is ordered, telephone information to customers 03.08.2017, 11:36 a.m. ==> Caller Pfl. Sebastian, Tel.: 0338179111 ==> LQC measurement error CRP 04.08.2017, 12:04 o'clock ==> Caller Schw. Anett, tel.: 03381791614 ==> LQC error CRP persists despite new batch | Replacement device | AQT90Flex exchanged back, with original configuration and network data recorded or entered, QC CRP i.O.
No | Visit | Complaint | RLTD | Frank Watson on 01983 534801 | Found that test cartridges were catching on the underside of the storage wheel cover. Adjusted height of storage wheel cover to stop test cartridges catching. Tested operation in service and by running 2 blank sample tests - all passed, no errors occurred. | Adjusted position of storage wheel cover. | SWO-347682 | Adjustment | HW | Storage Module | storage wheel cover | 24 | Frank Watson on 01983 534801 | Found that test cartridges were catching on the underside of the storage wheel cover. Adjusted height of storage wheel cover to stop test cartridges catching. Tested operation in service and by running 2 blank sample tests - all passed, no errors occurred. | Adjusted position of storage wheel cover.
No | Visit | Complaint | RIND | SET_SAFE22 error come. | check the system message & found SET SAFE_22 error. | Run the Safe position & found Storage wheel failed & Cartridge failed,Clean the storage wheel gear,clean the opto sensor in the gear wheel,clean the cartridge barcode reader,restart the instrument & it became ready. | SWO-347446 | Cleaned | HW | Storage Module | storage wheel, hub |  | SET_SAFE22 error come. | check the system message & found SET SAFE_22 error. | Run the Safe position & found Storage wheel failed & Cartridge failed,Clean the storage wheel gear,clean the opto sensor in the gear wheel,clean the cartridge barcode reader,restart the instrument & it became ready.
No | Visit | Complaint | RSCH | Rep | Sporadisch Fehler 1278 | Keine Aktion ausgeführt Servicedump rausgezogen für RMED | SWO-347664 | Defect | HW | Inlet module | Barcode reader, Inlet |  | Rep | Sporadic Error 1278 | No action taken Service dump pulled out for RMED
No | Visit | Complaint | RSCH | Rep | Verlangt Reinigung wie am 22.03.2018 SWO-326467 | Pumpenköpfe untereinander ausgetauscht. Kein Fehler feststellbar Pumpe 2ausgetauscht Restliche Pumpen untereinander ausgetauscht. Techline eröffnen und auf Antwort RMED warten Funktionskontrolle aktuell i.o | SWO-347633 | Defect | HW | Inlet module | Barcode reader, Inlet |  | Rep | Requires cleaning as on 22.03.2018 SWO-326467 | pump heads. No error detectable Pump 2exchanged Remaining pumps exchanged with each other. Techline open and wait for response RMED Functional control currently i.o
No | Visit | Complaint | RAME | Battery disconnected error. | Inspection of the battery revealed one of the cables has come out of the connector. Customer getting general error on Sensor ser# 8466159. Worked with Mike P. Over phone unable to locate Sensor warranty. Went into Brea office with Carlos Miranda & Lauretta Martin. Turns out warranty wasn't updated in our system yet but Lauretta resolved. | Replaced power supply that includes battery cable. Requesting customer service to send out a replacement TC84 Sensor. | SWO-347268 | [5 empty] | Battery disconnected error. | Inspection of the battery revealed one of the cables has come out of the connector. Customer getting general error on Sensor ser# 8466159. Worked with Mike P. Over phone unable to locate Sensor warranty. Went into Brea office with Carlos Miranda & Lauretta Martin. Turns out warranty wasn't updated in our system yet but Lauretta resolved. | Replaced power supply that includes battery cable. Requesting customer service to send out a replacement TC84 Sensor.
No | Visit | Complaint | RCHN | 断电后导致无法开机 | 断电后导致无法开机，借902-970一个 | 更换902-970后修复by tim | SWO-346824 | [5 empty] | Failure to power on after power failure | Failed to power up after power failure, borrow 902-970 one | Repair by tim after replacement 902-970
No | Visit | Complaint | RCHN | LAC,GLU定标不过 | LAC,GLU定标不过，借903-581一个 | 更换903-581后修复by tim | SWO-346704 | Defect | HW | Basic Unit | PCB, Sensor interface | Glu/Lac | LAC, GLU calibration but | LAC, GLU calibration, borrow 903-581 a | Repair by tim after replacement 903-581
No | Visit | Complaint | RCHN | 漏电流 0963 | 漏电流 0963，借902-980，903-199，903-581各一个。 | 更换903-183后修复by tim | SWO-346811 | Defect | HW | Basic Unit | PCB, Sensor interface | & hemolyzer, leak current | Drain current 0963 | Leak current 0963, borrow 902-980, 903-199, 903-581 each. | Repair by tim after replacement 903-183
No | Visit | Complaint | RCHN | 定标终止 | 5684，借14107（944-214）一个 | 重新定标 | SWO-348020 | Defect | SW | Analyzer Control Module | Motor control | 5684 | Bid termination | 5684, borrow 14107 (944-214) one | Re-calibration
No | Visit | Complaint | RCHN | 光学单元超范围 | 定标还是不通过，加借944-212（13914），944-394（13849）一张，另外还报错1235，加借905-908一个 | 更换905-908后修复(By Ran) Temp sensor was broken. | SWO-346851 | Defect | HW | Process Module | Sensor, BnP | 1235 | Optical unit over range | Bid or not pass, add 944-212 (13914), 944-394 (13849) one, plus 1235, plus 905-908 one | 更换905-908后修复(By Ran) Temp sensor was broken.
No | Visit | Complaint | RIND | Service Visit. | Need to replace the SP9. | Replace the SP9 and automatic calibration take by the analyzer but air detection calibration failed again and again, analyze the sensor cassette, no solution inserted during the calibration from the solution pack in to the sensor cassette. Open the back panel of the analyzer and saw that roller pump tube not connected. Reconnect it and close the analyzer, Now again give the manual calibration, successfully done. Analyzer handed over to staff in well working condition. | SWO-349507 | [5 empty] | Service Visit. | Need to replace the SP9. | Replace the SP9 and automatic calibration take by the analyzer but air detection calibration failed again and again, analyze the sensor cassette, no solution inserted during the calibration from the solution pack in to the sensor cassette. Open the back panel of the analyzer and saw that roller pump tube not connected. Reconnect it and close the analyzer, Now again give the manual calibration, successfully done. Analyzer handed over to staff in well working condition.
No | Visit | Complaint | RCHN | PH定标不过 | PH定标不过，借945-614一个 |  | SWO-348005 | Assisted Replacement | Electrode | Electrode | pH | 0 | PH calibration is not | PH calibration is not, borrow 945-614 a
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | PH,CO2定标不过 | 更换测试卡后修复 | SWO-354957 | [5 empty] | Test card, reagent package failure | PH, CO2 calibration but | Repair after replacing test card
No | Visit | Complaint | RCHN | 机器死机 | 机器死机，需要借件910-343、836-166、944-216（14050）、944-212（14205） | 更换910-343重新定标后修复by tim | SWO-347483 | Revisit required | HW | Analyzer Control Module | CPU |  | Machine crash | Machine crashes, need to borrow 910-343, 836-166, 944-216 (14050), 944-212 (14205) | Repair by tim after replacement 910-343 recalibration
No | Visit | Complaint | RCHN | 定标卡无法扫描 | 定标卡无法扫描，借902-924,903-584 2个 | 更换902-924,903-584 $2个后修复by tim | SWO-347465 | Defect | HW | Cover and Chassis | Barcode reader, display |  | Bid card cannot be scanned | The bid card cannot be scanned, borrowing 902-924, 903-584 2 | Replace 902-924,903-584 $2 after repair by tim
No | Visit | Complaint | RDEU | permanente tHb Fehler auch nach neu Kalibrierungen | Permanent tHb Fehler trotz neukalibration | Fehleranalyse, hemolyzer sowie Pinch Valve erneuert, Maintainance parts erneuert, Kalibrierung und Qualitätskontrolle i. O. | SWO-348366 | Defect | HW | Oxi module | Hemolyzer | tHb | permanente tHb Fehler auch nach neu Kalibrierungen | Permanent tHb Fehler trotz neukalibration | Fault analysis, hemolyzer and pinch valve renewed, Maintainance parts renewed, calibration and quality control in the first place
No | Visit | Complaint | RAME | Thb calibration is drifting | Intermittent Thb drifting on calibration | Replace the Spectrometer will open techline case if this doesn’t fix the problem | SWO-348773 | Defect | HW | Oxi module | Spectrometer | tHb | Thb calibration is drifting | Intermittent Thb drifting on calibration | Replace the Spectrometer will open techline case if this doesn’t fix the problem
No | Visit | Complaint | RDEU | 07.05.2018; Fr, Mengesdorf; Tel.: 062150335589; tHb Kal Fehler | OXI Spectrum Missmatch Oximetrie-Problem (QC und tHb KAL) | Oximetrie Test durchgeführt ( Cravitation absorb-Fehler) - Analysator gespült - Erneuerung Schlauch an Quetschventil - Erneuerung Hemolysator notwendig - Rückmontage und Funktionsüberprüfung etc. ohne Beanstandung (Erneuerung SP, da QC Verletzung pH) ! | SWO-348468 | Defect | HW | Oxi module | Hemolyzer | 581 | 07.05.2018; Fr, Mengesdorf; Tel.: 062150335589; tHb Kal Fehler | OXI Spectrum Missmatch Oximetrie-Problem (QC und tHb KAL) | Oximetry test performed (cravitation absorb error) - analyzer flushed - renewal hose to squeegee valve - renewal hemolyser necessary - reassembly and function verification etc. without complaint (renewal SP, since QC violation pH) !
Yes | 8 | Y | RPAC | Blank | Failed SC |  | D44E744F68402350 | Defect | HW | Basic Unit | Power Supply |  | Blank | Failed SC
Yes | 4 | Y | RPAC | Blank | Failed SC |  | D44E75A068072354 | Defect | HW | Basic Unit | Power Supply |  | Blank | Failed SC
Yes | 6 | Y | RPAC | Blank | Failed SC |  | D44E7955685C2358 | Defect | HW | Basic Unit | Power Supply |  | Blank | Failed SC
No | Visit | Complaint | RFRA | Erreur code 2 après un arrêt brutal. | En accord avec le laboratoire, récupération ABL pour changer carte CF, installation (Windows 7 et logiciel ABL), réinstallation lundi matin. | Chgt carte CF, installation Win7, logiciel ABL. Configuration Windows. | SWO-348149 | Defect | HW | DMS_PC | CF Card | error code 2 | Error code 2 after a sudden stop. | In agreement with the laboratory, ABL recovery to change CF card, installation (Windows 7 and ABL software), resettlement on Monday morning. | Chgt CF card, Win7 installation, ABL software. Windows setup.
No | Visit | Complaint | RDEU | laute Geräusche nach SC Wechsel während Kal und auch Messung | laute Geräusche nach SC Wechsel während Kal und auch Messung | SC-Fach überprüft, i.O., Fluid Programme im Service durchgeführt, LP entnommen und Fenster im oberen Teil gereinigt, LP montiert, Kal, QC | SWO-349285 | Assisted Replacement | SP | SP | Pump tubes | 0 | loud noises after SC change during Kal and also measurement | loud noises after SC change during Kal and also measurement | SC tray checked, i.O., fluid programs performed in service, LP removed and window cleaned in the upper part, LP mounted, Kal, QC
No | Visit | Complaint | RDEU | Probleme beim Wechsel der Einlassanschlussdichtung | Probleme beim Wechsel der Einlassanschlussdichtung, bzw. dem anschließenden Wiedereinsetzen der Einlasseinheit. Seit dem 1.5. 22:34 Uhr wiederkehrend Fehler 1401 Fehler: Einlassdichtungshalter . In dem Falle führt dies dann bei Probenmessungen zu Fehler 1190 Einlass geöffnet . Der Kunde spricht außerdem extrem lange und häufige Kalibrierzeiten bei Glucose und Lactat nach Betriebsmittelwechsel an. | Überprüfung und Kalibration der Einlasseinheit im Servicemenü. Anschließend ordnungsgemäße Funktion. Der Kunde wird im Betrieb beobachten, ob die Fehler 1190 sowie 1401 wieder auftreten und sich melden. Überprüfung des Gerätes mittels Fluid-Transport-Programmen: Keine Fehler festgestellt. Kalibrierzeiten Glucose & Lactat: Am 24.04. wurde zuletzt die Sensorkassette gewechselt: Parameter nach unter einer Stunde grün (lt. Handbuch kann dies bis 4 Std. dauern). Weitere Kalbrierprobleme bei Lactat in der Nacht 20. $21. 4. beruhen auf einem leergelaufenen Lösungspack. Hier war Lactat ohne Sensorkassettenwechsel nach ca. 3 Std. wieder messbereit. | SWO-348349 | Revisit required | HW | Auto Inlet module | Auto Inlet module | 0 | Problems changing the inlet connection seal | Problems when changing the inlet connection seal or the subsequent re-insertion of the inlet unit. Since May 1, 10:34 p.m. Recurring error 1401 Error: Inlet seal holder . In this case, this then leads to error 1190 inlet opened during sample measurements. The customer also addresses extremely long and frequent calibration times for glucose and lactate after the change of equipment. | Checking and calibration of the inlet unit in the service menu. Then proper function. The customer will monitor in operation whether errors 1190 and 1401 reoccur and report. Checking the device using fluid transport programs: No errors detected. Calibration times Glucose & Lactate: On April 24th the sensor cassette was last changed: parameters after less than one hour green (according to manual this can take up to 4 hours). Further calcifering problems with lactate on the night of 20. 21. 4. are based on an empty solution pack. Here lactate was again ready for measurement after approx. 3 hrs without sensor cassette change.
No | Visit | Complaint | RDEU | 04.05.2018, Frau Tagscherer, Tel: 023195057224: Mess-differenzen zu hauseigenen BZ Geräten bei Patientenmessungen von 40-50mg% | Gerät überprüft und gereinigt, Gerät im Serviceprogramm überprüft, | Sensorkassette erneuert, Kaltstartphase abgewartet, Kals und QCs durchgeführt | SWO-348255 | Assisted Replacement | SC | SC | unknown | 0 | 04.05.2018, Mrs Tagscherer, Tel: 023195057224: Measurement differences to in-house BZ devices for patient measurements of 40-50mg% | device checked and cleaned, device checked in the service program, | Sensor cassette renewed, cold start phase awaited, kals and QCs performed
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | PH报N $E，PCO2报C $E，CL报N $E，NA,K,CA偏高 | 更换测试卡正常。 | SWO-348821 | [5 empty] | Test card, reagent package failure | PH报N /E, PCO2报C /E, CL报N /E, NA, K, CA偏高 | The replacement test card is OK.
No | Visit | Complaint | RIND | Calibration failure for all parameters except pO2 after multiple calibration attempts at Surya Hospital, Pune | Need to check the SC and SP. | Tried reinstalling both SC and SP, tried running heparin on the SC, but same prob is observed. So need to replace the SC. | SWO-351047 | [5 empty] | Calibration failure for all parameters except pO2 after multiple calibration attempts at Surya Hospital, Pune | Need to check the SC and SP. | Tried reinstalling both SC and SP, tried running heparin on the SC, but same prob is observed. So need to replace the SC.
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | pco2值高 | 更换测试卡后修复 | SWO-355636 | Assisted Replacement | SC | SC | pCO2 | 0 | Test card, reagent package failure | High pco2 value | Repair after replacing test card
No | Visit | Complaint | RESP | ERRORES DE CAL EN LAC. | ERRORES DE CALIBRACION DE LACTATO | TRAS PRUEBAS Y CAMBIO DE SENSORES SE DETERMINA SUSTITUCION DE LA PCB SENSORT INTERFACE. SE EFECTUAN PRUEBAS DE FUNCIONAMIENTO SIENDO ESTAS SATISFACTORIAS. | SWO-348551 | Defect | HW | Basic Unit | PCB, Sensor interface | Lac | CAL ERRORS IN LAC. | LACTATE CALIBRATION ERRORS | AFTER TESTS AND CHANGE OF SENSORS, REPLACEMENT OF THE SENSORT INTERFACE PCB IS DETERMINED. OPERATING TESTS ARE EFFECTED BY BEING THESE SATISFACTORY.
No | Visit | Complaint | RCHN | 机器死机 | 重启死机，借636-571一个。 |  | SWO-348118 | Defect | HW | Analyzer Control Module | CPU |  | Machine crash | Restart the crash, borrow 636-571 one.
No | Visit | Complaint | RCHN | 报错1999 | 报错1999，借836-166一个 | 更换836-166一个正常。(By Ran) | SWO-348019 | Defect | HW | Process Module | rubber disk | 1999 | Error 1999 | Misquoted 1999, loan 836-166 one | Replace 836-166 one normal. (By Ran)
No | Visit | Complaint | RCHN | 报错1068 | 报错1068，借903-210一个 | 更换903-210后修复(By Ran) | SWO-348044 | Defect | HW | Process Module | PCB, Cup dryer | 1068 | 1068 error | 1068, borrow edith 903-210 | Repair after replacement 903-210 (By Ran)
No | Visit | Complaint | RAME | Customer states unit rebooted by itself and now has a status 10 error. | Verified customer's stated complaint. Unit rebooted due to the batteries in the UPS it's connected to being depleted, causing the analyzer to reboot. Analyzer will not pass the startup screen. | Reinstalled Windows 7 and analyzers software version 6.18 and transfered the user settings from their other working analyzer. Customer is aware that all data will be lost. Recommended that customer contacts biomed to replace UPS batteries. | SWO-348739 | Defect | User error | UPS | UPS | 0 | Customer states unit rebooted by itself and now has a status 10 error. | Verified customer's stated complaint. Unit rebooted due to the batteries in the UPS it's connected to being depleted, causing the analyzer to reboot. Analyzer will not pass the startup screen. | Reinstalled Windows 7 and analyzers software version 6.18 and transfered the user settings from their other working analyzer. Customer is aware that all data will be lost. Recommended that customer contacts biomed to replace UPS batteries.
No | Visit | Complaint | RDEU | 07.05.2018; Fr. Mengesdorf - Tel.: 062150335589; Na und Ca Drift Fehler | KAL Driften Elektrolyte | Allgemeine Überprüfung - Erneuerung Referenz Elektrode - Funktionsüberprüfung etc. Bemerkung: Referenz Membrane LOT 2252 | SWO-348470 | Assisted Replacement | Electrode | Membrane | Reference | 0 | 07.05.2018; Ms. Mengesdorf - Tel.: 062150335589; Na and Ca Drift Error | KAL Drifting Electrolytes | General verification - Renewal reference electrode - function check etc. Note: Reference Membrane LOT 2252
No | Visit | Complaint | RDEU | KAL Driften Elektrolyte | KAL Driften Elektrolyte (sehr häufig $ hauptsächlich Na und Ca) | Allgemeine Überprüfung - Erneuerung Referenz Elektrode - Funktionsüberprüfung etc. | SWO-348469 | Assisted Replacement | Electrode | Electrode | Reference | 0 | KAL Drifting Electrolytes | KAL Drifting Electrolytes (very common - mainly Na and Ca) | General verification - Renewal reference electrode - function check etc.
No | Visit | Complaint | RAME | ctHB low | ctHB low | Adjusted lamp assembly. Performed full system check since this is a beta site for the new abl 800 flex. | SWO-349421 | Adjustment | HW | Oxi module | Lamp unit | ctHB | ctHB low | ctHB low | Adjusted lamp assembly. Performed full system check since this is a beta site for the new abl 800 flex.
No | Visit | Complaint | RAME | ctHB low | ctHB low | Recalibrated oxi module assembly. Also performed full system check since this is a new beta site with abl 800 flex units. | SWO-349422 | Adjustment | HW | Oxi module | Lamp unit | 0 | ctHB low | ctHB low | Recalibrated oxi module assembly. Also performed full system check since this is a new beta site with abl 800 flex units.
No | Visit | Complaint | RDEU | Erst nach Neustart druckt der Drucker wieder, bereits bei der Installation gab es dieses Problem | Fehler tritt nur sporadisch auf das PC-Modul hat keine Verbesserung gebracht, da auch Drucker u Softwareinstallation in Schwerin keine Verbesserung brachten, habe ich das BIOS betrachtet | USB Geschwindigkeit von highspeed auf full $low gesetzt | SWO-349516 | Defect | HW | DMS_PC | CPU | 0 | only after reboot the printer prints again, this problem was already in place during installation | Error occurs only sporadically the PC module did not bring any improvement, since also printer u software installation in Schwerin did not bring any improvement, I looked at the BIOS | USB speed set from highspeed to full $low
No | Visit | Complaint | RAME | Printer randomly not printing | Found that power save was turned on in windows | Turned off power save on USB ports in windows. | SWO-349359 | Setup | SW | Windows | Power save | 0 | Printer randomly not printing | Found that power save was turned on in windows | Turned off power save on USB ports in windows.
No | Visit | Complaint | RDEU | Fehlerbehebung erforderlich | Windows 7 und WLAN-Adapter installieren | Installation mehrmals durchgeführt ==> kein Erfolg, WLAN-Verbindung wird trotz guten Signals nicht aufgebaut Fehlersuche brachte ebenfalls keinen Erfolg | SWO-355546 | Revisit required | SW | His/Lis | Wireless connection | 0 | Troubleshooting required | Install Windows 7 and Wi-Fi Adapter | Installation performed several times ==> no success, Wi-Fi connection is not established despite good signal Troubleshooting also did not succeed
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | NA,K质控偏高，CL质控偏低 | 更换测试卡 | SWO-349457 | [5 empty] | Test card, reagent package failure | NA, K quality control is high, CL quality control is low | Replace the test card
No | Visit | Complaint | RIND | Calibration failure for all parameters except pO2 after multiple calibration attempts at Surya Hospital, Pune | SC replacement. | As requested by the technicians of the Lab Department, went to change the SC and given training to the new staff on SC and SP replacement. Also run 2 samples, machine is working fine. | SWO-351048 | [5 empty] | Calibration failure for all parameters except pO2 after multiple calibration attempts at Surya Hospital, Pune | SC replacement. | As requested by the technicians of the Lab Department, went to change the SC and given training to the new staff on SC and SP replacement. Also run 2 samples, machine is working fine.
No | Visit | Complaint | RCHN | 客户反映报告中的碱剩余机器显示和报告不符 | 客户反映报告中的碱剩余机器显示和报告不符，报告打印浪费纸， | 检查设置发现是LIS问题，客户联系LIS修改，将客户的ABL90自动打印取消，给删除打印排队序列 | SWO-350839 | No fault found | SW | His/Lis | His/Lis | 0 | Customer reflects in report ingale residual machine display and report ing-match | The customer reflects that the alkali residual machine display and report does not match in the report, and reports print waste paper, | Check the settings found to be LIS issues, customers contact LIS modification, the customer's ABL90 automatic printing cancellation, to remove the print queuing sequence
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 不均匀的空气 1544 | 更换试剂包正常。 | SWO-349508 | Assisted Replacement | SP | SP | SP | 1544 | Test card, reagent package failure | Uneven air 1544 | The replacement kit is OK.
No | Visit | Complaint | RESP | Temperatura de sustrato fuera de rango | Error de temperatura de sustrato. Requiere sustitución de la sensor interface PCB | Se sustituyo la sensor interface board según protocolo del fabricante. Finalmente se comprobó correcto funcionamiento del equipo mediante QC. | SWO-349064 | Defect | HW | Basic Unit | PCB, Sensor interface | temp | Substrate temperature out of range | Substrate temperature error. Requires replacement of PCB interface sensor | The interface board sensor is replaced according to the manufacturer's protocol. The equipment was finally found to be functioning correctly using QC.
Yes | Visit | Complaint | RAME | Network connection error. | Unit will not connect to Aqure. | Tried resetting metrics. No help. Performed complete software installation including synapse 16 and llht registry. | SWO-349296 | Defect | HW | Basic Unit | Power Supply |  | Network connection error. | Unit will not connect to Aqure. | Tried resetting metrics. No help. Performed complete software installation including synapse 16 and llht registry.
No | Visit | Complaint | RDEU | Autocheckmodul defekt |  |  | SWO-349370 | Defect | HW | Auto Check Module | Auto Check Module | 0 | Autocheck module defective
No | Visit | Complaint | RSCH | Rep. Einlass | Tropft aus Einlass während Spülung | Einlass-Nippel überprüft i.o Ablass-Schläuche überprüft i.o Ablass-Pumpenschlauch i.o Einlassdichtung ersetzt, tropft nicht mehr i.o LS und Pumpenkal. durchgeführt i.o | SWO-349152 | Defect | HW | Inlet module | Inlet gasket | 0 | Rep. Intake | Drips from inlet during rinsing | Inlet nipple checked i.o Drain hoses checked i.o Drain pump hose i.o Inlet seal replaced, no longer drips i.o LS and pump scale. carried out i.o
No | Visit | Complaint | RAME | All oxy failing | Spectrometer trigger error, neon lamp error and neon intensity error. | Replaced the oxy module. | SWO-349373 | Defect | HW | Oxi module | Lamp unit | neon | All oxy failing | Spectrometer trigger error, neon lamp error and neon intensity error. | Replaced the oxy module.
No | Visit | Complaint | RAME | CoHb not correlating off by 0.5 % Melissa Nason CB:816-407-4555 X6355 | Checked data logs, found customers comparisons and noted the COHb is off by at least .5 on each run. | Performed tHb calibration in the service mode on both analyzers. Ran test samples. Able to get comparisons to match their older current analyzer and .4 from new analyzer comparison. Obtained a fresher sample and was able to get good repeatable realists. Suspect age of samples to be playing a factor for repeatable results on COHb. | SWO-348763 | User error | preanalytical | blood sample | coHb | 0 | CoHb not correlating off by 0.5 % Melissa Nason CB:816-407-4555 X6355 | Checked data logs, found customers comparisons and noted the COHb is off by at least .5 on each run. | Performed tHb calibration in the service mode on both analyzers. Ran test samples. Able to get comparisons to match their older current analyzer and .4 from new analyzer comparison. Obtained a fresher sample and was able to get good repeatable realists. Suspect age of samples to be playing a factor for repeatable results on COHb.
No | Visit | Complaint | RAME | Oxi hardware error Denise Gursky CB:330-344-6912 | Error codes 0251 Halogen Lamp intensity too high, 0579 Oximetry hardware problem. | Adjusted Lamp Unit to recommended parameters. Normal operations restored. | SWO-349347 | Adjustment | HW | Oxi module | Lamp unit | halogen | Oxi hardware error Denise Gursky CB:330-344-6912 | Error codes 0251 Halogen Lamp intensity too high, 0579 Oximetry hardware problem. | Adjusted Lamp Unit to recommended parameters. Normal operations restored.
Yes | Visit | Complaint | RKOR | Sample barcode error |  | Sample barcode check. Barcode check. position check. Opration training. | SWO-350218 | Defect | HW | Basic Unit | Power Supply |  | Sample barcode error |  | Sample barcode check. Barcode check. position check. Opration training.
No | Visit | Complaint | RKK | 溶液パック取り外し時、自動で手前に排出されない | アクチュエーターモーター交換にて改善致しました。 | アクチュエーターモーター交換にて改善致しました。 | SWO-349406 | Defect | HW | Basic Unit | Actuator module | 0 | When the solution pack is removed, it is not automatically discharged to the front. | We have improved the actuator motor replacement. | We have improved the actuator motor replacement.
No | Visit | Complaint | RDEU | 09.05.18 14:00 Frau Junker Tel.023415343 , Pumpenkalibrationsfehler nach einsetzten eines neuen LP´s | LP geliefert | LP geliefert und eingesetzt. Funktion OK. | SWO-350495 | Assisted Replacement | SP | SP | SP | pump tube | 09.05.18 14:00 Mrs Junker Tel.023415343 , Pump calibration error after inserting a new LP | LP supplied | LP is supplied and used. Function OK.
No | Visit | Complaint | RDEU | 09.05.2018 08:50 Schwester Katrin Tel. 023234982187 , Chipdaten des LP können nicht gelesen werden | Chipdaten nicht lesbar. Kunden telefonisch angeleitet, das Lösungspack zu erneuern sowie die Kontakte zu reinigen. Dadurch wurde das Lösungspack zwar erkannt, aber Rückruf des Kunden, da weiterhin Intervention erforderlich (Vakuumfehler, inhomogene Spüllösung). | Lösungspack erneut initialisiert, Einlassdichtung sowie Einlassnadel aus Kundenbestand erneuert. Dadurch Fehler beseitigt. Funktionskontrolle: I.O. | SWO-349760 | Defect | HW | Inlet module | Inlet gasket | 0 | 05/09/2018 08:50 Sister Katrin Tel. 023234982187 , Chip data of the LP cannot be read | Chip data is not readable. customer by phone to renew the solution pack and clean the contacts. This detected the solution pack, but called back the customer, as intervention still required (vacuum error, inhomogeneous flushing solution). | Solution pack reinitialized, inlet seal and inlet needle renewed from customer stock. This eliminates errors. Functional control: I.O.
Yes | 9 | Y | RPAC | Blank | Sensor cassette Calibration failure on Glucose |  | D44DA1D3F25F0F48 | Defect | HW | Basic Unit | Power Supply |  | Blank | Sensor cassette Calibration failure on Glucose
No | Visit | Complaint | RDEU | Flüssigkeitstransportfehler 0508 | Einlaßdichtung defekt | Gerät überprüft, Einlaß korrekt positioniert, Einlaßdichtung erneuert, | SWO-349588 | Defect | HW | Auto Inlet module | Inlet gasket | 508 | Flüssigkeitstransportfehler 0508 | Inlet seal defective | device checked, inlet correctly positioned, inlet seal renewed,
No | Visit | Complaint | RDEU | Probleme beim Wechsel der Einlassanschlussdichtung | Folgeeinsatz nach Ersatzteilbeschaffung. Nach Rücksprache mit dem Kunden ist es seit dem Einsatz vom 7.5. erneut zu Abbrüchen der Messung auf Grund des Fehlers 1190 Einlass geöffnet gekommen. Der Fehler 1190 kann seit dem 15.4. zunächst sporadisch, dann immer häufiger beobachtet werden. | Erneuerung des Einlassmoduls. Funktionskontrolle: I.O. | SWO-349761 | Defect | HW | Auto Inlet module | Auto Inlet module | 0 | Problems changing the inlet connection seal | Follow-up use after spare parts procurement. After consultation with the customer, the measurement has been canceled again since the operation of 7.5. due to the error 1190 admission. Since April 15, error 1190 can be observed sporadically, then more and more frequently. | Renewal of the inlet module. Functional control: I.O.
No | Visit | Complaint | RKK | 検体ＰＨ値が低い | 検体ＰＨ値 7.12 7.05 | カセット、溶液交換 | SWO-350137 | [5 empty] | Low sample PH value | Sample PH value 7.12 7.05 | Cassettes, solution exchange
No | Visit | Complaint | RCHN | 客戶反映sensor定標失敗, 和公司demo機交換sensor後定標OK. 工程師協助換膜之後測試仍然無法通過定標,確認為sensor故障. | 客戶反映sensor定標不過 | 工程師協助換膜後仍無法通過定標,確認為sensor本身問題. | SWO-349433 | [5 empty] | Customers report sensor calibration failure, and the company demo machine exchange sensor after calibration OK. The test after the engineer assisted in the film change still failed the calibration and confirmed the sensor fault. | Customer reflects sensor calibration but | Engineers can not pass the calibration after the film change, identified as the sensor itself problem.
No | Visit | Complaint | RIND | Electrolytes Not calibrating. | Electrolytes not OK. | Reference electrode cleaning done. Chambers cleaning done. pH electrode cleaning done. Analyser Serviced and calibrated properly. Handed over in a fine working condition. | SWO-350605 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway | ref | Electrolytes Not calibrating. | Electrolytes not OK. | Reference electrode cleaning done. Chambers cleaning done. pH electrode cleaning done. Analyser Serviced and calibrated properly. Handed over in a fine working condition.
Yes | Visit | Complaint | RCHN | 自动质控无法上传 | 自动质控无法上传，需要借件910-310 | 更换910-310后修复(By Ran) | SWO-348258 | Defect | HW | Basic Unit | Power Supply |  | Automatic QC cannot upload | Automatic quality control can not be uploaded, need to borrow 910-310 | Repair after replacement 910-310 (By Ran)
No | Visit | Complaint | RDEU | 09.05.2018, 11:01 Uhr, H. Welle, Tel. 063120381271: QC Verletzung bei tHb, pO2 und Gluc. | 581 Oxi Spectrum Missmatch bei QC QC Fehler Elektrolyte | Oximetrie Test ohne Beanstandung - Halogen Lampe eingestellt - tHb KAL - Membranierung und Reinigung Referenz Elektrode etc. - mehrfache Kalibrationen durchgeführt - QC trat in Level 3 mit Oxi Spectrum Missmatch Fehler auf (bei Wiederholung ohne Beanstandung) ! Ersatzteile-Bestellung erfolgt ! | SWO-349812 | Revisit required | HW | Oxi module | Spectrometer | 0 | 09.05.2018, 11:01 Uhr, H. Welle, Tel. 063120381271: QC Verletzung bei tHb, pO2 und Gluc. | 581 Oxi Spectrum Missmatch bei QC QC Fehler Elektrolyte | Oximetry test without complaint - halogen lamp set - tHb KAL - membraneing and cleaning reference electrode etc. - multiple calibrations performed - QC occurred in level 3 with Oxi Spectrum Missmatch error (in repetition without complaint) ! Spare parts order is made !
No | Visit | Complaint | RAME | Oxi issues on QC. He would like to schedule a service call. | Inspection of the oxi system revealed the hemolyzer leaked QC fluid onto the Oxi PCB which then had a burned smell. The tube valve works but seems hesitant at times, replaced for diagnostic purposes. | Replaced oxi pinch valve, Oxi PCB, & hemolyzer. | SWO-349400 | Defect | HW | Oxi module | Hemolyzer | cuvette | Oxi issues on QC. He would like to schedule a service call. | Inspection of the oxi system revealed the hemolyzer leaked QC fluid onto the Oxi PCB which then had a burned smell. The tube valve works but seems hesitant at times, replaced for diagnostic purposes. | Replaced oxygen pinch valve, oxygen PCB, & Hemolyzer.
No | Visit | Complaint | RDEU | 09.05.2018 Herr Kammermeier 08161243540 meldet einen Einlasskalibrierfehler. | 0537: Oxi Liquid Sensor Error. 0676: Inlet Error - Sensitivity out of Range. Einlasskalibrierfehler Kunde hat Einlassdichtung schon erneuert. | Einlassnadel ist nicht verbogen. Zum Test eingebaut. Fehler weiterhin vorhanden. Neue Einlassnadel wieder eingebaut. Neustart brachte keinen Erfolg. Inlet Modul ausgebaut und gereinigt. Antriebswelle des Inlet Motors geölt. 5x Inlet Kalibration im Service durchgeführt - jetzt ok. Luftweg über V7 gespült. Initialisierungsphase wird nun kpl. durchgeführt. LS Justage - ok. Kalibrationen und QC’s vorgenommen. | SWO-349671 | Defect | HW | Inlet module | Spinlet | 0 | 09.05.2018 Mr. Kammermeier 08161243540 reports an inlet calibration error. | 0537: Oxi Liquid Sensor Error. 0676: Inlet Error - Sensitivity out of Range. Einlasskalibrierfehler Kunde hat Einlassdichtung schon erneuert. | Inlet needle is not bent. Built for testing. Errors still present. New inlet needle reinstalled. Restart did not succeed. Inlet module removed and cleaned. drive shaft of the Inlet engine. 5x Inlet calibration performed in service - now ok. Airway via V7. Initialization phase is now kpl. Performed. LS Adjustment - ok. calibrations and QC's.
No | Visit | Complaint | RSCH | Rep. Einlass | Tropft aus Einlass während Spülung | Im Service-Menu Einlass Kalibriert i.o Spülung durchgeführt, ohne Erfolg Einlassdichtung ersetzt, tropft nicht mehr aus Einlass i.o System i.o | SWO-349698 | Defect | HW | Inlet module | Inlet gasket | 0 | Rep. Intake | Drips from inlet during rinsing | In the service menu inlet Calibrated i.o Rinse performed, without success Inlet seal replaced, no longer drips from inlet i.o System i.o
No | Visit | Complaint | RIND | Service Visit. | User select the TnT while no test cartridge of TnT inserted by user. | Check the activity log, after observation go to service program solve the interference of storage wheel, run safe position and check the punch up retracted in solution pack, solve the interference. Now exit from the service program, after start up analyzer is in ready mode, check the system clean status, found ok. Insert the TnT test cartridge and analyzer handed over to staff in well working condition. | SWO-350872 | Assisted Replacement | Consumables | Test Kit | cartridge, TnT |  | Service Visit. | User select the TnT while no test cartridge of TnT inserted by user. | Check the activity log, after observation go to service program solve the interference of storage wheel, run safe position and check the punch up retracted in solution pack, solve the interference. Now exit from the service program, after start up analyzer is in ready mode, check the system clean status, found ok. Insert the TnT test cartridge and analyzer handed over to staff in well working condition.
No | Visit | Complaint | RDEU | 09.05.2018 15:30 Fr. Krebs 017634355914: Setsafe22 Fehler | D-Dimer Kalibration wird abgebrochen, Processunit 24 | Gerät demontiert, Cups aus dem Gerät entfernt, D-Dimer erneut kalibriert, Abbruch, neue Kal. Kassetten für Montag bestellt, | SWO-349966 | Defect | Consumables | Solution Pack | stack |  | 05/09/2018 15:30 Fr. Krebs 017634355914: Setsafe22 Error | D-Dimer calibration is aborted, Processunit 24 | Device dismantled, cups removed from the device, D-dimer recalibrated, cancellation, new cal. Cassettes ordered for Monday,
Yes | 4 | Y | RPAC | Blank | Failed SC |  | D4563D312C178BD0 | Defect | HW | Basic Unit | Power Supply |  | Blank | Failed SC
No | Visit | Complaint | RKK | SC・SP接続不良 多発 | SC・SP接続不良 多発 | 装置分解 SPセンサー位置調整 各動作確認 | SWO-357380 | [5 empty] | SC/SP connection failure frequently | SC/SP connection failure frequently | Equipment decomposition SP sensor position adjustment Check each operation
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 无法阅读试剂包 | 更换试剂包修复 | SWO-350633 | [5 empty] | Test card, reagent package failure | Unable to read the reagent package | Replacement reagent kit repair
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | LAC定标不过 0378 | 更换测试卡后修复 | SWO-356184 | Assisted Replacement | SC | SC | Lac | 0 | Test card, reagent package failure | LAC calibration is no more than 0378 | Repair after replacing test card
No | Visit | Complaint | RCHN | 报错1999.61.22 | 加借920-775一个。 | Use 920-775 to replace the rubber disk(By Ran) | SWO-348976 | Defect | HW | Process Module | rubber disk | 1999 | Error 1999.61.22 | Borrow 920-775 one. | Use 920-775 to replace the rubber disk(By Ran)
Yes | Visit | Complaint | RCHN | 定标失败 1024 | 定标失败 1024，借903-364，903-236各一个。 | 更换903-364，903-236各一个正常。(By Ran) | SWO-349608 | Defect | HW | Basic Unit | Power Supply |  | Bid Failure 1024 | Bid failure 1024, borrow 903-364, 903-236 each. | Replace 903-364, 903-236 each normal. (By Ran)
No | Visit | Complaint | RCHN | 八楼的两台机器都出现问题,5663&4036&2873 | 八楼的两台机器都出现问题,5663&4036&2873， | 检查看两台仪器，进样盘，孵育盘，CKMB的条码切割异常，扫不到测试卡，重新粘贴安装。 | SWO-350848 | Defect | Consumables | Test Kit | Cartridge, CKMB |  | Problems with both machines on the eighth floor, 5663-4036-2873 | Both machines on the eighth floor have problems, 5663-4036-2873, | Check the two instruments, the sample disk, the incubation disk, CKMB's bar code cutting abnormal, can not sweep the test card, repaste installation.
No | Visit | Complaint | RCHN | 报HCT值不能被测量 | 报HCT值不能被测量，借903-329，636-561，903-209各一个 | 更换636-561，903-329后修复(By Ran) | SWO-349585 | Defect | HW | Aspiration Module | Needle Wash Unit |  | Report HCT value sit | The HCT value cannot be measured, one of 903-329, 636-561, 903-209 | Replacement 636-561,903-329 post-repair (By Ran)
Yes | Visit | Complaint | RCHN | 报错1031 | 报错1031，需要借件903-364 | 更换903-364一个正常。(By Ran) | SWO-347408 | Defect | HW | Basic Unit | Power Supply |  | 1031 | 1031, need to borrow 903-364 | Replace 903-364 one normal. (By Ran)
No | Visit | Complaint | RCHN | 报错1031 | 报错1031，借903-364一个 | 更换903-364后修复(By Ran) | SWO-348954 | Defect | HW | Process Module | Optical Unit | 1031 | 1031 | 1031, loan 903-364 one | Repair after replacement 903-364 (By Ran)
No | Visit | Complaint | RCHN | 报错5784 | 报错5784 | 更换测试卡后修复 | SWO-350177 | Defect | Consumables | Test Kit | Cartridge, Broken | 5784 | 5784 | 5784 | Repair after replacing test card
No | Visit | Complaint | RLTD | Scanner error on board | Scanner error | Replaced and tested new scanner. Passed. Took files off machine - software issues and other 2 AQT90s on site (R59N6 & R94N2)as high photo count errors on board. RMED to investigate issues. | SWO-350283 | Defect | HW | Storage Module | Barcode reader, storage wheel |  | Scanner error on board | Scanner error | Replaced and tested new scanner. Passed. Took files off machine - software issues and other 2 AQT90s on site (R59N6 & R94N2)as high photo count errors on board. RMED to investigate issues.
No | Visit | Complaint | RKOR | Inlet probe bent. | . | Inlet probe replace. | SWO-351118 | Defect | HW | Inlet module | Probe bend | 0 | Inlet probe bent. | . | Inlet probe replace.
Yes | 6 | Y | RPAC | Blank | Failed SC |  | D4560F962C498BCC | Defect | HW | Basic Unit | Power Supply |  | Blank | Failed SC
No | Visit | Complaint | RLTD | Sensor bay won't close after SC replacement | Customer reported that sensor bay door tries to close but doesn't. This is after a sensor cassette replacement. Looks to close slightly before reopening and stays open. Site visit to investigate. | Arrived on site and attended analyser. Found analyser with sp and sc removed and I service menus. Attempted to run close bay program but this did not close the door. Motor running but door reopens and stays open. Interface pcb bay light not on when door open. Manually wound door closed and then ran open program. Door opened but would not close again. Powered down and took wet section out of analyser. Checked and reseated cables for wetsection from main au board. Found fpc cable from au board to interface pcb was not seated correctly. Securing clip was open and cable at an angle in connector. Refitted this and secured. Refitted all parts and retested. Refitted all consumables. Sensor door opened, fitted cassette and then door automatically closed. Also, bay light now switched on. Cable was issue. Analyser did the full startup and Conditioning. Ran cal and qcs and these both passed. | SWO-350824 | Adjustment | HW | Basic Unit | cable | 0 | Sensor bay won't close after SC replacement | Customer reported that sensor bay door tries to close but doesn't. This is after a sensor cassette replacement. Looks to close slightly before reopening and stays open. Site visit to investigate. | Arrived on site and attended analyser. Found analyser with sp and sc removed and I service menus. Attempted to run close bay program but this did not close the door. Motor running but door reopens and stays open. Interface pcb bay light not on when door open. Manually wound door closed and then ran open program. Door opened but would not close again. Powered down and took wet section out of analyser. Checked and reseated cables for wetsection from main au board. Found fpc cable from au board to interface pcb was not seated correctly. Securing clip was open and cable at an angle in connector. Refitted this and secured. Refitted all parts and retested. Refitted all consumables. Sensor door opened, fitted cassette and then door automatically closed. Also, bay light now switched on. Cable was issue. Analyser did the full startup and Conditioning. Ran cal and qcs and these both passed.
Yes | Visit | Complaint | RDEU | 17.04.2018 13:15Uhr Teis 02611377310: Der Einlass schließ nicht bzw. fährt auch nicht mehr aus. | Service Dump gezogen. | An EVARA verschickt. | SWO-351049 | Re classified RMED | SD request from RMED | 0 | 0 | 0 | 17.04.2018 13:15h Teis 02611377310: The entrance does not close or does not stop. | Service Dump pulled. | Sent to EVARA.
No | Visit | Complaint | RKK | SC・SP接続不良 多発 | SC・SP接続不良 多発 | 修理完了装置返却 | SWO-357381 | [5 empty] | SC/SP connection failure frequently | SC/SP connection failure frequently | Return repair completion equipment
No | Visit | Complaint | RIND | Service Visit. | Analyzer shows solution pack is not connected. | Re-install the solution pack again, but its hanging and no message came for installation complete. After cancelling analyzer shows solution pack is not connected while solution pack is connected. Need to replace the solution pack. Change the solution pack and install the sensor cassette, successfully done. Analyzer handed over to staff in well working condition. | SWO-350543 | [5 empty] | Service Visit. | Analyzer shows solution pack is not connected. | Re-install the solution pack again, but its hanging and no message came for installation complete. After cancelling analyzer shows solution pack is not connected while solution pack is connected. Need to replace the solution pack. Change the solution pack and install the sensor cassette, successfully done. Analyzer handed over to staff in well working condition.
No | Visit | Complaint | RAME | Customer reporting WDC data indicates problem with sO2 running above peer range. | reviewed WDC data, confirmed sO2 trending high along with FCOHb and MetHb dropping. | 5 $10 $2018 - Inspected analyzer, ran complete Oxi test, failed for lamp (neon intensity low @ 46%) and hemolyzer (impedance shape, phase Reg error). Reran service tHB, no impact on Oxi test, still failing. Ordered Oxi module, will return next day to complete repairs. 5 $11 $2018 - Replaced complete Oxi module, performed all associated adjustments and test, no failures. Ran VKR7 kit for customer due to Oxi failure and up coming proficiency test. | SWO-351638 | Defect | HW | Oxi module | Oxi module | 0 | Customer reporting WDC data indicates problem with sO2 running above peer range. | reviewed WDC data, confirmed sO2 trending high along with FCOHb and MetHb dropping. | 5 $10 $2018 - Inspected analyzer, ran complete Oxi test, failed for lamp (neon intensity low @ 46%) and hemolyzer (impedance shape, phase Reg error). Reran service tHB, no impact on Oxi test, still failing. Ordered Oxi module, will return next day to complete repairs. 5 $11 $2018 - Replaced complete Oxi module, performed all associated adjustments and test, no failures. Ran VKR7 kit for customer due to Oxi failure and up coming proficiency test.
No | Visit | Complaint | RDEU | 04.05.2018; Hr. Henneboele; Tel.: 02093643880; Na und Ca Kal Fehler 04.05.2018; 11:40 Uhr; Hr. Henneboele; Tel.:02093644504; weiterhin Na und Ca Driftfehler | Na bei Kalibrationen mit ?, Statusfehler, | Elektrode ersetzt und Gerät kalibriert, | SWO-350715 | Assisted Replacement | Electrode | Electrode | Na | 0 | 04.05.2018; Mr. Henneboele; Tel.: 02093643880; Na and Ca Kal Error 04.05.2018; 11:40 a.m.; Mr. Henneboele; Tel.:0 2093644504; continue Na and Ca drift errors | Na for calibrations with ?, status errors, | replaced by the electrode and calibrated the device,
No | 2 | Y | RPAC | Blank | Drifting after Cleaning program on Electrolytes |  | D44FD0BEDE2C1DFB | User error | User error | user error | user error | 0 | Blank | Drifting after Cleaning program on Electrolytes
No | Visit | Complaint | RDEU | 11.05.2018, Sr. Eisel, Tel: 021513245158: Fehler bei p02 | pO2 Nullstromfehler, QC Fehler 210 | pO2 Kanal kalibriert, Kal und QC wiederholt, Funktion OK. | SWO-351030 | Adjustment | HW | pH/BG Module | pO2 | 0 | 11.05.2018, Sr. Eisel, Tel: 021513245158: Error with p02 | pO2 Nullstromfehler, QC Fehler 210 | pO2 channel calibrated, Kal and QC repeated, function OK.
No | Visit | Complaint | RBNL | AQT herkent niet dat er een buis geplaatst wordt. | Bij plaatsen buis is te zien dat de barcodereader niet werkt (geen rode laser waarneembaar). Vervangen van de barcodereader lost het probleem niet op. Hierna de FPC kabel van de barcodereader vervangen, nu werkt de barcode reader weer zoals het hoort. Originele barcode reader terug geplaatst, deze werkt ook zoals het hoort. AQT opgestart en Testmeting gedaan —>oke. Systeem controle uitgevoerd —> oke. | FPC kabel van de barcode reader van de inlet vervangen. | SWO-350617 | Defect | HW | Inlet module | Barcode reader, Inlet |  | AQT does not recognise that a tube is inserted. | When inserting tube it is seen that the barcode reader does not work (no red laser observable). Replacing the barcode reader does not solve the problem. After this, replace the FPC cable from the barcode reader, now the barcode reader works again as it should. Original barcode reader placed back, this also works as it should. AQT booted and Test measurement done — > okay. System Check performed — > okay. | FPC kabel van de barcode reader van de inlet vervangen.
Yes | Visit | Complaint | RAME | PCO2 is not passing calibrations. Two sensor cassette already have been replaced. | PCO2 keeps failing calibrations. | The customer already replaced two sensor cassettes and, one solution pack. I check the analyzer by running the sensor simulator test. Found the sensor interface pcb is at fault, not passing the electrical noise check. Replaced a new sensor interface pcb and, a new sensor cassette from my stock. Ran the fluid transport tests, passing within specifications. | SWO-350981 | Defect | HW | Basic Unit | Power Supply |  | PCO2 is not passing calibrations. Two sensor cassette already have been replaced. | PCO2 keeps failing calibrations. | The customer already replaced two sensor cassettes and, one solution pack. I check the analyzer by running the sensor simulator test. Found the sensor interface pcb is at fault, not passing the electrical noise check. Replaced a new sensor interface pcb and, a new sensor cassette from my stock. Ran the fluid transport tests, passing within specifications.
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | CL质控失败 | 更换测试卡后修复 | SWO-354407 | [5 empty] | Test card, reagent package failure | CL QUALITY CONTROL FAILED | Repair after replacing test card
No | Visit | Complaint | RCHN | 蠕动板破裂 | 蠕动泵破裂，借902-971一个 | 更换902-971后修复(By Ran) | SWO-350572 | [5 empty] | Creep plate rupture | Creep pump rupture, borrow 902-971 one | Post-replacement 902-971 repair (By Ran)
Yes | Visit | Complaint | RCHN | 机器蓝屏 | 机器蓝屏，借903-198一个（上海缺货，只有从北办发货） | 更换910-343、903-198重新定标后修复(By Ran) | SWO-350615 | Defect | HW | Basic Unit | Power Supply |  | Machine blue screen | Machine blue screen, borrow 903-198 one (Shanghai out of stock, only from north to ship) | Replace 910-343, 903-198 Relabeling Repair (By Ran)
No | Visit | Complaint | RCHN | BNP结果偏高 | BNP结果偏高，借944-258（14147）一张。 | 重新定标后修复(By Ran) Recal , then OK. | SWO-349643 | Assisted Replacement | Consumables | Test Kit | cartridge, BNP |  | BNP results are high | The BNP result was high, with a 944-258 (14147) one. | Re-calibration repair (By Ran) Recal , then OK.
No | Visit | Complaint | RAME | PCO2 is failing QC running high out side the insert range. Before repair it was running low | After replacement of the gas module and measuring chamber pCO2 is now running high outside the insert package ranges 7745 lot 775 range is 36.8 to 42.8 coming in at 45.9 7755 lot 736 range is 20-24 coming in at 24.1 7735 Lot 761 range is 60.9 - 70.9 coming in at 70.6 | Turned off power and checked both valves they look good. Replaced the blue oring on the ph electrode; Remembraned the ref--may have found a hair line crack on the ref electrode--if problem continues will order Ref electrode; Remembraned the PO2, Grey tubing came loose which causes numerous errors on start up. PCO2 is now at 65.1 for 7735, 39.2 for 7745 and 22.5 for 7755, right on the mean | SWO-351070 | Assisted Replacement | Electrode | Electrode | pCO2 | 0 | PCO2 is failing QC running high out side the insert range. Before repair it was running low | After replacement of the gas module and measuring chamber pCO2 is now running high outside the insert package ranges 7745 lot 775 range is 36.8 to 42.8 coming in at 45.9 7755 lot 736 range is 20-24 coming in at 24.1 7735 Lot 761 range is 60.9 - 70.9 coming in at 70.6 | Turned off power and checked both valves they look good. Replaced the blue oring on the ph electrode; Remembraned the ref--may have found a hair line crack on the ref electrode--if problem continues will order Ref electrode; Remembraned the PO2, Grey tubing came loose which causes numerous errors on start up. PCO2 is now at 65.1 for 7735, 39.2 for 7745 and 22.5 for 7755, right on the mean
No | Visit | Complaint | RCHN | HCT值无法测定 | HCT值无法测定，借903-329一个 | 更换903-329后修复(By Ran) | SWO-351018 | Defect | HW | Aspiration Module | Needle Wash Unit |  | HCT value cannot be measured | HCT value can not be determined, borrow 903-329 one | Repair after replacement 903-329 (By Ran)
No | Visit | Complaint | RCHN | 报错2887 | 报错2887 | 重新清理处理盘后修复 by jesson | SWO-351180 | Cleaned | HW | Storage Module | Cartridge lock | 2887 | 2887 | 2887 | Fix by jesson after recleaning the processing disk
No | Visit | Complaint | RCHN | 血氧检测出错 0581 | 血氧检测出错 0581，借910-274一个。 | 机器未修复，加借902-833一个(By Ran) | SWO-350627 | Defect | HW | Oxi module | Oxi module | 0 | Blood oxygen test error 0581 | Blood oxygen test error 0581, borrow 910-274 one. | Machine not repaired, plus 902-833 (By Ran)
Yes | Visit | Complaint | RCHN | 血氧检测出错 0581 | 一次未修复，加借902-833一个 | 更换备用的902-833后修复(By Ran) | SWO-351091 | Re classified RMED | duplicate upSWO | 0 | 0 | 0 | Blood oxygen test error 0581 | One unrepaired, plus 902-833 one | Replacement of spare 902-833 post-fix (By Ran)
No | Visit | Complaint | RDEU | Glu und Lac Fehler in Kal | Glu undLac Status viel zu hoch empf zu niedrig | Materialfehler!!! SC reinstaliiert Status IO Empf zu niedrig, SC getauscht und reklamiert, man kal und QC vermessen | SWO-351845 | Assisted Replacement | SC | SC | Glu/Lac | 0 | Glu and Lac Errors in Kal | Glu andLac status far too high empf too low | Material defects!!! SC pureized status IO empf too low, SC swapped and complained, one scale and QC measured
No | Visit | Complaint | RDEU | QC-Verletzung FHbF und sO2 | Seit 5.5.18 QC in Level C Parameter FHbF erhöht Seit 11.5.18 QC in Level B Parameter sO2 und FHbF erhöht QC in Level A ohne Beanstandung Kalibrationen ohne Beanstandung Keine Aktivitäten-Meldungen | Oximetrie Test im Service ohne Beanstandung - Analysator demontiert und Erneuerung Quetschventil - Rückmontage und Funktionsüberprüfung etc. und tHb KAL | SWO-351561 | Defect | HW | Oxi module | Pinch valve | 0 | QC violation FHbF and sO2 | Since 5.5.18 QC in Level C parameter FHbF increased Since 11.5.18 QC in Level B parameter sO2 and FHbF increased QC in Level A without complaint Calibrations without complaint No activity messages | Oximetry test in service without complaint - Analyzer disassembled and renewal crushing valve - Reassembly and function check etc. and tHb KAL
No | Visit | Complaint | RAME | Leak/ Pressure Vac error | Pressure test failing vacuum value out of range Pump cal failing Possible air in line or clot Flushed analyzer Ran service e tests Fluid flow tests all passing after flush and re install in service | Flush and re install after service tests completed for fluid flow | SWO-352335 | Cleaned | Fluidic pathway | Fluidic pathway | clot | 0 | Leak/ Pressure Vac error | Pressure test failing vacuum value out of range Pump cal failing Possible air in line or clot Flushed analyzer Ran service e tests Fluid flow tests all passing after flush and re install in service | Flush and re install after service tests completed for fluid flow
Yes | Visit | Complaint | RDEU | Analysator steht im Fehler | Analysator steht im Fehler | Analysator steht im Fehler. Lösungspack und Sensorkassette getauscht. Startphase abgewartet. Kein Lösungspack mehr vorhanden. Frau Gerstmayr mitgeteilt dass Techniker am nächsten früh nochmals kommt und einen Lösungspack mitbringt. Analysator nicht messbereit. | SWO-351379 | Re classified RMED | Logistical SWO | SP | 0 | 0 | Analyzer is in error | Analyzer is in error | Analyzer is in error. Solution pack and sensor cartridge. start-up phase. No more solution packs available. Ms. Gerstmayr informs that technicians will come again the next time and bring a solution pack. Analyzer not ready to measure.
No | Visit | Complaint | RLTD | Battery reading zero % Scarlett Willet 0203-4483479 | Battery reading zero % | Battery reading zero % and when main power removed analysers shuts down. Removed battery and tried to reset battery but still shows zero % after half hour of being on main power. New battery required. Part on order. | SWO-351600 | Revisit required | HW | Battery module | Battery module | 0 | Battery reading zero % Scarlett Willet 0203-4483479 | Battery reading zero % | Battery reading zero % and when main power removed analysers shuts down. Removed battery and tried to reset battery but still shows zero % after half hour of being on main power. New battery required. Part on order.
No | Visit | Complaint | RDEU | Intervention erforderlich (Einlass schließt nicht, bleibt in Spritzenposition stehen) | Fehler von Sonntag 13.5.18: Einlass schließt nicht richtig, bleibt in Kapillarposition stehen (Einlass fährt zu, versucht zu schließen, schafft dies nicht und fährt zurück in die Kapillarposition --> Kunde hat schon die Einlassnadel, Einlassdichtung und Lösungspack erneuert --> Einlass fährt wieder zu, bleibt dennoch in Intervention erforderlich mit Flüssigkeits-Transportfehler Fehlersuche --> Einlassnadel, Einlassdichtung, Einlassmodul, Lösungspack, Quetschventil, Analysator im Service überprüfen | Einlassmodul, Einlassdichtung und Nadel überprüft und durchgespült, Einlassmodul wieder angebracht, Analysator im Service (Adjust Manifold, Pressuretest, Pumpenkal., LS-Justage, SP Check, Clot detect $remove) überprüft --> Servicetest OK, Analysator neu gestartet, Startphase abgewartet, System und Funktion i.O. | SWO-351447 | Adjustment | HW | Auto Inlet module | Probe Cal | 0 | Intervention required (inlet does not close, remains in syringe position) | Error of Sunday 13.5.18: Inlet does not close properly, stops in capillary position (entrance closes, tries to close, does not manage this and drives back to the capillary position --> Customer has already renewed the inlet needle, inlet seal and solution pack --> Inlet shuts down again, but remains required in intervention with liquid transport error Troubleshooting --> Check inlet needle, inlet seal, inlet module, solution pack, crush valve, analyzer in service | Inlet module, inlet seal and needle checked and rinsed, inlet module reattached, analyzer in service checked (Adjust Manifold, Pressuretest, Pump Calco., LS-Adjustment, SP Check, Clot detect $remove) --> Service test OK, Analyzer restarted, start-up phase awaited, system and function in the
Yes | 5 | Y | RPAC | Blank | Instructions for Use |  | D4561F592C128BCE | Re classified RMED | IFU | 0 | 0 | 0 | Blank | Instructions for Use
No | 1 | Y | RPAC | Blank |  |  | D453C450403BC5D3 | Defect | HW | Basic Unit | Pump rotor | 0 | Blank
No | Visit | Complaint | RESP | BASE DE DATOS CORRUPTA | CF CARD ERROR | SUSTITUCION DE CF CARD POR CORRUPCION DE DATOS EN LA MISMA, CONFIGURACION DE TODO EL SISTEMA Y PRUEBAS DE FUNCIONAMIENTO. CALIBRADO Y VEDIFICADO CON Q.C. | SWO-351411 | [5 empty] | CORRUPTA DATABASE | CF CARD ERROR | SUBSTITUTION OF CF CARD FOR DATA CORRUPTION IN IT, CONFIGURATION OF ALL SYSTEM AND OPERATION TESTS. CALIBRATED AND VEDIFIED WITH Q.C.
No | Visit | Complaint | RESP | BASE DE DATOS CORRUPTA | ERROR CF CARD | SUSTITUCION DE CF CARD POR CORRUPCION DE DATOS EN LA MISMA, CONFIGURACION DE TODO EL SISTEMA Y PRUEBAS DE FUNCIONAMIENTO. CALIBRADO Y VEDIFICADO CON Q.C. | SWO-352482 | [5 empty] | CORRUPTA DATABASE | ERROR CF CARD | CF CARD REPLACEMENT FOR DATA CORRUPTION IN IT, CONFIGURATION OF ALL SYSTEM AND TESTS Operation. CALIBRATED AND VEDIFIED WITH Q.C.
No | Visit | Complaint | RAME | PO2 fails on L1 and L4 external QC. | SC lot# 254207. SP lot# 255142 experience on the other Instrument at this location. The same lot numbers of consumables is being used on both instruments starting in April. This was when the issue was first noticed. | Perform sensor Zero in service diagnostics. Perform zero channels in service diagnostics. Would monitor instruments and have feedback from the customer. Instrument is functional and all tests are successful. | SWO-351682 | [5 empty] | PO2 fails on L1 and L4 external QC. | SC lot# 254207. SP lot# 255142 experience on the other Instrument at this location. The same lot numbers of consumables is being used on both instruments starting in April. This was when the issue was first noticed. | Perform sensor Zero in service diagnostics. Perform zero channels in service diagnostics. Would monitor instruments and have feedback from the customer. Instrument is functional and all tests are successful.
No | Visit | Complaint | RCHN | 报错1031 | 报错1031，借944-212（13914）,944-394（14070），944-258（14051）各一个 |  | SWO-351195 | Defect | HW | Process Module | Optical Unit | 1031 | 1031 | 1031, loan 944-212 (13914), 944-394 (14070), 944-258 (14051) each
Yes | Visit | Complaint | RCHN | 报错1031 | 报错1031，借903-364一个（广办自提） |  | SWO-351196 | Re classified RMED | same day upSWO |  |  |  | 1031 | 1031, borrow 903-364 one (wide-office self-mentioned)
No | Visit | Complaint | RCHN | 报错1031 | 报错1031，需要借903-364、定标卡：944-212（13841）、944-214（13886）、944-258（13918）和944-394（13954）、944-222（13628）缺货、944-216（13484）暂时不借了 |  | SWO-351271 | Defect | HW | Process Module | Optical Unit | 1031 | 1031 | For error 1031, 903-364, bid card: 944-212 (13841), 944-214 (13886), 944-258 (13918) and 944-394 (13954), 944-222 (13628) 944-216 (13484) temporarily unborrowed
No | Visit | Complaint | RCHN | 进样门的夹断裂 | 进样门的夹断裂，需要借件903-584一个 |  | SWO-351236 | Defect | HW | Inlet module | Tube holder |  | The clip of the inlet door is broken. | The clip of the inlet door is broken and requires a loan of 903-584.
Yes | Visit | Complaint | RIND | Display is not working | Checked with new display but found that touch panel is not working. | replaced touch panel from old display. Then checked display. It is working fine | SWO-354444 | Defect | HW | Basic Unit | Power Supply |  | Display is not working | Checked with new display but found that touch panel is not working. | replaced touch panel from old display. Then checked display. It is working fine
Yes | Visit | Complaint | RCHN | 血氧质控错误 | 血氧质控错误，借902-833一个 |  | SWO-351319 | Defect | HW | Basic Unit | Power Supply |  | Blood oxygen control error | Blood oxygen control error, borrow 902-833 one
No | Visit | Complaint | RCHN | ca数值偏高 | ca数值偏高，需要借件902-755一个 |  | SWO-351225 | Defect | HW | Spinlet module | heater | 0 | Ca value is high | Ca value is high, need to borrow 902-755 one
No | Visit | Complaint | RSCH | Rep Probenproblem | Seit 1:28 Proben Abbruch Medizintechnik First Level Support | Ventile und Flüssigkeitsweg kontrolliert i.o LS Justage ausgeführt 6 Messungen i.o Funktionskontrolle i.o | SWO-351444 | Adjustment | HW | Basic Unit | LS | 0 | Rep Probenproblem | Since 1:28 samples demolition Medical Technology First Level Support | Valves and fluid path controlled i.o LS Adjustment Executed 6 Measurements i.o Functional control i.o
No | Visit | Complaint | RDEU | 14.05.2018, Herr magiera, Tel. 059311524482: Kalibrationsfehler bei pCO2 das ganze Wochenende, Herr Magiera hat alles gereinigt und neu membraniert. Kal Daten schwanken zwischen 80-105% Empfindlichkeit. | Zeitweise pCO2 Driftfehler | Gerät überprüft, Meßkammer und Luftauslass gesäubert, Elektroden kontrolliert, Dichtungsring pH und Pumpenschlauch pH $BG Modul ersetzt, Kal und QC o.k. | SWO-351849 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway | 0 | 14.05.2018, Mr. magiera, tel. 059311524482: Calibration error at pCO2 all weekend, Mr. Magiera cleaned and re-logged everything. Kal data vary between 80-105% sensitivity. | Intermittent pCO2 drift error | Device checked, measuring chamber and air outlet cleaned, electrodes checked, sealing ring pH and pump hose pH $BG module replaced, Kal and QC ok.
No | Visit | Complaint | RAME | Rinse errors. | Checked data logs. Observed fluid flow during a rinse cycle. Found that rinse was not running through the system at all. Checked all valves and tubing for cracks or breaks. Found dual valve housing nipple on the pH module to be broken | Replaced dual valve housing. | SWO-351724 | Defect | HW | pH/BG Module | Dual Valve | 0 | Rinse errors. | Checked data logs. Observed fluid flow during a rinse cycle. Found that rinse was not running through the system at all. Checked all valves and tubing for cracks or breaks. Found dual valve housing nipple on the pH module to be broken | Replaced dual valve housing.
No | Visit | Complaint | RFRA | Erreur 5784 SETSAFE22 | Erreur 5784 SETSAFE22 | Cartouche coincée car ailette avant cassée. Dégagement cartouche ok. Ailette retrouvée à bord et éliminée. Test service ok Réglage lecteur code barre ok Redémarrage ok. Présentation trouble shooting et du menu service aux utilisateurs. | SWO-351567 | Defect | Consumables | Test Kit | Cartridge, Broken |  | Error 5784 SETSAFE22 | Error 5784 SETSAFE22 | Stuck cartridge because broken front wing. Clearing cartridge ok. Wing found on board and eliminated. Test service ok Setting drive barcode ok Restart ok. Introducing trouble shooting and menu service to users.
No | Visit | Complaint | RDEU | 09.05.2018 15:30 Fr. Krebs 017634355914: Setsafe22 Fehler | Cartridge Transport fehlerhaft | Geräte Funktionen geprüft, Storagewheel Einheit demontiert und geprüft, Cartridge Slider justiert, Funktionskontrolle ok | SWO-351645 | Adjustment | HW | Storage Module | Cartridge slider |  | 05/09/2018 15:30 Fr. Krebs 017634355914: Setsafe22 Error | Cartridge transport faulty | Devices functions checked, Storagewheel unit disassembled and tested, Cartridge slider adjusted, function control ok
No | Visit | Complaint | RDEU | Kalpack wurde ersetzt, pO2 fehlerhaft, Kalpackwechsel erneut gefordert | Fehlersuche, Service aufgerufen und diverse Tests durchgeführt, Einlaßdichtung bewegt sich nur schwer in Ausgangslage zurück | Dichtung mit Wasser benetzt - Bewegung wieder leichtgängig, Dichtung zu Messkammer korrekt befestigt, Drucktest i.O., mehrere Kal, fällige tHb Kal durchgeführt, QC | SWO-352500 | Cleaned | HW | Inlet module | Inlet gasket | 0 | Kalpack has been replaced, pO2 faulty, calpack change again required | Troubleshooting, service called and various tests carried out, inlet seal hardly moves back in initial position | Seal wetted with water - movement again smooth,seal to measuring chamber correctly attached, pressure test i.O., multiple kal, due tHb Kal performed, QC
No | Visit | Complaint | RDEU | Analysator steht im Fehler | Erneute Fehlersuche. | Mehrmals Lösungspack erneuert. Mehrmals im Sensorkassette erneuert. Kalibrierungen durchgeführt. Lac und Glu Fehler ( Reaktionsfehler). Frau Gerstmayr mitgeteilt das eine neue Sensor Interface PCB bestellt ist und diese eingebaut wird sobald sie kommt. Sie wird morgen nochmals eine neue Sensorkassette einsetzen vorausgesetzt die jetzt eingelegte Kassette Stabilisiert sich nicht und das versprochene Betriebsmaterial wird morgen geliefert. QCs für alle Level durchgeführt. ABL bis auf Glu und Lac messbereit. | SWO-352165 | Assisted Replacement | SC | SC | Glu/Lac | 0 | Analyzer is in error | Re-troubleshooting. | Renewed several times solution pack. Renewed several times in the sensor cartridge. calibrations. Lac and Glu error (reaction error). Ms. Gerstmayr announced that a new Sensor Interface PCB has been ordered and this will be installed as soon as it arrives. It will use a new sensor cartridge again tomorrow, provided that the cassette that is now inserted does not stabilize and the promised operating material will be delivered tomorrow. QCs for all levels. ABL ready to be measured except for Glu and Lac.
Yes | 2 | Y | RPAC | Blank | Annual service required on the analyser. |  | D454E386B6525F42 | Re classified RMED | annual service | 0 | 0 | 0 | Blank | Annual service required on the analyser.
No | Visit | Complaint | RIND | sensor cassette not installing properly. also zeroing also not ok. | Found that the there is no changes in value in diagnostic mode after placing zeroing, | Checked and found that the Data signal board and Sensor board Faulty. need to replace sensor board and Data Signal Board. | SWO-353685 | [5 empty] | sensor cassette not installing properly. also zeroing also not ok. | Found that the there is no changes in value in diagnostic mode after placing zeroing, | Checked and found that the Data signal board and Sensor board Faulty. need to replace sensor board and Data Signal Board.
No | Visit | Complaint | RIND | Leakage from the back of solution pack on ABL80 analyzer at Surya Hospital, Pune | Need to check all the tubes for any clots. | Found that all the tubes are blocked so cleaned all the tubes, along with the waste pump and valve, reconnected everything, installed both SP and SC, machine is working fine. | SWO-354235 | [5 empty] | Leakage from the back of solution pack on ABL80 analyzer at Surya Hospital, Pune | Need to check all the tubes for any clots. | Found that all the tubes are blocked so cleaned all the tubes, along with the waste pump and valve, reconnected everything, installed both SP and SC, machine is working fine.
No | Visit | Complaint | RCHN | 报错1031 | 报错1031，加借903-364一个 | 更换903-364后修复 by jesson | SWO-351838 | Defect | HW | Process Module | Optical Unit | 1031 | 1031 | 1031, plus 903-364 one | Repair by jesson after replacing 903-364
No | Visit | Complaint | RCHN | 定标数据丢失 | 定标数据丢失，借944-212（14213缺货）944-258（14051）各一张 | 重新定标后修复 by jesson | SWO-351329 | User error | User error | user error | user error |  | Target data loss | Bid data loss, borrow 944-212 (14213 out of stock) 944-258 (14051) each one | Fix by jesson after re-calibration
No | Visit | Complaint | RAME | Receiving inhomogenous sample error | Received inhomogenous sample error | Reviewed logs, checked and cleaned measuring chambers, replaced ph $bg pump tubing, replaced ph electrode seal ring. After verifying proper operations on the analyzer, I ran 10 blood samples to show the inhomogenous sample error was cleared. | SWO-352237 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway | 0 | Receiving inhomogenous sample error | Received inhomogenous sample error | Reviewed logs, checked and cleaned measuring chambers, replaced ph $bg pump tubing, replaced ph electrode seal ring. After verifying proper operations on the analyzer, I ran 10 blood samples to show the inhomogenous sample error was cleared.
Yes | Visit | Complaint | RDAN | Pco2 fejl | Ventil tjek pH $BG modul - ok Tjekket inserts grænser - ok pH pakning tjekket - ok | Foreslog at genindlæse blå og grøn QC | SWO-352563 | Defect | HW | Basic Unit | Power Supply |  | PCO2 Error | Valve Check pH /BG module-OK Checked inserts boundaries-OK pH Packing checked-ok | Suggested reloading blue and green QC
Yes | Visit | Complaint | RIND | Electrolyte on parameter bar turns yellow. | Checked and found that the status is out of range | found that the ref electrode is faulty, Lot No- 70-07. Need to replace Reference electrode. | SWO-353683 | Defect | HW | Basic Unit | Power Supply |  | Electrolyte on parameter bar turns yellow. | Checked and found that the status is out of range | found that the ref electrode is faulty, Lot No- 70-07. Need to replace Reference electrode.
Yes | Visit | Complaint | RDEU | 15.05.2018, 09:20 Uhr, Sr. Sandra, Tel. 02037332328: Drucker defekt; schlechtes Druckbild | Gerät überprüft und gereinigt | Druckkopf erneuert, Gerät neu gestartet, Startphase abgewartet, Kals und QCs durchgeführt | SWO-352086 | Defect | HW | Basic Unit | Power Supply |  | 15.05.2018, 09:20, Sr. Sandra, tel. 02037332328: printer defective; poor print image | Device checked and cleaned | Print head renewed, device restarted, start-up phase awaited, Kals and QCs performed
Yes | 5 | Y | RCZE | Blank | Zlomeny pin pre nasunutie hadicky |  | D4550519B66066BF | Defect | HW | Basic Unit | Power Supply |  | Blank | Zlomeny pin pre nasunutie hadicky
No | Visit | Complaint | RDEU | 09.05.2018, 11:01 Uhr, H. Welle, Tel. 063120381271: QC Verletzung bei tHb, pO2 und Gluc. | Oxi Spectrum missmatch. taucht sporadisch bei Qc und Vollblutmessungen auf. Glu und Hb driften sehr häufig. | Glu und Lac Dichtungsringe fehlen. Dichtungsringe ersetzt. Hämolysator ausgetauscht. Lampenheit justiert. | SWO-352248 | Defect | HW | Oxi module | Hemolyzer | 0 | 09.05.2018, 11:01 Uhr, H. Welle, Tel. 063120381271: QC Verletzung bei tHb, pO2 und Gluc. | Oxi Spectrum missmatch. appears sporadically at Qc and thoroughbred measurements. Glu and Hb drift very frequently. | Glu and Lac sealing rings are missing. sealing rings. hemolyser. Lamp-light adjustment.
No | Visit | Complaint | RAME | There is a blue line running down the left side of the monitor from top to bottom. | None necessary | Replaced display monitor to fix blue line on the screen. | SWO-352383 | Defect | HW | DMS_PC | Display | 0 | There is a blue line running down the left side of the monitor from top to bottom. | None necessary | Replaced display monitor to fix blue line on the screen.
No | Visit | Complaint | RAME | Intermittently occurring error 355 temp error. | Reviewed data logs, confirmed error 355. Performed thermostatting diagnostics, no errors observed at this time. | Suspect intermittent measuring light $heat PCB, replaced PCB. Performed all applicable adjustments and tested, no errors observed at this time. Validated analyzer, Crea failing for QC and calibration. Re-membraned CREA electrodes and replaced Cal 1, 2 and cleaning solution to resolve. | SWO-352099 | Defect | HW | pH/BG Module | Illumination PCBA | 0 | Intermittently occurring error 355 temp error. | Reviewed data logs, confirmed error 355. Performed thermostatting diagnostics, no errors observed at this time. | Suspect intermittent measuring light $heat PCB, replaced PCB. Performed all applicable adjustments and tested, no errors observed at this time. Validated analyzer, Crea failing for QC and calibration. Re-membraned CREA electrodes and replaced Cal 1, 2 and cleaning solution to resolve.
No | Visit | Complaint | RIND | Analyzer not initializing at time when it showing loading data file. | Checked and found that the software got corrupted | tried to install Application Software, but it doesn't work. As the analyzer is with Windows7, so asked Mr. prakatheesh to send the Windows 7. | SWO-353684 | Revisit required | HW | DMS_PC | HDD | 0 | Analyzer not initializing at time when it showing loading data file. | Checked and found that the software got corrupted | tried to install Application Software, but it doesn't work. As the analyzer is with Windows7, so asked Mr. prakatheesh to send the Windows 7.
No | Visit | Complaint | RKK | エラー1278 CUP洗浄フロー確認が警告値を超えました が頻発 | エラー1278 CUP洗浄フロー確認が警告値を超えました が頻発 | エラー1278 CUP洗浄フロー確認が警告値を超えました が頻発 | SWO-351772 | Defect | HW | Wet section | Fluidic pathway | 1278 | Error 1278 CUP cleaning flow confirmation exceeded warning value, but frequently | Error 1278 CUP cleaning flow confirmation exceeded warning value, but frequently | Error 1278 CUP cleaning flow confirmation exceeded warning value, but frequently
No | Visit | Complaint | RAME | Error 1310 response error on multiple cassettes | Found lac response error, activity shows glu had response error 1310 | Replaced the sensor interface board | SWO-352757 | Defect | HW | Basic Unit | PCB, Sensor interface | Glu/Lac | Error 1310 response error on multiple cassettes | Found lac response error, activity shows glu had response error 1310 | Replaced the sensor interface board
No | Visit | Complaint | RLTD | David Simmonds Tel 0191 2829240 | Found analyser failing pressure test on vacuum value, also failing clot detect on SC and Oxi. Reseat inlet connector gasket (grey) as it was slightly proud, try a different inlet gasket, still not passing tests. Checked Oxi pinch valve and tube, both ok. Fitted a new SC and tests were still failing, and then swapped SP from working analyser to faulty one, still failing pressure test but clot detect tests now passing. Finally decide to fit a new inlet connector gasket (black) and all tests passed. | Fit new inlet connection gasket. Perform calibrations and QCs, left analyser conditioning as Glu sensitivity was low. | SWO-353136 | Defect | HW | Auto Inlet module | connector gasket | 0 | David Simmonds Tel 0191 2829240 | Found analyser failing pressure test on vacuum value, also failing clot detect on SC and Oxi. Reseat inlet connector gasket (grey) as it was slightly proud, try a different inlet gasket, still not passing tests. Checked Oxi pinch valve and tube, both ok. Fitted a new SC and tests were still failing, and then swapped SP from working analyser to faulty one, still failing pressure test but clot detect tests now passing. Finally decide to fit a new inlet connector gasket (black) and all tests passed. | Fit new inlet connection gasket. Perform calibrations and QCs, left analyser conditioning as Glu sensitivity was low.
No | Visit | Complaint | RLTD | Faye 02075804400 ex 10867 Silva 02075804400 ex 10867 | Intervention required message on screen 1216 lifetime in analyser exceeded for sensor cassette was firstly installed 23 $3 $18 at 10.16 | Removed cassette and installed cassette run 1214-277 this was installed 22 $4 $18 at 13.48 and will expire on the 23 $5 $18 cassette is working . It was showing CI calibration sensitivity errors but this working ok. The only parameter that are not working are glucose K,Ca and lactate due to sensitivity error probably because the cassette has been left out of the analyser to long. Customer will fit new cassette when they arrive. | SWO-352812 | User error | SC | SC | lifetime | 0 | Faye 02075804400 ex 10867 Silva 02075804400 ex 10867 | Intervention required message on screen 1216 lifetime in analyser exceeded for sensor cassette was firstly installed 23 /3 /18 at 10.16 | Removed cassette and installed cassette run 1214-277 this was installed 22 $4 $18 at 13.48 and will expire on the 23 $5 $18 cassette is working . It was showing CI calibration sensitivity errors but this working ok. The only parameter that are not working are glucose K,Ca and lactate due to sensitivity error probably because the cassette has been left out of the analyser to long. Customer will fit new cassette when they arrive.
No | Visit | Complaint | RCHN | 溶血器单元故障 | 溶血器单元故障，借902-980一个 | 902-980的管路接头断裂，更换902-980后修复 by jesson | SWO-351312 | Defect | HW | Oxi module | Hemolyzer | 0 | Hemolytic unit failure | Hemolytic unit failure, borrow 902-980 one | 902-980 pipe connector broken, replace 902-980 after repair by jesson
No | 3 | Y | RPAC | Blank | software upgrade to v1.2 FAN 915-380 |  | D45611922C625074 | [5 empty] | Blank | software upgrade to v1.2 FAN 915-380
No | 3 | Y | RPAC | Blank | software upgrade to v1.2 FAN915-380 |  | D4561A312C245081 | [5 empty] | Blank | software upgrade to v1.2 FAN915-380
No | 3 | Y | RPAC | Blank | Software upgrade to v1.2 FAN915-380 |  | D4561EF82C2B5089 | [5 empty] | Blank | Software upgrade to v1.2 FAN915-380
No | 3 | Y | RPAC | Blank | software upgrade to v1.2 FAN915-380 |  | D456217C2C60508D | [5 empty] | Blank | software upgrade to v1.2 FAN915-380
No | 3 | Y | RPAC | Blank | software upgrade to v1.2 FAN915-380 |  | D45617682C20507D | [5 empty] | Blank | software upgrade to v1.2 FAN915-380
No | 3 | Y | RPAC | Blank | software upgrade to v1.2 FAN915-380 |  | D456153A2C155079 | [5 empty] | Blank | software upgrade to v1.2 FAN915-380
No | 3 | Y | RPAC | Blank | Software upgrade to v1.2 FAN915-380 |  | D4561C892C135085 | [5 empty] | Blank | Software upgrade to v1.2 FAN915-380
Yes | Visit | Complaint | RCHN | 显示屏花屏 | 显示屏花屏 借636-571一个 | 更换636-571后修复 by jesson | SWO-351310 | Defect | HW | Basic Unit | Power Supply |  | Display Flower Screen | Display Flower Screen Borrowing 636-571 One | Repair by jesson after replacement 636-571
No | Visit | Complaint | RCHN | 报错1031 | 报错1031 借903-364一个 | 更换903-364后修复 by jesson | SWO-351939 | Defect | HW | Process Module | Optical Unit | 1031 | 1031 | False 1031 Borrow 903-364 one | Repair by jesson after replacing 903-364
Yes | Visit | Complaint | RCHN | 系统软件定标数据丢失 | 系统软件定标数据丢失，借944-212（14251）944-248（14249）944-214（14220）缺货 | 重新定标后修复 by jesson | SWO-351889 | Defect | HW | Basic Unit | Power Supply |  | System software calibration data loss | System software calibration data loss, 944-212 (14251)944-248 (14249)944-214 (14220) out of stock | Fix by jesson after re-calibration
No | Visit | Complaint | RAME | Crea and liquid sensor errors | Crea errors no longer present when FSE arrived. Liquid sensor error present. | Upon investigation it was discovered that waste tube 2 was on air vent 3. Tube was moved to the proper position and normal operations were restored. | SWO-352926 | User error | User error | user error | user error | interchanging of tubes | Crea and liquid sensor errors | Crea errors no longer present when FSE arrived. Liquid sensor error present. | Upon investigation it was discovered that waste tube 2 was on air vent 3. Tube was moved to the proper position and normal operations were restored.
No | Visit | Complaint | RAME | Lac being changed about every 2 weeks would like fse to visit. Contact: Mike Reay - 360-493-5181 | Lactate has been requiring membrane change approximately every two weeks. New electrode and different membrane lot did not change performance. Interrogation of data logs show precipitous drop of sensitivity at 11-14 days and L-3 QC drops out of range. | Left current electrode and membrane in place that was installed 5-7. Replaced EL $MET PCB, loaded SW, calibrated and ran QC. Sensitivity came in nominally higher than prior calibrations and QC results were on or near mean. | SWO-352839 | Defect | HW | El/Met Module | PCBA | Kapton | Lac being changed about every 2 weeks would like fse to visit. Contact: Mike Reay - 360-493-5181 | Lactate has been requiring membrane change approximately every two weeks. New electrode and different membrane lot did not change performance. Interrogation of data logs show precipitous drop of sensitivity at 11-14 days and L-3 QC drops out of range. | Left current electrode and membrane in place that was installed 5-7. Replaced EL $MET PCB, loaded SW, calibrated and ran QC. Sensitivity came in nominally higher than prior calibrations and QC results were on or near mean.
No | Visit | Complaint | RDEU | 16.05.2018, 10:36 Uhr, Fr. Neiswirth, Tel. 089857911451: Bruchstelle zwischen pH/BG- und Lyte Modul | ABL nicht messbereit | Fehlerbehebung erfolgt im Zuge der Halbjahreswartung, System- und Funktionskontrolle in Ordnung | SWO-352802 | Defect | HW | Basic Unit | pH/BG Module | 0 | 16.05.2018, 10:36 a.m., Ms. Neiswirth, tel. 089857911451: breaking point between pH/BG and Lyte module | ABL not ready for measurement | Troubleshooting takes place in the course of half-year maintenance, system and function control in order
No | Visit | Complaint | RIND | Lact Query & Cl Level 1 & 2 QC was getting failed | Checked the data logs Lact calibration sensitivity was showing high cleaned lact electrode & chamber got stable but when run QC cl was getting failed level 1 was getting low value & level 2 was getting high was replaced new electrode & checked then also same problem replaced new Cl membrane & checked then also QC was getting failed level 1 was low & level 2 was high cross checked reference electrode & checked then also same problem crosschecked cl electrode from working machine then it was passing found problem membrane | Replaced new lot cl membrane & performed QC level 1 & 2 values was within the range | SWO-353910 | Assisted Replacement | Electrode | Membrane | Cl | 0 | Lact Query & Cl Level 1 & 2 QC was getting failed | Checked the data logs Lact calibration sensitivity was showing high cleaned lact electrode & chamber got stable but when run QC cl was getting failed level 1 was getting low value & level 2 was getting high was replaced new electrode & checked then also same problem replaced new Cl membrane & checked then also QC was getting failed level 1 was low & level 2 was high cross checked reference electrode & checked then also same problem crosschecked cl electrode from working machine then it was passing found problem membrane | Replaced new lot cl membrane & performed QC level 1 & 2 values was within the range
No | Visit | Complaint | RDEU | QVC Grenzverletzung bei CRP in Level gelb und rot. | Frau Schweger; Tel.: 061529862440 Systemreinigung durchgeführt, Mehrmals QC Level rot und blau gemessen. Immer QC Grenzverletzung. Messung leicht oberhalb der Grenze. Vmtl. keine Verschmutzung, da AQT erst seit April in Betrieb. | Keine neue QC CRP Charge vorhanden, bestellt, aber nicht geliefert, auch kein neues Testkit CRP vorhanden. FST Einsatz notwendig. | SWO-352782 | Revisit required | Consumables | QC | CRP |  | QVC Grenzverletzung bei CRP in Level gelb und rot. | Mrs Schweger; Phone: 061529862440 System cleaning performed, measured multiple QC levels red and blue. Always QC border violation. Measurement slightly above the limit. Vmtl. no pollution, since AQT has only been in operation since April. | No new QC CRP batch available, ordered but not delivered, also no new test kit CRP available. FST use necessary.
No | 12 | Y | RCZE | Blank |  |  | D458AB9718560A70 | Defect | HW | Basic Unit | Power Supply |  | Blank
No | Visit | Complaint | RDEU | Nur mit mechanischer Hilfe zu schließen | Sc lid closed indicator gebrochen | 924-609 getauscht in betrieb genommen und Deckel mehrfach auf Funktion überprüft. | SWO-353590 | Defect | HW | Basic Unit | detector plate | 0 | Only to close with mechanical help | Sc lid closed indicator broken | 924-609 replaced in operation and lid checked several times for function.
No | Visit | Complaint | RAME | If draw time is missing seconds on the time stamps result wont cross over. | Checked the data logs. The seconds needed to be removed from the time set up. | Went into windows and made the correction too the time. | SWO-353493 | Setup | SW | Windows | time | 0 | If draw time is missing seconds on the time stamps result wont cross over. | Checked the data logs. The seconds needed to be removed from the time set up. | Went into windows and made the correction too the time.
No | Visit | Complaint | RAME | Ph drifting out on QC B | Ph drifting out on QC B, sometimes you can rerun QC’s and it will come back in | Replaced SP, ran QC’s, all back in green. Will return to replace the sensor interface PCB.. Returned and replaced sensor interface PCB, ran cals and QC all green | SWO-352863 | Defect | HW | Basic Unit | PCB, Sensor interface | 0 | Ph drifting out on QC B | Ph drifting out on QC B, sometimes you can rerun QC’s and it will come back in | Replaced SP, ran QC’s, all back in green. Will return to replace the sensor interface PCB.. Returned and replaced sensor interface PCB, ran cals and QC all green
No | Visit | Complaint | RLTD | Christina or Joanne 01312422542 | Removed consumables and found SP had red dye on pump tubing | Replaced SP with new and replaced Inlet Gasket with new at request of customer | SWO-353198 | Assisted Replacement | SP | SP | Pump tubes | 0 | Christina or Joanne 01312422542 | Removed consumables and found SP had red dye on pump tubing | Replaced SP with new and replaced Inlet Gasket with new at request of customer
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | Lac定标不过 1310 | 更换测试卡后修复 | SWO-353061 | Defect | HW | Basic Unit | Power Supply |  | Test card, reagent package failure | Lac calibration is no more than 1310 | Repair after replacing test card
No | 6 | Y | RPAC | Blank | software upgrade to v1.2 FAN915-380 |  | D4575F4CA231509A | [5 empty] | Blank | software upgrade to v1.2 FAN915-380
No | 6 | Y | RPAC | Blank | software upgrade to v1.2 FAN915-380 |  | D4574FC9A2235096 | [5 empty] | Blank | software upgrade to v1.2 FAN915-380
No | 2 | Y | RPAC | Blank | software upgrade to v1.2 FAN915-380 |  | D457641AA24C509E | [5 empty] | Blank | software upgrade to v1.2 FAN915-380
No | Visit | Complaint | RCHN | PCT定标错误 | PCT定标错误 944-394（14071）一个 | 重新定标后修复 by jesson | SWO-351852 | Assisted Replacement | Consumables | Test Kit | Cal cartridge, PCT |  | PCT calibration error | PCT calibration error 944-394 (14071) one | Fix by jesson after re-calibration
Yes | Visit | Complaint | RCHN | 报错1029 | 报错1029 | 更换903-364后修复 by jesson | SWO-353651 | Defect | HW | Basic Unit | Power Supply |  | 1029 error | 1029 error | Repair by jesson after replacing 903-364
No | Visit | Complaint | RCHN | 报错1031 | 报错1031 借903-364一个 | 更换903-364后修复 by jesson | SWO-351940 | Defect | HW | Process Module | Optical Unit | 1031 | 1031 | False 1031 Borrow 903-364 one | Repair by jesson after replacing 903-364
No | Visit | Complaint | RDEU | 09.05.2018 14:15Uhr Frau Fleischer 0017616658311: Alle zwei Tage muss ein neuer Reagenzienpack installiert werden. Nach jeder Neuinstallation muss das Gerät wegen Fehler neu gebootet werden. | verschiedene Fehler sind in den vergangenen Wochen aufgetreten 3x 5684 Rotateproccesswheel 40 1x 4726 Openstorage 24 1x 2877 cartridge 40 | Innenraum gesäubert insbesondere Proccesswheel - Test i.O. anschließend Halbjahreswartung durchgeführt | SWO-353454 | Cleaned | HW | Process Module | Process Wheel | 5684 | 09.05.2018 14:15Pm Mrs Fleischer 0017616658311: A new reagent pack must be installed every two days. After each reinstallation, the device must be rebooted due to errors. | various errors have occurred in recent weeks 3x 5684 Rotateproccesswheel 40 1x 4726 Openstorage 24 1x 2877 cartridge 40 | Interior cleaned in particular Proccesswheel - Test i.O. subsequently carried out half-year maintenance
No | Visit | Complaint | RDEU | QVC Grenzverletzung bei CRP in Level gelb und rot. | LQC Parameter CRP Oberer Grenzwert in Level 1 und 2 überschritten D-dimer zielstrebig und TNT okay, aber unter Zielwert | Allgemeine Überprüfung - Abfallstrecke mit Hypochlorite Solution gespült - Erneuerung Einlassnadel etc. - Funktionsüberprüfung (keine Fehlerbeseitigung) - Service Dump gezogen und nach Zentrale verschickt - Information Labor und Frau Löwenkamp (Radiometer) Klärung folgt ! Langzeit-Abschaltung durchgeführt - Neustart - Labor misst LQC (CRP) nochmals nach ! | SWO-353550 | Defect | Consumables | QC | CRP |  | QVC Grenzverletzung bei CRP in Level gelb und rot. | LQC parameter CRP Upper limit exceeded in levels 1 and 2 D-dimer purposeful and TNT okay, but below target | General inspection - Waste section with hypochlorite solution flushed - Renewal inlet needle etc. - Functional check (no error removal) - Service dump pulled and sent to headquarters - Information Laboratory and Mrs. Löwenkamp (radiometer) Clarification follows ! Long-term shutdown performed - Restart - Laboratory measures LQC (CRP) again !
No | Visit | Complaint | RKK | ﾕｰｻﾞがSC3つ交換したが、Glu,Lacの初期化が通らない。 また、ｵｷｼｺﾝﾌﾟﾘｰﾄﾃｽﾄにてﾍﾓﾗｲｻﾞｰ項目がﾌｪｰﾙ。 | 代替機設置 | 代替機設置 | SWO-354438 | Revisit required | 0 | 0 | analyzer replaced | Glu/Lac | The user replaced three SC,, glu, lac initialization does not pass. In addition, the hemolizer item fails in the oxycomplete test. | Installation of alternative machines | Installation of alternative machines
No | Visit | Complaint | RFRA | Erreur Calibration Glucose et Lactate persistante | Pbl de calibration pour le lactate | Changement de la carte , changement de la cassette ok | SWO-356038 | Assisted Replacement | SC | SC | Glu/Lac | 0 | Glucose And Persistent Lactate Calibration Error | Pbl calibration for lactate | Change of map, change of cassette ok
No | 7 | Y | RPAC | Blank | Aborted the sample intermittently |  | D4589A0318090389 | Defect | HW | Auto Inlet module | Auto Inlet module | 0 | Blank | Aborted the sample intermittently
No | 4 | Y | RPOL | Blank | Wymieniono pogiêt¹ przez u¿ytkonika ig³ê, uszczelkê wlotu. Testy serwisowe. Aparat sprawny. |  | D45883BF1863E612 | Defect | HW | Auto Inlet module | Probe bend | 0 | Blank | They are replaced by a needle, a seal in the inlet. Service tests. The camera is functional.
No | Visit | Complaint | RDEU | 11.05.2018, Frau Sellmann MT, tel. 0401818828816: LP lässt sich nicht richtig entfernen, bereits mehrfach konnte der Anwender das LP nur unter Neustart und starkem rukeln am LP entfernen. Es gab dann ein lautes unangenehmes Geräusch | LP lässt sich nur schwierig entfernen, verklemmt sich und macht laute Geräusche | Fehleranalyse, Stepper Motor in Actuator Unit inkl. Zugstange und waste con. Unit erneuert, Lösungspack mehrfach wiederholt Aus und wieder eingesetzt i. O. , Kalibrierung und Qualitätskontrolle i. O. | SWO-353893 | Defect | HW | Basic Unit | Actuator motor | 0 | 11.05.2018, Mrs. Sellmann MT, tel. 040181888816: LP cannot be removed properly, already several times the user could only remove the LP under restart and strong rusting on the LP. There was then a loud unpleasant noise | LP is difficult to remove, gets stuck and makes loud noises | Fault analysis, Stepper motor in Actuator Unit incl. drawbar and waste con. Unit renewed, solution pack repeatedly re-inserted i.o., calibration and quality control in the
No | Visit | Complaint | RDEU | Patientenbarcode kann nicht gelesen werden. | Patientenbarcode kann nicht gelesen werden. | Eintrag in roma.ini ergänzt. Danach hat der Scan funktioniert. Kalibrierung durchgeführt. QCs für,alle Level durchgeführt. ABL wieder messbereit. | SWO-353900 | Setup | SW | Roma.ini | Barcode | 0 | Patient bar code cannot be read. | Patient bar code cannot be read. | Entry in roma.ini. After that, the scan worked. calibration performed. QCs for,all levels performed. ABL ready to be measured again.
No | Visit | Complaint | RAME | Customer keeps getting replace inlet gasket. |  |  | SWO-354098 | Defect | HW | Auto Inlet module | Inlet gasket | 0 | Customer keeps getting replace inlet gasket.
Yes | Visit | Complaint | RCRO | CCN/PER | Error code 964 Leak current in relation to solution pack detected | All parameters | SWO-354524 | Defect | HW | Basic Unit | Power Supply |  | CCN/PER | Error code 964 Leak current in relation to solution pack detected | All parameters
No | Visit | Complaint | RIND | SC not calibrating and Solution not coming. | SC not calibrating and Solution not coming. | SP cross checked for Solution not coming but same problem. Analyser Opened and tubings cleaned and flushed. Againg try to calibrate but same problem coming. Cross check with Valve set for another analyser and found working well. Need to replace Valve set complete. Sending TO and PER for that. | SWO-354209 | [5 empty] | SC not calibrating and Solution not coming. | SC not calibrating and Solution not coming. | SP cross checked for Solution not coming but same problem. Analyser Opened and tubings cleaned and flushed. Againg try to calibrate but same problem coming. Cross check with Valve set for another analyser and found working well. Need to replace Valve set complete. Sending TO and PER for that.
No | Visit | Complaint | RDEU | sO2 besteht QC Messung nicht | Geräterücklieferung zum Kunden | Gerät geliefert. QA Zyklus wird von Frau Meixner durchgeführt. Reparaturersatzgerät abgeholt. | SWO-354151 | [5 empty] | sO2 does NOT exist QC measurement | Device return to customer | device. QA cycle is carried out by Ms. Meixner. repair replacement device.
No | Visit | Complaint | RCHN | 无法开机 | 无法开机 借903-663一个 | 用库存机上的903-663修复 by jesson | SWO-353054 | [5 empty] | Cannot power on | Can't turn on, borrow 903-663 one. | Fix by jesson with 903-663 on the stock machine
No | Visit | Complaint | RAME | Error 2108 – “Possible fluid pathway blockage” | Error 2108 – “Possible fluid pathway blockage”. Took analyzer apart and found liquid sensor had been leaked on by cut in tubing. | Replace liquid sensor. Cleaned and lubed pinch valve. Replaced parts from PM kit. | SWO-354135 | [5 empty] | Error 2108 – “Possible fluid pathway blockage” | Error 2108 – “Possible fluid pathway blockage”. Took analyzer apart and found liquid sensor had been leaked on by cut in tubing. | Replace liquid sensor. Cleaned and lubed pinch valve. Replaced parts from PM kit.
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | LAC,GLU定标不过 | 更换测试卡后修复 | SWO-353682 | [5 empty] | Test card, reagent package failure | LAC, GLU calibration but | Repair after replacing test card
No | Visit | Complaint | RCHN | 报错5658 | 报错5658 借944-216（14099）一张 | 重新定标正常。by jesson | SWO-353769 | Cleaned | HW | Process Module | Incubator wheel | 5658 | 5658 error | 5658 Loan 944-216 (14099) one | Recalibration is normal. by jesson
No | Visit | Complaint | RCHN | 报错1024 | 报错1024,借944-394一张（13955） | 重新定标正常。by jesson | SWO-353115 | Cleaned | HW | Process Module | Process Wheel | 1024 | 1024 error | 1024, loan 944-394 (13955) | Recalibration is normal. by jesson
Yes | Visit | Complaint | RIND | pH sample measurement error. | All the sample measurement query for pH results. | Found pH and Na sample measurement error for results, replaced the Na membrane, cleaned the pH electrode and measuring chamber calibrated found both parameter results are OK, need to be performed the sample measurement and check, it’s under observation. | SWO-353887 | [5 empty] | pH sample measurement error. | All the sample measurement query for pH results. | Found pH and Na sample measurement error for results, replaced the Na membrane, cleaned the pH electrode and measuring chamber calibrated found both parameter results are OK, need to be performed the sample measurement and check, it’s under observation.
No | Visit | Complaint | RIND | Gasket is leaking | Gasket is leaking | Removed the gasket dipped it in hypo for half an hour then cleaned the sample inlet path it was obstacled by cotton peice so cleaned it then given pump calibartion and liquid sensor calibartion , after that tested the sample its working fine | SWO-354071 | Cleaned | Fluidic pathway | Inlet module | Inlet gasket |  | Gasket is leaking | Gasket is leaking | Removed the gasket dipped it in hypo for half an hour then cleaned the sample inlet path it was obstacled by cotton peice so cleaned it then given pump calibartion and liquid sensor calibartion , after that tested the sample its working fine
No | Visit | Complaint | RDAN | 17-05-2018 Hej Rasmus Som du kan se herunder har vi tidligere haft udfordringer med den ABL der står i gymnastiksalen, Nørre Alle 51. Vi klarede det med en genstart men nu er der desværre dukket et tilsvarende problem op. Der blev meldt om printerfejl og her til morgen var der sort skærm. Vi prøvede med en windows genstart som igen har løst problemet men der er jo tilsyneladende et problem som gerne skulle fjernes. Håbe du kan hjælpe med det 18-04-2018 Hej Rasmus Vi har lidt problemer med ABL på Nørre Alle. Den kommer en gang imellem med et vindue ”Crystal Reports: Database Error” ODC error [Microsoft][ODBC Text Driver] Cannot update. Database or object is read only. Man skal manuelt bede om en udskrift og printeren skriver langsomt og hakkende ud. Ligeledes har vi oplevet at skærmbillede med resultater ser underligt ud. Nu er skærmen også gået helt i stå – det er maskinen dog ikke Håber du kan hjælpe med det Jens Jung Nielsen PhD AC-TAP Institut for Idræt og Ernæring Integrativ Fysiologi Københavns Universitet Universitetsparken 13 2100 København Ø Tlf 35321574 Mobil 93565318 jjnielsen@nexs.ku.dk | I loggen er der registret printerfejl flere gange: 14-03-2018 21-03-2018 29-03-2018 15-05-2018 17-05-2018 | Jeg har skiftet printeren og den er testet og fundet ok. Membran for Ca er skiftet. Der er kørt 2 punkts kal. Der er kørt tHb kal. Der er kørt QCer. Alt er fundet OK. | SWO-353951 | Defect | HW | DMS_PC | Printer Module |  | 17-05-2018 Hi Rasmus As you can see below, we have previously had challenges with the ABL that is in the gym, Nørre Alle 51. We did it with a reboot but now, unfortunately, a similar problem has emerged. The printer error was reported and this morning there was a black screen. We tried with a Windows reboot which again has solved the problem but apparently there is a problem that would have to be removed. Hope you can help with it 18-04-2018 Hi Rasmus We have some problems with ABL at Nørre Alle. It comes once in a while with a window "Crystal Reports: Database Error" ODC error [Microsoft] [ODBC Text Driver] Cannot update. Database or object is read only. You have to manually ask for a printout and the printer writes slowly and chosh out. Likewise, we have experienced that screen shot with results look strange. Now the screen has also gone to a standstill – but the machine is not Hope you can help with it Jens Jung Nielsen PhD AC-TAP Department of Sport and Nutrition Integrative Physiology University University Park 13 2100 Copenhagen Ø Tel 35321574 Mobil 93565318 jjnielsen@nexs.ku.dk | In the log is where the registry Printerfejl several walking: 14-03-2018 21-03-2018 29-03-2018 15-05-2018 17-05-2018 | I changed the printer and it has been tested and found OK. Membrane for Ca is switched. Two-point Kal has been run. A tHb Kal has been run. QCer has been run. Everything is found OK.
No | Visit | Complaint | RAME | Electrolytes drifting after nightly cleaning | Electrolytes drift after nightly cleaning cycle. | Replaced reference electrode per Service Bulletin. | SWO-354095 | Assisted Replacement | Electrode | Electrode | Reference | CAPA-00733 | Electrolytes drifting after nightly cleaning | Electrolytes drift after nightly cleaning cycle. | Replaced reference electrode per Service Bulletin.
No | Visit | Complaint | RAME | Electrolytes drifting after nightly cleaning | Electrolytes drifting after nightly cleaning. | Replaced reference electrode per Service Bulletin. | SWO-354096 | Assisted Replacement | Electrode | Electrode | Reference | CAPA-00733 | Electrolytes drifting after nightly cleaning | Electrolytes drifting after nightly cleaning. | Replaced reference electrode per Service Bulletin.
No | Visit | Complaint | RIND | Analyser Calibrating again and again and Na and Cl value coming high for all readings. | Analyser Calibrating again and again and Na and Cl value coming high for all readings. | Inlet probe and Gasket cleaned. Electrode chambers cleaned. Electrode cleaning procedure done. Tubing flushed and cleaned properly. Analyser Cal setup done. Analyser Serviced and calibrated properly. Handed over in a fine working condition. | SWO-354207 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | Analyser Calibrating again and again and Na and Cl value coming high for all readings. | Analyser Calibrating again and again and Na and Cl value coming high for all readings. | Inlet probe and Gasket cleaned. Electrode chambers cleaned. Electrode cleaning procedure done. Tubing flushed and cleaned properly. Analyser Cal setup done. Analyser Serviced and calibrated properly. Handed over in a fine working condition.
No | Visit | Complaint | RAME | High Crea Check borderline failing high | Customer reported High Crea Check borderline failing high. Other 2 instruments they were instructed to replace the electrodes. | No Corrective action yet. Reached out to team for next step and collected data files and service dump. | SWO-354128 | Revisit required | Electrode | Electrode | crea | HCC | High Crea Check borderline failing high | Customer reported High Crea Check borderline failing high. Other 2 instruments they were instructed to replace the electrodes. | No Corrective action yet. Reached out to team for next step and collected data files and service dump.
No | Visit | Complaint | RAME | Electrolytes drifting after nightly cleaning | Electrolytes drifting after nightly cleaning. | Replaced reference electrode per Service Bulletin. | SWO-354097 | Assisted Replacement | Electrode | Electrode | Reference | CAPA-00733 | Electrolytes drifting after nightly cleaning | Electrolytes drifting after nightly cleaning. | Replaced reference electrode per Service Bulletin.
No | Visit | Complaint | RIND | Analyzer not initializing at time when it showing loading data file. | Software corrupted | Removed ABL800 application and reinstalled SW v6.17. Software installation completed still ABL800 is not initializing. | SWO-354192 | Revisit required | HW | DMS_PC | HDD |  | Analyzer not initializing at time when it showing loading data file. | Software corrupted | Removed ABL800 application and reinstalled SW v6.17. Software installation completed still ABL800 is not initializing.
No | Visit | Complaint | RDEU | AQT bricht sporadisch Proben ab und will neu gestartet werden Fehler 3998 Lock_24 | AQT macht Fehler erneut. | AQT zur Reparatur abgeholt. | SWO-354149 | Revisit required | HW | Storage Module | rubber disk | 3998 | AQT sporadically cancels samples and wants to be restarted Error 3998 Lock_24 | AQT makes mistakes again. | AQT for repair.
No | Visit | Complaint | RAME | Prep Radiometer owned system for demo & Inspect oxi hardware errors 1016, 1018, 1019 for Nathan Dacuycuy | Inspection of the oxi system revealed several internal leaks at and around the hemolyzer. Hemolyzer failing complete oxi test along with spectrometer. | Replaced hemolyzer | SWO-353574 | Defect | HW | Oxi module | Hemolyzer | leak | Prep Radiometer owned system for demo & Inspect oxi hardware errors 1016, 1018, 1019 for Nathan Dacuycuy | Inspection of the oxi system revealed several internal leaks at and around the hemolyzer. Hemolyzer failing complete oxi test along with spectrometer. | Replaced hemolyzer
No | Visit | Complaint | RCHN | LAC定标不过 | LAC定标不过 借903-581 | 更换903-581一个正常。by jesson | SWO-353707 | Defect | HW | Basic Unit | PCB, Sensor interface | Lac | LAC calibration is not | LAC calibration but borrow 903-581 | Replace the 903-581 one normal. by jesson
No | Visit | Complaint | RCHN | GLU,LAC定标失败 | GLU,LAC定标失败，更换两张测试卡还是不行。借903-581一个。 |  | SWO-354205 | Defect | HW | Basic Unit | PCB, Sensor interface | Glu/Lac | GLU, LAC calibration failed | GLU, LAC calibration failed, replacing two test cards is still not possible. Borrow 903-581 one.
No | Visit | Complaint | RCHN | BNP传感器线断裂 1235 | BNP传感器线断裂 1235，借905-908一个。 | 更换905-908一个正常。by jesson | SWO-354176 | Defect | HW | Process Module | Sensor, BnP | 1235 | BNP sensor line break 1235 | BNP sensor line break 1235, borrow 905-908 one. | Replace 905-908 one normal. by jesson
No | 27 | Y | RPOL | Blank | Wymieniono uszczelkê wlotu a nastêpnie kasetkê sensorow¹. Testy aparatu. Analizator sprawny technicznie. |  | D45C73AF7A472FB1 | Defect | HW | Inlet module | Inlet gasket |  | Blank | The inlet seal and then the sensor cassette are listed. Camera tests. Technically efficient analyzer.
Yes | 6 | Y | RPAC | Blank | Moved analyser from old maternity ward to new Maternity ward |  | D45C9AB37A089D44 | [5 empty] | Blank | Moved analyser from old maternity ward to new Maternity ward
Yes | 6 | Y | RPAC | Blank | Transfer from Maternity to |  | D45C99A17A479D43 | [5 empty] | Blank | Transfer from Maternity to
No | Visit | Complaint | RLTD | Judith Gibbons on 0191 2829240 | Site visit to investigate and repair analyser vacuum pressure fail. | Arrived on site and powered down analyser. Removed wet section from analyser and inspected tubings on fluid path. Found signs of leaked fluid under pinch tube. Fitted new waste and pinch tube. When inspecting the removed pinch tube I found a pinhole in the side wall where valve action takes place. Refitted all parts and powered up the analyser. Analyser started 15 minute startup. Startup completed. Ran calibration and qcs to Confirm analyser functionality. Analyser passed all well. No further action required. | SWO-354564 | Defect | HW | Oxi module | Pinch valve tube |  | Judith Gibbons on 0191 2829240 | Site visit to investigate and repair analyser vacuum pressure fail. | Arrived on site and powered down analyser. Removed wet section from analyser and inspected tubings on fluid path. Found signs of leaked fluid under pinch tube. Fitted new waste and pinch tube. When inspecting the removed pinch tube I found a pinhole in the side wall where valve action takes place. Refitted all parts and powered up the analyser. Analyser started 15 minute startup. Startup completed. Ran calibration and qcs to Confirm analyser functionality. Analyser passed all well. No further action required.
No | Visit | Complaint | RCHN | O2漂移超范围 | O2漂移超范围，借902-972,903-334各一个 |  | SWO-354501 | [5 empty] | O2 Drift Over range | O2 Drift Overrange, borrow 902-972, 903-334 each
No | Visit | Complaint | RAME | Customer states that solution packs are having to be changed out ~ every 2 weeks and they are having to order additional solution packs over their PCA shipments. | Checked settings in analyzer, all set correctly. Downloaded logs & DB in case of CQC submission becomes necessary. | Reset database. Transferred settings from 315380 analyzer. Replaced solution pack due to database reset. | SWO-354644 | [5 empty] | Customer states that solution packs are having to be changed out ~ every 2 weeks and they are having to order additional solution packs over their PCA shipments. | Checked settings in analyzer, all set correctly. Downloaded logs & DB in case of CQC submission becomes necessary. | Reset database. Transferred settings from 315380 analyzer. Replaced solution pack due to database reset.
No | Visit | Complaint | RAME | Customer states that solution packs are having to be changed out ~ every 2 weeks and they are having to order additional solution packs over their PCA shipments. | Checked settings in analyzer, all set correctly. Downloaded logs & DB in case of CQC submission becomes necessary. | Reset database. Manually set up analyzer. Replaced solution pack due to database reset. | SWO-354608 | [5 empty] | Customer states that solution packs are having to be changed out ~ every 2 weeks and they are having to order additional solution packs over their PCA shipments. | Checked settings in analyzer, all set correctly. Downloaded logs & DB in case of CQC submission becomes necessary. | Reset database. Manually set up analyzer. Replaced solution pack due to database reset.
No | Visit | Complaint | RIND | SC not installing Leak found. | Sensor Board to replace. | Sensor Board replaced. Zeroing calibration done. Analyser Serviced and calibrated properly. Handed over in a fine working condition. | SWO-355722 | [5 empty] | SC not installing Leak found. | Sensor Board to replace. | Sensor Board replaced. Zeroing calibration done. Analyser Serviced and calibrated properly. Handed over in a fine working condition.
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 除了po2其他参数都不过 | 更换测试卡后修复 | SWO-356658 | Assisted Replacement | SC | SC | pO2 |  | Test card, reagent package failure | Except for po2 other parameters are no more than | Repair after replacing test card
No | Visit | Complaint | RESP | Error de fuga detectada | fuga detectada. Tubo de la válvula picado. | Se sustituyo tubo de válvula, aguja, tubo del desperdicio y tubo de entrada. Finalmente se sustituyo SCI y se comprobó correcto funcionamiento mediante QC. | SWO-354552 | Defect | HW | Oxi module | Pinch valve tube |  | Leak error detected | escape detected. Chopped valve tube. | Valve tube, needle, waste tube and inlet tube are replaced. SCI was finally replaced and properly functioned using QC.
Yes | Visit | Complaint | RCHN | 报错 1031 | 报错 1031，借903-364一个 |  | SWO-354422 | [5 empty] | Error 1031 | Error 1031, loan 903-364 one
No | Visit | Complaint | RCHN | 报错5297 | 报错5297，借903-371一个。 |  | SWO-354481 | Defect | HW | Process Module | Punch Up Unit | 5297 | 5297 error | 5297, loan 903-371 one.
No | Visit | Complaint | RCHN | 定标终止 | 定标终止，借944-214（14220）一张 |  | SWO-354430 | Assisted Replacement | Consumables | Test Kit | Cal cartridge, Myo |  | Bid termination | Bid termination, loan 944-214 (14220) one
Yes | Visit | Complaint | RKK | Neon number too high頻発 | Neon number too high頻発 | OXIモジュール引き上げ。代替えモジュール設置。 | SWO-354814 | [5 empty] | Neon number too high頻発 | Neon number too high頻発 | OXI module raise. Installed an alternative module.
No | Visit | Complaint | RIND | nothing is appearing on display, it is blank. | by replacing CPU it is found that the CPU unit is faulty. | Fixed the problem by placing CPU from demo analyzer, Need to replace the CPU unit. | SWO-355001 | Defect | HW | DMS_PC | CPU |  | nothing is appearing on display, it is blank. | by replacing CPU it is found that the CPU unit is faulty. | Fixed the problem by placing CPU from demo analyzer, Need to replace the CPU unit.
Yes | Visit | Complaint | RIND | Option key installation & Query in the patient result. | Check the activity log & patient result log & found query in the patient result & in the option menu found Demo key will be expired. | option key installed successfully.Clean the electrode chamber.Clean the electrode module pump tubes,then run few samples & got the desired result,handover the instrument to the customer. | SWO-357305 | Re classified RMED | Option Key |  |  |  | Option key installation & Query in the patient result. | Check the activity log & patient result log & found query in the patient result & in the option menu found Demo key will be expired. | option key installed successfully. Clean the electrode chamber. Clean the electrode module pump tubes,then run few samples & got the desired result,handover the instrument to the customer.
No | Visit | Complaint | RIND | Service Visit. | Need to observe the analyzer. | Check the analyzer pro-actively, no sample aborted during the observation days, no aborted sample, service issue has resolved, analyzer handed over to staff in well working condition. | SWO-356971 | No fault found | No fault found | No fault found | No fault found |  | Service Visit. | Need to observe the analyzer. | Check the analyzer pro-actively, no sample aborted during the observation days, no aborted sample, service issue has resolved, analyzer handed over to staff in well working condition.
No | Visit | Complaint | RAME | Analyzer does not read the smart chip on the sensor cassette. | Checked the data logs and verified the problem. | Replaced the sensor interface board. Set the IP address to work with customers network. | SWO-355308 | Defect | HW | Basic Unit | PCB, Sensor interface |  | Analyzer does not read the smart chip on the sensor cassette. | Checked the data logs and verified the problem. | Replaced the sensor interface board. Set the IP address to work with customers network.
No | Visit | Complaint | RDEU | Erst nach Neustart druckt der Drucker wieder, bereits bei der Installation gab es dieses Problem | BIOS-Einstellung für USB von highspeed auf eine Stufe niedriger gestellt, danach lief das Gerät vermutlich 10 Tage störungsfrei, so die Auskunft der MTA | Ursprüngliches PC-Modul wieder eingebaut, BIOS bearbeitet, 3 Kal durchgeführt, Kal-Fehler Glu Status außer Bereich, Glu abgemeldet, Meldung an die Med.-Technik zwecks remote support Labor ab Mittag unbesetzt | SWO-355117 | Defect | SW | Windows | Printer Module |  | only after reboot the printer prints again, this problem was already in place during installation | BIOS setting for USB from highspeed to a lower level, after which the device probably ran without interference for 10 days, according to the MTA | Original PC module reinstalled, BIOS edited, 3 cal performed, Kal error Glu status out of range, Glu logged out, message to Med. technology for remote support Laboratory vacant from noon
No | Visit | Complaint | RDEU | Fehlerbehebung erforderlich | WLAN-Verbindung funktioniert noch nicht, trotz Win7-Upgrade ==> Einstellungen für WLAN-Verbindung überprüfen | Einstellungen überprüft und die Einstellung fürDatenverbindungsprotokolle $-standards korrigiert | SWO-355547 | Setup | SW | His/Lis | Wireless connection |  | Troubleshooting required | Wi-Fi connection is not working yet, despite Win7 upgrade ==> Check Wi-Fi connection settings | Settings is checked and the setting for data connection protocols is corrected by the
No | Visit | Complaint | RIND | Service Visit/PM Visit. | Sample not aspirate by the anlyzer, need to check, need to do preventive maintenance. | Remove the sensor cassette and solution pack from the analzyzer, open the back panel as well as front panel of the analyzer, remove the dust from the various parts and check the firm connections and electronics contact. Remove the signal data board, remove all the tubings and replace with the new one, check the roller motor and wheel, tigten the screw of rolloer wheel. Close and connect all the parts same as before, insert the solution pack and install the sensor cassette, successfully done. User do the sample, measurement is coming and satisfactory declare by the user. Analyzer handed over to user in well working condition. | SWO-357028 | [5 empty] | Service Visit/PM Visit. | Sample not aspirate by the anlyzer, need to check, need to do preventive maintenance. | Remove the sensor cassette and solution pack from the analzyzer, open the back panel as well as front panel of the analyzer, remove the dust from the various parts and check the firm connections and electronics contact. Remove the signal data board, remove all the tubings and replace with the new one, check the roller motor and wheel, tigten the screw of rolloer wheel. Close and connect all the parts same as before, insert the solution pack and install the sensor cassette, successfully done. User do the sample, measurement is coming and satisfactory declare by the user. Analyzer handed over to user in well working condition.
No | Visit | Complaint | RAME | SC failing to Initialize | SC failed to initialize. Continually fails pO2. Have tried multiple SC and SP. | Put a PM kit on the analyzer and changed the sensor connector board, spine board and analog board. | SWO-354727 | Defect | HW | Basic Unit | PCB, Sensor connector assembly |  | SC failing to Initialize | SC failed to initialize. Continually fails pO2. Have tried multiple SC and SP. | Put a PM kit on the analyzer and changed the sensor connector board, spine board and analog board.
No | Visit | Complaint | RESP | Fuga detectada | Fuga detectada. Inlet Gasket deteriorada. | Se sustituyo inlet gasket y se comprobó correcto funcionamiento del equipo mediante QC. | SWO-354956 | Defect | HW | Auto Inlet module | Inlet gasket |  | Leak detected | Leak detected. Inlet Gasket deteriorated. | Inlet gasket is replaced and equipment operation was proven correctly using QC.
No | Visit | Complaint | RCHN | 报错1999 | 顽固性报错1999.61.22，借903-198、903-029、580-097各一个 | 备件未使用，清洁机器后修复 by Andrew | SWO-355026 | Cleaned | HW | Process Module | Incubator wheel | 1999.61.22 | Error 1999 | Recalcitrant error 1999.61.22, borrowing 903-198, 903-029, 580-097 each | Spare parts not used, repair by Andrew after cleaning machine
No | Visit | Complaint | RCHN | 报错5784和5689 | 报错5784和5689 | 取出断裂的测试卡后修复 by Andrew | SWO-354853 | Defect | Consumables | Test Kit | Cartridge, Broken | 5784 | 5784 and 5689 | 5784 and 5689 | Fix by Andrew after removing broken test card
No | Visit | Complaint | RDEU | AQT bricht sporadisch Proben ab und will neu gestartet werden Fehler 3998 Lock_24 | AQT Auslieferung | AQT ausgeliefert, LQC MCC1, LQC CRP und LQC D-Dimmer gemessen. | SWO-355228 | Defect | HW | Storage Module | rubber disk | 3998 | AQT sporadically cancels samples and wants to be restarted Error 3998 Lock_24 | AQT Delivery | AQT, LQC MCC1, LQC CRP and LQC D-Dimmer measured.
Yes | 8 | Y | RPAC | Blank | Failed SC |  | D4605957DC1F202E | [5 empty] | Blank | Failed SC
No | Visit | Complaint | RDEU | 25.04.2018 15:00Uhr Sr.Laura 06213836237: Die Uhrzeit hat sich nicht auf Sommerzeit umgestellt | Automatische Zeitumstellung nicht erfolgt | Einstellungen für automatische Zeitumstellung überprüft und korrigiert - Funktionsüberprüfung | SWO-355794 | Setup | SW | Time setting | Time setting |  | 25.04.2018 15:00h Sr.Laura 06213836237: The time has not changed to daylight saving time | Automatic time change not made | Settings for automatic time change checked and corrected - Function check
No | Visit | Complaint | RIND | Analyser is not working | 0 cycles was remains in Solution pack | Changed Solution pack. after that analyser was start working. | SWO-357184 | [5 empty] | Analyser is not working | 0 cycles was remains in Solution pack | Changed Solution pack. after that analyser was start working.
No | Visit | Complaint | RCHN | GLU灵敏度超范围 0378 | GLU灵敏度超范围 0378 借903-581一个 | 更换903-581后修复 by Andrew | SWO-354861 | Defect | HW | Basic Unit | PCB, Sensor interface |  | GLU Sensitivity Over range 0378 | GLU Sensitivity Over range 0378 Borrowing 903-581 one | Repair by Andrew after replacement 903-581
No | Visit | Complaint | RESP | Conexion de equipo a red y configuración Instalación de reactivos sp y sc | Conexion de equipo a red y configuración Instalación de SP y Sc Chequeo de datos y sacados reportes defectuosos | Conexion de equipo a red y configuración Instalación de SP y Sc Chequeo de datos y sacados reportes defectuosos | SWO-356285 | Assisted Replacement | SP | SP | unknown |  | Equipment-to-network connection and configuration Installing reagents sp and sc | Equipment-to-network connection and configuration Installing SP and Sc Data check and faulty reports | Equipment-to-network connection and configuration Installing SP and Sc Data check and faulty reports
No | Visit | Complaint | RESP | Fallo en ph y mensajes de solución Nº2 no homogénea | Fallo en ph y mensajes de solución Nº2 no homogénea | Se detectó puerta de entrada defectuosa, se procedió a sustituirla. Tras esto, se realizaron test de servicio satisfactoriamente. Se calibró equipo y se comprobó correcto funcionamiento mediante Qc´s. | SWO-355908 | Defect | HW | Auto Inlet module | Inlet gasket |  | Failure in ph and non-homogeneous solution messages No.2 | Failure in ph and non-homogeneous solution messages No.2 | Faulty entrance door was detected, replaced. After this, service tests were successfully performed. Equipment was graded and proper operation was tested using Qc's.
No | Visit | Complaint | RIND | Service Visit. | There is two clot in the sensor cassette. | Remove the sensor cassette and flush with the distilled water and sodium heparin, successfully removed, connect the SC9 and give the manual two point calibration, successfully done. Analyzer handed over to staff in well working condition. | SWO-357036 | [5 empty] | Service Visit. | There is two clot in the sensor cassette. | Remove the sensor cassette and flush with the distilled water and sodium heparin, successfully removed, connect the SC9 and give the manual two point calibration, successfully done. Analyzer handed over to staff in well working condition.
No | Visit | Complaint | RIND | Analyzer showing barometer error. | Checked Barometer temperature in service menu and found that it was out side range. | Cleaned fan filter and then reinstalled it.Also checked in service menu Temp was 37. Then given 2 point calibration. Unit is working fine. | SWO-357299 | Cleaned | HW | Basic Unit | Fan Filter |  | Analyzer showing barometer error. | Checked Barometer temperature in service menu and found that it was out side range. | Cleaned fan filter and then reinstalled it. Also checked in service menu Temp was 37. Then given 2 point calibration. Unit is working fine.
No | Visit | Complaint | RDEU | 22.05.2018, Herr Vollmer (MT), Tel: 051218941817: ABL fährt selbständig runter und startet neu. AQC werden nicht gemacht | Temperaturfehler im AC Modul >50 Grad | AC Modul demontiert, gereinigt und geprüft, Korrosionsbelag am Flachbandkabel (6 polig zum Termistor) entfernt, Ersatzteile bestellt, Funktionskontrolle ok | SWO-355739 | Revisit required | HW | Analyzer Control Module | termister PCB |  | 22.05.2018, Mr. Vollmer (MT), Tel: 051218941817: ABL shuts down and restarts. AQC are not made | Temperature error in AC module >50 degrees | AC module dismantled, cleaned and tested, corrosion covering on the ribbon cable (6 poles to the termistor) removed, spare parts ordered, function control ok
No | Visit | Complaint | RAME | Customer reporting error 251. | Verified error 251. Performed applicable service diagnostics, Halogen lamp output above limits. | Adjusted Halogen lamp, performed required adjustments, and tests. Ran complete Oxi test, passed. No further issues observed or noted. | SWO-356154 | Adjustment | HW | Oxi module | Lamp unit | halogen | Customer reporting error 251. | Verified error 251. Performed applicable service diagnostics, Halogen lamp output above limits. | Adjusted Halogen lamp, performed required adjustments, and tests. Ran complete Oxi test, passed. No further issues observed or noted.
No | Visit | Complaint | RLTD | Dave Gardner rang direct to myself | Replaced cracked mixer tray in sample mixer unit. Verified operation on restart - all ok. | Replaced sample mixer tray. | SWO-356486 | Defect | HW | Basic Unit | Mixer Tray | crack | Dave Gardner rang direct to myself | Replaced cracked mixer tray in sample mixer unit. Verified operation on restart - all ok. | Replaced sample mixer tray.
No | Visit | Complaint | RFRA | Calibration et CQ en échec sur paramètres d'oxymétrie | Test OXY en mode service. | Remplacement de l'hemolyseur. Calibration, cal tHb, et passage CQ. | SWO-356456 | Defect | HW | Oxi module | Hemolyzer |  | Failed calibration and QC on oxymetry parameters | OXY test in service mode. | Replacement of the hemolyseur. Calibration, cal tHb, and passage CQ.
No | Visit | Complaint | RDEU | 24.05.2018 9:00 Sr. Conny 09071578229: Software bleibt in der Offline -Phase hängen | Fehlersuche vor Ort | Sicherung des Analysators erstellt. Analysator überprüft. Sensorkassette neu eingelegt und dann erneuert. Startphase und Kalibrierungen durchgeführt. QCs für alle Level durchgeführt. ABL wieder messbereit. | SWO-356311 | Defect | SW | SW | off/on | offline | 24.05.2018 9:00 Sr. Conny 09071578229: Software gets stuck in the offline phase | Troubleshooting on site | backup of the analyzer. Analyzer checked. Sensor cartridge reinserted and then renewed. start-up phase and calibrations. QCs for all levels. ABL ready to be measured again.
No | Visit | Complaint | RCHN | 泵轮故障 | 泵轮故障，借902-971一个 | 更换902-971后修复 by Andreew | SWO-355579 | [5 empty] | Pump wheel failure | Pump wheel failure, borrow 902-971 one | Repair by Andreew after replacing 902-971
No | Visit | Complaint | RAME | Blood samples with 292 and 581 errors | Checked data logs Errors codes 292 and 581 | Replaced Hemolyzer assembly and pinch valve. Gather data logs . Open Tech-line case. Ran tHb calibration. Ran system cycles. Ran quality controls. Ran blood samples | SWO-356710 | [5 empty] | Blood samples with 292 and 581 errors | Checked data logs Errors codes 292 and 581 | Replaced Hemolyzer assembly and pinch valve. Gather data logs . Open Tech-line case. Ran tHb calibration. Ran system cycles. Ran quality controls. Ran blood samples
No | Visit | Complaint | RAME | Low MetHb on cap survey | Found the cavitation absorb value out of range on the service test. Also found some 581 errors in the patient results | Replaced hemolyzer assembly and ran THB calibration. This is a warranty repair and the hemolyzer has been decontaminated for return to Radiometer | SWO-356555 | [5 empty] | Low MetHb on cap survey | Found the cavitation absorb value out of range on the service test. Also found some 581 errors in the patient results | Replaced hemolyzer assembly and ran THB calibration. This is a warranty repair and the hemolyzer has been decontaminated for return to Radiometer
No | Visit | Complaint | RCHN | 屏幕不亮 | 屏幕不亮，借910-342一个 | 更换910-342后修复 by Andrew | SWO-355521 | Defect | HW | DMS_PC | CPU | user interface | The screen is not lit | The screen is not bright, borrow 910-342 one | Fix by Andrew after replacing 910-342
No | Visit | Complaint | RCHN | 报错5784 | 报错5784，借944-258（13691）一张 | 重新定标后修复 by Andrew | SWO-355619 | Assisted Replacement | Consumables | Test Kit | Cartridge, NT-PROBNP |  | 5784 | 5784, loan 944-258 (13691) one | Fix by Andrew after re-calibration
No | Visit | Complaint | RCHN | 报错1024 | 报错1024 借944-212 （14137）一张 | 重新定标后修复 by Andrew | SWO-355524 | Defect | HW | Process Module | Optical Unit | 1024 | 1024 error | Error 1024 Borrow 944-212 (14137) a | Fix by Andrew after re-calibration
No | Visit | Complaint | RCHN | 光学单元故障 | 光学单元故障，借903-364一个 | 更换903-364后修复 by Andrew | SWO-355037 | Defect | HW | Process Module | Optical Unit |  | Optical unit failure | Optical unit failure, borrowing 903-364 one | Repair by Andrew after replacing 903-364
No | Visit | Complaint | RCHN | 报错5785 | 报错5785 借903-030 925-016 944-258（14147） | 更换903-030 925-016 重新定标后修复 by Andrew | SWO-354907 | Defect | HW | Storage Module | Cartridge lock | 5785 | 5785 error | Error 5785 Borrowed 903-030 925-016 944-258 (14147) | Replace 903-030 925-016 Repair by Andrew after re-calibration
No | Visit | Complaint | RCHN | 测试卡门卡住 | 测试卡门卡住，借580-098一个 | 更换580-098后修复 by Andrew | SWO-355650 | Defect | HW | Cover and Chassis | Cartridge door |  | Test Carmen stuck | Test Carmen stuck, borrow 580-098 one | Repair by Andrew after replacement 580-098
No | Visit | Complaint | RCHN | AU断开，无法连接 | AU断开，无法连接。借903-198一个 | 未使用备件，仪器重启后又自动连接上了 by Andrew | SWO-356307 | Defect | SW | Analyzer Control Module | CPU |  | AU disconnected, unable to connect | The AU is disconnected and cannot be connected. Borrow 903-198 one | Spare parts are not used, the instrument is restarted and automatically connected to by Andrew
No | Visit | Complaint | RCHN | 报错5684 | 报错5684 借944-220（14154）一张 | 重新定标后修复 by Andrew | SWO-354856 | Defect | Consumables | Test Kit | Cartridge, D-Dimer |  | 5684 error | 5684 Borrow 944-220 (14154) a | Fix by Andrew after re-calibration
No | Visit | Complaint | RIND | frequent calibration error. | found problem with SC and SP. | Checked the instrument for error. Did zeroing for sensor connector.calibrated but it is showing same error.re-installed the SP and SC. but problem not rectified. | SWO-357048 | [5 empty] | frequent calibration error. | found problem with SC and SP. | Checked the instrument for error. Did zeroing for sensor connector.calibrated but it is showing same error.re-installed the SP and SC. but problem not rectified.
Yes | Visit | Complaint | RCHN | 无法启动 | 无法启动 借910-274 910-321各一个 | 更换910-274 910-321后修复 by Andrew | SWO-355031 | [5 empty] | Unable to start | Unable to start Borrowing one each of 910-274 910-321 | Repair by Andrew after replacing 910-274 910-321
No | Visit | Complaint | RKK | QC測定にて電解質（Cl、Na）の値が不安定（範囲外頻発） | QC測定にて電解質（Cl、Na）の値が不安定（範囲外頻発） | PH $BG、EL $Metモジュール各測定チャンバー洗浄。Ref電極交換。PH電極シーリングリング交換。タンパク除去実施。Glu、Lacメンブラン定期交換。エアーライン洗浄。FLEXQモジュールポジション調整実施。 | SWO-356664 | Assisted Replacement | Electrode | Electrode | Reference |  | The value of the electrolyte (Cl, Na) is unstable (frequently outside the range) in QC measurement. | The value of the electrolyte (Cl, Na) is unstable (frequently outside the range) in QC measurement. | PH $BG, EL $Met module each measurement chamber cleaning. Ref electrode replacement. PH electrode sealing ring replacement. Protein removal was carried out. Glu, Lac Membrane regular exchange. Air line cleaning. FlexQ module position adjustment.
No | Visit | Complaint | RDEU | KAL Fehler Elektrolyte / QC Verl. Na | Trotz Memranierungen etc. KAL Fehler Elektrolyte QC Verletzung Natrium | Allgemeine Überprüfung - Erneuerung Referenz Elektrode - mehrfache Kalibrationen durchgeführt - Funktionsüberprüfung etc. ohne Beanstandung | SWO-356454 | Assisted Replacement | Electrode | Electrode | Reference |  | KAL Error Electrolytes / QC Verl. well | Despite memranations etc. KAL Error Electrolytes QC Violation Sodium | General check - renewal reference electrode - multiple calibrations performed - function check etc. without complaint
No | Visit | Complaint | RDEU | KAL Fehler Elektrolyte / QC Verl. Na | KAL Fehler Elektrolyte $ QC Verletzung Na | Allgemeine Überprüfung - Erneuerung Referenz Elektrode - Funktionsüberprüfung etc. ohne Beanstandung | SWO-356455 | Assisted Replacement | Electrode | Electrode | Reference |  | KAL Error Electrolytes / QC Verl. well | KAL Error Electrolytes - QC Violation Na | General inspection - Renewal reference electrode - function check etc. without complaint
No | Visit | Complaint | RAME | High Crea Check borderline failing high | Unit is close to failing high crea check, recommended to replace electrodes and membranes along with electrode connector caps. While looking over unit found measuring chamber to have a possible crack, replaced as a precaution. | Replaced electrodes, membranes, electrode connector caps and measuring chamber | SWO-356032 | User expectation | Specifications | QC | High Crea Check |  | High Crea Check borderline failing high | Unit is close to failing high crea check, recommended to replace electrodes and membranes along with electrode connector caps. While looking over unit found measuring chamber to have a possible crack, replaced as a precaution. | Replaced electrodes, membranes, electrode connector caps and measuring chamber
No | Visit | Complaint | RIND | Analyzer not initializing at time when it showing loading data file. | Windows 7 and ABL800 SW v6.17 to be installed. | Installed Windows 7 and SW v6.17 successfully. Set up performed again. | SWO-356975 | Defect | SW | Windows | Windows |  | Analyzer not initializing at time when it showing loading data file. | Windows 7 and ABL800 SW v6.17 to be installed. | Installed Windows 7 and SW v6.17 successfully. Set up performed again.
No | Visit | Complaint | RDEU | 25.05.2018 16:30Uhr Herr Anton 03455573172: Die Station meldet Oximetriefehler, Gerinnsel erkannt | Vakuumfehler Leckage erkannt LP- und Einlassdichtungswechsel ohne Erfolg | Quetschventil schaltet Spülung der Fließwege mit und ohne SC ohne Probleme Nochmaliger Wechsel der Einlassdichtung ohne Erfolg Austausch Quetschventil und Wechsel der Einlassanschlussdichtung => keine Besserung Wechsel der Einlassnadel aus Kundenbestand => keine Besserung => weiterhin Vacuum Value zu hoch LP reinstalliert => keine Besserung => Lösungen können nicht gepumpt werden LP ersetzt (gleiches Lot wie die 2 LPs zuvor BL-50) => im ersetzten LP Pumpenschlauch rötlich verfärbt => Fehler bleibt Jahreswartungskit verbaut Geräteneustart => Startphase i.O. In Kal noch Lac mit Reaktionsfehler => mehrmals Kal nachgeholt => keine Besserung SC Reinstallierung gestartet Info an Station und Hr. Anton | SWO-357016 | Defect | HW | Basic Unit | PM kit |  | 25.05.2018 16:30Mr Anton 03455573172: The station reports oximetry errors, clots detected | Vacuum error Leakage detected LP and inlet seal changes without success | Squeeze valve switches Flushing of flow paths with and without SC without problems Another change of the inlet seal without success Replacement of crushing valve and change of inlet connection seal => no improvement Change of the inlet needle from customer stock => no improvement => still Vacuum Value too high LP reinstalled => no improvement => Solutions cannot be pumped LP replaced (same lot as the 2 LPs previously BL-50) => reddish discolored in the replaced LP pump hose => Error remains Annual maintenance kit installed Device restart => Start phase in the In Kal still Lac with reaction error => several times Kal made up => no improvement SC reinstall started Info at Station and Mr. Anton
No | Visit | Complaint | RLTD | Adam on 01772 523007 | Site visit to investigate printer intermittently stopping printing. | Arrived on site and took backup of analyser setup and data. Took service dump for Denmark to investigate. Powered down and fitted new cpu and flash card. Changed sybase for Windows 7. Installed windows 7. Installed app software. Changed date, time and regional settings to correct ones. Committed changes to drive and restarted. Loaded consumables again to get access to main screen and startup. Restored customer data and setup. Rebooted. Analyser restarted with 50 min startup. | SWO-357113 | Defect | HW | DMS_PC | CPU |  | Adam on 01772 523007 | Site visit to investigate printer intermittently stopping printing. | Arrived on site and took backup of analyser setup and data. Took service dump for Denmark to investigate. Powered down and fitted new cpu and flash card. Changed sybase for Windows 7. Installed windows 7. Installed app software. Changed date, time and regional settings to correct ones. Committed changes to drive and restarted. Loaded consumables again to get access to main screen and startup. Restored customer data and setup. Rebooted. Analyser restarted with 50 min startup.
No | Visit | Complaint | RDEU | 02.05.2018, 11:10 Uhr, Fr. Rosebrook, Tel. 05191800184: Einlass, Einlassdichtung prüfen 17.05.2018 14:15 Frau Reusmann 05191800594 Einlass, Einlassdichtung prüfen; Einlassdichtung ersetzen 24.05.2018 15:00 Sr.Heike 05191800181: Das Problem besteht weiterhin, ständig Fehlermeldung bei Kapillareingabe. | System überprüft => Bügelventil schwergänig , Einlasssonde verbogen | Bügelventil gangbar gemacht und Sonde wieder gerade gebogen , Kalibration und Qc durchgeführt | SWO-356770 | Adjustment | HW | Auto Inlet module | Probe bend |  | 02.05.2018, 11:10 a.m., Ms. Rosebrook, tel. 05191800184: inlet, check inlet seal 17.05.2018 14:15 Mrs. Reusmann 05191800594 Inlet, check inlet seal; Replace inlet seal 24.05.2018 15:00 Sr.Heike 05191800181: The problem persists, constantly error message for capillary input. | System checked => Handle valve heavy-duty, inlet probe bent | Ironing valve made manable and probe again straight bent, calibration and Qc performed
No | Visit | Complaint | RFRA | Lors du remplacement de la cassette capteur impossible de fermé la Bay unit | Mode service Bay unit ne se ferme pas. | Démontage réglage de la sécurité fermeture de la Bay unit. Vérification ouverture et fermeture Bay unit. OK Vérification: en mode service fluide transport, agitateur. OK Installation cassette: calibration et CQ pour Lac ne passent pas. | SWO-356831 | Adjustment | HW | Basic Unit | Sensor Cassette Door |  | When replacing the sensor cassette impossible to close the Bay unit | Mode service Bay unit does not close. | Disassembly setting the closure of the Bay unit. Check opening and closing Bay unit. Ok Check: in fluid service mode transport, agitator. Ok Installation cassette: calibration and CQ for Lake do not pass.
No | Visit | Complaint | RCHN | GLU,LAC一直飘移错误 | GLU,LAC一直飘移错误，借903-581一个 |  | SWO-356914 | Defect | HW | Basic Unit | PCB, Sensor interface | Glu/Lac | GLU, LAC has been drifting errors | GLU, LAC has been drifting errors, borrow 903-581 one
No | Visit | Complaint | RCHN | LAC定标不过 | LAC定标不过 ，借903-581一个 | 更换903-581后修复 by Andrew | SWO-355499 | Defect | HW | Basic Unit | PCB, Sensor interface | Lac | LAC calibration is not | LAC calibration is not, borrow 903-581 one | Repair by Andrew after replacement 903-581
No | Visit | Complaint | RCHN | 报错5684 | 报错5684 | 清理处理盘传感器。 by Andrew | SWO-356716 | Cleaned | HW | Storage Module | Storage wheel | 5684 | 5684 error | 5684 error | Clean the processing disk sensor. by Andrew
No | Visit | Complaint | RIND | Printer is not working | Opened instrument and checked printer connection.Also done reset of analyzer and restart it. | Printer is not working after that also.Neeed to change printer unit. | SWO-357318 | [5 empty] | Printer is not working | Opened instrument and checked printer connection. Also done reset of analyzer and restart it. | Printer is not working after that also. Neeed to change printer unit.
Yes | Visit | Complaint | RFRA | Lors de la mesure en 95 ul instabilité Lactate En mode 195 ul pas d'alarme d'instabilité | regularisation stock suite à expésdition electrode de lactate | N $A | SWO-357034 | [5 empty] | When measuring in 95 ul instability Lactate In 195 ul mode no instability alarm | stock regularization following lactate electrode expésdition | No$A
Yes | Visit | Complaint | RDEU | 24.05.2018, Sr. Sarah, Tel: 02037330: Ansprachefehler bei pH und E´lyte | keine Ref.elektrode im Gerät, Gerät ist nicht Einsatzbereit, | neue Ref.elektrode eingesetzt, Gerät kalibriert, | SWO-356967 | [5 empty] | 24.05.2018, Sr. Sarah, Tel: 02037330: Ansprachefehler bei pH und E´lyte | no ref.electrode in the device, device is not ready for use, | new Ref.electrode, device calibrated,
No | Visit | Complaint | RAME | Cal ver sample type unavailable | Inspection of the system revealed the cal ver is corrupted. | 5-24-18 Reset Radiometer default setup values, same issue. Fri 5-25-18 Reset SW version to clean 6.17 the reloaded setup manually. | SWO-356603 | Setup | SW | Roma.ini | Cal ver |  | Cal ver sample type unavailable | Inspection of the system revealed the cal ver is corrupted. | 5-24-18 Reset Radiometer default setup values, same issue. Fri 5-25-18 Reset SW version to clean 6.17 the reloaded setup manually.
Yes | Visit | Complaint | RDEU | 24.05.2018, 09:19 Uhr, Sr. Wibke, Tel. 021714092693: Kal und QC Fehler bei pCO2 | S. Case | Fehlersuche: „Frosted Window“ der Messkammer fehlt und ist nicht auffindbar, wird bestellt und nachgerüstet. Kalibrierungen seit vorgeschlagener Reinigungsaktion gestern Nachmittag alle ok, QC Messung konnte aufgrund von nicht nachgefülltem AC Karussell nicht automatisch stattfinden. Karussell aufgefüllt, QC Messung wiederholt, Endkontrolle ok. | SWO-356919 | Re classified RMED | installation report |  |  |  | 24.05.2018, 09:19, Sr. Wibke, tel. 021714092693: Kal and QC error in pCO2 | S. Case | Troubleshooting: "Frosted Window" of the measuring chamber is missing and cannot be found, is ordered and retrofitted. Calibrations since proposed cleaning action yesterday afternoon all ok, QC measurement could not take place automatically due to unfilled AC carousel. Carousel replenished, QC measurement repeated, final check ok.
Yes | Visit | Complaint | RCHN | 温度不稳定 | 温度不稳定，借903-334一个 | 更换备用903-334后修复 | SWO-356813 | [5 empty] | Temperature instability | Temperature instability, borrow 903-334 a | Repair after replacing spare 903-334
No | Visit | Complaint | RCHN | 血氧单元模块定标不过 | 血氧单元模块定标不过，借902-980一个 |  | SWO-357347 | Revisit required | HW | Oxi module | Hemolyzer | spectrometer 30/5 | Blood oxygen unit module calibration is not | Blood oxygen unit module calibration is however, borrow 902-980 a
No | Visit | Complaint | RCHN | 报错1310，0210，0211 | 报错1310，0210，0211，借903-581一个 |  | SWO-357334 | Defect | HW | Basic Unit | PCB, Sensor interface | Responce | Error 1310,0210,0211 | 1310,0210,0211, loan 903-581 one
No | Visit | Complaint | RCHN | 液路板错误 | 液路板错误，借903-644一个 | 更换903-644后修复 By Andrew | SWO-356690 | Defect | HW | Basic Unit | PCB, Wet section |  | Liquid plate error | Liquid road plate error, borrow 903-644 one | Fix By Andrew after replacing 903-644
No | Visit | Complaint | RCHN | LAC定标不过 | LAC定标不过 借903-581一个 | 更换903-581后修复 By Andrew | SWO-356815 | Defect | HW | Basic Unit | PCB, Sensor interface | Lac | LAC calibration is not | LAC calibration but borrow 903-581 one | Fix By Andrew after replacing 903-581
No | 3 | Y | RCZE | Blank |  |  | D46598D4B4210A87 | Setup | SW | His/Lis | His/Lis |  | Blank
Yes | Visit | Complaint | RFRA | Pb de configuration de NetOp | Contrôle de la configuration | Réinstallation de Windows et de l’application de l’ABL. Configuration des connections. Essais | SWO-357849 | [5 empty] | NetOp configuration pb | Configuration control | Reinstalling Windows and the ABL app. Setting up connections. Tests
No | Visit | Complaint | RDEU | 24.05.2018 9:00 Sr. Conny 09071578229: Software bleibt in der Offline -Phase hängen 25.05.2018 08:34 Herr Richter 09071578229: Die Software hat sich wieder aufgehangen und das Gerät ist nicht meßbereit. 28.05.2018, 08:23 Uhr, H. Kummer, Tel. 09071578229: Problem besteht weiter. ABL nicht messbereit. | Fehlersuche vor Ort | Bestellte CPU und Netzteil eingebaut. Bei Installation von Software keine Verbindung zu weit section herstellen können. Alte cou und Netzteil wieder eingebaut. Welt section PCB bestellt. Herrn Kummer darüber informiert das ABL nicht einsatzbereit ist. | SWO-357776 | Revisit required | HW | Basic Unit | PCB, Wet section |  | 24.05.2018 9:00 Sr. Conny 09071578229: Software gets stuck in the offline phase 25.05.2018 08:34 Mr. Richter 09071578229: The software has hung up again and the device is not ready to measure. 28.05.2018, 08:23, H. Kummer, tel. 09071578229: Problem persists. ABL not ready to measure. | Troubleshooting on site | Ordered CPU and power supply installed. Cannot connect too far when installing software. Old cou and power supply reinstalled. World section PCB ordered. Mr Kummer informs the ABL is not operational.
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | lac定标失败 | 更换测试卡后修复 | SWO-357682 | [5 empty] | Test card, reagent package failure | lac calibration failed | Repair after replacing test card
No | Visit | Complaint | RESP | Error calibración | Fallo en pO2 | Se procede a comprobar registros y al cacmbio de SP. Se realizan varias Cal. así como Qc´s. Funcionamiento correcto. | SWO-357723 | Assisted Replacement | SP | SP | pO2 | cal | Calibration error | Failure in pO2 | Records and SP cacmbio are checked. Several Cals. as well as Qc's. Proper Operation are performed.
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 参比电极错误 1361 | 更换测试卡后修复 | SWO-357602 | Assisted Replacement | SC | SC | Reference | 1361 | Test card, reagent package failure | Reference electrode error 1361 | Repair after replacing test card
No | Visit | Complaint | RESP | pCO2 Fuera de rango | pCO2 Fuera de rango. Requiere sustitucion SCI. | Se sustituyo SCI y se comprobo correcto funcionamiento del equipo mediante QC. | SWO-357845 | Assisted Replacement | SC | SC | pCO2 |  | pCO2 Out of range | pCO2 Out of range. Requires SCI replacement. | SCI is replaced and the equipment is compatible with QC.
Yes | 2 | Y | RPAC | Blank | software upgrade to v1.2 FAN915-380 |  | D46565FBB43BF582 | [5 empty] | Blank | software upgrade to v1.2 FAN915-380
No | Visit | Complaint | RCHN | 显示屏下方扫码器无法识别条码 | 显示屏下方扫码器无法识别条码 | 显示屏下方连接计算机单元的排线松了，重新连接后修复by tim | SWO-357977 | Adjustment | HW | Cover and Chassis | cable, display |  | The code sweeper below the display does not recognize the barcode | The code sweeper below the display does not recognize the barcode | The line connecting the computer unit below the display is loose and the by-tim is repaired after reconnection
No | Visit | Complaint | RCHN | CA质控偏高 | CA质控偏高，借902-824一个 |  | SWO-357536 | Defect | HW | El/Met Module | El/Met Module | Ca | CA quality control is high | CA quality control is high, borrow 902-824 one
No | Visit | Complaint | RKK | QCにてOxiのMetHb範囲外 | QCにてOxiのMetHb範囲外 | ヘモライザー故障の為、交換を実施しました。 | SWO-357906 | Defect | HW | Oxi module | Hemolyzer |  | Oxi's MetHb out of range at QC | Oxi's MetHb out of range at QC | Replacement was carried out for the hemolizer failure.
No | Visit | Complaint | RIND | pH Values not Reliable. | pH values not reliable. | Electrode service program done. Chambers cleaned and service properly. Analyser Serviced and calibrated properly. Handed over in a fine working condition. | SWO-359673 | Cleaned | Fluidic pathway | pH/BG Module | pH |  | pH Values not Reliable. | pH values not reliable. | Electrode service program done. Chambers cleaned and service properly. Analyser Serviced and calibrated properly. Handed over in a fine working condition.
No | Visit | Complaint | RDEU | 09.05.2018 12:30 Herr Dietz DDimer Proben werden abgebrochen | Bei DDimer unerwartete Messabbrüche bei Blutproben. | Fehler analysiert. Es gibt keine konkreten Anhaltspunkte warum TnT immer funktionierte, DDimer aber abgebrochen wurde. 12 Testmessungen mit Pufferlösung ( Blanktest ) und QC Lösung durchgeführt. Alle Messungen TnT und DDimer sind ohne Störung durchgelaufen. Weitere Tests mit Blutproben werden von HarHo durchgeführt . | SWO-357995 | Revisit required |  |  |  | mail | 05/09/2018 12:30 Mr Dietz DDimer Samples are cancelled | In DDimer unexpected drop-outs in blood samples. | Error analyzed. There is no concrete evidence as to why TnT always worked, but DDimer was canceled. 12 test measurements performed with buffer solution ( blank test ) and QC solution. All measurements TnT and DDimer have passed without interference. Further tests with blood samples are carried out by HarHo .
No | Visit | Complaint | RFRA | Pb d'initialisation du module process unit. | Erreur 5257. Process unit 24. Blocage mécanique descente module optic | bloquage mécanique du process unit avec le ressort de la cup wash. Rotation du ressort du la cup Wash 1 $2 tour. Essais ok. Redémarrage ok. Passage 3 blancs TnI. Ok. <0,007 Calibration TnI lot 13991 car contrôle niv 2 ok mais trop haut pour les bornes Cq client. Passage CQ TnI niv 1 et 2 OK | SWO-357880 | User error | HW | Process Module | Wash unit | Spring | Pb to boot the process module unit. | Error 5257. Process unites 24. Mechanical blocking descent module optic | mechanical bloquation of the process unites with the spring of the cup wash. Rotation of the spring of the cup Wash 1 $2 round. Try it ok. Restart ok. Passage 3 white TnI. Ok. 0.007 Calibration TnI lot 13991 because control niv 2 ok but too high for the customer Cq terminals. Passage CQ TnI niv 1 and 2 OK
No | Visit | Complaint | RDEU | 25.05.2018 16:30 Fr. Lener 08856936370: Systemfehler Inkubation-Wheel | Systemfehler Inkubation-Wheel, Cartridge lock fehler | Gerät mehrmals ein und ausgeschaltet, Safe Position Run all gemacht, Cartridge lock failed, Storage Wheel und Inkubator Wheel aus gebaut, Staub entfernt und gereinigt, alle Sensoren gereinigt, Gerät eingeschaltet, Startphase, Fehler weiterhin vorhanden, Cartridge Lock PCB defekt, Ohne Erfolg, Ersatzteile bereits bestellt | SWO-357987 | Revisit required | HW | Storage Module | PCB, Cartridge lock |  | 25.05.2018 16:30 Fr. Lener 08856936370: System error incubation wheel | System error incubation wheel, cartridge lock error | Device multiple times on and off, Safe Position Run all made, Cartridge lock failed, Storage Wheel and Incubator Wheel built out, dust removed and cleaned, all sensors cleaned, device turned on, start phase, error still present, Cartridge Lock PCB defective, Without success, spare parts already ordered
No | Visit | Complaint | RBNL | SETSAFE_22 Error | AQT90 komt in FOUT en geeft daarbij de volgende error codes: 5784: SAFESET_22 5689: ROTATESTORAGEWHEEL_24 | Bij nader onderzoek blijkt dat een van de zwarte houdertjes van het storagewheel los te zijn geworden. Zodoende vast heb gezet vervolgens heb gecontroleerd, OK. | SWO-357646 | Defect | HW | Storage Module | Storage wheel | Cartridge guide | SETSAFE_22 Error | AQT90 comes in ERROR and gives the following error codes: 5784: SAFESET_22 5689: ROTATESTORAGEWHEEL_24 | Upon further examination it appears that one of the black carts of the storagewheel has become loose. So fixed have put then have checked, OK.
Yes | 3 | Y | RPAC | Blank | pO2 low on QC |  | D467D6BDA0121CCE | [5 empty] | Blank | pO2 low on QC
Yes | 3 | Y | RPAC | Blank | Replaced failed SC |  | D46662312A35B2A1 | [5 empty] | Blank | Replaced failed SC
No | Visit | Complaint | RDEU | 24.05.2018 9:00 Sr. Conny 09071578229: Software bleibt in der Offline -Phase hängen 25.05.2018 08:34 Herr Richter 09071578229: Die Software hat sich wieder aufgehangen und das Gerät ist nicht meßbereit. 28.05.2018, 08:23 Uhr, H. Kummer, Tel. 09071578229: Problem besteht weiter. ABL nicht messbereit. | Fehlerbehebung | Alte Wretsection PCB R2007N0090 ausgebaut und neue Wetsection PCB R001N008 eingebaut weil Analysator immer offline und keine Verbindung zur AU hatte. Nach Einbau war Verbindung zur AU wieder hergestellt. Softwar Update durchgeführt weil PCB noch alte Software hatte. Startphase abgewartet. Analysator wieder messbereit. (S sind noch nicht alle Parameter auf grün , Herr Kummer wird aber noch ein paar mal Kalibrieren und abwarten ob sich die Sensorkassette amortisiert. Ansonsten wird er die Sensorkassete tauschen. ) | SWO-358172 | Defect | HW | Basic Unit | PCB, Wet section |  | 24.05.2018 9:00 Sr. Conny 09071578229: Software gets stuck in the offline phase 25.05.2018 08:34 Mr. Richter 09071578229: The software has hung up again and the device is not ready to measure. 28.05.2018, 08:23, H. Kummer, tel. 09071578229: Problem persists. ABL not ready to measure. | Troubleshooting | Old Wretsection PCB R2007N0090 removed and new Wetsection PCB R001N008 installed because analyzer always had offline and no connection to the AU. After installation, connection to the AU was restored. Softwar update because PCB still had old software. start-up phase. Analyzer ready to measure again. (S are not yet all parameters on green , but Mr. Kummer will calibrate a few more times and wait if the sensor cartridge pays for itself. Otherwise, he will replace the sensor cash. )
No | Visit | Complaint | RKOR | Sample clot error | . | Liquid flow line clean. Sensor cassette clean. Solution pack clean. Waste line clean. 2point cal work. | SWO-362727 | [5 empty] | Sample clot error | . | Liquid flow line clean. Sensor cassette clean. Solution pack clean. Waste line clean. 2point cal work.
Yes | Visit | Complaint | RCRO | CCN/PER | Error 1218 System Cycle failed | Glu | SWO-368970 | [5 empty] | CCN/PER | Error 1218 System Cycle failed | Glu
No | Visit | Complaint | RAME | Customer reports possible fluid blockage error after changing solution pack. | Possible blockage | Reset solution pack, removed cassette and solution pack rebooted instrument , found corrupt system cycle file acknowledge file error, installed solution pack and cassette ran external qc | SWO-358277 | [5 empty] | Customer reports possible fluid blockage error after changing solution pack. | Possible blockage | Reset solution pack, removed cassette and solution pack rebooted instrument , found corrupt system cycle file acknowledge file error, installed solution pack and cassette ran external qc
No | Visit | Complaint | RCHN | lac定标失败 | lac定标失败，需要借件903-581 | 更换903-581备件后修复by tim | SWO-357638 | Defect | HW | Basic Unit | PCB, Sensor interface | Lac | lac calibration failed | Lac bid failed, need to borrow 903-581 | Repair by tim after replacement of 903-581 spare parts
No | Visit | Complaint | RCHN | 报错5684 | 报错5684 借944-216（14259）一张 | 重新定标后修复by tim | SWO-357610 | Defect | Sw | Analyzer Control Module | PC | 5684 | 5684 error | 5684 Borrow 944-216 (14259) a | Fix by tim after re-calibration
No | Visit | Complaint | RCHN | 报错1024 | 报错1024，借944-394（13955）一个 | 重新定标后修复by tim | SWO-358087 | Defect | HW | Process Module | Optical Unit | 1024 | 1024 error | Report error 1024, borrow 944-394 (13955) one | Fix by tim after re-calibration
No | Visit | Complaint | RCHN | 报错4035、4036 | 报错4035、4036 | 重新清理传感器和删除数据后修复by tim | SWO-357996 | Adjustment | HW | Aspiration Module | Sensor, inlet wheel | 4035 | Error 4035, 4036 | Error 4035, 4036 | Fix by tim after cleaning up the sensor and deleting the data
No | 8 | Y | RCZE | Blank | chybnì detekuje hladinu odpadu |  | D467640FA05012BF | Defect | HW | Basic Unit | Waste detector |  | Blank | Detects the waste level incorrectly
No | Visit | Complaint | RIND | Option key not installed properly. | Project key need to be installed for Glucose parameter | Installed the project key but still Glu parameter not activated,replace the iButton with another one,reinstall the option keys & all the parameter successfully. | SWO-358418 | Defect | HW | Analyzer Control Module | HW key |  | Option key not installed properly. | Project key need to be installed for Glucose parameter | Installed the project key but still Glu parameter not activated,replace the iButton with another one,reinstall the option keys & all the parameter successfully.
No | Visit | Complaint | RFRA | pCO2 rendue trop basse sur cet appareil pour tous les niveaux de CQ, mais surtout pour le niveau le plus élevé (niveau 4) Rétrospectivement, cela était déjà le cas à l’issue des comparaisons inter-automates , même si acceptable « à la limite ». Nous sortons des bornes Radiometer pour le niveau 4. Un changement de membrane pCO2 et plusieurs calibrations n’ont rien corrigé. | Résultat pco2 trop bas sur les niveaux de contrôles | Vérification module gaz , changement des bouteilles des gaz, démontage des vannes d'air et de rinçage , changement des membranes, changement des tubing du module pH, changement de la chambre de mesure , changement joint pH et cl-,changement rotor A. Calibration ok passage des contrôles(légère augmentation des contrôles) , récupération du service dump pour expertise . | SWO-357895 | Defect | HW | pH/BG Module | measuring chamber | pCO2 | pCO2 made too low on this device for all levels of CQ, but especially for the highest level (level 4) In retrospect, this was already the case after the inter-automate comparisons, even if acceptable "at the limit". We're leaving the Radiometer terminals for Level 4. A change in pCO2 membrane and several calibrations did not correct anything. | Pco2 result too low on control levels | Gas module check, gas cylinder change, dismantling of air valves and rinsing, changing membranes, changing the tubing of the pH module, changing the measuring chamber, changing pH and cl-, changing rotor A. Calibration ok passing controls (slight increase in controls), recovery of the dump service for expertise.
No | Visit | Complaint | RDEU | 25.05.2018 16:30 Fr. Lener 08856936370: Systemfehler Inkubation-Wheel | Systemfehler Inkubation-Wheel, Cartridge lock fehler | Cartridge Lock PCB und Cable FPC 6-pole, 140mm erneuert, Startphase, LQC | SWO-358569 | Defect | HW | Storage Module | PCB, Cartridge lock |  | 25.05.2018 16:30 Fr. Lener 08856936370: System error incubation wheel | System error incubation wheel, cartridge lock error | Cartridge Lock PCB and Cable FPC 6-pole, 140mm renewed, start phase, LQC
No | Visit | Complaint | RFRA | Difficulté à aspirer un échantillon (seringue). Lorsque le client insère la seringue dans le site d'introduction, il est obligé de reculé un peu la seringue pour que l'aspiration se fasse. Il a déjà remplacé plusieurs fois de Pack de Solution mais toujours le même problème. Site d'introduction ou carte de détection du site certainement défectueux. | Difficulté aléatoire à aspirer | Changement lors de la maintenançe de la carte Inlet et du module du site introduction Calibration ok Validation cf rapport swo de la maintenance préventive | SWO-358820 | Defect | HW | Auto Inlet module | PCB, Auto inlet | HALL Element | Difficulty sucking up a sample (syringe). When the client inserts the syringe into the introductory site, he is forced to step back the syringe a little for suction to be done. It has already replaced several times solution pack but still the same problem. Introductory site or site detection card certainly defective. | Random difficulty to suck | Change when maintaining the Inlet map and the introduction site module Calibration ok Validation cf swo report of preventative maintenance
No | Visit | Complaint | RBNL | Via account manager gehoord dat er met name in het hoge gebied wat afwijking is in de glucose resultaten tussen beide instrumenten voor SKML metingen. Geen data meegenomen. | vergelijk met bloedgasspuiten ABL90 1 ABL90 2 1 7,1 8 2 3,7 3,5 3 4,7 4,5 4 3,8 3,7 5 3,6 3,6 Laatste glu calibratie status 398pA en gevoeligheid 816,0 pA $mMol. Graag langer vergelijk maken. | Connector pennetjes schoongemaakt. Simulatortest met succes uitgevoerd. Opnieuw opgestart en laten kalibreren. GLU status nu 1037 pA en gev. 847,2 pA $mMol. | SWO-358624 | No fault found | performance |  | Glu |  | Through account Manager heard that especially in the high area there is some deviation in glucose results between the two instruments for SKML measurements. No data included. | Compare with blood gas syringes ABL90 1 ABL90 2 1 7.1 8 2 3.7 3.5 3 4.7 4.5 4 3.8 3.7 5 3.6 3.6 Latest Glu Calibration status 398pA and sensitivity 816.0 pA /mMol. Like to make comparison longer. | The Connector is cleaned. Simulator test performed successfully. Restarted and have it calibrated. GLU status now 1037 pA and GeV. 847.2 pA $mMol.
No | Visit | Complaint | RDEU | 28.05.2018, Herr Leien (MT), Tel: 03761444720: seit dem Installation von Windows 7 Rev.2 zeigt der ABL 90 andere Grenzen und Mittelwert für Na gegenüber dem Laborprogramm | Einstellungen überprüfen | Roma.ini überprüft ==> Eintrag für Fixe Zielwerte in Roma.ini ergänzt | SWO-359100 | Setup | SW | Roma.ini | Fixed assigned values |  | 28.05.2018, Mr. Leien (MT), Tel: 03761444720: since the installation of Windows 7 Rev.2, the ABL 90 shows other limits and averages for Na compared to the laboratory program | Check settings | Roma.ini reviewed ==> Entry for fixes target values in Roma.ini added
Yes | Visit | Complaint | RDEU | Fehlerbehebung erforderlich | WLAN-Verbindung war wieder abgebrochen | Einstellungen in Windows für Surf Wireless Micro USB Adapter überprüft ==> unter Configuration $Advanced Propertyp suf Wireless Mode und Value auf IEEE 802.11b $g $n eingestellt ==> WLAN-Verbindung zum Netzwerk wieder o.k. In Datei Aktivitäten befinden sich Einträge, dass die Verbindung sehr häufig verloren und wieder hergestellt wird ==> Herr Lein wird mit der IT-Abteilung die Routereinstellungen überprüfen (Herr Lempe war nicht erreichbar) | SWO-359101 | Re classified RMED | same day upSWO |  | same day UpSWO | see SWO-359100 | Troubleshooting required | Wi-Fi connection was canceled again | Check settings in Windows for Surf Wireless Micro USB Adapter ==> under Configuration $Advanced Propertyp suf Wireless Mode and Value set to IEEE 802.11b $g $n ==> Wi-Fi connection to the network again OK. In file activities there are entries that the connection is lost and re-established very often ==> Mr. Lein will check the router settings with the IT department (Mr. Lempe was not reachable)
No | Visit | Complaint | RCHN | 血氧单元模块损坏 | 血氧单元模块损坏，借903-197一个 | 更换了903-197后修复by tim | SWO-357949 | Defect | HW | Oxi module | spectrometer | Hemolyzer 26/5 | Damage to the blood oxygen unit module | Damage to the blood oxygen unit module, borrowed 903-197 one | Replaced 903-197 after repair by tim
No | Visit | Complaint | RCHN | 报错2167，卡盘 | 需要借件944-216（14099）、944-258（14051）、944-220（14087）、944-214（14107） | 重新定标后修复by tim | SWO-358116 | Assisted Replacement | HW | Storage Module | sensor, cup |  | 2167, card | Borrowed 944-216 (14099), 944-258 (14051), 944-220 (14087), 944-214 (14107) | Fix by tim after re-calibration
No | Visit | Complaint | RCHN | 报错1031 | 报错1031，借903-364,944-216（13940），944-220（13785），944-214（13886），944-258（13280），944-212（13982）各一个 | 更换903-364后重新定标修复by tim | SWO-357619 | Defect | HW | Process Module | Optical Unit | 1031 | 1031 | 1031, borrow 903-364,944-216 (13940), 944-220 (13785), 944-214 (13886), 944-258 (13280), 944-212 (13982) each | Re-calibration repair by tim after replacing 903-364
No | Visit | Complaint | RCHN | 报1031 | 报1031，借903-364,944-216（13940），944-220（13785），944-214（13886），944-258（13280），944-212（13982）各一个 | 更换903-364后重新定标后修复by tim | SWO-357620 | Defect | HW | Process Module | Optical Unit | 1031 | Reported 1031 | Reported 1031, lending 903-364,944-216 (13940), 944-220 (13785), 944-214 (13886), 944-258 (13280), 944-212 (13982) each | Repair by tim after re-calibration after replacement of 903-364
No | Visit | Complaint | RCHN | 报错2167 | 报错2167，借636-539,944-216（13707），944-214（13800），944-258（13642）各一个 | 更换636-539，重新定标后修复 by tim | SWO-358040 | Defect | HW | Storage Module | sensor, cup |  | 2167 error | 2167, loan 636-539, 944-216 (13707), 944-214 (13800), 944-258 (13642) each | Replace 636-539, fix by tim after recalibration
No | Visit | Complaint | RCHN | 仓门故障 | 仓门故障，借903-198一个 | 更换了903-198后修复by tim | SWO-357969 | Defect | HW | Analyzer Control Module | PCB, Main Board |  | Warehouse failure | Warehouse door failure, borrow 903-198 one | Replaced 903-198 after repair by tim
Yes | Visit | Complaint | RAME | Customer reported no readings for O2, hemoglobin and other oximetry values. | Oximetry error no patient result. Error 791. | Replaced pump tubing | SWO-358899 | Defect | HW | Fluidic manifold | Pump tubes |  | Customer reported no readings for O2, hemoglobin and other oximetry values. | Oximetry error no patient result. Error 791. | Replaced pump tubing
No | Visit | Complaint | RAME | PCO2 failing calibrations | PCO2 failing calibrations. PCO2 is failing with sensitivity of 0.3%. Inspected electrode with not issues(broken glass membrane). | Replaced the PCO2 electrode under warranty (install date 8 $2017). Several calibrations (sensitivity 94%) Calibrations and QC to verify performance. | SWO-358930 | Assisted Replacement | Electrode | Electrode | pCO2 |  | PCO2 failing calibrations | PCO2 failing calibrations. PCO2 is failing with sensitivity of 0.3%. Inspected electrode with not issues(broken glass membrane). | Replaced the PCO2 electrode under warranty (install date 8 $2017). Several calibrations (sensitivity 94%) Calibrations and QC to verify performance.
Yes | Visit | Complaint | RKK | ｷｬﾘﾌﾞﾚｰｼｮﾝ時にtHbのﾄﾞﾘﾌﾄｴﾗｰ発生 | Ｃａｌ時にtHbドリフト | 修理完了 機器入替 | SWO-359706 | Re classified RMED | logistic SWO |  |  | mail | THb drift error occurs during calibration | Cal when tHb drift | Repair complete Equipment replacement
No | Visit | Complaint | RDEU | 30.05.18 Pfleger Bernhard Tel.0210389921190 , Neon Lamp too high, Trigger failed 30.05.2018,Frau Hagedorn, tel: 0212645645432: Fehler besteht weiterhin | Neon number too low | Bei Ankunft alles grün, OXI Complete Test durchgeführt neon Intensität zu niedrig, (30% wert über 50 wäre bestanden)neues Oxi Modul eingebaut thb kal durchgeführt und bestanden,oxi complete Test durchgeführt und bestanden Service thb durchgeführt,und kal Verifizierung durchgeführt, 2 Punkt kal und QCs bestanden | SWO-359672 | Defect | HW | Oxi module | Oxi module |  | 30.05.18 Caregiver Bernhard Tel.0210389921190 , Neon Lamp too high, Trigger failed 30.05.2018,Mrs. Hagedorn, tel: 0212645645432: Error persists | Neon number too low | On arrival all green, OXI Complete test performed neon intensity too low, (30% worth over 50 would pass)new Oxi module installed thb kal carried out and passed,oxi complete test performed and passed service thb performed,and kal verification performed, 2 point kal and QCs passed
Yes | Visit | Complaint | RDEU | AQT bricht sporadisch Proben ab und will neu gestartet werden Fehler 3998 Lock_24 29.05.2018, 11:35 Uhr, Fr. Hierner, Tel. 08921777735: Fehler ist wieder aufgetreten. | Fehler 3998 Lock 24 | 2 Sensoren von Storage Wheel und Inkubator Wheel Lock erneuert, Startphase, Fehler verschwunden, LQC vorgenommen.ok | SWO-359748 | [5 empty] | AQT sporadically cancels samples and wants to be restarted Error 3998 Lock_24 29.05.2018, 11:35 a.m., Fr. Hierner, tel. 0892177735: Error has occurred again. | Error 3998 Lock 24 | 2 sensors from Storage Wheel and Incubator Wheel Lock renewed, start phase, bug gone, LQC made.ok
No | Visit | Complaint | RAME | tHb proficiency issues | Reviewed logs and inspected instrument. | No action taken. Exported data logs. | SWO-359549 | No fault found |  |  |  | tHb | tHb proficiency issues | Reviewed logs and inspected instrument. | No action taken. Exported data logs.
No | Visit | Complaint | RAME | Wifi connection unstable | Captured data and tested | None necessary | SWO-359630 | No fault found | SW |  | WIFI |  | Wifi connection unstable | Captured data and tested | None necessary
No | Visit | Complaint | RAME | Customer getting oxi errors on patient sample error codes are 292, 581, 688 | Found that they are getting oxi error still | Replaced hemolyzer on analyzer to correct the issue. | SWO-360494 | Defect | HW | Oxi module | Hemolyzer | 581 | Customer getting oxi errors on patient sample error codes are 292, 581, 688 | Found that they are getting oxi error still | Replaced hemolyzer on analyzer to correct the issue.
Yes | Visit | Complaint | RAME | SC was installed and startup not completed. | Customer had to replace Sensor cassette. Analyzer did not complete SC startup and anlayzer SW closed and analyzer stuck at Windows screen. | Had the customer reboot the system, Analyzer rebooted the system and startup properly intiated. | SWO-359558 | Re classified RMED | phone support |  | customer reboot system |  | SC was installed and startup not completed. | Customer had to replace Sensor cassette. Analyzer did not complete SC startup and anlayzer SW closed and analyzer stuck at Windows screen. | Had the customer reboot the system, Analyzer rebooted the system and startup properly intiated.
No | Visit | Complaint | RFRA | Erreur 1176. Intervention nécessaire. | Erreur 1176. Intervention nécessaire. Dl OXI >254 | Remplacement du détecteur OXI. Pas d'amélioration. Probablement problème sur hémolyseur. Enlèvement machine ce jour pour expertise. | SWO-359847 | Revisit required |  |  | LS sensor oxi | see SWO-359848 | Error 1176. Intervention required. | Error 1176. Intervention required. Dl OXI - 254 | Replacing the OXI detector. No improvement. Probably a problem with a hemolyser. Removal machine this day for expertise.
Yes | Visit | Complaint | RCHN | 电池没电 | 电池没电 431-018C一个 | 更换431-018C后修复by tim | SWO-358543 | [5 empty] | The battery's dead. | Battery no power 431-018C one | Repair by tim after replacement 431-018C
No | Visit | Complaint | RCHN | 显示器屏幕不亮 | 显示器屏幕不亮，需要借件910-342 | 更换910-342后修复by tim | SWO-358104 | Defect | HW | DMS_PC | User interface module |  | The display screen is not lit | The display screen is not bright, 910-342 is required | Repair by tim after replacing 910-342
Yes | Visit | Complaint | RCHN | 定标时BNP定标卡断裂 | 定标时BNP定标卡断裂，借944-258（14051）一个 | 重新定标后修复by tim | SWO-358622 | [5 empty] | BNP calibration card break stakes at calibration time | BNP calibration card break, borrow 944-258 (14051) one | Fix by tim after re-calibration
No | Visit | Complaint | RCHN | 定标失败，试管支架破裂 | 定标失败，试管支架破裂,需要再加借944-216（14259）、944-214（14220）、944-220（14239）、944-248（14249）、944-394（14070）各一张 | 重新定标后正常by tim | SWO-358136 | Defect | HW | Aspiration Module | Tube holder |  | Failed calibration, test tube support ruptured | Failed to set the standard, the test tube bracket ruptured, need to borrow 944-216 (14259), 944-214 (14220), 944-220 (14239), 944-248 (14249), 944-394 (14070) each | Normal by tim after re-calibration
No | Visit | Complaint | RKK | pCO2 QC測定値がバラつく、Glu CAｌでドリフトが大きい | pCO2 QC測定値がバラつく、Glu CAｌでドリフトが大きい | インレットモジュール 分解・清掃・潤滑 ｐH $BG・El $Met各バルブ分解・清掃・潤滑 エアライン・測定チャンバー清掃 各部調整実施 CAL・QC確認 | SWO-363281 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | PCO2 QC readings are rose, glu CAl drifts high | PCO2 QC readings are rose, glu CAl drifts high | Inlet module disassembly, cleaning and lubrication pH $BG, El $Met each valve disassembly, cleaning, lubrication Airline measurement chamber cleaning Each part adjustment is carried out. CAL/QC Confirmation
Yes | Visit | Complaint | RIND | calibration error. | NA | Cleaned the instrument inner and outer dust. cleaned the process wheel .performed the system clean and system check found values within the range.performed parameter calibration found calibration got passed, did sample found results with in the range. Now instrument working satisfactorily. | SWO-360228 | [5 empty] | calibration error. | In | Cleaned the instrument inner and outer dust. cleaned the process wheel .performed the system clean and system check found values within the range.performed parameter calibration found calibration got passed, did sample found results with in the range. Now instrument working satisfactorily.
No | Visit | Complaint | RDEU | Hotline - Probenabbruch 1067 | Verschluss vom LP öffnete nicht vollständig - abhängig vom LP Fehler RotateProcessWheel_40 mehrere Cups lagen unten im Fach vom LP | Grundplatte vom Gerät demontiert Haken mechanisch justiert mit 2. LP Verschluss kontrolliert Gerät wieder zusammengebaut und QC durchgeführt | SWO-359727 | Defect | Consumables | SP | Cup stack |  | Hotline - Probenabbruch 1067 | Closure of the LP did not open completely - depending on the LP Error RotateProcessWheel_40 several cups were located at the bottom of the compartment of the LP | Base plate disassembled from the device Hook mechanically adjusted controlled with 2nd LP closure Device reassembled and QC performed
No | Visit | Complaint | RFRA | Erreur 1176. Intervention nécessaire. | Erreur 1176. Intervention nécessaire. Erreur présente malgré remplacement du détecteur. | Démontage de l'hemolyseur. Un débris est présent dans la partie transparente de l'hemolyseur au niveau du DL. | SWO-359848 | Cleaned | Fluidic pathway | Fluidic pathway | clot | See SWO-359847 | Error 1176. Intervention required. | Error 1176. Intervention required. Error present despite replacement of the detector. | Disassembly of the hemolyseur. Debris is present in the transparent part of the hemolyseur at the DL.
No | Visit | Complaint | RKK | 検体測定時高混濁後異常を表示する | 状態確認 | ヘモライザー交換実施 | SWO-360357 | Defect | HW | Oxi module | Hemolyzer | turbidity | Displays abnormalities after high turbidity during sample measurement | Status Check | Hemolyzer replacement
No | Visit | Complaint | RIND | Analysis is disable. | Check solution pack cycles and SC remaining test. SP cycles was completed need to replcaed | Replaced Solution pack. | SWO-360225 | User error | SP | SP | Expired SP |  | Analysis is disable. | Check solution pack cycles and SC remaining test. SP cycles was completed need to replcaed | Replaced Solution pack.
Yes | Visit | Complaint | RAME | Error 292 and 581 Contact is Callie at 601-815-4957 | Downloaded data logs for tech line case. Ran Oxi service test and downloaded data. opened tech line case. | Attended Escalation call and determine course of action. Ordered all needed items. | SWO-359585 | [5 empty] | Error 292 and 581 Contact is Callie at 601-815-4957 | Downloaded data logs for tech line case. Ran Oxi service test and downloaded data. opened tech line case. | Attended Escalation call and determine course of action. Ordered all needed items.
No | Visit | Complaint | RIND | SC not calibrating and Solution not coming. | Manifold unit to replace. | Manifold Unit Replaced. Analyser Serviced and calibrated properly. Handed over in a fine working condition. | SWO-360215 | Defect | HW | Basic Unit | Valve |  | SC not calibrating and Solution not coming. | Manifold unit to replace. | Manifold Unit Replaced. Analyser Serviced and calibrated properly. Handed over in a fine working condition.
Yes | Visit | Complaint | RIND | electrolytes results are high in all samples. | Checked by zeroing procedure and found that it is ok. Then cross checked with other lab ABL80 samples. | Need to change sensor cassette. | SWO-359867 | [5 empty] | electrolytes results are high in all samples. | Checked by zeroing procedure and found that it is ok. Then cross checked with other lab ABL80 samples. | Need to change sensor cassette.
No | Visit | Complaint | RCHN | 光学单元超出范围 | 光学单元超出范围，需要借件903-364 | 更换903-364后修复by tim | SWO-358129 | Defect | HW | Process Module | Optical Unit |  | Optical unit out of range | Optical unit out of range, need to borrow 903-364 | Repair by tim after replacing 903-364
No | Visit | Complaint | RCHN | 进样盘扫描器扫不了 | 进样盘扫描器扫不了 | 更换910-339后修复by tim | SWO-359731 | Defect | HW | Aspiration Module | Barcode reader, Inlet |  | The sample disk scanner can't be swept. | The sample disk scanner can't be swept. | Repair by tim after replacing 910-339
No | Visit | Complaint | RCHN | 定标中止，报错1061 | 定标中止，报错1061，需要借件944-258（14051） |  | SWO-359643 | User error | User error | User error | User error |  | Bid suspension, error 1061 | Bid suspension, error 1061, need to borrow 944-258 (14051)
No | Visit | Complaint | RCHN | 定标中止，报错1993.-1 | 定标中止，报错1993.-1,需要借件944-394（14170） | 重新定标后修复by tim | SWO-358614 | Defect | SW | Analyzer Control Module |  |  | Bid suspension, error 1993.-1 | Bid suspension, error 1993.1, need to borrow 944-394 (14170) | Fix by tim after re-calibration
No | Visit | Complaint | RAME | Wont pass QC, getting an error Contact is Craig Noel at 314-517-2389 PO# CE206787 | Error 0741 autocheck ampoules house failure. | Replaced defective carrousel tilt stepping motor. | SWO-360365 | Defect | HW | Autocheck Module | Stepper motor | 0741 | Wont pass QC, getting an error Contact is Craig Noel at 314-517-2389 PO# CE206787 | Error 0741 autocheck ampoules house failure. | Replaced defective carrousel tilt stepping motor.
No | Visit | Complaint | RKK | pH Cal error | pH Cal error | pH電極より液漏れを起こしていたため、測定チャンバー掃除。ｐHシーリングリング交換。 | SWO-360341 | Defect | Electrode | Electrode | Sealing ring, pH |  | pH Cal error | pH Cal error | Because the liquid leakage was caused from the pH electrode, the measurement chamber cleaning. pH sealing ring replacement.
No | Visit | Complaint | RAME | Ph calibration failure, errors 375,376,377,378 | Follow up to SWO-358550. | Installed replacement pH electrode, performed calibration until 2PT passed. Ran QC, all levels in range. | SWO-361653 | Assisted Replacement | Electrode | Electrode | pH | revisit | Ph calibration failure, errors 375,376,377,378 | Follow up to SWO-358550. | Installed replacement pH electrode, performed calibration until 2PT passed. Ran QC, all levels in range.
No | Visit | Complaint | RAME | Neon Lamp intensity too low | Customer reported Neon lamp too low error 316, 258, Oxi Temp error, 259. Performed applicable service diagnostics, Complete Oxi test failed for Neon Lamp Intensity @ 25%, position @ -1.000. At the conclusion of the Oxi module testing I began to observe numerous temp error for various module, 354,259,270,355, 943, 356. AU disconnect $reconnect errors, 643, 642. Suspect possible marginal power supply as cause of multiple errors. Also observed leak detected error. | Replaced Oxi module, performed all applicable tests. Complete Oxi test now passing, performed service and user tHb cal. Replace power supply as possible cause of numerous errors linked with AU disconnection $reconnect errors. No further temp or AU errors observed after PS was replaced. Leak detected error caused by using worn reused inlet gasket. replaced with new inlet gasket. Recommended customer replace the inlet gasket periodically. They are currently cleaning them by soaking in an unknown cleaner . During final validation Crea failed L2 QC Crea A and calibration sensitivity 3. Replaced all solutions and remembraned Crea A & B. | SWO-359642 | Defect | HW | Oxi module | Oxi module |  | Neon Lamp intensity too low | Customer reported Neon lamp too low error 316, 258, Oxi Temp error, 259. Performed applicable service diagnostics, Complete Oxi test failed for Neon Lamp Intensity @ 25%, position @ -1.000. At the conclusion of the Oxi module testing I began to observe numerous temp error for various module, 354,259,270,355, 943, 356. AU disconnect /reconnect errors, 643, 642. Suspect possible marginal power supply as cause of multiple errors. Also observed leak detected error. | Replaced Oxi module, performed all applicable tests. Complete Oxi test now passing, performed service and user tHb cal. Replace power supply as possible cause of numerous errors linked with AU disconnection $reconnect errors. No further temp or AU errors observed after PS was replaced. Leak detected error caused by using worn reused inlet gasket. replaced with new inlet gasket. Recommended customer replace the inlet gasket periodically. They are currently cleaning them by soaking in an unknown cleaner . During final validation Crea failed L2 QC Crea A and calibration sensitivity 3. Replaced all solutions and remembraned Crea A & B.
No | Visit | Complaint | RDEU | Mischerträger Schlupf | Sporadisch Meldung 882 Mischerträger Schlupf erkannt . | Das FlexQ-Modul wird nicht genutzt. FlexQ-Modul diskonnektiert. Funktionskontrolle: I.O. | SWO-359889 | Defect | HW | FlexQ | FlexQ |  | Mixer carrier slip | Sporadic message 882 mixer carrier slip detected. | The FlexQ module is not in use. FlexQ module discounted. Functional control: I.O.
No | Visit | Complaint | RAME | THb failing calibration | Lamp was above. 4000 in the service test. Found lamp tightened down | Adjusted lamp to 3198, ran a complete oxy test, Ran all test | SWO-359588 | Adjustment | HW | Oxi module | Lamp unit |  | THb failing calibration | Lamp was above. 4000 in the service test. Found lamp tightened down | Adjusted lamp to 3198, ran a complete oxy test, Ran all test
No | Visit | Complaint | RIND | HIS/LIS connection error. | HIS $LIS connection error,data not able transmit. | Check the LAN cable connection & found that it was not connected properly ,So reconnect the LAN connection & put the server IP address manually & then data transmitted successfully. | SWO-360129 | User error | HW | Analyzer Control Module | PC |  | HIS/LIS connection error. | HIS /LIS connection error,data not able transmit. | Check the LAN cable connection & found that it was not connected properly ,So reconnect the LAN connection & put the server IP address manually & then data transmitted successfully.
No | Visit | Complaint | RKOR | Sample data error | . | re sampling!! | SWO-364244 | User error | preanalytical | preanalytical | preanalytical |  | Sample data error | . | re sampling!!
Yes | Visit | Complaint | RIND | Po2 not calibrating | Po2 not calibrating | As chekced the Po2 sense values it was not in the range cleaned the po2 electrode with brush and refixed calibrated but still same problem repeated , it was new membrane was in the po2 electrode , so suspected the po2 electrode so needs to becross chekced the po2 electrode with new one | SWO-360406 | [5 empty] | Po2 not calibrating | Po2 not calibrating | As chekced the Po2 sense values it was not in the range cleaned the po2 electrode with brush and refixed calibrated but still same problem repeated , it was new membrane was in the po2 electrode , so suspected the po2 electrode so needs to becross chekced the po2 electrode w ith new one
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 不均匀的冲洗溶液 1324 | 更换试剂包后修复 | SWO-360288 | Assisted Replacement | SP | SP | FT | 1324 | Test card, reagent package failure | Uneven flushing solution 1324 | Repair after replacement of reagent kit
No | Visit | Complaint | RCHN | 漏电流 | 漏电流，借902-980一个 |  | SWO-360297 | Defect | HW | Oxi module | Hemolyzer | Leak current | Leak current | Leak current, borrow 902-980 one
No | Visit | Complaint | RDEU | Hotline - Probenabbruch 1067 | Trennsteg vom Storage Wheel war ausserhalb der Halterung nach Einsetzen einer Cass. verklemmte sich das Rad mit der Meldung 5689 Rotate Storage Wheel 24 | Steg wieder eingesetzt Positionen des Storage Wheels und des Cardridge Sliders im Service überprüft | SWO-360321 | Defect | HW | Storage Module | Storage wheel | cartridge guide | Hotline - Probenabbruch 1067 | Separation bar from the Storage Wheel was outside the bracket after inserting a cass. the wheel jammed with the message 5689 Rotate Storage Wheel 24 | Bridge reinstated Checked the position of the Storage Wheel and Cardridge Slider in the service
No | Visit | Complaint | RLTD | Mark 01895828757 | Battery icon showing battery at zero % with mains power connected. When mains disconnected Battery works ok but battery icon shows zero %. | Discharged battery for 15 minuets. Removed and resets battery status using jumper connector. Reinstalled and now battery showing charging up. | SWO-360954 | Adjustment | HW | Battery module | Battery module | reset | Mark 01895828757 | Battery icon showing battery at zero % with mains power connected. When mains disconnected Battery works ok but battery icon shows zero %. | Discharged battery for 15 minuets. Removed and resets battery status using jumper connector. Reinstalled and now battery showing charging up.
Yes | 6 | Y | RPAC | Blank | Solution pack failed |  | D46E2618EE419F37 | [5 empty] | Blank | Solution pack failed
No | Visit | Complaint | RSCH | Rep 1190 | Bringt Fehler 1190 bei Messungen. Seit SW Upgrade verhäuft | Einlass nocheinmal kalibriert Servicedump gezogen Probemessungen mit kapilaren i.o Funktionskontrolle i.o Kein Fehler festgestellt | SWO-360844 | Adjustment | HW | Auto Inlet module | Probe Cal |  | Rep 1190 | Brings error 1190 for measurements. Since SW Upgrade, it has piled up | Inlet calibrated again Service dump pulled Test measurements with kapilaren i.o Functional control i.o No error detected
No | Visit | Complaint | RFRA | Erreur 1176. Intervention nécessaire. | Remise en route de l'ABL90 sur site. | Procédure de mise en route 50 minutes. Calibration 2 points, calibration Thb et passage des Cq 3 niveaux ok. | SWO-360816 | Adjustment | SW | SW | off/on | 1176 | Error 1176. Intervention required. | ABL90 is restarted on site. | Start-up 50 minutes. 2-point calibration, Thb calibration and passage of Cq 3 levels ok.
No | Visit | Complaint | RSAF | Call Service engineer, analyzer blocked and unable to do back flush | Unable to do back flush as sample path is blocked, worked on flushing the analyzer segment by segment and located blockage in the inlet probe, | Replaced inlet probe, done complete fluid transport test and passed with new inlet probe, | SWO-367104 | Defect | HW | Auto Inlet module | Probe | clot | Call Service engineer, analyzer blocked and unable to do back flush | Unable to do back flush as sample path is blocked, worked on flushing the analyzer segment by segment and located blockage in the inlet probe, | Replaced inlet probe, done complete fluid transport test and passed with new inlet probe,
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 客户反映有个试剂包用了五天就完了，经查看，溶液2剩余不足，可能Sendx在写入信息时溶液2写入数量小于标称容量 | 待技术支持分析数据后给予质保 | SWO-360461 | [5 empty] | Test card, reagent package failure | Customer reported that there is a reagent package took five days to finish, after viewing, solution 2 is insufficient, may Sendx when writing information solution 2 write quantity less than the nominal capacity | Warranty to be given after technical support analyzes data
No | Visit | Complaint | RDEU | die RiLiBÄK wird nicht kontinuierlich eingehalten | QC-Einstellungen korrekt, L2 OK, L1 pCO2 sehr häufig zu hoch, L3 pO2 zu tief | keine mit positivem Ergebnis, Erneuerung SP ohne Besserung neues QC-Material bestellen, QC nur schwer in den Bereich zu bekommen | SWO-360799 | [5 empty] | the RiLiBÄK is not continuously adhered to | QC settings correct, L2 OK, L1 pCO2 very often too high, L3 pO2 too deep | none with positive result, renewal SP without improvement order new QC material, QC difficult to get into the area
No | Visit | Complaint | RIND | pO2 values not OK. | pO2 Values Not Ok. | pO2 mambrane asked. Till the time. tried with reinstalling and cleaning of electrodes and chembers. Electrode cleaning procedure and Inlet probe and inlet gasked cleaned. pO2 mam,brane replaced and calibrated. Analyser Serviced and calibrated properly. Handed over in a fine working condition. | SWO-363852 | Assisted Replacement | Electrode | Membrane | pO2 |  | pO2 values not OK. | pO2 Values Not Ok. | pO2 mambrane asked. Till the time. tried with reinstalling and cleaning of electrodes and chembers. Electrode cleaning procedure and Inlet probe and inlet gasked cleaned. pO2 mam,brane replaced and calibrated. Analyser Serviced and calibrated properly. Handed over in a fine working condition.
No | Visit | Complaint | RIND | Blank Display, nothing is visible on the display. | Need to change the CPU unit | Changed the CPU unit now the analyzer working satisfactorily. | SWO-361208 | Defect | HW | DMS_PC | CPU |  | Blank Display, nothing is visible on the display. | Need to change the CPU unit | Changed the CPU unit now the analyzer working satisfactorily.
Yes | Visit | Complaint | RDAN | Lene ringer, meddeler, at hun ikke kan få tHb på plads - måler for lavt. | . | Cuvette renset med protein remover, kalibreret, tHb kalibreret, 4xQC, ok. | SWO-360793 | [5 empty] | Lene rings, announces that she cannot get tHb in place-gauge too low. | . | Cuvette renset med protein remover, kalibreret, tHb kalibreret, 4xQC, ok.
No | Visit | Complaint | RDEU | 04.06.2018 10:15Uhr Frau Kirdudu 071562037655: Das Gerät meldet Intervention erforderlich. | Keine Messung oder Kal möglich, da Spülfehler oder Lös zu früh an LS | Gerät überprüft. BG Valves ausgebaut und gereinigt. Funktion ok. Messkammer gereinigt Dichtigkeitskontrolle ok. Pump und LS Kalibration durchgeführt ok. Fluidic Modul Valves in Ordnung. Fehler beseitigt Kalibration wieder möglich. PO2 Elektrode Empf. Zu tief Nullwert zu hoch. PO2 Elektrode ist def. Elektrode aus Kundenbestand Einkalibriert Kalibration und QC durchgeführt und in Ordnung | SWO-361913 | Assisted Replacement | Electrode | Electrode | pO2 |  | 06/04/2018 10:15AM Ms Kirdudu 071562037655: The device reports intervention required. | No measurement or calving possible, as flushing errors or loess too early on LS | device checked. BG Valves removed and cleaned. Function ok. Measuring chamber cleaned tightness control ok. Pump and LS calibration performed ok. Fluidic module Valves fine. Error eliminates calibration possible again. PO2 Electrode Empf. Too low zero value too high. PO2 electrode is def. Electrode from customer stock Calibrated Calibration and QC carried out and in order
No | Visit | Complaint | RAME | Neon intensity too low. | Error still present when I arrived. | I tried to adjust the lamp several times, but it would not pass the oxy test. I will need to order a replacement lamp and return to replace it. Staff has been informed. | SWO-360965 | Revisit required | HW | Oxi module | Lamp unit | neon | Neon intensity too low. | Error still present when I arrived. | I tried to adjust the lamp several times, but it would not pass the oxy test. I will need to order a replacement lamp and return to replace it. Staff has been informed.
Yes | Visit | Complaint | RLTD | Frank Watson on 01983 534810 | Removed covers and retrieved both wheels and one securing pin for the inlet insert, unable to locate 2nd securing pin. Spare securing pin supplied from engineer stock. Re-assembled and refitted both inlet wheel inserts. Analyser was also not recognising removal of tubes and getting stuck on this screen. Cause found to be blood splash on the tube sensor. Cleaned sensor and tested - now operating and recognising tube removal. Analyser ready for use. | Found and refitted parts to inlet wheel inserts. Cleaned tube sensor. | SWO-360741 | [5 empty] | Frank Watson on 01983 534810 | Removed covers and retrieved both wheels and one securing pin for the inlet insert, unable to locate 2nd securing pin. Spare securing pin supplied from engineer stock. Re-assembled and refitted both inlet wheel inserts. Analyser was also not recognising removal of tubes and getting stuck on this screen. Cause found to be blood splash on the tube sensor. Cleaned sensor and tested - now operating and recognising tube removal. Analyser ready for use. | Found and refitted parts to inlet wheel inserts. Cleaned tube sensor.
No | Visit | Complaint | RDEU | keine Remoteverbindung möglich | keine Remoteverbindung möglich | IP Adressen überprüft und IP eingetragen, Gerät hat Verbindung und ist online. | SWO-361488 | Setup | SW | SW | Remote | IP-address | No remote connection possible | No remote connection possible | IP addresses checked and IP entered, device has connection and is online.
No | Visit | Complaint | RDEU | Labor meldet fehlerhafte QC-Werte bei der Auswertung Zielwerte der Auswertung stimmen nicht mit den Zielwerten im Gerät überein | Datenbank geprüft Zielwerte der Auswertungsdatei ( Exel ) stimmen nicht mit den Zielwerten am ABL überein | Eintragung fixe Zielwerte fehlt --> Eintragung angepasst und fixe Zielwerte aktiviert | SWO-361932 | Setup | SW | Roma.ini | Fixed assigned values |  | Laboratory reports faulty QC values during evaluation Target values of the evaluation do not match the target values in the device | Database checked Target values of the evaluation file ( Exel ) do not match the target values on the ABL | Registration fixed target values are missing --> Registration adjusted and fixed target values activated
No | Visit | Complaint | RLTD | Jennifer 07872417539 | Analyser passing fluid transport tests and entered start up. At end of start analyser showing leak detected. Inspected the pinch valve tubing which showed signs of a leak. | Replaced the pinch valve tubing. | SWO-361391 | Defect | HW | Oxi module | Pinch valve tube |  | Jennifer 07872417539 | Analyser passing fluid transport tests and entered start up. At end of start analyser showing leak detected. Inspected the pinch valve tubing which showed signs of a leak. | Replaced the pinch valve tubing.
Yes | 1 | Y | RPAC | Blank | Software Upgrade analyser to Version V3.4. |  | D46F817D6415D878 | [5 empty] | Blank | Software Upgrade analyser to Version V3.4.
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 全参数定标失败 lac ph pco2 po2 na ca cl hct 定标不过 | 更换测试卡后修复 | SWO-361197 | [5 empty] | Test card, reagent package failure | Full parameter calibration failure lac ph pco2 po2 na ca cl hct calibration but | Repair after replacing test card
No | Visit | Complaint | RIND | All Parameter is inactive. | First Checked Calibration Report to identify error. find out that sensor Cassette is not working. | Removed sensor cassette and flush it with heparin solution. then keep it in refrizarator for hydration around 2 hr. then removed it and installed again properly. but still some parameter like pH, Na, K, Ca and Cl was inactive. removed solution pack and again reinstalled it but still parameter was inactive. need to replaced Sensor Cassette. | SWO-363015 | [5 empty] | All Parameter is inactive. | First Checked Calibration Report to identify error. find out that sensor Cassette is not working. | Removed sensor cassette and flush it with heparin solution. then keep it in refrizarator for hydration around 2 hr. then removed it and installed again properly. but still some parameter like pH, Na, K, Ca and Cl was inactive. removed solution pack and again reinstalled it but still parameter was inactive. need to replaced Sensor Cassette.
No | Visit | Complaint | RCHN | 显示器无法显示 | 显示器无法显示 | 重新连接电源线后修复(By Ran) Reconnected the cable. | SWO-361175 | [5 empty] | The monitor can't display | The monitor can't display | Repair (By Ran) Reconnect the cable after reconnecting the power cord.
No | Visit | Complaint | RCHN | 气体流量超范围 | 气体流量超范围，借903-111一个。 | 更换903-111一个正常。(by chuanwei) | SWO-360506 | [5 empty] | Gas flow over range | The gas flow is over range, borrowing 903-111. | Replace 903-111 one normal. (by chuanwei)
No | Visit | Complaint | RCHN | GLU,LAC定标失败 | GLU,LAC定标失败 ，借903-581一个。 | 更换903-581后修复(By Ran) | SWO-360387 | Defect | HW | Basic Unit | PCB, Sensor interface | Glu/Lac | GLU, LAC calibration failed | GLU, LAC calibration failed , borrow edith of 903-581. | Repair after replacement 903-581 (By Ran)
No | Visit | Complaint | RCHN | 无法读取测试卡扫描器 | 无法读取测试卡扫描器 | 调整位置 。（by chuanwei） | SWO-361348 | Defect | HW | Storage Module | Barcode reader, storage wheel |  | Failed to read test card scanner | Failed to read test card scanner | Adjust the position. (by chuanwei)
No | Visit | Complaint | RCHN | 定标失败 1031 | 定标失败 1031，借903-364一个。 | 更换903-364备件后修复(by chuanwei) | SWO-360527 | Defect | Hw | Process Module | Optical Unit | 1031 | Bid Failure 1031 | Bid failure 1031, borrowing 903-364 one. | Repair after replacement of 903-364 spare parts (by chuanwei)
No | Visit | Complaint | RIND | No printout | Checked and found printer unit cover is broken. | Printer unit requires replacement to rectify the error. | SWO-363569 | [5 empty] | No printout | Checked and found printer unit cover is broken. | Printer unit requires replacement to rectify the error.
Yes | 7 | Y | RPAC | Blank | PO2 zero current error |  | D46F80AB642DD041 | [5 empty] | Blank | PO2 zero current error
No | Visit | Complaint | RIND | Option key installation & continuous leakage from gasket. | Option key need to install & continuous leakage from gasket. | Clean the gasket & waste pump tubes,run the LS adjust & pump calibration,leakage was stopped,then install the option key successfully. | SWO-363866 | Cleaned | HW | Inlet module | Inlet gasket |  | Option key installation & continuous leakage from gasket. | Option key need to install & continuous leakage from gasket. | Clean the gasket & waste pump tubes,run the LS adjust & pump calibration,leakage was stopped,then install the option key successfully.
Yes | Visit | Complaint | RDEU | 04.06.2018, 8:45 Uhr, Anrufer Herr Kammermeier, Tel.: 08161/243540: Betriebsmittel können nicht gescannt werden | Fehlersuche vor Ort | Demo keys abgelaufen, Neue Demo-keys installiert. Laufen am 08.09.2018 ab. Mehrere Kalibrierungen und QCs für alle Level durchgeführt. ABL wieder messbereit und an Conworx angebunden. Herrn Kammermeier und Herrn Lubig darüber informiert | SWO-361581 | Re classified RMED | demo key expired |  |  |  | 04.06.2018, 8:45 a.m., Caller Mr. Kammermeier, Tel.: 08161/243540: Equipment cannot be scanned | Troubleshooting on site | Demo keys expired, New demo keys installed. Run on 08.09.2018. Multiple calibrations and QCs performed for all levels. ABL ready for measurement again and connected to Conworx. Mr Kammermeier and Mr Lubig informed
No | Visit | Complaint | RCHN | po2定标失败，报错0408 | po2定标失败，报错0408，需要借件945-613和903-652 | 清洁气体阀。(by chuanwei) | SWO-361176 | Defect | HW | pH/BG Module | PCB, ph/BG |  | po2 bid failure, error 0408 | Po2 bid failure, error 0408, need to borrow 945-613 and 903-652 | Clean the gas valve. (by chuanwei)
No | Visit | Complaint | RKK | キャリブレーションが安定しない | 状態確認 | 部品交換実施 | SWO-363190 | Defect | HW | gas module | Flow regulator |  | Calibration is not stable | Status Check | Parts replacement
No | Visit | Complaint | RAME | Neon intensity too low. | Returned with the new lamp unit to replace suspected faulty unit. | I replaced the lamp unit and performed subsequent procedures according to manufacturer. Took a couple of adjustments but it passed the oxy test along with the tHb calibration. I also had to run several calibrations afterwards but the device settled into operating ranges. I ran the QC as well. | SWO-361658 | Defect | HW | Oxi module | Lamp unit | neon | Neon intensity too low. | Returned with the new lamp unit to replace suspected faulty unit. | I replaced the lamp unit and performed subsequent procedures according to manufacturer. Took a couple of adjustments but it passed the oxy test along with the tHb calibration. I also had to run several calibrations afterwards but the device settled into operating ranges. I ran the QC as well.
No | Visit | Complaint | RDEU | 05.06.2018 11:30uhr Frau Glomdick 0212645645434: Das Gerät bricht die Kal ab | Leckage Fehler | PO2 Membran beschädigt (leer) Messkammer gereinigt, PO2 Neu membraniert, 2 Punkt kal und Qc's bestanden | SWO-361377 | Assisted Replacement | Electrode | Membrane | pO2 |  | 06/05/2018 11:30am Ms Glomdick 0212645645434: The device breaks off the calf | Leakage error | PO2 membrane damaged (empty) measuring chamber cleaned, PO2 re-membraneed, 2 point kal and Qc's passed
No | Visit | Complaint | RDEU | Beim Scannen wird eine Ziffer dazu erfunden, die muss man von Hand löschen, dann erscheinen auch die Pat. Daten. Bitte um Problembehebung | Gerät überprüft, Bar Code ist nicht am Gerät konfiguriert. | Barcode Konfiguriert, Neustart, Funktionsprüfung i.O. Anwender informiert. | SWO-361561 | Setup | SW | Basic Unit | Barcode reader | mail | During scanning, a digit is invented, which must be deleted by hand, then the pat also appear. Data. Please fix the problem | Device checked, bar code is not configured on the device. | Barcode Configured, restart, function check i.O. informed users.
No | Visit | Complaint | RDEU | 04.06.2018 16:20Uhr Frau Schenk 09519801422: keine Netzverbindung 04.06.2018; 16:45 Uhr; Anruferin Frau Schenk Tel.: 09519801422: Fehler ist nach AQT und Serverneustart noch immer vorhanden | Siehe auch SWO-360692 --> hier wurde bei der Wartung noch etwas mit ASTM.net überprüft und da wurde die IP-Adresse am Gerät zum testen verstellt und nach dem Zurückstellen vergessen den Speicherbefehl auszuführen | Netzwerkkonfiguration auf Windows-Ebene abgeändert, Speicherbefehl ausgeführt und Analysator neu gestartet --> danach war die averbindung zu Link2Lab wieder hergestellt, System und Funktion i.O. | SWO-361278 | User error | HW | Analyzer Control Module | CPU | IP-address | 06/04/2018 16:20pm Mrs Schenk 09519801422: no mains connection 04.06.2018; 4:45 p.m.; Caller Mrs. Schenk Tel.: 09519801422: Error still exists after AQT and server restart | See also SWO-360692 --> here a little bit was checked with ASTM.net during maintenance and since the IP address on the device was adjusted for testing and after the reset forgot to execute the memory command | Network configuration changed at the Windows level, memory command executed and analyzer restarted --> after that the connection to Link2Lab was restored, system and function i.O.
No | Visit | Complaint | RDEU | 28.05.2018; Fr. Sannemohn; Tel.: 06213836751; Hbf und sO2 QC Messfehler 28.05.2018; 14:35 Uhr; Anruferin Fr. Salomon; Tel.: 06213833392; Wieder Fehler 1034 Oxi Hardware Problem | 1020 $ 1034 Oxi Hardware Problem | Allgemeine Überprüfung - Analysator gespült - Oximetrie Test durchgeführt (Neon Position Fehler = -10,00) - Erneuerung Hemolysator-Einheit - Erneuerung Einlassdichtung - Funktionsüberprüfung, Oximetrie Test, Kalibrationen und Qualitätskontrollen ohne Beanstandung | SWO-362084 | Defect | HW | Oxi module | Hemolyzer |  | 28.05.2018; Ms. Sannemohn; Tel.: 06213836751; Hbf and sO2 QC measurement error 28.05.2018; 2:35 p.m.; Caller Ms. Salomon; Tel.: 06213833392; Again Error 1034 Oxi Hardware Problem | /1020 1034 Oxy Hardware Problem | General review - Analyzer rinsed - Oximetry test performed (neon position error = -10.00) - Renewal hemolyser unit - Renewal inlet seal - Function check, oximetry test, calibrations and quality controls without complaint
No | Visit | Complaint | RSCH | Rep. Einlass | Messungen werden sporadisch abgebrochen Meldung: keine Probe während Ansaugung erkannt In Aktivitäten: 1235 Proneninspiration gescheitert 1189 FTC abgebrochen, Fehler LS Zustand Seit 05.01.2018 ca. 60x | Im Service-Menu Fluid alles i.o Drei Testmessungen i.o Bei vierter Messung konnte Fehler reproduziert werden Gerät erkennt nicht bei Spritzenposition dass Aspiriert werden soll, obwohl Spritze bis ganz nacht hinten geschoben wurde. Schiebt man die Spritze etwas zurück und wieder nach hinten, erst danach saugt das Gerät an. Im Service-Menu unter Miscellaneuous werden Einlass-Positionen korrekt angezeigt, auch Aspirierung-Position Ersatzgerät von MedTech in Abteilung installiert i.o Kunden Gerät wird in Thalwil überprüft Inlet position detector PCB ersetzt i.o 10 Testmessungen i.o Weitere Messungen werden am 07.06.2018 durchgeführt | SWO-362153 | Defect | HW | Auto Inlet module | PCB, Auto inlet | HALL Element | Rep. Intake | Measurements are aborted sporadically Message: no sample detected during suction In activities: 1235 Proneninspiration failed 1189 FTC Aborted, Error LS State Since 05.01.2018 approx. 60x | In the service menu Fluid everything i.o Three test measurements i.o Errors could be reproduced in fourth measurement Device does not detect when syringe position is placed that aspirated, although syringe was pushed back until all night. If you push the syringe back a little and back again, only then sucks in the device. In the service menu under Miscellaneuous, inlet positions are correctly displayed, also aspiration position Replacement device installed by MedTech in department i.o Customer device is checked in Thalwil Inlet position detector PCB replaces i.o 10 test measurements i.o Further measurements will be carried out on 07.06.2018
No | Visit | Complaint | RLTD | Jenny Forsett 0203 447 0310 | C02 results out side reportable range. | Checked logs found blood gases results out of reportable rang above 14.7. Advised customer that they can change range and make it higher but it will not be validated so they will need to validate the change them self. Customer will validate the change they make. | SWO-362318 | No fault found | No fault found | No fault found | No fault found | out of reportable range | Jenny Forsett 0203 447 0310 | C02 results out side reportable range. | Checked logs found blood gases results out of reportable rang above 14.7. Advised customer that they can change range and make it higher but it will not be validated so they will need to validate the change them self. Customer will validate the change they make.
No | Visit | Complaint | RFRA | ERR 1401 POSITION JOINT D'INTRODUCTION | Contrôle du site d'introduction | Remplacement du joint de connexion et du joint d'introduction $Contrôle du site d'introdcution dans le menu service $Contrôle de toutes les positions du site d'introduction $Calibration du site d'introduction $Resserage des vis du moteur de la vis sans fin du site d'introduction. | SWO-362059 | Adjustment | HW | Auto Inlet module | Inlet drive gear |  | ERR 1401 POSITION JOINT D'INTRODUCTION | Control of the introductory site | Replacing the connection seal and the introduction seal $Controf the introdcution site in the service menu $Controf all positions of the introduction site $Calibration of the introduction site $Resserage screws of the site's endless screw engine introduction.
Yes | Visit | Complaint | RIND | All Parameter is inactive. | Sensor cassette and solution pack was need to replaced. | Replaced sensor Cassette with new one. but still PCo2 was inactive. after that I have given 2 point calibration. but still pCO2 was inactive. then replaced new solution pack. after that PCO2 was active. handed over to fine working condition. | SWO-363027 | [5 empty] | All Parameter is inactive. | Sensor cassette and solution pack was need to replaced. | Replaced sensor Cassette with new one. but still PCo2 was inactive. after that I have given 2 point calibration. but still pCO2 was inactive. then replaced new solution pack. after that PCO2 was active. handed over to fine working condition.
No | Visit | Complaint | RIND | Electrolyte results not reliable | Checked calibration and patient logs for possible error messages. | Flushed sensor cassette with NS solution. Performed zeroing of all channels successfully. 2 point calibration performed without errors. | SWO-363604 | [5 empty] | Electrolyte results not reliable | Checked calibration and patient logs for possible error messages. | Flushed sensor cassette with NS solution. Performed zeroing of all channels successfully. 2 point calibration performed without errors.
Yes | Visit | Complaint | RAME | Error 292 and 581 Contact is Callie at 601-815-4957 | Patient samples are experiencing errors 581 & 688, sometimes along with a 292. Checked the line voltage in the dept and it was found to be 120.4 VAC,Ran 15 samples of blood with the analyzer in the point of care dept with no errors. Took the analyzer back to the cath lab to test there. | The line voltage in the cath lab was 121 VAC, continued running blood samples and was able to duplicate the errors only after 8 successful samples with no errors. I intentionally introduced air in the samples and was able to duplicate the errors this way. Also the temperature in the cath lab was 16C. At this time the data logs have been collected and will be forwarded to Matt Rutrough. The CO-OX board will be replaced to rule it out while Matt Rutrough reviews the data | SWO-361838 | [5 empty] | Error 292 and 581 Contact is Callie at 601-815-4957 | Patient samples are experiencing errors 581 & 688, sometimes along with a 292. Checked the line voltage in the dept and it was found to be 120.4 VAC,Ran 15 samples of blood with the analyzer in the point of care dept with no errors. Took the analyzer back to the cath lab to test there. | The line voltage in the cath lab was 121 VAC, continued running blood samples and was able to duplicate the errors only after 8 successful samples with no errors. I intentionally introduced air in the samples and was able to duplicate the errors this way. Also the temperature in the cath lab was 16C. At this time the data logs have been collected and will be forwarded to Matt Rutrough. The CO-OX board will be replaced to rule it out while Matt Rutrough reviews the data
No | Visit | Complaint | RKK | 温度エラーにて、基板交換 |  |  | SWO-361826 | [5 empty] | Replacement of board due to temperature error
No | Visit | Complaint | RAME | CO-OX control board failure | Checked data logs for errors, intermittent Oxi control board failure | Removed wet section and replaced CO-OX control board, spine board and power cable, performed Tikhonov matrix | SWO-362480 | [5 empty] | CO-OX control board failure | Checked data logs for errors, intermittent Oxi control board failure | Removed wet section and replaced CO-OX control board, spine board and power cable, performed Tikhonov matrix
No | Visit | Complaint | RKK | SP認識しない。 SPとラッチの間に紙を挟むと認識する。 | SPの押し込み不足 | SP固定ﾗｯﾁの受け部の位置を調整。 | SWO-364092 | [5 empty] | SP not recognized. Recognize when you sandwich the paper between the SP and the latch. | Insufficient SP push | Adjusted the position of the receiving part of the SP securing latch.
No | Visit | Complaint | RAME | Customer gets intermittent fluid blockage errors with question marks on oximetry values when they run external QC | Checked data logs, found errors on external QC | Replaced hemolyzer for troubleshooting purposes. Ran tHb calibration. | SWO-362350 | [5 empty] | Customer gets intermittent fluid blockage errors with question marks on oximetry values when they run external QC | Checked data logs, found errors on external QC | Replaced hemolyzer for troubleshooting purposes. Ran tHb calibration.
No | Visit | Complaint | RCHN | 报错5784 | 报错5784，需要借件580-098和903-031 | 更换580-098和903-031正常。(by chuanwei) | SWO-360421 | Defect | HW | Process Module | Process Wheel | 5784 | 5784 | 5784, 580-098 and 903-031 | Replacement s580-098 and 903-031 are normal. (by chuanwei)
No | Visit | Complaint | RCHN | 报错1024 | 报错1024,需要借件903-364和944-212（14374） | 更换903-364一个，重新定标正常。(by chuanwei) | SWO-361266 | Defect | HW | Process Module | Optical Unit | 1024 | 1024 error | 1024, need to borrow 903-364 and 944-212 (14374) | Replace the 903-364 one and re-label normally. (by chuanwei)
No | Visit | Complaint | RCHN | 报错5784 | 报错5784 | 取出卡盘的测试卡。（by chuanwei） | SWO-361863 | Defect | Consumables | Test Kit | Cartridge, Broken | 5784 | 5784 | 5784 | Remove the test card of the card tray. (by chuanwei)
No | Visit | Complaint | RDEU | die RiLiBÄK wird nicht kontinuierlich eingehalten | heute war die QC i.O. | Aktuell keine, Vermutung: Die QC-Ergebnisse sind bekanntlich von der Temperatur abhängig, die im Laborraum stark schwankt, da es keine Klimaanlage gibt (unter kalifornischen Bedingungen mit Vollklimatisierung sicherlich kein Problem) | SWO-362524 | [5 empty] | the RiLiBÄK is not continuously adhered to | heute war die QC i.O. | Currently no, guess: The QC results are known to depend on the temperature, which varies greatly in the laboratory room, as there is no air conditioning (certainly no problem under California conditions with full air conditioning)
No | Visit | Complaint | RKK | Naドリフト | Naドリフト | メンブラン交換しても安定せず、Ref電極交換 | SWO-361801 | Assisted Replacement | Electrode | Electrode | Reference |  | Na Drift | Na Drift | It is not stable even if I change the membrane, and the Ref electrode exchange
No | Visit | Complaint | RDEU | 06.06.10:45 Uhr, Herr Küchelbacher, Tel.: 06213831024: ABL zeigt verschieden Fehlermeldungen an | Herr Küchelbacher berichtet telefonisch : - QC-Flüssigkeit wird nicht ins pH $BG Modul transportiert und pO2-Elektrode entweicht Elektrolyt-Lösung aus Membran ! Bei Eintreffen meldet Analysator Kalibrations und QC-Fehler bei pH, pO2, pCO2 und Elektrolyten (teilweise keine Werte) 501 Spüllösung nicht angesaugt im pH $BG Modul innerhalb Zeitrahmen (seit 4.6.18 , ca. 13 Uhr) | Überprüfung Messventile pH $BG Modul im Service (ohne Beanstandung) - Überprüfung Messkammer und Elektroden im pH $BG Modul - pO2 Elektrode mechanisch defekt (komplette Elektrodenspitze nicht vorhanden $abgerissen) - Erneuerung pO2 Elektrode - Funktionsüberprüfung etc. ohne Beanstandung ! | SWO-362261 | Assisted Replacement | Electrode | Electrode | pO2 |  | 06.06.10:45, Mr. Küchelbacher, Tel.: 06213831024: ABL displays various error messages | Mr. Küchelbacher reports by phone : - QC liquid is not transported to the pH /BG module and pO2 electrode escapes electrolyte solution from membrane ! Upon arrival, analyser reports calibration and QC errors in pH, pO2, pCO2 and electrolytes (sometimes no values) 501 Rinsing solution not sucked in pH /BG module within time frame (since 4.6.18 , approx. 1 p.m.) | Check measuring valves pH $BG module in service (without complaint) - Inspection measuring chamber and electrodes in the pH $BG module - pO2 electrode mechanically defective (complete electrode tip not available $abgerissen) - renewal pO2 electrode - function check etc. without Complaint!
No | Visit | Complaint | RDEU | KAL-Driften Elektrolyte | KAL Fehler (KAL-Driften Elektrolyte) | Frau Salomon (Gerätebeauftragte) hat bereits mehrfach Kontakte gereinigt und Membranierungen durchgeführt (mehrfacher telefonischen Kontakt) - KAL-Driften Elektrolyte treten trotzdem immer wieder auf - allgemeine Überprüfung - Erneuerung Referenz Elektrode - mehrfache Kalibrationen durchgeführt - Funktionsüberprüfung etc. | SWO-362123 | Assisted Replacement | Electrode | Electrode | Reference |  | KAL Driftelectrodes Electrolytes | KAL Error (KAL Drift Electrolytes) | Ms. Salomon (device officer) has already cleaned contacts and performed membranes several times (multiple telephone contact) - KAL-Driften electrolytes still occur again and again - general review - renewal reference electrode - multiple Calibrations performed - function check, etc.
No | 28 | Y | RPNZ | Blank | 1178: Pump calibration error detected. Solution packs BE-01 faulty. |  | D472092D505AB346 | Assisted Replacement | SP | SP | Pump cal | 1178 | Blank | 1178: Pump calibration error detected. Solution packs BE-01 faulty.
No | Visit | Complaint | RLTD | Battery reading zero % Scarlett Willet 0203-4483479 | Battery error reading zero % | Battery disconnected and resets battery module refitted and the battery icon is now showing 100 charge and backing up analysers ok. Analyser has also been connected to a UPS . | SWO-362916 | Adjustment | HW | Battery module | Battery module | reset | Battery reading zero % Scarlett Willet 0203-4483479 | Battery error reading zero % | Battery disconnected and resets battery module refitted and the battery icon is now showing 100 charge and backing up analysers ok. Analyser has also been connected to a UPS .
No | Visit | Complaint | RAME | Customer analyzer getting status code 2. | Customer called stating that after rebooting analyzer got status code 2 | Reloaded software on analyzer with new database. Transferred all setting back on to analyzer. | SWO-363043 | Defect | SW | RMED SW | RMED SW | Status code 2 | Customer analyzer getting status code 2. | Customer called stating that after rebooting analyzer got status code 2 | Reloaded software on analyzer with new database. Transferred all setting back on to analyzer.
Yes | 2 | Y | RCZE | Blank |  |  | D471BFF75059DA71 | Re classified RMED | Demo anlyzer |  |  |  | Blank
Yes | Visit | Complaint | RIND | K failed | Checked calibration data and found K failed. | Flushed sensor cassette with NS solution. Performed zeroing of all channels successfully. 2 point calibration performed, still K failed. | SWO-363605 | [5 empty] | K failed | Checked calibration data and found K failed. | Flushed sensor cassette with NS solution. Performed zeroing of all channels successfully. 2 point calibration performed, still K failed.
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | lac glu响应超出范围，报错0378 | 更换测试卡后修复 | SWO-363213 | [5 empty] | Test card, reagent package failure | lac glu response out of range, error 0378 | Repair after replacing test card
Yes | Visit | Complaint | RCHN | 试管夹断裂 | 试管夹断裂，需要借件903-584一个 | 更换903-584一个正常。(by chuanwei) | SWO-361273 | [5 empty] | Test tube clip break | Test tube clip broken, need to borrow 903-584 one | Replace the 903-584 one normal. (by chuanwei)
No | Visit | Complaint | RCHN | 显示器线断裂 | 温度感应器线断裂，需要借件905-908 | 更换905-908后修复(by chuanwei) | SWO-361934 | Defect | HW | Process Module | Sensor, BnP |  | Display line break | Temperature sensor line breaks, needs to borrow 905-908 | Repair after replacement 905-908 (by chuanwei)
Yes | Visit | Complaint | RIND | Po2 is low for QC | REPLACEMENT OF Po2 electrode | Replaced the po2 electrode and given calibration also cl- had problem so removed the cl- electrode then cleaned the puss and given protien removal calibarted to raedy mode | SWO-363812 | [5 empty] | Po2 is low for QC | REPLACEMENT OF Po2 electrode | Replaced the po2 electrode and given calibration also cl- had problem so removed the cl- electrode then cleaned the puss and given protien removal calibarted to raedy mode
No | Visit | Complaint | RDAN | Marianne melder om mange fejl med bla. for lave OXI værdier(tHb=0,2). | Afprøvet og undersøgt. Log efterset. OXI komplet test. Hans Peter har skiftet slangeklemme ventil samt oxi slange. | Inletprobe, preheater, luftventiler, div. slanger og pakringe. Slutafprøvet. | SWO-362842 | Defect | HW | Oxi module | pinch valve |  | Marianne reports many bugs with bla. For low OXI values (tHb = 0.2). | Tested and investigated. Log checked. OXI complete test. Hans Peter has changed the hose clamp valve as well as the oxi hose. | Inletprobe, preheater, air vents, misc. hoses and packing rings. Final tested.
No | Visit | Complaint | RSCH | Rep Ventilator | Analysator macht lautes Geräusch | Ventilator Auto Check Modul ausgetauscht Gerinnsel pH Elektrode entfernt Funktionskntrolle i.o | SWO-362821 | Defect | HW | Auto Check Module | Fan |  | Rep Ventilator | Analyzer makes loud noise | Fan Auto Check Module replaced Clot pH electrode removed Functional knob s.o.
No | 1 | Y | RPOL | Blank | Aparat zg³asza³ b³¹d sprzêtowy oksymetrii. Regulacja lampy halogenowej. Kalibracja tHb od serwisu. Testy. Aparat sprawny. |  | D471A0005002749F | Adjustment | HW | Oxi module | Lamp unit | halogen | Blank | The camera has reported hardware hardware symmetry. Halogen lamp adjustment. THb Calibration from service. Tests. The camera is functional.
No | Visit | Complaint | RDAN | Marianne melder, at ABL'en er utæt. | Undersøgt | Wasteslanger udskiftet, slutafprøvet. | SWO-362846 | Defect | HW | Basic Unit | Waste tubes |  | Marianne reports that the Abbey is leaking. | Examined | Waste hoses replaced, final tested.
No | Visit | Complaint | RDEU | 07.06.2018, 08:59 Uhr, Fr. Jannick, Tel. 0211904361209: Messung wird nach 2 Min. abgebrochen | Verunreinigung im Storage Wheel. | Storage Wheel Unit demontierung gereinigt, Sensoren gereinigt, Punch UP gereinigt, QC und Blutige Probenmessung durchgeführt. | SWO-363492 | Cleaned | HW | Storage Module | Storage wheel |  | 07.06.2018, 08:59, Fr. Jannick, Tel. 0211904361209: Measurement is aborted after 2 min. | Contamination in the Storage Wheel. | Storage Wheel Unit deassembly cleaned, sensors cleaned, Punch UP cleaned, QC and Bloody sample measurement performed.
No | Visit | Complaint | RIND | QC pO2 values high in Level -2 | found problem with QC lot | Checked the instrument for error, checked the sensor connector, wet section solution flow path found flow path fine. performed three times each QC level-2 and Level-3.found problem with QC level 2 ,but QC level -3 all results within the range. found instrument working satisfactorily. | SWO-363884 | Defect | HW | Basic Unit | PCB, Sensor connector assembly |  | QC pO2 values high in Level -2 | found problem with QC lot | Checked the instrument for error, checked the sensor connector, wet section solution flow path found flow path fine. performed three times each QC level-2 and Level-3.found problem with QC level 2 ,but QC level -3 all results within the range. found instrument working satisfactorily.
No | Visit | Complaint | RIND | Drain space is leaking | Waste tubes is blocked | checked the analyser there was leakakge in the wste unit , so cleannend the waste tubes and pump tubes and soluton moving path all the parts were chekced and refixed itand checked for leakage is there or not , after that no leakag found instrument working fine | SWO-363816 | [5 empty] | Drain space is leaking | Waste tubes is blocked | checked the analyser there was leakakge in the wste unit , so cleannend the waste tubes and pump tubes and soluton moving path all the parts were chekced and refixed itand checked for leakage is there or not , after that no leakag found instrument working fine
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | K结果偏高，测试卡初始化失败 | 连续更换2张测试卡后修复 | SWO-365409 | [5 empty] | Test card, reagent package failure | K result is high, test card initialization failed | Repair after 2 consecutive test cards
No | Visit | Complaint | RCHN | 报错1024 | 报错1024，需要借件944-212（14308） | 清洁处理盘，更换903-364一个，重新定标正常。(by chuanwei) | SWO-362591 | Defect | HW | Process Module | Optical Unit | 1024 | 1024 error | 1024, 944-212 (14308) | Clean the handling disc, replace 903-364 one, re-label normal. (by chuanwei)
No | Visit | Complaint | RCHN | 存储盘螺丝松动 | 存储盘螺丝松动,有异响，需要借件903-378一个 | 加借903-030,636-452各一个.（by chuanwei） | SWO-362474 | Defect | HW | Storage Module | Cartridge lock |  | The storage tray screws are loose | Storage tray screw loose, noise, need to borrow 903-378 one | Borrow 903-030, 636-452 each. (by chuanwei)
No | Visit | Complaint | RCHN | 存储盘螺丝松动 | 加借903-030，636-452各一个。 | 更换903-030后修复 by Andrew | SWO-363422 | Revisit required | HW | Storage Module | Cartridge lock |  | The storage tray screws are loose | Borrow 903-030, 636-452 each. | Repair by Andrew after replacement 903-030
No | Visit | Complaint | RFRA | Problèmes Po2 Résultats non cohérent avec l'oximetre de pouls Résultats non repetable lorsqu'ils sont passés deux fois d'affilée | Vérification complète | Nettoyage d s contacts de la cassette de capteur $ $ Réinstallation des consommable $ $ passage de CQ $ $ essais: bons | SWO-363577 | [5 empty] | Po2 problems Results not consistent with pulse oximetre Unrepeatable results when they have gone twice in a row | Full check | Cleaning d s contacts of sensor cassette $$ Reinstallation of consumables $$ passage of CQ $$ trials: vouchers
No | Visit | Complaint | RDEU | 07.06.2018, 11:03 Uhr, Fr. Bretting, Tel. 030940113280: QC bei Crea A zu tief | Reinigungsarbeiten, Fehlersuche | Dekontamination, Messkammer Crea und Elektrodenkontakte gereinigt, Crea membraniert, korrekte Kal abgewartet, Kal und QK ok | SWO-364061 | Cleaned | HW | Met II module | Connector |  | 06/07/2018, 11:03 AM, Ms Bretting, tel. 030940113280: QC at Crea A | Cleaning work, troubleshooting | Decontamination, measuring chamber Crea and electrode contacts cleaned, Crea membraneed, correct calf waited, Kal and QK ok
Yes | 6 | Y | RCZE | Blank |  |  | D476A0CE2832ADF1 | [5 empty] | Blank
No | Visit | Complaint | RAME | Lamp intensity low error 257,258 | Lamp intensity low, neon lamp errors -- ran complete oxy test hemolyzer and lamp failed | Replaced lamp and was able to adjust to 3200 but neon lamp failed; Replaced hemolyzer and hemolyzer voltage still low--put the original hemolyzer back in the machine; ordered a oxy pcb and power supply for tomorrow -- Replaced the oxy PCB part not in the CRM part number for oxy pcb is 903-675--pCO2 Failed QC level 3 too high, remembraned PCO2 | SWO-363178 | Defect | HW | Oxi module | PCB, oxi |  | Lamp intensity low error 257,258 | Lamp intensity low, neon lamp errors -- ran complete oxy test hemolyzer and lamp failed | Replaced lamp and was able to adjust to 3200 but neon lamp failed; Replaced hemolyzer and hemolyzer voltage still low--put the original hemolyzer back in the machine; ordered a oxy pcb and power supply for tomorrow -- Replaced the oxy PCB part not in the CRM part number for oxy pcb is 903-675--pCO2 Failed QC level 3 too high, remembraned PCO2
No | Visit | Complaint | RIND | Oximetry error. | Found halogen intensity very high. | Adjusted the halogen intensity. performed complete oximetry test , 2 point calibration and samples found instrument working satisfactorily. | SWO-363885 | Adjustment | HW | Oxi module | Lamp unit | halogen | Oximetry error. | Found halogen intensity very high. | Adjusted the halogen intensity. performed complete oximetry test , 2 point calibration and samples found instrument working satisfactorily.
No | Visit | Complaint | RBNL | Vals positieve TnT | Vorige week,werd er een vals positief patientstaal gemeten voor TnT | Het proceswiel + wasstation werden nagekeken en gereinigd. Hierna het filter gedemonteerd en gereinigd. Niets abnormaals kunnen vaststellen. Systeemcontrole OK 3 levels MC gecontroleerd en deze waren allemaal goed. Er werd eveneens een SD genomen voor eventueel verder onderzoek bij de fabrikant. Waarschijnlijk heeft er een micro-stolsel gezeten in de testcup tijdens de meting waardoor het resultaat foutief beïnvloed werd. | SWO-363529 | Cleaned | HW | Process Module | Process Wheel |  | False positive TnT | Last week, a false positive patient sample was measured for TnT | The process wheel + wash station were checked and cleaned. Then disassembled and cleaned the filter. Cannot establish anything abnormal. System Check OK 3 levels MC checked and these were all good. An SD was also taken for any further investigation by the manufacturer. Probably a micro-clot was seated in the Testcup during the measurement which caused the result to be wrongly influenced.
No | Visit | Complaint | RDEU | Mitteilung vom Kunden das dieses Gerät nicht funktioniert | Gerät überprüft , Fehlerlog überprüft | keine Fehler in den letzten Tagen aufgetreten, PCT Testkit auf Wunsch hinterlegt | SWO-364399 | No fault found | No fault found | No fault found | No fault found |  | Notification from the customer that this device is not working | Device checked , error log checked | no errors occurred in the last days, PCT test kit stored on request
Yes | Visit | Complaint | RCHN | 报错 4587 | 报错 4587，借903-198一个。 | 更换903-198一个正常。(By Ran) | SWO-363336 | [5 empty] | Error 4587 | 4587, borrowed 903-198 one. | Replace 903-198 one normal. (By Ran)
No | Visit | Complaint | RCHN | 温度上不去 1020 | 温度上不去 1020,借636-652，903-128各一个。 6月10日加借 903-138一个。 |  | SWO-363904 | Defect | HW | Incubator Module | PCB, Incubator | 1020 | Temperature not 1020 | Temperatures do not go up to 1020, borrow 636-652, 903-128 each. On June 10th, 903-138 was added.
No | Visit | Complaint | RAME | Halogen lamp to high | Oxi issue | Ajusté la lampe | SWO-363968 | Adjustment | HW | Oxi module | Lamp unit | halogen | Halogen lamp to high | Oxygen issue | Fitted lamp
No | Visit | Complaint | RLTD | Adam 01772523007 | Sensor bay door would not close. Nathan had previously attended this analyser for the same issue and found a loose cable. | The analyser was powered down and wet section removed. The sensor bay door was manually closed where some resistance was met toward the bottom. The open $close stepper motor was replaced with new and analyser restarted. In the service menu the door was open and closed again, the error was still present. Analyser powered down and again and the wet section PCB was replaced. Service menu entered and door opened and closed with no issues. Exited service menu and went through the replace sensor cassette process again with no issues. | SWO-364880 | Defect | HW | Basic Unit | PCB, Wet section | SC door | Adam 01772523007 | Sensor bay door would not close. Nathan had previously attended this analyser for the same issue and found a loose cable. | The analyser was powered down and wet section removed. The sensor bay door was manually closed where some resistance was met toward the bottom. The open $close stepper motor was replaced with new and analyser restarted. In the service menu the door was open and closed again, the error was still present. Analyser powered down and again and the wet section PCB was replaced. Service menu entered and door opened and closed with no issues. Exited service menu and went through the replace sensor cassette process again with no issues.
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 血红蛋白定标失败 0210 | 更换试剂包后修复 | SWO-373161 | [5 empty] | Test card, reagent package failure | Hemoglobin calibration failure 0210 | Repair after replacement of reagent kit
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | GLU,LAC定标不过 0210 | 更换测试卡后修复 | SWO-364711 | [5 empty] | Test card, reagent package failure | GLU, LAC calibration but 0210 | Repair after replacing test card
No | Visit | Complaint | RSCH | Reparatur Austausch | Einlass offen bei Messung 1190 Seit SW Upgrade 3.4 vermehrt aufgetreten | Analysator zur Reparatur ausgetauscht Leihgerät installiert und konfiguriert sowie Personal Einweisung Funktionskontrolle i.o | SWO-364120 | Revisit required | HW | Auto Inlet module | PCB, Auto inlet | Analyzer removed | Repair Replacement | Inlet open at measurement 1190 Increased since SW Upgrade 3.4 | Analyzer replaced for repair rental device is installed and configured, as well as personnel instruction Functional control i.o
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 无法加入试剂包 1165，0378,0210,1227 | 连续更换1张测试卡和1个试剂包后修复 | SWO-365408 | [5 empty] | Test card, reagent package failure | Reagent pack 1165,0378,0210,1227 | Repair after continuous replacement of 1 test card and 1 reagent kit
Yes | Visit | Complaint | RCHN | 缓冲液安装失败 | 缓冲液安装失败 | 更换试剂包后修复 | SWO-364827 | [5 empty] | Buffer installation failed | Buffer installation failed | Repair after replacement of reagent kit
No | Visit | Complaint | RCHN | 无法开机 | 无法开机，借910-343一个 |  | SWO-364024 | Defect | HW | Analyzer Control Module | CPU |  | Cannot power on | Cannot power on, borrow 910-343 one
No | Visit | Complaint | RDAN | Marianne melder om mange fejl med bla. for lave OXI værdier(tHb=0,2). | Undersøgt. Har gjort det igen. | OXI -R-Lånemodul 902-833R1773N0007 isat, service tHb kalibreret, slutafprøvet. | SWO-364449 | Revisit required | HW | Oxi module | Oxi module |  | Marianne reports many bugs with bla. For low OXI values (tHb = 0.2). | Examined. Have done it again. | OXI -R-Lånemodul 902-833R1773N0007 isat, service tHb kalibreret, slutafprøvet.
No | Visit | Complaint | RDEU | 04.05.2018; Hr. Henneboele; Tel.: 02093643880; Na und Ca Kal Fehler 04.05.2018; 11:40 Uhr; Hr. Henneboele; Tel.:02093644504; weiterhin Na und Ca Driftfehler 11.06.2018; 11:15 Uhr; Hr. Hennboehle; Tel.: 02093643880; nach der Na und Ca Driftfehler besteht weiterhin | E'lyte kalibrationsfehler, zu hohe Driften, beide Airventile kleben im Ventilkörper, | Fließwege und Messkammer gereinigt, Ventile aus den Messmodulen ausgebaut und gereinigt, Gerät kalibriert, | SWO-364489 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | 04.05.2018; Mr. Henneboele; Tel.: 02093643880; Na and Ca Kal Error 04.05.2018; 11:40 a.m.; Mr. Henneboele; Tel.:0 2093644504; continue Na and Ca drift errors 11.06.2018; 11:15 a.m.; Mr. Hennboehle; Tel.: 02093643880; after the Na and Ca drift error, there is still | E'lyte calibration error, too high drifts, both airvalves stick in the valve body, | Flow paths and measuring chamber cleaned, valves removed and cleaned from the measuring modules, device calibrated,
No | Visit | Complaint | RAME | Analyzer froze and wont boot-up | Screen locking up during boot. Installed cf card with no change. Ordered new 945 computer. | Replaced below parts. Re-installed original cf card. | SWO-364734 | Defect | HW | DMS_PC | CPU |  | Analyzer froze and wont boot-up | Screen locking up during boot. Installed cf card with no change. Ordered new 945 computer. | Replaced below parts. Re-installed original cf card.
Yes | 3 | Y | RCZE | Blank |  |  | D4781B309E5DDA74 | Re classified RMED | planned SWO |  |  |  | Blank
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 多参数定标不过 | 更换测试卡后修复 | SWO-364855 | [5 empty] | Test card, reagent package failure | Multi-parameter calibration but | Repair after replacing test card
No | Visit | Complaint | RAME | Consumable errors | Issues with recently with replacing SP’s and SC’s documented in action logbook on 5-25 and 6-01. There have been others. SP’s not recognized. SC’s PO2 won’t calibrate and or QC won’t come in. Proactive onsite service is needed. | Perform cleaning program. Replace parts from PM kit. Set Zero in service. No issues. | SWO-364678 | [5 empty] | Consumable errors | Issues with recently with replacing SP’s and SC’s documented in action logbook on 5-25 and 6-01. There have been others. SP’s not recognized. SC’s PO2 won’t calibrate and or QC won’t come in. Proactive onsite service is needed. | Perform cleaning program. Replace parts from PM kit. Set Zero in service. No issues.
Yes | Visit | Complaint | RCHN | 定标失败 | 定标失败，借944-216（14259），944-214（14322）各一张。 | 重新定标后修复 by Andrew | SWO-363941 | [5 empty] | Bid failed | The bid failed, borrowing 944-216 (14259), 944-214 (14322) each. | Fix by Andrew after re-calibration
No | Visit | Complaint | RCHN | 进样针弯折 | 进样针弯折，借636-549,903-171（不可借用，已告知） | 更换636-549后修复 by Andrew | SWO-364135 | Defect | HW | Aspiration Module | Sensor, needle |  | Into the sample needle bend | Sample needle bend, borrow 636-549, 903-171 (not borrowable, informed) | Repair by Andrew after replacing 636-549
No | Visit | Complaint | RCHN | 定标终止 | 定标终止，借944-258（14391）一个 | 重新定标后修复 by Andrew | SWO-364834 | Assisted Replacement | Consumables | Test Kit | Cal cartridge, NT-proBNP |  | Bid termination | Bid termination, loan 944-258 (14391) one | Fix by Andrew after re-calibration
No | Visit | Complaint | RIND | Service Visit. | Calibration has failed of the parameters except pO2. | Try to reinitiate the sensor cassette after flushing with de-ionized water, not successfull, the elctrical response and no end points errorr shows in the calibration screen, Need to install the new sensor cassette. Install the new SC9, successfull done, analyzer handed over to staff in well working condition. | SWO-366522 | [5 empty] | Service Visit. | Calibration has failed of the parameters except pO2. | Try to reinitiate the sensor cassette after flushing with de-ionized water, not successfull, the elctrical response and no end points errorr shows in the calibration screen, Need to install the new sensor cassette. Install the new SC9, successfull done, analyzer handed over to staff in well working condition.
No | Visit | Complaint | RAME | Cal Gas 2 leaking | Error 637, gas 2 cylinder low pressure. Performed service diagnostics gas test, gas 2 failed test. | replaced gas 2 flow regulator. Reran gas 2 diagnostic test, passed, observed no further failures. | SWO-366103 | Defect | HW | gas module | Flow regulator | 637 | Cal Gas 2 leaking | Error 637, gas 2 cylinder low pressure. Performed service diagnostics gas test, gas 2 failed test. | replaced gas 2 flow regulator. Reran gas 2 diagnostic test, passed, observed no further failures.
No | Visit | Complaint | RAME | Having to replace the glu memebrane more frequently to keep cals and QCs green. She would like to schedule a service call. | Suspect sample volume has increased since going live causing accelerated wear on membranes | Collected data from data logs for review. No other intervention performed. Examination of patient data logs reveals since going live this analyzer is averaging approximately 100 samples a day, significantly higher than the suggested 1 month change interval for analyzers running 40 samples or less per day. Suggested customer use a 20 day change interval for impacted tests as a starting point to determine what will work best for their work load. | SWO-366225 | Assisted Replacement | Electrode | Membrane | Glu |  | Having to replace the glu memebrane more frequently to keep cals and QCs green. She would like to schedule a service call. | Suspect sample volume has increased since going live causing accelerated wear on membranes | Collected data from data logs for review. No other intervention performed. Examination of patient data logs reveals since going live this analyzer is averaging approximately 100 samples a day, significantly higher than the suggested 1 month change interval for analyzers running 40 samples or less per day. Suggested customer use a 20 day change interval for impacted tests as a starting point to determine what will work best for their work load.
No | Visit | Complaint | RKK | ﾌﾟﾘﾝﾀｰを印字したとき左に紙が寄っていく。 | ﾌﾟﾘﾝﾀｰ印字の斜行 | プリンターヘッド交換、印字確認。 | SWO-365367 | Defect | HW | Printer Module | Printer head |  | When i print the printer, the paper comes up to the left. | Printer printing diagonally | Printer head replacement, print confirmation.
No | Visit | Complaint | RAME | Rinse Error code 0508 | Rinse error 508 - pH rinse valve not functioning. | Replaced the rinse valve and ran calibrations and QC. All tests passed. | SWO-365204 | Defect | HW | pH/BG Module | Valve | 0508 | Rinse Error code 0508 | Rinse error 508 - pH rinse valve not functioning. | Replaced the rinse valve and ran calibrations and QC. All tests passed.
No | Visit | Complaint | RAME | Windows settings | none necessary | Made adjustment in the windows control panel for military time to be displayed. | SWO-365321 | Setup | SW | Windows | Time setting |  | Windows settings | none necessary | Made adjustment in the windows control panel for military time to be displayed.
No | Visit | Complaint | RDEU | 12.06.2018; Fr. Thul; Tel.: 068977991360; roter Bildschirm | Bildschirm nimmt andere Farbe an wenn er bewegt wird. Touchscreen ohne Funktion | Touchscreen getauscht. Neustart. | SWO-365439 | Revisit required | HW | Cover and Chassis | display |  | 12.06.2018; Ms. Thul; Tel.: 068977991360; red screen | Screen takes on a different color when it is moved. Touchscreen without function | Touchscreen. Restart.
No | Visit | Complaint | RDEU | 11.06.2018/10:00 Hr. Schmidt (MT) kann den ABL nicht neu starten, Bildschirm bleibt schwarz | Fehlersuche Am Ausgang des Netzteils kommen keine Spannungen an Sicherungen geprüft | Netzteil ausgetauscht, Einlaufverfahren, SC90 gewechselt | SWO-366228 | Defect | HW | Basic Unit | Power Supply |  | 11.06.2018/10:00 Mr. Schmidt (MT) cannot restart the ABL, screen remains black | Troubleshooting There are no voltages at the output of the power supply Fuses checked | Power supply replaced, inlet method, SC90 changed
No | Visit | Complaint | RDEU | 13.06.2018 8:45 Fr. Höhne 040790202953: Interventionsmeldung LP | LP Tausch wird ständig verlangt. | Fehlersuche am Analysator durchgeführt. Jahreswartungkitt verbaut. Keine Besserung. Alle Fließwege gespült. Weg zum Hemolyzer verstopft. Hemolyzer ausgebaut und durch neuen Ersetzt. Analysator hochgefahren. System wieder einsatzbereit. | SWO-365614 | Defect | HW | Oxi module | Hemolyzer | clot | 13.06.2018 8:45 Fr. Höhne 040790202953: Intervention report LP | LP exchange is constantly required. | Troubleshooting is performed on the analyzer. annual maintenance kitt installed. No improvement. All flow paths rinsed. Way to the Hemolyzer clogged. Hemolyzer expanded and replaced with new ones. analyzer. system is ready for use again.
No | 1 | Y | RPAC | Blank | Replaced borrowed SC |  | D4798A2D145E1A3D | Assisted Replacement | SC | SC | unknown |  | Blank | Replaced borrowed SC
Yes | 3 | Y | RPAC | Blank | SC |  | D48AC78588527D62 | Re classified RMED | same day upSWO |  |  |  | Blank | SC
No | Visit | Complaint | RAME | Customer call stating analyzer will not automatically log off | Found that with trying to correct printing issue and setting the printer to select printer from list stopped the analyzer from automatically logging off when a sample was ran unless you hit start button. | Unchecked the box to select print on analyzer has corrected the logging off problem. I will have customer monitor to see if printing problem comes back. | SWO-365926 | Defect | SW | RMED SW | Setup | log off | Customer call stating analyzer will not automatically log off | Found that with trying to correct printing issue and setting the printer to select printer from list stopped the analyzer from automatically logging off when a sample was ran unless you hit start button. | Unchecked the box to select print on analyzer has corrected the logging off problem. I will have customer monitor to see if printing problem comes back.
No | Visit | Complaint | RDEU | 5.06.2018 15:40Uhr VonKanen 023195057411: Es gab Abweichungen bei tHb zum Labor. Die Werte schwanken. 13.06.2018, 10:22 Uhr, H. Söderberg, Tel. 023195057413: tHb Werte zu niedrig | Gerät überprüft und gereinigt, Oxi Complet Test durchgeführt | Hemolyzer erneuert, Gerät neu gestartet, tHb Kal und QCs durchgeführt | SWO-366463 | Defect | HW | Oxi module | Hemolyzer |  | 5.06.2018 15:40H VonKanen 023195057411: There were deviations at tHb to the laboratory. The values fluctuate. 13.06.2018, 10:22, H. Söderberg, tel. 023195057413: tHb values too low | Device checked and cleaned, Oxi Complet test performed | Hemolyzer renewed, device restarted, tHb Kal and QCs performed
No | Visit | Complaint | RIND | PCo2, Na, Ca parameter is not active. | Pco2 Na, Ca parameter was inactive. | Remove Sensor cassette then keep it for hydration. zeroing analyser with the help of zeroing cassette. then again reinstall sensor cassette. after that all parameter was active. and analyser start working properly. handed over in fine working condition. | SWO-367422 | [5 empty] | PCo2, Na, Ca parameter is not active. | Pco2 Na, Ca parameter was inactive. | Remove Sensor cassette then keep it for hydration. zeroing analyser with the help of zeroing cassette. then again reinstall sensor cassette. after that all parameter was active. and analyser start working properly. handed over in fine working condition.
No | Visit | Complaint | RFRA | Blocage | Blocage fluidique | Tuyaux pince entre le médico-légal et la cassette, démontage ok remise en place correctement du tubing.Cylce de qualité ok passage cq ok | SWO-365822 | [5 empty] | freezing | Fluid blockage | Pipes pinch between the forensic and the cassette, disassembly ok properly re-establishing the tubing. Cylce quality ok passage cq ok ok
No | Visit | Complaint | RCHN | BNP温度传感器故障 | BNP温度传感器故障 借905-908一个 | 更换905-908后修复 by Andrew | SWO-364888 | Defect | HW | Process Module | Sensor, BnP |  | BNP temperature sensor failure | BNP temperature sensor failure borrowed 905-908 one | Repair by Andrew after replacement 905-908
No | Visit | Complaint | RCRO | Persistent tHb calibration error. | Provjera sustava. Kiveta u Oxymetru vjerojatno ima pukotinu ili je kontaminirana. | Zamjena hemolizatora. Servisna tHb kalibracija. | SWO-366277 | Defect | HW | Oxi module | Hemolyzer |  | Persistent tHb calibration error. | System check. The Custer in Oxymeter probably has a leak or it's contaminated. | Zamjena hemolizatora. Servisna tHb kalibracija.
Yes | 11 | Y | RPAC | Blank | Supplied and fitted l pump tube. |  | D479627A140B50BF | [5 empty] | Blank | Supplied and fitted l pump tube.
Yes | Visit | Complaint | RAME | Neon Error on analyzer, oximetry down Contact is Joanne at 925-295-4629 | Error 316 neon lamp. Per manual will replace oxy pcb. Part number 903-675 replaced not in the CRM. | Replaced oxy pcb; replaced lamp loaded software 6.18 performed annual maintenance in conjunction with repair will document on a Separate SWO for tracking purposes -- neon lamp error still occurring -- on site replace complete oxy module | SWO-366081 | [5 empty] | Neon Error on analyzer, oximetry down Contact is Joanne at 925-295-4629 | Error 316 neon lamp. Per manual will replace oxy pcb. Part number 903-675 replaced not in the CRM. | Replaced oxy pcb; replaced lamp loaded software 6.18 performed annual maintenance in conjunction with repair will document on a Separate SWO for tracking purposes -- neon lamp error still occurring -- on site replace complete oxy module
No | Visit | Complaint | RDEU | 08.06.2018; Hr. Mageira; Tel.: 059311524482; hoher QC Verbrauch, Kal Fehler | Zeitweise Ca Rilibäkverletzung Immer wieder Driftfehler bei den Elektrolyten | Gerät überprüft, QC Einstellungen kontrolliert, Meßkammer und Luftauslass gesäubert, Ca, Ref. und Glu membraniert, Ref Elektrode ausgetauscht | SWO-366362 | Assisted Replacement | Electrode | Electrode | Reference |  | 08.06.2018; Mr. Mageira; Tel.: 059311524482; high QC consumption, Kal error | Intermittent Ca Rilibäk injury Time and again drift faults in the electrolytes | Device checked, QC settings checked, measuring chamber and air outlet cleaned, Ca, Ref. and Glu membraned, ref electrode replaced
No | Visit | Complaint | RDEU | 12.06.2018; Fr. Thul; Tel.: 068977991360; roter Bildschirm | Bildschirmschoner ist aktiv. Touchscreen ohne Funktion. Auf dem Bildschirm Wechseln die Farbe, wenn sie bewegt wird. | Bildschirm kompl. getauscht. Bewegung in alle Richtungen bringt keine Veränderung. Touchscreen ist wieder bedienbar. | SWO-365787 | Defect | HW | Cover and Chassis | display |  | 12.06.2018; Ms. Thul; Tel.: 068977991360; red screen | Screen saver is active. Touch screen without function. On the screen, change color when it is moved. | Screen compl. Exchanged. Movement in all directions brings no change. Touchscreen can be operated again.
Yes | 4 | Y | RPAC | Blank | Pressure test vacuum error |  | D47AB86A8A62F9C7 | [5 empty] | Blank | Pressure test vacuum error
Yes | 6 | Y | RPAC | Blank | Inlet gasket due for replace |  | D47ABD2E8A62411E | [5 empty] | Blank | Inlet gasket due for replace
Yes | 1 | Y | RCZE | Blank |  |  | D47F78E76252DA85 | Re classified RMED | planned SWO |  |  |  | Blank
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 测试卡初始化失败 | 现场指导老师操作后解决 | SWO-366858 | [5 empty] | Test card, reagent package failure | Test card initialization failed | On-site guidance teacher after operation to solve
No | Visit | Complaint | RCHN | po2质控不过，报错210 | po2质控不过，报错210 | 重新清洁管路后修复by Andrew | SWO-366226 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway | pO2 | po2 QC but, misreported 210 | po2 QC but, misreported 210 | Repair by Andrew after recleaning lines
No | Visit | Complaint | RCHN | 卡盘，报错5784 | 卡盘，报错5784 | 更换测试卡后修复 | SWO-366807 | Defect | Consumables | Test Kit | Cartridge, Broken | 5784 | Card, 5784 | Card, 5784 | Repair after replacing test card
No | Visit | Complaint | RDEU | die RiLiBÄK wird nicht kontinuierlich eingehalten | QC 1+ verwendet | Rilibäk-Grenzen werden nicht sicher erreicht, Kunde testet in den nächsten Tagen das QC-Material | SWO-366819 | [5 empty] | the RiLiBÄK is not continuously adhered to | QC 1+ used | Rilibäk limits are not reached safely, customer tests QC material in the next few days
Yes | Visit | Complaint | RIND | Gasketleaking | Gasket leaking | cleaned the gasket clott was there in the gasket then reinserted it but still lil bit leakag was there so interchanged the pumptubes then obsrevd for few samples and calibartion it was working fine | SWO-367327 | [5 empty] | Gasketleaking | Gasket leaking | cleaned the gasket clott was there in the gasket then reinserted it but still lil bit leakag was there so interchanged the pumptubes then obsrevd for few samples and calibartion it was working fine
No | Visit | Complaint | RKK | リークがあります | リークがあります | PH電極シーリングリング外れ、Glu電極シーリングリングのずれが原因。PH電極シーリングリング取り付け、Glu電極シーリングリング及びメンブラン再取り付け、Refメンブラン再取り付け、エアーライン洗浄。 | SWO-367534 | Defect | Electrode | Electrode | sealing ring | Sealing ring, Glu | There's a leak. | There's a leak. | Ph electrode sealing ring is out of bounds and glu electrode sealing ring is misaligned. PH electrode sealing ring mounting, Glu electrode sealing ring and membrane re-installation, Ref membrane re-installation, air line cleaning.
No | Visit | Complaint | RDAN | Marianne melder om mange fejl med bla. for lave OXI værdier(tHb=0,2). 11-6-2018: OXI Lånemodul isat. | Undersøgt. Trykprøvet mm. Prøvet med dummyer istedet for Crea elektroder. | Ventil 7+8 udskiftet. Pakring på Cl dummy skiftet, membran på Ca skiftet. Kørt 28 x kapillær 125ul uden fejl. | SWO-366493 | No fault found | No fault found | No fault found | No fault found |  | Marianne reports many bugs with bla. For low OXI values (tHb = 0.2). 11-6-2018: OXI loan module inserted. | Examined. Pressure tested mm. tested with dummies instead of Crea electrodes. | Valve 7 + 8 replaced. Packing on the Cl dummy shift, membrane on the Ca shift. Run 28 x Capillary 125ul without error.
No | Visit | Complaint | RDEU | 25.05.2018 16:30 Fr. Lener 08856936370: Systemfehler Inkubation-Wheel | Optical Unit erneuern | Optical Unit erneuert,Kalibration Nicht bestanden, Ohne Erfolg, Neu Optical Unit defekt, Ersatzteile bereits bestellt. | SWO-366919 | Revisit required | HW | Process Module | Optical Unit |  | 25.05.2018 16:30 Fr. Lener 08856936370: System error incubation wheel | Renewing the Optical Unit | Optical Unit renewed,calibration Not passed, Without success, New Optical Unit defective, spare parts already ordered.
No | Visit | Complaint | RDEU | Bei der Auslieferung und Inbetriebnahme sind Probleme mit der Systemüberprüfung aufgetreten. Intensität Blitzlampe, abweichung >13 % | Optische Einheit defekt. | Optische Einheit erneuert, Gerät kalibriert und konfiguriert. Systemüberprüfung nach Erneuerung der optischen Einheit OK. | SWO-367499 | Defect | HW | Process Module | Optical Unit | 1031 | System inspection issues have arisen during delivery and commissioning. Intensity flash lamp, deviation >13% | Optical unit defective. | Optical unit renewed, device calibrated and configured. System check after renewal of optical unit OK.
Yes | Visit | Complaint | RAME | Analyzer was left unplugged and now will not hold a charge | Has battery cell fault | Replaced Battery unit | SWO-366764 | [5 empty] | Analyzer was left unplugged and now will not hold a charge | Has battery cell fault | Replaced Battery unit
No | Visit | Complaint | RDEU | 15.06.2018 12:00Uhr Herr Varhy 03078824079 (Station 03078822040): Der Einlass ist defekt, ein Teil ist abgebrochen | Fehlersuche, Einlaßmodul nicht mehr funktionsfähig - Nadel komplett verbogen und Einlaßdichtungshalter abgebrochen - ein Rest steckt noch im Modul, weist eindeutig auf Fehlbedienung eines Anwenders hin | Einlaß abgebaut und Reststück des Halters entfernt, Nadel entfernt, beides ersetzt, zusammengesetzt und autom. Fehlerbehebnung abgewartet, Gerät startet erneut Startphase, Kal und QC | SWO-367540 | User error | HW | Inlet module | Inlet gasket |  | 15.06.2018 12:00pm Mr Varhy 03078824079 (Station 03078822040): The inlet is defective, one part is broken | Troubleshooting, inlet module no longer functional - needle completely bent and inlet seal holder broken - a remainder is still in the module, clearly indicates incorrect operation of a user | Inlet dismantled and leftover piece of the holder removed, needle removed, both replaced, assembled and autom. Error lifting waited, device restarts start phase, Kal and QC
No | Visit | Complaint | RAME | CAP cross over material METHb results are higher than mean | CAP cross over material METHb results are 1% higher than target. METHb blood sample comparisons are approximately 1% higher than peer results . All diagnostics, QC and VK results are in range. All setup is configured correctly and inspection shows instrument in good condition throughout. | Replaced hemolyzer without change of METHb blood results. Replaced spectrometer and METHb blood results were within .2% of peer or closer. All OXI parameters results are within Radiometer comparison specifications. QC results are in range. Customer waiting to run CAP SO solutions at a later date. Instrument operating normally. | SWO-367236 | Defect | HW | Oxi module | spectrometer | MetHb | CAP cross over material METHb results are higher than mean | CAP cross over material METHb results are 1% higher than target. METHb blood sample comparisons are approximately 1% higher than peer results . All diagnostics, QC and VK results are in range. All setup is configured correctly and inspection shows instrument in good condition throughout. | Replaced hemolyzer without change of METHb blood results. Replaced spectrometer and METHb blood results were within .2% of peer or closer. All OXI parameters results are within Radiometer comparison specifications. QC results are in range. Customer waiting to run CAP SO solutions at a later date. Instrument operating normally.
No | Visit | Complaint | RDEU | 01.05.2018 09:35uhr Herr Feldt 043116972333: Der Kunde hatte die Bereiche für pH, Co2 und pO2 erweitert, weil bei einer Patientenmessung keine Werte angezeigt wurden. | Ausdruck zeigt keine Werte die Außerhalb der Range sind. | Roma ini in den Sytem Setttings geändert. Und Kundengespräch durchgeführt. | SWO-366953 | Setup | SW | Roma.ini | reportable range |  | 01.05.2018 09:35o Mr Feldt 043116972333: The customer had expanded the ranges for pH, Co2 and pO2 because no values were displayed during a patient measurement. | Expression does not show values that are outside the range. | Roma ini changed to the system Setttings. And customer talk conducted.
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | PCO2质控不过 | 更换测试卡后修复 | SWO-366816 | [5 empty] | Test card, reagent package failure | PCO2 is not the only | Repair after replacing test card
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 所有参数定标失败 | 更换测试卡后修复 | SWO-372521 | [5 empty] | Test card, reagent package failure | Failed to calibrate all parameters | Repair after replacing test card
Yes | Visit | Complaint | RCHN | 报错1031 | 报错1031，借903-364一个 | 更换903-364后修复 by Andrew | SWO-366800 | [5 empty] | 1031 | 1031, loan 903-364 one | Repair by Andrew after replacing 903-364
Yes | Visit | Complaint | RSAF | Unit manager called to inform us regarding the TCM being non usable as it has a monitor error 08000200. | Plexus released new software for the TCM5 to be installed. | Installed newly released software at workshop, tested unit ok, unit delivered to customer. | SWO-368473 | [5 empty] | Unit manager called to inform us regarding the TCM being non usable as it has a monitor error 08000200. | Plexus released new software for the TCM5 to be installed. | Installed newly released software at workshop, tested unit ok, unit delivered to customer.
No | Visit | Complaint | RAME | Has issues with calibrations. | Issues with Ca,NA, Ref. Suspect customer has ECMO built up in instrument. Inspection did not show anything. | Customer is going to perform a protein removal once a day and decontamination once a week. | SWO-367373 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway | user training | Has issues with calibrations. | Issues with Ca,NA, Ref. Suspect customer has ECMO built up in instrument. Inspection did not show anything. | Customer is going to perform a protein removal once a day and decontamination once a week.
No | Visit | Complaint | RDEU | 25.05.2018 16:30 Fr. Lener 08856936370: Systemfehler Inkubation-Wheel | Kalibration gescheitert (TnI), Werte zu hoch (650-750), 3x über 700. | Optical Unit erneuert, Process Wheel erneuert, Kalibrationen vorgenommen, LQC MCC2 vorgenommen. | SWO-367300 | Defect | HW | Process Module | Optical Unit |  | 25.05.2018 16:30 Fr. Lener 08856936370: System error incubation wheel | Calibration failed (TnI), values too high (650-750), 3x over 700. | Optical Unit renewed, Process Wheel renewed, calibrations made, LQC MCC2 made.
No | Visit | Complaint | RDEU | Bei der Auslieferung und Inbetriebnahme sind Probleme mit der Systemüberprüfung aufgetreten. Intensität Blitzlampe, abweichung >13 % | Ddimer Kal zur Inbetriebnahme des Parameters erforderlich. | DDimere Kal durchgeführt, QC durchgeführt. | SWO-367502 | Assisted Replacement | Consumables | Test Kit | Cartridge, D-Dimer |  | System inspection issues have arisen during delivery and commissioning. Intensity flash lamp, deviation >13% | Ddimer Kal required to commission the parameter. | DDimere Kal performed, QC performed.
No | Visit | Complaint | RCHN | K报阻抗错误，温度下层超出范围 | K报阻抗错误，温度下层超出范围，借903-581一个 | 自行更换903-581后修复 by Andrew | SWO-366918 | Defect | HW | Basic Unit | PCB, Sensor interface | K | K reported impedance error, temperature lower layer out of range | K reported impedance error, temperature lower layer out of range, borrowed 903-581 one | Fix by Andrew after replacing 903-581
No | Visit | Complaint | RCHN | 断电后定标失败 | 断电后定标失败，需要借件944-214（14322） | 重新定标后后修复 by Andrew | SWO-366188 | Assisted Replacement | Consumables | Test Kit | Cal cartridge, Myo |  | Failed to bid after power failure | Failed to bid after power failure, 944-214 (14322) required | Fix by Andrew after re-calibration
No | Visit | Complaint | RKK | ディスプレイ表示しない。電源は入っている。 | ディスプレイ表示しない | 代替機設置、機器引上げ。 | SWO-369573 | Revisit required | HW | analyzer replaced | analyzer replaced | display | Do not display the display. The power is on. | Hide display | Replacement machine installation, equipment pulling up.
No | Visit | Complaint | RCHN | 定标中断 | 定标中断，借944-214（14322）和944-220（14401）各一张 | 重新定标后修复 by Andrew | SWO-367448 | Assisted Replacement | Consumables | Test Kit | Cal cartridge, Myo |  | Calibration interruption | Bid interruption, one each for 944-214 (14322) and 944-220 (14401) | Fix by Andrew after re-calibration
No | Visit | Complaint | RDEU | Keine Ansaugung während Probe erkannt | Fehleranalyse keine Ansaugung während Probe erkannt | Einlassdichtung und Anschlussdichtung erneuert(Kundenmaterial), Servicetests durchgeführt, Probemessung durchgeführt i.O., tHb Kalibrierung durchgeführt, neue Sensorkassette eingesetzt da Kaliumsensor einen Ansprachefehler hatte. Qualitätskontrollen werden von Station überwacht. | SWO-367568 | Defect | HW | Inlet module | connector gasket |  | No suction detected during sample | Error analysis no suction detected during sample | Inlet seal and connection seal renewed (customer material), service tests carried out, sample measurement carried out in the o.o., tHb calibration performed, new sensor cassette inserted because potassium sensor had an address error. Quality controls are monitored by the station.
No | Visit | Complaint | RLTD | Caller Name: SUSAN GOURLAY Telephone No 0191 28 29 240 Prob / Mess HAD RED LIGHTS ON THE GLUCOSE & LACTATE FOR 5 DAYS-SAID TO CHANGE THE SENSOR CASSETTE-DONE SO 5 TIMES BUT NOT SOLVING THE FAULT-PLEASE RING | On inspection of the pins it was found that one had been pushed up and it had not returned to the correct position. | The analyser was powered down and the wet section removed. The pins were cleaned and ensured that they sprung back into the correct position. Analyser was then re assembled and powered on. After start up and calibrations sens had begun to decrease. Suspect faulty sensor interface PCB. A visit will be made tomorrow to replace this. | SWO-367739 | Revisit required | HW | Basic Unit | PCB, Sensor interface | Glu/Lac | Caller Name: SUSAN GOURLAY Telephone No 0191 28 29 240 Prob / Mess HAD RED LIGHTS ON THE GLUCOSE & LACTATE FOR 5 DAYS-SAID TO CHANGE THE SENSOR CASSETTE-DONE SO 5 TIMES BUT NOT SOLVING THE FAULT-PLEASE RING | On inspection of the pins it was found that one had been pushed up and it had not returned to the correct position. | The analyser was powered down and the wet section removed. The pins were cleaned and ensured that they sprung back into the correct position. Analyser was then re assembled and powered on. After start up and calibrations sens had begun to decrease. Suspect faulty sensor interface PCB. A visit will be made tomorrow to replace this.
No | Visit | Complaint | RSCH | Reparatur Austausch | Fehler 1190 während Messungen Fehler Vorort nicht reproduzierbar | Analysator bei RSCH repariert Auto Inlet PCB ausgetauscht Funktionskontrolle i.o Analysator auf Station wieder installiert und IT Anpassungen vorgenommen | SWO-367610 | Defect | HW | Auto Inlet module | PCB, Auto inlet |  | Repair Replacement | Error 1190 during measurements Error suburb not reproducible | Analyzer repaired at RSCH Auto Inlet PCB replaced Functional control i.o Analyzer reinstalled on station and IT adjustments made
No | 3 | Y | RCZE | Blank | Prístroj vydáva zvuk - vysoký tón |  | D47FD9396244DA88 | Defect | HW | Basic Unit | PCB, Wet section |  | Blank | The device issues sound-high tone
No | Visit | Complaint | RDEU | Pat Daten werden, bei Eingabe der Pat ID nicht aus der lokalen Datenbank übernommen | Gerät überprüft | Einstellungen der Pat-Abfrageoption angepasst, | SWO-368015 | Setup | SW | RMED SW | Patient ID |  | Pat data is not taken from the local database when entering the Pat ID | Device checked | adjusted the settings of the Pat query option,
No | Visit | Complaint | RIND | suddenly analyzer display got into blank. | checked the analyzer and found issue with CPU unit . | cross checked with stand by CPU board now display working ,we need to replace CPU board for this issue. | SWO-370368 | [5 empty] | suddenly analyzer display got into blank. | checked the analyzer and found issue with CPU unit . | cross checked with stand by CPU board now display working ,we need to replace CPU board for this issue.
No | Visit | Complaint | RIND | Report coming with ? | Report coming with ? | Found Ca 7.4 is not ok. Recalibrated and found all well. Sampling done. Cleaning done for inlet gasket, inlet tube, inlet probe and electrode and chambers. Analyser Serviced and calibrated properly. Handed over in a fine working condition. | SWO-370345 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | Report coming with ? | Report coming with ? | Found Ca 7.4 is not ok. Recalibrated and found all well. Sampling done. Cleaning done for inlet gasket, inlet tube, inlet probe and electrode and chambers. Analyser Serviced and calibrated properly. Handed over in a fine working condition.
Yes | Visit | Complaint | RAME | Rinse and cleaning errors | Pump rotors not spinning | Replaced pump rotor shaft rubber gaskets to allow rotors to seat properly | SWO-368083 | [5 empty] | Rinse and cleaning errors | Pump rotors not spinning | Replaced pump rotor shaft rubber gaskets to allow rotors to seat properly
No | Visit | Complaint | RIND | Printer not working, paper is not rolling out from printer. | checked and found that the printer head gears were broken, it happens due to mishandling by the user. | tried to fix the gears but they has been broken so it is not possible to repair the printer. printer need to be replace, printer need to be purchase by the customer. For the time being a printer has been fixed on standby till customer will purchase the printer. | SWO-369093 | Defect | HW | Printer Module | Printer head |  | Printer not working, paper is not rolling out from printer. | checked and found that the printer head gears were broken, it happens due to mishandling by the user. | tried to fix the gears but they has been broken so it is not possible to repair the printer. printer need to be replace, printer need to be purchase by the customer. For the time being a printer has been fixed on standby till customer will purchase the printer.
No | Visit | Complaint | RSCH | Rep. FlexQ | Sporadisch Fehler 882 Mischerträger Schlupf erkannt FlexQ bereits mehrmals kalibriert | Mischerträger ersetzt=> Garantie FlexQ kalibriert i.o Testmessung alle drei Positionen i.o | SWO-367797 | Defect | HW | FlexQ | Mixer Tray | 882 | Rep. FlexQ | Sporadic Error 882 Mixer carrier slip detected FlexQ calibrated several times | Mixer carrier replaced=> Warranty FlexQ calibrates i.o Test measurement of all three positions i.o
No | Visit | Complaint | RSCH | Rep. FlexQ | Erneut Fehler 882 Mischerträger Schlupf erkannt Mischerträger bereits 1x ersetzt | Mischerträger erneut ersetzt=> Garantie FlexQ kalibriert i.o Testmessungen alle drei Positionen i.o | SWO-367798 | Defect | HW | FlexQ | Mixer Tray |  | Rep. FlexQ | Again Error 882 Mixer carrier slip detected Mixer carrier already replaced 1x | Mixer carrier replaced=> Warranty FlexQ calibrates i.o Test measurements of all three positions i.o
No | Visit | Complaint | RAME | Error in multiple samples 655: Sample in pH/BG upper liquid sensor detected prematurely during aspiration 293: Detected FHbF is greater than 20 % of ctHb. Oxi parameters are automatically corrected. 476: Measurement unstable 664: A problem was detected during sample transport through the analyzer. Contact: John -734-246-6037 | Per phone conversation found issue was pH sealing print. | Replaced pH ring and cleaned and checked measuring chambers. Found pO2 sealing ring was missing, replaced, | SWO-367934 | User error | Electrode | Electrode | Sealing ring |  | Error in multiple samples 655: Sample in pH/BG upper liquid sensor detected prematurely during aspiration 293: Detected FHbF is greater than 20 % of ctHb. Oxi parameters are automatically corrected. 476: Measurement unstable 664: A problem was detected during sample transport through the analyzer. Contact: John -734-246-6037 | Per phone conversation found issue was pH sealing print. | Replaced pH ring and cleaned and checked measuring chambers. Found pO2 sealing ring was missing, replaced,
No | 5 | Y | RCZE | Blank | nejde nakalibrovat pH, ani po precisteni, dekontaminaci, ani po vymene ref membrany chyby 376,377,378,379 |  | D4811056D80C3FB4 | Revisit required | Electrode | Electrode | pH |  | Blank | No calibration of pH, even after precisteni, decontamination, even after disinfers ref Membrany error 376,377,378,379
No | Visit | Complaint | RDEU | 07.06.2018, 08:59 Uhr, Fr. Jannick, Tel. 0211904361209: Messung wird nach 2 Min. abgebrochen 17.06.2018, 08:40 Uhr, Fr. Wiegand, Tel. 0211904361209: Messung wird abgebrochen, Bildschirm flackert, wenn er bewegt wird, Scanner scannt nicht die komplette Pat. ID. | Patientenmessungen werden nach kurzer Zeit abgebrochen mit den Fehlern 5689, 5691 und 5694 Rotatestoragewheel 24, 28 und 40. Cartridge werden nicht alle erkannt. | Fließwege mit Hypochlorid gereinigt, Storagewheel ausgebaut und gereinigt, Sensoren gereinigt und Motoren gefettet, Scheibe des Cartridge Barcodescanners gerissen - nicht auf Lager, Scheibe entfernt und Barcode neu abgeglichen. Cartridge werden wieder alle erkannt, Patienten-Barcode Scanner gereingt, Systemüberprüfung, Systemreinigung und QC’s durchgeführt. Neue Abdeckung mit Scheibe bestellt. | SWO-368932 | Cleaned | HW | Storage Module | Storage wheel | 5689, 5691, 5694 | 07.06.2018, 08:59, Fr. Jannick, Tel. 0211904361209: Measurement is aborted after 2 min. 17.06.2018, 08:40, Ms. Wiegand, tel. 0211904361209: Measurement is aborted, screen flickers when it is moved, scanner does not scan the complete pat. Id. | Patient measurements are aborted after a short time with errors 5689, 5691 and 5694 Rotatestoragewheel 24, 28 and 40. Cartridges are not all recognized. | Flow paths cleaned with hypochloride, storagewheel removed and cleaned, sensors cleaned and motors greased, disc of the cartridge barcode scanner ripped - not in stock, disc removed and barcode rebalanced. Cartridges are again recognized all, patient barcode scanner setherin, system verification, system cleaning and QC's performed. New cover with disc ordered.
No | Visit | Complaint | RDEU | 043116972333 Hr. Feldt So2 bei QC außerhalb Fehler wieder aufgetreten | QC SO2 Level B außer Bereich. Trotz Tausch des LP Fehler bleibt bestehen. | Hemolyzer getauscht. | SWO-368342 | Defect | HW | Oxi module | Hemolyzer | sO2 | 043116972333 Mr. Feldt So2 reappeared at QC outside error | QC SO2 Level B out of range. Despite exchanging the LP error, it remains. | Hemolyzer.
No | Visit | Complaint | RDEU | 043116972333 Hr. Feldt So2 bei QC außerhalb Fehler wieder aufgetreten | System überprüft => Hemolyser defekt | Hemolyser erneuert , Kalibration und Qc durchgeführt | SWO-368700 | Defect | HW | Oxi module | Hemolyzer | sO2 | 043116972333 Mr. Feldt So2 reappeared at QC outside error | System checked => Hemolyser defective | Hemolyser renewed , calibration and Qc performed
No | Visit | Complaint | RLTD | Caller Name: SUSAN GOURLAY Telephone No 0191 28 29 240 Prob / Mess HAD RED LIGHTS ON THE GLUCOSE & LACTATE FOR 5 DAYS-SAID TO CHANGE THE SENSOR CASSETTE-DONE SO 5 TIMES BUT NOT SOLVING THE FAULT-PLEASE RING | Faulty sensor interface PCB | Powered analyser down and removed wet section to allow replacement of the PCB. After initial start up the sensor cassette was replaced and left to condition. | SWO-368541 | Defect | HW | Basic Unit | PCB, Sensor interface | Glu/Lac | Caller Name: SUSAN GOURLAY Telephone No 0191 28 29 240 Prob / Mess HAD RED LIGHTS ON THE GLUCOSE & LACTATE FOR 5 DAYS-SAID TO CHANGE THE SENSOR CASSETTE-DONE SO 5 TIMES BUT NOT SOLVING THE FAULT-PLEASE RING | Faulty sensor interface PCB | Powered analyser down and removed wet section to allow replacement of the PCB. After initial start up the sensor cassette was replaced and left to condition.
Yes | 9 | Y | RPAC | Blank | Sensor cassette failure (1361) - Internal reference electrode error in sensor cassette. |  | D48880F69C42948D | [5 empty] | Blank | Sensor cassette failure (1361) - Internal reference electrode error in sensor cassette.
Yes | Visit | Complaint | RLTD | Rob 01772523007 | site visit to collect fresh service dump for RMED investigation | arrived on site and attended analyser. took service dump to send to RMED for investigation and feedback on resolution. | SWO-368625 | Re classified RMED | planned SWO | SD for RMED |  |  | Rob 01772523007 | site visit to collect fresh service dump for RMED investigation | arrived on site and attended analyser. took service dump to send to RMED for investigation and feedback on resolution.
Yes | Visit | Complaint | RLTD | Rob 01772523007 | site visit to collect fresh service dump for RMED investigation | arrived on site and attended analyser. took service dump to send to RMED for investigation and feedback on resolution. | SWO-368627 | Re classified RMED | planned SWO | SD for RMED |  |  | Rob 01772523007 | site visit to collect fresh service dump for RMED investigation | arrived on site and attended analyser. took service dump to send to RMED for investigation and feedback on resolution.
No | Visit | Complaint | RIND | Received a complaint from the Biomedical Engineer that, leakage from the back of SC on ABL80 analyzer at Inamdar Hospital, Pune | Need to check the Sensor Board for any damage. | Kept the machine OFF for 30 minutes, tried drying the connector board, after that reinstalled a new SC, machine was giving the 'SC temperature not stable' message. So installed the SC on the backup ABL80 analyzer. Need to replace the Sensor Connector Board. | SWO-370900 | [5 empty] | Received a complaint from the Biomedical Engineer that, leakage from the back of SC on ABL80 analyzer at Inamdar Hospital, Pune | Need to check the Sensor Board for any damage. | Kept the machine OFF for 30 minutes, tried drying the connector board, after that reinstalled a new SC, machine was giving the 'SC temperature not stable' message. So installed the SC on the backup ABL80 analyzer. Need to replace the Sensor Connector Board.
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | PCO2,NA,LAC,GLU定标不过 0210 | 更换试剂包后修复 | SWO-373163 | [5 empty] | Test card, reagent package failure | PCO2, NA, LAC, GLU calibration but 0210 | Repair after replacement of reagent kit
No | Visit | Complaint | RAME | Ça |  |  | SWO-369354 | Assisted Replacement | Electrode | Membrane | Ca |  | it
No | Visit | Complaint | RDEU | 04.05.2018; Hr. Henneboele; Tel.: 02093643880; Na und Ca Kal Fehler 04.05.2018; 11:40 Uhr; Hr. Henneboele; Tel.:02093644504; weiterhin Na und Ca Driftfehler 11.06.2018; 11:15 Uhr; Hr. Hennboehle; Tel.: 02093643880; nach der Na und Ca Driftfehler besteht weiterhin 18.06.2018; 18:15 Uhr; Sr. Birgit; Tel.: 02093643850; der Ca Kal Fehler ist wieder aufgetreten | Sporadische Driften | Daten ausgewertet, pH und K Driften auch vorhanden aber innerhalb der Herstellergrenzen. E1001 ersetzt, ABL mehrfach kalibriert, Funktion z.Zt. OK. | SWO-368920 | Assisted Replacement | Electrode | Electrode | Reference | CAPA-00733 | 04.05.2018; Mr. Henneboele; Tel.: 02093643880; Na and Ca Kal Error 04.05.2018; 11:40 a.m.; Mr. Henneboele; Tel.:0 2093644504; continue Na and Ca drift errors 11.06.2018; 11:15 a.m.; Mr. Hennboehle; Tel.: 02093643880; after the Na and Ca drift error, there is still 18.06.2018; 6:15 p.m.; Sr. Birgit; Tel.: 02093643850; the Ca Kal error has reoccurred | Sporadic drifts | Data evaluated, pH and K drift also present but within the manufacturer's limits. E1001 replaced, ABL calibrated several times, function currently OK.
No | Visit | Complaint | RKK | オキシキャルエラー | Halogen lamp voltage too low Halogen lamp intensity too high | Replacement lamp unit | SWO-368960 | Defect | HW | Oxi module | Lamp unit | halogen | Oxycal error | Halogen lamp voltage too low Halogen lamp intensity too high | Replacement lamp unit
No | Visit | Complaint | RAME | OXI lamp high | Oxi lamp high 4095 | Adjusted lamp to 3300 and ran service THB cal. | SWO-368641 | Adjustment | HW | Oxi module | Lamp unit |  | OXYGEN Lamp High | Oxygen Lamp High 4095 | Adjusted lamp to 3300 and ran service THB cal.
No | Visit | Complaint | RDEU | 19.06.2018, Frau Dohm, Tel: 06946086121: setsafe 22; AQT macht sehr laute Geräusche | Gerät überprüft, Alle Safepositionen werden im Service Menü erreicht, prosesswheel, storagewheel, getestet, alles i.O Gerät nach Service initialisiert und ist Messbereit. Blank Messungen durchgeführt, alles i.O | Setsafe 22 und Prosesswheel 40 tauchten gleichzeitig 2 mal bis jetzt auf. Servicedump erstellt, ggf. Weitergabe nach CPH. Kunden informiert. | SWO-368642 | No fault found | No fault found | No fault found | No fault found |  | 19.06.2018, Mrs Dohm, Tel: 06946086121: setsafe 22; AQT makes very loud noises | Device checked, All safe positions are reached in the service menu, prosesswheel, storagewheel, tested, all i.O device initialized by service and is ready for measurement. Blank measurements performed, all i.O | Setsafe 22 and Prosesswheel 40 appeared at the same time 2 times until now. Servicedump created, if necessary, propagation to CPH. informed customers.
Yes | Visit | Complaint | RKK | 装置使用中にエラー発生 | 状態確認 | デバイスによる清掃作業 | SWO-373323 | [5 empty] | Error occurring while using device | Status Check | Cleaning with devices
No | 1 | Y | RPAC | Blank | Fixed Assinged QC ranges enabled |  | D48338C9C45B65C6 | Setup | SW | RMED SW | Fixed assigned values |  | Blank | Fixed Assinged QC ranges enabled
Yes | 7 | Y | RPAC | Blank | pCO2 low sensitivities |  | D483703EC4571A29 | [5 empty] | Blank | pCO2 low sensitivities
No | Visit | Complaint | RDEU | 5.06.2018 15:40Uhr VonKanen 023195057411: Es gab Abweichungen bei tHb zum Labor. Die Werte schwanken. 13.06.2018, 10:22 Uhr, H. Söderberg, Tel. 023195057413: tHb Werte zu niedrig 19.06.2018, 11.45Uhr, Herr Söderbeck,Tel: 023195057410: weiterhn tHb Werte zu niedrig. | Gerät überprüft | Quetschventil erneuert, Gerät neu gestartet, Startphase abgewartet | SWO-369105 | Defect | HW | Oxi module | Pinch valve |  | 5.06.2018 15:40H VonKanen 023195057411: There were deviations at tHb to the laboratory. The values fluctuate. 13.06.2018, 10:22, H. Söderberg, tel. 023195057413: tHb values too low 19.06.2018, 11.45am, Mr. Söderbeck,Tel: 023195057410: continue tHb values too low. | Device checked | Squeeze valve renewed, device restarted, start-up phase awaited
No | Visit | Complaint | RDEU | Fehlerbehebung erforderlich | Meldung in Aktivitäten: LIS $KIS-Verbindung verloren und LIS $KIS-Verbindung hergestellt mehrmals innerhalb einer Minute bei Verbindung über WLAN LAN-Verbindung ohne Probleme | WLAN-Verbindung unter Windows stabil IT-Abteilung des Hauses hat bestätigt, dass die WLAN-Verbindung am Router ebenfalls stabil ist Servicedump gezogen | SWO-369160 | Revisit required | SW | SW | Wireless connection |  | Troubleshooting required | Message in activities: LIS /KIS connection lost and LIS /KIS connection established several times within a minute when connected over Wi-Fi LAN connection without problems | Wi-Fi connection stable on Windows IT department of the house has confirmed that the Wi-Fi connection on the router is also stable Service dump pulled
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 报错1305 | 更换测试卡后修复 | SWO-368943 | [5 empty] | Test card, reagent package failure | 1305 error | Repair after replacing test card
No | Visit | Complaint | RAME | SP leaking | Customer installed SP yesterday mid day. Came into work this morning and noticed pool of red $pink fluid collecting under analyzer. I did more cleaning of where leak occurred. Appeared that the SP was leaking internally. Waste manifold was fine. | Replaced SP. Ran several system cycles and QC. No leaking noted. | SWO-369324 | [5 empty] | SP leaking | Customer installed SP yesterday mid day. Came into work this morning and noticed pool of red /pink fluid collecting under analyzer. I did more cleaning of where leak occurred. Appeared that the SP was leaking internally. Waste manifold was fine. | Replaced SP. Ran several system cycles and QC. No leaking noted.
Yes | Visit | Complaint | RCHN | 定标底部外露，光学单元超出范围，报错1024和1031 | 定标底部外露，光学单元超出范围，报错1024和1031,需要借件903-364和903-236 | 更换903-364和903-236后修复 by jesson | SWO-368334 | [5 empty] | The bottom of the calibration is exposed, the optical unit is out of range, and the errors 1024 and 1031 are reported | The bottom of the calibration is exposed, the optical unit is out of range, the error 1024 and 1031 are reported, the borrowings are 903-364 and 903-236 | Repair by jesson after replacing 903-364 and 903-236
No | Visit | Complaint | RCHN | CL电极定标不稳 | CL电极定标不稳，借945-617一个 | 用库存945-617的修复 by jesson | SWO-368895 | Assisted Replacement | Electrode | Electrode | Cl |  | CL electrode calibration is not stable | CL electrode calibration is not stable, borrow 945-617 one | Fix with stock 945-617 by jesson
Yes | 13 | Y | RPAC | Blank | Annual Service |  | D482553D4E24B075 | [5 empty] | Blank | Annual Service
Yes | Visit | Complaint | RDEU | 30.05.2018; Hr. Uhl; Tel.: 07221912179; Uhrzeit verstellt sich immer wieder 14.06.2018 H. Dürr 07221912178: Anwender bittet weiterhin um einen FSE-Einsatz am Gerät | Uhrzeit verstellt sich | Uhrzeit gemäß Anleitung in der Windows Konfig. angepasst | SWO-370163 | [5 empty] | 30.05.2018; Mr. Uhl; Tel.: 07221912179; Time is constantly dislocated 14.06.2018 H. Dürr 07221912178: User continues to ask for FSE use on the device | Time is changed | Time according to the instructions in the Windows Config. Adapted
No | Visit | Complaint | RDEU | Akku Anzeige auf 0% |  |  | SWO-370438 | Defect | HW | Battery module | Battery module |  | Battery display to 0%
Yes | 3 | Y | RPAC | Blank | Replaced failed SC |  | D484CCD63A441A8D | [5 empty] | Blank | Replaced failed SC
Yes | 8 | Y | RPAC | Blank | Sensor cassette pCO2 calibration failure. |  | D48C510BFE058490 | [5 empty] | Blank | Sensor cassette pCO2 calibration failure.
Yes | 5 | Y | RCZE | Blank |  |  | D48332B2C400DA95 | Re classified RMED | planned SWO |  |  |  | Blank
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 病人结果与临床不符 | 更换测试卡后修复 | SWO-369596 | [5 empty] | Test card, reagent package failure | Patient results are not clinically consistent | Repair after replacing test card
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 试剂包漏电流 | 更换试剂包后修复 | SWO-374133 | [5 empty] | Test card, reagent package failure | Reagent packet leakage current | Repair after replacement of reagent kit
No | Visit | Complaint | RCHN | 报错1352 | 报错1352 借903-199一个 | 更换903-199备件后修复 by jesson | SWO-368356 | Defect | HW | Basic Unit | Cassette retainer |  | 1352 error | Error 1352 Borrowing 903-199 one | Repair by jesson after replacement of 903-199 spare parts
Yes | Visit | Complaint | RCHN | 光单故障 | 光单故障，借903-364一个 | 更换了903-364后修复 by jesson | SWO-364956 | [5 empty] | Light single fault | Light single fault, borrow 903-364 one | Repaired by jesson after replacing 903-364
No | Visit | Complaint | RCHN | 报错1223 1259 | 报错1223 1259 借903-329一个 | 更换903-329后修复 by jesson | SWO-369024 | Defect | HW | Aspiration Module | Needle Wash Unit | 1223 | Error 1223 1259 | Error 1223 1259 Borrow 903-329 one | Repair by jesson after replacing 903-329
No | Visit | Complaint | RCHN | 报错1068 | 报错1068 借903-210一个 | 更换903-210一个正常。by jesson | SWO-368962 | Defect | HW | Process Module | PCB, Cup dryer | 1068 | 1068 error | False 1068 Borrow 903-210 one | Replace the 903-210 one normal. by jesson
No | Visit | Complaint | RCHN | 样本无法固定 | 样本无法固定 借903-584 2个 | 更换903-584 2个 后修复 by jesson | SWO-369039 | Defect | HW | Aspiration Module | Tupe holder |  | Sample cannot be pinned | Samplecan cannot be fixed borrowing 903-584 2 | Replace 903-584 2 after repair by jesson
No | Visit | Complaint | RCHN | 最大暗视野测试超范围 1028 | 最大暗视野测试超范围 1028,借903-364一个。 | 更换903-364备件后修复 by jesson | SWO-368350 | Defect | HW | Process Module | Optical Unit | 1028 | Maximum dark field test over range 1028 | Maximum dark field test over range 1028, borrowing 903-364 one. | Repair by jesson after replacement of 903-364 spare parts
No | Visit | Complaint | RCHN | 报错1035 | 报错1035 借903-364一个 | 更换备用的903-364后修复 by jesson | SWO-368995 | Defect | HW | Process Module | Optical Unit | 1035 | 1035 error | False 1035 Borrow 903-364 one | Repair by after replacing the spare 903-364 jesson
No | Visit | Complaint | RKK | オキシキャルエラー | オキシキャルエラー Neon lamp intensity too low Neon number too low サービスオキシテストにてNeon intensity failed | ネオンランプ交換 | SWO-372080 | Defect | HW | Oxi module | Lamp unit | Neon low | Oxycal error | Oxycal error Neon lamp intensity too low Neon number too low Neon intensity failed in service oxytest | Neon Lamp Exchange
No | 3 | Y | RPAC | Blank | Lactate QC6+ level 3 issues. |  | D48C4048FE62ABCA | Assisted Replacement | Electrode | Electrode | Lactate |  | Blank | Lactate QC6+ level 3 issues.
Yes | 1 | Y | RPAC | Blank |  |  | D484CEE83A52E88B | [5 empty] | Blank
No | Visit | Complaint | RLTD | Lorriane Hambleton 01324 566137 | Inlet module broken | Replaced inlet module with new | SWO-370619 | Defect | HW | Inlet module | Inlet module |  | Lorriane Hambleton 01324 566137 | Inlet module broken | Replaced inlet module with new
No | Visit | Complaint | RBNL | Sensor bay sluit niet meer bij het vervangen van een sensorcassette. | Sensor bay sluit niet na vervangen van sensorcassette. Bay wil dichtgaan, maar laatste cm blijft open, waarna bay helemaal open gaat staan. Gedemonteerd en het lipje dat van de bovenste klep in de sensor zit iets rechter gebogen. Na montage en opstart werkt alles weer normaal. Sensor cassette vervangen en ABL opgestart. Gekalibreerd en QC gemeten —> oke. | Lip van klep die in sensor zit rechtgebogen | SWO-370471 | Adjustment | HW | Basic Unit | Sensor door |  | Sensor Bay no longer closes when replacing a sensor cassette. | Sensor Bay does not close after replacing sensor cassette. Bay wants to go close, but last cm remains open, after which Bay will be completely open. And the tab that sits from the top flap in the sensor slightly right bent. After assembly and start-up everything works normally again. Replace the Sensor cassette and ABL. Calibrated and QC measured — > okay. | Lip of valve that sits in sensor straightened
No | Visit | Complaint | RDEU | Station wünscht Überprüfung des Gerätes In letzter Zeit vermehrt LP, SC und Einlassdichtungswechsel | Geräteüberprüfung | Überprüfung Dateiarchiv Betriebsmittel => in der letzten Woche mehrmals LP und Einlassdichtung gewechselt => Credit Claims ausgedruckt Überprüfung Einlassdichtung, LP und SC über Service Reinigung Flow Selector Einlassnadel leicht verbogen => Nadel gerichtet Durchführung alle Fluid Programme => keine Fehler Info an Personal zwecks Handhabung der Reinigung der Einlassdichtung und Wiederverwendung und Erstellung der Reklamationsmeldung | SWO-370615 | Adjustment | HW | Auto Inlet module | Probe bend |  | Station wants to check the device Recently, LP, SC and inlet seal changes have increased | Geräteüberprüfung | Checking File Archive Equipment => Changed LP and Inlet Seal several times in the last week => Credit Claims printed Inspection inlet seal, LP and SC via service Cleaning Flow Selector Inlet needle slightly bent => Needle directed Execution of all fluid programs => no errors Info to personnel for handling the cleaning of the inlet seal and reuse and preparation of the complaint notification
No | Visit | Complaint | RFRA | 1271 Echec aspiration échantillon. | Problème aspiration | Chgt carte détection aiguille. Test avec 2 gazométries | SWO-370482 | Defect | HW | Auto Inlet module | PCB, Auto inlet | HALL Element | 1271 Sample suction failure. | Aspiration problem | Chgt map needle detection. Test with 2 gasmetry
No | Visit | Complaint | RFRA | L’aspiration ne démarre pas lors de la présentation de la seringue même jusque-là butée. | Problème d'aspiration | Chgt carte détection aiguille. Test avec 2 gazométries | SWO-370591 | Defect | HW | Auto Inlet module | PCB, Auto inlet | HALL Element | The suction does not start when the syringe is presented even before it has been stumbled. | Aspiration problem | Chgt map needle detection. Test with 2 gasmetry
Yes | 5 | Y | RCZE | Blank | Nefunkèný prenos výsledkov do LIS |  | D4848A6F3A26DA9D | Re classified RMED | planned SWO |  |  |  | Blank | Failure to transfer results to LIS
No | Visit | Complaint | RDEU | HB Abweichungen bei den Messungen | Gerät überprüft und gereinigt | Quetschventil erneuert, Startphase abgewartet | SWO-370629 | Defect | HW | Oxi module | Pinch valve |  | HB Deviations in measurements | Device checked and cleaned | Squeeze valve renewed, start phase awaited
No | 3 | Y | RCZE | Blank |  |  | D484BDF83A4DDAA5 | Setup | SW | RMED SW | Parameter |  | Blank
Yes | Visit | Complaint | RCHN | 开机报错 | 开机报错，需要借件903-228和902-972 | 更换备用903-228和902-972后修复 by jesson | SWO-369549 | [5 empty] | Boot error | Boot error, need to borrow 903-228 and 902-972 | Repair by jesson after replacing spare 903-228 and 902-972
No | Visit | Complaint | RAME | PO2 failing external QC on level 2 and 3 Customer dissatisfied with analyzer | Checked data logs. Found errors. | Ran the zeroing cassette to zero channels. Ran system cycle and then QC level 2. Analyzer QC levels 2 and 3 now within range. | SWO-370804 | [5 empty] | PO2 failing external QC on level 2 and 3 Customer dissatisfied with analyzer | Checked data logs. Found errors. | Ran the zeroing cassette to zero channels. Ran system cycle and then QC level 2. Analyzer QC levels 2 and 3 now within range.
No | Visit | Complaint | RCHN | 报错1034,1020 | 报错1034,1020，借903-197一个 | 更换903-197一个正常。by jesson | SWO-369727 | Defect | HW | Oxi module | Hemolyzer |  | Error 1034,1020 | 1034,1020, loan 903-197 one | Replace 903-197 one normal. by jesson
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | PCO2定标不过 0210 | 更换测试卡后修复 | SWO-373164 | [5 empty] | Test card, reagent package failure | PCO2 calibration is no more than 0210 | Repair after replacing test card
No | 1 | Y | RPAC | Blank | Software Upgrade TCM5 to Version V1.2, FAN 915-380. |  | D484AA793A137697 | [5 empty] | Blank | Software Upgrade TCM5 to Version V1.2, FAN 915-380.
No | 1 | Y | RPAC | Blank | Software Upgrade unit to Version V1.2, FAN 915-380. |  | D484A4BC3A567694 | [5 empty] | Blank | Software Upgrade unit to Version V1.2, FAN 915-380.
Yes | Visit | Complaint | RCHN | 报错5658 | 报错5658 | 重新清洁了5个传感器和3个转轮 by jesson | SWO-370412 | [5 empty] | 5658 error | 5658 error | Re-cleaning 5 sensors and 3 wheels by jesson
Yes | Visit | Complaint | RCHN | 打印机故障 | 打印机故障 借910-275一个 | 更换910-275后修复 by jesson | SWO-369773 | [5 empty] | Printer failure | Printer failure Borrowing 910-275 one | Repair by jesson after replacement 910-275
Yes | Visit | Complaint | RIND | Pco2 not calibarting | Pco2 not calibarting | checked the analyser calibartion status pco2 sense values were below range so removed the electrode and cleaned the puss and chamber with gause then given calibartion then also pco2 failed to calibrate, then changed the pco2 membarne to chekcking purpose then after that also poc2 not came to raedy mode, so suspected with pco2 electrode that needs to be checked | SWO-370760 | [5 empty] | Pco2 not calibarting | Pco2 not calibarting | checked the analyser calibartion status pco2 sense values were below range so removed the electrode and cleaned the puss and chamber with gause then given calibartion then also pco2 failed to calibrate, then changed the pco2 membarne to chekcking purpose then after that also poc2 not came to raedy mode, so suspected with pco2 electrode that needs to be checked
Yes | Visit | Complaint | RLTD | John Meaeher Tel: 0782 333 3757 | “Parameters not installed”. | Software license keys were installed, but suspect corruption. Reordered & installed license keys. Cleaned electrode chambers. Removed & cleaned contacts for glucose, lactate & sodium. | SWO-370413 | Re classified RMED | planned SWO |  |  |  | John Meaeher Tel: 0782 333 3757 | “Parameters not installed”. | Software license keys were installed, but suspect corruption. Reordered & installed license keys. Cleaned electrode chambers. Removed & cleaned contacts for glucose, lactate & sodium.
No | Visit | Complaint | RAME | 0509 Rinse error. Clot in system. | Inspection of the fluidics revealed the pH $BG V7 vent valve rubber membrane is torn allowing fluid to get into the valve. Customer states this started with a clot in the pH chamber that they were able to cleared out of system. I found the pH lower sealing ring detached in the bottom of the pH well. Crea & Lac initially weren't passing, Lac electrode was remembraned and needed to equilibrate. | Replaced the pH $BG V7 vent valve. Cleaned and inspected the pH & ref chambers. Replaced the pH lower sealing ring. Lac membrane was replaced. | SWO-370277 | Defect | HW | pH/BG Module | pH/BG Valve |  | 0509 Rinse error. Clot in system. | Inspection of the fluidics revealed the pH /BG V7 vent valve rubber membrane is torn allowing fluid to get into the valve. Customer states this started with a clot in the pH chamber that they were able to cleared out of system. I found the pH lower sealing ring detached in the bottom of the pH well. Crea & Lac initially weren't passing, Lac electrode was remembraned and needed to equilibrate. | Replaced the pH $BG V7 vent valve. Cleaned and inspected the pH & ref chambers. Replaced the pH lower sealing ring. Lac membrane was replaced.
No | Visit | Complaint | RKK | Po2zero high(236mmHg） | Po2zero high(236mmHg） | 回路洗浄 メンブラン交換実施 | SWO-371163 | Assisted Replacement | Electrode | Membrane | pO2 |  | Po2zero high(236mmHg） | Po2zero high(236mmHg） | Circuit cleaning membrane replacement
No | Visit | Complaint | RCHN | 定标无法通过 | 定标无法通过，借945-377一个 | 更换了945-377后修复 by jesson | SWO-370497 | [5 empty] | The bid failed to pass | The bid cannot be passed, borrow 945-377 one | Repaired by jesson after replacing 945-377
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | CL,PCO2,PH,NA定标不过 1319,1315,1321 | 更换测试卡后修复 | SWO-371050 | Assisted Replacement | SC | SC | SC |  | Test card, reagent package failure | CL, PCO2, PH, NA calibration is nothing but 1319,1315,1321 | Repair after replacing test card
No | Visit | Complaint | RCHN | LAC,GLU定标不过 | LAC,GLU定标不过 借891-051 32个 | 更换891-051后修复 by jesson‘ | SWO-369661 | Defect | HW | Basic Unit | Sensor contact |  | LAC, GLU calibration but | LAC, GLU calibration but borrow 891-051 32 | Repair by jesson' after replacing 891-051
No | Visit | Complaint | RCHN | 加热传感器错误 | 加热传感器错误，需要借件903-581一个 | 更换了903-581后修复 by jesson | SWO-370315 | Defect | HW | Basic Unit | PCB, Sensor interface | Temp | Heating sensor error | Heating sensor error, need to borrow 903-581 one | Repair edited after replacing 903-581 by jesson
No | Visit | Complaint | RCHN | 报错5257 | 报错5257 借944-394 （14411）一张 | 重新定标后修复 by jesson | SWO-369709 | Defect | HW | Process Module | Sensor, Process Unit | 5257 | 5257 error | Error 5257 Borrow 944-394 (14411) a | Fix by jesson after re-calibration
No | Visit | Complaint | RCHN | 报错1029 | 报错1029，需要借件903-364和903-236 | 更换了903-364和903-236后修复 by jesson | SWO-370403 | Defect | HW | Process Module | Optical Unit | 1029 | 1029 error | 1029, need to borrow 903-364 and 903-236 | Repair edited after replacing 903-364 and 903-236 by jesson
No | Visit | Complaint | RAME | Rinse error and missing option keys | Inspection of the system revealed there are rinse errors above option key issue. | Replaced pH BG and El Met rinse and vent valves and Fluidic block V6 & V7 valves. | SWO-369609 | Defect | HW | pH/BG Module | pH/BG Valve |  | Rinse error and missing option keys | Inspection of the system revealed there are rinse errors above option key issue. | Replaced pH BG and El Met rinse and vent valves and Fluidic block V6 & V7 valves.
No | Visit | Complaint | RCHN | 报错0529，液体传感器失败 | 报错0529，液体传感器失败，需要借件903-199和902-925各一个 | 更换903-199和902-925备件后修复 | SWO-370905 | Defect | HW | Basic Unit | Cassette retainer |  | Error 0529, liquid sensor failed | Error 0529, liquid sensor failure, need to borrow 903-199 and 902-925 each one | Repair after replacement of 903-199 and 902-925 spare parts
No | Visit | Complaint | RCHN | 测量时po2无法识别 | 测量时血氧值无法测量，需要借件902-980 | 更换了902-980后修复 by jesson | SWO-370501 | Defect | HW | Oxi module | Hemolyzer |  | Po2 is not recognized when measuring | Blood oxygen value cannot be measured when measured, 902-980 is required | Repaired by jesson after replacing 902-980
No | Visit | Complaint | RCHN | 报错1024 | 报错1024，借件903-236 |  | SWO-370970 | Defect | HW | Process Module | Optical Unit | 1024 | 1024 error | 1024, loan 903-236
Yes | Visit | Complaint | RCHN | 报错1999 | 报错1999，需要借件903-113 | 更换903-113后修复 by jesson | SWO-370496 | [5 empty] | Error 1999 | Reported error 1999, need to borrow 903-113 | Repair by jesson after replacement 903-113
No | Visit | Complaint | RCHN | 报错376576 | 报错376576，需要借件902-823和945-603 |  | SWO-370969 | Defect | HW | pH/BG Module | pH/BG Module |  | Error 376576 | 376576, need to borrow 902-823 and 945-603
No | Visit | Complaint | RLTD | Christine Robertson Tel: 0131 242 2601 | Analyser failing calibration and QCs for pC02 and lactate, sensitivity not rising above 5 % | Customer already replaced SC and SP with new Replaced SC with new again New Interface PCB required, ordered to arrive tomorrow | SWO-371503 | Revisit required | HW | Basic Unit | PCB, Sensor interface | pCO2, Lac | Christine Robertson Tel: 0131 242 2601 | Analyser failing calibration and QCs for pC02 and lactate, sensitivity not rising above 5 % | Customer already replaced SC and SP with new Replaced SC with new again New Interface PCB required, ordered to arrive tomorrow
No | Visit | Complaint | RLTD | Gary Taylor Tel: 01604 54 56 11 | Analyser in intervention screen asking for solution pack change, unable to pump solutions. | Entered service, ran fluid transport tests failing vacuum, opened sensor bay, liquid viable under pinch valve tubing. Replaced pinch valve tubing, now ok. | SWO-371641 | Defect | HW | Oxi module | Pinch valve tube |  | Gary Taylor Tel: 01604 54 56 11 | Analyser in intervention screen asking for solution pack change, unable to pump solutions. | Entered service, ran fluid transport tests failing vacuum, opened sensor bay, liquid viable under pinch valve tubing. Replaced pinch valve tubing, now ok.
No | Visit | Complaint | RDEU | 25.06.2018, 08:58 Uhr, Fr. Fischer, Tel. 023195057600: Blutprobe wird nicht angesaugt | Gerät überprüft und gereinigt | Einlassdichtung gereinigt, mehrfache Testmessungen durchgeführt, Gerät iO. | SWO-371345 | Cleaned | HW | Auto Inlet module | Inlet gasket |  | 25.06.2018, 08:58, Ms. Fischer, tel. 023195057600: blood sample is not sucked in | Device checked and cleaned | Inlet seal cleaned, multiple test measurements carried out, device iO.
No | Visit | Complaint | RDEU | Benötigt viel Zeit für die Flüssigkeitssensorkalibrierung, häufige Meldung : mögliche Verstopfung der Fließwege | Flüssigkeitssensor defekt | Fehlersuche, Austausch Flüssigkeitssensor, Funktionskontrolle | SWO-371338 | [5 empty] | Requires a lot of time for liquid sensor calibration, frequent message: possible blockage of the flow paths | Liquid sensor defective | Troubleshooting, replacement of liquid sensor, function control
No | Visit | Complaint | RCHN | 断电后无法开机 | 断电后无法开机，借903-381一个 |  | SWO-371072 | [5 empty] | Unable to power on after power failure | Cannot power on after power failure, borrow 903-381 one
Yes | Visit | Complaint | RCHN | 报错1031 | 报错1031，借903-364一个 |  | SWO-371078 | [5 empty] | 1031 | 1031, loan 903-364 one
No | Visit | Complaint | RIND | Creatinine low | Check Calibration reports, Senstivity of Crea is low | Protien removal is run, 2point calibration done,Calibration report shows senstivity is ok,. Sample result is satisfactory | SWO-374703 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | Creatinine low | Check Calibration reports, Senstivity of Crea is low | Protien removal is run, 2point calibration done,Calibration report shows senstivity is ok,. Sample result is satisfactory
No | Visit | Complaint | RDEU | 25.06.2018, 10:22 Uhr, fr. Bretting, Tel. 030940113280: Oxymetrie Hardware Fehler | Reinigungsarbeiten, Fehlersuche | OxiHardwarefehler ist beseitig nach Neustart, Kontrolle CreaElektroden und Messkammern, Kontrolle Lösungstransport, Luftblasen aus Elektroden entfernt, Kalibrierlösungen erneuert, Crea deaktiviert, Elektroden bestellt, weitere Reparatur am Mittwoch 27.6.2018, Gerät bis auf Crea ok | SWO-372163 | Revisit required | Electrode | Membrane | Crea |  | 25.06.2018, 10:22, fr. Bretting, tel. 030940113280: Oxymetry Hardware Error | Cleaning work, troubleshooting | OxiHardware error is removed after restart, control creaelectrodes and measuring chambers, control solution transport, air bubbles removed from electrodes, calibration solutions renewed, Crea deactivated, electrodes ordered, further repair on Wednesday 27.6.2018, device except for Crea ok
No | Visit | Complaint | RDEU | Hfo meldet bei Geräteinbetriebnahme beim Kunde für Auslieferungen eine Fehler 5768 | Einbau der bestellten Ersatzteile | Tausch der Storage PCB, Punch Ort Motor & Stepper Motor, Startphase abgewartet QC Level 1&2 mit aufgetautem QC Material durchgeführt | SWO-371106 | Defect | HW | Storage Module | PCB, Storage Module | 5768 | Hfo reports an error 5768 to the customer for deliveries when device is put into operation | Installation of the ordered spare parts | Replacement of Storage PCB, Punch Ort Motor & Stepper Motor, start-up phase awaited QC Level 1&2 performed with thawed QC material
No | Visit | Complaint | RDEU | 07.06.2018, 08:59 Uhr, Fr. Jannick, Tel. 0211904361209: Messung wird nach 2 Min. abgebrochen 17.06.2018, 08:40 Uhr, Fr. Wiegand, Tel. 0211904361209: Messung wird abgebrochen, Bildschirm flackert, wenn er bewegt wird, Scanner scannt nicht die komplette Pat. ID. | Gehäuse am Bracode Scanne (Storage Wheel ) defekt. | Gehäuse für Barcode Laser am Storage Wheel erneuert, Barcode Scanner justiert und gereinigt. | SWO-371454 | Defect | HW | Storage Module | Barcode reader, storage wheel |  | 07.06.2018, 08:59, Fr. Jannick, Tel. 0211904361209: Measurement is aborted after 2 min. 17.06.2018, 08:40, Ms. Wiegand, tel. 0211904361209: Measurement is aborted, screen flickers when it is moved, scanner does not scan the complete pat. Id. | Housing on the Bracode Scanne (Storage Wheel) defective. | Housing for barcode laser on the storage wheel renewed, barcode scanner adjusted and cleaned.
No | Visit | Complaint | RDEU | Barcodes werden nur sporadisch im Inlet erkannt 05.02.2018; Fr. Reither; Tel.: 077619386740; Röhrchen für QC und Systemreinigung werden nicht mehr erkannt, AQT verlangt schon wieder eine Systemreinigung. | Sampler Erkennung fehlerhaft | Gerät wurde über Zentrale überprüft und instandgesetzt, Gerät ausgeliefert, einkalibriert und LQC durchgeführt | SWO-371341 | Defect | HW | Aspiration Module | Barcode reader, Inlet |  | Barcodes are only detected sporadically in the inlet 05.02.2018; Ms. Reither; Tel.: 077619386740; Tubes for QC and system cleaning are no longer recognized, AQT requires system cleaning again. | Sampler detection faulty | Device was checked and repaired via control panel, device delivered, calibrated and LQC performed
No | Visit | Complaint | RDEU | Akku Anzeige auf 0% | Akku zeigt dauerhaft auf 0% An | Akku gewechselt Gerät in betrieb genommen und überprüft | SWO-372673 | Defect | HW | Battery module | Battery module | 0% | Battery display to 0% | Battery permanently points to 0% on | Battery changed device put into operation and checked
No | Visit | Complaint | RDAN | tHb fejl | QC tjek | Komplet OXI test udført Kalibrering tHb kalibrering QC | SWO-372146 | Defect | HW | Oxi module | Hemolyzer | sO2, Hbf | tHb fejl | QC Check | Complete OXI test performed Calibration tHb Calibration Qc
No | Visit | Complaint | RAME | pO2(T) issues | PO2(T) not showing up on patient results. Enabled OXI parameters on result in question. None of those are out of reportable ranges. HbF is not enabled. Can not turn on HBF until VK kit is run to validate new parameter. | Could not correctly issue today. Will order a VK kit to come back tomorrow and validate HBF. | SWO-372319 | Revisit required | SW | RMED SW | Oxi parameter |  | pO2(T) issues | PO2(T) not showing up on patient results. Enabled OXI parameters on result in question. None of those are out of reportable ranges. HbF is not enabled. Can not turn on HBF until VK kit is run to validate new parameter. | Could not correctly issue today. Will order a VK kit to come back tomorrow and validate HBF.
Yes | 6 | Y | RPAC | Blank | Leak detected |  | D48B272488060D05 | [5 empty] | Blank | Leak detected
Yes | 5 | Y | RPAC | Blank | Replaced failed SC |  | D48C0EDCFE0ECD29 | [5 empty] | Blank | Replaced failed SC
No | Visit | Complaint | RLTD | Christine Robertson Tel: 0131 242 2601 | New sensor interface PCB required from previous visit | Fitted new sensor interface PCB | SWO-372140 | Defect | HW | Basic Unit | PCB, Sensor interface | pCO2, Lac | Christine Robertson Tel: 0131 242 2601 | New sensor interface PCB required from previous visit | Fitted new sensor interface PCB
Yes | 2 | Y | RCZE | Blank | Riešené po telefóne bez nápravy chyby Analyzátor bol den bez SP. |  | D489BF2E120AC646 | Re classified RMED | phone support |  |  |  | Blank | Solved by phone without correcting error analyzer was a day without SP.
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 病人结果与临床不符 | 更换测试卡后修复 | SWO-374954 | [5 empty] | Test card, reagent package failure | Patient results are not clinically consistent | Repair after replacing test card
No | Visit | Complaint | RAME | SP installed and now has pink fluid on counter | Arrived onsite and checked all tubing. Flushed tubing and reassembled. Found piece of solution pack still on nipple cause the new solution pack not to seat properly. | Removed plastic solution pack piece. | SWO-372444 | [5 empty] | SP installed and now has pink fluid on counter | Arrived onsite and checked all tubing. Flushed tubing and reassembled. Found piece of solution pack still on nipple cause the new solution pack not to seat properly. | Removed plastic solution pack piece.
No | Visit | Complaint | RCHN | 处理盘复位失败 | 处理盘复位失败，借580-098,636-571各一个 | 更换580-098,636-571 后修复by tim | SWO-371189 | Defect | HW | Process Module | Process Wheel |  | Failed to process disk reset | Failed to handle disk reset, borrow 580-098, 636-571 each | Replace 580-098,636-571 after repair by tim
Yes | Visit | Complaint | RIND | Sampling Error | Air in sample,insufficient sample | Demonstrated how to remove air from sample, sample volume should be proper | SWO-374706 | Re classified RMED | User error | User error | User error |  | Sampling Error | Air in sample,insufficient sample | Demonstrated how to remove air from sample, sample volume should be proper
Yes | Visit | Complaint | RDEU | KAL Fehler Na, Glu und Lactat | KAL Fehler Na, Glu und Lactat (Hotline und Medizintechnik) Bei Eintreffen KAL und QC Fehler Glucose Natrium Status und Empfindlichkeitswert grenzwertig $ ständig auftretende Lactat-Driften Referenz und Lactat bereits membraniert | Kontakte gereinigt etc. - Membranierung Glucose und Natrium - mehrfache Kalibrationen (379 Kalibrierung instabil und Nullstromfehler bei Glucose) - Nullstromkalibrierung - Kalibrierung Glucose instabil und Natrium grenzwertig - Erneuerung Referenz Elektrode - Glucose instabil - Erneuerung Glucose Elektrode (keine Fehlerbehebung) - Rücktausch Glucose Elektrode - Messkammer mit Proteinentfernung gereinigt - Reinigung durchgeführt - mehrfache Kalibrationen durchgeführt - Funktionsüberprüfung etc. | SWO-372313 | [5 empty] | KAL Error Na, Glu and Lactate | KAL Error Na, Glu and Lactate (Hotline and Medical Technology) On arrival KAL and QC error glucose Sodium status and sensitivity value borderline - permanent lactate drifts Reference and lactate already membraneed | Contacts purified etc. - membraneing glucose and sodium - multiple calibrations (379 calibration unstable and zero current error in glucose) - zero current calibration - calibration glucose unstable and sodium borderline - renewal reference electrode - glucose unstable - renewal glucose electrode (no troubleshooting) - exchange glucose electrode - measuring chamber cleaned with protein removal - cleaning performed - multiple calibrations performed - function check etc.
No | Visit | Complaint | RAME | Lac being changed about every 2 weeks would like fse to visit. Contact: Mike Reay - 360-493-5181 | Lactate membrane is lasting about two weeks. EL $MET module optical filter was damaged. | Replaced filter and verified operation. Calibrated and ran QC instrument operating normally. | SWO-372440 | Defect | HW | El/Met Module | Kapton Filter |  | Lac being changed about every 2 weeks would like fse to visit. Contact: Mike Reay - 360-493-5181 | Lactate membrane is lasting about two weeks. EL /MET module optical filter was damaged. | Replaced filter and verified operation. Calibrated and ran QC instrument operating normally.
No | Visit | Complaint | RPOL | Brak kalibracji pH | Sprawdzenie aparatu | Wymiana elektrody | SWO-372112 | Assisted Replacement | Electrode | Electrode | pH |  | No pH calibration | Camera Check | Replacing the electrode
Yes | Visit | Complaint | RHUN | Blank Data | 1184 hibakód jelezve, oldat, felszívótű tartó csere nem oldja meg a helyzetet --- pinch valve működése bizonytalan, és inkább rossz | OXI modul alsó tartó burkolat cseréje, pinch valve csere, szerviz tesztek elvégzése, init, kalibráció, QC mérés | SWO-372662 | [5 empty] | Blank Data | 1184 error code indicated, solution, needle holder replacement does not resolve the situation --- Pinch valve operation is uncertain and rather bad | OXY module Lower mount casing replacement, pinch valve replacement, service tests, init, calibration, QC measurement
No | Visit | Complaint | RDEU | 27.06.2018 08:20Uhr Frau Kraus 02088814181: Das Gerät verlangt ständig neue Aktionen | Ableitströme erkannt, Gerinsel erkannt, LP def. | Einl.dichtung ersetzt, SC ersetzt, Schläuche kontrolliert, LP ersetzt, Gerät kalibriert und QC durchgeführt, | SWO-372746 | Assisted Replacement | SP | SP | Leak Current |  | 27.06.2018 08:20pm Mrs Kraus 02088814181: The device constantly demands new actions | Discharge currents detected, Gerinsel detected, LP def. | Single seal replaced, SC replaced, hoses inspected, LP replaced, device calibrated and QC performed,
No | Visit | Complaint | RAME | pO2(T) issues | VK kit material for HbF needs to be run to validate parameter to see if temp correction works. | Went to FedEx to pick up VK kit via overnight shipment and then to hospital. Ran VK material. Will process results later. | SWO-373181 | Revisit required | SW | RMED SW | Oxi parameter |  | pO2(T) issues | VK kit material for HbF needs to be run to validate parameter to see if temp correction works. | Went to FedEx to pick up VK kit via overnight shipment and then to hospital. Ran VK material. Will process results later.
Yes | 9 | Y | RPAC | Blank | pO2 failing QC lvl 3 |  | D48B1630883A59E2 | [5 empty] | Blank | pO2 failing QC lvl 3
Yes | Visit | Complaint | RDEU | QC grenzwertig hoch im Level 1 bei K und im Level 2 bei Na | Data Logs durch Zentrale angefordert | DataLogs Kal, QC, Aktivitäten und Systemüberprüfungen gezogen für Versand nach RDEU TS | SWO-372914 | Re classified RMED | planned SWO | Data collection |  |  | QC borderline high in level 1 at K and in level 2 at Na | Data logs requested by headquarters | DataLogs Kal, QC, activities and system checks drawn for shipping to RDEU TS
No | Visit | Complaint | RAME | leaking from the solution pack compartment | SP leaking, waste manifold contaminated. | Replaced the SP, waste pump and manifold. Ran calibrations and QC. All tests passed and no visible signs of leaking | SWO-372887 | [5 empty] | leaking from the solution pack compartment | SP leaking, waste manifold contaminated. | Replaced the SP, waste pump and manifold. Ran calibrations and QC. All tests passed and no visible signs of leaking
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | NA定标失败，1319；不均匀3溶液，1539 | 更换试剂包后修复 | SWO-372717 | [5 empty] | Test card, reagent package failure | NA calibration failure, 1319; | Repair after replacement of reagent kit
No | Visit | Complaint | RCHN | 测试卡不能被读取，HCT不能被正确测量 | 测试卡不能被读取，HCT不能被正确测量 | 清洁仪器后修复by tim | SWO-372606 | Cleaned | HW | Storage Module | Barcode reader, storage wheel |  | The test card cannot be read and the HCT cannot be measured correctly | The test card cannot be read and the HCT cannot be measured correctly | Repair by tim after cleaning the instrument
No | Visit | Complaint | RCHN | 测试卡存储舱门无法打开 | 测试卡存储舱门无法打开，借903-198一个 | 更换903-198后修复by tim | SWO-371219 | Defect | HW | Analyzer Control Module | PCB, Main Board |  | Test card storage door cannot be opened | Test card storage door can not be opened, borrow 903-198 one | Repair by tim after replacement 903-198
No | Visit | Complaint | RCHN | 报错5689,5784 | 报错5689,5784 | 更换测试卡后修复 | SWO-372715 | Defect | Consumables | Test Kit | Cartridge, Broken | 5689, 5784 | 56895784 | 56895784 | Repair after replacing test card
No | Visit | Complaint | RFRA | Configuration soft | Configuration soft | Configuration. | SWO-373022 | [5 empty] | Configuration soft | Configuration soft | Configuration.
No | Visit | Complaint | RDEU | 25.06.2018, 10:22 Uhr, fr. Bretting, Tel. 030940113280: Oxymetrie Hardware Fehler | Reinigungsarbeiten, Fehlersuche | Crea wieder aktiviert, korrekte Kal abgewartet, QK bei Crea weiterhin zu niedrig in allen Leveln, Crea-Elektroden ersetzt mit MembranLot: 800, korrekte Kal abgewartet, QK ok | SWO-372616 | Assisted Replacement | Electrode | Membrane | Crea |  | 25.06.2018, 10:22, fr. Bretting, tel. 030940113280: Oxymetry Hardware Error | Cleaning work, troubleshooting | Crea reactivated, correct kal waited, QK at Crea continues to be too low in all levels, Crea electrodes replaced with MembranLot: 800, correct calf waited, QK ok
Yes | Visit | Complaint | RIND | Pco2 not calibarting | Reolacemento of new pco2 electrode | Replaced the new pco2 electrode and calibrated the analyser to raedy mode analysed the sampe tested qc | SWO-374748 | [5 empty] | Pco2 not calibarting | Reolacemento of new pco2 electrode | Replaced the new pco2 electrode and calibrated the analyser to raedy mode analysed the sampe tested qc
No | Visit | Complaint | RAME | Leaking f rom the inlet gasket. | Leaking from inlet gasket. Found blood clot in the waste tubing below the inlet gasket. Found tubing nipple on the tubing organizer damaged. | Flushed the blood clot from the waste tubing and replaced the tubing organizer. | SWO-373134 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway | Clot | Leaking f rom the inlet gasket. | Leaking from inlet gasket. Found blood clot in the waste tubing below the inlet gasket. Found tubing nipple on the tubing organizer damaged. | Flushed the blood clot from the waste tubing and replaced the tubing organizer.
No | Visit | Complaint | RSCH | Rep. Ref. | Immer wieder Ca Drift Kunde bereits Ref. und Ca membraniert | Messkammer i.o Connector i.o Ref. Elektrode ersetzt i.o Alte Lot = 8803 => Garantie Kal. und QC i.o System i.o | SWO-372624 | Assisted Replacement | Electrode | Electrode | Reference | CAPA-00733 | Rep. Ref. | Again and again Ca Drift Customer already Ref. and Ca membraned | Measuring chamber i.o Connector i.o Ref. Electrode replaces i.o Old lot = 8803 => Warranty Kal. and QC i.o System i.o
No | Visit | Complaint | RDEU | Na wird zu hoch gemessen | Überprüfung der Elektrodenkanäle | Na Elektrode erneuert. | SWO-372801 | Assisted Replacement | Electrode | Electrode | Na |  | Na is measured too high | Checking the electrode channels | Na electrode renewed.
No | Visit | Complaint | RKK | サンプル測定が中止 | サンプル測定が中止する | メインPCB交換にて改善しました。 | SWO-372565 | Defect | HW | Analyzer Control Module | PCB, Main Board |  | Sample measurement stopped | Sample measurement aborts | Improved by main PCB replacement.
No | Visit | Complaint | RAME | pO2(T) issues | Return on site to resolve PO2(T) issue. | Activate parameter HbF. Enabled FO2Hb, FCOHb, FCOHb and HBF on dynamic parameters. | SWO-373751 | Setup | SW | RMED SW | Oxi parameter |  | pO2(T) issues | Return on site to resolve PO2(T) issue. | Activate parameter HbF. Enabled FO2Hb, FCOHb, FCOHb and HBF on dynamic parameters.
Yes | 4 | Y | RCZE | Blank | Kovarova ma kontaktovala ze jej nejde na tento pristroj dialkovka, na aqt aro a abl90 neo ide, vraj mali vypadok, it vraj konrolovalo vsetko. |  | D48C3FB0FE26140F | Re classified RMED | phone support |  |  |  | Blank | Kovarova contacted me that it does not go to this machine Dialkovka, on Aqt Aro and abl90 Neo IDE, supposedly had a fall, it supposedly was congratalo all.
No | Visit | Complaint | RESP | NO INICIALIZA ERRORES DE OXI | PROBLEMAS AL INICIALIZAR. | SE DETECTA PROBLEMAS AL REINICIO NO MOVIENDOSE LOS REACTIVOS CON FLUIDEZ, ASI MISMO SE DESMONTA SISTEMA DE HEMOLIZACION NOTANDOSE HEMOLIZADOR OBSTRUIDO, SE SUSTITUYE ESTE, SE EFECTUAN AJUSTES Y CALIBRACIONES. CALIBRADO Y VERIFICADO CON Q.C. | SWO-373800 | [5 empty] | DOES NOT INITIALIZE OXI ERRORS | PROBLEMS WHEN INITIALIZING. | PROBLEMS ARE DETECTED BY THE REASSEMBLY NOT MOVING THE REACTIVES WITH FLUIDITY, IT IS DISASSEMBLED HEMOLIZATION SYSTEM NOTTO BE UNBLOCKED HEMOLIZADOR, SUBSTITUTE THIS, ADJUSTMENTS AND CALIBRATIONS ARE EFFECTED. CALIBRATED AND VERIFIED WITH Q.C.
No | Visit | Complaint | RDEU | 27.06.2018, Sr. Jana, Tel: 03646197704: Datum springt immer wieder auf 25.08.2018 | Fehlersuche ==> Datum unter Windows falsch (26.08.2018) | Datum unter Windows eingestellt Neue SC installiert, da Meldung das SC verfallen und diese läßt sich nicht mehr installieren | SWO-373832 | [5 empty] | 27.06.2018, Sr. Jana, Tel: 03646197704: Date jumps again and again to 25.08.2018 | Troubleshooting ==> Date on Windows incorrect (08/26/2018) | Date set on Windows New SC installed because message expires and it can no longer be installed
No | Visit | Complaint | RAME | SP leaking | Follow up needed after recent service. | Perform onsite visit to check status of analyzer. | SWO-373837 | [5 empty] | SP leaking | Follow up needed after recent service. | Perform onsite visit to check status of analyzer.
No | Visit | Complaint | RCHN | HCT值不能被正确测定 | HCT值不能被正确测定 | 调整针位置后修复by tim | SWO-374234 | Adjustment | HW | Aspiration Module | Needle adjustment |  | HCT value cannot be measured correctly | HCT value cannot be measured correctly | Fix by tim after adjusting needle position
No | Visit | Complaint | RIND | Failure of po2 | Po2 failure. | Flush out the cassette, 2point calibration suceesfully run.Po2 is Ready. Run a sample result is satisfactory | SWO-374763 | [5 empty] | Failure of po2 | Po2 failure. | Flush out the cassette, 2point calibration suceesfully run. Po2 is Ready. Run a sample result is satisfactory
Yes | Visit | Complaint | RPOL | Szkolenie personelu z zakresu obsługi analizatora . | Szkolenie personelu | Szkolenie personelu | SWO-373486 | Re classified RMED | Training | Training | Training |  | Training of personnel in the field of analyzer operation. | Staff training | Staff training
Yes | Visit | Complaint | RLTD | Gordon Nickel on 01896 826245 | Analyser failing again after REF electrodes were previously replaced | Data logs and service dump collected and TechLine case to be raised | SWO-373740 | Re classified RMED | planned SWO | Data collection | Data collection |  | Gordon Nickel on 01896 826245 | Analyser failing again after REF electrodes were previously replaced | Data logs and service dump collected and TechLine case to be raised
No | Visit | Complaint | RAME | Missing old datalog files in Archived datalogs | Looked in data logs, found that on June 8, 2018 on the one analyzer and June 9, 2018 on other analyzer calibration log was archived. So I looked and searched for archives for calibration. They were not in analyzers. | Went through with customer the archive setup and how to open archives and how to download datalogs and a WDC. | SWO-373890 | No fault found | No fault found | No fault found | No fault found |  | Missing old datalog files in Archived datalogs | Looked in data logs, found that on June 8, 2018 on the one analyzer and June 9, 2018 on other analyzer calibration log was archived. So I looked and searched for archives for calibration. They were not in analyzers. | Went through with customer the archive setup and how to open archives and how to download datalogs and a WDC.
No | Visit | Complaint | RPOL | Nie działa oxi | Sprawdzenie aparatu | Wymiana oxi | SWO-373778 | Defect | HW | Oxi module | Oxi module |  | Oxi does not work | Camera Check | Exchange Oxi
No | Visit | Complaint | RFRA | Mesure annulée | Contrôle des rotations des différentes roues | Retrait des cartouches qui empêchaient la rotation correcte de la « storage wheel ». Réglage du lecteur code barres des cartouche. Essais | SWO-373764 | Defect | Consumables | Test Kit | Cartridge, Broken |  | Cancelled measure | Controlling the rotations of different wheels | Removing cartridges that prevented the correct rotation of the "storage wheel." Setting the cartridge barcode reader. Tests
No | Visit | Complaint | RDEU | 29.06.2018 8:45 Fr. Salomon 06213831157: Oxi-Harwarefehler | KAL und QC ohne Beanstandung Bei Messung 581 Oxi Spectrum Missmatch 292 Trübung zu hoch | Allgemeine Überprüfung - Analysator gespült - Oximetrie Test ohne Beanstandung - Ventil Überprüfung etc. - Service Tests $ Justagen durchgeführt - Kalibration, tHb KAL und QC durchgeführt - Hemolysator mit Proteinentfernung gespült - Funktionsüberprüfung (jede 2. Messung mit Fehlermeldung) - Erneuerung Hemolysator - Funktionsüberprüfung etc. | SWO-374358 | Defect | HW | Oxi module | Hemolyzer |  | 29.06.2018 8:45 Fr. Salomon 06213831157: Oxi-Harware error | KAL and QC without complaint When measuring 581 Oxi Spectrum Missmatch 292 turbidity too high | General check - Analyzer rinsed - Oximetry test without complaint - Valve review etc. - Service tests performed - Calibration, tHb KAL and QC performed - Hemolysator rinsed with protein removal - Functional check (every 2nd measurement with Error message) - Renewal Hemolysator - Function check, etc.
No | Visit | Complaint | RDEU | 12.06.2018; Hr. Hasanspahic; Tel.: 020917254201; ABL verlangt regelmäßig eine Spülung | Gerinsel im Oximodul erkannt, Verstopfung in den Fließwegen erkannt, immer wieder Wartungsarbeiten, Dichtungen sollen getauscht werden, Fehler tritt sporadisch auf, | LP getausch SC getausvht, Einl.dichtung und Nadel ersetzt, Fließwege gereinigt, alle Teile wieder zurück getauscht, Gerät im Serviceprogramm getestet, alle Werte sind o.k., | SWO-374600 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway | Clot | 12.06.2018; Mr Hasanspahic; Tel.: 020917254201; ABL regularly requires flushing | Gerinsel detected in the oximodul, constipation detected in the flow paths, repeated maintenance work, seals to be replaced, error occurs sporadically, | LP exchange SC getausvht, inserting seal and needle replaced, flow paths cleaned, all parts replaced again, device tested in service program, all values are OK,
No | Visit | Complaint | RLTD | Connie emailed while I was on leave to report the battery was displaying 0% despite being plugged in to the mains | The battery is displaying 0% charge even though its plugged into the mains and there is no battery error present. | Removed the mains power for 15 minutes to see if the charge cycle would start again when the mains was reconnected. I reset the battery module PCB as well and after being on charge for about 15 minutes the analyser started to display the battery charge percentage. I monitored the display and it slowly started to increase so it looks like the problem has been resolved. | SWO-375017 | Adjustment | HW | Battery module | Battery module | Reset | Connie emailed while I was on leave to report the battery was displaying 0% despite being plugged in to the mains | The battery is displaying 0% charge even though its plugged into the mains and there is no battery error present. | Removed the mains power for 15 minutes to see if the charge cycle would start again when the mains was reconnected. I reset the battery module PCB as well and after being on charge for about 15 minutes the analyser started to display the battery charge percentage. I monitored the display and it slowly started to increase so it looks like the problem has been resolved.
No | Visit | Complaint | RAME | PCO2 failing calibration, sensitivity <15%. | Confirmed reported issues, pCO2 calibration failing due to sensitivity <15%. Customer tried 2 sensor cassettes (946-008 R0772N048). Both SCs failed due to the same issue. Performed visual inspection and service diagnostics, ran sensor simulator test, no failures observed. Suspect problematic sensor cassettes. | Replaced sensor cassette with 946-005R1305. | SWO-374729 | Assisted Replacement | SC | SC | pCO2 | Cal | PCO2 failing calibration, sensitivity <15%. | Confirmed reported issues, pCO2 calibration failing due to sensitivity <15%. Customer tried 2 sensor cassettes (946-008 R0772N048). Both SCs failed due to the same issue. Performed visual inspection and service diagnostics, ran sensor simulator test, no failures observed. Suspect problematic sensor cassettes. | Replaced sensor cassette with 946-005R1305.
No | Visit | Complaint | RAME | pCo2 Sensitivity low : <15% | Confirmed reported issues, pCO2 calibration failing due to sensitivity <15%. Customer tried 2 sensor cassettes (946-008 R0773N191). Both SCs failed due to the same issue. Performed visual inspection and service diagnostics, ran sensor simulator test, no failures observed. Suspect problematic sensor cassettes. | Replaced sensor cassette with 946-005R1305. | SWO-374728 | Assisted Replacement | SC | SC | pCO2 | Cal | pCo2 Sensitivity low : <15% | Confirmed reported issues, pCO2 calibration failing due to sensitivity <15%. Customer tried 2 sensor cassettes (946-008 R0773N191). Both SCs failed due to the same issue. Performed visual inspection and service diagnostics, ran sensor simulator test, no failures observed. Suspect problematic sensor cassettes. | Replaced sensor cassette with 946-005R1305.
No | Visit | Complaint | RLTD | The barcode reader is not working. It was agreed with the customer to attend site on Friday 29th June to resolve this. | The barcode reader was not working and turning the analyser off and on failed to resolve this. The barcode reader was removed and all cable connections were secure. | Barcode reader was replaced. | SWO-374405 | Defect | HW | Basic Unit | Barcode Reader |  | The barcode reader is not working. It was agreed with the customer to attend site on Friday 29th June to resolve this. | The barcode reader was not working and turning the analyser off and on failed to resolve this. The barcode reader was removed and all cable connections were secure. | Barcode reader was replaced.
No | Visit | Complaint | RDEU | 27.06.2018, Sr. Jana, Tel: 03646197704: Datum springt immer wieder auf 25.08.2018 | Fehlersuche ==> Keine QC-Messwerte Oximetrie THb-Kalibrierung wird abgebrochen | Fehlerdiagnose durchgeführt ==> kein Fehler erkennbar THb-Kalibrierung wiederholt | SWO-374354 | [5 empty] | 27.06.2018, Sr. Jana, Tel: 03646197704: Date jumps again and again to 25.08.2018 | Troubleshooting ==> No QC readings oximetry THb calibration is canceled | Fault diagnosis performed ==> no error recognizable THb calibration repeated
No | Visit | Complaint | RCHN | 报错1061 | 报错5222，借944-216（14294）一个 | 重新定标后修复 by tim | SWO-373180 | Assisted Replacement | Consumables | Test Kit | Cartridge, CKMB | 5222 | 1061 error | 5222, loan 944-216 (14294) one | Fix by tim after re-calibration
No | Visit | Complaint | RDEU | 28.06.2018 16:45Uhr Sr.Jenny/Pfl. Mario 030940113662: Neonlampenfehler | Fehlersuche, Archiv zeigt immer wieder Spannungsfehler bei der Neonlampe, | Gerät abgeschalten und geöffnet, AU Sicherungen entfernt und Korrossion entfernt, Steckverbindungen ebenfalls überprüft, Neustart, Kal und QC | SWO-375010 | Cleaned | HW | Basic Unit | Analyzer Control Module |  | 28.06.2018 16:45h Sr.Jenny/Pfl. Mario 030940113662: Neon lamp error | Troubleshooting, archive shows again and again voltage errors in the neon lamp, | Device switched off and opened, AU fuses removed and corrossion removed, connectors also checked, restart, Kal and QC
Yes | Visit | Complaint | RIND | Blank | Check the system.check the logs.check the calibration logs. Check the ca parameter sensitivity. Check the membrane & electrode. | Checking the membrane & replace the membrane. Calibrate multiple times. Replace the ref membrane Also but still the status issue in ca parameter. Need to check the ref electode. | SWO-374428 | [5 empty] | Blank | Check the system.check the logs.check the calibration logs. Check the ca parameter sensitivity. Check the membrane & electrode. | Checking the membrane & replace the membrane. Calibrate multiple times. Replace the ref membrane Also but still the status issue in ca parameter. Need to check the ref electode.
No | Visit | Complaint | RAME | Hypochlorite Detected in analyzer. | Unit giving errors for detecting hypochlorite in analyzer. | Replaced all necessary membranes. Unit working properly. | SWO-374469 | User error | Electrode | Membrane | Hypochlorite |  | Hypochlorite Detected in analyzer. | Unit giving errors for detecting hypochlorite in analyzer. | Replaced all necessary membranes. Unit working properly.
No | Visit | Complaint | RPOL | Błąd Oxi na pacjentach. Testy serwisowe ok. | Testy serwisowe | Wymiana modułu Oxi | SWO-374875 | Defect | HW | Oxi module | Oxi module |  | Oxi error on patients. Service tests approx. | Service Tests | Replacing the Oxi module
No | Visit | Complaint | RAME | Customer called with issues pertaining to oximetry hardware errors codes 0251 and 0549 | Halogen Lamp intensity too high, resulting in incorrect oximetry readings during QC. | Adjusted the halogen lamp intensity to an acceptable level within its predetermined range. Replaced Na and Cl membranes for customer. Normal operations restored and verified. | SWO-374518 | Adjustment | HW | Oxi module | Lamp unit | halogen | Customer called with issues pertaining to oximetry hardware errors codes 0251 and 0549 | Halogen Lamp intensity too high, resulting in incorrect oximetry readings during QC. | Adjusted the halogen lamp intensity to an acceptable level within its predetermined range. Replaced Na and Cl membranes for customer. Normal operations restored and verified.
No | Visit | Complaint | RSCH | Rep Fluid | Fehler 1278: Flüssigkeitsmenge oberhalb Warngrenze | Alle Pumpenköpfe ausgetauscht Fitting und LS Sensor (LS Sensor BS) ausgetauscht Reinigung ausgeführt Patch installiert Service Dump gezogen Funktionskontrolle aktuell i.o | SWO-374224 | Defect | HW | Wet section | Wet section | 1278 | Rep Fluid | Error 1278: Amount of liquid above warning limit | Replaced all pump heads Fitting and LS Sensor (LS Sensor BS) replaced Cleaning done Patch installed Service Dump Pulled Functional control currently i.o
No | Visit | Complaint | RCHN | lac定标不过 | LAC定标不过 借903-581一个 | 更换903-581后修复by tim | SWO-373392 | Defect | HW | Basic Unit | PCB, Sensor interface | Lac | lac calibration is not | LAC calibration but borrow 903-581 one | Repair by tim after replacement 903-581
No | Visit | Complaint | RCHN | CO2灵敏度超范围 | CO2灵敏度超范围，借945-612一个 | 更换945-612后修复by tim | SWO-374053 | Assisted Replacement | Electrode | Electrode | pCO2 |  | CO2 sensitivity over range | CO2 sensitivity over range, borrowed 945-612 one | Repair by tim after replacement 945-612
No | Visit | Complaint | RIND | nothing is appearing on display, it is blank. | need to change the CPU | Changed the CPU, now analyzer working satisfactorily. | SWO-375061 | Defect | HW | DMS_PC | CPU |  | nothing is appearing on display, it is blank. | need to change the CPU | Changed the CPU, now analyzer working satisfactorily.
No | Visit | Complaint | RKOR | Vaccume error | clot error | SP SC clean. Gsaket hold, connector, Inlet replace. Liquid flow line clean. (DM 동행) | SWO-379700 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway | Clot | Vaccume error | clot error | SP SC clean. Gsaket hold, connector, Inlet replace. Liquid flow line clean. (DM 동행)
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 病人结果与临床不符 | 更换测试卡后修复 | SWO-374955 | [5 empty] | Test card, reagent package failure | Patient results are not clinically consistent | Repair after replacing test card
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 漏电流 0963 | 更换试剂包后修复 | SWO-377023 | Assisted Replacement | SP | SP | Leak Current | In Analyzer | Test card, reagent package failure | Drain current 0963 | Repair after replacement of reagent kit
Yes | Visit | Complaint | RCHN | 报错4794 | 报错4794 借636-549一个 | 更换636-549后修复by tim | SWO-374746 | [5 empty] | 4794 error | 4794 Loan 636-549 one | Repair by tim after replacing 636-549
No | Visit | Complaint | RCHN | 报错1031 | 报错1031 借903-364一个 |  | SWO-374857 | Defect | HW | Process Module | Optical Unit | 1031 | 1031 | False 1031 Borrow 903-364 one
No | Visit | Complaint | RCHN | 报错5784 | 报错5784 借636-539 903-138 902-924 944-220（14087）944-268（13747）缺货 | 更换636-539 903-138 902-924 重新定标后修复by tim | SWO-374734 | Defect | HW | Incubator Module | PCB, Incubator |  | 5784 | 5784 Borrow636-539 903-138 902-924 944-220 (14087) 944-268 (13747) Out of Stock | Replace 636-539 903-138 902-924 Repair by tim after recalibration
No | Visit | Complaint | RDEU | Lösungen können nicht gepumpt werden | Lösungen können nicht gepumpt werden | Serviceprogramme durchgeführt, keine Auffälligkeiten, aufgef. Verschleißteile ersetzt, Kal, QC | SWO-376611 | Defect | HW | Basic Unit | PM Kit | Unable to pump | Solutions cannot be pumped | Solutions cannot be pumped | service programs, no abnormalities, unreported. Wear parts replaced, Kal, QC
No | Visit | Complaint | RAME | Error 0378 lactate calibration sensitivity out of range. | Customer replaced four sensor cassettes and the solution pack but did not help. | Replaced the sensor interface PCB. | SWO-377054 | Defect | HW | Basic Unit | PCB, Sensor interface | Lac | Error 0378 lactate calibration sensitivity out of range. | Customer replaced four sensor cassettes and the solution pack but did not help. | Replaced the sensor interface PCB.
No | Visit | Complaint | RDEU | Kein Fernzugriff | Kein Fernzugriff Netop Host startet nicht | Netop Host manuell gestartet (startet nicht) - Applikationssoftware geladen (keine Fehlerbehebung) - Betriebssystem und Software geladen - Neuinstallation Betriebsmittel - Touchscreen Kalibration - gespeicherte Daten wieder hergestellt - Funktionsüberprüfung etc. Ersatzanalysator aus Labor ausgestellt - Momentan keine Anbindung - Komplette Software musste nochmals installiert werden - SC und SP nochmals installiert ! | SWO-376794 | Revisit required | SW | RMED SW | Remote |  | No remote access | No remote access Netop Host does not start | Netop Host manually started (does not start) - application software loaded (no troubleshooting) - operating system and software loaded - reinstallation equipment - touchscreen calibration - stored data restored - function verification etc. Replacement analyzer issued from laboratory - Currently no connection - Complete software had to be reinstalled - SC and SP reinstalled !
Yes | 7 | Y | RPAC | Blank | Pressure vacuum test error |  | D491300DD65D88AA | [5 empty] | Blank | Pressure vacuum test error
Yes | 3 | Y | RPAC | Blank | Calibration failure on Lactate |  | D49159A8D65563FA | [5 empty] | Blank | Calibration failure on Lactate
Yes | 3 | Y | RPAC | Blank | SW Upgrade to V3.4 |  | D491584ED62EB7DE | Re classified RMED | planned SWO |  |  |  | Blank | SW Upgrade to V3.4
No | Visit | Complaint | RIND | Service Visit. | SC calibration failed. | Reinitiate the SC, not successful, check the sensor board, scale and sludge had formed on connectors, do the zeroing, not successful, need to change the sensor board. | SWO-377592 | [5 empty] | Service Visit. | SC calibration failed. | Reinitiate the SC, not successful, check the sensor board, scale and sludge had formed on connectors, do the zeroing, not successful, need to change the sensor board.
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | PH,CO2,O2,NA,K,CA,CL定标不过 | 更换测试卡后修复 | SWO-377017 | [5 empty] | Test card, reagent package failure | PH,CO2,O2,NA,K,CA,CL定标不过 | Repair after replacing test card
No | Visit | Complaint | RFRA | Vérification suite pb de surconsommation de solution | Vérification de l'analyseur | Vérification complète $ $ essais en mode service $ $ vérification des zéros $ $ contrôle des configuration $ $ remise en route $ $ essais bons | SWO-376730 | [5 empty] | Check after pb of solution overconsumption | Analyzer check | Full verification $$ trials in service mode $$ check of zero $$ configuration control $$reroute $$ trials vouchers
No | Visit | Complaint | RFRA | Pb de cal pO2 et Glu | Contrôle des canaux de électrodes | Remise à zéro des canaux électrodes. Essais | SWO-376669 | [5 empty] | Pb of cal pO2 and Glu | Control of electrode channels | Reset the electrode channels. Tests
No | Visit | Complaint | RFRA | Désactivation du firewall et test connexion | Désactivation du firewall et test connexion | Désactivation du firewall et test connexion | SWO-376670 | [5 empty] | Firewall deactivation and login test | Firewall deactivation and login test | Firewall deactivation and login test
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | CO2定标不过 0378 | 更换测试卡后修复 | SWO-375131 | [5 empty] | Test card, reagent package failure | CO2 calibration is no more than 0378 | Repair after replacing test card
No | 3 | Y | RPAC | Blank | FAN 915-380 |  | D49161EAD60E83FA | [5 empty] | Blank | FAN 915-380
No | 1 | Y | RPAC | Blank | FAN 915-380 |  | D49140A2D60F83F1 | [5 empty] | Blank | FAN 915-380
No | Visit | Complaint | RCHN | 读取条码失败 | 读取条码失败，借910-304,905-960各一个 |  | SWO-374979 | Defect | HW | Storage Module | Barcode reader, storage wheel |  | Failed to read barcode | Failed to read barcodes, borrowing 910-304,905-960 each
Yes | Visit | Complaint | RCHN | 报错3415 | 报错3415，需要借件903-329一个 |  | SWO-374999 | [5 empty] | Error 3415 | 3415, need to borrow 903-329 one
No | Visit | Complaint | RCHN | 报错5784 | 报错5784，借903-030一个 |  | SWO-375007 | Defect | HW | Storage Module | PCB, Cartridge lock |  | 5784 | 5784, loan 903-030 one
No | Visit | Complaint | RCHN | 报错1031 | 报错1031 借903-364一个 |  | SWO-374941 | Defect | HW | Process Module | Optical Unit | 1031 | 1031 | False 1031 Borrow 903-364 one
No | Visit | Complaint | RSCH | Rep. FlexQ | Erneut Fehler 0882 Mischerträger bereits 2x ersetzt Fehler kommt mind. 1x am Tag | Komplettes FlexQ durch Leihmodul ersetzt i.o FlexQ kalibriert i.o Testmessung i.o | SWO-378052 | Defect | HW | FlexQ | FlexQ |  | Rep. FlexQ | Error again 0882 Mixer carrier already replaced 2x Error comes at least 1x a day | Complete FlexQ replaced by rental module i.o FlexQ calibrates i.o Test measurement i.o
No | Visit | Complaint | RKK | 測定にてｐH応答異常、多発 | 測定にてｐH応答異常、多発 | pH電極交換。 ｷｬﾙ、QCﾁｪｯｸ、OK。 | SWO-377025 | Assisted Replacement | Electrode | Electrode | pH |  | PH response abnormality by measurement, multiple | PH response abnormality by measurement, multiple | pH electrode replacement. Cal, QC check, OK.
No | Visit | Complaint | RDEU | 02.07.2018 Pfl. Kai 06652987150 meldet, dass seit dem Wochenende der ABL einen pCO2 QC Fehler zeigt. | Po2-Fehler in Level C | Anwender hatte bereits mehrfach Lösungspack und Kassette getauscht. Lösungspack gegen vorhandenen des Notfallgerätes getauscht. Fehler war danach mitgewandert. Lösungspack reklamiert. Gerät überprüft. | SWO-377614 | Assisted Replacement | SP | SP | pCO2 | Level C | 02.07.2018 Pfl. Kai 06652987150 reports that since the weekend the ABL shows a pCO2 QC error. | Po2 error in level C | User had already exchanged solution packs and cassettes several times. solution pack for existing emergency equipment. Mistakes were then immigrated. solution pack. device checked.
No | Visit | Complaint | RAME | Thb calibration is drifting | Thb drifting and inhomogeneous cal 3 | Customer had replaced the solution pack and all was working ok $ performed total Oxi test ok $ performed Thb cal | SWO-378506 | Assisted Replacement | SP | SP | Cal 3 | tHb drift | Thb calibration is drifting | Thb drifting and inhomogeneous cal 3 | Customer had replaced the solution pack and all was working ok $ performed total Oxi test ok $ performed Thb cal
No | Visit | Complaint | RDEU | Kein Fernzugriff | Kein Fernzugriff | Analysator messbereit - Touchscreen Kalibration - Netzwerkanbindung - Funktionsüberprüfung etc. Backup Analysator zurück Labor | SWO-377183 | Setup | SW | RMED SW | Remote |  | No remote access | No remote access | Analyzer ready for measurement - Touchscreen calibration - Network connection - Function check, etc. Backup Analyzer Back Lab
No | Visit | Complaint | RAME | Intermittent error 1324 inhomogeneous rinse solution. | Checked analyzer fluid flow. Found bent inlet probe. Checked liquid sensors, pump and pinch valve. All checked okay. | Replaced bent inlet probe and flushed analyzer. | SWO-377055 | Defect | HW | Inlet module | Probe bend |  | Intermittent error 1324 inhomogeneous rinse solution. | Checked analyzer fluid flow. Found bent inlet probe. Checked liquid sensors, pump and pinch valve. All checked okay. | Replaced bent inlet probe and flushed analyzer.
No | Visit | Complaint | RFRA | Err 292-581 | Contrôle module oxymétrie (spectrophotomètre et hémolyseur) | Contrôle menu service $Test oxy complète,constate erreur cavité hémolyseur $Remplacement de l'hemolyseur $Test oxy complète ok $Tests fluid transport ok $Calibration de la tHB $Calibration et passage des CQ $Passage échantillons ok. | SWO-377378 | Defect | HW | Oxi module | Hemolyzer |  | Err 292-581 | Oxymetry module control (spectrophotometer and hemolyser) | Control menu service $Test complete oxy, finds error cavity hemolyseur $Remplacement hemolyseur $Test oxy complete ok $Tests fluid transport ok $Calibration of the tHB $Calibration and passage of cq $Passage samples ok.
No | Visit | Complaint | RLTD | Please call Clare Latham on 01553 613 780 or 07890 715 176 | Failing to print | Saved setup and backup, rebooted analyser, now ok. Attempted to get analyser connected to Live Connect | SWO-378048 | Adjustment | SW | RMED SW | RMED SW | print | Please call Clare Latham on 01553 613 780 or 07890 715 176 | Failing to print | Saved setup and backup, rebooted analyser, now ok. Attempted to get analyser connected to Live Connect
No | Visit | Complaint | RIND | Service Visit. | After providing stand by sensor board, pO2 value is not coming in the sample measurement. | Check the consumable in other outright analyzer, same problem occurs, check the frc cable connection, replace the signal data board and do the zeroing of the analyzer, found ok, again reinitiate the consumable in the analyzer but same problem occurs. Provide stand by analyzer and install the consumable, sample analysis do by the staff, successfully done, need to recheck the analyzer with other consumable. | SWO-379368 | [5 empty] | Service Visit. | After providing stand by sensor board, pO2 value is not coming in the sample measurement. | Check the consumable in other outright analyzer, same problem occurs, check the frc cable connection, replace the signal data board and do the zeroing of the analyzer, found ok, again reinitiate the consumable in the analyzer but same problem occurs. Provide stand by analyzer and install the consumable, sample analysis do by the staff, successfully done, need to recheck the analyzer with other consumable.
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 试剂包漏电流。报错 0964 | 更换试剂包后修复 | SWO-377027 | [5 empty] | Test card, reagent package failure | The reagent package leakage current. Error 0964 | Repair after replacement of reagent kit
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 试剂包漏电流 | 更换试剂包后修复 | SWO-377037 | [5 empty] | Test card, reagent package failure | Reagent packet leakage current | Repair after replacement of reagent kit
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | NA,PH,CO2,CL定标不过 | 更换测试卡后修复 | SWO-377743 | Assisted Replacement | SC | SC | SC | Na, pH, pCO2, Cl | Test card, reagent package failure | NA, PH, CO2, CL calibration but | Repair after replacing test card
No | Visit | Complaint | RCHN | 报错1024 | 上机后仍报错1024，加借903-364一个 | 更换903-364后修复(By Ran) | SWO-374057 | Defect | HW | Process Module | Optical Unit | 1024 | 1024 error | On board still reported error 1024, plus 903-364 one | Repair after replacement 903-364 (By Ran)
No | Visit | Complaint | RCHN | 干燥前温度超出范围，报错1068 | 干燥前温度超出范围，报错1068,需要借件903-210 | 更换903-210备件后修复(By Ran) | SWO-375088 | Defect | HW | Process Module | PCB, Cup dryer | 1068 | Temperature out of range before drying, error 1068 | Temperature out of range before drying, error 1068, need to borrow 903-210 | Repair after replacement of 903-210 spare parts (By Ran)
No | Visit | Complaint | RKK | SCドアセンサー認識不良 | センサー汚れによる開閉不良。 | 客先感情も考慮しセンサー交換実施。 改善。 | SWO-377058 | [5 empty] | SC door sensor recognition failure | Poor opening and closing due to sensor stains. | Sensor exchange was implemented taking into account customer sentiment. Improve.
No | Visit | Complaint | RAME | Technician reporting the time is shift back by 1 hour every week (Sunday). | Reviewed time setting within windows. Observed DST was not activated although it had been activated initially several months ago. | Reactivated DST. Verified correct time zone, confirmed timeserver is set to sync with UCLA timeserver. Rebooted analyzer | SWO-378295 | Setup | SW | Windows | Timesetting |  | Technician reporting the time is shift back by 1 hour every week (Sunday). | Reviewed time setting within windows. Observed DST was not activated although it had been activated initially several months ago. | Reactivated DST. Verified correct time zone, confirmed timeserver is set to sync with UCLA timeserver. Rebooted analyzer
No | Visit | Complaint | RAME | autocheck mechanical error | Completed inspection of data logs and found communication errors and rebooting occurred last 2 days. Autocheck errors normal operating procedures during restarting. | None necessary | SWO-377742 | Setup | SW | Autocheck Module | Reboot |  | autocheck mechanical error | Completed inspection of data logs and found communication errors and rebooting occurred last 2 days. Autocheck errors normal operating procedures during restarting. | None necessary
No | Visit | Complaint | RSAF | Storage door closes half way and goes on error state, analyzer requires restart | Door gets stuck halfway and skew - spring almost loose and stuck on the side. | Checked unit and cleaned guide rails for storage door and tightened pulling spring of storage door - tested ok. | SWO-377596 | Adjustment | HW | Cover and Chassis | Cartridge door |  | Storage door closes half way and goes on error state, analyzer requires restart | Door gets stuck halfway and skew - spring almost loose and stuck on the side. | Checked unit and cleaned guide rails for storage door and tightened pulling spring of storage door - tested ok.
No | Visit | Complaint | RDEU | 02.07.2018 Frau Rheinfelder 095491881 meldet Einen Fehler undefined_lookup_string 1999 | Halterung Cartridge herausgesprungen->neu eingesetzt | Fehlersuche, Safe Positions Probleme bei Rotate Storage Wheel (SetSafe Error), Abdeckung Storage Wheel abgebaut, Halterung Cartridge hat sich gelöst ist aber nicht defekt $gebrochen, Halterung angebracht, Rückbau, Neustart und Startphase ok Im Anschluss HJW | SWO-377323 | Adjustment | HW | Storage Module | Storage wheel | 1999 | 02.07.2018 Ms Rheinfelder 095491881 reports an error undefined_lookup_string 1999 | Bracket Cartridge jumped out->re-inserted | Troubleshooting, Safe Positions Problems with Rotate Storage Wheel (SetSafe Error), Cover Storage Wheel Dismantled, Bracket Cartridge has been solved but is not defective $gebrochen, bracket attached, dismantling, restart and start-up phase ok Following HJW
No | Visit | Complaint | RKK | 溶液パック交換後、パックの読み取り不良にて、作動せず。再度交換画面になる。 | 溶液交換対応 | 溶液パック交換 | SWO-378135 | Assisted Replacement | SP | SP | SP |  | After the solution pack replacement, the pack is not read poorly and does not operate. It becomes the replacement screen again. | Solution exchange support | Solution pack replacement
No | Visit | Complaint | RDEU | 03.07.2018 11:30Uhr Herr Dr.Reich 015254785507: Das Gerät meldet Pumpen-Kal-Fehler | Inhomogene spüllösung, Kal. Fehler Flow selektor, LS Justage Fehler | Einlass überprüft, manuell gespült, Quetschventilschlauch gewechselt, interne Tests durchgeführt. Mehrfach Pumpenkalibrierung durchgeführt entweder grenzwertig hoch oder niedrig. pumpenkranz überprüft. LP gewechselt, installiert und Probe Test durchgeführt. | SWO-379746 | Defect | HW | Oxi module | Pinch valve tube |  | 07/03/2018 11:30AM Mr. Reich 015254785507: The device reports pump-kal errors | Inhomogeneous rinsing solution, Kal. Error Flow Selector, LS Adjustment Error | Inlet checked, manually rinsed, squeezing valve hose changed, internal tests carried out. Multiple pump calibration performed either borderline high or low. pump wreath. LP changed, installed and sample test performed.
No | Visit | Complaint | RIND | Instrumenent has only white screen on display. | Opended instrument and checked power supply ,CF card and CPU board. | Power supply is working ok but need to check with CF card or CPU unit. | SWO-379500 | [5 empty] | Instrumenent has only white screen on display. | Opended instrument and checked power supply ,CF card and CPU board. | Power supply is working ok but need to check with CF card or CPU unit.
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | CO2定标失败 | 连续更换2张测试卡后修复 | SWO-378890 | [5 empty] | Test card, reagent package failure | CO2 calibration failure | Repair after 2 consecutive test cards
No | Visit | Complaint | RCHN | 报错1035 | 报错1035 借903-364一个 | 更换903-364后修复(By Ran) | SWO-377774 | Defect | HW | Process Module | Optical Unit | 1035 | 1035 error | False 1035 Borrow 903-364 one | Repair after replacement 903-364 (By Ran)
No | Visit | Complaint | RCHN | 扫不到条形码 | 扫不到条形码，借902-924一个 | 更换902-924后修复(by chuanwei) | SWO-377234 | Defect | HW | Analyzer Control Module | Barcode reader, display |  | Can't sweep barcodes | Can't sweep the barcode, borrow 902-924 a | Repair after replacement 902-924 (by chuanwei)
No | Visit | Complaint | RKK | Gluキャリブレーション時の電極が不安定（応答不良）頻発。Glu電極交換後復帰。検体測定、QC確認。 | Gluキャリブレーション時の電極が不安定（応答不良）頻発。 | Glu電極交換後復帰。検体測定、QC確認。 | SWO-378870 | Assisted Replacement | Electrode | Electrode | Glu |  | Electrodes during Glu calibration are frequently unstable (unresponsive). Returned after glu electrode replacement. Sample measurement, QC confirmation. | Electrodes during Glu calibration are frequently unstable (unresponsive). | Returned after glu electrode replacement. Sample measurement, QC confirmation.
No | Visit | Complaint | RAME | LIS connectivity issue | Assistance with LIS connectivity. | Provided assistance, will return to verify analyzer transmits data. | SWO-378821 | Setup | His/Lis | His/Lis | His/Lis |  | LIS connectivity issue | Assistance with LIS connectivity. | Provided assistance, will return to verify analyzer transmits data.
Yes | 5 | Y | RPAC | Blank | Software Upgrade analyser to Version V3.4, FAN915-384. |  | D49503B8380922A2 | [5 empty] | Blank | Software Upgrade analyser to Version V3.4, FAN915-384.
Yes | 4 | Y | RPAC | Blank | Replaced failed SC |  | D494F8AC381803E3 | [5 empty] | Blank | Replaced failed SC
Yes | 4 | Y | RPAC | Blank | Software upgrade to v3.4 FAN915-384 |  | D495032B3860E3CF | Re classified RMED | planned SWO |  |  |  | Blank | Software upgrade to v3.4 FAN915-384
No | Visit | Complaint | RAME | Sample not aspirating | Failed to aspirated samples error occurred few times in the past week. | Replaced Auto inlet PCB. Ran All service checks. Calibrations and QC accepted. | SWO-378914 | Defect | HW | Auto Inlet module | PCB, Auto inlet |  | Sample not aspirating | Failed to aspirated samples error occurred few times in the past week. | Replaced Auto inlet PCB. Ran All service checks. Calibrations and QC accepted.
No | Visit | Complaint | RIND | Service Visit. | pO2 parameter deactivate again and again with the auto calibration taken by the analyzer, Solution pack installed by the user on 27 june, till now the solution pack having 24 cycles remaining. | Do the zeroing of the analyzer, in between the procedure, after set zero the value of pO2 increasing and not stable, check the sensor board of the analyzer but physically found ok. Try with the proper working sensor board, connect with the analyzer, auto calibration taken by the analyzer which is successfull, need to observe the analyzer for 15 days. | SWO-379358 | [5 empty] | Service Visit. | pO2 parameter deactivate again and again with the auto calibration taken by the analyzer, Solution pack installed by the user on 27 june, till now the solution pack having 24 cycles remaining. | Do the zeroing of the analyzer, in between the procedure, after set zero the value of pO2 increasing and not stable, check the sensor board of the analyzer but physically found ok. Try with the proper working sensor board, connect with the analyzer, auto calibration taken by the analyzer which is successfull, need to observe the analyzer for 15 days.
No | Visit | Complaint | RCHN | 无法连接试剂包 | 无法连接试剂包，需要借件903-363 | 备件未使用，垫了A4纸，机器恢复使用(By Ran) | SWO-377028 | [5 empty] | Unable to connect the reagent package | Cannot connect the reagent package, 903-363 is required | Spare parts not used, A4 paper padd edgy, machine resumed (By Ran)
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 试剂包未连接 | 更换试剂包后修复 | SWO-378354 | [5 empty] | Test card, reagent package failure | Reagent package not connected | Repair after replacement of reagent kit
Yes | Visit | Complaint | RIND | Solution pack is not installing. | 2 point calibration failed while installing solution pack | Reset solution pack and reinstalled but V $C check was not performed. try second available SP but this one also failed.need to replaced with new one. | SWO-379350 | [5 empty] | Solution pack is not installing. | 2 point calibration failed while installing solution pack | Reset solution pack and reinstalled but V $C check was not performed. try second available SP but this one also failed.need to replaced with new one.
No | Visit | Complaint | RCHN | 报错2168-442 | 报错2168-442，借636-539一个 | 更换了636-539后修复(By Ran) | SWO-378461 | Defect | HW | Incubator Module | sensor, cup | 2168 | 2168-442 | 2168-442, loan 636-539 one | Repair after replacement of 636-539 (By Ran)
Yes | Visit | Complaint | RIND | El/Met pump error | Checked the data logs El $Met air value was getting failed found block in air value | Cleaned El $Met air membrane value initialization & 2 point calibration done no error found machine is working fine | SWO-379110 | Cleaned | Fluidic pathway | El/Met Module | Valve, air |  | El/Met pump error | Checked the data logs El /Met air value was getting failed found block in air value | Cleaned El $Met air membrane value initialization & 2 point calibration done no error found machine is working fine
No | Visit | Complaint | RAME | Qc Ampules not reading | Qc Ampules not reading. Found broken ribbon cable | Replaced ribbon cable | SWO-379493 | Defect | HW | Autocheck Module | FPC Cable |  | Qc Ampules not reading | Qc Ampules not reading. Found broken ribbon cable | Replaced ribbon cable
No | Visit | Complaint | RDEU | 28.06.2018 16:45Uhr Sr.Jenny/Pfl. Mario 030940113662: Neonlampenfehler | Reinigungsarbeiten, Fehlersuche | OxiControllerBoard ersetzt, OxiCompleteTest ok, korrekte Kal abgewartet, QK ok | SWO-378421 | Adjustment | HW | Oxi module | Lamp unit | neon | 28.06.2018 16:45h Sr.Jenny/Pfl. Mario 030940113662: Neon lamp error | Cleaning work, troubleshooting | OxiControllerBoard replaced, OxiCompleteTest ok, correct scale waited, QK ok
No | Visit | Complaint | RAME | Ca Out of range high on QC. The customer indicated that it is an on-going issue. | Observed Membrane change on 7 $2, all indicators passed but QC L2 & L3 bean to fail just above range reported on 7 $5. | Removed and inspected CA electrode, tapped to clear membrane, clean connects in electrode connector cap. Reinitiated calibration cycles. After several 2 PT calibration cycles CA passed cal and QC with no further issues. | SWO-379858 | Cleaned | HW | El/Met Module | Electrode contact |  | Ca Out of range high on QC. The customer indicated that it is an on-going issue. | Observed Membrane change on 7 /2, all indicators passed but QC L2 & L3 bean to fail just above range reported on 7 /5. | Removed and inspected CA electrode, tapped to clear membrane, clean connects in electrode connector cap. Reinitiated calibration cycles. After several 2 PT calibration cycles CA passed cal and QC with no further issues.
No | Visit | Complaint | RAME | Chloride has been going out of range high on QC all levels. After remembraning it will be good for a while and then goes out high again. The Cl membrane last around 1.5 months to 2 months. | Checked data logs. Found errors. Unable to resolve issue. Need more supporting data. | Will get a service dump and check into WDC for peer group ranges and results. I have brought in Keith Clark to help with looking at the data. Will report back to Kathy when we have a better idea of what could be the reason for high Chloride results. After running the service dump program the analyzer gave a status code 2 error. The database had a corrupt file and in the process of copying the database, it gave the error. Reloaded software. Upgraded to software 6.18. | SWO-378837 | Revisit required | Electrode | Membrane | Cl | QC issues | Chloride has been going out of range high on QC all levels. After remembraning it will be good for a while and then goes out high again. The Cl membrane last around 1.5 months to 2 months. | Checked data logs. Found errors. Unable to resolve issue. Need more supporting data. | Will get a service dump and check into WDC for peer group ranges and results. I have brought in Keith Clark to help with looking at the data. Will report back to Kathy when we have a better idea of what could be the reason for high Chloride results. After running the service dump program the analyzer gave a status code 2 error. The database had a corrupt file and in the process of copying the database, it gave the error. Reloaded software. Upgraded to software 6.18.
No | Visit | Complaint | RDEU | 02.07.2018 : Frau Hannisky 072314982779Das Gerät meldet keine Lösung erkannt 03.07.2018 08:00Uhr Frau Rauleder 072314982777: Das Gerät verlangt wieder einen neuen LP 03.07.2018, 08:41 Uhr, H. Gauss, Tel. 072314989621: LP Wechsel wird verlangt 06.07.2018; 13:10 Uhr; Hr. Gauss; Tel.: 072314989625; ständig fallen andere Parameter aus | Lösungspack wird nicht akzeptiert. Diverse Parameter fallen immer raus. Auftrag durch Herrn Gauß | Fehleranalyse durchgeführt. Kontakte der Solutionspack PCB sind leicht eingedrückt, -erneuert. Eingedrückte Kontakte kommen meist durch nich richtig heruntergedrückten Lösungspack. Oxi Vale Test durchgeführt, - keine Auffälligkeiten. ABL90 gespült und Dichtigkeitstest durchgeführt, -alles o.k. Vorsorglich Teile im Umfang eine Jahreswartung erneuert. Einlaufverfahren durchgeführt. Anstandslose Einlaufphase bestanden, -alles wieder grün tHB Kal. fällig seit Juni 2017, tHb Kal. durchgeführt, -o.k. Noch zwei QC ausgeführt, Gerät bleibt stabil. - o.k. Anwender kam mit Blutproben: Kapillare wurden nicht gemischt und es wurde kein Gerinselfänger verwendet. -Spontane Fehlmessung durch unsachgemäßes Ansetzen der Kapillare. Weiterhin werden Spritzen von B+D verwendet und ungemischt mit viel Luft im Sampler dem ABL eingegeben. Diensthabende Hebammen nochmal auf die Präanalytik aufmerksam gemacht. Es müssen auf jeden Fall die Gerinselfänger verwendet werden und das Blut in den Samplern richtig gemischt werden. Auf die Vorteile der sPico Spritze hingewiesen, da das Gerät sehr wahrscheinlich einige Störungen durch geronnenes Blut im ABL Messsystem hat. | SWO-381137 | Defect | HW | Basic Unit | PCB, Solution Pack PCB Id |  | 02.07.2018 : Ms. Hannisky 072314982779The device does not report a solution detected 03.07.2018 08:00h Mrs Rauleder 072314982777: The device requires a new LP again 03.07.2018, 08:41, H. Gauss, Tel. 072314989621: LP Change is required 06.07.2018; 1:10 p.m.; Mr. Gauss; Tel.: 072314989625; constantly other parameters fail | Solution pack is not accepted. Various parameters always fall out. Commissioned by Mr. Gauss | error analysis. Contacts of the Solutionspack PCB are easily pressed in, renewed. Pressed contacts usually come through not really pressed down solution pack. Oxi Vale test performed - no abnormalities. ABL90 rinsed and leak test performed, -all ok. As a precautionary measure, parts in the scope of an annual maintenance renewed. run-in procedures. Decent run-in phase passed, -everything green again tHB Kal. due since June 2017, tHb Kal. performed, -o.k. Two more QC executed, device remains stable. - ok. Users came with blood samples: Capillaries were not mixed and no gerinsel catcher was used. -Spontaneous mismeasurement due to improper setting of the capillary. Furthermore, syringes from B+D are used and entered unmixed with a lot of air in the sampler of the ABL. Serving midwives again drew attention to the pre-analysis. In any case, the Gerinsel catchers must be used and the blood in the samplers must be mixed correctly. The advantages of the sPico syringe are noted as the device most likely has some interference from the tanned blood in the ABL measuring system.
No | Visit | Complaint | RAME | WIRELESS IS NOT WORKING, WII NOT CONNECT TO AQURE | 7 $5 $18 WIRELESS IS NOT WORKING, WII NOT CONNECT TO AQURE. I SWAPPED THE DONGLE WITH ANOTHER 90 AND IT IS NOT CONNECTING AND THE OTHER 90 IS NOW NOW CONNECTING. WILL ORDER WIRELESS ADAPTOR AND SENT TO PAULA. WILL CHECK WITH HER TOMORROW. 7 $6 $18 PAUL SENT ME A TEXT AND STATED THAT SHE RECEIVED THE WIFI ADAPTER, PUT IT ON AND THE ANALYZER IS NOW CONNECTING. | PAULA INSTALLED THE NEW WIFI ADAPTER | SWO-378752 | Defect | SW | SW | Wifi | Adapter | WIRELESS IS NOT WORKING, WII NOT CONNECT TO AQURE | 7 /5 /18 WIRELESS IS NOT WORKING, WII NOT CONNECT TO AQURE. I SWAPPED THE DONGLE WITH ANOTHER 90 AND IT IS NOT CONNECTING AND THE OTHER 90 IS NOW NOW CONNECTING. WILL ORDER WIRELESS ADAPTOR AND SENT TO PAULA. WILL CHECK WITH HER TOMORROW. 7 /6 /18 PAUL SENT ME A TEXT AND STATED THAT SHE RECEIVED THE WIFI ADAPTER, PUT IT ON AND THE ANALYZER IS NOW CONNECTING. | PAULA INSTALLED THE NEW WIFI ADAPTER
Yes | 4 | Y | RPAC | Blank | Consumable failure - unable to read chip |  | D499F76510618AA6 | [5 empty] | Blank | Consumable failure - unable to read chip
No | Visit | Complaint | RFRA | Configuration lecteur Z pour le sauvegarde automatique. |  |  | SWO-379183 | Setup | SW | SW | SW | Backup | Z drive setup for automatic backup.
No | Visit | Complaint | RLTD | Mansour Tel 0151 67851111 07917 | Site visit to replace battery with new battery module. | Replaced battery with new, rebooted analyser and waited for restart to complete. On e restarted I tested that analyser battery error had cleared. Battery charged at good steady rate. Disconnected the ups from circuit and ran analyser on battery only. Battery depleted at good rate. Applied power directly from the mains supply and analyser engaged mains supply and began recharging battery again. Ran calibration and qc which all passed. Analyser back running correctly again. | SWO-379334 | Defect | HW | Battery module | Battery module |  | Mansour Tel 0151 678511111 07917 | Site visit to replace battery with new battery module. | Replaced battery with new, rebooted analyser and waited for restart to complete. On e restarted I tested that analyser battery error had cleared. Battery charged at good steady rate. Disconnected the ups from circuit and ran analyser on battery only. Battery depleted at good rate. Applied power directly from the mains supply and analyser engaged mains supply and began recharging battery again. Ran calibration and qc which all passed. Analyser back running correctly again.
Yes | Visit | Complaint | RCRO | CCN/PER | Error 1305 Temperature malfunction | All parameters | SWO-387425 | [5 empty] | CCN/PER | Error 1305 Temperature malfunction | All parameters
Yes | Visit | Complaint | RCHN | 温度错误 | 温度错误，借903-183一个 | 更换903-183后修复(By Ran) | SWO-377823 | [5 empty] | Temperature error | Temperature error, borrow 903-183 one | Repair after replacement 903-183 (By Ran)
No | Visit | Complaint | RCHN | CKMB定标失败 | CKMB定标失败,需要借件944-216（14567） | 重新定标正常。Recal, then OK(By Ran) | SWO-378996 | Assisted Replacement | Consumables | Test Kit | Cartridge, CKMB |  | CKMB calibration failed | CKMB bid failed, need to borrow 944-216 (14567) | Recalibration is normal. Recal, then OK (By Ran)
Yes | Visit | Complaint | RCHN | 报错1068 | 报错1068 借944-212（14280）944-216（14259）各一张 | 重新定标后修复(By Ran) Recal ,then OK | SWO-377906 | [5 empty] | 1068 error | Error 1068 Borrowing 944-212 (14280)944-216 (14259) each | Re-calibration repair (By Ran) Recal, then OK
Yes | Visit | Complaint | RIND | Lac calibration failure even after recent membrane replacement on ABL800 analyzer at Ruby Hall Clinic, Pune | Need to check the procedure which was followed while changing the membrane, also need to check the membrane for any leakage. | After repeating the calibration 2 times, the sensitivity of Lac was found to be decreasing. So after checking the membrane, it was observed that there is a leakage from the bottom side of the membrane, so replaced the membrane. Done calibration, run IQC and run one sample control. | SWO-379570 | Assisted Replacement | Electrode | Membrane | Lac | User Error ? | Lac calibration failure even after recent membrane replacement on ABL800 analyzer at Ruby Hall Clinic, Pune | Need to check the procedure which was followed while changing the membrane, also need to check the membrane for any leakage. | After repeating the calibration 2 times, the sensitivity of Lac was found to be decreasing. So after checking the membrane, it was observed that there is a leakage from the bottom side of the membrane, so replaced the membrane. Done calibration, run IQC and run one sample control.
No | Visit | Complaint | RSCH | Rep. Na Testmessung | Kunde meldet abweichung zu Gerät Gebs Betroffener Parameter : Na MedTech am 05.07.2018 vor Ort, Gerinnsel in pH entdeckt, O-Ring ersetzt Danach meldet Kunde erneut Abweichung | 06.07.2018 vor Ort erneut pH Messkammer mit Blutresten pH Elektrode ersetzt i.o Alte Lot = YM05 => Garantie Lac Messkammer auch mit Blutresten Lac fehlte O-Ring, montiert Membranenwartung Lac durchgeführt i.o Ref. Messkammer mit kristallisierten resten Membranenwartung Ref. durchgeführt i.o Kal. und QC i.o Vier Testmessungen zu Ersatzgerät MedTech durchgeführt i.o Abweichungen im Rahmen i.o System i.o | SWO-379272 | Revisit required | Electrode | Electrode | Sealing ring | Sealing ring, pH | Rep. Na Testmessung | Customer reports deviation to device Gebs Affected parameter : Na MedTech on site on 05.07.2018, Clots discovered in pH, O-ring replaced After that, customer reports deviation again | 06.07.2018 on site again pH measuring chamber with blood residues pH electrode replaces i.o Old Lot = YM05 => Warranty Lac measuring chamber also with blood residues Lac missing O-ring, mounted Membrane maintenance Lac carried out i.o Ref. Measuring chamber with crystallized remnants Membrane Maintenance Ref. carried out i.o Kal. and QC i.o Four test measurements for replacement device MedTech carried out i.o Deviations within the framework of i.o. System i.o
No | Visit | Complaint | RAME | Waste detector error 0618. | Found waste detector to be defective. | Replaced the waste detector. | SWO-379427 | Defect | HW | Basic Unit | Waste detector |  | Waste detector error 0618. | Found waste detector to be defective. | Replaced the waste detector.
No | Visit | Complaint | RSCH | Rep. FlexQ | Vermehrt Proben Problem FlexQ bricht Messungen ab | FlexQ Schlitten ist schräg Lässt sich nicht Kalibrieren Komplettes FlexQ durch Leihmodul ersetzt i.o FlexQ kalibriert i.o FlexQ wird in Thalwil repariert | SWO-379274 | Defect | HW | FlexQ | FlexQ |  | Rep. FlexQ | Increases sample problem FlexQ cancels measurements | FlexQ slide is oblique Cannot be calibrated Complete FlexQ replaced by rental module i.o FlexQ calibrates i.o FlexQ is repaired in Thalwil
No | Visit | Complaint | RSCH | Rep. Na Testmessung | Kunde meldet abweichung zu Gerät Neo Betroffener Parameter : Na MedTech am 05.07.2018 vor Ort, Gerinsell in pH entdeckt, O-Ring ersetzt Ref. Membrane erneuert, auch mit Blutresten Gerinnsel in pCO2 gereinigt | 06.07.2018 vor Ort Erneut pH Messkammer mit Blutresten pH Elektrode ersetzt i.o Alte Lot = YT04 => Garantie Kal. und QC i.o Vier Vergleichsmessungen zu MedTech durchgeführt i.o Abweichungen im Rahmen i.o System i.o | SWO-379273 | Assisted Replacement | Electrode | Membrane | Reference | Blood contamination | Rep. Na Testmessung | Customer reports deviation to device Neo Affected parameter : Na MedTech on site on 05.07.2018, Gerinsell discovered in pH, O-ring replaced Ref. Membrane renewed, even with blood residues Clots cleaned in pCO2 | 06.07.2018 on site Again pH measuring chamber with blood residues pH electrode replaces i.o Old lot = YT04 => Warranty Kal. and QC i.o Four comparative measurements carried out for MedTech i.o Deviations within the framework of i.o. System i.o
No | Visit | Complaint | RPOL | Wyciek płynów z uszczelki wlotu | Przeciek z uszczelki wlotu | Wymiana uszczelki wlotu oraz wężyka ściekowego zatkanego skrzepem | SWO-379150 | Cleaned | HW | Basic Unit | Waste tubes | Clot | Fluid leakage from Inlet gasket | Leak from Inlet gasket | The inlet gasket and the drip hose replacement
No | Visit | Complaint | RKOR | SC, SP istall error, calibration error | . | Liquid flow line clean. SC SP replace. Calibration Sample test | SWO-379701 | [5 empty] | SC, SP istall error, calibration error | . | Liquid flow line clean. SC SP replace. Calibration Sample test
No | Visit | Complaint | RCHN | 软件打不开 | 软件打不开 software error | 重新安装AQT90应用软件(By Ran) | SWO-378972 | Defect | SW | RMED SW | RMED SW |  | The software won't open. | Software can't open software error | Reinstall the AQT90 Application (By Ran)
No | Visit | Complaint | RCHN | 进样支架断裂，报错1223和1260 | 进样支架断裂，报错1223和1260，需要借件903-584*2和903-329 | 更换903-584*2和903-329正常。(By Ran) | SWO-378905 | Defect | HW | Aspiration Module | Needle Wash Unit | 1223 | The sample bracket broke, and the errors were 1223 and 1260. | The sample bracket is broken, 1223 and 1260 are reported, and a loan is required 903-584 x2 and 903-329 | Replacement s903-584 x2 and 903-329 are normal. (By Ran)
No | Visit | Complaint | RHUN | Szivárgás hiba | Szivárgás hiba | Bemeneti tűtartó csere, tű kiegyenesítése, szerviz tesztek, kalibrálás, QC mérés. | SWO-379668 | Defect | HW | Inlet module | Probe bend |  | Leakage Error | Leakage Error | Inlet needle holder replacement, needle straightening, service tests, calibration, QC measurement.
No | Visit | Complaint | RCHN | 温度-B | 温度-B,借902-972一个。 | 更换库存902-972后修复 by Andrew | SWO-379522 | [5 empty] | Temperature-B | Temperature-B, borrow edi-902-972. | Repair by Andrew after replacing stock 902-972
No | Visit | Complaint | RCHN | 温度-B | 温度-B,借902-972一个。 | 更换902-972后修复 by Andrew | SWO-379521 | [5 empty] | Temperature-B | Temperature-B, borrow edi-902-972. | Repair by Andrew after replacement 902-972
No | Visit | Complaint | RCHN | 打印机无法打印 | 打印机无法打印 ，借910-286一个。 |  | SWO-379634 | Defect | HW | Analyzer Control Module | Printer |  | The printer can't print | The printer can't print , borrow edith of 910-286.
No | Visit | Complaint | RCHN | 定标失败 | 定标失败，需要借件944-220（14239） | 重新定标正常。(By Ran) Recal,then OK | SWO-378935 | Assisted Replacement | Consumables | Test Kit | Cartridge, D-Dimer |  | Bid failed | Failed bid, 944-220 (14239) required | Recalibration is normal. (By Ran) Recal, then OK
No | Visit | Complaint | RCHN | 进样舱门无法开合 | 进样舱门无法开合，借903-224一个 | 更换903-224一个正常。(By Ran) | SWO-379006 | Defect | HW | Cover and Chassis | Door, Inlet |  | The intake door cannot be opened and closed | The sample door cannot be opened and closed, borrow edges 903-224 | Replace the 903-224 one normal. (By Ran)
Yes | 11 | Y | RPAC | Blank | Software upgrade to V3.4 |  | D49B235E860F2BDF | [5 empty] | Blank | Software upgrade to V3.4
Yes | Visit | Complaint | RPOL | Reklamacje |  | Sprawdzenie analizatora pod kątem ostatnich wymian odczynników i czujników. Zgłoszenie kasety czujników do reklamacji. | SWO-380177 | [5 empty] | Complaints |  | Inspect the analyzer for recent reagent and sensor exchanges. Application of the sensor cassette to the complaint.
Yes | 17 | Y | RPAC | Blank | Software Upgrade to V3.4 |  | D49B224F861A2BDC | Re classified RMED | planned SWO |  |  |  | Blank | Software Upgrade to V3.4
No | 4 | Y | RCZE | Blank |  |  | D49A17F6100FE23D | Adjustment | HW | Oxi module | tHb Cal |  | Blank
No | Visit | Complaint | RAME | The liquid sensors values are out of acceptable range. | Liquid sensor values are out of acceptable range. | Flushed the analyzer. Replaced the inlet gasket holder. Rebooted the analyzer and it fixed the issue. Customer will be receiving credit for a sensor cassette as it failed due to the analyzer issue. | SWO-380399 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway | Liquid sensor | The liquid sensors values are out of acceptable range. | Liquid sensor values are out of acceptable range. | Flushed the analyzer. Replaced the inlet gasket holder. Rebooted the analyzer and it fixed the issue. Customer will be receiving credit for a sensor cassette as it failed due to the analyzer issue.
Yes | 2 | Y | RPAC | Blank | Software Upgrade analyser to Version V3.4, FAN915-384. |  | D499F7471041CF8D | Re classified RMED | planned SWO |  |  |  | Blank | Software Upgrade analyser to Version V3.4, FAN915-384.
No | Visit | Complaint | RKOR | Sample clot error, Inlet probe bent. | Inlet probe bent, Liquid flonw line clean. | Liquid flow line clean. Probe replace. User Training. | SWO-383339 | Defect | HW | Auto Inlet module | Probe bend |  | Sample clot error, Inlet probe bent. | Inlet probe bent, Liquid flonw line clean. | Liquid flow line clean. Probe replace. User Training.
No | Visit | Complaint | RKOR | Sample clot error | . | Liquid flow line clean, SP replace. User training. | SWO-383340 | Cleaned | Fluidic pathway | Fluidic pathway | Clot |  | Sample clot error | . | Liquid flow line clean, SP replace. User training.
Yes | Visit | Complaint | RIND | Solution pack is not installing. | SP and SC Need to replaced with new one. | Replaced solution pack with new one but PCO2 parameter failed to calibrate. changed sensor cassette with new one. installed successfully. handed over in fine working condition | SWO-382564 | [5 empty] | Solution pack is not installing. | SP and SC Need to replaced with new one. | Replaced solution pack with new one but PCO2 parameter failed to calibrate. changed sensor cassette with new one. installed successfully. handed over in fine working condition
No | Visit | Complaint | RCHN | 锂子响应错误 | 锂子响应错误，借903-581一个 | 更换测试卡后修复 | SWO-379780 | Defect | HW | Basic Unit | PCB, Sensor interface |  | Lithium sub-response error | Lithium sub response error, borrow 903-581 one | Repair after replacing test card
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | CO2，CL定标不过 0210,1227 | 更换测试卡后修复 | SWO-379890 | [5 empty] | Test card, reagent package failure | CO2, CL calibration is no more than 0210, 1227 | Repair after replacing test card
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | pco2定标失败 | 连续更换2张测试卡后修复 | SWO-380070 | Assisted Replacement | SC | SC | pCO2 |  | Test card, reagent package failure | pco2 calibration failure | Repair after 2 consecutive test cards
No | Visit | Complaint | RCHN | 光学强度超范围1031，屏幕发黄 | 光学强度超范围1031，屏幕发黄,借903-364，636-571各一个。 | 更换903-364，636-571后修复 by Andrew | SWO-379590 | Defect | HW | Process Module | Optical Unit | 1031 | Optical strength over range 1031, screen yellowing | Optical intensity over range 1031, screen yellowing, borrow 903-364, 636-571 each. | Fix by Andrew after replacing 903-364, 636-571
No | Visit | Complaint | RCHN | HCT值不能被测得 | HCT值不能被测得，借903-329一个 |  | SWO-379838 | Defect | HW | Aspiration Module | Needle Wash Unit |  | HCT value cannot be measured | HCT value cannot be measured, borrowing 903-329 one
No | Visit | Complaint | RDEU | 28.06.2018 16:45Uhr Sr.Jenny/Pfl. Mario 030940113662: Neonlampenfehler 04.07.2018 10:00 Uhr Pfl. Heiko teilt mit, dass der genannte Fehler im letzten Nachtdienst wieder aufgetreten ist | Reinigungsarbeiten, Fehlersuche | Kompletter OxiTest fehlerhaft, Lampeneinheit nochmals justiert, Kal, Service-tHb-Kal, weitere Kal abgewartet, Oximetriewerte deaktiviert, Material aus Geschäftsstelle geholt (11:00-13:30), Oximetriewerte aktiviert, Leihmodul (Spektrometer) eingebaut, Kundenmodul zur Werkstattreparatur abgeholt, Inbetriebnahme, kompletter OxiTest 6.5.2018 $13:00 Telefonisch gelöst durch Anwender (empfohlener Membranwechsel Glu $Lac $Ref)ok, Kal ok, tHb-Kal ok, weitere korrekte Kal abgewartet, QK ok, Spektrometer muß bestellt werden | SWO-380673 | Revisit required | HW | Oxi module | Oxi module |  | 28.06.2018 16:45h Sr.Jenny/Pfl. Mario 030940113662: Neon lamp error 04.07.2018 10:00 am Pfl. Heiko announces that this error occurred again in the last night service | Cleaning work, troubleshooting | Complete OxiTest faulty, lamp unit adjusted again, Kal, service-tHb-kal, further cal awaiting, oximetry values deactivated, material fetched from office (11:00-13:30), oximetry values activated, rental module (spectrometer) built-in, customer module for workshop repair picked up, commissioning, complete OxiTest 6.5.2018 $13:00 By phone solved by user (recommended membrane change Glu $Lac $Ref)ok, Kal ok, tHb-Kal ok, further correct cal waiting, QK ok, spectrometer must be ordered
No | Visit | Complaint | RDEU | Probleme bei Kal | Fehlersuche vor Ort | Driften bei Kal. Ventile aus EleMet und pHBG Modul ausgebaut und gereinigt. Mehreren Kalibrierungen durchgeführt. Keine driften mehr veststellbar. QCs für alle Level durchgeführt. ABL wieder der messbereit. | SWO-381021 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway | Valves of El/Met + pH/BG Module | Problems with Kal | Troubleshooting on site | Drifting at Kal. Valves made of EleMet and pHBG module removed and cleaned. Multiple calibrations performed. No more drifting. QCs for all levels. ABL again ready to measure.
No | Visit | Complaint | RIND | Service Visit. | Start the analyzer, there is multiple error shows in system message, and also sample aborted again and again because of no leading air segment in to the inlet section. | Check the analyzer properly, replace the tubings and inlet gaskit according to the requirement of the analyzer, after this procedure enter in to the service program, adjust the liquid sensor, do the pump calibration, not succesful, increse the 10 ul and again do the pump calibration, pH $Bg module pump ok, again increase 10 ul and do the pump calibration successfully done, remove the dust from the various part of the analyzer, try to convince the technician about the interpretention and relation between the bicarbonate, pH and pCO2. Analyzer handed over to staff in well working condition. | SWO-386386 | Defect | HW | Pump tubes | Fluid manifold |  | Service Visit. | Start the analyzer, there is multiple error shows in system message, and also sample aborted again and again because of no leading air segment in to the inlet section. | Check the analyzer properly, replace the tubings and inlet gaskit according to the requirement of the analyzer, after this procedure enter in to the service program, adjust the liquid sensor, do the pump calibration, not succesful, increse the 10 ul and again do the pump calibration, pH $Bg module pump ok, again increase 10 ul and do the pump calibration successfully done, remove the dust from the various part of the analyzer, try to convince the technician about the interpretention and relation between the bicarbonate, pH and pCO2. Analyzer handed over to staff in well working condition.
No | Visit | Complaint | RDAN | Anne Marie ringede ang. ny ABL på Intensiv (Reklamation). Hun siger at ABL'en ikke giver besked om fuld spildbeholder..... Defekt føler? Aftalt besøg mandag, men ring lige først. | Undersøgt. | Spildføler udskiftet, slutafprøvet. | SWO-380067 | Defect | HW | Basic Unit | Waste detector |  | Anne Marie called Ang. New ABL on intensive (complaint). She says that the Abbey does not notify the full waste container..... Defect feel? Agreed visit Monday, but call right first. | Examined. | Waste sensor replaced, final tested.
Yes | Visit | Complaint | RIND | Interventation Required | Check the Replacement schedule. | Rinse bottle was empty,Change the rinse bottle. Run 2point calibration | SWO-382909 | Re classified RMED | Curtesy Call | Curtesy Call | Curtesy Call |  | Interventation Required | Check the Replacement schedule. | Rinse bottle was empty,Change the rinse bottle. Run 2point calibration
No | Visit | Complaint | RDEU | 09.07.2018 10:55 Uhr Herr Pethe | Fehler 5694 | Gerät demontiert und gereinigt, Küvettentransport geprüft, Funktionskontrolle ok | SWO-380695 | Cleaned | HW | Storage Module | Storage wheel | 5694 | 07/09/2018 10:55 AM Mr Pethe | Error 5694 | Device dismantled and cleaned, cuvette transport tested, function control ok
No | Visit | Complaint | RKK | 溶液パックを認識しない | 確認 | 持参した別ロット溶液パックに交換して改善。 | SWO-382684 | Defect | Consumables | Solution Pack | Chip, Solution Pack |  | Do not recognize the solution pack | Check | Brought another lot solution pack to replace and improve.
No | 16 | Y | RPNZ | Blank | Advice line |  | D49D7A72722F769B | Cleaned | Fluidic pathway | Fluidic pathway | Clot |  | Blank | Advice line
No | Visit | Complaint | RKOR | Pump calibration error | Anlayzer Status check Re-insert the Solution pack Solution pack check -> ok Check the Inlet gasket with holder $ inlet probe -> Failed Replace the inlet probe -> OK | Replace the inlet probe | SWO-383351 | Defect | HW | Inlet module | Probe |  | Pump calibration error | Anlayzer Status check Re-insert the Solution pack Solution pack check -> ok Check the Inlet gasket with holder / inlet probe -> Failed Replace the inlet probe -> OK | Replace the inlet probe
No | Visit | Complaint | RAME | Unable to logon to user or service accounts | Unable to logon to any user accounts including service account. Suspect user database is corrupted. | Uninstalled and reinstalled Application Software v3.4, loaded configuration. Analyzer functioning properly. | SWO-381172 | Defect | SW | SW | Logon | Re install SW | Unable to logon to user or service accounts | Unable to logon to any user accounts including service account. Suspect user database is corrupted. | Uninstalled and reinstalled Application Software v3.4, loaded configuration. Analyzer functioning properly.
Yes | Visit | Complaint | RCRO | CCN/PER | Error 1361 Internal reference electrode error in sensor cassette | All parameters | SWO-387419 | [5 empty] | CCN/PER | Error 1361 Internal reference electrode error in sensor cassette | All parameters
Yes | Visit | Complaint | RCHN | 数据无法传输到系统 | 数据无法传输到系统 | 重新安装LIS系统后修复 by Andrew | SWO-380590 | [5 empty] | Data cannot be transferred to the system | Data cannot be transferred to the system | Fix by Andrew after reinstalling the LIS system
No | Visit | Complaint | RAME | instructed to reboot analyzer 2 sp and 1 sc still with fluid blockage. Vanessa Moore - 828-399-0286 | Fluid blockage error. Found clot in waste connector line. Small clear tubing was disconnected from manifold. | Flushed clot from system. Reconnected clear tubing to manifold. Ran 4 system cycles. Performed cleaning procedure. | SWO-381028 | [5 empty] | instructed to reboot analyzer 2 sp and 1 sc still with fluid blockage. Vanessa Moore - 828-399-0286 | Fluid blockage error. Found clot in waste connector line. Small clear tubing was disconnected from manifold. | Flushed clot from system. Reconnected clear tubing to manifold. Ran 4 system cycles. Performed cleaning procedure.
Yes | Visit | Complaint | RCRO | CCN/PER | Error 1218 System Cycle Failed | Glu, Hct | SWO-387420 | [5 empty] | CCN/PER | Error 1218 System Cycle Failed | Glu, Hct
No | Visit | Complaint | RCHN | 夹阀断裂 | 夹阀断裂，需要借件903-129一个 | 更换903-129后修复 by Andrew | SWO-380080 | Defect | HW | Oxi module | Pinch valve |  | Clip valve break | Clip valve break, need to borrow 903-129 one | Repair by Andrew after replacing 903-129
No | Visit | Complaint | RAME | Ca sensitively high. Wouldn’t calibrate | Ca sensitivity high. Found missing seal rings on Glucose and PO2 causing Reference and Ph chamber to fill with blood. | Replaced Glucose and pO2 seal rings. Replaced reference membrane and pH electrode. Multiple 2 point cals to stabilize | SWO-381081 | User error | Electrode | Electrode | Sealing ring |  | Ca sensitively high. Wouldn’t calibrate | Ca sensitivity high. Found missing seal rings on Glucose and PO2 causing Reference and Ph chamber to fill with blood. | Replaced Glucose and pO2 seal rings. Replaced reference membrane and pH electrode. Multiple 2 point cals to stabilize
No | Visit | Complaint | RDEU | 10.07.2018, Frau Eilers-Gladrow, tel: 05315952045: Oxi Hardware Fehler | Light Intensity too high | Oximetriemodul geprüft, Lampeneinheit justiert, Funktionskontrolle ok | SWO-381404 | Adjustment | HW | Oxi module | Lamp unit | halogen | 10.07.2018, Mrs. Eilers-Gladrow, tel: 05315952045: Oxi Hardware Error | Light Intensity too high | Oximetry module tested, lamp unit adjusted, function control ok
No | Visit | Complaint | RSCH | Blank | Erreur de conditionnement pH-GDS sporadique (garantie) | Remplacé les conduites fluidiques (kit semi-annuel) e.o Nettoyage des chambres de mesures e.o Calibration e.o contrôle de qualité e.o | SWO-381105 | Cleaned | Fluidic pathway | Fluidic pathway | pH/BG Module |  | Blank | Sporadic pH-GDS conditioning error (guarantee) | Replaced fluid pipes (semi-annual kit) e.o Cleaning of e.o measuring chambers Calibration e.o quality control e.o
No | Visit | Complaint | RAME | Leak detected error | Customer was having leak detected error 663 after replacing reference membrane. | I did find a leak coming from the tube leading from the inlet module to the pH $BG module. I reconnected the tube to take care of the leak but continued to get the error during initialization. Talking with John Scott, I disabled the leak detection program in the service setup menu and the issue went away. There are no more leaks in the system and it is operating as it should. | SWO-381136 | Defect | HW | Basic Unit | pH/BG inlet tube |  | Leak detected error | Customer was having leak detected error 663 after replacing reference membrane. | I did find a leak coming from the tube leading from the inlet module to the pH $BG module. I reconnected the tube to take care of the leak but continued to get the error during initialization. Talking with John Scott, I disabled the leak detection program in the service setup menu and the issue went away. There are no more leaks in the system and it is operating as it should.
No | Visit | Complaint | RDEU | 5.6.2018 08:25Uhr Frau Otto 051916023550: Das Gerät macht ungewöhnlich, quitschende Geräusche | System überprüft => Pumperad 2 quietscht | Pumpenschlauch erneuert , CRP Qc durchgeführt | SWO-380848 | Defect | HW | Process Module | Pump tubes |  | 5.6.2018 08:25pm Mrs Otto 051916023550: The device makes unusual, squealing noises | System checked => Pump wheel 2 squeaks | Pump hose renewed, CRP Qc performed
No | Visit | Complaint | RDEU | 26.06.2018 Herr Betsch 015204767315 meldet, dass die Anbindung nicht funktioniert. | Datenverbindung nicht möglich | IP 10.200.112.181 Sub: 255.255.255.0 GW 10.200.112.110 Alle Einstellungen sind korrekt. Es müsste lediglich dem KFH die Mac- Adresse des neuen Gerätes mitgeteilt werden damit es frei geschaltet werden kann. | SWO-381814 | Setup | SW | SW | Wireless connection | MAC adress | 26.06.2018 Mr. Betsch 015204767315 reports that the connection is not working. | Data connection not possible | IP 10.200.112.181 Sub: 255.255.255.0 GW 10.200.112.110 All settings are correct. All that would have to be communicated to the KFH is the Mac address of the new device in order for it to be unlocked.
No | Visit | Complaint | RDEU | 10.07.2018 Frau Rathmer 025639126985 meldet Crea Fehler bei ihrem DEMO-Gerät. | Crea Empf. Fehler, Flüssigkeitstransportfehler | Gerät überprüft, Fließwege gespült, Einlaß gereinigt, Verbrauchsmittel reinstalliert Keine Besserung | SWO-381828 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway | Crea, not solved | 10.07.2018 Ms. Rathmer 025639126985 reports Crea error with her DEMO device. | Crea Empf. Errors, fluid transport errors | Device checked, flow paths rinsed, inlet cleaned, consumables reinstalled No improvement
No | Visit | Complaint | RDEU | Einlassdichtung undicht | Flüssigkeitstransport fehlerhaft | Fehleranalyse, Einlassdichtung erneuert, Kalibrierung und Qualitätskontrolle i. O. | SWO-381813 | Defect | HW | Auto Inlet module | Inlet gasket |  | Inlet seal leaks | Liquid transport faulty | Fault analysis, inlet seal renewed, calibration and quality control in the
No | Visit | Complaint | RBNL | USB stick(s) worden niet herkend | USB stick(s) worden niet herkend waardoor er geen Status overzicht gemaakt kan worden. | Bij aankomst een andere USB stick aangesloten en deze functioneert direct en is het mogelijk een Status overzicht te maken. Via de taskmanager de verkenner gestart en ook hier is de USB -stick zichtbaar. Wel is er op de achtergrond nog het Found new hardware scherm zichtbaar. Dit is gekomen door een eerste keer aansluiten van een nieuwe USB stick. USB sticks van de klant aangesloten en deze functioneren nu ook. Oorzaak: onbekend. | SWO-381611 | Adjustment | HW | DMS_PC | USB port |  | USB Stick (s) not recognized | USB stick (s) are not recognized, so no Status overview can be created. | A different USB stick is connected on arrival and it functions directly and it is possible to create a Status overview. Through the TaskManager the explorer started and also here the USB stick is visible. However, the Found new hardware screen is still visible in the background. This has come through a first time connecting a new USB stick. USB sticks from the customer and they now function as well. Cause: Unknown.
Yes | Visit | Complaint | RCHN | 无法连接试剂包 | 无法连接试剂包，借902-989、903-228各一个 | 重新连接线路板后修复 by Andrew | SWO-380674 | [5 empty] | Unable to connect the reagent package | Cannot connect the reagent package, borrow 902-989, 903-228 each | Fix by Andrew after reconnecting the circuit board
No | Visit | Complaint | RIND | Service Visit. | Sensor cassette calibration failed and due to leakage of sensor cassette sensor board has damaged. | Need to replace the sensor board as well as sensor cassette. | SWO-381662 | [5 empty] | Service Visit. | Sensor cassette calibration failed and due to leakage of sensor cassette sensor board has damaged. | Need to replace the sensor board as well as sensor cassette.
No | Visit | Complaint | RAME | GIVING TEMP ERROR 1303 DURING SYSTEM CYCLE. | GIVING TEMP ERROR 1303 DURING SYSTEM CYCLE. TRIED ZERO CHECK AND TEMP DID NOT GO UP. SOME PARTS WERE SENT TO FEDEX AND SOME TO MY HOME. HAD TO GO BACK AND GET PARTS | REPLACED SIGNAL DATA BOARD AND SENSOR BOARD. RAN ZERO CHECK | SWO-381769 | [5 empty] | GIVING TEMP ERROR 1303 DURING SYSTEM CYCLE. | GIVING TEMP ERROR 1303 DURING SYSTEM CYCLE. TRIED ZERO CHECK AND TEMP DID NOT GO UP. SOME PARTS WERE SENT TO FEDEX AND SOME TO MY HOME. HAD TO GO BACK AND GET PARTS | REPLACED SIGNAL DATA BOARD AND SENSOR BOARD. RAN ZERO CHECK
No | Visit | Complaint | RAME | device is displaying possible fluid blockage during system cycle. | Device was encountering a possible fluid blockage error during system cycle. | I checked all of the tubing and connections to make sure there were no kinks or bad connections. I also checked the cable connections for the hemolyzer and spectrometer. The issue was not resolved after reassembly so I performed the liquid sensor calibration and service test. The service test indicated that the spectrometer was experiencing excessive noise and had a low offset as well. I ran the service test several times with no improvment. Joe Bondarenko will come out to the site tomrrow 7-12-18 and replace the hemolyzer as I was not in posession of one. Staff have been notified. POC staff were gone by the time I departed. Verbal signature used. | SWO-381922 | [5 empty] | device is displaying possible fluid blockage during system cycle. | Device was encountering a possible fluid blockage error during system cycle. | I checked all of the tubing and connections to make sure there were no kinks or bad connections. I also checked the cable connections for the hemolyzer and spectrometer. The issue was not resolved after reassembly so I performed the liquid sensor calibration and service test. The service test indicated that the spectrometer was experiencing excessive noise and had a low offset as well. I ran the service test several times with no improvment. Joe Bondarenko will come out to the site tomrrow 7-12-18 and replace the hemolyzer as I was not in posession of one. Staff have been notified. POC staff were gone by the time I departed. Verbal signature used.
No | Visit | Complaint | RCHN | GLU 定标失败 | GLU 定标失败 借903-581一个 | 更换903-581一个正常。 by Andrew | SWO-381400 | Defect | HW | Basic Unit | PCB, Sensor interface | Glu | GLU bid failure | GLU bid failure borrowed 903-581 one | Replace the 903-581 one normal. by Andrew
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 压力蒸空错误 | 更换试剂包后修复 | SWO-381386 | Assisted Replacement | SP | SP | Pressure Test |  | Test card, reagent package failure | Pressure steaming error | Repair after replacement of reagent kit
No | Visit | Complaint | RCHN | 定标失败 | 定标失败 借944-220（14531）一个 |  | SWO-381482 | Assisted Replacement | Consumables | Test Kit | Cartridge, D-Dimer |  | Bid failed | Bid failure Borrows 944-220 (14531) one
Yes | Visit | Complaint | RFRA | Configuration date et heure. | Configuration date et heure. | Configuration date et heure. Remplacement des consommables suite changement date et heure. Calibration 2 poins ok. | SWO-381745 | [5 empty] | Date and time set-up. | Date and time set-up. | Date and time set-up. Replacement of consumables following change date and time. Calibration 2 punches ok.
No | Visit | Complaint | RIND | Printer unit of ABL9 is not working | Checked software version and upgraded to 1.1.1. Then reconnected printer.Instrument is showing printer door is open.Checked cable of door lock sensor. It is damaged. | Need to replace printer unit. | SWO-381407 | [5 empty] | Printer unit of ABL9 is not working | Checked software version and upgraded to 1.1.1. Then reconnected printer. Instrument is showing printer door is open. Checked cable of door lock sensor. It is damaged. | Need to replace printer unit.
No | Visit | Complaint | RKK | Neon lump intensity too low, Neon number too lowが発生。 納入後１年未満。 oxiモジュールを代替えに交換し様子見。 |  |  | SWO-382301 | Defect | HW | Oxi module | Lamp unit | Neon | Neon lump intensity too low, Neon number too low. Less than one year after delivery. The oxi module was replaced with an alternative.
No | Visit | Complaint | RKK | ｐO2応答異常 | pO2応答異常 | ｐO2電極交換 | SWO-390922 | Assisted Replacement | Electrode | Electrode | pO2 | mail: what does abnormal mean (outside spec?) | pO2 response abnormality | pO2 response abnormality | pO2 electrode replacement
No | Visit | Complaint | RKK | ミキサーエラー。 | ミキサートレイユニット交換にて改善 | ミキサートレイユニット交換にて改善 | SWO-384772 | Defect | HW | FlexQ | Mixer Tray |  | Mixer error. | Improved by mixer tray unit replacement | Improved by mixer tray unit replacement
No | Visit | Complaint | RAME | Analyzer is having rinse error 508 | Rinse error present when arrived onsite. Noticed flat membrane for pH $BG module was damaged | Replaced air flat membrane for pH $BG module. All calibrations and 3 Levels of QC passed as per manufacturer specification | SWO-381807 | Defect | HW | pH/BG Module | Valve, air | Membrane, 508 | Analyzer is having rinse error 508 | Rinse error present when arrived onsite. Noticed flat membrane for pH /BG module was damaged | Replaced air flat membrane for pH $BG module. All calibrations and 3 Levels of QC passed as per manufacturer specification
Yes | Visit | Complaint | RSCH | Blank | Aucune détection d’ouverture du volet aspiration d’échantillons (sporadique) | Remplacé le cable FPC 6pole du module d’aspiration e.o Test de la section fluidique e.o Test échantillon sang e.o Calibration e.o CQ e.o | SWO-382820 | [5 empty] | Blank | No detection of opening of sample suction flap (sporadic) | Replaced the FPC 6pole cable of the e.o suction module Testing the fluidic section e.o Blood sample test e.o Calibration e.o CQ e.o
Yes | Visit | Complaint | RAME | ls transport failure replace sp, inlet gasket, flush sc replaced lac acid failed. contact: Sharmaine Cabelin - 515-241-5320 | Lactate response error. | Problem was corrected on case CAS-369398-T8N8 SWO-377054. | SWO-382478 | Re classified RMED | No Visit |  |  | INGZI mail from JGIBB | ls transport failure replace sp, inlet gasket, flush sc replaced lac acid failed. contact: Sharmaine Cabelin - 515-241-5320 | Lactate response error. | Problem was corrected on case CAS-369398-T8N8 SWO-377054.
No | Visit | Complaint | RAME | Leak detected. | Found machine with multiple errors under fluid transport test screen | Performed all calibrations under fluid transport screen. I had to perform the solution pack test twice, but all errors went away. Performed start up of analyzer. Got multiple calibration errors. Tried calibrating many times. Replaced Sensor cassette. All errors went away except PO2. Replaced solution pack with a pack that had previously been removed from analyzer for troubleshooting. PO2 calibrations came in and analyzer was green. | SWO-381940 | Defect | SP | SP | Leak detected |  | Leak detected. | Found machine with multiple errors under fluid transport test screen | Performed all calibrations under fluid transport screen. I had to perform the solution pack test twice, but all errors went away. Performed start up of analyzer. Got multiple calibration errors. Tried calibrating many times. Replaced Sensor cassette. All errors went away except PO2. Replaced solution pack with a pack that had previously been removed from analyzer for troubleshooting. PO2 calibrations came in and analyzer was green.
No | Visit | Complaint | RLTD | Customer reported a very noisy fan. | PCU fan noisy, bearing failed. | Replaced fan, now running with no noise. Tested printer after reconnect. Replaced with part number 580-418, not recognised in swo lines. | SWO-382886 | Defect | HW | DMS_PC | Fan | Noise | Customer reported a very noisy fan. | PCU fan noisy, bearing failed. | Replaced fan, now running with no noise. Tested printer after reconnect. Replaced with part number 580-418, not recognised in swo lines.
No | Visit | Complaint | RCHN | 病人结果与临床不符 | 病人结果与临床不符 | 与老师解释，并重新做质控，全部通过 by Andrew | SWO-381988 | [5 empty] | Patient results are not clinically consistent | Patient results are not clinically consistent | explain with the teacher, and redo the qc, all by By Andrew
No | Visit | Complaint | RDEU | Benötigt viel Zeit für die Flüssigkeitssensorkalibrierung, häufige Meldung : mögliche Verstopfung der Fließwege | Oximetrieventil schaltet schlecht | Sporadisch sind die Hb Werte zu niedrig, Fehlersuche, Austausch Oximetrieventil, Funktionskontrolle | SWO-382320 | [5 empty] | Requires a lot of time for liquid sensor calibration, frequent message: possible blockage of the flow paths | Oximetry valve switches badly | Sporadically, the Hb values are too low, troubleshooting, replacement of oximetry valve, function control
No | Visit | Complaint | RIND | ABL80 instrument has problem in Na and K results. | Checked calibration data and consumables replacement data. Also checked patients data.Need to check with Quality control | Checked with quality control and found that it was out of range. Need to replace SC. | SWO-382589 | [5 empty] | ABL80 instrument has problem in Na and K results. | Checked calibration data and consumables replacement data. Also checked patients data. Need to check with Quality control | Checked with quality control and found that it was out of range. Need to replace SC.
No | Visit | Complaint | RAME | device is displaying possible fluid blockage during system cycle. | Fluid blockage error 2106 | Rebooted analyzer, checked liquid sensor status, checked pinch valve function, checked co-ox status in service. Spectrometer hardware was failing. Replaced hemolyzer, pinch valve, tubing. Readjusted waste tubing. Reimbursed customer for 1 solution pack and 1 sensor cassette | SWO-382623 | [5 empty] | device is displaying possible fluid blockage during system cycle. | Fluid blockage error 2106 | Rebooted analyzer, checked liquid sensor status, checked pinch valve function, checked co-ox status in service. Spectrometer hardware was failing. Replaced hemolyzer, pinch valve, tubing. Readjusted waste tubing. Reimbursed customer for 1 solution pack and 1 sensor cassette
No | Visit | Complaint | RAME | device is displaying possible fluid blockage during system cycle. | Fluid blockage error 2106 | Replaced all parts in annual pm kit. | SWO-382627 | [5 empty] | device is displaying possible fluid blockage during system cycle. | Fluid blockage error 2106 | Replaced all parts in annual pm kit.
No | Visit | Complaint | RCHN | 电源烧毁 | 电源烧毁 借910-285一个 | 更换910-285一个后故障仍未排除，加借910-342一个。 by Andrew | SWO-381388 | Defect | HW | DMS_PC | CPU |  | Power burned | Power burned, borrow 910-285 one. | Replace 910-285 one after the fault is still not resolved, plus 910-342 one. by Andrew
Yes | Visit | Complaint | RCHN | 电源烧毁 | 更换910-285一个后故障仍未排除，加借910-342一个。 | 更换910-285和910-342后修复 by Andrew | SWO-382169 | Re classified RMED | same day upSWO |  |  |  | Power burned | Replace 910-285 one after the fault is still not resolved, plus 910-342 one. | Fix by Andrew after replacing 910-285 and 910-342
No | Visit | Complaint | RCHN | 屏幕花屏 | 屏幕花屏,借910-298一个。 | 更换910-298后修复 by Andrew | SWO-381385 | Defect | HW | DMS_PC | display |  | Screen flower screen | Screen flower screen, borrow 910-298 one. | Repair by Andrew after replacing 910-298
No | Visit | Complaint | RCHN | 报错5297 | 报错5297，借636-536,836-166各一个 | 更换636-536,836-166后修复 by Andrew | SWO-381374 | Defect | HW | Process Module | sensor, cup | 5297 | 5297 error | 5297, borrow 636-536, 836-166 each | Repair by Andrew after replacing 636-536, 836-166
No | Visit | Complaint | RKK | SC初期化時に結構か頻度で温度エラー。 SCロットを替えても変わらず。 | 装置引き上げ | 代替機設置 | SWO-382761 | [5 empty] | Temperature error at the time of SC initialization. It does not change even if the SC lot is changed. | Equipment pulling up | Installation of alternative machines
No | Visit | Complaint | RFRA | Difficultés à installer les consommables | Vérification complète de l'analyseur, de l'historique des calibrations et des activités | Nettoyage fluidique $ $ nettoyage des contacts de la cassette $ $ essais divers $ $ réinstallation de la cassette $ $ configuration du ticket patient | SWO-382361 | [5 empty] | Difficulties in installing consumables | Full verification of the analyzer, calibration history and activities | Fluid cleaning $$ cleaning of contacts of the cassette $$ various trials $ $ reinstallation of the cassette $$ configuration of the patient ticket
No | Visit | Complaint | RCHN | LAC定标失败 | LAC定标失败，借945-619一个 | 更换945-619后修复 by Andrew | SWO-380766 | Assisted Replacement | Electrode | Electrode | Lac |  | LAC calibration failure | LAC bid failed, borrow 945-619 one | Repair by Andrew after replacement 945-619
No | Visit | Complaint | RLTD | Gordon Nickel on 01896 826245 | Ca drift on firs calibration after cleaning | Reference electrodes sent direct to customer to fit, this has been completed and analysers have not failed calibration since Customer to monitor in the coming weeks Old electrodes removed to be sent back due warranty return | SWO-382899 | Assisted Replacement | Electrode | Electrode | Reference | Phone Call | Gordon Nickel on 01896 826245 | Ca drift on firs calibration after cleaning | Reference electrodes sent direct to customer to fit, this has been completed and analysers have not failed calibration since Customer to monitor in the coming weeks Old electrodes removed to be sent back due warranty return
No | Visit | Complaint | RIND | Service Visit. | In the analyzer the insert cartridge door is open, need to close. | Go to service program and give the command of safe position, successfully done, storage wheel door close, exit from the service program after start up analyzer shows replacement, need to calibrate the PCT, calibrate the PCT lot with proper procedure, successfully done analyzer handed over to staff in well working condition. | SWO-382592 | Defect | SW | Analyzer Control Module | Analyzer Logic |  | Service Visit. | In the analyzer the insert cartridge door is open, need to close. | Go to service program and give the command of safe position, successfully done, storage wheel door close, exit from the service program after start up analyzer shows replacement, need to calibrate the PCT, calibrate the PCT lot with proper procedure, successfully done analyzer handed over to staff in well working condition.
No | Visit | Complaint | RFRA | 1401 Erreur joint d'introduction régulièrement. | Joint d'introduction ne coulisse pas, | Nettoyage sonde, le joint coulisse mieux. Contrôle en mode service: ok. Nettoyage: à l'hypochlorite du circuit fluidique, Vanne hémolyseur. Remplacement: Joint d'introduction, Joint connexion, tubulure vanne oxi. Test service: Fluid transport, Oxi $pO2. Calibration, passage CQ: A, B et C. A la demande du laboratoire: chgt Cassette qui ne termine. | SWO-383061 | Defect | HW | Auto Inlet module | Inlet gasket | 1401 | 1401 Joint introductory error regularly. | Introductory seal does not slide, | Cleaning probe, the joint slides better. Control in service mode: ok. Cleaning: to the hypochlorite of the fluid circuit, Valve hemolyseur. Replacement: Introductory joint, joint connection, tube valve oxi. Test service: Fluid transport, Oxi $pO2. Calibration, passage CQ: A, B and C. At the request of the laboratory: chgt Cassette which does not finish.
No | Visit | Complaint | RFRA | Pb d’écran tactile | Contrôle de l’écran tactile | Changement de l’écran tactile. Essais | SWO-382961 | Defect | HW | DMS_PC | display |  | Touch screen pb | Touch screen control | Changing the touch screen. Tests
No | Visit | Complaint | RPOL | Błąd uszczelki wlotu | Sprawdzenie igły. | Wymiana igły na nową. Testy serwisowe - ok. | SWO-382837 | Defect | HW | Auto Inlet module | Probe |  | Inlet GASKET Error | Check the needle. | Replacing the needle with a new one. Service Tests-OK.
No | Visit | Complaint | RKOR | Analyzer Status is red | No test remains in Solution pack -> Replace the Solution pack | Replace the Solution pack | SWO-383366 | [5 empty] | Analyzer Status is red | No test remains in Solution pack -> Replace the Solution pack | Replace the Solution pack
No | Visit | Complaint | RAME | instructed to reboot analyzer 2 sp and 1 sc still with fluid blockage. Vanessa Moore - 828-399-0286 | Found evidence of another clot in waste lines. | Installed PM kit and flushed hemolyzer. Checked functionality of pinch valve. Ran system cycles with no further errors. | SWO-383246 | [5 empty] | instructed to reboot analyzer 2 sp and 1 sc still with fluid blockage. Vanessa Moore - 828-399-0286 | Found evidence of another clot in waste lines. | Installed PM kit and flushed hemolyzer. Checked functionality of pinch valve. Ran system cycles with no further errors.
No | Visit | Complaint | RCHN | PH,PCO2定标失败 | PH,PCO2定标失败 | 排出测试卡，试剂包内的气泡后修复 by Andrew | SWO-382772 | [5 empty] | PH, PCO2 calibration failed | PH, PCO2 calibration failed | Fix by Andrew after draining the test card and fixing the bubbles in the reagent package
No | Visit | Complaint | RCHN | 接口板腐蚀 | 接口板腐蚀,借902-972一个。 | 拆库存902-972后修复 by Andrew | SWO-382193 | [5 empty] | Interface plate corrosion | Interface plate corrosion, borrow 902-972 one. | Repair by Andrew after dismantling stock 902-972
No | Visit | Complaint | RAME | device is displaying possible fluid blockage during system cycle. | Fluid blockage error and spectrometer hardware status failure | Attempted to perform cleaning procedure. Was not able to get past cleaning cassette requirement. Performed system cycle and co-ox service test. Will run cleaning procedure through analyzer on Monday. Opened CQC case and called tech support. | SWO-383230 | [5 empty] | device is displaying possible fluid blockage during system cycle. | Fluid blockage error and spectrometer hardware status failure | Attempted to perform cleaning procedure. Was not able to get past cleaning cassette requirement. Performed system cycle and co-ox service test. Will run cleaning procedure through analyzer on Monday. Opened CQC case and called tech support.
No | Visit | Complaint | RCHN | 屏幕触摸不灵敏 | 屏幕触摸不灵敏，需要借件910-298一个 | 更换910-298后修复 by Andrew | SWO-382173 | Defect | HW | DMS_PC | display |  | Screen touch is not sensitive | Screen touch is not sensitive, need to borrow 910-298 a | Repair by Andrew after replacing 910-298
No | Visit | Complaint | RCHN | 条码扫描器故障 | 条码扫描器故障 借910-304一个 | 更换910-304后修复 by Andrew | SWO-381973 | Defect | HW | Storage Module | Barcode reader, storage wheel |  | Barcode scanner failure | Barcode scanner failure borrowed 910-304 one | Repair by Andrew after replacing 910-304
No | Visit | Complaint | RCHN | 报错3412 | 报错3412，借903-112一个 | 更换903-112后修复 by Andrew | SWO-381995 | Defect | HW | Aspiration Module | PCB, Wet section | 3412 | Error 3412 | 3412, loan 903-112 one | Fix by Andrew after replacing 903-112
No | Visit | Complaint | RCHN | 不能开机 | 不能开机，借910-343一个 | 更换910-343后修复 by Andrew | SWO-381486 | Defect | HW | Analyzer Control Module | CPU |  | Cannot power on | Cannot power on, borrow 910-343 one | Fix by Andrew after replacing 910-343
Yes | Visit | Complaint | RCHN | AQT90机器故障，卡盘 | 测试卡卡盘 | 取出断脚测试卡，Tni-14374，剩余杯位14个 | SWO-384187 | [5 empty] | AQT90 machine failure, chuck | Testing the kaka plate | Take out the broken foot test card, Tni-14374, 14 remaining cup slots
No | Visit | Complaint | RDEU | Glu. Instabil. Kalibrationdrift. | Glu. driftet sehr häufig bei der Kalibration. | Elektrode wurde bereits vom Anwender membraniert. Kallösung S1820 mit der Lot: BD50 ausgebaut und gegen eine neuere Charge ersetzt. Kalibration mehrfach wiederholt. | SWO-383382 | Assisted Replacement | Consumables | Solution | Cal 1 |  | Glu. Unstable. Calibration drift. | Glu. drifts very often during calibration. | Electrode has already been membraneed by the user. Cal solution S1820 with the lot: BD50 removed and replaced with a newer batch. calibration repeated several times.
No | Visit | Complaint | RAME | 0618-waste detector errors are present and 65% full is displayed when an empty bottle is present. | When checking the waste detector in service mode, the T_ambient voltage is observed outside of the acceptable range and a waste status error is generated. Also, the waste detector is observed connected to j10 pins 3 & 4 of the bottle illumination pcb instead of pins 1 & 2. | Connected the waste detector to pins 1 & 2 of the bottle illumination pcb and successfully reran the waste detection program in service. The T_ambient voltage is within the acceptable range and no errors are generated. Analyzer is within acceptable operating specifications. | SWO-383186 | Adjustment | HW | Basic Unit | Waste detector | cable positioned wrong | 0618-waste detector errors are present and 65% full is displayed when an empty bottle is present. | When checking the waste detector in service mode, the T_ambient voltage is observed outside of the acceptable range and a waste status error is generated. Also, the waste detector is observed connected to j10 pins 3 & 4 of the bottle illumination pcb instead of pins 1 & 2. | Connected the waste detector to pins 1 & 2 of the bottle illumination pcb and successfully reran the waste detection program in service. The T_ambient voltage is within the acceptable range and no errors are generated. Analyzer is within acceptable operating specifications.
No | Visit | Complaint | RDEU | 12.07.2018 Sr. Anja 0211904361530 meldet, dass der Bidlschirm flackert und keine Messung möglich ist. | Kein Fehler festgestellt. Am 12.07.2018 an das Gerät den Fehler :“Rotatestoragewheel_24“ angezeigt. | Storage Wheel demontiert und gereinigt, Lichtsensoren gereinigt, Punch Up am Storage Wheel demontiert und gereinigt. Gerät mit Testkassetten bestückt und QC durchgeführt. | SWO-383483 | Cleaned | HW | Storage Module | Storage wheel |  | 12.07.2018 Sr. Anja 0211904361530 reports that the bidl screen is flickering and no measurement is possible. | No error detected. On 12/07/2018, the error :"Rotatestoragewheel_24" was displayed to the device. | Storage Wheel dismantled and cleaned, light sensors cleaned, Punch Up disassembled and cleaned on the Storage Wheel. equipped with test cassettes and performed QC.
No | 2 | Y | RCZE | Blank |  |  | D4A2B2944A1B75E0 | Defect | HW | Process Module | Process Wheel | CRP level 1 | Blank
Yes | Visit | Complaint | RDEU | QVC Grenzverletzung bei CRP in Level gelb und rot. | Analysator repariert in -R-Zentrale, geliefert ins Technikbüro Frankfurt. Analysator noch verpack, ausgepackt. Sehr hohe Rüstzeit durch Stau, ausgelöst durch Autobahnsperrung, bis drei Stunden Wartezeit bis Polizei aus Stau befreit. | Analysator zum Anwender gebracht, in Funktion genommen und festgestellt, dass nur die Parameter CRP und D-Dimer verwendet werden können. Bei TnT ist nicht die richtige Los Nummer kalibriert. Kalibrationskassette über Logistik Abteilung in CPH bestellt. Anwender kalibrieren bei Lieferung der TnT Testkassette führen die Qualitätskontrolle durch. Bei Erfolg melden sie sich und das Leihgerät kann abgeholt werden. | SWO-383056 | Re classified RMED |  |  |  |  | QVC Grenzverletzung bei CRP in Level gelb und rot. | Analyzer repaired in -R headquarters, delivered to the Frankfurt technical office. Analyzer still packaged, unpacked. Very high set-up time due to traffic jams, triggered by motorway closure, up to three hours waiting time until police free from traffic jams. | Analyzer to the user, put into operation and determined that only the parameters CRP and D-Dimer can be used. TnT does not calibrate the correct lot number. Calibration cassette ordered via logistics department in CPH. Users calibrating when the TnT test cartridge is delivered perform the quality control. If successful, they report and the rental device can be picked up.
No | Visit | Complaint | RDEU | LP muss immer wieder ersetzt werden, diverse Fließfehler, keine Messungen möglich, | Fließfehler, keine Messung möglich, Vacuumfehler, Verbindungsschlauch zw. Messkammer und Hemolysator def. | Fließwege gereinigt, Schlauch ersetzt, Gerät kalibriert, | SWO-383478 | Defect | HW | Basic Unit | Tubes |  | LP has to be replaced again and again, various flow errors, no measurements possible, | Flow error, no measurement possible, vacuum error, connection hose between Measuring chamber and hemolysator def. | Flow paths cleaned, hose replaced, device calibrated,
No | Visit | Complaint | RKOR | Sample clot error, SP smart chip error | . | Liquid flow line clean. SP replace. Gasket hold replace. Gasket module clean. Service mode check. Calibration QC check. Sample test. | SWO-383348 | Cleaned | Fluidic pathway | Fluidic pathway | Clot |  | Sample clot error, SP smart chip error | . | Liquid flow line clean. SP replace. Gasket hold replace. Gasket module clean. Service mode check. Calibration QC check. Sample test.
Yes | Visit | Complaint | RCHN | 机器花屏 | 机器花屏，需要借件910-320和910-274 | 更换910-320和910-274后修复 by Andrew | SWO-381469 | Defect | HW | DMS_PC | Power supply |  | Machine Flower Screen | Machine flower screen, need to borrow 910-320 and 910-274 | Fix by Andrew after replacing 910-320 and 910-274
No | Visit | Complaint | RSCH | Pikett Rep. Oxi | Kunde telefonierte am 13.07.2018 um 16:50 tHb Drift Danach QC Level 1 FMetHb und FO2Hb oberhalb $unterhalb Kontrollbereich Kunde wünscht ein Pikett-Einsatz am 14.07.2018 Kunde wurde über zusätzliche Kosten informiert | Vor Ort zusätzlich QC Level 3 Na zu hoch Ref. Membrane erneuert i.o Na Messkammer mit resten von Watte-Stäbchen, gereinigt i.o Komplettes Oxi-Modul ersetzt i.o Nachfüllung, LS Justage und Pumpen Kal. durchgeführt i.o Kal., QC und tHb i.o System i.o Kunden Oxi-Modul wird in Thalwil repariert | SWO-383605 | Defect | HW | Oxi module | Oxi module |  | Pikett Rep. Oxygen | Customer phoned on 13/07/2018 at 16:50 tHb Drift After that QC Level 1 FMetHb and FO2Hb above /unterhalb control area Customer wishes a picket operation on 14.07.2018 Customer has been informed of additional costs | On-site additional QC Level 3 Na too high Ref. Membrane renewed i.o Na measuring chamber with remnants of cotton swabs, cleaned i.o Complete Oxi module replaces i.o Refilling, LS Adjustment and Pumps Kal. carried out i.o Kal., QC and tHb i.o System i.o Customer Oxi module is repaired in Thalwil
No | Visit | Complaint | RAME | SP errors | Fluid errors will not start up | Flushed with hot water and ran Diag, all passed | SWO-383398 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | SP errors | Fluid errors will not start up | Flushed with hot water and ran Diag, all passed
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | na定标失败 | 更换测试卡后修复 | SWO-383442 | [5 empty] | Test card, reagent package failure | na bid failed | Repair after replacing test card
No | Visit | Complaint | RCHN | 接口板腐蚀 | 接口板腐蚀，借902-972一个 | 更换902-972修复 by Andrew | SWO-382811 | [5 empty] | Interface plate corrosion | Interface plate corrosion, borrow 902-972 one | Replace 902-972 Repair by Andrew
Yes | 3 | Y | RPAC | Blank | Software Upgrade. |  | D4A44547C0485D37 | Re classified RMED | planned SWO |  | SW |  | Blank | Software Upgrade.
Yes | 6 | Y | RPAC | Blank | Software Upgrade analyser to Version V 3.4, FAN915-384. |  | D4A2F1B84A0F60FA | Re classified RMED | planned SWO |  | SW |  | Blank | Software Upgrade analyser to Version V 3.4, FAN915-384.
No | Visit | Complaint | RDEU | 10.07.2018 Frau Rathmer 025639126985 meldet Crea Fehler bei ihrem DEMO-Gerät. | Crea Kal Fehler | Bertriebsmittel geliefert und installiert | SWO-383665 | Assisted Replacement | SC | SC | Crea | Cal | 10.07.2018 Ms. Rathmer 025639126985 reports Crea error with her DEMO device. | Crea Kal Error | Means of operation supplied and installed
Yes | 1 | Y | RPAC | Blank | Clot catchers-replacement |  | D4A2C4A74A56403A | Re classified RMED | Logistical upSWO |  |  |  | Blank | Clot catchers-replacement
No | Visit | Complaint | RAME | device is displaying possible fluid blockage during system cycle. | Analyzer needs cleaning procedure performed. Fluid blockage error and spectrometer hardware failure. | Checked status of analyzer. Performed cleaning procedure. Performed liquid sensor adjust, pinch valve test and Co-ox service test. | SWO-383921 | [5 empty] | device is displaying possible fluid blockage during system cycle. | Analyzer needs cleaning procedure performed. Fluid blockage error and spectrometer hardware failure. | Checked status of analyzer. Performed cleaning procedure. Performed liquid sensor adjust, pinch valve test and Co-ox service test.
No | Visit | Complaint | RCHN | 质控对比 | 老师怀疑机器检测钠离子不准 | 质控对比后解决by tim | SWO-384168 | No fault found | No fault found | No fault found | QC |  | Quality control comparison | Teacher suspects machine sodium ions are not allowed to detect | Quality control comparison after solving by tim
No | Visit | Complaint | RCHN | 系统检查失败 1031 | 系统检查失败 1031,借903-364一个。 | 借的903-364有质量问题， 加借 903-364一个。此单关单by tim | SWO-383638 | Defect | HW | Process Module | Optical Unit | 1031 | System check failed 1031 | System check failed 1031, borrowing 903-364 one. | Borrowing 903-364 has quality problems, plus 903-364 one. This single single by tim
No | Visit | Complaint | RIND | Printer unit of ABL9 is not working | Opened instrument and removed old printer unit. | Replaced printer unit and done setup in software. Then checked with print out. It is working fine. | SWO-386518 | [5 empty] | Printer unit of ABL9 is not working | Opened instrument and removed old printer unit. | Replaced printer unit and done setup in software. Then checked with print out. It is working fine.
No | Visit | Complaint | RDEU | halogen intensity too high | Halogenlampen Intensität zu hoch. Auf Ersatzteil in der GS gewartet. Nicht geliefert worden. | In den Serviceeinstellungen so lange die Halogenlampe justiert bis die Werte brauchbar waren. Mehrere Kalibrationen durchgeführt. Und danach QC. System wieder einsatzbereit. | SWO-383580 | Adjustment | HW | Oxi module | Lamp unit | Revisit, Spare parts? | halogen intensity too high | Halogen lamps intensity too high. Waiting for spare part in the GS. Not delivered. | In the service settings, the halogen lamp is adjusted until the values are usable. Several calibrations performed. And then QC. system is ready for use again.
No | Visit | Complaint | RAME | Doesn't recognize Rinse solution barcode | Analyzer doesn't recognize Rinse solution barcode. I believe it is a software key issue. Worked with Mark Z. And RMED to get new key for instrument. | Replaced software keys with new keys from RMED | SWO-384827 | Defect | HW | Analyzer Control Module | SW key |  | Doesn't recognize Rinse solution barcode | Analyzer doesn't recognize Rinse solution barcode. I believe it is a software key issue. Worked with Mark Z. And RMED to get new key for instrument. | Replaced software keys with new keys from RMED
No | Visit | Complaint | RDAN | Mette Grønbæk klager over tHb 0-punkts drift. | Afprøvet og undersøgt. Halogen pæres holdeskrue overspændt . | Holdeskrue løsnet, lampe justeret, kalibreret, slutafprøvet. | SWO-383831 | Adjustment | HW | Oxi module | Lamp unit |  | Mette Grønbæk complains about tHb 0-point operation. | Tested and investigated. Halogen peed Screw overstretched. | Keep the screw loose, lamp aligned, calibrated, final tested.
No | Visit | Complaint | RDEU | Glu. Instabil. Kalibrationdrift. | Glu und Lac. sehr instabil. Drift fehler bei jede 2 Kalibration | Elektroden membranierung und austausch der Kallösung bringt keine Veränderung. Neue Elektroden eingebaut. | SWO-383952 | Assisted Replacement | Electrode | Electrode | Glu | Glu+Lac | Glu. Unstable. Calibration drift. | Glu and Lac. very unstable. Drift error at every 2 calibrations | Electrode membraneation and replacement of the calendsolution does not bring any change. New electrodes installed.
Yes | Visit | Complaint | RIND | External printer error. | External printer error. | Found problem with LAN connection, coordinated with IT dept, they came and resolved the issue, now printer is working. | SWO-385849 | Re classified RMED | His/Lis |  |  |  | External printer error. | External printer error. | Found problem with LAN connection, coordinated with IT dept, they came and resolved the issue, now printer is working.
No | Visit | Complaint | RDEU | Fehler 1279: Gemessene Flüssigkeitsmenge war oberhalb der Warngrenze Gerät verlangt öfter das Reinigungsverfahren | AQT verlangt weiterhin sporadische Reinigung | Manifold-Unit erneuert. Systemüberprüfung und Qualitätskontrollen durchgeführt. Gerät überprüft. | SWO-384196 | Defect | HW | Wet section | Manifold | 1279 | Error 1279: Measured liquid quantity was above the warning limit Device requires the cleaning procedure more often | AQT continues to require sporadic cleaning | Manifold unit. system inspection and quality control. device checked.
No | Visit | Complaint | RAME | No sample detected during aspiration | No sample detected during aspiration. Errors 0331 most of the time and sometimes 1271. | Take apart analyzer for inspection. While apart, replaced parts from PM kit. Cleaned and lubricated pinch valve. Performed all fluidic tests in service program. Also turned off auto temp correction in misc setup for customer. | SWO-384528 | Defect | HW | Basic Unit | PM Kit | 331 | No sample detected during aspiration | No sample detected during aspiration. Errors 0331 most of the time and sometimes 1271. | Take apart analyzer for inspection. While apart, replaced parts from PM kit. Cleaned and lubricated pinch valve. Performed all fluidic tests in service program. Also turned off auto temp correction in misc setup for customer.
No | Visit | Complaint | RDEU | Gerät verliert nach einige Zeit die Pat-Profile | Gerät überprüft | Geräteeinstellungen der Quelle der Pat-Profile auf 365 Tage angepasst, tHb Kal durchgeführt | SWO-384360 | Adjustment | HW | Oxi module | tHb Cal |  | Device loses pat profiles after some time | Device checked | Device settings of the source of the Pat profiles adjusted to 365 days, tHb Kal performed
No | Visit | Complaint | RLTD | Connie emailed in to say that the fan on the analyser was making a lot of noise. | The noise was coming from the CPU fan and when I removed the cover to investigate I found 2 pieces of paper on top of the fan. | I removed the paper from the fan and now the noise has gone, the fan blades must have been catching on the paper and this was making the noise. | SWO-384301 | Cleaned | HW | DMS_PC | Fan, CPU |  | Connie emailed in to say that the fan on the analyser was making a lot of noise. | The noise was coming from the CPU fan and when I removed the cover to investigate I found 2 pieces of paper on top of the fan. | I removed the paper from the fan and now the noise has gone, the fan blades must have been catching on the paper and this was making the noise.
No | Visit | Complaint | RDEU | 17.7.2018/10:15 Hr. Dr. Sina: Einlaß des ABL steht offen und rattert bei Versuch zu schließen | Einlass bewegt sich nicht, Zahnräder stehen falsch, | Einl.system im Serviceprogramm neu kalibriert, Einl.modul wieder eingebaut, Gerät kalibriert, | SWO-384564 | Adjustment | HW | Auto Inlet module | Auto Inlet module | gear-wheels | 17.7.2018/10:15 Mr. Dr. Sina: Admission of the ABL is open and rattles at attempt to close | Inlet does not move, gears are wrong, | recalibrated in the service program, reinstalled single module, calibrated device,
No | Visit | Complaint | RLTD | Haval Ozgon Tel: 07737 457222 | All three Analysers with reported issues were working upon my arrival. The advice given was to carry out more regular cleaning of the inlet gaskets and ensure all staff are trained correctly and emphasize the importance of mixing samples. They are using Radiometer syringes as well as BD. | Ensure inlet gaskets are clean and samples are mixed correctly. | SWO-391073 | User error | preanalytical | preanalytical | preanalytical |  | Haval Ozgon Tel: 07737 457222 | All three Analysers with reported issues were working upon my arrival. The advice given was to carry out more regular cleaning of the inlet gaskets and ensure all staff are trained correctly and emphasize the importance of mixing samples. They are using Radiometer syringes as well as BD. | Ensure inlet gaskets are clean and samples are mixed correctly.
No | Visit | Complaint | RLTD | Haval Ozgon Tel: 07737 457222 (POC Mgr) | Checked data logs and service programs. Removed inlet and checked connector gasket. Printer is working ok, despite the reported issue. Analyser software is on 3.3 MR1 and will upgraded to V3.4 | Cleaned inlet gasket, replaced and tested. | SWO-385027 | Defect | SW | RMED SW | Auto Inlet module | 3.4 Upgrade | Haval Ozgon Tel: 07737 457222 (MGR POC) | Checked data logs and service programs. Removed inlet and checked connector gasket. Printer is working ok, despite the reported issue. Analyser software is on 3.3 MR1 and will upgraded to V3.4 | Cleaned inlet gasket, replaced and tested.
No | Visit | Complaint | RLTD | Haval Ozgon Tel: 07737 457222 | All three Analysers with reported issues were working upon my arrival. The advice given was to carry out more regular cleaning of the inlet gaskets and ensure all staff are trained correctly and emphasize the importance of mixing samples. They are using Radiometer syringes as well as BD. | Ensure inlet gaskets are clean and samples are mixed correctly. | SWO-391078 | Defect | SW | RMED SW | Auto Inlet module | 3.4 Upgrade | Haval Ozgon Tel: 07737 457222 | All three Analysers with reported issues were working upon my arrival. The advice given was to carry out more regular cleaning of the inlet gaskets and ensure all staff are trained correctly and emphasize the importance of mixing samples. They are using Radiometer syringes as well as BD. | Ensure inlet gaskets are clean and samples are mixed correctly.
No | Visit | Complaint | RAME | Battery not holding charge. After 30 secs once disconnect from the outlet analyzer turns off | The battery isn't holding a charge. | Replaced the battery. | SWO-384715 | [5 empty] | Battery not holding charge. After 30 secs once disconnect from the outlet analyzer turns off | The battery isn't holding a charge. | Replaced the battery.
No | Visit | Complaint | RAME | Battery not holding charge. After 30 secs once disconnect from the outlet analyzer turns off | Original SWO-384715 closed. Battery needs to be replaced. | Replaced the battery. | SWO-387263 | [5 empty] | Battery not holding charge. After 30 secs once disconnect from the outlet analyzer turns off | Original SWO-384715 closed. Battery needs to be replaced. | Replaced the battery.
No | Visit | Complaint | RKOR | Multiple parameter calibration error | . | Liquid flow line clean. SP SC replace. | SWO-386707 | [5 empty] | Multiple parameter calibration error | . | Liquid flow line clean. SP SC replace.
No | Visit | Complaint | RDEU | Benötigt viel Zeit für die Flüssigkeitssensorkalibrierung, häufige Meldung : mögliche Verstopfung der Fließwege | Hemolyzerleistung zu niedrig | Oft zu niedrige Oximetriewerte, Austausch Hemolyzer, Funktionskontrolle | SWO-384380 | [5 empty] | Requires a lot of time for liquid sensor calibration, frequent message: possible blockage of the flow paths | Hemolyzer performance too low | Often too low oximetry values, replacement hemolyzer, function control
Yes | Visit | Complaint | RCRO | CCN/PER | Error 1218 System Cycle Failed | Glu | SWO-387421 | [5 empty] | CCN/PER | Error 1218 System Cycle Failed | Glu
No | 0 | Y | RPAC | Blank | Client has advised this new tcpO2 sensor is failing calibrations and zero point testing with calibration gas. |  | D4A6662BAC209794 | [5 empty] | Blank | Client has advised this new tcpO2 sensor is failing calibrations and zero point testing with calibration gas.
No | Visit | Complaint | RCHN | 抽取样本失败 | 抽取样本失败，借903-129一个 | 更换903-129后修复by tim | SWO-383647 | Defect | HW | Oxi module | Pinch valve |  | Failed to extract samples | Failed to take samples, borrowing 903-129 one | Repair by tim after replacement 903-129
No | Visit | Complaint | RESP | Se observa fallo de estado en pO2 durante Cal. | Fallo en pO2 | Se comprueban registros, se observa que el cliente ha sustituido CE, sin solventarse el problema, se procede intenta sustituir SP debido a que le quedan pocos ciclos, pero este da error sin que el cliente disponga de ningún otro, por lo que no se puede solventar el error al quedarse sin material | SWO-385633 | Assisted Replacement | SP | SP | pO2 Cal |  | State failure is observed in pO2 during Cal. | Failure in pO2 | Records are checked, it is observed that the customer has replaced CE, without solving the problem, it is necessary to replace SP because it has few cycles left, but this fails without the client having any other, so that the error cannot be solved to the remainder without material
Yes | Visit | Complaint | RCHN | 光学单元超出范围，仪器无法开机 | 光学单元超出范围，仪器无法开机,需要借件910-343和903-364各一个 | 更换903-364一个正常。by tim | SWO-383435 | [5 empty] | The optical unit is out of range and the instrument does not power on | Optical unit out of range, the instrument can not be switched on, need to borrow 910-343 and 903-364 each one | Replace 903-364 one normal. by tim
Yes | Visit | Complaint | RCHN | 报错 1999 | 报错 1999,借636-541一个。 | 更换636-541一个正常。by tim | SWO-383641 | [5 empty] | Error 1999 | Error 1999, borrowed 636-541 one. | Replace 636-541 one normal. by tim
No | Visit | Complaint | RCHN | 做样本报1070，更换空气过滤网 | 做样本报1070 | 更换空气过滤网 | SWO-385675 | Defect | HW | Process Module | Cup dryer | 1070 | Do a sample of this newspaper 1070, replace the air filter | Do a sample of this newspaper 1070 | Replace the air filter
No | Visit | Complaint | RDAN | Vi har netop lige nu problemer med stor drift på begge vores OP-apparater, også glucose. | Afprøvet og undersøgt. | Luftventil og pumpeslange udskiftet, slutafprøvet. | SWO-384462 | Defect | HW | El/Met Module | Valve, air |  | We just now have problems with great operation on both of our UP-appliances, including glucose. | Tested and investigated. | Air valve and pump hose replaced, final tested.
No | Visit | Complaint | RKK | DMSが立ち上がらない | DMSが立ち上がらない HDDの認識が遅いときがある | HDD交換 | SWO-384786 | Defect | HW | DMS_PC | HDD |  | DMS does not stand up | DMS does not stand up Hdd recognition may be slow | HDD switch
No | Visit | Complaint | RSCH | Rep. Ref./Oxi | Immer wieder Ca und Na Drift QC pH und Elektrolyten schwankend Ref. Elektrode bereits 2x ersetzt Oxi-Modul in Thalwil getestet alles i.o Power Supply Kundengerät ersetzen Ref. Membrane Lot = R2255 | Power Supply ersetzt i.o => Garantie Messkammer pH $BG-Modul ersetzt i.o => Garantie Ref. Electrode connector ersetzt i.o => Garantie Ref. Elektrode ersetzt i.o Alte Lot = 9803 => Garantie Membranenwartung Cl, Na, K und Ca durchgeführt i.o O-Ring pH ersetzt Danach Ansprechfehler und Empf. 63% O-Ring erneut ersetzt, ohne Erfolg pH Elektrode ersetzt i.o Alte Lot = BD01 => Garantie Kal., QC und tHb i.o System i.o | SWO-384492 | Defect | HW | DMS_PC | PSUP |  | Rep. Ref./Oxi | Again and again Ca and Na Drift QC pH and electrolytes fluctuating Ref. Electrode already replaced 2x Oxi module in Thalwil tested everything i.o Power Supply Customer Device Replacement Ref. Membrane Lot = R2255 | Power Supply replaces i.o => Warranty Measuring chamber pH $BG module replaces i.o => Warranty Ref. Electrode connector replaces i.o => Warranty Ref. Electrode replaces i.o Old lot = 9803 => Warranty Membrane maintenance Cl, Na, K and Ca carried out i.o O-ring pH replaced After that, contact error and empf. 63% O-ring replaced again, to no avail pH electrode replaces i.o Old Lot = BD01 => Warranty Kal., QC and tHb i.o System i.o
No | Visit | Complaint | RAME | System in hold without explanation or cause | Reviewed all data logs and spoke in depth to Tammy . Cause not readily identified, suspect electrode window accidentally jostling during cleaning. Causing hold | Secured window and will check with customer in one week | SWO-385941 | User error | User error | User Error | jostling electrode window during cleaning |  | System in hold without explanation or cause | Reviewed all data logs and spoke in depth to Tammy . Cause not readily identified, suspect electrode window accidentally jostling during cleaning. Causing hold | Secured window and will check with customer in one week
Yes | Visit | Complaint | RDEU | 16.07.2018; Hr. Smyrra (Med. Technik); Tel.: 023031068221; Demokey läuft ab | Gerät überprüft | Option Keys eingespielt, Gerät neu gestartet, Startphase abgewartet, Proteinentfernung durchgeführt Es fehlt ein Elektrlytkey, neuen Demokey bestellt | SWO-384359 | Re classified RMED | Demokey expired |  |  |  | 16.07.2018; Mr. Smyrra (Med. Technik); Tel.: 023031068221; Demokey expires | Device checked | Option Keys played in, device restarted, start-up phase awaited, protein removal performed Missing an electriftkey, new demokey ordered
No | Visit | Complaint | RKK | crea キャリブレーションが取れない | crea ｷｬﾘﾌﾞﾚｰｼｮﾝが取れない | POWER SUPPLYﾎﾞｰﾄﾞ代替交換 | SWO-384159 | Assisted Replacement | Electrode | Electrode | Crea | Calibration | Crea calibration is not possible | Crea calibration is not possible | POWER SUPPLY Board Replacement
No | Visit | Complaint | RAME | 0618-waste detector errors are present and 65% full is displayed when an empty bottle is present. | 0618 error has returned. After the previous repair, the 0618 error was observed generated after several calibrations. | Replaced the waste detector and performed the waste check procedure in service mode. No errors are present during the waste check and the T_ambient voltage is 290, which is in the acceptable range of (100-500). Also, the slope for (T_low-T_high) is 253, which is in accordance with the figure described in the service manual. Ran multiple calibrations to allow normal filling of the waste bottle. No errors were generated. Analyzer is within acceptable operating specifications. | SWO-384624 | Defect | HW | Basic Unit | Waste detector | 618 | 0618-waste detector errors are present and 65% full is displayed when an empty bottle is present. | 0618 error has returned. After the previous repair, the 0618 error was observed generated after several calibrations. | Replaced the waste detector and performed the waste check procedure in service mode. No errors are present during the waste check and the T_ambient voltage is 290, which is in the acceptable range of (100-500). Also, the slope for (T_low-T_high) is 253, which is in accordance with the figure described in the service manual. Ran multiple calibrations to allow normal filling of the waste bottle. No errors were generated. Analyzer is within acceptable operating specifications.
No | Visit | Complaint | RDEU | Hardware Fehler bei Patientenmessungen bei der Oximetrie , Hemolyser Fequenz Error | System überprüft => Hemolyser defekt | Hemolyser erneuert, Kalibration und Qc durchgeführt | SWO-385380 | Defect | HW | Oxi module | Hemolyzer | 581 | Hardware error in patient measurements in oximetry , Hemolyser Fequenz Error | System checked => Hemolyser defective | Hemolyser renewed, calibration and Qc performed
No | Visit | Complaint | RKOR | Pump calibration error | Re-insert the solution pack Clean & Re-insert the inlet gasket with holder Perform the fluid translate check Replace the inlet gasket with holder $ Adjust the tubing valve | Replace the inlet gasket with holder $ Adjust the tubing valve | SWO-386690 | Defect | HW | Inlet module | Inlet gasket |  | Pump calibration error | Re-insert the solution pack Clean & Re-insert the inlet gasket with holder Perform the fluid translate check Replace the inlet gasket with holder / Adjust the tubing valve | Replace the inlet gasket with holder $ Adjust the tubing valve
No | Visit | Complaint | RDEU | 18.07.2018, Herr blenkert, tel. 020917253391: Ableitstrom im Analysator erkannt | Gerät überprüft | Neues Lösungspack eingesetzt, Startphase abgewartet, | SWO-385528 | Assisted Replacement | SP | SP | Leak Current | 963 | 18.07.2018, Mr. blenkert, tel. 020917253391: Discharge current detected in the analyzer | Device checked | New solution pack used, start-up phase awaited,
No | Visit | Complaint | RKOR | Leak current in relation to Solution pack detected | Re-insert the solution pack Perform the fluid translate check | Perform the fluid translate check | SWO-386680 | Adjustment | SP | SP | Leak Current | Re insert | Leak current in relation to Solution pack detected | Re-insert the solution pack Perform the fluid translate check | Perform the fluid translate check
Yes | Visit | Complaint | RDEU | Lösungspack entriegelt nicht | Lösungspack entriegelt nicht, lässt sich nur manuell auswerfen | LP versucht im Serviceprogramm zu entriegeln - nicht möglich kein mechanisches Geräusch von Actuator-Einheit Actuator Einheit ausgebaut, Spindel nicht an Messinggabel befestigt! Spindel befestigt und gefettet, Actuator Einheit eingebaut, SCi erneuert Gerät kalibriert und überprüft | SWO-385123 | [5 empty] | Solution pack does not unlock | Solution pack does not unlock, can only be ejected manually | LP tries to unlock in the service program - not possible No mechanical noise from Actuator unit Actuator unit removed, spindle not attached to brass fork! Spindle attached and greased, Actuator unit built-in, SCi renewed Device calibrated and checked
No | Visit | Complaint | RCHN | 打印机盖板无法打开 | 打印机盖板无法打开 借903-001一个 | 更换903-001后修复by tim | SWO-383480 | [5 empty] | Printer cover does not open | Printer cover can't be opened Borrowing 903-001 one | Repair by tim after replacement 903-001
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | K,NA,CA,CL定标失败 | 更换测试卡正常。 | SWO-385035 | [5 empty] | Test card, reagent package failure | K, NA, CA, CL calibration failed | The replacement test card is OK.
No | Visit | Complaint | RAME | device is displaying possible fluid blockage during system cycle. | Fluid blockage error 2106 and Spectrometer hardware failure. | Replaced manifold unit. Power analyzer back up and performed CO-OX service test and system cycle. Still receiving same errors. Replaced CO-OX board. Powered analyzer back up and performed CO-OX service test and system cycle. No more errors present. Ran tHb cal, calibration and qc. | SWO-385285 | [5 empty] | device is displaying possible fluid blockage during system cycle. | Fluid blockage error 2106 and Spectrometer hardware failure. | Replaced manifold unit. Power analyzer back up and performed CO-OX service test and system cycle. Still receiving same errors. Replaced CO-OX board. Powered analyzer back up and performed CO-OX service test and system cycle. No more errors present. Ran tHb cal, calibration and qc.
No | Visit | Complaint | RCHN | 系统检查失败 1031 | 之前借的903-364有质量问题， 1031,加借 903-364一个。 | 更换903-364后修复by tim | SWO-384343 | Defect | HW | Process Module | Optical Unit | 1031 | System check failed 1031 | The previous loan of 903-364 had quality problems, 1031, plus 903-364 one. | Repair by tim after replacing 903-364
No | Visit | Complaint | RCHN | 定标失败 | 定标失败，需要借件944-216（14616） | 重新定标正常(5784)。by tim | SWO-384921 | Assisted Replacement | Consumables | Test Kit | Cal Cartridge, CKMB |  | Bid failed | Failed bid, 944-216 (14616) required | Re-calibration normal (5784). by tim
No | Visit | Complaint | RCHN | 放试管夹卡槽断裂 | 放试管夹卡槽断裂，借903-584一个 | 更换了903-584后修复by tim | SWO-384152 | Defect | HW | Aspiration Module | Tube holder |  | The test tube clip card slot is broken | The test tube clip card slot is broken, borrowed 903-584 one | Replaced 903-584 after repair by tim
No | Visit | Complaint | RAME | Inlet leak | Leaking out of inlet caused by blockage | Cleared blockage from waste tubing. Tested ok | SWO-385373 | Cleaned | Fluidic pathway | Fluidic pathway | Waste tubes |  | Inlet leak | Leaking out of inlet caused by blockage | Cleared blockage from waste tubing. Tested ok
No | Visit | Complaint | RDEU | Sporadisch Driftalarm bei Ca und Na. War schon vor Wartung vorhanden. | Sporadisch Driftalarm bei Ca und Na. War schon vor Wartung vorhanden. Bei Wartung durchgeführte Maßnahmen brachten nur für einen Tag Besserung. | Referenzelektrode muß erneuert werden. Herrn Koch Referenzelektrode ausgehändigt, da ich noch im Aufwachraum die Jahreswartung noch nicht fertiggestellt hatte. Austausch der Elektrode übernimmt Herr Koch. Herr Christof Koch wird Kalibrationen und QC’s vornehmen. Defekte Elektrode mitgenommen. | SWO-385175 | Assisted Replacement | Electrode | Electrode | Reference | CAPA-00733 | Sporadic drift alarm at Ca and Na. Was already present before maintenance. | Sporadic drift alarm at Ca and Na. Was already present before maintenance. Maintenance measures only brought improvement for one day. | Reference electrode must be renewed. Mr. Koch handed over reference electrode, as I had not yet completed the annual maintenance in the wake-up room. Replacing the electrode is done by Mr. Koch. Mr Christof Koch will make calibrations and QC's. Defective electrode.
No | Visit | Complaint | RAME | Unit leaking underneath the electrode chamber. | Unit leaking from chamber. Found bad pH seal ring and chamber was extremely dirty. | Replaced measuring chamber and pH sealing ring. | SWO-384514 | Defect | HW | pH/BG Module | measuring chamber |  | Unit leaking underneath the electrode chamber. | Unit leaking from chamber. Found bad pH seal ring and chamber was extremely dirty. | Replaced measuring chamber and pH sealing ring.
No | Visit | Complaint | RIND | gasket leakage | gasket leaking | clott was found in the gasket cleaned the gasket and pump calibration done then calibrated the analyser to ready mode then observed for few samples , working fine | SWO-386275 | Cleaned | Fluidic pathway | Fluidic pathway | Inlet gasket | Clot | gasket leakage | gasket leaking | clott was found in the gasket cleaned the gasket and pump calibration done then calibrated the analyser to ready mode then observed for few samples , working fine
No | Visit | Complaint | RDEU | Glu. Instabil. Kalibrationdrift. | Glu und Lac. Instabil bei Kal und QC | Alle Elektroden in der EL $ Met- Modul neu bezogen. Alle Ventilen überprüft. Kalibration mehrfach duchgeführt. QC und Kal in Ordnung | SWO-385501 | Assisted Replacement | Electrode | Membrane | Glu | Glu+Lac | Glu. Unstable. Calibration drift. | Glu and Lac. Unstable in Kal and QC | All electrodes in the EL's Met module are re-sourced. Checked all valves. calibration several times. QC and Kal all right
Yes | 1 | Y | RPAC | Blank | Software upgrade V3.4 |  | D4A6AC62AC518002 | Re classified RMED | planned SWO |  | SW |  | Blank | Software upgrade V3.4
No | Visit | Complaint | RKOR | Leak error | . | SC, SP replace. Analyzer restart Leak detecting pass. | SWO-386713 | Revisit required | SP | SP | Leak detected |  | Leak error | . | SC, SP replace. Analyzer restart Leak detecting pass.
No | Visit | Complaint | RIND | Temperature not stable | Checked the machine found leaking from new sensor cassette & sensor board was defective | Need to replace sensor board | SWO-385574 | [5 empty] | Temperature not stable | Checked the machine found leaking from new sensor cassette & sensor board was defective | Need to replace sensor board
No | Visit | Complaint | RAME | Customer reported analyzer will not initialize cassette or pass system cycles. Glucose parameter shut off. | Customer reported multiple consumables replaced for glucose. Glucose failing level C8301 high out of range on system cycle. Sensor cassettes taking over 1 hour to initialize. | Zero all channels with zeroing cassette. Restart system cycles with glucose parameter on. Still failing. Replaced solution pack with newer expiration date. Restart system cycle. System Cycles failed again. Replaced sensor cassette. Restart initialization. Long initialization resulted in PCO2 and Glucose failures. Reinstalled original sensor cassette leaving Glucose disabled. | SWO-385366 | [5 empty] | Customer reported analyzer will not initialize cassette or pass system cycles. Glucose parameter shut off. | Customer reported multiple consumables replaced for glucose. Glucose failing level C8301 high out of range on system cycle. Sensor cassettes taking over 1 hour to initialize. | Zero all channels with zeroing cassette. Restart system cycles with glucose parameter on. Still failing. Replaced solution pack with newer expiration date. Restart system cycle. System Cycles failed again. Replaced sensor cassette. Restart initialization. Long initialization resulted in PCO2 and Glucose failures. Reinstalled original sensor cassette leaving Glucose disabled.
No | Visit | Complaint | RSCH | Blank | Divers erreurs Hct-tHb-PCO2 lors du remplacement 1XSP ensuite 1XSC | Remplacé le kit de maintenance e.o Nettoyage de la section fluidique e.o Obstruction système par sang e.o Cycle de qualité e.o | SWO-385746 | [5 empty] | Blank | Various Hct-tHb-PCO2 errors when replacing 1XSP after 1XSC | Replaced the e.o maintenance kit Cleaning the fluidsection e.o Blood system obstruction e.o E.o quality cycle
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | PO2定标不过 0210 | 更换测试卡后修复 | SWO-388113 | [5 empty] | Test card, reagent package failure | PO2 calibration is no more than 0210 | Repair after replacing test card
No | Visit | Complaint | RCHN | 血氧硬件故障 | 血氧硬件故障 借 903-197 902-980各一个 | 更换903-197和902-980后修复by tim | SWO-384808 | Defect | HW | Oxi module | Spectrometer |  | Blood oxygen hardware failure | Blood oxygen hardware failure Borrowing 903-197 902-980 each | Repair by tim after replacing 903-197 and 902-980
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 不均匀的QC2溶液，1541 | 更换试剂包后修复 | SWO-386643 | [5 empty] | Test card, reagent package failure | Uneven QC2 solution, 1541 | Repair after replacement of reagent kit
No | Visit | Complaint | RESP | No pueden cerrar la entrada de muestras, parece que hay alguna pieza partida | Chequeo de módulo inlet completo observando que han partido la inlet gasket dentro del módulo y la sonda esta doblada. | Sustitución de sonda e inlet gasket del stock del cliente | SWO-387571 | Defect | HW | Inlet module | Probe bend |  | They can't close the sample entry, it looks like there's some split piece | Full inlet module check noting that the inlet gasket has split into the module and the probe is bent. | Catheter and inlet gasket replacement of customer stock
No | Visit | Complaint | RCHN | 条码扫描器故障 | 报错3050,0002，借903-029,636-558,910-304各一个 | 更换903-029,636-558,910-304后修复by tim | SWO-384214 | Defect | HW | Storage Module | Barcode reader, storage wheel | 3050, 0002 | Barcode scanner failure | 3050,0002, borrow 903-029, 636-558, 910-304 each | Replace 903-029, 636-558, 910-304 after repair by tim
No | Visit | Complaint | RCHN | 进样针堵塞3412 | 进样针堵塞3412 | 疏通进样针正常。by tim | SWO-385998 | Cleaned | HW | Aspiration Module | Needle Assembly | 3412 | Sample needle clogging 3412 | Sample needle clogging 3412 | The dredging into the sample needle is normal. by tim
No | Visit | Complaint | RCHN | 报错0002 | 报错0002 借903-198一个 | 更换903-198一个正常。by tim | SWO-384848 | Defect | HW | Analyzer Control Module | PCB, Main Board | 0002 | Report error 0002 | Error 0002 Borrowing 903-198 one | Replace 903-198 one normal. by tim
No | Visit | Complaint | RIND | Sample abortted | check blockege | clean inlet gassket, waste tube, run 2 point calibration ,run a sample result is satisfactory. Given training how to procees sample ,how to check clotted sample | SWO-386570 | Cleaned | Fluidic pathway | Fluidic pathway | Inlet gasket | Clot | Sample abortted | check blockege | clean inlet gassket, waste tube, run 2 point calibration ,run a sample result is satisfactory. Given training how to procees sample ,how to check clotted sample
No | Visit | Complaint | RAME | Unusual fluid build-up in pH measuring chamber | Fluid present in pH electrode bay. | Cleaned electrode bay, replaced old, blue sealing ring with black new style. | SWO-385897 | Defect | Electrode | Electrode | Sealing ring |  | Unusual fluid build-up in pH measuring chamber | Fluid present in pH electrode bay. | Cleaned electrode bay, replaced old, blue sealing ring with black new style.
No | Visit | Complaint | RAME | Intermittent fluid in pH electrode bay | No fluid currently in pH electrode bay, but customer has seen it in the past. | Cleaned electrode bay, replaced old, blue sealing ring with black new style. | SWO-385895 | Defect | Electrode | Electrode | Sealing ring |  | Intermittent fluid in pH electrode bay | No fluid currently in pH electrode bay, but customer has seen it in the past. | Cleaned electrode bay, replaced old, blue sealing ring with black new style.
No | 3 | Y | RCZE | Blank |  |  | D4A68C9EAC167713 | Assisted Replacement | Electrode | Electrode | pCO2 |  | Blank
Yes | Visit | Complaint | RCHN | 新机开箱后发现少一个备件，电解质模块的塑料板没有 | 新机开箱后发现少一个备件，电解质模块的塑料板没有，借924-054一个 | Andrew会通过Techline向RMED申请发一个给经销商，仪器不影响使用，故直接关单处理by tim | SWO-384284 | Re classified RMED | OOB |  |  |  | One spare part was found after the new machine was opened, and the plastic plate of the electrolyte module was not | The new machine was found to be less than one spare part, the electrolyte module of the plastic plate did not, borrowed 924-054 a | Andrew will apply to RMED through Techline for a send to the dealer, the instrument does not affect the use, so directly close the bill processing by tim
Yes | 13 | Y | RPAC | Blank | Analyzer constantly asking to replace gasket and solution pack |  | D4A7ED35225841D6 | [5 empty] | Blank | Analyzer constantly asking to replace gasket and solution pack
No | Visit | Complaint | RAME | Customer states unit having Intermittent Inhomogeneous Blood Samples. | Upon arrival, found unit having Intermittent Inhomogeneous Samples only on Blood. | Replaced all above parts listed. Ran all service Oximetry Alignment tests. Ran all service Fluid Flow Tests. Unit now Failing Glucose and Lactate. Tried numerous Sensor cassettes. Will return with Different Lot Sensor Cassette to Replace. Returned 07 $21. Upon arrival, Unit currently working. Checked and Ran a blood . Unit functioning to Specifications. | SWO-386488 | Defect | HW | Oxi module | Pinch Valve |  | Customer states unit having Intermittent Inhomogeneous Blood Samples. | Upon arrival, found unit having Intermittent Inhomogeneous Samples only on Blood. | Replaced all above parts listed. Ran all service Oximetry Alignment tests. Ran all service Fluid Flow Tests. Unit now Failing Glucose and Lactate. Tried numerous Sensor cassettes. Will return with Different Lot Sensor Cassette to Replace. Returned 07 $21. Upon arrival, Unit currently working. Checked and Ran a blood . Unit functioning to Specifications.
No | Visit | Complaint | RDEU | 20.7.2018/11:45 Fr. Krüger (02611377364) meldet einen nicht bedienbaren Bildschirm 20.07.2018. 12.05, Frau Kröber, Tel: 02611377364: ABL fordert erneute den Wechsel der Einlassdichtung 20.07.2018 12:20 Fr. Grüber 026130405676: Interventionsmeldung ELD nicht quittierbar | Sichtkontrolle | Einlassnadel ersetzt, erfolglos versucht das Gerät zu spülen. Hemolyzer und SC sind dicht und unbrauchbar, Hemolyzer und SC ersetzt, Neustart, oxy test, thb Kalibration, Kalibration, QC- Kontrolle | SWO-387677 | Defect | HW | Oxi module | Hemolyzer | Clot | 20.7.2018/11:45 Ms. Krüger (02611377364) reports an unusable screen 20.07.2018. 12.05, Mrs. Kröber, Tel: 02611377364: ABL again calls for the change of inlet seal 20.07.2018 12:20 Fr. Grüber 026130405676: Intervention notification ELD not apprescable | Visual inspection | replaced inlet needle, unsuccessfully tries to flush the device. Hemolyzer and SC are dense and unusable, Hemolyzer and SC are replaced, restart, oxy test, thb calibration, calibration, QC control
No | Visit | Complaint | RLTD | Claire Bailey Tel: 01708 435000 |  |  | SWO-386323 | User error | User error | User Error | front casing cracked | pushed on the floor by patient with mental health issues | Claire Bailey Tel: 01708 435000
No | Visit | Complaint | RKOR | Leak error | . | SP replace ( Lot cheange) | SWO-386714 | Assisted Replacement | SP | SP | Leak detected |  | Leak error | . | SP replace ( Lot cheange)
No | Visit | Complaint | RAME | Customer reported analyzer will not initialize cassette or pass system cycles. Glucose parameter shut off. | Multiple consumable failures for Glu. | Consumable troubleshooting for Glu calibration failures. I was able to use a failing SC (Glu and PCO2) after 24 hours of hydration on another analyzer with passing results. The SC was flushed with saline solution and left to hydrate overnight. Due to analyzer accessibility and department work load, I am unable to troubleshoot the analyzer is question. Staff will at their convenience swap the SCs between the two analyzers and record results for future investigation. I will follow up with Cathy on Monday morning to discuss the next step required to resolve this issue. | SWO-386409 | [5 empty] | Customer reported analyzer will not initialize cassette or pass system cycles. Glucose parameter shut off. | Multiple consumable failures for Glu. | Consumable troubleshooting for Glu calibration failures. I was able to use a failing SC (Glu and PCO2) after 24 hours of hydration on another analyzer with passing results. The SC was flushed with saline solution and left to hydrate overnight. Due to analyzer accessibility and department work load, I am unable to troubleshoot the analyzer is question. Staff will at their convenience swap the SCs between the two analyzers and record results for future investigation. I will follow up with Cathy on Monday morning to discuss the next step required to resolve this issue.
No | Visit | Complaint | RIND | Solution pack disabled | Solution pack disabled | Reset solution pack. Solution pack yet not install, reset the analyzer ,install solution pack ,solution pack install sucessfully. | SWO-386573 | [5 empty] | Solution pack disabled | Solution pack disabled | Reset solution pack. Solution pack yet not install, reset the analyzer ,install solution pack ,solution pack install sucessfully.
Yes | Visit | Complaint | RCHN | 接口板腐蚀 | 接口板腐蚀，借902-972一个 | 更换902-972后修复by tim | SWO-385461 | [5 empty] | Interface plate corrosion | Interface plate corrosion, borrow 902-972 one | Repair by tim after replacement 902-972
No | Visit | Complaint | RCHN | 接口板腐蚀 | 接口板腐蚀，借902-972一个 | 更换902-972后修复by tim | SWO-385475 | [5 empty] | Interface plate corrosion | Interface plate corrosion, borrow 902-972 one | Repair by tim after replacement 902-972
Yes | Visit | Complaint | RCHN | 报错1999 | 报错1999 借636-549一个 | 更换636-549后修复by tim | SWO-384891 | [5 empty] | Error 1999 | Misquoted 1999 Borrowing 636-549 one | Repair by tim after replacing 636-549
No | Visit | Complaint | RCHN | 测试卡断裂 | 测试卡断裂 | 更换测试卡，清洁仪器后修复 | SWO-386636 | Defect | Consumables | Test Kit | Cartridge, Broken |  | Test card break | Test card break | Replace the test card and repair the instrument after cleaning
No | Visit | Complaint | RESP | No detecta los packs de soluciones |  |  | SWO-387572 | Defect | Consumables | Solution Pack | Chip, Solution Pack |  | Doesn't detect solution packs
No | Visit | Complaint | RCHN | 报错1068 | 报错1068 ,1070借848-037一个 | 更换848-037后修复by tim | SWO-384887 | Defect | HW | Process Module | Membrane Pump | 1068, 1070 | 1068 error | Error 1068 , 1070 borrowed 848-037 one | Repair by tim after replacement 848-037
Yes | Visit | Complaint | RDEU | Met. Hb Wert bei Ringversuch zu tief, sporadisch ungenügend Probe | Met. Hb Wert bei Ringversuch zu tief, sporadisch ungenügend Probe Oxi-Ventil schwergängig | Proteinentfernung durchgeführt Oxi Complete Test failed, Oxi-Ventil erneuert tHb kalibriert, Gerät kalibriert und überprüft | SWO-386169 | [5 empty] | Met. Hb value in ring test too deep, sporadic insufficient sample | Met. Hb value in ring test too deep, sporadic insufficient sample Oxi valve difficult to handle | Protein removal performed Oxi Complete Test failed, Oxi valve renewed tHb calibrated, device calibrated and checked
No | Visit | Complaint | RAME | System in hold without explanation or cause | Passed all service checks | Lock down electrode cover and will monitor | SWO-386726 | No fault found | HW | Basic Unit | Electrode cover |  | System in hold without explanation or cause | Passed all service checks | Lock down electrode cover and will monitor
Yes | Visit | Complaint | RKOR | Data transmission error | IP address check Re-insert the IP address $ Turn-off the firewall $ Save the system in CMD | Re-insert the IP address $ Turn-off the firewall $ Save the system in CMD | SWO-386693 | Re classified RMED | OOB | Setup IP address |  |  | Data transmission error | IP address check Re-insert the IP address / Turn-off the firewall / Save the system in CMD | Re-insert the IP address $ Turn-off the firewall $ Save the system in CMD
No | Visit | Complaint | RDEU | 12.07.2018 Sr. Anja 0211904361530 meldet, dass der Bidlschirm flackert und keine Messung möglich ist. | Testkassette blockiert das Storage Wheel. | Monitor Kabel überprüft und Befestigt. Storage Wheel demontiert, Testkassetten Förderung demontiert und gereinigt. Alle Lichtschranken gereinigt. Ballbearing Ring erneuert. | SWO-386964 | Defect | Consumables | Test Kit | Cartridge, Broken |  | 12.07.2018 Sr. Anja 0211904361530 reports that the bidl screen is flickering and no measurement is possible. | Test cartridge blocks the storage wheel. | Monitor cable checked and attached. Storage Wheel dismantled, test cartridges conveyed and cleaned. All light barriers cleaned. Ballbearing Ring renewed.
No | Visit | Complaint | RIND | Service Visit. | Critical count is 220, analyzer shows Error. | In activity log there is an error of calibrate 13, enter in to service program, check the incubato wheel and punch out section not found ok, open the front panel and back panel and check the punch out motor and incubator wheel, remove the dust from it and in punch out motor change the position of motor and calibrate in the service programe, successfully done, close all the parts same as before, analyzer handed over to staff in well working condition. | SWO-386471 | Cleaned | HW | Incubator Module | sensor, cup |  | Service Visit. | Critical count is 220, analyzer shows Error. | In activity log there is an error of calibrate 13, enter in to service program, check the incubato wheel and punch out section not found ok, open the front panel and back panel and check the punch out motor and incubator wheel, remove the dust from it and in punch out motor change the position of motor and calibrate in the service programe, successfully done, close all the parts same as before, analyzer handed over to staff in well working condition.
No | Visit | Complaint | RIND | Cartridge door locked | Checked cartridge door and also for solution pack. | Logon into service program and completed cartridge lock checking.Then installed solution pack. Unit is working ok | SWO-386217 | Adjustment | HW | Cover and Chassis | Cartridge door |  | Cartridge door locked | Checked cartridge door and also for solution pack. | Logon into service program and completed cartridge lock checking. Then installed solution pack. Unit is working ok
No | Visit | Complaint | RAME | Battery error 1560, intervention required, unable to proceed with sample testing. | Confirmed battery error. Performed applicable service diagnostics, charging rate stuck on 92%, in a constant charging mode, cell voltage for all batteries out of range. | Replaced battery module. Verified performance in service diagnostics. No issues observed or noted. Analyzer restarted with no issues. | SWO-386648 | Defect | HW | Battery module | Battery module | 1560 | Battery error 1560, intervention required, unable to proceed with sample testing. | Confirmed battery error. Performed applicable service diagnostics, charging rate stuck on 92%, in a constant charging mode, cell voltage for all batteries out of range. | Replaced battery module. Verified performance in service diagnostics. No issues observed or noted. Analyzer restarted with no issues.
No | Visit | Complaint | RAME | Notified by CAS on site that the battery is not charging as determined by screen icon. | Confirmed battery icon on main screen indicates 0% charge. Had been at 100% within the the last several days. Service diagnostics indicates the battery is charging, all cell voltages are within specifications. | Customer is in the middle of an extensive validation process. Rather than risk a possible failure causing the analyzer to require intervention I have disconnected the battery to prevent possible battery related failure. I will order a replacement and revisit customer when convenient. Disconnecting the battery will have no impact on analyzer operation in any manner. | SWO-386649 | Revisit required | HW | Battery module | Battery module | temporarily disconnected battery | Notified by CAS on site that the battery is not charging as determined by screen icon. | Confirmed battery icon on main screen indicates 0% charge. Had been at 100% within the the last several days. Service diagnostics indicates the battery is charging, all cell voltages are within specifications. | Customer is in the middle of an extensive validation process. Rather than risk a possible failure causing the analyzer to require intervention I have disconnected the battery to prevent possible battery related failure. I will order a replacement and revisit customer when convenient. Disconnecting the battery will have no impact on analyzer operation in any manner.
No | Visit | Complaint | RIND | Frequently instrument showing air detected and calibration fail. | found problem with sensor connector. sometimes its calibrating but it giving error. | presently installed the standby instrument. installed the SP and SC. performed samples found instrument working fine. | SWO-386662 | [5 empty] | Frequently instrument showing air detected and calibration fail. | found problem with sensor connector. sometimes its calibrating but it giving error. | presently installed the standby instrument. installed the SP and SC. performed samples found instrument working fine.
No | Visit | Complaint | RCHN | 电路板错误，ph,co2,o2定标不过 | 电路板错误，ph,co2,o2定标不过，借902-972和903-334各一个 |  | SWO-386545 | [5 empty] | Board error, ph, co2, o2 calibration but | Board error, ph, co2, o2 calibration but, borrow 902-972 and 903-334 each one
No | Visit | Complaint | RCHN | OXI光谱不匹配，浑浊度过高 | OXI光谱不匹配，浑浊度过高，借902-980一个 | 更换902-980后修复 by jesson | SWO-386185 | Defect | HW | Oxi module | Hemolyzer |  | OXI spectrum mismatch, turbidity over high | OXI spectrum mismatch, turbidity over high, borrow 902-980 one | Repair by jesson after replacement 902-980
No | Visit | Complaint | RCHN | 定标失败 1031 | 定标失败 1031,借903-364一个。 | 更换903-364后修复 by jesson | SWO-385692 | Defect | HW | Process Module | Optical Unit | 1031 | Bid Failure 1031 | Bid failure 1031, borrowing 903-364 one. | Repair by jesson after replacing 903-364
No | Visit | Complaint | RCHN | 接口板腐蚀 | 接口板腐蚀，借902-972一个 | 更换902-972后修复by tim | SWO-386517 | [5 empty] | Interface plate corrosion | Interface plate corrosion, borrow 902-972 one | Repair by tim after replacement 902-972
No | Visit | Complaint | RCHN | 打印机故障，po2结果打问号 | 打印机故障，po2结果打问号，借910-286和903-581各一个 |  | SWO-386661 | Defect | HW | Basic Unit | PCB, Sensor interface | pO2 | Printer failure, po2 results question mark | Printer failure, po2 results question mark, borrow 910-286 and 903-581 each one
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 测试卡参比电极错误 1361 | 更换测试卡后修复 | SWO-388112 | [5 empty] | Test card, reagent package failure | Test card reference electrode error 1361 | Repair after replacing test card
Yes | Visit | Complaint | RCHN | 报错1031 | 报错1031，借903-364一个 | 更换903-364后修复by tim | SWO-386587 | [5 empty] | 1031 | 1031, loan 903-364 one | Repair by tim after replacing 903-364
No | Visit | Complaint | RKK | 装置メンテナンス必要 | 装置メンテナンス必要 | 流体回路フラッシュ実施 インレット分解・調整 ポンプローター取り付け直し 各部調整・動作確認 CAL・QC確認 | SWO-387563 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | Equipment maintenance required | Equipment maintenance required | Fluid circuit flushing Inlet decomposition and adjustment Reseating the pump rotor Each part adjustment and operation confirmation CAL/QC Confirmation
No | Visit | Complaint | RDEU | 23.07.2018 Herr Soest 020172384285 meldet, das seit zwei Wochen ein Lac/Glu Fehler besteht. | Kein Fehler festzustellen, kal und QC erforderlich. | Gerät kalibriert und QC durchgeführt. Blutprobe wurde gemessen und OK. Blutprobe wurde mit einem Sarstedt Probennehmer gemessen, wo im Vorfeld der Konus der Spritze bei Eingabe nicht zurück gezogen wurde. Wodurch das verdrängte Blut über die Einlassdichtung gelaufen ist. | SWO-388069 | No fault found | No fault found | No fault found | No fault found | Glu/Lac QC | 23.07.2018 Mr. Soest 020172384285 reports that there has been a Lac/Glu error for two weeks. | No error to determine kal and QC required. | calibrated and QC performed. Blood sample was measured and OK. Blood sample was measured with a Sarstedt sampler, where in the run-up to the cone the syringe was not pulled back upon input. Which caused the displaced blood to run over the inlet seal.
No | Visit | Complaint | RDEU | 23.07.2018; Fr. Alilovioic; Tel.: 02093643545; Probenabbrüche, Einlass schließt nicht richtig | Proben werden teilw. abgebrochen, Proben werden nicht erkannt, | Nadel gereinigt und gerichtet, Einl.dichtung gereinigt, Probemessungen durchgeführt, Einlass kalibriert, | SWO-387413 | Revisit required | HW | Auto Inlet module | Auto Inlet module |  | 23.07.2018; Ms. Alilovioic; Tel.: 02093643545; Sample aborts, inlet does not close properly | Samples are taken in a part. aborted, samples are not recognized, | Needle cleaned and directed, sealing cleaned, sample measurements carried out, inlet calibrated,
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 温度错误 | 更换测试卡后修复 | SWO-386875 | [5 empty] | Test card, reagent package failure | Temperature error | Repair after replacing test card
No | Visit | Complaint | RAME | Customer reported analyzer will not initialize cassette or pass system cycles. Glucose parameter shut off. | Ongoing testing of sensor cassettes. Long initialization times. Glucose failing. | Continue test available sensor cassettes. Problems found with two that will not pass system cycles for glucose. Customer to request credits. | SWO-387324 | [5 empty] | Customer reported analyzer will not initialize cassette or pass system cycles. Glucose parameter shut off. | Ongoing testing of sensor cassettes. Long initialization times. Glucose failing. | Continue test available sensor cassettes. Problems found with two that will not pass system cycles for glucose. Customer to request credits.
No | Visit | Complaint | RAME | Customer reported analyzer will not initialize cassette or pass system cycles. Glucose parameter shut off. | Continuing Glu errors Consumable troubleshooting Glu is intermittently failing on the same SC. | Troubleshooting the SC errors found the failures follow the SC from analyzer to analyzer. Replaced the Sc with another passing SC from another analyzer for confirmation. I will continue to monitor the analyzer remotely to confirm proper operation. | SWO-387308 | [5 empty] | Customer reported analyzer will not initialize cassette or pass system cycles. Glucose parameter shut off. | Continuing Glu errors Consumable troubleshooting Glu is intermittently failing on the same SC. | Troubleshooting the SC errors found the failures follow the SC from analyzer to analyzer. Replaced the Sc with another passing SC from another analyzer for confirmation. I will continue to monitor the analyzer remotely to confirm proper operation.
No | Visit | Complaint | RIND | Analyser showing the V/c failed . | Verification check is not getting passed . Sensor cassette initialization is not passing and taking a lot of time . Showing air detected message intermittently . | performed zeroing calibration . Replaced the sensor board . 902-972 – Sensor board Cleaned the inner surfaces . Re installed sensor cassette and solution pack . | SWO-391548 | [5 empty] | Analyser showing the V/c failed . | Verification check is not getting passed . Sensor cassette initialization is not passing and taking a lot of time . Showing air detected message intermittently . | performed zeroing calibration . Replaced the sensor board . 902-972 – Sensor board Cleaned the inner surfaces . Re installed sensor cassette and solution pack .
No | Visit | Complaint | RIND | Service Visit. | Aanlyzer message shows open circuit in between the installation of sensor cassette. | Keep the SC in to 2-8 degree Celsius for 30 minutes, after 30 minutes flush the sensor cassette with the de-ionized water and installed it, successfully done analyzer handed over to staff in well working condition. | SWO-391352 | [5 empty] | Service Visit. | Aanlyzer message shows open circuit in between the installation of sensor cassette. | Keep the SC in to 2-8 degree Celsius for 30 minutes, after 30 minutes flush the sensor cassette with the de-ionized water and installed it, successfully done analyzer handed over to staff in well working condition.
No | Visit | Complaint | RAME | SP leaking | Customer reported the analyzer had a fluid blockage message, walked her through a rinse and reseating the SP. error still present. She then replaced the SP and the analyzer began to leak from the rear of the SP. | Removed wet section and inspected all tubings for blockage, flushed hemolyzer, no blockages were found. Removed and cleaned waste manifold. Cleaned analyzer thoroughly from leak. Checked waste port on the back of sp and massaged the valve a couple of times,Re-installed same SP and tested . Found analyzer is no longer leaking. Tested several times and found no leaking. Left customer a spare SP as she used her last one and tech support did not send the replacement I had ordered. | SWO-387334 | [5 empty] | SP leaking | Customer reported the analyzer had a fluid blockage message, walked her through a rinse and reseating the SP. error still present. She then replaced the SP and the analyzer began to leak from the rear of the SP. | Removed wet section and inspected all tubings for blockage, flushed hemolyzer, no blockages were found. Removed and cleaned waste manifold. Cleaned analyzer thoroughly from leak. Checked waste port on the back of sp and massaged the valve a couple of times,Re-installed same SP and tested . Found analyzer is no longer leaking. Tested several times and found no leaking. Left customer a spare SP as she used her last one and tech support did not send the replacement I had ordered.
No | Visit | Complaint | RCHN | TNI定标失败 | TNI定标失败 借944-212（14308）一张 | 重新定标后修复 by jesson | SWO-386788 | Cleaned | HW | Process Module | Process Wheel |  | TNI calibration failed | TNI bid failure borrowed 944-212 (14308) a | Fix by jesson after re-calibration
No | Visit | Complaint | RIND | frequently calibration error. | Found problem with sensor cassette. | Reinstalled the sensor cassette and solution pack but problem not rectified found problem with sensor cassette. installed the standby instrument. | SWO-391598 | [5 empty] | frequently calibration error. | Found problem with sensor cassette. | Reinstalled the sensor cassette and solution pack but problem not rectified found problem with sensor cassette. installed the standby instrument.
No | Visit | Complaint | RIND | Regularly error in QC for electrolytes. | Found Cl- and Na+ is out of range. | Check the machine clean all the chambers, run protein removal, cleaning and 2 point calibration but still the problem is same. Change Ref membrane and run 2 point calibration. After the calibration Na+ and Ca2+ status going high, check the ref connecter and run calibration but still the status of electrolytes high. Change new ref electrode and run 2 point calibration again, Now machine start working and Run QC successfully. | SWO-388101 | Assisted Replacement | Electrode | Electrode | Reference | CAPA-00733 | Regularly error in QC for electrolytes. | Found Cl- and Na+ is out of range. | Check the machine clean all the chambers, run protein removal, cleaning and 2 point calibration but still the problem is same. Change Ref membrane and run 2 point calibration. After the calibration Na+ and Ca2+ status going high, check the ref connecter and run calibration but still the status of electrolytes high. Change new ref electrode and run 2 point calibration again, Now machine start working and Run QC successfully.
No | Visit | Complaint | RAME | QC is being aborted | Autocheck module is aborting qc vials. | Checked probe, autocheck carousel alignment, and autocheck module parts to make sure nothing was preventing the carousel from rotating correctly. Cleaned around and under carousel. Qc samples are no longer aborting, Customer will continue to monitor. | SWO-387309 | Revisit required | HW | Autocheck Module | Autocheck Module | Cleaned | QC is being aborted | Autocheck module is aborting qc vials. | Checked probe, autocheck carousel alignment, and autocheck module parts to make sure nothing was preventing the carousel from rotating correctly. Cleaned around and under carousel. Qc samples are no longer aborting, Customer will continue to monitor.
No | Visit | Complaint | RIND | pH parameter query. | check the system. check the logs. check the calibration logs. | Replace the ref membrane. clean the ph electrode. replace the o-ring & outer ring. Clean the chamber & tubing. still the ph query need to cross check the ph electrode. | SWO-386808 | Revisit required | Electrode | Electrode | pH |  | pH parameter query. | check the system. check the logs. check the calibration logs. | Replace the ref membrane. clean the ph electrode. replace the o-ring & outer ring. Clean the chamber & tubing. still the ph query need to cross check the ph electrode.
No | Visit | Complaint | RDEU | 23.07.2018 Frau Ebert 062215634988 meldet, dass sich eine Testcass verklemmt hat. | Setsafe 22 (5784) | Analysator Überprüfung - Testkassette steckt mit abgebrochener Führungsnase unter Puncher Unit - Testkassette entfernt - Neue Testkassetten installiert - Funktionsüberprüfung etc. (Systemreinigung, 3 mal LQC und Testmessung) ohne Beanstandung | SWO-387126 | Defect | Consumables | Test Kit | Cartridge, Broken | 5784 | 23.07.2018 Ms. Ebert 062215634988 reports that a test cassestation has become stuck. | Setsafe 22 (5784) | Analyzer review - Test cartridge is stuck with aborted guide nose under Puncher Unit - Test cartridge removed - New test cassettes installed - function check etc. (system cleaning, 3 times LQC and test measurement) without complaint
No | Visit | Complaint | RKOR | Inhomogeneous rinse solution | RE-insert the solution pack Fiuld transport check | RE-insert the solution pack | SWO-391684 | Assisted Replacement | SP | SP | FT | Rinse | Inhomogeneous rinse solution | RE-insert the solution pack Fiuld transport check | RE-insert the solution pack
Yes | 1 | Y | RPAC | Blank | Software Upgrade. |  | D4AD134AFA3267EE | Re classified RMED | planned SWO |  | SW |  | Blank | Software Upgrade.
No | Visit | Complaint | RDEU | 23.07.2018; Fr. Alilovioic; Tel.: 02093643545; Probenabbrüche, Einlass schließt nicht richtig 23.07.2018 Fr. Alilovioic 02093643545: gleicher Fehler wie gestern ist erneut aufgetreten | Pat.messungen werden zeitweise abgebrochen, Proben werden nicht richtig angesaugt, Einl.system ist nicht o.k. | Eine kompl. Wartung durchgeführt, Ventil ausgebaut und gereinigt, Gerät kalibriert, Probemessungen durchgeführt, wieder Abbrüche, Einlass bestellt, | SWO-389830 | Defect | HW | Auto Inlet module | Auto Inlet module |  | 23.07.2018; Ms. Alilovioic; Tel.: 02093643545; Sample aborts, inlet does not close properly 23.07.2018 Ms. Alilovioic 02093643545: same error as yesterday has occurred again | Pat measurements are temporarily aborted, samples are not sucked in properly, Einl.system is not OK. | A complete Maintenance carried out, valve removed and cleaned, device calibrated, test measurements carried out, demolitions again, inlet ordered,
No | Visit | Complaint | RSAF | Customer doesn't want to go to machine to give details, wants FSE to come. | Error 1011: Missing blank cal | Calibration done and restarted analyzer. QC levels A,B and C done. | SWO-387846 | Adjustment | SW | RMED SW | off/on | 1011 | Customer doesn't want to go to machine to give details, wants FSE to come. | Error 1011: Missing blank cal | Calibration done and restarted analyzer. QC levels A,B and C done.
No | Visit | Complaint | RAME | FMET is not comparing to the ABL 820 | The instrument was compared to the 820 and the 90 at Kitimat. Patient comparisons were not correlating. | The hemolyser was replaced. | SWO-387936 | Defect | HW | Oxi module | Hemolyzer | MetHb | FMET is not comparing to the ABL 820 | The instrument was compared to the 820 and the 90 at Kitimat. Patient comparisons were not correlating. | The hemolyser was replaced.
No | Visit | Complaint | RFRA | Commande ( Aqure ) accès rapide ne fonctionne pas | Investigations effectuées par ingénieur d’application lors de l’installation du dispositif | Récupération de l’ensemble des données $ $ arrêt du dispositif $ $ ré-installation Windows 7 + soft en version 3.3MR1 ok $ $ redémarrage à froid $ $ paramétrage Windows ok $ $ réinstallation de l’ensemble des sauvegardes ok $ $ test de connexion ok : fonction accès rapide opérationnelle. | SWO-389917 | Setup | SW | RMED SW | AQURE |  | Order (Aqure) fast access does not work | Investigations carried out by application engineer during the installation of the device | Recovery of all the data $$ device shutdown $$$ re-installing Windows 7 - app in version 3.3MR1 ok $$$cold restart $$ Windows setting ok $$ reinstall of all backups ok $$ login test ok: rap access function operational ide.
No | Visit | Complaint | RIND | Analysis disable | Check Solution pack cycle. cycle was finished. need to replaced | Replaced solution pack. | SWO-391448 | [5 empty] | Analysis disable | Check Solution pack cycle. cycle was finished. need to replaced | Replaced solution pack.
No | Visit | Complaint | RCHN | 数据库损坏 | 数据库损坏 借933-441一个 | 更换933-441后修复 by jesson | SWO-386777 | [5 empty] | Database corruption | Database corruption, borrow 933-441 one. | Repair by jesson after replacement 933-441
No | Visit | Complaint | RCHN | 光学单元超出范围，报错1031 | 光学单元超出范围，报错1031，需要借件903-364和944-212（14251） | 更换903-364后重新定标修复 by jesson | SWO-386762 | Defect | HW | Process Module | Optical Unit | 1031 | Optical unit out of range, error 1031 | Optical unit out of range, error 1031, need to borrow 903-364 and 944-212 (14251) | Re-calibration repair by jesson after replacing 903-364
Yes | 12 | Y | RPAC | Blank | Bi Annual Service |  | D4AD025AFA4CC313 | Re classified RMED | planned SWO |  |  |  | Blank | Bi Annual Service
No | Visit | Complaint | RAME | No connection to Health Connect, unable to remote in. | Confirmed the network plug was active in the analyzer. Suspected a problem with the cable and or Modtap connection. I found the network cable not securely plugged into the correct modtap connector. I tested each available modtap connector. Determined only one was active, B, the blue connector. The cable was not completely plugged into an active modtap. | Securely plugged the network cable into the active modtap connector. Verified an active connection. Verified with Health Connect Support who tested access remotely. | SWO-390042 | User error | His/Lis | Connection | Cable not plugged |  | No connection to Health Connect, unable to remote in. | Confirmed the network plug was active in the analyzer. Suspected a problem with the cable and or Modtap connection. I found the network cable not securely plugged into the correct modtap connector. I tested each available modtap connector. Determined only one was active, B, the blue connector. The cable was not completely plugged into an active modtap. | Securely plugged the network cable into the active modtap connector. Verified an active connection. Verified with Health Connect Support who tested access remotely.
No | Visit | Complaint | RDEU | Kunde hat immer wieder Kal Fehler mit pH + Elyte | Ref Elektrode defekt. | Ref Elektrode erneuert. | SWO-388071 | Assisted Replacement | Electrode | Electrode | Reference | CAPA-00733 | Customer has again and again Kal error with pH + Elyte | Ref electrode defective. | Ref electrode renewed.
No | Visit | Complaint | RFRA | Dérive calibration | Contrôle des vannes, de l’alignement du FlexQ et de l’étanchéité des chambres de mesures | Nettoyage et lubrification de toutes les vannes. Essais | SWO-387617 | Cleaned | HW | ABL | Valves |  | Drift calibration | Control of valves, FlexQ alignment and sealing of measuring chambers | Cleaning and lubricating all valves. Tests
No | Visit | Complaint | RKK | サンプルホルダーの落下 | サンプルホルダーの交換を実施しました。 | サンプルホルダーの交換を実施しました。 | SWO-390316 | Defect | HW | Inlet module | Tube holder |  | Falling sample holder | The sample holder was replaced. | The sample holder was replaced.
No | Visit | Complaint | RLTD | Nicola Tel: 01582 497991 | The customer has had to replace the inlet due to a user having tried to reseat it themselves. POCT have tested and it is processing sample ok, but they are concerned that it may not be sitting in the correct position. | Removed inlet and checked internally. All mechanisms are ok. Re-installed all parts and fitted inlet to the Analyser. Tested and is functioning well. | SWO-391224 | No fault found | HW | Auto Inlet module | Auto Inlet module | 1391 | Nicola Tel: 01582 497991 | The customer has had to replace the inlet due to a user having tried to reseat it themselves. POCT have tested and it is processing sample ok, but they are concerned that it may not be sitting in the correct position. | Removed inlet and checked internally. All mechanisms are ok. Re-installed all parts and fitted inlet to the Analyser. Tested and is functioning well.
No | Visit | Complaint | RLTD | Sensor cassette slow start up for Glu | The customer has put three SC's in this Analyser of different LOT's, but they are still having calibration problems with Glu & Lac. | Cleaned the sensor by contact pins with alcohol. The SC still hasn't recovered or shown any improvement. Replaced the SC and is now functioning correctly. | SWO-389998 | Assisted Replacement | SC | SC | Glu/Lac |  | Sensor cassette slow start up for Glu | The customer has put three SC's in this Analyser of different LOT's, but they are still having calibration problems with Glu & Lac. | Cleaned the sensor by contact pins with alcohol. The SC still hasn't recovered or shown any improvement. Replaced the SC and is now functioning correctly.
No | Visit | Complaint | RLTD | Battery is showing 0% | Battery is displaying 0%. Checked battery cells in service programs and all cells are good. | Removed battery, removed the battery jumper to reset. Replaced the jumper and battery. Charge cycle has been reset and is now showing 100%. | SWO-389989 | Defect | HW | Battery module | Battery module | 0% | Battery is showing 0% | Battery is displaying 0%. Checked battery cells in service programs and all cells are good. | Removed battery, removed the battery jumper to reset. Replaced the jumper and battery. Charge cycle has been reset and is now showing 100%.
No | Visit | Complaint | RFRA | Le site d'introduction ne se ferme pas complètement sur le pack, il reste un peu en avant. Dans le menu service, monté du site en position seringue sans problème, position closed sans problème ; monté en position capillaire sans problème, appui sur closed, le site ne redescend pas, l'appareil demande à calibrer le site d'intro. Appuie sur calibrate, le site redescend mais pas complètement sur le pack. En mode normal, appuie sur Seringue, le site monte et se referme sans problème, appui sur capilaire le site monte mais ne descend pas complètement sur le pack de solution. Demander Audrey LAMBRE ; poste 49678 | Tests mouvements de l’inlet dans menu de service : en capillaire le retour en position « closed »n’est pas fluide. | Arrêt du dispositif $ $ démontage de l’ensemble « Motion mechanism » $ $ inspection de tous les éléments composant ce mécanisme ok $ $ inspection de le « curve wheel »ok $ $ remontage de l’ensemble ok $ $ remise en route du dispositif ( démarrage à froid ) $ $ tests dans menu technique du mouvement du module inlet dans toutes les positions ok | SWO-389919 | Adjustment | HW | Auto Inlet module | Inlet drive gear |  | The introductory site does not completely close on the pack, it stays a little ahead. In the service menu, mounted from the site in the syringe position without problem, position closed without problem; mounted in the hair position without problem, support on closed, the site does not go down, the device asks to calibrate the site of intro. Press on calibrate, the site goes down but not completely on the pack. In normal mode, press Seringue, the site rises and closes without problem, support on capilaire the site goes up but does not go down completely on the solution pack. Ask Audrey LAMBRE; 49678 | Inlet movement tests in service menu: in capillaries the return to the "closed" position is not fluid. | Stopping the device $$ dismantling of the whole "Motion mechanism" $$ inspection of all the components of this mechanism ok $$ inspection of the "curve wheel"ok $$ reassembly of the whole ok $$ restarted device (cold start)$$ tests da ns technical menu of inlet module movement in all positions ok
No | Visit | Complaint | RFRA | Commande ( Aqure ) accès rapide ne fonctionne pas | Investigations effectuées par ingénieur d’application lors de l’installation du dispositif | Récupération de l’ensemble des données $ $ arrêt du dispositif $ $ ré-installation Windows 7 + soft en version 3.3MR1 ok $ $ redémarrage à froid $ $ paramétrage Windows ok $ $ réinstallation de l’ensemble des sauvegardes ok $ $ test de connexion ok : fonction accès rapide opérationnelle. | SWO-389918 | Setup | SW | RMED SW | AQURE |  | Order (Aqure) fast access does not work | Investigations carried out by application engineer during the installation of the device | Recovery of all the data $$ device shutdown $$$ re-installing Windows 7 - app in version 3.3MR1 ok $$$cold restart $$ Windows setting ok $$ reinstall of all backups ok $$ login test ok: rap access function operational ide.
Yes | 1 | Y | RPAC | Blank | Software Upgrade the analyser to Version V3.4, FAN915-384. |  | D4AE227D70557BF9 | Re classified RMED | planned SWO |  | SW |  | Blank | Software Upgrade the analyser to Version V3.4, FAN915-384.
No | Visit | Complaint | RSAF | PO2 reading low on level B, Sol pack and SC replaced | Sol pack already replaced by customer, Calibrations ok, service PO2 test ok, possible environment issue. | Replaced SC, Calibration ok, PO2 still reading out of range. Data logs taken . | SWO-390093 | Revisit required | SP | SP | pO2 QC |  | PO2 reading low on level B, Sol pack and SC replaced | Sol pack already replaced by customer, Calibrations ok, service PO2 test ok, possible environment issue. | Replaced SC, Calibration ok, PO2 still reading out of range. Data logs taken .
No | Visit | Complaint | RAME | Customer reported analyzer will not initialize cassette or pass system cycles. Glucose parameter shut off. | Sensor cassette that was replaced yesterday failed system cycle early in the morning. Contacted Area Specialist to discuss plans to continue troubleshooting ongoing glucose failures. | Replaced sensor connector board and waste tubing set. Zero channels with set zero cassette. Restart system cycles. Download data files if needed for technical support CQC. Replaced sensor cassette. Solution pack running low, replaced pack. Sensor cassette slow to initialize. Ran tHB calibration with pass results. Restart system cycles. | SWO-387881 | [5 empty] | Customer reported analyzer will not initialize cassette or pass system cycles. Glucose parameter shut off. | Sensor cassette that was replaced yesterday failed system cycle early in the morning. Contacted Area Specialist to discuss plans to continue troubleshooting ongoing glucose failures. | Replaced sensor connector board and waste tubing set. Zero channels with set zero cassette. Restart system cycles. Download data files if needed for technical support CQC. Replaced sensor cassette. Solution pack running low, replaced pack. Sensor cassette slow to initialize. Ran tHB calibration with pass results. Restart system cycles.
No | Visit | Complaint | RCHN | 零点卡校正时氧的零点值不在范围内 | 零点卡校正时氧的零点值不在范围内 借902-972一个 | 更换902-972后修复 by jesson' | SWO-386770 | [5 empty] | Zero value of oxygen is out of range when zero card correction | Zero-point value of oxygen is not in range when zero card correction borrows 902-972 one | Repair by after replacement 902-972 Jesson'
No | Visit | Complaint | RCHN | 开机白屏 | 开机白屏 | 给医院送用备用机。此仪器送到维修间送修。 | SWO-390999 | [5 empty] | Power-on white screen | Power-on white screen | Send a spare machine to the hospital. This instrument is sent to the repair room for repair.
Yes | Visit | Complaint | RESP | fuga detectada | Fuga detectada. | Se realizo kit de tubos anual. Se sustituyo SP defectuoso. Se calibró equipo y se comprobó correcto funcionamiento mediante QC´s. Equipo operativo. | SWO-390530 | [5 empty] | leak detected | Leak detected. | Annual pipe kit was made. Defective SP is replaced. Equipment was graded and proper operation was tested by QC's. Operating Equipment.
No | Visit | Complaint | RCHN | 报错4025 | 报错4025，借944-394（14170）一个 | 重新定标后修复 by jesson | SWO-386924 | Assisted Replacement | Consumables | Test Kit | Cartridge, pct | 4025 | 4025 error | 4025, loan 944-394 (14170) one | Fix by jesson after re-calibration
No | Visit | Complaint | RCHN | 定标失败 | 定标失败，借905-908和903-364各一个 | 更换905-908和903-364后修复 by jesson | SWO-387406 | Defect | HW | Process Module | Optical Unit |  | Bid failed | Bid failure, borrow 905-908 and 903-364 each | Repair by jesson after replacing 905-908 and 903-364
No | Visit | Complaint | RCHN | 报错1035 | 报错1035 借903-364一个 | 需加借定标卡，此单关单 | SWO-388038 | Revisit required | HW | Process Module | Optical Unit | 1035 | 1035 error | False 1035 Borrow 903-364 one | A bid card needs to be borrowed, this single
No | Visit | Complaint | RESP | Fallo de iones, Lac y Crea. | Según indica Ma José Díaz fallan los iones, Lac y Crea. | Se le indica que cambíe la membrana de Ref y deje estabilizar el equipo. Pese a ello sigue persistiendo el problema, por lo que se realiza cambio de la membranas de los iones y Crea. Según indicaciones de la Sra. Díaz sigue sin solucionarse, así que tras descartar posibilidades se localiza el problema en el reactivo, siendo el mismo sustituido y subsanada la incidencia. Limpieza espèjo AutoQC y rellenado del mismo. Verificado con calibración y QC. | SWO-388216 | Assisted Replacement | Consumables | Solutions | Cal |  | Ion failure, Lac and Crea. | According to Ma José Díaz, the ions, Lac and Crea, fail. | You are instructed to change ref's membrane and allow the equipment to stabilize. However, the problem persists, so the ion membranes and Crea is changed. According to Mrs Díaz's instructions, the problem is still unresolved, so after ruling out the problem in the reagent, the same being replaced and remedied. AutoQC spjo cleaning and filling it. Certified with calibration and QC.
No | Visit | Complaint | RKK | ミキサーエラー頻発 | ミキサーエラー頻発 | FLEXQモジュール引き上げ、代替えモジュール設置。Glu、Lac、Ref、Cl、Ca、Naメンブラン定期交換。タンパク除去実施。エアーライン洗浄、PH $BGモジュールのフラットメンブランバルブ分解整備 | SWO-390234 | Defect | HW | Basic Unit | FlexQ Module |  | Mixer errors frequently | Mixer errors frequently | FlexQ module pull-up and replacement module installed. Glu, Lac, Ref, Cl, Ca, Namenbran regular replacement. Protein removal was carried out. Flat membrane valve decomposition and maintenance of air line cleaning and PH $BG modules
No | Visit | Complaint | RSCH | Rep. Ref. | Immer wieder Ca Drift | Ref. Elektrode ersetzt i.o Material von MedTech Alte Lot = 6808 => Garantie Kal. und QC i.o System i.o | SWO-390355 | Assisted Replacement | Electrode | Electrode | Reference | mail: what is material from MedTech + suggest to look at measurment chamber (defects?) | Rep. Ref. | Again and again Ca Drift | Ref. Electrode replaces i.o Material from MedTech Old lot = 6808 => Warranty Kal. and QC i.o System i.o
No | Visit | Complaint | RAME | QC is being aborted | Analyzer is failing Autocheck carousel | Replaced FPC Cable. Replaced autocheck module PCB. Updated Autocheck software to be compatible with analyzer software. | SWO-390109 | Defect | HW | Autocheck Module | FPC Cable |  | QC is being aborted | Analyzer is failing Autocheck carousel | Replaced FPC Cable. Replaced autocheck module PCB. Updated Autocheck software to be compatible with analyzer software.
No | Visit | Complaint | RIND | Leaking from inlet gasket | Checked the machine found clot inlet gasket and waste tube | Cleaned inlet gasket and waste tube | SWO-391403 | Cleaned | HW | Inlet module | Inlet gasket |  | Leaking from inlet gasket | Checked the machine found clot inlet gasket and waste tube | Cleaned inlet gasket and waste tube
No | Visit | Complaint | RESP | Pierde líquido por debajo | Chequeo de neumatica observando que el tuno número 6 se desconecta | Limpieza de circuito neumatico y desobstrucción | SWO-390453 | Cleaned | HW | Wet section | Wet section | tube 6 | Loses liquid below | Tire check noting that tuno number 6 is disconnected | Tire circuit cleaning and obstruction
No | Visit | Complaint | RDEU | 06.02.2018 Pfl. Ismael 09671930672 meldet, dass der AQT täglich eine Systemreinnigung aufruft. 08.02.2018 9:15 H. Zinkel 09671930244: weiterhin fordert das Gerät zu Systemreinigungen auf | CRP LQC Fehler Kunden AQT aus TB zurück zu Kunden verbracht und in Betrieb genommen | AQT90 aufgestellt, angeschaltet, gewartet bis der AQT Betriebstemperatur hat, Systemüberprüfung gestartet, Betriebsmittel aufgefüllt, LQC für jeden Parameter durchgeführt, IT Anbindung mittels IT-Abteilung der Asklepios aus Bad Abbach durchgeführt, LQC eingelesen welche noch nicht im AQT registriert waren, Kunden informiert, dass sie LQC für den Parameter CRP bestellen müssen - ist nichts mehr vor Ort, Patientenmessung abgewartet und überprüft ob Übertragung in Patientenakte funktioniert, System- und Funktionskontrolle in Ordnung | SWO-390317 | Cleaned | HW | Wet section | Wet section | Flow restriction | 06.02.2018 Pfl. Ismael 09671930672 reports that the AQT calls a system clean-up every day. 08.02.2018 9:15 H. Zinkel 09671930244: the device continues to request system cleaning | CRP LQC Error Customers AQT from TB moved back to customers and put into operation | AQT90 set up, switched on, waited until the AQT has operating temperature, system check started, equipment replenished, LQC performed for each parameter, IT connection carried out via IT department of Asklepios from Bad Abbach, LQC read in which not yet registered in The AQT, customers informed that they need to order LQC for the CRP parameter - nothing is on site anymore, patient measurement awaits and checks whether transmission in patient file works, system and function control is in order
No | 2 | Y | RPAC | Blank | Preasure test vacuum error |  | D4B0CE685C0366DA | Assisted Replacement | SP | SP | FT |  | Blank | Preasure test vacuum error
No | Visit | Complaint | RFRA | Suite a remplacement de cassette 100 et 300 test problème de calibration Lactate ( erreur Sensibilité) | Test de la fluidique OK Ajustement pompe OK Pressure Test OK Ajustement Manifold OK Clot $détection OK Simulator Test OK | inversion cassette de capteur entre ABL 1 laboratoire et ABL90 SAUV . Erreur de calibration du à cassette de capteur défectueuse (N° série: 1288-138) Installation cassette de capteur provenant ABL 1 Laboratoire OK sur ABL90 SAUV | SWO-390734 | Assisted Replacement | SC | SC | Lac |  | Following a replacement cassette 100 and 300 test calibration problem Lactate (error Sensitivity) | Fluidic testing OK OK pump fit Pressure Test OK Adjustment Manifold OK Clot /detection OK Simulator Test OK | sensor cassette reversal between ABL 1 laboratory and ABL90 SAUV . Faulty sensor cassette calibration error (Serial Number: 1288-138) Installation sensor cassette from ABL 1 Lab OK on ABL90 SAUV
No | Visit | Complaint | RAME | SP going empty prematurely, today says 39 events eft in pack and 30 day expiry is at aug 5 Contact is Dave at 507-537-9179 | Called to confirm issue with customer. The last two solution packs seemed to run out over a week before the expiration date $typical time frame. | I called the customer and was informed that he had found out they were running extra competency samples on this past solution pack. I explained that running a higher than normal sample count will drain the solution pack more quickly, also if the device had to run extra calibrations over the month to correct issues it can contribute to this issue as well. Customer will monitor the new solution pack that is going to be used and report back if the issue persists. | SWO-390847 | [5 empty] | SP going empty prematurely, today says 39 events eft in pack and 30 day expiry is at aug 5 Contact is Dave at 507-537-9179 | Called to confirm issue with customer. The last two solution packs seemed to run out over a week before the expiration date /typical time frame. | I called the customer and was informed that he had found out they were running extra competency samples on this past solution pack. I explained that running a higher than normal sample count will drain the solution pack more quickly, also if the device had to run extra calibrations over the month to correct issues it can contribute to this issue as well. Customer will monitor the new solution pack that is going to be used and report back if the issue persists.
No | Visit | Complaint | RCHN | FTC-105 定标卡无法扫描 | FTC-105 定标卡无法扫描 借903-329 902-924各一个 | 更换903-329 902-924后修复 by jesson | SWO-388097 | Defect | HW | Aspiration Module | Needle Wash Unit |  | FTC-105 Calibration Card Cannot Be Scanned | FTC-105 Bid Card Cannot Scan Borrowing 903-329 902-924 One | Repair by jesson after replacing 903-329 902-924
No | Visit | Complaint | RCHN | 处理包无法识别 | 处理包无法识别，借903-036一个 | 重新调整了处理包芯片位置后修复 by jesson | SWO-388046 | Defect | HW | Cover and Chassis | PCB, SP |  | Processing package is not recognized | Processing package is not recognized, borrow 903-036 one | Fix by jesson after re-locating the processing package chip
Yes | Visit | Complaint | RLTD | Blank | The demo keys are due to expire on 27 $07 $18. | New demo keys installed and these expire 28 $10 $18. | SWO-390512 | Re classified RMED | Demokey expired |  | Demo key |  | Blank | The demo keys are due to expire on 27 /07 /18. | New demo keys installed and these expire 28 $10 $18.
Yes | 5 | Y | RCZE | Blank |  |  | D4AF6706E61B1F2D | Re classified RMED | planned SWO |  | SW |  | Blank
No | Visit | Complaint | RDEU | 26.07.18 09:45 Herr Thomas Küchelbacher 06213832674 wünscht aufgrund personelle Engpass eine Gerätewartung und Fehlerbeseitigung durch FST (kostenpflichtig) | Einlass tropft tHB KAL fällig (Keine automatische QC) | Abfallstrecke gespült - Proteinentfernung - allgemeine Überprüfung - Glucose in QC instabil - Dichtring Glucose Elektrode fehlt (erneuert und Membranierung) - Elektrodenkontakt gereinigt - Messkammer Glucose mit Proteinentferner gereinigt - Kalibration durchgeführt, tHb KAL und QC ! | SWO-390635 | Defect | Electrode | Electrode | Sealing ring | Sealing ring, Glu | 26.07.18 09:45 Mr. Thomas Küchelbacher 06213832674 wants equipment maintenance and troubleshooting by FST due to staff shortages | Inlet drips tHB KAL due (No automatic QC) | Waste section rinsed - protein removal - general check - glucose in QC unstable - sealing ring glucose electrode is missing (renewed and membraneization) - electrode contact cleaned - measuring chamber glucose cleaned with protein remover - calibration performed, tHb KAL and QC !
No | Visit | Complaint | RIND | PO2 Level 1 QC was getting failed | Checked the data logs po2 Level 1 QC was getting failed | Perfomed decomtinamtion cleaned all the chambers performed qc po2 value was within the range | SWO-391435 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway | pO2 QC | PO2 Level 1 QC was getting failed | Checked the data logs po2 Level 1 QC was getting failed | Perfomed decomtinamtion cleaned all the chambers performed qc po2 value was within the range
No | Visit | Complaint | RIND | Parameter could not be measured | Checked the data logs tHB & so2 query for all the sample and parameter could not be measured | Performed complete oximetry test and 2 point calibration done processed the sample was ok | SWO-391414 | Adjustment | HW | Oxi module | Oxi module | Revisit | Parameter could not be measured | Checked the data logs tHB & so2 query for all the sample and parameter could not be measured | Performed complete oximetry test and 2 point calibration done processed the sample was ok
Yes | Visit | Complaint | RDEU | 19.07.2018 Frau Wilhelm0832396390 meldet, dass der Bildschirm dunkel ist. | Bildschirm dunkel | Cable CPU unit $Display erneuert, LQC vorgenommen. | SWO-390605 | [5 empty] | 19.07.2018 Ms Wilhelm0832396390 reports that the screen is dark. | Screen dark | Cable CPU unit $Display renewed, LQC made.
No | Visit | Complaint | RKK | ショートサンプルエラーが多発する。 | ショートサンプルエラーが多発する | ニードルウォッシュユニット交換を実施。 | SWO-390312 | Defect | HW | Aspiration Module | Needle Wash Unit |  | Short sample errors occur frequently. | Frequent short sample errors | Needle wash unit replacement was carried out.
No | Visit | Complaint | RKOR | Vacuume error | . | SC, SP replace, Gasket hold, connector replace, Liquid flow line clean. Valve tube position adjust. Analyzer restart. | SWO-391665 | Defect | HW | Basic Unit | PM Kit |  | Vacuume error | . | SC, SP replace, Gasket hold, connector replace, Liquid flow line clean. Valve tube position adjust. Analyzer restart.
No | Visit | Complaint | RAME | Glu Response Error | Looked in data logs, found glucose and lactate response error. | Cleaned sensor contacts, also replaced sensor cassette, also while I had it apart I did Annual maintenance. Wrote annual maintenance on separate service work order. | SWO-391517 | Cleaned | HW | Basic Unit | Sensor contact | Glu/Lac | Glu Response Error | Looked in data logs, found glucose and lactate response error. | Cleaned sensor contacts, also replaced sensor cassette, also while I had it apart I did Annual maintenance. Wrote annual maintenance on separate service work order.
No | Visit | Complaint | RFRA | Lors du remplacement de la cassette de capteur relance de l'analyseur avec au redémarrage erreur code 2 puis écran noir | Relance analyseur | suppression de la base de données restauration base de données vierge redémarrage analyseur Restauration base de données du 25 $07 $2018 relance analyseur : erreur calibration pCO2 (sensibilité hors gamme basse ) installation nouvelle cassette de capteur transmission ticket pour procédure de remboursement ° cassette de capteur défectueux via Navette interne Essai ok | SWO-391177 | Defect | SW | RMED SW | Database | Error code 2 | When replacing the sensor cassette restarts the analyzer with reboot error code 2 then black screen | Relaunch analyzer | deletion of the database restoration blank database reboot analyzer Restoration database of $25 $07 $2018 relaunch analyzer: pCO2 calibration error (low-range sensitivity) installation new sensor cassette ticket transmission for refund procedure - defective sensor cassette via Internal Shuttle Try ok
No | Visit | Complaint | RDEU | 27.07.2018; Fr. Anici; Tel.: 08912183306; Oxy Messfehler 27.07.2018; 11:00 Uhr; Fr. Anici; Tel.: 08912183306; weiterhin Oxy Messfehlerin | sporadisch Oxi Spectrum Mismatch (Fehler 0581 + 0292) bei Blutmessung | Gerät manuell gespült Hemolyzer sowie Oxi Ventilschlauch erneuert tHb kalibriert, Gerät kalibriert und überprüft Blutmessungen durchgeführt - ok | SWO-391211 | Defect | HW | Oxi module | Hemolyzer | 581 | 27.07.2018; Ms. Anici; Tel.: 08912183306; Oxy measurement error 27.07.2018; 11 a.m.; Ms. Anici; Tel.: 08912183306; further Oxy measurement error | sporadically Oxi Spectrum Mismatch (error 0581 + 0292) for blood measurement | Device manually rinsed Hemolyzer and Oxi valve hose renewed tHb calibrated, device calibrated and checked Blood measurements performed - ok
No | Visit | Complaint | RAME | Customer reported analyzer will not initialize cassette or pass system cycles. Glucose parameter shut off. | Consumable troubleshooting for Glu errors. | Assisted primary engineer with continuing Glu errors on analyzer. Delivered analog board to Mike for next level service if Glu continues to fail. We replaced the SC with one off another analyzer that has been stable for two days now. I will continue to monitor the analyzer remotely. | SWO-391487 | [5 empty] | Customer reported analyzer will not initialize cassette or pass system cycles. Glucose parameter shut off. | Consumable troubleshooting for Glu errors. | Assisted primary engineer with continuing Glu errors on analyzer. Delivered analog board to Mike for next level service if Glu continues to fail. We replaced the SC with one off another analyzer that has been stable for two days now. I will continue to monitor the analyzer remotely.
No | Visit | Complaint | RAME | Customer reported analyzer will not initialize cassette or pass system cycles. Glucose parameter shut off. | Ongoing glucose failures on system cycles. Checked data logs for errors. | Test several sensor cassettes. Was able to get glucose parameter to pass with latest sensor cassette replacement. Continue to monitor analyzer performance. | SWO-391504 | [5 empty] | Customer reported analyzer will not initialize cassette or pass system cycles. Glucose parameter shut off. | Ongoing glucose failures on system cycles. Checked data logs for errors. | Test several sensor cassettes. Was able to get glucose parameter to pass with latest sensor cassette replacement. Continue to monitor analyzer performance.
No | Visit | Complaint | RCHN | PH,CL定标不过 | PH,CL定标不过 借902-972一个 | 更换902-972备件后修复 jesson | SWO-390438 | [5 empty] | PH, CL calibration but | PH, CL calibration but borrow 902-972 one | Repair jesson after replacement of 902-972 spare parts
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | pco2 cl定标失败 | 更换测试卡后修复 | SWO-391550 | [5 empty] | Test card, reagent package failure | pco2 cl bid failure | Repair after replacing test card
No | Visit | Complaint | RIND | Service Visit. | Need to replace the sensor board and do the preventive maintenance. | Shutdown the analyzer, open the analyzer front panel as well as back panel, remove the dust from the various part from the analyzer, replace the sensor board, remove the manifold and flush it with hypochlorite and reconnect, connect all the part same as before, start the analyzer, install the solution pack and sensor cassette, V $C failed, do the zeroing, successfully done, install the sensor cassette successfully done, analyzer handed over to staff in well working condition. | SWO-391271 | [5 empty] | Service Visit. | Need to replace the sensor board and do the preventive maintenance. | Shutdown the analyzer, open the analyzer front panel as well as back panel, remove the dust from the various part from the analyzer, replace the sensor board, remove the manifold and flush it with hypochlorite and reconnect, connect all the part same as before, start the analyzer, install the solution pack and sensor cassette, V $C failed, do the zeroing, successfully done, install the sensor cassette successfully done, analyzer handed over to staff in well working condition.
Yes | Visit | Complaint | RIND | Preventive maintenance visit. | Need to do preventive maintenance. | Remove the sensor cassette and solution pack from the analyzer, shut down the analyzer, open the back panel and front panel, remove the dust from the various part of the analyzer, check the firm connection found ok. connect all the parts same as before, close the panel and start the analyzer and connect the sensor cassette and solution pack, successfully done, analyzer handed over to staff in well working condition. | SWO-391283 | Re classified RMED | planned SWO |  |  |  | Preventive maintenance visit. | Need to do preventive maintenance. | Remove the sensor cassette and solution pack from the analyzer, shut down the analyzer, open the back panel and front panel, remove the dust from the various part of the analyzer, check the firm connection found ok. connect all the parts same as before, close the panel and start the analyzer and connect the sensor cassette and solution pack, successfully done, analyzer handed over to staff in well working condition.
No | Visit | Complaint | RCHN | 接口板腐蚀 | 接口板腐蚀，借902-972一个 | 更换902-972后修复 by jesson | SWO-390221 | [5 empty] | Interface plate corrosion | Interface plate corrosion, borrow 902-972 one | Repair by jesson after replacement 902-972
No | Visit | Complaint | RCHN | 报错5297 | 报错5297 借903-236 836-166 903-215各一个 | 更换了836-166 903-215备件后修复 由于是间歇性故障，需要继续观察 by jesson | SWO-390492 | Defect | HW | Process Module | Process Wheel | 5297 | 5297 error | 5297 Borrowing 903-236 836-166 903-215 each | After replacing 836-166 903-215 spare parts repairdued due to intermittent failure, it is necessary to continue to observe by jesson
No | Visit | Complaint | RCHN | 报错1035 | 报错1035，借944-212（14153缺货），944-214（14322），944-220（13566），944-258（14249），944-216（14259），944-394（14411）各一个 | 更换光单，重新定标后修复 by jesson | SWO-390251 | Defect | HW | Process Module | Optical Unit | 1035 | 1035 error | 1035, 944-212 (14153 out of stock), 944-214 (14322), 944-220 (13566), 944-258 (14249), 944-216 (14259), 944-394 (14411) each | Replace the light sheet and fix by jesson after recalibration
Yes | Visit | Complaint | RIND | frequently calibration error. | sensor cassette failure. | installed the new sensor cassette. performed QC and Samples all results found within the range. Found instrument working fine. | SWO-391599 | [5 empty] | frequently calibration error. | sensor cassette failure. | installed the new sensor cassette. performed QC and Samples all results found within the range. Found instrument working fine.
No | Visit | Complaint | RDEU | Fehler0883: Mischerfehler bei Patientenprobe | Fehler0883: Mischerfehler bei Patientenprobe | Mischerträger fährt nicht die richtige Position Justage Mischerträger bzw. FlexQ Modul Funktionskontrolle | SWO-391170 | Adjustment | HW | Basic Unit | FlexQ Module |  | Error0883: Mixer error in patient sample | Error0883: Mixer error in patient sample | Mixer carrier does not drive the correct position Adjustment mixer carrier or FlexQ module Functional control
No | Visit | Complaint | RAME | Customer says LAC has drifted outside of their control range on qc. Calibrations are passing. | Found cracked measuring chamber on EL $MET module. | Replaced cracked measuring chamber. Performed zero current calibration. Cleaned electrodes and inspected for O-Rings. Performed necessary calibrations. The measuring chamber was contaminated with blood, so I did not save to return as a warranty part. | SWO-390914 | Defect | HW | El/Met Module | measuring chamber | Lac | Customer says LAC has drifted outside of their control range on qc. Calibrations are passing. | Found cracked measuring chamber on EL /MET module. | Replaced cracked measuring chamber. Performed zero current calibration. Cleaned electrodes and inspected for O-Rings. Performed necessary calibrations. The measuring chamber was contaminated with blood, so I did not save to return as a warranty part.
No | Visit | Complaint | RAME | pCo2 out on QC S7735 | Arrived onsite and spoke with Teresa about QC issues. After reviewing PC02 QC plot, it was determined that a lot change took place and the QC was on the target mean. Teresa will adjust the ranges for the new lot. | Range adjusted. | SWO-391524 | User error | Consumables | QC | Range adjustment |  | pCo2 out on QC S7735 | Arrived onsite and spoke with Teresa about QC issues. After reviewing PC02 QC plot, it was determined that a lot change took place and the QC was on the target mean. Teresa will adjust the ranges for the new lot. | Range adjusted.
No | Visit | Complaint | RAME | PCO2 QC flagged | Leak at pCO2, PO2, and pH electrodes. | Replace seal rings for pCO2, PO2, and pH electrodes. Clean measuring chambers. | SWO-391394 | Defect | Electrode | Electrode | Sealing ring | Sealing ring, pCO2 | PCO2 QC flagged | Leak at pCO2, PO2, and pH electrodes. | Replace seal rings for pCO2, PO2, and pH electrodes. Clean measuring chambers.
No | Visit | Complaint | RAME | Customer ran linearity and crew was high. | Analyzer showing high crea. | Cleaned measuring chambers and ran test sample. Still high, replaced membranes and cleaning solution. Ran cals until drift settled and ran cleaning. Ran crea again and it was within spec. | SWO-391396 | Assisted Replacement | Electrode | Membrane | Crea | CAPA-00790 | Customer ran linearity and crew was high. | Analyzer showing high crea. | Cleaned measuring chambers and ran test sample. Still high, replaced membranes and cleaning solution. Ran cals until drift settled and ran cleaning. Ran crea again and it was within spec.
No | Visit | Complaint | RIND | Parameter could not be measured | Checked the data logs found all sample query when we checked the result oximetry parameter values was not reported query in all the oxi parameter could not be measured performed protien removal complete oximetry test hemolyzer test was getting failed switch OFF the machine ON the machine initialization and 2 point calibrtion done sample processed then also same problem oxi paramter could not be measured found problem with oxi module | Need to replace oxi module | SWO-391418 | Defect | HW | Oxi module | Oxi module |  | Parameter could not be measured | Checked the data logs found all sample query when we checked the result oximetry parameter values was not reported query in all the oxi parameter could not be measured performed protien removal complete oximetry test hemolyzer test was getting failed switch OFF the machine ON the machine initialization and 2 point calibrtion done sample processed then also same problem oxi paramter could not be measured found problem with oxi module | Need to replace oxygen module
No | Visit | Complaint | RCHN | 定标失败，灵敏度错误 | 定标失败，灵敏度错误 借 945-737一个。 | 更换945-737备件后修复 | SWO-391117 | [5 empty] | Bid failed, sensitivity error | Bid failed, sensitivity error Borrow 945-737 one. | Repair after replacement of 945-737 spare parts
No | Visit | Complaint | RCHN | 开机白屏 | 开机白屏，借903-365一个。 | 更换903-365后修复 by jesson | SWO-391000 | [5 empty] | Power-on white screen | Power on the white screen, borrow 903-365 one. | Repair by jesson after replacement 903-365
No | Visit | Complaint | RCHN | 血氧不出值 | 血氧不出值 借902-980一个 | 更换902-980后修复 by jesson | SWO-390369 | Defect | HW | Oxi module | Hemolyzer |  | Blood oxygen is not worth. | Blood oxygen does not have a value, borrow 902-980 one | Repair by jesson after replacement 902-980
No | Visit | Complaint | RAME | Intermittent error code 0408 Zero current error. | Confirmed reported error 408, zero current error. Review data logs, first began on 7 $16. Intermittently occurring. | Performed applicable service diagnostics. No issues observed. Performed zero current adjustment procedure. Reran 2PT calibration cycle until stable and passing. | SWO-391778 | No fault found | HW | pH/BG Module | zero current | 0408 | Intermittent error code 0408 Zero current error. | Confirmed reported error 408, zero current error. Review data logs, first began on 7 /16. Intermittently occurring. | Performed applicable service diagnostics. No issues observed. Performed zero current adjustment procedure. Reran 2PT calibration cycle until stable and passing.
No | Visit | Complaint | RKOR | Sample clot error | . | Liquid flow line clean. Inlet module clean. Service mode check. | SWO-391666 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway | clot | Sample clot error | . | Liquid flow line clean. Inlet module clean. Service mode check.
No | Visit | Complaint | RKK | リンスエラー | リンスエラー。凝固血液検体を測定後から発生したとのこと。 | インレットガスケット内に血液クロットあり、除去するもインレットフロー妨害発生。インレットガスケット交換、測定チャンバー掃除、排液ライン洗浄。 | SWO-392035 | Defect | HW | Inlet module | Inlet gasket |  | Rinse error | Rinse error. It is said that it occurred after the measurement of the coagulation blood sample. | There is a blood clot in the inlet gasket, and even the inlet flow disturbance occurs to be removed. Inlet gasket replacement, measuring chamber cleaning, drainage line cleaning.
Yes | Visit | Complaint | RAME | Received call for analyzer failing to pass some information to hospital network. | Found analyzer appears to be configured properly. | Verified analyzer settings. Will be returning to talk to LIS and view analyzer transmitting. | SWO-392299 | [5 empty] | Received call for analyzer failing to pass some information to hospital network. | Found analyzer appears to be configured properly. | Verified analyzer settings. Will be returning to talk to LIS and view analyzer transmitting.
No | Visit | Complaint | RAME | Customer having trouble with pC02 not coming in. | Found that customer had replaced 2 SC and 2 SP still getting the error for pC02. | Replace SC and SP and back flushed the analyzer. | SWO-392318 | Assisted Replacement | SC | SC | pCO2 |  | Customer having trouble with pC02 not coming in. | Found that customer had replaced 2 SC and 2 SP still getting the error for pC02. | Replace SC and SP and back flushed the analyzer.
No | Visit | Complaint | RCHN | 定标不过 | 定标不过 | 更换备用件945-377后修复 | SWO-391890 | [5 empty] | The calibration is not the | The calibration is not the | Repair after replacing spare parts 945-377
No | Visit | Complaint | RCHN | 处理包一处卡扣异常导致处理包不能正常取出来 | 处理包一处卡扣异常导致处理包不能正常取出来 | 破拆处理 | SWO-391944 | Defect | Consumables | SP | SP | stack | Processing package one snap exception causes the processing package can not be removed properly | Processing package one snap exception causes the processing package can not be removed properly | Demolition treatment
No | Visit | Complaint | RDEU | 12.07.2018 Frau Waldschmidt 023729082223 meldet, das die Pumpoen Kal gescheitert ist. 30.07.2018 Frau Arents 023729082220 meldet einen Pumpen-Kal-Fehler. |  |  | SWO-392397 | Revisit required | HW |  | Pump tubes |  | 12.07.2018 Ms. Waldschmidt 023729082223 reports that the Pumpoen Kal has failed. 07/30/2018 Ms Arents 023729082220 reports a pump-calf error.
No | Visit | Complaint | RAME | System in hold without explanation or cause | Data logs not accessible due to patient in room, customer suggested pm visit | None necessary | SWO-392426 | No fault found |  |  |  | nyt (planlagt) besøg? | System in hold without explanation or cause | Data logs not accessible due to patient in room, customer suggested pm visit | None necessary
Yes | 4 | Y | RCZE | Blank |  |  | D4B5BDFF3406C66E | Re classified RMED | planned SWO |  |  |  | Blank
No | Visit | Complaint | RFRA | Pb de mesure de l’oxy | Contrôle de la fluidique et de l’oxy | Changement de l’hemolyseur. Cal tHb. Essais | SWO-392807 | Defect | HW | Oxi module | Hemolyzer |  | Pb measuring oxy | Fluidic and oxy control | Changing the hemolyseur. Cal tHb. Tests
No | Visit | Complaint | RLTD | Software not booting up. | Error code 2status on boot up. New analyser not in use at present. Customer also required that the show value out of reportable range to turn on. On all the new analyser. | Reload window7 and analyser software ok. Reinstalled setup ok. Analyser starting up ok. Line add to the Roma ini file for the show value out of reportable rang on 8 ABL 90. | SWO-392845 | Defect | SW | RMED SW | Database | status code 2 | Software not booting up. | Error code 2status on boot up. New analyser not in use at present. Customer also required that the show value out of reportable range to turn on. On all the new analyser. | Reload window7 and analyser software ok. Reinstalled setup ok. Analyser starting up ok. Line add to the Roma ini file for the show value out of reportable rang on 8 ABL 90.
Yes | Visit | Complaint | RIND | waste value leaking | waste value was leaking | checked the analyser waste value was leaking so removed the waste value and cleaned the all blood stain and tried , but then also it was leaking, so changed the waste value and calibrated to raedy mode | SWO-394523 | [5 empty] | waste value leaking | waste value was leaking | checked the analyser waste value was leaking so removed the waste value and cleaned the all blood stain and tried , but then also it was leaking, so changed the waste value and calibrated to raedy mode
No | Visit | Complaint | RCHN | 温度不稳定 | 温度不稳定，借903-334,902-972各一个 | 更换903-334,902-972后修复(by chuanwei) | SWO-391805 | [5 empty] | Temperature instability | Temperature instability, borrow 903-334,902-972 each | Replacement 903-334,902-972 post-repair (by chuanwei)
No | Visit | Complaint | RCHN | 报错1031，4548 | 报错1031，借903-028，903-364,902-924,944-220（14087），944-394（13915缺货）各一个 | 更换903-028，903-364,902-924后重新定标后修复(by chuanwei) | SWO-391881 | Defect | HW | Process Module | Optical Unit | 1031 | 10314548 | 1031, lend ingress 903-028, 903-364, 902-924,944-220 (14087), 944-394 (13915 out of stock) each | Replace 903-028, 903-364, 902-924 after re-calibration repair (by chuanwei)
No | Visit | Complaint | RKK | 装置から異音がする。 日付/時刻がずれる | 装置から異音 | 装置引き上げ | SWO-403585 | [5 empty] | There is a noise coming from the device. Date/Time Shifts | Noise from the device | Equipment pulling up
No | Visit | Complaint | RAME | QC is not being read | Ampoule stepping motor is not functioning during qc aspiration. | Replaced ampoule stepping motor. | SWO-392913 | Defect | HW | Autocheck Module | Stepper motor |  | QC is not being read | Ampoule stepping motor is not functioning during qc aspiration. | Replaced ampoule stepping motor.
No | 5 | Y | RPAC | Blank | Supply New Glucose Electrode |  | D4B586F634064F94 | Assisted Replacement | Electrode | Electrode | glu |  | Blank | Supply New Glucose Electrode
No | Visit | Complaint | RAME | Customer says LAC has drifted outside of their control range on qc. Calibrations are passing. | Sent customer new lot number of membranes to try (lot 3915) This didn't help problem. Ordered new EL $MET module board and AU board. | Replaced EL $MET module PCB. Performed zero current calibration and multiple calibrations. This didn't have any affect, so I removed new part and put in original. Performed multiple decontamination programs and reprogrammed. Machine in came in after that. | SWO-392920 | User expectation | Specifications | QC | Electrode | Lac | Customer says LAC has drifted outside of their control range on qc. Calibrations are passing. | Sent customer new lot number of membranes to try (lot 3915) This didn't help problem. Ordered new EL /MET module board and AU board. | Replaced EL $MET module PCB. Performed zero current calibration and multiple calibrations. This didn't have any affect, so I removed new part and put in original. Performed multiple decontamination programs and reprogrammed. Machine in came in after that.
No | Visit | Complaint | RAME | Customer describing problem with frequent Glucose 1 PT cal failure. | Observed leak out of pH measuring chamber well. Ref chamber also had filled with fluid. | Cleaned Measuring chamber, ph electrode and sleeve. Cleaned Ref electrode and replaced membrane. Completed 2 PT cal cycles until values became stable. Final 1 and 2 PT calibration stable, QC in range. | SWO-392385 | Assisted Replacement | Electrode | Membrane | Reference |  | Customer describing problem with frequent Glucose 1 PT cal failure. | Observed leak out of pH measuring chamber well. Ref chamber also had filled with fluid. | Cleaned Measuring chamber, ph electrode and sleeve. Cleaned Ref electrode and replaced membrane. Completed 2 PT cal cycles until values became stable. Final 1 and 2 PT calibration stable, QC in range.
Yes | Visit | Complaint | RIND | Service Visit. | Touch screen is not working properly. | Take Spirit on gauss piece and wipe the screen, now it is working, need to install the software, but beacause of sample rush staff and pathologist request to come 6pm. | SWO-394756 | Re classified RMED | same day upSWO |  |  |  | Service Visit. | Touch screen is not working properly. | Take Spirit on gauss piece and wipe the screen, now it is working, need to install the software, but beacause of sample rush staff and pathologist request to come 6pm.
No | Visit | Complaint | RIND | Service Visit. | Need to install touch screen software rufus2.11. | Try to install the software, analyzer boot from the pendrive, but the process of installation was not successfully done, again do the installation with the proper procedure, failed to install, clean the touch screen with spirit, if further problem arrive need to replace the touch screen. | SWO-394760 | Cleaned | HW | DMS_PC | Touch screen |  | Service Visit. | Need to install touch screen software rufus2.11. | Try to install the software, analyzer boot from the pendrive, but the process of installation was not successfully done, again do the installation with the proper procedure, failed to install, clean the touch screen with spirit, if further problem arrive need to replace the touch screen.
No | Visit | Complaint | RLTD | NOT ACCEPTING THE CRP NUMBER Habib - 07580 691 855 | Checked data logs and service programs. Scanned all cartridges in service and nothing was picked up. Opened cartridge door and scanned again, now all cartridges have been recognised. Cartridges on board were not being recognised due to the calibration being invalidated by the customer. | Removed cartridge scanner, cleaned and adjusted. Analyser is ready for use. | SWO-392652 | Cleaned | HW | Storage Module | Barcode reader, storage wheel |  | NOT ACCEPTING THE CRP NUMBER Habib - 07580 691 855 | Checked data logs and service programs. Scanned all cartridges in service and nothing was picked up. Opened cartridge door and scanned again, now all cartridges have been recognised. Cartridges on board were not being recognised due to the calibration being invalidated by the customer. | Removed cartridge scanner, cleaned and adjusted. Analyser is ready for use.
Yes | 13 | Y | RCZE | Blank |  |  | D4B83C2D20379D60 | [5 empty] | Blank
Yes | Visit | Complaint | RDEU | 11.06.2018/10:00 Hr. Schmidt (MT) kann den ABL nicht neu starten, Bildschirm bleibt schwarz | Funktioniert ohne Beanstandung | Fehler behoben | SWO-393582 | [5 empty] | 11.06.2018/10:00 Mr. Schmidt (MT) cannot restart the ABL, screen remains black | Works without complaint | Fixed a bug
No | Visit | Complaint | RKK | キャルにてｔHbが・・・。オキシハードウェア問題。 | ケース詳細の通り | 引上げ修理。液回路内における気泡混入の疑い。液回路内の全洗浄及びチューブ類の全交換。 | SWO-395658 | Defect | HW | Basic Unit | PM Kit |  | THb in the car. Oxy hardware problem. | As the case details | Raise repairs. Suspected of air bubbles mixing in the liquid circuit. Total cleaning and total replacement of tubes in the liquid circuit.
No | Visit | Complaint | RKK | Cal3溶液に気泡があります | Cal3溶液に気泡があります | 回路洗浄実施 インレットプローブ曲がり、詰まりチェック→OK ホルダ付インレットガスケット交換（院内在庫使用） | SWO-398519 | Defect | HW | Inlet module | Inlet gasket |  | There are bubbles in the Cal3 solution. | There are bubbles in the Cal3 solution. | Circuit cleaning Inlet probe bend, jam check → OK Inlet gasket replacement with holder (in-hospital stock use)
No | Visit | Complaint | RAME | Sample mixer running slowly | Sample mixer intermittently running slowly | Replaced mixer and verified operation. Instrument operating normally. | SWO-393415 | Defect | HW | Basic Unit | Sample mixer |  | Sample mixer running slowly | Sample mixer intermittently running slowly | Replaced mixer and verified operation. Instrument operating normally.
Yes | 6 | Y | RPAC | Blank | Consumable failure - Sensor cassette calibration failures |  | D4B6DDD7AA537E45 | [5 empty] | Blank | Consumable failure - Sensor cassette calibration failures
No | 6 | Y | RCZE | Blank |  |  | D4B8395C20419D5D | Defect | HW | Basic Unit | PCB, Wet section |  | Blank
No | Visit | Complaint | RLTD | Please call Claire Latham on 01553 613 780 | Sample mixer not mixing. | Sample mixer intermittently working, cleaned mixer tray, tested cables, issue remained. Replaced sample mixer, now ok. | SWO-393398 | Defect | HW | Basic Unit | Sample mixer |  | Please call Claire Latham on 01553 613 780 | Sample mixer not mixing. | Sample mixer intermittently working, cleaned mixer tray, tested cables, issue remained. Replaced sample mixer, now ok.
No | Visit | Complaint | RSAF | PO2 reading low on level B, Sol pack and SC replaced | RMED gave feedback with regards pO2 act issues | Move analyzer to a more ventilated room, moved analyzer to Sluice room. Left analyzer to stabilize and calibrate, done calibration after a hour of analyzer being moved ok all levels of qc's passed | SWO-394998 | User error | User error | User Error | Ventilation | pO2 QCB low | PO2 reading low on level B, Sol pack and SC replaced | RMED gave feedback with regards pO2 act issues | Move analyzer to a more ventilated room, moved analyzer to Sluice room. Left analyzer to stabilize and calibrate, done calibration after a hour of analyzer being moved ok all levels of qc's passed
No | Visit | Complaint | RAME | Battery will not hold charge. | Battery does not hold charge when analyzer is unplugged. | Replaced battery pack. | SWO-393489 | [5 empty] | Battery will not hold charge. | Battery does not hold charge when analyzer is unplugged. | Replaced battery pack.
No | Visit | Complaint | RAME | Scanner is malfunctioning and not scanning sometimes | Scanner is not scanning properly and will fail intermittently | Replaced internal scanner and checked connections. | SWO-393491 | [5 empty] | Scanner is malfunctioning and not scanning sometimes | Scanner is not scanning properly and will fail intermittently | Replaced internal scanner and checked connections.
No | Visit | Complaint | RAME | Jessica Olson 320-769-4393 ext 2193 Customer called stating that CO2 is out on linearity kit. | Customer was having issues with pC02 reading slightly high on both levels one and 2 of Linearity. | I arrived and saw the printed information that they were having slightly high pC02 readings. I looked though the data logs and found that the Linearity kit was being ran off of an old manual QC pC02 calibration which was out of range. The device was also displaying that levels 1, 2, and 3 manual QC needed to be ran before proceeding. I ran all 3levels of QC and they all passed within range. I ran the remainder of the VK kit that the staff was using and all remaining vials for all levels, including Hct were within the specified range for the lot. I spoke with Jessica on site and went over the process for running Linearity and explained to always make sure the device is ready for analysis before running Linearity. I will also speak with the lab manager when she is available. Device is ready to use and I left a VK kit for them to run again. | SWO-393497 | [5 empty] | Jessica Olson 320-769-4393 ext 2193 Customer called stating that CO2 is out on linearity kit. | Customer was having issues with pC02 reading slightly high on both levels one and 2 of Linearity. | I arrived and saw the printed information that they were having slightly high pC02 readings. I looked though the data logs and found that the Linearity kit was being ran off of an old manual QC pC02 calibration which was out of range. The device was also displaying that levels 1, 2, and 3 manual QC needed to be ran before proceeding. I ran all 3levels of QC and they all passed within range. I ran the remainder of the VK kit that the staff was using and all remaining vials for all levels, including Hct were within the specified range for the lot. I spoke with Jessica on site and went over the process for running Linearity and explained to always make sure the device is ready for analysis before running Linearity. I will also speak with the lab manager when she is available. Device is ready to use and I left a VK kit for them to run again.
No | Visit | Complaint | RAME | 581 and 688 errors. | Checked the data logs and verified the problem. | Replaced the hemolyzer and the pinch valve | SWO-393534 | [5 empty] | 581 and 688 errors. | Checked the data logs and verified the problem. | Replaced the hemolyzer and the pinch valve
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 参比电极错误 1361 | 更换测试卡后修复 | SWO-393046 | Assisted Replacement | SC | SC | Reference | 1361 | Test card, reagent package failure | Reference electrode error 1361 | Repair after replacing test card
No | Visit | Complaint | RCHN | 打印机无法打印 | 打印机无法打印，需要借件910-286 | 更换910-286后修复（by chuanwei） | SWO-392541 | Defect | HW | Cover and Chassis | Printer |  | The printer can't print | The printer cannot print and needs a loan 910-286 | Repair after replacement 910-286 (by chuanwei)
No | Visit | Complaint | RCHN | 报错1068 | 报错1068，需要借件903-210 | 更换903-210后修复(by chuanwei) | SWO-392536 | Defect | HW | Process Module | PCB, Cup dryer | 1068 | 1068 error | 1068, need to borrow 903-210 | Repair after replacement 903-210 (by chuanwei)
No | Visit | Complaint | RDEU | Nach der Wartung besteht der Analysator die Qualitätskontrolle Level 2 nicht im Parameter LAC. Ist immer außerhalb der Rilibäk. 01.08.2018 9:45 Fr. Langenreher 0401818854383: weiterhin Lac-Probleme | Lactat Problem in QC level 2 | Fehleranalyse, Lactat Elektrode erneuert und membraniert, Kalibrierung und Qualitätskontrolle durchgeführt, i. O. | SWO-393221 | User expectation | Specifications | QC | Rilibak |  | After maintenance, the analyzer does not pass the Quality Control Level 2 in the LAC parameter. Is always outside the Rilibäk. 01.08.2018 9:45 Ms. Langenreher 0401818854383: Still Lac problems | Lactat Problem in QC level 2 | Fault analysis, lactate electrode renewed and membraneed, calibration and quality control carried out, i.e.
No | Visit | Complaint | RKK | 電解質・pCO2のQCのばらつき |  |  | SWO-395889 | No fault found |  |  |  | mail, again: change of measurement chamber pH/BG module | Qc variation s of electrolyte and pCO2
No | Visit | Complaint | RAME | Customer describing problem with frequent Glucose 1 PT cal failure. | Continued intermittent Glucose cal failure $drift along with Lac. | Removed and cleaned measuring chamber and all electrode connectors in Met I module. Performed Zero current adjustment procedure, reverified setting. Replace critically low cal 1 solution bottle. Performed multiple 2 PT and 1 PT calibration cycles. No failures observed once module stabilized after service. | SWO-393103 | Cleaned | HW | El/Met Module | Electrode Connector |  | Customer describing problem with frequent Glucose 1 PT cal failure. | Continued intermittent Glucose cal failure /drift along with Lac. | Removed and cleaned measuring chamber and all electrode connectors in Met I module. Performed Zero current adjustment procedure, reverified setting. Replace critically low cal 1 solution bottle. Performed multiple 2 PT and 1 PT calibration cycles. No failures observed once module stabilized after service.
No | Visit | Complaint | RIND | Blank Results in pO2 | Checked and found that the PO2 showing Blank results | Cleaned the PO2 chamber and decontaminate the PO2 membrane, Performed the cleaning procedure and 2 point calibration. Now the analyzer working satisfactorily | SWO-394727 | Cleaned | HW | pH/BG Module | measuring chamber | PO2 | Blank Results in pO2 | Checked and found that the PO2 showing Blank results | Cleaned the PO2 chamber and decontaminate the PO2 membrane, Performed the cleaning procedure and 2 point calibration. Now the analyzer working satisfactorily
No | Visit | Complaint | RDEU | Intensität der Blitzlampe ist gescheitert | Das optische Einheit ist defekt. Intensität der Blitzlampe ist grenzwertig | Das optische Einheit ersetzt. Alle Parameter neu kalibriert. Neustart. | SWO-393663 | Defect | HW | Process Module | Optical Unit |  | Intensity of the flash light has failed | The optical unit is defective. Intensity of the flash lamp is borderline | Replaces the optical unit. All parameters are recalibrated. Restart.
No | Visit | Complaint | RPOL | Podłączenie do Netop. | Podłączenie do zdalnego nadzoru. IP adres 192.168.13.200 | Kalibracja hemoglobiny, instalacja zdalnego dostępu Netop. Szkolenie personelu z obsługi oprogramowania Netop. | SWO-393281 | Setup | SW | RMED SW | Remote |  | Connection to Netop. | Remote monitoring connection. IP Address 192.168.13.200 | Hemoglobin calibration, remote access installation Netop. Training of personnel with the support of Netop software.
No | Visit | Complaint | RLTD | Carole Shanahan 0667184361 - Valve pressure error - ED | Valve pressure error - ED | Changed the pinch valve tubing on board and flushed out device X 3. Ran all fluid service tests. All passed. | SWO-393800 | Defect | HW | Oxi module | Pinch valve tube |  | Carole Shanahan 0667184361 - Valve pressure error - ED | Valve pressure error - ED | Changed the pinch valve tubing on board and flushed out device X 3. Ran all fluid service tests. All passed.
Yes | 8 | Y | RPAC | Blank | Upgrade analyser to have Windows 7 platform. |  | D4B83A692025F261 | [5 empty] | Blank | Upgrade analyser to have Windows 7 platform.
No | Visit | Complaint | RFRA | Ne détecte pas lorsque la seringue est enfoncée. | Erreur détection seringue de tps en tps | Chgt carte détection seringue avant aspiration. Calibration Passage CQ. | SWO-394073 | Defect | HW | Auto Inlet module | PCB, Auto inlet | Hall | Do not detect when the syringe is pressed. | Tps syringe detection error | Chgt syringe detection card before suction. Calibration Passage CQ.
No | Visit | Complaint | RDEU | 23.07.2018; Fr. Alilovioic; Tel.: 02093643545; Probenabbrüche, Einlass schließt nicht richtig 23.07.2018 Fr. Alilovioic 02093643545: gleicher Fehler wie gestern ist erneut aufgetreten | Immer wieder Probenabbrüche, Einl.dichtung geht nicht richtig zurück, nach Einbau Leckage erkannt, | Neues Einl.modul eingebaut und kalibriert, Gerät kalibriert nicht, Leckagefehler, altes Modul wieder eingebaut, | SWO-393826 | Revisit required | HW | Auto Inlet module | Auto Inlet module | Cap spring | 23.07.2018; Ms. Alilovioic; Tel.: 02093643545; Sample aborts, inlet does not close properly 23.07.2018 Ms. Alilovioic 02093643545: same error as yesterday has occurred again | Time and again sample breaks, single sealing does not return properly, after installation leakage detected, | New single module built-in and calibrated, device does not calibrate, leakage error, old module reinstalled,
Yes | Visit | Complaint | RCHN | 接口板接触不良 | 接口板接触不良，借902-972一个 | 更换902-972后修复(by chuanwei) | SWO-393100 | [5 empty] | Poor interface board contact | Poor interface board contact, borrow 902-972 one | Repair after replacement 902-972 (by chuanwei)
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 质控不过 | 更换测试卡后修复 | SWO-393692 | [5 empty] | Test card, reagent package failure | But the quality control is not | Repair after replacing test card
No | Visit | Complaint | RCHN | 无法连接试剂包 | 无法连接试剂包 | 更换试剂包后修复 | SWO-393778 | [5 empty] | Unable to connect the reagent package | Unable to connect the reagent package | Repair after replacement of reagent kit
No | Visit | Complaint | RIND | parameter failure | Lac failure | flush out the sensor cassette , re-install ,calibration not passed ,need to change sensor cassette. | SWO-394783 | [5 empty] | parameter failure | Lac failure | flush out the sensor cassette , re-install ,calibration not passed ,need to change sensor cassette.
No | Visit | Complaint | RCHN | 定标不过 | 定标不过，借945-737一个 | 更换945-737后修复(by chuanwei) | SWO-392516 | [5 empty] | The calibration is not the | But, borrow 945-737. | Repair after replacement 945-737 (by chuanwei)
No | Visit | Complaint | RESP | Erro 1223 - Erro de medição de Hct | Erro 1223 $1257 - valor de Hct não pode ser determinado corretamente | Verificação do registro de dados. Descontaminacao do circuito de fluidos. Limpeza com tubo de limpeza. Realização de controle de qualidade e de várias amostras para comprovação do correcto funcionamento. Verificou-se que duas amostras voltaram a dar o mesmo erro, precisamente as que tinham dado antes da intervenção, conclui-se que o problema será da amostra. | SWO-393871 | User error | HW | Aspiration Module | Needle Wash Unit | 1223 | Error 1223-HCT measurement error | Error /1223 1257-Hct value cannot be determined correctly | Checking the data record. Decontamination of the fluid circuit. Cleaning with cleaning tube. Conducting quality control and several samples to prove the correct operation. It was found that two samples again gave the same error, precisely those that had given before the intervention, it is concluded that the problem will be the sample.
No | Visit | Complaint | RCHN | 软件不匹配 | 软件不匹配 借910-343 944-258（14591）944-212（14558）944-220（14531） | 更换910-343，重新定标 | SWO-393162 | Revisit required | HW | Analyzer Control Module | CPU |  | Software mismatch | Software Mismatch Borrowing 910-343 944-258 (14591) 944-212 (14558) 944-220 (14531) | Replace 910-343, re-bid
No | Visit | Complaint | RLTD | Anita Casey 01865226513 | The electrode connectors had previously been checked by Anita but I checked them again and they looked ok, I cleaned the contacts with a fibreglass pencil just to make sure a good contact was being made. I checked the electrode updatings while the 2 point calibration was running and Crea was only giving -1umol $L on Cal2 solution and the bottle label says it should be 190uM, when I checked the Cal2 bottle the foil had been removed from the dosi-cap but it hadn’t been turned round and had the powder dissolved into the solution. | I activated a new Cal2 bottle and installed it on the analyser the 2 point calibration was then run and it passed. | SWO-393912 | User error | Consumables | Calibration solution | DosiCap |  | Anita Casey 01865226513 | The electrode connectors had previously been checked by Anita but I checked them again and they looked ok, I cleaned the contacts with a fibreglass pencil just to make sure a good contact was being made. I checked the electrode updatings while the 2 point calibration was running and Crea was only giving -1umol /L on Cal2 solution and the bottle label says it should be 190uM, when I checked the Cal2 bottle the foil had been removed from the dosi-cap but it hadn’t been turned round and had the powder dissolved into the solution. | I activated a new Cal2 bottle and installed it on the analyser the 2 point calibration was then run and it passed.
No | Visit | Complaint | RIND | Electrolyte vale and Ca++ (7.4) results are not correct. | Checked and found that the blood clots in pH chamber, Ca++ (7.4) values not correct. | Cleaned the pH chamber and reference chamber, performed cleaning and 2 point calibration which got passed. Now the analyzer working satisfactorily. For values of Ca++ (7.4). If the pH value falls below 7.4 then the value of Ca++ (7.4) will not calculate. | SWO-394729 | Cleaned | HW | pH/BG Module | measuring chamber | clot | Electrolyte vale and Ca++ (7.4) results are not correct. | Checked and found that the blood clots in pH chamber, Ca++ (7.4) values not correct. | Cleaned the pH chamber and reference chamber, performed cleaning and 2 point calibration which got passed. Now the analyzer working satisfactorily. For values of Ca++ (7.4). If the pH value falls below 7.4 then the value of Ca++ (7.4) will not calculate.
Yes | Visit | Complaint | RAME | Intermittent error code 0408 Zero current error. | Intermittent error 408, Zero current error. Observed erratic fluctuating elevated pO2 zero values on 2PT calibration. Also observed pCO2 2PT sensitivity passing but not as stable as expected. Performed all applicable service diagnostics , no errors indicated. Suspect possible pH $BG PCB adjustment needed, possible intermittent marginal failure of PCB or gas valve. | Replaced pump tubing, black tubing to gas valves, checked and verified power supply and AC supply was not cause, confirmed pH $BG PCB was not cause. Will return to continue troubleshooting. | SWO-395109 | Re classified RMED | HW | gas module | Valve 1 | ran out of time | Intermittent error code 0408 Zero current error. | Intermittent error 408, Zero current error. Observed erratic fluctuating elevated pO2 zero values on 2PT calibration. Also observed pCO2 2PT sensitivity passing but not as stable as expected. Performed all applicable service diagnostics , no errors indicated. Suspect possible pH /BG PCB adjustment needed, possible intermittent marginal failure of PCB or gas valve. | Replaced pump tubing, black tubing to gas valves, checked and verified power supply and AC supply was not cause, confirmed pH $BG PCB was not cause. Will return to continue troubleshooting.
No | Visit | Complaint | RDEU | pH Elektrode muss überprüft werden. |  | pH Elektrode instabil bei der QC Kontrolle.Gerät kalibriert und QC durchgeführt. | SWO-394348 | Assisted Replacement | Electrode | Electrode | pH |  | pH electrode must be checked. |  | pH electrode unstable in the QC control.Device calibrated and QC performed.
Yes | 9 | Y | RPAC | Blank | Upgrade ABL90 to Windows 7 Platform. |  | D4D3F72F44326BB0 | [5 empty] | Blank | Upgrade ABL90 to Windows 7 Platform.
No | Visit | Complaint | RFRA | Demande souvent changement joint d'introduction | Aiguille cintrée ->chgt aiguille. Contrôle fuite d'air: pas de fuite. Vérification en mode service fluid transport. | Chgt aiguille Re contrôle en mode service fluid transport. Formation utilisateur dans le mode service. | SWO-394548 | Defect | HW | Auto Inlet module | Probe bend |  | Request often change attached introduction | Curved needle - 'chgt needle'. Air leak control: no leakage. Check in fluid transport service mode. | Chgt needle Re control in fluid transport service mode. User training in service mode.
No | Visit | Complaint | RDEU | 23.07.2018; Fr. Alilovioic; Tel.: 02093643545; Probenabbrüche, Einlass schließt nicht richtig 23.07.2018 Fr. Alilovioic 02093643545: gleicher Fehler wie gestern ist erneut aufgetreten | Proben werden zeitweise abgebrochen, Einl.dichtung fährt nicht zurück, | neues Einl.modul eingebaut, kalibriert u.Probemessungen durchgeführt, | SWO-394587 | Defect | HW | Auto Inlet module | Auto Inlet module | Cap spring | 23.07.2018; Ms. Alilovioic; Tel.: 02093643545; Sample aborts, inlet does not close properly 23.07.2018 Ms. Alilovioic 02093643545: same error as yesterday has occurred again | Samples are temporarily aborted, inl. sealing does not return, | new single module installed, calibrated and sample measurements carried out,
No | Visit | Complaint | RFRA | Erreur site ouvert lors de répétabilité. | Tests mouvements de l’inlet dans menu de service effectuées par téléphone avec Mr AZIZI technicien bio : en capillaire le retour en position « closed » ne se fait pas. | Arrêt du dispositif $ $ démontage de l’ensemble « Motion mechanism » $ $ inspection de tous les éléments composants ce mécanisme : couplage moteur $ vis sans fin défectueux $ $ changement complet du « motion mechanism » ok $ $ remontage de l’ensemble ok $ $ remise en route du dispositif ( démarrage à froid ) $ $ tests dans menu technique du mouvement du module « inlet » dans toutes les positions ok | SWO-394515 | Defect | HW | Auto Inlet module | Auto Inlet module | Cap spring | Site error opened during repeatability. | Inlet movement tests in service menu performed by phone with Mr AZIZI organic technician: in capillaries the return to the position "closed" is not done. | Device shutdown $$ dismantling of the "Motion mechanism" set $$ inspection of all components this mechanism: engine coupling $ screws endlessly defective $$ complete change of the "motion mechanism" ok $$ reassembly of the set ok $$ rebate in ro device ute (cold start) $ $ tests in technical menu of the "inlet" module movement in all positions ok
Yes | 3 | Y | RCZE | Blank |  |  | D4B95EAA96209D63 | Re classified RMED | planned SWO |  |  |  | Blank
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | pco2定标失败，0210，0378 | 更换测试卡后修复 | SWO-399075 | [5 empty] | Test card, reagent package failure | pco2 bid failure, 0210,0378 | Repair after replacing test card
No | Visit | Complaint | RESP | Erro 1223 - Erro de medição de Hct | Erro de | Substituição da unidade de lavagem com célula de Hct. Realização de kit anual e kit de tubos. Ajustes e calibração da agulha e célula de Hct de acordo com o procedimento. Inicialização. Verificação do correto funcionamento com controle de qualidade e execução de uma amostra. | SWO-394555 | Defect | HW | Aspiration Module | Needle Wash Unit | 1223 | Error 1223-HCT measurement error | Error | Replacement of the washing unit with HCT cell. Conducting annual kit and tube kit. Adjustments and calibration of the needle and Hct cell according to the procedure. Startup. Verification of the correct functioning with quality control and execution of a sample.
Yes | 7 | Y | RCZE | Blank |  |  | D4BE4B016E2A374E | [5 empty] | Blank
No | Visit | Complaint | RAME | Analyzer will not properly Boot | 8-2-18 Inspection of the system revealed a boot disk error. Bios voltages screen shows +12 V is at 7.4V. | 8-2-18 replaced power supply, same issue, ordering parts for 8-3-18. On 8-3-18 replaced the CPU, same issue. Negotiated with the customer to return Mon 8-6-18 to replace hard drive & do a full SW installation. | SWO-394208 | Revisit required | HW | DMS_PC | Power supply |  | Analyzer will not properly Boot | 8-2-18 Inspection of the system revealed a boot disk error. Bios voltages screen shows +12 V is at 7.4V. | 8-2-18 replaced power supply, same issue, ordering parts for 8-3-18. On 8-3-18 replaced the CPU, same issue. Negotiated with the customer to return Mon 8-6-18 to replace hard drive & do a full SW installation.
Yes | Visit | Complaint | RAME | Intermittent error code 0408 Zero current error. | Intermittent error 408, Zero current error. Observed erratic fluctuating elevated pO2 zero values on 2PT calibration. Also observed pCO2 2PT sensitivity passing but not as stable as expected. Performed all applicable service diagnostics , no errors indicated. Suspect possible pH $BG PCB adjustment needed, possible intermittent marginal failure of PCB or gas valve. | Replaced pump tubing, black tubing to gas valves, checked and verified power supply and AC supply was not cause, confirmed pH $BG PCB was not cause. Will return to continue troubleshooting. | SWO-395108 | Re classified RMED | HW | gas module | Valve 1 | ran out of time | Intermittent error code 0408 Zero current error. | Intermittent error 408, Zero current error. Observed erratic fluctuating elevated pO2 zero values on 2PT calibration. Also observed pCO2 2PT sensitivity passing but not as stable as expected. Performed all applicable service diagnostics , no errors indicated. Suspect possible pH /BG PCB adjustment needed, possible intermittent marginal failure of PCB or gas valve. | Replaced pump tubing, black tubing to gas valves, checked and verified power supply and AC supply was not cause, confirmed pH $BG PCB was not cause. Will return to continue troubleshooting.
No | Visit | Complaint | RAME | Waste is leaking out the syringe port. | Found broken inlet nipple for capillary | Replaced inlet nipple for capillary, flushed both side. Problem resolved. All calibrations and 3 Levels of QC passed as per manufacturer specification | SWO-394699 | Defect | HW | Inlet module | Nipple kit |  | Waste is leaking out the syringe port. | Found broken inlet nipple for capillary | Replaced inlet nipple for capillary, flushed both side. Problem resolved. All calibrations and 3 Levels of QC passed as per manufacturer specification
No | Visit | Complaint | RIND | Sample aborted. | Check the patient result log & found sample aborted. | Remove the gasket & found clot in it,Clean the gasket with distilled water & then run few samples & got the desired ,Handover the instrument to the customer. | SWO-394745 | Cleaned | HW | Inlet module | Inlet gasket | clot | Sample aborted. | Check the patient result log & found sample aborted. | Remove the gasket & found clot in it,Clean the gasket with distilled water & then run few samples & got the desired ,Handover the instrument to the customer.
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | CO2定标不过 0378 | 更换测试卡后修复 | SWO-398461 | Assisted Replacement | SC | SC | pCO2 | Cal | Test card, reagent package failure | CO2 calibration is no more than 0378 | Repair after replacing test card
No | Visit | Complaint | RCHN | 软件故障，报错0001 | 工程师货号报错，借910-343一个 | 更换910-343后 ，定标失败，需重新加借定标卡，故此单先关单(By Ran) | SWO-394379 | Revisit required | HW | Analyzer Control Module | CPU |  | Software failure, error 0001 | Engineer's number misreported, borrowed 910-343 one | After replacing 910-343, the bid failed and the bid card needs to be re-borrowed, so this single first clearance (By Ran)
No | Visit | Complaint | RCHN | 光学强度超范围，报错1031 | 光学强度超范围，报错1031，需要借件903-364、944-216（14456）、944-258（14485）缺货 | 更换903-364后重新定标后修复(By Ran) | SWO-393653 | Defect | HW | Process Module | Optical Unit | 1031 | Optical strength over range, error 1031 | Optical strength over range, error 1031, need to borrow 903-364, 944-216 (14456), 944-258 (14485) out of stock | Re-fix after replacement of 903-364 (By Ran)
No | Visit | Complaint | RAME | Intermittent error code 0408 Zero current error. | Intermittent error 408, Zero current error. Observed erratic fluctuating elevated pO2 zero values on 2PT calibration. Also observed pCO2 2PT sensitivity passing but not as stable as expected. Performed all applicable service diagnostics , no errors indicated. Suspect possible pH $BG PCB adjustment needed, possible intermittent marginal failure of PCB or gas valve. | After extensive troubleshooting determined issue to be leaking gas 1 valve causing zero error. Replaced valve, performed service diagnostic testing. No further calibration failures, all testing is now passing. | SWO-395107 | Defect | HW | gas module | Valve 1 |  | Intermittent error code 0408 Zero current error. | Intermittent error 408, Zero current error. Observed erratic fluctuating elevated pO2 zero values on 2PT calibration. Also observed pCO2 2PT sensitivity passing but not as stable as expected. Performed all applicable service diagnostics , no errors indicated. Suspect possible pH /BG PCB adjustment needed, possible intermittent marginal failure of PCB or gas valve. | After extensive troubleshooting determined issue to be leaking gas 1 valve causing zero error. Replaced valve, performed service diagnostic testing. No further calibration failures, all testing is now passing.
No | Visit | Complaint | RCHN | 接口板腐蚀 | 接口板腐蚀 借902-972一个 |  | SWO-394846 | [5 empty] | Interface plate corrosion | Interface board corrosion borrows 902-972 one
Yes | Visit | Complaint | RCHN | 报错5684 | 报错5684，借636-533一个。 | 更换636-533后修复(By Ran) | SWO-394382 | [5 empty] | 5684 error | 5684, 636-533 one. | Repair after replacement 636-533 (By Ran)
No | Visit | Complaint | RCHN | 进样支架断裂 | 进样支架断裂 借903-584 2个 | 更换903-584 2个后修复(By Ran) | SWO-394738 | Defect | HW | Inlet module | Tube holder |  | The sample bracket breaks | The sample bracket is broken, borrowed 903-584 2 | Replace 903-584 2 post-repair (By Ran)
No | Visit | Complaint | RCHN | 软件故障，报错0001 | 加借定标卡 944-212（14558）944-220（14531） 944-258（14591）各一张 | 重新定标后修复 Reinstalled the software. Recal, then OK.(By Ran) | SWO-394718 | Defect | HW | Analyzer Control Module | CPU |  | Software failure, error 0001 | Plus Bid Card 944-212 (14558)944-220 (14531) 944-258 (14591) each | 重新定标后修复 Reinstalled the software. Recal, then OK. (By Ran)
Yes | Visit | Complaint | RKK | 気泡エラー Oxiエラー パック脱着異音 | 気泡エラー Oxiエラー パック脱着異音 | イヤリーキット交換 アクチュエーターモーター交換 | SWO-396295 | [5 empty] | Bubble Error Oxi Error Pack Desorption Noise | Bubble Error Oxi Error Pack Desorption Noise | Early Kit Replacement Actuator motor replacement
No | Visit | Complaint | RDEU | 15.06.2018 12:00Uhr Herr Varhy 03078824079 (Station 03078822040): Der Einlass ist defekt, ein Teil ist abgebrochen | Anwenderfehler -> wieder Einlaßdichtung abgebrochen + Nadel verbogen | Anwenderfehler -> Einlaßnadel u. -dichtung gew.; Startphase abgewartet; Kal. + QC i.O. | SWO-395518 | Defect | HW | Inlet module | Inlet module | broken | 15.06.2018 12:00pm Mr Varhy 03078824079 (Station 03078822040): The inlet is defective, one part is broken | User error -> again inlet seal broken off + needle bent | User error -> inlet needle and seal; start-up phase awaited; Kal. + QC in the first place
Yes | Visit | Complaint | RDEU | Kunde bittet um kurzfristige Lieferung einer SC und eines LP | LP und SC ausgeliefert | . | SWO-395054 | Re classified RMED | Logistical upSWO |  |  |  | Customer asks for short-term delivery of a SC and an LP | LP and SC delivered | .
No | Visit | Complaint | RKK | 先週の機械入れ替えからミキサー作動していない | ケース詳細の通り | ミキサーﾌﾗｯﾄｹｰﾌﾞﾙの抜け | SWO-395668 | Adjustment | HW | Basic Unit | Mixer Tray | Cable | The mixer hasn't been working since last week's machine replacement. | As the case details | Mixer flat cable missing
Yes | 8 | Y | RPAC | Blank | Pre Install checks |  | D4BDC67DF844B8CB | [5 empty] | Blank | Pre Install checks
No | Visit | Complaint | RLTD | Lynn 01582497991 | Analyser failing to start up. | Replaced power supply, now ok. | SWO-396445 | Defect | HW | Power Supply | Power supply |  | Lynn 01582497991 | Analyser failing to start up. | Replaced power supply, now ok.
Yes | 8 | Y | RPAC | Blank | SC failure. |  | D4BD68CCF83E2DFF | [5 empty] | Blank | SC failure.
No | Visit | Complaint | RLTD | Failing Oximetry QC’s and pressure tests. | Failing Oximetry QC’s and pressure tests. | Opened sensor bay, fluid under pinch valve tubing. Replaced pinch valve tubing, now ok. | SWO-395306 | Defect | HW | Oxi module | Pinch valve tube |  | Failing Oximetry QC’s and pressure tests. | Failing Oximetry QC’s and pressure tests. | Opened sensor bay, fluid under pinch valve tubing. Replaced pinch valve tubing, now ok.
No | Visit | Complaint | RKOR | pO2 QC3 solution out of range. | . | Service mode check => OK pO2 complet test => OK Gasket module clean. Sensor cassette Ref. air clean. Solution pack replace. QC check => OK. | SWO-395303 | Assisted Replacement | SP | SP | pO2 | QC C | pO2 QC3 solution out of range. | . | Service mode check => OK pO2 complet test => OK Gasket module clean. Sensor cassette Ref. air clean. Solution pack replace. QC check => OK.
Yes | Visit | Complaint | RKOR | pO2 QC error | . | Gasket module clean. Service mode check, SP replace. SC clean. | SWO-398301 | Re classified RMED | same day upSWO |  |  |  | pO2 QC error | . | Gasket module clean. Service mode check, SP replace. SC clean.
No | Visit | Complaint | RDEU | 31.07.2018 Frau Geisler 052177279211 meldet, dass die festen Zielwerte am ABL nicht funktionieren. 02.08.2018 Frau Geisler 052177279211 fragt wann der Rückruf durch den Techniker erfolgt. | WG Fixe Zielwerte | Gerät überprüft , Roma.ini geprüft. Fixe Zielwerte sind bereits eingetragen. Haken für Rilibäk ist auch gesetzt. | SWO-396533 | No fault found | SW | RMED SW | Roma.ini | fixed assigned values | 31.07.2018 Ms. Geisler 052177279211 reports that the fixed target values on the ABL do not work. 02.08.2018 Ms. Geisler 052177279211 asks when the callback is made by the technician. | WG Fixed target values | Device checked, Roma.ini checked. Fixed target values are already entered. Hook for Rilibäk is also set.
No | Visit | Complaint | RIND | VERIFICATION CHECKED FAILED | verification check fail replace SC | Zeroing is correct but calibration is failed, check this SC & SP on another working analyzer calibration passed . | SWO-395159 | [5 empty] | VERIFICATION CHECKED FAILED | verification check fail replace SC | Zeroing is correct but calibration is failed, check this SC & SP on another working analyzer calibration passed .
No | Visit | Complaint | RDEU | tHb-Kal meldete mögliche Verstopfung | tHb-Kal wiederholt, i.O. | Berichtsprotokoll und QC-Plan angepasst, Touch Screen kalibriert anschließend beim Neustart Coox-Board nicht erkannt, Gerät neu gestartet, Datenbank zerstört, Gerät neu gestartet, SC neu initialisiert und Konfiguration erneut vorgenommen... | SWO-395271 | [5 empty] | tHb-Kal reported possible constipation | tHb-Kal wiederholt, i.O. | report protocol and QC plan, Touch Screen then calibrates Coox board undetected when restarting, restarted device, destroys database, restarts device, re-initializes SC, and reconfigures...
No | Visit | Complaint | RCHN | 进样口支架断裂 | 进样口支架断裂 借903-584 一个 | 更换903-584一个正常。 by Andrew | SWO-395052 | Defect | HW | Aspiration Module | Tube holder |  | The inlet bracket is broken | The inlet bracket is broken, borrowed 903-584. | Replace the 903-584 one normal. by Andrew
No | Visit | Complaint | RCHN | 定标终止报错4036 | 定标终止报错4036 借944-394（14500）一张 |  | SWO-394969 | Defect | SW | Analyzer Control Module | Analyzer SW | 4036 | Bid Termination Error 4036 | Bid termination error 4036 Borrow 944-394 (14500) one
No | Visit | Complaint | RAME | Georgene states the instrument is leaking from side of analyzer | Leaking from measuring chamber front trough. | Found sealing ring for pH was half on causing it to leak. Once I reseat the ring no more leaking. While onsite inlet probe was bent really bad. Upon speaking With Jamie, we will cover the replacement of probe this time but next customer will get charge. She said she will order some for their inventory | SWO-395498 | Defect | Electrode | Electrode | Sealing ring | pH | Georgene states the instrument is leaking from side of analyzer | Leaking from measuring chamber front trough. | Found sealing ring for pH was half on causing it to leak. Once I reseat the ring no more leaking. While onsite inlet probe was bent really bad. Upon speaking With Jamie, we will cover the replacement of probe this time but next customer will get charge. She said she will order some for their inventory
No | Visit | Complaint | RSCH | Re. Air Ventil | Luft im System Fehler 500, 509, 508 | Ventile pH $BG-Modul ausgebaut Air Membrane gerissen, ersetzt i.o Kal. und QC Level 1 i.o System i.o | SWO-395283 | Defect | HW | pH/BG Module | Valve, membrane |  | Re. Air Valve | Air in the system Error 500, 509, 508 | Valves pH $BG module removed Air membrane ripped, replaces i.o Kal. and QC Level 1 i.o System i.o
No | Visit | Complaint | RAME | Crea results below reportable ranges on level 1 Contact: Arthur Cassidy - 617-7327403 | Bottom of both Crea membranes are pink. | Cleaned and replaced the crea membranes and sealing rings. | SWO-395571 | Assisted Replacement | Electrode | Membrane | Crea |  | Crea results below reportable ranges on level 1 Contact: Arthur Cassidy - 617-7327403 | Bottom of both Crea membranes are pink. | Cleaned and replaced the crea membranes and sealing rings.
No | Visit | Complaint | RKK | バーコードリーダー 動作不良 | バーコードリーダー 動作不良 レーザーがすぐに消えてしまう | バーコードリーダー交換 | SWO-397605 | Defect | HW | Basic Unit | barcode reader |  | Barcode reader malfunction | Barcode reader malfunction The laser disappears quickly. | Bar code reader exchange
No | Visit | Complaint | RIND | Na measurement unstable | Checked patient data logs. | It is observed that Na measurement is now stable after change of liquid heparin lot in the ICU. | SWO-395635 | No fault found | No fault found | No fault found | No fault found |  | Na measurement unstable | Checked patient data logs. | It is observed that Na measurement is now stable after change of liquid heparin lot in the ICU.
Yes | 1 | Y | RPAC | Blank | Prepare and Setup analyser in readiness for install. |  | D4BD609CF848B3D1 | [5 empty] | Blank | Prepare and Setup analyser in readiness for install.
No | Visit | Complaint | RAME | Screen blank | Screen blank - Cycled power and no response. Monitor working on other analyzers. | Ordered replacement DMS | SWO-395585 | Revisit required | HW | DMS_PC | CPU |  | Screen blank | Screen blank - Cycled power and no response. Monitor working on other analyzers. | Ordered replacement DMS
No | Visit | Complaint | RIND | storage door stucking, not opening properly. | Storage door not opening properly. | checked and found that the corner of storage door is broken which help in sliding the door. Need to replace the Storage door. | SWO-395078 | Defect | HW | Cover and Chassis | Storage door |  | storage door stucking, not opening properly. | Storage door not opening properly. | checked and found that the corner of storage door is broken which help in sliding the door. Need to replace the Storage door.
No | Visit | Complaint | RIND | customer calibrated new lot then its failed with error optical intensity out of range. | checked the analyzer and found optical unit failure error while calibrated new lot. | checked the analyzer and performed system check and its showing ok . but optical intensity is near to higher side 8 .we need optical unit for this case. | SWO-395051 | Defect | HW | Process Module | Optical Unit |  | customer calibrated new lot then its failed with error optical intensity out of range. | checked the analyzer and found optical unit failure error while calibrated new lot. | checked the analyzer and performed system check and its showing ok . but optical intensity is near to higher side 8 .we need optical unit for this case.
No | 61 | Y | RPNZ | Blank | Solution pack replacement due to inhomogeneous errors. |  | D4C0134FE451A721 | Assisted Replacement | SP | SP | FT |  | Blank | Solution pack replacement due to inhomogeneous errors.
Yes | 61 | Y | RPNZ | Blank | Sensor cassette replaced due to (378) pCO2 calibration sensitivity below range |  | D4BE9ED76E3E2431 | Re classified RMED | same day upSWO |  |  |  | Blank | Sensor cassette replaced due to (378) pCO2 calibration sensitivity below range
No | Visit | Complaint | RAME | LEAK DETECTED IN ANALYZER | 8-3-18 0963-LEAK DETECTED IN ANALYZER. I WORKED WITH PAULA OVER THE PHONE. SHE CAME IN THIS MORNING AND THE ANALYZER WAS GIVING THE LEAK ERROR. STRIED INLET GASKET AND CONNECTOR. SHE TRIED SEVERAL SOLUTIONS PACK AND SENSOR CASSETTES. I HAD HER SWAP BOTH WITH A 90 THAT WAS WORKING AND STILL SAME ERROR. 8-6-18 ON SITE. DID NOT SEE ANY LEAKS. PUT ON PM KIT. I USED 2 KITS. I PUT A HOLE IN THE WASTE TUBE. REPLACED SENSOR CASSETTE WITH CUSTOMER STOCK. WILL LEAVE OPEN AND CHECK WITH PAULA TOMORROW AFTER IT DOES THE MIDNIGHT CHECK. | 8 $6 $18 PUT ON YEARLY SERVICE KIT AND WASTE CONNECTOR. 8 $7 $18 CHECKED WITH PAULA AND THE ANALYZER IS STILL GREEN. WILL CLOSE SWO. | SWO-395252 | Defect | HW | Basic Unit | PM Kit |  | LEAK DETECTED IN ANALYZER | 8-3-18 0963-LEAK DETECTED IN ANALYZER. I WORKED WITH PAULA OVER THE PHONE. SHE CAME IN THIS MORNING AND THE ANALYZER WAS GIVING THE LEAK ERROR. STRIED INLET GASKET AND CONNECTOR. SHE TRIED SEVERAL SOLUTIONS PACK AND SENSOR CASSETTES. I HAD HER SWAP BOTH WITH A 90 THAT WAS WORKING AND STILL SAME ERROR. 8-6-18 ON SITE. DID NOT SEE ANY LEAKS. PUT ON PM KIT. I USED 2 KITS. I PUT A HOLE IN THE WASTE TUBE. REPLACED SENSOR CASSETTE WITH CUSTOMER STOCK. WILL LEAVE OPEN AND CHECK WITH PAULA TOMORROW AFTER IT DOES THE MIDNIGHT CHECK. | 8 $6 $18 PUT ON YEARLY SERVICE KIT AND WASTE CONNECTOR. 8 $7 $18 CHECKED WITH PAULA AND THE ANALYZER IS STILL GREEN. WILL CLOSE SWO.
No | Visit | Complaint | RDEU | Fr. Pigulla bemängelt das nicht alle im Haus verwendeten Pat.ID Barcodes erkannt werden | Überprüfung der Barcodeerkenneung | Barcode mit führenden a o. d werden lt. Aussage Laborpersonal fehlerfrei erkannt, nur der auf der Stat. erstellt nicht ->Test ja ist so; Haus EDV versucht Barcodebezeichnung zur ermitteln ->Stat. Barcode= 2aus5, für LaborBarcode = Codebar; entsprechende Änderungen in Roma.ini neben den Standarteinstellungen vorgenommen, übrige Eintragungen entfernt, gesichert; Gerät Netz aus $ ein Test i.O.; Test: aller bei Kunden vorliegenden Pat ID Barcodetypen: werden erkannt und übertragen ->Labor MA beobachten Geräteverhalten und melden sich ggf. | SWO-395935 | Setup | SW | RMED SW | Roma.ini | barcode | Ms. Pigulla complains that not all Pat.ID barcodes used in the house are detected | Checking barcode verification | Barcode with leading a o. d are detected correctly according to statement laboratory personnel, only the one on the stat. does not create ->test yes is so; House EDP tries to determine barcode name ->Stat. Barcode= 2out5, for LaboratoryBarcode = Codebar; corresponding changes made in Roma.ini in addition to the default settings, other entries removed, secured; Device network from a test i.O.; Test: of all Pat ID barcode types available to customers: are detected and transmitted ->Laboratory MA observe device behavior and report if necessary.
No | Visit | Complaint | RKOR | System temperature out of range / Intervention Required screen and Start up screen are repeated. | Re-insert the sensor cassette Check the temperature in service mode Replace the sensor cassette | Replace the sensor cassette | SWO-398376 | Assisted Replacement | SC | SC | Thermistor |  | System temperature out of range / Intervention Required screen and Start up screen are repeated. | Re-insert the sensor cassette Check the temperature in service mode Replace the sensor cassette | Replace the sensor cassette
No | Visit | Complaint | RIND | Verification check failed | Verification check failed | Replace sensor board ,check zeroing , O2falctuation in zeroing is unsatable , need to change singal data board. | SWO-398279 | [5 empty] | Verification check failed | Verification check failed | Replace sensor board ,check zeroing , O2falctuation in zeroing is unsatable , need to change singal data board.
No | Visit | Complaint | RCHN | LAC定标失败 | LAC定标失败，借903-581一个 | 更换903-581备件后修复 | SWO-395086 | Defect | HW | Basic Unit | PCB, Sensor interface | Lac | LAC calibration failure | LAC bid failed, borrow 903-581 one | Repair after replacement of 903-581 spare parts
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | lac定标失败，报错1310 | 更换测试卡后修复 | SWO-395669 | Re classified RMED | same day upSWO |  |  |  | Test card, reagent package failure | Lac bid failed, error 1310 | Repair after replacing test card
No | Visit | Complaint | RDEU | Gerät meldet kein Akku installiert | Gerät meldet das Kein Akku vorhanden ist | Gerät überprüft, Akku ausgebaut, kontrolliert und erneuert | SWO-395754 | [5 empty] | Device does not report battery installed | Device reports the no battery is present | Device checked, battery removed, controlled and renewed
No | Visit | Complaint | RDEU | Fernwartung nicht voll funktionsfähig | Fehlersuche --> Analysatoreinstellungen überprüft | Analysatoreinstellungen überprüft, konfiguriert und über die Ferneartung getestet, System und Funktion i.O. | SWO-395973 | Revisit required | SW | RMED SW | 6.18 |  | Remote maintenance not fully functional | Troubleshooting --> Analyzer settings checked | Analyzer settings checked, configured and tested remotely, system and function i.O.
No | Visit | Complaint | RFRA | Drive des contrôles et des mesures sur iono | Echange de l'électrode de référence | Echange de l'électrode de référence | SWO-395789 | Assisted Replacement | Electrode | Electrode | Reference |  | Drive des controls et des mesures sur iono | Reference electrode exchange | Reference electrode exchange
No | Visit | Complaint | RAME | Neon lamp error | The customer was getting a neon lamp error | I checked the lamp and ran the complete test. I adj the H lamp and ran cal's | SWO-396583 | Adjustment | HW | Oxi module | Lamp unit |  | Neon lamp error | The customer was getting a neon lamp error | I checked the lamp and ran the complete test. I adj the H lamp and ran cal's
No | Visit | Complaint | RAME | Screen blank | Analyzer screen not working at all. | Replaced analyzer DMS module. Analyzer now operating properly. | SWO-396312 | Defect | HW | DMS_PC | CPU |  | Screen blank | Analyzer screen not working at all. | Replaced analyzer DMS module. Analyzer now operating properly.
No | Visit | Complaint | RPOL | Nie można usunąć pakietu roztworu podczas wymiany | Testy serwisowe | Wymiana silnika odpowiedzialnego za blokowanie i zwalnianie pakietu roztworów. Testy serwisowe. | SWO-395888 | Defect | HW | Basic Unit | Actuator unit |  | The solution package cannot be removed during replacement | Service Tests | Replacement of the engine responsible for blocking and releasing the solution package. Service tests.
No | Visit | Complaint | RDAN | Frederik Boe Hansen skriver følgende: ABL'en har egentlig drillet lidt siden vi fik den, men den har selv kunnet klare ærterne - indimellem med lidt hjælp. Den har hyppigt meldt fejl på føtal hæmoglobin, da QC-plot var fejlet. Efter noget tid har den hyppigt meldt fejl på hb også. Som regel har vi ved en rensning og/eller kalibrering og/eller et par QC-plot fået maskinen til at melde alle analyseresultater OK. Siden seneste væskepakke og kassette blev isat har den dog meldt konstant fejl på føtal hæmoglobin, hæmoglobin og nu også saturation, hvoraf sidstnævnte desværre er et af de parametre, som vi er meget interesseret i. Vi har løbende forsøgt at køre et par QC-plot, kalibrere og rensninger, men det er ikke lykkedes os at få maskinen til at melde grøn /klar på de førnævnte parametre. | Afprøvet og undersøgt. | OXI ventil udskiftet, cuvette renset i protein remover, slutafprøvet. | SWO-396575 | Cleaned | Fluidic pathway | Fluidic pathway | Oxi module |  | Frederik Boe Hansen writes the following: The Abbey has actually teaped a little since we got it, but it has been able to cope with the vents itself-sometimes with a little help. It has frequently reported errors in foetal hemoglobin, as the QC plot was failed. After some time it has frequently reported bugs on HB too. As a rule, we have, by cleaning and/or calibration and/or a few QC plots, been given the machine to report all analytical results OK. Since the last liquid package and cassette was inserted, it has reported constant errors on fetal hemoglobin, haemoglobin and now also saturation, the latter unfortunately being one of the parameters in which we are very interested. We have continually attempted to run a few QC plots, Calibration and cleanings, but we have not managed to get the machine to report Green/clear on the aforementioned parameters. | Tested and investigated. | OXI valve replaced, cuvette cleaned in protein remover, final tested.
No | Visit | Complaint | RDEU | 06.08.2018; Fr. Renken; Tel.: 030400372391; Systemtemperatur außer Bereich - mit Schw. Eva telefoniert und erklärt das bei zu warmer Umgebung das Gerät nicht mehr zuverlässig funktioniert(siehe Handbuch S.341), Kühlung unbedingt empfohlen 08.08.18: Hr. Schwerdt meldet erneut Probleme mit Temperatur - Nachfrage auf Station ergab Thermistorfehler trotz neuer Kassette und LP | Fehlersuche, Substarttemp. außer Bereich, Kassette ist neu, LP ebenfalls, Service aufgerufen und div. Test durchgeführt, Thermistor mehrfach kalibriert - sporadisch scheitert dieses Kalibrierung wenn man den Service verlässt, Kaltstart durchgeführt, Startphase scheitert wiederholt mit Thermistorfehler | Sensor erneut ersetzt und Startphase abgewartet, Kal, tHb Kal und QC | SWO-397050 | Assisted Replacement | SC | SC | Thermistor |  | 06.08.2018; Ms. Renken; Tel.: 030400372391; System temperature out of range - calls with Schw. Eva and explains that if the environment is too warm, the device no longer functions reliably (see manual p.341), cooling is strongly recommended 08.08.18: Mr. Schwerdt reports problems with temperature again - demand on station revealed thermistor error despite new cassette and LP | Troubleshooting, SubstartTemp. except area, cassette is new, LP also called service and various test performed, thermistor calibrated several times - sporadically this calibration fails when you leave the service, cold start is performed, start phase fails repeatedly with Thermistor error | Sensor replaced again and start-up phase awaited, Kal, tHb Kal and QC
Yes | Visit | Complaint | RSAF | Intermittent Qc level B reading lower than statistical range. | Investigated all possible route causes and ran sensor simulator test, all ok. Electronically analyzer checks out. | Flushed the entire sample path and sensor cassette but did not correct problem, Installed new sensor cassette as the Status values for pC02 and other parameters on calibrations was low. New sensor cassette and level B Qc is is on the mean for pC02. | SWO-397146 | [5 empty] | Intermittent Qc level B reading lower than statistical range. | Investigated all possible route causes and ran sensor simulator test, all ok. Electronically analyzer checks out. | Flushed the entire sample path and sensor cassette but did not correct problem, Installed new sensor cassette as the Status values for pC02 and other parameters on calibrations was low. New sensor cassette and level B Qc is is on the mean for pC02.
No | Visit | Complaint | RLTD | Natalie 02088695402 | Battery icon showing zero % battery | Discharge battery for 15 minuets removed battery and reset. Reinstalled and battery icon now showing battery charging %. | SWO-396812 | Adjustment | HW | Battery module | Battery module | 0% | Natalie 02088695402 | Battery icon showing zero % battery | Discharge battery for 15 minuets removed battery and reset. Reinstalled and battery icon now showing battery charging %.
No | Visit | Complaint | RAME | Customer says the sensor cassette wouldn’t pass system cycles because of PO2. Now they are getting temperature errors. | Customer says new sensor cassette won’t pass system cycles for PO2. Now getting temperature errors. Upon further investigation, found QC material on the backside of the sensor cassette and on the sensor connector board. | Replaced sensor connector board. Submitted credit claim for the new sensor cassette. Replaced sensor cassette with customer supply. Performed zero check. | SWO-397327 | [5 empty] | Customer says the sensor cassette wouldn’t pass system cycles because of PO2. Now they are getting temperature errors. | Customer says new sensor cassette won’t pass system cycles for PO2. Now getting temperature errors. Upon further investigation, found QC material on the backside of the sensor cassette and on the sensor connector board. | Replaced sensor connector board. Submitted credit claim for the new sensor cassette. Replaced sensor cassette with customer supply. Performed zero check.
No | Visit | Complaint | RCHN | 误操作导致定标终止 | 误操作导致定标终止，借944-220（14429）一张 | 重新定标后修复 by Andrew | SWO-395132 | User error | Consumables | Test Kit | Cal |  | Misoperation causes bid termination | Misoperation caused the calibration to terminate, with a 944-220 (14429) | Fix by Andrew after re-calibration
No | Visit | Complaint | RDEU | Fernwartung nicht voll funktionsfähig | Nachtrag zu SWO-395973 --> erneut ist die Fernwartung nur eingeschränkt $ nicht möglich --> Windows-Einstellungen überprüft --> Windows-Firewall aktiviert sich nach Neustart wieder, obwohl diese vorher deaktiviert wurde --> Problem hängt mit Analysator-Software Version 6.18 zusammen --> diese Einstellungen wurden zusammen mit Michael Kozelj (telefonisch) durchgegangen und konfiguriert!!! | Windows-Einstellungen (Firewall, Inbound Regeln (NetOP)) überprüft und abgeändert, am Analysator mehrmals Neustart durchgeführt und Einstellungen erneut überprüft und getestet, | SWO-396500 | Defect | SW | RMED SW | 6.18 |  | Remote maintenance not fully functional | Addendum to SWO-395973 --> again remote maintenance is only limited --> Windows settings checked --> Windows Firewall activates again after reboot, even though it was previously disabled --> Problem related to analyzer software version 6.18 --> these settings were passed through and configured together with Michael Kozelj (by phone!!! | Windows settings (Firewall, Inbound Rules (NetOP)) checked and modified, rebooted on the analyzer several times, and re-checked and tested settings,
No | Visit | Complaint | RKK | pH/電解質項目QC範囲外。Refメンブラン交換後も変わらず。Ref電極交換。 | pH $電解質項目QC範囲外。Refメンブラン交換後も変わらず。 | Ref電極交換。 | SWO-403547 | Assisted Replacement | Electrode | Electrode | Reference | CAPA-00733 | PH/electrolyte item QC out of range. Ref Menbran remains the same after the exchange. Ref electrode replacement. | pH / electrolyte item QC out of range. Ref Menbran remains the same after the exchange. | Ref electrode replacement.
No | Visit | Complaint | RKK | 夜中のクリーニングのキャルでpH/電解質項目でドリフトする。Ref電極交換。 | 夜中のクリーニングのキャルでpH $電解質項目でドリフトする。 | Ref電極交換。 | SWO-403553 | Assisted Replacement | Electrode | Electrode | Reference | CAPA-00733 | Drift in the pH/electrolyte item in the car of cleaning in the middle of the night. Ref electrode replacement. | Drift in the pH / electrolyte item in the car of cleaning in the middle of the night. | Ref electrode replacement.
No | Visit | Complaint | RKOR | Data transmisson error | IP address, Setting cechk Remove and add the new connetion 1 Change the IP address | Remove and add the new connetion 1 Change the IP address | SWO-398380 | Setup | SW | Analyzer Control Module | Analyzer Control Module |  | Data transmisson error | IP address, Setting cechk Remove and add the new connetion 1 Change the IP address | Remove and add the new connetion 1 Change the IP address
No | Visit | Complaint | RAME | Time reverting back 1 hour cannot adjust in the analyzer | Time shifts in analyzer after 3.4 upgrade Time server sync not holding correct time Rrmeddstservice reactivated to auto and service started after upgrade | Set service to stop and manual Re setup time server | SWO-397585 | Setup | SW | SW | Time server |  | Time reverting back 1 hour cannot adjust in the analyzer | Time shifts in analyzer after 3.4 upgrade Time server sync not holding correct time Rrmeddstservice reactivated to auto and service started after upgrade | Set service to stop and manual Re setup time server
No | Visit | Complaint | RLTD | Mark Harper Tel: 01895 828757 | Failing pressure test 1063 and leak detected. Customer has replaced inlet gasket and probe and fitted new consumable on. | Pressure test system and found tubing valve leaking. Removed wet section and replaced valve tube. Run all service test all passed ok. Calibration and QC passing apart from lactate failing due to response error. | SWO-397417 | Defect | HW | Oxi module | Pinch valve tube |  | Mark Harper Tel: 01895 828757 | Failing pressure test 1063 and leak detected. Customer has replaced inlet gasket and probe and fitted new consumable on. | Pressure test system and found tubing valve leaking. Removed wet section and replaced valve tube. Run all service test all passed ok. Calibration and QC passing apart from lactate failing due to response error.
Yes | 1 | Y | RPAC | Blank | ABL90 Flex installation |  | D4C116265A1CBB8A | Re classified RMED | planned SWO |  |  |  | Blank | ABL90 Flex installation
No | Visit | Complaint | RLTD | leak detected error code 1184 / steven devroe - 01642 285381 | Site visit to investigate and repair leak detected error. | Arrived on site and attended analyser. Found it in intervention required screen. Analyser showing leak detected. Entered service and found vacuum pressure test failed at 0.96 Powered down analyser and removed wet section. Found evidence of leaked fluid under pinch valve $tube. Cleaned leaked fluid and replaced pinch tube as this was clearly perforated as shown by the crystallised solution around the leaked area. Checked pinch valve and this all looked in good condition. Refitted all parts and powered up the analyser. Analyser started 55 minute startup. Analyser quickly went to 15 minute startup. When analyser got to 5 seconds, it was showing an inhomogeneous gas error and tried to rerun the failed qc. After retry the analyser requested a flush. Removed gasket and replaced this with new. Checked probe and this looked curved. Replaced this with new as a precaution. Refitted parts and ran flush procedure. Cleaned flow selector window as there was blood staining around it and the optic. Refitted all parts and retested. Analyser now passing. Rerun the failed qc and this now passes. Noticed battery was showing 0%. Powered down the analyser, and reset battery using internal jumper. Refitted and repowered the analyser. Battery now displaying correct value. Analyser running well again. | SWO-397148 | Defect | HW | Oxi module | Pinch valve tube |  | leak detected error code 1184 / steven devroe - 01642 285381 | Site visit to investigate and repair leak detected error. | Arrived on site and attended analyser. Found it in intervention required screen. Analyser showing leak detected. Entered service and found vacuum pressure test failed at 0.96 Powered down analyser and removed wet section. Found evidence of leaked fluid under pinch valve $tube. Cleaned leaked fluid and replaced pinch tube as this was clearly perforated as shown by the crystallised solution around the leaked area. Checked pinch valve and this all looked in good condition. Refitted all parts and powered up the analyser. Analyser started 55 minute startup. Analyser quickly went to 15 minute startup. When analyser got to 5 seconds, it was showing an inhomogeneous gas error and tried to rerun the failed qc. After retry the analyser requested a flush. Removed gasket and replaced this with new. Checked probe and this looked curved. Replaced this with new as a precaution. Refitted parts and ran flush procedure. Cleaned flow selector window as there was blood staining around it and the optic. Refitted all parts and retested. Analyser now passing. Rerun the failed qc and this now passes. Noticed battery was showing 0%. Powered down the analyser, and reset battery using internal jumper. Refitted and repowered the analyser. Battery now displaying correct value. Analyser running well again.
No | Visit | Complaint | RAME | Customer called stating pCO2 will not calibrate. | Found that pCO2 is out on calibration for low sensitivity. | Did fluidics checks, and replaced SC on analyzer. | SWO-397433 | Assisted Replacement | SC | SC | pCO2 | Cal | Customer called stating pCO2 will not calibrate. | Found that pCO2 is out on calibration for low sensitivity. | Did fluidics checks, and replaced SC on analyzer.
No | Visit | Complaint | RSAF | Customer called and asked thast we come and instsall a new inlet probe as there exsisting one is badly bent and cant be straightended. Martin was informed and gave the go ahead for repair. |  |  | SWO-397155 | Defect | HW | Auto Inlet module | Probe bend |  | Customer called and asked thast we come and instsall a new inlet probe as there exsisting one is badly bent and cant be straightended. Martin was informed and gave the go ahead for repair.
No | Visit | Complaint | RIND | Lactate calibration failed. | Lactate calibration failure. | Replace the SC80 with the new one,the calibration take place around 50mins but still Lactate failed again,Flush the SC80 with NS & again reinstall it ,but still lactate was failed again,active the lactate parameter but after calibration Lactate failed again,Run the zeroing check,replace the sensor board,But still lactate not passed,Replace the SC80 with the new one,run the calibration but pco2,Na,Lac got failed again,again replace the SC80 & after successful calibration instrument become ready. | SWO-398252 | [5 empty] | Lactate calibration failed. | Lactate calibration failure. | Replace the SC80 with the new one,the calibration take place around 50mins but still Lactate failed again,Flush the SC80 with NS & again reinstall it ,but still lactate was failed again,active the lactate parameter but after calibration Lactate failed again,Run the zeroing check,replace the sensor board,But still lactate not passed,Replace the SC80 with the new one,run the calibration but pco2,Na,Lac got failed again,again replace the SC80 & after successful calibration instrument become ready.
No | Visit | Complaint | RIND | Parameter failure | CA is failure. | check zeroing , O2 is unstable, flus out the cassette ,re-install the cassette ,run 2point calibration ,calibration passed | SWO-398284 | [5 empty] | Parameter failure | CA is failure. | check zeroing , O2 is unstable, flus out the cassette ,re-install the cassette ,run 2point calibration ,calibration passed
No | Visit | Complaint | RCHN | 报错1999 | 报错1999 | 清洁孵玉盘后修复 by Andrew | SWO-397020 | Cleaned | HW | Incubator Module | PCB, Incubator |  | Error 1999 | Error 1999 | Repair by Andrew after cleaning hatching jade plate
No | Visit | Complaint | RDEU | 09.08.2018; Sr. Janine; Tel.: 021318882018; Probenproblem | Probenproblem bei allen Messungen bei den Parametern pH + Abgeleitete und Glu. 0476 pH Messung instabil und 0500 Spüllösung vom pH $BG Modul erkannt, wenn nicht erwartet. Glu Empfindlichkeit „- - - - -„. | Air Ventil Membran im pH $BG Moful durchschlagen - ersetzt, Ventile gereinigt und gefettet. Glu Empfindlichkeit weiterhin nicht ok. Glu membraniert, mehrfach Kalibriert und QC’s durchgeführt. QC Lvl bei Lac zu niedrig, Lac Membran erneuert, mehrfach Kalibriert und QC wiederholt. | SWO-397960 | Defect | HW | pH/BG Module | Valve, membrane |  | 09.08.2018; Sr. Janine; Tel.: 021318882018; Probenproblem | Sample problem with all measurements with the parameters pH + Derived and Glu. 0476 pH measurement unstable and 0500 rinsing solution detected by pH /BG module if not expected. Glu sensitivity "- - - - -". | Air valve membrane in pH $BG Moful - replaced, valves cleaned and greased. Glu sensitivity still not ok. Glu membraned, multiple calibrated and QC's performed. QC Lvl at Lac too low, Lac membrane renewed, multiple calibrated and QC repeated.
No | Visit | Complaint | RAME | QC is being aborted | QC began aborting again. Ampoules get opened but not aspirated. | Checked carousel and holder function. Replaced carousel from customer stock. Verified and re-adjusted autocheck module and inlet module. Made sure there was no obstruction for inlet probe when aspirating QC. Ran qc afterwards and all 3 levels ran without error. | SWO-397462 | Defect | HW | Autocheck module | Carousel |  | QC is being aborted | QC began aborting again. Ampoules get opened but not aspirated. | Checked carousel and holder function. Replaced carousel from customer stock. Verified and re-adjusted autocheck module and inlet module. Made sure there was no obstruction for inlet probe when aspirating QC. Ran qc afterwards and all 3 levels ran without error.
No | Visit | Complaint | RIND | Calibration error | Checked the data logs po2 query in all the sample in new machine cleaned po2 electrode and chamber performed zero current calibration & 2-point calibration checked the calibration po2 zero value for stable for few calibration after few calibration po2 zero value was decreasing | Replaced po2 membrane | SWO-397997 | Assisted Replacement | Electrode | Membrane | pO2 |  | Calibration error | Checked the data logs po2 query in all the sample in new machine cleaned po2 electrode and chamber performed zero current calibration & 2-point calibration checked the calibration po2 zero value for stable for few calibration after few calibration po2 zero value was decreasing | Replaced po2 membrane
No | 7 | Y | RCZE | Blank | Mineraly nepravidelne driftuju, tiez v QC 1,2,3 neprechadzaju. Nepomohlo ani premembranovanie jednotlivych mineralov a ref elektrody. |  | D4C0FD8A5A4403FB | Assisted Replacement | Electrode | Electrode | Reference | CAPA-00733 | Blank | Mineraly irregular Driftuju, Tiez v QC 1, 2, 3 Neprechadzaju. It did not help even Premembranovanie jednotlivych Mineralov and ref electrodes.
No | Visit | Complaint | RCHN | 新机质控Lac不准确，报错219 | 新机质控Lac不准确，报错219,需要借件902-824一个 | 更换902-824备件后修复 by Andrew | SWO-395621 | Defect | HW | El/Met Module | El/Met Module |  | New machine quality control Lac inaccurate, reported error 219 | New machine quality control Lac is not accurate, 219 wrong, need to borrow 902-824 one | Repair by Andrew after replacement of 902-824 spare parts
Yes | 1 | Y | RPAC | Blank | ABL825 installation |  | D4C22980D0499FEB | Re classified RMED | installation report |  |  |  | Blank | ABL825 installation
Yes | Visit | Complaint | RHUN | 379-es hibakód GLu, Lac. Na érzékenysége néhe 95 alatt és kérdőjelezi. | Glu, Lac kalibrálási hiba | Glu, Lac elektróda és membrán csere. Na elektróda csere. Kalibrálás, QC mérés. | SWO-397315 | Re classified RMED | OOB |  |  |  | Error code 379 is GLu, Lac. Na's sensitivity is under 95 and question. | Glu, Lac kalibrálási hiba | Glu, Lac electrode and membrane replacement. Na electrode replacement. Calibration, QC measurement.
Yes | Visit | Complaint | RAME | Clot in unit | Clot in unit | Cleared clot $ performed calibration $ Lactate failed ordered sensor for Monday | SWO-398422 | [5 empty] | Clot in unit | Clot in unit | Cleared clot $ performed calibration $ Lactate failed ordered sensor for Monday
No | Visit | Complaint | RFRA | Appel du 10 août. Le client se plaint de remplacer la k7 de capteur trop souvent suite à des pb de PCo2. | Mise en place du sensor simulator pour vérifier les signaux électroniques. | PCO2, lac et créa en erreur sur le simulateur. Ajustement des contacts de la carte sensor board pour une bonne connexion avec la cassette sensor. Test simulateur ok Redémarrage. Remplacement du détenir manifold suite erreur sélecteur de flux. Redémarrage. Essais ok Réinstallation de la k7 sensor. Ok. Calibration et passage Cq ok. | SWO-398613 | Adjustment | HW | Basic Unit | Sensor contact |  | Appeal of August 10. The customer complains about replacing the sensor k7 too often following PCo2 pbs. | Setting up the sensor simulator to check electronic signals. | PCO2, lake and mis-created on the simulator. Adjusting the sensor board's contacts for a good connection with the sensor cassette. Simulator test ok Restart. Replacing the hold manifold following flow selector error. Restart. Tests ok Reinstallation of the k7 sensor. Ok. Calibration and passage Cq ok.
Yes | Visit | Complaint | RDEU | STK Prüfung | STK Prüfung fällig | STK Prüfung durchgeführt, Messung ok. | SWO-398418 | Re classified RMED | planned SWO |  |  | mail | STK Exam | STK exam due | STK test performed, measurement ok.
No | Visit | Complaint | RAME | ABL looses connection when moved | Suspect ini file needs update | Reloaded win 7 and changed ini file, power settings, radiance reconnect, and loaded new driver for wifi. | SWO-397630 | Setup | SW | Windows | Wifi |  | ABL looses connection when moved | Suspect ini file needs update | Reloaded win 7 and changed ini file, power settings, radiance reconnect, and loaded new driver for wifi.
No | Visit | Complaint | RSAF | Lac failing |  |  | SWO-397935 | Assisted Replacement | SC | SC | Lac | Sensitivity | Lac failing
Yes | 5 | Y | RPAC | Blank | sc failure |  | D4C2578CD05CB57F | [5 empty] | Blank | sc failure
Yes | Visit | Complaint | RSAF | Oxi liquid sensor failed to calibrate | Damaged LS during hemolyzer replacement | Replaced oxi liquid sensor, start up tests ok. | SWO-398151 | Re classified RMED | same day upSWO |  |  |  | Oxy Liquid sensor failed to calibrated | Damaged LS during hemolyzer replacement | Replaced Oxygen liquid Sensor, start up tests OK.
No | Visit | Complaint | RSAF | Hemolyzer to be replaced on receiving PO number, | Hemolyzer replaced by Joe from Medhold, and struggled with numerous errors after replacing Hemolyzer unit | Stripped and inspected all possible causes for error and correctly aligned actuator spindal shaft for solution pack lock, Oxi liquid sensor to be replaced as ribbon cable is faulty, | SWO-398590 | Defect | HW | Oxi module | Hemolyzer |  | Hemolyzer to be replaced on receiving PO number, | Hemolyzer replaced by Joe from Medhold, and struggled with numerous errors after replacing Hemolyzer unit | Stripped and inspected all possible causes for error and correctly aligned actuator spindal shaft for solution pack lock, Oxi liquid sensor to be replaced as ribbon cable is faulty,
Yes | Visit | Complaint | RDEU | STK Prüfung | STK Prüfung fällig | STK Prüfung durchgeführt, Messung ok. | SWO-398417 | Re classified RMED | planned SWO |  |  | mail | STK Exam | STK exam due | STK test performed, measurement ok.
No | Visit | Complaint | RCHN | 触摸屏故障 | 触摸屏故障，借552-039一个 | 更换552-039一个正常。 | SWO-397639 | [5 empty] | Touch screen failure | Touch screen fault, borrow 552-039 a | Replace 552-039 one normal.
No | Visit | Complaint | RCHN | po2零点校准失败 | po2零点校准失败，借903-334一个 | 更换903-334一个正常。 | SWO-397062 | [5 empty] | Po2 zero calibration failed | Po2 zero calibration failed, borrow 903-334 one | Replace 903-334 one normal.
No | Visit | Complaint | RCHN | 打印机打印不清晰 | 打印机打印不清晰 借903-001一个 | 更换了打印纸。备件没有更换。 | SWO-396966 | [5 empty] | Printer printing is not clear | Printer printing is not clear borrowing 903-001 one | The paper has been replaced. Spare parts have not been replaced.
No | Visit | Complaint | RCHN | 气泡探测定标失败，pco2，hct定标失败 | 气泡探测定标失败，pco2，hct定标失败，借902-972一个 | 更换902-972后修复 | SWO-397128 | [5 empty] | Bubble detection calibration failed, pco2, hct calibration failed | Bubble detection calibration failure, pco2, hct calibration failure, borrow 902-972 one | Repair after replacement 902-972
No | Visit | Complaint | RCHN | 泵定标出错 | 泵定标出错，借903-165和580-401各一个 | 更换903-165和580-401各一个正常。 | SWO-397087 | Defect | HW | Basic Unit | Pump motor |  | Pump calibration error | Pump calibration error, borrow 903-165 and 580-401 each one | Replace the 903-165 and 580-401 each one normal.
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 液体传输错误0508 | 更换试剂包正常。 | SWO-397881 | Assisted Replacement | SP | SP | FT | 508 | Test card, reagent package failure | Liquid transfer error 0508 | The replacement kit is OK.
No | Visit | Complaint | RESP | Medidas con interrogante en las hemoglobinas | Hemolizador defectuoso | Sustitucion del hemolizador y de tubos Chequeo por servicio Cal, qc y muestras | SWO-398055 | Defect | HW | Oxi module | Hemolyzer |  | Measurements with questioning in hemoglobins | Faulty hemolizer | Substitution of hemolizer and tubes Check for service Cal, qc and samples
Yes | Visit | Complaint | RCHN | 报错 1031 | 报错1031，借903-364一个。 | 更换903-364备件后修复 by Andrew | SWO-396443 | [5 empty] | Error 1031 | 1031, borrow 903-364 one. | Repair by Andrew after replacement of 903-364 spare parts
No | Visit | Complaint | RCHN | 报错5784 | 报错5784 | 调节进样门螺丝后修复 by jesson | SWO-397729 | Adjustment | HW | Cover and Chassis | Door, Inlet | 5784 | 5784 | 5784 | Repair by jesson after adjusting the screws in the sample door
No | Visit | Complaint | RAME | Potassium results from correlation too high | Reproduction of results not duplicating error. Running blood samples on two like analyzers result in matching results. Cause of compliant unknown. | None necessary | SWO-397633 | No fault found | No fault found | No fault found | No fault found |  | Potassium results from correlation too high | Reproduction of results not duplicating error. Running blood samples on two like analyzers result in matching results. Cause of compliant unknown. | None necessary
No | Visit | Complaint | RAME | SMART error on HD | SMART error on HD. Found header on HD corrupted. Unable to fix | Replaced hd. Reloaded windows 7 and system software. Reloaded all customer settings. | SWO-399084 | Defect | HW | DMS_PC | HDD |  | SMART error on HD | SMART error on HD. Found header on HD corrupted. Unable to fix | Replaced hd. Reloaded windows 7 and system software. Reloaded all customer settings.
No | Visit | Complaint | RAME | Inhomogeneous sample errors | Inhomogeneous sample errors | Replaced probe and inlet tubing. | SWO-399083 | Defect | HW | Inlet module | Probe |  | Inhomogeneous sample errors | Inhomogeneous sample errors | Replaced probe and inlet tubing.
No | Visit | Complaint | RKK | Naドリフト大及びGluアンステ | 代替pHモジュールと代替電解質モジュールの取り付けを実施しました。 引き上げ修理 | 代替pHモジュールと代替電解質モジュールの取り付けを実施しました。 引き上げ修理 | SWO-402204 | Defect | HW | El/Met Module | El/Met Module |  | Na Drift Large &amp; Glu Anste | The installation of an alternative pH module and an alternative electrolyte module was implemented. Raise repair | The installation of an alternative pH module and an alternative electrolyte module was implemented. Raise repair
No | Visit | Complaint | RIND | glucose parameter showing question mark in patient report. | checked the analyzer and found glucose membrane and lactate sensitivities out of range. | replaced glucose and lactate membrane and performed 2point calibration checked the 2point calibration results all parameters are in good working condition analyzer working fine. | SWO-397741 | Assisted Replacement | Electrode | Membrane | Glu | Glu/Lac | glucose parameter showing question mark in patient report. | checked the analyzer and found glucose membrane and lactate sensitivities out of range. | replaced glucose and lactate membrane and performed 2point calibration checked the 2point calibration results all parameters are in good working condition analyzer working fine.
No | Visit | Complaint | RIND | Lactate calibration failed. | Lactate calibration failure. | Flush SC80 with normal Saline water,Kept it on fridge,then after 30 min kept it on room temperature ,then install the cassette but after 70% during calibration it got failed again,Then again repeat the procedure 2times but every time it got failed,Need to replace the SC80. | SWO-398248 | [5 empty] | Lactate calibration failed. | Lactate calibration failure. | Flush SC80 with normal Saline water,Kept it on fridge,then after 30 min kept it on room temperature ,then install the cassette but after 70% during calibration it got failed again,Then again repeat the procedure 2times but every time it got failed,Need to replace the SC80.
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | PO2定标不过 0375 | 更换测试卡后修复 | SWO-400906 | [5 empty] | Test card, reagent package failure | PO2 calibration is no more than 0375 | Repair after replacing test card
No | Visit | Complaint | RCHN | 扫描器不亮 | 扫描器不亮,借903-029一个。 |  | SWO-398207 | Defect | HW | Storage Module | PCB, cup transport front |  | The scanner is not on | The scanner is not lit, borrowing 903-029 one.
No | Visit | Complaint | RCHN | 报错1225 | 报错1225，借903-364一个 | 更换903-364一个正常。 by jesson | SWO-397013 | Defect | HW | Process Module | Optical Unit | 1225 | 1225 error | 1225, loan 903-364 one | Replace 903-364 one normal. by jesson
No | Visit | Complaint | RAME | I.T settings configuration after HD change | I.T settings configuration after HD change. Worked with hospital I.t to get instrument back on their network | I.T settings configuration after HD change. Worked with hospital I.t to get instrument back on their network | SWO-399085 | Setup | His/Lis | His/Lis | His/Lis |  | I.T settings configuration after HD change | I.T settings configuration after HD change. Worked with hospital I.t to get instrument back on their network | I.T settings configuration after HD change. Worked with hospital I.t to get instrument back on their network
No | Visit | Complaint | RCHN | 经常报温度错误 | 经常报温度错误,借902-972一个 | 更换测试卡正常。by jesson | SWO-398255 | [5 empty] | Regularly report temperature errors | Frequently reported temperature errors, borrow 902-972 a | The replacement test card is OK. by jesson
No | Visit | Complaint | RCHN | 停电后无法传输历史数据 | 停电后无法传输历史数据,借910-342一个。 |  | SWO-398367 | Defect | HW | DMS_PC | CPU |  | Unable to transfer historical data after power outage | It is not possible to transmit historical data after the power failure, borrowing 910-342 one.
No | Visit | Complaint | RCHN | 报错1225 | 之前的光学单元 报1024.再加借903-364一个。 |  | SWO-398359 | Defect | HW | Process Module | Optical Unit | 1024 | 1225 error | The previous optical unit reported 1024. Add 903-364.
No | Visit | Complaint | RKK | 田代Drより時間外当番に連絡。朝から中検のNaと比較して10m㏖/L程度高く出ている修理依頼。 | QCレベル２でNa範囲外。 | Ref,Naメンブラン交換。 | SWO-398570 | Assisted Replacement | Electrode | Membrane | Reference |  | Dr. Tashiro contacted me on duty after hours. The repair request which has come out high about 10mmol / L compared with Na of the examination in the morning. | Na out of range at QC level 2. | Ref, Na men bran replacement.
No | Visit | Complaint | RIND | Application software error. | found application software DB file and .ini file got corrupted. | Reinstalled the application software and setup the report formats and other settings as per customer requirements.Did calibration and samples all results found within the range. found within the range. | SWO-401389 | Defect | SW | RMED SW | ini file |  | Application software error. | found application software DB file and .ini file got corrupted. | Reinstalled the application software and setup the report formats and other settings as per customer requirements. Did calibration and samples all results found within the range. found within the range.
Yes | Visit | Complaint | RAME | Analyzer receiving 581 errors in NICU. Replacing proactively. | Analyzer receiving 581 errors. | Replacing hemolyzer | SWO-398965 | [5 empty] | Analyzer receiving 581 errors in NICU. Replacing proactively. | Analyzer receiving 581 errors. | Replacing hemolyzer
No | Visit | Complaint | RFRA | Fuite détectée. Demande de remplacement du joint malgré joint d'introduction remplacé. | Fuite détectée. | Circuit fluidique bouché. Débouchage conduit gauche de la bay unit. Essais flush ok. Test service : vacuum Value = 1 remplacement aiguille + kit maintenance idem. Retour machine pour expertise. | SWO-398615 | Defect | HW | Basic Unit | PM Kit |  | Leak detected. Request to replace the joint despite the introduction seal replaced. | Leak detected. | Clogged fluid circuit. Uncorking leads left from the bay unit. Try flush ok. Test service: vacuum Value - 1 needle replacement - maintenance kit ditto. Back machine for expertise.
Yes | Visit | Complaint | RFRA | Fuite détectée. Demande de remplacement du joint malgré joint d'introduction remplacé. | Vaccum Value =1. | Vaccum Value =1. Pb d'étanchéité sur plusieurs joints d'introduction.: échange ok. Réglage sensor module $ bay unit. Simulator test ok mais Zpo2 énergy proche de 3. Remplacement de la carte sensor interface Pcb Réglage et essais ok. Redémarrage et test routine. | SWO-398616 | Re classified RMED | same day upSWO |  |  |  | Leak detected. Request to replace the joint despite the introduction seal replaced. | Vaccum Value =1. | Vaccum Value 1. Waterproofing pb on several introductory seals.: exchange ok. Sensor module setting $ bay unit. Simulator test ok but Zpo2 energy close to 3. Replacing the pcb interface sensor map Setting and testing ok. Restart and test routine.
No | Visit | Complaint | RAME | Too many recent failures. Would like FSE to check out | Flushed the analyzer, change inlet gasket with holder and SP X3 | In service found it failing pressure test Replaced inlet gasket with holder | SWO-398924 | Defect | HW | Inlet module | Inlet gasket |  | Too many recent failures. Would like FSE to check out | Flushed the analyzer, change inlet gasket with holder and SP X3 | In service found it failing pressure test Replaced inlet gasket with holder
Yes | 8 | Y | RPAC | Blank | inlet not closing - error 1190 |  | D4C5F8DA325CED1A | [5 empty] | Blank | inlet not closing - error 1190
No | Visit | Complaint | RLTD | Caller Name: ANYA Telephone No 01604 634 700 EXTN 5342 Prob / Mess LEAK DETECTED - ERROR 1342 - PLEASE CALL | On opening the sensor bay it could be seen that the oximetry pinch tube was leaking as there were crystals underneath it, also the inlet probe had a bend in it. | Carried out the routine service which was due on SWO:398680 | SWO-398681 | Defect | HW | Oxi module | Pinch valve tube | 1342 | Caller Name: ANYA Telephone No 01604 634 700 EXTN 5342 Prob / Mess LEAK DETECTED - ERROR 1342 - PLEASE CALL | On opening the sensor bay it could be seen that the oximetry pinch tube was leaking as there were crystals underneath it, also the inlet probe had a bend in it. | Carried out the routine service which was due on SWO:398680
No | 3 | Y | RPAC | Blank | Failed consumable |  | D4CAFCB90A3829CA | Assisted Replacement | SC | SC | unknown |  | Blank | Failed consumable
Yes | 3 | Y | RPAC | Blank | Replacement of failed SC |  | D4C78078A83EB587 | [5 empty] | Blank | Replacement of failed SC
No | Visit | Complaint | RAME | Unit will not clear intervention required screen | Intervention required message pressure error. | Tested fluidic operation plug between waste and Sc, plug between Sc and inlet probe. Replaced hemolyzer and inlet probe. Next issue was with vacuum error. Suspect issue with pinch valve or inlet gasket holder. Parts to be ordered. | SWO-400130 | Revisit required | HW | Oxi module | Pinch valve |  | Unit will not clear intervention required screen | Intervention required message pressure error. | Tested fluidic operation plug between waste and Sc, plug between Sc and inlet probe. Replaced hemolyzer and inlet probe. Next issue was with vacuum error. Suspect issue with pinch valve or inlet gasket holder. Parts to be ordered.
No | Visit | Complaint | RAME | 0581 on Patient Samples | Reviewed patient data logs and found multiple 581,292 and 688 errors. QC is passing fine. Ran service test and found all levels passed. | Removed pinch valve and found it to be stuck, replaced pinch valve as well as hemolyzer due to the distance to this facility 4.5 hours one way. Reassembled analyzer and ran THB cal, system cycles and manual QC, had customer test analyzer with some green top blood as well. | SWO-399052 | [5 empty] | 0581 on Patient Samples | Reviewed patient data logs and found multiple 581,292 and 688 errors. QC is passing fine. Ran service test and found all levels passed. | Removed pinch valve and found it to be stuck, replaced pinch valve as well as hemolyzer due to the distance to this facility 4.5 hours one way. Reassembled analyzer and ran THB cal, system cycles and manual QC, had customer test analyzer with some green top blood as well.
No | Visit | Complaint | RIND | Parameter failure | Checked and found sensor board is corroded and requires replacement. | Replaced sensor board (902-972). Performed zero adjustment successfully. 2 point calibration performed without errors. | SWO-399410 | [5 empty] | Parameter failure | Checked and found sensor board is corroded and requires replacement. | Replaced sensor board (902-972). Performed zero adjustment successfully. 2 point calibration performed without errors.
No | 8 | Y | RPAC | Blank | Touch screen |  | D4C63A1F3203EF73 | Defect | HW | DMS_PC | Touch screen |  | Blank | Touch screen
No | Visit | Complaint | RAME | QC is being aborted | Analyzer is intermittently aborting qc. | Replaced friction detector unit, inlet probe and inlet tube. Ran multiple levels of qc for diagnostic purpose. All passed without any abortion. | SWO-398972 | Defect | HW | Inlet module | Friction detection |  | QC is being aborted | Analyzer is intermittently aborting qc. | Replaced friction detector unit, inlet probe and inlet tube. Ran multiple levels of qc for diagnostic purpose. All passed without any abortion.
No | Visit | Complaint | RDEU | pH/Lyte-Drift- und QK-Fehler auffällig seit letzter Wartung | Div. Drift- und QK-Fehler bei pH $Lyten, Ref.-Elektrode ist frisch membraniert | Reinigungsarbeiten, Fehlersuche, Messkammern gereinigt, Ventilfunktionen von Fluidic- und MessModulen kontrolliert, Lösungstransport kontrolliert, Pumpenschläuche der Messmodule vom Kundenmaterial ersetzt, PumpenKal, LS-Justage, mehrere korrekte Kal abgewartet, QK ok | SWO-399257 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | pH/Lyte drift and QK errors conspicuous since last maintenance | Div. Drift and QK error in pH /Lyten, ref. electrode is freshly membraneed | Cleaning work, troubleshooting, measuring chambers cleaned, valve functions controlled by fluidic and measurement modules, solution transport controlled, pump hoses of the measuring modules replaced by customer material, pumpsKal, LS-adjustment, several correct scales awaiting, QK Ok
No | Visit | Complaint | RIND | Sample problem. | Checked the system for found problem with pump calibration.it Offset flow value is near to high. | Did calibration for inlet ,LS and pump, now values are fine,did 2 point calibration ,QC and samples all results found within the range. Now instrument working satisfactorily | SWO-401416 | Adjustment | HW | Basic Unit | LS |  | Sample problem. | Checked the system for found problem with pump calibration.it Offset flow value is near to high. | Did calibration for inlet ,LS and pump, now values are fine,did 2 point calibration ,QC and samples all results found within the range. Now instrument working satisfactorily
No | Visit | Complaint | RIND | customer calibrated new lot then its failed with error optical intensity out of range. | checked the analyzer and performed system check its failed optical intensity. | replaced new optical unit and calibrated factory reference and performed system check all are ok. performed new lot calibration now analyzer working fine . | SWO-400264 | Defect | HW | Process Module | Optical Unit |  | customer calibrated new lot then its failed with error optical intensity out of range. | checked the analyzer and performed system check its failed optical intensity. | replaced new optical unit and calibrated factory reference and performed system check all are ok. performed new lot calibration now analyzer working fine .
Yes | 6 | Y | RPAC | Blank | Installation |  | D4C858F01E629235 | [5 empty] | Blank | Installation
No | Visit | Complaint | RDEU | 14.08.2018; Fr. Moritz; Tel.: 06213836540; Touchscreen ist dunkel | Überprüfung des Analysators. | Analysator ausgeschaltet und neu gestartet. Alle Module erhalten Spannung. Analysator führt aber keine Aktivitäten durch, durch die Messeinheit, vermutlich ist die CPU Einheit defekt. Neue per Express bestellt, auch Netzteil und Display unit. | SWO-399492 | Revisit required | HW | DMS_PC | CPU |  | 14.08.2018; Ms. Moritz; Tel.: 06213836540; Touchscreen is dark | Checking the analyzer. | Analyzer turned off and restarted. All modules receive voltage. However, analyzer does not perform any activities through the measurement unit, presumably the CPU unit is defective. New ordered by express, also power supply and display unit.
No | Visit | Complaint | RDEU | 13.08., 12:30 Uhr, Frau Pauscmann, Tel.: 062150335589 OXI Probleme bei Kal. - kein Nullpunkt bei Kal. gefunden | Überprüfung des Analysators im Service, Oxi Complete Test durchgeführt. Das Ergebnis zeigt, dass die LED Peak intensity extrem hoch ist, 34812 pA, anstatt bis 300 pA. Auch die cavitation absorb ist außrt Bereich. | Tausch des Hemolyzer Unit. Nach Neustart und erfolgter Kalibrationen sind Werte in Ordnung. tHb Kalibration durchgeführt, in Ordnung. Einlasssonde krumm, erneuert, auch Einlassdichtung. | SWO-399491 | Defect | HW | Oxi module | Hemolyzer |  | 13.08., 12:30 p.m., Mrs. Pauscmann, Tel.: 062150335589 OXI Problems at Kal. - no zero point at Kal. Found | Verification of the analyzer performed in the service, Oxi Complete Test. The result shows that the LED peak intensity is extremely high, 34812 pA, instead of up to 300 pA. The cavitation absorb is also an outside area. | Swap of the Hemolyzer Unit. After rebooting and calibration, values are fine. tHb calibration performed, all right. Inlet probe crooked, renewed, also inlet seal.
No | Visit | Complaint | RIND | 2 point calibration failed | Checked and found sensor board was corroded. | Replaced sensor board, zero adjustment performed successfully. 2 point calibration performed successfully. | SWO-399441 | [5 empty] | 2 point calibration failed | Checked and found sensor board was corroded. | Replaced sensor board, zero adjustment performed successfully. 2 point calibration performed successfully.
No | Visit | Complaint | RIND | Capillary sample result issue. | Check the patient result & found capillary sample value high & few parameter Not derived. | Check the capillary tubes & found that they are not using the capillary tube with mixer ,told them to use our capillary tube & before run mix the sample properly to get proper result then Run a capillary sample & got the desired value. | SWO-401421 | [5 empty] | Capillary sample result issue. | Check the patient result & found capillary sample value high & few parameter Not derived. | Check the capillary tubes & found that they are not using the capillary tube with mixer ,told them to use our capillary tube & before run mix the sample properly to get proper result then Run a capillary sample & got the desired value.
No | Visit | Complaint | RIND | Service Visit. | SC is not installing because of V $C failed. | Try to reset the SP and again install the SC, not successful, try with other SP, not successful, do the zeroing of the analyzer successfully done, install the SP and SC, successfully done, analyzer handed over to staff in well working condition. | SWO-401208 | [5 empty] | Service Visit. | SC is not installing because of V /C failed. | Try to reset the SP and again install the SC, not successful, try with other SP, not successful, do the zeroing of the analyzer successfully done, install the SP and SC, successfully done, analyzer handed over to staff in well working condition.
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | pco2定标失败，0378 | 更换测试卡后修复 | SWO-399094 | Assisted Replacement | SC | SC | pCO2 | 378 | Test card, reagent package failure | pco2 bid failure, 0378 | Repair after replacing test card
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 不均匀的定标2.3溶液 1034 1089 | 连续更换2个试剂包后修复 | SWO-399767 | Assisted Replacement | SP | SP | FT | 1034, 1089 | Test card, reagent package failure | Uneven calibration 2.3 solution 1034 1089 | Repair after 2 reagent kits in a row
No | Visit | Complaint | RCHN | 孵育盘卡盘，报错5297 | 孵育盘卡盘，报错5297 | 重新清洁了孵育盘后修复 by jesson | SWO-399764 | Cleaned | HW | Incubator Module | PCB, Incubator | 5297 | Incubating plate chuck, 5297 | Incubating plate chuck, 5297 | Repair by jesson after re-cleaning the incubation plate
No | Visit | Complaint | RSCH | Rep Elektrolyte | Na und pH Instabil bei QC | pH Elektrodendichtung ausgetauscht Referenz Elektrode ausgetauscht Funktionskontrolle i.o Kunde beobachtet und gibt innerhalb der nächsten 2 Wochen Feedback | SWO-399131 | User expectation | Electrode | Electrode | QC |  | Rep Electrolyte | Na und pH Unstable bei QC | pH Electrode seal replaced Reference electrode replaced Functional control i.o Customer observes and gives feedback within the next 2 weeks
No | Visit | Complaint | RAME | Various Oxi neon lamp hardware errors. | FROM PREVIOUS REMOTE SUPPORT T $S: Neon number and neon intensity too low errors. Remote investigation to confirm errors. Performed complete OXI test in service with lamp failures. Lamp is failing: Neon intensity at 24% (range >=50%) Neon position -1.000 (range 0.100 - 2.000) Primary engineer is in route for service. Lamp is suspect and will need replaced. | Replaced and calibrated lamp unit, 902-458. Calibration and Quality Checks completed. Errors no longer present, normal operations restored. | SWO-399042 | Defect | HW | Oxi module | Lamp unit |  | Various Oxi neon lamp hardware errors. | FROM PREVIOUS REMOTE SUPPORT T /S: Neon number and neon intensity too low errors. Remote investigation to confirm errors. Performed complete OXI test in service with lamp failures. Lamp is failing: Neon intensity at 24% (range >=50%) Neon position -1.000 (range 0.100 - 2.000) Primary engineer is in route for service. Lamp is suspect and will need replaced. | Replaced and calibrated lamp unit, 902-458. Calibration and Quality Checks completed. Errors no longer present, normal operations restored.
No | Visit | Complaint | RSCH | Crea QC | Crea QC sehr instabil auf allen Level seit 24.07.18 Kunde hat schon mehrfach (2 Schachteln R0806) Crea neu membraniert ohne Erfolg | Crea A+B Elektrode ausgetauscht und neu membraniert pH Elektrode Dichtung ausgetauscht da undicht Funktionskontrolle und QC i.o (Foto im Anhang) | SWO-399251 | Assisted Replacement | Electrode | Electrode | Crea |  | Crea QC | Crea QC very unstable at all levels since 24.07.18 Customer has already re-membraneed Crea several times (2 boxes R0806) without success | Crea A+B electrode replaced and re-membraneed pH electrode seal replaced there leaking Functional control and QC i.o (photo in appendix)
No | Visit | Complaint | RFRA | Prévoir le changement de la nées le Washington unit | Suite au erreurs de çalcul de l'hématocrite, changement de la needle wash unit | Changement de la nées le Wash unit, calibration en mode service , passage des cq | SWO-399371 | Defect | HW | Aspiration Module | Needle Wash Unit |  | Anticipating the change of the born Washington unites | Following the errors of hematocrite, change of the needle wash unit | Change of the wash unit, calibration in service mode, passage of cq
No | Visit | Complaint | RLTD | Anna Morton 01708435000 ext 6672 or ext 2621 | The Analyser was turned off at 21.42 12 $08 $18. When it was switched back on the time and date were set to 31 $12 $2006, causing environment errors and the Analyser cartridges not to be recognised Unsure whether the Analyser was turned off correctly. | Retrieved service dump and inspected the data. The customer will monitor Analyser performance. | SWO-399228 | Defect | HW | Analyzer Control Module | CPU | Time Difference, mail: more information | Anna Morton 01708435000 ext 6672 or ext 2621 | The Analyser was turned off at 21.42 12 /08 /18. When it was switched back on the time and date were set to 31 /12 /2006, causing environment errors and the Analyser cartridges not to be recognised Unsure whether the Analyser was turned off correctly. | Retrieved service dump and inspected the data. The customer will monitor Analyser performance.
No | Visit | Complaint | RSCH | Rep | Automatische Sicherung nicht mehr möglich. Fehler | USB Stick ausgetauscht da defekt Funktion i.o Sicherung i.o | SWO-399129 | Defect | HW | External HW | USB Stick |  | Rep | Automatic backup is no longer possible. Error | USB stick replaced as defective Function i.o Fuse i.o
Yes | 35 | Y | RPAC | Blank | pO2 was failing multiple levels of QC with error codes 589 and 90. pO2 was passing calibration but failing QC. Earlier error 1084 Inhomogenous Cal 3 has also been noted. |  | D4CEB10A6C43A76A | [5 empty] | Blank | pO2 was failing multiple levels of QC with error codes 589 and 90. pO2 was passing calibration but failing QC. Earlier error 1084 Inhomogenous Cal 3 has also been noted.
No | Visit | Complaint | RAME | Unable to logon to user or service accounts | Userlist has been disabled again. | Reloaded Application Software. Copied configuration from another analyzer onsite. Disabled Centralized User Management. | SWO-400105 | Defect | SW | RMED SW | RMED SW | Logon | Unable to logon to user or service accounts | Userlist has been disabled again. | Reloaded Application Software. Copied configuration from another analyzer onsite. Disabled Centralized User Management.
No | Visit | Complaint | RDEU | 15.08.2018 08:40Uhr Frau Luckner 0376144721: Die Patientendatenübertragung funktioniert nur sporadisch | Überprüfung der Einstellungen | Abfrage der Daten nach Eingabe Patientennummer erfolgt Es gibt Messungen, bei denen die Daten nicht in der ITS ankommen Laut Aussage der IT-Abteilung fehlen im Datensatz vom Radiometeranalysator bei den Rohdatensätzen Felder | SWO-400070 | Setup | SW | External SW | His/Lis | not yet solved and that’s the reason why the SWO is set to „Waiting for internal information“. Our FSE is in contact with an external provider of lab IT to solve the problem. | 08/15/2018 08:40PM Ms Luckner 0376144721: Patient data transfer only works sporadically | Checking the settings | Query of the data after entering patient number is carried out There are measurements where the data does not arrive in the ITS According to the IT department, fields are missing from the record of the radiometer analyzer in the raw records
No | Visit | Complaint | RDEU | Erst nach Neustart druckt der Drucker wieder, bereits bei der Installation gab es dieses Problem | auf Vorschlag LUTRO Masseverbindung zum Drucker untersucht, keine Auffälligkeiten, was den ohmschen Widerstand abgeht, Spannungsversorgung auch i.O. hinsichtlich stat. Gleichspannungen | keine, Fehler tritt nur sporadisch auf | SWO-400368 | Revisit required | SW | RMED SW | printer |  | only after reboot the printer prints again, this problem was already in place during installation | on proposal LUTRO mass connection to the printer investigated, no abnormalities, which emanates the ohmic resistance, power supply also i.O. with regard to stat. DC voltages | none, errors occur only sporadically
No | Visit | Complaint | RLTD | Please call Denise Young on 0203 299 3222 | Inlet probe is slightly bent. Inlet gasket is very dirty and encrusted with dry blood. Solution pack was not activated correctly. Failing with fluid transport errors due to a clot in the fluid pathway. | Straightened inlet probe and thoroughly cleaned the inlet gasket. Activated the solution pack and performed intensive manual flush on the fluid pathway. Analyser is now functioning correctly but pCO2 and Lac are failing calibrations and QC's. Advised the customer to replace the SC. | SWO-400043 | Cleaned | Fluidic pathway | Fluidic pathway | Clot |  | Please call Denise Young on 0203 299 3222 | Inlet probe is slightly bent. Inlet gasket is very dirty and encrusted with dry blood. Solution pack was not activated correctly. Failing with fluid transport errors due to a clot in the fluid pathway. | Straightened inlet probe and thoroughly cleaned the inlet gasket. Activated the solution pack and performed intensive manual flush on the fluid pathway. Analyser is now functioning correctly but pCO2 and Lac are failing calibrations and QC's. Advised the customer to replace the SC.
No | Visit | Complaint | RAME | Unit will not clear intervention required screen | Pinch valve and Pm kit were ordered to resolve issue with vacuum error. | Performed Pm and replaced pinch valve. | SWO-400131 | Defect | HW | Oxi module | Pinch valve |  | Unit will not clear intervention required screen | Pinch valve and Pm kit were ordered to resolve issue with vacuum error. | Performed Pm and replaced pinch valve.
No | Visit | Complaint | RCHN | 报错1031 | 报错1031 借903-364一个 | 更换903-364备件后修复 by jesson | SWO-398592 | Defect | HW | Process Module | Optical Unit | 1031 | 1031 | False 1031 Borrow 903-364 one | Repair by jesson after replacement of 903-364 spare parts
No | Visit | Complaint | RAME | Leak test failed | Fuite détecté faire test pression aucune remplacé membrane po2 et tube de la probe | Installer nouveau joint introduction | SWO-400172 | Defect | HW | Inlet module | Inlet Tube |  | Leak test failed | Leak detected do pressure test no replaced membrane po2 and probe tube | Install new joint introduction
No | Visit | Complaint | RAME | Lactate membranes being replaced every 2 weeks | Lactate membranes being replaced every 2 weeks. Verified instrument and electrode | Lactate membranes being replaced every 2 weeks. Verified instrument and electrode . Performed zero calibration check in service. Installed new membranes. Performed several 2- point calibrations. Note: sample counter shows very high volume at this account. | SWO-402876 | Assisted Replacement | Electrode | Membrane | Lac |  | Lactate membranes being replaced every 2 weeks | Lactate membranes being replaced every 2 weeks. Verified instrument and electrode | Lactate membranes being replaced every 2 weeks. Verified instrument and electrode . Performed zero calibration check in service. Installed new membranes. Performed several 2- point calibrations. Note: sample counter shows very high volume at this account.
No | Visit | Complaint | RAME | Lactate only lasting 7 days: Level 3 2sd range:9.0-11.6 result:8.0 | Lactate membranes being replaced every 2 weeks. Verified instrument and electrode . Performed zero calibration check in service. Installed new membranes. Performed several 2- point calibrations. Note: sample counter shows very high volume at this account. | Lactate membranes being replaced every 2 weeks. Verified instrument and electrode . Performed zero calibration check in service. Installed new membranes. Performed several 2- point calibrations. Note: sample counter shows very high volume at this account. | SWO-402877 | Assisted Replacement | Electrode | Membrane | Lac |  | Lactate only lasting 7 days: Level 3 2sd range:9.0-11.6 result:8.0 | Lactate membranes being replaced every 2 weeks. Verified instrument and electrode . Performed zero calibration check in service. Installed new membranes. Performed several 2- point calibrations. Note: sample counter shows very high volume at this account. | Lactate membranes being replaced every 2 weeks. Verified instrument and electrode . Performed zero calibration check in service. Installed new membranes. Performed several 2- point calibrations. Note: sample counter shows very high volume at this account.
No | Visit | Complaint | RAME | QC is at different than scheduled Contact: Macoully 617-732-7403 | QC is at different than scheduled. Checked qc auto schedule, good for every 4 hrs. | Backed up all data and gave to lab director. Reinstalled all qc levels and set up all qc schedules. | SWO-402878 | Setup | SW | RMED SW | QC schedule |  | QC is at different than scheduled Contact: Macoully 617-732-7403 | QC is at different than scheduled. Checked qc auto schedule, good for every 4 hrs. | Backed up all data and gave to lab director. Reinstalled all qc levels and set up all qc schedules.
No | Visit | Complaint | RAME | Customer complaints about sensitivity and stability of creatinine during QC ampoules and samples measurements. | Creatinine sensitivities and stability issue. QC measurements are dropping down during continuous measurements. The crea membrane does not last for 14 days and the QC failure happening earlier than 14 days with very low number of samples. | Create a tech line case for this issue. | SWO-399941 | Assisted Replacement | Electrode | Membrane | Crea |  | Customer complaints about sensitivity and stability of creatinine during QC ampoules and samples measurements. | Creatinine sensitivities and stability issue. QC measurements are dropping down during continuous measurements. The crea membrane does not last for 14 days and the QC failure happening earlier than 14 days with very low number of samples. | Create a tech line case for this issue.
No | Visit | Complaint | RAME | System communication fails to send | Worked with IT to resolve communication issues | Copied and restored Roma.ini and it resolved the rest | SWO-400832 | Setup | SW | RMED SW | ini file |  | System communication fails to send | Worked with IT to resolve communication issues | Copied and restored Roma.ini and it resolved the rest
No | Visit | Complaint | RKK | 患者検体測定時にオキシスペクトラムミスマッチが多発し測定できず。 | 機器引き上げ | 代替え機設置 | SWO-402171 | Revisit required | Analyzer removed | Oxi module | Hemolyzer | 581 | Oxyspectrum mismatch is frequent ly and can not be measured when measuring patient specimens. | Equipment pulling up | Installation of alternative machines
No | Visit | Complaint | RDEU | Erst nach Neustart druckt der Drucker wieder, bereits bei der Installation gab es dieses Problem | No | Printer Unit erneuert, Zeit wird zeigen, ob es die Ursache behebt | SWO-400524 | Defect | SW | RMED SW | printer |  | only after reboot the printer prints again, this problem was already in place during installation | No | Printer Unit renewed, time will show if it fixes the cause
No | Visit | Complaint | RFRA | 1401 Erreur joint d'introduction | Mode service ok Le joint d'introduction ne revient pas bien à sa position initiale (aiguille tordue, sale ou le joint d'introduction accroche contre l'aiguille). Dans ce cas présent, cela venait de l'aiguille. | Chgt du module inlet complet. Contrôle en fonctionnement ok. Constations sur place pendant utilisation sur une personne: la limaille en place dans le capillaire, mais n'est pas agitée, pas d' aimant sur place (d'où échantillon non homogène). Non utilisation du filtre à caillot. Serait-il possible d'installer l'ABL de backup à la place de l'ABL néonatale ? Pour vérifier si cela vient de l'ABL ou de l'utilisation. | SWO-400627 | Defect | HW | Auto Inlet module | Auto Inlet module |  | 1401 Joint Introductory Error | Service mode ok The introductory seal does not go back well to its original position (twisted, dirty needle or introductory seal hangs against needle). In this case, it came from the needle. | Chgt of the full inlet module. Control in operation ok. On-site observations during use on a person: the chip in place in the capillary, but is not agitated, no magnet on the spot (hence non-homogeneous sample). No use of the clot filter. Would it be possible to install the backup ABL in place of the neonatal ABL? To check if this comes from the ABL or use.
No | Visit | Complaint | RSAF | 2x SP rejected by analyzer | Crystals formed in Inlet gasket w.holder, | Cleaned inlet gasket and reinstalled SP. startup tests completed. PCO2, Gluc and Lac sensitivity low, customer will monitor and replace SC when necessary. | SWO-401052 | User error | SP | SP | solution pack replacement |  | 2x SP rejected by analyzer | Crystals formed in Inlet gasket w.holder, | Cleaned inlet gasket and reinstalled SP. startup tests completed. PCO2, Gluc and Lac sensitivity low, customer will monitor and replace SC when necessary.
No | Visit | Complaint | RFRA | ERR STATUS 2/PAS DE DEMARRAGE DE LA MACHINE | Roma.db $CF card | Constate err status 2 et pas de démarrage de la machine $Restauration du Roma.db empty impossible $Tenetative de réinstation de W7 impossible de formater la CF card car défectuause $Pas de CF card dispo $Retour atelier pour préparation d'une CF card $Instatlion de W7 $du patch sécurité $Instalation de l'application Radiometer $ | SWO-400945 | Revisit required | HW | DMS_PC | CF card |  | MACHINE'S ERR STATUS 2/PAS | Roma.db /CF card | Finds err status 2 and no machine start $Restauration of Roma.db empty impossible $Tenetative reinstation of W7 impossible to format the CF card because defective $Pas of CF card available $Retour workshop for preparation of a CF card $Instatlion of W7$ the security patch $Instalation of the Radiometer app $
No | Visit | Complaint | RFRA | Écran noir /pas de démarrage de la machine |  | Remplacement de la Cf card $installation de W7 ,du patch sécurité,ainsi que de l'application Radiometer.Impossible de restaurer les données sauvegardées sur le serveur $ Configuration des adresses IP $Connexion à Aqure $ | SWO-400967 | Defect | HW | DMS_PC | CF card |  | Black screen / no machine start |  | Replacing W7's $installation cf card, security patch, and Radiometer.Can't restore data backed up on the server $ Setting up IP addresses $Connexion at Aqure $
No | Visit | Complaint | RCHN | pco2灵敏度偏低 | pco2灵敏度偏低，借903-581一个 | 更换903-581后修复 | SWO-399098 | Defect | HW | Basic Unit | PCB, Sensor interface | pCO2 | low sensitivity of pco2 | pco2 sensitivity is low, borrow903-581 one | Repair after replacement 903-581
No | Visit | Complaint | RCHN | 液体传感器出错 | 液体传感器出错 借903-199一个 | 更换库存903-199后修复 | SWO-399750 | Defect | HW | Basic Unit | Cassette retainer | LS | Liquid sensor error | Liquid sensor error borrows 903-199 one | Repair after replacement of stock 903-199
Yes | Visit | Complaint | RCHN | 报错1223和1267 | 报错1223和1267,需要借件903-329一个 | 更换903-329后修复 by jesson | SWO-399795 | [5 empty] | 1223 and 1267 | 1223 and 1267, need to borrow 903-329 one | Repair by jesson after replacing 903-329
No | Visit | Complaint | RIND | Electrolyte parameter calibration failed. | Checked reference membrane and found it was empty. | Replaced Ref membrane and then given 2 point calibration.Checked with samples and found it is ok. | SWO-401119 | Assisted Replacement | Electrode | Membrane | Reference |  | Electrolyte parameter calibration failed. | Checked reference membrane and found it was empty. | Replaced Ref membrane and then given 2 point calibration. Checked with samples and found it is ok.
No | Visit | Complaint | RAME | Ca failing calibrations Status high 21.19 | Cal data hi on ca++ 21.19 Ph low 7.080 Suspect ref membrane or contaminate in chamber Ph o ring stuck to bottom also contaminates in ph ref chamber Inspected ca chamber electrode and membrane | Clean electrodes replace o ring on ph clean chambers re membrane ref electrode | SWO-400841 | Defect | Electrode | Electrode | Sealing ring, pH |  | Ca failing calibrations Status high 21.19 | Cal data hi on ca++ 21.19 Ph low 7.080 Suspect ref membrane or contaminate in chamber Ph o ring stuck to bottom also contaminates in ph ref chamber Inspected ca chamber electrode and membrane | Clean electrodes replace o ring on ph clean chambers re membrane ref electrode
No | Visit | Complaint | RDEU | pH/Lyte-Drift- und QK-Fehler auffällig seit letzter Wartung | pH $Lyte-Driften | Reinigungsarbeiten, Fehlersuche, alle Membranventile demontiert, gereinigt und wieder montiert, erneut Lösungstransport kontrolliert, nochmals alle betroffenen Elektroden incl. E1001 kontrolliert, E1001 muß ersetzt werden, Lieferung durch RDEU-VSCNO (CarKe), Kal und QK ok | SWO-400471 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | pH/Lyte drift and QK errors conspicuous since last maintenance | pH /Lyte operation | Cleaning work, troubleshooting, all diaphragm valves dismantled, cleaned and reassembled, solution transport checked again, all affected electrodes incl. E1001 checked again, E1001 must be replaced, delivery by RDEU-VSCNO (CarKe), Kal and QK Ok
Yes | Visit | Complaint | RAME | po2 Zero high on Calibrations | Deliver PO2 and pick. Up old. Electrode | Picked up old electrode will send back. to Westlake | SWO-402032 | Re classified RMED | Logistical upSWO |  |  |  | po2 Zero high on Calibrations | Deliver PO2 and pick. Up old. Electrode | Picked up old electrode will send back. to Westlake
No | Visit | Complaint | RDEU | Gerät kalibriert nicht. | Fehler 5689 Rotatestorafewheel_24 Das Storagewheel hat nicht rotiert, da ein Cartridge Guide For Storage Wheel defekt war und dadurch zwei der Kassetten blockiert haben. | Erneuerung des defekten Teils. Durchführung der TnT-Kalibration fehlerfrei. | SWO-400533 | Defect | HW | Storage Module | Cartridge guide | 5689 | Device does not calibrate. | Error 5689 Rotatestorafewheel_24 The storage wheel did not rotate because a Cartridge Guide For Storage Wheel was defective, blocking two of the cartridges. | Renewal of the defective part. Execution of the TnT calibration error-free.
Yes | Visit | Complaint | RAME | Error codes 1169, 0508 and 1324 | Checked fluid transport in service Vacuum value out of range high | Replaced inlet connector gasket | SWO-401271 | [5 empty] | Error codes 1169, 0508 and 1324 | Checked fluid transport in service Vacuum value out of range high | Replaced inlet connector gasket
No | Visit | Complaint | RDEU | 14.08.2018; Fr. Moritz; Tel.: 06213836540; Touchscreen ist dunkel | Lieferung CPU verzögert. Bereits geliefert wurde Display und Netzteil. Im Haus unterwegs zu weiteren Service. | Einbau eines Netzteils und eines Displays. Nach Neustart bleibt Display weiterhin dunkel. Analysator startet gar nicht auf DOS Ebene. Fehler Ursache liegt an CPU. Nach Rücksprache mit Logistik eilige Lieferung aus Dänemark über Redline. Analysator nicht betriebsbereit. | SWO-401152 | Revisit required | HW | DMS_PC | CPU |  | 14.08.2018; Ms. Moritz; Tel.: 06213836540; Touchscreen is dark | Delivery CPU delayed. The display and power supply have already been delivered. In the house on the way to further service. | Installation of a power supply and a display. After rebooting, the display remains dark. Analyzer does not start at THE DOS level. Error Cause is due to CPU. After consultation with logistics, urgent delivery from Denmark via Redline. Analyzer not ready for operation.
No | Visit | Complaint | RAME | Analyzer will not power up after turning it off. Customer did not use temporary shutdown procedure, rather the customer shut it down simply by flipping the power switch. | Analyzer will not power up after being turned off. Upon investigation, it was discovered that two of the three fuses in the back of the analyzer were not lit. Fuses were deemed good after a resistance check on both, however, a voltage check showed that there was no voltage to or from the fuses. It was determined that the power supply was faulty. | Replaced power supply (910-285). Analyzer powered up correctly, normal operations restored. | SWO-401202 | Defect | HW | Power Supply | Power supply |  | Analyzer will not power up after turning it off. Customer did not use temporary shutdown procedure, rather the customer shut it down simply by flipping the power switch. | Analyzer will not power up after being turned off. Upon investigation, it was discovered that two of the three fuses in the back of the analyzer were not lit. Fuses were deemed good after a resistance check on both, however, a voltage check showed that there was no voltage to or from the fuses. It was determined that the power supply was faulty. | Replaced power supply (910-285). Analyzer powered up correctly, normal operations restored.
No | Visit | Complaint | RLTD | David Mather Tel: 01689 865318 | Replaced faulty sensor interface pcb and replaced Sensor Cassette with new one supplied by Radiometer as warranty replacement. Verified operation of analyser on restart, all parameters including Lactate now working correctly. | Replaced sensor interface pcb. | SWO-401107 | Defect | HW | Basic Unit | PCB, Sensor interface | Lac | David Mather Tel: 01689 865318 | Replaced faulty sensor interface pcb and replaced Sensor Cassette with new one supplied by Radiometer as warranty replacement. Verified operation of analyser on restart, all parameters including Lactate now working correctly. | Replaced sensor interface pcb.
No | Visit | Complaint | RLTD | Contact: Sam Khadaroo Tel: 01823 343 140 | On arrival the analyser was going through its start-up which it completed successfully but a pressure test error did flag up towards the end and then it cleared. Once the startup was complete I checked the fluid transport service programs and the pressure test vacuum value was a bit high. | The inlet gasket was replaced as they didn’t have one to replace last night but this didn’t make much difference to the vacuum value. Then I changed the pinch valve tube and this brought the vacuum value down nearer the middle value. As half then service kit had been fitted and the annual service is due in about a month the rest of the kit was fitted and the annual service completed. The inlet probe was also replaced as it was bent. | SWO-401016 | Defect | HW | Oxi module | Pinch valve tube |  | Contact: Sam Khadaroo Tel: 01823 343 140 | On arrival the analyser was going through its start-up which it completed successfully but a pressure test error did flag up towards the end and then it cleared. Once the startup was complete I checked the fluid transport service programs and the pressure test vacuum value was a bit high. | The inlet gasket was replaced as they didn’t have one to replace last night but this didn’t make much difference to the vacuum value. Then I changed the pinch valve tube and this brought the vacuum value down nearer the middle value. As half then service kit had been fitted and the annual service is due in about a month the rest of the kit was fitted and the annual service completed. The inlet probe was also replaced as it was bent.
No | Visit | Complaint | RLTD | On arrival on unit to do software upgrade advised by staff analyser not working. | Tried different solution pack, QC errors still present. Cleaned hemolyser with hypochlorite solution. Performed tHb calibration. Repeated quality controls - now QC passed for all parameters. | Cleaned hemolyser and performed tHb calibration. | SWO-401108 | Cleaned | Fluidic pathway | Fluidic pathway | Oxi module |  | On arrival on unit to do software upgrade advised by staff analyser not working. | Tried different solution pack, QC errors still present. Cleaned hemolyser with hypochlorite solution. Performed tHb calibration. Repeated quality controls - now QC passed for all parameters. | Cleaned hemolyser and performed tHb calibration.
Yes | 3 | Y | RPAC | Blank | failed SP |  | D4CAF9040A5F29C5 | [5 empty] | Blank | failed SP
No | Visit | Complaint | RSAF | 2x SP rejected by analyzer | SC not replaced as adviced | Replaced faulty SC, startup test ok. | SWO-401326 | Assisted Replacement | SC | SC | unknown |  | 2x SP rejected by analyzer | SC not replaced as adviced | Replaced faulty SC, startup test ok.
No | Visit | Complaint | RFRA | ERR STATUS 2/PAS DE DEMARRAGE DE LA MACHINE | SOFTWARE | Mise en place de la CF card préparée à l'atelier $Mise en route calibration et passage des CQ ok $ Restauration de la sauvegarde $Test passage échantillon ok $ | SWO-400952 | Defect | HW | DMS_PC | CF card |  | MACHINE'S ERR STATUS 2/PAS | SOFTWARE | Setting up the CF card prepared at the workshop $Mise on the way calibration and passing the CQ ok $ Restoration of the backup $Test sample pass ok $
No | Visit | Complaint | RSAF | Inlet gasket and SP replaced still requires SP replacement | Faulty solution pack- fluid transport tests done with activities count not changing at 677. | Replaced Sol pack startup tests ok. | SWO-401325 | Assisted Replacement | SP | SP | FT |  | Inlet gasket and SP replaced still requires SP replacement | Faulty solution pack- fluid transport tests done with activities count not changing at 677. | Replaced Sol pack startup tests ok.
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | PCO2定标不过 0210 | 更换测试卡后修复 | SWO-402160 | [5 empty] | Test card, reagent package failure | PCO2 calibration is no more than 0210 | Repair after replacing test card
No | Visit | Complaint | RCHN | 做样本频繁提示样本量不足 | 做样本频繁提示样本量不足 | 进样针定标，HCT检查和清洁，做空套管装处理包缓冲液做32个样本，绿头管子血样20个样本，紫头管子血样12个样本，空套管和紫头管没出现问题，绿头管出现5个样本量不足。判断:绿头管标称4ml，实际大部分抽样3ml以内，可能负压原因吸不上样本。让科室用紫头管抽样操作。 | SWO-401486 | User error | HW | Aspiration Module | pre analytical | Short sample | Do samplefrequent prompt sample size is insufficient | Do samplefrequent prompt sample size is insufficient | Sample needle calibration, HCT inspection and cleaning, empty casing treatment package buffer to do 32 samples, green tube blood sample 20 samples, purple tube blood sample 12 samples, empty casing and purple tube no problem, green tube appeared 5 sample size is insufficient. Judgement: Green tube nominal 4 ml, the actual majority of the sampling within 3 ml, may be negative pressure reasons can not suck the sample. Let the department use purple tube sampling operation.
No | Visit | Complaint | RCHN | 报错5784 | 报错5784 | 调整进样门位置后修复 by jesson | SWO-400895 | Adjustment | HW | Cover and Chassis | Door, Inlet | 5784 | 5784 | 5784 | Fix by jesson after adjusting the position of the sample door
No | Visit | Complaint | RCHN | 报错5777 | 报错5777 | 重新安装试剂包后修复 | SWO-401080 | Assisted Replacement | Consumables | Solution Pack | solution pack replacement | 5777 | 5777 error | 5777 error | Repair after reinstalling rehousing kit
No | Visit | Complaint | RIND | Software upgradation & Parameter failure | PH,po2,cl failure | Flush out the casstte with normal sline,re install it , run 2point calibration,calibration passed. upgraded the new software V 1.1.1. | SWO-401376 | [5 empty] | Software upgradation & Parameter failure | PH,po2,cl failure | Flush out the casstte with normal sline,re install it , run 2point calibration,calibration passed. upgraded the new software V 1.1.1.
No | Visit | Complaint | RSAF | Sensor remembrances but always calibrating | Intermittent error code 05100020 on startup | Cable error, monitor tested with sensor from different monitor. Software upgrade done. Sensor to be replaced. | SWO-401328 | [5 empty] | Sensor remembrances but always calibrating | Intermittent error code 05100020 on startup | Cable error, monitor tested with sensor from different monitor. Software upgrade done. Sensor to be replaced.
No | Visit | Complaint | RSCH | Rep. Drift Toleranzen | Drift Kalibrierung = 0 Alle Parameter mit ? | Drift Toleranzen waren auf 0 gesetzt Drift Toleranzen angepassr i.o Kal. und QC i.o System i.o | SWO-401054 | Setup | SW | RMED SW | Drift Tolerances |  | Rep. Drift Tolerances | Drift Calibration = 0 All parameters with ? | Drift tolerances were set to 0 Drift Tolerances customizable i.o Kal. and QC i.o System i.o
No | Visit | Complaint | RDEU | Kunde meldet iterative Driften bei Elektrolyten | Iterative Driften bei Elektrolyten | Kunde hat bereits auf Anraten ( im Laufe des Vormittags des Servicetages ) bereits die Ref-Elektrode lt. Aktivitätenlog erneuert, pro-aktiv Dichtungsring an der pH-Elektrode erneuert und Messkammer gereinigt, Kunden empfohlen bei weiteren Auffälligkeiten noch die pH-Elektrode zu ersetzen, letzten 6 Kalibrierungen sowie QC seit den Maßnahmen unauffällig | SWO-401110 | Cleaned | HW | pH/BG Module | measuring chamber |  | Customer reports iterative drifts in electrolytes | Iterative drifting in electrolytes | On advice ( in the course of the morning of the service day ) the ref electrode has already been renewed according to activities log, pro-active sealing ring at the pH electrode and cleaned measuring chamber, customers recommended in case of further abnormalities still the pH electrode, last 6 calibrations as well as QC since the measures inconspicuous
No | Visit | Complaint | RHUN | Glu, Lac kalibrálási hiba | Glu, Lac kal.hiba. | Glu, Lac membrán csere. Kalibrálás, QC mérés. | SWO-401082 | Assisted Replacement | Electrode | Membrane | Glu | Glu/Lac | Glu, Lac kalibrálási hiba | Glu, Lac kal.hiba. | Glu, Lac membrane replacement. Calibration, QC measurement.
No | Visit | Complaint | RKK | バッテリーエラーが検出されました。バッテリーセルにおけるエラー。バッテリーケーブル抜き差しでも変わらず。バッテリー接続外し作業終了。バッテリー入荷次第交換予定。 | バッテリーエラーが検出されました。バッテリーセルにおけるエラー。 | バッテリーケーブル抜き差しでも変わらず。バッテリー接続外し作業終了。バッテリー入荷次第交換予定。 | SWO-401545 | Revisit required | HW | Battery module | Battery module |  | A battery error was detected. Error in battery cell. Even the battery cable is not changed. The battery is disconnected and the work is finished. It will be replaced as soon as the battery is in stock. | A battery error was detected. Error in battery cell. | Even the battery cable is not changed. The battery is disconnected and the work is finished. It will be replaced as soon as the battery is in stock.
Yes | Visit | Complaint | RSAF | Solution pack depleted | Customer reported that their solution pack is depleted and they don't have stock. | Replaced with a new solution pack. | SWO-403401 | [5 empty] | Solution pack depleted | Customer reported that their solution pack is depleted and they don't have stock. | Replaced with a new solution pack.
No | Visit | Complaint | RCHN | 报错1028 1035 | 报错1028 1035 借903-364一个 | 更换903-364后修复 by jesson | SWO-400415 | Defect | HW | Process Module | Optical Unit | 1035 | Error 1028 1035 | Error 1028 1035 Borrowing 903-364 one | Repair by jesson after replacing 903-364
No | Visit | Complaint | RCHN | 频繁自动关机 | 频繁自动关机，借910-321一个（上海缺货，周一广办发货） |  | SWO-401442 | Revisit required | HW | DMS_PC | CPU |  | Frequent automatic shutdowns | Frequent automatic shutdown, borrowing 910-321 one (Shanghai out of stock, Monday wide delivery)
No | Visit | Complaint | RSCH | Samstageinsatz. Fehlrhafte Betriebsmittel | Analysator blockiert. Telefonisch Longterm Shutdown nicht möglich (Processwheel Fehler) | Analysator geöffnet. Testkassetten Manuel entfernt da einen CRP Kassette Produktionsfehler. (Lasche für den Auswurf fehlt.) Alle Kontrollen auf Kundenwunsch ausgeführt um die Funktion zu Prüfen Funktionskntrolle i.o | SWO-401444 | Defect | Consumables | Test Kit | Cartridge, Broken |  | Saturday. Incorrect equipment | Analyzer blocked. Telephone Longterm Shutdown Not Possible (Processwheel Error) | Analyzer open. Test cartridges Manuel removes a CRP cassette production error. (The flap for the ejection is missing.) All checks carried out on customer request in order to check the function Functional knob s.o.
Yes | Visit | Complaint | RAME | test | testing | tested | SWO-402070 | Re classified RMED | Test SWO |  |  |  | test | testing | tested
No | Visit | Complaint | RDEU | 20.08.2018, Herr Nüssel, Tel: 094194417751: AQC Fehler bei FO2Hb und MetHb Level B und C. 20.08.2018 Herr Nüssel , Tel: 094194417751: AQC Fehler bei FO2Hb und MetHb Level B und C besteht weiterhin trotz tHb Kalibration | AQC Fehler bei FO2Hb und MetHb Level B und C besteht weiterhin. Ventilschlauch ist undicht. | Backflush durchgeführt. Ventilschlauch erneuert. Fehler bleibt vorhanden. Hemolyzer erneuert. Oxi Complete Test ok. tHb Kalibration (99,7%), Kalibration und QC’s vorgenommen. | SWO-401942 | Defect | HW | Oxi module | Hemolyzer | O2Hb, MetHb | 20.08.2018, Mr. Nüssel, Tel: 094194417751: AQC Errors at FO2Hb and MetHb Level B and C. 20.08.2018 Mr. Nüssel , Tel: 094194417751: AQC errors at FO2Hb and MetHb Level B and C persist despite tHb calibration | AQC errors at FO2Hb and MetHb Level B and C still exist. Valve hose is leaking. | Backflush. Valve hose renewed. Error remains. Hemolyzer. Oxi Complete Test ok. tHb calibration (99.7%), calibration and QC's.
Yes | 11 | Y | RPAC | Blank | SC failure |  | D4CF0C166C4C33A1 | [5 empty] | Blank | SC failure
Yes | 8 | Y | RPAC | Blank | Calibration failure on Lactate |  | D4CF1D216C13EEF7 | [5 empty] | Blank | Calibration failure on Lactate
No | Visit | Complaint | RLTD | Keymea contacted us whilst on site. | Re-attached front Solution Pack guide, but still loose. | Return visit required with correct size screw to complete repair. | SWO-404015 | Revisit required | HW | Basic Unit | SP guide |  | Keymea contacted us whilst on site. | Re-attached front Solution Pack guide, but still loose. | Return visit required with correct size screw to complete repair.
No | Visit | Complaint | RAME | SP going empty prematurely, today says 39 events eft in pack and 30 day expiry is at aug 5 Contact is Dave at 507-537-9179 | On site visit to retrieve the files requested by Sendx | Copied the database and the data log files to send to Sendx for evaluation regarding the solution packs running empty too early. I explained the status of the situation to Amy, the lab manager just to make things clear. I also replaced the solution pack and sensor cassette with brand new units. The device is ready for use. | SWO-402042 | [5 empty] | SP going empty prematurely, today says 39 events eft in pack and 30 day expiry is at aug 5 Contact is Dave at 507-537-9179 | On site visit to retrieve the files requested by Sendx | Copied the database and the data log files to send to Sendx for evaluation regarding the solution packs running empty too early. I explained the status of the situation to Amy, the lab manager just to make things clear. I also replaced the solution pack and sensor cassette with brand new units. The device is ready for use.
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 1361参比电极错误，pco2定标失败0378 | 更换2张测试卡后修复 | SWO-401745 | [5 empty] | Test card, reagent package failure | 1361 Reference electrode error, pco2 calibration failure 0378 | Repair after replacing 2 test cards
No | Visit | Complaint | RCHN | 升级后数据丢失 | 升级后数据丢失 借944-258（14381）一张 缺货 |  | SWO-401612 | User error | Software | Analyzer Control Module | CPU | SW upgrade | Data loss after upgrade | Post-upgrade data loss Borrowing 944-258 (14381) a out of stock
Yes | Visit | Complaint | RCHN | 报错1035 | 报错1035，借903-364一个（广州自提） |  | SWO-401592 | [5 empty] | 1035 error | 1035, borrow edgy 903-364 (Guangzhou self-mention)
Yes | Visit | Complaint | RCHN | 报错1035 | 借944-212（13719缺货），944-220（13612）各一个 |  | SWO-401593 | [5 empty] | 1035 error | Loan 944-212 (13719 out of stock), 944-220 (13612) each
No | Visit | Complaint | RKK | 装置から異音がする。 日付/時刻がずれる | 装置から異音 | 装置返却 | SWO-403591 | [5 empty] | There is a noise coming from the device. Date/Time Shifts | Noise from the device | Return of equipment
Yes | 8 | Y | RPAC | Blank | Compensate 2 boxes of Autocheck 6+ QC |  | D4CF165F6C4DEEF3 | Re classified RMED | Logistical upSWO |  |  | Compensate for higher amout of QCs needed - FAN | Blank | Compensate 2 boxes of Autocheck 6+ QC
Yes | Visit | Complaint | RDEU | pH/Lyte-Drift- und QK-Fehler auffällig seit Ende Juni @CarKe: Bitte 21. o. 22.8.2018 eine E1001 mit SWO für diesen ABL an Frau Bretting liefern! | Materiallieferung | Ref.-Elektrode an Fr. Bretting übergeben | SWO-401779 | Re classified RMED | Logistical upSWO |  |  |  | pH/Lyte drift and QK errors conspicuous since end of June @CarKe: Please deliver an E1001 with SWO for this ABL to Mrs. Bretting 21. o. 22.8.2018! | Materiallieferung | Ref.-Electrode handed over to Ms. Bretting
No | Visit | Complaint | RAME | Analyzer will not load software | On arrival found Corrupted Roma.ini | Tried copying the Roma.ini from another instrument next to it did not work. Rebooted multiple times still nothing. Tried uninstalling Kira 6.18 but no option to remove it. Manually launched setup.exe for 6.18 using usb and reinstall the software and made sure to uncheck keeping the old files and everything came up just fine. All the old data was still there. All calibrations and 3 Levels of QC passed as per manufacturer specification | SWO-402110 | Defect | SW | RMED SW | ini file | 6.18 | Analyzer will not load software | On arrival found Corrupted Roma.ini | Tried copying the Roma.ini from another instrument next to it did not work. Rebooted multiple times still nothing. Tried uninstalling Kira 6.18 but no option to remove it. Manually launched setup.exe for 6.18 using usb and reinstall the software and made sure to uncheck keeping the old files and everything came up just fine. All the old data was still there. All calibrations and 3 Levels of QC passed as per manufacturer specification
No | Visit | Complaint | RSCH | Crea QC | CREA QC ausserhalb Bereich | Crea mit neuer Lot membraniert. Aktuell QC Top Neue Kal 1 & 2 Lösung | SWO-401740 | Assisted Replacement | Electrode | Membrane | Crea | CAPA-00789 | Crea QC | CREA QC out of range | Crea membraneed with new solder. Current QC Top New Kal 1 & 2 solution
No | Visit | Complaint | RESP | La imagen de la pantalla vibra | Se ha roto algún contacto en el cable | Se sustituye cable | SWO-401661 | Defect | HW | Analyzer Control Module | cable, display |  | The screen image vibrates | Some contact broke on the cable | Cable is replaced
No | Visit | Complaint | RDEU | 20.08.2018, Frau Dorn, tel. 06946086125: ständiger Neustart | Gerät überprüft, Fehler Enviroment 3264. Sevice Test geprüft, Temp ist i.O Datum und Uhrzeit ist am 17.08.2018 um 04:02 umgestellt worden. > 31.12.2006, 23:02 Uhr. | Datum und Uhrzeit korrigiert, Time Server aktualisiert. Startphase, Systemüberprüfung i.O | SWO-401889 | Adjustment | SW | Analyzer Control Module | CPU, time | 3264 | 20.08.2018, Mrs. Dorn, tel. 06946086125: permanent reboot | Device checked, error Enviroment 3264. Sevice test tested, Temp is i.O The date and time changed on 17/08/2018 at 04:02. > 31.12.2006, 23:02. | Fixed the date and time, Time Server updated. Start-up phase, system check i.O
No | Visit | Complaint | RIND | Calibration failed after cal adjustment started. | Checked and found in service program that storage wheel failed. Due to cartridge stucked. | Checked and removed the stucked cartridge, as the cartridge got broken. | SWO-401675 | Defect | Consumables | Test Kit | Cartridge, Broken |  | Calibration failed after cal adjustment started. | Checked and found in service program that storage wheel failed. Due to cartridge stucked. | Checked and removed the stucked cartridge, as the cartridge got broken.
No | Visit | Complaint | RDEU | 15.08.2018; Fr. Bode; Tel.: 01726414309; D-dimer Kal wird 2x abgebrochen | Mit Frau Krebs gesprochen, die D-Dimer Kalibration sollen später durchgeführt werden | Mit Frau Krebs gesprochen, die D-Dimer Kalibration sollen später durchgeführt werden | SWO-402962 | User error | Consumables | Test Kit | Cal, D-Dimer |  | 15.08.2018; Ms. Bode; Tel.: 01726414309; D-dimer Kal is aborted 2x | Spoke to Ms Cancer, D-Dimer calibration to be carried out later | Spoke to Ms Cancer, D-Dimer calibration to be carried out later
No | Visit | Complaint | RAME | Lines on the display | Screen displayed colored lines. | Replaced the display. | SWO-402039 | [5 empty] | Lines on the display | Screen displayed colored lines. | Replaced the display.
No | Visit | Complaint | RDEU | 14.08.2018; Fr. Moritz; Tel.: 06213836540; Touchscreen ist dunkel | Touch Screen dunkel | Erneuerung und Umbau CPU-Einheit - SC Erneuerung nötig (verfallen ) - KAL Fehler pCO2 (Empfindlichkeitswert 10%) und QC Fehler pCO2 und Oximetrie - Analysator gespült (pCO2 Fehler bleibt bestehen) - Funktionsüberprüfung und Servicechecks etc. KAL Fehler pCO2 weiterhin ( Anwenderin möchte noch abwarten, ob sich pCO2 regeneriert und erneuert gegebenenfalls SC erneut ) Neue MAC-Adr. 00-04-5F-9F-76-EA | SWO-402430 | Defect | HW | DMS_PC | CPU |  | 14.08.2018; Ms. Moritz; Tel.: 06213836540; Touchscreen is dark | Touch Screen dunkel | Renewal and conversion CPU unit - SC renewal necessary (forfeit) - KAL error pCO2 (sensitivity value 10%) and QC error pCO2 and oximetry - analyzer flushed (pCO2 error remains) - function check and service checks etc. KAL error pCO2 continues (user still wants to wait and see if pCO2 regenerates and renews SC if necessary ) New MAC-Adr. 00-04-5F-9F-76-EA
No | Visit | Complaint | RKK | 患者検体測定時にオキシスペクトラムミスマッチが多発し測定できず。 | 修理完了 | 機器戻し | SWO-402176 | Defect | HW | Oxi module | Hemolyzer | 581 | Oxyspectrum mismatch is frequent ly and can not be measured when measuring patient specimens. | It's done. | Equipment return
No | Visit | Complaint | RFRA | Erreur de Calibration de la Temperature de la cassette, malgré le remplacement de celle ci | Thermistor recalibration en mode service. | Remplacement de la carte sensor board . Remise en route, erreur de la base de données (statuts code 5). Pertes de l'ensembles des données patients, Calibration, contrôles et activité. Pas de sauvegardes disponible. Remplacement de la cassette. Calibration et passage CQ | SWO-402700 | Defect | HW | Basic Unit | PCB, Sensor interface |  | Error in Calibration of the Temperature of the Cassette, despite the replacement of the cassette | Thermistor recalibration en mode service. | Replacing the sensor board card. Restarted, database error (code 5 statutes). Losses of patient data sets, calibration, controls and activity. No backups available. Replacing the tape. Calibration and passage CQ
Yes | 8 | Y | RPAC | Blank | Calibration failure on Lactate |  | D4D05AE1E20F02CD | [5 empty] | Blank | Calibration failure on Lactate
No | Visit | Complaint | RLTD | Mary Tel: 01642 282676 | Clot suspected in sensor Cassette.All consumables had been replaced and analyser had been flushed several times during the night. | On arrival analyser was showing Clot suspected in SC and OXI modules. Fluid transport tests were carried out and clot detect $remove failed and SP failed for rinse. The analyser was flushed and no restriction was felt. Fitted service kit and flushed analyser thoroughly. On start up the fluid transport tests were carried out again with no failures. Calibrations and QCs passed. A sample was put through the analyser to test functionality. | SWO-402502 | Defect | HW | Basic Unit | PM Kit |  | Mary Tel: 01642 282676 | Clot suspected in sensor Cassette.All consumables had been replaced and analyser had been flushed several times during the night. | On arrival analyser was showing Clot suspected in SC and OXI modules. Fluid transport tests were carried out and clot detect $remove failed and SP failed for rinse. The analyser was flushed and no restriction was felt. Fitted service kit and flushed analyser thoroughly. On start up the fluid transport tests were carried out again with no failures. Calibrations and QCs passed. A sample was put through the analyser to test functionality.
Yes | 2 | Y | RPAC | Blank | Software upgrade to version 3.4MR1 |  | D4D07910E262C982 | Re classified RMED | planned SWO |  |  |  | Blank | Software upgrade to version 3.4MR1
No | Visit | Complaint | RLTD | Ellis Dickson Tel: 07427 609947 | Inlet probe bent and gasket damaged on further investigation. | Replaced both gasket and probe needle. Full pressure test passed and cals $qc’s passed. Analyser Okay. | SWO-402576 | Defect | HW | Auto Inlet module | Probe bend | 1391 | Ellis Dickson Tel: 07427 609947 | Inlet probe bent and gasket damaged on further investigation. | Replaced both gasket and probe needle. Full pressure test passed and cals $qc’s passed. Analyser Okay.
No | Visit | Complaint | RSCH | Kal Backlog Fehler auf mehrere Parameter. | Kal Backlog Fehler auf mehrere Parameter. | Allgemeine Kontrolle ok. Kunde hat Sensorkassette ausgetauscht, Probleme mit PCO2, nochmals Ausgetauscht alle Elektrolyten Kal Backlog Fehler, Sensorkassette ausgebaut Stifte überprüft und gereinigt. Sensorkassette nochmals installiert ohne Erfolg. Neue Sensorkassette eingebaut ok. System i.o. Kal. + QC i.o. | SWO-402438 | Assisted Replacement | SC | SC | Multi | backlog | Kal Backlog error on multiple parameters. | Kal Backlog error on multiple parameters. | General control ok. Customer has replaced sensor cartridge, problems with PCO2, again replaced all electrolytes Kal Backlog error, sensor cartridge removed pins checked and cleaned. Sensor cartridge reinstalled without success. New sensor cartridge built-in ok. System i.o. Kal. + QC i.o.
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 废液口堵塞 | 更换试剂包后修复 | SWO-402858 | [5 empty] | Test card, reagent package failure | Waste mouth blockage | Repair after replacement of reagent kit
No | Visit | Complaint | RAME | Customer says they flushed a clot out of sensor cassette, now its leaking from the waste port. | Customer tried to run QC this morning. PO2 still out of range on new sensor cassette. | Reactivated PO2 parameter and ran system cycle. PO2 is still out of range. Will order a new sensor cassette for overnight. | SWO-402798 | [5 empty] | Customer says they flushed a clot out of sensor cassette, now its leaking from the waste port. | Customer tried to run QC this morning. PO2 still out of range on new sensor cassette. | Reactivated PO2 parameter and ran system cycle. PO2 is still out of range. Will order a new sensor cassette for overnight.
No | Visit | Complaint | RAME | Customer says they flushed a clot out of sensor cassette, now its leaking from the waste port. | Checked waste port, sensor cassette, fluid pathway. Flushed all tubing. | Performed zero calibration after flushing all lines. Still leaking. Replaced all pumps, gaskets, and tubings. Also replaced sensor cassette with customer supply. Will submit for a credit on remaining tests. | SWO-402209 | [5 empty] | Customer says they flushed a clot out of sensor cassette, now its leaking from the waste port. | Checked waste port, sensor cassette, fluid pathway. Flushed all tubing. | Performed zero calibration after flushing all lines. Still leaking. Replaced all pumps, gaskets, and tubings. Also replaced sensor cassette with customer supply. Will submit for a credit on remaining tests.
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 响应错误 1310 | 更换测试卡正常。 | SWO-404420 | Assisted Replacement | SC | SC | Glu/Lac | 1310 | Test card, reagent package failure | Response Error 1310 | The replacement test card is OK.
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | GLU,LAC定标不过 0375 0378 1396 1340 1346 | 更换测试卡、试剂包后修复 | SWO-402861 | [5 empty] | Test card, reagent package failure | GLU, LAC calibration is no more than 0375 0378 1396 1340 1346 | Repair after replacing test cards and reagent kits
No | Visit | Complaint | RCHN | 报错1259,1258,1223 | 报错1259,1258,1223，借903-329一个 | 更换903-329后修复by tim | SWO-401651 | Defect | HW | Aspiration Module | Needle Wash Unit | 1259, 1258, 1223 | Error 1259,1258,1223 | 1259,1258,1223, loan 903-329 one | Repair by tim after replacement 903-329
No | Visit | Complaint | RCHN | 报错 5784 | 报错 5784,借 944-212*14402 一张 | 重新定标后修复by tim | SWO-402286 | Defect | SW | Analyzer Control Module | CPU | 5784 | Error 5784 | Error 5784, borrowed 944-212 x 14402 one | Fix by tim after re-calibration
No | Visit | Complaint | RIND | Parameter failure | Parameter failure | Flush out the cassette re-install it, parameter not ready , install a new sensor cassette ,calibration passed. | SWO-404878 | [5 empty] | Parameter failure | Parameter failure | Flush out the cassette re-install it, parameter not ready , install a new sensor cassette ,calibration passed.
No | Visit | Complaint | RDEU | 20.08.2018; Hr. Mageira; Tel.: 059311524482; Upper Liquid Sensor Error, Inlet LS Fehler | Gasventil defekt | Gerät überprüft, Gasventile ausgebaut, El $Met Gasventilmembran ersetzt, pH $BG Gasventil ausgetauscht, Kal und QC o.k. | SWO-402442 | Defect | HW | pH/BG Module | Valve, air |  | 20.08.2018; Hr. Mageira; Tel.: 059311524482; Upper Liquid Sensor Error, Inlet LS Fehler | Gas valve defective | Device checked, gas valves removed, El $Met gas valve membrane replaced, pH $BG gas valve replaced, Kal and QC ok.
No | Visit | Complaint | RDEU | Geräusche beim anfahren der Endposition des FlexQ | FlexQ Geräusche beim anfahren der Endposition | Gerät überprüft, FlexQ kontrolliert, Abgleich durchgeführt | SWO-402441 | Adjustment | HW | FlexQ | FlexQ |  | Noise when approaching the end position of the FlexQ | FlexQ noises when starting the end position | Device checked, FlexQ checked, matching performed
Yes | 2 | Y | RPAC | Blank | 6 Mth service. |  | D4D03C88E2577121 | [5 empty] | Blank | 6 Mth service.
No | Visit | Complaint | RAME | Leak detected error code 1342 | Found inlet gasket with holder not seated properly on solution pack. Customer replaced inlet gasket with holder when error message first occurred with out success. Pressure test failing even with original gasket on unit. THB calibration due. | Replaced inlet gasket with holder. Ran all fluid transport diagnostics with pass results. Ran tHB calibration with pass results. Startup restarted. Calibrations failing PCO2 sensitivity after restart. Sensor cassette near end of life. Replaced sensor cassette. Wait for startup to complete. Lac slow to calibrate. Ran additional calibrations until settled. | SWO-403335 | Defect | HW | Inlet module | Inlet gasket | 1342 | Leak detected error code 1342 | Found inlet gasket with holder not seated properly on solution pack. Customer replaced inlet gasket with holder when error message first occurred with out success. Pressure test failing even with original gasket on unit. THB calibration due. | Replaced inlet gasket with holder. Ran all fluid transport diagnostics with pass results. Ran tHB calibration with pass results. Startup restarted. Calibrations failing PCO2 sensitivity after restart. Sensor cassette near end of life. Replaced sensor cassette. Wait for startup to complete. Lac slow to calibrate. Ran additional calibrations until settled.
Yes | 3 | Y | RPAC | Blank | Replaced SC x 2 |  | D4D2BF8CCE16121B | [5 empty] | Blank | Replaced SC x 2
No | Visit | Complaint | RAME | Customer called stating analyzer will not boot up getting blue screen with white writing. | Found that analyzer coming up with blue screen | Disconnected everything from analyzer and rebooted the analyzer corrected the issue then had to wait for pCO2 to come in. | SWO-403374 | Adjustment | SW | Windows | Blue screen |  | Customer called stating analyzer will not boot up getting blue screen with white writing. | Found that analyzer coming up with blue screen | Disconnected everything from analyzer and rebooted the analyzer corrected the issue then had to wait for pCO2 to come in.
Yes | Visit | Complaint | RAME | Possible Fluid Blockage errors | Fluid blockage error | Tried pumping fluids using diagnostic feature. It didn’t work so took the wet section out and found small tube came off the nozzle. Cut small piece in the beginning and reseat it. All calibrations and 3 Levels of QC passed as per manufacturer specification | SWO-403485 | [5 empty] | Possible Fluid Blockage errors | Fluid blockage error | Tried pumping fluids using diagnostic feature. It didn’t work so took the wet section out and found small tube came off the nozzle. Cut small piece in the beginning and reseat it. All calibrations and 3 Levels of QC passed as per manufacturer specification
No | Visit | Complaint | RCHN | 响应错误 | 响应错误 | 更换903-581后修复by tim | SWO-403596 | Defect | HW | Basic Unit | PCB, Sensor interface | Glu/Lac, 1310 | Wrong response | Wrong response | Repair by tim after replacement 903-581
No | Visit | Complaint | RCHN | 报错4548 | 报错4548 借903-028一个 | 更换903-028后修复by tim | SWO-402180 | Defect | HW | Cover and Chassis | PCB, Top | 4548 | 4548 error | 4548 Loan 903-028 one | Repair by tim after replacement 903-028
No | Visit | Complaint | RCHN | 卡盘 | 卡盘 借580-410一个 | 更换580-410后修复by tim | SWO-401741 | Defect | HW | Storage Module | Motor, wheellock |  | Chuck | Card, borrow 580-410 one. | Repair by tim after replacement 580-410
No | Visit | Complaint | RCHN | 样本量失败 | 样本量失败，需要借件905-953和903-329各一个 | 备件未使用，清洁机器，插拔线路后修复by tim | SWO-401780 | Cleaned | HW | Aspiration Module | Needle Wash Unit |  | Sample size failed | Sample size failure, need to borrow 905-953 and 903-329 each one | Spare parts not used, cleaning machine, repair by tim after wiring is plugged in
No | Visit | Complaint | RDEU | Fehlermeldung 316 Nein number Tool Löw und 258 Neon Lamp Intensität too Löw | Fehler Neon Lamp Intensits 258 und Neon number too low 363 | Oxi complete Test durchgeführt Fehler bei Neon Lamp Intensität bei 48% - Neon Lamp ersetzt, keine Änderung - Oxi PCB ersetzt Neon Lamp Intensität bei 51%, Oxi Modul ersetzt Neon Lamp Intensity bei 85%, Kal und QC durchgeführt | SWO-403627 | Defect | HW | Oxi module | Oxi module | 316, 258 | Error message 316 No number Tool Löw and 258 Neon Lamp Intensity too Löw | Fehler Neon Lamp Intensits 258 und Neon number too low 363 | Oxi complete test performed error at Neon Lamp Intensity at 48% - Neon Lamp replaced, no change - Oxi PCB replaces Neon Lamp Intensity at 51%, Oxi Module replaces Neon Lamp Intensity at 85%, Kal and QC performed
Yes | 9 | Y | RPAC | Blank | inlet not lowering after capilary sample |  | D4D283DDCE05BB82 | [5 empty] | Blank | inlet not lowering after capilary sample
No | Visit | Complaint | RLTD | Linda Stocks Tel. 07950 892250 | Error found to be due to a blockage in the right hand sensor bay pipe towards the inlet. Sensor bay had to be removed before it was possible to clear the blockage. Blockage substance not observed, assumed eventually dissolved with water being flushed through. Re-assembled analyser and tested operation, all now ok. | Removed blockage from sample pathway. | SWO-404014 | Cleaned | Fluidic pathway | Fluidic pathway | Pump Cal | Clot | Linda Stocks Tel. 07950 892250 | Error found to be due to a blockage in the right hand sensor bay pipe towards the inlet. Sensor bay had to be removed before it was possible to clear the blockage. Blockage substance not observed, assumed eventually dissolved with water being flushed through. Re-assembled analyser and tested operation, all now ok. | Removed blockage from sample pathway.
No | Visit | Complaint | RLTD | Keymea contacted us whilst on site. | Removed actuator unit from analyser and refitted front Solution Pack guide using a slightly larger screw. Front guide now secured correctly. Also failing with intermittent inhomogeneous solution errors, fault traced to damaged inlet gasket. Fitted new inlet gasket and restarted. Analyser now working correctly. | Refitted front guide for solution pack. Replaced inlet gasket. | SWO-404016 | Defect | HW | Basic Unit | SP guide |  | Keymea contacted us whilst on site. | Removed actuator unit from analyser and refitted front Solution Pack guide using a slightly larger screw. Front guide now secured correctly. Also failing with intermittent inhomogeneous solution errors, fault traced to damaged inlet gasket. Fitted new inlet gasket and restarted. Analyser now working correctly. | Refitted front guide for solution pack. Replaced inlet gasket.
No | Visit | Complaint | RAME | Connectivity issues. | Unit can't be pinged. Unit will not connect to Aqure. | Adjusted settings to allow for ping.. Still no connection to Aqure. Reseated oxi PCB cables to clear unit not passing tHb zero. Upon recheck settings reverted. Reset again. | SWO-404030 | [5 empty] | Connectivity issues. | Unit can't be pinged. Unit will not connect to Aqure. | Adjusted settings to allow for ping.. Still no connection to Aqure. Reseated oxi PCB cables to clear unit not passing tHb zero. Upon recheck settings reverted. Reset again.
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 所有离子定标失败，1318，1319,1315,1320,1321 | 更换测试卡后修复 | SWO-403612 | [5 empty] | Test card, reagent package failure | All ion calibration failures, 1318,1319,1315,1320,1321 | Repair after replacing test card
No | Visit | Complaint | RESP | Fuga de líquidos parte trasera | Se detecta fuga de líquidos parte trasera por obstrucción de tubo N6 bomba P4 | Sustitución de tubos y bomba así como pruebas de funcionamiento, y limpieza de la zona. Calibrado y verificado con QC | SWO-403878 | Defect | HW | Wet section | Tube #6 | blocked | Leakage of liquids rear | Liquid leakage rear edifold from N6 tube obstruction P4 pump | Replacement of tubes and pump as well as operational tests, and cleaning of the area. Calibrated and verified with QC
No | Visit | Complaint | RCHN | ca定标失败，报错5684 | 报错5684，需要借件944-214（14220） | 重新固定处理盘卡杯的卡槽后修复by tim | SWO-402840 | Defect | Consumables | Solution Pack | SP, stacking | 5684 | ca bid failure, error 5684 | 5684, need to borrow 944-214 (14220) | Fix by tim after refixing the card slot of the disc cup
No | Visit | Complaint | RKK | ミキサーエラー頻発 | ミキサーエラー頻発 | FLEXQモジュール修理完了の為、返却。代替えモジュール引き上げ。FLEXQモジュール調整実施。Glu、Lac、Refメンブラン及びインレットガスケット定期交換。 | SWO-404159 | Defect | HW | FlexQ | Mixer Tray |  | Mixer errors frequently | Mixer errors frequently | Returned for flexq module repair completion. Alternative module raise. FlexQ module adjustment implemented. Glu, Lac, Ref Membrane and Inlet gasket supine replacement.
No | Visit | Complaint | RAME | The pH calibrations are intermittently drifting. Found pH chamber full of dried blood and dried blood outside the measuring chamber. | Liquid leaking from the front overflow spout on the measuring chamber module. Found the pH and pO2 electrode sealing gaskets missing and the pCO2 sealing gasket worn. | Replaced the missing gaskets. Cleaned the spillages from all the modules. Found the Hb lamp out of ideal adjustment. Adjust to 3402. PO2 fails calibrations with high zero value. Replaced pO2 membrane. | SWO-404019 | Defect | Electrode | Electrode | Sealing ring | pH, pO2 | The pH calibrations are intermittently drifting. Found pH chamber full of dried blood and dried blood outside the measuring chamber. | Liquid leaking from the front overflow spout on the measuring chamber module. Found the pH and pO2 electrode sealing gaskets missing and the pCO2 sealing gasket worn. | Replaced the missing gaskets. Cleaned the spillages from all the modules. Found the Hb lamp out of ideal adjustment. Adjust to 3402. PO2 fails calibrations with high zero value. Replaced pO2 membrane.
No | Visit | Complaint | RAME | PH calibration failing due to drift. | Reviewed data logs, inspected pH electrode, oring, measuring chamber. Inspected remaining components for the pH $BG module. Suspect possible problem with pH sealing oring. | Replaced pH electrode sealing oring. Cleaned all electrode connector caps. Ran multiple calibration cycles, drift stabilized, 2PT and 1 PT calibration passing. | SWO-403500 | Defect | Electrode | Electrode | Sealing ring | ph | PH calibration failing due to drift. | Reviewed data logs, inspected pH electrode, oring, measuring chamber. Inspected remaining components for the pH /BG module. Suspect possible problem with pH sealing oring. | Replaced pH electrode sealing oring. Cleaned all electrode connector caps. Ran multiple calibration cycles, drift stabilized, 2PT and 1 PT calibration passing.
No | Visit | Complaint | RPOL | Cieknie z aparatu. | Wizualna ocena | Założono wężyk ściekowy, który spadł. Wymiana zgiętej igły oraz uszczelki wlotu. Wymiana elektrody ref oraz membran: ref, Glu i Lac zgodnie z harmonogramem. Odbiałczanie, kalibracja tHb. | SWO-403716 | Defect | HW | Basic Unit | Waste tube |  | Liquid from the camera. | Visual rating | The sewage hose was assumed to have fallen. Replacing the bent needle and inlet gasket. Ref electrode replacement and membrane: REF, Glu and Lac according to schedule. Deleukation, calibration of tHb.
Yes | Visit | Complaint | RAME | Mixer issues | Mixer not working | Found mixer cable inserted upside down after pm. Fixed issue. | SWO-404719 | [5 empty] | Mixer issues | Mixer not working | Found mixer cable inserted upside down after pm. Fixed issue.
No | Visit | Complaint | RDEU | Kunde meldet das Seit einigen Tagen die Kalibration der SC scheitert. | Kalibrationen werden nicht bestanden. | Fehlersuche am Analysator durchgeführt. Mehrere Pins von dem Sensorinterface Bord sind schief. PCB getauscht. Analysator wieder hochgefahren und Kalibrationen $ QC durchgeführt. Neue SC wieder eingesetzt. Startphase abgewartet. Mehrere Kalibrationen durchgeführt und danach die QC. | SWO-404314 | Defect | HW | Basic Unit | PCB, Sensor interface | contacts bend | Customer reports that the calibration of the SC fails for a few days. | Calibrations are not passed. | Troubleshooting is performed on the analyzer. Several pins from the sensor interface board are skewed. PCB. Analyzer restarted and calibrations performed by QC. New SC reinstated. start-up phase. Several calibrations were performed and then the QC.
No | Visit | Complaint | RDEU | 08.08.2018 10:20 Fr. Judäa 06217591319: Interventionsmeldung Flüssigkeiten können nicht gepumpt werden 09.08.2018, 11:45 Uhr ==> Anrufer Fr. Bianka Judäa, Tel.: 0621759131 ==> Analysator befindet sich weiterhin in der Fehlerbehebungsroutine 13.08.2018, 14.20Uhr, Frau Judea, Tel: 06217591319: weiterhin Wechselschleife- die Anwender haben bereits alle getauscht | 1169 Lösungen können nicht gepumpt werden Bestellung SP und SC eingetroffen SC und SP Erneuerung notwendig | Einlass Überprüfung (Einlassnadel und Einlassdichtung bereits von Anwenderin erneuert) - Rollenkranz etc. Überprüfung - Erneuerung SP und SC - Überprüfung Flüssigkeitstransport im Service - Funktionsüberprüfung | SWO-404437 | Assisted Replacement | SP | SP | FT | 1169 | 08.08.2018 10:20 Fr. Judea 06217591319: Intervention notification Liquids cannot be pumped 09.08.2018, 11:45 ==> Caller Ms. Bianka Judea, tel.: 0621759131 ==> Analyzer is still in the troubleshooting routine 13.08.2018, 14.20h, Mrs. Judea, Tel: 06217591319: continue alternating loop - the users have already changed all | 1169 solutions cannot be pumped Order SP and SC arrived SC and SP renewal necessary | Inlet check (inlet needle and inlet seal already renewed by user) - roller wreath etc. Review - Renewal SP and SC - Inspection of liquid transport in service - Functional check
No | Visit | Complaint | RAME | ANALYZER WILL NOT POWER UP. SUSPECTED CPU ERROR. | ANALYZER APPEARS TO NOT POWER UP CPU $ SCREEN. | REPLACED CPU, POWER SUPPLY AND WET SECT CONTROL BOARD. DETERMINED 5V POWER BEING DRAWN DOWN TO CPU AND NOT ALLOWING IT OR WET SECT BOARD TO POWER PROPERLY. DISCONNECTED PRINTER ASSBLY AND RESOLVED POWER UP ISSUES. PRINTER TO BE REPLACED BY DIFFERENT FSE THIS MONDAY. ANALYZER IS NOT QUALIFIED AT THIS TIME. | SWO-404783 | Revisit required | HW | Printer Module | Printer |  | ANALYZER WILL NOT POWER UP. SUSPECTED CPU ERROR. | ANALYZER APPEARS TO NOT POWER UP CPU / SCREEN. | REPLACED CPU, POWER SUPPLY AND WET SECT CONTROL BOARD. DETERMINED 5V POWER BEING DRAWN DOWN TO CPU AND NOT ALLOWING IT OR WET SECT BOARD TO POWER PROPERLY. DISCONNECTED PRINTER ASSBLY AND RESOLVED POWER UP ISSUES. PRINTER TO BE REPLACED BY DIFFERENT FSE THIS MONDAY. ANALYZER IS NOT QUALIFIED AT THIS TIME.
No | Visit | Complaint | RDEU | 26.06.2018 Herr Betsch 015204767315 meldet, dass die Anbindung nicht funktioniert. | Netzwerkverbindung ist weiterhin nicht möglich | Das Gerät via ASTMNet verbunden. Verbindung ist fehlerfrei. Folgende Adresse am Gerät eingestellt: 10.200.112.181 255.255.255.0 10.200.112.110 Serveradresse 10.100.200.50 Das Gerät kann nicht von außen angepingt werden . Fehler liegt definitiv am Netzwerk. Diese Einsatz wird berechnet. | SWO-404488 | No fault found | SW | External SW | Network |  | 26.06.2018 Mr. Betsch 015204767315 reports that the connection is not working. | Network connection is still not possible | The device is connected via ASTMNet. Connection is error-free. The following address set on the device: 10.200.112.181 255.255.255.0 10.200.112.110 Server address 10.100.200.50 The device cannot be pinged from the outside. Error is definitely due to the network. This bet is calculated.
No | 3 | Y | RPAC | Blank | SC troubleshooting |  | D4D4142A441F4BEE | Assisted Replacement | SC | SC | Lac | Calibration | Blank | SC troubleshooting
No | Visit | Complaint | RLTD | Sally on 02392 286000 ex 3232 | Checked analyser logs, calibrations and quality controls have been passing without errors. Checking patient result log for today - 67 samples have been run. Of these: 54 samples have been performed without any errors 7 had ? on one or more parameters, mostly K+ response or out of range errors. 2 were Inhomogeneous - air bubble in sample 4 aborted - unable to aspirate sample. Analyser is working correctly. Discussed with Sally Veness possible reasons for the failed samples, including samples potentially clotting if not taken directly into heparinized sampling device and also discussed possibilities of hemolysing the sample when transferring from syringe to capillary. There could also be other pre-analytical reasons that are contributing. | None required - Analyser performing correctly. | SWO-404552 | User error | preanalytical | preanalytical | Sample handling |  | Sally on 02392 286000 ex 3232 | Checked analyser logs, calibrations and quality controls have been passing without errors. Checking patient result log for today - 67 samples have been run. Of these: 54 samples have been performed without any errors 7 had ? on one or more parameters, mostly K+ response or out of range errors. 2 were Inhomogeneous - air bubble in sample 4 aborted - unable to aspirate sample. Analyser is working correctly. Discussed with Sally Veness possible reasons for the failed samples, including samples potentially clotting if not taken directly into heparinized sampling device and also discussed possibilities of hemolysing the sample when transferring from syringe to capillary. There could also be other pre-analytical reasons that are contributing. | None required - Analyser performing correctly.
No | Visit | Complaint | RKOR | Inlet probe bent | . | Inlet probe position adjust. | SWO-404955 | Adjustment | HW | Auto Inlet module | Probe bend |  | Inlet probe bent | . | Inlet probe position adjust.
No | Visit | Complaint | RLTD | Peter Williams Tel: 01243 788122 x: 33591 - please call after 1pm | Investigated data logs, pCO2 Quality controls were failing high from 15 $08 to 20 $08. 3 new Sensor cassettes and one in use SC from another analyser tried with no change noted, different solution pack also tried. Then suddenly on re-insertion of Sensor Cassette already tried at 16:50 on 20 $08 the Quality controls then were ok. Checked barometer, no error found. Ran Sensor simulator test to verify channels - all channels passed. Refitted Sensor Cassette and verified performance - initial SC substrate temperature errors experienced, resolved by recalibration thermistor in service programs. Calibrations passed, Quality Controls passed and pCO2 QC results mid range. Unable to explain earlier QC failures on pCO2, so data taken for review by Radiometer Copenhagen. | Analyser tested and no fault found. | SWO-404431 | Adjustment | SC | SC | Thermistor | recalibrated | Peter Williams Tel: 01243 788122 x: 33591 - please call after 1pm | Investigated data logs, pCO2 Quality controls were failing high from 15 /08 to 20 /08. 3 new Sensor cassettes and one in use SC from another analyser tried with no change noted, different solution pack also tried. Then suddenly on re-insertion of Sensor Cassette already tried at 16:50 on 20 /08 the Quality controls then were ok. Checked barometer, no error found. Ran Sensor simulator test to verify channels - all channels passed. Refitted Sensor Cassette and verified performance - initial SC substrate temperature errors experienced, resolved by recalibration thermistor in service programs. Calibrations passed, Quality Controls passed and pCO2 QC results mid range. Unable to explain earlier QC failures on pCO2, so data taken for review by Radiometer Copenhagen. | Analyser tested and no fault found.
No | Visit | Complaint | RCHN | 废液口堵塞 | 废液口堵塞 | 清洁管路后修复by tim | SWO-404220 | [5 empty] | Waste mouth blockage | Waste mouth blockage | Repair by tim after cleaning lines
No | Visit | Complaint | RAME | Customer states that solution packs are having to be changed out ~ every 2 weeks and they are having to order additional solution packs over their PCA shipments. | During service call for a separate issue, the customer mentioned that they are back to changing out the solution packs every 2 weeks and they should be lasting longer. In May a service call was made for the same complaint. Analyzer was reset to factory defaults and manually set up. | Due to this being the second complaint, a Customer Quality Case will be opened. Saved data logs and data base for submission to manufacturer. | SWO-404663 | [5 empty] | Customer states that solution packs are having to be changed out ~ every 2 weeks and they are having to order additional solution packs over their PCA shipments. | During service call for a separate issue, the customer mentioned that they are back to changing out the solution packs every 2 weeks and they should be lasting longer. In May a service call was made for the same complaint. Analyzer was reset to factory defaults and manually set up. | Due to this being the second complaint, a Customer Quality Case will be opened. Saved data logs and data base for submission to manufacturer.
No | Visit | Complaint | RCHN | 报错 1035 | 报错 1035,借944-394*14412一张。 | 重新定标正常。by tim | SWO-402957 | Defect | HW | Process Module | Optical Unit | 1035 | Error 1035 | 1035, borrow 944-394 x 14412. | Recalibration is normal. by tim
No | Visit | Complaint | RCHN | 错误1999,4036, | 1999，4036 | 微调进样盘 | SWO-404267 | Adjustment | HW | Aspiration Module | Inlet Wheel | 4036 | Error 1999,4036, | 1999，4036 | Fine-tuning the sample disk
No | Visit | Complaint | RCHN | 报错 1031，4548 | 报错 1031，4548,借903-028，903-364各一个。 | 更换903-028和903-364后修复by tim | SWO-402997 | Defect | HW | Process Module | Optical Unit | 1031 | Error 1031,4548 | Error 1031,4548, borrow 903-028, 903-364 each. | Repair by tim after replacing 903-028 and 903-364
No | Visit | Complaint | RCHN | 报错1031，1024 | 报错1031，1024,借 903-364一个，944-212*14642，944-214*14465，944-216*14616，944-258*14738，944-394*14412 | 更换903-364一个，重新定标正常。by tim | SWO-402954 | Defect | HW | Process Module | Optical Unit | 1031, 1024 | Error 1031,1024 | Error 1031,1024, borrowing 903-364, 944-212 x 14642, 944-214 x 14465, 944-216 x 14616, 944-258 x 14738, 944-394 x 1412 | Replace the 903-364 one and re-label normally. by tim
No | Visit | Complaint | RDEU | 24.08.2018, Herr Markosch MT/ Sr. Claudia , Tel: 02064422000: AutoCheck Ampullen werden nicht alle gelesen | QC Ampullen werden nicht zuverlässig gelesen, Na QC Fehler Level 1 | Auto QC Modul ausgebaut und gereinigt, Einlassnadel stark verbogen - Einlassnadel ersetzt, Einlassdichtung aus Kundenbestand ersetzt, Lüftungskanal gespült, Messkammern gereinigt, Kontakt des Ref Konnektors oxidiert - Kontakte des Konnektors und der Ref Elektrode gereinigt, Kal und QC durchgeführt, | SWO-404923 | Cleaned | HW | Autocheck module | Autocheck Module | Revisit? | 24.08.2018, Mr. Markosch MT/ Sr. Claudia , Tel: 02064422000: AutoCheck ampoules are not all read | QC ampoules are not read reliably, Na QC error level 1 | Auto QC module removed and cleaned, inlet needle strongly bent - inlet needle replaced, inlet seal replaced by customer stock, ventilation duct flushed, measuring chambers cleaned, contact of the ref connector oxidized - contacts of the connector and the ref electrode cleaned , Kal and QC performed,
No | Visit | Complaint | RIND | Calibration aborted | Checked the data logs calibration was getting aborted found pH $BG air valve was defective | Need to replace pH $BG air valve | SWO-404718 | Defect | HW | pH/BG Module | Valve, air |  | Calibration aborted | Checked the data logs calibration was getting aborted found pH /BG air valve was defective | Need to replace pH $BG air valve
No | Visit | Complaint | RAME | Cl out on QC | Cl out on cals and QC. Ref electrode not fully seated and pH electrode contaminated and chamber. | Remembraned Ref electrode and replaced the sealing ring on the pH electrode. Cleaned the chamber. Ran calibrations and all levels of QC. All tests passed. | SWO-404797 | Assisted Replacement | Electrode | Membrane | Reference |  | Cl out on QC | Cl out on cals and QC. Ref electrode not fully seated and pH electrode contaminated and chamber. | Remembraned Ref electrode and replaced the sealing ring on the pH electrode. Cleaned the chamber. Ran calibrations and all levels of QC. All tests passed.
No | Visit | Complaint | RAME | Oximetry Errors 258 and 316 Neon Errors Contact is Aida or Cheryl at 312-864-7408 | Found OXI hardware error 579. | Once I ran complete OXI test found lamp unit failed. Replaced entire OXI module. Ran lamp adjustment. Complete OXI test. All calibrations and 3 Levels of QC passed as per manufacturer specification | SWO-404814 | Defect | HW | Oxi module | Oxi module | 258, 316 | Oximetry Errors 258 and 316 Neon Errors Contact is Aida or Cheryl at 312-864-7408 | Found OXI hardware error 579. | Once I ran complete OXI test found lamp unit failed. Replaced entire OXI module. Ran lamp adjustment. Complete OXI test. All calibrations and 3 Levels of QC passed as per manufacturer specification
No | Visit | Complaint | RAME | OBF: Neon intensity error | Neon lamp position failed | Replaced oximetry module, installed calibrated ran thb cal verified | SWO-404736 | Defect | HW | Oxi module | Oxi module |  | OBF: Neon intensity error | Neon lamp position failed | Replaced oximetry module, installed calibrated ran thb cal verified
No | Visit | Complaint | RDEU | 24.08.2018; Fr. Schupfner; Tel.: 09471705455; Totalausfall | AQT lässt sich einschalten. Inlet Door schließt nicht korrekt. | Inlet door ausgebaut und justiert. Systemüberprüfung und Systemreinigung ok. LQC‘s (1x pro Parameter) durchgeführt - ok. | SWO-404634 | Adjustment | HW | Cover and Chassis | Door, Inlet |  | 24.08.2018; Ms. Schupfner; Tel.: 09471705455; Total failure | AQT can be switched on. Inlet Door does not close correctly. | Inlet door expanded and adjusted. System inspection and system cleaning ok. LQC's (1x per parameter) performed - ok.
No | Visit | Complaint | RCHN | 试剂包无法读取 | 试剂包无法读取，借903-158一个 | 更换903-158 后修复by tim | SWO-404167 | Defect | HW | Basic Unit | PCB Solution Pack ID |  | Reagent package cannot be read | Reagent package can not be read, borrow 903-158 one | Repair by tim after replacing 903-158
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | LAC,GLU超限度 CO2,CL定标超响应范围 | 连续更换2张测试卡后修复 | SWO-404886 | [5 empty] | Test card, reagent package failure | LAC, GLU over limit CO2, CL calibration over-response range | Repair after 2 consecutive test cards
Yes | Visit | Complaint | RCHN | 报错1031 | 报错1031，借944-212（14466）一个 | 重新定标后修复by tim | SWO-403561 | Defect | HW | Process Module | Optical Unit | 1031 | 1031 | 1031, loan 944-212 (14466) one | Fix by tim after re-calibration
No | Visit | Complaint | RAME | Analyzer on hold with error message leak detected. | Leaking out the ph chamber | Found oring on ph not seated--installed oring correctly which fixed the leaking issue. PO2 and pco2 both due remembraning -- remembraned both electrodes; QC schedule for new lot not scheduled correctly, updated the QC schedule. THb cal over due, Ran tHb cal. Ran QC and lactate failing low on level 3-- remembraned Lactate electrode -- Ca and Lac having drifts on cals, remembraned the Ref and the Ca. Talked to Amanda about letting the Lactate stabilize over night. Michelle is working tomorrow--I will follow up with Michelle tomorrow morning | SWO-404790 | Defect | Electrode | Electrode | Sealing ring, pH |  | Analyzer on hold with error message leak detected. | Leaking out the ph chamber | Found oring on ph not seated--installed oring correctly which fixed the leaking issue. PO2 and pco2 both due remembraning -- remembraned both electrodes; QC schedule for new lot not scheduled correctly, updated the QC schedule. THb cal over due, Ran tHb cal. Ran QC and lactate failing low on level 3-- remembraned Lactate electrode -- Ca and Lac having drifts on cals, remembraned the Ref and the Ca. Talked to Amanda about letting the Lactate stabilize over night. Michelle is working tomorrow--I will follow up with Michelle tomorrow morning
No | Visit | Complaint | RAME | instructed to reboot analyzer | Neon lamp error | Replaced oxy pcb and lamp unit -- updated QC schedule new lot not scheduled, removed old lot from the schedule, remembraned PO2 and PO2 both due, ran complete oxy test and tHb. Lamp adjusted to 3223 | SWO-404794 | Defect | HW | Oxi module | Lamp unit | neon lamp | instructed to reboot analyzer | Neon lamp error | Replaced oxy pcb and lamp unit -- updated QC schedule new lot not scheduled, removed old lot from the schedule, remembraned PO2 and PO2 both due, ran complete oxy test and tHb. Lamp adjusted to 3223
No | Visit | Complaint | RCHN | 接口板故障 | 接口板故障，借902-972一个 |  | SWO-404967 | [5 empty] | Interface board failure | Interface board failure, borrow 902-972 one
No | Visit | Complaint | RCHN | 报错0581 | 报错0581，借902-980一个 |  | SWO-404960 | Defect | HW | Oxi module | Hemolyzer | 581 | Error 0581 | Error 0581, loan 902-980 one
No | Visit | Complaint | RCHN | 报错2168 | 报错2168，借903-138一个 | 更换903-138后修复by tim | SWO-404393 | Defect | HW | Incubator Module | PCB, Incubator | 2168 | 2168 error | 2168, loan 903-138 one | Repair by tim after replacement 903-138
No | Visit | Complaint | RAME | ANALYZER WILL NOT POWER UP. SUSPECTED CPU ERROR. | Returned to resume the repair hoping that a replacement printer will fix the issue. Put the printer on, issue remained. Tried new printer power harness cable and usb connector btw cpu and printer, issue remained. Tried second printer. Issue remained. Got on three-way call with Dennis and Matt. Suggested disconnecting all cables and connectors at wet section pcb, leaving printer connected. Issue remained. Reconnected all cables and connectors at wet section pcb, power up analyzer with printer disconnected. Work well. Tried reconnecting printer power harness, analyzer shuts down and left fuse on wet section board off. | Informed both Pablo and Debbie about status, analyzer will run without printer until next course of action has been relayed to FSE, HupChew. | SWO-405515 | Revisit required | HW | DMS_PC | CPU |  | ANALYZER WILL NOT POWER UP. SUSPECTED CPU ERROR. | Returned to resume the repair hoping that a replacement printer will fix the issue. Put the printer on, issue remained. Tried new printer power harness cable and usb connector btw cpu and printer, issue remained. Tried second printer. Issue remained. Got on three-way call with Dennis and Matt. Suggested disconnecting all cables and connectors at wet section pcb, leaving printer connected. Issue remained. Reconnected all cables and connectors at wet section pcb, power up analyzer with printer disconnected. Work well. Tried reconnecting printer power harness, analyzer shuts down and left fuse on wet section board off. | Informed both Pablo and Debbie about status, analyzer will run without printer until next course of action has been relayed to FSE, HupChew.
No | Visit | Complaint | RDEU | 24.08.18 14:25 Herr Bär Tel. 03447/52-3007 , Trübung bzw Oxi miss Match Error bei Patientenmessungen | Fehlersuche ==> Fehler 0581: Oxi spectrum mismatch bei Patientenproben | Hemolyser unit ausgetauscht | SWO-405463 | Defect | HW | Oxi module | Hemolyzer | 581 | 24.08.18 14:25 Mr Bär Tel. 03447/52-3007 , turbidity or Oxi miss Match Error in patient measurements | Troubleshooting ==> Error 0581: Oxi spectrum mismatch in patient samples | Hemolyser unit replaced
No | Visit | Complaint | RDEU | 27.08.2018 Herr Feldt 043116972333 meldet einen pO2 Kal Fehler. 27.08.2018, 11:58 Uhr, H. Feldt, Tel. 043116972333: pO2 Kal Fehler 27.08.2018 14:00 Herr Feldt 043116972333 Fehler bleibt bestehen. | System überprüft => Sensor Interface PCB defekt | Sensor Interface PCB wird bestellt | SWO-405350 | Revisit required | HW | Basic Unit | PCB, Sensor interface | pO2 | 27.08.2018 Mr Feldt 043116972333 reports a pO2 Kal error. 27.08.2018, 11:58, H. Feldt, tel. 043116972333: pO2 Kal Error 27.08.2018 14:00 Mr Feldt 043116972333 Error remains. | System checked => Sensor Interface PCB defective | Sensor Interface PCB is ordered
No | Visit | Complaint | RAME | Connectivity issues. | Customer has questions regarding setup. | Went through setup screens with customer. Could not answer all questions. Contacted Lyle Simpson for assistance and meeting setup. Contacted Steve McClain for assistance with Aqure license issue. | SWO-405546 | [5 empty] | Connectivity issues. | Customer has questions regarding setup. | Went through setup screens with customer. Could not answer all questions. Contacted Lyle Simpson for assistance and meeting setup. Contacted Steve McClain for assistance with Aqure license issue.
No | Visit | Complaint | RCHN | 定标失败 | 偶发5684错误导致定标失败 借944-258（13918）一张 | 重新定标后修复（by chuanwei） | SWO-405148 | Defect | SW | Process Module | Process Wheel | 5684 | Bid failed | Occasional 5684 error leads to bid failure Borrowing 944-258 (13918) | Re-calibration repair (bywei chuan)
No | Visit | Complaint | RIND | SC calibration failed frequently | Checked analyzer and SC. We had taken failure data to check it but it is showing 0 number of test remaining for new SC. | Need to check further with Sendex. Sent all data to Lakshmikanth. | SWO-408063 | [5 empty] | SC calibration failed frequently | Checked analyzer and SC. We had taken failure data to check it but it is showing 0 number of test remaining for new SC. | Need to check further with Sendex. Sent all data to Lakshmikanth.
No | Visit | Complaint | RAME | Called to have service go and replace hemolyzer | Refer to SWO-405558 | Replaced hemolyzer. Adjusted lamp unit. Ran complete OXI test and all passed. Ran multiple cals and All calibrations and 3 Levels of QC passed as per manufacturer specification. Tested the hemolyzer with level 3 VK kit. All good Refer to SWO-405558 | SWO-405565 | Defect | HW | Oxi module | Hemolyzer | 581 | Called to have service go and replace hemolyzer | Refer to SWO-405558 | Replaced hemolyzer. Adjusted lamp unit. Ran complete OXI test and all passed. Ran multiple cals and All calibrations and 3 Levels of QC passed as per manufacturer specification. Tested the hemolyzer with level 3 VK kit. All good Refer to SWO-405558
No | Visit | Complaint | RDEU | FlexQ Modul arbeitet nicht richtig | FlexQ Modul bleibt beim fahren hängen und macht Geräusch. Teile aus Modul gefallen. | Metllwinkel mit Lager des Sampler Tray hatten sich gelöst und sind heraus gefallen. Lager wieder eingebaut. FlexQ im Service überprüft. Kalibrationen und QC’s vorgenommen. | SWO-405221 | Defect | HW | FlexQ | FlexQ | sampler tray loose | FlexQ module is not working properly | FlexQ module gets stuck while driving and makes noise. Parts from module. | Metllwinkel with bearing of the sampler tray had dissolved and fell out. Bearing reinstalled. FlexQ in the service. calibrations and QC's.
No | Visit | Complaint | RAME | Customer told sales person Bill about this issue | Came onsite and found OXI mismatch error when running level 3 cal verification | Replaced hemolyzer. Adjusted lamp unit. Ran complete OXI test and all passed. Ran multiple cals and All calibrations and 3 Levels of QC passed as per manufacturer specification. Tested the hemolyzer with level 3 VK kit. All good | SWO-405557 | Defect | HW | Oxi module | Hemolyzer |  | Customer told sales person Bill about this issue | Came onsite and found OXI mismatch error when running level 3 cal verification | Replaced hemolyzer. Adjusted lamp unit. Ran complete OXI test and all passed. Ran multiple cals and All calibrations and 3 Levels of QC passed as per manufacturer specification. Tested the hemolyzer with level 3 VK kit. All good
Yes | 0 | Y | RPAC | Blank | Analyser installation |  | D4DCA4857E4E53F4 | Re classified RMED | installation report |  |  |  | Blank | Analyze installation
No | Visit | Complaint | RKOR | Broken the tube holder | Replace the tube holder Perform the longterm shutdown Insert the solution pack -> Failed Solution pack check -> OK Re-insert the SP -> Failed Replace the SP | Replace the tube holder $ Replace the solution pack | SWO-408158 | Defect | HW | Inlet module | Tube holder |  | Broken the tube holder | Replace the tube holder Perform the longterm shutdown Insert the solution pack -> Failed Solution pack check -> OK Re-insert the SP -> Failed Replace the SP | Replace the tube holder $ Replace the solution pack
No | Visit | Complaint | RKK | 引き上げ修理。フロントパネルを代替に交換 | 引き上げ修理。フロントパネルを代替に交換 | 引き上げ修理。フロントパネルを代替に交換 | SWO-405854 | Defect | HW | Cover and Chassis | Door, Inlet |  | Raise repairs. Replacing the front panel as an alternative | Raise repairs. Replacing the front panel as an alternative | Raise repairs. Replacing the front panel as an alternative
Yes | Visit | Complaint | RDEU | Beim Aufstellen des Analysators wird festgestellt, dass die Inlet Türchen nicht gleichmäßig aufgehen. | Analysator bei Anwender aufgestellt, Konfiguration des Analysators, Durchführung der Qualitätskontrolle. Festgestellt, dass die Inlet doors sich nicht gleichmäßig öffnen und schließen. | Keine Ersatzteile auf dem Fahrzeug. Entscheidung den alten Analysator noch nicht abzuziehen, erst wenn der neue Analysator repariert ist. Ersatzteile bestellt, Analysator in Betrieb. Versucht die vorder Abdeckung durch die alte zu erneuern, Analysator zeigte aber nicht nachvollziehendes Verhalten, daher Entscheidung erst Reparatur, dann Abzug des alten Analysators. | SWO-405440 | Re classified RMED | OOB | Cover and Chassis | Door, Inlet |  | When setting up the analyzer, it is determined that the inlet doors do not rise evenly. | Analyzer set up at user, configuration of the analyzer, execution of quality control. Determined that the inlet doors do not open and close evenly. | No spare parts on the vehicle. Do not decide to withdraw the old analyzer until the new analyzer has been repaired. Spare parts ordered, analyzer in operation. Tried to renew the front cover by the old, but analyzer showed not to understand behavior, therefore decision first repair, then deduction of the old analyzer.
No | Visit | Complaint | RPOL | Mechaniczne uszkodzenie mocowania pakietu roztworów | Uszkodzenie mocowania pakietu roztworów | Nowa śruba mocująca uchwyt pakietu roztworów. | SWO-405392 | Defect | HW | Basic Unit | Mounting Base | screw | Mechanical damage to clamping of the solution package | Damage to the attachment of the solution package | New screw to mount the handle of the solution package.
No | Visit | Complaint | RAME | ANALYZER WILL NOT POWER UP. SUSPECTED CPU ERROR. | Returned with loaner abl90 to identify root cause by swapping parts btw analyzers. | CPU was causing current to be overdrawn. Put on cpu from loaner unit and original printer. Analyzer boots up properly. CPU, pn 910-342 on Order. ETA 9 $2. Will return then to replace it. Printer is currently disconnected until part arrives. | SWO-406128 | Revisit required | HW | DMS_PC | CPU |  | ANALYZER WILL NOT POWER UP. SUSPECTED CPU ERROR. | Returned with loaner abl90 to identify root cause by swapping parts btw analyzers. | CPU was causing current to be overdrawn. Put on cpu from loaner unit and original printer. Analyzer boots up properly. CPU, pn 910-342 on Order. ETA 9 $2. Will return then to replace it. Printer is currently disconnected until part arrives.
No | Visit | Complaint | RAME | Multiple errors | Customer having ongoing errors and struggles with consumables failing. | Performed maintenance and ran diagnostics and calibrations. All tests passed. | SWO-406598 | Defect | HW | Basic Unit | PM Kit |  | Multiple errors | Customer having ongoing errors and struggles with consumables failing. | Performed maintenance and ran diagnostics and calibrations. All tests passed.
No | Visit | Complaint | RDEU | 27.08.2018 Herr Feldt 043116972333 meldet einen pO2 Kal Fehler. 27.08.2018, 11:58 Uhr, H. Feldt, Tel. 043116972333: pO2 Kal Fehler 27.08.2018 14:00 Herr Feldt 043116972333 Fehler bleibt bestehen. | pO2 Kalibrier Fehler | Fehleranalyse, Sensor Interface PCB erneuert, Kalibrierung und Qualitätskontrolle i.O. | SWO-405750 | Defect | HW | Basic Unit | PCB, Sensor interface | pO2 | 27.08.2018 Mr Feldt 043116972333 reports a pO2 Kal error. 27.08.2018, 11:58, H. Feldt, tel. 043116972333: pO2 Kal Error 27.08.2018 14:00 Mr Feldt 043116972333 Error remains. | pO2 Calibration Error | Error analysis, sensor interface PCB renewed, calibration and quality control in the
No | Visit | Complaint | RFRA | Bilirubine plus élevée sur les échantillons que sur l'ABL90flex | Vérification complète | Test complet oximetrie $ $ Nettoyage de l'hemoliseur à l'hypochlorite $ $ relance $ $ calibration thb $ $ passage des contrôles $ $ essais avec du sang | SWO-406841 | Cleaned | Fluidic pathway | Fluidic pathway | Oxi | Bil | Bilirubin higher on samples than on ABL90flex | Full check | Complete test oximetrie $$ Cleaning of the hemoliser to the hypochlorite $$raise $$ calibration thb $$ passage of controls $$ tests with blood
Yes | 4 | Y | RPNZ | Blank | Installation on site |  | D4E52D74B83E9DFE | Re classified RMED | planned SWO |  |  |  | Blank | Installation on site
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 试剂包无法连接 液体量不足 | 连续更换2个试剂包后修复 | SWO-406311 | [5 empty] | Test card, reagent package failure | Reagent pack can't be connected Insufficient liquid | Repair after 2 reagent kits in a row
No | Visit | Complaint | RCHN | CO2灵敏度定标不过 | CO2灵敏度定标不过，借903-581一个 | 更换903-581后修复(by chuanwei) | SWO-405192 | Defect | HW | Basic Unit | PCB, Sensor interface | pCO2 | CO2 sensitivity calibration is however | CO2 sensitivity calibration, however, by 903-581 a | Repair after replacement 903-581 (by chuanwei)
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 漏电流 | 更换试剂包后修复 | SWO-407564 | [5 empty] | Test card, reagent package failure | Leak current | Repair after replacement of reagent kit
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | CA定标不过 1320 CO2定标不过0378 | 更换测试卡后修复 | SWO-406316 | Assisted Replacement | SC | SC | pCO2/Ca | 378/1320 | Test card, reagent package failure | CA calibration but 1320 CO2 calibration is no more than 0378 | Repair after replacing test card
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | CA定标不过 1320 CO2定标不过 0378 CO2定标不过 0378 | 连续更换3张测试卡后修复 | SWO-407585 | Assisted Replacement | SC | SC | pCO2/Ca | 378/1320 | Test card, reagent package failure | CA calibration is no more than 1320 CO2 calibration is no more than 0378 CO2 calibration is no more than 0378 | Repair after 3 consecutive test cards
No | Visit | Complaint | RCHN | 客户反映pH测量结果不准，有异议 | 客户反映pH测量结果不准，有异议 | 客户使用肝素抗凝的试管抽取动脉血，然后再用用注射器分出部分来做样本，客户说以前也是一直这样做的，他们不想改变。通过两天跟踪，发现他们并不是每个病人抽取的样本容量都是足量的，他们的试管是5ml的负压管，有时候抽的就2ml，把这个问题解释给了主任。 | SWO-408137 | User error | preanalytical | preanalytical | Sample handling |  | Customer reflects pH measurement results are not allowed, there are objections | Customer reflects pH measurement results are not allowed, there are objections | Customers use heparin-resistant test tubes to extract arterial blood, and then use syringes to separate parts for samples, which the customer says have been doing before, and they don't want to change. After two days of follow-up, it was found that they were not the sample size taken by each patient is sufficient, their test tube is 5 ml of negative pressure tube, sometimes pumped on 2 ml, the problem was explained to the director.
No | Visit | Complaint | RCHN | 血氧检测失败 | 血氧检测失败，借902-980一个 | 更换902-980后修复（by chuanwei） | SWO-405199 | Defect | HW | Oxi module | Hemolyzer | 581 | Failed blood oxygen test | Blood oxygen test failed, borrow 902-980 one | Repair after replacement 902-980 (by chuanwei)
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 漏电流 | 更换试剂包后修复 | SWO-407565 | [5 empty] | Test card, reagent package failure | Leak current | Repair after replacement of reagent kit
No | Visit | Complaint | RCHN | 报错1999.37.222 | 报错1999.37.322 借903-175 903-034 636-540各一个 | 未更换借用配件，更换进样针模块后的皮带后修复（by chuanwei） | SWO-405138 | Defect | HW | Aspiration Module | PCB, needle | 1999.37 | Error 1999.37.222 | Error 1999.37.322 Borrowing 903-175 903-034 636-540 each | Unreplaceable borrowed accessory, belt repair after replacement of the sample needle module (by chuanwei)
No | Visit | Complaint | RDEU | 28.08.2018 09:00Uhr Brigitte Lenzen 02151324423: Spülfehler | 0509 und 0510 Oberer El $Met und pH $BG Flüssigkeitsfühler hat keine Spüllösung von oben erhalten. TopPlug never reached pH $BG upper Liquid Sensor. | Ventile ausgebaut und gereingt, Ventil Membran im pH $BG Modul verschlissen - ersetzt. Spülung und LS justage durchgeführt. Kalibriert und QC’s durchgeführt. | SWO-405863 | Cleaned | Fluidic pathway | Fluidic pathway | pH/BG Module | 509, 510 | 28.08.2018 09:00H Brigitte Lenzen 02151324423: Flushing error | 0509 and 0510 Upper El /Met and pH /BG liquid sensor has not received a rinsing solution from above. TopPlug never reached pH /BG upper Liquid Sensor. | Valves removed and rimmed, valve membrane worn in pH $BG module - replaced. rinse and LS adjustment. Calibrated and QC's performed.
No | Visit | Complaint | RAME | OXI results are giving errors running VK kit L-3 only | OXI results are giving 581 error running VK kit L-3 only. OXI diagnostic, calibration results are normal and QC is in control. L-3 VK solution has been problematic for some instruments and a substitution of QC 6+ L-3 is recommended. The only effect of the substitution is that the METHb upper reportable range will be reduced to 9-11% from 29-31% | Ran QC 6+ lot 103 and all results are in control. Completed linearity report and sent to Kelly K. for review. Instrument operating normally. | SWO-407442 | No fault found | HW | Oxi module | Oxi module | 581 | OXYGEN results are giving errors running VK Kit L-3 only | OXI results are giving 581 error running VK kit L-3 only. OXI diagnostic, calibration results are normal and QC is in control. L-3 VK solution has been problematic for some instruments and a substitution of QC 6+ L-3 is recommended. The only effect of the substitution is that the METHb upper reportable range will be reduced to 9-11% from 29-31% | Ran QC 6+ lot 103 and all results are in control. Completed linearity report and sent to Kelly K. for review. Instrument operating normally.
No | Visit | Complaint | RAME | OXI results are giving errors running VK kit L-3 only | OXI results are giving 581 error running VK kit L-3 only. OXI diagnostic, calibration results are normal and QC is in control. L-3 VK solution has been problematic for some instruments and a substitution of QC 6+ L-3 is recommended. The only effect of the substitution is that the METHb upper reportable range will be reduced to 9-11% from 29-31% | Ran QC 6+ lot 103 and all results are in control. Completed linearity report and sent to Kelly K. for review. Instrument operating normally. | SWO-407441 | No fault found | HW | Oxi module | Oxi module | 581 | OXYGEN results are giving errors running VK Kit L-3 only | OXI results are giving 581 error running VK kit L-3 only. OXI diagnostic, calibration results are normal and QC is in control. L-3 VK solution has been problematic for some instruments and a substitution of QC 6+ L-3 is recommended. The only effect of the substitution is that the METHb upper reportable range will be reduced to 9-11% from 29-31% | Ran QC 6+ lot 103 and all results are in control. Completed linearity report and sent to Kelly K. for review. Instrument operating normally.
No | Visit | Complaint | RKK | Naドリフト大及びGluアンステ | 電解質項目ドリフトにて対応 | 接続UPSにてバッテリー切り替えが頻繁に行われていた。 切り替えの際に10Vくらいドロップする事も確。 切り替えの要因は大元の電源不良か？？ 電源を壁から直接取るように変更。 念の為、AU BDと電解質PCB交換も実施 | SWO-408821 | Defect | HW | El/Met Module | PCB |  | Na Drift Large &amp; Glu Anste | Compatible with electrolyte item drift | The battery switching was done frequently in the connection UPS. It is also certain to drop about 10V at the time of switching. Is the reason for the switching due to the power failure of the original?? Changed to take power directly from the wall. AU BD and electrolyte PCB replacement
No | Visit | Complaint | RAME | When they go into the patient report setup there are no available items to choose from. She needs to add operator. | Operator was missing on patient reports | Found operator was not missing it was all the way in bottom. But it wasn’t showing them when logging on. Copied set up from old instruments to new instrument and that resolved the issue. No test run due did not disturbed measuring chamber | SWO-406183 | Setup | SW | RMED SW | RMED SW |  | When they go into the patient report setup there are no available items to choose from. She needs to add operator. | Operator was missing on patient reports | Found operator was not missing it was all the way in bottom. But it wasn’t showing them when logging on. Copied set up from old instruments to new instrument and that resolved the issue. No test run due did not disturbed measuring chamber
No | Visit | Complaint | RIND | Na measurement unstable | Follow up, checked ABL800 data logs, Ca status out of limit. | Replaced Ca membrane, performed 2 point calibration successfully. Liquid heparin vial given to customer for ABG sampling. | SWO-406151 | Assisted Replacement | Electrode | Membrane | Ca | Status out of range | Na measurement unstable | Follow up, checked ABL800 data logs, Ca status out of limit. | Replaced Ca membrane, performed 2 point calibration successfully. Liquid heparin vial given to customer for ABG sampling.
No | Visit | Complaint | RAME | Leak current detected (963 & 964) | 8 $29 $18 0964-LEAK CURRENT IN SOLUTION PACK DETECTED. CARL HAS REPLACED NUMEROUS SPS AND A SC. ERROR STILL COMING UP. I WENT TO FEDEX TO PICK UP PARTS. I RAN SC SIMULATOR. CHECKED FOR LEAKS. PUT ON PM KIT. WILL CHECK WITH CARL TOMORROW. 8 $30 $18 0964 RETURNED. CARL REBOOTED AND ERROR WENT AWAY. CAME IN TO GATHER DATA LOGS AND SERVICE DUMP FOR CQC CASE. | 8 $29 $18 RAN SIMULATOR TEST AND ALL PASSED. REPLACED SP AND SC WITH ONES I BROUGHT. PUT ON PM KIT. PRINTED OUT CREDIT FORM. | SWO-406701 | Assisted Replacement | SP | SP | Leak current |  | Leak current detected (963 & 964) | 8 /29 /18 0964-LEAK CURRENT IN SOLUTION PACK DETECTED. CARL HAS REPLACED NUMEROUS SPS AND A SC. ERROR STILL COMING UP. I WENT TO FEDEX TO PICK UP PARTS. I RAN SC SIMULATOR. CHECKED FOR LEAKS. PUT ON PM KIT. WILL CHECK WITH CARL TOMORROW. 8 /30 /18 0964 RETURNED. CARL REBOOTED AND ERROR WENT AWAY. CAME IN TO GATHER DATA LOGS AND SERVICE DUMP FOR CQC CASE. | 8 $29 $18 RAN SIMULATOR TEST AND ALL PASSED. REPLACED SP AND SC WITH ONES I BROUGHT. PUT ON PM KIT. PRINTED OUT CREDIT FORM.
No | Visit | Complaint | RAME | Calibrations PO2 Contact: Judy Wright - 412-232-7823 | PCO2 is reading low sensitivity on calibrations. This was second cassette replaced by customer for same issue. Couldn't find customers QC adapter while onsite, customer will locate. | Replaced sensor cassette. Restart calibrations. Wait for sensor to initialize. Ran tHB calibration with pass results. | SWO-406875 | Assisted Replacement | SC | SC | pCO2 |  | Calibrations PO2 Contact: Judy Wright - 412-232-7823 | PCO2 is reading low sensitivity on calibrations. This was second cassette replaced by customer for same issue. Couldn't find customers QC adapter while onsite, customer will locate. | Replaced sensor cassette. Restart calibrations. Wait for sensor to initialize. Ran tHB calibration with pass results.
No | Visit | Complaint | RAME | Customer called stating that analyzer keeps giving fluid flow error. | Customer keeps getting clot detected when they replace consumables | Flushed analyzer and reinstalled sc and sp. I then had to replace sc due to pCO2 calibration error. Then had to replace it again for LAC error. After all that it finally came in. | SWO-411937 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway | Clot | Customer called stating that analyzer keeps giving fluid flow error. | Customer keeps getting clot detected when they replace consumables | Flushed analyzer and reinstalled sc and sp. I then had to replace sc due to pCO2 calibration error. Then had to replace it again for LAC error. After all that it finally came in.
No | Visit | Complaint | RLTD | Tina (Senior Midwife) called to say that there is an issue with the analyser and has been playing up all weekend. It keeps asking for sensor cassette replacement which she has now done twice and it hasnt made a differance. She informed me that this serial number is on delivery suite and NICU also have the same issue happening therefore im not sure that we have these analysers mapped in the correct locations. Tina gave me R0150N018 as Delivery Suite. Tina 01234 795805 | Customer reported that the analyser was failing calibration and had changed the sensor cassette twice. | Customer training issue, customer had indeed replaced the sensor twice, once because it had time expired, and the second time because they had tried to install an out of date sensor, the third cassette had taken a while to pass calibration. When I arrived on site I just ran the last QC to get the analyser in green. Gave the customer some basic training whilst on site. Whilst there I carried out Live Connect trouble shooting, getting a new IP address from ITC and setting up ipv4 and ie settings, now show green in gateway. | SWO-406788 | User error | SC | SC | Training | Expired SC | Tina (Senior Midwife) called to say that there is an issue with the analyser and has been playing up all weekend. It keeps asking for sensor cassette replacement which she has now done twice and it hasnt made a differance. She informed me that this serial number is on delivery suite and NICU also have the same issue happening therefore im not sure that we have these analysers mapped in the correct locations. Tina gave me R0150N018 as Delivery Suite. Tina 01234 795805 | Customer reported that the analyser was failing calibration and had changed the sensor cassette twice. | Customer training issue, customer had indeed replaced the sensor twice, once because it had time expired, and the second time because they had tried to install an out of date sensor, the third cassette had taken a while to pass calibration. When I arrived on site I just ran the last QC to get the analyser in green. Gave the customer some basic training whilst on site. Whilst there I carried out Live Connect trouble shooting, getting a new IP address from ITC and setting up ipv4 and ie settings, now show green in gateway.
No | Visit | Complaint | RCHN | 报错 1999.5.209 | 报错 1999.5.209,扫描不到条码,借902-978，910-304，902-924各一个。 | 更换902-978，910-304，902-924后修复（by chuanwei） | SWO-405755 | Defect | HW | Cover and Chassis | Barcode reader |  | Error 1999.5.209 | Error 1999.5.209, can not scan the bar code, borrow 902-978, 910-304, 902-924 each one. | Replace 902-978, 910-304, 902-924 post-repair (by chuanwei)
No | Visit | Complaint | RCHN | 报错1024 | 报错1024，借944-214（14465）一张 | 清洁后重新定标修复（by chuanwei） | SWO-405165 | Cleaned | HW | Process Module | Process Wheel | 1024 | 1024 error | 1024, borrow 944-214 (14465) one | Re-calibration repair after cleaning (by chuanwei)
No | Visit | Complaint | RKK | Na QC範囲外 | Na QC範囲外 | Ref電極交換 | SWO-409344 | Assisted Replacement | Electrode | Electrode | Reference |  | Na QC out of range | Na QC out of range | Ref electrode replacement
No | Visit | Complaint | RAME | pCO2 QC for L-1 is out of control low. | pCO2 QC for L-1 is out of control low and all levels are below mean. Multiple membrane changes have not improved result performance. | Replaced pCO2 electrode and membrane. Equilibrated, calibrated and ran QC. Instrument QC results are in control and instrument is operating normally | SWO-407443 | Assisted Replacement | Electrode | Electrode | pCO2 |  | pCO2 QC for L-1 is out of control low. | pCO2 QC for L-1 is out of control low and all levels are below mean. Multiple membrane changes have not improved result performance. | Replaced pCO2 electrode and membrane. Equilibrated, calibrated and ran QC. Instrument QC results are in control and instrument is operating normally
No | Visit | Complaint | RAME | Connect analyzer HDMI output to customer’s Hisense television. | Connect analyzer HDMI output to customer’s Hisense television. | Created Hisense television profile in the Intel HD Graphics Control Panel within the Windows OS. Tested. | SWO-406913 | Setup | SW | Windows | TV profile | TV profile within Windows to use HDMI output | Connect analyzer HDMI output to customer’s Hisense television. | Connect analyzer HDMI output to customer’s Hisense television. | Created Hisense television profile in the Intel HD Graphics Control Panel within the Windows OS. Tested.
Yes | 2 | Y | RPNZ | Blank | Installation on NBU site |  | D4E84B51A43C9E27 | Re classified RMED | planned SWO |  |  |  | Blank | Installation on NBU site
No | Visit | Complaint | RSAF | Aborted calibration of TNT. Customer borrowed from closest site and re calibrated and passed. Cal cartridge to be replaced | No fault found. Cartridge Error 24 | Customer borrowed same lot calibrator from another lab and reran calibration all ok. Cleaned storage wheel and cartridge slider units photo sensors. | SWO-406406 | Defect | Consumables | Test Kit | Cartridge, Broken | 2872 | Aborted calibration of TNT. Customer borrowed from closest site and re calibrated and passed. Cal cartridge to be replaced | No fault found. Cartridge Error 24 | Customer borrowed same lot calibrator from another lab and reran calibration all ok. Cleaned storage wheel and cartridge slider units photo sensors.
No | Visit | Complaint | RESP | 1184 fuga detectada | 1184 fuga detectada. | Se reviso el equipo, se limpio junta de entrada y se comprobó correcto funcionamiento mediante QC. | SWO-409853 | Defect | HW | Inlet module | Inlet gasket | 1184 | 1184 leak detected | 1184 leak detected. | The equipment was checked, the input gasket was cleaned and proper operation was checked by QC.
Yes | Visit | Complaint | RAME | Need to add some ini files for Sunquest interface | No data is transmitted to Sunquest | Added needed lines to ini file Connection established and sent data Assist with QC study | SWO-407524 | Re classified RMED | planned SWO |  |  |  | Need to add some ini files for Sunquest interface | No data is transmitted to Sunquest | Added needed lines to ini file Connection established and sent data Assist with QC study
Yes | Visit | Complaint | RAME | He'd to make ini file changes for Sunquest interface | No data is transferring to Sunquest Settings transferred from other analyzer | Add needed changes to ini file Checked connection in activity log, established Send button present | SWO-407522 | Re classified RMED | planned SWO |  |  |  | He'd to make ini file changes for Sunquest interface | No data is transferring to Sunquest Settings transferred from other analyzer | Add needed changes to ini file Checked connection in activity log, established Send button present
Yes | Visit | Complaint | RAME | HUP checklist | HUP checklist | Completed internal $external lists | SWO-407504 | Re classified RMED | planned SWO |  |  |  | HUP checklist | HUP checklist | Completed internal $external lists
Yes | Visit | Complaint | RAME | HUP checklist | HUP Checklist | Completed internal $external lists | SWO-407502 | Re classified RMED | planned SWO |  |  |  | HUP checklist | HUP Checklist | Completed internal $external lists
Yes | Visit | Complaint | RAME | HUP checklist | HUP checklist | Completed internal $external checklist | SWO-407498 | Re classified RMED | planned SWO |  |  |  | HUP checklist | HUP checklist | Completed internal $external checklist
No | Visit | Complaint | RLTD | Connie on her mobile 07989107397 | It was attempted to move the analyser database and to install a blank database, however on doing so SQL was in use and preventing modifications to the database. Analyser software v3.4 installed and this failed to resolve the issue and it was still not possible to modify the database file. A fresh install of Windows 7 was performed, date and time and region settings were changed. Touch screen calibrations performed. Radiometer software v3.4 installed onto the analyser. The setup from the labour ward was copied and installed onto the analyser. The report layout for capillary sampling was modified to mimic a report provided by the customer. | Windows 7 installed. Radiometer SW v3.4 installed. | SWO-407483 | Defect | SW | RMED SW | not loading |  | Connie on her mobile 07989107397 | It was attempted to move the analyser database and to install a blank database, however on doing so SQL was in use and preventing modifications to the database. Analyser software v3.4 installed and this failed to resolve the issue and it was still not possible to modify the database file. A fresh install of Windows 7 was performed, date and time and region settings were changed. Touch screen calibrations performed. Radiometer software v3.4 installed onto the analyser. The setup from the labour ward was copied and installed onto the analyser. The report layout for capillary sampling was modified to mimic a report provided by the customer. | Windows 7 installed. Radiometer SW v3.4 installed.
No | Visit | Complaint | RAME | Oxi 581 errors | 581 errors on QC and spectrometer failed the oxi test for spectrometer noise | Replaced spectrometer. Had issues with the sp so the sp needed to be replaced before testing could be completed. Will complete this service on 8 $31 $18 as I need to service the 825 here as well. | SWO-407593 | Revisit required | HW | Oxi module | Hemolyzer | 581 | Oxygen 581 errors | 581 errors on QC and spectrometer failed the oxi test for spectrometer noise | Replaced spectrometer. Had issues with the sp so the sp needed to be replaced before testing could be completed. Will complete this service on 8 $31 $18 as I need to service the 825 here as well.
No | Visit | Complaint | RCHN | PO2质控偏高 | PO2质控偏高，借903-334一个 | 更换903-334后修复(by chuanwei) | SWO-405708 | [5 empty] | PO2 QUALITY CONTROL IS HIGH | PO2 HIGH QC, borrow 903-334 one | Repair after replacement 903-334 (by chuanwei)
No | Visit | Complaint | RAME | Connectivity issues. | Analyzer not connecting to network | Found comma instead of period in network settings. Corrected issue. | SWO-407341 | [5 empty] | Connectivity issues. | Analyzer not connecting to network | Found comma instead of period in network settings. Corrected issue.
No | Visit | Complaint | RAME | Analyzer is connected to RALS, but it is not transferring data. | astm.ini file for RALS was not installed | Installed current astm.ini file, and this resolved the problem. | SWO-407413 | [5 empty] | Analyzer is connected to RALS, but it is not transferring data. | astm.ini file for RALS was not installed | Installed current astm.ini file, and this resolved the problem.
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | pco2灵敏度超范围，0378 | 更换测试卡后修复 | SWO-408358 | [5 empty] | Test card, reagent package failure | pco2 sensitivity over range, 0378 | Repair after replacing test card
Yes | 0 | Y | RPAC | Blank | Deliver interface cables and grounding cables. |  | D4DB7E4808485291 | Re classified RMED | Logistical upSWO |  |  |  | Blank | Deliver interface cables and grounding cables.
Yes | 1 | Y | RPAC | Blank | Deliver interface cables and grounding cables. |  | D4DB7E4808485291 | Re classified RMED | same day upSWO |  |  |  | Blank | Deliver interface cables and grounding cables.
No | Visit | Complaint | RCHN | FTC-106 | FTC-106 | 清洁HCT模块后修复 Cleaned the HCT module.(By Ran) | SWO-407169 | Cleaned | HW | Aspiration Module | Needle Wash Unit | FTC-106 | FTC-106 | FTC-106 | Repair Cleaned the HCT module after cleaning the HCT module. (By Ran)
No | Visit | Complaint | RCHN | 存储盘卡盘 | 存储盘卡盘 | 更换测试卡后修复 | SWO-407049 | Defect | Consumables | Test Kit | Cartridge, Broken | 5784 | Memory disk card | Memory disk card | Repair after replacing test card
No | Visit | Complaint | RCHN | AU板的脚弯了，门关不紧，屏幕故障 | AU板的接口针脚弯了，门关不紧，2168错误，借636-574、903-198、886-638、886-639各一个 | 更换636-574、903-138、886-639后修复（by chuanwei） | SWO-405726 | Defect | HW | Analyzer Control Module | PCB, Main Board |  | AU board's feet bent, door closed tight, screen failure | The interface of the AU board is bent, the door is not tight, 2168 errors, borrow 636-574, 903-198, 886-638, 886-639 each | Replacement 636-574, 903-138, 886-639 post-repair (by-chuanwei)
No | Visit | Complaint | RCHN | CL定标失败 | CL定标失败，借945-617一个 |  | SWO-407113 | Assisted Replacement | Electrode | Electrode | Cl |  | CL calibration failed | CL bid failure, borrow 945-617 one
No | Visit | Complaint | RDEU | Touch Screen schlecht bedienbar | Touch Screen schlecht bedienbar Meldung per e-mail ! | Nach Neustart und Reinigung Touchscreen keine Beanstandung ! Allgemeine Überprüfung und Funktionsüberprüfung ohne Beanstandung ! | SWO-407246 | Cleaned | HW | DMS_PC | Touch screen |  | Touch screen poorly operated | Touch screen poorly operated Message by e-mail ! | After reboot and cleaning touchscreen no complaint ! General check and function check without complaint !
No | Visit | Complaint | RSCH | Rep. Elektrolyten | Abteilung hat sich über sprunghaft erhöhte K Werte des gleichen Patienten beschwert Vort Ort ist aufgefallen, dass Ref. Messkammer nach Messung resp. Spülung kleine Reste von Blut bleiben Immer wieder Ca und K Drift Seit Wartung nicht besser geworden (SWO-400660) | Ref. Elektrode ersetzt i.o Alte Lot = 9813 => Garantie Kal. und QC i.o Vergleichsmessungen mit ABL835 MedTech durchgeführt i.o System i.o | SWO-407234 | Assisted Replacement | Electrode | Electrode | Reference |  | Rep. Electrolyte | Department has complained about skyrocketing K values of the same patient complained On the ground, it has been noticed that Ref. Measuring after measurement or rinsing small residues of blood remain Again and again Ca and K Drift Not improved since maintenance (SWO-400660) | Ref. Electrode replaces i.o Old lot = 9813 => Warranty Kal. and QC i.o Comparative measurements carried out with ABL835 MedTech i.o System i.o
No | Visit | Complaint | RCHN | NA定标失败 | NA定标失败 借945-618一个 | 拆库存机上的945-618后修复（By Ran） | SWO-406325 | Assisted Replacement | Electrode | Electrode | Na |  | NA calibration failure | NA calibration failure borrowed 945-618 one | 945-618 post-repair on the stock-removal machine (By Ran)
No | Visit | Complaint | RAME | Ph sealing ring sticking to bottom of measuring chamber | Ph sealing ring sticks to bottom of the measuring chamber when ph electrode is removed | Replaced ph lower sealing ring | SWO-407473 | Defect | Electrode | Electrode | Sealing ring, pH |  | Ph sealing ring sticking to bottom of measuring chamber | Ph sealing ring sticks to bottom of the measuring chamber when ph electrode is removed | Replaced ph lower sealing ring
No | Visit | Complaint | RPOL | Neon lamp too low, neon numer too low | Awaria modułu oxi | Wymiana modułu oxi | SWO-407194 | Defect | HW | Oxi module | Oxi module | neon lamp | Neon lamp too low, neon numer too low | Module Oxi Failure | Replacing the Oxi module
No | Visit | Complaint | RAME | Customer is testing transmission of patient data but is not getting data to cross over. | Roma settings need to be configured to allow for serial connection data transmission | Adjusted Roma configuration file to activate the serial port. Customer’s Lantronix box is attached and a connection is established. Packet failed to send messages are generated. This indicates that the analyzer is now connected and attempting to send data but the host is not available. Customer will contact hospital IT to reboot the interface. A follow up call will be made on the next business day. | SWO-408000 | Setup | SW | RMED SW | Roma.ini |  | Customer is testing transmission of patient data but is not getting data to cross over. | Roma settings need to be configured to allow for serial connection data transmission | Adjusted Roma configuration file to activate the serial port. Customer’s Lantronix box is attached and a connection is established. Packet failed to send messages are generated. This indicates that the analyzer is now connected and attempting to send data but the host is not available. Customer will contact hospital IT to reboot the interface. A follow up call will be made on the next business day.
No | Visit | Complaint | RAME | pCO2 failed linearity for level 3 | Verified pCO2 failure on level VK kit, reordered VK kit and re ran. | All values are within package insert. | SWO-407831 | No fault found | No fault found | VK kit | pCO2 | Installation | pCO2 failed linearity for level 3 | Verified pCO2 failure on level VK kit, reordered VK kit and re ran. | All values are within package insert.
No | Visit | Complaint | RFRA | Le client signale de nombreux problème d'aspiration de l'échantillon . Lors de la présentation de la seringue il n'y a pas d'aspiration qui est effectué Mme Lopez | La carte Auto Inler PCB ne tient pas bien sur son support. | Chgt carte Auto Inler PCB. Contrôle en fonctionnement ok | SWO-411751 | Defect | HW | Auto Inlet module | PCB, Auto inlet | HALL | The customer reports many sample suction problems. When presenting the syringe there is no suction that is carried out Mme Lopez | The Auto Inler PCB card doesn't hold up well on its support. | Chgt Auto Inler PCB card. Control in operation ok
No | Visit | Complaint | RESP | No inicializa | Se efectúan limpieza y sustitución de entrada y sonda de aspiración. | Es necesario desmontaje de válvula y limpieza de la misma así como tubo de válvula para su sustitución, se monta todo el sistema y se efectúan test de servicio, así como calibraciones y Q.C. | SWO-407804 | Cleaned | HW | Oxi module | Pinch valve |  | Doesn't initialize | Cleaning and replacement of inlet and suction probe is performed. | Valve disassembly and valve cleaning are required as well as valve tube for replacement, the entire system is assembled and service tests are performed, as well as calibrations and Q.C.
No | Visit | Complaint | RKOR | Probe bent, position error | . | Inlet probe replace. Gasket hold replace. Position adjust. User training. | SWO-408200 | Defect | HW | Auto Inlet module | Probe bend |  | Probe bent, position error | . | Inlet probe replace. Gasket hold replace. Position adjust. User training.
No | Visit | Complaint | RAME | Oxi 581 errors | 581 errors on internal QC continued, replaced SP twice before being able to run oxi test. Found spectrometer noise test still fails after spectrometer replacement, voltage high out of range. Also po2 test failed the zpo2 energy portion low. Flushed analyzer thoroughly also before replacing SP the second time. Also errors 1324 in homogeneous rinse. | Replaced hemolyzer and attempted to test again, analyzer is still requesting a flush or a SP with a 1343 unable to pump solutions error. Will replace the SP once more and leave it to calibrate. Will contact Hupchew Wong about following up on this analyzer Tuesday 9 $4 $18, this analyzer is not yet active as it is a new installation. I realized all the spare SPs we have are lot number BQ-58, I have had an issue with this lot at more than one location this week where it has taken several SP until the analyzer will work. Still have error 1324 in homogeneous rinse, call not resolved today Hupchew Wong will take up on Tuesday 9 $4 $18 | SWO-408696 | Revisit required | HW | Oxi module | Hemolyzer | 581 | Oxygen 581 errors | 581 errors on internal QC continued, replaced SP twice before being able to run oxi test. Found spectrometer noise test still fails after spectrometer replacement, voltage high out of range. Also po2 test failed the zpo2 energy portion low. Flushed analyzer thoroughly also before replacing SP the second time. Also errors 1324 in homogeneous rinse. | Replaced hemolyzer and attempted to test again, analyzer is still requesting a flush or a SP with a 1343 unable to pump solutions error. Will replace the SP once more and leave it to calibrate. Will contact Hupchew Wong about following up on this analyzer Tuesday 9 $4 $18, this analyzer is not yet active as it is a new installation. I realized all the spare SPs we have are lot number BQ-58, I have had an issue with this lot at more than one location this week where it has taken several SP until the analyzer will work. Still have error 1324 in homogeneous rinse, call not resolved today Hupchew Wong will take up on Tuesday 9 $4 $18
No | Visit | Complaint | RCHN | 报错1117 | 报错1117，借903-334,902-972各一个 | 清理接口板后修复(By Ran) Cleaned the 902-972, then OK. | SWO-406430 | [5 empty] | 1117 error | 1117, lend ingress 903-334,902-972 each | Clean interface board post-fix (By Ran) Cleaned the 902-972, then OK.
No | Visit | Complaint | RIND | Service Visit. | Need to run control in to the analyzer. | Run the control L1, L2 and L3 , consult with the pathologist regarding the control values, pathologist satisfies with the values, there is no issue in to the analyzer, analyzer handed over to staff in well working condition. | SWO-411538 | [5 empty] | Service Visit. | Need to run control in to the analyzer. | Run the control L1, L2 and L3 , consult with the pathologist regarding the control values, pathologist satisfies with the values, there is no issue in to the analyzer, analyzer handed over to staff in well working condition.
No | Visit | Complaint | RCHN | 夹紧阀不灵活 | 夹紧阀不灵活，借903-129一个 | 更换903-129后修复(By Ran) | SWO-407006 | Defect | HW | Oxi module | Pinch valve |  | The clamping valve is not flexible | The clamping valve is inflexible, with a 903-129 | Repair after replacement 903-129 (By Ran)
No | Visit | Complaint | RCHN | 仪器无法启动 | 仪器无法启动，借903-198,636-551各一个 | 更换903-198,636-551后修复(By Ran) | SWO-406975 | Defect | HW | Analyzer Control Module | PCB, Main Board |  | The instrument does not start | The instrument does not start, borrow 903-198, 636-551 each | Replacement 903-198,636-551 post-repair (By Ran)
No | Visit | Complaint | RAME | PCO2 is running high on level 3 QC. Customer range is 20.0-24.6 and results are higher than 25. | Analyzer parameter bar was green upon arrival. Verified errors in data logs. Removed all electrodes, no air bubbles on electrodes, all o-rings present. | Cleaned all measuring chambers, removed and reinstalled pCO2 membrane. | SWO-407955 | Cleaned | Fluidic pathway | Fluidic pathway | pH/BG Module |  | PCO2 is running high on level 3 QC. Customer range is 20.0-24.6 and results are higher than 25. | Analyzer parameter bar was green upon arrival. Verified errors in data logs. Removed all electrodes, no air bubbles on electrodes, all o-rings present. | Cleaned all measuring chambers, removed and reinstalled pCO2 membrane.
No | Visit | Complaint | RAME | Customer reporting frequent intermittent Oximetry error with very low tHb results. | Reviewed data logs and patient results. For patient results noted error 581 Oxi measuring error and error 292 turbidity too high. Observed similar failure with QC, all levels. Inspected Oximetry module, fluidics, and performed applicable service diagnostics. No immediate failure reported by diagnostics. Suspect intermittent pinch valve, possible issue with pinch valve tube. Also noted the solution pump tube was almost off the black nipple on the fluidic module panel possibly causing fluid flow issues. | Replaced the Oxi pinch valve and pinch tube, Pushed the solution pump tube completely onto the black nipple. Performed calibration and QC, no further issues observed. Asked tech to run several patient samples, no errors reported. | SWO-407555 | Defect | HW | Oxi module | Valve, pinch valve |  | Customer reporting frequent intermittent Oximetry error with very low tHb results. | Reviewed data logs and patient results. For patient results noted error 581 Oxi measuring error and error 292 turbidity too high. Observed similar failure with QC, all levels. Inspected Oximetry module, fluidics, and performed applicable service diagnostics. No immediate failure reported by diagnostics. Suspect intermittent pinch valve, possible issue with pinch valve tube. Also noted the solution pump tube was almost off the black nipple on the fluidic module panel possibly causing fluid flow issues. | Replaced the Oxi pinch valve and pinch tube, Pushed the solution pump tube completely onto the black nipple. Performed calibration and QC, no further issues observed. Asked tech to run several patient samples, no errors reported.
No | Visit | Complaint | RAME | Error 0332 pH/BG module not filled. Aborting samples | Analyzer aborting capillary samples. Found seal ring from PO2 electrode missing and seal ring for pH electrode leaking. | Replace seal rings for pH and pCO2 electrodes. Clean measuring chambers in pG $BG module. | SWO-407932 | Defect | Electrode | Electrode | Sealing ring, pH |  | Error 0332 pH/BG module not filled. Aborting samples | Analyzer aborting capillary samples. Found seal ring from PO2 electrode missing and seal ring for pH electrode leaking. | Replace seal rings for pH and pCO2 electrodes. Clean measuring chambers in pG $BG module.
Yes | Visit | Complaint | RAME | Error 0332 pH/BG module not filled. Aborting samples | Maintenance assist with Charles Reisinger. | Assisted Charles Reisinger with maintenance and troubleshooting. | SWO-407933 | Re classified RMED | same day upSWO |  |  |  | Error 0332 pH/BG module not filled. Aborting samples | Maintenance assist with Charles Reisinger. | Assisted Charles Reisinger with maintenance and troubleshooting.
No | Visit | Complaint | RAME | Unstable electrolytes | Confirmed issue when onsite. User has tried multiple cals, electrolytes like NA, Ca and K failed calibration with drift error. | Found electrode contact of Ref electrode dirty. Cleaned with alcohol and then performed two cals. All electrolytes now passing with minimal drift. | SWO-407520 | Cleaned | HW | Electrode | Reference |  | Unstable electrolytes | Confirmed issue when onsite. User has tried multiple cals, electrolytes like NA, Ca and K failed calibration with drift error. | Found electrode contact of Ref electrode dirty. Cleaned with alcohol and then performed two cals. All electrolytes now passing with minimal drift.
No | Visit | Complaint | RIND | analyzer showing error code calibrate_42. | checked the analyzer and found its showing calibrate_42 error. | cleaned the incubator wheel and tried to calibrate it and its failed with error calibrate_42 . checked punch out and punch up and process wheel are fine. but still incubator wheel is not calibrating.so we need incubator control PCB for this case. | SWO-408221 | Revisit required | HW | Incubator Module | PCB, Incubator | calibrate_42 | analyzer showing error code calibrate_42. | checked the analyzer and found its showing calibrate_42 error. | cleaned the incubator wheel and tried to calibrate it and its failed with error calibrate_42 . checked punch out and punch up and process wheel are fine. but still incubator wheel is not calibrating.so we need incubator control PCB for this case.
No | Visit | Complaint | RPOL | Niejednorodny roztwór płuczący (czujniki LS). | Krzywa igła. | Wyprostowanie igły. Testy serwisowe ok. | SWO-407787 | Defect | HW | Auto Inlet module | Probe bend |  | Non-homogenous flushing solution (LS sensors). | Needle curve. | Straighting the needle. Service tests approx.
No | Visit | Complaint | RAME | PCO2 and Cl failing calibration | Cl and pCO2 failing cals and QC | Replaced sensor cassette and solution pack -- Lac is sensitivity low on calibrations-- letting analyzer calibrate for 4 hours | SWO-408056 | Assisted Replacement | SC | SC | pCO2 |  | PCO2 and Cl failing calibration | Cl and pCO2 failing cals and QC | Replaced sensor cassette and solution pack -- Lac is sensitivity low on calibrations-- letting analyzer calibrate for 4 hours
No | Visit | Complaint | RAME | ANALYZER WILL NOT POWER UP. SUSPECTED CPU ERROR. | REPLACE CPU ON UNIT AFTER PREVIOUS WORK. | REPLACED THE CPU. THE ANALYZER WILL NOW STAY ON WITH THE PRINTER PLUGGED IN. HOWEVER, COULD NOT GET ANALYZER TO ACTUALLY PRINT. DISCUSSED WITH HUPCHEW WONG, AREA SPECIALIST FSE. BELIEVE NEW PRINTER IS REQUIRED. | SWO-408741 | Defect | HW | DMS_PC | CPU |  | ANALYZER WILL NOT POWER UP. SUSPECTED CPU ERROR. | REPLACE CPU ON UNIT AFTER PREVIOUS WORK. | REPLACED THE CPU. THE ANALYZER WILL NOW STAY ON WITH THE PRINTER PLUGGED IN. HOWEVER, COULD NOT GET ANALYZER TO ACTUALLY PRINT. DISCUSSED WITH HUPCHEW WONG, AREA SPECIALIST FSE. BELIEVE NEW PRINTER IS REQUIRED.
No | Visit | Complaint | RCHN | 电路开路，温度不稳定 | 电路开路，温度不稳定，需要借件903-334和902-972 |  | SWO-408128 | [5 empty] | Circuit open, temperature unstable | Circuit open, temperature instability, need to borrow 903-334 and 902-972
No | Visit | Complaint | RCHN | 处理盘被子底部漏水 | 处理盘被子底部漏水，借903-236一个 | 更换903-236备件后修复(By Ran) | SWO-407614 | Defect | HW | Process Module | Process Wheel | Spring | Water leaks at the bottom of the processing tray quilt | The bottom of the handle quilt is leaking, borrow 903-236 one | Repair after replacement of 903-236 spare parts (By Ran)
No | Visit | Complaint | RCHN | 报错1999 | 报错1999，借836-166一个 | 更换836-166备件后修复(By Ran) | SWO-407069 | Defect | HW | Incubator Module | Drive wheel |  | Error 1999 | Misquoted 1999, loan 836-166 one | Repair after replacement of 836-166 spare parts (By Ran)
No | Visit | Complaint | RCHN | 报错1999.5.209 | 报错1999.5.209，借902-978一个（从西安仓库发货） | 更换902-978备件后修复(By Ran) | SWO-407709 | Defect | HW | Aspiration Module | Pump, cavro |  | Error 1999.5.209 | False 1999.5.209, loan 902-978 one (shipped from Xi'an warehouse) | Repair after replacement of 902-978 spare parts (By Ran)
Yes | Visit | Complaint | RCHN | 报错1999.5.209 | 报错1999.5.209，借958-201一个 | 更换了958-201后修复(By Ran) | SWO-407711 | Re classified RMED | same day upSWO |  |  |  | Error 1999.5.209 | Error 1999.5.209, loan 958-201 | Post-replacement 958-201 repair (By Ran)
No | Visit | Complaint | RCHN | 机器一直重启，报错5684和5297 | 机器一直重启，报错5684和5297,需要借件903-236 |  | SWO-408112 | Defect | HW | Process Module | Process Wheel | 5297, 5684 | The machine has been restarting, reporting errors 5684 and 5297 | The machine has been restarting, the error 5684 and 5297, need to borrow 903-236
No | Visit | Complaint | RCHN | 报错5689 | 报错5689，借944-258（14820）一个 | 重新定标后修复 Recal, then OK(By Ran) | SWO-407616 | Assisted Replacement | Consumables | Test Kit | cartridge, cal | 5689 | 5689 error | 5689, loan 944-258 (14820) one | Fix Recal, then OK (By Ran) after re-calibration
No | Visit | Complaint | RAME | Glucose frequently failing during 2PT cal due to drift, QC L2 above range. | Reviewed data logs, reported problem appears to have begun at approx 8 $31 @ 11:29 during QC L2 analysis. Calibration also exhibiting sporadic drift errors for Glucose as high as 412 down to -34. Inspected EL $MET module, noted evidence of past overflow of one or more chambers due to dried blood in overflow path and dried blood at base of module. Discovered the Glucose electrode was missing the lower sealing oring. Membrane was stuck to electrode tip, possibly damaged. Most likely accidentally removed during membrane change and not replaced. It's not clear when the oring went missing. The activity log indicates several recent changes: 8 $31, 8 $27 & 8 $21. Probable cause of sporadic fluctuating glucose results. | Installed missing Glucose electrode lower sealing oring. As precautionary measure replaced measuring chamber, rinse and air valve for EL $MET I module, replaced pump tube, grey tubes between modules. Also replaced probe, blue tube and inlet gasket holder as precautionary measure. Ran multiple cal cycles until results stabilized. Ran multiple QC cycles until results recover in range. Will monitor. Glucose electrode may have sustained damage due to missing oring as last membrane was stuck to electrode tip. | SWO-408139 | Defect | HW | Electrode | Glu |  | Glucose frequently failing during 2PT cal due to drift, QC L2 above range. | Reviewed data logs, reported problem appears to have begun at approx 8 /31 @ 11:29 during QC L2 analysis. Calibration also exhibiting sporadic drift errors for Glucose as high as 412 down to -34. Inspected EL /MET module, noted evidence of past overflow of one or more chambers due to dried blood in overflow path and dried blood at base of module. Discovered the Glucose electrode was missing the lower sealing oring. Membrane was stuck to electrode tip, possibly damaged. Most likely accidentally removed during membrane change and not replaced. It's not clear when the oring went missing. The activity log indicates several recent changes: 8 /31, 8 /27 & 8 /21. Probable cause of sporadic fluctuating glucose results. | Installed missing Glucose electrode lower sealing oring. As precautionary measure replaced measuring chamber, rinse and air valve for EL $MET I module, replaced pump tube, grey tubes between modules. Also replaced probe, blue tube and inlet gasket holder as precautionary measure. Ran multiple cal cycles until results stabilized. Ran multiple QC cycles until results recover in range. Will monitor. Glucose electrode may have sustained damage due to missing oring as last membrane was stuck to electrode tip.
No | Visit | Complaint | RDEU | KAL und QC Fehler Lactat | KAL und QC Fehler Lactat | Neumembranierung Lactat - KAlibrationen bis Status Grün - QC durchführen Anwender meldet sich falls Problem weiterhin auftritt ! | SWO-408225 | Assisted Replacement | Electrode | Membrane | Lac | Cal + QC | KAL and QC Error Lactate | KAL and QC Error Lactate | Remembraneation Lactate - KAlibrations to Status Green - QC Perform User logs in if problem persists !
No | Visit | Complaint | RAME | Glucose frequently failing during 2PT cal due to drift, QC L2 above range. | Continuation from SWO-408139. Errors observed 378, cal sensitivity drift, 379, cal unstable, response fault. | Removed and inspected and cleaned all measuring module valves, No issues observed. Replaced reference electrode membrane. Noted error 408, zero current error, performed zero current adjustment. Testing indicates improvement in operation but still occasional drift or QC out of range error. Have requested assistance from Area Specialist to join me on 9 $4 to continue troubleshooting. | SWO-409451 | Revisit required | HW | Basic Unit | Power supply |  | Glucose frequently failing during 2PT cal due to drift, QC L2 above range. | Continuation from SWO-408139. Errors observed 378, cal sensitivity drift, 379, cal unstable, response fault. | Removed and inspected and cleaned all measuring module valves, No issues observed. Replaced reference electrode membrane. Noted error 408, zero current error, performed zero current adjustment. Testing indicates improvement in operation but still occasional drift or QC out of range error. Have requested assistance from Area Specialist to join me on 9 $4 to continue troubleshooting.
No | Visit | Complaint | RKK | リークがあります | リークがあります | 各メンブラン交換後より症状発生との事。Caメンブラン先端チップなしの為、メンブラン再交換。排液ポンプチューブ、溶液ポンプチューブ定期交換。エアーライン洗浄。 | SWO-409410 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway | 663 | There's a leak. | There's a leak. | The thing with the symptom occurrence from after each membrane exchange. For ca-menbran tip without tip, menbran re-exchange. Regular replacement of drainage pump tubes and solution pump tubes. Air line cleaning.
No | Visit | Complaint | RKK | 突然、画面が真っ黒になり、装置が動かなくなった。メインの電源スイッチON/OFFで復帰した。 溶液パックを交換すると、流体系のエラーが多発する。 | 突然、画面が真っ黒になり、装置が動かなくなった。メインの電源スイッチON $OFFで復帰した。 溶液パックを交換すると、流体系のエラーが多発する。 | 機器返却 | SWO-414946 | Adjustment | Power OFF/ON | Power OFF/ON | Power OFF/ON |  | All of a sudden, the screen went black and the device stopped working. The main power switch on/off was restored. When the solution pack is replaced, the error of the flow system occurs frequently. | All of a sudden, the screen went black and the device stopped working. The main power switch on the /OFF returned. When the solution pack is replaced, the error of the flow system occurs frequently. | Return equipment
Yes | Visit | Complaint | RCRO | CCN/PER | Error 1117 Verification check (V $C) | Multi parameter | SWO-410391 | [5 empty] | CCN/PER | Error 1117 Verification check (V /C) | Multi parameter
Yes | Visit | Complaint | RCHN | 干燥前温度超出范围，报错1999和1068 | 干燥前温度超出范围，报错1999和1068，需要借件836-166、903-210、944-214（14460） | 更换836-166 903-210 重新定标后修复 by Andrew | SWO-408127 | Defect | HW | Process Module | PCB, cup dryer | 1068 | Temperature out of range before drying, error 1999 and 1068 | Temperature swelled before drying, errors 1999 and 1068, requiring borrowings of 836-166, 903-210, 944-214 (14460) | Replace 836-166 903-210 After re-calibration repair by Andrew
No | Visit | Complaint | RIND | Service Visit. | Need to replace the touch screen display. | Shut down the analyzer, replace the new touch screen instead of old one, start the analyzer, analyzer handed over to staff in well working condition. | SWO-411394 | Defect | HW | DMS_PC | Touch screen |  | Service Visit. | Need to replace the touch screen display. | Shut down the analyzer, replace the new touch screen instead of old one, start the analyzer, analyzer handed over to staff in well working condition.
Yes | 2 | Y | RPAC | Blank | 2 failed SC's |  | D4E1A4A0563CC2F3 | [5 empty] | Blank | 2 failed SC's
Yes | 7 | Y | RPAC | Blank | 2 failed SC |  | D4E1AE875622C2F9 | [5 empty] | Blank | 2 failed SC
No | Visit | Complaint | RLTD | Analyser routinely fails to print and is only resolved with a reboot. | Analyser routinely fails to print and is only resolved with a reboot. | Backed up customer data and settings, loaded a fresh install of version 3.3 MR1. Reloaded customer setup and data. Upgraded to version 3.4, see software upgrade SWO. Customer to monitor, may need to return for a full Windows re-install. | SWO-408900 | Defect | SW | RMED SW | RMED SW | Printer | Analyser routinely fails to print and is only resolved with a reboot. | Analyser routinely fails to print and is only resolved with a reboot. | Backed up customer data and settings, loaded a fresh install of version 3.3 MR1. Reloaded customer setup and data. Upgraded to version 3.4, see software upgrade SWO. Customer to monitor, may need to return for a full Windows re-install.
No | Visit | Complaint | RSAF | Oxi hardware error | Failing calibrations and QC's on oxi values only, done complete oxi test and fluid transport test and all passed | Done complete oxi test and fluid transport test and all passed, straightened bent inlet probe, both calibrations and QC's passed | SWO-409610 | Cleaned | Fluidic pathway | Fluidic pathway | Hemolyzer | 1340, 1034, 375 | Oxi hardware error | Failing calibrations and QC's on oxi values only, done complete oxi test and fluid transport test and all passed | Done complete oxi test and fluid transport test and all passed, straightened bent inlet probe, both calibrations and QC's passed
No | Visit | Complaint | RDEU | Nach Sensorkassettentausch Temperaturfehler D | Getauschte Sensorkassette war intern undicht, dadurch trat Korrosion an der Anschlussbuchse auf | Fehlersuche, Reinigung der Kontakte, Bestellung Ersatzteil, Funktionskontrolle Fehlerhafte Sensorkassette : 945-727 Lot 259089 | SWO-408943 | [5 empty] | After sensor cassette replacement temperature error D | Replaced sensor cartridge was leaking internally, causing corrosion to occur on the connection socket | Troubleshooting, cleaning of contacts, ordering spare part, function control Faulty sensor cartridge: 945-727 Lot 259089
No | Visit | Complaint | RCHN | 电极灵敏度错误 | 电极灵敏度错误 | 更换945-377一个正常。 by Andrew | SWO-408980 | [5 empty] | Electrode sensitivity error | Electrode sensitivity error | Replace the 945-377 one normal. by Andrew
No | Visit | Complaint | RCHN | 触摸屏故障 | 触摸屏故障，借910-324一个 | 更换910-324后修复 by Andrew | SWO-408426 | Defect | HW | DMS_PC | Touch screen |  | Touch screen failure | Touch screen fault, borrow 910-324 one | Repair by Andrew after replacing 910-324
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | NA,K,CA,CL,PH定标不过 1318,1319,1315,1320 | 更换测试卡后修复 | SWO-408799 | [5 empty] | Test card, reagent package failure | NA,K,CA,CL,PH定标不过 1318,1319,1315,1320 | Repair after replacing test card
No | Visit | Complaint | RCHN | 报错1035 | 报错1035，需要借件903-364一个 | 更换903-364后修复 by Andrew | SWO-408727 | Defect | HW | Process Module | Optical Unit | 1035 | 1035 error | 1035, need to borrow 903-364 one | Repair by Andrew after replacing 903-364
No | Visit | Complaint | RCHN | 报错1999,1031 | 报错1999,1031，借902-978各一个 | 更换903-364，902-978后修复 by Andrew | SWO-408833 | Defect | HW | Process Module | Optical Unit | 1031 | Error 1999,1031 | Error 1999, 1031, loan 902-978 each | Fix by Andrew after replacing 903-364, 902-978
No | Visit | Complaint | RAME | Obstruction in tubing | Obstruction in tubing. Unable to clear. | Retubed all instrument. | SWO-410202 | Defect | HW | Basic Unit | Tubes |  | Obstruction in tubing | Obstruction in tubing. Unable to clear. | Retubed all instrument.
No | Visit | Complaint | RAME | Level 3 on linearity has Oxi spectrum error | Looked in data logs and found level 3 on cal verification not showing any results level 1 and 2 did. | Replaced oximetry module, did complete oximetry test, did THB in service and as user and as sample. | SWO-409251 | Defect | HW | Oxi module | Oxi module |  | Level 3 on linearity has Oxi spectrum error | Looked in data logs and found level 3 on cal verification not showing any results level 1 and 2 did. | Replaced oximetry module, did complete oximetry test, did THB in service and as user and as sample.
Yes | Visit | Complaint | RAME | Error 0332 pH/BG module not filled. Aborting samples | Maintenance assist with Charles Reisinger. | Assisted Charles Reisinger. | SWO-409115 | Re classified RMED | same day upSWO |  |  |  | Error 0332 pH/BG module not filled. Aborting samples | Maintenance assist with Charles Reisinger. | Assisted Charles Reisinger.
No | Visit | Complaint | RAME | Error 0332 pH/BG module not filled. Aborting samples | Found inlet probe bent | Replace inlet probe and inlet tubing. Perform inlet calibration. | SWO-409123 | Defect | HW | Inlet module | Probe bend |  | Error 0332 pH/BG module not filled. Aborting samples | Found inlet probe bent | Replace inlet probe and inlet tubing. Perform inlet calibration.
No | Visit | Complaint | RAME | Customer is getting OXI hardware error on patient sample | Found customer is getting error 269. | Found bad hemolyzer. Replaced it. Adjusted lamp unit. Ran complete OXI test. All passed. Ran multiple calibrations. All passed then ran complete tHB calibrations in service. All passed. All calibrations and 3 Levels of QC passed as per manufacturer specification | SWO-409329 | Defect | HW | Oxi module | Hemolyzer | 269 | Customer is getting OXI hardware error on patient sample | Found customer is getting error 269. | Found bad hemolyzer. Replaced it. Adjusted lamp unit. Ran complete OXI test. All passed. Ran multiple calibrations. All passed then ran complete tHB calibrations in service. All passed. All calibrations and 3 Levels of QC passed as per manufacturer specification
No | Visit | Complaint | RIND | Touch not activating according do touch. | found touch not working properly and sometimes not activating. | Instrument restated, checked the monitor connection and tighten the connection. tried for calibration but it not working.connected the external mouse. With mouse instrument working fine. Touch screen need to be replace with new. | SWO-411413 | Defect | HW | DMS_PC | Touch screen | OOB | Touch not activating according do touch. | found touch not working properly and sometimes not activating. | Instrument restated, checked the monitor connection and tighten the connection. tried for calibration but it not working.connected the external mouse. With mouse instrument working fine. Touch screen need to be replace with new.
No | 4 | Y | RPNZ | Blank | Installation - precision check. |  | D4E206D6562AFAB9 | Adjustment | HW | Oxi module | tHb Cal | Service mode | Blank | Installation - precision check.
No | Visit | Complaint | RIND | Calibration not accepted | Checked calibration data and found that it was ok but when we insert test cartridge instrument show error. | Need to re calibrate instrument with new cal cartridge | SWO-409385 | Assisted Replacement | Consumables | Test Kit | cartridge, cal |  | Calibration not accepted | Checked calibration data and found that it was ok but when we insert test cartridge instrument show error. | Need to re calibrate instrument with new cal cartridge
No | Visit | Complaint | RPOL | Błąd uruchomienia oprogramowania | Błąd oprogramowania podczas uruchomienia analizatora | Usunięcie bazy danych , instalacja oprogramowania 6.18 | SWO-408965 | Defect | SW | RMED SW | Booting | 6.18 | Software startup error | Software error when running analyzer | Database removal, software installation 6.18
No | Visit | Complaint | RESP | Sol QC2 no homogénea | Sol QC2 no homogénea | Se realizó kit anual de tubos. Tras esto, se realizaron tests de servicio sin observarse mas anomalías. Se calibró equipo y se comprobó correcto funcionamiento mediante QC´s. Equipo operativo. | SWO-409884 | Defect | HW | Basic Unit | PM Kit | inhomogeneous QC2 | Non-homogeneous QC2 sun | Non-homogeneous QC2 sun | Annual tube kit was made. After this, service tests were performed without any further anomalies. Equipment was graded and proper operation was tested by QC's. Operating Equipment.
No | Visit | Complaint | RDEU | KAL Fehler pCO2 | KAL Fehler pCO2 (Empfindlichkeitswert 21%) QC Fehler pCO2 und Cl (Werte im Ziel) ! SP und SC bereits erneuert Jahreswartung durchgeführt | Erneuerung SC - Startphase - pCO2 KAL Fehler bleibt bestehen Abklärung und Bestellung erfolgt ! Kunde informiert ! | SWO-409832 | Revisit required | SC | SC | pCO2 | Cal | KAL Error pCO2 | KAL Error pCO2 (sensitivity 21%) QC error pCO2 and Cl (values in target) ! SP and SC already renewed Annual maintenance carried out | Renewal SC - Start phase - pCO2 KAL error remains Clarification and ordering takes place ! Customer informed !
No | Visit | Complaint | RAME | Customer is testing transmission of patient data but is not getting data to cross over. | Follow up with customer indicated that patient data is still not transferring. On-site IS personnel, is able to ping the lantronix boxes successfully. Further investigation revealed that the analyzer had the incorrect Lantronix box attached. | Attached Lantronix box 10.170.234.99 and was able to successfully transmit patient results. Analyzer is within acceptable operating specifications. | SWO-409321 | Setup | His/Lis | His/Lis | His/Lis | Lantronix box | Customer is testing transmission of patient data but is not getting data to cross over. | Follow up with customer indicated that patient data is still not transferring. On-site IS personnel, is able to ping the lantronix boxes successfully. Further investigation revealed that the analyzer had the incorrect Lantronix box attached. | Attached Lantronix box 10.170.234.99 and was able to successfully transmit patient results. Analyzer is within acceptable operating specifications.
No | Visit | Complaint | RAME | ANALYZER WILL NOT POWER UP. SUSPECTED CPU ERROR. | Returned to continue work on malfunctioned printer. Customer had problem loading new solution pack. Grinding noise noticed. | Adjusted printer cable and settings. Reloaded printer driver. Printed numerous cal reports successfully. Entered service program to reset solution pack encoder, then reloaded solution pack successfully. Performed all tests in fluidic transport section, all passed. Per Debbie, analyzer is not reporting electrolytes and EI $Met (both Glu & Lac) yet so these parameters are turned off or disabled. Just pH and oxy panels are remaining. | SWO-409924 | Adjustment | HW | Printer Module | Printer Cable |  | ANALYZER WILL NOT POWER UP. SUSPECTED CPU ERROR. | Returned to continue work on malfunctioned printer. Customer had problem loading new solution pack. Grinding noise noticed. | Adjusted printer cable and settings. Reloaded printer driver. Printed numerous cal reports successfully. Entered service program to reset solution pack encoder, then reloaded solution pack successfully. Performed all tests in fluidic transport section, all passed. Per Debbie, analyzer is not reporting electrolytes and EI $Met (both Glu & Lac) yet so these parameters are turned off or disabled. Just pH and oxy panels are remaining.
Yes | 5 | Y | RCZE | Blank |  |  | D4E2FE41CC36E87B | Re classified RMED | planned SWO |  |  |  | Blank
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | CA离子传感器错误 1070 | 更换测试卡后修复 | SWO-411614 | [5 empty] | Test card, reagent package failure | CA ion sensor error 1070 | Repair after replacing test card
Yes | Visit | Complaint | RCHN | 氖灯电压超范围 | 氖灯电压超范围,借902-980一个。 | 更换902-980一个正常。by Andrew | SWO-408915 | [5 empty] | Xenon voltage over range | Xenon voltage over range, borrow 902-980 one. | Replace the 902-980 one normal. by Andrew
No | Visit | Complaint | RCHN | 光学强度超范围 | 光学强度超范围，借903-364一个 | 更换备用903-364后修复 by Andrew | SWO-409432 | Defect | HW | Process Module | Optical Unit | 1031 | Optical strength over range | Optical strength over range, borrowed 903-364 a | Repair by Andrew after replacing spare 903-364
No | Visit | Complaint | RCHN | 孵育盘PCB板损坏 2168 | 孵育盘PCB板损坏 2168 借 903-138一个。 | 更换库存903-138后修复 by Andrew | SWO-408986 | Defect | HW | Incubator Module | PCB, incubator | 2168 | Incubation disk PCB board damage 2168 | Incubation disk PCB board damage 2168 Borrow 903-138 one. | Repair by Andrew after replacing stock 903-138
No | Visit | Complaint | RCHN | 报错3050 | 报错3050，借944-394（14421）一张 | 重启分析仪解决3050报错，重新定标后修复 by Andrew | SWO-408296 | Defect | SW | RMED SW | Corrupt xml | 3050 | 3050 error | 3050, borrow 944-394 (14421) one | Restart analyzer to resolve 3050 errors, fix by Andrew after recalibration
No | Visit | Complaint | RCHN | 屏幕不亮 | 屏幕不亮，借910-298 886-639各一个 | 更换910-298 886-639各一个正常 by Andrew | SWO-408305 | Defect | HW | Analyzer Control Module | display |  | The screen is not lit | The screen is not bright, borrow 910-298 886-639 each | Replace 910-298 886-639 each a normal by Andrew
No | Visit | Complaint | RSCH | Rep. Ref. | Immer wieder Ca Drift um 15Uhr und 23Uhr pH und Ref. Messkammer mit Blutresten (Siehe Foto) | Messkammer Ref. und pH gereinigt i.o O-Ring pH ersetzt i.o Ref. Elektrode ersetzt i.o Alte Lot = 9813 => Garantie Kal. und QC i.o System i.o | SWO-410490 | Revisit required | Electrode | Electrode | Reference | mail: Which action solved the problem?None. Techline: RSCH 822368646 - Try change valves and tubes and inlet gasket | Rep. Ref. | Again and again Ca Drift at 3pm and 11pm pH and Ref. Measuring chamber with blood residues (See photo) | Measuring chamber Ref. and pH cleaned i.o O-ring pH replaces i.o Ref. Electrode replaces i.o Old lot = 9813 => Warranty Kal. and QC i.o System i.o
No | Visit | Complaint | RSCH | Rep. Ref. | Immer wieder Ca Drift Vermehrt um 23Uhr und 15Uhr | Ref. Elektrode ersetzt i.o Kal. und QC i.o System i.o | SWO-410492 | Assisted Replacement | Electrode | Electrode | Reference | Ca Drift | Rep. Ref. | Again and again Ca Drift Increased at 11pm and 3pm | Ref. Electrode replaces i.o Kal. and QC i.o System i.o
No | Visit | Complaint | RSCH | Rep. Ref. | Immer wieder Ca, Na Drift Vermehrt um 23Uhr und 15Uhr Messkammer Ref. und pH mit Blut (siehe Foto) | Messkammer Ref. und pH gereinigt i.o O-Ring pH ersetzt i.o Ref. Elektrode ersetzt i.o Alte Lot = 6809 => Garantie Kal. und QC i.o System i.o | SWO-410491 | Revisit required | Electrode | Electrode | Reference | mail: Which action solved the problem? None. RSCH 822348168 | Rep. Ref. | Again and again Ca, Na Drift Increased at 11pm and 3pm Measuring chamber Ref. and pH with blood (see photo) | Measuring chamber Ref. and pH cleaned i.o O-ring pH replaces i.o Ref. Electrode replaces i.o Old lot = 6809 => Warranty Kal. and QC i.o System i.o
Yes | Visit | Complaint | RAME | Glucose frequently failing during 2PT cal due to drift, QC L2 above range. | No investigation needed. | Assisted local FSE Carl Levine. | SWO-409304 | Re classified RMED | same day upSWO |  |  |  | Glucose frequently failing during 2PT cal due to drift, QC L2 above range. | No investigation needed. | Assisted local FSE Carl Levine.
No | Visit | Complaint | RAME | Glucose frequently failing during 2PT cal due to drift, QC L2 above range. | Ref: SWO-408139, continued troubleshooting with Area Specialist assisting. | Discovered Calibration solution 1 sealing oring in fluid module for screw-in nipple appears out of place, may be leaking. Remove and reinstalled correctly. Suspect power supply may be fluctuating contributing to erratic performance. Replace power supply. Subsequent testing demonstrated improved performance and stability. Calibration and QC passing. Will monitor. | SWO-409452 | Defect | HW | Basic Unit | Power supply |  | Glucose frequently failing during 2PT cal due to drift, QC L2 above range. | Ref: SWO-408139, continued troubleshooting with Area Specialist assisting. | Discovered Calibration solution 1 sealing oring in fluid module for screw-in nipple appears out of place, may be leaking. Remove and reinstalled correctly. Suspect power supply may be fluctuating contributing to erratic performance. Replace power supply. Subsequent testing demonstrated improved performance and stability. Calibration and QC passing. Will monitor.
No | Visit | Complaint | RAME | Customer called and reported that the device was leaking. They could not easily identify the location of the leak. | I saw that the device was leaking from the crea module from the bottom-most dummy electrode chamber drip tray. They were also getting rinse errors by the time I arrived. Error 751, 602, 652. | I reseated the dummy electrode in the chamber that was causing the leak and was able to stop the leak. The rinse related errors also disappeared when the device ran a 1-point calibration. I also performed a pump calibration and liquid sensor calibration. I ran a 2-point calibration along with QC measurements with no leaking or issues. | SWO-410024 | Adjustment | Electrode | Electrode | Dummy | 751, 602, 652 | Customer called and reported that the device was leaking. They could not easily identify the location of the leak. | I saw that the device was leaking from the crea module from the bottom-most dummy electrode chamber drip tray. They were also getting rinse errors by the time I arrived. Error 751, 602, 652. | I reseated the dummy electrode in the chamber that was causing the leak and was able to stop the leak. The rinse related errors also disappeared when the device ran a 1-point calibration. I also performed a pump calibration and liquid sensor calibration. I ran a 2-point calibration along with QC measurements with no leaking or issues.
No | Visit | Complaint | RDAN | Jan melder, at KBA 14 har fejl 791. | Afprøvet og undersøgt. | Flowvej renset med protein remover, reagens pumpeslange(og de andre) udskiftet, slutafprøvet. | SWO-409588 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway | 791 | Jan reports that KBA 14 has error 791. | Tested and investigated. | Flow path purified with protein remover, reagent pump hose (and the others) replaced, final tested.
No | Visit | Complaint | RAME | On going drift on Electrolytes | Electrolyte drift $ intermittent on cals. | Found both analyzers had the clear waste line partly pinched. They were routed out of position possibly during transport to new lab. Also replaced electrode tubes do to excessive stretching with client stock, cleaned measuring chambers and inspected each, after replacements unit still had drift, changed PH lower sealing ring and analyzer stabilized | SWO-409965 | Revisit required | HW | pH/BG Module | Electrode Connector |  | On going drift on Electrolytes | Electrolyte drift / intermittent on cals. | Found both analyzers had the clear waste line partly pinched. They were routed out of position possibly during transport to new lab. Also replaced electrode tubes do to excessive stretching with client stock, cleaned measuring chambers and inspected each, after replacements unit still had drift, changed PH lower sealing ring and analyzer stabilized
Yes | Visit | Complaint | RDEU | Standardisierte Konfiguration zur Uhrzeit Einstellung (Anbindung Timeserver ) erforderlich. | Timerserver Anbindung erforderlich. | Gerät konfiguriert, neugestartet und QC durchgeführt. | SWO-409959 | Re classified RMED | planned SWO |  |  |  | Standardized configuration at the time setting (connection time server ) required. | Timer server connection required. | device configured, restarted and QC performed.
No | Visit | Complaint | RDEU | KAL Fehler pCO2 | KAL Fehler pCO2 Empfindlichkeitswert 23% | Im Service SC und SP entfernt - Neustart durchgeführt und danach SC und SP wieder installiert - Nach Initialisierung keine Fehlerbehebung ! Erneuerung SC - Funktionsüberprüfung etc. Reklamation durchgeführt ! | SWO-410550 | Assisted Replacement | SC | SC | pCO2 | Cal | KAL Error pCO2 | KAL Error pCO2 Sensitivity 23% | Removed in Service SC and SP - Restart and then reinstall SC and SP - No troubleshooting after initialization ! Renewal SC - Function check etc. Complaint executed !
No | Visit | Complaint | RAME | Data transfer issue via serial port | The flagged parameters under patient results does not come across the LIS system. | Port communication functionality tested by using Docklight software. Test result shows that all parameter are being transferred to LIS expect the parameters which there in no value for them on the patient test results. Analyzer does not send the message error to LIS. | SWO-410736 | Revisit required | SW | External SW | His/Lis |  | Data transfer issue via serial port | The flagged parameters under patient results does not come across the LIS system. | Port communication functionality tested by using Docklight software. Test result shows that all parameter are being transferred to LIS expect the parameters which there in no value for them on the patient test results. Analyzer does not send the message error to LIS.
Yes | 1 | Y | RPAC | Blank | Software upgrade V3.4 |  | D4E420A9420F5DC7 | [5 empty] | Blank | Software upgrade V3.4
No | Visit | Complaint | RFRA | Site d'introduction bloqué ouvert, lors de passage d'échantillon ou tout autre manipulation du site, impossible de le fermer. Dans le menu service, impossible de faire redescendre le site, l’appareil a été éteint et rallumé plusieurs fois, le problème est toujours présent. Mme BOUDEHEN - 02.97.06.80.03 | Il manque le circlips de maintien sur le module Inlet. Un nouveau module sera renvoyé sous garantie. | Remise en place du module Inlet. Test en mode service OK. Remise en route. | SWO-410526 | Defect | HW | Auto Inlet module | Auto Inlet module | OOB? | Blocked introductory site open, when passing sample or any other manipulation of the site, impossible to close it. In the service menu, impossible to bring down the site, the device has been turned off and turned on several times, the problem is still present. Ms BOUDEHEN - 02.97.06.80.03 | It lacks the holding circumlips on the Inlet module. A new module will be returned under warranty. | Refurbishment of the Inlet module. Testing in OK service mode. Restarted.
No | Visit | Complaint | RKOR | Clot detected in sensor cassette, OXI module | Flduid transport check Manual flushing with hypoclorite solution, water | Manual flushing with hypoclorite solution, water | SWO-411509 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway | Clot | Clot detected in sensor cassette, OXI module | Flduid transport check Manual flushing with hypoclorite solution, water | Manual flushing with hypoclorite solution, water
No | Visit | Complaint | RAME | E1324-inhomogeneous rinse solution | Confirmed issue when onsite. Analyzer was in UIR mode for sp replacement. | Installed new sp and allowed analyzer to perform startup and cal which later passed. | SWO-410665 | Assisted Replacement | SP | SP | FT | 1324 | E1324-inhomogeneous rinse solution | Confirmed issue when onsite. Analyzer was in UIR mode for sp replacement. | Installed new sp and allowed analyzer to perform startup and cal which later passed.
No | Visit | Complaint | RCHN | 接口板腐蚀 | 接口板腐蚀，借902-972一个 | 更换902-972后修复 by Andrew | SWO-409444 | [5 empty] | Interface plate corrosion | Interface plate corrosion, borrow 902-972 one | Repair by Andrew after replacement 902-972
No | Visit | Complaint | RKK | 再起動時アナログボード不良 | 再起動時アナログボード不良 | CFC交換 イメージバージョン 2.0.0.0⇒2.0.0.1 | SWO-413281 | [5 empty] | Bad analog board on reboot | Bad analog board on reboot | CFC Replacement Image version 2.0.0.0 ⇒ 2.0.0.1
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 参比电极错误和pco2超出范围，报错1361和0378 | 连续更换2张测试卡后修复 | SWO-410258 | [5 empty] | Test card, reagent package failure | Reference electrode error and pco2 out of range, reporterror 1361 and 0378 | Repair after 2 consecutive test cards
No | Visit | Complaint | RCHN | 光学强度超范围 1031 | 光学强度超范围 1031,借903-364一个。定标卡公司有。 | 更换903-364后修复 by Andrew | SWO-409576 | Defect | HW | Process Module | Optical Unit | 1031 | Optical strength over range 1031 | Optical strength over range 1031, borrowed 903-364 one. The bid card company has. | Repair by Andrew after replacing 903-364
Yes | Visit | Complaint | RCHN | BNP温度传感器错误 | BNP温度传感器错误，借905-908一个 | 更换905-908后修复 by Andrew | SWO-410330 | Re classified RMED | same day upSWO |  |  |  | BNP temperature sensor error | BNP temperature sensor error, borrow 905-908 one | Repair by Andrew after replacement 905-908
No | Visit | Complaint | RCHN | BNP定标错误 | BNP定标错误，借944-258（14480）一个 | 重新定标正常。 by Andrew | SWO-409417 | Defect | SW | RMED SW | Motor, process wheel | 5684 | BNP calibration error | BNP calibration error, borrow 944-258 (14480) one | Recalibration is normal. by Andrew
Yes | Visit | Complaint | RCHN | 进样针定标错误 | 进样针定标错误，借902-781一个 | 更换902-781一个正常。 by Andrew | SWO-409548 | [5 empty] | Sample pin calibration error | Sample pin calibration error, borrow 902-781 one | Replace 902-781 one normal. by Andrew
No | Visit | Complaint | RSCH | Rep. Ref. | Immer wieder Ca Drift Vermehrt um 23Uhr und 15Uhr Ref. Elektrode bereits 1x ersetzt | Ref. Elektrode ersetzt i.o Alte Lot =8803 = > Garantie Elektrode von MedTech Kal. und QC i.o System i.o | SWO-410497 | Assisted Replacement | Electrode | Electrode | Reference | Ca Drift | Rep. Ref. | Again and again Ca Drift Increased at 11pm and 3pm Ref. Electrode already replaced 1x | Ref. Electrode replaces i.o Old lot =8803 = > Warranty Electrode from MedTech Kal. and QC i.o System i.o
No | Visit | Complaint | RDEU | 5.09.2018 17:00Uhr Frau Memec 060213832358: Das Gerät hat sich abgeschaltet und kann nicht mehr eingeschaltet werden | Analysator eingeschaltet und hochgefahren (ohne Problem) | Initialisierung - Überprüfung Netzteil im Service (Oxi Complete Test) - Allgemeine Überprüfung - Leckagefehler, KAL und QC Fehler (Membranierungen notwendig) - Erneuerung Dichtring Lactat (fehlte) und Dichtring pH (festgeklebt) - Funktionsüberprüfung etc. ohne Beanstandung | SWO-410548 | Defect | Electrode | Electrode | Sealing ring, Lac |  | 5.09.2018 17:00pm Ms Memec 060213832358: The device has turned off and can no longer be switched on | Analyzer switched on and powered up (without problem) | Initialization - Inspection power supply in service (Oxi Complete Test) - General inspection - Leakage error, KAL and QC errors (membranes necessary) - Renewal sealing ring lactate (missing) and sealing ring pH (fixed) - Functional check etc. without complaint
No | Visit | Complaint | RFRA | Changement des bouteilles de gaz récurrent | Fuite sur le dual valve | Changement du dual valve, des 2 vannes | SWO-410501 | Defect | HW | Gas module | Dual valve |  | Changing recurrent gas cylinders | Leak on the dual valve | Change of dual valve, 2 valves
No | Visit | Complaint | RAME | Drifting Electrolytes on calibration | Drifting Electrolytes on calibration. Intermittent $ on going problem | Suspect cal solution is moving in chamber or not clearing out of the chamber completely , replaced PH lower sealing ring | SWO-410141 | Defect | Electrode | Electrode | Sealing ring, pH |  | Drifting Electrolytes on calibration | Drifting Electrolytes on calibration. Intermittent / on going problem | Suspect cal solution is moving in chamber or not clearing out of the chamber completely , replaced PH lower sealing ring
Yes | Visit | Complaint | RSCH | NA immer wieder drift | Na immer wieder drift | Na Elektrode an Frau Baumann abgegeben. 1 Woche zum testen Sie gibt bescheid. | SWO-410426 | Re classified RMED | Logistical upSWO |  |  |  | NA drift again and again | Na again and again drift | Na electrode to Ms. Baumann. 1 week to test She gives notice.
No | Visit | Complaint | RAME | MetHb is off by 0.6-0.8 from other 2 analyzers on blood check | Customer stated this analyzer is showing a low MetHb compared to other 2 analyzers. Jerry Kelly assisted. | Adjusted lamp, did complete oximetry, did THB in service and as user and as sample, customer did a few samples and Met is still off by 0.6. Replaced hemolyzer, adjusted lamp, did complete oximetry, did a THB in service and as user and as sample. | SWO-410068 | No fault found | HW | Oxi module | Hemolyzer |  | MetHb is off by 0.6-0.8 from other 2 analyzers on blood check | Customer stated this analyzer is showing a low MetHb compared to other 2 analyzers. Jerry Kelly assisted. | Adjusted lamp, did complete oximetry, did THB in service and as user and as sample, customer did a few samples and Met is still off by 0.6. Replaced hemolyzer, adjusted lamp, did complete oximetry, did a THB in service and as user and as sample.
No | Visit | Complaint | RAME | Customer was reporting results with ?'s | The customer call and wanted the analyzer checked out after the user was reporting results that were questioned. I talked with the main user and they were not notified that the analyzer was questioning the results and went ahead and used the results anyway. | I checked out the analyzer and found where blood leaked out the Ph $BG mod. This comes from a clot under the Ph or the Ph seal was not tight. The main user said that when they came to trouble shoot, the Ph O-ring stuck to the bottom of the measuring chamber. They put the O-ring back on and cleaned the chamber. I ran blood samples with no errors. The analyzer is running correctly. I reminded the customer again not to report any results with a ? On them. | SWO-411235 | Defect | Electrode | Electrode | Sealing ring, pH |  | Customer was reporting results with ?' s | The customer call and wanted the analyzer checked out after the user was reporting results that were questioned. I talked with the main user and they were not notified that the analyzer was questioning the results and went ahead and used the results anyway. | I checked out the analyzer and found where blood leaked out the Ph $BG mod. This comes from a clot under the Ph or the Ph seal was not tight. The main user said that when they came to trouble shoot, the Ph O-ring stuck to the bottom of the measuring chamber. They put the O-ring back on and cleaned the chamber. I ran blood samples with no errors. The analyzer is running correctly. I reminded the customer again not to report any results with a ? On them.
No | Visit | Complaint | RAME | On going drift on Electrolytes | Drifting Electrolytes on calibration | Problem was resolved for a few hours then returned at midnight. Same issue and same interval, returned on site and replaced BG Chamber and EL $Met Chamber , also replaced upper fluid valves on BG mod and checked upper valves on ELMET module, replaced REF membrane with new lot #R2338, and REF connector cap. | SWO-410754 | Defect | HW | pH/BG Module | Electrode Connector |  | On going drift on Electrolytes | Drifting Electrolytes on calibration | Problem was resolved for a few hours then returned at midnight. Same issue and same interval, returned on site and replaced BG Chamber and EL $Met Chamber , also replaced upper fluid valves on BG mod and checked upper valves on ELMET module, replaced REF membrane with new lot #R2338, and REF connector cap.
No | 4 | Y | RPNZ | Blank | Installation |  | D4E4728B42553E29 | Setup | SW | RMED SW | RMED SW |  | Blank | Installation
No | Visit | Complaint | RFRA | erreur 3414 : FTC/105 | Intervention suite à investigations effectuées par téléphone le 05 $09 | Ouverture du dispositif et test pompes P4 $ P5 : surpression au niveau du tuyau 8 pompe P5 ( tuyau sorti de son support ) $ $ recherche de la cause du bouchage : vérification pompe ok $ $ démontage et flush du tuyau 8 ok $ $ flush coté needle Wash unit impossible $ $ démontage de la needle Wash unit et débouchage ok $ $ remontage de l’ensemble ok $ $ redémarrage du dispositif ok $ $ contrôle système ok $ $ passage d'un tube de nettoyage $ $ sauvegarde toutes données Ok $ $ sauvegarde de la configuration ok | SWO-410417 | Cleaned | HW | Aspiration Module | Needle Wash Unit | 3414 | erreur 3414 : FTC/105 | Intervention following telephone investigations on /05.09 | Device opening and testing P4 $P5 pumps: overpressure at pipe 8 P5 pump (pipe out of its holder) $ $ search for the cause of the plug: pump check ok $$ dismantling and flush of the pipe $8 ok $ flush rated needle Wash unit impossible $$ needle Wash unit and ok $$reuniting of the whole ok $$reboot of the device ok $$control system ok $$ passage of a cleaning tube $$ backup all data Ok $$ backup of the configuration ok
No | Visit | Complaint | RPOL | Moduł Auto-check | Testy serwisowe. | Wymiana modułu Auto-check. Testy serwisowe. | SWO-410390 | Defect | HW | Autocheck module | Autocheck Module |  | Auto-Check Module | Service tests. | Auto-Check module replacement. Service tests.
No | Visit | Complaint | RKOR | Calibration is out of range [pCO2, Cl] | Perform the calibration, QC A,B,C | Perform the calibration, QC A,B,C | SWO-411504 | Assisted Replacement | SC | SC | pCO2 | Cal | Calibration is out of range [pCO2, Cl] | Perform the calibration, QC A,B,C | Perform the calibration, QC A,B,C
No | Visit | Complaint | RAME | Lactate and glucose is failing calibration | Looked in data logs, getting glucose and lactate response errors, customer already replaced 3 sensor cassettes and a solution pack. Called Keith Clark for assistance and he recommends replacing sensor interface PCB. | Replaced replaced sensor interface board and sensor cassette and solution pack. | SWO-410781 | Defect | HW | Basic Unit | PCB, Sensor interface | Glu+Lac | Lactate and glucose is failing calibration | Looked in data logs, getting glucose and lactate response errors, customer already replaced 3 sensor cassettes and a solution pack. Called Keith Clark for assistance and he recommends replacing sensor interface PCB. | Replaced replaced sensor interface board and sensor cassette and solution pack.
No | Visit | Complaint | RFRA | Mr MORCILLO - Tél : 06.88.17.79.05 - Labo tél : 04.67.33.72.79 Appelé le client quand le FSE arrive sur le site. L'appareil reste bloqué sur l'initialisation du software. L'appareil a été éteint et rallumé, toujours le problème. | ABL90 Flex Plus ne démarre pas correctement. Bloquer sur lancement Application ROMA sur la 1ière ligne status... Le fichier Roma.db corrompue et fichier sur sauvegarde automatique sur réseau est corrompue aussi. | Chargement Soft 3.3MR1 OK. erreur Out of Memory sur ROMA . Pbm sur Roma.db (même sur fichier de sauvegarde réseau). Relance avec Roma.db vierge OK (Manip fmwg OK) sur carte Flash W7 du client. Lancement application ROMA OK. Changement des consommables OK. Calibration OK, Passage de tous les CQ OK. Vérification connection AQURE OK. Tech de Labo Mr MORCILLO passera la Cal Thb dès que possible OK. NB : Client ne souhaite pas la maintenance la veille d'un week end. Rapport envoyé à s-morcillo@chu-montpellier.fr | SWO-411291 | Defect | SW | RMED SW | Roma.ini |  | Mr MORCILLO - Tel: 06.88.17.79.05 - Lab Tel: 04.67.33.72.79 Called the customer when the ESF arrives on the site. The device remains stuck on the software's initialization. The device was turned off and turned on, still the problem. | ABL90 Flex Plus doesn't start properly. Block on launch Application ROMA on the 1st line status... The corrupted Roma.db file and automatic network backup file is corrupted too. | Soft Loading 3.3MR1 OK. Out of Memory error on ROMA . Pbm on Roma.db (even on network backup file). Relaunch with Roma.db virgin OK (Manip fmwg OK) on customer's Flash W7 card. Launch application ROMA OK. Change of consumables OK. Calibration OK, Passage of all CQ OK. AQURE connection check OK. Lab tech Mr MORCILLO will pass the Cal Thb as soon as possible OK. NB: Customer does not want maintenance on the eve of a weekend. Report sent to s-morcillo@chu-montpellier.fr
No | Visit | Complaint | RDEU | 07.09.2018, Frau Mai, Tel. 02611371990: Einlass schließt nicht mehr | Dichtung ersetzt, Einlassonde gerichtet. | Startphase QC und Kalibration, pCO2 empf. bei30% mehrere Kalibrationen>>SC ersetzt. Der Anwender wird sich melden wen es nach der Startphase noch Problemme geben sollte. | SWO-411173 | Defect | HW | Auto Inlet module | Inlet gasket |  | 07.09.2018, Mrs. May, Tel. 02611371990: Entrance no longer closes | Seal replaced, inlet probe directed. | Start phase QC and calibration, pCO2 empf. several calibrations>>SC. The user will contact you if there should be any problems after the start-up phase.
No | Visit | Complaint | RAME | Barcode Reader not reading barcodes Contact is Marketta Latsen at 843-792-5715 | Found that barcode lights up scans and beeps but does not fill in. | I tried rebooting the analyzer and checked the connections. I Have barcode reader on order. Replaced barcode reader. | SWO-411879 | Defect | HW | Basic Unit | Barcode reader |  | Barcode Reader not reading barcodes Contact is Marketta Latsen at 843-792-5715 | Found that barcode lights up scans and beeps but does not fill in. | I tried rebooting the analyzer and checked the connections. I Have barcode reader on order. Replaced barcode reader.
No | Visit | Complaint | RAME | Internal Battery Issues . battery indicator not visible well on battery power. | The analyzer Battery $Electrical power indicator icon is showing the analyzer is running only on battery power even though the instrument is connected to the power outlet. The customer has confirmed the electrical outlet is working fine. | I inspected the power cord and, confirmed it was properly plugged in the back of the analyzer. I checked the battery voltages and, they are charging within the specified ranges. Last, I performed a cold reboot and, now the icon is showing, properly indicating the analyzer is plugged in the wall outlet and, it is also showing the battery charge percentage. To confirm the analyzer is working properly, I disconnected the analyzer off the wall jack and, the instrument immediately switched to battery power. I reconnected the analyzer to the electrical outlet again and, the screen displayed the power outlet icon and, the battery percentage again. The battery was low at 44% but it gradually charged up to 100%. | SWO-411570 | Adjustment | HW | Battery module | Battery module | Charging percentage | Internal Battery Issues . battery indicator not visible well on battery power. | The analyzer Battery /Electrical power indicator icon is showing the analyzer is running only on battery power even though the instrument is connected to the power outlet. The customer has confirmed the electrical outlet is working fine. | I inspected the power cord and, confirmed it was properly plugged in the back of the analyzer. I checked the battery voltages and, they are charging within the specified ranges. Last, I performed a cold reboot and, now the icon is showing, properly indicating the analyzer is plugged in the wall outlet and, it is also showing the battery charge percentage. To confirm the analyzer is working properly, I disconnected the analyzer off the wall jack and, the instrument immediately switched to battery power. I reconnected the analyzer to the electrical outlet again and, the screen displayed the power outlet icon and, the battery percentage again. The battery was low at 44% but it gradually charged up to 100%.
No | Visit | Complaint | RAME | Oxi 581 errors | Confirmed issue when onsite. Internal QC level 1 would fail consistently and in service program, spectrometer noise would fail continuously. | Replaced hemolyzer, power supply and wet section pcb. Spectrometer noise now passed. | SWO-411290 | Defect | HW | Oxi module | Hemolyzer | 581 | Oxygen 581 errors | Confirmed issue when onsite. Internal QC level 1 would fail consistently and in service program, spectrometer noise would fail continuously. | Replaced hemolyzer, power supply and wet section pcb. Spectrometer noise now passed.
No | Visit | Complaint | RKK | アナログボード不良 | アナログボード不良 | CFカード交換 設定引き継ぎ CAL確認 | SWO-412280 | [5 empty] | Analog board bad | Analog board bad | CF card exchange Setting takeover CAL confirmation
No | Visit | Complaint | RKK | アナログボード不良 | アナログボード不良 | CFカード交換 設定引き継ぎ CAL確認 | SWO-412278 | [5 empty] | Analog board bad | Analog board bad | CF card exchange Setting takeover CAL confirmation
No | Visit | Complaint | RESP | Sensor92. No funciona | Comprobación del sensor con pacientes, comprobación del equipo con sensor del tosca. | Una vez verificado el defecto de funcionamiento del sensor, solicitamos uno nuevo en garantía. Se calibra y verificamos el correcto funcionamiento. | SWO-411651 | [5 empty] | Sensor92. No Works | Sensor check with patients, equipment check with cough sensor. | Once the sensor has been verified, we request a new one under warranty. It is calibrated and verify the correct operation.
Yes | Visit | Complaint | RCHN | 光学单元超出范围，定标失败，报错1031、1024、1025 | 光学单元超出范围，定标失败，报错1031、1024、1025，需要借件903-364、944-212（14757） | 更换光单 | SWO-410284 | Defect | HW | Process Module | Optical Unit | 1024, 1031 | Optical unit out of range, calibration failure, error 1031, 1024, 1025 | Optical unit out of range, calibration failure, error 1031, 1024, 1025, need to borrow 903-364, 944-212 (14757) | Replace the light sheet
No | Visit | Complaint | RCHN | 冲洗温度丢失 | 冲洗温度丢失，借905-908一个 | 更换905-908一个正常。 by Andrew | SWO-409362 | Defect | HW | Process Module | Sensor, BnP | 1235 | Flushing temperature is lost | Flushing temperature lost, borrow 905-908 one | Replace 905-908 one normal. by Andrew
No | Visit | Complaint | RCHN | Punch up顶不上来 5297 | Punch up顶不上来 5297 ,借903-031一个。 | 更换903-031后修复 by Andrew | SWO-410335 | Defect | HW | Process Module | PCB, cup transport rear | 5297 | Punch up can't top up 5297 | Punch up can't top up 5297 , borrow edith of 903-031. | Fix by Andrew after replacing 903-031
No | 20 | Y | RCZE | Blank |  |  | D4E906B21A1357EB | Defect | HW | Oxi module | Pinch valve |  | Blank
Yes | Visit | Complaint | RCHN | 测量报错，光度不稳定，报错0581 | 测量报错，光度不稳定，报错0581,借902-833一个 | 更换902-833后修复 by Andrew | SWO-410329 | [5 empty] | Measurement error, luminus, error 0581 | Measurement error, luminometer instability, error 0581, borrow902-833 one | Repair by Andrew after replacement 902-833
No | Visit | Complaint | RDEU | QC Messungen sind bei allen 3 Leveln Sprunghaft. Obere Grenze, nächste QC Mittig. | Analysatorüberprüfung. Ventilmembran am pH $ BG Modul klebt im Ventilblock. Na Bezugsmaterial R0812 vor Ort. Kal 1 im ABL YN04 | Membrane gelöst und Ventilblock gereinigt. Na Elektrode mit anderem Bezugsmaterial (R0786) bespannt. Kal 1 Lösung (YW02) erneuert. System- und Funktionskontrolle. | SWO-411053 | Cleaned | HW | pH/BG Module | Valve membrane |  | QC measurements are jumping at all 3 levels. Upper limit, next QC Middle. | Analyzer check. Valve membrane on the pH BG module sticks in the valve block. Na reference material R0812 on site. Kal 1 in ABL YN04 | membrane dissolved and valve block cleaned. Na electrode covered with other reference material (R0786). Kal 1 solution (YW02). System and function control.
Yes | 3 | Y | RPAC | Blank | Tube between inlet and BG module was a bit moist |  | D4E55B22B825DB3D | [5 empty] | Blank | Tube between inlet and BG module was a bit moist
No | Visit | Complaint | RIND | Nt-ProBNP Calibration invalid. | NT-proBNP Calibration got Invalid. | found problem with 8.9 software, followed service tip and calibrated instrument for NT-proBNP and D-dimer parameter found results are with in the range found instrument is functional. | SWO-411425 | Assisted Replacement | Consumables | Test Kit | Cartridge, cal |  | Nt-ProBNP Calibration invalid. | NT-proBNP Calibration got Invalid. | found problem with 8.9 software, followed service tip and calibrated instrument for NT-proBNP and D-dimer parameter found results are with in the range found instrument is functional.
No | Visit | Complaint | RAME | Analyzer not networking with LIS after software upgrade to version 3.4. The analyzer ID is missing. | Found incorrect ASTM.ini file installed. Correct file deleted during software update. | Installed correct RALs ASTM.ini file. Performed QC measurement and ran patient sample. Information now crosses to the LIS. | SWO-411354 | User Error, FSE | SW | RMED SW | RMED SW | User Error, FSE (SW upgrade) | Analyzer not networking with LIS after software upgrade to version 3.4. The analyzer ID is missing. | Found incorrect ASTM.ini file installed. Correct file deleted during software update. | Installed correct RALs ASTM.ini file. Performed QC measurement and ran patient sample. Information now crosses to the LIS.
No | Visit | Complaint | RAME | Not Networking with LIS | Found incorrect ASTM.ini file installed. Correct file deleted during software update. | Installed correct RALs ASTM.ini file. Performed QC measurement but unable to test LIS connection. | SWO-411490 | User Error, FSE | SW | RMED SW | RMED SW | User Error, FSE (SW upgrade) | Not Networking with LIS | Found incorrect ASTM.ini file installed. Correct file deleted during software update. | Installed correct RALs ASTM.ini file. Performed QC measurement but unable to test LIS connection.
No | Visit | Complaint | RAME | Not Networking with LIS | Found incorrect ASTM.ini file installed. Correct file deleted during software update. | Installed correct RALs ASTM.ini file. Performed QC measurement but unable to test LIS connection. | SWO-411491 | User Error, FSE | SW | RMED SW | RMED SW | User Error, FSE (SW upgrade) | Not Networking with LIS | Found incorrect ASTM.ini file installed. Correct file deleted during software update. | Installed correct RALs ASTM.ini file. Performed QC measurement but unable to test LIS connection.
No | Visit | Complaint | RAME | Not Networking with LIS | Found incorrect ASTM.ini file installed. Correct file deleted during software update. | Installed correct RALs ASTM.ini file. Performed QC measurement but unable to test LIS connection. | SWO-411492 | User Error, FSE | SW | RMED SW | RMED SW | User Error, FSE (SW upgrade) | Not Networking with LIS | Found incorrect ASTM.ini file installed. Correct file deleted during software update. | Installed correct RALs ASTM.ini file. Performed QC measurement but unable to test LIS connection.
Yes | Visit | Complaint | RCHN | 氖灯电压超范围 | 氖灯电压超范围，借903-165（被903-644替代） 903-197 各一个 | 更换903-197,903-644后修复by Andrew | SWO-410263 | [5 empty] | Xenon voltage over range | Xenon voltage over range, borrowed 903-165 (replaced by 903-644) 903-197 each | Repair by Andrew after replacing 903-197, 903-644
No | Visit | Complaint | RCHN | 报错1031 | 报错1031 借903-364一个 | 更换903-364后修复 by Andrew | SWO-410231 | Defect | HW | Process Module | Optical Unit | 1031 | 1031 | False 1031 Borrow 903-364 one | Repair by Andrew after replacing 903-364
No | Visit | Complaint | RCHN | 报错5684 | 报错5684，借944-258（14820），944-212（14280）各一个 |  | SWO-411434 | Defect | SW | RMED SW | Motor, process wheel | 5684 | 5684 error | 5684, loan 944-258 (14820), 944-212 (14280) each
No | Visit | Complaint | RCHN | 报错1031 | 报错1031，借903-364,905-908,903-210,944-258（13956），944-394（144411），944-212（14402）各一个 |  | SWO-411433 | Defect | HW | Process Module | Optical Unit | 1031 | 1031 | 1031, lend 903-364, 905-908, 903-210, 944-258 (13956), 944-394 (144411), 944-212 (14402) each
No | Visit | Complaint | RCHN | 定标失败 1024 | 定标失败 1024,借944-220（13764）一张 | 重新定标后修复 by Andrew | SWO-410353 | Defect | HW | Process Module | Optical Unit | 1024 | Bid Failure 1024 | Bid failure 1024, borrow 944-220 (13764) one | Fix by Andrew after re-calibration
No | Visit | Complaint | RIND | Blank display. | checked and found that the display is blank nothing is coming on the screen even the analyzer not booting | checked by changing the CPU unit after changing the CPU unit the analyzer working. So we need to change the CPU unit. Need to replace the CPU unit. | SWO-414857 | Defect | HW | DMS_PC | CPU | Revisit? | Blank display. | checked and found that the display is blank nothing is coming on the screen even the analyzer not booting | checked by changing the CPU unit after changing the CPU unit the analyzer working. So we need to change the CPU unit. Need to replace the CPU unit.
No | Visit | Complaint | RIND | Oximetry hardware problem shows | Check patient results in which So2 & tHb is missing.Check system message in which Neon number too low error is showing. | Run complete OXI test to pin out the error and found that Fneon is very low.Check neon lamp which have milky colour and found problem in neon lamp which need to replace to resolve the problem. | SWO-411514 | Revisit required | HW | Oxi module | Lamp unit | neon lamp | Oximetry hardware problem shows | Check patient results in which So2 & tHb is missing. Check system message in which Neon number too low error is showing. | Run complete OXI test to pin out the error and found that Fneon is very low. Check neon lamp which have milky colour and found problem in neon lamp which need to replace to resolve the problem.
No | Visit | Complaint | RCHN | 报错1024 | 报错1024，借903-364,944-212（14402） | 更换903-364重新定标后修复 by Andrew | SWO-411056 | Defect | HW | Process Module | Optical Unit | 1024 | 1024 error | 1024, loan 903-364,944-212 (14402) | Repair by Andrew after replacement 903-364 recalibration
No | Visit | Complaint | RCHN | 光单故障 | 光单故障，借903-364,944-216（14696），944-214（14465），944-258（14485），944-212（14642）各一个 |  | SWO-411544 | Defect | HW | Process Module | Optical Unit | 1031 | Light single fault | Light single fault, borrow 903-364, 944-216 (14696), 944-214 (14465), 944-258 (14485), 944-212 (14642) each
No | Visit | Complaint | RAME | The inlet will not close after sample | The inlet gasket would go the whole way down | The inlet probe was bent. I reformed the probe and will have the customer order one | SWO-412277 | Adjustment | HW | Inlet module | Probe bend |  | The inlet will not close after sample | The inlet gasket would go the whole way down | The inlet probe was bent. I reformed the probe and will have the customer order one
No | Visit | Complaint | RAME | When the qc barcodes were being programmed during the instrument installation the bar codes were not scanned at all. | Part follow up. During the installation of this instrument, I found the bar code reader not scanning the qc barcodes. | Replaced the barcode scanner. Tested the new barcode by scanning qc barcodes and, they were successfully scanned. | SWO-412261 | Defect | HW | Basic Unit | Barcode reader |  | When the qc barcodes were being programmed during the instrument installation the bar codes were not scanned at all. | Part follow up. During the installation of this instrument, I found the bar code reader not scanning the qc barcodes. | Replaced the barcode scanner. Tested the new barcode by scanning qc barcodes and, they were successfully scanned.
No | 3 | Y | RCZE | Blank | pristroj neprijma solution pack |  | D4E9ABA81A520FBD | Assisted Replacement | SP | SP | Leak current |  | Blank | pristroj neprijma solution pack
No | 2 | Y | RCZE | Blank |  |  | D4E94B1C1A111E16 | Defect | HW | Auto Inlet module | Inlet gasket |  | Blank
No | Visit | Complaint | RFRA | ERR STATUS 2 |  | Constate err status2 lors du démarrage de la base $Restauration de la base vide impossible $Réinstallation de l'appalication Radiometer impossible $Formatage de la CF CARD impossible $Installation d'une nouvelle CFCARD impossible d'installer W7 $Programmation d'un intervention demain avec nouvelle CFCARD et PC unit. | SWO-412339 | Revisit required | HW | DMS_PC | CF card | status code 2 | ERR STATUS 2 |  | Finds err status2 when starting the base $Restauration the empty base impossible $Rinstallation of the Radiometer installation impossible $Formatage of the CF CARD impossible $Installation of a new CFCARD impossible to install W7$ Programming an intervention tomorrow with new CFCARD and PC unit.
No | 5 | Y | RPNZ | Blank | Solution pack replacement due to 964: leak current in relation to solution pack detected. |  | D4E9BE5C1A3EAE3D | Assisted Replacement | SP | SP | Leak current | 964 | Blank | Solution pack replacement due to 964: leak current in relation to solution pack detected.
Yes | 6 | Y | RPNZ | Blank | Installation on TeKuiti site |  | D4F61DF6B61484A4 | Re classified RMED | installation report |  |  |  | Blank | Installation on TeKuiti site
No | Visit | Complaint | RAME | QC failed on oxi parameters | All oxy failing QC on all levels- Lamp failing complete oxy test. | Replaced the hemolyzer, lamp, spectrometer and oxy pcb. | SWO-412133 | Defect | HW | Oxi module | Lamp unit |  | QC failed on Oxi parameters | All oxy failing QC on all levels- Lamp failing complete oxy test. | Replaced the hemolyzer, lamp, spectrometer and oxy pcb.
No | Visit | Complaint | RAME | pCO2 out of range high on QC. | Checked data logs, found errors on QC level 3. Checked measuring chambers, found pH chamber to be bloody and o-ring sticks to the bottom of the chamber when you pull the pH electrode out. | Replaced pH o-ring. Cleaned chambers and ran calibrations. Tried QC level 3, still out too high. Changed Gas 2 bottle, tried QC level 3. Still out too high. Removed PCO2 electrode and found small air bubble. Tapped it out, ran calibrations. Sensitivity came up to 98% on 2 point calibrations. Reran QC level 3, now within range. | SWO-412255 | Adjustment | Electrode | Electrode | pCO2 | Air bubble | pCO2 out of range high on QC. | Checked data logs, found errors on QC level 3. Checked measuring chambers, found pH chamber to be bloody and o-ring sticks to the bottom of the chamber when you pull the pH electrode out. | Replaced pH o-ring. Cleaned chambers and ran calibrations. Tried QC level 3, still out too high. Changed Gas 2 bottle, tried QC level 3. Still out too high. Removed PCO2 electrode and found small air bubble. Tapped it out, ran calibrations. Sensitivity came up to 98% on 2 point calibrations. Reran QC level 3, now within range.
No | Visit | Complaint | RIND | Parameter Issue | Parameter was yellow | Check ythe calibration report senstiviy ok, found barcode not read of cleaninig solution, read the barcode to analyzer, 2point calibration given , parameter active. | SWO-414951 | User Error | Consumables | Solution | Cleaning solution | Barcode not scanned | Parameter Issue | Parameter was yellow | Check ythe calibration report senstiviy ok, found barcode not read of cleaninig solution, read the barcode to analyzer, 2point calibration given , parameter active.
Yes | Visit | Complaint | RAME | Customer was reporting results with ?'s | To collect data for a CQC | I downloaded the data and will put it in a case | SWO-412276 | Re classified RMED | planned SWO |  |  |  | Customer was reporting results with ?' s | To collect data for a CQC | I downloaded the data and will put it in a case
No | Visit | Complaint | RIND | Hazy print out quality. | Hazy print out coming,affecting the patient result. | Clean the printer head with alcohol swab,Then check the print out but still it was not coming properly.Need to replace the printer head. | SWO-412315 | Revisit required | HW | Printer Module | Printer Head |  | Hazy print out quality. | Hazy print out coming,affecting the patient result. | Clean the printer head with alcohol swab,Then check the print out but still it was not coming properly. Need to replace the printer head.
No | Visit | Complaint | RDEU | Beim Aufstellen des Analysators wird festgestellt, dass die Inlet Türchen nicht gleichmäßig aufgehen. | Analysator in Betrieb. Ersatzteile geliefert, einige fehlen. Abstandslehre und passender kleiner Torx Schraubenzieher fehlt, nicht geliefert. | Ausbau der gesamten Mechanik und der Antriebsmotoren mit den Inlet doors aus dem alten Analysator, gereinigt und Einbau in den neuen Analysator. Nach Neustart und Aufwärmphase Test der Funktion, in Ordnung. Die Inlet doors lassen sich ohne Beanstandung öffnen und schließen. Qualitätskontrolle durchgeführt. | SWO-412113 | Defect | HW | Cover and Chassis | Door, Inlet |  | When setting up the analyzer, it is determined that the inlet doors do not rise evenly. | analyzer in operation. Spare parts delivered, some are missing. Distance gauge and matching small Torx screwdriver is missing, not supplied. | Removal of the entire mechanics and drive motors with the inlet doors from the old analyzer, cleaned and installed in the new analyzer. After restart and warm-up test of the function, all right. The inlet doors can be opened and closed without complaint. quality control carried out.
No | Visit | Complaint | RPOL | Błąd transportu płynów | Wygięta igła. | Wyprostowanie igły. Testy przepływowe ok. Uruchomienie aparatu ok. | SWO-411872 | Adjustment | HW | Auto Inlet module | Probe bend |  | Fluid Transport Error | Bent Needle. | Straighting the needle. Flow tests OK. Start camera OK.
No | Visit | Complaint | RAME | Analyzer not networking with LIS after software upgrade to version 3.4. The analyzer ID is missing. | Analyzer not transmitting to RALs after upgrading to 3.4 version software. Jerry Kelly, Mike Pellittieri, Brian Vazquez, Emmett Williams, and Ryan Sedlacek assisted via phone and email. | Reinstalled RALS astm.ini file and upgraded software again. Still did not transmit. Downgraded software to version 3.3 MR1 and was able to get system working. During software upgrade to version 3.4, the Windows bombed out because the Service Pack was corrupt so Windows needed to be replaced. Analyzer unable to read my Windows thumb drive. Will return tomorrow with another Windows thumb drive. | SWO-411977 | Defect | SW | Windows | Service Pack corrupt | User Error, FSE (SW upgrade) | Analyzer not networking with LIS after software upgrade to version 3.4. The analyzer ID is missing. | Analyzer not transmitting to RALs after upgrading to 3.4 version software. Jerry Kelly, Mike Pellittieri, Brian Vazquez, Emmett Williams, and Ryan Sedlacek assisted via phone and email. | Reinstalled RALS astm.ini file and upgraded software again. Still did not transmit. Downgraded software to version 3.3 MR1 and was able to get system working. During software upgrade to version 3.4, the Windows bombed out because the Service Pack was corrupt so Windows needed to be replaced. Analyzer unable to read my Windows thumb drive. Will return tomorrow with another Windows thumb drive.
No | Visit | Complaint | RAME | Analyzer reported to not power up | After receiving parts, headed to facility. Turned power and mixer does have blue light, which indicates some power. | Replaced power supply to isolate issue. Ordered CPU, CF card and a display for overnight and will return. | SWO-412272 | Revisit required | HW | Battery module | Battery module |  | Analyzer reported to not power up | After receiving parts, headed to facility. Turned power and mixer does have blue light, which indicates some power. | Replaced power supply to isolate issue. Ordered CPU, CF card and a display for overnight and will return.
No | Visit | Complaint | RLTD | Rhiannon Tel. 01823 346747 | I replaced the sensor cassette with a test cassette to see if this would get to the required temperature and the Substrate Temperature did come in but the Bottom Temperature was still below limits. It did start to slowly rise but the Bottom Power was at 0% so it looks like the Bottom Temperature was just increasing due to the ambient temperature. | I rebooted the analyser and it went through its start-up with no temperature errors present. After the startup I went in to the service programs to check the temperatures and all were ok. I also checked the operation of the inlet after the problems yesterday and it all appeared ok. | SWO-412456 | Adjustment | Power OFF/ON | Power OFF/ON | Power OFF/ON | Temperature | Rhiannon Tel. 01823 346747 | I replaced the sensor cassette with a test cassette to see if this would get to the required temperature and the Substrate Temperature did come in but the Bottom Temperature was still below limits. It did start to slowly rise but the Bottom Power was at 0% so it looks like the Bottom Temperature was just increasing due to the ambient temperature. | I rebooted the analyser and it went through its start-up with no temperature errors present. After the startup I went in to the service programs to check the temperatures and all were ok. I also checked the operation of the inlet after the problems yesterday and it all appeared ok.
No | 1 | Y | RPAC | Blank | Calibration failure on pCO2, K, Na |  | D4EA5FF5902CB204 | Assisted Replacement | SC | SC | pCO2, K, Na |  | Blank | Calibration failure on pCO2, K, Na
No | Visit | Complaint | RLTD | Claire Bailey Tel: 01708 435000 | Analyser front casing is damaged. | Replaced Analyser front casing (886-960). Analyser has passed calibrations and QC. | SWO-412559 | Defect | HW | Basic Unit | Front casing |  | Claire Bailey Tel: 01708 435000 | Analyser front casing is damaged. | Replaced Analyser front casing (886-960). Analyser has passed calibrations and QC.
No | Visit | Complaint | RAME | Customer called stating analyzer failing calibration for pCO2. | Analyzer pCO2 calibration value was 29% | Replaced SC on analyzer. | SWO-412864 | Assisted Replacement | SC | SC | pCO2 | Cal Sens | Customer called stating analyzer failing calibration for pCO2. | Analyzer pCO2 calibration value was 29% | Replaced SC on analyzer.
No | Visit | Complaint | RLTD | Gareth called to say that there was a battery error on the analyser. | 13XX - Battery Error Detected came up when the analyser was moved to its department, a UPS was fitted while waiting for a new battery module. | New battery module fitted. | SWO-412591 | Defect | HW | Battery module | Battery module |  | Gareth called to say that there was a battery error on the analyser. | 13XX - Battery Error Detected came up when the analyser was moved to its department, a UPS was fitted while waiting for a new battery module. | New battery module fitted.
No | Visit | Complaint | RFRA | Err status2 |  | Préparation de la cfcard à l'atelier $Réinstallation de W7 $Réinstallation de l'application Radiometer $Restauration de la configuration $Impossible de restaurer la base de données $Se met en err status 2 lors du chargement de la sauvegarde récupérée sur le réseau $Remise en route $Tests FLUID TRANSPORT ok $Calibration et passage des CQ. | SWO-413120 | Defect | HW | DMS_PC | CF card | status code 2 | Err status2 |  | Preparing the cfcard at the workshop $R installation of W7 $R installation of the Radiometer application $Restauration the configuration $Impossible restore the database $Se err status 2 when loading the backup recovered on the network $Remise en route $Tests FLUID TRANSPORT ok $Calibration and passage of the CQ.
No | Visit | Complaint | RSCH | Rep | Diverse Parameter nicht mehr Messbar Diverse Sensor Ansprachefehler Verstopfung ! Kunde misst nicht mit Radiometer Probenehmer. | Sensorkassette ausgetauscht Initialisierung 1 Stunde Kal nach 90min i.o QC i.o | SWO-412167 | Assisted Replacement | SC | SC | SC | Clot | Rep | Various parameters no longer measurable Various sensor speech errors Constipation! Customer does not measure with radiometer sampler. | Sensor cartridge replaced Initialization 1 hour Kal after 90min i.o QC i.o
No | Visit | Complaint | RCHN | 连续更换3张测试卡后仍然多参数定标失败 | 连续更换3张测试卡后仍然多参数定标失败，借902-972一个 | 更换902-972后修复 by jesson | SWO-411705 | [5 empty] | After replacing 3 test cards in a row, multi-parameter calibration failed | After replacing 3 test cards in a row, the multi-parameter calibration failed, borrowing 902-972 one | Repair by jesson after replacement 902-972
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 定标频繁，试剂包损耗太快 | 到达后发现今早安装了测试卡，由于装卡用将近一个小时，护士长急得不行，多次打电话抱怨，护士长意思是他们就早上集中抽样做，抽的样本都坏掉了，想在要一台机器。试剂包经过上次建议休眠，这个试剂包已经做了192个样本，还有17个溶液2 ，建议继续启用休眠，同时和荣诚沟通建议给加一台机器。 | SWO-412782 | [5 empty] | Test card, reagent package failure | Frequent calibration, reagent package loss too fast | Upon arrival, I found that the test card was installed this morning, because the card was used for nearly an hour, the head nurse was too anxious, repeatedly called to complain, the head nurse meant that they concentrated in the morning sampling, the samples were broken, want to get a machine. The reagent package has been recommended for hibernation, which has 192 samples and 17 solutions 2, and recommends continuing to enable sleep, while communicating with Rongcheng suggesting adding a machine.
No | Visit | Complaint | RCHN | PO2病人结果偏高，接口板腐蚀 | PO2病人结果偏高，接口板腐蚀借902-972一个。 | 更换902-972备件后修复 by jesson | SWO-411726 | [5 empty] | PO2 patient results are high, interface plate corrosion | PO2 patient results are high, interface plate corrosion borrowed 902-972 one. | Repair by jesson after replacement of 902-972 spare parts
No | Visit | Complaint | RAME | Possible fluid blockage errors and purple solution is stuck in cassette window. | Possible blockage error. Found analyzer could not pull solution 1 from the solution pack. All other solutions checked okay. | Reseated the solution pack. | SWO-412981 | [5 empty] | Possible fluid blockage errors and purple solution is stuck in cassette window. | Possible blockage error. Found analyzer could not pull solution 1 from the solution pack. All other solutions checked okay. | Reseated the solution pack.
No | Visit | Complaint | RCHN | 报错0581 | 报错0581 借902-980一个 | 更换902-980后修复 by jesson | SWO-411615 | Defect | HW | Oxi module | Hemolyzer | 581 | Error 0581 | Error 0581 Borrowing 902-980 one | Repair by jesson after replacement 902-980
No | Visit | Complaint | RCHN | 科室反映数值不准 | 科室反映数值不准 | 检查并清洁管路，怀疑是样本有凝血造成结果偏高。by jesson | SWO-412453 | User Error | preanalytical | preanalytical | Wet section | Clot | Section reflects the value is not allowed | Section reflects the value is not allowed | Check and clean the lines, suspecting that the sample has a clotting result resulting in high results. by jesson
No | Visit | Complaint | RCHN | 试管夹破损 | 试管夹破损，借903-584一个 | 更换903-584后修复 by jesson | SWO-412318 | Defect | HW | Aspiration Module | Tube holder |  | The test tube clip is broken | The test tube clip is broken, borrow 903-584 one | Repair by jesson after replacing 903-584
Yes | Visit | Complaint | RCHN | 报错1067 | 报错1067，借944-258（14485缺货）一个 | 上次维修，漏借一个批号的定标卡，重新定标后修复 by jesson | SWO-412329 | Re classified RMED | SWO on wrong analyzer |  |  | Mail from Daniel Yang to INGZI | 1067 error | False 1067, borrow 944-258 (14485 out of stock) one | Last repair, debit a lot number of calibration card, re-calibration after repair by jesson
No | Visit | Complaint | RAME | 531 Ph liqui sensor error | Error 0531 pH $BG upper liquid sensor error and error 0946 conditioning error in MetII electrode module. | Replaced defective pH $BG module gas valve. | SWO-412980 | Defect | HW | pH/BG Module | Valve, air | 531 | 531 Ph liqui sensor error | Error 0531 pH /BG upper liquid sensor error and error 0946 conditioning error in MetII electrode module. | Replaced defective pH $BG module gas valve.
No | Visit | Complaint | RAME | Intermittent, frequent error 791, no Oximetry results, parameter could not be measured. | Reviewed data logs and applicable service diagnostics. Noted under Oximetry testing, the neon lamp failed due to lamp intensity below specification. | Replaced neon lamp, performed applicable adjustments and testing, no further failures observed. | SWO-412778 | Defect | HW | Oxi module | Lamp unit | neon lamp | Intermittent, frequent error 791, no Oximetry results, parameter could not be measured. | Reviewed data logs and applicable service diagnostics. Noted under Oximetry testing, the neon lamp failed due to lamp intensity below specification. | Replaced neon lamp, performed applicable adjustments and testing, no further failures observed.
No | Visit | Complaint | RDEU | 10.09.2018; Sr. Franziska; Tel.: 03432218452; ABL bootet nach einem Stromausfall nicht emhr hoch, Status Code 9 | Gerät bootet nicht, bleibt im Startvorgang hängen Fehler besteht seit Stromabschaltung $ Notstromtest | DMS Software erneut installiert, dabei Haken für Roma.ini Upgrade entfernt -> Neustart, Sprache eingestellt, Helios Cco Dateien wieder eingespielt, Netzwerkeinstellungen geprüft-> einspricht Daten im CRM Kal & QC Daten gemäß RDEU Vorgaben gesichert u. dokumentiert | SWO-412485 | Defect | SW | RMED SW | Roma.ini |  | 10.09.2018; Sr. Franziska; Tel.: 03432218452; ABL does not boot up after a power failure, status code 9 | Device does not boot, gets stuck in the boot process Error persists since power shutdown - Emergency current test | DMS software reinstalled, removing hook for Roma.ini upgrade -> Restart, voice set, Helios Cco files replayed, network settings checked-> indentdata in CRM Kal & QC data secured according to RDEU specifications and documented
No | Visit | Complaint | RHUN | mérés közben, új minta beadásakor a minta ajtó nem záródik be, mérés megszakad, visszaállítás folyamatban, újra indítás nem segít | minta ajtó záródása néha bizonytalan, nyitva marad, a készülék megszakítja a folyamatban lévő mérést | minta ajtó ellenőrzése, szerviz tesztek többszöri megismétése, inicializálás, 2 teszt mérés egymás utáni indítással, LQC mérés | SWO-412427 | Defect | HW | Cover and Chassis | Door, Inlet |  | During measurement, when a new sample is given, the sample door does not close, measurement is interrupted, reset is in progress, Re-booting does not help | Pattern door closure is sometimes uncertain, remains open, the device interrupts the current measurement | Sample door check, repeated service tests, initializer, 2 test measurement by consecutive start, LQC measurement
No | Visit | Complaint | RKOR | FTC_105 : P1 - P5 tesf failed | Flow test -> P1 - P5 Failed Re-install the pump, tubing Clean the neddle wash block Clean the all flow path Flow test -> P1 - P5 Failed Air trap check -> empty Replace the solution pack | Air trap is empty. Replace the solution pack | SWO-414968 | Defect | Consumables | Solution Pack | Valves | FTC_105 | FTC_105 : P1 - P5 tesf failed | Flow test -> P1 - P5 Failed Re-install the pump, tubing Clean the neddle wash block Clean the all flow path Flow test -> P1 - P5 Failed Air trap check -> empty Replace the solution pack | Air trap is empty. Replace the solution pack
No | Visit | Complaint | RESP | Equipo no enciende y emite un pitido fuerte y continuo | Chequeo de Power supply, cf card y modulo pc | Reinicio completo, chequeo de conexiones Se procede a petición de piezas | SWO-416373 | Revisit required | HW | DMS_PC | CPU |  | Equipment does not turn on and emits a loud and continuous beep | Check power supply, cf card and pc module | Full reboot, connection check Proceeded at the request of parts
No | Visit | Complaint | RAME | Analyzer not networking with LIS after software upgrade to version 3.4. The analyzer ID is missing. | Need to program analyzer with Windows software. | Unable to get thumb drive to program the analyzer. Attempted to reload the thumb drive. | SWO-413019 | Revisit required | SW | Windows | thumb drive | User Error, FSE (SW upgrade) | Analyzer not networking with LIS after software upgrade to version 3.4. The analyzer ID is missing. | Need to program analyzer with Windows software. | Unable to get thumb drive to program the analyzer. Attempted to reload the thumb drive.
No | Visit | Complaint | RAME | Analyzer not networking with LIS after software upgrade to version 3.4. The analyzer ID is missing. | Analyzer not transmitting to RALs after upgrading to 3.4 version software. | Created Windows programming drive. Reinstalled Windows, ROMA, data logs and RALS astm.ini file onto the analyzer and upgraded software again. Performed calibrations and ran QC to test transmission, and the analyzer is now transmitting as it should. Phone help from Jim Gibbons and Sean Hecht. | SWO-413592 | Defect | SW | Windows | thumb drive | User Error, FSE (SW upgrade) | Analyzer not networking with LIS after software upgrade to version 3.4. The analyzer ID is missing. | Analyzer not transmitting to RALs after upgrading to 3.4 version software. | Created Windows programming drive. Reinstalled Windows, ROMA, data logs and RALS astm.ini file onto the analyzer and upgraded software again. Performed calibrations and ran QC to test transmission, and the analyzer is now transmitting as it should. Phone help from Jim Gibbons and Sean Hecht.
No | Visit | Complaint | RAME | Analyzer not networking with LIS after software upgrade to version 3.4. The analyzer ID is missing. | Analyzer not transmitting to RALs after upgrading to 3.4 version software. | Reinstalled Windows, ROMA, data logs and RALS astm.ini file onto the analyzer and upgraded software again. Performed calibrations and ran QC to test transmission, and the analyzer is now transmitting as it should. | SWO-413625 | Defect | SW | Windows | Windows | User Error, FSE (SW upgrade) | Analyzer not networking with LIS after software upgrade to version 3.4. The analyzer ID is missing. | Analyzer not transmitting to RALs after upgrading to 3.4 version software. | Reinstalled Windows, ROMA, data logs and RALS astm.ini file onto the analyzer and upgraded software again. Performed calibrations and ran QC to test transmission, and the analyzer is now transmitting as it should.
No | Visit | Complaint | RCHN | 未连接电极 | 未连接电极 借902-778一个 | 更换902-778后修复 | SWO-411667 | [5 empty] | Electrodes not connected | Unconnected electrode borrows 902-778 one | Repair after replacement 902-778
No | Visit | Complaint | RCHN | LAC,GLU定标失败 | LAC,GLU定标失败，0376，借903-581一个。 | 更换903-581后修复 by jesson | SWO-412489 | Defect | HW | Basic Unit | PCB, Sensor interface | Glu+Lac | LAC, GLU bid failed | LAC, GLU bid failure, 0376, borrowed 903-581 one. | Repair by jesson after replacement 903-581
No | Visit | Complaint | RCHN | 报错1031 | 报错1031 借903-364一个，944-258(14591) 944-220(14531) 944-394(14412) 944-212(14558) 944-216（14567）944-214（14460）各一个 | 更换903-364，重新定标后修复by jesson | SWO-411601 | Defect | HW | Process Module | Optical Unit | 1031 | 1031 | 1031 Borrowing 903-364 one, 944-258 (14591) 944-220 (14531) 944-394 (14412) 944-212 (14558) 944-216 (14567) 944-214 (14441) | Replace 903-364, fix by jesson after recalibration
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 无法读取处理包 | 更换试剂包后修复 | SWO-413719 | Defect | Consumables | Solution Pack | Chip |  | Test card, reagent package failure | The processing package could not be read | Repair after replacement of reagent kit
No | Visit | Complaint | RCHN | 报错5297,1024 | 报错5297,1024，借903-702一个 | 更换备用机备件903-702后修复 by jesson | SWO-412451 | Defect | HW | Process Module | Process Wheel | 1024 | 52971024 | 5297,1024, loan 903-702 one | Repair by jesson after replacing spare parts 903-702
No | Visit | Complaint | RCHN | 进样门关闭打开卡,开不开,关不上,新装机,刚刚做十几个样本就这个 | 进样门关闭打开卡,开不开,关不上 | 调整进样门组件 | SWO-413406 | Adjustment | HW | Cover and Chassis | Door, Inlet |  | The sample door closed open card, open, closed, new installation, just do a dozen samples on this | The sample door closes the open card, opens, can't close | Adjust the sample door assembly
No | Visit | Complaint | RCHN | 新机器进样门开关不利索 | 新机器进样门开关不利索，进样门左侧滑轨空隙太大， | 进样门左侧滑轨空隙太大，用胶布贴了一下，把间隙限制小 | SWO-416910 | Adjustment | HW | Cover and Chassis | Door, Inlet | Gear rack | New machine feed door switch adverse rope | The new machine into the sample door switch adverse rope, the left rail gap on the left side of the sample door is too large, | The left rail gap on the left side of the sample door is too large, with a adhesive cloth paste, limit the gap to a small
No | Visit | Complaint | RDEU | 12.09.2018, 08:03 Uhr, H. Adebar, Tel. 06213833528: Kal Fehler bei Lyten | KAL Driften Elektrolyte | Allgemeine Überprüfung - Erneuerung Referenz Elektrode - Funktionsüberprüfung etc. | SWO-413294 | Assisted Replacement | Electrode | Electrode | Reference |  | 12.09.2018, 08:03, H. Adebar, Tel. 06213833528: Kal error at Lyten | KAL Drifting Electrolytes | General verification - Renewal reference electrode - function check etc.
No | Visit | Complaint | RAME | MetHb is off by 0.6-0.8 from other 2 analyzers on blood check | Customer is still having 0.5-0.6 difference on MetHb from Reference analyzer. | Adjusted lamp to 3319, did complete oximetry test, failed for neon intensity, reading 40% and it should be over 50%. Replaced lamp adjusted to 3323 Did Complete oximetry test, did THB in service and as user and as sample. | SWO-413656 | No fault found | HW | Oxi module | Lamp unit | neon lamp | MetHb is off by 0.6-0.8 from other 2 analyzers on blood check | Customer is still having 0.5-0.6 difference on MetHb from Reference analyzer. | Adjusted lamp to 3319, did complete oximetry test, failed for neon intensity, reading 40% and it should be over 50%. Replaced lamp adjusted to 3323 Did Complete oximetry test, did THB in service and as user and as sample.
No | Visit | Complaint | RAME | MetHb is off by 0.6-0.8 from other 2 analyzers on blood check | Customer is still having 0.5-0.6 difference on MetHb from Reference analyzer. | Adjusted lamp to 3315, did complete oximetry test, did THB in service and as user and as sample. | SWO-413654 | No fault found | HW | Oxi module | Lamp unit |  | MetHb is off by 0.6-0.8 from other 2 analyzers on blood check | Customer is still having 0.5-0.6 difference on MetHb from Reference analyzer. | Adjusted lamp to 3315, did complete oximetry test, did THB in service and as user and as sample.
No | Visit | Complaint | RAME | Customer seen CoHb values during tHb cal verification test. | CoHb value was low when customer ran tHb Calibration solution as sample test. | Performed the tHb cal in Service mode, it gave a passed tests results. Ran tHb Calibration verification and the numbers for all oxymetry parameters met the numbers from insert sheet. Customer was trained how to run tHb Calibration in Service mode in case if they need it in future. Analyzer is up and running ready for real sample testing. | SWO-413556 | Adjustment | HW | Oxi module | tHb Cal | Service mode | Customer seen CoHb values during tHb cal verification test. | CoHb value was low when customer ran tHb Calibration solution as sample test. | Performed the tHb cal in Service mode, it gave a passed tests results. Ran tHb Calibration verification and the numbers for all oxymetry parameters met the numbers from insert sheet. Customer was trained how to run tHb Calibration in Service mode in case if they need it in future. Analyzer is up and running ready for real sample testing.
No | Visit | Complaint | RDEU | 05.09.2018 Herr Lein 03761444 meldet laute Geräusche. | Fehlersuche ==> laute Geräusche beim Ent- und Verriegeln des Lösungspacks | Stepper motor actuator erneuert | SWO-414021 | Defect | HW | Basic Unit | Actuator unit |  | 05.09.2018 Mr. Lein 03761444 reports loud noises. | Troubleshooting ==> loud noises when unlocking and locking the solution pack | Stepper motor actuator renewed
No | Visit | Complaint | RAME | Analyzer giving inhomogeneous solution errors. | Found inlet probe to be bent as well as inlet gasket to have large crack. | Replaced probe and gasket with customer stocked parts. | SWO-414018 | Defect | HW | Inlet module | Inlet gasket |  | Analyzer giving inhomogeneous solution errors. | Found inlet probe to be bent as well as inlet gasket to have large crack. | Replaced probe and gasket with customer stocked parts.
No | Visit | Complaint | RAME | Not Networking with LIS | Analyzer not transmitting to RALs after upgrading to 3.4 version software. | Reinstalled Windows, ROMA, data logs and RALS astm.ini file onto the analyzer and upgraded software again. Performed calibrations and ran QC. Level B oximetry is out of range but the calibrations do pass. Will check with the customer in the morning to make sure it is transmitting as it should. | SWO-413649 | Defect | SW | Windows | Windows | User Error, FSE (SW upgrade) | Not Networking with LIS | Analyzer not transmitting to RALs after upgrading to 3.4 version software. | Reinstalled Windows, ROMA, data logs and RALS astm.ini file onto the analyzer and upgraded software again. Performed calibrations and ran QC. Level B oximetry is out of range but the calibrations do pass. Will check with the customer in the morning to make sure it is transmitting as it should.
No | Visit | Complaint | RAME | Analyzer reported to not power up | Unit does not power up properly and doesn't t operate when unplugged (has battery) | Replaced battery assembly. Ran and test battery | SWO-414288 | Defect | HW | Battery module | Battery module |  | Analyzer reported to not power up | Unit does not power up properly and doesn't t operate when unplugged (has battery) | Replaced battery assembly. Ran and test battery
No | Visit | Complaint | RFRA | Mme MAZARS ou Mme LOPEZ - 04.67.33.59.22 Depuis début Juillet, le site d'introduction reste bloqué ouvert lors du passage d'un échantillon. Le joint reste enfoncé et ne revient pas. | Je joint d'introduction ne revient pas bien. Mode service : contrôle position seringue, capillaire, position sonde enfoncée. Ok Problème non constaté. | Le laboratoire me signale sonde tordue, pourquoi ? La sonde ne peut pas se tordre toute seule, Possibilité, si la seringue bouge pendant l'aspiration. Chgt module inlet. Nettoyer de tps en tps la sonde pour que le joint d'introduction coulisse mieux. Contrôle en fonctionnement avec utilisatrice. | SWO-413897 | Defect | HW | Auto Inlet module | Auto Inlet module | Spring | Ms MAZARS or Mrs LOPEZ - 04.67.33.59.22 Since the beginning of July, the introductory site remains blocked open when passing a sample. The seal remains pressed and does not return. | I joint introduction does not come back well. Service mode: syringe position control, capillary, pressed probe position. Ok Unrecognized problem. | The lab reports me twisted probe, why? The probe can't twist itself, Possibility, if the syringe moves during suction. Chgt module inlet. Clean tps in tps the probe so that the introductory seal slides better. Control in operation with user.
No | Visit | Complaint | RDEU | 12.09.2018 Pfl. Sören 05191800184 meldet, dass der ABL immer wieder ein neues LP haben möchte. 12.09.2018 15:00 Pfl. Sören 05191800184 Internvention erforderlich, mehrfaches durchspülen hat nicht gebracht, keine Betriebsmittel mehr vorrätig. | System überprüft => Lp defekt Lp aus GS geholt da Kunde keines mehr hat | LP erneuert, Kalibration und Qc durchgeführt | SWO-413832 | Assisted Replacement | SP | SP | FT |  | 12.09.2018 Pfl. Sören 05191800184 announces that the ABL wants to have a new LP again and again. 12.09.2018 15:00 Pfl. Sören 05191800184 Internvention required, multiple flushing has not brought, no more equipment in stock. | System checked => Lp defective Lp fetched from GS since customer has none anymore | LP renewed, calibration and Qc performed
No | Visit | Complaint | RAME | Battery Error | Battery Error | Replaced the battery with a new one. | SWO-414117 | Defect | HW | Battery module | Battery module |  | Battery Error | Battery Error | Replaced the battery with a new one.
No | Visit | Complaint | RDEU | am Radiometer-Messgerät Werte um 1.5 % (Probanden auch Nicht-Raucher, als Ref.-Wert werden bis zu 1 % akzeptiert (Thomas)), ca. 50 – 70 % höhere Ergebnisse als am Messgerät von Roche im Labor, QC in Ordnung Tech-Line eröffnet | No | Cleaning tool gestartet, Start cleaning procedure, Meldung insert regulär SP, aber das steckt! Somit komme ich in der procedure nicht weiter, also Vorgang abgebrochen und Gerät wieder in Messbereitschaft gebracht, ServiceHandbuch bestätigt Vorgehensweise | SWO-415043 | [5 empty] | on the radiometer meter values by 1.5% (subjects also non-smokers, as ref.-value up to 1% are accepted (Thomas)), approx. 50 – 70% higher results than on the Roche measuring instrument in the laboratory, QC fine Tech-Line opens | No | Cleaning tool started, start cleaning procedure, message insert regular SP, but that's in! Thus I can't get any further in the procedure, so process aborted and device brought back into measurement readiness, ServiceManual confirms procedure
No | Visit | Complaint | RCHN | 打印机接口的探灯有问题 | 打印机接口的针脚损坏，需要借件910-286 | 更换910-286后修复 by jesson | SWO-413125 | Defect | HW | Printer Module | Printer |  | There is a problem with the searchlight on the printer interface | The pin of the printer interface is damaged and requires a loan of 910-286 | Repair by jesson after replacement 910-286
No | Visit | Complaint | RCHN | 报错1031 | 报错1031，借944-212（14558）、944-214（14322、14465）、944-216（14616）、944-258（14702）、944-220（14531）、944-394（14170）各一张 | 更换903-364后报错1024，加借903-364一个，944-212（14558）一张。by jesson | SWO-412394 | Defect | HW | Process Module | Optical Unit | 1031 | 1031 | 1031, borrowing 944-212 (14558), 944-214 (14322, 14465), 944-216 (14616), 944-258 (14702), 944-220 (14531), 944-394 (14170) each | After replacing 903-364, report error 1024, add 903-364 one, 944-212 (14558) one. by jesson
No | Visit | Complaint | RCHN | 试管卡槽开裂 | 试管卡槽开裂,借903-584两个。 | 更换903-584两个正常。by jesson | SWO-413149 | Defect | HW | Aspiration Module | Tube holder |  | Test tube card groove cracking | The test tube card groove cracked, borrowing 903-584 two. | Replace 903-584 both normal. by jesson
No | Visit | Complaint | RCHN | 报错5684 | 报错5684，借944-212（14757）一个 缺货 | 重新定标后修复 by jesson | SWO-413203 | Defect | SW | RMED SW | Motor, process wheel | 5684 | 5684 error | 5684, loan 944-212 (14757) one out of stock | Fix by jesson after re-calibration
No | Visit | Complaint | RCHN | 报错1020 | 报错1020，需要借件636-652 | 更换636-652后修复 by jesson | SWO-413071 | Defect | HW | Incubator Module | Cable | 1020 | 1020 | 1020, 636-652 for a loan | Repair by jesson after replacing 636-652
No | Visit | Complaint | RDEU | Flüssigkeit läuft aus dem Gerät | Gerät überprüft und gereinigt | Cl neu membraniert, unteren Dichtungsring bei der Lac Elektrode ersetzt, Lac neu membraniert, Gas2 erneuert, tHb Kal, Kals und QCs durchgeführt | SWO-413906 | Defect | Electrode | Membrane | Cl | Mempack missing | Liquid runs out of the device | Device checked and cleaned | Cl re-membraneed, lower sealing ring replaced at lac electrode, Lac re-membraneed, Gas2 renewed, tHb Kal, Kals and QCs performed
No | Visit | Complaint | RAME | Customer states that all oximetry is out and cannot be measured at this time. | Verified customer's stated issue and discovered it was happening intermittently after the night before. Upon troubleshooting the unit, a clot was discovered in the pH chamber. | Cleaned out the ph $BG chamber and the pH electrode. Flushed the unit from the inlet on, replaced the waste tubings and remembraned the Ref electrode. Verified the lamp status and adjusted it in service mode. The unit then was put through several 2 point calibrations until all the electrolyte and oximetry parameters passed. Once all parameters were green, all the levels of QC were run. 9 $14 Followed up with customer to ensure analyzer was still functioning properly. Customer reports no issues at this time. | SWO-414149 | Cleaned | Fluidic pathway | Fluidic pathway | Oxi module | Clot | Customer states that all oximetry is out and cannot be measured at this time. | Verified customer's stated issue and discovered it was happening intermittently after the night before. Upon troubleshooting the unit, a clot was discovered in the pH chamber. | Cleaned out the ph $BG chamber and the pH electrode. Flushed the unit from the inlet on, replaced the waste tubings and remembraned the Ref electrode. Verified the lamp status and adjusted it in service mode. The unit then was put through several 2 point calibrations until all the electrolyte and oximetry parameters passed. Once all parameters were green, all the levels of QC were run. 9 $14 Followed up with customer to ensure analyzer was still functioning properly. Customer reports no issues at this time.
No | Visit | Complaint | RAME | Customer states analyzer is giving oximetry measuring errors and all oxi is out on patient results and QC. | Verified customer's stated issue. Upon looking at the data logs it was discovered that the errors present indicated an issue with the halogen lamp. It was reading too low. | Adjusted the lamp, performed a complete oxi test and ran a tHb calibration both in service and user mode with the tHb results passing at this time. After running a few 2 point calibrations the Lac membrane had to be replaced due to it not passing level 2 and 3 on QC. 9 $14 Followed up with customer to ensure analyzer was still functioning properly. Customer reports no issues at this time. | SWO-414145 | Adjustment | HW | Oxi module | Lamp unit |  | Customer states analyzer is giving oximetry measuring errors and all oxi is out on patient results and QC. | Verified customer's stated issue. Upon looking at the data logs it was discovered that the errors present indicated an issue with the halogen lamp. It was reading too low. | Adjusted the lamp, performed a complete oxi test and ran a tHb calibration both in service and user mode with the tHb results passing at this time. After running a few 2 point calibrations the Lac membrane had to be replaced due to it not passing level 2 and 3 on QC. 9 $14 Followed up with customer to ensure analyzer was still functioning properly. Customer reports no issues at this time.
No | Visit | Complaint | RDEU | Immer wieder Kal Driften bei Na, immer der gleiche Drift Wert von 4 fast immer Nachts | Immer wieder Drift bei Na | Ref Elektrode ersetzt, Kal und QC durchgeführt, Anwender beobachtet und meldet sich bei weiteren Problemen | SWO-413797 | Assisted Replacement | Electrode | Electrode | Reference | CAPA-00733 | Again and again Kal Driften at Na, always the same drift value of 4 almost always at night | Drift again and again at Na | Ref electrode replaced, Kal and QC performed, users observed and reports further problems
No | Visit | Complaint | RAME | Analyzer not running qc | QC not running | Found out that customer tried to change pump tubes. They did not connect it properly. Used wrong tube for solution pump. Waste tube damaged on capillary side. Redid the tubing. Replaced reference membrane due to QC failing on pH and Electrolyte. Cleaned measuring chamber for pH. Replaced with sealing ring for po2. All calibrations and 3 Levels of QC passed as per manufacturer specification | SWO-413686 | User Error | HW | pH/BG Module | Pump tubes | Wrong type | Analyzer not running qc | QC not running | Found out that customer tried to change pump tubes. They did not connect it properly. Used wrong tube for solution pump. Waste tube damaged on capillary side. Redid the tubing. Replaced reference membrane due to QC failing on pH and Electrolyte. Cleaned measuring chamber for pH. Replaced with sealing ring for po2. All calibrations and 3 Levels of QC passed as per manufacturer specification
No | Visit | Complaint | RIND | Oximetry hardware problem shows | Check system message in which Neon number too low error shows and calibration also shows error of Oximetry hardware problem. | Run complete OXI test in which neon position is failed.Check by replacing neon lamp but still neon position failed.Then check spectrometer and found the problem in spectrometer which need to replace to resolve the problem. | SWO-414885 | Revisit required | HW | Oxi module | Oxi module |  | Oximetry hardware problem shows | Check system message in which Neon number too low error shows and calibration also shows error of Oximetry hardware problem. | Run complete OXI test in which neon position is failed. Check by replacing neon lamp but still neon position failed. Then check spectrometer and found the problem in spectrometer which need to replace to resolve the problem.
Yes | Visit | Complaint | RAME | Downgrade ABL 837 to 835 | Downgrade ABL 837 to 835 | Downgrade ABL 837 to 835. Installed and ran new reagents and QC. Conducted analyzer training. | SWO-414058 | Re classified RMED | planned SWO |  |  |  | Downgrade ABL 837 to 835 | Downgrade ABL 837 to 835 | Downgrade ABL 837 to 835. Installed and ran new reagents and QC. Conducted analyzer training.
No | Visit | Complaint | RDEU | Oximetrie Hardwarefehler 0315 | Oximetrie Hardwarefehler 0315 | Gerät überprüft, Oximetrietest durchgeführt, Oxi Modul ausgetauscht, Kal, tHb Kal und QC o.k. Gerät auf und online gestellt | SWO-413777 | Defect | HW | Oxi module | Oxi module | 315 | Oximetrie Hardwarefehler 0315 | Oximetrie Hardwarefehler 0315 | Device checked, oximetry test performed, Oxi module replaced, Kal, tHb Kal and QC o.k. Device put on and online
No | Visit | Complaint | RKOR | TnI Calibration error | 1024 TnI Calibration error 내부의 청소 문제라고 판단 | Clean inner tube Clean reference of process wheel. Clean reflector bottom of process wheel. | SWO-421411 | Cleaned | HW | Process Module | Process Wheel | 1024 | TnI Calibration error | 1024 TnI Calibration error I think it's a cleaning problem inside | Clean inner tube Clean reference of process wheel. Clean reflector bottom of process wheel.
No | Visit | Complaint | RAME | Ca is failing calibration and qc |  |  | SWO-414545 | No fault found | No fault found | No fault found | No fault found | Ca | Ca is failing calibration and qc
No | 4 | Y | RCZE | Blank | Jehla se nevraci do sve puvodni polohy po nasati vzorku. |  | D4EE2FA2F25BB912 | Defect | HW | Auto Inlet module | Probe bend |  | Blank | Do not drill the needle into the water position after the sample has been saturated.
Yes | Visit | Complaint | RSAF | Printer jam, not printing properly |  |  | SWO-414467 | Re classified RMED | duplicate upSWO |  |  |  | Printer jam, not printing properly
No | Visit | Complaint | RSAF | Printer damaged | Missing brass ring from the printer roller. | Replaced with a new printer unit due to the missing brass Ring | SWO-415094 | Defect | HW | Printer Module | Printer |  | Printer damaged | Missing brass ring from the printer roller. | Replaced with a new printer unit due to the missing brass Ring
No | Visit | Complaint | RCHN | 报错508冲洗错误 | 报错508冲洗错误，借902-980一个 | 902-980内部管路堵塞，无法冲洗。更换902-980一个正常。by jesson | SWO-413714 | Defect | HW | Oxi module | Hemolyzer | Clot | Report error 508 flushing error | Reported error 508 flushing error, borrow 902-980 one | The internal lines of the 902-980 are blocked and cannot be flushed. Replace the 902-980 one normal. by jesson
No | Visit | Complaint | RCHN | 报错1031 | 报错1031,加借903-364一个。944-212（14558）一张。 | 更换903-364备件和重新定标后修复 by jesson | SWO-413771 | Defect | HW | Process Module | Optical Unit | 1031 | 1031 | 1031, plus 903-364 one. 944-212 (14558) one. | Replacement of 903-364 spare parts and re-calibration repair by jesson
No | Visit | Complaint | RCHN | 报错1070 | 报错1070，需要借件903-210一个 | 更换903-210后修复 by jesson | SWO-413108 | Defect | HW | Process Module | PCB, cup dryer | 1070 | 1070 | 1070, need to borrow 903-210 one | Repair by jesson after replacement 903-210
No | Visit | Complaint | RCHN | 报错1024和1068 | 报错1024和1068，需要借件944-212（14402） | 清洁灰尘，在维修菜单中重新校准温度。重新定标后修复 by jesson | SWO-413068 | Cleaned | HW | Process Module | Optical Unit | 1024, 1068 | Errors 1024 and 1068 | 1024 and 1068, need to borrow 944-212 (14402) | Clean the dust and recalibrate the temperature in the service menu. Fix by jesson after re-calibration
No | Visit | Complaint | RIND | PCO2 calibration error & patient result abnormal. | Check the calibration log & found sensitivity of Pco2 was low. | Check the Pco2 electrode & found blood clot inside the chamber,Clean the clot,Then run the protein removal & cleaning few times,Then run the 2point calibration & it got passed successfully. | SWO-414861 | Cleaned | Fluidic pathway | pH/BG Module | pCO2 | Clot | PCO2 calibration error & patient result abnormal. | Check the calibration log & found sensitivity of Pco2 was low. | Check the Pco2 electrode & found blood clot inside the chamber,Clean the clot,Then run the protein removal & cleaning few times,Then run the 2point calibration & it got passed successfully.
No | Visit | Complaint | RIND | Hazy print out quality. | Hazy print out coming. | Replace the printer head with the new one,Then check the print out & it was working properly. | SWO-414855 | Defect | HW | Printer Module | Printer Head |  | Hazy print out quality. | Hazy print out coming. | Replace the printer head with the new one,Then check the print out & it was working properly.
No | Visit | Complaint | RIND | Cl- sensitivity low. | Cl- sensitivity goes down automatically after 4p.m calibration.Patient results are also abnormal. | Check the Cl- membrane & clean the Cl- electrode,Run the protein removal & 2 point calibration few times but still the sensitivity was down,Then bring a new Cl- membrane & then run 2point calibration 2times & it got passed successfully. | SWO-414858 | Assisted Replacement | Electrode | Membrane | Cl | Sensitivity low | Cl- sensitivity low. | Cl- sensitivity goes down automatically after 4p.m calibration. Patient results are also abnormal. | Check the Cl- membrane & clean the Cl- electrode,Run the protein removal & 2 point calibration few times but still the sensitivity was down,Then bring a new Cl- membrane & then run 2point calibration 2times & it got passed successfully.
No | Visit | Complaint | RAME | Customer reported they are having problems running QC. | Found sensor cassettes failed. One pCO2 low sensitivity. Second one low sensitivity on both Glu and Lac. Follow up with customer. From Friday visit. | Install 3rd sensor cassette. | SWO-416305 | Assisted Replacement | SC | SC | pCO2 | Low Sens | Customer reported they are having problems running QC. | Found sensor cassettes failed. One pCO2 low sensitivity. Second one low sensitivity on both Glu and Lac. Follow up with customer. From Friday visit. | Install 3rd sensor cassette.
Yes | Visit | Complaint | RSAF | No more activities left | Solution pack ran out of cycles | Replaced with a new solution pack | SWO-415095 | Re classified RMED | Logistical upSWO |  |  |  | No more activities left | Solution pack ran out of cycles | Replaced with a new solution pack
No | Visit | Complaint | RCHN | PCO2定标失败 0378 | PCO2定标失败 0378，借 903-581一个。 | 先后更换3张测试卡，故障现象均为PCO2定标不通过，更换903-581一个正常。by jesson | SWO-414443 | Defect | HW | Basic Unit | PCB, Sensor interface | pCO2 | PCO2 calibration failure 0378 | PCO2 bid failure 0378, borrowing 903-581 one. | Has replaced 3 test cards, the failure phenomenon is PCO2 calibration does not pass, replace 903-581 a normal. by jesson
No | Visit | Complaint | RIND | Hazy print out coming out. | Hazy print out coming out. | Replace the paper rolls with the new one to cross check but same problem persist.Reconnect all the connector to printer PCB,But still hazy print out coming out,Need to replace the print head for further check. | SWO-415816 | Revisit required | HW | Printer Module | Printer Head |  | Hazy print out coming out. | Hazy print out coming out. | Replace the paper rolls with the new one to cross check but same problem persist. Reconnect all the connector to printer PCB,But still hazy print out coming out,Need to replace the print head for further check.
No | Visit | Complaint | RKOR | Sample clot error | . | SP clean. Gasket module clean. Gasket hold replace. Service mode check. Calibration QC check. User training. | SWO-414989 | User Error | preanalytical | preanalytical | Clot |  | Sample clot error | . | SP clean. Gasket module clean. Gasket hold replace. Service mode check. Calibration QC check. User training.
No | Visit | Complaint | RKOR | Sample clot error | . | Liquid flow line clean. Service mode check. Cal QC check. | SWO-419983 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway | Clot | Sample clot error | . | Liquid flow line clean. Service mode check. Cal QC check.
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | HCT,CL,NA N $E,N $C | 更换测试卡后修复 | SWO-416510 | [5 empty] | Test card, reagent package failure | HCT,CL,NA N /E,N /C | Repair after replacing test card
No | Visit | Complaint | RCHN | 光学强度超范围 1031，定标卡扫描失败 | 光学强度超范围 1031，定标卡扫描失败,借903-364，902-924各一个。 | 更换903-364，902-924后修复（By Ran） | SWO-414458 | Defect | HW | Process Module | Optical Unit | 1031 | Optical strength over range 1031, bid card scan failed | Optical strength over range 1031, calibration card scanning failure, borrow 903-364, 902-924 each one. | Replacement 903-364,902-924 post-repair (By Ran)
No | Visit | Complaint | RCHN | 触摸屏无反应 | 触摸屏无反应,借910-324一个。 |  | SWO-414930 | Defect | HW | Analyzer Control Module | display |  | Touch screen does not respond | Touch screen does not respond, borrow 910-324 one.
No | Visit | Complaint | RDEU | 14.09.2018; Sr. Rebecca Hünefeld?; Tel.: 0698505250; Probenabbrüche | Intervention erforderlich Leckage erkannt Messungen wurden zuvor abgebrochen | Allgemeine Überprüfung - Verkrustungen an SP entfernt- Erneuerung Schlauch an Quetschventil - Service Justagen etc. durchgeführt - Analysator gespült - Erneuerung Einlass und Anschlussdichtung - Alle Tests bestanden, trotzdem Leckage erkannt - Erneuerung SP (Verunreinigungen an SP Manifold kommen wieder) - Probenaspiration gescheitert oder keine Oximetrie Werte (521 Inhomogene Probe - Quetschventil Überprüfung und Schmierung) - Funktionsüberprüfung etc. ohne Beanstandung | SWO-415591 | Adjustment | HW | Oxi module | Pinch valve |  | 14.09.2018; Sr. Rebecca Hünefeld?; Tel.: 0698505250; Sample aborts | Intervention needed Leakage detected Measurements have been canceled before | General check - Encrustations on SP removed- Renewal hose performed on squeezing valve - Service adjustment etc. - Analyzer rinsed - Renewal inlet and connection seal - All tests passed, nevertheless leakage detected - Renewal SP ( Impurities at SP Manifold come back) - Sample aspiration failed or no oximetry values (521 inhomogeneous sample - crushing valve inspection and lubrication) - functional check etc. without complaint
No | Visit | Complaint | RESP | Equipo no enciende y emite un pitido fuerte y continuo | Chequeo de módulo PC escuchando un pitido fuerte al iniciar | Se procede a sustitución de módulo pc completo | SWO-416382 | Defect | HW | DMS_PC | CPU |  | Equipment does not turn on and emits a loud and continuous beep | Checking PC module listening to a loud beep at startup | Proceeds to replace full pc module
No | Visit | Complaint | RAME | pCO2 sensitivity errors and Lac calibration errors. | Arrived on site and verified that analyzer was giving sensitivity errors for pCO2 causing LAC calibration errors. Verified that unit pCO2 sensitivity would not go above 20 during calibrations often only going as high as 16. | Replaced sensor cassette. Analyzer now operating properly. | SWO-415626 | Assisted Replacement | SC | SC | pCO2 | Sensitivity | pCO2 sensitivity errors and Lac calibration errors. | Arrived on site and verified that analyzer was giving sensitivity errors for pCO2 causing LAC calibration errors. Verified that unit pCO2 sensitivity would not go above 20 during calibrations often only going as high as 16. | Replaced sensor cassette. Analyzer now operating properly.
No | Visit | Complaint | RLTD | Analyser routinely fails to print and is only resolved with a reboot. | Software causing further issues which have increased since initial printer issue. | Loaded new flash card, with Windows loaded off site, issues remained. Fitted new CPU, loaded version 3.4, and customer setup and data logs. Analyser immediately shows vast improvement, operating as expected, with no freezing, no lag between selections and no buffering. Loaded regional windows settings. Customer to monitor | SWO-415922 | Defect | HW | DMS_PC | CPU |  | Analyser routinely fails to print and is only resolved with a reboot. | Software causing further issues which have increased since initial printer issue. | Loaded new flash card, with Windows loaded off site, issues remained. Fitted new CPU, loaded version 3.4, and customer setup and data logs. Analyser immediately shows vast improvement, operating as expected, with no freezing, no lag between selections and no buffering. Loaded regional windows settings. Customer to monitor
Yes | 0 | Y | RPAC | Blank | Pre Delivery service |  | D4F4730D40254E83 | [5 empty] | Blank | Pre Delivery service
No | Visit | Complaint | RCHN | 灵敏度错误 | 灵敏度错误，借945-737一个 |  | SWO-415031 | [5 empty] | Sensitivity error | Sensitivity error, borrow 945-737 one
No | Visit | Complaint | RCHN | 灵敏度错误 | 灵敏度错误，借945-737一个 |  | SWO-415030 | [5 empty] | Sensitivity error | Sensitivity error, borrow 945-737 one
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 试剂包泄漏 1342 | 更换试剂包后修复 | SWO-415770 | [5 empty] | Test card, reagent package failure | Reagent Pack Leak 1342 | Repair after replacement of reagent kit
No | Visit | Complaint | RAME | Cl issues | Seal ring missing from Lactate electrode. Leak from pH electrode. | Replace seal rings for Lactate and pH electrodes. | SWO-415551 | Defect | Electrode | Electrode | Sealing ring, pH |  | Cl issues | Seal ring missing from Lactate electrode. Leak from pH electrode. | Replace seal rings for Lactate and pH electrodes.
No | Visit | Complaint | RAME | PH and electrolytes drifting. | Found leak at pH electrode. | Replace seal ring for pH electrode. Remembrane reference electrode. | SWO-415618 | Defect | Electrode | Electrode | Sealing ring, pH |  | PH and electrolytes drifting. | Found leak at pH electrode. | Replace seal ring for pH electrode. Remembrane reference electrode.
No | Visit | Complaint | RIND | Parameter could not be measured | Checked the data logs found all sample query when we checked the result oximetry parameter values was not reported query in all the oxi parameter could not be measured performed protien removal complete oximetry test hemolyzer test was getting failed switch OFF the machine ON the machine initialization and 2 point calibrtion done sample processed then also same problem oxi paramter could not be measured found problem with oxi module | Replaced OXI Module | SWO-417471 | Defect | HW | Oxi module | Oxi module |  | Parameter could not be measured | Checked the data logs found all sample query when we checked the result oximetry parameter values was not reported query in all the oxi parameter could not be measured performed protien removal complete oximetry test hemolyzer test was getting failed switch OFF the machine ON the machine initialization and 2 point calibrtion done sample processed then also same problem oxi paramter could not be measured found problem with oxi module | Replaced OXYGEN Module
Yes | 10 | Y | RPAC | Blank | Replacement of AQT90 with another AQT90 |  | D4F223DF5463D02A | Re classified RMED | planned SWO |  |  | Techline: RPAC 819319389 | Blank | Replacement of AQT90 with another AQT90
No | Visit | Complaint | RFRA | Suite à la mise en service de l'analyseur IA signale qu'il est impossible de retirer les vis du capot de la phase avant | Colis non livre. Contact Ciblex pour voir où ils sont. En attente d'un retour De CIBLEX | Repasser changer le capot arrière. | SWO-416093 | Revisit required | HW | Cover and Chassis | Screws, Back cover | OOB | Following the commissioning of the AI analyzer, it is not possible to remove the screws from the hood of the front phase | Non-book parcels. Contact Targetx to see where they are. Waiting for a return from CIBLEX | Ironing change the rear hood.
Yes | Visit | Complaint | RAME | ANALYZER WILL NOT POWER UP. SUSPECTED CPU ERROR. | Write off CPU and power supply used for tshooting | Write off CPU and power supply used for tshooting | SWO-416307 | Re classified RMED | administration |  |  |  | ANALYZER WILL NOT POWER UP. SUSPECTED CPU ERROR. | Write off CPU and power supply used for tshooting | Write off CPU and power supply used for tshooting
No | Visit | Complaint | RFRA | Défaut Lactates Trois cassettes de capteurs changées d'affilé suite problèmes de calibrations lactates | Vérification complète | Attente accès à l'analyseur Remplacement de la carte sensor interface $ $ relance $ $ réinstallation du pack de solution $ $ mise en route $ $ calibrations $ $ améliorations mais toujours en défaut sur les Lactates $ $ attendre 2 3H et si pas de changement, remplacer cassette | SWO-416317 | Defect | HW | Basic Unit | PCB, Sensor interface | Lac | Lactates defect Three cassettes of sensors changed in a row following lactate calibration problems | Full check | Waiting for access to the analyzer Replacement of the sensor interface card $$ raise $$ reinstallation of the solution pack $$ start $$ calibrations $$ upgrades but still in default on the Lactates $$ wait 2 3H and if no change, replace cassette
No | Visit | Complaint | RLTD | Inlet stuck out Lucy 07736833628 | Inlet stuck out. Inlet cover missing | Arrived on site and found inlet closed down and analyser working. The inlet cover is missing and unable to find. Removed back cover and checked inlet motor assembly and checked screws are tight on clutch ok. Test operation in service mode working ok. Run dummy blood test and inlet opening and closing ok. | SWO-416261 | No fault found | No fault found | Auto Inlet module | Auto Inlet module | Inlet open | Inlet stuck out Lucy 07736833628 | Inlet stuck out. Inlet cover missing | Arrived on site and found inlet closed down and analyser working. The inlet cover is missing and unable to find. Removed back cover and checked inlet motor assembly and checked screws are tight on clutch ok. Test operation in service mode working ok. Run dummy blood test and inlet opening and closing ok.
No | Visit | Complaint | RFRA | CQ B et C non ok sur valeurs d’oxymetrie. (ThB / So2 / ...) CQ A ok. | Pbm de valeur d’oxymetrie sur CQ B et C revient ce matin. (Pbm sur Sat $ hbe et met légèrement hors gamme) | Changement hemolyseur ok. Passage calibration thB ok. Calibration ok, passage de tous les CQ ok. Tous les paramètres disponibles en vert ok. NB: maintenance effectuée sur SWO | SWO-416364 | Defect | HW | Oxi module | Hemolyzer | QC | CQ B and C not ok on oxymetrie values. (ThB / So2 / ...) CQ A ok. | Oxymetrie value pbm on CQ B and C returns this morning. (Pbm on Sat / hbe and slightly off-range) | Change hemolyseur ok. Passage calibration thB ok. Calibration ok, passage of all CQ ok. All settings available in green ok. NB: maintenance performed on SWO
No | Visit | Complaint | RESP | Equipamento não funciona ligado a bateria. | Verificação de não funcionamento do equipamento ligado a bateria | Subsituição da bateria em garantia e verificação de funcionamento OK | SWO-416221 | [5 empty] | Equipment does not work on battery. | Checking of non-functioning of the equipment connected to the battery | Battery replacement in warranty and operation check OK
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | PCO2定标失败 0378 | 更换测试卡正常。 | SWO-415883 | [5 empty] | Test card, reagent package failure | PCO2 calibration failure 0378 | The replacement test card is OK.
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | PCO2定标失败，PH灵敏度超范围 1316，0378 | 更换三张测试卡正常。 | SWO-415876 | [5 empty] | Test card, reagent package failure | PCO2 calibration failure, PH sensitivity over range 1316,0378 | Replacing the three test cards is OK.
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 液体传输失败，报错1249 | 更换试剂包后修复 | SWO-416465 | Assisted Replacement | SP | SP | FT | 1249 | Test card, reagent package failure | Liquid transmission failed, error 1249 | Repair after replacement of reagent kit
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | lAC定标失败 0210 | 更换测试卡正常。 | SWO-415878 | [5 empty] | Test card, reagent package failure | lAC calibration failure 0210 | The replacement test card is OK.
No | Visit | Complaint | RAME | The threads on. The gas 2 were bad and messed up the reg | Gas2 reg Was damaged due to bad threads on Gas 2 bottles | I replaced the gas 2 reg and checked the bottles. There was a part on the threads that was damaged | SWO-416395 | Defect | HW | Gas module | Gas bottle | Gas regulator damaged | The threads on. The gas 2 were bad and messed up the reg | Gas2 reg Was damaged due to bad threads on Gas 2 bottles | I replaced the gas 2 reg and checked the bottles. There was a part on the threads that was damaged
No | Visit | Complaint | RIND | Calibration aborted | Checked the data logs calibration was getting aborted found pH $BG air valve was defective | Replaced pH $BG air valve | SWO-417474 | Defect | HW | pH/BG Module | Valve, air |  | Calibration aborted | Checked the data logs calibration was getting aborted found pH /BG air valve was defective | Replaced pH $BG air valve
No | Visit | Complaint | RDEU | FlexQ mit gelegentlichen Funktionsstörungen | FlexQ mit gelegentlichen Funktionsstörungen | FlexQ gereinigt und initialisiert - Funktionsüberprüfung | SWO-415921 | Cleaned | HW | FlexQ | FlexQ |  | FlexQ with occasional malfunctions | FlexQ with occasional malfunctions | FlexQ cleaned and initialized - function check
No | Visit | Complaint | RDEU | KAL Driften Elektrolyte | KAL Driften Elektrolyte Elektroden bereits mehrfach membraniert | Allgemeine Überprüfung - Erneuerung Referenz Elektrode - Funktionsüberprüfung etc. | SWO-415920 | Assisted Replacement | Electrode | Electrode | Reference |  | KAL Drifting Electrolytes | KAL Drifting Electrolytes Electrodes already membraneed several times | General verification - Renewal reference electrode - function check etc.
No | Visit | Complaint | RDEU | 17.09.2018; Hr. Knoll; Tel.: 091318542349; Crea Kal Fehler | Fehler: Crea-Kal.Fehler, Empf. außer Bereich (Empf. 1 = 1,5 pA $uM; Empf. 3 = 587%) Fehlersuche --> Dateiarchive (Kal., QC, Aktivitäten, Betriebsmittel), Crea-Elektroden, Kal.-Lösungen, Ventile im Service, Messkammern, Verstärkerkappen überprüft --> E8088 Crea A (YN02) und E8089 Crea B (YN02) defekt --> wurden vom Kunden schon mehrmals membraniert --> Berechnung der Elektroden gemäß Auftragsnummer 30262899 | E8088 (BR02), E8089 (BS01) eingebaut und E8088 (YN02), E8089 (YN02) ausgebaut, mehrere Kal., QC, System und Funktion i.O. | SWO-415826 | Assisted Replacement | Electrode | Electrode | Crea |  | 17.09.2018; Mr Knoll; Tel.: 091318542349; Crea Kal Error | Error: Crea-Kal.Error, Empf. out range (1 = 1.5 pA /uM; Empf. 3 = 587%) Troubleshooting --> File archives (Cal., QC, activities, equipment), Crea electrodes, cal. solutions, valves in service, measuring chambers, amplifier caps checked --> E8088 Crea A (YN02) and E8089 Crea B (YN02) defective --> have been membraned several times by the customer --> Calculation of electrodes according to order number 30262899 | E8088 (BR02), E8089 (BS01) built-in and E8088 (YN02), E8089 (YN02) removed, several cal., QC, system and function i.O.
No | Visit | Complaint | RDEU | AU nicht mit DMS verbunden / Temp. fehler | Au nicht mit DMS verbunden . Temp. fehler | Lüfter ist sauber. Neustart bringt keine Besserung. Softwareinstallation bleibt bei Au hängen. Sichrungen vom AU Board herausgenommen und durchgemessen . Deren Sockel nachgebogen. Neustart. Startphase. Kal und QC in Ordnung. | SWO-416193 | Adjustment | HW | Basic Unit | PCB, analyzer control |  | AU not connected to DMS / Temp. Error | Au not connected to DMS . Temp. Error | Fan is clean. Restart brings no improvement. Software installation hangs at Au. Sightings removed from the AU board and measured through . Their base bent. Restart. Start. Kal and QC all right.
No | Visit | Complaint | RKOR | Timing belt is broken | Calibration, Quality Control, Activity logs check Flow path check Consumables status check QC plots check Tube holder check All position safe check Pipette is failed timing belt is broken Replace the pipette module | Replace the pipette module | SWO-420051 | Defect | HW | Aspiration Module | Timing belt |  | Timing belt is broken | Calibration, Quality Control, Activity logs check Flow path check Consumables status check QC plots check Tube holder check All position safe check Pipette is failed timing belt is broken Replace the pipette module | Replace the pipette module
No | Visit | Complaint | RDEU | 13.09.2018, Herr Kamp, Tel: 06946086123: Messung abgebrochen bei einer Probe | Gerät in Ramen der Wartung überprüft, keine Fehler festgestellt. Alle Sensoren funktionieren. | Keine Fehler feststellbar. Seit dem Abbruch sind ca 20 Proben ohne Abbruch gelaufen. | SWO-416574 | No fault found | No fault found | No fault found | No fault found | 1067 | 13.09.2018, Mr. Kamp, Tel: 06946086123: Measurement aborted on a sample | Device checked in ramen maintenance, no errors detected. All sensors work. | No errors detectable. Since the demolition, about 20 samples have been run without demolition.
No | Visit | Complaint | RIND | Instrument showing error and calibration aborted. | Found process unit not moving properly. adjusted the process unit screws. | Adjusted the base and checked the all mechanical movements.found all movements found fine. Did calibration and samples all results within the range instrument working fine. | SWO-418447 | Adjustment | HW | Process Module | Process Unit | 5257 | Instrument showing error and calibration aborted. | Found process unit not moving properly. adjusted the process unit screws. | Adjusted the base and checked the all mechanical movements.found all movements found fine. Did calibration and samples all results within the range instrument working fine.
No | Visit | Complaint | RIND | calibration failed with error codes temperature error before drying and back ground outliers and cal adjust outliers failure. | checked the analyzer and cal adjust results conformed optical unit and cup dryer failure. and inlet door also not working properly. | we need to replace optical unit and cup dryer and inlet door for this case. | SWO-416523 | Defect | HW | Process Module | PCB, cup dryer |  | calibration failed with error codes temperature error before drying and back ground outliers and cal adjust outliers failure. | checked the analyzer and cal adjust results conformed optical unit and cup dryer failure. and inlet door also not working properly. | we need to replace optical unit and cup dryer and inlet door for this case.
No | Visit | Complaint | RDEU | 17.09.2018, Frau Rauch(Labor), Tel: 09471705456: Einlass schließt nicht mehr | Inlet Door schließt nicht richtig | Führung für Inlet Door erneuert Inlet Door justiert Systemreinigung und Systemüberprüfung LQC gemessen | SWO-416072 | Defect | HW | Cover and Chassis | Cartridge door | 5784 | 17.09.2018, Mrs Rauch (Laboratory), Tel: 09471705456: Entrance no longer closes | Inlet Door does not close properly | Guided tour for Inlet Door renewed Inlet Door adjusted System cleaning and system verification LQC measured
No | Visit | Complaint | RPOL | Przeciek | Sprawdzenie uszczelki wlotu. Testy serwisowe. | Wymiana uszczelki wlotu na nową. | SWO-416241 | Defect | HW | Inlet module | Inlet gasket |  | Leak | Check Inlet gasket. Service tests. | Replacing the inlet gasket with a new one.
Yes | Visit | Complaint | RPOL | Błąd odczytu pakietu odczynnikowego | Uszkodzenie płyty reagent PCB 903-036 | Wymiana płyty reagent PCB | SWO-415943 | Defect | Consumables | Solution Pack | Chip | misalignment | Defuncting Packet Read error | Uszkodzenie płyty reagent PCB 903-036 | Wymiana płyty reagent PCB
Yes | Visit | Complaint | RKOR | open / close module error - When the module is closed, it closed to the middle and reopen | Removal and reassemble the open $ close module | Removal and reassemble the open $ close module | SWO-420059 | [5 empty] | open / close module error - When the module is closed, it closed to the middle and reopen | Removal and reassemble the open / close module | Removal and reassemble the open $ close module
No | Visit | Complaint | RAME | Customer states that the analyzer keeps on giving FUTURE results when testing patients in real time. | Verified customer's stated issue. | Adjusted the analyzers Windows time, checked off the proper time zone and adjusted it for daylight saving time. | SWO-417047 | Setup | SW | External SW | Time setting |  | Customer states that the analyzer keeps on giving FUTURE results when testing patients in real time. | Verified customer's stated issue. | Adjusted the analyzers Windows time, checked off the proper time zone and adjusted it for daylight saving time.
Yes | Visit | Complaint | RDEU | Innerer Gehäuse-Bruch an Wetsection Befestigungen | Innerer Gehäuse-Bruch an Wetsection Befestigungen Feststellung bei Jahreswartung | Spezialkleber aus Medizintechnik entnommen - Verschraubungen gehäuseseitig verklebt und Wetsection-Block montiert und verschraubt ! Eventuelle Bestellung Gehäuse (Klärung erfolgt) ! Bemerkung: Keine Sturzspuren sichtbar und keine Beschädigungen bei Anwendern bekannt! | SWO-416782 | Re classified RMED | planned SWO |  | yearly maintenance |  | Inner housing break to Wetsection fasteners | Inner housing break to Wetsection fasteners Determination for annual maintenance | Special adhesive taken from medical technology - screw connections glued on the case side and wetsection block mounted and screwed ! Possible order housing (clarification done) ! Note: No traces of fall visible and no damage known to users!
No | Visit | Complaint | RAME | Customer called stating that analyzer not connecting to Aqure since moving it. | Found that since the customer temporarily moved it to another area of the hospital and back to original area have not been able to connect to Aqure. | I worked with their It department to get the analyzer connected to Aqure it is plugged into the right port has the right IP address for Aqure and is pulling the right IP address according to their IT department. The IT department thinks it something on their end. They are looking into it. | SWO-418841 | No fault found | SW | External SW | Network | Aqure | Customer called stating that analyzer not connecting to Aqure since moving it. | Found that since the customer temporarily moved it to another area of the hospital and back to original area have not been able to connect to Aqure. | I worked with their It department to get the analyzer connected to Aqure it is plugged into the right port has the right IP address for Aqure and is pulling the right IP address according to their IT department. The IT department thinks it something on their end. They are looking into it.
No | Visit | Complaint | RAME | Analyzer showing co-ox board failure | Analyzer showing coox board failure. | Removing analyzer to return to Westlake. | SWO-417087 | [5 empty] | Analyzer showing co-ox board failure | Analyzer showing coox board failure. | Removing analyzer to return to Westlake.
No | Visit | Complaint | RAME | Analyzer showing coox board failure. | Analyzer showing coox board failure | Removed analyzer to return to Westlake. | SWO-417089 | [5 empty] | Analyzer showing coox board failure. | Analyzer showing coox board failure | Removed analyzer to return to Westlake.
No | Visit | Complaint | RCHN | 开机不启动 | 开机不启动，借910-342一个 | 更换910-342后修复(By Ran) | SWO-415112 | Defect | HW | DMS_PC | CPU |  | Power on does not start | Power on does not start, borrow 910-342 one | Repair after replacement 910-342 (By Ran)
No | Visit | Complaint | RCHN | 光学强度超范围 | 光学强度超范围，借903-364一个 | 更换903-364后修复(By Ran) | SWO-415179 | Defect | HW | Process Module | Optical Unit | 1031 | Optical strength over range | Optical strength over range, borrowed 903-364 a | Repair after replacement 903-364 (By Ran)
No | Visit | Complaint | RDEU | 28.06.2018 16:45Uhr Sr.Jenny/Pfl. Mario 030940113662: Neonlampenfehler 04.07.2018 10:00 Uhr Pfl. Heiko teilt mit, dass der genannte Fehler im letzten Nachtdienst wieder aufgetreten ist | bestelltes OxiModul wg. Spektrometerfehler ersetzen | OxiModul (alt: 902-833R2155N010) ersetzt, Software ok, kompletter OxiTest ok, anschließend Halbjahreswartung | SWO-416777 | Defect | HW | Oxi module | Oxi module | neon lamp; mail: tried to change leon lamp? | 28.06.2018 16:45h Sr.Jenny/Pfl. Mario 030940113662: Neon lamp error 04.07.2018 10:00 am Pfl. Heiko announces that this error occurred again in the last night service | ordered OxiModul wg. Replace spectrometer errors | OxiModule (old: 902-833R2155N010) replaced, software ok, complete OxiTest ok, then half-year maintenance
No | Visit | Complaint | RAME | OXI values not showing | No OXI values. Intermittent but low thb | I checked the pinch valve and tubing. I ran many samples with out an issue | SWO-417800 | No fault found | HW | Oxi module | Pinch valve |  | OXI values not showing | No OXI values. Intermittent but low thb | I checked the pinch valve and tubing. I ran many samples with out an issue
No | Visit | Complaint | RAME | Intermittent, frequent error 791, no Oximetry results, parameter could not be measured. | Reoccurrence of error 791 (parameter could not be measured due to insufficient or inhomogeneous sample) reported previous week. Reviewed data logs, performed service diagnostics. Problem reoccurred 9 $18 @ approx 4:30AM. Error 791 reported intermittently but frequently on all levels of QC and sporadic patient sample. Performed complete Oximetry test, no failures, valve test passed. Suspect intermittent failure within Oxi module. | Replaced current Oxi module with used module obtained at home office temporarily until a new module is available for replacement. Performed all applicable adjustment, complete Oxi Test, pass, Service and User tHb cal. Will return ASAP with a new Oxi module. | SWO-417058 | Revisit required | HW | Oxi module | Oxi module | 791 | Intermittent, frequent error 791, no Oximetry results, parameter could not be measured. | Reoccurrence of error 791 (parameter could not be measured due to insufficient or inhomogeneous sample) reported previous week. Reviewed data logs, performed service diagnostics. Problem reoccurred 9 /18 @ approx 4:30AM. Error 791 reported intermittently but frequently on all levels of QC and sporadic patient sample. Performed complete Oximetry test, no failures, valve test passed. Suspect intermittent failure within Oxi module. | Replaced current Oxi module with used module obtained at home office temporarily until a new module is available for replacement. Performed all applicable adjustment, complete Oxi Test, pass, Service and User tHb cal. Will return ASAP with a new Oxi module.
Yes | Visit | Complaint | RAME | Glucose frequently failing during 2PT cal due to drift, QC L2 above range. | While investigating cause of intermittent Oxi error 791, LAC & GLUC began to exhibit intermittent increase in calibration drift similar to the issue described in SWO-409452. Observed a bubble in the GLUC electrode measuring well during the final rinse cycle, possible cause. Also suspect possible electronic issue with EL $MET module PCB. | Replace GLUC membrane due to problematic membrane in place not allowing fluid (bubbles) to flush cleanly from measuring chamber. Bubbles were observable thru the view window for the GLUC electrode measuring chamber well. Replace CA membrane due to frequent manipulation during service intervention. Replaced EL $MET module PCB. Performed zero current adjustment. Performed all checks and tests to verify proper operation. No further drift error (408) observed at this time. | SWO-417059 | Re classified RMED | same day upSWO |  |  | ; mail: if problem occurs again: try cleaning of fluidic pathway and changing valves in El/Met module | Glucose frequently failing during 2PT cal due to drift, QC L2 above range. | While investigating cause of intermittent Oxi error 791, LAC & GLUC began to exhibit intermittent increase in calibration drift similar to the issue described in SWO-409452. Observed a bubble in the GLUC electrode measuring well during the final rinse cycle, possible cause. Also suspect possible electronic issue with EL /MET module PCB. | Replace GLUC membrane due to problematic membrane in place not allowing fluid (bubbles) to flush cleanly from measuring chamber. Bubbles were observable thru the view window for the GLUC electrode measuring chamber well. Replace CA membrane due to frequent manipulation during service intervention. Replaced EL $MET module PCB. Performed zero current adjustment. Performed all checks and tests to verify proper operation. No further drift error (408) observed at this time.
Yes | Visit | Complaint | RAME | Crea QC Lv 1 out after changing range per FAN 915-387 | Getting service dump and data logs for CQC case. | Downloaded logs and service dump. | SWO-417413 | Re classified RMED | planned SWO |  |  |  | Crea QC Lv 1 out after changing range per FAN 915-387 | Getting service dump and data logs for CQC case. | Downloaded logs and service dump.
No | Visit | Complaint | RIND | Lactate electorde is faulty as we changed the new membarne and cross checked with other machine also. | Found the lactate electrode is faulty,cross checked with other machine also same error ocuurs,checked with new membranes also same error occurs. | Want to change the Lactate Electrode. | SWO-416932 | Revisit required | Electrode | Electrode | Lac |  | Lactate electorde is faulty as we changed the new membarne and cross checked with other machine also. | Found the lactate electrode is faulty,cross checked with other machine also same error ocuurs,checked with new membranes also same error occurs. | Want to change the Lactate Electrode.
No | Visit | Complaint | RKK | AQT90にて、画面がちらつくという症状が発生。ディスプレイ交換しても改善せず。ケーブル交換で改善。 | AQT90にて、画面がちらつくという症状が発生。ディスプレイ交換しても改善せず。 | ケーブル交換実施。 | SWO-419251 | Defect | HW | Analyzer Control Module | Cable, display |  | AQT90, the symptom that the screen flickers occurs. It does not improve even if the display is replaced. Improved with cable replacement. | AQT90, the symptom that the screen flickers occurs. It does not improve even if the display is replaced. | Cable replacement.
Yes | Visit | Complaint | RFRA | Mise en place appareil de prêt. Thierry Guillou valide ok. | Mise en Appareil de prêt ok à la place de AQT90 Uzès 838R0325N010. | Re installation config ok. Calibration tnt , Ddimmer et CRP ok. Passage de CQ tnt niveau 1 et 2 ok MCC1 : 30,1 $31,1 $29,8 (19,4-38,4) MCC2 : 154 $145 $141 (120-184) Lancement CQ 1 et 2 Ddimmer et CRP ok. Client lancera qualif demain. | SWO-416938 | Re classified RMED | installation report |  |  |  | Setting up loan device. Thierry Guillou validates ok. | Putting in loan device ok in place of AQT90 Uzès 838R0325N010. | Re installation config ok. Tnt calibration, Ddimmer and CRP ok. Passage of CQ tnt level 1 and 2 ok MCC1: $30.1 $31.1 $29.8 (19.4-38.4) MCC2: $154 $145 $141 (120-184) Launch CQ 1 and 2 Ddimmer and CRP ok. Customer will launch Qualifying tomorrow.
No | Visit | Complaint | RPOL | Problemu przepływowe. Codzienne płukanie analizatora. | Oględziny uszczelki wlotu. | Wymiana uszczelki na nową. | SWO-416589 | Defect | HW | Inlet module | Inlet gasket |  | Flow problem. Daily rinsing of the analyser. | Inspection of Inlet gasket. | Replacing the gasket with a new one.
No | Visit | Complaint | RPOL | Uszkodzona drukarka | Przerywany wydruk | Wymiana drukarki | SWO-417692 | Defect | HW | Printer Module | Printer Module |  | Defective printer | Intermittent print | Replacing the printer
No | Visit | Complaint | RAME | Customer called that they did not get their WDC report. | Found that the data was not uploaded into qa portal | I downloaded and uploaded all the data from their 14 analyzers to the qa portal. | SWO-418835 | Adjustment | SW | RMED SW | QA Portal |  | Customer called that they did not get their WDC report. | Found that the data was not uploaded into qa portal | I downloaded and uploaded all the data from their 14 analyzers to the qa portal.
No | 4 | Y | RCZE | Blank |  |  | D4F639A5B636AC12 | Setup | SW | His/Lis | His/Lis |  | Blank
No | Visit | Complaint | RCHN | pco2质控不通过 | pco2质控不通过，借903-334一个 | 更换903-334备件后修复(By Ran) | SWO-417138 | [5 empty] | pco2 QC does not pass | pco2 QC does not pass, borrow903-334 one | Repair after replacement of 903-334 spare parts (By Ran)
Yes | Visit | Complaint | RIND | lac failed in SC | Checked SC and found there was error for hydration in SC | Cleaned SC with distilled water and kept in fridge for 30 minutes. Then reinstalled SC and given calibration but still error is there. Need to replaced SC. | SWO-418275 | [5 empty] | lac failed in SC | Checked SC and found there was error for hydration in SC | Cleaned SC with distilled water and kept in fridge for 30 minutes. Then reinstalled SC and given calibration but still error is there. Need to replaced SC.
No | Visit | Complaint | RIND | Na Level 3 QC was getting failed | Checked the data log found Na Level 3 QC was getting failed found problem wth SC | Replaced new SC performed QC values wiithin the range | SWO-417496 | [5 empty] | Na Level 3 QC was getting failed | Checked the data log found Na Level 3 QC was getting failed found problem wth SC | Replaced new SC performed QC values wiithin the range
No | Visit | Complaint | RAME | Sensor cassette will not initialize. Fails initialization at 60% for Glucose. | Sensor cassette will not initialize. Fails initialization at 60% for Glucose. Error code 1218. Customer has tried three times to initialize. | Hand-carried replacement sensor cassette for delivery. Set Zero on analyzer and checked pinch valve operation. Installed new sensor cassette for customer. Sensor cassette initialized and passed all calibration and Quality Checks. Normal operations restored. | SWO-417716 | [5 empty] | Sensor cassette will not initialize. Fails initialization at 60% for Glucose. | Sensor cassette will not initialize. Fails initialization at 60% for Glucose. Error code 1218. Customer has tried three times to initialize. | Hand-carried replacement sensor cassette for delivery. Set Zero on analyzer and checked pinch valve operation. Installed new sensor cassette for customer. Sensor cassette initialized and passed all calibration and Quality Checks. Normal operations restored.
No | Visit | Complaint | RCHN | 电极故障 | 电极故障 | 更换945-737后修复 | SWO-417879 | [5 empty] | Electrode failure | Electrode failure | Repair after replacement of 945-737
No | Visit | Complaint | RCHN | 屏幕故障 | 屏幕故障，借910-298,910-343,636-571各一个 | 更换910-343,636-571备件后修复(By Ran) | SWO-415792 | Defect | HW | Analyzer Control Module | CPU |  | Screen failure | Screen failure, borrow 910-298, 910-343, 636-571 each | Repair after replacement of 910-343,636-571 spare parts (By Ran)
No | Visit | Complaint | RAME | Our Clinical Specialist Jon Curry is concerned about the precision study performed, where he is finding qc level 2 failing high outside the required range. | During the precision studies it was discovered that the COHB for qc level 2 was running high. Visit to inspect the oximetry system. | Checked and, performed the service Complete Oximetry test, passing within specifications. Confirmed the lamp value to be at 3380 and, it is within specifications. Performed the service ctHb calibration. Found pCO2 not passing calibrations due to low sensitivity of 84%. Replaced a new pCO2 membrane. | SWO-417782 | Adjustment | HW | Oxi module | tHb Cal | mail: was COHB outside RMED spec or customer spec? | Our Clinical Specialist Jon Curry is concerned about the precision study performed, where he is finding qc level 2 failing high outside the required range. | During the precision studies it was discovered that the COHB for qc level 2 was running high. Visit to inspect the oximetry system. | Checked and, performed the service Complete Oximetry test, passing within specifications. Confirmed the lamp value to be at 3380 and, it is within specifications. Performed the service ctHb calibration. Found pCO2 not passing calibrations due to low sensitivity of 84%. Replaced a new pCO2 membrane.
No | Visit | Complaint | RAME | Intermittent, frequent error 791, no Oximetry results, parameter could not be measured. | Follow up to SWO-417058. Suspect intermittent internal Oxi module failure. | Replaced the Oxi module, completed all adjustments and testing. All calibration and QC within specifications, no errors reported or observed. | SWO-417151 | Defect | HW | Oxi module | Oxi module | 791 | Intermittent, frequent error 791, no Oximetry results, parameter could not be measured. | Follow up to SWO-417058. Suspect intermittent internal Oxi module failure. | Replaced the Oxi module, completed all adjustments and testing. All calibration and QC within specifications, no errors reported or observed.
No | Visit | Complaint | RIND | Leak current detected. | Check the system logs & calibration logs. | Check the rinse valve & liquid sensor. inlet & el $met liquid sensors are faulty. check the rinse valve there is leakage in valve causing error. | SWO-417872 | Defect | HW | El/Met Module | Valve |  | Leak current detected. | Check the system logs & calibration logs. | Check the rinse valve & liquid sensor. inlet & el $met liquid sensors are faulty. check the rinse valve there is leakage in valve causing error.
No | Visit | Complaint | RDEU | 19.09.2018, 08:38 Uhr, Fr. Gerg, Tel. 080417927616: Keine CRP Messung möglich | Fehlersuche vor Ort | Incubation Wheel ausgebaut. Sensoren gereinigt und Motoren gefettet. Systemreinigung durchgeführt ->i.O. QC D-dimer Level 1 und QC Multicheck Level 2 durchgeführt. -> i.O. QCs wurden nicht abgebrochen und ohne Fehlermeldungen durchgeführt. AQT wieder messbereit. | SWO-417337 | Defect | Consumables | Solution Pack | Chip, Solution Pack | 1052 | 19.09.2018, 08:38, Ms. Gerg, tel. 080417927616: No CRP measurement possible | Troubleshooting on site | Incubation Wheel. sensors cleaned and motors greased. System cleaning performed ->i.O. QC D-dimer Level 1 and QC Multicheck Level 2. -> i.O. QCs were not canceled and performed without error messages. AQT ready to measure again.
No | Visit | Complaint | RDEU | 19.09.2018 Herr Magira 059311524482 meldet, das der QC Wert für Trop. seit der Installation ansteigt. | MCC3 TnT außerhalb der Grenzen | Gerät überprüft, Kontrolle wiederholt, weiterhin grenzwertig zu hoch, MCC3 wird ersetzt | SWO-417865 | Defect | Consumables | QC | QC, TnT |  | 19.09.2018 Mr Magira 059311524482 reports that the QC value for Trop. has increased since installation. | MCC3 TnT out of bounds | Device checked, control repeated, still borderline too high, MCC3 is replaced
No | Visit | Complaint | RDEU | 26.06.2018 Herr Betsch 015204767315 meldet, dass die Anbindung nicht funktioniert. 24.08.2018; Hr. Markus Mann; Tel.: 06102359713; es besteht immer noch keine Aqure Verbindung 24.08.2018; Hr. Betsch; Tel.: 015204767315;IP Adressen stimmt nicht | Absprache mit Herrn Bett. Die Netzwerkverbindung ist weiterhin gestört und nicht möglich | Das Gerät mit den gleichen Konfiguration mitgenommen und zum KFH Dialyse gebracht und getestet. Die Verbindung war sofort da und die Synchronisation der Daten haben sofort angefangen zu laufen. Netzwerkkabel auch entsprechend getestet. Es liegt definitiv am Netzwerk und nicht am Gerät. | SWO-418491 | No fault found | SW | External SW | Wifi |  | 26.06.2018 Mr. Betsch 015204767315 reports that the connection is not working. 24.08.2018; Mr. Markus Mann; Tel.: 06102359713; there is still no Aqure connection 24.08.2018; Mr. Betsch; Tel.: 015204767315;IP addresses are not correct | Consultation with Mr. Bett. The network connection is still disrupted and not possible | The device with the same configuration was taken with you and taken to KFH dialysis and tested. The connection was there immediately and the synchronization of the data started to run immediately. network cable also tested accordingly. It's definitely down to the network, not the device.
No | Visit | Complaint | RDEU | 26.06.2018 Herr Betsch 015204767315 meldet, dass die Anbindung nicht funktioniert. 24.08.2018; Hr. Markus Mann; Tel.: 06102359713; es besteht immer noch keine Aqure Verbindung 24.08.2018; Hr. Betsch; Tel.: 015204767315;IP Adressen stimmt nicht | Die Netzwerkverbindung ist weiterhin gestört und nicht möglich | Das Gerät mit dem gleichen Konfiguration mitgenommen, zum KFH Dialyse gebracht und getestet. Die Verbindung war sofort da und die Synchronisation der Daten haben sofort angefangen zu laufen. Netzwerkkabel auch entsprechend getestet. Es liegt definitiv am Netzwerk und nicht am Gerät. | SWO-420827 | Defect | SW | External SW | Network |  | 26.06.2018 Mr. Betsch 015204767315 reports that the connection is not working. 24.08.2018; Mr. Markus Mann; Tel.: 06102359713; there is still no Aqure connection 24.08.2018; Mr. Betsch; Tel.: 015204767315;IP addresses are not correct | The network connection is still disrupted and not possible | The device with the same configuration, brought to KFH dialysis and tested. The connection was there immediately and the synchronization of the data started to run immediately. network cable also tested accordingly. It's definitely down to the network, not the device.
No | Visit | Complaint | RDEU | Probleme mit Kal und QC | Fehlersuche vor Ort | Frau Wolfinger gebeten nochmals den LP zu tauschen Beineintreffen war der ABL auf grünem Ampelstatus und messbereit. Trotzdem den ABL nochmals überprüft. Kalibrierung durchgeführt -> i.O. QCs für alle Level durchgeführt ->i.O. Patientenmessung durchführen lassen. Messung ohne Fehler. ABL wieder messbereit. | SWO-418122 | No fault found | No fault found | No fault found | No fault found |  | Probleme mit Kal und QC | Troubleshooting on site | Ms. Wolfinger asked to swap the LP again The ABL was on green traffic light status and ready to be measured. Nevertheless, the ABL has been re-examined. Calibration performed -> i.O. QCs performed for all levels ->i.O. patient measurement. Measurement without errors. ABL ready to be measured again.
No | Visit | Complaint | RSAF | PCO2sensitivity error on 2x SC | Faulty PCO2 sensor in cassette | Replaced faulty SC, Start up tests ok. | SWO-418355 | Assisted Replacement | SC | SC | pCO2 | Sensitivity | PCO2sensitivity error on 2x SC | Faulty PCO2 sensor in cassette | Replaced faulty SC, Start up tests ok.
Yes | Visit | Complaint | RAME | Customer getting oxi errors on patient sample error codes are 292, 581, 688 | To get photo of print outs for qc | I had to get photos of qc print outs for SWO-360494 because they did not attach the first time back in may. Also talked with the customer on how the analyzers are doing and showed new person they hire the basics and how to run qc. | SWO-418843 | Re classified RMED | planned SWO |  |  |  | Customer getting oxi errors on patient sample error codes are 292, 581, 688 | To get photo of print outs for qc | I had to get photos of qc print outs for SWO-360494 because they did not attach the first time back in may. Also talked with the customer on how the analyzers are doing and showed new person they hire the basics and how to run qc.
Yes | 2 | Y | RPAC | Blank | Training |  | D4F713782C2A8BA7 | Re classified RMED | planned SWO |  |  |  | Blank | Training
No | Visit | Complaint | RLTD | Adam 02082164934 | Printer cases damaged broken unit has been dropped. | Printer case cracked and not closing properly and not printing. Replaced printer unit and test ok. Printer not covered under warranty customer damage. Cost for printer 910-286 £475 + vat. Please supply perches order number . Checked analyser for other damage none found. | SWO-418127 | User Error | HW | Printer Module | Printer Module | Analyzer dropped | Adam 02082164934 | Printer cases damaged broken unit has been dropped. | Printer case cracked and not closing properly and not printing. Replaced printer unit and test ok. Printer not covered under warranty customer damage. Cost for printer 910-286 £475 + vat. Please supply perches order number . Checked analyser for other damage none found.
No | Visit | Complaint | RAME | Intermittent OXI Spectrum Mismatch Error | Troubleshot to a faulty hemolyzer, and pinch valve. | Replaced the hemolyzer and pinch valve. | SWO-418326 | [5 empty] | Intermittent OXI Spectrum Mismatch Error | Troubleshot to a faulty hemolyzer, and pinch valve. | Replaced the hemolyzer and pinch valve.
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | O2偏高 病人结果与临床不符 | 更换测试卡后修复 | SWO-420754 | [5 empty] | Test card, reagent package failure | O2 high Patient results are not clinically consistent | Repair after replacing test card
No | Visit | Complaint | RCHN | PO2病人结果偏高，接口板腐蚀 | O2偏高 | 更换903-334后修复by tim | SWO-420758 | [5 empty] | PO2 patient results are high, interface plate corrosion | O2 high | Repair by tim after replacing 903-334
No | Visit | Complaint | RCHN | 电极故障 | 电极故障 |  | SWO-417881 | [5 empty] | Electrode failure | Electrode failure
No | Visit | Complaint | RCHN | Na定标不过 | Na定标不过，借903-581一个 | 更换903-581后修复(by chuanwei) | SWO-417196 | Defect | HW | Basic Unit | PCB, Sensor interface | Na | Na calibration is no more than | Na calibration, however, borrow 903-581 a | Repair after replacement 903-581 (by chuanwei)
No | Visit | Complaint | RCHN | 数据库损坏 | 数据库损坏，借944-216（14696），944-214（14583），944-212（14732）各一个 | 重新定标后修复(by chuanwei) | SWO-416637 | Defect | SW | RMED SW | RMED SW | Database corruption | Database corruption | Database corruption, borrowing 944-216 (14696), 944-214 (14583), 944-212 (14732) each | Re-calibration repair (bywei chuan)
No | Visit | Complaint | RCHN | 进样门卡住，报错5784 | 进样门卡住，报错5784，需要借件903-028 | 更换903-028后修复(by chuanwei) | SWO-417147 | Defect | HW | Cover and Chassis | PCB, cabinet top | 5784 | The entrance door is stuck, and the error 5784 is reported. | The sample door is stuck, the error 5784, need to borrow 903-028 | Repair after replacement 903-028 (by chuanwei)
No | Visit | Complaint | RIND | Service Visit. | SP and SC is invalid due to change in date and time shows in display. | Check the date and time of the analyzer, it is showing back date, Reset the solution pack and remove the SC, enter the present date and time in the analyzer, install the SP and SC successfully done, analyzer handed over to staff in well working condition. | SWO-418346 | [5 empty] | Service Visit. | SP and SC is invalid due to change in date and time shows in display. | Check the date and time of the analyzer, it is showing back date, Reset the solution pack and remove the SC, enter the present date and time in the analyzer, install the SP and SC successfully done, analyzer handed over to staff in well working condition.
No | Visit | Complaint | RFRA | Capteur 54-72171, 54-72167 et 54-72168 défecteux. Affiche erreur 0500010 sur les TCM5. | Discussion avec Mme Aimiard et Mr Duguet par rapport à ces capteurs défectueux. | Remplacement des capteurs sous garantie. | SWO-420442 | [5 empty] | Sensor 54-72171, 54-72167 and 54-72168 defective. Error 0500010 display on TCM5. | Discussion with Ms. Aimiard and Mr. Duguet about these defective sensors. | Replacement of sensors under warranty.
No | Visit | Complaint | RCHN | 显示屏故障导致机器无法开机 | 显示屏故障导致机器无法开机，借910-320一个 | 更换910-320后修复(by chuanwei) | SWO-416685 | Defect | HW | DMS_PC | display |  | A faulty display causes the machine to fail to power on | Display failure causes the machine to fail to power on, borrow 910-320 one | Repair after replacement 910-320 (by chuanwei)
No | Visit | Complaint | RKK | pCO2ﾒﾝﾌﾞﾗﾝ交換後から感度ｴﾗｰとなり電極故障。保証期間の為、新品交換。 | 感度32% | 電極交換。 | SWO-419249 | Assisted Replacement | Electrode | Electrode | pCO2 | Sensitivity | It becomes a sensitivity error after the pCO2 membrane exchange and the electrode failure. Because of the warranty period, i.e. new replacement. | Sensitivity 32% | Electrode replacement.
No | Visit | Complaint | RDEU | 23.08.2018 09:45Uhr Herr Glienk 03334691503: Das Gerät läuft nicht zuverlässig | PumpenKalFehler (923), diverse Drift- & QK-Fehler bei versch. Parametern, sporadisch Einlaß undicht | Reinigungsarbeiten, Fehlersuche, einmal komplett Halbjahreswartungsmaterial ersetzt, alle Messkammern und Elektroden kontrolliert, Startphase, Kal und QK ok | SWO-418320 | Defect | HW | Inlet module | Inlet gasket | Half year maintenance | 23.08.2018 09:45Pm Mr Glienk 03334691503: The device is not running reliably | PumpsKalError (923), various drift & QK errors in case of Parameters, sporadic inlet leaks | Cleaning work, troubleshooting, once completely half-year maintenance material replaced, all measuring chambers and electrodes checked, start phase, Kal and QK ok
No | Visit | Complaint | RAME | Error 537, Oxi liquid sensor error, Error 532, EL/MET lower liquid sensor error. | Suspect issues with either liquid sensor or fluid flow and air purging from blue vent port. | Cleaned blue vent port for valve 8. Replaced solution pump tubing. Reran liquid Sensor and pump calibration. No Further errors observed. Investigated all tubing. Found the rinse tube connected to the rinse connector $nipple had become loose, most likely not allowing for a consistent intake of rinse solution without air bubbles or micros bubbles, probable cause of error 537 & 532. I have completed annual maintenance as noted on SWO-417787. | SWO-417788 | Adjustment | HW | Fluidic manifold | Rinse Tube | 532, 537 | Error 537, Oxi liquid sensor error, Error 532, EL/MET lower liquid sensor error. | Suspect issues with either liquid sensor or fluid flow and air purging from blue vent port. | Cleaned blue vent port for valve 8. Replaced solution pump tubing. Reran liquid Sensor and pump calibration. No Further errors observed. Investigated all tubing. Found the rinse tube connected to the rinse connector $nipple had become loose, most likely not allowing for a consistent intake of rinse solution without air bubbles or micros bubbles, probable cause of error 537 & 532. I have completed annual maintenance as noted on SWO-417787.
No | Visit | Complaint | RFRA | Perte de cal | Perte de cal tni. Vérification des messages ok. Suite au pbm n'a plus de K7 de cal de lot14732 Tni. | Update en version 8.11 pour régler pbm de perte de cal et ralentissement. Mise en place sauvegarde automatique tous les jours à 3h sur clé usb OK. Pas de CQ pour valider. Passage d’un tube test en tni ok. | SWO-418293 | User Error | SW | RMED SW | Calibration data lost |  | Loss of cal | Loss of cal tni. Checking messages ok. Following the pbm no longer has a K7 lot cal14732 Tni. | Update in version 8.11 to adjust pbm of cal loss and slowdown. Setting up automatic backup every day at 3am on OK USB key. No CQ to validate. Passing a test tube in tni ok.
No | Visit | Complaint | RSAF | Replaced SC and Gluc and lac fail QC | Error occurred with new SC | Replace faulty SC, QC Level A still out on Gluc and Lac. Possible Faulty Sol pack fault. Replaced solution pack and all levels ok. | SWO-418517 | Assisted Replacement | SC | SC | Glu/Lac |  | Replaced SC and Gluc and lac fail QC | Error occurred with new SC | Replace faulty SC, QC Level A still out on Gluc and Lac. Possible Faulty Sol pack fault. Replaced solution pack and all levels ok.
No | Visit | Complaint | RKK | ﾋﾞﾘﾙﾋﾞﾝ測定不能。度重なるトラブルのため点検依頼。 | 確認 | 代替機設置。 | SWO-418541 | Revisit required | HW | Oxi module | Bilirubin | 713; Install loan analyzer | Bilirubin measurement not able. The inspection request for repeated trouble. | Check | Installation of alternative machines.
No | Visit | Complaint | RCHN | PO2病人结果偏高，接口板腐蚀 | PO2偏高，与临床不符，借903-334一个 | 更换903-334后修复(by chuanwei) | SWO-417847 | [5 empty] | PO2 patient results are high, interface plate corrosion | PO2 high, not clinical, borrowed 903-334 a | Repair after replacement 903-334 (by chuanwei)
No | Visit | Complaint | RAME | Solution pack rejects due to mechanical error | Dirty contacts on sol pack | Cleaned contacts and restarted system. | SWO-419200 | Cleaned | HW | Basic Unit | PCB, Solution Pack PCB Id |  | Solution pack rejects due to mechanical error | Dirty contacts on sol pack | Cleaned contacts and restarted system.
No | Visit | Complaint | RAME | Liquid sensor failure | Could not confirm failure when onsite. Analyzer was unplugged and transferred to adjacent office. Waited for an hour, after which analyzer came up to Ready state with cals and QCs passed. Activity log showed analyzer was rebooted on Sep 21 1022pm, afterwards, started having UIR (user intervention required) messages. Liquid sensor messages started appearing at 1023pm. Use has tried replacing sensor cassette and solution pack, changed out inlet gasket holder and flushing to no avail. | Advised customer to order and install a ups (uninterrupted power supply) with line conditioning feature to avoid power surge that could disturb analyzer and monitor its effect. | SWO-419113 | No fault found | No fault found | No fault found | LS |  | Liquid sensor failure | Could not confirm failure when onsite. Analyzer was unplugged and transferred to adjacent office. Waited for an hour, after which analyzer came up to Ready state with cals and QCs passed. Activity log showed analyzer was rebooted on Sep 21 1022pm, afterwards, started having UIR (user intervention required) messages. Liquid sensor messages started appearing at 1023pm. Use has tried replacing sensor cassette and solution pack, changed out inlet gasket holder and flushing to no avail. | Advised customer to order and install a ups (uninterrupted power supply) with line conditioning feature to avoid power surge that could disturb analyzer and monitor its effect.
No | Visit | Complaint | RLTD | Mark 07939069137 | Battery icon showing zero % battery . | Discharge battery for 15 minuets. Removed module and rest battery jumpers. Reinstalled and battery icon now shows battery charge. | SWO-418814 | Adjustment | HW | Battery module | Battery module | 0% | Mark 07939069137 | Battery icon showing zero % battery . | Discharge battery for 15 minuets. Removed module and rest battery jumpers. Reinstalled and battery icon now shows battery charge.
No | Visit | Complaint | RFRA | Erreur de site d'introduction // base de données corrompue// erreur vide de pression Erreur apparue ce week-end // toujours présente malgré remplacement joint d'intro et pack de solution | Vérification complète | Test en mode service $ $ remplacement du joint d'introduction $ $ retrait et remise en place du site d'introduction , calibration $ $ réinstallation du pack de solution $ $ mise en route (1H) $ $ calibrations $ $ passage des contrôles $ $ essais: pco2 en défaut sur la calibrations, attendre 2h et si ça ne s'arrange pas, changer la cassette | SWO-418963 | Defect | HW | Auto Inlet module | Inlet gasket |  | Introductory site error // corrupted database// empty pressure error Error appeared this weekend // still present despite joint replacement of intro and solution pack | Full check | Test in service mode $$ replacement of the introductory seal $$ withdrawal and re-establishment of the introductory site, calibration $$ reinstallation of the solution pack $$ start (1H) $ calibrations $ $ passage of controls $$ trials: pco2 in default on the calibrations, wait 2 hours and if it doesn't work out, change the cassette
No | Visit | Complaint | RCHN | 检测结果异常 | 检测结果异常借903-702 944-216（14696）944-214（14465）944-268（14343） 944-220（14531）各一个 |  | SWO-418628 | Revisit required | HW | Process Module | Process Wheel |  | Abnormal detection results | Test results abnormal lynis903-702 944-216 (14696) 944-214 (14465) 944-268 (14343) 944-220 (14531) each
No | Visit | Complaint | RDEU | 20.09.2018 8:40Uhr Herr Adepar 0621 383-3528: Ständig Lac und Cl-Fehler | KAL Driften und QC Verletzung Lactat | Allgemeine Überprüfung - Kontakte etc. gereinigt - Erneuerung Dichtring Lactat (fehlte) - Membranierung mit neuer Membran-Charge (Station 10-4 mit gleichem Problem) - Funktionsüberprüfung etc. | SWO-419046 | Defect | Electrode | Electrode | Sealing ring, Lac |  | 20.09.2018 8:40h Mr Adepar 0621 383-3528: Constant Lac and Cl errors | KAL Drifting and QC Violation Lactate | General check - Contacts etc. cleaned - Renewal sealing ring lactate (missing) - membraneing with new membrane batch (station 10-4 with same problem) - functional check etc.
No | Visit | Complaint | RIND | Service Visit. | Check the calibration result and need to clean the pH electrode. | Check the pH chamber and electrode, clean it with hypochlorite, successfully done, do the inlet calibration, ls adjust, pump calibration, protein removal and cleaning and give the 2-point calibration, successfully done, analyzer handed over to staff in well working condition. | SWO-421822 | Cleaned | Fluidic pathway | Fluidic pathway | pH/BG Module |  | Service Visit. | Check the calibration result and need to clean the pH electrode. | Check the pH chamber and electrode, clean it with hypochlorite, successfully done, do the inlet calibration, ls adjust, pump calibration, protein removal and cleaning and give the 2-point calibration, successfully done, analyzer handed over to staff in well working condition.
No | Visit | Complaint | RAME | Customer called to report OXI hardware error. | Came on site to find neon intensity low. Neon number | After running complete OXI test, found lamp unit failed. Tested with used lamp unit for troubleshooting purpose and it passed. Got a new lamp unit from John S. Installed it. Adjusted lamp unit, ran complete OXI test and all passed. Ran couple of cals once passed ran tHB cal in service and all passed. All calibrations and 3 Levels of QC passed as per manufacturer specification | SWO-419219 | Defect | HW | Oxi module | Lamp unit |  | Customer called to report OXI hardware error. | Came on site to find neon intensity low. Neon number | After running complete OXI test, found lamp unit failed. Tested with used lamp unit for troubleshooting purpose and it passed. Got a new lamp unit from John S. Installed it. Adjusted lamp unit, ran complete OXI test and all passed. Ran couple of cals once passed ran tHB cal in service and all passed. All calibrations and 3 Levels of QC passed as per manufacturer specification
No | Visit | Complaint | RFRA | Suite à la mise en service de l'analyseur IA signale qu'il est impossible de retirer les vis du capot de la phase avant | Aucune | Remplacement du capot arrière. | SWO-419117 | Defect | HW | Cover and Chassis | Screws, Back cover | OOB | Following the commissioning of the AI analyzer, it is not possible to remove the screws from the hood of the front phase | No | Replacement of the rear hood.
No | Visit | Complaint | RDEU | 18.09.2018 10:00 Conny Löwenkamp 01737127139: Barcodescanner arbeitet nicht einwandfrei | Barcode der Monovetten werdem sporadisch nicht erkannt. | Barcodehinter dem Inlet mit Cleaner gereinigt und entstaubt. LQC Multicheck2 und CRP 1 sind in Ordnung. Mehrere EDTA- Röhrchen ausprobiert. Barcodereader in Ordnung. | SWO-418949 | Cleaned | HW | Aspiration Module | Barcode reader, Inlet |  | 18.09.2018 10:00 Conny Löwenkamp 01737127139: Barcode scanner does not work properly | barcode of the monovettes is not recognized sporadically. | Barcode behind the inlet cleaned with cleaner and dusted. LQC Multicheck2 and CRP 1 are fine. Several EDTA tubes tried out. Barcode reader in order.
No | Visit | Complaint | RDEU | Sensor Bay geht nur halb zu. | Kein fälliger SC Wechsel möglich. | Gerät überprüft. Oberteil stößt an SC an. Wet Section ausgebaut. Alle Komponenten sind fest. Keine Justagemöglichkeit erkennbar. Unter Manipulation SC Wechsel durchgeführt. Weitere Rep. Vor Ort nicht möglich. Kalibration und QC durchgeführt und in Ordnung | SWO-419642 | Adjustment | HW | Basic Unit | Sensor Cassette Door |  | Sensor Bay is only half-closed. | No due SC change possible. | device checked. Top bumps into SC. Wet Section. All components are fixed. No adjustment discernible. Under manipulation SC change performed. More Rep. Not possible on site. Calibration and QC carried out and in order
No | Visit | Complaint | RFRA | disk boot failure | ... | Réinstallation complète Windows 7 et application $ $ configuration Windows et programme abl90 $ $ chargement des données et de la configuration ( sur la clef du client) $ $ vérification du ticket et du Roma $ $ l'analyseur ne fonctionne qu'avec la clef USB à bord $ $ carte CF à remplacer, le technicien repassera cette semaine $ $ lac non fonctionnel depuis installation version 3.4 | SWO-419841 | Defect | HW | DMS_PC | CF card |  | disk boot failure | ... | Complete Windows 7 reinstallation and $$ Windows configuration app and abl90 program $$ data loading and configuration (on the customer key) $ $ ticket check and Roma $$ analyzer only works with the USB key on board $$$ ca RTE CF to replace, the technician will return this week $$ non-functional lake since installation version 3.4
No | Visit | Complaint | RFRA | Le paramètre Po2(T) non rendu dans l'analyse | Glucose non rendu car PO2 trop faible. Présence de deux types de seringue Radiometer et BD. Les seringues BD sont utilisées pour le cordon artériel, ce qui n'est pas recommandé. | 1-Création d 'une touche sonde courte. 2-Remise en place valeur par défaut avec suppression des hors gammes 3-Pour rendre les glucoses élevés, il est spécifié que la PO2 doit être supérieure à 3,3 kPa pour avoir toute la gamme enregistrable. 4-Mise a jour logiciel version 3.4. Essai : bon 5-Modification fichier roma.ini pour affichage des N° de travail. | SWO-419431 | No fault found | No fault found | No fault found | Glu | pO2 low | The Po2(T) setting not rendered in the analysis | Glucose not rendered because PO2 too low. Presence of two types of Radiometer and comics syringe. BD syringes are used for the arterial cord, which is not recommended. | 1-Creating a short probe key. 2-Re-up default with off-range removal 3-To make glucose high, it is specified that PO2 must be greater than 3.3 kPa to have the full range registerable. 4-Software update version 3.4. Try: good 5-Change file roma.ini for display of work numbers.
No | Visit | Complaint | RAME | Solution pack problem. | Analyzer SP multiple errors due to air contamination. | Replaced tubing during annual maintenance. SWO-419756 | SWO-419754 | [5 empty] | Solution pack problem. | Analyzer SP multiple errors due to air contamination. | Replaced tubing during annual maintenance. SWO-419756
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 定标不过 1315 | 更换测试卡后修复 | SWO-420745 | [5 empty] | Test card, reagent package failure | No more than 1315 | Repair after replacing test card
No | Visit | Complaint | RCHN | 处理包接触点被撞坏 | 处理包接触点被撞坏，借903-036一个 | 更换903-036后修复by tim | SWO-419352 | Defect | HW | Cover and Chassis | PCB, Solution Pack |  | The handling bag contact point was damaged | Treatment package contact point was damaged, borrow 903-036 one | Repair by tim after replacement 903-036
No | Visit | Complaint | RIND | Lactate sensitivity low. | Lactate sensitivity low. | Replaced the Lactate membrane & after 2point calibration it became ok but after that when run the sample it got aborted continuously.Found El $Met module chamber filled with liquid & overflowing,Flush the el $met module path but still leakage persist.Check all the electrodes 7 membrane for any clot but nothing found.Flush whole liquid path with distilled water,Clean the gasket properly.Run the pump calibration & LS adjust.It got passed successfully.Then run protein removal ,cleaning & the 2point calibration ,It got passed successfully.Then run few samples & got the desired result. | SWO-422145 | Assisted Replacement | Electrode | Membrane | Lac | Sensitivity low | Lactate sensitivity low. | Lactate sensitivity low. | Replaced the Lactate membrane & after 2point calibration it became ok but after that when run the sample it got aborted continuously. Found El $Met module chamber filled with liquid & overflowing,Flush the el $met module path but still leakage persist. Check all the electrodes 7 membrane for any clot but nothing found. Flush whole liquid path with distilled water,Clean the gasket properly. Run the pump calibration & LS adjust. It got passed successfully. Then run protein removal ,cleaning & the 2point calibration ,It got passed successfully. Then run few samples & got the desired result.
No | Visit | Complaint | RAME | Errors neon lamp intensity low and neon # too low | Parameter bar yellow failed qc oximetry hardware error Error 258 and 316 present neon lamp intensity and neon number too low | Replaced oximetry module | SWO-419885 | Defect | HW | Oxi module | Lamp unit | Neon number too low; mail: tried to adjust the lamp unit? | Errors neon lamp intensity low and neon # too low | Parameter bar yellow failed qc oximetry hardware error Error 258 and 316 present neon lamp intensity and neon number too low | Replaced oximetry module
No | Visit | Complaint | RDEU | 12.09.2018; KorLo; Tel.: 01737127139; Touchscreen reagiert nicht richtig, SW hängt sich auf | Es wurde berichtet, dass der Bildschirm beim bewegen ausfällt. | Flachbandkabel vom CPU- Einheit zum Bildschirm kontrolliert und ausgetauscht. | SWO-419711 | Defect | HW | Analyzer Control Module | Cable, display |  | 12.09.2018; Korlo; Tel.: 01737127139; Touchscreen does not respond correctly, SW hangs on | It has been reported that the screen fails while moving. | Tape cable from the CPU unit to the screen is checked and replaced.
No | Visit | Complaint | RDEU | Diverse QC und Kal. Fehler, trotz mehrfachem LP und SC wechseln. | Kunde berichtet von diversen QC und Kal. Fehlern trotz mehrfachem Wechsel von LP und SC | ABL gespült, Teile im Umfang einer Jahreswartung erneuert. Einlassonde erneuert. Einlaufverfahren und tHb Kal. durchgeführt. ABL90 Reklamations-Flyer ausgehändigt. Demo sPico , Anwender sollen in nächster Zeit mal sPico verwenden und in Mischer legen. Kapillaren sollen nun mit Mischerstäbchen gemischt werden. | SWO-420787 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway | Clot? | Various QC and Kal. Error, despite changing lp and SC multiple times. | Customer reports from various QC and Kal. Errors despite multiple changes of LP and SC | ABL rinsed, parts renewed in the scope of an annual maintenance. Inlet probe renewed. inlet method and tHb Kal. Performed. ABL90 Complaint Flyer. Demo sPico , users should use sPico in the near future and put them in mixers. Capillaries are now to be mixed with mixer rods.
No | Visit | Complaint | RAME | Hold with no restart | Found pinch tubing barely connected probably due to over pressure of flush. Hemolyzer now questionable. | Replaced all tubings and hemolyzer . Completed startup with calibration passing. | SWO-420702 | Defect | HW | Oxi module | Hemolyzer |  | Hold with no restart | Found pinch tubing barely connected probably due to over pressure of flush. Hemolyzer now questionable. | Replaced all tubings and hemolyzer . Completed startup with calibration passing.
No | Visit | Complaint | RLTD | Please call Steve on 01206 742 493 | Status code 2 error | Attempted to rename files as per status code 2 manual solution, however every time I tried to delete old file it told me it was in use, attempted to reload software but it would just reboot and not acknowledge new software. Attempted to reload Windows software, again analyser came up with a number of errors pertaining to flash card. Replaced flash card and reloaded windows 7 software, and analyser software. Copied setup from a nearby analyser, confirmed with POCT that this was ok. Pushed users via Aqure. | SWO-420595 | Defect | HW | DMS_PC | CF card | status code 2 | Please call Steve on 01206 742 493 | Status code 2 error | Attempted to rename files as per status code 2 manual solution, however every time I tried to delete old file it told me it was in use, attempted to reload software but it would just reboot and not acknowledge new software. Attempted to reload Windows software, again analyser came up with a number of errors pertaining to flash card. Replaced flash card and reloaded windows 7 software, and analyser software. Copied setup from a nearby analyser, confirmed with POCT that this was ok. Pushed users via Aqure.
No | Visit | Complaint | RLTD | Joseph 02034470329 | Analyser booting up and loading programs ok. The front screen gives you log on box and you can log on ok. Screen says analyser not ready and the only button you can press in the data logs. | Checked all connections on WS PCB on and connection to CPU ok. When CPU is disconnected from SW and reconnected the screen then changes and it gives you the option to replaced SC and the screen works as normal. Suspect CPU or WS is at fault parts on order. | SWO-420438 | Revisit required | SW | RMED SW | reinstall |  | Joseph 02034470329 | Analyser booting up and loading programs ok. The front screen gives you log on box and you can log on ok. Screen says analyser not ready and the only button you can press in the data logs. | Checked all connections on WS PCB on and connection to CPU ok. When CPU is disconnected from SW and reconnected the screen then changes and it gives you the option to replaced SC and the screen works as normal. Suspect CPU or WS is at fault parts on order.
Yes | 1 | Y | RPAC | Blank | 1066 Temperature in sensor cassette substrate out of range |  | D4FD4FEB7A20BE22 | [5 empty] | Blank | 1066 Temperature in sensor cassette substrate out of range
No | Visit | Complaint | RLTD | Intervention required unable to pump solutions. Andrew Masters | Intervention required unable to pump solutions. Front support catch for SP found to be lose screw . | Removed SP and flushed system from Oxi to waste no clot fond. Flushed SC and this is ok. Flushed from right hand cassette sensor bay port out to probe and found this to be blocked. Cleaned blockage and refitted all part. Run all service test passed no errors. Refitted screw and checked support bracket is well fixed in place. | SWO-420475 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway | Clot | Intervention required unable to pump solutions. Andrew Masters | Intervention required unable to pump solutions. Front support catch for SP found to be lose screw . | Removed SP and flushed system from Oxi to waste no clot fond. Flushed SC and this is ok. Flushed from right hand cassette sensor bay port out to probe and found this to be blocked. Cleaned blockage and refitted all part. Run all service test passed no errors. Refitted screw and checked support bracket is well fixed in place.
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | pco2灵敏度超范围，0378 | 更换测试卡后修复 | SWO-420846 | Assisted Replacement | SC | SC | pCO2 |  | Test card, reagent package failure | pco2 sensitivity over range, 0378 | Repair after replacing test card
No | Visit | Complaint | RCHN | 温度超出范围，报错1221 | 温度超出范围，报错1221，需要借件903-199 | 更换903-199后修复by tim | SWO-419424 | Defect | HW | Basic Unit | Cassette retainer | 1221 | Temperature out of range, error 1221 | Temperature out of range, error 1221, need to borrow 903-199 | Repair by tim after replacement 903-199
No | Visit | Complaint | RCHN | 电源损坏 | 电源损坏 借910-285一个 | 更换910-285后修复by tim | SWO-420152 | Defect | HW | Power Supply | Power supply |  | Power damage | Power damage, borrow 910-285 one. | Repair by tim after replacement 910-285
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | SO2,CHB质控失败 0090 | 更换测试卡 | SWO-422183 | [5 empty] | Test card, reagent package failure | SO2, CHB QC Failure 0090 | Replace the test card
No | Visit | Complaint | RCHN | 测试卡无法扫描 | 测试卡无法扫描 借910-304一个 | 更换910-304后修复by tim | SWO-418741 | Defect | HW | Storage Module | Barcode reader, storage wheel |  | Test card can't be scanned | Test card can't scan borrowing 910-304 one | Repair by tim after replacing 910-304
Yes | Visit | Complaint | RCHN | 报错1031 | 报错1031，借903-364一个 | 更换903-364后修复by tim | SWO-415843 | Defect | HW | Process Module | Optical Unit | 1031 | 1031 | 1031, loan 903-364 one | Repair by tim after replacing 903-364
No | Visit | Complaint | RCHN | 报错2872 | 报错2872，借580-097一个 | 更换580-097后修复by tim | SWO-419414 | Defect | HW | Storage Module | Motor, cartridge slider | 2872 | 2872 | 2872, 580-097 one | Repair by tim after replacement 580-097
No | Visit | Complaint | RCHN | 定标卡无法扫描 | 定标卡无法扫描 借902-924一个 | 更换902-924后修复by tim | SWO-418780 | Defect | HW | Cover and Chassis | Barcode reader, display |  | Bid card cannot be scanned | The bid card cannot be scanned borrowing 902-924 one | Repair by tim after replacement 902-924
No | Visit | Complaint | RIND | SC9 was unable to use with the remaining test. | Found that SC unable use due to date $time change,But customer has not change it,Instrument was automatically off due to power cut & after that SC9 was unable to use | Check the current date $time & found that it was previous date,change it to the current date& time & now remaining test of SC9 was showing.Handover the instrument to the customer. | SWO-422155 | [5 empty] | SC9 was unable to use with the remaining test. | Found that SC unable use due to date /time change,But customer has not change it,Instrument was automatically off due to power cut & after that SC9 was unable to use | Check the current date $time & found that it was previous date,change it to the current date& time & now remaining test of SC9 was showing. Handover the instrument to the customer.
Yes | Visit | Complaint | RAME | Maria called to advise that the demo keys have expired and the analyzer is no longer allowing calibrations. | Missing option keys | Installed option keys | SWO-421794 | Re classified RMED | Demokey expired |  | Demo key |  | Maria called to advise that the demo keys have expired and the analyzer is no longer allowing calibrations. | Missing option keys | Installed option keys
No | Visit | Complaint | RAME | Lac out on cals and qc | Lac sensitivity out of range drift out of range Customer replaced membranes with 2 lots Pa $mm high running around 1600 passing previous calibrations a few months back 800-1065 pa $mm Suggest changing electrode Replaced electrode and membrane Calibrations passing 3 calibrations after a zero current cal was run Qc all in with pre established ranges Stayed for correlations 2 hours | Replacement of lac electrode Zero current cal Recommend re establishing 2 SD for lac with the new electrode | SWO-420740 | Assisted Replacement | Electrode | Electrode | Lac |  | Lac out on cals and qc | Lac sensitivity out of range drift out of range Customer replaced membranes with 2 lots Pa /mm high running around 1600 passing previous calibrations a few months back 800-1065 pa /mm Suggest changing electrode Replaced electrode and membrane Calibrations passing 3 calibrations after a zero current cal was run Qc all in with pre established ranges Stayed for correlations 2 hours | Replacement of lac electrode Zero current cal Recommend re establishing 2 SD for lac with the new electrode
Yes | Visit | Complaint | RAME | Customer called to report OXI hardware error. | Follow up to last visit. | Came onsite with Jamie to visit the customer as well as check up on all the instruments. All working good. No QC as instrument not disturbed. Drove to airport to pick up Jamie and then to customer | SWO-421393 | Re classified RMED | planned SWO |  |  |  | Customer called to report OXI hardware error. | Follow up to last visit. | Came onsite with Jamie to visit the customer as well as check up on all the instruments. All working good. No QC as instrument not disturbed. Drove to airport to pick up Jamie and then to customer
No | Visit | Complaint | RAME | Failed MetHb correlation with reference analyzers | Customer having problems correlating with old analyzer. | Adjusted Lamp to 3288. Did Complete oximetry. Failed for neon lamp. Replaced Oximetry module. Adjusted lamp to 3218. Did complete oximetry, all passed. Did THB in service and as user and as sample. | SWO-420664 | Defect | HW | Oxi module | Oxi module |  | Failed MetHb correlation with reference analyzers | Customer having problems correlating with old analyzer. | Adjusted Lamp to 3288. Did Complete oximetry. Failed for neon lamp. Replaced Oximetry module. Adjusted lamp to 3218. Did complete oximetry, all passed. Did THB in service and as user and as sample.
No | Visit | Complaint | RLTD | Jon Jenkins-Waud reports that the analyser has not performed well during validation on MetHb and tHb. | In order to asses if the analyser’s performance is not as it should be, a sample was ran on Lab A, Lab B and the SCBU new analyser. The results were as follows, (results in order of Lab A, Lab B, SCBU new), tHB - 140, 143, 139, MetHb - 0.9, 0.8, 0.9. These appear comparable. All three analysers then had a protein removal, a complete Oxi test was ran and passed and the halogen lamps were adjusted close to 3400 if required. A service and user tHb calibration was ran and passed. The same sample as prior to testing was then ran on the analysers and the results are as follows, (results in order of Lab A, Lab B and SCBU new). tHb - 141, 141,139. MetHb - 1.2, 0.9, 0.6. QC S7735 Lot769 was ran on all analysers and the results were very good comparisons. | At this moment in time no error can be identified with the analyser. It was noted that the sample volumes used for comparison were not the same. | SWO-420500 | No fault found | HW | Oxi module | Oxi module |  | Jon Jenkins-Waud reports that the analyser has not performed well during validation on MetHb and tHb. | In order to asses if the analyser’s performance is not as it should be, a sample was ran on Lab A, Lab B and the SCBU new analyser. The results were as follows, (results in order of Lab A, Lab B, SCBU new), tHB - 140, 143, 139, MetHb - 0.9, 0.8, 0.9. These appear comparable. All three analysers then had a protein removal, a complete Oxi test was ran and passed and the halogen lamps were adjusted close to 3400 if required. A service and user tHb calibration was ran and passed. The same sample as prior to testing was then ran on the analysers and the results are as follows, (results in order of Lab A, Lab B and SCBU new). tHb - 141, 141,139. MetHb - 1.2, 0.9, 0.6. QC S7735 Lot769 was ran on all analysers and the results were very good comparisons. | At this moment in time no error can be identified with the analyser. It was noted that the sample volumes used for comparison were not the same.
No | Visit | Complaint | RAME | Cleaning solution bottle leaking. | Cleaning solution bottle leaking. | Instructed customer to replace with new, part # 944-126. | SWO-420659 | Assisted Replacement | Solution | Solution | Cleaning solution | Leakage | Cleaning solution bottle leaking. | Cleaning solution bottle leaking. | Instructed customer to replace with new, part # 944-126.
No | Visit | Complaint | RIND | Service Visit. | Test cartridge barcode is not read by barcode reader. | Open the front panel of the analyzer, remove the dust from the barcode, successfully done, close the panel, analyzer handed over to staff in well working condition. | SWO-421807 | Cleaned | HW | Storage Module | Barcode reader, storage wheel |  | Service Visit. | Test cartridge barcode is not read by barcode reader. | Open the front panel of the analyzer, remove the dust from the barcode, successfully done, close the panel, analyzer handed over to staff in well working condition.
No | Visit | Complaint | RDEU | Akku Fehler 1559 unter Systemmeldungen | Akku defekt Fehler 1559 | Akku ersetzt, Einlassnadel justiert, Kal und QC durchgeführt | SWO-420961 | Defect | HW | Battery module | Battery module | 1559 | Battery Error 1559 under System Messages | Battery defective Error 1559 | Battery replaced, inlet needle adjusted, Kal and QC performed
Yes | 2 | Y | RPAC | Blank | 378 calibration sensitivity out of range, 1319 cal back log error Na, 1321 cal back log error |  | D4FEBB5CF01D18FB | [5 empty] | Blank | 378 calibration sensitivity out of range, 1319 cal back log error Na, 1321 cal back log error
No | Visit | Complaint | RFRA | Fuite détectée | Contrôle de l’étanchéité de la machine, de la vanne oxy | Changement de tuyau de la vanne oxy. Essais | SWO-420928 | Defect | HW | Oxi module | Pinch valve tube |  | Leak detected | Control of the sealing of the machine, the oxy valve | Change of pipe from the oxy valve. Tests
No | Visit | Complaint | RFRA | Err base de données status 2 | Pc $cf card | Démontage et remplacement de la CPU UNIT UIM945 $ Installation de Windows7 $Installation du patch $Installation de l'application Radiometer $Configuration IP machine $Restauration de la configuration de la machine $Remise du lecteur Z $Remise en route $ Calibration $Calibration tHB $passage des CQ | SWO-423739 | Defect | HW | DMS_PC | CPU | status code 2 | Err status database 2 | Pc /cf card | UnIT UIM945 $ UnIT CPU removal and replacement Installing Windows7 $Installation the $Installation patch of the Radiometer $Configuration IP machine application $Restauration the $Remise configuration of the Z player $Remise en route $ THB $Calibration calibration $passage CQ
No | 3 | Y | RPNZ | Blank | Lactate failing calibrations (0375: calibration status out of limits) |  | D4FECF55F03CBDB2 | Assisted Replacement | SC | SC | Lac | mail: PER system | Blank | Lactate failing calibrations (0375: calibration status out of limits)
No | Visit | Complaint | RCHN | 开机卡机 | 开机卡机，借903-381一个 | 更换903-381后修复by tim | SWO-419310 | [5 empty] | Power-up card machine | Power-up card, borrow 903-381 one | Repair by tim after replacement 903-381
No | Visit | Complaint | RIND | K, Na, pH, pCO2 and pO2 calibration failure on ABL80 anlyzer at Inamdar Hospital, Pune | Need to check the SC as well as SP. | Tried reinstalling both SC and SP, tried 2-point calibration after that, machine is still not calibrating these parameters. Replaced the cassette, done calibration. Machine is working fine. | SWO-422161 | [5 empty] | K, Na, pH, pCO2 and pO2 calibration failure on ABL80 anlyzer at Inamdar Hospital, Pune | Need to check the SC as well as SP. | Tried reinstalling both SC and SP, tried 2-point calibration after that, machine is still not calibrating these parameters. Replaced the cassette, done calibration. Machine is working fine.
No | Visit | Complaint | RAME | Failing glucose. Contact: Susan 615-574-0600 | Glucose not in. Also arrived on site and leak under analyzer reported | Replaced various tubings and manifold from PM kit. Bottom of SP appears to have small crack in structure installed account replacement Sensor Cassette. Also printed documents to replace SP | SWO-421454 | [5 empty] | Failing glucose. Contact: Susan 615-574-0600 | Glucose not in. Also arrived on site and leak under analyzer reported | Replaced various tubings and manifold from PM kit. Bottom of SP appears to have small crack in structure installed account replacement Sensor Cassette. Also printed documents to replace SP
No | Visit | Complaint | RCHN | 无法读取或者写入传感器数据1123 | 无法读取或者写入传感器数据1123,借 903-581一个。 | 更换903-581后修复by tim | SWO-420147 | Defect | HW | Basic Unit | PCB, Sensor interface | 1123 | Cannot read or write sensor data 1123 | Sensor data 1123 could not be read or written, borrowing 903-581 one. | Repair by tim after replacement 903-581
Yes | Visit | Complaint | RCHN | 温度报警 | 温度报警，借903-210一个 | 更换903-210后修复by tim | SWO-419970 | Defect | HW | Process Module | PCB, cup dryer | 1068 | Temperature alarm | Temperature alarm, borrow 903-210 one | Repair by tim after replacement 903-210
No | Visit | Complaint | RIND | Pco2 sensitivity and status out of range. | Status and sensitivity is out in calibration. | Check the machine, clean the chambers and run 2 point calibration. But still there is same problem exist. Change Pco2 membrane and run 2 point calibration but still no changes. Than change pco2 electrode from another machine. And run 2 point calibration successfully. Now machine is working fine and ready to use. | SWO-422193 | Assisted Replacement | Electrode | Electrode | pCO2 | Sensitivity and Status | Pco2 sensitivity and status out of range. | Status and sensitivity is out in calibration. | Check the machine, clean the chambers and run 2 point calibration. But still there is same problem exist. Change Pco2 membrane and run 2 point calibration but still no changes. Than change pco2 electrode from another machine. And run 2 point calibration successfully. Now machine is working fine and ready to use.
No | 4 | Y | RPAC | Blank | Cl issues |  | D4FE9BC5F00F4DE7 | Assisted Replacement | Electrode | Electrode | Cl | Lot 802 | Blank | Cl issues
No | Visit | Complaint | RAME | Failed MetHb correlation with reference analyzers | Customer having problem correlating MetHB to Old analyzer. | Adjusted lamp 3263, then did complete oximetry test, failed for hemolyzer, phase regulator error. Replaced hemolyzer. Adjusted lamp to 3276, then did complete oximetry test, all passed. Did THB calibration in service as user and as sample. | SWO-420668 | Defect | HW | Oxi module | Hemolyzer |  | Failed MetHb correlation with reference analyzers | Customer having problem correlating MetHB to Old analyzer. | Adjusted lamp 3263, then did complete oximetry test, failed for hemolyzer, phase regulator error. Replaced hemolyzer. Adjusted lamp to 3276, then did complete oximetry test, all passed. Did THB calibration in service as user and as sample.
No | Visit | Complaint | RHUN | mérés közben, új minta beadásakor a minta ajtó nem záródik be, mérés megszakad, visszaállítás folyamatban, újra indítás nem segít 09.26-án a hiba újra jelentkezik, 2x-es újraindítás után a gép működik | - | minta ajtók pozíciójának ellenőrzése és után állítása, cabinet PCB csere, szerviz tesztek elvégzése | SWO-421021 | Defect | HW | Cover and Chassis | PCB, cabinet top |  | During measurement, when a new sample is checked, the sample door is not closed, measurement is interrupted, restoration is in progress, re-booting does not help 09.26 The error occurs again, after 2x reboot the machine is working | - | Sample door position check and after adjustment, cabinet PCB replacement, performing service tests
No | Visit | Complaint | RKK | HCTが測定出来ない事があった。 | HCTが測定出来ない事があった。 | クリーニング後も変わらずとの事で、 ウォッシュブロック交換及び排液フィルター清掃実施。 暫く様子見です。 | SWO-420834 | Defect | HW | Aspiration Module | Needle Wash Unit |  | HCT could not be measured. | HCT could not be measured. | In the thing that it does not change even after cleaning, Wash block replacement and drainage filter cleaning. It is a wait for a while.
No | Visit | Complaint | RPOL | Błąd przepływów , błąd kalibracji pompy. | Błąd przepływów spowodowany skrzepem w analizatorze | Przepłukanie toru pomiarowego . Wymiana igły . | SWO-421240 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway | Clot | Error flows, pump calibration error. | Flow error caused by clot in analyzer | Flush the measuring track. Replacing the needle.
Yes | Visit | Complaint | RDEU | 27.09.2018 11:00 Fr. Judäa 06217591319: Anwenderin möchte Netzwerkanbindung | LIS $KIS Anbindung | LIS $ KIS Anbindung ohne feste IP- Adresse Funktionsüberprüfung | SWO-421605 | Re classified RMED | planned SWO |  |  |  | 27.09.2018 11:00 Fr. Judea 06217591319: User wants network connection | LIS /KIS connection | LIS - KIS connection without fixed IP address Functional verification
No | Visit | Complaint | RLTD | Joseph 02034470329 | Checked all connections on WS PCB on and connection to CPU ok. When CPU is disconnected from SW and reconnected the screen then changes and it gives you the option to replaced SC and the screen works as normal. Suspect CPU or WS is at fault parts on order. | Replaced SW pcb installed software and after reboot the screen then came up asking for consumable. Rebooted analyser again and screen returned to intervention as before and the only key available is the data logs. Fitted new cpu using flash card form original cpu. Analysers started up and allowed you to install consumable. Again after second reboot the screen returned to intervention as before with data log key only available. Traced fault to corrupt data files. Wiped flash card and reinstalled windows and analyser software refitted originally parts. Analyser starts up and consumable installed ok. Setup reinstalled and analyser up and running ok. Data from analyser had already been sent to Aqure. | SWO-421930 | Defect | SW | RMED SW | reinstall |  | Joseph 02034470329 | Checked all connections on WS PCB on and connection to CPU ok. When CPU is disconnected from SW and reconnected the screen then changes and it gives you the option to replaced SC and the screen works as normal. Suspect CPU or WS is at fault parts on order. | Replaced SW pcb installed software and after reboot the screen then came up asking for consumable. Rebooted analyser again and screen returned to intervention as before and the only key available is the data logs. Fitted new cpu using flash card form original cpu. Analysers started up and allowed you to install consumable. Again after second reboot the screen returned to intervention as before with data log key only available. Traced fault to corrupt data files. Wiped flash card and reinstalled windows and analyser software refitted originally parts. Analyser starts up and consumable installed ok. Setup reinstalled and analyser up and running ok. Data from analyser had already been sent to Aqure.
No | Visit | Complaint | RCHN | 无法读取试剂包智能芯片 | 无法读取试剂包智能芯片 借903-158一个 | 更换903-158后修复by tim | SWO-420755 | Defect | HW | Basic Unit | PCB, Solution Pack PCB Id | 1152 | Unable to read the reagent package smart chip | Unable to read reagent package smart chip borrowed 903-158 one | Repair by tim after replacement 903-158
Yes | 2 | Y | RPAC | Blank | Software upgrade V1.2.1 |  | D4FFDDAC662CF25C | Re classified RMED | planned SWO |  |  |  | Blank | Software upgrade V1.2.1
Yes | 2 | Y | RPAC | Blank | Software upgrade V1.2.1 |  | D4FFDA14665EF258 | Re classified RMED | planned SWO |  |  |  | Blank | Software upgrade V1.2.1
No | Visit | Complaint | RCHN | 报错1068 测试卡无法扫描 | 报错1068 测试卡无法扫描 借903-210 910-304各一个 | 更换903-210 910-304后修复by tim | SWO-420018 | Defect | HW | Process Module | PCB, cup dryer | 1068 | Error 1068 test card cannot be scanned | Error 1068 test card can not be scanned borrowing 903-210 910-304 each | Repair by tim after replacing 903-210 910-304
No | Visit | Complaint | RCHN | 存储盘卡盘 5784 | 存储盘卡盘 5784 | 更换测试卡正常。by tim | SWO-422085 | Defect | SW | Analyzer Control Module | RMED SW | 5784 | Memory Tray 5784 | Memory Tray 5784 | The replacement test card is OK. by tim
Yes | Visit | Complaint | RIND | pH & PCO2 error in calibration | checked and found that the problem in pH and PCO2 | done the zeroing, but no change in the calibration result. tried with new distribution board but still same problem was there as the same problem also came earlier even changing the solution pack and the cassette. lastly changed the solution pack the analyzer starts working fine. | SWO-422154 | [5 empty] | pH & PCO2 error in calibration | checked and found that the problem in pH and PCO2 | done the zeroing, but no change in the calibration result. tried with new distribution board but still same problem was there as the same problem also came earlier even changing the solution pack and the cassette. lastly changed the solution pack the analyzer starts working fine.
No | Visit | Complaint | RBNL | Errormelding 1029 op DDIM meting. | Bij DDIM meting werd error 1029 vermeld. Verder zijn er geen foutmeldingen aanwezig in de AQT. Reference cup van het Processwheel schoongemaakt vlgs voorschrift. Hierna een blank meting uitgevoerd, deze is oke. Testpatient gemeten en nogmaals blank meting uitgevoerd, deze zijn beiden oke. Alle levels QC gemeten —> oke. | reference cup in processwheel schoongemaakt. | SWO-421600 | Cleaned | HW | Process Module | Process Wheel | 1029 | Error message 1029 on DDIM metering. | At DDIM metering, error 1029 was mentioned. Furthermore, no error messages are present in the AQT. Reference Cup of the processwheel cleaned ACC prescription. After this a blank measurement is performed, this one is okay. Testpatient measured and again carried out blank measurement, these are both okay. All levels QC measured — > okay. | Reference Cup in Processwheel cleaned.
No | Visit | Complaint | RAME | Sensor bay is not closed | SC door not closing. ref SWO CTSL 280818-1 | Ran all service checks. Opened and closed sensor door multiple times and it forze once. Analyzer seemd to be wokring ok . Calibrations and QC accepted. | SWO-432565 | Adjustment | HW | Basic Unit | Sensor door |  | Sensor bay is not closed | SC door not closing. ref SWO CTSL 280818-1 | Ran all service checks. Opened and closed sensor door multiple times and it forze once. Analyzer seemd to be wokring ok . Calibrations and QC accepted.
No | Visit | Complaint | RCHN | 接口板腐蚀 | 接口板腐蚀，借902-972一个 | 更换902-972后修复by tim | SWO-420973 | [5 empty] | Interface plate corrosion | Interface plate corrosion, borrow 902-972 one | Repair by tim after replacement 902-972
No | Visit | Complaint | RCHN | 接口板腐蚀 | 接口板腐蚀，借902-972一个 | 更换902-972后修复by tim | SWO-421555 | [5 empty] | Interface plate corrosion | Interface plate corrosion, borrow 902-972 one | Repair by tim after replacement 902-972
No | Visit | Complaint | RCHN | GLU,LAC定标不过0378 | GLU,LAC定标不过0378 借903-581一个 | 更换903-581一个正常。by tim | SWO-420768 | Defect | HW | Basic Unit | PCB, Sensor interface | Glu/Lac | GLU, LAC calibration but 0378 | GLU, LAC calibration but 0378 borrowing 903-581 one | Replace the 903-581 one normal. by tim
No | Visit | Complaint | RCHN | 升级错误 | 升级错误，需要借件910-329 | 更换910-329后，发现CPU无法识别U盘，需要重新借件910-342一个by tim | SWO-420794 | Defect | HW | DMS_PC | CPU | SW upgrade | Upgrade error | Upgrade error, need to borrow 910-329 | After replacing 910-329, it was found that the CPU could not recognize the USB stick and needed to borrow a by-tim of 910-342.
Yes | Visit | Complaint | RCHN | 升级错误 | 升级错误，之前借的配件发生CPU无法识别U盘 需要借件910-342 |  | SWO-422142 | Re classified RMED | same day upSWO |  |  |  | Upgrade error | Upgrade error, previously borrowed accessories occurred CPU unrecognizable USB stick need sloan 910-342
No | Visit | Complaint | RCHN | 报错4548 | 报错4548，借903-028一个 | 更换903-028后修复by tim | SWO-420955 | Defect | HW | Cover and Chassis | PCB, cabinet top | 4548 | 4548 error | 4548, loan 903-028 one | Repair by tim after replacement 903-028
Yes | Visit | Complaint | RCHN | 定标失败 5784 | 在镇江市第三人民医院检验科装机，报错定标失败 5784,借 944-258（14178）一张。 | 重新定标正常。by tim | SWO-421598 | [5 empty] | Bid failure 5784 | In Zhenjiang City, the Third People's Hospital Inspection Department installed, reported the miscalibration failure 5784, borrowed 944-258 (14178) one. | Recalibration is normal. by tim
No | Visit | Complaint | RCHN | 报错5784 | 报错5784 借580-098一个 | 更换580-098一个正常。by tim | SWO-421426 | Defect | HW | Process Module | Motor, process wheel | 5784 | 5784 | False 5784 Borrow 580-098 one | Replace 580-098 one normal. by tim
No | Visit | Complaint | RPOL | Glu i Lac nierzetelny na kalibracji | Błąd formatowania w module El $Met. | Wymiana wężyka pompowego w module El $Met. Testy serwisowe ok. | SWO-422176 | Defect | HW | El/Met Module | Pump tubes | 677 | Glu and Lac unreliable for calibration | Formatting error in module El /Met. | Replacing the pump hose in the El $Met module. Service tests approx.
Yes | 18 | Y | RPNZ | Blank | Daylight savings time update. |  | D50244A15234BDB6 | [5 empty] | Blank | Daylight savings time update.
No | Visit | Complaint | RCHN | 屏幕不亮 | 屏幕不亮 借910-342一个 |  | SWO-423680 | Defect | HW | DMS_PC | CPU | Display | The screen is not lit | The screen doesn't light, borrow 910-342 one.
No | Visit | Complaint | RCHN | 测试卡条码扫码器无法扫码 | 测试卡条码扫码器无法扫码 | 调整了扫码器扫描位置后修复 by Andrew | SWO-424024 | Adjustment | HW | Storage Module | Barcode reader, storage wheel |  | Test card barcode sweeper can't sweep code | Test card barcode sweeper can't sweep code | Fix by Andrew after adjusting the scanner scan position
No | Visit | Complaint | RKOR | cartridgelock_52, The original solution pack is released and not be reinserted | All position check punch up -> Failed process wheel, incubator wheel -> failed Restart the analyzer -> all position oK The original solution pack is released and not be reinserted-> re-insert the SP -> falied Replace the sp cartridge is broken => replace the cartridge | Restart the anlayzer, replace the solution pack | SWO-423713 | Defect | Consumables | Test Kit | Cartridge, Broken |  | cartridgelock_52, The original solution pack is released and not be reinserted | All position check punch up -> Failed process wheel, incubator wheel -> failed Restart the analyzer -> all position oK The original solution pack is released and not be reinserted-> re-insert the SP -> falied Replace the sp cartridge is broken => replace the cartridge | Restart the anlayzer, replace the solution pack
Yes | Visit | Complaint | RAME | QC Errors/ Flag sent to Patient Chart | Found patient sample in question, per IT team I checked the output of the analyzer. The analyzer sent exactly what the screen had shown | I sent the outcome to the I T department. As aqure needed to be addressed | SWO-427882 | Re classified RMED | Courtesy |  |  | Refer to mailXXX | QC Errors/ Flag sent to Patient Chart | Found patient sample in question, per IT team I checked the output of the analyzer. The analyzer sent exactly what the screen had shown | I sent the outcome to the I T department. As aqure needed to be addressed
No | Visit | Complaint | RDEU | 24.09.2018; Hr. Wolter; Tel.: 03334692061; pCO2 Kal Fehler 24.09.2018; 11.00 Uhr; Hr. Wolter; Tel.: 03334692061; pCO2 Kal Fehler | pCO2-Fehler QC Level C | gl. SC erneut initialisiert; Startphase abgewartet -> weiter Fehler; neue SC eingesetzt; Startphase -> alles i.O.; Fehlerprotokolle erstellt | SWO-424428 | Assisted Replacement | SC | SC | pCO2 |  | 24.09.2018; Mr. Wolter; Tel.: 03334692061; pCO2 Kal Error 24.09.2018; 11 a.m.; Mr. Wolter; Tel.: 03334692061; pCO2 Kal Error | pCO2-Fehler QC Level C | Gl. SC reinitialized; Start-up phase awaited -> further error; new SC; Start-up phase -> everything in the world; Error logs created
Yes | 11 | Y | RPAC | Blank | 1066 temperature in sensor cassette substrate out of range 1326 sensor thermistor reclabibration falield - thermistor mal function |  | D503E7AFC81C5CD0 | [5 empty] | Blank | 1066 temperature in sensor cassette substrate out of range 1326 sensor thermistor reclabibration falield - thermistor mal function
No | Visit | Complaint | RLTD | Emma on 01823 346747 | On arrival the issue with the inlet not closing had been resolved but I was unable to find out how. | I checked the screws on the motor clutch to make sure they were fully tightened and they did appear to be although I was able to tighten them a small amount. If the problem reoccurs then leave the analyser as it is and put a call in so that I can come back and see it in its fault state. | SWO-424208 | No fault found | HW | Auto Inlet module | Auto Inlet module |  | Emma on 01823 346747 | On arrival the issue with the inlet not closing had been resolved but I was unable to find out how. | I checked the screws on the motor clutch to make sure they were fully tightened and they did appear to be although I was able to tighten them a small amount. If the problem reoccurs then leave the analyser as it is and put a call in so that I can come back and see it in its fault state.
Yes | 4 | Y | RCZE | Blank | Pri merani ci striekacka alebo kapilary zostane vstup vysiet a nezasunie sa. Nie pri kazdej vzorke.Inak analyzator sa tvari ok, kal ok QC ok. |  | D503B088C80C96B4 | [5 empty] | Blank | When the tube or capillary is measured, the entry shall remain in the sowing and not inserted. Not for the sample. Otherwise the analyator is Tvari OK, kal ok QC OK.
Yes | 1 | Y | RPAC | Blank | 378- cal sensitivity out of range, 1070 sensor response error |  | D503E23BC8285CCC | [5 empty] | Blank | 378- cal sensitivity out of range, 1070 sensor response error
Yes | Visit | Complaint | RLTD | RED BITS ON ALL ELECTROLITES AND STILL SAYING THERES A PROBLEM GEMMA LINDSEY - 01206 742 857 | Oximetry pinch valve tubing failed. | Carried out annual service visit as this is due this month. All parts and labour on annual service work order. SWO-424193 | SWO-424192 | [5 empty] | RED BITS ON ALL ELECTROLITES AND STILL SAYING THERES A PROBLEM GEMMA LINDSEY - 01206 742 857 | Oximetry pinch valve tubing failed. | Carried out annual service visit as this is due this month. All parts and labour on annual service work order. SWO-424193
No | Visit | Complaint | RLTD | SAYS LEAK ON SCREEN CODE 1342 LEAK DETECTED - Shaun Quinlivan - 01708 435 0000 EXT 2621 | Checked data logs and service programs. Opened SC bay and removed OXI tubing. The tubing has split and is the cause of the error. SC has run out of tests. The SC will be replaced by the customer once the delivery has arrived. | Replaced OXI tubing and tested. All fluid transport tests have passed. | SWO-423960 | Defect | HW | Oxi module | Pinch valve tube | 1342 | SAYS LEAK ON SCREEN CODE 1342 LEAK DETECTED - Shaun Quinlivan - 01708 435 0000 EXT 2621 | Checked data logs and service programs. Opened SC bay and removed OXI tubing. The tubing has split and is the cause of the error. SC has run out of tests. The SC will be replaced by the customer once the delivery has arrived. | Replaced OXI tubing and tested. All fluid transport tests have passed.
No | Visit | Complaint | RDEU | Fehlerbehebung erforderlich | Einstellungen und Daten geprüft Mögliche Störquellen wurden bei der IT abgefragt | ASUS USB AC51 nach RDEU-Anweisung installiert --> ABL schreibt in der Datenbank im Sekundentakt, dass er die Verbindung zu KIS $LIS hergestellt. bzw verloren hat. --> ABL wieder per LAN angeschlossen --> ABL nochmals neu gestartet, da die Meldungen 0133 $0134 weiterhin ständig aktualisiert werden --> Verbindung hergestellt Daten gesichert --> Meldung erfolgt an TS WLAN-Verbindung ist stabil , Access-Point im selben Raum wie der ABL. Keine Störquellen im Umkreis Kunde ist informiert | SWO-424028 | Setup | SW | RMED SW | Wifi |  | Troubleshooting required | Settings and data checked Possible sources of interference were queried during IT | ASUS USB AC51 installed according to RDEU statement --> ABL writes in the database every second that it connects to KIS $LIS. or lost. --> ABL reconnected via LAN --> ABL restarted again, as the messages 0133 .0134 are still constantly updated --> Connection established Data backed up --> Message is sent to TS Wi-Fi connection is stable, access point in the same room as the ABL. No sources of interference in the vicinity Customer is informed
Yes | 1 | Y | RPAC | Blank | sensor cassette failure - 1066 temperature in sensor cassette out of range |  | D503DFEBC8055CC7 | [5 empty] | Blank | sensor cassette failure - 1066 temperature in sensor cassette out of range
No | Visit | Complaint | RAME | Aqure Connection Error | Analyzer will not connect to Aqure. Verified wireless adapter is working properly. | Reloaded windows 7 on the analyzer to correct the issue. | SWO-424239 | Defect | SW | Windows | Windows | Win 7 | Aqure Connection Error | Analyzer will not connect to Aqure. Verified wireless adapter is working properly. | Reloaded windows 7 on the analyzer to correct the issue.
Yes | 1 | Y | RPAC | Blank | Analyer set up for installaiton |  | D503C204C84BF848 | Re classified RMED | planned SWO |  |  |  | Blank | Analyer set up for installaiton
No | Visit | Complaint | RIND | Lactate calibration failed. | Lactate sensitivity in the Calibration was very low. | Flush the sensor cassette with distilled water,Kept it on fridge for 15mins.Then keep it at room temp for 10mins,then reinstall the cassette,but calibration takes almost half an hour & lactate failed again.Replace the sensor cassette with the new one & after successful calibration it became ready. | SWO-426421 | [5 empty] | Lactate calibration failed. | Lactate sensitivity in the Calibration was very low. | Flush the sensor cassette with distilled water,Kept it on fridge for 15mins. Then keep it at room temp for 10mins,then reinstall the cassette,but calibration takes almost half an hour & lactate failed again. Replace the sensor cassette with the new one & after successful calibration it became ready.
No | Visit | Complaint | RIND | Qc out of range | Qc out of range | AS customer were using Bio rad qc and for Abl 9 machine bio rad doesnt have values so suggested the customer to purchase our qc | SWO-426356 | [5 empty] | Qc out of range | Qc out of range | AS customer were using Bio rad qc and for Abl 9 machine bio rad doesnt have values so suggested the customer to purchase our qc
No | Visit | Complaint | RAME | Calibration erreur | Pco2 calibration erreur vérifier canal Pco2 ,électrode Brisée | Remplacé électrode Pco2 sur garantie | SWO-424322 | Defect | Electrode | Electrode | pCO2 | Electrode broken | Error calibration | Pco2 calibration error check pun canal Pco2 ,broken electrode | Replaced Pco2 electrode on warranty
No | 11 | Y | RPAC | Blank | pCO2 reading consistently low on QC. |  | D503DD5EC80A7B2A | Defect | HW | Gas module | Gas tubing | pCO2 low QC | Blank | pCO2 reading consistently low on QC.
No | Visit | Complaint | RDEU | 01.10.2018 08:30Uhr Herr Glienke 03334691503: Das Gerät meldet häufig Driftalarm bei Elektrolyte | spor. Lytefehle | Fehlersuche; Gerät aus; alle Ventile ausgebaut u. gereinigt; div. Kontakte d. Elektrodenkappen poliert; Oxi-Hakenventil geschmiert; Kal. = pO2-Nullstromfehler; Nullstromkal. im Serviceprogr. durchgef.; Kal. weiter Fehler; pO2 neu membraniert; Ref.-Elektrode (alt: 98-09) u. Gas2 gew.; Kal. + QC i.O. | SWO-424427 | Assisted Replacement | Electrode | Electrode | Reference |  | 01.10.2018 08:30Am Mr Glienke 03334691503: The device frequently reports drift alarm for electrolytes | Sports. Lytefehle | Troubleshooting; device off; all valves removed and cleaned; div. Contacts of electrode caps polished; Oxi hook valve lubricated; Kal. = pO2 zero current error; Zero current cal. in the service progr. through it; Kal. further errors; pO2 re-membraneed; Ref.-electrode (old: 98-09) and gas2. Kal. + QC in the first place
No | Visit | Complaint | RDAN | Dorthe skriver, at ABLen drypper periodisk. Hun har skiftet inletpakning. | Afprøvet. | Inletprobe og arm udskiftet, slutafprøvet | SWO-424175 | Defect | HW | Inlet module | Inlet probe |  | Dorthe writes that the ABLen drips periodically. She has changed the inlets. | Tested. | Inletprobe and arm replaced, final tested
No | Visit | Complaint | RIND | Printer not workinig | Printer is not givinig print out, Found that liver is movinig free. | Take out the printer from the machine, try to fix it but the liver is free and the wheel which make liver movevable is also broken. | SWO-426862 | Revisit required | HW | Printer Module | Printer Head |  | Printer not workinig | Printer is not givinig print out, Found that liver is movinig free. | Take out the printer from the machine, try to fix it but the liver is free and the wheel which make liver movevable is also broken.
No | Visit | Complaint | RAME | Customer states that Glu stopped calibrating properly despite Re-membraning the electrode. | Verified customer's stated issue. The calibration logs revealed that the electrodes sensitivity dropped very low and there were many electrode response errors. Checked all the electrodes in the El $Met chamber to ensure that they all had their respective sealing rings. All the electrodes had their sealing rings. Removed the Glu electrode and cleaned it using Cal 1 solution and re-membraned the electrode. Removed the Glu electrode connector and discovered a few synthetic strands jammed by the connectors contacts. | Cleaned the Glu electrode connector thoroughly, replaced the Cal 1 solution and Rinse solution that were very low and performed several 2 point calibrations. | SWO-424263 | Cleaned | HW | El/Met Module | Electrode Connector | Glu | Customer states that Glu stopped calibrating properly despite Re-membraning the electrode. | Verified customer's stated issue. The calibration logs revealed that the electrodes sensitivity dropped very low and there were many electrode response errors. Checked all the electrodes in the El /Met chamber to ensure that they all had their respective sealing rings. All the electrodes had their sealing rings. Removed the Glu electrode and cleaned it using Cal 1 solution and re-membraned the electrode. Removed the Glu electrode connector and discovered a few synthetic strands jammed by the connectors contacts. | Cleaned the Glu electrode connector thoroughly, replaced the Cal 1 solution and Rinse solution that were very low and performed several 2 point calibrations.
No | Visit | Complaint | RAME | Patient results not sending | Checked all data logs | Could not find analyzer issues so I called I.T.who reset port and system now normal | SWO-425239 | Defect | SW | External SW | His/Lis |  | Patient results not sending | Checked all data logs | Could not find analyzer issues so I called I.T.who reset port and system now normal
No | Visit | Complaint | RDAN | Ventil fejl? | Ventil tjek | Oxi ventil skiftet Software opdateret til V. 3.4 | SWO-424504 | Defect | HW | Oxi module | Pinch valve | 1063 | Valve error? | Valve Check | OXI Valve Switch Software updated to V. 3.4
No | Visit | Complaint | RKK | OXI項目測定エラー | 測定時スペクトルミスマッチ他 OXI項目エラー多発 | ヘモライザー、スペクトロメーター、ピンチバルブ各交換、 | SWO-429447 | Defect | HW | Oxi module | Hemolyzer |  | OXI item measurement error | Measurement spectral mismatch and oxi item error frequent | Hemolyzers, spectrometers, pinch valve supts,
No | Visit | Complaint | RSAF | Customer installed 2nd Sensor cassette but still Na not calibrating, | Checked sample path for blockages and inlet probe and gasket for any defects with no errors found, checked solution and found that it's not activated correctly. | Correctly activated solution pack, 2 x calibration ok and all QC's ok. | SWO-424577 | User Error | SP | SP | SP | Activation | Customer installed 2nd Sensor cassette but still Na not calibrating, | Checked sample path for blockages and inlet probe and gasket for any defects with no errors found, checked solution and found that it's not activated correctly. | Correctly activated solution pack, 2 x calibration ok and all QC's ok.
No | Visit | Complaint | RDEU | 01.10.2018 11:20Uhr Frau Hermann 02037330: pCO2-Fehler bei QC Level3 | CO2 Fehler beim QC Level C, Lösungspack nicht o.k. | LP ersetzt, neues LP BS 55, Gerät kalibriert und QC durchgeführt, | SWO-424630 | Assisted Replacement | SP | SP | pCO2 | QC C | 01.10.2018 11:20Uhr Frau Hermann 02037330: pCO2-Fehler bei QC Level3 | CO2 Fehler beim QC Level C, Lösungspack nicht o.k. | LP replaced, new LP BS 55, device calibrated and QC performed,
No | Visit | Complaint | RESP | El equipo no arranca. Parece que la fuente de alimentación no funciona. | Sustitucion power supply | Se sustituye power supply y al arrancar solo enciende PC y esta sin linea con wet section, se chequea PCB y se detecta F1 fundido sustituyendose este y volviendose a fundir. No se pueden hacer mas pruebas por falta de material. Se prepara equipo para reparar en nuestros talleres de Madrid. | SWO-424609 | Revisit required | Analyzer removed | Analyzer removed | Analyzer removed |  | The equipment won't start. Looks like the power supply isn't working. | Power supply substitution | Power supply is replaced and when booting only turns on PC and it is unlined with wet section, PCB check is checked and fused F1 is detected replacing it and melting again. No further testing can be done for lack of material. Equipment is prepared to repair in our workshops in Madrid.
No | Visit | Complaint | RLTD | Error reported by Linda Stocks on arrival to site. | Checked inlet probe, perfectly straight. Inlet gasket ok, as new fitted by customer at around 12:30 today. Performed manual flush to check sample pathway, all clear. Checked solution pack activated correctly , yes. Tested solutions in service now passed. Restarted analyser, all ok except pO2 failing QC level C. Advised customer to replace Solution Pack. | No fault found with analyser. Recommended replacing Solution Pack due to pO2 QC error. | SWO-424826 | No fault found | SP | SP | pO2 | QC C | Error reported by Linda Stocks on arrival to site. | Checked inlet probe, perfectly straight. Inlet gasket ok, as new fitted by customer at around 12:30 today. Performed manual flush to check sample pathway, all clear. Checked solution pack activated correctly , yes. Tested solutions in service now passed. Restarted analyser, all ok except pO2 failing QC level C. Advised customer to replace Solution Pack. | No fault found with analyser. Recommended replacing Solution Pack due to pO2 QC error.
Yes | Visit | Complaint | RLTD | Peter Williams Tel: 01243 788122 x: 33591 - please call after 1pm | Cassette sensor interface pcb replacement complete. | Sensor interface pcb replaced, calibration and qcs passed. Interface pcb will be sent back as per tech line request. | SWO-424575 | Re classified RMED | planned SWO |  |  | Techline: RLTD 820636942 | Peter Williams Tel: 01243 788122 x: 33591 - please call after 1pm | Cassette sensor interface pcb replacement complete. | Sensor interface pcb replaced, calibration and qcs passed. Interface pcb will be sent back as per tech line request.
No | Visit | Complaint | RLTD | Battery indicator showing 0% but working ok | Battery indicator 0% | Removed battery module and reset jumper, still no good Replaced battery module with new Key Operator training carried out with perfusionists 1 $10 $18 | SWO-424601 | Defect | HW | Battery module | Battery module | 0% | Battery indicator showing 0% but working ok | Battery indicator 0% | Removed battery module and reset jumper, still no good Replaced battery module with new Key Operator training carried out with perfusionists 1 $10 $18
Yes | 1 | Y | RPNZ | Blank | Preinstall check out and set up on site in Waikato POCT office - destined for install at Waikato biochemistry laboratory |  | D508094CA018E791 | Re classified RMED | planned SWO |  |  |  | Blank | Preinstall check out and set up on site in Waikato POCT office - destined for install at Waikato biochemistry laboratory
Yes | 1 | Y | RPNZ | Blank | Preinstall set up and check out on site in POCT office - destined for Tokoroa |  | D5080B73A05BE794 | Re classified RMED | planned SWO |  |  |  | Blank | Preinstall set up and check out on site in POCT office - destined for Tokoroa
Yes | 1 | Y | RPNZ | Blank | Presintall check out and set up at Waikato POCT office destined for Thames site. |  | D50B526D8C611C8F | Re classified RMED | installation report |  |  | Move to OOB trending | Blank | Presintall check out and set up at Waikato POCT office destined for Thames site.
Yes | 2 | Y | RPNZ | Blank | Preinstall check out and set up on site in Waikato POCT office - destined for install at Taumarunui |  | D50DA6797829F153 | Re classified RMED | installation report |  |  | Move to OOB trending | Blank | Preinstall check out and set up on site in Waikato POCT office - destined for install at Taumarunui
No | Visit | Complaint | RAME | Possible fluid blockage errors and purple solution is stuck in cassette window. | Found analyzer would not pump solution one. | Checked manifold, pump and tubing. All checked okay. Reseated solution pack several times before it would pump solution one. | SWO-424904 | [5 empty] | Possible fluid blockage errors and purple solution is stuck in cassette window. | Found analyzer would not pump solution one. | Checked manifold, pump and tubing. All checked okay. Reseated solution pack several times before it would pump solution one.
No | Visit | Complaint | RAME | Leaking from below solution pack. | Analyzer leaking from below solution pack. | Cleaned leakage from the analyzer and counter top. Flushed the solution pack waste port. PO2 reading above insert limits on Qualicheck 5+. Replaced the sensor cassette. | SWO-424903 | [5 empty] | Leaking from below solution pack. | Analyzer leaking from below solution pack. | Cleaned leakage from the analyzer and counter top. Flushed the solution pack waste port. PO2 reading above insert limits on Qualicheck 5+. Replaced the sensor cassette.
No | Visit | Complaint | RDEU | 01.10.2018 12:25Uhr Herr Kräuter 022678897048: Das Gerät meldet bei der QC-Messung fast immer Oxi-Spectrumfehler | Dateien gesichtet | Dateien gesichtet, Oxy Modul ersetzt, Oxy complete test, tHb Kalibration Qualitätskontrolle | SWO-424534 | Defect | HW | Oxi module | Oxi module | 581 | 01.10.2018 12:25pm Mr Herbs 022678897048: The device almost always reports Oxi spectrum error during QC measurement | Files spotted | Files spotted, Oxy module replaced, Oxy complete test, tHb calibration quality control
Yes | Visit | Complaint | RSCH | Rep. Ref. | Immer wieder Ca und Na Drift | Ref. Elektrode ersetzt i.o Alte Lot = 6807 => Garantie Kal. und QC i.o System i.o | SWO-424928 | [5 empty] | Rep. Ref. | Again and again Ca and Na Drift | Ref. Electrode replaces i.o Old lot = 6807 => Warranty Kal. and QC i.o System i.o
No | 4 | Y | RCZE | Blank | nahlasen problem s QC |  | D504CA843E56D026 | Defect | Electrode | Electrode | Sealing ring, pH |  | Blank | nahlasen problem s QC
No | Visit | Complaint | RSCH | System blockiert | Vermehrt Corruptxml File Analysator blockiert immer wieder und muss neugestartet werden | CF Card ausgetauscht Windows neu installiert 3x DMS neu installiert Alle Daten zurückgespielt Funktionskontrolle i.o | SWO-424587 | Defect | SW | RMED SW | RMED SW | Corrupt xml | System blocked | Increases Corruptxml File Analyzer blocks again and again and needs to be restarted | CF Card replaced Windows reinstalled 3x DMS reinstalled All data played back Functional control i.o
No | Visit | Complaint | RDEU | 02.10.2018 12:50Uhr Herr Päuler 020172382715: Das Türchen für die Meßkassetten schließt nicht komplett | Cartridgedoor öffnet und schließt nicht: Eine Feder ist von der Tür ausgehakt. | Feder an Cartridgedoor befestigt. Funktionskontrolle: I.O. | SWO-425163 | Adjustment | HW | Cover and Chassis | Door, Storage | Spring | 02.10.2018 12:50pm Mr. Päuler 020172382715: The door for the measuring cassettes does not close completely | Cartridgedoor does not open or close: A spring is hooked out of the door. | spring attached to Cartridgedoor. Functional control: I.O.
Yes | Visit | Complaint | RFRA | Nous avons rencontré un problème potentiellement grave avec l'AQT90 le samedi 29 septembre. L'AQT90 installé en dernier aux urgences le 11/07/2018 et qui est censé remplacer celui de l'institut de cardiologie en octobre a comme numéro de série R048N001. Ce samedi dernier, nous avons rendu une troponine T à 2900ng/L sur celui-ci et le patient a été orienté en cardiologie. Le cardiologue demande une 2ème troponine T qui sort <10 ng/L sur l'AQT90 de l'Institut de cardiologie. Cette ambiguité a crée un doute auprès du médecin des urgences qui a demandé à la technicienne de repasser le premier prélèvement sur le premier AQT (installé au tout début du protocole) et le résultat sort <10 ng/L. Le biologiste face à cette discordance demande d'envoyer le prélèvement au laboratoire et le résultat sort négatif sur le cobas... Nous avons tout vérifié sur les CQ du jour qui sont bons sur l'AQT en question: - le 29/09/18: TnT-14343=150 (120-184) sur MCC2 et TnT-14343 = 5040 (4270-5890) sur MCC3 - le 30/09/18: TnT-14343=30 (19.4-38.4) sur MCC1 et TnT-14343= 4990 (4270-5890) sur MCC3 Le dimanche 30/09/18 vers 18h30 ce même AQT demande un redémarrage 3 fois et au final non opérationnel et nous nous sommes retrouvés entrain de charger en toute urgence le seul 1er AQT qu'on a... Pourriez vous nous apporter des explications concernant le problème qu'on a eu ce samedi svp et surtout demander une intervention en toute urgence sur cet automate qui n'est pas opérationnel depuis hier soir?? P.S: Je refuse que cet automate soit mis à l'Institut de cardiologie sans révision complète et sérieuse ou même remplacement par un autre. | Contrôle du cup wash unit, du needle wash unit, de l’aiguille, de l’incubator wheel, de la process wheel | Nettoyage du filtre, de la process wheel. Récupération du service dump. Essais | SWO-424555 | Re classified RMED | planned SWO |  |  | Techline: RFRA 823475368 | We encountered a potentially serious problem with the AQT90 on Saturday, September 29. The AQT90, which was last installed in the ER on 11/07/2018 and is supposed to replace the Heart Institute in October, has the serial number R048N001 as its serial number. This past Saturday, we returned a T troponin to 2900ng/L on it and the patient was referred to cardiology. The cardiologist asks for a 2nd troponin T that comes out on the AQT90 of the Heart Institute. This ambiguity created doubt among the emergency physician who asked the technician to retake the first sample on the first AQT (installed at the very beginning of the protocol) and the result comes out '10 ng/L' The biologist in the face of this discordance asks to send the sample to the laboratory and the result comes out negative on the cobas... We checked everything on the CQ of the day which are good on the AQT in question: - 29/09/18: TnT-14343-150 (120-184) on MCC2 and TnT-14343 - 5040 (4270-5890) on MCC3 - 09/30/18: TnT-14343-30 (19.4-38.4) on MCC1 and TnT-14343 4990 (4270-5890) on MCC3 On Sunday 30/09/18 around 6:30 pm the same AQT asks for a restart 3 times and in the end non-operational and we found ourselves urgently loading the only 1st AQT that we have... Could you explain the problem we had this Saturday please and especially request an emergency intervention on this automaton that is not operational since last night?? P.S: I refuse to allow this automaton to be put at the Heart Institute without a complete and serious overhaul or even replacement by another. | Control of the cup wash unit, the needle wash unit, the needle, the incubator wheel, the process wheel | Cleaning the filter, the process wheel. Recovery of the dump service. Tests
No | Visit | Complaint | RPOL | Klient zgłosił potrzebę częstej wymiany uszczelki wlotu | Weryfikacja pracy analizatora | Zakrzywiona igła aspirująca - wymiana, dodatkowo wymiana uszczelki wlotu | SWO-424861 | Defect | HW | Auto Inlet module | Probe bend |  | Customer reported the need for frequent replacement of the inlet gasket | Validate Analyzer Work | Curved aspirating needle-replacement, additional inlet gasket replacement
No | Visit | Complaint | RSCH | Re. Einlass | Einlassdichtung springt nach Messung nicht (oder nur langsam) nach vorne zurück Mit Kunde am telefon Enlasshalter ersetzt, ohne Erfolg Einlass-Modul ebenfalls mit Kunde am Tel. ausgebaut und wieder eingebaut, ohne Erfolg | Dichtung von Einlassdichtungshalter war nicht richtig eingesteckt Dichtung richtig montiert i.o Testmessung i.o System i.o | SWO-425069 | User Error | HW | Inlet module | Inlet gasket |  | Re. Intake | Inlet seal jumps after measurement not (or slowly) back to the front Replaced with customer on the phone enlass holder, to no avail Entrance module also extended with customer on tel. and reinstalled, to no avail | Sealing of inlet seal holder was not plugged in properly Seal correctly mounted i.o Test measurement i.o System i.o
No | Visit | Complaint | RAME | Printer has become inoperative 1-2 times per week. | Printer errors are not showing in activity logs. Frequent LIS errors are occurring however - 0131 failure to send packet connection 1 | Removed auto sent from connection 1 (not being used anyway). Customer to monitor and advise after 7-14 days if printer issues have re-occurred. | SWO-425256 | Adjustment | SW | SW | His/Lis | printer | Printer has become inoperative 1-2 times per week. | Printer errors are not showing in activity logs. Frequent LIS errors are occurring however - 0131 failure to send packet connection 1 | Removed auto sent from connection 1 (not being used anyway). Customer to monitor and advise after 7-14 days if printer issues have re-occurred.
No | Visit | Complaint | RFRA | disk boot failure | .... | Changement de la carte CF OK $ $ réinstallation complète Windows 7 et application version 3.4 ok $ $ configuration Windows et programme ok $ $ chargement et sauvegarde toutes données et configuration ok $ $ la manipulation demandée pour résoudre le problème de lactate n’a pas fonctionné. | SWO-425107 | Defect | HW | DMS_PC | CF card |  | disk boot failure | .... | Change of the CF OK card $$ full windows 7 reinstallation and application version 3.4 ok $Windows configuration and program ok $$ loading and backup all data and configuration ok $$ the manipulation requested to solve the problem of lactate n 'didn't work.
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 气泡探测定标失败，CO2，PH,K,NA,CA,CL,LAC定标不过，验证未执行 1117 | 更换3张测试卡2个试剂包后修复 | SWO-426570 | [5 empty] | Test card, reagent package failure | Bubble detection calibration failure, CO2, PH, K, NA, CA, CL, LAC calibration, verification did not perform 1117 | Repair after replacing 3 test cards 2 kits
No | Visit | Complaint | RAME | Na+ QC CV's are high | Na+ QC CV's are high since August when instruments came online. Instrument inspection shows no sign of abnormal operation. | Replaced Na+ and reference electrode's. Calibrated and ran QC 5 times overnight. QC data plot is improved and stable. Instrument operating normally. | SWO-425616 | User expectation | Electrode | Electrode | Electrode | Na QC high CV | Na+ QC CV's are high | Na+ QC CV's are high since August when instruments came online. Instrument inspection shows no sign of abnormal operation. | Replaced Na+ and reference electrode's. Calibrated and ran QC 5 times overnight. QC data plot is improved and stable. Instrument operating normally.
No | Visit | Complaint | RAME | Na+ QC CV's are high | Na+ QC CV's are high since August when instruments came online. Instrument inspection shows no sign of abnormal operation. | Replaced Na+ and reference electrode's. Calibrated and ran QC 5 times overnight. QC data plot is improved and stable. Instrument operating normally. | SWO-425617 | User expectation | Electrode | Electrode | Electrode | Na QC high CV | Na+ QC CV's are high | Na+ QC CV's are high since August when instruments came online. Instrument inspection shows no sign of abnormal operation. | Replaced Na+ and reference electrode's. Calibrated and ran QC 5 times overnight. QC data plot is improved and stable. Instrument operating normally.
No | Visit | Complaint | RKK | Na QC範囲外 | Na QC範囲外 メンブラン交換しても変化が見られない | Ref電極交換 | SWO-425470 | Assisted Replacement | Electrode | Electrode | Reference | Na QC | Na QC out of range | Na QC out of range There is no change even if I change the membrane. | Ref electrode replacement
No | Visit | Complaint | RAME | Leak Currents Detected. | Leak current detected. Problem due to leaking inlet gasket due to broken nipple. | Clean up spill. Replace flap detector PCB. | SWO-425104 | Defect | HW | Inlet module | Nipple kit | Leak Current? | Leak Currents Detected. | Leak current detected. Problem due to leaking inlet gasket due to broken nipple. | Clean up spill. Replace flap detector PCB.
No | Visit | Complaint | RAME | Neon lamp intensity too low. All oxi parameters questioned. | Via service program found Neon lamp intensity >50% failure. | Replaced lamp unit. Performed necessary adjustments including THB service cal. | SWO-425147 | Defect | HW | Oxi module | Lamp unit | Neon Lamp | Neon lamp intensity too low. All oxi parameters questioned. | Via service program found Neon lamp intensity >50% failure. | Replaced lamp unit. Performed necessary adjustments including THB service cal.
No | Visit | Complaint | RIND | Instrument showing error and calibration aborted. | Found wash unit spring in reverse position, it is hitting to inlet base plate when it is moving down. | Corrected the wash unit spring and checked the process unit movements.found all movements fine without obstruction. found instrument working fine. | SWO-426389 | User Error, FSE | HW | Process Module | Wash Unit | Spring | Instrument showing error and calibration aborted. | Found wash unit spring in reverse position, it is hitting to inlet base plate when it is moving down. | Corrected the wash unit spring and checked the process unit movements.found all movements fine without obstruction. found instrument working fine.
Yes | Visit | Complaint | RLTD | Mohammed 01895828757 | Battery error 1559 one or more cells failing. | Replaced faulty battery module and test ok. | SWO-425533 | [5 empty] | Mohammed 01895828757 | Battery error 1559 one or more cells failing. | Replaced faulty battery module and test ok.
No | Visit | Complaint | RLTD | Muhammad on 01895 828757 | Not printing results and logging off to quick. Analyser had been rebooted. | Found log off time set to 1 minuets. Reset to 3 minuets . Patient sample run and results printed out ok. Check printer connections and remade connections ok. Patient sample run and results printed ok. | SWO-425534 | Adjustment | SW | RMED SW | RMED SW |  | Muhammad on 01895 828757 | Not printing results and logging off to quick. Analyser had been rebooted. | Found log off time set to 1 minuets. Reset to 3 minuets . Patient sample run and results printed out ok. Check printer connections and remade connections ok. Patient sample run and results printed ok.
Yes | 6 | Y | RPNZ | Blank | Set up and installation on site |  | D50B50268C381C8C | Re classified RMED | installation report |  |  | Move to OOB trending | Blank | Set up and installation on site
No | Visit | Complaint | RLTD | Analyser had been dropped and smashed printer | Printer damaged | Printer replaced and tested. | SWO-425541 | User Error | HW | Printer Module | Printer | Analyzer dropped | Analyser had been dropped and smashed printer | Printer damaged | Printer replaced and tested.
No | Visit | Complaint | RLTD | Issue reported by Haval Ozgun. | Checked sample success rate on analyser since setup corrected on 02 $10 $2018. 171 samples run. 149 samples ok, 87%. 13 aborted, 7.6%. Found some ABL8xx clot catchers near the analyser which were contributing to failures. These clot catchers have been removed. 9 samples with sample transport errors, 5.3%. Compared sample success rate with old ABL835 analyser for the month of September. 2143 samples run. 1926 samples ok, 90%. 112 aborted, 5.2%. 105 samples with sample transport errors, 4.9%. Similar rates of sample failures between new ABL90 and old ABL835, ABL90 analyser performance is ok. | No fault found with analyser. ABL8xx clot catchers removed. Follow up visit to made in approximately 2 weeks time to further check sample success rate. | SWO-425680 | User Error | preanalytical | preanalytical | capillary | Clot catcher ABL8XX | Issue reported by Haval Ozgun. | Checked sample success rate on analyser since setup corrected on 02 /10 /2018. 171 samples run. 149 samples ok, 87%. 13 aborted, 7.6%. Found some ABL8xx clot catchers near the analyser which were contributing to failures. These clot catchers have been removed. 9 samples with sample transport errors, 5.3%. Compared sample success rate with old ABL835 analyser for the month of September. 2143 samples run. 1926 samples ok, 90%. 112 aborted, 5.2%. 105 samples with sample transport errors, 4.9%. Similar rates of sample failures between new ABL90 and old ABL835, ABL90 analyser performance is ok. | No fault found with analyser. ABL8xx clot catchers removed. Follow up visit to made in approximately 2 weeks time to further check sample success rate.
Yes | Visit | Complaint | RIND | pH & PCO2 error in calibration | Checked zeroing and found correct. | Arranged new SP from CP's office. Replaced SP, 2 point calibration performed successfully. | SWO-426351 | [5 empty] | pH & PCO2 error in calibration | Checked zeroing and found correct. | Arranged new SP from CP's office. Replaced SP, 2 point calibration performed successfully.
No | Visit | Complaint | RHUN | Bemeneti elszívó csonk eltört | Bemeneti elszívó csonk eltört | Alkatrész cseréje, Membráncserék (Ca, K, Cl), Kalibrációk, QC mérés | SWO-425374 | Defect | HW | Inlet module | Nipple kit | For inlet tubing | Inlet exhaust manifold broken | Inlet exhaust manifold broken | Replacement of parts, membrane substitutions (Ca, K, Cl), calibration, QC measurement
No | Visit | Complaint | RDEU | 20.09.2018 8:40Uhr Herr Adepar 0621 383-3528: Ständig Lac und Cl-Fehler 04.10.2018 Herr Adebar 06213833528 meldet, das weiterhin Glu/Lac Fehler bestehen. | KAL Driften Lactat und Glucose Reinigung wird häufig abgebrochen | Überprüfung Messkammer und Ventile - Erneuerung Wartungskit (siehe Case) - Memranierung und mehrfache Kalibrationen (Lactat mit Driften und Instabilität $ Empfindlichkeitswert schwankend) - Erneuerung Lactat Elektrode - Funktionsüberprüfung etc. | SWO-425686 | Assisted Replacement | Electrode | Electrode | Lac | Drift | 20.09.2018 8:40h Mr Adepar 0621 383-3528: Constant Lac and Cl errors 04.10.2018 Mr. Adebar 06213833528 reports that there are still Glu/Lac errors. | KAL Drifting Lactate and Glucose Cleaning is often cancelled | Checking measuring chamber and valves - Renewal maintenance kit (see case) - memranation and multiple calibrations (lactate with drifts and instability - sensitivity value fluctuating) - Renewal lactate electrode - function check etc.
No | Visit | Complaint | RAME | System clot detected after flush and cassette /solution pack replacement | System errors seem to be related to pinch valve | Lubricated valve and cleaned excess with caned air. | SWO-426309 | Cleaned | HW | Oxi module | Pinch valve | Clot | System clot detected after flush and cassette /solution pack replacement | System errors seem to be related to pinch valve | Lubricated valve and cleaned excess with caned air.
No | Visit | Complaint | RDEU | 08.08.2018 10:20 Fr. Judäa 06217591319: Interventionsmeldung Flüssigkeiten können nicht gepumpt werden 09.08.2018, 11:45 Uhr ==> Anrufer Fr. Bianka Judäa, Tel.: 0621759131 ==> Analysator befindet sich weiterhin in der Fehlerbehebungsroutine 13.08.2018, 14.20Uhr, Frau Judea, Tel: 06217591319: weiterhin Wechselschleife- die Anwender haben bereits alle getauscht | Intervention erforderlich Lösungen können nicht gepumpt werden | Allgemeine Überprüfung - Analysator gespült - Erneuerung Einlassdichtung, Anschlussdichtung und Einlassnadel - Quetschventil geschmiert - Erneuerung Schlauch an Quetschventil - Funktionsüberprüfung etc. (plus einmal Einlassdichtung als Ersatz) Bemerkung : KAL und QC-Fehler, da Sensor defekt (ist Anwenderin bekannt) Inhomogene Kal (2 $3) Lösungen wird gemeldet. 1539 $1538. Nach Austausch von neuer Anschlussdichtung wird neuer Sensor eingesetzt. Warte auf Rückmeldung von Anwender nach Sensortausch. Wenn weiterhin Meldung: Inhomogene Kal Lösungen $ Lösung kann nicht gepumpt werden auftritt bittet Kunde um Gerätetausch!!!! | SWO-425993 | Defect | HW | Basic Unit | PM Kit |  | 08.08.2018 10:20 Fr. Judea 06217591319: Intervention notification Liquids cannot be pumped 09.08.2018, 11:45 ==> Caller Ms. Bianka Judea, tel.: 0621759131 ==> Analyzer is still in the troubleshooting routine 13.08.2018, 14.20h, Mrs. Judea, Tel: 06217591319: continue alternating loop - the users have already changed all | Intervention needed Solutions cannot be pumped | General inspection - Analyzer rinsed - Renewal inlet seal, connection seal and inlet needle - Lubricated crushing valve - Renewal hose to crush valve - function check etc. (plus once inlet seal as replacement) Note : KAL and QC error because sensor is defective (known to user) Inhomogeneous Kal (2 x 3) solutions are reported. 1539 .1538. After replacing the new connection seal, a new sensor is used. Waiting for feedback from users after sensor replacement. If still message: Inhomogeneous Kal Solutions can't be pumped, solution can't be pumped asks customer to exchange equipment!!!!
No | Visit | Complaint | RFRA | Perte du lecteur réseau et du chemin de destination pour la sauvegarde automatique | Aucune | Essais remise en place lecteur réseau. Impossible d'enregistrer les modifs. Nouvelle inter prévu la semaine prochaine pour re instal complète. | SWO-426292 | Revisit required | HW | DMS_PC | CPU |  | Loss of network drive and destination path for automatic backup | No | Tests re-establish network reader. Can't save the changes. New inter scheduled for next week for full instal.
Yes | Visit | Complaint | RDEU | 04.10.2018 13:55Uhr Frau Eilers-Gladrow 05315952045: Bei vielen Geräten zeigt sich der selbe Fehler - pCO2 QC-Verletzung im LevelA und C | pCO2 Messwerte oberhalb der QC Grenzen im Level A $C | Datensicherung erstellt | SWO-425908 | Re classified RMED | planned SWO | data collection | data collection |  | 04.10.2018 13:55PM Ms Eilers-Gladrow 05315952045: Many devices show the same error - pCO2 QC violation in LevelA and C | pCO2 readings above the QC limits in level A /C | Data backup created
No | Visit | Complaint | RIND | Service Visit. Doctors having problem with the measurement of pCO2, and its calculated values. | Check the analyzer, observe the analyzer, all the parameters are green and analyzer status is also green. Need to analyze the same patient sample on another ABL800 analyzer. | Need to analyze the sample on another ABL800 analyzer. QC is also passed. | SWO-430068 | Revisit required | HW | pH/BG Module | pH/BG Module |  | Service Visit. Doctors having problem with the measurement of pCO2, and its calculated values. | Check the analyzer, observe the analyzer, all the parameters are green and analyzer status is also green. Need to analyze the same patient sample on another ABL800 analyzer. | Need to analyze the sample on another ABL800 analyzer. QC is also passed.
No | Visit | Complaint | RKK | Na QC範囲外 | Na QC範囲外 電解質もばらつきあり | 比較電極交換 | SWO-425975 | Assisted Replacement | Electrode | Electrode | Reference |  | Na QC out of range | Na QC out of range Electrolytes are also varied. | Comparative electrode replacement
No | Visit | Complaint | RDEU | 19.09.2018 Herr Magira 059311524482 meldet, das der QC Wert für Trop. seit der Installation ansteigt. | MCC3 Lot 09163 TnT außerhalb Grenzen | Gerät überprüft, MCC3 Lot 01183 eingelesen und vermessen | SWO-426172 | Defect | Consumables | QC | QC, TnT |  | 19.09.2018 Mr Magira 059311524482 reports that the QC value for Trop. has increased since installation. | MCC3 Lot 09163 TnT outside limits | Device checked, MCC3 Lot 01183 read in and measured
No | Visit | Complaint | RPOL | Ekran dotykowy - uszkodzony | Rozmowa z klientem, weryfikacja | Wymiana ekranu, testowanie, kalibracja | SWO-426044 | Defect | HW | DMS_PC | Touch screen |  | Touchscreen-damaged | Customer interview, Verification | Screen replacement, testing, calibration
No | Visit | Complaint | RCHN | 接口板腐蚀 | 接口板腐蚀,需要借件902-972和903-334各一个 |  | SWO-426445 | [5 empty] | Interface plate corrosion | Interface plate corrosion, need to borrow 902-972 and 903-334 each one
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 定标失败 | 更换测试卡 | SWO-427616 | Assisted Replacement | SC | SC | SC |  | Test card, reagent package failure | Bid failed | Replace the test card
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | CO2灵敏度超范围 0378 | 更换测试卡后修复 | SWO-427310 | [5 empty] | Test card, reagent package failure | CO2 sensitivity over range 0378 | Repair after replacing test card
No | Visit | Complaint | RCHN | 光学单元故障 | 光学单元故障，需要借件903-364一个，需要借件944-212（14642）944-216（14696）944-220（14429）944-258（14738）944-394（14500） | 更换903-364和重新定标后，定标仍有问题，需要加借定标卡944-214（14583） by Andrew | SWO-425911 | Defect | HW | Process Module | Optical Unit | mail: error code? | Optical unit failure | Optical unit failure, need to borrow 903-364 one, need to borrow 944-212 (14642) 944-216 (14696) 944-220 (14429) 944-258 (14738) 944-394 (14500) | After replacing 903-364 and re-marking, there is still a problem with the calibration and the bid card 944-214 (14583) by Andrew
Yes | Visit | Complaint | RCHN | 光学单元故障 | 光学单元故障，需要加借944-214（14583） |  | SWO-426447 | Re classified RMED | same day upSWO |  |  |  | Optical unit failure | Optical unit failure, need to borrow 944-214 (14583)
No | Visit | Complaint | RCHN | 存储盘卡盘，报错5784 | 存储盘卡盘，报错5784 | 重新安装测试卡后修复 by Andrew | SWO-426380 | Defect | Consumables | Test Kit | Cartridge, Broken | 5784 | Memory disk card, error 5784 | Memory disk card, error 5784 | Fix by Andrew after reinstalling the test card
No | Visit | Complaint | RIND | Service Visit. Doctors having problem with the measurement of pCO2, and its calculated values. | Need to replace the chamber and Oxi module of the analyzer. | Replace the pH $BG full module and Oxi module of the analyzer, upgrade the software version of the analyzer, successfully done, now compare and cross verify the sample in both the analyzer, approximate same reading is arriving and doctors now satisfied with their own analyzer reading, successfully completed the task, analyzer handed over to staff in well working condition. | SWO-430072 | Defect | HW | pH/BG Module | pH/BG Module | Oxi module | Service Visit. Doctors having problem with the measurement of pCO2, and its calculated values. | Need to replace the chamber and Oxi module of the analyzer. | Replace the pH $BG full module and Oxi module of the analyzer, upgrade the software version of the analyzer, successfully done, now compare and cross verify the sample in both the analyzer, approximate same reading is arriving and doctors now satisfied with their own analyzer reading, successfully completed the task, analyzer handed over to staff in well working condition.
No | Visit | Complaint | RIND | Calibrate_13 | Checked and found dust accumulated inside AQT90. | Performed thorough cleaning of AQT90. Restarted the instrument. | SWO-426487 | Cleaned | HW | Storage Module | sensor, cup | calibrate_13 | Calibrate_13 | Checked and found dust accumulated inside AQT90. | Performed thorough cleaning of AQT90. Restarted the instrument.
No | Visit | Complaint | RLTD | Mario Dumic Tel: 0208 296 4772 | Checked operation of oximetry pinch valve and condition of pinch tube, both ok. Carried out repeat tHb calibration with customer for training. tHb calibration passed, analyser all ok. | Repeated failed tHb calibration. | SWO-427196 | User Error | Training | Training | tHb Cal |  | Mario Dumic Tel: 0208 296 4772 | Checked operation of oximetry pinch valve and condition of pinch tube, both ok. Carried out repeat tHb calibration with customer for training. tHb calibration passed, analyser all ok. | Repeated failed tHb calibration.
No | Visit | Complaint | RDEU | LP steckt im Gerät fest | Lösungspack verhakt | Lösungspack entriegelt und richtig aktiviert, Lösungspack eingesetzt Startphase, | SWO-427406 | User Error | SP | SP | SP | Activation | LP is stuck in the device | Solution pack hooked | Solution pack unlocked and properly activated, solution pack used start phase,
No | Visit | Complaint | RAME | Display black. Analyzer does nothing. | Display black. Analyzer does nothing. Analyzer goes through Bios and screen goes black, no activity. Tried replacing CF card and reinstall software. Win 7 hanger at beginning of installation. CPU is suspect. | Replaced CPU. Install original CXF card. Flash BIOS on CPU. All good after that. Had to replaced SC as it has been sitting since last Friday. Sent credit form for SC to Amy. She will process. | SWO-427081 | Defect | HW | DMS_PC | CPU |  | Display black. Analyzer does nothing. | Display black. Analyzer does nothing. Analyzer goes through Bios and screen goes black, no activity. Tried replacing CF card and reinstall software. Win 7 hanger at beginning of installation. CPU is suspect. | Replaced CPU. Install original CXF card. Flash BIOS on CPU. All good after that. Had to replaced SC as it has been sitting since last Friday. Sent credit form for SC to Amy. She will process.
No | Visit | Complaint | RDEU | Hotline Einsatz-Gerät steht im Intervention erforderlich | Hotline Einsatz-Gerät steht im Intervention erforderlich | Nadel und Einlassdichtung aus Kundenbestand ersetzt, Fluid Transport Service-Programme durchgeführt, nach Wechsel alles OK, Kal, PCO2 im Level 3 zu hoch, SC nochmals installiert, Kal, QC Level 3 erneut zu hoch, auch im Level 1 an oberer Grenze, Wechsel LP, Kal, QC erneut PCO2-Fehler in der QC, Wechsel SC (LotIn Absprache mit Kunden Fehlerfreie Kal und QC nicht abgewartet, Analysator läuft über Nacht | SWO-427181 | Defect | HW | Inlet module | Inlet gasket |  | Hotline deployment device is required in intervention | Hotline deployment device is required in intervention | Needle and inlet seal replaced from customer stock, Fluid Transport service programs carried out, after changing everything OK, Kal, PCO2 in level 3 too high, SC reinstalled, Kal, QC Level 3 again too high, also in level 1 at upper limit, change LP, Kal, QC again PCO2 error in the QC, Change SC (LotIn agreement with customers error-free Kal and QC not waited, analyzer runs overnight
No | Visit | Complaint | RDEU | Sensor Bay geht nur halb zu. | Sensor Bay war fehlerhaft justiert | Wet Section ausgebaut. Justage der Sensor Bay durchgeführt, und mit Sensor Justagetool überprüft. Startphase abgewartet Kalibration und QC durchgeführt und in Ordnung | SWO-426956 | Adjustment | HW | Basic Unit | Sensor door |  | Sensor Bay is only half-closed. | Sensor Bay was incorrectly adjusted | Wet Section. Adjustment of the Sensor Bay performed, and checked with Sensor Adjustment Tool. Start-up phase awaited Calibration and QC carried out and in order
Yes | 1 | Y | RPAC | Blank | ABL90 Flex Pre-Delivery Checklist has been completed on this installtion as per Radiomater Protocol. |  | D50C84D8020137BA | [5 empty] | Blank | ABL90 Flex Pre-Delivery Checklist has been completed on this installtion as per Radiomater Protocol.
No | Visit | Complaint | RIND | Verification Check Failed | Verification Check failed | Check zeroinig ,O2,Lac Unstable, need to change sensor board , signal data board. | SWO-426925 | [5 empty] | Verification Check Failed | Verification Check failed | Check zeroinig ,O2,Lac Unstable, need to change sensor board , signal data board.
No | Visit | Complaint | RAME | SC verification check failed | Troubleshot to a faulty analog board or sensor connector board. | Replaced the analog board and sensor connector board. | SWO-427236 | [5 empty] | SC verification check failed | Troubleshot to a faulty analog board or sensor connector board. | Replaced the analog board and sensor connector board.
No | Visit | Complaint | RCHN | 测试卡无法扫描 | 测试卡无法扫描 | 调整扫描器位置后修复 by jesson | SWO-426767 | Adjustment | HW | Storage Module | Barcode reader, storage wheel |  | Test card can't be scanned | Test card can't be scanned | Fix by jesson after adjusting scanner position
Yes | Visit | Complaint | RCHN | 测试卡损坏 5784 | 测试卡损坏 5784 | 更换测试卡 | SWO-426589 | [5 empty] | Damaged test card 5784 | Damaged test card 5784 | Replace the test card
No | Visit | Complaint | RFRA | Logiciel actif mais aucune réponse de l'automate. | Logiciel actif mais aucune réponse de l'automate. | Remplacement de la carte alimentation-->idem. Retour machine atelier. | SWO-427646 | [5 empty] | Active software but no response from the automaton. | Active software but no response from the automaton. | Replacing the power card--idem. Back machine workshop.
No | Visit | Complaint | RDEU | 08.10.2018 08:10Uhr Pfl thomas 06213831024: Nach dem Befüllen vom QC-Karussell werden nicht alle Ampullen erkannt. | AutoCheck - Erkennung fehlerhaft | fällige Jahreswartung durchgeführt | SWO-427032 | Cleaned | HW | Autocheck module | Autocheck Module |  | 08.10.2018 08:10Uhr Pfl thomas 06213831024: After filling the QC carousel, not all ampoules are detected. | AutoCheck - Detection faulty | due annual maintenance carried out
No | Visit | Complaint | RDEU | KAL Driften Na und Ca | KAL Driften Na und Ca | Allgemeine Überprüfung - Intensive Reinigung Referenz Elektrodendeckel, Kontakte und Messkammer ( eventuelle Kontamination) - Membranierung und Funktionsüberprüfung etc. Weitere Beobachtung durch Labor ! | SWO-427030 | Cleaned | HW | pH/BG Module | Electrode Connector |  | KAL Drifting Na and Ca | KAL Drifting Na and Ca | General inspection - Intensive cleaning reference electrode lid, contacts and measuring chamber (possible contamination) - membraneing and function verification etc. Further observation by laboratory !
No | Visit | Complaint | RDEU | KAL Driften Na und Ca | KAL Driften Na und Ca | Allgemeine Überprüfung - Intensive Reinigung Referenz Elektrodendeckel, Kontakte und Messkammer ( eventuelle Kontamination) - Membranierung und Funktionsüberprüfung etc. Weitere Beobachtung durch Labor ! | SWO-427029 | Cleaned | HW | pH/BG Module | Electrode Connector |  | KAL Drifting Na and Ca | KAL Drifting Na and Ca | General inspection - Intensive cleaning reference electrode lid, contacts and measuring chamber (possible contamination) - membraneing and function verification etc. Further observation by laboratory !
No | Visit | Complaint | RDEU | KAL Driften Na und Ca | KAL Driften Na und Ca | Allgemeine Überprüfung - Intensive Reinigung Referenz Elektrodendeckel, Kontakte und Messkammer ( eventuelle Kontamination) - Membranierung und Funktionsüberprüfung etc. Weitere Beobachtung durch Labor ! | SWO-427031 | Cleaned | HW | pH/BG Module | Electrode Connector |  | KAL Drifting Na and Ca | KAL Drifting Na and Ca | General inspection - Intensive cleaning reference electrode lid, contacts and measuring chamber (possible contamination) - membraneing and function verification etc. Further observation by laboratory !
No | Visit | Complaint | RDEU | 08.10.2018, 11:45 Uhr, Fr. Fritz, Tel. 0611577470: Bootfehler | Applikation startet nicht Data Upgrade Failure 9 | Datensicherung via DasiTool. Erneute Installation der Software Version 6.18, ohne Datenbankübernahme und Kommunikationsdateien. Weiterhin Statuscode 9. Roma.ini nicht gefüllt, Datei leer. Roma.ini und roma.db (+.bak) gelöscht. Applikation nochmals voll installiert. (Software wurde insgesamt 3mal installiert) Kundenkonfiguration wiederhergestellt. CCO Login Dateien wiederhergestellt. Anbindung an Aqure QC eingelesen. Kalibrationen durchgeführt QC ohne Beanstandung | SWO-427092 | Defect | SW | RMED SW | Roma.ini | 6.18 | 08.10.2018, 11:45, Fr. Fritz, Tel. 0611577470: Boat error | Application does not start Data Upgrade Failure 9 | Data backup via DasiTool. Reinstall software version 6.18, without database transfer and communication files. Continue status code 9. Roma.ini not filled, file empty. Roma.ini and roma.db (+.bak). application fully installed again. (Software has been installed 3 times in total) Customer configuration restored. CCO login files restored. Connection to Aqure QC. Calibrations carried out QC without complaint
No | Visit | Complaint | RSAF | 2pt cal result for tHb has no reading | Faulty neon lamp | Replaced Neon lamp. Lamp unit status check ok. 2pt cal and QC ok. | SWO-427233 | [5 empty] | 2pt cal result for tHb has no reading | Faulty neon lamp | Replaced Neon lamp. Lamp unit status check ok. 2pt cal and QC ok.
No | Visit | Complaint | RDEU | 09.10.2018; Hr. Kastner; Tel.: 033124134872; SC Fach schließt nicht | SC Fach schließt nicht richtig | Analysator ausgeschaltet, Nasse Section demontiert, verriegelungsmechanismus nachjustiert, montiert, Analysator eingeschaltet, SC ersetzt, da Verfallsdatum überschritten, Kal, Lac QC-Fehler, SC muss weiterhin temperieren, Kunden informiert, alle anderen Parameter Kal und QC i.O. | SWO-427891 | Adjustment | HW | Basic Unit | Sensor door |  | 09.10.2018; Mr. Kastner; Tel.: 033124134872; SC Fach does not close | SC fach does not close properly | Analyzer switched off, Wet Section disassembled, Locking mechanism readjusted, mounted, Analyzer switched on, SC replaced as expiry date exceeded, Kal, Lac QC error, SC must continue to temper, inform customers, all other parameters Kal and QC i.O.
No | Visit | Complaint | RDEU | 08.10.2018; Hr. Siebert; Tel.: 056217955811; Intervention erforderlich, Leckage erkannt 08.10.2018 13:20Uhr Herr Siebert 056217955811: Weiterhin Leckagefehler | Fehlersuche Systemstatus- und Betriebsmitteldatei überprüft Schlauchsystem gespült Ventilschlauch, Einlaßdichtung, Dichtung und Schlauch ersetzt | SP90 BS-60 durch SP90 BS-53 ersetzt neue SC90 LOT826-53 installiert Kal und QC | SWO-427632 | Defect | HW | Basic Unit | PM Kit | 1184 | 08.10.2018; Mr. Siebert; Tel.: 056217955811; Intervention required, leakage detected 08.10.2018 13:20pm Mr Siebert 056217955811: Leakage error | Troubleshooting System status and equipment file checked Hose system flushed Valve hose, inlet seal, seal and hose replaced | SP90 BS-60 replaced by SP90 BS-53 new SC90 LOT826-53 installed Kal and QC
No | Visit | Complaint | RLTD | Alex Hesse hardware failure on the oximeter 0208 6721255 bleep 7069 | Ran complete oximetry test in service, failing in Lamp Status for LED peak intensity, value 45392.00 pA [limit 80.00 - 300.00]. Replaced faulty hemolyser unit. Repeated complete oximetry test - all points passed. Performed tHb calibration. Performed quality controls - passed. | Replaced hemolyser unit. | SWO-427820 | Defect | HW | Oxi module | Hemolyzer |  | Alex Hesse hardware failure on the oximeter 0208 6721255 bleep 7069 | Ran complete oximetry test in service, failing in Lamp Status for LED peak intensity, value 45392.00 pA [limit 80.00 - 300.00]. Replaced faulty hemolyser unit. Repeated complete oximetry test - all points passed. Performed tHb calibration. Performed quality controls - passed. | Replaced hemolyser unit.
No | Visit | Complaint | RAME | QC failing in Oxi values | QC failing in Oxi values | Ran total Oxi check and everything passed $Reset hemolyzer status $Ran calibration and Thb cal | SWO-427866 | Adjustment | HW | Oxi module | tHb Cal |  | QC failing in Oxi values | QC failing in Oxi values | Ran Total oxygen Check and everything passed $Reset hemolyzer status $Ran calibration and THB cal
Yes | 2 | Y | RPNZ | Blank | Installation on site at Taumarunui Emergency department. |  | D50DA5CB7830F152 | Re classified RMED | installation report |  |  | Move to OOB trending | Blank | Installation on site at Taumarunui Emergency department.
No | Visit | Complaint | RDEU | 04.10.2018 15:45 H. Poser 03371699708: Kunde bittet um weitere Unterstützung bei der Einbindung hausinterner IP-Adressen auf Windowsebene | Netzwerkadressen verbleibend speichern | Netzwerkeinstellungen verbleibend gespeichert mit ewf-Befehl, Einweisungen Anwender vorgenommen, Kal, QC | SWO-427726 | Setup | SW | External SW | Network |  | 04.10.2018 15:45 H. Poser 03371699708: Customer asks for further assistance in integrating in-house IP addresses at the Windows level | Store network addresses to remain | Network settings remaining stored with ewf command, instruction users made, Kal, QC
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | CO2定标不过 0378 | 更换测试卡后修复 | SWO-431824 | Re classified RMED | installation report |  |  | Move to OOB trending | Test card, reagent package failure | CO2 calibration is no more than 0378 | Repair after replacing test card
No | Visit | Complaint | RCHN | 进样口支架断裂，报错4794 | 进样口支架断裂，报错4794.需要借件903-584*2 | 更换903-584*2后修复 by jesson | SWO-426643 | Defect | HW | Aspiration Module | Tube holder | 4794 | The inlet bracket broke, and the error was 4794. | The inlet bracket is broken, and the error is 4794. The loan is required 903-584 x 2 | Fix by jesson after replacing 903-584 x2
No | Visit | Complaint | RAME | On-board printer has stopped working. | Customer reporting onboard printer not printing. Suspect this is an intermittent issue, unable to reproduce during on site visit. | Proactively replaced printer module. Tested, printing properly, no failures observed. | SWO-427365 | Defect | HW | DMS_PC | Printer Module |  | On-board printer has stopped working. | Customer reporting onboard printer not printing. Suspect this is an intermittent issue, unable to reproduce during on site visit. | Proactively replaced printer module. Tested, printing properly, no failures observed.
No | Visit | Complaint | RAME | Crea not calibrating | Customer reporting CREA Sens 3 below range. | Removed electrode, tapped to remove bubbles, reran 2 PT calibration until Sens 3 came in range. Ran QC, all parameters within specifications. | SWO-427686 | Cleaned | Electrode | Membrane | Crea | Air bubble | Crea not calibrating | Customer reporting CREA Sens 3 below range. | Removed electrode, tapped to remove bubbles, reran 2 PT calibration until Sens 3 came in range. Ran QC, all parameters within specifications.
No | Visit | Complaint | RAME | Failing thb calibration when ran a s a patient Cohb - 3.9( 3.6-3.8) Contact: Adam Alatian - 818-719-4510 | Customer is unable to obtain a passing tHb calibration calibration verification check, FCOHb above range. | Performed service diagnostics. No failures observed. Adjusted halogen lamp from nominal 595 to 573, performed completed Oxi test, passed, performed service tHb calibration, passed, performed user tHb calibration and cal check, all passed, results on the mean.No further issues observed. | SWO-428071 | Adjustment | HW | Oxi module | Lamp unit | halogen | Failing thb calibration when ran a s a patient Cohb - 3.9( 3.6-3.8) Contact: Adam Alatian - 818-719-4510 | Customer is unable to obtain a passing tHb calibration calibration verification check, FCOHb above range. | Performed service diagnostics. No failures observed. Adjusted halogen lamp from nominal 595 to 573, performed completed Oxi test, passed, performed service tHb calibration, passed, performed user tHb calibration and cal check, all passed, results on the mean. No further issues observed.
No | Visit | Complaint | RAME | Error 579, Oximetry hardware problem. | Review data logs, observed tHb pA values was missing from 2 PT calibration. Performed service diagnostics, complete Oxi test failed due to neon lamp out of range. | Replaced lamp unit. Performed required adjustments. All parameters and test now functioning per specifications, no failures observed. | SWO-427685 | Defect | HW | Oxi module | Lamp unit | 579 | Error 579, Oximetry hardware problem. | Review data logs, observed tHb pA values was missing from 2 PT calibration. Performed service diagnostics, complete Oxi test failed due to neon lamp out of range. | Replaced lamp unit. Performed required adjustments. All parameters and test now functioning per specifications, no failures observed.
No | Visit | Complaint | RDEU | 24.09 2018 14:00 Frau Lam 0684111622792 immer wieder Kal Fehler Lyte und pH ( Referenzelekrtode ) Frau Lam hätte gerne einen Rückruf und Einsatz weitere Geräte mit Problemen Einlasnadel etc. | PH Elyte und Lac fehler | Ref. Lac und pCO2 Elektroden getauscht. Einlassnadel ausgetauscht. Kal und Qc in Ordnung | SWO-427452 | Assisted Replacement | Electrode | Electrode | Reference |  | 24.09 2018 14:00 Mrs. Lam 0684111622792 again and again Kal Error Lyte and pH ( Reference Elekrtode ) Ms. Lam would like to have a recall and use of other devices with problems lasing needle etc. | PH Elyte and Lac error | Ref. Lac and pCO2 electrodes. Inlet needle replaced. Kal and Qc all right
No | Visit | Complaint | RSCH | Roma.ini non corretto e errore calibrazione elettroliti. | Dati non corretti nel Roma.ini e errori di calibrazioni su diversi parametri. | Reinstallata Roma.ini completa, sostituito elettrodo di riferimento e gommino elettrodo PH. | SWO-428156 | Assisted Replacement | Electrode | Electrode | Reference |  | Rome.ini incorrect and electrolyte calibration error. | Incorrect data in Roma.ini and calibration errors on different parameters. | Reinstalled Roma.ini complete, replaced reference electrode and PH electrode.
No | Visit | Complaint | RPOL | Neon lamp number too low. | Testy serwisowe. | Wymiana modułu OXI. Kalibracja hemoglobiny. | SWO-427485 | Defect | HW | Oxi module | Oxi module | Neon number too low | Neon lamp number too low. | Service tests. | Replacing the OXI module. Hemoglobin calibration.
No | Visit | Complaint | RDEU | 08.10.2018; Hr. Steidler; Tel.: 07161642720; ABL kann nicht angepingt werden |  | Gerät überprüft Konfiguration...Ping jetzt ok. Time Server Anbindung jetzt ok. Dasi durchgeführt. | SWO-428527 | No fault found | SW | RMED SW | RMED SW |  | 08.10.2018; Mr. Steidler; Tel.: 07161642720; ABL cannot be pinged |  | Device Checks Configuration... Ping ok now. Time Server connection now ok. Dasi performed.
Yes | Visit | Complaint | RDEU | 10.10.2018, 10:19 Uhr, Fr. Stockeler, Tel. 02421599: Kd. benötigt Probennehmer | No | 2VE Safe Pico Aspirator geliefert, Einweisung durchgeführt, Fragen beantwortet, et. Scanner gewünscht | SWO-428389 | Re classified RMED | Logistical upSWO |  |  |  | 10.10.2018, 10:19 a.m., Ms. Stockeler, tel. 02421599: Kd. needs sampler | No | 2VE Safe Pico Aspirator Supplied, Instruction Performed, Questions Answered, et. Scanner desired
No | 8 | Y | RPAC | Blank | FAILS TO READ SC CORRECTLY |  | D50F29CDEE0300DB | [5 empty] | Blank | FAILS TO READ SC CORRECTLY
Yes | 4 | Y | RPAC | Blank | ANNUAL sERVICE |  | D50F2946EE2E00D7 | [5 empty] | Blank | ANNUAL sERVICE
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | na定标失败和测试卡初始化失败 | 连续更换2张测试卡后修复 | SWO-428157 | [5 empty] | Test card, reagent package failure | na calibration failure and test card initialization failure | Repair after 2 consecutive test cards
No | Visit | Complaint | RAME | Unit leaking from SC port. | Unit leaking at SC port. SC has restricted flow. May be due to a clotted sample during method comparison testing. | Flushed SC. Reseated waste pump to insure proper evacuation of samples. | SWO-428707 | [5 empty] | Unit leaking from SC port. | Unit leaking at SC port. SC has restricted flow. May be due to a clotted sample during method comparison testing. | Flushed SC. Reseated waste pump to insure proper evacuation of samples.
No | Visit | Complaint | RCHN | 报错3412 连接片断裂 | 报错3412 903-329的电缆连接接头损坏 借580-096 910-304 903-329各一个 | 该机保养时发现903-329的电缆接头损坏，4号试剂泵电机不能转动，测试卡条码不能被扫描。更换580-096 910-304 903-329备件后修复 by jesson | SWO-427498 | Defect | HW | Process Module | Needle Wash Unit | 3412 | Report error 3412 connection piece break | Error 3412 903-329 cable connector damage Borrow580-096 910-304 903-329 each | The machine maintenance found that 903-329 cable connector damage, No. 4 reagent pump motor can not be turned, test card bar code can not be scanned. Repair by jesson after replacing 580-096 910-304 903-329 spare parts
No | Visit | Complaint | RCHN | 温度加热超出范围，报错1068 | 温度加热超出范围，报错1068，需要借件903-210一个 | 更换903-210备件后修复 by jesson | SWO-427437 | Defect | HW | Process Module | PCB, cup dryer | 1068 | Temperature heating out of range, error 1068 | Temperature heating out of range, reported error 1068, need to borrow 903-210 a | Repair by jesson after replacement of 903-210 spare parts
No | Visit | Complaint | RIND | pH Measurement unstable for samples frequently. | Found blood clots at pH electrode bottom 'O' Ring. | Cleaned the pH Electrode and performed 2 point calibration ,QC and samples all results found within the range.found instrument working fine. | SWO-430056 | Cleaned | Fluidic pathway | pH/BG Module | pH/BG Module | Clots | pH Measurement unstable for samples frequently. | Found blood clots at pH electrode bottom 'O' Ring. | Cleaned the pH Electrode and performed 2 point calibration ,QC and samples all results found within the range.found instrument working fine.
No | Visit | Complaint | RKK | QCにて電解質（K,Cl）がバラツク。時々、Clが上限範囲外。Kが下限範囲外になる | QCにて電解質（K,Cl）がバラツク。時々、範囲外になる | 代替器に入替。 | SWO-435806 | Revisit required | Unknown | Unknown | Unknown | K, Cl QC | In QC, electrolyte (K, Cl) is baratzk. Sometimes Cl is out of the limit. K is outside the lower limit range. | In QC, electrolyte (K, Cl) is baratzk. sometimes goes out of range | Replaced with an alternative.
No | Visit | Complaint | RAME | Lac out on cals and qc | After customer maintenance lac will calibrate but qc is out Customer has tried a few membranes from same lot same box With no resolve. Lac range on level 3 qc 87-107 Lac is running 81-84 Cleaned measuring chamber with hypochlorite and rinsed with water Ran rinses and re membranes lactate Glucose sensitivity dropped so I re membranes glu Ca was running high on qc so I re membraned After multiple calibrations qc is passing lac running 93-94 on level 3 | Cleaned measuring chamber and re membrane of electrode allowing at least 3 one points and 2 two point calibrations before running qc after a glu $lac membrane change Also after discussing with lab techs Ronald and Sozit from night shift they are using kimwipes to dry electrodes after rinsing. This can damage electrodes. I have trained both on proper membrane maintenance. | SWO-428068 | User Error | Electrode | Electrode | Lac | use of kimwipes | Lac out on cals and qc | After customer maintenance lac will calibrate but qc is out Customer has tried a few membranes from same lot same box With no resolve. Lac range on level 3 qc 87-107 Lac is running 81-84 Cleaned measuring chamber with hypochlorite and rinsed with water Ran rinses and re membranes lactate Glucose sensitivity dropped so I re membranes glu Ca was running high on qc so I re membraned After multiple calibrations qc is passing lac running 93-94 on level 3 | Cleaned measuring chamber and re membrane of electrode allowing at least 3 one points and 2 two point calibrations before running qc after a glu $lac membrane change Also after discussing with lab techs Ronald and Sozit from night shift they are using kimwipes to dry electrodes after rinsing. This can damage electrodes. I have trained both on proper membrane maintenance.
No | Visit | Complaint | RAME | Bent probe | Probe is bent | Replace probe $Replaced pump tubings for customer | SWO-428774 | Defect | HW | Inlet module | Probe bend |  | Bent probe | Probe is bent | Replace probe $Replaced pump tubings for customer
Yes | Visit | Complaint | RPOL | Dodatkowe szkolenie | Szkolenie z zakresu wydruku oraz profilów pacjenta. | Szkolenie personelu. | SWO-428464 | Re classified RMED | planned SWO | Training | Training |  | Additional Training | Print training and patient profiles. | Staff training.
No | Visit | Complaint | RPOL | Błąd płukania | Pęknięta membrana płaska na zaworze górnym w module pH $bg | Wymiana membrany | SWO-428350 | Defect | HW | pH/BG Module | Valve membrane |  | Rinse error | Cracked flat membrane on top valve in pH module /bg | Replacing the membrane
No | Visit | Complaint | RAME | Hold with no restart | Possible valve erractic behavior | Replaced pinch valve and studied clot detection in system. Ran several samples to validate. | SWO-429403 | Defect | HW | Oxi module | Pinch valve |  | Hold with no restart | Possible valve erractic behavior | Replaced pinch valve and studied clot detection in system. Ran several samples to validate.
No | Visit | Complaint | RLTD | Program cannot continue on screen, status code error 2 Lilian Odoch 0203 4470326 | Program cannot continue on screen, status code error 2 Data base corrupted. Unable to identify the cause of error. | Unable to restore corrupted database. install clean databases and analyser starting up ok. Reinstalled setup information ok. SC and SP reinstalled ok. Calibration and QC ok. | SWO-429175 | Defect | SW | RMED SW | Database | status code 2 | Program cannot continue on screen, status code error 2 Lilian Odoch 0203 4470326 | Program cannot continue on screen, status code error 2 Data base corrupted. Unable to identify the cause of error. | Unable to restore corrupted database. install clean databases and analyser starting up ok. Reinstalled setup information ok. SC and SP reinstalled ok. Calibration and QC ok.
No | Visit | Complaint | RLTD | Analyser routinely fails to print and is only resolved with a reboot. | Follow up visit to install Windows 7 | Upgraded bios, then windows 7 software, loaded 3.4 and customer settings and data logs, setup windows 7 regional settings. | SWO-429909 | Defect | SW | Windows | Windows | printer | Analyser routinely fails to print and is only resolved with a reboot. | Follow up visit to install Windows 7 | Upgraded bios, then windows 7 software, loaded 3.4 and customer settings and data logs, setup windows 7 regional settings.
No | Visit | Complaint | RSAF | Faulty hemolyzer | Oxi hardware error | Replaced with a new hemolyzer unit.ran thb cal successfully. The analyzer is fully functional. | SWO-429409 | Defect | HW | Oxi module | Hemolyzer |  | Faulty hemolyzer | Oxi hardware error | Replaced with a new hemolyzer unit.ran thb cal successfully. The analyzer is fully functional.
No | Visit | Complaint | RAME | Oxi Errors | Oxi errors during internal and external QC. Can absorb failure during oxi complete test. | Replaced the hemolyzer. Will return hemolyzer as WAR part. | SWO-429267 | Defect | HW | Oxi module | Hemolyzer | 1012/1034 | Oxygen Errors | Oxygen errors during internal and external QC. Can absorb failure during oxygen complete test. | Replaced the hemolyzer. Will return hemolyzer as WAR part.
No | Visit | Complaint | RHUN | Képernyő kalibrációs hibák | Képernyő kalibrációs hibák | A műszer alapos kivizsgálást igényel, ezért a jelengivel megegyező típusú csereműszer került kilyezésre (316098). iButton átcserélve az adküldés zavartalan biztosítása érdekében Csereműszer beüzemelése, kalibrálás, QC mérés | SWO-429103 | [5 empty] | Screen calibration Errors | Screen calibration Errors | The instrument has to be thoroughly investigated, so the same type of replacer has been added (316098). iButton switched to ensure uninterrupted delivery Commissioning of the replacement instrument, calibration, QC measurement
No | Visit | Complaint | RAME | Frequent analog PCB failure. | Confirmed analog PCB failure. | Replaced analog PCB.QC | SWO-429611 | [5 empty] | Frequent analog PCB failure. | Confirmed analog PCB failure. | Replaced analog PCB. QC
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | PCO2,PH定标失败 | 更换测试卡 | SWO-431755 | [5 empty] | Test card, reagent package failure | PCO2, PH calibration failure | Replace the test card
No | Visit | Complaint | RCHN | 光学强度超范围 报错1031 | 光学强度超范围 报错1031 借903-364一个 | 更换903-364后修复 by jesson | SWO-427459 | Defect | HW | Process Module | Optical Unit | 1031 | Optical strength over range reported error 1031 | Optical strength over range error 1031 borrowed 903-364 one | Repair by after replacing 903-364 jesson
No | Visit | Complaint | RCHN | 测试卡卡盘 | 测试卡卡盘 | 更换测试卡后修复 | SWO-429035 | Defect | Consumables | Test Kit | Cartridge, Broken |  | Testing the kaka plate | Testing the kaka plate | Repair after replacing test card
No | Visit | Complaint | RCHN | 报错1024 | 报错1024，借903-364,944-212（14774）各一个 |  | SWO-428914 | Defect | HW | Process Module | Optical Unit | 1024 | 1024 error | 1024, borrow 903-364,944-212 (14774) each
No | Visit | Complaint | RCHN | 测试卡断裂 | 测试卡断裂 | 更换测试卡后修复 | SWO-429024 | Defect | Consumables | Test Kit | Cartridge, Broken |  | Test card break | Test card break | Repair after replacing test card
No | Visit | Complaint | RFRA | Logiciel actif mais aucune réponse de l'automate. | Remise en route machine sur site | Démarrage. Calibration 2 points ok. | SWO-429126 | [5 empty] | Active software but no response from the automaton. | Machine re-start on site | Start. Calibration 2 points ok.
No | Visit | Complaint | RLTD | Not calibrating | Not calibrating | Found the customer had been using the wrong membrane kit for the 84 sensor they have for the TCM5. They are using the blue ring membrane kit for a 92 sensor. Advised customer to order the correct membrane kit for the 84 sensor part number 905-871. Run through membrane training using on screen video. | SWO-429176 | [5 empty] | Not calibrating | Not calibrating | Found the customer had been using the wrong membrane kit for the 84 sensor they have for the TCM5. They are using the blue ring membrane kit for a 92 sensor. Advised customer to order the correct membrane kit for the 84 sensor part number 905-871. Run through membrane training using on screen video.
No | Visit | Complaint | RLTD | Unable to download data. David Ruaux reported 11.10.18. | Unable to download via Visi-Download. Checked settings, all ok. Checked patient data on the monitor and is OK to download to USB. Battery was flat. Charged battery until it had reached 50% to carry out reprogramming. TCM5 has some damage to casing to the bottom right of the monitor underneath. | Reprogrammed monitor with V1.20 software. Successfully downloaded patient data to PC. | SWO-429303 | [5 empty] | Unable to download data. David Ruaux reported 11.10.18. | Unable to download via Visi-Download. Checked settings, all ok. Checked patient data on the monitor and is OK to download to USB. Battery was flat. Charged battery until it had reached 50% to carry out reprogramming. TCM5 has some damage to casing to the bottom right of the monitor underneath. | Reprogrammed monitor with V1.20 software. Successfully downloaded patient data to PC.
No | Visit | Complaint | RLTD | Unable to download data. | Unable to download data. There are no communication lights appearing on customer PC hub when connected via the Ethernet cable. Rear USB ports are also unresponsive. Reprogrammed with V1.2 software, but no improvement has been made. | Monitor needs to be returned back to base for investigation of USB and Ethernet ports. | SWO-429314 | [5 empty] | Unable to download data. | Unable to download data. There are no communication lights appearing on customer PC hub when connected via the Ethernet cable. Rear USB ports are also unresponsive. Reprogrammed with V1.2 software, but no improvement has been made. | Monitor needs to be returned back to base for investigation of USB and Ethernet ports.
Yes | Visit | Complaint | RDEU | Kalibrierprobleme in Na und Ca | Referenzelektrode muss erneuert werden | Referenzelektrode bei Medizintechnik abgegeben. Wird durch S. Topel eingebaut. | SWO-429554 | Re classified RMED | Logistical upSWO |  |  | Only delivery, mail from BORWE to INGZI "AW: SWO-429554" | Calibration problems in Na and Ca | Reference electrode must be renewed | reference electrode in medical technology. Installed by S. Topel.
No | Visit | Complaint | RKK | QCでNa範囲外 | QCでNa範囲外 | 比較電極交換 | SWO-428958 | Assisted Replacement | Electrode | Electrode | Reference |  | Na out of range in QC | Na out of range in QC | Comparative electrode replacement
No | Visit | Complaint | RKK | 電解質QC範囲外 | 電解質（Na Ca）QC範囲外 メンブラン交換後も同症状 | Ref電極交換 | SWO-429045 | Assisted Replacement | Electrode | Electrode | Reference |  | Electrolyte QC out of range | Electrolyte (Na Ca) QC out of range Same symptoms even after the menbran exchange | Ref electrode replacement
No | Visit | Complaint | RIND | Print out not coming. | When touch the Print tab instrument gets hanged & print out also not coming. | Install the printer driver software,check the print out & it was coming out properly,install the application software.After that handover the instrument to the customer. | SWO-430069 | Defect | SW | External SW | Printer driver |  | Print out not coming. | When touch the Print tab instrument gets hanged & print out also not coming. | Install the printer driver software,check the print out & it was coming out properly,install the application software. After that handover the instrument to the customer.
No | Visit | Complaint | RAME | PCO2 level 1 is running lower that peer 2 SD. She is reading at 65.9 and the peer 2 SD range is 65.9 to 70.5. | PCO2 level 1 is running lower that peer 2 SD. She is reading at 65.9 and the peer 2 SD range is 65.9 to 70.5. | Replaced pco2 electrode with dummy electrode. Then did three 2pt Cals and ran QC, level 1 went from 64.3 to 68.5. Replaced Reference membrane with lot R2333. Then did three 2pt Cals and ran QC, level 1 is still at 68.5. | SWO-428749 | Assisted Replacement | Electrode | Membrane | Reference |  | PCO2 level 1 is running lower that peer 2 SD. She is reading at 65.9 and the peer 2 SD range is 65.9 to 70.5. | PCO2 level 1 is running lower that peer 2 SD. She is reading at 65.9 and the peer 2 SD range is 65.9 to 70.5. | Replaced pco2 electrode with dummy electrode. Then did three 2pt Cals and ran QC, level 1 went from 64.3 to 68.5. Replaced Reference membrane with lot R2333. Then did three 2pt Cals and ran QC, level 1 is still at 68.5.
No | Visit | Complaint | RKOR | Hct value cannot be correctly determined | Perform the system cleaning | Perform the system cleaning | SWO-430192 | Revisit required | HW | Process Module | Needle Wash Unit |  | Hct value cannot be correctly determined | Perform the system cleaning | Perform the system cleaning
No | Visit | Complaint | RDEU | Fehler 3264 Errors related to power, temperature, or Time/date missmatctch(>100h bettween internal pc clock and factory clock) | Fehler 3264, Solution Problem | Fehleranalyse, Uhrzeit und Datum neu eingestellt und abgespeichert, Verriegelung des SP gelöst und SP erneuert, Qualitätskontrolle durchgeführt (siehe gleichzeitige Wartung) | SWO-429076 | Adjustment | SW | Analyzer Control Module | RMED SW | Time Setting; Potential SP User Error | Fehler 3264 Errors related to power, temperature, or Time/date missmatctch(>100h bettween internal pc clock and factory clock) | Error 3264, Solution Problem | Error analysis, time and date re-set and saved, locking of the SP solved and SP renewed, quality control performed (see simultaneous maintenance)
No | Visit | Complaint | RIND | Inlet door not closing during analysis | Checked and found that Inlet door is not closing. | It was found that inlet door mechanism jammed after closing the front cover. Removed two metal spacers under the screws for left $right panels. | SWO-429596 | Adjustment | HW | Cover and Chassis | Door, Inlet |  | Inlet door not closing during analysis | Checked and found that Inlet door is not closing. | It was found that inlet door mechanism jammed after closing the front cover. Removed two metal spacers under the screws for left $right panels.
No | Visit | Complaint | RPOL | Błąd temperatury AutoCheck | Testy serwisowe. | Wymiana modułu AutoCheck. Wymiana szybki skanera wewnętrznego. | SWO-429151 | Defect | HW | Autocheck module | Autocheck Module | 736 | AutoCheck Temperature Error | Service tests. | AutoCheck module replacement. Quick internal scanner replacement.
No | Visit | Complaint | RDEU | Akku-Anzeige 0% | Akku-Anzeige 0% | Erneuerung Akku-Einheit - Funktionsüberprüfung | SWO-429723 | Defect | HW | Battery module | Battery | 0% | Battery display 0% | Battery display 0% | Renewal battery unit - function check
No | Visit | Complaint | RDEU | Akku defekt | Akku-Anzeige 0% | Erneuerung Akku-Einheit - Funktionsüberprüfung | SWO-429819 | Defect | HW | Battery module | Battery | 0% | Battery defective | Battery display 0% | Renewal battery unit - function check
No | Visit | Complaint | RDEU | Akku defekt / Bestellung bereits erfolgt | Akku Anzeige 0% | Erneuerung Akku-Einheit und Funktionsüberprüfung etc. | SWO-429711 | Defect | HW | Battery module | Battery | 0% | Battery defective / order already made | Battery display 0% | Renewal battery unit and function verification etc.
Yes | Visit | Complaint | RAME | Hold with no restart | Onsite checkup on ABL90 system repair | None necessary | SWO-430031 | Re classified RMED | planned SWO |  |  | mail refer to XXX | Hold with no restart | Onsite checkup on ABL90 system repair | None necessary
No | Visit | Complaint | RFRA | Erreur d'étalonnage de pompe | Vérification complète | Tests en mode service -> pression en défaut $ $ débouchage à la seringue-> blocage $ $ débouchage de l'aiguille à l'hypochlorite $ $ tests en mode service -> ok $ $ remise en route (1h) $ $ calibration $ $ passage des contrôles $ $ formation utilisateur (calibration thb) $ $ essais bons | SWO-429825 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | Pump calibration error | Full check | Tests in service mode - 'failure pressure $ $ unblocking at the syringe -$ $ uncorking the needle to the hypochlorite $$ tests in service mode - ok $$ restarted ($1h) $ calibration $ passage of controls $ $ training user (calibration t hb)$$ good trials
No | Visit | Complaint | RLTD | Liz 01412 377 676 | Analyser failed level 3 QC for co-oximetry parameters | Re ran level 3 QC and passed OK | SWO-429843 | No fault found | HW | Oxi module | Oxi module | QC C | Liz 01412 377 676 | Analyser failed level 3 QC for co-oximetry parameters | Re ran level 3 QC and passed OK
No | Visit | Complaint | RKOR | Inlet position error | . | Service mode check, Gasket module position adjust. | SWO-429744 | Adjustment | HW | Auto Inlet module | Auto Inlet module | Inlet Cal | Inlet position error | . | Service mode check, Gasket module position adjust.
No | Visit | Complaint | RCHN | 无法识别测试卡 | 无法识别测试卡 借903-334 902-972各一个 | 更换903-334 902-972后修复 by jesson | SWO-429040 | [5 empty] | The test card is not recognized | Unrecognizable test card Borrowing 903-334 902-972 each | Repair by jesson after replacing 903-334 902-972
No | Visit | Complaint | RCHN | 传感器错误，报错4794 | 传感器错误，报错4794，需要借件636-549 | 更换636-549后修复 by jesson | SWO-428879 | Defect | HW | Aspiration Module | Sensor, needle | 4794 | Sensor error, error 4794 | Sensor error, error 4794, need to borrow 636-549 | Repair by jesson after replacing 636-549
No | Visit | Complaint | RDEU | KAL Driften Elektrolyte | KAL Driften Elektrolyte | Überprüfung Referenz Elektroden-Bereich (keine Verunreinigungen etc.) - Erneuerung Referenz Elektrode - Funktionsüberprüfung etc. | SWO-429710 | Assisted Replacement | Electrode | Electrode | Reference |  | KAL Drifting Electrolytes | KAL Drifting Electrolytes | Check reference electrode area (no impurities etc.) - renewal reference electrode - function check etc.
No | Visit | Complaint | RAME | 581 oximetry errors during Linearity | 581 error during Linearity. ***Error occurred while the customer was running the linearity and wasted a kit. Please approve a replacement VK-R7 Linearity kit for the customer.*** | Replaced Hemolyzer and calibrations. ***Error occurred while the customer was running the linearity and wasted a kit. Please approve a replacement VK-R7 Linearity kit for the customer.*** | SWO-430035 | Defect | HW | Oxi module | Hemolyzer | 581 | 581 oximetry errors during Linearity | 581 error during Linearity. ***Error occurred while the customer was running the linearity and wasted a kit. Please approve a replacement VK-R7 Linearity kit for the customer.*** | Replaced Hemolyzer and calibrations. ***Error occurred while the customer was running the linearity and wasted a kit. Please approve a replacement VK-R7 Linearity kit for the customer.***
No | Visit | Complaint | RKOR | Hct value cannot be correctly determined | Perform the needle wash unit $ tube #9, pump 4, pump 5 cleaning | Perform the needle wash unit $ tube #9, pump 4, pump 5 cleaning | SWO-430194 | Cleaned | HW | Process Module | Needle Wash Unit |  | Hct value cannot be correctly determined | Perform the needle wash unit / tube #9, pump 4, pump 5 cleaning | Perform the needle wash unit $ tube #9, pump 4, pump 5 cleaning
No | Visit | Complaint | RDEU | 12.10.2018, Frau Kreis, Tel: 071121840626: AQT im Halt mit 5784-Setsafe 22 | Rotate Storage Wheel 24 | Haltegriffe an Messkassette abgebrochen. Defekte Kassette entfernt. Abgebrochenes Teil nicht auffindbar. Wiederinbetriebnahme. Kein QC vorrätig. | SWO-430326 | Defect | Consumables | Test Kit | Cartridge, Broken | Ears | 12.10.2018, Mrs. Kreis, Tel: 071121840626: AQT in stop with 5784-Setsafe 22 | Rotate Storage Wheel 24 | Handles on measuring cassette aborted. Defective cartridge removed. Cancelled part cannot be found. Recommissioning. No QC in stock.
No | Visit | Complaint | RPOL | Błąd skanera kodów | Nie czytaj odmów kreskowych | Wymiana skanera | SWO-430600 | Defect | HW | Basic Unit | Barcode reader |  | Code scanner Error | Do not read bar ODMs | Replacing the scanner
No | Visit | Complaint | RCHN | 无法读取测试卡智能芯片 | 无法读取测试卡智能芯片,借 903-581一个。 | 更换903-581后修复 by jesson | SWO-429666 | Defect | HW | Basic Unit | PCB, Sensor interface | Chip | Failed to read test card smart chip | The test card smart chip could not be read, borrowing 903-581 one. | Repair by jesson after replacement 903-581
No | Visit | Complaint | RCHN | 处理盘错误 5257 | 处理盘错误 5257,借636-550一个。 | 更换636-550后修复 by jesson | SWO-428993 | Defect | HW | Process Module | Sensor, process unit | 5257 | Processing disk error 5257 | Processing disk error 5257, borrowing 636-550 one. | Repair by jesson after replacement 636-550
No | Visit | Complaint | RIND | Service Visit. Doctors having problem with the measurement of pCO2, and its calculated values. | Need to place the other analyzer for cross verify the sample. | Install the analyzer, there is many complication in the analyzer installation because of software corrupt and other issues and also analyzer is not switch on from more then 6 months. Successful install the analyzer, do the same sample on both analyzer, there is a different result of pCO2 measurement, its around lagging 15-20mmhg, Observe the analyzer sample processing there is a fault found in the pCO2 chamber, the base diameter of pCO2 chamber has increased, so the sample is not keep ideal volume at the base of the electrode and ideally measured by the pCO2 electrode, hence because of this there is lagging of 15-20 mmHg. Need to replace the chamber. | SWO-430070 | Defect | HW | pH/BG Module | measuring chamber | pCO2 | Service Visit. Doctors having problem with the measurement of pCO2, and its calculated values. | Need to place the other analyzer for cross verify the sample. | Install the analyzer, there is many complication in the analyzer installation because of software corrupt and other issues and also analyzer is not switch on from more then 6 months. Successful install the analyzer, do the same sample on both analyzer, there is a different result of pCO2 measurement, its around lagging 15-20mmhg, Observe the analyzer sample processing there is a fault found in the pCO2 chamber, the base diameter of pCO2 chamber has increased, so the sample is not keep ideal volume at the base of the electrode and ideally measured by the pCO2 electrode, hence because of this there is lagging of 15-20 mmHg. Need to replace the chamber.
No | Visit | Complaint | RKOR | Oxi error | . | Oxi module clean. Hemolyzer module clean. Oxi test. Cal QC check Sample test. | SWO-430183 | Cleaned | HW | Oxi module | Oxi module |  | Oxi error | . | Oxygen module CLEAN. Hemolyzer module CLEAN. Oxygen test. Cal QC Check Sample test.
No | Visit | Complaint | RCHN | 报错3624 | 报错3624，借903-215一个 | 更换903-215后修复 by jesson | SWO-430093 | Defect | HW | Incubator Module | PCB, incubator |  | Error 3624 | 3624, loan 903-215 one | Repair by jesson after replacement 903-215
No | Visit | Complaint | RDEU | Labor meldet Fehler 0964 Ableitstrom im Anschluss des SP erkannt | Meldung Ableitstrom im Anschluss LP erkannt tritt sporadisch auf $ nach Systemüberprüfung ABL z.Z. ohne Fehlermeldung meßbereit | Fehlersuche Daten überprüft , Kontrolle LP , Ansaugstutzen Reklamation SP90 Geräteneustart $ Warmstart durchgeführt keine Fehlermeldung Kunde wird ABL weiter beobachten und ggf. LP bei erneutem Auftreten der Meldung erneuern | SWO-430399 | Assisted Replacement | SP | SP | Leak current |  | Laboratory reports error 0964 leakage current detected following SP | Message dissipation current detected in the follow-up LP occurs sporadically on the system after system verification ABL ready to be measured without error message | Troubleshooting Data checked , control LP , suction nozzle Complaint SP90 Device restart - Warm start performed No error message Customer will continue to monitor ABL and renew LP if the message reappears
No | 4 | Y | RPAC | Blank | cal error present |  | D51546E63C157F13 | Assisted Replacement | SP | SP | Cal |  | Blank | cal error present
No | Visit | Complaint | RLTD | Haval Ozgun 07737 457222 | Analyser found to be blocked in the right hand sensor bay pipe. Tried to flush just this part of the analyser, unable to flush, totally blocked. Sensor bay unit removed from the analyser and fresh attempt made to flush and also to clear blockage with a thin wire. Still unable to remove blockage and nothing at all being flushed out, so cannot say what has caused the block. As unable to clear blockage, fitted replacement Sensor Bay unit. Restarted analyser - calibration ok. | Cleared blockage by replacing the Sensor Bay Unit. | SWO-431294 | Cleaned | HW | Basic Unit | Sensor Bay Unit |  | Haval Ozgun 07737 457222 | Analyser found to be blocked in the right hand sensor bay pipe. Tried to flush just this part of the analyser, unable to flush, totally blocked. Sensor bay unit removed from the analyser and fresh attempt made to flush and also to clear blockage with a thin wire. Still unable to remove blockage and nothing at all being flushed out, so cannot say what has caused the block. As unable to clear blockage, fitted replacement Sensor Bay unit. Restarted analyser - calibration ok. | Cleared blockage by replacing the Sensor Bay Unit.
Yes | Visit | Complaint | RLTD | Haval Ozgun 07737 457222 | Supplied replacement inlet cover part no. 886-916 x 1. | Supplied replacement cover. | SWO-431494 | Re classified RMED | Logistical upSWO |  |  |  | Haval Ozgun 07737 457222 | Supplied replacement inlet cover part no. 886-916 x 1. | Supplied replacement cover.
No | Visit | Complaint | RESP | El equipo no arranca. Parece que la fuente de alimentación no funciona. | Instalación | Instalación y puesta en marcha tras reparación en Madrid, transporte y sustitución de equipamiento prestado. Pruebas de funcionamiento. Calibrado y verificado con QC | SWO-430708 | Defect | HW | Power Supply | Power supply |  | The equipment won't start. Looks like the power supply isn't working. | Installation | Installation and commissioning after repair in Madrid, transport and replacement of borrowed equipment. Operational tests. Calibrated and verified with QC
No | Visit | Complaint | RCHN | 液路堵塞 | 液路堵塞 | 使用冲洗工具后修复by tim | SWO-431054 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | Liquid road blockage | Liquid road blockage | Fix by tim after using the flushing tool
Yes | Visit | Complaint | RCHN | 进样针超出次数 | 进样针超出次数 | 更换后正常by tim | SWO-431053 | Re classified RMED | planned SWO |  |  | see mail from Yang, Daniel CNDYA "SWO-431053" | Sample needle out of time | Sample needle out of time | Normal by tim after replacement
No | Visit | Complaint | RFRA | Air détecté lors de démo dès le 2iem jour. | Air détecté au bout de 12 analyse. Système neuf en démo (remplacement abl80) | Changement kit maintenance ok. Calibration ok. Plus de msg d’erreur. Système en démo depuis 2 jours. Pas de facturation ok. NB : Attente fin des consultations pour faire intervention. | SWO-430657 | [5 empty] | Air detected during demo on the 2nd day. | Air detected after 12 scans. New system in demo (abl80 replacement) | Change maintenance kit ok. Calibration ok. No more error msg. System in demo for 2 days. No billing ok. NB: Waiting for the end of consultations to intervene.
No | Visit | Complaint | RDEU | KAL Driften Elektrolyte | KAL Driften Elektrolyte QC Verletzung Na | Referenz Elektroden-Bereich intensiv gereinigt (evt. Kontaminiert $ Messkammer mit Proteinentferner gereinigt) - Neumembranierung etc. - Funktionsüberprüfung etc. | SWO-430742 | Cleaned | Electrode | Electrode | Reference |  | KAL Drifting Electrolytes | KAL Drifting Electrolytes QC Violation Na | Reference electrode area intensively cleaned (evt. Contaminated - Measuring chamber cleaned with protein remover) - re-membraneing etc. - function check, etc.
No | Visit | Complaint | RDEU | 12.10.2018; Fr. Schönebeck; Tel.: 024616202892; Display fällt teilweise aus | Sicht- und Funktionskontrolle | Displaykabel ersetzt, weiterhin aber keine Wackelfreie Verbindung zum Monitor gewährleistet. Erneuter Besuch mit Ersatzteil 910-298 erforderlich. | SWO-430799 | Revisit required | HW | Cover and Chassis | Display |  | 12.10.2018; Ms. Schönebeck; Tel.: 024616202892; Display partially fails | Visual and functional control | display cable, but still does not ensure a wobbly connection to the monitor. Revisit with spare part 910-298 required.
No | Visit | Complaint | RKOR | Sample clot error | . | Liquid flow line clean. Service mode check. User training. | SWO-433890 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway | Clot | Sample clot error | . | Liquid flow line clean. Service mode check. User training.
Yes | 5 | Y | RPNZ | Blank | Set up at CWMH Fiji for demo $trial. |  | D516E7CBB2252A89 | Re classified RMED | installation report |  |  |  | Blank | Set up at CWMH Fiji for demo /trial.
Yes | 2 | Y | RPAC | Blank | Failed SP |  | D516BB39B20916B4 | [5 empty] | Blank | Failed SP
No | Visit | Complaint | RLTD | Analyser showing clot, which trainers are unable to clear. | Analyser failing fluid transport tests on clot detect | Flushed analyser a number of times, no resistance felt. Analyser passes all fluid transport tests in service, but then 15 minutes into 55 min startup it goes into intervention required. Tried a number of things, ended up replacing sensor cassette, which resolved issue. | SWO-432817 | Assisted Replacement | SC | SC | FT |  | Analyser showing clot, which trainers are unable to clear. | Analyser failing fluid transport tests on clot detect | Flushed analyser a number of times, no resistance felt. Analyser passes all fluid transport tests in service, but then 15 minutes into 55 min startup it goes into intervention required. Tried a number of things, ended up replacing sensor cassette, which resolved issue.
No | Visit | Complaint | RDEU | Nach Sensorkassettentausch Temperaturfehler D | Kontaktproblem am connector board | Austausch connector Board, Funktionskontrolle | SWO-431429 | [5 empty] | After sensor cassette replacement temperature error D | Kontaktproblem am connector board | Replacement connector board, function control
No | Visit | Complaint | RCHN | 屏幕白色 | 屏幕白色，借910-298一个 | 更换910-298后修复by tim | SWO-430452 | Defect | HW | Basic Unit | Display |  | Screen white | Screen white, borrow 910-298 one | Repair by tim after replacement 910-298
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | PO2定标失败，质控失败，初始化失败0210，1340，0375 | 更换测试卡试剂包正常。 | SWO-433436 | Assisted Replacement | SC | SC | pO2 | Cal | Test card, reagent package failure | PO2 calibration failure, QUALITY control failure, initialization failure 0210,1340,0375 | Replacing the test card kit is OK.
No | Visit | Complaint | RCHN | 老师对试剂包的使用人份数有疑问，总觉得次数不够 | 老师对试剂包的使用人份数有疑问，总觉得次数不够 | 下载数据，解释仪器在什么情况下会消耗杯位by tim | SWO-432704 | User Error | Consumables | Solution Pack | Solution Pack | Inconsistent mental model | The teacher has questions about the number of people who use the reagent kit, and always feels that the number is not enough | The teacher has questions about the number of people who use the reagent kit, and always feels that the number is not enough | Download the data to explain when the instrument will consume cup bit sup
No | Visit | Complaint | RCHN | 老师对试剂包的使用人份数有疑问，总觉得次数不够 | 老师对试剂包的使用人份数有疑问，总觉得次数不够 | 下载数据，解释仪器在什么情况下会消耗杯位by tim | SWO-432700 | User Error | Consumables | Solution Pack | Solution Pack | Inconsistent mental model | The teacher has questions about the number of people who use the reagent kit, and always feels that the number is not enough | The teacher has questions about the number of people who use the reagent kit, and always feels that the number is not enough | Download the data to explain when the instrument will consume cup bit sup
No | Visit | Complaint | RCHN | 定标底物外露检查失败 1024 | 定标底物外露检查失败 1024,借光学单元，使用备用903-364R0304N015 | 更换备用光单后，做1024的检测结果均无法通过，加借一个903-364 | SWO-430462 | Revisit required | HW | Process Module | Optical Unit | 1024 | Calibration substrate exposure check failed 1024 | Bid substrate exposure check failed 1024, borrowed optical unit, using backup 903-364R0304N015 | After replacing the spare light sheet, the test results of 1024 can not pass, plus a 903-364
No | Visit | Complaint | RCHN | 报错1024 | 报错1024。需要借件903-364，和944-212（14774） | 收到的光单有问题，现勉强使用第一次借用的光单，重新定标后修复by tim | SWO-430481 | Revisit required | HW | Process Module | Optical Unit | 1024 | 1024 error | 1024. Borrowed 903-364, and 944-212 (14774) | There is a problem with the light sheet received, and now barely using the first borrowed light sheet, re-calibration after repair by tim
No | Visit | Complaint | RAME | Neon number too low & Oxi hardware error. | Neon lamp intensity failed complete oxi test. | Mon: Replaced lamp, same issue. Tue: replaced oxi module. | SWO-431092 | Defect | HW | Oxi module | Oxi module | 258, 579 | Neon number too low & OXI hardware error. | Neon lamp intensity failed complete oxi test. | Mon: Replaced lamp, same issue. Tue: Replaced oxygen module.
No | Visit | Complaint | RAME | Rinse Error : Possible blockage | Rinse errors 0751, 0510 and 059. Checked valves and tubing. Found defective pH $BG module gas valve. | Replaced pH $BG module gas valve. Replaced missing lactate electrode sealing ring. | SWO-432160 | Defect | HW | pH/BG Module | Valve, membrane |  | Rinse Error : Possible blockage | Rinse errors 0751, 0510 and 059. Checked valves and tubing. Found defective pH /BG module gas valve. | Replaced pH $BG module gas valve. Replaced missing lactate electrode sealing ring.
No | Visit | Complaint | RAME | PCO2 status, sensitivity, drift, and unstable errors during calibrations. | PCO2 failing calibrations. Found electrode connector not staying securely locked down on electrode. Problem continues after replacing electrode connector. | Replace electrode connector. Problem not corrected. Will return with pCO2 electrode. | SWO-431639 | Revisit required | Electrode | Electrode | pCO2 |  | PCO2 status, sensitivity, drift, and unstable errors during calibrations. | PCO2 failing calibrations. Found electrode connector not staying securely locked down on electrode. Problem continues after replacing electrode connector. | Replace electrode connector. Problem not corrected. Will return with pCO2 electrode.
No | Visit | Complaint | RDEU | Fehler 3264 Errors related to power, temperature, or Time/date missmatctch(>100h bettween internal pc clock and factory clock) 15.10.2018, Frau Otto, Tel: 051916023925: erneut Messabbrüche am AQT mit 1067 und 3050- CorruptXML | erneut Messabbrüche am AQT mit 1067 und 3050- CorruptXML | Fehleranalyse, Sicherungen an AU Board geprüft und gereinigt, BIOS Batterie geprüft, PC Board defekt-> neu bestellt | SWO-431512 | Revisit required | HW | Cover and Chassis | PCB, Main Board | 3264 | Fehler 3264 Errors related to power, temperature, or Time/date missmatctch(>100h bettween internal pc clock and factory clock) 15.10.2018, Frau Otto, Tel: 051916023925: erneut Messabbrüche am AQT mit 1067 und 3050- CorruptXML | again measurement interruptions on the AQT with 1067 and 3050- CorruptXML | Error analysis, fuses on AU board checked and cleaned, BIOS battery checked, PC board defective-> reordered
No | Visit | Complaint | RFRA | Pb de rotation de la storage wheel | Contrôle de la storage wheel et des éléments pouvant bloquer la rotation | Mise en place d’une rondelle pour surélever le cartrige lock. Essais | SWO-431289 | Adjustment | HW | Storage Module | Cartridge lock | Potential FSE User Error | Pb of storage wheel rotation | Control of the storage wheel and items that can block the rotation | Placing a washer to raise the cartrige lock. Tests
No | Visit | Complaint | RDEU | 15.10.2018 17:35Uhr Frau Heinrich 09661520331: Die neue Cartridge wurde eingelegt, aber hat sich verklemmt und hängt fest. | AQT steht in Fehler Kal. Kassette steckt unter PunchDown ins Incubatorwheel fest - ein der vorderen Nasen der Kal. Kassette ist abgebrochen | AQT aufgeschraubt, verklemmte Kal. Kassette entfernt, AQT im Service überprüft - alles wieder in Ordnung, Kal. Kassette für NT-proBNP gestartet, System- und Funktionskontrolle in Ordnung | SWO-431867 | Defect | Consumables | Test Kit | Cartridge, Broken | Nose | 15.10.2018 17:35pm Mrs Heinrich 09661520331: The new cartridge has been inserted, but has become jammed and hangs. | AQT is in error Kal. Cassette is stuck under PunchDown in the Incubatorwheel - one of the front noses of the Kal. Cassette is canceled | AQT screwed on, jammed Kal. Cassette removed, AQT checked in service - everything back in order, Kal. Cassette started for NT-proBNP, system and function control in order
No | Visit | Complaint | RDEU | Fenster für Parameterkassetten geht nicht auf | Fenster für Parameterkassetten geht nicht auf | Vordere Abdeckung geöffnet, ausgehakte Feder an Fenster befestigt, Abdeckung wieder befestigt, Fkt.test i.O. | SWO-431372 | Adjustment | HW | Cover and Chassis | Door, Storage | Spring | Window for parameter cassettes does not open | Window for parameter cassettes does not open | Front cover open, hooked spring attached to windows, cover reattached, Fkt.test i.O.
No | Visit | Complaint | RAME | Battery failed at install | Battery had failed on install | Replaced battery assembly. Waited until battery charged past 50%. Also checked other 8 analyzers on site for any issues | SWO-432561 | Defect | HW | Battery module | Battery module |  | Battery failed at install | Battery had failed on install | Replaced battery assembly. Waited until battery charged past 50%. Also checked other 8 analyzers on site for any issues
No | Visit | Complaint | RSAF | Needle bent and inlet gasket leaking | Needle bent and worn inlet gasket | Replaced inlet gasket w holder and inlet probe. Replaced empty solution pack. Start up test ok. | SWO-431913 | Defect | HW | Auto Inlet module | Probe bend |  | Needle bent and inlet gasket leaking | Needle bent and worn inlet gasket | Replaced inlet gasket w holder and inlet probe. Replaced empty solution pack. Start up test ok.
No | Visit | Complaint | RCHN | 科室老师反馈LAC结果不出值 | 科室老师反馈LAC结果不出值 | 检测接口板和信号板，发现并无问题。考虑是科室老师的操作问题或者样本问题，并对其解释 by tim | SWO-432005 | [5 empty] | Section teacher feedback LAC results are not worth | Section teacher feedback LAC results are not worth | Check the interface board and signal board and find no problems. Consider the operation or sample problem of the department teacher and explain it by tim
No | Visit | Complaint | RAME | Analyzer leaking from the bottom. | Customer reported that the analyzer has been leaking even after replacing the solution pack twice. Found analyzer very messy with QC from the leak, cleaned up all of the QC internally and externally. | Inspected the SP that was removed and determined it was the problem, the waste block and tubings are fine. Installed new SP and ran system cycles and manual QC, found all levels passed. | SWO-432498 | [5 empty] | Analyzer leaking from the bottom. | Customer reported that the analyzer has been leaking even after replacing the solution pack twice. Found analyzer very messy with QC from the leak, cleaned up all of the QC internally and externally. | Inspected the SP that was removed and determined it was the problem, the waste block and tubings are fine. Installed new SP and ran system cycles and manual QC, found all levels passed.
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | PH,PCO2，NA,CL定标不过 1315 1316 1319 1321 | 更换测试卡后修复 | SWO-436567 | [5 empty] | Test card, reagent package failure | PH, PCO2, NA, CL but 1315 1316 1319 1321 | Repair after replacing test card
No | Visit | Complaint | RPOL | Błąd sensora | Analiza problemu, rozmowa z klientem | Deinstalacja w celu diagnozy usterki w serwisie | SWO-432246 | [5 empty] | Sensor error | Problem analysis, customer conversation | Uninstallation to diagnose a defect in the service
No | Visit | Complaint | RCHN | 试管夹滚轮脱落，定标卡条码器不能扫描 | 试管夹滚轮脱落，定标卡条码器不能扫描，借903-584和902-924各一个 | 更换903-584 902-924后修复by tim | SWO-431912 | Defect | HW | Inlet module | Tube holder |  | Test tube clip roller spleus off, calibration card barcode cannot be scanned | Test tube clip roller spleus off, calibration card barcode can not be scanned, borrow 903-584 and 902-924 each one | Repair by tim after replacing 903-584 902-924
No | Visit | Complaint | RCHN | 试管夹损坏 | 试管夹损坏，借903-584一个 | 更换903-584后修复by tim | SWO-431160 | Defect | HW | Inlet module | Tube holder |  | Damage to test tube clip | Damaged tube clips, borrow edi-903-584 | Repair by tim after replacement 903-584
No | Visit | Complaint | RCHN | 定标底物外露检查失败 1024 | 定标底物外露检查失败，更换备用光单后，做1024的检测结果均无法通过，加借一个903-364 | 更换903-364后修复by tim | SWO-431214 | Defect | HW | Process Module | Optical Unit | 1024 | Calibration substrate exposure check failed 1024 | The calibration substrate exposure inspection failed, replacing the spare light sheet, do 1024 test results can not pass, plus a 903-364 | Repair by tim after replacing 903-364
No | Visit | Complaint | RCHN | TNI定标不过 报错3410 | TNI定标不过 报错3410 借944-212（14868）一张 | 重新定标后修复by tim | SWO-430454 | Revisit required | Consumables | Test Kit | Cartridge, cal |  | TNI calibration but misreported 3410 | TNI calibration but error 3410 loan 944-212 (14868) one | Fix by tim after re-calibration
No | Visit | Complaint | RAME | Na+ QC results CV's are high | Na+ QC results CV's are high | Replaced Na+ and reference electrodes. Calibrated and ran QC multiple times overnight. Results look good, instrument operating normally. | SWO-433008 | User expectation | Electrode | Electrode | Reference |  | Na+ QC results CV's are high | Na+ QC results CV's are high | Replaced Na+ and reference electrodes. Calibrated and ran QC multiple times overnight. Results look good, instrument operating normally.
No | Visit | Complaint | RAME | Customer reported issues with Calcium calibration after replacing Reference electrode and Ca membrane once. | Upon arrival found customer did not replaced reference electrode, and having issues with multiple electrolytes. | Found pH sealing was half on. Reseat the pH sealing ring. Did couple of 2 points. Noticed no changes to multiple electrolytes. Replaced reference electrode and membrane from customer stock. Tested it with multiple 2 point and all went green. All calibrations and 3 Levels of QC passed as per manufacturer specification | SWO-432472 | Defect | Electrode | Electrode | Sealing ring, pH |  | Customer reported issues with Calcium calibration after replacing Reference electrode and Ca membrane once. | Upon arrival found customer did not replaced reference electrode, and having issues with multiple electrolytes. | Found pH sealing was half on. Reseat the pH sealing ring. Did couple of 2 points. Noticed no changes to multiple electrolytes. Replaced reference electrode and membrane from customer stock. Tested it with multiple 2 point and all went green. All calibrations and 3 Levels of QC passed as per manufacturer specification
No | Visit | Complaint | RKK | 電解質 CALドリフト多発 | Na,Ca,Cl のドリフト多発 | Ref電極交換 | SWO-431970 | Assisted Replacement | Electrode | Electrode | Reference |  | Electrolyte CAL drift frequent | Frequent drift of Na, Ca, and Cl | Ref electrode replacement
No | Visit | Complaint | RIND | Calibration aborted | Checked and found inlet flow impeded. | Cleaned waste line, replaced waste pump tube. 2 point calibration performed successfully. | SWO-433821 | Cleaned | Fluidic pathway | Fluidic pathway | Waste tube |  | Calibration aborted | Checked and found inlet flow impeded. | Cleaned waste line, replaced waste pump tube. 2 point calibration performed successfully.
No | Visit | Complaint | RAME | I was notified that QC level 1 was continuously running although it was passing all parameters. | I was informed that one of the analyzers was continuously running QC level 1. | Staff tried rebooting the analyzer and that seemed to correct it. I loaded a set up from another analyzer to reset the QC schedule as well and see if that helps. Issue will be monitored. | SWO-432343 | Defect | SW | RMED SW | RMED SW | QC schedule | I was notified that QC level 1 was continuously running although it was passing all parameters. | I was informed that one of the analyzers was continuously running QC level 1. | Staff tried rebooting the analyzer and that seemed to correct it. I loaded a set up from another analyzer to reset the QC schedule as well and see if that helps. Issue will be monitored.
No | Visit | Complaint | RAME | PCO2 status, sensitivity, drift, and unstable errors during calibrations. | Return with replacement electrode. | Replacing pCO2 electrode corrected problem. Electrode under warranty. | SWO-432278 | Assisted Replacement | Electrode | Electrode | pCO2 |  | PCO2 status, sensitivity, drift, and unstable errors during calibrations. | Return with replacement electrode. | Replacing pCO2 electrode corrected problem. Electrode under warranty.
No | Visit | Complaint | RAME | Neon lamp too low error | Neon lamp too low | I replaced the lamp and made all the adjustments. All cal's passed with no errors | SWO-432880 | Defect | HW | Oxi module | Lamp unit | Neon lamp | Neon lamp too low error | Neon lamp too low | I replaced the lamp and made all the adjustments. All cal's passed with no errors
No | Visit | Complaint | RKOR | MIXWHEEL_25 | Automatic recovery was successful | Automatic recovery was successful | SWO-433915 | Defect | SW | Analyzer Control Module | RMED SW | Motor control | MIXWHEEL_25 | Automatic recovery was successful | Automatic recovery was successful
No | Visit | Complaint | RDEU | 12.10.2018; Fr. Schönebeck; Tel.: 024616202892; Display fällt teilweise aus | Fehler bekannt | Display ersetzt, Funktionskontrolle erfolgreich, weiter mit Halbjahreswartung | SWO-431955 | Defect | HW | Cover and Chassis | Display |  | 12.10.2018; Ms. Schönebeck; Tel.: 024616202892; Display partially fails | Errors known | Display replaced, function control successful, further with half-year maintenance
No | Visit | Complaint | RDEU | Fehler 3264 Errors related to power, temperature, or Time/date missmatctch(>100h bettween internal pc clock and factory clock) 15.10.2018, Frau Otto, Tel: 051916023925: erneut Messabbrüche am AQT mit 1067 und 3050- CorruptXML | erneut Messabbrüche am AQT mit 1067 und 3050- CorruptXML | Neues Pc Modul eingebaut, Funktionskontrolle i. O. , Qualitätskontrollen durchgeführt i. O. | SWO-431960 | Defect | HW | Cover and Chassis | PCB, Main Board | 3264 | Fehler 3264 Errors related to power, temperature, or Time/date missmatctch(>100h bettween internal pc clock and factory clock) 15.10.2018, Frau Otto, Tel: 051916023925: erneut Messabbrüche am AQT mit 1067 und 3050- CorruptXML | again measurement interruptions on the AQT with 1067 and 3050- CorruptXML | New Pc module built-in, function control as a rule, quality controls carried out in the whole
No | Visit | Complaint | RIND | Service Visit. | System message shows restart the analyzer but after restart again the analyzer is going in error. | Open the front panel of the analyzer, remove the dust from the incubator wheel, and storage wheel sensor, Start the analyzer, successfully done, analyzer handed over to staff in well working condition. | SWO-433830 | Cleaned | HW | Incubator Module | sensor, cup | calibrate_13 | Service Visit. | System message shows restart the analyzer but after restart again the analyzer is going in error. | Open the front panel of the analyzer, remove the dust from the incubator wheel, and storage wheel sensor, Start the analyzer, successfully done, analyzer handed over to staff in well working condition.
No | Visit | Complaint | RPOL | Niepowodzenie transportu płynów | Weryfikacja pracy analizatora | Wymiana uszczelki połączenia i uszczelki wlotu. Weryfikacja transportu płynów. | SWO-432244 | Defect | HW | Basic Unit | PM Kit | 1271 | Fluid transport Failure | Validate Analyzer Work | Replacing the connection gasket and inlet seals. Verification of fluid transport.
No | Visit | Complaint | RESP | Pide continuamente cambio de reactivos | Chequeo de circuito neumático y reactivos | Sustitución de sp, Gasket, inlet probé | SWO-436832 | Assisted Replacement | SP | SP | FT |  | Continuous calls for reagent change | Pneumatic and reactive circuit check | Replacing sp, Gasket, inlet tested
Yes | 5 | Y | RCZE | Blank | Geciova dala vypoved, stara sa o analyzator vrchna p. Vrankova |  | D518F6FE9E4C6140 | Re classified RMED | planned SWO |  |  |  | Blank | Geciova gave vypoved, old with O analyzator Vrchna p. Vrankova
No | Visit | Complaint | RSAF | Pump cal error | Pump cal error. Probe blocked. | Performed manual back flush. Unblocked the probe. Verified that the analyzer is fully functional. | SWO-432909 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | Pump cal error | Pump cal error. Probe blocked. | Performed manual back flush. Unblocked the probe. Verified that the analyzer is fully functional.
No | Visit | Complaint | RLTD | Analyser failing to pump solutions across three different solution pack, solution varied between gas and QC’s. | Analyser failing to pump solutions across three different solution pack, solution varied between gas and QC’s. | Checked fluid system, no resistance felt when flushed. This is a new analyser and has not ran any samples yet, it has been sitting in the lab since initial installation, has been up and running for a few days now and has just started to have these issues. Passes all service tests apart from solution pack check, failing on air. Will have to return tomorrow to investigate further. | SWO-435263 | Revisit required | HW | Basic Unit | Actuator unit |  | Analyser failing to pump solutions across three different solution pack, solution varied between gas and QC’s. | Analyser failing to pump solutions across three different solution pack, solution varied between gas and QC’s. | Checked fluid system, no resistance felt when flushed. This is a new analyser and has not ran any samples yet, it has been sitting in the lab since initial installation, has been up and running for a few days now and has just started to have these issues. Passes all service tests apart from solution pack check, failing on air. Will have to return tomorrow to investigate further.
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | LAC定标林敏度偏低 | 更换测试卡后修复 | SWO-435644 | [5 empty] | Test card, reagent package failure | LOW SENSITIVITY OF LAC CALIBRATION FOREST | Repair after replacing test card
No | Visit | Complaint | RCHN | 开不了机，触摸屏破损 | 开不了机，触摸屏破损，借910-285,910-324各一个 | 更换910-285,910-324备件后修复by tim | SWO-431931 | Defect | HW | Power Supply | Power supply |  | Can't open the machine, touch screen is broken | Unable to open, touch screen broken, borrow 910-285, 910-324 each one | Repair by tim after replacement of 910-285,910-324 spare parts
No | Visit | Complaint | RCHN | 固定试剂包的夹子松动 | 固定试剂包的夹子松动，借886-786、886-776、905-959各一个 | 更换886-786、886-776、905-959后修复by tim | SWO-431123 | Defect | HW | Basic Unit | Actuator unit |  | The clip of the fixing reagent bag is loose | The clip of the fixed reagent bag is loose, one each of the 886-786, 886-776, 905-959 | Repair by tim after replacing 886-786, 886-776, 905-959
No | Visit | Complaint | RCHN | 试管夹损坏 | 试管夹损坏，借903-584一个 | 更换903-584备件后修复by tim | SWO-431928 | Defect | HW | Inlet module | Tube holder |  | Damage to test tube clip | Damaged tube clips, borrow edi-903-584 | Repair by tim after replacement of 903-584 spare parts
No | Visit | Complaint | RCHN | AQT 出現error1031 | 出現error1031 | 更換光單 | SWO-433062 | Defect | HW | Process Module | Optical Unit | 1031 | AQT Error1031 | Appearance error1031 | Replace the light sheet
No | Visit | Complaint | RCHN | 报错1024 | 加借定标卡944-212（14774）一张 | 重新定标后修复by tim | SWO-432519 | Defect | HW | Process Module | Optical Unit | 1024 | 1024 error | Plus Loan Bid Card 944-212 (14774) One | Fix by tim after re-calibration
No | Visit | Complaint | RIND | Calibration error | Checked the data log pH,Na,Ca calibration drift out of range replaced new reference membrane and checked then also same problem crosschecked reference electrode found working confirmed reference electrode is defective | Need to replace reference electrode | SWO-434075 | Revisit required | Electrode | Electrode | Reference |  | Calibration error | Checked the data log pH,Na,Ca calibration drift out of range replaced new reference membrane and checked then also same problem crosschecked reference electrode found working confirmed reference electrode is defective | Need to replace reference electrode
No | Visit | Complaint | RAME | Customer is reporting a reoccurrence of an Oximetry error neon lamp intensity too low. | Reviewed data logs and service diagnostics. Ran complete Oxi test, neon lamp failed, intensity @ 19%, must be >=50%. Suspect possible issue with power supply causing repeated failure of Oximetry components. Also asked customer to investigate power coming from UPS to ensure it is not a possible contributor of the frequent issues. | Replaced lamp unit & power supply. Completed all required adjustments and checks. All tests now passing, no further errors observed. Asked Diana Crary to check with UCLA Biomed or Powervar and request that the power coming from the Powervar UPS to the analyzer be monitored for 2 to 3 weeks for any fluctuations or disturbances that may be contributing to the frequent sporadic failure of electronics in this analyzer. | SWO-432583 | Defect | HW | Oxi module | Lamp unit | Neon lamp | Customer is reporting a reoccurrence of an Oximetry error neon lamp intensity too low. | Reviewed data logs and service diagnostics. Ran complete Oxi test, neon lamp failed, intensity @ 19%, must be >=50%. Suspect possible issue with power supply causing repeated failure of Oximetry components. Also asked customer to investigate power coming from UPS to ensure it is not a possible contributor of the frequent issues. | Replaced lamp unit & power supply. Completed all required adjustments and checks. All tests now passing, no further errors observed. Asked Diana Crary to check with UCLA Biomed or Powervar and request that the power coming from the Powervar UPS to the analyzer be monitored for 2 to 3 weeks for any fluctuations or disturbances that may be contributing to the frequent sporadic failure of electronics in this analyzer.
No | Visit | Complaint | RDEU | 18.10.2018; Fr. Leppla; Tel.: 06232221450; Setsafe_22 | Setsafe 22 | Storage Wheel wieder im Gang gesetzt. Führungschiene war herausgefallen. Flachbandkabel der Hct- Sensor war ebenfalls heraus gefallen. Rollenpumpen P4 und P5 ausgetauscht. Systemüberprüfung. | SWO-433061 | Adjustment | HW | Storage Module | Cartridge guide |  | 18.10.2018; Ms. Leppla; Tel.: 06232221450; Setsafe_22 | Setsafe 22 | Storage Wheel back in motion. Leadership rail had fallen out. Ribbon cable of the Hct sensor had also fallen out. roller pumps P4 and P5. System check.
No | Visit | Complaint | RKK | 圧力吸引テストエラーが多発する。流体系のエラーが多発する。QCレベル3でpO2範囲外 | 機器状況確認 | 本体引き上げ、代替機設置作業 | SWO-437254 | Revisit required | Analyzer replaced | Analyzer replaced | FT | 1063 | Frequent pressure suction test errors occur frequently. There are a lot of errors of the flow system. QC level 3 out of pO2 range | Equipment status check | Body pulling and installation work of replacement machines
No | Visit | Complaint | RDEU | 19.10.2018; Sr. Silvia; Tel.: 0376544265; Intervention erforderlich, Ableitstrom SC erkannt | Fehlersuche ==> Fehler 0963: Ableitstrom im Analysator erkannt | Verschleißteile, Hemolyser unit und Quetschventil erneuert | SWO-433968 | Defect | HW | Oxi module | Hemolyzer | 963 | 19.10.2018; Sr. Silvia; Tel.: 0376544265; Intervention required, leake stream SC detected | Troubleshooting ==> Error 0963: Leakage current detected in the analyzer | Wear parts, hemolyser unit and crushing valve renewed
No | Visit | Complaint | RLTD | Program cannot continue on screen, status code error 2 Lilian Odoch 0203 4470326 Data upgrade failed, status code 2 - Anthony Pacittict on 07894 149683 19/10/18 | Data upgrade failed status code 2 | Reformatted the flash card,reloaded windows 7. Loaded analyser software and reinstalled analyser setup. Reinstalled consumable analyser start up ok. Calibration,QC ok. | SWO-433722 | Defect | SW | RMED SW | Status code 2 | status code 2 | Program cannot continue on screen, status code error 2 Lilian Odoch 0203 4470326 Data upgrade failed, status code 2 - Anthony Pacittict on 07894 149683 19/10/18 | Data upgrade failed status code 2 | Reformatted the flash card,reloaded windows 7. Loaded analyser software and reinstalled analyser setup. Reinstalled consumable analyser start up ok. Calibration,QC ok.
No | Visit | Complaint | RDEU | 18.10.2018; Fr. Hegerich; Tel.: 091318539119; tHb Kal ist fällig | tHb-Kal. fällig | tHb-Kal. durchgeführt, tHb-Kal. i.O. | SWO-435191 | User Error | HW | Oxi module | tHb Cal |  | 18.10.2018; Ms. Hegerich; Tel.: 091318539119; tHb Kal is due | tHb-Kal. Due | tHb-Kal. durchgeführt, tHb-Kal. i.O.
No | Visit | Complaint | RFRA | Erreur de joint d’introduction 2 fois en 10 jours. | Erreur 1189 erreur changement d’état DL uniquement sur capillaire. Demande de passage SAV pour vérification avant week-end end car pbm déjà survenu semaine dernière. Avec changement du joint d’introduction. | Vérification mécanique du module inlet ok. Ressort ok. Joint d’intro humidifié ok. (Neuf du matin) Changement en prévention du joint de connection (consommable) Vérification des erreurs des derniers jours : erreur 1189 d’aspiration sur passage patient capillaire. Pas d’erreur lors des calibrations ou contrôles. Pack solution neuf du matin, fait du bruit (a surveiller) Nettoyage fluidique ok. Passage calibration ok. Passage de tous les CQ ok. Tous les paramètres disponibles ok. Passage d’un capillaire avec infirmière ok. Pas d’erreur. Mail à j.aurran@chu-nimes.fr | SWO-433870 | User Error | preanalytical | preanalytical | preanalytical | 1189 | Introduction seal error 2 times in 10 days. | Error 1189 dL state change error only on capillary. Request for passage SAV for verification before weekend because pbm already occurred last week. With change of the introductory seal. | Mechanical check of the inlet module ok. Spring ok. Intro humidified joint ok. (Nine in the morning) Change in connection seal prevention (consumable) Checking the errors of the last days: 1189 suction error on patient hair passage. No errors during calibrations or controls. Pack new solution in the morning, makes noise (to watch) Fluid cleansing ok. Passage calibration ok. Passage of all CQ ok. All available settings ok. Passing a capillary with nurse ok. No mistake. Mail to j.aurran@chu-nimes.fr
No | Visit | Complaint | RAME | Customer called stating battery is at 0% charge. | Found that battery is showing 0% charge | Reset the battery using the reset jumper which corrected the issue. | SWO-433715 | Adjustment | HW | Battery module | Battery module | 0% | Customer called stating battery is at 0% charge. | Found that battery is showing 0% charge | Reset the battery using the reset jumper which corrected the issue.
No | Visit | Complaint | RFRA | ERR STATUS 2 avec écran noir | Contrôle de la CF card et du PC | Reinstallation de Windows7 impossible $Installation d'un nouveau disque dur avec installation de Windows7 impossible $Commande d'un PC complet pour intervention lundi ou mardi. | SWO-434364 | Revisit required | HW | DMS_PC | CPU | status code 2 | ERR STATUS 2 with black screen | Control of the CF card and the PC | Reinstalling Windows7 impossible $Installation a new hard drive with Windows7 installation impossible $Commande a full PC for intervention Monday or Tuesday.
No | Visit | Complaint | RLTD | Sophia Harrow 0203 2998238 ext 38652 | Checked calibration and quality control logs - all calibrations and quality controls are ok. Ran comparison tests for 2 blood samples between analysers in Christine Brown & Frank Stansil CCUs and this analyser. Results for pO2 compared well. Sample 1. Frank Stansil pO2 = 8.19 kPa Lonsdale pO2 = 8.09 kPa Sample 2 Christine Brown pO2 = 11.5 kPa Lonsdale pO2 = 11.5 kPa Analyser is performing correctly, comparison blood gas results confirm accuracy of pO2 measurements. Also performed due tHb calibration. Copies of the sample results are attached to this report and will be emailed to Sophia Harrow together with this report. | No fault found, analyser working correctly. | SWO-433693 | No fault found | No fault found | No fault found | pO2 |  | Sophia Harrow 0203 2998238 ext 38652 | Checked calibration and quality control logs - all calibrations and quality controls are ok. Ran comparison tests for 2 blood samples between analysers in Christine Brown & Frank Stansil CCUs and this analyser. Results for pO2 compared well. Sample 1. Frank Stansil pO2 = 8.19 kPa Lonsdale pO2 = 8.09 kPa Sample 2 Christine Brown pO2 = 11.5 kPa Lonsdale pO2 = 11.5 kPa Analyser is performing correctly, comparison blood gas results confirm accuracy of pO2 measurements. Also performed due tHb calibration. Copies of the sample results are attached to this report and will be emailed to Sophia Harrow together with this report. | No fault found, analyser working correctly.
No | Visit | Complaint | RLTD | Jepsy Thomas 07414 920186 | Checked inlet module - inlet probe still partly bent, straightened completely. Checked inlet gasket and inlet movement all ok. Carried out manual flush - sample pathway is totally clear. Restarted analyser, start up ok without errors. Checked fluid transport in service screen, pump calibration, pressure test, liquid sensors and clot detection values all mid range. Analyser is working correctly. | Straightened bent inlet probe. | SWO-433692 | Adjustment | HW | Auto Inlet module | Probe bend |  | Jepsy Thomas 07414 920186 | Checked inlet module - inlet probe still partly bent, straightened completely. Checked inlet gasket and inlet movement all ok. Carried out manual flush - sample pathway is totally clear. Restarted analyser, start up ok without errors. Checked fluid transport in service screen, pump calibration, pressure test, liquid sensors and clot detection values all mid range. Analyser is working correctly. | Straightened bent inlet probe.
No | Visit | Complaint | RLTD | Issue raised by Jepsy Thomas in POC for Neonatal Unit whilst on site. | Checked analyser performance, calibrations and quality controls all ok. Obtained blood sample from General ICU. Performed analysis of this sample on General ICU ABL90 and Neonatal ABL90. Electrolyte results including Na showed no difference. Analyser deemed to be working correctly. | Comparative test with blood sample showed no fault present. | SWO-435145 | No fault found | SC | SC | Na |  | Issue raised by Jepsy Thomas in POC for Neonatal Unit whilst on site. | Checked analyser performance, calibrations and quality controls all ok. Obtained blood sample from General ICU. Performed analysis of this sample on General ICU ABL90 and Neonatal ABL90. Electrolyte results including Na showed no difference. Analyser deemed to be working correctly. | Comparative test with blood sample showed no fault present.
No | Visit | Complaint | RCHN | 接口板腐蚀 | 接口板腐蚀，借902-972一个 | 备件未使用，更换测试卡后修复by tim | SWO-431858 | [5 empty] | Interface plate corrosion | Interface plate corrosion, borrow 902-972 one | Spare parts not used, repair by tim after replacement test card
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 测试卡断脚 | 更换测试卡后修复 | SWO-433993 | [5 empty] | Test card, reagent package failure | Test card broken foot | Repair after replacing test card
No | Visit | Complaint | RCHN | 测试卡漏液，导致接口板腐蚀 | 测试卡漏液，导致接口板腐蚀,需要借件902-972 | 更换902-972后修复by tim | SWO-431907 | [5 empty] | Test card leakage, resulting in interface plate corrosion | Test card leakage, resulting in interface plate corrosion, need to borrow 902-972 | Repair by tim after replacement 902-972
No | Visit | Complaint | RDEU | 18.10.2018, Sr. Juliana, Tel: 0393869780: LP lässt sich nicht installieren | Windowsdatum falsch | Fehlersuche, Datumskorrektur auf Windowsebene, Installation von Lösungspack und Sensorkassette, Funktionskontrolle | SWO-433372 | [5 empty] | 18.10.2018, Sr. Juliana, Tel: 0393869780: LP cannot be installed | Windows date incorrect | Troubleshooting, Date Correction at the Windows Level, Installation of Solution Pack and Sensor Cassette, Function Control
No | Visit | Complaint | RCHN | 测试卡验证失败 | 测试卡验证失败 | 拆库存机的903-334后修复（By Ran） | SWO-434023 | [5 empty] | Test card validation failed | Test card validation failed | 903-334 post-repair of the stock-removal machine (By Ran)
No | Visit | Complaint | RAME | SC verification check failed on multiple cassettes | Troubleshot to a faulty sensor connector board or analog board. | Replaced the sensor connector board and the analog board, and this resolved the problem. | SWO-433716 | [5 empty] | SC verification check failed on multiple cassettes | Troubleshot to a faulty sensor connector board or analog board. | Replaced the sensor connector board and the analog board, and this resolved the problem.
No | Visit | Complaint | RCHN | lac和glu定标失败，报错1310 | lac和glu定标失败，报错1310,需要借件903-581一个 | 更换903-581后修复by tim | SWO-431876 | Defect | HW | Basic Unit | PCB, Sensor interface | Glu/Lac | lac and glu bid failed, error 1310 | lac and glu bid failure, error 1310, need to borrow 903-581 one | Repair by tim after replacement 903-581
No | Visit | Complaint | RCHN | 测试卡不能正常打开 | 测试卡不能正常打开，借925-017一个 | 更换925-017一个正常。by tim | SWO-431126 | Defect | HW | Basic Unit | Sensor door |  | The test card does not open properly | The test card does not open properly, borrow 925-017 one | Replace 925-017 one normal. by tim
No | Visit | Complaint | RCHN | 测试卡不能正常打开 | 测试卡不能正常打开，借925-017一个 | 更换925-017一个正常。by tim | SWO-431134 | Defect | HW | Basic Unit | Sensor door |  | The test card does not open properly | The test card does not open properly, borrow 925-017 one | Replace 925-017 one normal. by tim
No | Visit | Complaint | RCHN | 测试卡扫描失败 | 测试卡扫描失败,借910-304一个。 | 更换库存机910-304一个修复。by tim | SWO-433832 | Defect | HW | Storage Module | Barcode reader, storage wheel |  | Test card scan failed | The test card scan failed, borrowing 910-304 one. | Replace the stock machine 910-304 with a repair. by tim
No | Visit | Complaint | RCHN | 仪器死机导致定标失败 | 仪器死机导致定标失败,借944-268（14070）一张。对应批号缺货 | 重新定标正常。by tim | SWO-432733 | Assisted Replacement | Consumables | Test Kit | Cartridge, cal |  | Instrument crash causes calibration failure | The instrument crash caused the calibration to fail, borrowing 944-268 (14070). Out of stock for corresponding lot number | Recalibration is normal. by tim
No | Visit | Complaint | RCHN | 报错3410 | 报错3410,需要借件 944-212（14868） | 重新定标后修复by tim | SWO-431899 | Assisted Replacement | Consumables | Test Kit | Cartridge, cal |  | Error 3410 | 3410, need to borrow 944-212 (14868) | Fix by tim after re-calibration
No | Visit | Complaint | RCHN | BNP传感器损坏 报错1235 | BNP传感器损坏 报错1235 借905-908一个 | 更换905-908一个正常。by tim | SWO-433312 | Defect | HW | Process Module | Sensor, BnP | 1235 | BNP sensor damage 1235 | BNP sensor damage 1235 debit 905-908 one | Replace 905-908 one normal. by tim
No | Visit | Complaint | RCHN | 报错5689 | 报错5689，借903-029和636-532各一个 | 测试卡卡住的原因。把测试卡拿出来调整了一下位置。by tim | SWO-433261 | Defect | Consumables | Test Kit | Cartridge, Broken |  | 5689 error | 5689, borrowing 903-029 and 636-532 each | Test the reason for the card's stay. Take out the test card and adjust the position. by tim
No | Visit | Complaint | RAME | Ca Sensitivity @107.00 Kate - 708-216-3428 | Upon arrival found instrument went running sensitivity was down to 99% | It took more than 8 hours and used electrode for it to work. Put new electrode from customer stock and used membrane I ordered for troubleshooting purpose and after two cals everything went green and sensitivity was down to 99%. Possible customer had bad box of membrane. All calibrations and 3 Levels of QC passed as per manufacturer specification | SWO-433754 | Assisted Replacement | Electrode | Membrane | Ca |  | Ca Sensitivity @107.00 Kate - 708-216-3428 | Upon arrival found instrument went running sensitivity was down to 99% | It took more than 8 hours and used electrode for it to work. Put new electrode from customer stock and used membrane I ordered for troubleshooting purpose and after two cals everything went green and sensitivity was down to 99%. Possible customer had bad box of membrane. All calibrations and 3 Levels of QC passed as per manufacturer specification
No | Visit | Complaint | RDEU | 17.10.2018; Hr. Jutrzenka; Tel.: 076127024670; Parameter nicht installiert | Demo Keys abgelaufen | Demo Keys aufgespielt, Gerät einkalibriert, QC durchgeführt, i.O. | SWO-433429 | User Error, FSE | Demo key | Demo key | Demo key |  | 17.10.2018; Mr. Jutrzenka; Tel.: 076127024670; Parameters not installed | Demo Keys expired | Demo keys played, device calibrated, QC performed, i.O.
No | Visit | Complaint | RFRA | Inspection SAV | Inspection SAV suite à résultat discordant | Inspection SAV selon procédure ok. Renvoi de toutes les infos ok. Passage de CQ MCC1 et 2 sur TnT 2fois ok. | SWO-433869 | No fault found | No fault found | No fault found | No fault found |  | SAV inspection | Inspection sAV following discordant result | Inspection SAV according to procedure ok. Referral of all the info ok. Passage of CQ MCC1 and 2 on TnT 2times ok.
No | Visit | Complaint | RPOL | Błąd transportu płynów | Błąd transportu płynów. Błąd selektora przepływów. | Czyszczenie selektora przepływów . Usunięcie zatoru w module inlet. Aktualizacja oprogramowania 3.4 .Kalibracja hemoglobiny | SWO-433476 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway | Clot | Fluid Transport Error | Fluid transport error. Flow selector error. | Cleaning the flow selector. Removal of the congestion in the inlet module. 3.4 Software update. Hemoglobin Calibration
No | Visit | Complaint | RCHN | CA,NA离子定标失败 | CA,NA离子定标失败，借902-972一个 | 更换902-972一个正常。by tim | SWO-432591 | [5 empty] | CA, NA ion calibration failed | CA, NA ion calibration failed, borrowed 902-972 one | Replace 902-972 one normal. by tim
No | Visit | Complaint | RCHN | 混匀错误，报错4035、4036 | 混匀错误，报错4035、4036，需要借件886-730、924-824各一个 | 更换886-730、924-824各一个正常。by tim | SWO-432726 | Defect | HW | Inlet module | Inlet Wheel | 4035, 4036 | Mixing error, report error 4035, 4036 | Mixing error, report error 4035, 4036, need to borrow 886-730, 924-824 each one | Replace 886-730, 924-824 each normal. by tim
No | Visit | Complaint | RCHN | 干燥前温度超出范围，报错代码：1068 | 干燥前温度超出范围，报错代码：1068，需要借件903-210一个 | 更换903-210一个正常。by tim | SWO-432717 | Defect | HW | Process Module | PCB, cup dryer | 1068 | Temperature out of range before drying, error code: 1068 | Temperature out of range before drying, error code: 1068, need to borrow 903-210 one | Replace the 903-210 one normal. by tim
No | Visit | Complaint | RIND | Rinse error. | Check the logs & activity logs. | Check the rinse & replace the rinse. calibration fail & rinse error. Check chamber & replace the ph ring. Replace the inlet tubing & check the LS. Rinse valve fail to work. Clean the valve & check the valve & run test. valve not working replace the valve & calibrate the system. | SWO-433931 | Defect | HW | Fluid manifold | Valve |  | Rinse error. | Check the logs & activity logs. | Check the rinse & replace the rinse. calibration fail & rinse error. Check chamber & replace the ph ring. Replace the inlet tubing & check the LS. Rinse valve fail to work. Clean the valve & check the valve & run test. valve not working replace the valve & calibrate the system.
No | Visit | Complaint | RAME | Analyzer failing LIS adjust. | Found liquid in fluidic module not flowing properly. | Cleaned all valves and tubing. Replaced inlet. | SWO-433817 | Defect | HW | Inlet module | Inlet module |  | Analyzer failing LIS adjust. | Found liquid in fluidic module not flowing properly. | Cleaned all valves and tubing. Replaced inlet.
No | Visit | Complaint | RIND | Setsafe_22 | Checked and found Inlet door failed in set safe program. | Realigned inlet door and run set safe program successfully. | SWO-433956 | Adjustment | HW | Cover and Chassis | Door, Inlet |  | Setsafe_22 | Checked and found Inlet door failed in set safe program. | Realigned inlet door and run set safe program successfully.
No | Visit | Complaint | RCHN | 试管支架损坏 | 试管支架损坏,借903-584一个。 |  | SWO-433929 | Defect | HW | Inlet module | Tube holder |  | Damage to test tube support | The test tube support is damaged, borrow 903-584 one.
No | Visit | Complaint | RCHN | 新机器进样门开关不利索 | 样本舱门开关不灵敏，需要借件886-638、886-639各一个 | 更换886-638、886-639各一个正常。by tim | SWO-432774 | Defect | HW | Cover and Chassis | Door, Inlet |  | New machine feed door switch adverse rope | Sample door switch is not sensitive, need to borrow 886-638, 886-639 each one | Replace 886-638, 886-639 each normal. by tim
No | Visit | Complaint | RDEU | 19.10.2018; Fr. Gerick; Tel.: 06213832504; Intervention erforderlich, Ableistrom SC erkannt 22.10.2018, 10:44 Uhr, Fr. Gerick, Tel. 06213832504: Der Fehler besteht weiter, noch kein Servicebesuch erfolgt. | Ableitströme erkannt | Allgemeine Überprüfung etc. - Erneuerung SP (Reklamation erfolgt) - Funktionsüberprüfung etc. | SWO-434551 | Assisted Replacement | SP | SP | Leak current | 963 | 19.10.2018; Ms. Gerick; Tel.: 06213832504; Intervention required, Ableistrom SC detected 22.10.2018, 10:44 a.m., Ms. Gerick, tel. 06213832504: The error persists, no service visit has yet taken place. | Discharge currents detected | General check etc. - Renewal SP (complaint is made) - Function check etc.
No | Visit | Complaint | RDEU | 22.10.2018, 08:41 Uhr, Fr. Jung, Tel. 089476036: Leckage erkannt | Leckage erkannt | Pinch Valve überprüft, Jahreswartung durchgeführt. Oxi Complete Test ok. ABL manuell gespült. tHb Kalibration, Kalibrationen und QC's vorgenommen | SWO-434417 | Defect | HW | Basic Unit | PM Kit | Leak detected | 22.10.2018, 08:41, Fr. Jung, Tel. 089476036: Leakage detected | Leakage detected | Pinch Valve checked, annual maintenance performed. Oxi Complete Test ok. ABL manually rinsed. tHb calibration, calibrations and QC's.
No | Visit | Complaint | RFRA | ERR STATUS 2 avec écran noir | PC | Préparation d'un PC unit à l'atelier $Installation de Windows 7 $Installation de l'application Radiometter $Installation du patch sécurité $Restauration de la configuration de la machine $Configuration des données informatique pour la connexion réseau $Modification du run cmd pour reconnexion au lecteur réseau $Impossible de d'importer la base de données $ Remise en route $Calibration et passage des CQ.Erreur de Calibration du Lac attendre 1heure pour la stabilisation sinon remplacer la cassette de capteurs. | SWO-434438 | Defect | HW | DMS_PC | CPU | status code 2 | ERR STATUS 2 with black screen | PC | Preparing a PC unites at the $Installation workshop of Windows 7 $Installation of the Radiometter app $Installation the security patch $Restauration machine configuration $Configuration computer data for the network connection $ Changing the run cmd to reconnect to the network drive $Impossible import the $database Restarted $Calibration and the passage of the CQ. Lake Calibration error wait 1 hour for stabilization if not replace the sensor cassette.
No | Visit | Complaint | RLTD | Analyser failing to pump solutions across three different solution pack, solution varied between gas and QC’s. | Analyser failing to pump solutions across three different solution pack, solution varied between gas and QC’s. Ran out of time on last visit. | Stripped down analyser and inspected, all looked ok, replaced service parts, issue still remained. On removing solution pack in service mode analyser makes a slightly louder sound than normal. Stripped down analyser again and found that the motor spindle which operates the solution pack release mechanism and engages the solution pack pump was incorrectly fitted and was catching on the case, see attached photo. Adjusted the spindle and clevis to the required length and retested in service mode, | SWO-435264 | Defect | HW | Basic Unit | Actuator unit | New Analyzer | Analyser failing to pump solutions across three different solution pack, solution varied between gas and QC’s. | Analyser failing to pump solutions across three different solution pack, solution varied between gas and QC’s. Ran out of time on last visit. | Stripped down analyser and inspected, all looked ok, replaced service parts, issue still remained. On removing solution pack in service mode analyser makes a slightly louder sound than normal. Stripped down analyser again and found that the motor spindle which operates the solution pack release mechanism and engages the solution pack pump was incorrectly fitted and was catching on the case, see attached photo. Adjusted the spindle and clevis to the required length and retested in service mode,
No | Visit | Complaint | RCHN | 进样门无法关闭，报错5784 | 进样门无法关闭，报错5784，需要借件903-028一个 |  | SWO-434116 | Defect | HW | Cover and Chassis | PCB, cabinet top | 5784 | The entrance door could not be closed, 5784 was reported. | The entrance door can not be closed, the error 5784, need to borrow 903-028 one
No | Visit | Complaint | RIND | Instrument frequently performing system check. | Found instrument operating temperature high | Checked the instrument for errors, performed the zeroing for sensor connector. Reinstalled the SP and SC,kept instrument under observation . Presently instrument working fine. | SWO-437920 | [5 empty] | Instrument frequently performing system check. | Found instrument operating temperature high | Checked the instrument for errors, performed the zeroing for sensor connector. Reinstalled the SP and SC,kept instrument under observation . Presently instrument working fine.
No | Visit | Complaint | RDEU | Med Tech Herr Stein meldet Oxymetrie Hardware Fehler und bittet um Reparatur. | Gerät überprüft, bei Kal Meldung; Halogen Lamp Intensity too high. | Service Test aufgerufen, Lampen Intensität auf 3161 eingestellt, Oxi Complete Test durchgeführt, Kalibriert, tHb Kal, QCs durchgeführt. Funktionsprüfung i.O Kunden informiert. | SWO-434705 | Adjustment | HW | Oxi module | Lamp unit | Halogen | Med Tech Mr. Stein reports oxymetry hardware errors and asks for repair. | Device checked, in case of Kal message; Halogen Lamp Intensity too high. | Service test called, lamp intensity set to 3161, Oxi Complete test performed, Calibrated, tHb Kal, QCs performed. Functional test in the informed customers.
No | Visit | Complaint | RDEU | 22.10.2018 9:00 Fr. Schneider 026211712031: Gerät steht in der Software-Hochlaufroutine | AU Sw fährt nicht hoch. | AQT mit genommen. | SWO-434980 | Revisit required | HW | Analyzer Control Module | CF card |  | 22.10.2018 9:00 Ms. Schneider 026211712031: Device is in the software boot routine | AU Sw does not drive up. | AQT.
No | Visit | Complaint | RBNL | Moet reagens pack verwisselen maar nieuwe pack lijkt geblokkeerd. Wil niet ver genoeg insteken. Ook oude pack nu problemen. | Beide packs van lot BR-69. Kan nu wel induwen, geen blokkade. Meerdere malen geprobeerd in servicemenu, blijft soms iets steken. Linkerzijde iets optillen. TNT calibratie gestart, lot14708. | Reagenspack geiïnstalleerd. | SWO-434501 | Defect | Consumables | Solution Pack | Solution Pack | Valve | Must swap reagent pack but new pack seems blocked. Do not want to enter far enough. Also old pack now problems. | Both packs of Lot BR-69. Can now be Induwen, no blockage. Tried several times in service menu, sometimes remains something sticking. Slightly on the left. TNT calibration started, lot14708. | Reagenspack geiïnstalleerd.
Yes | Visit | Complaint | RDEU | Oximetrie-Hardwareproblem | Oximetrie-Hardwareproblem | WSC-PCB erneuert. | SWO-435098 | Re classified RMED | installation report |  |  | Move to OOB trending | Oximetrie-Hardwareproblem | Oximetrie-Hardwareproblem | WSC-PCB erneuert.
No | 7 | Y | RCZE | Blank |  |  | D51F8CF7EC120261 | Adjustment | HW | Auto Inlet module | Probe bend |  | Blank
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 定标失败 | 更换试剂包后修复 | SWO-434913 | [5 empty] | Test card, reagent package failure | Bid failed | Repair after replacement of reagent kit
No | Visit | Complaint | RSCH | Mögliche Verstopfung, Luftblasenerkennung | Mögliche Verstopfung, Luftblasenerkennung, Sensorkassette nicht initialiesiert | Allgemeine Kontrolle ok. Alle Schläuche ers ok. Zero und Therm i.o. System i.o. Man. QA Zyklus i.o. | SWO-435032 | [5 empty] | Possible constipation, air bubble detection | Possible blockage, air bubble detection, sensor cartridge not initialized | General control ok. All the hoses he s ok. Zero and Therm i.o. System i.o. somebody. QA cycle i.o.
No | Visit | Complaint | RCHN | 报错1326 测试卡温度加热不够 | 报错1326 测试卡温度加热不够 | 更换备用件903-581后修复（By Ran） | SWO-435706 | Defect | HW | Basic Unit | PCB, Sensor interface | 1326 | Misquoted 1326 Test card temperature heating is not enough | Misquoted 1326 Test card temperature heating is not enough | Repair after replacement of spare parts 903-581 (By Ran)
No | Visit | Complaint | RCHN | 光学单元的高低管上的连接管断 | 光学单元的高低管上的连接管断,借905-908一个。 | 更换905-908后修复（By Ran） | SWO-434178 | Defect | HW | Process Module | Sensor, BnP | Potential FSE User Error | Connection tube break on the high and low tube of the optical unit | The connection tube on the high and low tube of the optical unit is broken, with a 905-908 one. | Repair after replacement 905-908 (By Ran)
No | Visit | Complaint | RCHN | 报错5784 测试卡断裂 | 报错5784 测试卡断裂 | 更换测试卡后修复 | SWO-434942 | Defect | Consumables | Test Kit | Cartridge, Broken |  | Reported error 5784 test card break | Reported error 5784 test card break | Repair after replacing test card
No | Visit | Complaint | RIND | Hazy print out coming out. | Hazy print out coming out | Replace the printer head with the new one,Check the print out & it was working properly. | SWO-437337 | Defect | HW | Printer Module | Printer Head |  | Hazy print out coming out. | Hazy print out coming out | Replace the printer head with the new one,Check the print out & it was working properly.
Yes | 6 | Y | RPAC | Blank | Annual Service |  | D51F40A3EC12FF9D | Re classified RMED | planned SWO |  |  |  | Blank | Annual Service
No | Visit | Complaint | RIND | Printer not workinig | Found that printer D lock is broken. | Need to change printer | SWO-438298 | Revisit required | HW | Printer Module | Printer Head |  | Printer not workinig | Found that printer D lock is broken. | Need to change printer
No | 3 | Y | RCZE | Blank | pCO2 mimo limit v kalibraciach aj QC |  | D51F7A9DEC305CCA | Assisted Replacement | Electrode | Electrode | pCO2 |  | Blank | pCO2 out limit in Calibraciach AJ QC
No | Visit | Complaint | RAME | Display is changing colors | Screen Changing colors. | Reseated cable and tighter the cable screws. Changed cables and disappears. Will order new cables. 910-309 Serial Number of new display 0807AH00176 | SWO-435384 | Defect | HW | Display | Display | mail: what changed? Return to RMED! | Display is changing colors | Screen Changing colors. | Reseated cable and tighter the cable screws. Changed cables and disappears. Will order new cables. 910-309 Serial Number of new display 0807AH00176
No | Visit | Complaint | RDEU | 22.10.2018 9:00 Fr. Schneider 026211712031: Gerät steht in der Software-Hochlaufroutine | Daten Gesichert, AU Installation wird abgebrochen>> Windows installiert, Au Installation wird abgebrochen>> CF Karte ersetzt>> Win. und AU SW installiert. Daten zurückgespielt, AQT SW fährt nicht mehr hoch, Win. und AU SW neu installiert, Konfiguration übertragen. LQC Degenblätter runtergeladen und eingescannt. | D-Dim. Kalibriert. | SWO-435184 | Defect | HW | Analyzer Control Module | CF card |  | 22.10.2018 9:00 Ms. Schneider 026211712031: Device is in the software boot routine | Data Secured, AU installation is canceled>> Windows installed, Au installation is canceled>> CF card replaced>> Win. and AU SW. Data played back, AQT SW no longer boots up, Win. and AU SW reinstalled, transfer configuration. LQC sheets downloaded and scanned. | D-Dim. Calibrated.
No | Visit | Complaint | RPOL | Zgłoszono awarie odczynników. | Zgłoszono awarie odczynników. | Zweryfikowanie sytuacji. Pakiet odczynników przekroczył okres ważności w analizatorze. Konieczna wymiana odczynników. | SWO-435189 | User Error | SP | SP | SP | Lifetime exceeded | Reagent failures have been reported. | Reagent failures have been reported. | Verify the situation. The reagent package exceeded the validity period in the analyzer. Replacement of reagents required.
No | Visit | Complaint | RPOL | Problem przepływowy w analizatorze. Zgięta igła aspiracyjna. | Błąd przepływowy w analizatorze wynikający z wygięcia igły aspiracyjnej. | Diagnoza problemu, wymiana igły aspiracyjnej, wymiana kasety sensorycznej, testy serwisowe. | SWO-435188 | Defect | HW | Auto Inlet module | Probe bend |  | Flow Problem in analyser. Bent aspiration needle. | Flow error in analyser resulting from aspiration needle bending. | Diagnosis of problem, replacement of aspiration needle, replacement of sensory cassette, service tests.
No | Visit | Complaint | RESP | Detectada corriente de fuga | Comprobación del estado de funcionamiento del equipo | Comprobación de tubos, módulo Oxi, módulo de entrada, se sustituye el tubo de válvula, la junta de entrada y el gasket, se comprueba el correcto funcionamiento del equipo. | SWO-435840 | Defect | HW | Basic Unit | PM Kit | 964 | Leak current detected | Checking the operating status of the equipment | Tube check, Oxi module, inlet module, valve tube, inlet joint and gasket are replaced, the correct operation of the equipment is checked.
Yes | 3 | Y | RPAC | Blank | Replacement of consumable failure |  | D52072EB622E850D | [5 empty] | Blank | Replacement of consumable failure
No | Visit | Complaint | RDAN | Det er et helt nyt apparat som er mindre en en uge gammel og som inden installation allerede debuterede med indgriben nødvendig hvor der skulle skylles igennem uden der har været kørt noget materiale. De har fra labs side gentagne gange måtte over at skylle igennem og også udskifte sensor kassette efter konferering med os om mulg fejl i sensor kassette. De er ret så irrerteret over situationen da det er et helt nyt apparat. | Ventil tjek Tryk test | OXI ventil og slange skiftet Inlet forbindelses pakning, inlet probe og inlet pakningsholder skiftet | SWO-435809 | Defect | HW | Basic Unit | PM Kit | 1350; CDR | It is a completely new apparatus which is less than a week old and, prior to installation, already debuted with the intervention needed where to rinse through without having been running any material. From the labs side, they repeatedly had to flush through and also replace the sensor cassette after consulting with us on MULG faults in the sensor cassette. They are quite as irrerated about the situation as it is a completely new apparatus. | Valve Check Pressure test | OXI Valve and hose change Inlet connection gasket, inlet probe and inlet gasket holder SHIFT
No | Visit | Complaint | RDEU | Fehlerbehebung erforderlich | im Rahmen von Test zur Lösungsfindung derWlan Anbindungsproblematik durch Hr. Fordran Kamm es zu Geräte Fehlermeldung Data upgrade Failed Status Code 2, Gerät startet Apllikationssoftwarenicht | Gerät mittels Restkopf heruntergefahren, vom Netzgetrennt und aus gelassen, Neustart -> keine Verbesserung, Applikationssoftware aus Win Explorer heraus neu installiert, alle Haken zur Datenerhaltung gesetzt, keine Verbesserung, erneut Applikationssoftware installiert dabei Datenbank & roma.ini gelöscht, Gerät startet, LP & SC ersetzt ( gleiches Material nochmal eingesetzt), Startphase abgewartet, Konfigurationsdaten aus Dasi Ordner wieder hergestellt Kal & QC Daten gemäß RDEU Anweisung gesichert und dokumentiert | SWO-436428 | Defect | SW | RMED SW | Roma.ini |  | Troubleshooting required | as part of a test to find the solution of thewlan connection problem by Mr. Fordran Kamm's to devices error message Data upgrade Failed Status Code 2, device does not start apllication software | Device shut down by residual head, disconnected from the mains and left, restart -> no improvement, application software reinstalled from Win Explorer, all hooks set for data preservation, no improvement, reinstalled application software deleted database & roma.ini, device starts, LP & SC replaced (same material used again), start phase waited, configuration data from Dasi folder restored Kal & QC data backed up and documented in accordance with RDEU instructions
No | Visit | Complaint | RAME | Customer ran capillary sample and there was an error. | Customer ran capillary sample and there was an issue. | Downloaded and reviewed data logs. Ran calibrations and service test to verify proper operation of analyzer. | SWO-436091 | No fault found | No fault found | No fault found | No fault found | Capillary | Customer ran capillary sample and there was an error. | Customer ran capillary sample and there was an issue. | Downloaded and reviewed data logs. Ran calibrations and service test to verify proper operation of analyzer.
No | Visit | Complaint | RLTD | Unable to pump solutions |  |  | SWO-436444 | Unknown | Unknown | Unknown | FT | mail: more info, what has been doe to solve the issue. | Unable to pump solutions
No | Visit | Complaint | RCHN | 多参数定标失败 | 多参数定标失败，需要借件902-972和903-334各一个 | 更换902-972和903-334后修复（By Ran） | SWO-434172 | [5 empty] | Multi-parameter calibration failed | Multi-parameter calibration failure, need to borrow 902-972 and 903-334 each one | Replacement 902-972 and 903-334 post-repair (By Ran)
No | Visit | Complaint | RCHN | BNP定标失败 | BNP定标失败借944-258（14178）一张。 | 重新定标后修复（By Ran） Recal，then OK。 | SWO-435046 | Defect | Consumables | Test Kit | Cartridge, Broken | 5689 | BNP bid failed | BNP bid failure borrowed 944-258 (14178). | Re-calibration repair (By Ran) Recal, then OK.
No | Visit | Complaint | RCHN | HCT值无法正确测量 | HCT值无法正确测量，借903-329一个 | 更换903-329后修复(By Ran) | SWO-434185 | Defect | HW | Aspiration Module | Needle Wash Unit |  | HCT value sitnoted correctly | HCT value sits not measured correctly, with 903-329 | Repair after replacement 903-329 (By Ran)
No | Visit | Complaint | RCHN | 温度错误，报错3624 | 温度错误，报错3624,需要借件636-537、903-138、902-478 | 更换636-537、903-138、902-478后修复(By Ran) | SWO-434190 | Defect | HW | Incubator Module | PCB, incubator | 3624 | Temperature error, error 3624 | Temperature error, error 3624, need to borrow 636-537, 903-138, 902-478 | Replacement 636-537, 903-138, 902-478 post-repair (By Ran)
Yes | Visit | Complaint | RLTD | Hannah McPeak called to say the ABL827M on Oncology states the demo keys are about to expire. 07887 248417 | The Demo key is about to expire which will stop Crea working. | Installed a new 3 month Demo key on the analyser. The rinse bottle was empty so this was replaced. | SWO-436784 | Re classified RMED | planned SWO |  |  |  | Hannah McPeak called to say the ABL827M on Oncology states the demo keys are about to expire. 07887 248417 | The Demo key is about to expire which will stop Crea working. | Installed a new 3 month Demo key on the analyser. The rinse bottle was empty so this was replaced.
No | Visit | Complaint | RKK | ミキサーエラー頻発（再発） | ミキサーエラー頻発（再発） | ミキサートレイモジュール再交換。FLEXQモジュール調整。エアーライン洗浄。 | SWO-437269 | Adjustment | HW | FlexQ | Mixer Tray | 883 | Mixer error supres (recurrence) | Mixer error supres (recurrence) | Mixer tray module re-replaced. FLEXQ module adjustment. Air line cleaning.
No | Visit | Complaint | RKK | QCにて電解質（K,Cl）がバラツク。時々、Clが上限範囲外。Kが下限範囲外になる | QCにて電解質（K,Cl）がバラツク。時々、範囲外になる | 修理完了器、返却。 | SWO-435808 | Unknown | Unknown | Unknown | Unknown | K, Cl QC | In QC, electrolyte (K, Cl) is baratzk. Sometimes Cl is out of the limit. K is outside the lower limit range. | In QC, electrolyte (K, Cl) is baratzk. sometimes goes out of range | Repair completion machine, return.
No | Visit | Complaint | RDEU | 24.10.2018, 12:26 Uhr ==> Anrufer Fr. Bernhardt, Tel.: 06213836751 ==> EL/MET-Lower LS Fehler | 0532 EL $MET Lower LS Error | Überprüfung LS im Service - Demontage EL $MET Modul - Erneuerung Liquid Sensor - Rückmontage und Funktionsüberprüfung etc. | SWO-436062 | Defect | HW | El/Met Module | LS |  | 24.10.2018, 12:26 o'clock ==> Caller Fr. Bernhardt, Tel.: 06213836751 ==> EL/MET-Lower LS Error | 0532 EL /MET Lower LS Error | Checking LS in service - Disassembly EL $MET module - Renewal Liquid Sensor - Backassembly and function verification etc.
No | Visit | Complaint | RAME | Ampoules not being recognized by auto check module | Auto check not recognizing QC ampoules. | Replaced FPC cable, worked again for little bit but went bad again. Patrick has Autocheck Module . He will stop by in the morning to replace it due to instrument is only 7months old and it’s under warranty. All cals passed. Ran manual QC and showed customer how to run manual QC to get by till morning. | SWO-436383 | Revisit required | HW | Autocheck module | FPC cable |  | Ampoules not being recognized by auto check module | Auto check not recognizing QC ampoules. | Replaced FPC cable, worked again for little bit but went bad again. Patrick has Autocheck Module . He will stop by in the morning to replace it due to instrument is only 7months old and it’s under warranty. All cals passed. Ran manual QC and showed customer how to run manual QC to get by till morning.
Yes | 5 | Y | RPAC | Blank | Reference electrode issues |  | D520A77C6259373B | Re classified RMED | Phone support |  |  | Only remote support and sendig of electrode, see e-mail from Dale Gassard | Blank | Reference electrode issues
No | 5 | Y | RCZE | Blank |  |  | D521DB25D85277DB | Defect | Consumables | Solution Pack | Solution Pack | Chip | Blank
No | Visit | Complaint | RIND | Rotate storage_24 error. | check the sysytem message & found Rotate storage_24 error | Clean the storage wheel gear & storage wheel sensor,Restart the the instrument & after start up it was became ready. | SWO-437954 | Cleaned | HW | Storage Module | storage wheel, hub |  | Rotate storage_24 error. | check the sysytem message & found Rotate storage_24 error | Clean the storage wheel gear & storage wheel sensor,Restart the the instrument & after start up it was became ready.
No | Visit | Complaint | RPOL | Sprawdzenie aparatu | Błąd kontroli jakości . | Sprawdzenie aparatu. Wymiana pakietu roztworów. | SWO-436891 | Assisted Replacement | SP | SP | SP |  | Camera Check | Quality control error. | Check the camera. Replacement of the solution package.
No | Visit | Complaint | RAME | Chris reports installation precision study failed | Found report errors | Forwarded to Jim gibbons for correction | SWO-437233 | User Error | Setup | RMED SW | QC Ranges |  | Chris reports installation precision study failed | Found report errors | Forwarded to Jim gibbons for correction
No | Visit | Complaint | RESP | De vez en cuando la PO2 de los controles las mide fuera de rango. | No se encuentra ningun fallo, la causa puede deberse a que el equipo permanece apagado mientras no se usa, y puede causar errores de deriva si se hace el control immediatamente despues de encenderlo. | Comprobar errores en el registro de datos, cambiar SC y SP, calibrar y probar con controles. | SWO-436636 | [5 empty] | Occasionally the PO2 of the controls measures them out of range. | No fault is found, the cause may be because the computer remains off while not in use, and may cause drift errors if the check is done immediately after turning it on. | Check for data logging errors, change SC and SP, calibrate and test with controls.
No | Visit | Complaint | RCHN | 漏液 | 漏液 | 清理管路以及接头后修复（Ran）Cleaned and reconnected the tubes。 | SWO-436420 | [5 empty] | Leak | Leak | Clean the lines and the connector after the repair (Ran) Cleanand and reconnect the tubes.
No | Visit | Complaint | RCHN | 历史传不出数据 | LIS传不出数据，借903-381一个 | 更换数据线正常。（By Ran） Changed the cable, then OK. | SWO-436547 | [5 empty] | History can't pass on data | LIS can't pass the data, borrow 903-381 one | The replacement data line is OK. (By Ran) Changed the cable, then OK.
No | Visit | Complaint | RCHN | 测试卡温度-A | 测试卡温度-A借902-972一个。 | 使用备用机备件902-972后修复(By Ran) | SWO-435810 | [5 empty] | Test card temperature-A | Test card temperature-A borrows 902-972 one. | Repair after use of spare parts 902-972 (By Ran)
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 不均匀的定标三溶液 1084 | 更换试剂包后修复 | SWO-438248 | Assisted Replacement | SP | SP | FT | 1084 | Test card, reagent package failure | Uneven calibration of three solutions 1084 | Repair after replacement of reagent kit
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | GLU质控高于统计范围 0089 | 更换测试卡试剂包正常。 | SWO-437352 | Assisted Replacement | SC | SC | Glu | QC | Test card, reagent package failure | GLU QUALITY CONTROL HIGHER THAN STATISTICAL RANGE 0089 | Replacing the test card kit is OK.
No | Visit | Complaint | RCHN | 新试剂包无法识别 | 新试剂包无法识别 | 更换试剂包后修复 | SWO-436668 | Defect | Consumables | Solution Pack | Solution Pack | Chip | New reagent package is not recognized | New reagent package is not recognized | Repair after replacement of reagent kit
No | Visit | Complaint | RCHN | LAC定标失败 | LAC定标失败，借945-619一个 | 更换945-619后修复（By Ran） | SWO-435722 | Assisted Replacement | Electrode | Electrode | Lac |  | LAC calibration failure | LAC bid failed, borrow 945-619 one | Repair after replacement 945-619 (By Ran)
No | Visit | Complaint | RIND | Parameter failure | Parameter were yellow. Senstivity of membrane is low. | Replaced the required membranes. Run two point calibration . Senstivity is in range. | SWO-438315 | Assisted Replacement | Electrode | Membrane | Unknown | mail: which membranes? | Parameter failure | Parameter were yellow. Senstivity of membrane is low. | Replaced the required membranes. Run two point calibration . Senstivity is in range.
No | Visit | Complaint | RDEU | Driftfehler Elyte, pH und Metabolite | Ref. Elektrode defekt, Einlaßnadel verbogen | Gerät überprüft, JW durchgeführt, Nadel aus Kundenbestand und Ref. Elektrode erneuert | SWO-436974 | Assisted Replacement | Electrode | Electrode | Reference |  | Drift error Elyte, pH and metabolites | Ref. Electrode defective, inlet needle bent | Device checked, JW performed, needle made of customer stock and ref. Electrode renewed
No | Visit | Complaint | RDEU | KAL Driften Na und Ca | Täglich auftretende KAL Driften Na und Ca QC ohne Beanstandung | Erneuerung Referenz Elektrode - Funktionsüberprüfung etc. | SWO-436850 | Assisted Replacement | Electrode | Electrode | Reference | mail: Error? Or preventive?; wait, look at datalogs first! | KAL Drifting Na and Ca | Daily KAL Drifts Na and Ca QC without complaint | Renewal Reference Electrode - Function Verification etc.
No | Visit | Complaint | RDEU | KAL Driften Na und Ca | Täglich auftretende KAL Driften Na und Ca QC ohne Beanstandung | Erneuerung Referenz Elektrode - Funktionsüberprüfung etc. | SWO-436849 | Assisted Replacement | Electrode | Electrode | Reference | mail: Error? Or preventive?; wait, look at datalogs first! | KAL Drifting Na and Ca | Daily KAL Drifts Na and Ca QC without complaint | Renewal Reference Electrode - Function Verification etc.
Yes | Visit | Complaint | RAME | Lac QC laevel 3 drifting low | No investigation required. | Supplemental SWO to SWO-437235 to include part usage not included on first SWO. | SWO-438229 | Re classified RMED | Not a visit |  |  | Only information added | Lac QC laevel 3 drifting low | No investigation required. | Supplemental SWO to SWO-437235 to include part usage not included on first SWO.