No | Visit | Complaint | RAME | Lac QC laevel 3 drifting low | Customer reported after regular lactate membrane change Lactate QC level 3 began recovering below range. Confirmed issue. Reviewed data logs, performed applicable service diagnostics, no immediate hardware failures detected. | Cleaned Lac measuring chamber, remembraned Lac with 2 different lots of membranes, replaced reference electrode as suggested by CAS. Performed multiple calibration cycles, unable to improve Lac level 3 QC recovery from approx. 81 to 85 to 91 to 94 where it had been recovering before maintenance and membrane change. I have requested assistance from Radiometer management and the support team designated for UCLA. Service team will follow up on 10 $25 $2018. | SWO-437235 | Revisit required | Electrode | Membrane | Lac |  | Lac QC laevel 3 drifting low | Customer reported after regular lactate membrane change Lactate QC level 3 began recovering below range. Confirmed issue. Reviewed data logs, performed applicable service diagnostics, no immediate hardware failures detected. | Cleaned Lac measuring chamber, remembraned Lac with 2 different lots of membranes, replaced reference electrode as suggested by CAS. Performed multiple calibration cycles, unable to improve Lac level 3 QC recovery from approx. 81 to 85 to 91 to 94 where it had been recovering before maintenance and membrane change. I have requested assistance from Radiometer management and the support team designated for UCLA. Service team will follow up on 10 $25 $2018.
No | Visit | Complaint | RAME | Blood buildup in the measuring chamber keeps coming back. Calibration failing. | Checked measuring chamber. Found it full of clots and possible crack on the side | Replaced measuring chamber for troubleshooting purposes. | SWO-437234 | Defect | HW | pH/BG Module | measuring chamber | mail: send measuring chamber back to RMED. ACLAR + THOPE | Blood buildup in the measuring chamber keeps coming back. Calibration failing. | Checked measuring chamber. Found it full of clots and possible crack on the side | Replaced measuring chamber for troubleshooting purposes.
No | Visit | Complaint | RAME | Ampoules not being recognized by auto check module | Qc ampoules are not recognized even after replacement of the 14 FPC cable by FSE, Rajan Kansara. | Replaced the AutoCheck module and verified proper recognition and running of QC ampoules. Replaced Crea membranes for customer and performed the necessary calibrations. Analyzer is within acceptable operating specifications. | SWO-436980 | Defect | HW | Autocheck module | Autocheck Module |  | Ampoules not being recognized by auto check module | Qc ampoules are not recognized even after replacement of the 14 FPC cable by FSE, Rajan Kansara. | Replaced the AutoCheck module and verified proper recognition and running of QC ampoules. Replaced Crea membranes for customer and performed the necessary calibrations. Analyzer is within acceptable operating specifications.
No | Visit | Complaint | RAME | Non lamp voltage errors and low neon lamp error | Reviewed logs found error 579 Oxi hardware problem: Neon number too low 316 and Neon lamp intensity too low 258. | Replaced lamp unit. | SWO-436356 | Defect | HW | Oxi module | Lamp unit | Neon lamp | Non lamp voltage errors and low neon lamp error | Reviewed logs found error 579 Oxi hardware problem: Neon number too low 316 and Neon lamp intensity too low 258. | Replaced lamp unit.
No | Visit | Complaint | RDEU | 25.10.2018, Frau Jauß, Tel: 03643573436: Kal Fehler bei pCO2, Na und Ca | QC Fehler in allen Lev pCO2, K, Ca, Glu Sensor Ansprachefehler Ca Anwender hat bereits mehrfach SC und SP erneuert | Daten überprüft Kalibrierung und QC durchgeführt Kontakte SC und an Sensor Interface PCB gereinigt Überprüfung im Service Programm | SWO-437510 | Cleaned | HW | Basic Unit | PCB, Sensor interface |  | 25.10.2018, Mrs. Jauß, Tel: 03643573436: Kal error at pCO2, Na and Ca | QC error in all Lev pCO2, K, Ca, Glu Sensor address error Ca User has already renewed SC and SP several times | Data checked Calibration and QC performed Contacts SC and cleaned to Sensor Interface PCB Review in the service program
No | Visit | Complaint | RKK | 圧力吸引テストエラーが多発する。流体系のエラーが多発する。QCレベル3でpO2範囲外 | 機器状況確認 |  | SWO-437259 | Unknown | Unknown | Unknown | FT | 1063; mail: what has been done sol solve the issue? | Frequent pressure suction test errors occur frequently. There are a lot of errors of the flow system. QC level 3 out of pO2 range | Equipment status check
No | Visit | Complaint | RDAN | Det er et helt nyt apparat som er mindre en en uge gammel og som inden installation allerede debuterede med indgriben nødvendig hvor der skulle skylles igennem uden der har været kørt noget materiale. De har fra labs side gentagne gange måtte over at skylle igennem og også udskifte sensor kassette efter konferering med os om mulg fejl i sensor kassette. De er ret så irrerteret over situationen da det er et helt nyt apparat. | - | Skiftet... | SWO-438296 | Defect | HW | Basic Unit | Waste connector |  | It is a completely new apparatus which is less than a week old and, prior to installation, already debuted with the intervention needed where to rinse through without having been running any material. From the labs side, they repeatedly had to flush through and also replace the sensor cassette after consulting with us on MULG faults in the sensor cassette. They are quite as irrerated about the situation as it is a completely new apparatus. | - | Changed...
No | Visit | Complaint | RAME | Customer correlation between analyzer has large differences between new and old instruments. | Elevated results of oxi parameters during correlation of blood samples for new instrument. No Methb results on samples. Spectrometer temperature error displayed on screen. | Replaced hemolyzer and spectrometer. Ran tHB and VK linearity kit. Ran 5 blood samples between the old and new instruments. THB results on the new instrument correlated with hematocrit values of the respective samples. All tests are successful. | SWO-437869 | [5 empty] | Customer correlation between analyzer has large differences between new and old instruments. | Elevated results of oxi parameters during correlation of blood samples for new instrument. No Methb results on samples. Spectrometer temperature error displayed on screen. | Replaced hemolyzer and spectrometer. Ran tHB and VK linearity kit. Ran 5 blood samples between the old and new instruments. THB results on the new instrument correlated with hematocrit values of the respective samples. All tests are successful.
No | Visit | Complaint | RCHN | GLU定标不过 | GLU定标不过，借903-581一个 | 更换903-581后修复（By Ran） | SWO-436548 | Defect | HW | Basic Unit | PCB, Sensor interface | Glu | GLU calibration is not | GLU calibration, borrow 903-581 a | Repair after replacement 903-581 (By Ran)
No | Visit | Complaint | RCHN | 血氧单元定标失败 | 血氧单元定标失败,借903-197一个。 | 更换903-197后修复（By Ran） | SWO-436552 | Defect | HW | Oxi module | Spectrometer |  | Failed to calibrate the blood oxygen unit | The blood oxygen unit failed to be calibrated, borrowing 903-197. | Post-replacement 903-197 repair (By Ran)
No | Visit | Complaint | RAME | THB failing calibration. Customer can't get it to pass. | Found 1 and 2 pt calibrations failing as well as THB cal. 0251 Halogen lamp intensity too high. 0579 oxi hardware problem. Adjusted lamp, but then got Neon lamp errors. | Analyzer was powered off when I got here so I rebooted analyzer and waited for it to come up. Performed the lamp adjustment procedure. Replaced the lamp. Still had errors, so I tried a power supply and oxi module pcb. This didn't help so I removed them. Installed a hemolyzer with the lamp unit and error went away. Upon restarting the analyzer I had valve failures on the PH $BG module causing cal aborted. I replaced the valves. I Performed necessary adjustments and calibrations to clear errors including THB Service cal | SWO-437186 | Defect | HW | Oxi module | Hemolyzer |  | THB failing calibration. Customer can't get it to pass. | Found 1 and 2 pt calibrations failing as well as THB cal. 0251 Halogen lamp intensity too high. 0579 oxi hardware problem. Adjusted lamp, but then got Neon lamp errors. | Analyzer was powered off when I got here so I rebooted analyzer and waited for it to come up. Performed the lamp adjustment procedure. Replaced the lamp. Still had errors, so I tried a power supply and oxi module pcb. This didn't help so I removed them. Installed a hemolyzer with the lamp unit and error went away. Upon restarting the analyzer I had valve failures on the PH $BG module causing cal aborted. I replaced the valves. I Performed necessary adjustments and calibrations to clear errors including THB Service cal
No | Visit | Complaint | RDEU | KAL Driften Na und Ca | Tägliche KAL Driften Elektrolyte QC ohne Beanstandung | Erneuerung Referenz Elektrode - Funktionsüberprüfung etc. | SWO-437584 | Assisted Replacement | Electrode | Electrode | Reference | mail: Error? Or preventive?; wait, look at datalogs first! | KAL Drifting Na and Ca | Daily KAL Drifting Electrolytes QC without complaint | Renewal Reference Electrode - Function Verification etc.
No | Visit | Complaint | RAME | Lac QC laevel 3 drifting low | Inspection of the analyzer unit found no mechanical issues. Note Lac S7855 level 3 QC lot# 105 ranges need to be re-evaluated. Lot to date range on 867 data points indicates a 2 SD range of 84-104 mg $dL. Customer's current range is 87-107 which could be a little on the high side since most data points fall below the current customer mean of 97. Note mfg insert range is 75-107 with a mean of 91. | Replaced Lac membrane lot# 3919 with lot# 3907 | SWO-437313 | Assisted Replacement | Electrode | Membrane | Lac |  | Lac QC laevel 3 drifting low | Inspection of the analyzer unit found no mechanical issues. Note Lac S7855 level 3 QC lot# 105 ranges need to be re-evaluated. Lot to date range on 867 data points indicates a 2 SD range of 84-104 mg /dL. Customer's current range is 87-107 which could be a little on the high side since most data points fall below the current customer mean of 97. Note mfg insert range is 75-107 with a mean of 91. | Replaced Lac membrane lot# 3919 with lot# 3907
No | Visit | Complaint | RAME | Inlet lower sensor failure | Inlet lower sensor failure | Replaced lower inlet liquid sensor and calibrated | SWO-437173 | Defect | HW | pH/BG Module | LS |  | Inlet lower sensor failure | Inlet lower sensor failure | Replaced lower inlet liquid sensor and calibrated
No | Visit | Complaint | RDEU | Nach der Wartung wird die Reinigung immer wieder gefordert | Service Dump zeigt FlowValue von 1840, 1857 mit Blut | Wie nach Rücksparche mit TS empfohlen P2 ersetzt, 2 Blankmessungen und eine Blutmessung, Servicedump ausgewertet, Blut 1408 und Blank 1389, 1310. Reinigungsprogramm ausgeführt, und abschliessend 2 Level PCT QC Messung | SWO-437462 | Defect | HW | Wet section | P2 | Flow restriction | After maintenance, cleaning is required again and again | Service Dump shows FlowValue from 1840, 1857 with blood | As replaced after recess with TS recommended P2, 2 blank measurements and a blood measurement, service dump evaluated, blood 1408 and blank 1389, 1310. Cleaning program executed, and finally 2 level PCT QC measurement
No | Visit | Complaint | RPOL | Wykryto asymetrię | Diagnoza problemu. | Oczyszczenie krućca powietrznego. | SWO-438149 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | Asymmetry detected | Diagnosis of the problem. | Cleaning the air crush.
No | Visit | Complaint | RPOL | Błędy pomiarowe modułu pH/BG. Problemy przepływowe | Diagnoza problemu. | Wymiana modułu pH $BG. Kalibrację, wymiana membrany pCO2. Kalibrację zera, czujników płynów oraz pomp. Testy serwisowe oraz na krwi pacjenta. | SWO-437592 | Defect | HW | pH/BG Module | pH/BG Module |  | Measurement errors of the pH/BG module. Flow problems | Diagnosis of the problem. | Replacing the pH module $BG. Calibrating, replacing the pCO2 membrane. Calibrating zero, fluid sensors and pumps. Service tests and the patient's blood.
No | Visit | Complaint | RCHN | 历史传不出数据 | LIS传不出数据，借903-381一个 | 更换903-381后修复 by Andrew | SWO-437972 | [5 empty] | History can't pass on data | LIS can't pass the data, borrow 903-381 one | Fix by Andrew after replacing 903-381
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | CO2定标不过 | 更换测试卡后修复 | SWO-439070 | [5 empty] | Test card, reagent package failure | CO2 calibration is not the | Repair after replacing test card
No | Visit | Complaint | RCHN | TNT定标失败 | TNT定标失败，借944-268（14707）一个 | 重新定标后修复(By Ran) Recal, then OK. | SWO-436432 | Assisted Replacement | Consumables | Test Kit | Cartridge, cal | 1068 | TNT calibration failure | TNT calibration failure, borrow 944-268 (14707) one | Re-calibration after repair (By Ran) Recal, then OK.
No | Visit | Complaint | RCHN | 间歇性仪器重启 | 间歇性仪器重启，借636-533两个、580-098和903-371各一个 |  | SWO-438067 | Defect | HW | Process Module | Motor, process wheel | 5297 | Intermittent instrument restart | Intermittent instrument restart, with 636-533 two, 580-098 and 903-371 each
No | Visit | Complaint | RCHN | TNI定标错误 | TNI定标错误，借944-212（14804）一个 | 重新定标后修复(By Ran) Recal, then OK. | SWO-435823 | Assisted Replacement | Consumables | Test Kit | Cartridge, cal | 5297 | TNI calibration error | TNI calibration error, borrow 944-212 (14804) one | Re-calibration after repair (By Ran) Recal, then OK.
No | Visit | Complaint | RCHN | 测试卡断裂 | 测试卡断裂 | 更换测试卡后修复 | SWO-447187 | Defect | Consumables | Test Kit | Cartridge, Broken |  | Test card break | Test card break | Repair after replacing test card
No | Visit | Complaint | RKK | インレットガスケットより液漏れ（インレットガスケット交換後も改善せず）、インレットガスケット下ﾆｯﾌﾟﾙ内の血液固着及びチューブ外れ ニップル内の洗浄、排液ﾗｲﾝ洗浄、排液ﾁｭｰﾌﾞ接続 | ケース詳細の通り | ニップル内の洗浄、排液ﾗｲﾝ洗浄、排液ﾁｭｰﾌﾞ接続 | SWO-438890 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | Leakage from inlet gasket (not improved after inlet gasket replacement), blood fixation and tube outof the nipple under the inlet gasket Cleaning in nipples, drainage line cleaning, drainage tube connection | As the case details | Cleaning in nipples, drainage line cleaning, drainage tube connection
No | Visit | Complaint | RPOL | Wyciek. | Diagnoza problemu. | Oczyszczenie analizatora. Kompletna wymiana owężowania oraz uszczelki wlotu, odkażenie toru pomiarowego. Kalibracja czujników płynów oraz pomp. Kalibracja tHb. Testy serwisowe. Tymczasowo niestabilne QC ze względu na odkażanie toru pomiarowego. | SWO-438151 | Defect | HW | Inlet module | Inlet gasket |  | Leak. | Diagnosis of the problem. | Cleaning the analyzer. Complete replacement of the incruising and inlet seals, the measurement track disinfection. Calibration of fluid sensors and pumps. THb Calibration. Service tests. Temporarily unstable QC due to disinfection of the measuring track.
No | Visit | Complaint | RDEU | 26.10.2018, Herr Hassels, Tel: 024215996996: AQC Fehler bei pCO2 Level A und C 26.10.208, Herr Hassels, Tel: 024215996996: weiterhin AQC Fehler bei pCO 29.10.2018; Hr. Hassels, Tel.: 024215996996; wieder pCO2 QC Fehler | Dateien gesichtet | Solutionpack ersetzt, eine SC und ein SP zusätzlich geliefert, QC Kontrolle Level C knapp im Bereich... | SWO-438570 | Assisted Replacement | SP | SP | pCO2 | QC A og C high | 26.10.2018, Mr Hassels, Tel: 024215996996: AQC Error at pCO2 Level A and C 26.10.208, Mr Hassels, Tel: 024215996996: continue AQC error with pCO 29.10.2018; Mr. Hassels, tel.: 024215996996; again pCO2 QC error | Files spotted | Solutionpack replaced, a SC and an SP additionally delivered, QC control level C just in the range of...
No | Visit | Complaint | RDEU | 26.09.2018 09:25Uhr Frau Gerstmayr: Heute habe ich erfahren, dass bei 3 Ringversuch das p02 nicht bestanden hat und wir somit keine Zertifikat mehr haben. Deshalb hat mich heute auch Chefarzt Dr. Hehl kontaktiert. Seit dem Absturz im Mai diesen Jahres ist das Gerät sehr störanfällig und außerdem geht ohne Zertifikat lt. Aussage von Herrn Wörner (Labor) eigentlich gar nichts. Wie ist das zu erklären bzw. muss ein Gerätetausch stattfinden? Ich bitte Sie mich baldmöglichst unter 08272/998171 telefonisch zu kontaktieren. 04.10.2018, 09:27 Uhr, Fr. Gerstmayer, Tel.: 08272998171: Ringversuch bei pO2 zum dritten Mal nicht bestanden. Labor und CA wollen ABL stilllegen. | Mit Laborleirung und Frau Krupkaz Ringversuch durchgeführt. | Vor Ringversuch 2 Punkt kal und QC für jedes vorhandene Level durchgeführt. Danach hat Frau Krupka den Ringversuch durchgeführt. U d die Ergebnisse Der Laborleitung übergeben. Mit Frau Krupka noch Probleme von Radiance geklärt damit sie wieder Zugriff auf Radiance hat. Browser hat Seite nicht anerkannt. Seite dem Browser hinzugefügt. Danach war Zugriff wieder ordnungsgemäß möglich. | SWO-439457 | No fault found | No fault found | No fault found | pO2 |  | 26.09.2018 09:25am Mrs. Gerstmayr: Today I learned that the p02 did not pass in 3 ring trials and we therefore no longer have a certificate. That is why dr. Hehl contacted me today. Since the crash in May of this year, the device is very prone to interference and, moreover, without a certificate, according to the statement of Mr. Wörner (Laboratory), nothing actually works. How can this be explained or does a device exchange have to take place? I would ask you to contact me by phone as soon as possible on 08272/998171. 04.10.2018, 09:27, Fr. Gerstmayer, Tel.: 08272998171: Ring test at pO2 failed for the third time. Labor and CA want to shut down ABL. | With laboratory and Ms. Krupkaz ring test performed. | Before ring test 2 point kal and QC performed for each existing level. Afterwards, Ms Krupka carried out the ring test. U d hand over the results to the laboratory management. With Ms. Krupka still problems of Radiance solved so that she has access to Radiance again. Browser has not recognized page. page added to the browser. After that, access was possible again properly.
No | Visit | Complaint | RLTD | Contact: Kingsley Tel: 0330 400 4333 EXT 2621 | Performed manual flushing several times. Replaced inlet gasket. straightened inlet probe. Tested Analyser with a test cassette and ran fluid checks. issue is still present. Performed service maintenance procedure and fitted service kit. The solution pack is the cause of the issue. | Advised the customer to replace the SP as soon one becomes available as this is the cause of the issue. | SWO-439796 | Assisted Replacement | SP | SP | FT | 1063 | Contact: Kingsley Tel: 0330 400 4333 EXT 2621 | Performed manual flushing several times. Replaced inlet gasket. straightened inlet probe. Tested Analyser with a test cassette and ran fluid checks. issue is still present. Performed service maintenance procedure and fitted service kit. The solution pack is the cause of the issue. | Advised the customer to replace the SP as soon one becomes available as this is the cause of the issue.
No | Visit | Complaint | RDEU | es werden nicht alle Barcodes erkannt | Patientenetiketten werden nicht gelesen | Fehlersuche; roma.ini angepaßt; Neustart; QC i.O. | SWO-438636 | Setup | SW | RMED SW | Roma.ini | barcode | not all barcodes are recognized | Patient labels are not read | Troubleshooting; roma.ini adapted; Restart; QC in the first
No | Visit | Complaint | RDEU | es werden nicht alle Barcodes erkannt | Patientenetiketten werden nicht gelesen | Fehlersuche; roma.ini angepaßt; Neustart; QC i.O. | SWO-438637 | Setup | SW | RMED SW | Roma.ini | barcode | not all barcodes are recognized | Patient labels are not read | Troubleshooting; roma.ini adapted; Restart; QC in the first
No | Visit | Complaint | RKK | アプリケーションエラー。 Romaシャットダウン。 | アプリケーションエラー発生。 | 代替機設置、引上げ。 | SWO-438186 | [5 empty] | Application error. Roma shutdown. | An application error occurred. | Installation and pulling of alternative machines.
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 多参数定标不过 | 更换测试卡后修复 | SWO-440430 | [5 empty] | Test card, reagent package failure | Multi-parameter calibration but | Repair after replacing test card
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | pco2定标失败，0378 | 更换2张测试卡后修复 | SWO-439843 | Assisted Replacement | SC | SC | pCO2 | 378 | Test card, reagent package failure | pco2 bid failure, 0378 | Repair after replacing 2 test cards
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | CO2定标不过 1227 0378 | 更换测试卡后修复 | SWO-442300 | Assisted Replacement | SC | SC | pCO2 | Sens | Test card, reagent package failure | CO2 calibration is no more than 1227 0378 | Repair after replacing test card
No | Visit | Complaint | RCHN | 条码扫描器故障 | 条码扫描器故障，借910-304一个 |  | SWO-438308 | Defect | HW | Storage Module | Barcode reader, storage wheel |  | Barcode scanner failure | Barcode scanner failure, borrow 910-304 one
No | Visit | Complaint | RCHN | 报错5784 | 报错5784，借903-117一个 |  | SWO-438203 | Defect | HW | Cover and Chassis | PCB, cabinet | 5784 | 5784 | 5784, loan 903-117 one
No | Visit | Complaint | RCHN | 无法识别处理包 | 无法识别处理包，借903-036一个 | 更换903-036一个正常。by Andrew | SWO-438064 | Defect | HW | Cover and Chassis | PCB, solution pack |  | The processing package is not recognized | The processing package is not recognized, and 903-036 is borrowed | Replace 903-036 one normal. by Andrew
No | Visit | Complaint | RCHN | Myo定标失效 | Myo定标失效，借944-214（14825）一个 |  | SWO-438236 | Assisted Replacement | Consumables | Test Kit | Cartridge, cal |  | Myo calibration failure | Myo calibration is invalid, borrow edgy 944-214 (14825) one
No | Visit | Complaint | RCHN | 测试卡断裂，5784 | 测试卡断裂，5784 | 更换测试卡后修复 | SWO-441659 | Defect | Consumables | Test Kit | Cartridge, Broken | 5784 | Test card break, 5784 | Test card break, 5784 | Repair after replacing test card
No | Visit | Complaint | RCHN | BNP测试卡断裂 | BNP测试卡断裂 | 更换测试卡后修复 | SWO-438294 | Defect | Consumables | Test Kit | Cartridge, Broken |  | BNP test card break | BNP test card break | Repair after replacing test card
No | Visit | Complaint | RAME | Analyzer display's OXI liquid sensor failure | Ran the liquid sensor check in service, and the OXI sensor failed. | Replaced the OXI liquid sensor, and this resolved the problem. | SWO-438803 | Defect | HW | Oxi module | LS |  | Analyzer display's OXI liquid sensor failure | Ran the liquid sensor check in service, and the OXI sensor failed. | Replaced the OXYGEN liquid sensor, and this resolved the problem.
No | Visit | Complaint | RAME | Power issue | Found instrument had been restarted several times in logs with no indication of the cause to shutdown. | Replaced power supply. | SWO-438830 | Defect | HW | DMS_PC | Power supply |  | Power issue | Found instrument had been restarted several times in logs with no indication of the cause to shutdown. | Replaced power supply.
No | Visit | Complaint | RDEU | Touch Screen ohne Funktion | Flüssigkeit in Touchscreen. Kundenverschulden. Hr. Prill (Medizintechnik) über Kosten informiert. | Touchscreen erneuert und kalibriert. Systemüberprüfung und Systemreinigung- ok. LQC Level1: 1x TnT, 1x CKMB, 1x D-dimer, 1x CRP und 1x ßhCG durchgeführt. LQC Level2: 1x TnT, 1x CKMB, 1x D-dimer, 1x CRP und 1x ßhCG durchgeführt. | SWO-438633 | Defect | HW | Cover and Chassis | display | Potential User Error | Touch screen without function | Liquid in touchscreen. Customer debt. Mr. Prill (medical technology) is informed about costs. | Touchscreen renewed and calibrated. System check and system cleaning- ok. LQC Level1: 1x TnT, 1x CKMB, 1x D-dimer, 1x CRP and 1x ßhCG performed. LQC Level2: 1x TnT, 1x CKMB, 1x D-dimer, 1x CRP and 1x ßhCG performed.
No | Visit | Complaint | RPOL | Błąd kontroli jakości pO2. QC3 za wysokie. | Sprawdzenie dzienników danych. | Wymiana pakietu roztworów. Pakiet zgłoszony do reklamacji. | SWO-438427 | Assisted Replacement | SP | SP | pO2 | QC 3 high | Quality control error pO2. QC3 too high. | Check the data logs. | Replacement of the solution package. The package submitted for the complaint.
No | Visit | Complaint | RAME | On-going Consumable replacement. FSE needs to schedule a in person visit. Sossity or Tabitha Nusbaum 319-369-7311- | Numerous Consumable Issues. | Performed electronic checkout and alignment following the annual preventive maintenance. All systems test okay except for the USB on the top of the analyzer. See SWO-438878 for PM and SWO-438883 for CPU replacement.. | SWO-438880 | No fault found | No fault found | No fault found | No fault found |  | On-going Consumable replacement. FSE needs to schedule a in person visit. Sossity or Tabitha Nusbaum 319-369-7311- | Numerous Consumable Issues. | Performed electronic checkout and alignment following the annual preventive maintenance. All systems test okay except for the USB on the top of the analyzer. See SWO-438878 for PM and SWO-438883 for CPU replacement..
No | Visit | Complaint | RDEU | 30.10.2018 Sr. Nadine 040790202992 hat gemeldet, dass sie immer wieder das LP tauschen muss. | Lösungen können nicht gepumpt werden | Fehleranalyse, siehe Wartungsmassnahmen im SWO-439367 | SWO-439366 | Defect | HW | Basic Unit | PM Kit |  | 30.10.2018 Sr. Nadine 040790202992 has announced that she has to change the LP again and again. | Solutions cannot be pumped | Fault analysis, see maintenance measures in SWO-439367
No | Visit | Complaint | RDEU | Grüne Messbereit-LED arbeitet nicht. | Grüne Messbereit-LED arbeitet nicht. | Inlet Position Detektor-PCB erneuert. | SWO-439107 | Defect | HW | Inlet module | PCB, Inlet |  | Green ready-to-measure LED does not work. | Green ready-to-measure LED does not work. | Inlet position detector PCB renewed.
No | Visit | Complaint | RDEU | 29.10.2018, Frau Thome, Tel: 02037330(Kreißsaal): Intervention erforderlich- Lösungen können nicht gepumpt werden 29.10.2018; 16:20 Fr. Thome; Tel.: 02037330; ; weiterhin Intervention erforderlich, Lösungen können nicht gepumpt werden. 29.10.2018; 18:10 Fr. Thome; Tel.: 02037330; ; weiterhin Intervention erforderlich, Lösungen können nicht gepumpt werden. | LP defekt. | Kunde hat Probleme nach LP erneuern. LP von der Station Lufu geholt, LP erneuert, Gerät kalibriert und QC durchgeführt. | SWO-439859 | Assisted Replacement | SP | SP | FT | 1343 | 29.10.2018, Mrs Thome, Tel: 02037330(Kreißsaal): Intervention required- Solutions cannot be pumped 29.10.2018; 16:20 Fr. Thome; Tel.: 02037330; ; intervention is still required, solutions cannot be pumped. 29.10.2018; 18:10 Fr. Thome; Tel.: 02037330; ; intervention is still required, solutions cannot be pumped. | LP defective. | Customer has problems renewing LP. LP fetched from the lufu station, LP renewed, device calibrated and QC performed.
No | Visit | Complaint | RLTD | Customer having numerous software issues, including failing to print, laggy software and freezes. | Customer having numerous software issues, including failing to print, laggy software and freezes. | Backed up customer data and customer settings. Reformatted flash card and fitted pc unit reloaded windows 7 software, applied regional settings and calibrated touch screen. Installed Radiometer operating software and reinstalled customer settings and data logs. Waited for analyser to finish 50 minutes startup. Customer reported that the issues still remained. | SWO-440079 | Revisit required | SW | RMED SW | Freeze |  | Customer having numerous software issues, including failing to print, laggy software and freezes. | Customer having numerous software issues, including failing to print, laggy software and freezes. | Backed up customer data and customer settings. Reformatted flash card and fitted pc unit reloaded windows 7 software, applied regional settings and calibrated touch screen. Installed Radiometer operating software and reinstalled customer settings and data logs. Waited for analyser to finish 50 minutes startup. Customer reported that the issues still remained.
Yes | Visit | Complaint | RIND | Lactate calibration error. | Check the system. | sc80 fail to calibrate the sc80. Conditioning the sc80 still the sc80 fail to calibrate. | SWO-441508 | [5 empty] | Lactate calibration error. | Check the system. | sc80 fail to calibrate the sc80. Conditioning the sc80 still the sc80 fail to calibrate.
No | Visit | Complaint | RCHN | 系统崩溃 | 系统崩溃,借 910-342一个。 | 更换U盘后重新安装系统修复 by Andrew | SWO-438372 | Defect | HW | DMS_PC | CPU |  | System crash | System crashes, borrowing 910-342 one. | Reinstall system repair by Andrew after replacing USB stick
No | Visit | Complaint | RDAN | Ved skift af væskepakke, opdages det, at der er aflejringer under ABL | Undersøge VP | Isat ny VP | SWO-439153 | [5 empty] | When changing the liquid package, it is detected that there are deposits under the ABL | Undersøge VP | Isat ny VP
No | Visit | Complaint | RAME | tHb QC results trending low | tHb QC results trending lower from mean. Multiple tHb calibrations have not changed trend. | Replaced hemolyzer per radiometer specifications. Calibrated and ran QC all results are in range. Ran all levels of QC multiple times overnight and performance is stable. Instrument operating normally. | SWO-439651 | Defect | HW | Oxi module | Hemolyzer |  | tHb QC results trending low | tHb QC results trending lower from mean. Multiple tHb calibrations have not changed trend. | Replaced hemolyzer per radiometer specifications. Calibrated and ran QC all results are in range. Ran all levels of QC multiple times overnight and performance is stable. Instrument operating normally.
No | Visit | Complaint | RDEU | Na und K. drift | Na und K. driften sehr häufig . Kalibration bringt nur sporadisch eine Verbesserung. Elektrode wurde bereits membraniert. | Ref. Elektrode ausgetauscht. Kalibration mehrfach wiederholt. Mehrere Dichtungsringe ausgetauscht. | SWO-439822 | Assisted Replacement | Electrode | Electrode | Reference | CAPA-733 | Na and K. drift | Na and K. drift very often. Calibration only sporadically improves. Electrode has already been membraneed. | Ref. electrode. calibration repeated several times. Replaced several sealing rings.
No | Visit | Complaint | RAME | Neon lamp intensity errors | Neon lamp errors 316 and 258. Complete Oxi test fail Hemolyzer for Impedance shape - Phase regulation error. Lamp failed for neon intensity 23% and neon position -1. | Replaced complete Oxi module. | SWO-439635 | Defect | HW | Oxi module | Oxi module | Neon lamp | Neon lamp intensity errors | Neon lamp errors 316 and 258. Complete Oxi test fail Hemolyzer for Impedance shape - Phase regulation error. Lamp failed for neon intensity 23% and neon position -1. | Replaced Complete Oxygen module.
No | Visit | Complaint | RDEU | pO2 und Oximetrieprobleme | Einlaßnadel verbogen, Einlaßdichtung defekt | Gerät überprüft, Einlaß gesäubert, Einlaßnadel und Dichtung ersetzt, Kal und QC | SWO-440004 | Defect | HW | Inlet module | Probe bend |  | pO2 and oximetry problems | Inlet needle bent, inlet seal defective | Device checked, inlet cleaned, inlet needle and seal replaced, Kal and QC
No | Visit | Complaint | RIND | Pco2 qc values low | Pco2 qc values low | checked the analyser and pco2 electrode puss formation was there so cleaned the pco2 electorde and calibated the analyser to raedy mode | SWO-441404 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway | pCO2 | Pco2 qc values low | Pco2 qc values low | checked the analyser and pco2 electrode puss formation was there so cleaned the pco2 electorde and calibated the analyser to raedy mode
No | Visit | Complaint | RFRA | Portes Inlet wheel: ouverture difficile. | Ouverture porte HS | Remplacement système crémaillère et portes. Essais ok en mode service. Test échantillon routine ok. Envoi du RI par E_mail à Mr Fevrier pour signature. | SWO-439818 | Defect | HW | Cover and Chassis | Door, Inlet |  | Inlet wheel doors: difficult opening. | Opening door HS | Replacing rack system and doors. Tests ok in service mode. Routine sample test ok. Sending the RI by e-mail to Mr. Fevrier for signature.
No | Visit | Complaint | RDEU | 30.10.2018, 10:00 Uhr, Sr. Christiane, Tel. 03634521530: Probenfenster schließt nicht; Bildschirm läßt sich nicht bedienen; Probenmessungen werden abgebrochen; AQT verlangt Neustart | Tür Einlaß schließt nicht korrekt Probenmessung abgebrochen | Fehlersuche, Justage Inlet Door, rechte Einlaßtür klemmt bzw. z.t. schwergängig, sollte erneuert werden Testmessung und QC Messung durchgeführt Systemreinigung - Fehlermeldung 3553 Incubatorwheelmove33, neue Leerkassette eingelegt, anschl.ok | SWO-439489 | Adjustment | HW | Cover and Chassis | Door, Inlet |  | 30.10.2018, 10:00 a.m., Sr. Christiane, tel. 03634521530: sample window does not close; Screen cannot be operated; Sample measurements are aborted; AQT requires restart | Door entrance does not close correctly Sample measurement aborted | Troubleshooting Justage Inlet Door, right entrance door clamps or partly difficult, should be renewed Test measurement and QC measurement performed System cleaning - Error message 3553 Incubatorwheelmove33, new empty cassette inserted, followed by ok
No | Visit | Complaint | RPOL | Wyciek. | Diagnoza problemu. | Kompletna wymiana owężowania ściekowego. Wymiana płaskiej membrany w manifolddzie fluidic module valves. Testy serwisowe. | SWO-439506 | Defect | HW | Fluid manifold | Valve, membrane |  | Leak. | Diagnosis of the problem. | Complete replacement of sewage. Replacing the flat membrane in the Manifolddzie fluidic module valves. Service tests.
Yes | Visit | Complaint | RPOL | Błędy pomiarowe modułu pH/BG. Problemy przepływowe | Niejednorodna próbka w module pHbBG | Użytkownik sam naprawił błąd przed moim przyjazdem nie informując mnie o tym fakcie. | SWO-439507 | Re classified RMED | problem solved by customer |  |  |  | Measurement errors of the pH/BG module. Flow problems | Nonuniform Sample in pHbBG module | The user has fixed the error before my arrival without telling me about this fact.
No | Visit | Complaint | RDEU | 020917253304 Hr. Altenberg Med Technik tHb Kal gescheitert | tHb-Kal. ist gescheitert, teilweise keine Hb Messung, Schlauch am Oximodul ist def. | Schlauch ersetzt, tHb-Kal. durchgeführt, Gerät kalibriert und QC durchgeführt, Meth Hb wird noch um 0,2g $% zu hoch im QC gemessen, Kunde meldet sich in 2 Tagen, | SWO-440872 | Defect | HW | Oxi module | Pinch valve tube |  | 020917253304 Mr. Altenberg Med Technique tHb Kal failed | tHb-Kal. has failed, partly no Hb measurement, hose on oximodul is def. | hose, tHb-Kal. performed, device calibrated and QC performed, meth Hb is still measured by 0.2g% too high in the QC, customer reports in 2 days,
No | Visit | Complaint | RCHN | 泵头坏了 | 泵头坏了，借902-971一个 | 更换902-971一个正常。 by Andrew | SWO-439150 | [5 empty] | The pump head is broken. | The pump head is broken, borrow 902-971 one. | Replace the 902-971 one normal. by Andrew
No | Visit | Complaint | RCHN | 报错5784 | 报错5784，加借一个903-198 | 更换903-198和903-117后修复 by Andrew | SWO-438986 | Defect | HW | Cover and Chassis | PCB, Main Board | Potential FSE User Error, 5784 | 5784 | 5784, plus a 903-198 | Repair by Andrew after replacing 903-198 and 903-117
No | Visit | Complaint | RCHN | LOCK-24 | LOCK-24，借903-215和888-654各一个 | 更换903-215和888-654后修复 by Anderw | SWO-439056 | Defect | HW | Incubator Module | Incubator wheel | 3998 | LOCK-24 | LOCK-24, one each for 903-215 and 888-654 | Repair by Anderw after replacing 903-215 and 888-654
No | Visit | Complaint | RCHN | 定标被中断 | 定标被中断，借944-214（14934）一个 | 重新定标后修复 by Andrew | SWO-439003 | Assisted Replacement | Consumables | Test Kit | Cartridge, cal |  | The calibration was interrupted | The bid was interrupted, borrowing 944-214 (14934) one | Fix by Andrew after re-calibration
No | Visit | Complaint | RAME | Sample in El/Met upper LS detected prematurely | Customer reports sample in El $Met upper LS detected prematurely during calibration and QC. Inspected membrane valves in the El $Met module. Air valve would not activate. Upon removal, it was also found to have a torn membrane. | Replaced Flat Membrane Valve (p $n 902-701). Fault with no leading air segment in inlet sensor arose. Ran protein removal and cleaning due to machine not being active for an extended time. Ran Service Inlet calibration, LS adjustment, and pump Cali. Ran three successful 2-point calibrations with no issue and all three levels of QC. Normal operations restored. | SWO-440199 | Defect | HW | El/Met Module | Valve, membrane | 0654 | Sample in El/Met upper LS detected prematurely | Customer reports sample in El /Met upper LS detected prematurely during calibration and QC. Inspected membrane valves in the El /Met module. Air valve would not activate. Upon removal, it was also found to have a torn membrane. | Replaced Flat Membrane Valve (p $n 902-701). Fault with no leading air segment in inlet sensor arose. Ran protein removal and cleaning due to machine not being active for an extended time. Ran Service Inlet calibration, LS adjustment, and pump Cali. Ran three successful 2-point calibrations with no issue and all three levels of QC. Normal operations restored.
No | Visit | Complaint | RDAN | Dorte skriver, at ABLen drypper periodisk. Hun har skiftet inletpakning. | Afprøvet | Inletpakning, probe og arm udskiftet, slutafprøvet. | SWO-439947 | Defect | HW | Inlet module | Probe holder |  | Dorte writes that the ABLen drips periodically. She has changed the inlets. | Tested | Inlets, probe and arm replaced, final tested.
No | Visit | Complaint | RIND | Printer have problem, its not printing properly. | NA | checked the printer head and cleaned the printer head but its not working.its need to replace with new. | SWO-439882 | Defect | HW | Printer Module | Printer Head |  | Printer have problem, its not printing properly. | In | checked the printer head and cleaned the printer head but its not working.its need to replace with new.
No | Visit | Complaint | RDAN | Dorte klager over, at den ny ABL drypper per. under rins. | Undersøgt | Intet pakning, probe og arm udskiftet, slutafprøvet. | SWO-439934 | Defect | HW | Inlet module | Probe holder |  | Dorte complains that the new ABL drips per. Under Rins. | Examined | No packing, probe and arm replaced, final tested.
No | Visit | Complaint | RIND | Inlet door not workining | Inlet door not workining properly | Check manually by pressinig introduce sample. Addjuset the screw from door pannel , then also problem not solved . Need to change the door pannel | SWO-448444 | Revisit required | HW | Cover and Chassis | Door, Inlet |  | Inlet door not workining | Inlet door not workining properly | Check manually by pressinig introduce sample. Addjuset the screw from door pannel , then also problem not solved . Need to change the door pannel
No | Visit | Complaint | RDEU | 30.10.2018; Hr. Brückner; Tel.: 0375514942; Oxy Spectrum mismatch Error bei Blutmessungen | Fehlersuche ==> bei Patientenmessungen Fehler 0581: Oxi spectrum mismatch | Hemolyser unit erneuert THb-Kalibrierung Testmessungen mit Patientenproben ==> o.k. | SWO-440572 | Defect | HW | Oxi module | Hemolyzer | 581 | 30.10.2018; Mr. Brückner; Tel.: 0375514942; Oxy Spectrum mismatch error in blood measurements | Troubleshooting ==> for patient measurements Error 0581: Oxi spectrum mismatch | Hemolyser unit renewed THb calibration Test measurements with patient samples ==> o.k.
No | Visit | Complaint | RKK | SC交換時毎回Lac感度低下となる。 | インターフェイスPCB故障 | インターフェイスPCB交換 | SWO-449335 | Defect | HW | Basic Unit | PCB, Sensor interface | Lac | Lac sensitivity decreases every time sc is replaced. | Interface PCB Failure | Interface PCB Replacement
No | Visit | Complaint | RAME | Customer has requested help with issues pertaining to inlet | Customer reported that inlet did not smoothly return to its down position. Customer also reported the presence of QC solution where the inlet gasket meets the solution pack. Inspect the inlet probe. Found it was not bent. Inspected inlet gasket. No issues found. Inspected inlet spring. Found it to be twisted and realigned it for proper use. | Corrected spring position. Normal operations restored. Calibration and QC completed. | SWO-440615 | Defect | HW | Inlet module | Inlet module |  | Customer has requested help with issues pertaining to inlet | Customer reported that inlet did not smoothly return to its down position. Customer also reported the presence of QC solution where the inlet gasket meets the solution pack. Inspect the inlet probe. Found it was not bent. Inspected inlet gasket. No issues found. Inspected inlet spring. Found it to be twisted and realigned it for proper use. | Corrected spring position. Normal operations restored. Calibration and QC completed.
No | Visit | Complaint | RAME | After changing sensor cassette analyzer had calibration issues, mainly due to time for calibration, but then oximetry values went out of sorts | Oximetry values way off. 581 spectrum mismatch errors arose suddenly, indicating a possible clot. | Troubleshot analyzer for clot. Performed clot detect and remove, automatic flush, and many manual flushes. Ran successful calibration and tHb calibration. All levels of internal and external QC passed. Normal operations restored. | SWO-440657 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | After changing sensor cassette analyzer had calibration issues, mainly due to time for calibration, but then oximetry values went out of sorts | Oximetry values way off. 581 spectrum mismatch errors arose suddenly, indicating a possible clot. | Troubleshot analyzer for clot. Performed clot detect and remove, automatic flush, and many manual flushes. Ran successful calibration and tHb calibration. All levels of internal and external QC passed. Normal operations restored.
No | Visit | Complaint | RDEU | Anruf von Frau Gärtner; Gerät steht in Intervention erforderlich, wurde schon alles getauscht, ohne Erfolg. Bitte um Reparatur. | Gerät überprüft. Gerinnsel im System. | Gerät rückwärts gespült, Service Tests durchgeführt. Alle Baugruppen separat ausgebaut und gespült, Sensorcassette Gerinsel nicht zu entfernen, SC ausgetauscht, Startphase 2 Std. I.O, Probemessung ergab wieder Gerinsel im System. Gerät nach Rücksprache mit Labor mitgenommen. | SWO-441536 | Revisit required | Analyzer replaced | Analyzer replaced | Analyzer replaced | clot | Call from Mrs Gärtner; Device is required in intervention, everything has already been exchanged, to no avail. Please repair. | device checked. Clots in the system. | Device flushed backwards, service tests performed. All assemblies separately removed and rinsed, Sensorcassette Gerinsel not to be removed, SC replaced, start phase 2 hrs I.O, sample measurement again yielded Gerinsel in the system. device after consultation with the laboratory.
No | Visit | Complaint | RLTD | Please call Claire on 01642 282676. Problem also with another ABL90 in another department which she said would be called through by cardiac theatre team so currently no functioning machine in dept. Thank you. | Keeps asking for inlet gasket to be replaced. | On arrival the analyser was on the replace inlet gasket screen, this was exited and service menu entered. It was showing leak detected and had failed vacuum pressure test. The inlet gasket and probe were removed and inspected with no damage seen. On exiting the service menu the analyser went into 55 minute start up but during this, looking at the inlet from another angle it could be seen that there was a curve in the inlet probe. This was replaced with the customers stock once the analyser had completed its start up. All fluid transport tests now passing. Advised customer on troubleshooting was inlet gasket replacement. | SWO-441225 | Defect | HW | Auto Inlet module | Probe bend |  | Please call Claire on 01642 282676. Problem also with another ABL90 in another department which she said would be called through by cardiac theatre team so currently no functioning machine in dept. Thank you. | Keeps asking for inlet gasket to be replaced. | On arrival the analyser was on the replace inlet gasket screen, this was exited and service menu entered. It was showing leak detected and had failed vacuum pressure test. The inlet gasket and probe were removed and inspected with no damage seen. On exiting the service menu the analyser went into 55 minute start up but during this, looking at the inlet from another angle it could be seen that there was a curve in the inlet probe. This was replaced with the customers stock once the analyser had completed its start up. All fluid transport tests now passing. Advised customer on troubleshooting was inlet gasket replacement.
No | Visit | Complaint | RCHN | 接口板腐蚀 | 接口板腐蚀，借902-972一个 | 更换902-972后修复by Andrew | SWO-438985 | [5 empty] | Interface plate corrosion | Interface plate corrosion, borrow 902-972 one | Repair by Andrew after replacement 902-972
No | Visit | Complaint | RCHN | PO2质控打问号 | PO2质控打问号，借 903-581一个。 | 更换903-581后修复 by Andrew | SWO-440467 | Defect | HW | Basic Unit | PCB, Sensor interface | pO2 | PO2 QC QUESTION MARK | PO2 QC QUESTION MARK, BORROW 903-581 ONE. | Repair by Andrew after replacement 903-581
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 测试卡检测到漏液，试剂包漏电流，报错1342、1678 | 更换一张测试卡和一个试剂包后修复 | SWO-440446 | Assisted Replacement | SP | SP | FT | 1342 | Test card, reagent package failure | Test card detected leakage, reagent package leakage current, reported error 1342, 1678 | Repair after replacing a test card and a reagent kit
No | Visit | Complaint | RCHN | 孵育盘故障，报错2912 | 孵育盘故障，报错2912 借636-574 903-138各一个 | 更换636-574 903-138备件后修复 by Andrew | SWO-439786 | Defect | HW | Incubator Module | PCB, incubator | 2912 | Incubation plate failure, reported error 2912 | Incubation plate failure, reported error 2912 borrowing 636-574 903-138 each | Repair by Andrew after replacing 636-574 903-138 spare parts
No | Visit | Complaint | RAME | Intermittent Electrolyte Drift Failures on 2-point calibrations | Customer reports intermittent, yet common, electrolyte failures during calibrations. | Could not replicate error, nor was any apparent fault found during troubleshooting. Replaced pH seal ring and cleaned contacts on the reference electrode. Will monitor for future issues. | SWO-440717 | Cleaned | HW | pH/BG Module | Electrode Connector | ref | Intermittent Electrolyte Drift Failures on 2-point calibrations | Customer reports intermittent, yet common, electrolyte failures during calibrations. | Could not replicate error, nor was any apparent fault found during troubleshooting. Replaced pH seal ring and cleaned contacts on the reference electrode. Will monitor for future issues.
No | Visit | Complaint | RAME | Electrolytes and Metabolytes Intermittently fail the 2-point calibration for drift | Troubleshot to a faulty valve on the pH $BG module | Replaced both valves on the pH $BG module, and this resolved the problem. | SWO-440646 | Defect | HW | pH/BG Module | Valve, membrane |  | Electrolytes and Metabolytes Intermittently fail the 2-point calibration for drift | Troubleshot to a faulty valve on the pH /BG module | Replaced both valves on the pH $BG module, and this resolved the problem.
No | Visit | Complaint | RAME | Temp error. PH/BG module | Error occurred after unit was moved. No errors on arrival. | Checked data logs. Reseated electrical connectors. | SWO-440677 | No fault found | No fault found | No fault found | No fault found |  | Temp error. PH/BG module | Error occurred after unit was moved. No errors on arrival. | Checked data logs. Reseated electrical connectors.
No | Visit | Complaint | RAME | Customer has requested help with issues pertaining to inlet | Concerns with inlet operability. | Pro-actively replaced inlet unit. Tested functionality. Op test sat. | SWO-441167 | Defect | HW | Inlet module | Inlet module |  | Customer has requested help with issues pertaining to inlet | Concerns with inlet operability. | Pro-actively replaced inlet unit. Tested functionality. Op test sat.
No | Visit | Complaint | RDEU | 29.10.2018 14:00 Fr. Richter ????: Interventionsmeldung Gerinsel in der Oximetrie | Gerät verlangt ständig Spülung und ZVBM Tausch bei eintreffen steht in Intervention erforderlich -> verlangt Gerätespülung | Austausch der kpl. internen Verschlauchung, Quetschventil zerlegt & gereinigt, Startphase abgewartet, Kal & QC i.O. Kal & QC Daten gemäß RDEU Anweisung gesichert & dokumentiert | SWO-441134 | Defect | HW | Basic Unit | PM Kit |  | 29.10.2018 14:00 Fr. Richter ????: Intervention report Gerinsel in oximetry | Device constantly requires flushing and ZVBM exchange when arrival is required in intervention -> requires equipment flushing | Replacement of the cpl. internal hose, crushing valve disassembled & cleaned, start phase awaited, Kal & QC i.O. Kal & QC data secured & documented according to RDEU instructions
No | Visit | Complaint | RDAN | OXI fejl | OXI test | Komplet oxi test Kalibrering tHb kalibrering | SWO-440982 | Adjustment | HW | Oxi module | tHb Cal | 581 | OXY Headl | OXYGEN Test | Complementary Oxi Test Calibration tHb Calibration
No | Visit | Complaint | RLTD | Angela Kumanoo Tel: 0203 447 0117 | The Analyser is on the main screen, but buttons are missing and nothing will allow anyone to proceed. Intervention Required is displayed. Restarted the Analyser and the same issue occurs. Powered down the Analyser, removed the Inlet and SP. Powered on and now the Analyser is in intervention required and able to carry out replacements. The analyser screen appears with no buttons after the automatic log off. SC is out of date and is the cause of the issue (LOT 946-010 R0608N196 20 $10 $2018) SP was not activated correctly. | Removed SC and informed the customer of the present issues. SC needs to replaced by the customer. No further action required. | SWO-441079 | user Error | SC | SC | Expired |  | Angela Kumanoo Tel: 0203 447 0117 | The Analyser is on the main screen, but buttons are missing and nothing will allow anyone to proceed. Intervention Required is displayed. Restarted the Analyser and the same issue occurs. Powered down the Analyser, removed the Inlet and SP. Powered on and now the Analyser is in intervention required and able to carry out replacements. The analyser screen appears with no buttons after the automatic log off. SC is out of date and is the cause of the issue (LOT 946-010 R0608N196 20 /10 /2018) SP was not activated correctly. | Removed SC and informed the customer of the present issues. SC needs to replaced by the customer. No further action required.
No | Visit | Complaint | RDEU | Anruf von Frau Gärtner; Gerät steht in Intervention erforderlich, wurde schon alles getauscht, ohne Erfolg. Bitte um Reparatur. | Nach Wartungsmassnahme besteht das Problem weiterhin. Gerinnsel in Sensorcassette erkannt, Wert liegt bei 1,38 (0,80-1,20) Peristaltik Pumpe ausgetauscht, Hemolyzer ausgetauscht, Wert stabil auf 1,31 gesunken. Sensor Bay mit Hypochlorite in beide Richtungen gereinigt, ohne Erfolg. | Sensor mit dünnem Drat gereinigt, Wert auf 1,05 stabil gesunken. Startphase i.O. Funktionsprüfung i.O Gerät an Labor übergeben. | SWO-441539 | Defect | SC | SC | clot |  | Call from Mrs Gärtner; Device is required in intervention, everything has already been exchanged, to no avail. Please repair. | After maintenance, the problem persists. Clot detected in sensor cassette, value is 1.38 (0.80-1.20) Peristaltic pump replaced, hemolyzer replaced, value dropped stable to 1.31. Sensor Bay cleaned with hypochlorite in both directions, to no avail. | Sensor cleaned with thin drat, value dropped to 1.05 stable. Start-up phase in the functional test i.O device to laboratory.
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 液体传输失败，1277 | 更换试剂包后修复 | SWO-441649 | Assisted Replacement | SP | SP | FT | 1277 | Test card, reagent package failure | Liquid transfer failure, 1277 | Repair after replacement of reagent kit
No | Visit | Complaint | RCHN | 样本量不足 | 样本量不足 | 疏通管路后修复 by Andrew | SWO-441045 | Cleaned | HW | Aspiration Module | Needle Wash Unit |  | Insufficient sample size | Insufficient sample size | Repair by Andrew after dredging lines
No | Visit | Complaint | RCHN | 定标失败 | 定标失败 借944-220（14921）一张 | 重新定标后修复 by Andrew | SWO-439878 | Assisted Replacement | Consumables | Test Kit | Cartridge, cal |  | Bid failed | Bid failure Borrows 944-220 (14921) | Fix by Andrew after re-calibration
No | Visit | Complaint | RCHN | 定标步骤被打断 1059 | 定标步骤被打断 1059 借944-394（14730）944-214（14322）各一张 | 重新定标后修复 by Andrew | SWO-440429 | Assisted Replacement | Consumables | Test Kit | Cartridge, cal | 1059 | The bid step was interrupted 1059 | The bid step was interrupted 1059 Borrowing 944-394 (14730) 944-214 (14322) each | Fix by Andrew after re-calibration
No | Visit | Complaint | RKK | 測定中に中止になってしまう。 | 機器測定中に中止になる | デモ機設置 | SWO-442235 | Unknown | Unknown | Unknown | Unknown |  | It becomes canceled during the measurement. | Abortduring equipment measurement | Demonstration machine set up
No | Visit | Complaint | RCHN | 条码扫描器无法扫描 | 条码扫描器无法扫描，需要借件902-924一个 | 更换902-924后修复 by Andrew | SWO-440431 | Defect | HW | Cover and Chassis | Barcode reader, display |  | Barcode scanner cannot be scanned | Barcode scanner cannot be scanned and requires a loan of 902-924 | Repair by Andrew after replacing 902-924
Yes | Visit | Complaint | RHUN | a készülék nem üzemképes, lefagy, telefonban újraindítás után kb 40 perc múlva ismét lefagyott, a kijelzőn 'nincs kapcsolat' kijelzés | 11.03. 18:30-tól 8db 641 újraindulás található a naplóban ------ kiérkezéskor 734-es hibakód a rendszer üzenetekben, a készülék újraindítás után rendben elindul, távoli elérés opció bekapcsolásakor az analizátor újraindul | analizátor sw újra telepítése, kalibráció, QC mérés *** TCP kábel ideiglenesen (mivel hálózati problémák vannak az osztályon) nincs csatlakoztatva | SWO-441543 | [5 empty] | The device is not working, freezes, and in the phone after a reboot is about 40 minutes later it crashed again, Display shows ' no connection ' | 11.03.18:30 from 8db 641 restart in log ------ On arrival, error code 734 is returned in system messages. The device starts after a reboot, remote access option will restart the analyzer | Analyzer SW re-installation, calibration, QC measurement *** TCP Cable temporarily (due to network problems in the class) not connected
No | Visit | Complaint | RCHN | TNI定标失败 | TNI定标失败 借944-212（14804）一张 | 重新定标正常。by jesson | SWO-441382 | Assisted Replacement | Consumables | Test Kit | Cartridge, cal | 4036 | TNI calibration failed | TNI bid failure borrowed 944-212 (14804) a | Recalibration is normal. by jesson
No | Visit | Complaint | RCHN | 自动质控模块机械错误，报错0740 | 自动质控模块机械错误，报错0740 借902-607一个 |  | SWO-441483 | Defect | HW | Autocheck module | Distribution board | 0740 | Automatic QUALITY CONTROL MODULE MECHANICAL ERROR, ERROR 0740 | Automatic QUALITY CONTROL MODULE MECHANICAL ERROR, REPORT ERROR 0740 BORROW902-607 ONE
No | Visit | Complaint | RAME | Battery is not charging, showing 0%. | Battery is not charging. | Checked and, confirmed the battery charging on the analyzer screen is showing 0%. Replaced a new battery. | SWO-442243 | Defect | HW | Battery module | Battery module | 0% | Battery is not charging, showing 0%. | Battery is not charging. | Checked and, confirmed the battery charging on the analyzer screen is showing 0%. Replaced a new battery.
No | Visit | Complaint | RPAC | Customer reported machine unable to pump solution. Machine stuck at Intervention required. | The reported unable to pump solution error has gone. pCO2 however has been failing calibration (sens too low) . | Fluid transport tests - passed Sensor channel tests - passed Replaced the sensor cassette and reinitialised. | SWO-442249 | No fault found | No fault found | FT | FT | SC replaced | Customer reported machine unable to pump solution. Machine stuck at Intervention required. | The reported unable to pump solution error has gone. pCO2 however has been failing calibration (sens too low) . | Fluid transport tests - passed Sensor channel tests - passed Replaced the sensor cassette and reinitialised.
No | Visit | Complaint | RAME | Rinse error | Review data logs, performed visual inspection, performed applicable service diagnostics. Observed possible debris in Blue air $gas valve for MET II module. | Confirmed debris on center of valve membrane causing leak, not allowing rinse to be aspirated. Replace both the Rinse and Air $Gas valves. Tested via diagnostics, initialization and calibration, no further errors observed. | SWO-445603 | Defect | HW | Met II module | Valve, membrane | 731, 751, 510 | Rinse error | Review data logs, performed visual inspection, performed applicable service diagnostics. Observed possible debris in Blue air /gas valve for MET II module. | Confirmed debris on center of valve membrane causing leak, not allowing rinse to be aspirated. Replace both the Rinse and Air $Gas valves. Tested via diagnostics, initialization and calibration, no further errors observed.
No | Visit | Complaint | RIND | Parameter yellow | Glucose parameter is yellow. | Check the calibration report senstiviy is low , clean the chamber ,run two point calibration , senstivity is low . Change the Glucose membrane , Run two point calibration , Senstivity is in range. | SWO-448446 | Assisted Replacement | Electrode | Membrane | Glu | Sens low | Parameter yellow | Glucose parameter is yellow. | Check the calibration report senstiviy is low , clean the chamber ,run two point calibration , senstivity is low . Change the Glucose membrane , Run two point calibration , Senstivity is in range.
No | Visit | Complaint | RIND | Pco2 qc low values | Pco2 qc values low | checked the analyser there was puss formation in the electrode and calibrated the analyser to ready mode and analysed the sample | SWO-448170 | Cleaned | Electrode | Electrode | pCO2 | QC low | Pco2 qc low values | Pco2 qc values low | checked the analyser there was puss formation in the electrode and calibrated the analyser to ready mode and analysed the sample
No | Visit | Complaint | RAME | Power outage | Analyzer is showing error Data Status Code 9 after power outage. | Exited Roma, deleted Roma.ini file that was corrupted, pasted a Roma.ini file from previous backup, rebooted analyzer, and validated. | SWO-442109 | Defect | SW | RMED SW | Roma.ini |  | Power outage | Analyzer is showing error Data Status Code 9 after power outage. | Exited Roma, deleted Roma.ini file that was corrupted, pasted a Roma.ini file from previous backup, rebooted analyzer, and validated.
No | Visit | Complaint | RAME | Analyzer is not loading software and showing data error | Analyzer is showing error Data Status Code 9 after power outage. | Exited Roma, deleted Roma.ini file that was corrupted, pasted a Roma.ini file from previous backup, rebooted analyzer, and validated. | SWO-442105 | Defect | SW | RMED SW | Roma.ini |  | Analyzer is not loading software and showing data error | Analyzer is showing error Data Status Code 9 after power outage. | Exited Roma, deleted Roma.ini file that was corrupted, pasted a Roma.ini file from previous backup, rebooted analyzer, and validated.
No | Visit | Complaint | RIND | Analyzer showing neon number too high. | Analyzer showing error neon number too high. | checked the analyzer and performed complete oximetry test and its failed lamp unit neon position 2.2 .we need lamp unit and oxi control PCB for this issue. | SWO-441753 | Revisit required | HW | Oxi module | Oxi module |  | Analyzer showing neon number too high. | Analyzer showing error neon number too high. | checked the analyzer and performed complete oximetry test and its failed lamp unit neon position 2.2 .we need lamp unit and oxi control PCB for this issue.
No | Visit | Complaint | RFRA | Erreur Portes des échantillons qui se bloque Erreur Ecran figé obligé de rebooter l'AQT | Erreur porte qui se bloque 5784 Satesafe_22 Erreur température d’incubation hors gamme | Vérification mécanique porte ok. Porte de gauche à l’air de plus bloquer. De temps en temps erreur en service. Démontage, nettoyage, lubrification et changement des courroies ok. Remise en place avec jeu suffisant ok. Plusieurs essais en service ouverture $ fermeture ok, plus d’erreurs. Rapport à envoyer par mail à Qualite@labo-barla.eu ok. Counter needle a 2300. | SWO-442018 | Adjustment | HW | Cover and Chassis | Door, Inlet |  | Error Sample Doors that hangs Frozen Screen Error Forced to Reboot AQT | Door Error That Crashes 5784 Satesafe-22 Out-of-range incubation temperature error | Mechanical check door ok. Door to the left looks more block. Occasionally error in service. Disassembly, cleaning, lubrication and changing belts ok. Put back up with enough play ok. Several trials in service opening $ closing ok, more errors. Report to be emailed to Qualite@labo-barla.eu okay. Counter needle has 2300.
No | Visit | Complaint | RAME | Pump errors | 1) Restricted flow in analyzer. 2)Error 1079: Sensor impedance error. | 1) Conducted analyzer backflush in sections and replaced tubing. 2) Replaced Sensor Cassette. | SWO-447010 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway | Clot | Pump errors | 1) Restricted flow in analyzer. 2)Error 1079: Sensor impedance error. | 1) Conducted analyzer backflush in sections and replaced tubing. 2) Replaced Sensor Cassette.
Yes | Visit | Complaint | RAME | Status code 7 | Data upgrade failed on boot-up. | Reloaded, OS, App SW, Setup and configured analyzer. | SWO-447005 | [5 empty] | Status code 7 | Data upgrade failed on boot-up. | Reloaded, OS, App SW, Setup and configured analyzer.
No | Visit | Complaint | RFRA | 1ère cassette HS, 2eme cassette erreur calibration pCO2. | Contrôle avec la cassette de test: ok. Intervertir les 2 cassettes: toujours erreur pCO2. Ok pour l'autre ABL90. Intervertir packs solution: toujours erreur pCO2. Ok pour l'autre ABL90 | Chgt carte sensor interface. Installation pack solution et cassette. Calibration, passage CQ. | SWO-445728 | Defect | HW | Basic Unit | PCB, Sensor interface | pCO2 | 1st cassette HS, 2nd cassette error calibration pCO2. | Control with the test cassette: ok. Switch the 2 cassettes: always error pCO2. Ok for the other ABL90. Intervertir packs solution: always error pCO2. Ok for the other ABL90 | Chgt map sensor interface. Installing solution pack and cassette. Calibration, CQ passage.
No | Visit | Complaint | RCHN | 孵玉盘转动错误 | 孵育盘转动错误，定标卡条码器无法扫描，借636-536和910-304各一个 | 更换636-536和910-304后修复 by jesson | SWO-441683 | Defect | HW | Incubator Module | Sensor | 5784 | Incubating jade plate rotation error | Incubation disk rotation error, calibration card barcode can not be scanned, borrow 636-536 and 910-304 each one | Repair by jesson after replacing 636-536 and 910-304
No | Visit | Complaint | RCHN | 仪器断电无法使用 | 仪器断电无法使用 | 来电后，机器正常启动。测量样本结果出现hct无法测量故障。重新启动仪器修复 by jesson | SWO-442298 | No fault found | No fault found | No fault found | No fault found |  | The instrument is not in use if it is powered | The instrument is not in use if it is powered | After the call, the machine starts normally. The measurement sample results have hct that cannot measure the fault. Restart the instrument to fix by jesson
Yes | Visit | Complaint | RIND | ABL9 Installation. | ABL9 installation required. | Complete installation of ABL9 with all accessories done. Calibration & patient setup checked & Found working satisfactory. | SWO-445608 | Re classified RMED | Installation |  |  |  | ABL9 Installation. | ABL9 installation required. | Complete installation of ABL9 with all accessories done. Calibration & patient setup checked & Found working satisfactory.
No | Visit | Complaint | RESP | Erro de calibração de bomba | 0793: Erro de assimetria da bomba de pH $GS | Verificação do registro de dados. Substituição dos tubos das bombas dos módulos dos eléctrodos. Verificação dos roletas das bombas. Inicialização. Calibração das bombas. Calibração e verificação do correto funcionamento com controle de qualidade. | SWO-445840 | Defect | HW | pH/BG Module | Pump tubes | 793 | Pump Calibration Error | 0793: /GS pH pump asymmetry Error | Checking the data record. Replacing the pump tubes of the electrode modules. Check the pumps ' rolets. Startup. Calibration of pumps. Calibration and verification of the correct operation with quality control.
No | Visit | Complaint | RAME | Analyzer receiving rinse error | Analyzer showing rinse error. Opened covers and checked all tubings. Reviewed activity log, to see if error is caused by solution replacement. Finding nothing I checked analyzer valves. | Found hole in pH $BG air valve. Replaced and ran calibrations and qc. | SWO-446096 | Defect | HW | pH/BG Module | Valve, air |  | Analyzer receiving rinse error | Analyzer showing rinse error. Opened covers and checked all tubings. Reviewed activity log, to see if error is caused by solution replacement. Finding nothing I checked analyzer valves. | Found hole in pH $BG air valve. Replaced and ran calibrations and qc.
No | Visit | Complaint | RDEU | 06.11.2018 Frau Schreiber 021513245135 meldet, dass seit letzter Woche ein pCO2 Fehler angezeigt wird. | Immer wieder pCO2 oder E'lytefehler, Status u. Empfl.fehler, Airventile in beiden Modulen klemmen, | Module ausgebaut und alle Ventile gereinigt, Gerät kalibriert und Qc's durchgeführt, | SWO-445867 | Cleaned | HW | pH/BG Module | Valve, air |  | 06.11.2018 Ms. Schreiber 021513245135 reports that a pCO2 error has been displayed since last week. | Time and again pCO2 or E'lyte error, status and recommendation errors, air valves in both modules clamp, | modules removed and all valves cleaned, device calibrated and Qc's performed,
No | Visit | Complaint | RDAN | Jane Smærup klager over tilbagevendende pco2 qc problemer. Ligger for lavt | Undersøgt. | Gasventil og pakring f. pH elektrode udskiftet. | SWO-445734 | Defect | HW | pH/BG Module | Valve, air |  | Jane Smærup complains about recurrent pCO2 QC problems. Is too low | Examined. | Gas valve and packaging F. pH electrode replaced.
No | Visit | Complaint | RIND | Showing status of all installed parameter keys as expired | Showing status of installed keys as expired | Installed the keys of required parameters and run calibration and samples | SWO-448331 | User Error, FSE | Demo key | Demo key | Demo key |  | Showing status of all installed parameter keys as expired | Showing status of installed keys as expired | Installed the keys of required parameters and run calibration and samples
No | Visit | Complaint | RIND | Printer have problem, its not printing properly. |  | Replaced the Printer and observed the print | SWO-448328 | Revisit required | HW | Printer Module | Printer Module |  | Printer have problem, its not printing properly. |  | Replaced the Printer and observed the print
Yes | Visit | Complaint | RPOL | Konfiguracja sieci oraz czytnika kodów kreskowych | Brak połączenia LIS $HIS. Błąd odczytu nowych kodów kreskowych 2 $5 | Konfiguracja analizatora . | SWO-445957 | Re classified RMED | planned SWO |  |  | see mail with SWO number reference | Network and Barcode Reader Setup | No LIS /HIS connection. Error reading new barcodes /2 5 | Analyzer configuration.
No | Visit | Complaint | RDEU | 29.10.2018 14:00 Fr. Richter ????: Interventionsmeldung Gerinsel in der Oximetrie | Gerät meldet weiterhin sporadische Gerinnsel im Oxymetriesystem erkannt und fordert zur Spülung auf Pat.proben werden langsam angesaugt | Gerät überprüft: im Seviceprog. kpl. Oxy Test fehlerfrei, O2 Test mit Fehler bei zpO2_Energy 2,75, Sensor Simuation gestartet: pO2 Fehler bei Rang & Lin. Prüfung Quetschventil -> lässt sich im aktivierten Zustand nicht aufziehen i.O.; Startphase abgewartet,Gerät zerlegt, keine Korrosion an den Kontakten feststellbar, Austausch der Sensor Interface PCB, Austausch der Hemolyzer Unit-> Warmstartphase abgewartet, div. Hilfsprogramme durchgeführt, im Serviceprog. kpl. Oxy Test & pO2 Test durchgeführt i.O., Kal & tHb Kal durchgeführt, pCO2 QC Fehler im Level A -> Kal & QC wiederholt danach QC i.O. Servicedump & Datalogs erstellt incl. Sicherung und Dokumentation, mehrere Patienteprobenmessungen durchgeführt i.O. | SWO-446676 | Defect | HW | Oxi module | Hemolyzer | Clot | 29.10.2018 14:00 Fr. Richter ????: Intervention report Gerinsel in oximetry | Device continues to report sporadic clots detected in oxymetry system and prompts for flushing Pat.samples are slowly sucked in | device checked: in the Seviceprog. Compl. Oxy Test error-free, O2 test with error at zpO2_Energy 2.75, Sensor Simuation started: pO2 error at Rank & Lin. Testing crushing valve -> cannot be woe when activated, i.O.; Start phase awaited, device disassembled, no corrosion detectable at the contacts, replacement of the sensor interface PCB, replacement of the Hemolyzer Unit-> warm start phase awaited, various auxiliary programs carried out, in the service program. Compl. Oxy Test & pO2 Test performed i.O., Kal & tHb Kal performed, pCO2 QC error in level A -> Kal & QC repeats after that QC in the following. Servicedump & Datalogs creates incl. backup and documentation, several patient sample measurements carried out in i.O.
No | Visit | Complaint | RFRA | Pb de démarrage de Windows | Tentative de récupération | Changement de la carte mère et de la compact flash. Réinstallation de la configuration et des paramètres de sauvegarde. Cal tHb. Essais | SWO-446651 | Defect | HW | DMS_PC | CPU | Windows | Windows Startup Pb | Attempted recovery | Change of the motherboard and the compact flash. Reinstall ingconfiguration and backup settings. Cal tHb. Tests
No | Visit | Complaint | RAME | Battery Showing 0% | Battery Showing 0% | Replaced battery from customers stock. | SWO-446801 | Defect | HW | Battery module | Battery module | 0% | Battery Showing 0% | Battery Showing 0% | Replaced battery from customers stock.
No | Visit | Complaint | RAME | Battery Showing 0% | Battery Showing 0% | Replaced battery from customer stock. | SWO-446714 | Defect | HW | Battery module | Battery module | 0% | Battery Showing 0% | Battery Showing 0% | Replaced battery from customer stock.
No | Visit | Complaint | RCHN | 测试卡卡盘 5784 | 测试卡卡盘 5784 | 更换测试卡正常。 | SWO-446481 | Defect | Consumables | Test Kit | Cartridge, Broken | Teeth; 5784 | Test Card 5784 | Test Card 5784 | The replacement test card is OK.
No | Visit | Complaint | RPOL | Błędy temperatury substratu | Weryfikacja pracy analizatora | Wymiana kasety sensorowego, weryfikacja parametrów technicznych, czyszczenie igieł sensorowych, testowanie, kalibracja. | SWO-446870 | Assisted Replacement | SC | SC | SC | Temp | Substrate Temperature Errors | Validate Analyzer Work | Replacement of the sensor cassette, verification of technical parameters, cleaning of sensor needles, testing, calibration.
No | Visit | Complaint | RPOL | Uszkodzona płyta reagent pack PCB | Błąd odczytu pakietu odczynnikowego | Wymiana uszkodzonej płyty reagent pack PCB . | SWO-446638 | Defect | HW | Cover and Chassis | PCB, solution pack | mail: Lot of SP | Uszkodzona płyta reagent pack PCB | Defuncting Packet Read error | Replacing the damaged disc Reagent pack PCB.
No | Visit | Complaint | RPOL | Wyciek z uszczelki wlotu. | Weryfikacja pracy analizatora | Wymiana uszczelki wlotu, oczyszczenie komory pomiarowej pH, Ref, testowanie, kalibracja, kontrola | SWO-446616 | Defect | HW | Inlet module | Inlet gasket |  | Leakage from inlet gasket. | Validate Analyzer Work | Inlet gasket replacement, pH measurement chamber Cleaning, REF, testing, calibration, inspection
No | Visit | Complaint | RKK | SC交換時に毎回Lacがエラーとなる。 | センサーインターフェイスPCB故障 | センサーインターフェイスPCB交換 | SWO-449322 | Defect | HW | Basic Unit | PCB, Sensor interface | Lac | Lac fails every time you replace SC. | Sensor interface PCB failure | Sensor Interface PCB Replacement
No | Visit | Complaint | RAME | Inlet gasket with holder sticking on probe when sampling Victoria at 813-396-6227 | Inlet gasket was sticking | Found the probe had a slight bend $Fixed problem $Also worked on getting live connect | SWO-447713 | Adjustment | HW | Inlet module | Probe bend |  | Inlet gasket with holder sticking on probe when sampling Victoria at 813-396-6227 | Inlet gasket was sticking | Found the probe had a slight bend $Fixed problem $Also worked on getting live connect
Yes | Visit | Complaint | RLTD | 0208 725 1503 Sarah Bond |  |  | SWO-447502 | Re classified RMED | same day upSWO |  |  |  | 0208 725 1503 Sarah Bond
No | Visit | Complaint | RLTD | 0208 725 1503 Sarah Bond | Database corrupted. | Renewed database and refreshed operators. Replaced solution pack in order to clear inconsistent db error. | SWO-448039 | Adjustment | SW | RMED SW | Database |  | 0208 725 1503 Sarah Bond | Database corrupted. | Renewed database and refreshed operators. Replaced solution pack in order to clear inconsistent db error.
No | Visit | Complaint | RAME | Expired SC and SP | Radiometer CAS unable to conduct training. | Replaced SP and SC and assisted training. | SWO-447631 | User Error | User Error | User Error | Consumables | lack of training | Expired SC and SP | Radiometer CAS unable to conduct training. | Replaced SP and SC and assisted training.
No | Visit | Complaint | RCHN | 泵管有气泡 | 泵管有气泡,借903-099一个。 | 在维修菜单中检测，发现只有在测试阀4的时候，测试卡的管路有气泡。更换903-099后修复 by jesson | SWO-446454 | [5 empty] | There are air bubbles in the pump tube | The pump tube has bubbles, borrow edith 903-099. | Tested in the service menu, it is found that only when testing valve 4, the test card line has air bubbles. Repair by jesson after replacement 903-099
No | Visit | Complaint | RCHN | 提示需要冲洗，需更换试剂包 | 提示需要冲洗，需更换试剂包 | 更换进样口垫片正常。by jesson | SWO-447815 | Defect | HW | Inlet module | Inlet gasket |  | Tips need to be flushed and reagent kits need to be replaced | Tips need to be flushed and reagent kits need to be replaced | Replacing the inlet gasket is correct. by jesson
No | Visit | Complaint | RCHN | 黑屏 | 黑屏,借903-175一个。 | 未更换备件，实际是电源线烧毁，更换电源线后修复 by jesson | SWO-446444 | Defect | HW | External HW | Power cable | 903-175 | Black screen | Black screen, borrow 903-175 one. | No replacement parts, actually the power cord burned, replace the power cord after repair by jesson
No | Visit | Complaint | RCHN | 处理盘弹簧片变形 5684 | 处理盘弹簧片变形 5684,借 903-236（被903-702替代）一个。 | 更换903-702后修复 by jesson | SWO-445686 | Defect | HW | Process Module | Process Wheel | 5684 | Processing disc spring sheet deformation 5684 | Process disc spring disc deformation 5684, borrow 903-236 (replaced by 903-702) one. | Repair by jesson after replacement of 903-702
No | Visit | Complaint | RDEU | Hotline - ständig wechselnde Driften bei Lyten und Lac | Fehlersuche, immer wieder Driften bei Lyten und teils anderen Parametern, alle Ventile in Messmodulen und FluidicChannel Unit überprüft und gereinigt, keine größeren Auffälligkeiten, Lösungspumpe arbeitet nicht korrekt - Pumpenschlauch ersetzt - kurze Aussetzer, während Kal dann komplett ausgefallen | MedTech stellte Modul von ausgemustertem Gerät zur Verfügung, Pumpe ersetzt - gleiches Phänomen, Modul komplett ersetzt, Neustart, mehrere Kal und QC | SWO-447570 | Defect | HW | El/Met Module | El/Met Module | Lac and Electrolyte Drift | Hotline - constantly changing drifts at Lyten and Lac | Troubleshooting, repeatedly drifting with lytes and partly other parameters, all valves checked and cleaned in measuring modules and FluidicChannel unit, no major abnormalities, solution pump does not work correctly - pump hose replaced - short dropouts while Kal then completely failed | MedTech provided module of retired device, pump replaced - same phenomenon, module completely replaced, restart, multiple Kal and QC
No | Visit | Complaint | RIND | Calibraiton error | Checked the data log crea calibration senstivity out of range | Removed Air bubbles in crea membrane & cleaned the chamber & checked equipment is functional | SWO-447917 | Cleaned | Electrode | Electrode | Crea | Air bubble | Calibraiton error | Checked the data log crea calibration senstivity out of range | Removed Air bubbles in crea membrane & cleaned the chamber & checked equipment is functional
No | Visit | Complaint | RIND | Leakage from capillary inlet. | Leakage from inlet & check the activity log found query in the capillary sample | Clean the gasket,run the protein removal,cleaning,Run the 2point calibration & no leakage found,Run the sample got the desired result. | SWO-448329 | Cleaned | HW | Inlet module | Inlet gasket |  | Leakage from capillary inlet. | Leakage from inlet & check the activity log found query in the capillary sample | Clean the gasket,run the protein removal,cleaning,Run the 2point calibration & no leakage found,Run the sample got the desired result.
No | Visit | Complaint | RIND | Printer have problem, its not printing properly. | its printing full black print ,without results. And checked the printer found thermal head continually heating. | Checked the printer connections and reconnected all connections.restated the analyzer found printer working fine. | SWO-448421 | Adjustment | HW | Printer Module | Printer Module |  | Printer have problem, its not printing properly. | its printing full black print ,without results. And checked the printer found thermal head continually heating. | Checked the printer connections and reconnected all connections.restated the analyzer found printer working fine.
No | Visit | Complaint | RAME | Temp error. PH/BG module | Recurring 512 temp errors. | Replaced pH $BG module. Programmed and set zero current on new module. | SWO-447683 | Defect | HW | pH/BG Module | pH/BG Module | 512 | Temp error. PH/BG module | Recurring 512 temp errors. | Replaced pH $BG module. Programmed and set zero current on new module.
Yes | Visit | Complaint | RDEU | TNT Kassetten werden nicht auf Anhieb gelesen. Es muss mehrmals versucht werden. | Analysator beim Anwender ausgeliefert. | Überprüfung des Testkassetten Barcode Scanners im Service. Einstellung verstellt, korrigiert. Alle Testkassetten werden erkannt. | SWO-447676 | Re classified RMED | installation report |  |  | Move to OOB trending | TNT cassettes are not read immediately. It must be tried several times. | analyzer at the user's end. | Checking the test cartridges barcode scanner in the service. Setting is adjusted, corrected. All test cartridges are detected.
No | Visit | Complaint | RPOL | Błąd modułu OXI | Diagnoza problemu. | Oczyszczenie toru pomiarowego, sprawdzenie modułu OXI, kalibracja tHb. | SWO-447597 | Cleaned | HW | Oxi module | Oxi module |  | OXI Module Error | Diagnosis of the problem. | Cleaning the measuring track, checking the OXI module, calibrating the tHb.
No | Visit | Complaint | RAME | Po2 low on qc 2 | Po2 low on qc level 2 Range 288-312 results 286-292 Customer has replace solution pack and sensor previously Customer has flushed analyzer I ran po2 complete check and solution pack check all ok I checked the po2 window above the sensor and found smudges After cleaning the window po2 jumped and stayed stable | Po2 window cleaning | SWO-448205 | Adjustment | HW | Basic Unit | pO2 Window |  | Po2 low on qc 2 | Po2 low on qc level 2 Range 288-312 results 286-292 Customer has replace solution pack and sensor previously Customer has flushed analyzer I ran po2 complete check and solution pack check all ok I checked the po2 window above the sensor and found smudges After cleaning the window po2 jumped and stayed stable | Po2 window cleaning
No | Visit | Complaint | RLTD | Anna 07984464577 | Analyser keeps running slow and locking up and now has status error code 2 | Unit has had software installed before but then start to run slow and locks up. Fitted new PC with flash card. Installed analyser software ok. Installed analyser setup same as others on ward ok. Analyser has a clean data base installed and consumable put on ok. Analyser started up ok, calibration and QC ok. | SWO-448059 | Defect | HW | DMS_PC | CPU | status code 2 | Anna 07984464577 | Analyser keeps running slow and locking up and now has status error code 2 | Unit has had software installed before but then start to run slow and locks up. Fitted new PC with flash card. Installed analyser software ok. Installed analyser setup same as others on ward ok. Analyser has a clean data base installed and consumable put on ok. Analyser started up ok, calibration and QC ok.
No | Visit | Complaint | RAME | From: Reiprish, Amy <Amy.Reiprish@uphs.upenn.edu> Sent: Thursday, November 8, 2018 4:14 PM To: Aitas, Tom TAITA <taitas@radiometercanada.com>; Le, Vy <Vy.Le@uphs.upenn.edu>; PANAK, RICHARD <Richard.Panak2@uphs.upenn.edu>; Iwankovitsch, James JRIWA <jiwankovitsch@radiometeramerica.com>; Scannell, Katelyn KATSC <Katelyn.Scannell@radiometeramerica.com> Subject: HUP OR ABL FLEX+ issue Tom Our analyzers we put into the or this week are reporting the results out as undefined. They are not posting to the patient chart. It is only the new or meters we are seeing this issue with. Amy Reiprish | From: Reiprish, Amy <Amy.Reiprish@uphs.upenn.edu> Sent: Thursday, November 8, 2018 4:14 PM To: Aitas, Tom TAITA <taitas@radiometercanada.com>; Le, Vy <Vy.Le@uphs.upenn.edu>; PANAK, RICHARD <Richard.Panak2@uphs.upenn.edu>; Iwankovitsch, James JRIWA <jiwankovitsch@radiometeramerica.com>; Scannell, Katelyn KATSC <Katelyn.Scannell@radiometeramerica.com> Subject: HUP OR ABL FLEX+ issue Tom Our analyzers we put into the or this week are reporting the results out as undefined. They are not posting to the patient chart. It is only the new or meters we are seeing this issue with. Amy Reiprish | On Nov 8, 2018, at 5:16 PM, Aitas, Tom TAITA <taitas@radiometercanada.com> wrote: Hi looks like the list of number I received was missing an extra 0 at the end. I’ve checked them from what is in AQURE. Kind Regards, Tom Aitas Cell: 905.512.1822 taitas@radiometercanada.com From: Vazquez, Brian BVAZQ Sent: Thursday, November 08, 2018 4:53 PM To: Iwankovitsch, James JRIWA <jiwankovitsch@radiometeramerica.com> Cc: Aitas, Tom TAITA <taitas@radiometercanada.com>; Riccardi, Robert RRICC <rriccardi@radiometeramerica.com> Subject: Re: HUP OR ABL FLEX+ issue Thanks. That’s how it should be on instrument, but there is also a rule in AQURE so it doesn’t sit as pending and automatically passes through to LIS. May be something different, but just didn’t know if the 092R was correct in rule. Brian Vazquez National Support Manager Radiometer America Inc. Sent from my iPhone On Nov 8, 2018, at 4:41 PM, Iwankovitsch, James JRIWA <jiwankovitsch@radiometeramerica.com> wrote: Brian – just confirmed with the customer, Enable Patient Result Approval is uncheck for the Silver ORs. All other analyzers (including EDs) it is checked. <image001.png> -----Original Message----- From: PANAK, RICHARD <Richard.Panak2@uphs.upenn.edu> Sent: Thursday, November 8, 2018 4:30 PM To: Le, Vy <Vy.Le@uphs.upenn.edu>; Aitas, Tom TAITA <taitas@radiometercanada.com>; Reiprish, Amy <Amy.Reiprish@uphs.upenn.edu>; Iwankovitsch, James JRIWA <jiwankovitsch@radiometeramerica.com>; Scannell, Katelyn KATSC <Katelyn.Scannell@radiometeramerica.com> Subject: RE: HUP OR ABL FLEX+ issue Hi Tom, Not sure if this helps, but this is what we are seeing in AQURE for the new OR meters. It's been like this for every patient since Monday when the meters were put into use. Richard M. Panak Jr., BS, MLS(ASCP)CM Point of Care Coordinator - Hospital of the University of Pennsylvania Mobile $Text: 215-606-8138 Office: 215-349-5879 (8:30AM-5PM, Mon-Fri) After Hours $On-Ca | SWO-447742 | Setup | SW | External SW | Aqure | Serial Number missing zero | From: Reiprish, Amy <Amy.Reiprish@uphs.upenn.edu> Sent: Thursday, November 8, 2018 4:14 PM To: Aitas, Tom TAITA <taitas@radiometercanada.com>; Le, Vy <Vy.Le@uphs.upenn.edu>; PANAK, RICHARD <Richard.Panak2@uphs.upenn.edu>; Iwankovitsch, James JRIWA <jiwankovitsch@radiometeramerica.com>; Scannell, Katelyn KATSC <Katelyn.Scannell@radiometeramerica.com> Subject: HUP OR ABL FLEX+ issue Tom Our analyzers we put into the or this week are reporting the results out as undefined. They are not posting to the patient chart. It is only the new or meters we are seeing this issue with. Amy Reiprish | From: Reiprish, Amy <Amy.Reiprish@uphs.upenn.edu> Sent: Thursday, November 8, 2018 4:14 PM To: Aitas, Tom TAITA <taitas@radiometercanada.com>; Le, Vy <Vy.Le@uphs.upenn.edu>; PANAK, RICHARD <Richard.Panak2@uphs.upenn.edu>; Iwankovitsch, James JRIWA <jiwankovitsch@radiometeramerica.com>; Scannell, Katelyn KATSC <Katelyn.Scannell@radiometeramerica.com> Subject: HUP OR ABL FLEX+ issue Tom Our analyzers we put into the or this week are reporting the results out as undefined. They are not posting to the patient chart. It is only the new or meters we are seeing this issue with. Amy Reiprish | On Nov 8, 2018, at 5:16 PM, Aitas, Tom TAITA <taitas@radiometercanada.com> wrote: Hi looks like the list of number I received was missing an extra 0 at the end. I’ve checked them from what is in AQURE. Kind Regards, Tom Aitas Cell: 905.512.1822 taitas@radiometercanada.com From: Vazquez, Brian BVAZQ Sent: Thursday, November 08, 2018 4:53 PM To: Iwankovitsch, James JRIWA <jiwankovitsch@radiometeramerica.com> Cc: Aitas, Tom TAITA <taitas@radiometercanada.com>; Riccardi, Robert RRICC <rriccardi@radiometeramerica.com> Subject: Re: HUP OR ABL FLEX+ issue Thanks. That’s how it should be on instrument, but there is also a rule in AQURE so it doesn’t sit as pending and automatically passes through to LIS. May be something different, but just didn’t know if the 092R was correct in rule. Brian Vazquez National Support Manager Radiometer America Inc. Sent from my iPhone On Nov 8, 2018, at 4:41 PM, Iwankovitsch, James JRIWA <jiwankovitsch@radiometeramerica.com> wrote: Brian – just confirmed with the customer, Enable Patient Result Approval is uncheck for the Silver ORs. All other analyzers (including EDs) it is checked. <image001.png> -----Original Message----- From: PANAK, RICHARD <Richard.Panak2@uphs.upenn.edu> Sent: Thursday, November 8, 2018 4:30 PM To: Le, Vy <Vy.Le@uphs.upenn.edu>; Aitas, Tom TAITA <taitas@radiometercanada.com>; Reiprish, Amy <Amy.Reiprish@uphs.upenn.edu>; Iwankovitsch, James JRIWA <jiwankovitsch@radiometeramerica.com>; Scannell, Katelyn KATSC <Katelyn.Scannell@radiometeramerica.com> Subject: RE: HUP OR ABL FLEX+ issue Hi Tom, Not sure if this helps, but this is what we are seeing in AQURE for the new OR meters. It's been like this for every patient since Monday when the meters were put into use. Richard M. Panak Jr., BS, MLS(ASCP)CM Point of Care Coordinator - Hospital of the University of Pennsylvania Mobile $Text: 215-606-8138 Office: 215-349-5879 (8:30AM-5PM, Mon-Fri) After Hours $On-Ca
No | Visit | Complaint | RLTD | 0208 725 1503 Sarah Bond | Copied set up file from Neuro Theatres analyser and uploaded to this analyser. Ensured all analyser set up is correct - all confirmed. Analyser ready for use. | Corrected analyser setup. | SWO-448116 | Setup | SW | RMED SW | Analyzer setup |  | 0208 725 1503 Sarah Bond | Copied set up file from Neuro Theatres analyser and uploaded to this analyser. Ensured all analyser set up is correct - all confirmed. Analyser ready for use. | Corrected analyser setup.
Yes | Visit | Complaint | RSCH | SC vom 5.11.18 noch nicht geliefert | Kunde hat am 5.11.SC bestellt und nicht erhalten. Es wurde als Übergang ein ABL825 installiert S $N 754R0582N001 | Kal 1 Lösung im KSSG geholt und in Herisau auf der IPS installiert. Funktionskontrolle i.o Kulant | SWO-447975 | [5 empty] | SC of 5.11.18 not yet delivered | Customer ordered on 5.11.SC and did not receive it. An ABL825 has been installed as a transition S /N 754R0582N001 | Kal 1 solution in the KSSG and installed in Herisau on the IPS. Functional control i.o Kulant
No | Visit | Complaint | RIND | Frequently calibration and cassette failures. | Found customer using sodium heparin 25000UI,due to this its effecting Ca++,Na+ and HCT values and calibration also. | Asked customer to use samplers PICO50,for uninterrupted breakdown of instrument. presently instrument working fine. | SWO-448422 | [5 empty] | Frequently calibration and cassette failures. | Found customer using sodium heparin 25000UI,due to this its effecting Ca++,Na+ and HCT values and calibration also. | Asked customer to use samplers PICO50,for uninterrupted breakdown of instrument. presently instrument working fine.
No | Visit | Complaint | RCHN | 无法开机 | 无法开机,借903-381，933-441各一个。 | 确定是CPU的故障。但是旧的CPU上的风扇拆不下来。需要加借636-520一个。 | SWO-447110 | [5 empty] | Cannot power on | Cannot be switched on, borrow 903-381, 933-441 each. | Determine that it is a CPU failure. But the fan on the old CPU can't be disassembled. Need to borrow 636-520 one.
Yes | Visit | Complaint | RCHN | 无法开机 | 确定是CPU的故障。但是旧的CPU上的风扇拆不下来。需要加借636-520一个。 | 更换636-520后修复 by jesson | SWO-447832 | Re classified RMED | same day upSWO |  |  |  | Cannot power on | Determine that it is a CPU failure. But the fan on the old CPU can't be disassembled. Need to borrow 636-520 one. | Repair by jesson after replacement 636-520
No | Visit | Complaint | RPOL | Błąd sensora | Weryfikacja pracy analizatora | Po weryfikacji pracy urządzenia w serwisie w Warszawie, wykonano czynności konserwacyjne i porównawcze. Testowanie urządzenia. Instalacja aparatu na oddziale, szkolenie personelu. | SWO-447986 | [5 empty] | Sensor error | Validate Analyzer Work | After verifying the operation of the machine in the service in Warsaw, maintenance and comparative activities were performed. Test the device. Installation of the camera at the ward, staff training.
Yes | Visit | Complaint | RCHN | 报错1031，测试卡无法扫描 | 报错1031，测试卡无法扫描 借903-364，910-304各一个，903-364使用库存光单（0243n005）故再开一个SWO | 更换903-364,910-304后修复 by jesson | SWO-447777 | [5 empty] | Error 1031, test card can't be scanned | Error 1031, test card can not scan borrowing 903-364, 910-304 each one, 903-364 using inventory light sheet (0243n005) so open another SWO | Fix by jesson after replacing 903-364, 910-304
No | Visit | Complaint | RSAF | Failing calibrations on pH and electrolytes and lactate and glucose low sense values | Checked replacement log and when problem started and found it started after replacement of cal1 & cal2 solutions on the 8 $11 $2018 | Correctly installed both cal1 & cal2 calibration solutions as they where installed cal1 in cal2 slot and cal2 installed in cal1 slot. | SWO-448114 | User Error | HW | Solution | Cal |  | Failing calibrations on pH and electrolytes and lactate and glucose low sense values | Checked replacement log and when problem started and found it started after replacement of cal1 & cal2 solutions on the 8 /11 /2018 | Correctly installed both cal1 & cal2 calibration solutions as they where installed cal1 in cal2 slot and cal2 installed in cal1 slot.
No | Visit | Complaint | RPOL | Błąd QC Hb | Błędy kalibracyjne oraz QC hemoglobiny | Kalibracja hemoglobiny | SWO-449754 | Adjustment | HW | Oxi module | tHb Cal |  | Błąd QC Hb | Calibration errors and hemoglobin QC | Hemoglobin Calibration
No | Visit | Complaint | RKK | 圧力テスト吸引エラー | 圧力テスト吸引エラー | 装置分解・各部点検 流体回路洗浄実施 溶液パック交換 CAL・QC確認 | SWO-451640 | Assisted Replacement | SP | SP | FT | 1063 | Pressure test suction error | Pressure test suction error | Equipment disassembly and inspection of each part Fluid circuit cleaning Solution pack replacement CAL/QC Confirmation
No | Visit | Complaint | RPAC | pH unstable | pH unstable | Replace Electrode | SWO-447744 | Assisted Replacement | Electrode | Electrode | pH |  | pH unstable | pH unstable | Replace Electrode
No | Visit | Complaint | RPOL | Błąd kalibracji pompy. | Diagnoza problemu. | Wymiana owężowania pompowego w module pH $BG. Wymiana uszczelki pod elektroda pH. Kalibrację oraz testy serwisowe. | SWO-448496 | Defect | HW | pH/BG Module | Pump tubes |  | Pump calibration error. | Diagnosis of the problem. | Replacing the pump concentration in the pH module $BG. Replacing the gasket with the pH electrode. Calibration and service tests.
No | Visit | Complaint | RAME | Data not crossing over to Sunquest | Data not crossing over to Sunquest | Sunquest corrected the baud. Results now cross. | SWO-449058 | No fault found | SW | External SW | His/Lis |  | Data not crossing over to Sunquest | Data not crossing over to Sunquest | Sunquest corrected the baud. Results now cross.
No | Visit | Complaint | RLTD | Connie Chilcott texted to ask me to contact her. | Connie informed me that there is a blue screen on the display with STOP followed by a line of numbers on it. It has been rebooted and it came back up with the same. | Visit required to either reinstall the software or replace the CPU unit. | SWO-449718 | Revisit required | HW | DMS_PC | CPU |  | Connie Chilcott texted to ask me to contact her. | Connie informed me that there is a blue screen on the display with STOP followed by a line of numbers on it. It has been rebooted and it came back up with the same. | Visit required to either reinstall the software or replace the CPU unit.
No | Visit | Complaint | RDEU | 12.11.2018 13:10Uhr Labor 020189783020: Trotz Wartungsmaßnahmen meldet das Gerät ständig Fehlermeldung | Kalibrationsfehler bei verschiedenen Parametern berichtet. Bei Eintreffen war das Gerät bereits seit ca. 1,5 Std. grün . Der Kunde hatte zuvor vor Lösungspack und Sensorkassette getauscht. Nach der üblichen Einlaufphase (bei Lac bspw. bis zu 4 Std.) waren die Kalibrationen und QCs fehlerfrei. | Keine Maßnahmen am Gerät erforderlich. Reklamation der fehlerhaften Verbrauchsmaterialien. Im Anschluss Durchführung eines Softwareupdates (s. separater Bericht). | SWO-449063 | No fault found | No fault found | No fault found | SC |  | 12.11.2018 13:10O.M. Laboratory 020189783020: Despite maintenance measures, the device constantly reports an error message | calibration error on various parameters. By time of arrival, the device had already been for about 1.5 hours. green. The customer had previously exchanged before the solution pack and sensor cartridge. After the usual run-in phase (e.g. up to 4 hrs for Lac), the calibrations and QCs were error-free. | No action required on the device. Complaint of defective consumables. Following a software update (see separate report).
No | Visit | Complaint | RLTD | Customer having numerous software issues, including failing to print, laggy software and freezes. | Customer still having poor analyser software performance. | Again tried to reload windows and 3.4, still having errors, further investigation required. | SWO-449657 | Revisit required | SW | RMED SW | Freeze |  | Customer having numerous software issues, including failing to print, laggy software and freezes. | Customer still having poor analyser software performance. | Again tried to reload windows and 3.4, still having errors, further investigation required.
No | Visit | Complaint | RKOR | pCO2 QC error | . | SC SP clean. Liquid flow line clean. SP replace SC replace. Cal QC check. | SWO-452353 | Assisted Replacement | SC | SC | pCO2 | QC | pCO2 QC error | . | SC SP clean. Liquid flow line clean. SP replace SC replace. Cal QC check.
No | Visit | Complaint | RCHN | 试管夹断裂 | 试管夹断裂 | 更换备用903-584*2后修复by tim | SWO-448815 | Defect | HW | Aspiration Module | Tube holder |  | Test tube clip break | Test tube clip break | Repair by tim after replacing spare 903-584 x 2
No | Visit | Complaint | RCHN | 屏幕开裂 | 屏幕开裂 | 更换备用的910-324后修复by tim | SWO-448819 | Defect | HW | Analyzer Control Module | display | Potentially User Error | Screen cracking | Screen cracking | Repair by tim after replacing spare 910-324
No | Visit | Complaint | RCHN | 定标底物外漏，1024 | 定标底物外漏，1024 借903-364一个 |  | SWO-448617 | Defect | HW | Process Module | Optical Unit | 1024 | Calibration substrate leakage, 1024 | Bid substrate leakage, 1024 borrow 903-364 one
No | Visit | Complaint | RFRA | Erreur configuration ticket | Configuration ticket | Teste de la configuration avec le médecin ok. | SWO-448936 | [5 empty] | Ticket configuration error | Configuration ticket | Test the setup with the doctor ok.
No | Visit | Complaint | RDEU | 05.11.2018, Herr Dietrich, Tel: 015256714506: License Key abgelaufen- Gerät ohne Funktion | Fehlersuche, License Key missing wird angezeigt - Software installiert, erneut versucht Keys zu installieren- ohne Erfolg, Basic Key entnommen und überprüft, erneut gestartet - keine Identifikation mehr auf dem Key | Basic Key defekt (367-683), AU (903-656) und Basic Key - bestellt | SWO-449857 | Revisit required | SW | RMED SW | Basic Key |  | 05.11.2018, Mr. Dietrich, Tel: 015256714506: License Key expired- Device without function | Troubleshooting, License Key missing is displayed - Software installed, retries to install keys - without success, Basic Key taken and checked, restarted - no more identification on the key | Basic Key defective (367-683), AU (903-656) and Basic Key - ordered
No | Visit | Complaint | RDEU | 12.11.2018, 11:05 Uhr, Fr. Mengesdorf, Tel. 062150335589: pO2 und pCO2 Empf. zu niedrig | Bei Eintreffen pCO2 Empfindlichkeitswert 83% Satus 30mmHg und pO2 QC Verletzung | Allgemeine Überprüfung - Reinigung Kontakte und Neu-Membranierung - Erneuerung pCO2 Elektrode (Rücktausch da keine Fehlerbeseitigung) - Gasüberprüfung im Service - Demontage Messventil Gaszufuhr - Erneuerung Ventilmembran (verschlissen) - Verschlauchung an pH $ BG Modul teilweise erneuert - Messkammern überprüft und gereinigt - Membranierungen - mehrfache Kalibrierung - Funktionsüberprüfung etc. KAL und QC Fehler bleiben bestehen ! Eventuelle Teilebestellung notwendig - Klärung folgt Labor- Ersatzanalysator aufgestellt Stationsleitung informiert | SWO-449187 | Revisit required | HW | pH/BG Module | Valve, air |  | 12.11.2018, 11:05 a.m., Ms. Mengesdorf, tel. 062150335589: pO2 and pCO2 Empf. too low | On arrival pCO2 sensitivity value 83% satus 30mmHg and pO2 QC violation | General inspection - Cleaning contacts and re-diaphragming - Renewal pCO2 electrode (replacement since no error removal) - Gas check in service - Disassembly of measuring valve gas supply - Renewal valve membrane (worn) - Hose on pH . BG module partially renewed - measuring chambers checked and cleaned - membraneings - multiple calibration - function check, etc. KAL and QC errors remain ! Possible part order necessary - clarification follows Laboratory replacement analyzer set up Station management informed
No | Visit | Complaint | RIND | Thb/so2 problem for patient results | Thb $so2 Problem for patient results | checked the analyser it was givien for few samples error 580 , oxymetry error so done complete oxymetry test and after that hemolyser on and off done for 2 to 3sec and after that calibrated the analyser to ready mode analysed the sample | SWO-452221 | Adjustment | HW | Oxi module | Oxi module | 580 | Thb/so2 problem for patient results | Thb /so2 Problem for patient results | checked the analyser it was givien for few samples error 580 , oxymetry error so done complete oxymetry test and after that hemolyser on and off done for 2 to 3sec and after that calibrated the analyser to ready mode analysed the sample
No | Visit | Complaint | RIND | Analyzer showing neon number too high. | Analyzer showing neon number too high error. | checked the analyzer and performed complete oximetry test its failed lamp unit neon position error its showing2.375 I replaced lamp unit and oxi control pcb still its showing same error. I need power supply to solve this issue. | SWO-449402 | Revisit required | HW | Oxi module | Oxi module |  | Analyzer showing neon number too high. | Analyzer showing neon number too high error. | checked the analyzer and performed complete oximetry test its failed lamp unit neon position error its showing2.375 I replaced lamp unit and oxi control pcb still its showing same error. I need power supply to solve this issue.
No | Visit | Complaint | RKK | インレットドアの動作が停止しサンプル取出しができなくなる。 |  |  | SWO-449350 | Adjustment | HW | Cover and Chassis | Door, Inlet |  | Operation of the inlet door is stopped and sample take-out can not be.
No | Visit | Complaint | RIND | PCT Test cartridge barcode not excepting | Pct test cartridge | Clean the storage wheel, clean the barcode sensor glass. Restart the analayzer. | SWO-452315 | Cleaned | HW | Storage Module | Barcode reader, storage wheel |  | PCT Test cartridge barcode not excepting | Pct test cartridge | Clean the storage wheel, clean the barcode sensor glass. Restart the analayzer.
No | Visit | Complaint | RFRA | Contrôles c en échec sur tout les paramètres | Erreur sur le contrôles c | Changement du pack , suppression des paramètres non utilisé pour l'utilisation , passage calibration et cq | SWO-450375 | Assisted Replacement | SP | SP | SP | QC C | C controls in check across all settings | Error on controls c | Changing the pack, removing settings not used for use, passing calibration and cq
No | Visit | Complaint | RKOR | QC pCO2 level1 is out of range | Analyzer Status check Calibration, Quality Control, Replacement, Events logs check Fiuld transport check Download and delete the data logs QC plots check Waste pump, manifold check Peform the zero check | Peform the zero check | SWO-452382 | [5 empty] | QC pCO2 level1 is out of range | Analyzer Status check Calibration, Quality Control, Replacement, Events logs check Fiuld transport check Download and delete the data logs QC plots check Waste pump, manifold check Peform the zero check | Peform the zero check
No | Visit | Complaint | RKOR | Multiple cla error | . | SC clean. Liquid flow line clean. SC replace. | SWO-452354 | [5 empty] | Multiple cla error | . | SC clean. Liquid flow line clean. SC replace.
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 漏电流 0963 | 更换试剂包后修复 | SWO-449486 | Assisted Replacement | SP | SP | Leak current | 963 | Test card, reagent package failure | Drain current 0963 | Repair after replacement of reagent kit
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | CO2定标不过 0378 | 更换测试卡后修复 | SWO-449453 | Assisted Replacement | SC | SC | pCO2 | 378 | Test card, reagent package failure | CO2 calibration is no more than 0378 | Repair after replacing test card
No | Visit | Complaint | RCHN | 光单强度超范围 | 光单强度超范围，借903-364一个 | 系统检查错误，测试没有问题，不用更换By tim | SWO-449386 | Defect | SW | RMED SW | System Check |  | Light single intensity over range | Light single intensity over range, borrow903-364 one | System check error, test no problem, do not replace By tim
No | Visit | Complaint | RCHN | 结果报1070，气流温度超范围 | 结果报1070，气流温度超范围 | 调整PCB板后修复by tim | SWO-449550 | Adjustment | HW | Process Module | PCB, cup dryer | 1070 | Results reported to 1070, airflow temperature over range | Results reported to 1070, airflow temperature over range | Fix by tim after adjusting PCB board
No | Visit | Complaint | RCHN | 检测结果与检验科有差异 | D二聚体检测结果与检验科有差异 | 重新做质控，都在可控范围内by tim | SWO-450207 | No fault found | No fault found | No fault found | No fault found | D-Dimer discr. Result; Techline 827968285 | There is a difference between the test results and the inspection department. | There is a difference between ddipolymer test results and inspection department | Re-do quality control, all within control by tim
No | Visit | Complaint | RDEU | 12.11.2018, 11:05 Uhr, Fr. Mengesdorf, Tel. 062150335589: pO2 und pCO2 Empf. zu niedrig | KAL und QC Fehler pCO2 QC Fehler pO2 | Allgemeine Überprüfung - Verschlauchung teilweise erneuert - Komplettes pH $BG Modul aus 754R2493R001 demontiert und im Analysator montiert - mehrfach KAL und QC durchgeführt (keine Fehlerbeseitigung) - Rücktausch pH $BG und Rückmontage etc. - Gasmodul teilweise demontiert und gereinigt (Druckminderer) etc. - mehrfach KAL durchgeführt - Empfindlichkeitswert pCO2 bei 87% - QC pCO2 in Level 1 und 2 zu tief (RiliBÄK Verletzung) pO2 in KAL und QC ohne Beanstandung - neue pCO2 zu Testzwecken installiert ! Bewertung morgen ! Labor- Ersatzanalysator bleibt aufgestellt und Ersatzteile per Express bestellt ! | SWO-449854 | Revisit required | HW | pH/BG Module | Valve, air |  | 12.11.2018, 11:05 a.m., Ms. Mengesdorf, tel. 062150335589: pO2 and pCO2 Empf. too low | KAL und QC Fehler pCO2 QC Fehler pO2 | General inspection - hose partially renewed - Complete pH $BG module from 754R2493R001 dismantled and mounted in the analyzer - carried out several times KAL and QC (no troubleshooting) - replacement pH $BG and reassembly etc. - gas module partially dismantled and gas module partially dismantled and cleaned (pressure reducer) etc. - performed multiple KAL - sensitivity value pCO2 at 87% - QC pCO2 in level 1 and 2 too deep (RiliBÄK violation) pO2 in KAL and QC without complaint - new pCO2 installed for testing purposes ! Review tomorrow ! Laboratory replacement analyzer remains set up and spare parts ordered by express !
No | Visit | Complaint | RDEU | 13.11.2018, 08:41 Uhr, Herr Nauheimer, Tel. 0611577375: Festplatte defekt; siehe beigefügter Auftrag | Netzteil defekt | Gerät ausgebaut, Netzteil getauscht. Startphase i.O, Kalibriert, QCs i.O | SWO-449688 | Defect | HW | DMS_PC | Power supply |  | 13.11.2018, 08:41, Mr. Nauheimer, tel. 0611577375: hard drive defective; see attached order | Power supply defective | Device removed, power supply replaced. Start-up phase i.O, Calibrated, QCs i.O
No | Visit | Complaint | RIND | Setsafe_22 | Not Operational and inlet door not working properly. | Clean the sensor of inlet door and run open close program in service option inlet door. But still inlet door is stacked and not closed properly. Change the both inlet door and check the in service option inlet door. Now inlet door is working properly. Machine is working fine and ready to use. | SWO-451007 | Defect | HW | Cover and Chassis | Door, Inlet | setsafe_22 | Setsafe_22 | Not Operational and inlet door not working properly. | Clean the sensor of inlet door and run open close program in service option inlet door. But still inlet door is stacked and not closed properly. Change the both inlet door and check the in service option inlet door. Now inlet door is working properly. Machine is working fine and ready to use.
No | Visit | Complaint | RDEU | Oximetrie-Hardwareproblem | Erneut Oxifehler. Erster Kalfehler 14.11. 08:42 | Zur weiteren Überprüfung WSC-PCB aus Ersatzgerät eingesetzt. Nach Neustart alle Parameter betriebsbereit. Oxi-Completetest in Ordnung. | SWO-450868 | Defect | HW | Basic Unit | PCB, Wet section |  | Oximetrie-Hardwareproblem | Again Oxi error. First calf error 14.11. 08:42 | For further verification, WSC-PCB from replacement device is used. After rebooting all parameters ready for operation. Oxi complete test in order.
Yes | Visit | Complaint | RCHN | 未连接分析仪 | 未连接分析仪，借903-381、933-441、905-914各一个 | 更换903-381、933-441、905-914后修复by tim | SWO-448689 | [5 empty] | The analyzer is not connected | Unconnected analyzer, one each for 903-381, 933-441, 905-914 | Repair by tim after replacing 903-381, 933-441, 905-914
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 不均匀的冲洗液 1081,1324 | 连续更换2个试剂包后修复 | SWO-450295 | Assisted Replacement | SP | SP | FT | 1081, 1324 | Test card, reagent package failure | Uneven flushing fluid 1081,1324 | Repair after 2 reagent kits in a row
No | Visit | Complaint | RCHN | 报错3412 | 报错3412 借902-977一个 | 更换902-977后修复by tim | SWO-449452 | Defect | HW | Wet section | Valve | 3412 | Error 3412 | Error 3412 Borrowing 902-977 one | Repair by tim after replacement 902-977
Yes | Visit | Complaint | RCHN | 报错1024 | 报错1024 借903-364 944-212（14995缺货）944-220（15003）944-258（14998缺货）各一个 | 更换903-364，重新定标后修复by Tim | SWO-449368 | [5 empty] | 1024 error | Error 1024 Borrowing 903-364 944-212 (14995 out of stock) 944-220 (15003) 944-258 (14998 out of stock) one each | Replace 903-364, fix by Tim after recalibration
No | Visit | Complaint | RIND | pH & PCO2 error in calibration | Interface board defective. | Replaced Interface board(903-623). Performed zeroing. 2 point calibration performed successfully. | SWO-451033 | [5 empty] | pH & PCO2 error in calibration | Interface board defective. | Replaced Interface board(903-623). Performed zeroing. 2 point calibration performed successfully.
No | 21 | Y | RCZE | Blank |  |  | D53B133F1031744D | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | Blank
No | Visit | Complaint | RAME | Crea failing calibration. | Reviewed data logs, performed applicable service diagnostics, no errors observed. Issue began upon remembraning CREA during regular maintenance. Sens 3 below specifications. Cleaned electrode connectors, measuring chamber, remembraned CREA A & B, sens 3 below specifications @ 43. Noted sens 3 had been marginal for several weeks. Suspect possible weak $worn electrodes. | Replaced CREA A & B electrodes with electrodes from ABL#3 (R2503N007), CREA is not used, parameter is disable, electrodes still in place. Rinsed clean and installed fresh membranes. Electrodes successfully calibrated, sens 3 @ 104%. | SWO-450188 | Assisted Replacement | Electrode | Electrode | Crea | Sens 3 | Crea failing calibration. | Reviewed data logs, performed applicable service diagnostics, no errors observed. Issue began upon remembraning CREA during regular maintenance. Sens 3 below specifications. Cleaned electrode connectors, measuring chamber, remembraned CREA A & B, sens 3 below specifications @ 43. Noted sens 3 had been marginal for several weeks. Suspect possible weak /worn electrodes. | Replaced CREA A & B electrodes with electrodes from ABL#3 (R2503N007), CREA is not used, parameter is disable, electrodes still in place. Rinsed clean and installed fresh membranes. Electrodes successfully calibrated, sens 3 @ 104%.
No | Visit | Complaint | RDEU | 12.11.2018, 11:05 Uhr, Fr. Mengesdorf, Tel. 062150335589: pO2 und pCO2 Empf. zu niedrig | Bei Eintreffen mit bestellten Teilen : QC- Messungen mit Meldung 0664 Probenproblem 0653 Probe bei Ansaugung vorzeitig im oberen pH $BG LS erkannt Empfindlichkeitswert pCO2 87% QC Fehler pCO2 und pO2 | Fehlerdiagnose - Test Inletmodul installiert (Rückmontage) - sämtliche Ventile demontiert, gereinigt und überprüft - Gas1 Ventil montiert (Rückmontage) - Analyzer Control PCB erneuert (Rückmontage $ DMS Not connected with AU) - Pressure transducer Unit installiert - mehrfache Kalibration durchgeführt - Funktionsüberprüfung etc. ohne Beanstandung Labor Austauschanalysator zurück Labor ! | SWO-450603 | Revisit required | HW | pH/BG Module | Valve, air |  | 12.11.2018, 11:05 a.m., Ms. Mengesdorf, tel. 062150335589: pO2 and pCO2 Empf. too low | Upon arrival with ordered parts: QC measurements with message 0664 Sample problem 0653 Sample detected prematurely in upper pH /BG LS during suction Sensitivity value pCO2 87% QC error pCO2 and pO2 | Fault diagnosis - Test inlet module installed (reassembly) - all valves disassembled, cleaned and checked - Gas1 valve mounted (reassembly) - Analyzer Control PCB renewed (reassembly - DMS Not connected with AU) - Pressure transducer unit installed - multiple Calibration carried out - function check etc. without complaint Laboratory replacement analyzer back laboratory !
No | Visit | Complaint | RDEU | 24.09 2018 14:00 Frau Lam 0684111622792 immer wieder Kal Fehler Lyte und pH ( Referenzelekrtode ) Frau Lam hätte gerne einen Rückruf und Einsatz weitere Geräte mit Problemen Einlasnadel etc. | PH und E $Lyte status fehler oder drift. | Ref. Elektrode ausgetauscht. | SWO-450413 | Assisted Replacement | Electrode | Electrode | Reference | CAPA-733 | 24.09 2018 14:00 Mrs. Lam 0684111622792 again and again Kal Error Lyte and pH ( Reference Elekrtode ) Ms. Lam would like to have a recall and use of other devices with problems lasing needle etc. | PH und E /Lyte status fehler oder drift. | Ref. electrode.
No | Visit | Complaint | RIND | Gasket leaking | Gasket leaking | cleaned the gasket and fixed it but leakage was there so there was block in the gasket tube that also cleaned then also problem not solved so while inspecting junction tube was found defective so changed the same and checked now leakage stopped working fine | SWO-452226 | Defect | HW | Inlet module | Inlet gasket |  | Gasket leaking | Gasket leaking | cleaned the gasket and fixed it but leakage was there so there was block in the gasket tube that also cleaned then also problem not solved so while inspecting junction tube was found defective so changed the same and checked now leakage stopped working fine
No | 3 | Y | RCZE | Blank |  |  | D53B1938100338FF | Adjustment | HW | Oxi module | Lamp unit | halogen | Blank
No | Visit | Complaint | RPOL | Głośne dźwięki | Sprawdzenie wentylatorów. | Wymiana wentylatora w module AutoCheck. | SWO-451613 | Defect | HW | Autocheck module | Fan |  | Loud sounds | Check the fans. | AutoCheck fan replacement.
No | Visit | Complaint | RESP | Les ha llegado un SC por errores en lactato y ahora no inicia dando error 1066 de temperatura de sustrato | Chequeo en servicio de equipo Chequeo de circuito neumático y accesorios Chequeo de consumibles | Sustitución de inlet Gasket, inlet probe y connector Sustitución de SC y SP Limpieza de circuito neumático Ejecución de programas de servicio | SWO-455254 | Assisted Replacement | SC | SC | Temp | 1066 | They got an SC for lactate errors and now it doesn't start by error1066 of substrate temperature | Team service check Pneumatic circuit check and accessories Check supplies | Substitution of gasket, inlet probe and connector Replacing SC and SP Pneumatic circuit cleaning Running service programs
Yes | Visit | Complaint | RESP | Cliente indica problemas de Cal. | Equipo da error en Cal. | Se indica a cliente que verifique registros. Se entra en servicio y se comprueba módulo OXi dando errores de Cavitation absorbance . Se da indicaciones al cliente para que reinicie el equipo. Se soluciona el problema con el reincio. | SWO-452133 | Re classified RMED | Phone support |  |  | see mail with SWO number reference | Client indicates Cal problems. | Team gives error in Cal. | The client is instructed to verify logs. It enters service and checks OXi module giving Cavitation absorbance errors. Prompts the client to restart the computer. Fixed the problem with reins.
No | Visit | Complaint | RDEU | 14.11.2018/17:00 Sr. Bärbel 0354629258 -> Spülung des ABL nicht möglich 14.11.2018 18:00 Sr. Bärbel 0354629258 Pumpenkal-Fehler | Aufforderung LP Wechsel aufgrund keine Lösungen gefördert, PcO2 Empfindlichkeitsfehler zu niedrig bei 38,2%, in der Kal | Serviceprogramme Fluid geprüft, alles i.O., Wechsel der SC (LotR0820N043) aus Kundenmaterial, da Fehler seit SC-Wechsel, PCO2 Empfindlichkeit in der Kal weiterhin sehr niedrig bei 14,2%, QC bei allen Parametern im Bereich, Abbruch der Arbeiten, Kunden informiert Datensicherung | SWO-451308 | Assisted Replacement | SC | SC | pCO2 | Sens | 14.11.2018/17:00 Sr. Bärbel 0354629258 -> Flushing of the ABL not possible 14.11.2018 18:00 Sr. Bärbel 0354629258 Pump scale error | Prompt LP change due to no solutions promoted, PcO2 sensitivity error too low at 38.2%, in the cal | Service programs Fluid checked, all in the o.o., change of SC (LotR0820N043) from customer material, since errors since SC change, PCO2 sensitivity in the Cal continues to be very low at 14.2%, QC at all parameters in the range, cancellation of work, customer informed Backup
No | Visit | Complaint | RDEU | Die Kundin beklagt mit dem neuen Gerät einen stark erhöhten Verbrauch an LP und SC 12.11.2018 09:30Uhr Frau Konrad 03832652416 | Werte fluidic transport ok, Ventil geprüft | Anschlussdichtung (Kundenmaterial) u Einlassdichtung erneuert, Kal durchgeführt | SWO-451323 | Defect | HW | Inlet module | Connector gasket |  | With the new device, the customer bemoans a greatly increased consumption of LP and SC 12/11/2018 09:30Pm Mrs Konrad 03832652416 | Values fluidic transport ok, valve tested | Connection seal (customer material) u inlet seal renewed, Kal carried out
No | Visit | Complaint | RDEU | Einlassdichtung rutscht nicht auf den Solutionpack-Konus nach einer Messung | Funktionskontrolle | Einlassmodul ersetzt, Funktionskontrolle Erfolgreich, Solutionpack ersetzt wg pCO2 Fehler Kalibration, QC-Kontrollen | SWO-451260 | Defect | HW | Inlet module | Inlet module |  | Inlet seal does not slide onto the solutionpack cone after a measurement | Funktionskontrolle | Inlet module replaced, function control Successful, Solutionpack replaces wg pCO2 error calibration, QC controls
No | Visit | Complaint | RLTD | Please call Anna 07984 464577 | Screen frozen saying intervention required no buttons on screen. | Analyser has already had PC replaced for same fault. When CPU disconnected from wet section the PC screen then comes active and keys are available and system message wet section not connected. Reconnect wet section to pc and consumable can be changed and all keys available. Analyser screen message SC needs replacement no more test. Replaced wet section board reinstalled windows and analyser software. Replaced SC and after start and wet section Oxi initialisations run test and Thb cal all passed ok. Part 903-644 not on CRM wet section control pcb. | SWO-451336 | Defect | HW | DMS_PC | PCB, wet section |  | Please call Anna 07984 464577 | Screen frozen saying intervention required no buttons on screen. | Analyser has already had PC replaced for same fault. When CPU disconnected from wet section the PC screen then comes active and keys are available and system message wet section not connected. Reconnect wet section to pc and consumable can be changed and all keys available. Analyser screen message SC needs replacement no more test. Replaced wet section board reinstalled windows and analyser software. Replaced SC and after start and wet section Oxi initialisations run test and Thb cal all passed ok. Part 903-644 not on CRM wet section control pcb.
No | Visit | Complaint | RFRA | Erreur calibration et QC, THB à zéro en résultat cal | Vérification complète | Remplacement de l'hemoliseur $ $ relance $ $ mise en route $ $ tests en mode service $ $ calibration thb $ $ passage des contrôles $ $ essais bons | SWO-452866 | Defect | HW | Oxi module | Hemolyzer |  | Error calibration and QC, THB to zero in cal result | Full check | Replacement of the hemoliser $$raise $$ start $$ tests in service mode $$ calibration thb $$ passage of controls $$ tests good
No | Visit | Complaint | RLTD | Connie Chilcott texted to ask me to contact her. | Blue screen with STOP followed by a series of numbers on the display, the analyser has been rebooted and it comes back with the same. | The CPU unit was replaced with a new one and the CF card from the old CPU fitted in to the new one. When the analyser was powered up it booted up ok and went into its start-up and completed it without any problems. The analyser hasn't connected to AQURE because the MAC address has changed, I left the new MAC address with Connie so she could give it to IT. | SWO-451285 | Defect | HW | DMS_PC | CPU |  | Connie Chilcott texted to ask me to contact her. | Blue screen with STOP followed by a series of numbers on the display, the analyser has been rebooted and it comes back with the same. | The CPU unit was replaced with a new one and the CF card from the old CPU fitted in to the new one. When the analyser was powered up it booted up ok and went into its start-up and completed it without any problems. The analyser hasn't connected to AQURE because the MAC address has changed, I left the new MAC address with Connie so she could give it to IT.
No | Visit | Complaint | RFRA | Le client a éteint l'appareil depuis l'interrupteur à l'arrière, lors de la remise en route, celui-ci affiche disk boot failure. Mme DELERTE - 04.50.83.22.00 | Disque boot faillée | Changement de la carte flash promet rage complet automate, test connections ok | SWO-451252 | Defect | HW | DMS_PC | CF card |  | The customer turned off the device from the switch at the rear, when it was restarted, it displays disk boot failure. Mrs DELERTE - 04.50.83.22.00 | Failed boot disc | Changing the flash card promises complete automaton rage, test connections ok
No | Visit | Complaint | RLTD | Range correction for crea | Site visit to change reportable ranges for crea. Customer had seen results not given as it was below 35. Analyser spec shows that it can be measured down to 10. | Arrived on site and attended analyser. Investigation showed that the default reportable range for crea is 35 - 900 Custimer said that they would like to measure down to 10 as this can be common and would also like to be able to report higher results than the default 900. Adjusted range to 10-1800 as per analyser measurement spec. | SWO-451092 | Setup | SW | RMED SW | Reportable range | Crea | Range correction for crea | Site visit to change reportable ranges for crea. Customer had seen results not given as it was below 35. Analyser spec shows that it can be measured down to 10. | Arrived on site and attended analyser. Investigation showed that the default reportable range for crea is 35 - 900 Custimer said that they would like to measure down to 10 as this can be common and would also like to be able to report higher results than the default 900. Adjusted range to 10-1800 as per analyser measurement spec.
No | Visit | Complaint | RLTD | Please call Frederick on 0208 401 3599 bleep 213 | Tested sample pathway for blockages by flushing with water, sample pathway found to be clear. Checked inlet probe, in good condition and straight. Tested fluid transport in service program, all tests passed. Restarted analyser, restarted ok except QCs failing on some oximetry parameters. Flushed oximetry module with hypochlorite solution to clean. Performed tHb calibration and repeated QCs. QCs now passing on all parameters. Analyser ready for use. | Cleaned oximetry module with hypochlorite solution. | SWO-451312 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | Please call Frederick on 0208 401 3599 bleep 213 | Tested sample pathway for blockages by flushing with water, sample pathway found to be clear. Checked inlet probe, in good condition and straight. Tested fluid transport in service program, all tests passed. Restarted analyser, restarted ok except QCs failing on some oximetry parameters. Flushed oximetry module with hypochlorite solution to clean. Performed tHb calibration and repeated QCs. QCs now passing on all parameters. Analyser ready for use. | Cleaned oximetry module with hypochlorite solution.
No | Visit | Complaint | RCHN | 无法开机 | 无法开机 借903-381 933-441 905-914 903-001各一个 | 更换903-381 933-441 905-914 903-001后修复by tim | SWO-450323 | [5 empty] | Cannot power on | Unturned Borrowing 903-381 933-441 905-914 903-001 each | Repair by tim after replacing 903-381 933-441 905-914 903-001
No | Visit | Complaint | RCHN | 无法识别试剂包 | 无法识别试剂包 借902-975一个 | 更换902-975后修复by tim | SWO-450326 | [5 empty] | Unable to identify the kit | Unrecognizable reagent package Borrows 902-975 one | Repair by tim after replacement 902-975
No | Visit | Complaint | RCHN | 电池没有蓄电能力 | 电池没有蓄电能力 借431-018C一个 | 更换431-018C后修复by tim | SWO-450949 | [5 empty] | The battery doesn't have the power to store electricity. | Battery does not have the capacity to store electricity, borrow 431-018C one | Repair by tim after replacement 431-018C
No | Visit | Complaint | RCHN | O2测得的值不准确 | O2测得的值不准确，借903-581一个 | 更换903-581后修复by tim | SWO-449369 | Defect | HW | Basic Unit | PCB, Sensor interface | pO2 | The value measured by O2 is inaccurate | O2 measured the value is not accurate, borrow 903-581 a | Repair by tim after replacement 903-581
No | Visit | Complaint | RDEU | 05.11.2018, Herr Dietrich, Tel: 015256714506: License Key abgelaufen- Gerät ohne Funktion 12.11.18: während Einsatz ermittelt das hier nicht Keys abgelaufen sind, sondern der Basic Key defekt ist „License Key missing“ Auftragsnummer: | Software key def. | Gerät aus; Service key (754R9000N0471) eingebaut; Gerät ein; POC-IT informiert; Startphase abgewartet; Kal. = pCO2-Fehler; pCO2 neu membraniert; Kal. + QC i.O. | SWO-451427 | Defect | SW | RMED SW | Basic Key |  | 05.11.2018, Mr. Dietrich, Tel: 015256714506: License Key expired- Device without function 12.11.18: while using it determines that no keys have expired here, but the basic key is defective "License Key missing" Order: | Software key def. | device off; Service key (754R9000N0471) built-in; device; POC-IT; start-up phase awaited; Kal. = pCO2 error; pCO2 re-membraneed; Kal. + QC in the first place
Yes | Visit | Complaint | RAME | Intermittent inhomogoneous mosamples and drifting on calibrations. | Maintenance assist with Charles Reisinger | Assisted Charles Reisinger | SWO-451414 | Re classified RMED | same day upSWO |  |  |  | Intermittent inhomogoneous mosamples and drifting on calibrations. | Maintenance assist with Charles Reisinger | Assisted Charles Reisinger
No | Visit | Complaint | RAME | Intermittent inhomogoneous mosamples and drifting on calibrations. | Did see air bubbles entering with samples and QCs. Not from syringes. | Replace all components in inlet pathway. Tested multiple samples and Did not detect any air bubbles. | SWO-451374 | Defect | HW | Inlet module | Inlet module |  | Intermittent inhomogoneous mosamples and drifting on calibrations. | Did see air bubbles entering with samples and QCs. Not from syringes. | Replace all components in inlet pathway. Tested multiple samples and Did not detect any air bubbles.
No | Visit | Complaint | RPAC | Correlations between Lab ABL, different to Mercy ABL ICU on the Na and Ca parameters. | The Ca and Na parameters, when correlated to the ABL8XX in the Lab, are consistantly the same. | The analyser was placed into the Hold mode and the El $Met module had teh Electrode Chamber replaced, as a pre-caution. The Ca and Na electrodes were also replaced. New membranes were placed onto these electrodes. The analyser was then allowed to re-calibrate all the parameters and once they had all stabilized, QC Levels 1 & 3 were performed. All values are within their limits. The Pathology will then perform more correlations with the Lab analyser in the coming weeks. | SWO-450834 | No fault found | No fault found | No fault found | No fault found | upSWO? mail: How much bias? | Correlations between Lab ABL, different to Mercy ABL ICU on the Na and Ca parameters. | The Ca and Na parameters, when correlated to the ABL8XX in the Lab, are consistantly the same. | The analyser was placed into the Hold mode and the El $Met module had teh Electrode Chamber replaced, as a pre-caution. The Ca and Na electrodes were also replaced. New membranes were placed onto these electrodes. The analyser was then allowed to re-calibrate all the parameters and once they had all stabilized, QC Levels 1 & 3 were performed. All values are within their limits. The Pathology will then perform more correlations with the Lab analyser in the coming weeks.
No | Visit | Complaint | RIND | Rinse error. | Check the system. | Check the LS & Valve. Check the inlet section. calibrate the system & calibration are in ranges. | SWO-452339 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | Rinse error. | Check the system. | Check the LS & Valve. Check the inlet section. calibrate the system & calibration are in ranges.
No | Visit | Complaint | RSCH | Blank | Erreur Créa 0953: Valeur de CQ mesurée en dessus de la gamme de contrôle (Créa A) pour level 1-3-4 malgré le remplacement 2X membrane Créa A-B et 2X Cal 1-2 Garantie | Calibration des préamplificateur Crea A-B e.o Remplacé les électrodes Créa A (lot BQ-03) et Créa B (lot BQ-03) e.o Garantie Calibration e.o contrôle de qualité e.o | SWO-451292 | Assisted Replacement | Electrode | Electrode | Crea | QC | Blank | Crea Error 0953: CQ value measured above the control range (Crea A) for level 1-3-4 despite replacing 2X membrane Crea A-B and 2X Cal 1-2 Guarantee | Crea A-B e.o preamplifier calibration Replaced the electrodes Créa A (lot BQ-03) and Créa B (lot BQ-03) e.o Guarantee Calibration e.o quality control e.o
No | Visit | Complaint | RDEU | Festeingetragene IP Adresse wird nach Geräteneustart nicht behalten | IP Adresse wird im Control Panel nach Geräteneustart gelöscht | Überprüfung Festplatte => Typ mSATA 3ME3 ATA Device => Flash-Speicher File Based Write Filter deaktiviert => Geräteneustart => IP Adresse 172.16.2.127 $255.255.0.0 $172.16.254.254 eingetragen und File Based Write Filter aktiviert => erneuter Geräteneustart => Netzwerkdaten werden behalten Gerät auf ITS aufgestellt => Netzwerkverbindung getestet Testmessung wird ordnungsgemäß an KIS $LIS weitergeleitet KIS $LIS Verbindung: Medidata; 172.16.2.124; Port 2001 | SWO-451102 | Defect | SW | RMED SW | IP address |  | Permanently registered IP address will not be retained after device restart | IP address is deleted in the Control Panel after device restart | Checking hard drive => Type mSATA 3ME3 ATA Device => Flash memory File Based Write Filter Disabled => Device Restart => IP Address 172.16.2.127 .255.255.0.0 .172.16.254.254 enabled File Based Write Filter => Restarting The Device => Retaining Network Data Device set up on ITS => Network connection tested Test measurement is properly forwarded to KIS $LIS KIS $LIS connection: Medidata; 172.16.2.124; Port 2001
No | Visit | Complaint | RFRA | Fuite sous l'AQT | Tuyau de la pompe P4 débranché. Circuit P4 bouché. | Nettoyage, débouchage du circuit P4. Contrôle du circuit P2, ok. Passage tube de nettoyage (x2). | SWO-450977 | Cleaned | HW | Wet section | Tubing | Waste P4 | Leak under the AQT | P4 pump hose unplugged. P4 circuit clogged. | Cleaning, unblocking the P4 circuit. Control of the P2 circuit, okay. Passage cleaning tube (x2).
No | Visit | Complaint | RKOR | LOCKINLET_24, MIXWHEEL_24 | Reinsert the Inlet wheel, Cut the cable tie at cable of pipette unit. | Reinsert the Inlet wheel, Cut the cable tie at cable of pipette unit. | SWO-452389 | Adjustment | HW | Aspiration Module | Inlet Wheel | Error 24 | LOCKINLET_24, MIXWHEEL_24 | Reinsert the Inlet wheel, Cut the cable tie at cable of pipette unit. | Reinsert the Inlet wheel, Cut the cable tie at cable of pipette unit.
No | Visit | Complaint | RDEU | 14.11.2018 11:00 Fr. Soll 02151363000: Probenfehler TN-T Temperatur | Sporadischer Temperatur Fehler bei Patientenmessungen „1068 Temperatur außer Bereich bei Trocknungsschritt“. | Find target temperature durchgeführt, Uhrzeit eingestellt, Systemüberprüfung durchgeführt, Systemreinigung durchgeführt, CRP und NT-proBNP QC durchgeführt. D-Dimer nur noch 1 Test vorhanden - neue Charge muss kalibriert werden. CRP QC Lvl 1 wird zu hoch und NT-proBNP Lvl 2 zu niedrig gemessen - Messung mit neuer Probe wiederholt. - ok. | SWO-452305 | Adjustment | HW | Process Module | Cup dryer | 1068 | 14.11.2018 11:00 Fr. Should 02151363000: Sample error TN-T Temperature | Sporadic temperature error in patient measurements "1068 temperature out of range at drying step". | Find target temperature performed, time set, system check performed, system cleaning performed, CRP and NT-proBNP QC performed. D-Dimer only 1 test available - new batch must be calibrated. CRP QC Lvl 1 is measured too high and NT-proBNP Lvl 2 too low - measurement repeated with new sample. - ok.
No | Visit | Complaint | RPOL | Błąd kalibracji pH oraz elektrolitów | Testy | Wymiana elektrody referencyjnej. | SWO-451615 | Assisted Replacement | Electrode | Electrode | Reference | CAPA-733 | PH and Electrolyte calibration error | Tests | Reference electrode replacement.
No | Visit | Complaint | RDEU | 16.11.2018 Frau Lauber 08964246146 meldet einen Fehler Data upgrade faild auf schwarzem Bildschirmhintergrund | Data upgrade failed - status code: 2 | Software-Version 3.4 installiert danach gleicher Fehler Software im Rahmen der FAN 915-384 erneut (clean) installiert siehe SWO-451890 | SWO-451889 | Defect | SW | RMED SW | Database | status code 2 | 16.11.2018 Ms. Lauber 08964246146 reports an error Data upgrade faild on black screen background | Data upgrade failed - status code: 2 | Software version 3.4 installed then the same error Software reinstalled as part of FAN 915-384 (clean) see SWO-451890
No | Visit | Complaint | RAME | Battery is no longer charging | Battery will not charge | Replaced battery assembly. Ran Analyzer startup and Calibration. | SWO-452241 | Defect | HW | Battery module | Battery module |  | Battery is no longer charging | Battery will not charge | Replaced battery assembly. Ran Analyzer startup and Calibration.
No | Visit | Complaint | RFRA | Appareil en ecran noir suite démarrage application -Erreur load Roma FlashCard et/ou PCunit à remplacer (attention n° de serie ne correspond pas à celui qui apparait sur leur Aqure(092R0152N032 vs 092R0151N032) à vérifier Intervention souhaitée tel: 0385910126 ou 0385446688 Demande transmise | Impossibilité de redémarrage soft | Chargement du soft en 3.3MR1 ,rechargement de la configuration client, test informatique ok passage d'un gazo ok | SWO-451972 | Defect | SW | RMED SW | RMED SW |  | Black screen device following start-up application -Error load Roma FlashCard and/or PCunit to replace (attention series does not correspond to the one that appears on their Aqure (092R0152N032 vs 092R0151N032) to check Desired intervention such as: 0385910126 or 0385446688 Request forwarded | Unable to restart softly | Loading the app in 3.3MR1 , recharging the customer configuration, computer test ok passing a gazo ok
No | Visit | Complaint | RAME | PCO2 reading above the insert limit on qualicheck 5+ S7730. | PCO2 reading above the insert limit on S7730 lot R0162. Customer replaced the solution pack and sensor cassette but did not help. Checked zero calibration and sensor cassette connection PC board. All checked okay. | Replaced the solution pack. | SWO-452169 | [5 empty] | PCO2 reading above the insert limit on qualicheck 5+ S7730. | PCO2 reading above the insert limit on S7730 lot R0162. Customer replaced the solution pack and sensor cassette but did not help. Checked zero calibration and sensor cassette connection PC board. All checked okay. | Replaced the solution pack.
Yes | Visit | Complaint | RCHN | PO2验证不通过 | PO2验证不通过，借902-972,903-334各一个 | 更换902-972后修复by tim | SWO-448687 | [5 empty] | PO2 validation fails | PO2 verification does not pass, borrow 902-972, 903-334 each | Repair by tim after replacement 902-972
No | Visit | Complaint | RCHN | 电池没有蓄电能力 | 电池没有蓄电能力 借431-018C一个 | 更换431-018C一个by tim | SWO-450952 | [5 empty] | The battery doesn't have the power to store electricity. | Battery does not have the capacity to store electricity, borrow 431-018C one | Replace 431-018C a by tim
No | Visit | Complaint | RCHN | 水晶柱断裂 | 水晶柱断裂 借924-591一个 | 更换924-591后修复by tim | SWO-450199 | Defect | HW | Wet section | Manifold |  | Crystal column break | Crystal column break, borrow 924-591 one. | Repair by tim after replacement 924-591
No | Visit | Complaint | RCHN | 测试卡损坏 | 测试卡损坏，具体部位是Teeth | 更换新测试卡后修复by tim | SWO-451638 | Defect | Consumables | Test Kit | Cartridge, Broken | Teeth; 5784 | Damaged test card | The test card is damaged, the specific location is Teeth | Fix by tim after replacing a new test card
Yes | Visit | Complaint | RCHN | 报错1024 | 报错3268，加借903-364一个 | 更换903-364后发现有新备件质量问题，需要重新借件903-364和903-702各一个by tim | SWO-450866 | [5 empty] | 1024 error | 3268, plus 903-364 one | New spare parts quality issues were found after replacement of 903-364, requiring a re-borrowing of 903-364 and 903-702 each by tim
No | Visit | Complaint | RDEU | Probenproblem | Probenproblem 0653 Probe bei Ansaugung vorzeitig am oberen pH $ BG Flüssigkeitsfühler erkannt | Fehlersuche - Erneuerung Luftventil pH $BG Modul - Funktionsüberprüfung mit anschließendem Wartungsdienst (siehe Case) | SWO-451875 | Defect | HW | pH/BG Module | Valve, air |  | Probenproblem | Sample problem 0653 Sample detected prematurely at the upper pH at the upper pH in case of suction | Troubleshooting - Renewal of air valve pH $BG module - Function check with subsequent maintenance service (see case)
No | Visit | Complaint | RCHN | PH电解质不过 | PH电解质不过，借945-603一个 | 更换945-603后修复by tim | SWO-450905 | Assisted Replacement | Electrode | Electrode | Reference | CAPA-733 | PH electrolyte is not the | PH electrolyte however, borrow 945-603 a | Repair by tim after replacement 945-603
No | Visit | Complaint | RAME | PCO2 level 1 dropped from 68.1 to 64.8 after Reference remembrane. | Looked in data logs and found that after remembraned the reference electrode the PCO2 started dropping on all 3 levels but level 1 was the worst. | Replaced PH oring and cleaned measuring chamber and move Reference electrode 120 degrees and rescanned and installed the Gas 1 and Gas 2 bottles. | SWO-451490 | Assisted Replacement | Electrode | Membrane | Reference |  | PCO2 level 1 dropped from 68.1 to 64.8 after Reference remembrane. | Looked in data logs and found that after remembraned the reference electrode the PCO2 started dropping on all 3 levels but level 1 was the worst. | Replaced PH oring and cleaned measuring chamber and move Reference electrode 120 degrees and rescanned and installed the Gas 1 and Gas 2 bottles.
No | Visit | Complaint | RAME | Rinse Errors | Rinse error | After troubleshooting found the gas $rinse valves to be the cause. Replaced valves for Ph $ Bg and El modules. | SWO-452116 | Defect | HW | pH/BG Module | Valve | Also El/Met Module | Rinse Errors | Rinse error | After troubleshooting found the gas $rinse valves to be the cause. Replaced valves for Ph $ Bg and El modules.
Yes | Visit | Complaint | RPOL | Analizator zastępczy Chrzanów Laboratorium. | Analizator zastępczy | Analizator zastępczy | SWO-451992 | Re classified RMED | same day upSWO |  |  | Loan analyzer, see mail with SWO number as reference | Substitute analyzer for the Chrzanów laboratory. | Replacement analyzer | Replacement analyzer
No | Visit | Complaint | RPOL | Błąd : analizator nie podłączony do DMS | Aparat nie podłącza się do DMS | Analizator niesprawny, zabrano do serwisu. | SWO-451991 | Revisit required | Analyzer replaced | Analyzer replaced | Analyzer replaced |  | Error: Analyzer not connected to DMS | The camera does not connect to DMS | Analyser not fit, taken to service.
No | Visit | Complaint | RCHN | 接口板腐蚀 | 接口板腐蚀,需要借件902-972一个 | 更换902-972后修复by tim | SWO-451533 | [5 empty] | Interface plate corrosion | Interface plate corrosion, need to borrow 902-972 one | Repair by tim after replacement 902-972
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | CO2定标灵敏度超范围 0378 | 连续更换2张测试卡后修复 | SWO-454120 | Assisted Replacement | SC | SC | pCO2 | Sens | Test card, reagent package failure | CO2 calibration sensitivity over range 0378 | Repair after 2 consecutive test cards
No | Visit | Complaint | RCHN | 无法正常识别定标卡和测试卡 | 无法正常识别定标卡和测试卡 | 更换910-304后修复by tim | SWO-452343 | Defect | HW | Storage Module | Barcode reader, storage wheel |  | The calibration and test cards are not properly identified | The calibration and test cards are not properly identified | Repair by tim after replacing 910-304
No | Visit | Complaint | RCHN | 仪器间接性无法使用，报错0643和0578 | 仪器间接性无法使用，报错0643和0578，需要借件902-847 | 更换902-847后修复by tim | SWO-451628 | Defect | HW | Basic Unit | PCB, analyzer control | 643, 578 | Instrument indirect ity can not be used, error0000s 0643 and 0578 | Instrument indirect can not be used, error 0643 and 0578, need to borrow 902-847 | Repair by tim after replacement 902-847
No | Visit | Complaint | RIND | Parameter are yellow, senstivity is low | Glucose , Creatnine parameter is yellow. | Check the calibration report, senstivity of membranes is low. Clean the chamber, run protein removal. run two point calibration, yet the parameter is yellow , senstivity is low. Need to change membranes. | SWO-452316 | Assisted Replacement | Electrode | Membrane | Crea | Also Glu; mail: how old were membranes? | Parameter are yellow, senstivity is low | Glucose , Creatnine parameter is yellow. | Check the calibration report, senstivity of membranes is low. Clean the chamber, run protein removal. run two point calibration, yet the parameter is yellow , senstivity is low. Need to change membranes.
No | Visit | Complaint | RIND | Oximetry hardware problem shows | Check for pending problem for which OXI parameters are missing and complete OXI test failed and found spectrometer error. | Replace complete OXI module,run LS adjust complete OXI test in which in hemolyzer shows phase regulation error shows and complete OXI test failed.Run 2-point calibration which is OK.Check some samples which is OK and no error for OXI parameters but need to check for phase regulation error. | SWO-452284 | Defect | HW | Oxi module | Oxi module |  | Oximetry hardware problem shows | Check for pending problem for which OXI parameters are missing and complete OXI test failed and found spectrometer error. | Replace complete OXI module,run LS adjust complete OXI test in which in hemolyzer shows phase regulation error shows and complete OXI test failed. Run 2-point calibration which is OK. Check some samples which is OK and no error for OXI parameters but need to check for phase regulation error.
No | Visit | Complaint | RCHN | 屏幕有黑点，pco2定标失败 | 屏幕有黑点，pco2定标失败，借903-239（被903-672代替）和903-344各一个 | 更换903-672和903-344后修复by tim | SWO-451597 | [5 empty] | The screen has a black dot, pco2 calibration fails | The screen has black dots, pco2 calibration fails, borrows 903-239 (replaced by 903-672) and 903-344 each | Repair by tim after replacing 903-672 and 903-344
No | Visit | Complaint | RCHN | 报错5648 | 报错5648,需要借件636-574和903-138各一个 | 更换636-574和903-138备件后修复by tim | SWO-452234 | Defect | HW | Incubator Module | PCB, incubator | 5648 | 5648 error | 5648, need to borrow 636-574 and 903-138 each | Repair by tim after replacement of 636-574 and 903-138 spare parts
Yes | Visit | Complaint | RLTD | Connect to AQURE & setup live connect | Connect to AQURE & setup live connect - CATH LAB machine | Set up IT connections on machines ahead of move to the wards. Lab to inform FSE date of move to ward for use. | SWO-452859 | Re classified RMED | planned SWO |  |  | see mail with title "RE: SWO-452859" | Connect to AQURE & setup live connect | Connect to AQURE & setup live connect - CATH LAB machine | Set up IT connections on machines ahead of move to the wards. Lab to inform FSE date of move to ward for use.
Yes | Visit | Complaint | RLTD | Connect to AQURE & setup live connect - 1A Ward | Connect to AQURE & setup live connect - 1A Ward | Set up IT connections on machines ahead of move to the wards. Lab to inform FSE date of move to ward for use. | SWO-452860 | Re classified RMED | planned SWO |  |  | see mail with title "RE: SWO-452859" | Connect to AQURE & setup live connect - 1A Ward | Connect to AQURE & setup live connect - 1A Ward | Set up IT connections on machines ahead of move to the wards. Lab to inform FSE date of move to ward for use.
No | Visit | Complaint | RDEU | Einlass offen, wird nicht geschlossen | Einlass in falscher Position, kein Gerätedefekt | Fehlersuche, Neujustierung des Einlasses, Funktionskontrolle | SWO-452945 | Adjustment | HW | Auto Inlet module | Auto inlet cal | 1391 | Entrance open, will not be closed | Inlet in wrong position, no device defect | Troubleshooting, re-adjustment of the inlet, function control
No | Visit | Complaint | RCHN | glu响应错误，报错1310 | glu响应错误，报错1310，需要借件903-581 | 更换903-581备件后修复（By Ran） | SWO-452272 | Defect | HW | Basic Unit | PCB, Sensor interface | Glu | glu response error, error 1310 | glu response error, error 1310, need to borrow 903-581 | Repair after replacement of 903-581 spare parts (By Ran)
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | NA,PO2定标不过 0375,1070,1319 | 连续更换2张测试卡和1个试剂包后修复 | SWO-454152 | Assisted Replacement | SP | SP | SP | pO2, Na | Test card, reagent package failure | NA, PO2 is not more than 0375,1070,1319 | Repair after 2 test cards and 1 reagent kit in a row
No | Visit | Complaint | RCHN | 光学单元灵敏度超范围1031 | 光学单元灵敏度超范围1031 借903-364一个 944-268（14707缺货），944-258（14867） |  | SWO-452508 | Cleaned | HW | Process Module | Process Wheel | 1029 | Optical unit sensitivity over range 1031 | Optical unit sensitivity over range 1031 borrowing 903-364 a 944-268 (14707 out of stock), 944-258 (14867)
Yes | Visit | Complaint | RCHN | 报错1024 | 更换903-364后发现有新备件质量问题，需要重新借件903-364和903-702各一个 | 更换903-364和903-702备件后修复(By Ran) | SWO-452219 | [5 empty] | 1024 error | New spare parts were found to be quality after replacement of 903-364, requiring a re-borrowing of 903-364 and 903-702 each | Repair after replacement of 903-364 and 903-702 spare parts (By Ran)
No | Visit | Complaint | RDEU | Hotline - immer wieder Driften bei Lyten, QC wird dadurch verletzt, wiederholte Kal löst das Problem aber bei nächster Kal wieder Fehler | Fehlersuche, sehr häufige Driften bei allen Lyten, Kal im Wechsel fehlerhaft, Ventile überprüft, Lösungstransport überprüft, Einlaßdichtung verstopft | alle Ventile oben gereinigt, beide V2 Ventile hingen fest, zusammengesetzt, LS und Pumpenkal, Ref Elektrode ersetzt, Lyte und Ref membraniert, Einlaßdichtung ersetzt, mehrere Kal, QC | SWO-454233 | Assisted Replacement | Electrode | Electrode | Reference | CAPA-733 | Hotline - repeatedly drifting at Lyten, QC is injured by this, repeated Kal solves the problem but again with next Kal error | Troubleshooting, very frequent drifting on all lytes, Kal alternating faulty, valves checked, solution transport checked, inlet seal clogged | all valves cleaned above, both V2 valves hung tight, composite, LS and pump scale, ref electrode replaced, lyte and ref membraneed, inlet seal replaced, several Kal, QC
No | Visit | Complaint | RPAC | Low bias on pCO2 QC since early October. | Low bias on pCO2 QC all levels since early October. pCO2 membrane replaced 2 $11 R0945. Found a lot of blood on reference electrode and chamber. | Cleaned chambers and replaced reference membrane to R2358. Replaced pH sealing ring (black). Checked isolation = 29. Clean, LS adjustment. Calibration and QC S7745 and S7735 now on mean. | SWO-452655 | Assisted Replacement | Electrode | Membrane | Reference |  | Low bias on pCO2 QC since early October. | Low bias on pCO2 QC all levels since early October. pCO2 membrane replaced 2 /11 R0945. Found a lot of blood on reference electrode and chamber. | Cleaned chambers and replaced reference membrane to R2358. Replaced pH sealing ring (black). Checked isolation = 29. Clean, LS adjustment. Calibration and QC S7745 and S7735 now on mean.
No | Visit | Complaint | RDAN | Dorte klager over tHb drifter Paa op til 27 nA. | Undersogt. | Lamp unit udskiftet opjusteret, service tHb kalibreret, slutafprovet. | SWO-452769 | Defect | HW | Oxi module | Lamp unit |  | Dorte complains about tHb drifter up to 27 nA. | Undersogt. | Lamp unit replaced upwards, service tHb calibrated, end of Provet.
Yes | Visit | Complaint | RIND | Sample door assembly error. | Check the system. | Sample inlet door assembly operational error. Door movement error & reset the connection. Still the movement error. need to replace the door assembly. | SWO-455854 | [5 empty] | Sample door assembly error. | Check the system. | Sample inlet door assembly operational error. Door movement error & reset the connection. Still the movement error. need to replace the door assembly.
No | Visit | Complaint | RKOR | Ca Calibration error | . | Calibration check, Service mode check. SC replace | SWO-455939 | Assisted Replacement | SC | SC | Ca | Cal | Ca Calibration error | . | Calibration check, Service mode check. SC replace
No | Visit | Complaint | RCHN | 条码扫描器没有亮光 | 条码扫描器没有亮光，借902-924一个 | 更换902-924后修复(By Ran) | SWO-452599 | Defect | HW | Basic Unit | Barcode reader |  | Bar code scanner does not have light | Bar code scanner has no light, borrow 902-924 a | Post-replacement 902-924 (By Ran)
No | Visit | Complaint | RCHN | 测试卡断裂，报错5784 | 测试卡断裂，报错5784，断裂部位（nose） | 更换测试卡后修复(By Ran) Catridge was broken. | SWO-453332 | Defect | Consumables | Test Kit | Cartridge, Broken | Nose | Test card break, 5784 error | Test card break, 5784, broken site (nose) | Repair after replacing the test card (By Ran) Catridge was broken.
Yes | Visit | Complaint | RCHN | 定标失败 | 定标失败，借944-220（15065）一张 | 重新定标后修复(By Ran)Recal, then OK. | SWO-452704 | [5 empty] | Bid failed | Failed to bid, borrow 944-220 (15065) | Re-calibration repair (By Ran) Recal, then OK.
No | Visit | Complaint | RDEU | 12.11.2018 08:50Uhr Adebar 049 621 383-3528: Hoher Drift bei den Elektrolyten 20.11.2018, HerrAdebar, Tel: 06213833528: weiterhin hohe Driften bei den Elektrolyten | KAL Driften Elektrolyte | Allgemeine Überprüfung - Erneuerung Referenz Elektrode - Funktionsüberprüfung etc. | SWO-453819 | Assisted Replacement | Electrode | Electrode | Reference | CAPA-733 | 12.11.2018 08:50Ade Adebar 049 621 383-3528: High drift in electrolytes 20.11.2018, HerrAdebar, Tel: 06213833528: still high drifts in electrolytes | KAL Drifting Electrolytes | General verification - Renewal reference electrode - function check etc.
No | Visit | Complaint | RAME | Leakage detected | Clot in inlet. | Flush clot from inlet. Remove electrodes and clean clot from Reference Electrode. | SWO-453920 | Cleaned | Fluidic pathway | Fluidic pathway | Clot |  | Leakage detected | Clot in inlet. | Flush clot from inlet. Remove electrodes and clean clot from Reference Electrode.
No | Visit | Complaint | RDEU | 19.11.2018, Frau Peters, tel. 093727001652: ständig Kal Fehler bei diversen Parameter | Analysator ist bei Service Messbereich mit grüner Ampel. Überprüfung der Kalibrationen und der Qualitätskontrollen. Es zeigt sich, dass bei den meisten Kalibrationen die Elektrolyt Parameter oft schwanken. pCO2 hatte auch vor kurzem extreme Werte. In der Qualitätskontroll Historie ist ersichtlich, dass Level rot und gelb zu niedrig und Level blau zu hoch gemessen werden. | Tausch der Referenz Elektrode und der pCO2 Elektrode. Messkammern gereinigt auch pH Elektrode. Analysator einkalibriert. Driften stabil und minimal. Beobachtung der Stabilität. Gewährleistung | SWO-454373 | Assisted Replacement | Electrode | Electrode | Reference | CAPA-733 | 19.11.2018, Mrs. Peters, tel. 093727001652: constantly Kal error with various parameters | Analyzer is at service measuring range with green traffic light. Verification of calibrations and quality controls. It turns out that in most calibrations the electrolyte parameters often fluctuate. pCO2 also recently had extreme values. Quality control history shows that levels are too low and levels are too low and levels are too high. | Exchange of the reference electrode and the pCO2 electrode. Measuring chambers also cleaned pH electrode. analyzer calibrated. Drifting stable and minimal. Observation of stability. Ensure
No | Visit | Complaint | RSCH | Rep. Undicht | Cleaningsolution undicht Y-Stück zu Abfall eingeklemmt Einlass bei Kapillare undicht pO2 Membrane undicht Driftet seit Auslieferung | Nippel Cleaningsolution gespült i.o Flasche ersetzt i.o Einlassdichtung ersetzt i.o Y-Stück richtig positioniert Membranenwartung pO2 durchgeführt i.o Kal. und QC Level 1 i.o System i.o 1x Cleaningsolution wird gutgeschrieben 1x pO2 Membrane | SWO-453425 | Adjustment | HW | Fluidic manifold | Y-piece |  | Rep. Leaking | Cleaningsolution leaks Y-piece to waste jammed Entrance at capillary leak pO2 membrane leaks Drifting since delivery | Nipple Cleaningsolution rinsed i.o Bottle replaces i.o Inlet seal replaces i.o Y-piece correctly positioned Membrane maintenance pO2 carried out i.o Kal. and QC Level 1 i.o System i.o 1x Cleaningsolution is credited 1x pO2 Membrane
No | Visit | Complaint | RDAN | Dorte klager over afbrudt 2 punks kalibrering : fejl 597 | Undersøgt. | Målemoduls ventiler udskiftet,slutafprøvet. | SWO-453489 | Defect | HW | pH/BG Module | Valve, membrane | 597 | Dorte complains of interrupted 2 punks calibration: Bug 597 | Examined. | Measuring module valves replaced, final tested.
No | Visit | Complaint | RIND | Storage wheel error. | Check the system message & found RotateStorage_24 | Clean the storage wheel & storage wheel sensor ,Restart the instrument & it became ready. | SWO-455793 | Cleaned | HW | Storage Module | sensor, storage wheel |  | Storage wheel error. | Check the system message & found RotateStorage_24 | Clean the storage wheel & storage wheel sensor ,Restart the instrument & it became ready.
No | Visit | Complaint | RDEU | Touch Screen ohne Funktion | Touch Screen ohne Funktion (Meldung durch Labor am 20.11.18) Bestellung von Display Unit und CPU am 20.11.18 erfolgt Display Unit nicht eingetroffen ! | Allgemeine Überprüfung - Erneuerung CPU (Rückmontage $ keine Fehlerbehebung) - Display Unit (930-353P1745N7313) defekt - aus Labor-Ersatz-Analysator (090R0927N004) Display Unit (930-353P1742N7008) demontiert und in Analysator installiert - Funktionsüberprüfung etc. ohne Beanstandung | SWO-454510 | Defect | HW | DMS_PC | display |  | Touch screen without function | Touch screen without function (message by laboratory on 20.11.18) Order of Display Unit and CPU on 20.11.18 takes place Display Unit not arrived ! | General Review - Renewal CPU (Reassembly no troubleshooting) - Display Unit (930-353P1745N7313) defective - from lab replacement analyzer (090R0927N004) Display Unit (930-353P1742N7008) dismantled and installed in analyzer - function check etc. without complaint
No | Visit | Complaint | RDEU | Verbindung Radiance unterbrochen | Netzwerkanbindung geprüft und hergestellt Nach Anruf von Herrn Ochs, Labor Staber Anwender hatte unterbrochene Verbindung nicht bemerkt erst durch Anruf von Herrn Ochs wurde dies an Radiometer gemeldet | Analysator hatte eigene IP verloren, diese eingetragen, Verbindung hergestellt, Datenabfrage funktioniert, Amwender darauf hingewiesen bei Problemen mit der Service Hotline Kontakt aufzunehmen Zählerstände im gesamten Sana Klinikum Hof aufgenommen | SWO-454630 | Setup | SW | RMED SW | His/Lis | IP | Connection Radiance interrupted | Network connection tested and established After a call from Mr. Ochs, Labor Staber User had not noticed interrupted connection only by calling Mr Ochs this was reported to radiometer | Analyzer had lost its own IP, entered it, established a connection, data retrieval works, Amwender pointed out to contact the service hotline if there were problems Counter readings recorded in the entire Sana Klinikum Hof
No | Visit | Complaint | RAME | In homogeneous errors | Customer replaced both SP & SC multiple times and still getting in homogeneous errors. | Replace the inlet gasket, connector gasket and internal tubing. Ran diagnostics and calibrations. | SWO-455631 | Defect | HW | Basic Unit | PM Kit |  | In homogeneous errors | Customer replaced both SP & SC multiple times and still getting in homogeneous errors. | Replace the inlet gasket, connector gasket and internal tubing. Ran diagnostics and calibrations.
Yes | Visit | Complaint | RHUN | a nyomtató díszítő sorokat nyomtat kizárólag, semmi karakter nem felismerhető.. | nyomtató szoftveres beállításokra nem reagál | nyomtató egység cseréje, LAC kalibrációs hiba miatt SC garanciális csere, készülék init, kalibráció, QC mérés | SWO-454433 | [5 empty] | The printer prints decorative lines only, no characters are recognizable.. | Printer does not respond to software settings | Replacing the printer unit, LAC calibration error due to SC warranty replacement, device init, calibration, QC measurement
No | Visit | Complaint | RAME | Printer is intermittently not working, has to reboot to print. Rick Panak - 215-606-8138 | 0216 General Printer Error | Reseated the Printer USB communication cables to analyzer, cycled power and tested. | SWO-454708 | Adjustment | HW | Printer Module | Printer Module | 216 | Printer is intermittently not working, has to reboot to print. Rick Panak - 215-606-8138 | 0216 General Printer Error | Reseated the Printer USB communication cables to analyzer, cycled power and tested.
No | Visit | Complaint | RAME | Printer is intermittently not working, has to reboot to print. | 0216 General Printer Error | Reseated the Printer USB communication cables to analyzer, cycled power and tested. Provided customer with October Service Work Orders. Sent National Tech Manager Service dumps. | SWO-454716 | Adjustment | HW | Printer Module | Printer Module | 216 | Printer is intermittently not working, has to reboot to print. | 0216 General Printer Error | Reseated the Printer USB communication cables to analyzer, cycled power and tested. Provided customer with October Service Work Orders. Sent National Tech Manager Service dumps.
Yes | Visit | Complaint | RPAC | pO2 calibration failures. Error code 1006 | Error 1006 persisting after change of sensor cassette. | Changed solution pack. Error 1006 persisting - pO2 failing calibration. Re-booted analyser and re-ran QC. Error 1006 persisting. Case transferred to service engineer. | SWO-454893 | Re classified RMED | installation report |  |  | Moved to OOB trending (see mail with SWO number title) | pO2 calibration failures. Error code 1006 | Error 1006 persisting after change of sensor cassette. | Changed solution pack. Error 1006 persisting - pO2 failing calibration. Re-booted analyser and re-ran QC. Error 1006 persisting. Case transferred to service engineer.
Yes | Visit | Complaint | RPAC | pO2 calibration failures. Error code 1006 | Error 1006. | Changed sensor cassette. Calibration still failing on pO2 (Error 1006). Changed solution pack. Calibration still failing on pO2 (Error1006). Case assigned to service engineer. | SWO-454889 | Re classified RMED | same day upSWO |  |  |  | pO2 calibration failures. Error code 1006 | Error 1006. | Changed sensor cassette. Calibration still failing on pO2 (Error 1006). Changed solution pack. Calibration still failing on pO2 (Error1006). Case assigned to service engineer.
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 液体传输失败 1249 | 更换试剂包后修复 | SWO-454815 | Assisted Replacement | SP | SP | FT | 1249 | Test card, reagent package failure | Liquid transfer failure 1249 | Repair after replacement of reagent kit
Yes | Visit | Complaint | RCHN | 老师投诉仪器的耗材经常断货 | 老师投诉仪器的耗材经常断货 | 已解决（By Ran）This customer was short of consumable. Their distributor didn't supply solution pack and sensor cassette in time. | SWO-454094 | Re classified RMED | Logistical upSWO |  |  |  | The teacher complained that the supplies of the instrument were often out of stock. | The teacher complained that the supplies of the instrument were often out of stock. | 已解决（By Ran）This customer was short of consumable. Their distributor didn't supply solution pack and sensor cassette in time.
No | Visit | Complaint | RCHN | 报错1068，干燥前温度超范围 | 报错1068，干燥前温度超范围 | 调整内部温度传感器位置(By Ran) Adjusted the temp sensor of cup dryer unit. | SWO-454159 | Adjustment | HW | Process Module | Cup dryer | 1068 | Error 1068, pre-drying temperature over range | Error 1068, pre-drying temperature over range | Adjust the position of the internal temperature sensor (By Ran) Adjusted the temp sensor of cup dryer unit.
No | Visit | Complaint | RDEU | 05.11.2018, Herr Dietrich, Tel: 015256714506: License Key abgelaufen- Gerät ohne Funktion 12.11.18: während Einsatz ermittelt das hier nicht Keys abgelaufen sind, sondern der Basic Key defekt ist „License Key missing“ Auftragsnummer: | SW key wechseln | Gerät aus; Key gewechselt; Gerät ein; Startphase abgewartet; POC-IT informiert; Kal. + QC i.O. | SWO-454703 | Defect | HW | Analyzer Control Module | License key |  | 05.11.2018, Mr. Dietrich, Tel: 015256714506: License Key expired- Device without function 12.11.18: while using it determines that no keys have expired here, but the basic key is defective "License Key missing" Order: | SW key change | device off; Key changed; device; start-up phase awaited; POC-IT; Kal. + QC in the first place
Yes | Visit | Complaint | RDEU | 21.11.2018, 11:12 Uhr, H. Stein, Tel.: 0611433260: Demo Option Key läuft bald ab | Finale Keys installieren. | Keys installiert, Update Type, Startphase, Kalibriert, Funktionsprüfung i.O | SWO-454724 | Re classified RMED | planned SWO |  |  |  | 21.11.2018, 11:12 a.m., H. Stein, Tel.: 0611433260: Demo Option Key expires soon | Install Final Keys. | Keys installed, Update Type, Start Phase, Calibrated, Functional Check i.O
No | Visit | Complaint | RDEU | 20.11.2018, Herr Kreuter( Labor), Tel: 022678897048: mehrere Messabbrüche seit Wochenende | O.g. Fehlermeldung (3553: INCUBATORWHEELMOVE_31) insgesamt drei mal seit 18.11.2018 | Fehlersuche: Verfahren im Rahmen der heutigen Wartung durchgeführt, keine mechanischen Fehler feststellbar! | SWO-454959 | Cleaned | HW | Incubator Module | Incubator wheel | 3553 | 20.11.2018, Mr. Kreuter( Laboratory), Tel: 022678897048: several measurement interruptions since the weekend | O.g. error message (3553: INCUBATORWHEELMOVE_31) three times in total since 18.11.2018 | Troubleshooting: Procedure carried out within the scope of today's maintenance, no mechanical faults detectable!
No | Visit | Complaint | RPOL | Klient zgłosił błąd lampy OXi | Błąd lampy neonowej. | Wymiana modułu OXi, cal i testy OXi w normie. | SWO-454416 | Defect | HW | Oxi module | Oxi module | Lamp Unit | Client reported error of OXi lamp | Neon lamp error. | Replacing OXi, Cal, and OXi tests in the standard.
No | Visit | Complaint | RSAF | New oxi liquid sensor to be installed by Medhold. | Oxi liquid sensor error | Liquid sensor LS3 to be installed as its not functioning, installed new liquid sensor for Oxi module LS3, started up and tested Liquid sensor and tested ok, calibrated but pCo2 not correcting, installed new Solution pack but sensor cassette still did not correct on Na, pCo2 and Glu and Lac, installed new sensor cassette and left analyzweer with Lactate calibration and Qc errors, Sensor cassette needs to stablize for 4 to 24hrs for the metabolites to stabilize. | SWO-454355 | Defect | HW | Oxi module | Liquid sensor |  | New oxi liquid sensor to be installed by Medhold. | Oxygen Liquid Sensor Error | Liquid sensor LS3 to be installed as its not functioning, installed new liquid sensor for Oxi module LS3, started up and tested Liquid sensor and tested ok, calibrated but pCo2 not correcting, installed new Solution pack but sensor cassette still did not correct on Na, pCo2 and Glu and Lac, installed new sensor cassette and left analyzweer with Lactate calibration and Qc errors, Sensor cassette needs to stablize for 4 to 24hrs for the metabolites to stabilize.
No | Visit | Complaint | RDEU | Touch Screen ohne Funktion | Touch Screen ohne Funktion | Erneuerung Touch Screen - Funktionsüberprüfung etc. | SWO-455068 | Defect | HW | DMS_PC | Touch screen |  | Touch screen without function | Touch screen without function | Renewal Touch Screen - Function check etc.
No | Visit | Complaint | RDEU | 20.11.2018, 08:21 Uhr, H. Wahl, Tel. 075419671082: Fehler Flüssigkeitssensor erkannt | Fehlersuche im Rahmen einer JW Fehler tritt morgens auf | LS 1 Inlet zeigt 254 ohne Änderung. Sensor Bay ersetzt. Inlet Probe neu LS Justage nun 69 ist ok. SP erneuert nach Startphase. Kalibration und QC durchgeführt und in Ordnung | SWO-455369 | Defect | HW | Basic Unit | Sensor Bay Unit | LS1 | 20.11.2018, 08:21, H. Wahl, Tel. 075419671082: Fault liquid sensor detected | Troubleshooting as part of a JW Error occurs in the morning | LS 1 Inlet shows 254 without change. Sensor Bay. Inlet Probe new LS Adjustment now 69 is ok. SP renews after start-up phase. Calibration and QC carried out and in order
No | Visit | Complaint | RKK | 圧力吸引テストエラー多発 | 圧力吸引テストエラー 多発 | 新品溶液パック交換対応 しかし新品溶液パック取付後気泡エラー発生。 ｲﾝﾚｯﾄｶﾞｽｹｯﾄ新品交換対応にてＯＫ | SWO-455910 | Defect | HW | Inlet module | Inlet gasket | 1063 | Frequent pressure suction test errors | Pressure suction test error frequent | New solution pack replacement support However, bubble error occurs after the installation of a new solution pack. Ok with new inlet gasket replacement
No | Visit | Complaint | RDEU | 29.10.2018 14:00 Fr. Richter ????: Interventionsmeldung Gerinsel in der Oximetrie 20.11.2018 14:20Uhr Dr.Holfertholfert 0352365225: Das Gerät meldet weiterhin Fehler Intervention erforderlich | Datenbank, Flüssigkeitsstrecke und Ventil geprüft | Flüssigkeitsstrecke gespült --> keine Behinderung im Fluss Quetschventil erneuert Neue SC installiert --> Neustart | SWO-455239 | Defect | HW | Oxi module | Pinch valve |  | 29.10.2018 14:00 Fr. Richter ????: Intervention report Gerinsel in oximetry Nov 20, 2018 2:20pm Dr.Holfertholfert 0352365225: The device continues to report error intervention required | Database, fluid section and valve tested | Liquid track rinsed --> no obstruction in the river Squeeze valve renewed New SC installed --> Restart
No | Visit | Complaint | RFRA | Réglage sauvegardes automatique | 22 $11: contrôle connexion AQURE - Modification paramétrage sauvegarde - Tests OK | Réglage sauvegardes automatique | SWO-455707 | Setup | SW | RMED SW | Backup | Aqure | Setting automatic backups | /22 /11: AQURE connection control - Backup setting change - OK Tests | Setting automatic backups
No | Visit | Complaint | RFRA | Err status 2 | PC unit | Démontage et remplacement du PC $Installation de W7 $Installation Bios UIM945 $installation de l'application Radiometer $Installation du patch 933-447-SinglePatchForKB4012215 $Restauration de la configuration récuperée sur le reseau $Calibration et passage des CQ. | SWO-455097 | Defect | HW | DMS_PC | CPU | status code 2 | Err status 2 | PC unit | Disassembly and replacement of W7 $Installation PC $Installation Bios UIM945 $installation of the Radiometer application $Installation patch 933-447-SinglePatchForKB4012215 $Restauration of the reclaimed configuration on the $Calibration network and passing of the CQ.
No | Visit | Complaint | RFRA | ABL90 en erreur de boot HDD avec écran qui ne répond pas pour pouvoir continuer le démarrage Erreur rapportée par le client Demande d'intervention transmise | PC module | Démontage et remplacement du pc $Mise en place de la CF card d'origine $Remise en route ok $erreur Calibration pCO2 $Glu $Lac $ Attendre 2heures pour stabilisation sinon remplacer cassette de capteurs. | SWO-455236 | Defect | HW | DMS_PC | CPU |  | ABL90 in HDD boot error with screen that does not respond to be able to continue the start Customer-reported error Request for intervention sent | PC module | Disassembly and replacement of pc $Mise in place of the original CF card $Remise en route ok $erreur PCO2 calibration $Glu $$Lac Wait 2 hours for stabilization if not replace cassette of sensors.
No | Visit | Complaint | RLTD | Battery indicator indicator not working | Mike fagan Hall had reported that analyser battery indicator light was not working. | Attended site and found that indicator on battery module was not working at all. Unplugged from mains and reconnected but this did not resolve. Battery module showing as present and charged. It powered the analyser when mains disconnected. Powered down and removed battery module. Found fpc cable from battery pcb to indicator pcb was not connected at led end. Refitted cable and put battery module back into analyser. Powered up and battery indicator led worked as normal. Tested battery function in service programs to confirm. No further action required. Also assisted customer with trying to get older analysers connected back to aqure system. Two abl90 in ICU had been disconnected from aqure when network change had been made. It had accidentally deleted the vlan that the medical devices were connected to. Abl90 now connecting again but older Siemens analyser in lab not connected. Customer IT are still working in this. Customer discussing whether to go live early with lab analyser since training almost completer and validation only required a few more samples. | SWO-455160 | Adjustment | HW | Battery module | Cable |  | Battery indicator indicator not working | Mike fagan Hall had reported that analyser battery indicator light was not working. | Attended site and found that indicator on battery module was not working at all. Unplugged from mains and reconnected but this did not resolve. Battery module showing as present and charged. It powered the analyser when mains disconnected. Powered down and removed battery module. Found fpc cable from battery pcb to indicator pcb was not connected at led end. Refitted cable and put battery module back into analyser. Powered up and battery indicator led worked as normal. Tested battery function in service programs to confirm. No further action required. Also assisted customer with trying to get older analysers connected back to aqure system. Two abl90 in ICU had been disconnected from aqure when network change had been made. It had accidentally deleted the vlan that the medical devices were connected to. Abl90 now connecting again but older Siemens analyser in lab not connected. Customer IT are still working in this. Customer discussing whether to go live early with lab analyser since training almost completer and validation only required a few more samples.
No | Visit | Complaint | RDEU | 21.11.2018, 09:05 Uhr, H. Hauck, Tel. 03461271203: QC Verletzung in Level C bei pCO2 21.11.2018 HanSc 10:18 Tel.: Alles wieder grün, Gerät ok. | QC Verletzung im Lev A pCO2, zuletzt auch im Lev C | Überprüfung Daten, Probleme mit QC traten mit der aktuellen SC auf SC erneuert ( aus VSC Bestand) alte SC reklamiert Einlaufverfahren | SWO-455202 | Assisted Replacement | SC | SC | pCO2 | QC A og C | 21/11/2018, 09:05 AM, H. Hauck, Tel. 03461271203: QC Violation in Level C at pCO2 21.11.2018 HanSc 10:18 Tel.: Everything green again, device ok. | QC injury in Lev A pCO2, most recently also in Lev C | Review data, problems with QC occurred with the current SC SC renewed (from VSC stock) old SC complained Entry procedure
No | Visit | Complaint | RAME | Interface issues with new analyzer to Meditech | Meditech interface not receiving data from ABL90 Flex Plus. | Reset communication on ABL90. Completed full testing with customer and all samples and data cross to Meditech interface. Copied setup for customer to USB drive. | SWO-455747 | Adjustment | SW | RMED SW | His/Lis |  | Interface issues with new analyzer to Meditech | Meditech interface not receiving data from ABL90 Flex Plus. | Reset communication on ABL90. Completed full testing with customer and all samples and data cross to Meditech interface. Copied setup for customer to USB drive.
No | Visit | Complaint | RCHN | 处理包安装失败 | 处理包安装失败 | 更换处理包修复 | SWO-455364 | Defect | Consumables | Solution Pack | Solution Pack |  | Processing package installation failed | Processing package installation failed | Replacement package repair
No | Visit | Complaint | RCHN | 报错1031，光学强度超范围，定标卡无法扫描 | 报错1031，光学强度超范围，定标卡无法扫描，借903-364 902-924 944-214（14934）944-216（15082）各一个 | 更换903-364 902-924并重新定标后修复(by chuanwei) | SWO-454091 | Defect | HW | Process Module | Optical Unit | 1031 | Error 1031, optical strength over range, bid card can not be scanned | Error 1031, optical strength over range, bid card can not be scanned, borrow 903-364 902-924 944-214 (14934) 944-216 (15082) one each | Replace 903-364 902-924 and re-fix (by chuanwei)
No | Visit | Complaint | RCHN | 试剂包无法读取 | 试剂包无法读取 | 连续更换2个试剂包后修复 | SWO-454838 | Defect | Consumables | Solution Pack | Solution Pack | Chip | Reagent package cannot be read | Reagent package cannot be read | Repair after 2 reagent kits in a row
No | Visit | Complaint | RCHN | 开机后系统无法使用 | 蠕动泵接头无开口导致开机后系统无法使用 | 重新连接注射器针头后修复（by chuanwei） | SWO-454926 | [5 empty] | The system cannot be used after power-on | No opening to the peristaltic pump connector causes the system to be unusable after power-on | Repair after reconnecting syringe needle (by chuanwei)
No | Visit | Complaint | RDEU | Mischerfehler | Mischerfehler 0883 tritt gelegentlich auf ! Wurde bei Wartungsdienst bereits festgestellt und FlexQ Modul neu kalibriert ! | Analysator heruntergefahren - Erneuerung FlexQ Modul - Erneuerung Nippelkit Einlassmodul (leicht beschädigt) - Funktionsüberprüfung etc. | SWO-455070 | Defect | HW | FlexQ Module | FlexQ Module | 883 | Mischerfehler | Mixer error 0883 occurs occasionally ! Has already been detected during maintenance service and FlexQ module has been recalibrated ! | Analyzer shut down - Renewal FlexQ module - Renewal nipple kit inlet module (slightly damaged) - function check etc.
No | Visit | Complaint | RFRA | Réglage sauvegardes automatique | 22 $11: contrôle connexion AQURE - Modification paramétrage sauvegarde - Tests OK | Réglage sauvegardes automatiques | SWO-455710 | Setup | SW | RMED SW | Aqure | Automatic backup | Setting automatic backups | /22 /11: AQURE connection control - Setting change backup - OK tests | Setting up automatic backups
No | Visit | Complaint | RFRA | L'appareil s'éteint tout seul lors du passage d'une calibration et de QC sur la Béta. Le client a fait un arrêt propre et un démarrage propre, relancer une calibration et l'appareil a redémarré. Tnt et Bnp aucun problème. Ce problème s'est produit 3 fois uniquement sur la Béta. Lot de la Béta : 14576 depuis le 29 août. 05.49.83.83.03 - Mme REDEMPT 06.42.40.63.54 - Astreinte labo | Historique des fichiers d'activité. | Sauvegarde toutes données. Mise à jour logicielle, passage à la version 8.11.300.0. Appareil qualifié conformément aux spécifications RADIOMETER. | SWO-455565 | Defect | SW | RMED SW | RMED SW | 3050 | The device shuts down on its own when a calibration and QC pass over the Beta. The customer made a clean stop and a clean start, restart a calibration and the device restarted. Tnt and Bnp no problem. This problem occurred 3 times only on the Beta. Beta Lot: 14576 since August 29. 05.49.83.83.03 - Ms REDEMPT 06.42.40.63.54 - Lab | Activity file history. | Backup all data. Software update, switch to version 8.11.300.0. Qualified device to RADIOMETER specifications.
No | Visit | Complaint | RFRA | Erreur aléatoire de fermeture des portes inlet. Malgré lubrification et réglage. | Vérification en mode service des mouvements mécaniques: ok qualite@labo-barla.eu | Démontage: optique, roue stockage cartouches, Dépoussiérage, nettoyage: Storage wheel, incubateur, needle wash unit, cup wash, capteurs et scanner. Graissage et huilage des pièces en mouvement. Déconnection, connexion des connecteurs informatique, électrique, opto. Test cup transport. Redémarrage routine, passage cleaning. | SWO-455182 | Defect | HW | Cover and Chassis | Door, Inlet | 5784 | Random inlet door closure error. Despite lubrication and setting. | Checking in mechanical motion service mode: ok qualite@labo-barla.eu | Disassembly: optical, wheel storage cartridges, Dust removal, cleaning: Storage wheel, incubator, needle wash unit, cup wash, sensors and scanner. Grease and oilmoving moving parts. Disconnection, connection of computer connectors, electrical, opto. Test cup transport. Routine restart, passage cleaning.
No | Visit | Complaint | RLTD | Alexandra - 02034470083 ex 70092 - analyser not turning on, mains checked, black screen | Not switching on. | Check fuses in mains socket ok. Removed battery and unit still not turning on. Suspect power supply is faulty part on order. | SWO-455743 | Revisit required | HW | Power Supply | Power supply |  | Alexandra - 02034470083 ex 70092 - analyser not turning on, mains checked, black screen | Not switching on. | Check fuses in mains socket ok. Removed battery and unit still not turning on. Suspect power supply is faulty part on order.
Yes | Visit | Complaint | RSAF | Monthly order delivered (only Sensor Cassette) no Solution pack or CCN replacement Solution pack also not delivered as arranged and will only arrive at customer sometime during the course of saterday. | Collect solution pack from stores and deliver before customer runs out of cycles on solution pack and damages sensor cassette. | Collected and delivered solution pack to customer. | SWO-455746 | [5 empty] | Monthly order delivered (only Sensor Cassette) no Solution pack or CCN replacement Solution pack also not delivered as arranged and will only arrive at customer sometime during the course of saterday. | Collect solution pack from stores and deliver before customer runs out of cycles on solution pack and damages sensor cassette. | Collected and delivered solution pack to customer.
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | PO2定标超范围 0375 | 更换测试卡正常。 | SWO-455810 | [5 empty] | Test card, reagent package failure | PO2 Calibration Over range 0375 | The replacement test card is OK.
No | Visit | Complaint | RIND | In samples electrolytes results are not satisfactory | Check analyzer status and calibration result which is OK.Check patient results which is also OK. | As reported check electrolytes results which are OK.Give basic training for sampling.In samples some parameters are missing with temperature missing and baro missing.In setting tick in temperature & Baro with retain. | SWO-455853 | [5 empty] | In samples electrolytes results are not satisfactory | Check analyzer status and calibration result which is OK. Check patient results which is also OK. | As reported check electrolytes results which are OK. Give basic training for sampling. In samples some parameters are missing with temperature missing and baro missing. In setting tick in temperature & Baro with retain.
No | Visit | Complaint | RDEU | Kal Fehler bei pH und den Elektrolyten | Referenzelektrode defekt | Fehlersuche, Referenzelektrode defekt->erneuert gegen LOT 08-16, mehrere Kalibrationen und QC jetzt ok | SWO-455731 | Assisted Replacement | Electrode | Electrode | Reference |  | Kal faults with pH and electrolytes | Reference electrode defective | Troubleshooting, reference electrode defective->renewed against LOT 08-16, several calibrations and QC now ok
No | Visit | Complaint | RDEU | zeitweise reagiert der Bildschirm nicht richtig, zeitw. ist der Bildschirm verschoben, | Touchscreen funktioniert nicht richtig, | ich habe einen anderen Touchscreen angeschlossen, der Kunde testet den Bildschirm und meldet sich eieder, | SWO-455649 | Defect | HW | DMS_PC | Touch screen |  | at times, the screen does not respond correctly, zeitw. the screen is moved, | Touchscreen is not working properly, | I have connected another touchscreen, the customer tests the screen and reports eieder,
No | Visit | Complaint | RIND | Capillary sample aborted. | Check the gasket found broken part of the capillary tube in the capillary inlet . | Remove the broken parts & clean the gasket,run the protein removal & cleaning .Run the capillary sample & got the desired result.Handover the instrument to the customer. | SWO-455800 | User Error | HW | Inlet module | Inlet gasket | broken capillary tube | Capillary sample aborted. | Check the gasket found broken part of the capillary tube in the capillary inlet . | Remove the broken parts & clean the gasket,run the protein removal & cleaning . Run the capillary sample & got the desired result. Handover the instrument to the customer.
No | Visit | Complaint | RPOL | Błąd UPS - „On battery”; „Overload” | Sprawdzenie UPS. | Wymiana UPS na nowy S $N 4B1824P34238. | SWO-455413 | Defect | HW | External HW | UPS |  | Błąd UPS - „On battery”; „Overload” | UPS check. | Replacing the UPS with the new S $N 4B1824P34238.
No | Visit | Complaint | RKOR | Interface error | . | Mirth check, AQURE check. Ping test, interafce PC error=> restart. | SWO-455950 | Adjustment | SW | His/Lis | His/Lis |  | Interface error | . | Mirth check, AQURE check. Ping test, interafce PC error=> restart.
No | Visit | Complaint | RCHN | 机器断电后，机器无法开机 | 机器断电后，机器无法开机,需要借件903-381、903-334、636-524 | 更换903-381、903-334、636-524正常。(by chuanwei) | SWO-454950 | [5 empty] | The machine does not power on when the machine is off | After the machine is powered off, the machine cannot be switched on, and the loan s903-381, 903-334, 636-524 is required | Replace 903-381, 903-334, 636-524 normal. (by chuanwei)
No | Visit | Complaint | RCHN | 未连接试剂包 | 未连接试剂包,借551-196，902-989各一个。 |  | SWO-455814 | [5 empty] | Unconnected reagent package | Unconnected reagent pack, borrow 551-196, 902-989 each.
No | Visit | Complaint | RCHN | CTHB,SO2检测值低于统计范围，质控B、C不通过 | 定标通过，CTHB,SO2检测值低于统计范围，质控B、C不通过 0090，0089，0588,借902-980，903-197各一个。 |  | SWO-455828 | Defect | HW | Oxi module | Hemolyzer |  | CTHB, SO2 detection value is lower than the statistical range, QC B, C does not pass | The standard pass, CTHB, SO2 detection value is lower than the statistical range, QC B, C does not pass 0090,0089,0588, borrow 902-980, 903-197 each.
No | Visit | Complaint | RCHN | 干燥前温度超范围 1068 | 干燥前温度超范围 1068,借903-210一个。 | 更换903-210一个正常。(by chuanwei) | SWO-455482 | Defect | HW | Process Module | PCB, cup dryer | 1068 | Over-range temperature before drying 1068 | The pre-drying temperature exceeds the range of 1068, borrowing 903-210 one. | Replace the 903-210 one normal. (by chuanwei)
No | Visit | Complaint | RCHN | 存储盘卡盘，定标失败 | 存储盘卡盘，定标失败，需要借件944-214（14465）,944-216（14616） | 重新定标正常。（by chuanwei） | SWO-454833 | Defect | Consumables | Test Kit | Cartridge, Broken | Potentially User Error | Memory disk card, bid failed | Memory disk card, bid failure, need to borrow 944-214 (14465), 944-216 (14616) | Recalibration is normal. (by chuanwei)
Yes | Visit | Complaint | RCHN | 孵育盘故障 | 孵育盘故障5658，借903-138一个 | 更换903-138一个正常。(by chuanwei) | SWO-455044 | [5 empty] | Incubation plate failure | Incubation plate fault 5658, borrow 903-138 one | Replace 903-138 one normal. (by chuanwei)
No | Visit | Complaint | RIND | Lactate calibration failure | Lactate calibration failed | Checked the activity log , found the error that hypochlorite detected at lactate and glucose , opened the membrane section checked the lactate membrane found liquid in membrane slot , cleaned it ,checked all the membranes ,changed the lactate membrane and then run Calibration three times , calibration passed then run QC , QC has come in range , Run samples instrument working fine customer satisfied with results. | SWO-456367 | Assisted Replacement | Electrode | Membrane | Lac |  | Lactate calibration failure | Lactate calibration failed | Checked the activity log , found the error that hypochlorite detected at lactate and glucose , opened the membrane section checked the lactate membrane found liquid in membrane slot , cleaned it ,checked all the membranes ,changed the lactate membrane and then run Calibration three times , calibration passed then run QC , QC has come in range , Run samples instrument working fine customer satisfied with results.
No | Visit | Complaint | RCHN | HCT无法测量 1258 | HCT无法测量 1258 | HCT的排线接触不良，重新连接排线。(by chuanwei) | SWO-455891 | Adjustment | HW | Aspiration Module | Needle Wash Unit | 1258 | HCT cannot measure 1258 | HCT cannot measure 1258 | The HCT has poor line contact and reconnects the row. (by chuanwei)
No | Visit | Complaint | RCHN | 触摸屏不灵敏 | 触摸屏不灵敏, 借910-298一个。 |  | SWO-455924 | Defect | HW | Cover and Chassis | display |  | The touch screen is not sensitive | The touch screen is not sensitive, borrow 910-298 one.
Yes | Visit | Complaint | RCHN | 显示器故障，仪器死机 | 显示器故障，仪器死机,借910-320一个。 |  | SWO-455920 | [5 empty] | Display failure, instrument crash | Monitor failure, instrument crash, borrow 910-320 one.
No | Visit | Complaint | RDEU | Leckage Fehler | Leckage Fehler | aufgef. Verschleißteile ersetzt, kein Erfolg, Wechsel des LP (Lot alt BX-50), Serviceprogramme Fluid Transport i.O., Kal, QC Datensicherung Kal QC | SWO-456677 | Assisted Replacement | SP | SP | FT |  | Leakage error | Leakage error | on the unraised. Wear parts replaced, no success, change of LP (Lot old BX-50), service programs Fluid Transport i.O., Kal, QC Data backup Kal QC
No | Visit | Complaint | RDEU | Oximetrie-Hardwareproblem | Oximetriehardwareproblem | Ersatzplatine zurückgebaut. Neue PCB eingesetzt. | SWO-456638 | Defect | HW | Basic Unit | PCB, wet section |  | Oximetrie-Hardwareproblem | Oximetriehardwareproblem | Replacement board. New PCBs.
No | Visit | Complaint | RLTD | Alexandra - 02034470083 ex 70092 - analyser not turning on, mains checked, black screen | Not switching on. Return to replace power supply. Inlet cover missing on arrival. | replace power supply and test ok. Inlet cover missing on arrival. | SWO-456451 | Defect | HW | Power Supply | Power supply |  | Alexandra - 02034470083 ex 70092 - analyser not turning on, mains checked, black screen | Not switching on. Return to replace power supply. Inlet cover missing on arrival. | replace power supply and test ok. Inlet cover missing on arrival.
No | Visit | Complaint | RDEU | Inhomogene Lösungen und Glu Messprobleme 23.11.2018, Frau Lehrer, Tel: 06931062724: weiterhin inhomogene Lösungen und pH Messprobleme | Clot detection zeigt erhöhten Wert im Sensor Bereich 106 (80-120). Verdacht auf Gerinnsel im Sensor Bay. | Sensor Bay ausgetauscht, Startphase i.O, Clot detektion liegt jetzt bei 102, alle Kal und QCs bestanden. Gerät ausgeliefert. | SWO-456916 | Defect | HW | Basic Unit | Sensor Bay Unit | Clot | Inhomogeneous solutions and Glu measurement problems 23.11.2018, Mrs. Lehrer, Tel: 06931062724: further inhomogeneous solutions and pH measurement problems | Clot detection shows increased value in sensor range 106 (80-120). Suspicion of clots in Sensor Bay. | Sensor Bay replaced, start phase i.O, Clot detection is now 102, all Kal and QCs passed. device.
No | Visit | Complaint | RPAC | pO2 calibration failures. Error code 1006 | Po2 calibration failure | Travelled to site. Verified error 1006. Po2 energy at 2.24. Changed Sensor interface board - performed complete oxi test- reading po2 energy 6. Tested again - all green. | SWO-456092 | Defect | HW | Basic Unit | PCB, Sensor interface | pO2, 1006 | pO2 calibration failures. Error code 1006 | Po2 calibration failure | Travelled to site. Verified error 1006. Po2 energy at 2.24. Changed Sensor interface board - performed complete oxi test- reading po2 energy 6. Tested again - all green.
No | Visit | Complaint | RCHN | 无法开机 | 无法开机 借903-663一个 | 更换备件903-663后发现机器还是无法使用，需要加借903-381一个by Andrew | SWO-456038 | [5 empty] | Cannot power on | Can't turn on, borrow 903-663 one. | Replacement of spare parts 903-663 found that the machine is still unusable, need to borrow 903-381 a by Andrew
No | Visit | Complaint | RCHN | 无法开机 | 无法开机，需要加借903-381一个 | 更换903-381后还是无法使用，需要加借903-228by Andrew | SWO-456078 | [5 empty] | Cannot power on | Cannot power on, need to borrow 903-381 one | Replacement 903-381 is still not available, need to borrow 903-228by Andrew
No | Visit | Complaint | RCHN | 无法开机 | 需要加借903-228 | 更换903-228故障未排除，加借903-334一个。 by Andrew | SWO-456156 | [5 empty] | Cannot power on | Need to borrow 903-228 | Replace the 903-228 fault not corrected, plus 903-334 one. by Andrew
No | Visit | Complaint | RCHN | 无法开机 | 更换903-228故障未排除，加借903-334一个。 | 更换903-334后修复 by Andrew | SWO-456261 | [5 empty] | Cannot power on | Replace the 903-228 fault not corrected, plus 903-334 one. | Repair by Andrew after replacing 903-334
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | pco2,po2定标失败，0378，0375 | 更换测试卡后修复 | SWO-456206 | Assisted Replacement | SP | SP | pCO2/pO2 |  | Test card, reagent package failure | pco2, po2 calibration failure, 0378,0375 | Repair after replacing test card
No | Visit | Complaint | RFRA | Sauvegarde systèmes plein | Calibration ok secretariat.drmenivale@orange.fr | Contrôle du zéro ok. Calibration toujours ok Problème erreur de résultats, non constaté. Pas d'ampoules de contrôle sur site pour vérifier. | SWO-456505 | [5 empty] | Backing up full systems | Calibration ok secretariat.drmenivale@orange.fr | Zero control ok. Calibration always ok Problem results error, not found. No on-site control bulbs to check.
No | Visit | Complaint | RAME | Neon Oximetry errors not allowing claibrations to complete | Verified customer's stated issue. | Removed the oxi module and checked the neon lamp cable and connector. Both cable and connector show no visible damage. Unplugged the unit from the outlet for 2 minutes and rebooted the unit. The unit rebooted and within a few seconds after initializing the error returned. Will order a lamp and oxi module PCB. | SWO-456719 | Revisit required | HW | Oxi module | Lamp unit | Neon lamp | Neon Oximetry errors not allowing claibrations to complete | Verified customer's stated issue. | Removed the oxi module and checked the neon lamp cable and connector. Both cable and connector show no visible damage. Unplugged the unit from the outlet for 2 minutes and rebooted the unit. The unit rebooted and within a few seconds after initializing the error returned. Will order a lamp and oxi module PCB.
No | Visit | Complaint | RSCH | Printer Zieht Papier nicht korrekt | Drucker zieht Papier immer schräg ein und verchruglet | Kein Fehler festgestellt Prophylaktisch Printer Kopf ausgetauscht Funktionskontrolle i.o | SWO-456309 | Defect | HW | Printer Module | Printer Head |  | Printer Pulls paper incorrectly | Printer always pulls paper in obliquely and verchruglet | No error detected Prophylactic Printer Head Replaced Functional control i.o
No | Visit | Complaint | RSAF | coox parameters not reading | Neon lamp intensity too low Neon number too low Neon position : -0.1 | Replaced neon lamp, and complete oxi test ok. 2pt cal tHb cal and qc ok. | SWO-456289 | Defect | HW | Oxi module | Lamp unit | Neon lamp | coox parameters not reading | Neon lamp intensity too low Neon number too low Neon position : -0.1 | Replaced neon lamp, and complete oxygen test OK. 2pt Cal tHb Cal and QC OK.
No | Visit | Complaint | RHUN | a készülék időnként újraindítást kér, a hiba ekkor megszűnik, majd hamarosan újra jelentkezik | storage wheel szerviz funkciók működnek, a készülék belsejében portalanítás szükséges | készülék belső tisztítása, szerviz tesztek elvégzése, init | SWO-456304 | Cleaned | HW | Storage Module | storage wheel, hub | 5694, 5784 | The device occasionally asks for a reboot, the error will disappear, and you will soon be re-logged | Storage Wheel service functions, the device must be dusting inside | Internal cleaning of the appliance, performing service tests, init
No | Visit | Complaint | RDEU | 26.11.2018 Frau Kathrin Götz 07351551226 meldet, dass der ABL immer wieder das LP wechseln soll. | Gerät überprüft, Leakage erkannt.LS Fehler.uam | Gerät zerlegt. Stenose in Hemolysator festgestellt. Pleura? Inlet Probe erneuert. Hemolysator ersetzt. Justage und Überprüfung durchgeführt. | SWO-457200 | Defect | HW | Oxi module | Hemolyzer | Clot | 26.11.2018 Ms. Kathrin Götz 07351551226 announces that the ABL is to change the LP again and again. | Device checked, leakage detected. LS Error.uam | device disassembled. stenosis in hemolysator. Pleura? Inlet sample renewed. Hemolysator. adjustment and verification.
No | Visit | Complaint | RDEU | Seit einigen Tagen trotz Spülungen und Tausch der Betriebsmittel verschiedene Oxi Hardware Fehler | Hemolyzer defekt | Fehlersuche, nochmals System gespült, und korrekte Installation der Betriebsmittel geprüft, der Kunde achtet auf die Präanalytik!, LP wurde bereits erneuert, Oxi Complete Test bei Hemolyzer durchgefallen, Überprüfung der Verbindung zu Spektrometer und Hemolyzer ok, Tausch des Hemolyzers, Neustart Startphase und Oxi Complete Test wird jetzt sofort bestanden Kal. und QC wieder ok, Blutmessung ohne Fehler | SWO-457268 | Defect | HW | Oxi module | Hemolyzer |  | For several days despite flushing and replacing the equipment various Oxi hardware errors | Hemolyzer defective | Troubleshooting, again system flushed, and correct installation of the equipment checked, the customer pays attention to the preanalysis!, LP has already been renewed, Oxi Complete Test at Hemolyzer failed, checking the connection to spectrometer and hemolyzer ok, exchange of the Hemolyzer, restart start phase and Oxi Complete test is now passed immediately Kal. and QC ok again, blood measurement without error
No | Visit | Complaint | RLTD | Please call Mark Spector 07939 069137 | Screen frozen with intervention required with log box in the middle. Logon box works and there are no other buttons. | Back up data and reloaded analyser software reinstall data. Analyser started up ok and now asking for cassette to be replaced as no more test available. New cassette in stalled start up ok . Calibration and QC ok. Rebooted analyser a couple of times starts up ok. | SWO-457462 | Defect | SW | RMED SW | RMED SW | re-install | Please call Mark Spector 07939 069137 | Screen frozen with intervention required with log box in the middle. Logon box works and there are no other buttons. | Back up data and reloaded analyser software reinstall data. Analyser started up ok and now asking for cassette to be replaced as no more test available. New cassette in stalled start up ok . Calibration and QC ok. Rebooted analyser a couple of times starts up ok.
No | Visit | Complaint | RDEU | 26.11.2018 15:45 Frau Anic 08912183306 -> Leckage bzw. Lösungen können nicht gepumpt werden 27.11.2018 08:12 Frau Anic 08912183306 Immer wieder Analysator spülen | Gerät steht in Intervention erforderlich Oxi Ventilschlauch undicht | Gerät manuell gespült Oxi Ventilschlauch erneuert ABL kalibriert und überprüft | SWO-457223 | Defect | HW | Oxi module | Pinch valve tube |  | 26.11.2018 15:45 Ms Anic 08912183306 -> Leakage or solutions cannot be pumped 27.11.2018 08:12 Ms Anic 08912183306 Rinse again and again | Device is required in intervention Oxi valve hose leaks | Device manually rinsed Oxi valve hose renewed ABL calibrated and checked
No | Visit | Complaint | RAME | No power | No power | Replaced the power supply | SWO-457567 | Revisit required | HW | DMS_PC | CPU |  | No power | No power | Replaced the power supply
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 气泡探测定标失败 | 更换测试卡后修复 | SWO-456961 | [5 empty] | Test card, reagent package failure | Bubble probe calibration failed | Repair after replacing test card
No | Visit | Complaint | RAME | Received random samples errors R/S, 292 and 581 | Unit failing cavitation absorb portion of service test. | Performed hemolyzer status update and rechecked service test results. All tests passed. | SWO-457478 | [5 empty] | Received random samples errors R/S, 292 and 581 | Unit failing cavitation absorb portion of service test. | Performed hemolyzer status update and rechecked service test results. All tests passed.
No | Visit | Complaint | RCHN | 进样支架断裂 | 进样支架断裂，需要借件903-584*2 | 更换903-584后修 by Andrew | SWO-456299 | Defect | HW | Aspiration Module | Tube holder | Potentially User Error | The sample bracket breaks | The sample bracket is broken and requires a loan of 903-584 x 2 | Replacement 903-584 repair by Andrew
No | Visit | Complaint | RESP | El cliente llama indicando que el AQT da el error 5784. Storage Wheel no funciona correctamente. | Chequeo de equipo, no pasa el test de safe position mostrando error 5784 Tras reinicio no entra en servicio quedándose bloqueado | Limpieza de circuito neumático y realización de kit Carga de software completo Sustitución de Cf card y carga completa de software | SWO-457164 | Cleaned | HW | Storage Module | storage wheel, hub | 5784 | The client calls indicating that the AQT gives error 5784. Storage Wheel is not working properly. | Team check, does not pass the safe position test showing error 5784 After restart does not go into service being blocked | Pneumatic circuit cleaning and kit ingescing Full software upload Cf card replacement and full software upload
No | Visit | Complaint | RCHN | 处理包上的PCB板损坏 | 处理包上的PCB板损坏，借903-036一个 | 更换903-036后修复 by Andrew | SWO-456118 | Defect | HW | Cover and Chassis | PCB, solution pack | SP suspected to be RC | Dealing with PCB board damage on package | Damage to the PCB board on the handling package, borrow 903-036 one | Repair by Andrew after replacement 903-036
Yes | Visit | Complaint | RBNL | Errormelding 1029 op DDIM meting. | Wederom errorcode 1029 op monstermetingen. Aangezien het een wederkerend probleem is, volgens Trouble shooting tree de optical unit vervangen. Na vervangen factory reference calibration uitgevoerd en alle testen gecalibreerd. Hierna QC gemeten --> oke. AQT meetklaar. | Optical unit vervangen. | SWO-457109 | [5 empty] | Error message 1029 on DDIM metering. | Again error code 1029 on sample measurements. Since it is a reconsidering problem, according to Trouble shooting tree the optical unit is replaced. After replacing factory reference calibration performed and all the tests calibrated. Hereafter QC measured--> okay. AQT measuring ready. | Optical Unit.
Yes | Visit | Complaint | RFRA | Contrôle système journalier en échec des l’installation. Erreur 1031 intensité optique hors gamme = 12,5 (de -10 à 10) | Vérification AQT90 et nettoyage. | Vérification ok. Nettoyage du module optique ok. Contrôle système En Échec. Même erreur. Erreur 1031 intensité optique hors gamme = 13,1 pour une limite entre -10 à 10. Demande de changement AQT90. Récupération ServiceDump ok. Et création techline. Rapport à envoyer par mail à Beaucaire@labosud.fr | SWO-457282 | Re classified RMED | installation report |  |  | Move to OOB trending | Check the installation's failed daily system. Error 1031 off-range optical intensity - 12.5 (from -10 to 10) | AQT90 check and cleaning. | Check ok. Cleaning the optical module ok. System Control In Failure. Same mistake. Error 1031 off-range optical intensity - 13.1 for a limit between -10 to 10. Request for change AQT90. Recovery ServiceDump ok. And techline creation. Report to be emailed to Beaucaire@labosud.fr
No | Visit | Complaint | RDEU | 29.10.2018 14:00 Fr. Richter ????: Interventionsmeldung Gerinsel in der Oximetrie 20.11.2018 14:20Uhr Dr.Holfertholfert 0352365225: Das Gerät meldet weiterhin Fehler Intervention erforderlich | Gerät geht erneut in geführte Fehlersuche mit Meldung 1352 Gerinnsel in Oxymetriesystem erkannt Kunde hat Gerät bei auftreten des Fehlers auch schon aus $eingeschaltet -> danach 5min Warmstart und Gerät war wieder messbereit bei eintreffen steht Gerät in Halt und fordert zur Spülung auf | Austausch der Sensor Bay Unit & des Liquide Sensor im Oxymodul, alle Daten gesichert incl. Service Dump, Rückmeldung technischer Probleme erstellt und an TS geschickt Startphase von 50 min abgewartete, bei 35 min Reststartphase schaltet Gerät in QC Messung, führt alle 3 Level durch und ist danach messbereit, Konfig. des Nachbargerätes gesichert und in Gerät eingespielt Kal & QC Daten gemäß RDEU Anweisung gesichert & dokumentiert, | SWO-458099 | Defect | HW | Basic Unit | Sensor Bay Unit | and LS 3 | 29.10.2018 14:00 Fr. Richter ????: Intervention report Gerinsel in oximetry Nov 20, 2018 2:20pm Dr.Holfertholfert 0352365225: The device continues to report error intervention required | Device goes back into guided troubleshooting with message 1352 clots detected in oxymetry system Customer has device even from /eingeschaltet -> after then 5min warm start and device was ready to measure again when the device arrives, it stops and prompts for flushing | Replacement of the Sensor Bay Unit & the Liquide Sensor in the Oxymodule, all data backed up incl. service dump, feedback of technical problems created and sent to TS Start phase of 50 min waited, at 35 min remaining start phase switches device in QC measurement, performs all 3 levels and is then ready to measure, Config. secured by the neighboring device and saved into the device Kal & QC data secured & documented in accordance with RDEU instructions,
No | Visit | Complaint | RLTD | Battery error detected 1559 Stacey Joyce 01708 43500 ext 2621 26.11.18 Telephone call from Sean Quinlivan re the battery error code 1559, please can you update him on progress. 27.11.18 Sean Quinlivan has called again today to advise that they are now locked out of the analyser, please call him on 01708 435 000 x: 2621 | Confirmed battery is faulty. | Replaced battery, checked charge cycle and cells. Analyser is ready for use. | SWO-457877 | Defect | HW | Battery module | Battery module | 1559 | Battery error detected 1559 Stacey Joyce 01708 43500 ext 2621 26.11.18 Telephone call from Sean Quinlivan re the battery error code 1559, please can you update him on progress. 27.11.18 Sean Quinlivan has called again today to advise that they are now locked out of the analyser, please call him on 01708 435 000 x: 2621 | Confirmed battery is faulty. | Replaced battery, checked charge cycle and cells. Analyser is ready for use.
No | Visit | Complaint | RESP | Error cal y qc en po2 Solucion de lavado no homogenea | Error cal y qc en po2 Solucion de lavado no homogenea | Limpieza del circuito Sustitucion de inlet gasket holder y gasket conector propiedad del cliente Chequeo con test de servicio Cambio de sp Cambio de sc Cal, qc | SWO-458213 | Assisted Replacement | SC | SC | pO2 | 521, 588, 979 | Error lime and qc in po2 Non-homogeneous washing solution | Error lime and qc in po2 Non-homogeneous washing solution | Cleaning the circuit Customer-owned inlet gasket holder and gasket connector replacement Check with service test Changing sp Changing sc Cal, qc
No | Visit | Complaint | RAME | Status 2 error/Battery issue | Error code 2 $Battery issue | Replaced software two times $ Replaced battery | SWO-459287 | Defect | SW | RMED SW | RMED SW | status code 2 | Status 2 error/Battery issue | Error code 2 /Battery issue | Replaced software two times $ Replaced battery
No | Visit | Complaint | RAME | In homogeneous errors | Getting gas 2 error - will not calibrate. | Ran sensor simulator test - passed. Replaced SP and ran calibrations. All tests passed. | SWO-459107 | Assisted Replacement | SP | SP | FT |  | In homogeneous errors | Getting gas 2 error - will not calibrate. | Ran sensor simulator test - passed. Replaced SP and ran calibrations. All tests passed.
No | Visit | Complaint | RAME | Failed proficiency testing. | Analyzer failed proficiency testing for pCO2. Need to run VK kit | Ran zeroing cassette on analyzer and verified proper operation on all channels. No issues found with zeroing cassette. Customer ran linearity and reported instrument is good. | SWO-458217 | [5 empty] | Failed proficiency testing. | Analyzer failed proficiency testing for pCO2. Need to run VK kit | Ran zeroing cassette on analyzer and verified proper operation on all channels. No issues found with zeroing cassette. Customer ran linearity and reported instrument is good.
No | Visit | Complaint | RCHN | 孵育盘故障，报错3557 | 孵育盘故障，报错3557,需要借件903-138一个 | 更换903-138后修复 by Andrew | SWO-456115 | Defect | HW | Incubator Module | PCB, incubator | 3557 | Incubation plate failure, error 3557 | Incubation plate failure, reported error 3557, need to borrow 903-138 one | Repair by Andrew after replacing 903-138
No | Visit | Complaint | RCHN | 孵育盘温度传感器故障，报错3564 | 孵育盘温度传感器故障，报错3564，借902-478一个 | 更换902-478一个正常。 by Andrew | SWO-456954 | Defect | HW | Incubator Module | Thermistor | 3564 | Incubation plate temperature sensor failure, error 3564 | Incubation disk temperature sensor failure, reported error 3564, borrowed 902-478 one | Replace 902-478 one normal. by Andrew
Yes | Visit | Complaint | RCHN | 光学强度超范围 | 光学强度超范围，借903-364一个 | 更换903-364后修复by jesson | SWO-456155 | [5 empty] | Optical strength over range | Optical strength over range, borrowed 903-364 a | Repair by jesson after replacement 903-364
No | Visit | Complaint | RCHN | 定标失败 | 定标失败，报错1029，借944-268（14707）一张 缺货 | 问其他经销商借了一张定标卡，重新定标后修复 by Andrew | SWO-456089 | Defect | SW | Process Module | Process Wheel |  | Bid failed | Failed to bid, error 1029, borrow 944-268 (14707) a out of stock | Ask the other dealer to borrow a calibration card to fix by Andrew after re-calibration
No | Visit | Complaint | RSCH | Rep Power Supply Neugerät | Gerät startet nach AQURE Installation nicht mehr | Diverse Test ausgeführt CPU Unit ausgetauscht ohne Erfolg $Zurückgetauscht (Sicherungen 5 A auf AU PCB plötzlich defekt. Austausch ohne Erfolg Power Supply ausgetauscht (grosser Kondensator aufgebläht) Geräte Initialisierung i.o | SWO-458090 | Defect | HW | Power Supply | Power supply |  | Rep Power Supply New Device | Device no longer starts after AQURE installation | Various test executed CPU Unit replaced without success $Zurück (Fuses 5 A to AU PCB suddenly defective. Exchange without success Power Supply replaced (large capacitor inflated) Devices Initialization i.o
No | Visit | Complaint | RDAN | Marianne klager over oxi fejl: 251+579 | Afprøvet, halogen lampe intensitet for høj. | Halogen lampe justeret, slutafprøvet. | SWO-457856 | Adjustment | HW | Oxi module | Lamp unit | halogen, 251, 579 | Marianne complains of the OXI error: 251 + 579 | Tested, halogen lamp intensity too high. | Halogen lamp adjusted, final tested.
No | Visit | Complaint | RLTD | Sean Quinlivan 01708 435000 x6672 |  |  | SWO-457887 | Adjustment | HW | Cover and Chassis | Bracket, Solution Pack |  | Sean Quinlivan 01708 435000 x6672
No | Visit | Complaint | RPAC | Flow error 1278 | Flow error 1278, forgot to add replacement pump head | Replaced pump head 1 on wetsection during annual pm | SWO-459193 | Defect | HW | Wet section | Pump head 1 | 1278 | Flow error 1278 | Flow error 1278, forgot to add replacement pump head | Replaced pump head 1 on wetsection during annual pm
No | Visit | Complaint | RDEU | Dunkelstromfehler bei Messung | Dunkelstromfehler bei Messung. | Optische Einheit erneuert. Alle installierten Parameter kalibriert. Qualitätskontrollen aller Parameter durchgeführt. | SWO-458246 | Defect | HW | Process Module | Optical Unit | 1035 | Dark current error during measurement | Dark current error during measurement. | Optical unit renewed. All installed parameters calibrated. quality control of all parameters carried out.
No | Visit | Complaint | RPOL | Błąd : analizator nie podłączony do DMS | Analizator po samoczynnym restarcie nie podłącza się do DMS | Wymiana karty CF. Reinstalacja Windows XP. | SWO-458193 | Defect | HW | DMS_PC | CF card |  | Error: Analyzer not connected to DMS | The automatic Restart Analyzer does not connect to DMS | Wymiana karty CF. Reinstalacja Windows XP.
No | Visit | Complaint | RPOL | Problem przepływowy. | Diagnoza problemu. | Oczyszczenie toru pomiarowego oraz krućca ściekowego. Testy serwisowe. Przygotowanie aparatów do przeniesienia. | SWO-458842 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | Flow Problem. | Diagnosis of the problem. | Cleaning of the measuring track and the sewage aggregate. Service tests. Preparation of the apparatus to be transferred.
No | Visit | Complaint | RDEU | 29.11.2018 8:00 Fr. Lahme 02418080946: Ableitstrom erkannt | Sichtkontrolle | Solutionpack Kontakte gereinigt, Einlassnadel ersetzt da geknickt, Lösungspack nochmals ersetzt da Leakvoltage 600mV mit neuem SP <200mV anschl. Update und Wartung | SWO-459125 | Assisted Replacement | SP | SP | Leak current |  | 29.11.2018 8:00 Fr. Lahme 02418080946: Leakflow detected | Visual inspection | Solutionpack contacts cleaned, inlet needle replaces there bent, solution pack replaced again since Leakvoltage 600mV with new SP 200mV connected. Update and maintenance
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 温度-B | 更换测试卡后修复 | SWO-467598 | [5 empty] | Test card, reagent package failure | Temperature-B | Repair after replacing test card
No | Visit | Complaint | RKK | アプリケーションエラー。 Romaシャットダウン。 | データベース圧縮の際アプリケーションエラー発生しRomaシャットダウン。 | CFカード交換実施。 言語変更、設定復元、CBM設定、キャル、QC、チケット印字確認・・良好 | SWO-458710 | [5 empty] | Application error. Roma shutdown. | Roma shutdown occurs in application error during database compression. | CF card exchange was implemented. Language change, setting restoration, CBM setting, cal, QC, ticket print confirmation, good
No | Visit | Complaint | RCHN | 电脑点不进去，无法开机 | 电脑点不进去，无法开机，借910-342一个 | 更换910-342后修复 by Andrew | SWO-457743 | Defect | HW | DMS_PC | CPU |  | The computer doesn't go in, can't turn it on. | Computer dot do not go in, can not turn on, borrow 910-342 one | Fix by Andrew after replacing 910-342
No | Visit | Complaint | RIND | Solution leakage. | Solution leake from above SP. | Found Sensor cassette got clotted due to that solution pump tube was opened form SC input because of this SP opt sensor and SP Spring board assemble also got damaged need to replace with new, customer is unhappy with analyzer performance, suggested us take back analyzer or provide new analyzer due to frequent brake down. We have taken approval for new analyzer by end of this week we will install new analyzer. | SWO-459984 | [5 empty] | Solution leakage. | Solution leake from above SP. | Found Sensor cassette got clotted due to that solution pump tube was opened form SC input because of this SP opt sensor and SP Spring board assemble also got damaged need to replace with new, customer is unhappy with analyzer performance, suggested us take back analyzer or provide new analyzer due to frequent brake down. We have taken approval for new analyzer by end of this week we will install new analyzer.
No | Visit | Complaint | RFRA | Echec de la calibration même après sortie mode Attente Cassette suspectée defectueuse En attente de disponibilité du client pour effectuer quelques tests | Erreur de calibration de la détection d’air | Plusieurs calibration et toujours même erreur. Lancement rinçage ok. Menu $activité $ rinçage. Calibration ok. ABL9 prêt ok plus d’erreur. Récupération des données pour remontés des infos via Techline ok. Problème à suivre . Procédure de dédommagement pour 1cassette et 1pack solution change à la demande du SAV ok. | SWO-459073 | [5 empty] | Failed calibration even after output Waiting mode Suspected defector tape Waiting for customer availability to perform some tests | Air detection calibration error | Several calibration and always same error. Launch rinse ok. Menu $activited $ rinse. Calibration ok. ABL9 ready ok more error. Recovering data to re-raise information via Techline ok. Problem to follow. Compensation procedure for 1cassette and 1pack solution changes at the request of the SAV ok.
No | Visit | Complaint | RLTD | Unable to download data. David Ruaux reported 11.10.18. |  |  | SWO-458960 | [5 empty] | Unable to download data. David Ruaux reported 11.10.18.
No | Visit | Complaint | RAME | Neon Oximetry errors not allowing claibrations to complete | Follow up to SWO-456719. Neon Oximetry errors not allowing claibrations to complete | Replaced lamp and performed all necessary adjustments in the service menu. | SWO-458421 | Defect | HW | Oxi module | Lamp unit | Neon lamp | Neon Oximetry errors not allowing claibrations to complete | Follow up to SWO-456719. Neon Oximetry errors not allowing claibrations to complete | Replaced lamp and performed all necessary adjustments in the service menu.
Yes | Visit | Complaint | RSCH | JW | ZW2 Rinnt aus Einlass trotz neuer Einlassdichtung | Zwischenwartung 2 ausgeführt tHb Kal Funktionskontrolle i.o | SWO-458657 | Re classified RMED | same day upSWO |  |  |  | JW | ZW2 Rinnt from inlet despite new inlet seal | Intermediate Maintenance 2 Executed tHb Kal Functional control i.o
No | Visit | Complaint | RSCH | Einlass undicht | Rinnt aus Einlass beim Spülen | Friction Detector ausgetauscht Einlassdichtung ausgetauscht Inlet Modul Kalibriert Funktionskontrolle i.o | SWO-458658 | Defect | HW | Inlet module | Friction detection |  | Entrance leaks | Rinnt from inlet during rinsing | Detector Friction replaced Inlet seal replaced Inlet Module Calibrated Functional control i.o
No | Visit | Complaint | RCHN | 两点定标，po2 cthb没有值 | 两点定标，po2 cthb没有值 0315,需要借件903-125和902-458 | 经确认不是903-125和902-458的问题，加借902-833一个。by Andrew | SWO-456906 | Revisit required | HW | Oxi module | Oxi module | 0315, 579, 310 | Two-point calibration, po2 cthb has no value | Two-point calibration, po2 cthb has no value of 0315, need to borrow 903-125 and 902-458 | It was confirmed that it was not a question of 903-125 and 902-458, plus 902-833. by Andrew
No | Visit | Complaint | RESP | Error FTC101 | Chequeo de neumática y SP | Limpieza a fondo de equipo, Wash block unit Realización de kit anual | SWO-460484 | Cleaned | HW | Aspiration Module | Needle Wash Unit |  | Error FTC101 | Pneumatic check and SP | Thorough team cleaning, Wash block unit Realization of annual kit
No | Visit | Complaint | RSCH | Rep Scanner | Test können nicht nachgefüllt werden da Analysator Barcode nicht lesen kann | Barcodereader gereinigt, konnte alle Test einwandfrei lesen Softwareupgrade V 8.11.300.0 ausgeführt Datenbackup ausgeführt Funktionskontrolle i.o | SWO-458712 | Cleaned | HW | Storage Module | Barcode reader, storage wheel |  | Rep Scanner | Test cannot be refilled because analyzer cannot read barcode | Barcode reader cleaned, could read all test correctly Software upgrade V 8.11.300.0 executed Data backup performed Functional control i.o
Yes | Visit | Complaint | RBNL | Defecte accu | Defecte batterij module wegens error 1559 heb vervangen. | Nieuwe batterij module heb geinstalleerd. | SWO-459851 | [5 empty] | Defective battery | Defective battery module Due to error 1559 have replaced. | New battery module installed.
No | Visit | Complaint | RAME | Customer getting battery errors. Bad Cell errors. | Had customer shut off analyzer and unplug from wall. Errors didn't go away. | Replaced battery. Performed all necessary start up calibrations. | SWO-459889 | Defect | HW | Battery module | Battery module |  | Customer getting battery errors. Bad Cell errors. | Had customer shut off analyzer and unplug from wall. Errors didn't go away. | Replaced battery. Performed all necessary start up calibrations.
No | Visit | Complaint | RAME | No power | Still powering down | Replaced computer module $Installed sensor | SWO-459288 | Defect | HW | DMS_PC | CPU |  | No power | Still powering down | Replaced computer module $Installed sensor
No | Visit | Complaint | RCHN | 接口板腐蚀 | 接口板腐蚀，借902-972一个 | 更换902-972后修复 by Andrew | SWO-459320 | [5 empty] | Interface plate corrosion | Interface plate corrosion, borrow 902-972 one | Repair by Andrew after replacement 902-972
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | LAC定标响应错误 1310 | 更换测试卡后修复 | SWO-460224 | Assisted Replacement | SC | SC | Lac | 1310 | Test card, reagent package failure | LAC calibration response error 1310 | Repair after replacing test card
No | Visit | Complaint | RCHN | 测试卡无法扫描 | 测试卡无法扫描，借903-112,910-304各一个 | 更换903-112,910-304后修复 by Andrew | SWO-458579 | Defect | HW | Storage Module | Barcode reader, storage wheel |  | Test card can't be scanned | Test cards cannot be scanned, borrowing 903-112, 910-304 each | Repair by Andrew after replacing 903-112, 910-304
Yes | Visit | Complaint | RCHN | 试管夹断裂 | 试管夹断裂，借903-584一个 | 更换903-584后修复 by Andrew | SWO-459330 | [5 empty] | Test tube clip break | Test tube clip broken, borrow 903-584 one | Repair by Andrew after replacing 903-584
No | Visit | Complaint | RDEU | Elyte Driftfehler | BG Messkammer stark verschmutzt | BG Messkammer demontiert und gereinigt, Elektroden geprüft und gereinigt, Funktionskontrolle ok | SWO-459562 | Cleaned | HW | pH/BG Module | pH/BG Module | electrolyte drift | A. M. Driftfehler | BG measuring chamber heavily polluted | BG measuring chamber dismantled and cleaned, electrodes tested and cleaned, function control ok
No | Visit | Complaint | RDAN | Birgit klager over pco2 på qc som blafre. | Undersøgt. | Pakning på pH elektrode udskiftet, luftbobler frigjort i pco2 membranen, slutafprøvet. | SWO-459547 | Defect | Electrode | Electrode | Sealing ring, pH |  | Birgit complains about pCO2 at QC as Blafre. | Examined. | Packaging on the pH electrode replaced, air bubbles released in the pCO2 membrane, final tested.
No | Visit | Complaint | RDEU | 30.11.2018 Frau Gerschau 017616658311 meldet, dass die meisten Proben abgebrochen werden. | 3553: Incubatorwheelmove 31 | Punch out Sensor gereinigt - Motor geprüft Incubator Wheel, inkl. Kugellager und Sensoren gereinigt | SWO-459790 | Cleaned | HW | Incubator Module | Incubator wheel | 3553 | 30.11.2018 Ms. Gerschau 017616658311 reports that most samples are cancelled. | 3553: Incubatorwheelmove 31 | Punch out sensor cleaned - motor tested Incubator Wheel, incl. ball bearings and sensors cleaned
No | Visit | Complaint | RSCH | AU Connect Failed | Fehler 2132: AUCONNECTTIMEOUT 2995: CLOSESTORAGE12 Inkonsistente SW | Gerät von Notfall ins Labor mitgenommen Systemtest ohne Fehler AU PCB ausgebaut und alle Sicherungen ausgetauscht (Goldkontakt) Funktionskontrolle i.o Gerät im Notfall wieder installiert | SWO-459505 | Defect | HW | Cover and Chassis | PCB, Main Board | 2132, 2995, fuses | AU Connect Failed | Error 2132: AUCONNECTTIMEOUT CLOSESTORAGE12 2995 Inconsistent SW | Device taken to the laboratory by emergency System test without errors AU PCB removed and all fuses replaced (gold contact) Functional control i.o Device reinstalled in case of emergency
No | Visit | Complaint | RCHN | 定标中止 | 定标中止，借944-220（15065）一张 |  | SWO-460008 | Assisted Replacement | Consumables | Test Kit | Cartridge, cal | 5257 | Bid abort | Bid suspension, borrow 944-220 (15065) one
No | Visit | Complaint | RCHN | 两点定标，po2 cthb没有值 报错0579、0310 | 氖灯数值偏高 0315。借902-833一个。 | 更换902-833后修复 by Andrew | SWO-458611 | Defect | HW | Oxi module | Oxi module | 0315, 579, 310 | Two-point calibration, po2 cthb no value error 0579, 0310 | The lamp value is 0315 higher. Borrow 902-833 one. | Repair by Andrew after replacement 902-833
No | Visit | Complaint | RCHN | GLU定标失败，没有数值 | GLU定标失败，没有数值，借903-581一个 |  | SWO-460093 | Defect | HW | Basic Unit | PCB, Sensor interface | Glu | GLU calibration failed, no value | GLU calibration failed, no value, borrow 903-581 one
No | Visit | Complaint | RAME | Multiple samples are flagging pH, PCO2, Ca, Cl | Customer getting flagged results. Also getting error 1070 response error. | Ran sensor simulator test Added greater &less than command line in the Roma file. Ran calibrations. | SWO-460854 | Setup | SW | RMED SW | Roma.ini | Reportable range | Multiple samples are flagging pH, PCO2, Ca, Cl | Customer getting flagged results. Also getting error 1070 response error. | Ran sensor simulator test Added greater &less than command line in the Roma file. Ran calibrations.
No | Visit | Complaint | RESP | Falló en transporte de fluidos y fuga detectada | El equipo da error de transporde de fluidos. Se aprecia que el SP está defectuoso. | Comprobado reactivo con otros dos equipos. Se cambia el SP, SC y se calibra la hemoglobina. Equipo correcto. | SWO-460597 | Assisted Replacement | SP | SP | FT |  | Failed in fluid transport and leak detected | The equipment gives fluid transporde error. It is appreciated that the SP is faulty. | Tested reactive with two other equipment. The SP, SC is changed and the hemoglobin is calibrated. Right equipment.
No | Visit | Complaint | RCHN | 更换电极膜后，定标仍无法通过，提示电极换膜 | 更换电极膜后，定标仍无法通过，提示电极换膜 借945-737一个 | 更换945-737一个正常。 | SWO-460360 | [5 empty] | After replacing the electrode film, the calibration still can not pass, prompting the electrode to change the film | After replacing the electrode film, the calibration still can not pass, prompt the electrode replacement film borrowed 945-737 a | Replace the 945-737 one normal.
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | LAC定标不过 | 更换测试卡后修复 | SWO-463257 | Assisted Replacement | SC | SC | Lac |  | Test card, reagent package failure | LAC calibration is not | Repair after replacing test card
Yes | 3 | Y | RCZE | Blank |  |  | D552A5B8D25CB85D | Re classified RMED | planned SWO |  |  | Phase 0, see e-mail | Blank
Yes | 9 | Y | RCZE | Blank | V zmysle instrukcii pre phase 0 |  | D552C92ED254A7E4 | Re classified RMED | planned SWO |  | Phase 0 | Lamp Unit | Blank | In terms of INSTRUCCII for phase 0
No | Visit | Complaint | RSCH | AU | Mehrmals Täglich Fehler 1278: Gemessene Flüssigkeitsmenge lag oberhalb der Warngrenze Analysator fordert mehrmals am Tag Reinigung Windows Blockiert z.t | Komplettes Datenbackup erstellt Softwareupgrade 8.11.300 ausgeführt AU Main Board ausgetauscht Analyzer Software nocheinmal installiert Windows neu installiert Konfiguration wieder hergestellt Reinigung ausgeführt $ Counter auf 0 Funktionskontrolle i.o | SWO-460459 | Cleaned | HW | Wet section | Wet section | Flow restriction; 1278 | The | Several times a day error 1278: Measured amount of liquid was above the warning limit Analyzer requires cleaning several times a day Windows Blocks z.t | Complete data backup created Software upgrade 8.11.300 executed AU Main Board replaced Analyzer software reinstalled Windows reinstalled Configuration restored Cleaning done - Counter to 0 Functional control i.o
Yes | Visit | Complaint | RDEU | 01.12.2018 14:30, RR Pfl. Buluar 05931-1521293, medizinische Aufnahme, Fehler 2168, Fehler in Startvorgang, zuvor ist bereits eine TropT Kalibrierung (alte Charge war verbraucht, neue Charge LOT 15133) mit Fehlermeldung 1061 gescheitert, Geräteneustart wurde 3x durchgeführt ohne Erfolg, FSE Einsatz erforderlich | System in Fehler 2168 calibrate 42 | Cups werden nicht in Abfallbehälter transportiert, Klappe öffnet sich nicht, Abfallbehälter erneuert und alle cups entfernt Testmessungen siehe HJW | SWO-461144 | Defect | Consumables | Solution Pack | Solution Pack | Stack | 01.12.2018 14:30, RR Pfl. Buluar 05931-1521293, medical recording, error 2168, error in startup process, previously a TropT calibration (old batch was used, new batch LOT 15133) failed with error message 1061, device restart was performed 3x without success, FSE use required | System in Fehler 2168 calibrate 42 | Cups are not transported in waste containers, flap does not open, waste bin is renewed and all cups removed test measurements see HJW
Yes | Visit | Complaint | RPOL | Analizator zastępczy Chrzanów Laboratorium. | Deinstalacja analizatora zastępczego. | Deinstalacja | SWO-460831 | Re classified RMED | planned SWO |  |  |  | Substitute analyzer for the Chrzanów laboratory. | Uninstallation of the replacement analyzer. | Deinstalacja
No | Visit | Complaint | RPOL | Klient zgłosił uszkodzenie igły aspirującej. | Wygięta igła. | Wymiana igły i uchwytu uszczelki wlotu. Pomiary techniczne w normie. | SWO-461100 | Defect | HW | Auto Inlet module | Probe bend |  | The customer reported damage to the aspirator needle. | Bent Needle. | Replacing the needle and the inlet gasket holder. Technical measurements in standard.
Yes | Visit | Complaint | RPOL | Zwrot analizatora z serwisu | Zwrot analizatora z serwisu | Konfiguracja analizatora. | SWO-460848 | Re classified RMED | planned SWO |  |  |  | Returning the analyzer from the site | Returning the analyzer from the site | Analyzer configuration.
No | Visit | Complaint | RPOL | Wykryto przeciek. | Błędy przepływów . Wykryto przeciek. Błąd temperatury kasety sensorowej. | Wymiana SP90 lot BX-60. Wymiana SC90 lot R0847. Sprawdzenie poprawności działania aparatu. | SWO-460452 | Assisted Replacement | SP | SP | FT |  | A leak has been detected. | Flow errors. A leak has been detected. Sensor cassette temperature error. | Replacement SP90 Lot BX-60. Replacing SC90 flight R0847. Validates the camera's performance.
No | Visit | Complaint | RPOL | Komputer wolno działa | Bardzo duża ilość kaset testowych w pamięci. | Wykasowanie z pamięci kaset starszych niż 10 dni. | SWO-461199 | Adjustment | SW | RMED SW | RMED SW | Delete cal records | Your computer is slow | Very large number of test cartridges in memory. | Erase cartridges older than 10 days.
No | Visit | Complaint | RLTD | Grace Creedon 00353 667184365 | Status code 2 error. | Reloaded analyser software. Status code 2 error remained. Loaded new database, reloaded software again. Ran calibration & QC’s. Pushed all operators from Aqure. Tested. | SWO-461240 | Defect | SW | RMED SW | Database | status code 2 | Grace Creedon 00353 667184365 | Status code 2 error. | Reloaded analyser software. Status code 2 error remained. Loaded new database, reloaded software again. Ran calibration & QC’s. Pushed all operators from Aqure. Tested.
Yes | 8 | Y | RCZE | Blank |  |  | D55462C4482563BD | Re classified RMED | planned SWO |  |  | see mail with title "RE: D55462C4482563BD - planned?" | Blank
No | Visit | Complaint | RFRA | Nous constatons depuis plusieurs semaines des arrêts à répétition d’un des deux ABL90 (092R0152N003/eq 1800288) de la réanimation chirurgicale de Lariboisière provoqués par des alertes demandant le changement du joint de prélèvement qui ne paraissent pas justifiées. Le fonctionnement correct de son miroir accentue l’impression de dysfonctionnement à répétition. Je souhaiterais que cet ABL soit examiné pour corriger son défaut. Je profite de cette demande pour que Radiometer nous précise les conditions d’accès à son SAV. Nous souhaiterions en particulier avoir un accusé de réception de nos messages. | Test inlet - sonde d’introduction tordue | Remplacement sonde introduction fourni par le laboratoire | SWO-461272 | Defect | HW | Auto Inlet module | Probe bend |  | We have been seeing for several weeks repeated shutdowns of one of the two ABL90s (092R0152N003/eq 1800288) of Lariboisière surgical resuscitation caused by alerts requesting the change of the sample joint that do not appear to be Justified. The correct operation of its mirror accentuates the impression of repeated malfunction. I would like this ABL to be reviewed to correct its flaw. I take advantage of this request to have Radiometer tell us the conditions of access to its Service. In particular, we would like to have an acknowledgement of our messages. | Inlet test - twisted introductory probe | Replacement introductory probe provided by the laboratory
No | Visit | Complaint | RAME | Fluid flow errors |  |  | SWO-461659 | Revisit required | HW | Basic Unit | Sensor Bay Unit | FT | Fluid flow errors
No | Visit | Complaint | RBNL | Verschillende sens en status errors op meerdere parameters. | Verschillende sens en status errors op meerdere parameters. | service menu - impedance - nok kuisen van de sensor pcb pinnen sensor simulator test -ok Cal - Nok vervangen sensor cassette cal -ok QC - 3levels | SWO-461293 | Assisted Replacement | SC | SC | Multi | Cal | Different Sens and status errors on multiple parameters. | Different Sens and status errors on multiple parameters. | service menu - impedance - nok kuisen van de sensor pcb pinnen sensor simulator test -ok Cal - Nok vervangen sensor cassette cal -ok QC - 3levels
No | Visit | Complaint | RFRA | Erreur cal à blanc - Fuite liquide cassette | Système neuf installé il y a une semaine. Pas encore servi. Fuite sous abl80. | Tuyau sur hemolyseur a sauté. Changement tuyau et mise en place de la bague de renfort ok. Changement des tuyaux d’évacuation ok. (Kit de maintenance en div) Flush cassette avec seringue ok (petit bouchage sonde) Flush hemolyseur ok. Relance abl80 ok. Calibration détecteur liquide ok Cal à blanc coox ok. Cycle de qualité x2 ok. Lancement de la procédure de nettoyage ok. Plus de fuite ok. Rapport par mail à qualite@labo-barla.eu toxicologie@labo-barla.eu | SWO-461590 | [5 empty] | Clear cal error - Liquid leak cassette | New system installed a week ago. Not served yet. Leak under abl80. | Pipe on hemolyseur jumped. Change hose and set up the reinforcement ring ok. Changing the drain pipes, okay. (Div maintenance kit) Flush cassette with syringe ok (small plug probe) Flush hemolyseur ok. Relaunch abl80 ok. Liquid detector calibration ok Cal to white coox ok. Quality cycle x2 ok. Launch of the cleaning procedure ok. No more leaks, okay. Report by email to qualite@labo-barla.eu toxicologie@labo-barla.eu
No | Visit | Complaint | RCHN | 无效的处理包 | 无效的处理包 | 更换试剂包后修复 | SWO-461111 | Defect | Consumables | Solution Pack | Chip, Solution Pack |  | Invalid processing package | Invalid processing package | Repair after replacement of reagent kit
No | Visit | Complaint | RCHN | 试管无法识别 | 试管无法识别，借903-186一个 | 清洁定标扫码器后修复 by jesson | SWO-460416 | Cleaned | HW | Aspiration Module | Barcode reader, Inlet |  | Test tube is not recognized | Test tube is not recognizable, borrow 903-186 one | Fix by after cleaning the calibration scanner jesson
No | Visit | Complaint | RDEU | Ausdruck nicht leserlich/ Nadel verbogen | Ausdruck ist nicht leserlich. Einlassnadel hat ein Knick in der Mitte | Druckeinheit ausgetauscht. Einlasssonde ausgetauscht. Neustart. | SWO-461781 | Defect | HW | Printer Module | Printer Head |  | Expression not legible/ Needle bent | Expression is not legible. Inlet needle has a crease in the middle | pressure unit. inlet probe. Restart.
No | Visit | Complaint | RDAN | Dorte klager over fejl på QC, hvilket viser sig at være ustabile målinger. | Fejl 664: Prøve-fejl Fejl 476: Ustabile måling Ingen synlige fejl. | Væskesensor justering samt pumpekalibrering. pH elektrodepakning - ok. | SWO-461499 | Adjustment | HW | pH/BG Module | LS | 664 | Dorte complains of errors on QC, which proves to be unstable measurements. | Error 664: Sample error Error 476: Unstable measurement No visible errors. | Sensor adjustment and pump calibration. pH electrodepaking-OK.
No | Visit | Complaint | RAME | Customer states that the analyzer is getting intermittent turbidity errors on patient samples. | Verified customer's stated issue. Performed a complete oxi test in service mode with all calibration results passing. Checked all tubings and discovered the pinch valve tubing was eroded. | Replaced the pinch valve tubing (provided by the customer) and performed all the fluidic transport tests in service mode and performed a tHb calibration. All fluid transport checks pass at this time. | SWO-462320 | Defect | HW | Oxi module | Pinch valve tube |  | Customer states that the analyzer is getting intermittent turbidity errors on patient samples. | Verified customer's stated issue. Performed a complete oxi test in service mode with all calibration results passing. Checked all tubings and discovered the pinch valve tubing was eroded. | Replaced the pinch valve tubing (provided by the customer) and performed all the fluidic transport tests in service mode and performed a tHb calibration. All fluid transport checks pass at this time.
Yes | Visit | Complaint | RDEU | 05.12.2018 13:15Uhr Frau Magner 040790202577: Das Gerät meldet Intervention erforderlich | Einlass verstellt, Gerät hat Verstopfung der Fliesswege | Fehleranalyse, Einlass kalibriert, Funktionskontrolle auf einzelne Probenpositionen i. O. , Gerät gespült, Nadel verstopft —> erneuert (Kundenmaterial), Kalibrierung und Qualitätskontrolle i. O. | SWO-462122 | [5 empty] | 05.12.2018 13:15pm Ms Magner 040790202577: The device reports intervention required | Inlet adjusted, device has blockage of the flow paths | Fault analysis, inlet calibrated, function control on individual sample positions i.o., device flushed, needle clogged —> renewed (customer material), calibration and quality control as a matter of time
Yes | Visit | Complaint | RBNL | Error 581 | Bij nader onderzoek blijkt dat de oximetrie tubevalve lek te zijn. | Defecte slangetje heb vervangen. | SWO-462299 | [5 empty] | Error 581 | Upon further examination it appears that the Oximetry tubevalve to be leak. | Defective hose have been replaced.
Yes | Visit | Complaint | RBNL | Defecte accu | Batterij icoontje staat constant op 0% . | Defecte batterij module heb vervangen en gecontroleerd. | SWO-462281 | [5 empty] | Defective battery | Battery icon is constantly at 0%. | Defective battery module have been replaced and checked.
No | Visit | Complaint | RAME | tHb cal failed | Found that thb failed blank cal | Rebooted analyzer corrected the issue | SWO-467628 | Adjustment | SW | RMED SW | RMED SW | Off/On | tHb cal failed | Found that thb failed blank cal | Rebooted analyzer corrected the issue
No | Visit | Complaint | RAME | Fluid flow errors |  |  | SWO-462315 | Defect | HW | Basic Unit | Sensor Bay Unit | FT | Fluid flow errors
Yes | Visit | Complaint | RAME | Proactive Service and Data Collection for Glucose Issues. | Customer reports recurring issues with Glucose. System cycles will not alleviate issues. Suspected issues with current sensor cassette. | Replaced sensor cassette. Issues went away with new sensor cassette. Spare sensor cassette of same LOT left for other analyzer. Conducted performance evaluation and ran zeroing cassette procedure. Collected data for further analysis. | SWO-462316 | [5 empty] | Proactive Service and Data Collection for Glucose Issues. | Customer reports recurring issues with Glucose. System cycles will not alleviate issues. Suspected issues with current sensor cassette. | Replaced sensor cassette. Issues went away with new sensor cassette. Spare sensor cassette of same LOT left for other analyzer. Conducted performance evaluation and ran zeroing cassette procedure. Collected data for further analysis.
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | PCO2定标灵敏度超范围 0378 0210 | 更换测试卡后修复 | SWO-464814 | Assisted Replacement | SC | SC | pCO2 | 378 | Test card, reagent package failure | PCO2 calibration sensitivity over range 0378 0210 | Repair after replacing test card
No | Visit | Complaint | RCHN | 测试卡无法扫描 | 测试卡无法扫描,借910-304一个 | 更换910-304一个正常。by jesson | SWO-461059 | Defect | HW | Storage Module | Barcode reader, storage wheel |  | Test card can't be scanned | Test card can not be scanned, borrow 910-304 one | Replace 910-304 one normal. by jesson
No | Visit | Complaint | RCHN | 进样支架故障 | 进样支架故障 借903-584*2个 | 更换2个903-584后修复 by jesson | SWO-460393 | Defect | HW | Aspiration Module | Tube holder |  | Sample bracket fault | Sample bracket fault borrowed 903-584 x 2 | Repair by jesson after replacing 2 903-584
No | Visit | Complaint | RCHN | 连接超时 3553 | 连接超时 3050,借910-373一个。 | 更换910-373一个正常。重新安装软件后正常 by jesson | SWO-460518 | Defect | HW | Analyzer Control Module | CF card | 3050 | Connection time-out 3553 | The connection times out 3050, borrowing 910-373 one. | Replace 910-373 one normal. Normal by jesson after reinstalling the software
No | Visit | Complaint | RCHN | 报错1999.61.22 | 报错1999.61.22，借836-166一个，903-373三个 | 更换836-166、903-373三个后修复by jesson | SWO-460425 | Defect | HW | Incubator Module | Drive wheel | 1999.61.22 | Error 1999.61.22 | Error 1999.61.22, loan 836-166 one, 903-373 three | Replace 836-166, 903-373 three post-fix by jesson
No | Visit | Complaint | RDEU | 05.12.2018, 08:40 Uhr, Fr. Hernsen, Tel. 028017102499: Oxy Hardware Problem | Immer wieder Fehler 269 Hemolyser Phase Regulation failed + Fehler 580 | Oxi Modul Oxi Complete test durchgeführt impedance shape failed, Oxi Modul ersetzt, bei Neustart Fehler Data Upgrade failed Datenbank beschädigt, Datenbank neu Ufgestzt Configuration aus der Sicherung eingespielt, Anwender informiert, Service tHb Kal Kal und THB Kal durchgeführt,, Kal und QC durchgeführt | SWO-462654 | Defect | HW | Oxi module | Oxi module | Phase error; 580, 269 | 05.12.2018, 08:40, Ms. Hernsen, tel. 028017102499: Oxy Hardware Problem | Again and again error 269 Hemolyser Phase Regulation failed + Error 580 | Oxi module Oxi Complete test performed impedance shape failed, Oxi module replaced, on restart error Data Upgrade failed database corrupted, database restarted Ufgestzt Configuration from backup infused, user informed, service tHb Kal Kal and THB Kal carried out, Kal and QC carried out
No | Visit | Complaint | RDEU | 04.12.2018 Sr. Wibke 021714092693 meldet, dass der Aqusdruck unleserlich ist. | Ausdruck unleserlich | Printer Head getauscht. | SWO-462260 | Defect | HW | Printer Module | Printer Head |  | 04.12.2018 Sr. Wibke 021714092693 reports that the Aqus print is illegible. | Expression illegible | Printer Head.
No | Visit | Complaint | RSCH | Rep OXI Hardware | Hemolyzer 0269 phase regulation failed Fehler seit Neu vorhanden | Hemolyzer ausgetauscht Oxi justiert tHb Kal Funktionskontrolle | SWO-461996 | Defect | HW | Oxi module | Hemolyzer | Phase error; 269 | Rep OXY Hardware | Hemolyzer 0269 phase regulation failed Errors since New present | Hemolyzer replaced Oxi adjusted tHb Kal Functional control
Yes | Visit | Complaint | RSCH | Rep OXI Hardware | Seit neu immer wieder Fehler 0882 Mischerträger Schlupf erkannt | Sicherung auf Mischerträger oxidiert. Ausgetauscht Funktionskontrolle aktuell i.o | SWO-462062 | Re classified RMED | same day upSWO |  |  |  | Rep OXY Hardware | Since new again and again error 0882 mixer carrier slip detected | insulated on mixer carrier. Exchanged Functional control currently i.o
No | Visit | Complaint | RKOR | Corruptxml error | Reboot the system | Reboot the system | SWO-463884 | Defect | SW | RMED SW | RMED SW | corrupt xml | Corruptxml error | Reboot the system | Reboot the system
No | Visit | Complaint | RFRA | Bug de temps en temps. Bruit porte chargement. | AQT90 neuf, en cours de reproductibilité. Vérification des connecteurs et porte chargement. | Nettoyage, huilage de la porte. Nettoyage, déconnection, reconnexion des connecteurs. Passage cleaning ok. | SWO-462322 | Adjustment | HW | Cover and Chassis | Door, Inlet |  | Bug from time to time. Noise carries loading. | AQT90 new, in the process of reproducibility. Check connectors and carry charging. | Cleaning, cleaning the door. Cleaning, disconnection, reconnecting connectors. Passage cleaning ok.
No | Visit | Complaint | RBNL | QC pas ok, fais le tHb cal - encore pas ok. | tHb cal 106% mauvais qc - sO2, FHbF, COHb | fais une autre tHb cal - 104% - QC ok test service - hemolyseur ok test de fuite - ok calibration, tHb cal, calibration, QC 3 niveaux - ok | SWO-462874 | Adjustment | HW | Oxi module | tHb Cal |  | QC not ok, do the tHb cal - still not ok. | tHb cal 106% mauvais qc - sO2, FHbF, COHb | make another cal tHb - 104% - QC ok test service - hemolyseur ok leak test - ok calibration, tHb cal, calibration, QC 3 levels - ok
Yes | Visit | Complaint | RPAC | Failed SC. Failing pararmeter Lactate. Manual flush and troubleshooting attended without success. SC replaced. | Failed SC. Failing pararmeter Lactate. Manual flush and troubleshooting attended without success. | SC replaced | SWO-462495 | Assisted Replacement | SC | SC | Lac | 1310 | Failed SC. Failing pararmeter Lactate. Manual flush and troubleshooting attended without success. SC replaced. | Failed SC. Failing pararmeter Lactate. Manual flush and troubleshooting attended without success. | SC replaced
Yes | Visit | Complaint | RIND | Preventive maintenance done. | Preventive maintenance required | Preventive maintenance carried by by clean the waste chamber,waste pump tubes,Clean the dust all electronics & mechanical parts,restart the instrument & after successful calibration instrument became ready. | SWO-462624 | Re classified RMED | planned SWO |  |  |  | Preventive maintenance done. | Preventive maintenance required | Preventive maintenance carried by by clean the waste chamber,waste pump tubes,Clean the dust all electronics & mechanical parts,restart the instrument & after successful calibration instrument became ready.
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | LAC定标失败，1225 | 更换测试卡后修复 | SWO-462542 | [5 empty] | Test card, reagent package failure | LAC calibration failure, 1225 | Repair after replacing test card
Yes | Visit | Complaint | RIND | ABL80 sensor board problem . | Hct error for every sample . | Replaced new SC80 , zeroing done . | SWO-463209 | [5 empty] | ABL80 sensor board problem . | Hct error for every sample . | Replaced new SC80 , zeroing done .
No | Visit | Complaint | RCHN | 未连接模块 | 未连接模块，借903-365一个 | 更换903-365后修复 | SWO-461866 | [5 empty] | The module is not connected | Unconnected module, borrowed 903-365 one | Repair after replacement 903-365
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 不均匀的定标3溶液 1539 | 更换试剂包正常。 | SWO-464169 | Assisted Replacement | SP | SP | FT | 1539 | Test card, reagent package failure | Uneven calibration 3 solution 1539 | The replacement kit is OK.
No | Visit | Complaint | RCHN | 连接废液口断裂废液口断裂 | 连接废液口断裂废液口断裂 | 更换924-591后修复 by jesson | SWO-463244 | Defect | HW | Wet section | Manifold |  | Break of the connection waste mouth | Break of the connection waste mouth | Repair by jesson after replacement 924-591
No | Visit | Complaint | RCHN | 无法开机 | 无法开机,借903-175，903-198各一个。 | 更换903-175后修复 by jesson | SWO-461152 | Defect | HW | Cover and Chassis | Power supply |  | Cannot power on | Cannot be switched on, borrow 903-175, 903-198 each. | Repair by jesson after replacement 903-175
No | Visit | Complaint | RCHN | 较准失败 2167 | 较准失败 2167借903-138一个。 | 更换903-138后修复 by jesson | SWO-461159 | Defect | HW | Incubator Module | PCB, incubator | 2167 | More than a quasi-failure 2167 | More than the quasi-failure 2167 borrowed 903-138 one. | Repair by jesson after replacement 903-138
No | Visit | Complaint | RCHN | 干燥前温度超范围，BNP传感器错误 | 干燥前温度超范围，BNP传感器错误，借905-908一个 | 更换905-908后修复 by jesson | SWO-462473 | Defect | HW | Process Module | Sensor, BnP | 1235 | Excess temperature before drying, BNP sensor error | Pre-drying temperature over range, BNP sensor error, borrowed 905-908 one | Repair by jesson after replacement 905-908
No | Visit | Complaint | RCHN | 干燥前温度超范围 | 干燥前温度超范围，借903-210一个 | 更换903-210后修复 by jesson | SWO-462503 | Defect | HW | Process Module | PCB, cup dryer | 1068 | Overrange temperature before drying | Over-range temperature before drying, borrow 903-210 one | Repair by jesson after replacement 903-210
No | Visit | Complaint | RDEU | 06.12.2018, 10:40 Uhr ==> Anrufer Fr. May, Tel.: 042161021232 ==> QC-Messfehler pO2 und pCO2 | pO2 Wert QC Fehler zu hoch gemessen122 | 2mal LP getauscht, manuell kalibriert QC 114 fast mittig | SWO-464400 | [5 empty] | 06.12.2018, 10:40 a.m. ==> Caller Fr. May, Tel.: 042161021232 ==> QC measurement error pO2 and pCO2 | pO2 Wert QC Fehler zu hoch gemessen122 | 2 times LP exchanged, manually calibrated QC 114 almost central
No | Visit | Complaint | RIND | Po2 QC values are coming low. | Clot in Po2 Membrane. | Clean the chamber and membrane then run the instrument its working good. | SWO-463395 | Cleaned | HW | pH/BG Module | measuring chamber | pO2; Clot | Po2 QC values are coming low. | Clot in Po2 Membrane. | Clean the chamber and membrane then run the instrument its working good.
No | Visit | Complaint | RPAC | Add TCP/IP Settings for communication. | Add TCP $IP Settings for communication. | Attend site and dismantle old abl that will not boot and remove hard disk. Put Hard disk in ziplock bag and freeze for 10 mins. Remove drive and re-install and boot old analysder. Note down ip settings and send copy to Jason (Radiometer). Attend new ABL and get TCP $IP Settings and test. Confirm Radiance connected ok and queue empty. Wait for machine to stabilize and runs QC's. | SWO-466169 | Setup | SW | His/Lis | His/Lis |  | Add TCP/IP Settings for communication. | Add TCP /IP Settings for communication. | Attend site and dismantle old abl that will not boot and remove hard disk. Put Hard disk in ziplock bag and freeze for 10 mins. Remove drive and re-install and boot old analysder. Note down ip settings and send copy to Jason (Radiometer). Attend new ABL and get TCP $IP Settings and test. Confirm Radiance connected ok and queue empty. Wait for machine to stabilize and runs QC's.
No | Visit | Complaint | RDEU | 05.12.2018 13:30 Fr. Bretting 030940113280: Kal-Fehler E`lythe hohe Driften | Wiederholt Lytedriften in der Kalibration | Proteinentfernungsprogramm durchgeführt, Messkammer gründlich gereinigt mit Aqua, Referenzelektrode zur Membran richtig zusammengeschoben, Kal, QC | SWO-463070 | Cleaned | HW | pH/BG Module | measuring chamber |  | 05.12.2018 13:30 Fr. Bretting 03094013280: Kal-Error E'lythe high drifts | Repeated lyte drifts in calibration | Protein removal program performed, measuring chamber thoroughly cleaned with aqua, reference electrode to the membrane correctly pushed together, Kal, QC
No | Visit | Complaint | RDEU | 14.11.2018 11:00 Fr. Soll 02151363000: Probenfehler TN-T Temperatur | Probenfehler bei Patientenmessungen beim Parameter Nt-probnp. | Cup dryer pcb erneuert - find Target Temperatur wird nicht bestanden. Temperatur zu niedrig (40 69-71). Steckverbindungen und Cup dryer pcb überprüft - Fehler besteht weiterhin. Alte Platine wieder eingebaut - Temperatur jetzt bei 53 aber weiterhin viel zu niedrig. Reparatur momentan nicht möglich. Frau Seuel ist damit einverstanden, dass das Gerät zur Reparatur mitgenommen wird und morgen ein Ersatz Gerät aufgestellt wird. | SWO-463411 | Revisit required | HW | Process Module | PCB, cup dryer |  | 14.11.2018 11:00 Fr. Should 02151363000: Sample error TN-T Temperature | Probenfehler bei Patientenmessungen beim Parameter Nt-probnp. | Cup dryer pcb renewed - find target temperature is not passed. Temperature too low (40 69-71). Connectors and Cup dryer pcb checked - error persists. Old board reinstalled - temperature now at 53 but still far too low. Repair is currently not possible. Mrs Seuel agrees that the device should be taken for repair and that a replacement device should be installed tomorrow.
No | Visit | Complaint | RDEU | 05.12.2018 14:30Uhr Herr Dietz 0627183277: LQC bei Tnt außerhalb Bereich | Anwender haben fünf mal mit der selben Qualitätskontrolle gemessen, Lot 01181. Drei mal war die Probe außerhalb über dem oberen Herstellergrenzen und RiliBäk Grenzen. Gemessen wird mit TNT Lot 15219 auch auf dem zweiten Analysator auf Station N3. Hier gibt es keine Grenzverletzungen, Qualitätskontroll Werte bei Level rot auch erhöht, aber nicht so stark. Kalibrationen sind in Ordnung, ohne Ausreißer. Anwender haben bereits Systemreinigung durchgeführt. | Reinigung der Fließwege durch Proteinreiniger im Service. Reinigung der Cup Wash Unit, des Process Wheels. Noch mal Systemreinigung durchgeführt. Service Dump gezogen, Flowwerte überprüft. Werte liegen an der oberen Grenze unterhalb von 1600. Pumpenköpfe P1 und P2 erneuert. Messung der Qualitätskontrolle mit anderer Los Nummer, 04181. Qualitätskontrolle ist innerhalb der Herstellergrenzen aber außerhalb der RiliBäk Grenzen. Weiteres Vorgehen in Absprache mit dem Technischen Support. | SWO-462941 | Cleaned | HW | Wet section | Wet section |  | 05.12.2018 14:30 Mr Dietz 0627183277: LQC at Tnt outside area | Users have measured five times with the same quality control, lot 01181. Three times the sample was outside the upper manufacturer's limits and RiliBäk limits. With TNT Lot 15219 is also measured on the second analyzer on station N3. Here there are no border violations, quality control values at level red also increased, but not so strong. Calibrations are fine, without outliers. Users have already performed system cleaning. | Cleaning of the flow paths by protein cleaners in the service. Cleaning the Cup Wash Unit, the Process Wheel. System cleaning once again. Service dump dragged, flow values checked. Values are at the upper limit below 1600. Pump heads P1 and P2 renewed. Measurement of quality control with other lot number, 04181. Quality control is within the manufacturer's limits but outside the RiliBäk limits. Further action in consultation with Technical Support.
No | Visit | Complaint | RKOR | Sample clot error | . | Liquid flow line clean. Gasket module clean. Service mode check. User training. | SWO-463848 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway | Clot | Sample clot error | . | Liquid flow line clean. Gasket module clean. Service mode check. User training.
No | Visit | Complaint | RKK | Kインピーダンスエエラー | Kインピーダンスエエラー | 機器引き上げ 持ち帰りチェック | SWO-463910 | Defect | HW | Basic Unit | PCB, Sensor interface | 1079 | K Impedance E-Error | K Impedance E-Error | Equipment pulling up Take-away check
No | Visit | Complaint | RFRA | Nous constatons depuis plusieurs semaines des arrêts à répétition d’un des deux ABL90 (092R0152N003/eq 1800288) de la réanimation chirurgicale de Lariboisière provoqués par des alertes demandant le changement du joint de prélèvement qui ne paraissent pas justifiées. Le fonctionnement correct de son miroir accentue l’impression de dysfonctionnement à répétition. Je souhaiterais que cet ABL soit examiné pour corriger son défaut. Je profite de cette demande pour que Radiometer nous précise les conditions d’accès à son SAV. Nous souhaiterions en particulier avoir un accusé de réception de nos messages. | Carte auto Inlet (détection ouverture $fermeture ) site introduction Sensor cassette bay | Démontage et remplacement de la carte auto Inlet avec nappe (détection ouverture $fermeture ) site introduction $Démontage et Remplacement de la sensor cassette bay.Calibration et passage des CQ | SWO-464222 | Defect | HW | Auto Inlet module | PCB, Auto inlet | hall element | We have been seeing for several weeks repeated shutdowns of one of the two ABL90s (092R0152N003/eq 1800288) of Lariboisière surgical resuscitation caused by alerts requesting the change of the sample joint that do not appear to be Justified. The correct operation of its mirror accentuates the impression of repeated malfunction. I would like this ABL to be reviewed to correct its flaw. I take advantage of this request to have Radiometer tell us the conditions of access to its Service. In particular, we would like to have an acknowledgement of our messages. | Inlet auto map (detection opening /fermeture) introduction site Sensor cassette bay | Disassembly and replacement of the Inlet auto card with tablecloth (detection opening $fermeture) site introduction $Demontage and replacement of the sensor cassette bay. Calibration and passage of CQ
No | Visit | Complaint | RESP | Solucion cal2 y 3 no homgenea | Solucion cal2 y 3 no homgenea | Limpieza del circuito con agua Chequeo por servicio Cambio del sp Cal, qc | SWO-463634 | Assisted Replacement | SP | SP | FT | 1539, 1538 | Solution cal2 and 3 nonhomgenea | Solution cal2 and 3 nonhomgenea | Cleaning the circuit with water Check for service Changing the sp Cal, qc
No | Visit | Complaint | RAME | Inlet gaskets not lasting. | Customer has had to change gaskets twice during a short interval. Out of reportable range results are not displayed. | Performed fluid transport programs. All passed. Set system settings to allow out of reportable range results to show. Could not get 24 hour clock setting to save. | SWO-463600 | Setup | SW | RMED SW | Roma.ini | Reportable range | Inlet gaskets not lasting. | Customer has had to change gaskets twice during a short interval. Out of reportable range results are not displayed. | Performed fluid transport programs. All passed. Set system settings to allow out of reportable range results to show. Could not get 24 hour clock setting to save.
No | Visit | Complaint | RBNL | printer print niet papier vervangen nok | printer print niet papier vervangen nok | pringting head - ok heropstarten toestel - ok | SWO-463438 | Adjustment | SW | RMED SW | RMED SW | Printer | Printer does not print Paper replace NOK | Printer does not print Paper replace NOK | Pringting Head-OK Reboot device-OK
No | Visit | Complaint | RAME | Analyzer is making noise during bootup and regular functioning of analyzer. | Cooling fan is making noise during analyzer power up. | Replaced cooling fan. Rebooted analyzer to make sure noise is no longer present. | SWO-463581 | [5 empty] | Analyzer is making noise during bootup and regular functioning of analyzer. | Cooling fan is making noise during analyzer power up. | Replaced cooling fan. Rebooted analyzer to make sure noise is no longer present.
No | Visit | Complaint | RCHN | cthb质控定标失败 | cthb质控定标失败 | 运行血红蛋白定标后，cthb质控通过 | SWO-463258 | Adjustment | HW | Oxi module | tHb Cal | 0089 | cthb QC failure | cthb QC failure | After running hemoglobin calibration, cthb QC passes
No | Visit | Complaint | RDEU | 07.12.2018 08:15Uhr Herr Haack +49 581 83-6050 : Das Gerät bootet nicht | Gerät startet nicht, roma.db roma .ini fehler Es wurde in der Nacht lt. Stationspersonal ein mehrfacher (23mal) Notstromtest durchgeführt, daraufhin war der Analysator nicht mehr startbereit | Fehleranalyse, Daten gesichert, Datenbank Rettung gescheitert, neue Anlaysator Software aufgesetzt, Gerätekonfiguration durchgeführt | SWO-463441 | User Error | HW | DMS_PC | CPU | Unintended Use | 07.12.2018 08:15Pm Mr Haack +49 581 83-6050 : The device does not boot | Device does not start, roma.db roma .ini error According to station staff, a multiple (23) emergency power test was carried out during the night, after which the analyzer was no longer ready to start | Error analysis, data backed up, database rescue failed, new anlaysator software set up, device configuration performed
No | Visit | Complaint | RIND | analysis is not processing. | SAMPLE ERROR: | check properly all possible faults like gasket inlet, inlet probe, inlet tube. clean everything properly. then try sampling but sample is not aspiring properly. need to replaced with new one. | SWO-463592 | Revisit required | HW | Inlet module | Preheater |  | analysis is not processing. | SAMPLE ERROR: | check properly all possible faults like gasket inlet, inlet probe, inlet tube. clean everything properly. then try sampling but sample is not aspiring properly. need to replaced with new one.
No | Visit | Complaint | RIND | Neon Lamp Failure | Checked the machine Oximetry Values got issue, and found that neon lamp got failure. | Need to change the Neon Lamp unit for proper working. | SWO-464654 | Revisit required | HW | Oxi module | Lamp unit | Neon lamp | Neon Lamp Failure | Checked the machine Oximetry Values got issue, and found that neon lamp got failure. | Need to change the Neon Lamp unit for proper working.
No | Visit | Complaint | RDEU | 14.11.2018 11:00 Fr. Soll 02151363000: Probenfehler TN-T Temperatur | Probenfehler bei Patientenmessungen beim Parameter NT-proBNP. | Leihgerät zur Reparatur Überbrückung aufgestellt - 838R0058N009. QC’s eingelesen und gemessen. | SWO-463412 | Adjustment | HW | Process Module | PCB, cup dryer |  | 14.11.2018 11:00 Fr. Should 02151363000: Sample error TN-T Temperature | Probenfehler bei Patientenmessungen beim Parameter NT-proBNP. | Rental device set up for repair bridging - 838R0058N009. QC read it and measured it.
No | Visit | Complaint | RIND | AQT90 system hang & error. | check the logs. | Solution pack fail to reload & system hang multiple time.Restart analyzer still fail to initialize. reset the cf card & check the logs. inlet door lock issue. machine under observation | SWO-463792 | Adjustment | SW | RMED SW | reinstall |  | AQT90 system hang & error. | check the logs. | Solution pack fail to reload & system hang multiple time. Restart analyzer still fail to initialize. reset the cf card & check the logs. inlet door lock issue. machine under observation
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | LAC,PH,K,NA,CL,LAC定标失败，1225 | 连续更换2张测试卡后修复 | SWO-465482 | [5 empty] | Test card, reagent package failure | LAC,PH,K,NA,CL,LAC定标失败，1225 | Repair after 2 consecutive test cards
No | Visit | Complaint | RCHN | 无法开机 | 无法开机，借903-381和902-970（903-663）各一个 | 原机上的电源指示灯不亮，更换903-663后，电源指示灯正常显示，但是机器仍不能正常启动，更换903-381后修复。 | SWO-462558 | [5 empty] | Cannot power on | Cannot be switched on, one each for 903-381 and 902-970 (903-663) | The power led on the original machine is not on, after replacing 903-663, the power led is displayed normally, but the machine still can not start normally, replace 903-381 after repair.
No | Visit | Complaint | RCHN | 温度过高 | 温度过高 借902-972 903-334各一个 | 更换903-334后修复 | SWO-463236 | [5 empty] | Too high temperature | Too high temperature Borrowing 902-972 903-334 each | Repair after replacement of 903-334
No | Visit | Complaint | RAME | Ca, K, Cl, and Na intermittently fail the 2-point calibration for drift. |  |  | SWO-463710 | Defect | HW | pH/BG Module | Valve, membrane | CRM: Troubleshot to a faulty valve on the pH/BG & El/Met modules. Replaced both valves on both modules, and this resolved the problem. (902-700 Membrane Valve x2) | Ca, K, Cl, and Na intermittently fail the 2-point calibration for drift.
No | Visit | Complaint | RCHN | 测试卡温度不稳定 | 加借备件903-334一个 |  | SWO-463821 | [5 empty] | The test card temperature is not stable | Borrow spare parts 903-334 one
No | Visit | Complaint | RCHN | 无法开机 | 无法开机 借910-285一个 |  | SWO-463820 | Defect | HW | Power Supply | Power supply |  | Cannot power on | Can't turn on, borrow 910-285 one.
No | Visit | Complaint | RCHN | 测试卡断裂 | 测试卡断裂，断裂部位（Teeth） | 更换测试卡后修复 by jesson | SWO-463823 | Defect | Consumables | Test Kit | Cartridge, Broken | Teeth; 5784 | Test card break | Test card break, break site (Teeth) | Fix by jesson after replacing test card
No | Visit | Complaint | RCHN | 光学强度超范围，报错1031 | 光学强度超范围，报错1031，借903-364,944-212（14868）,944-214（14934）各一个 |  | SWO-463830 | Defect | HW | Process Module | Optical Unit | 1031 | Optical strength over range, error 1031 | Optical strength over range, error 1031, borrowed 903-364, 944-212 (14868), 944-214 (14934) each
Yes | Visit | Complaint | RKK | QAキット交換作業 | QAパーツ交換作業 | 実施しました。 | SWO-463991 | [5 empty] | QA kit replacement work | QA parts replacement work | Was carried out.
No | Visit | Complaint | RKK | センサーカセットを交換して１、２時間後に、NaのQCの測定値が範囲外になった。（Naの感度が範囲外になっていた為） | 機器状況確認 訪問時には、自然復旧 | サービスモードにて、リーク電流値チェック。-9.27mV インレットプローブチェック。センサーカセットチェック。コンタクトプローブチェック。インレット $排液ライン洗浄。 インレットガスケット洗浄。 各部動作チェックOK。CALチェックOK。QCチェック（ABC）OK | SWO-464709 | No fault found | SC | SC | Na | 588 | After an hour or two after replacing the sensor cassette, na QC readings went out of range. (Because na's sensitivity was out of range) | Equipment status check During the visit, natural restoration | Leakcurrent value check in service mode. -9.27mV Inlet probe check. Sensor cassette check. Contact probe check. Inlet $ drainage line cleaning. Inlet gasket cleaning. Each part operation check OK. CAL check OK. QC Check (ABC) OK
Yes | Visit | Complaint | RKK | Kインピーダンスエエラー | Ｋインピーダンスエエラー | 機器返却 | SWO-464712 | Re classified RMED | planned SWO |  |  | Returning of Analyzer after repair in workshop, upSWO of Defect in SWO-463910 | K Impedance E-Error | K Impedance E-Error | Return equipment
Yes | Visit | Complaint | RFRA | Suite au problème erreur code 2 récurent sur les dispositifs demande de rendez vous service commerciale pour point technique avec le client . | Réunion avec chef de service, cadre bicetre paul brousse , biologiste | Point avant rendez-vous Point sur déroulement diagnostic Programmer maj 3.4 | SWO-464530 | Re classified RMED | planned SWO |  |  | Meeting | Following the error problem code 2 scour on the devices request inge dservice commercial for technical point with the customer . | Meeting with head of department, bicetre paul bush, biologist | Point before appointment Update on diagnostic progress Program maj 3.4
No | Visit | Complaint | RAME | Analyzer display is damaged and housing screws are no longer fasten in place. | Analyzer display is damaged and housing screws are no longer fasten in place. | Replaced display and tested. | SWO-464508 | Defect | HW |  | display | screws | Analyzer display is damaged and housing screws are no longer fasten in place. | Analyzer display is damaged and housing screws are no longer fasten in place. | Replaced display and tested.
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 测试卡漏液 | 更换测试卡后修复 | SWO-465455 | [5 empty] | Test card, reagent package failure | Test card leakage | Repair after replacing test card
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | pco2定标失败，0378 | 更换测试卡后修复 | SWO-465576 | Assisted Replacement | SC | SC | pCO2 | 378 | Test card, reagent package failure | pco2 bid failure, 0378 | Repair after replacing test card
No | Visit | Complaint | RIND | Calibration error | Checked the data logs found crea query in all the samples and calibration was not getting stable | Cleaned crea chamber | SWO-467066 | Cleaned | Fluidic pathway | Met II module | measuring chamber | Crea | Calibration error | Checked the data logs found crea query in all the samples and calibration was not getting stable | Cleaned crea chamber
No | Visit | Complaint | RAME | Printouts are about 95% solid black, nothing is readable. | Troubleshot to a bad printer | Replaced printer, and this resolved the problem | SWO-464536 | Defect | HW | Printer Module | Printer Module |  | Printouts are about 95% solid black, nothing is readable. | Troubleshot to a bad printer | Replaced printer, and this resolved the problem
No | Visit | Complaint | RDEU | Die Patientendaten sollen nicht zu Cobas IT 1000 verschickt werden. QC Daten sollen übermittelt werden. Nach der Installation war dieser Fehler nicht gegeben und trat plötzlich auf. | Die Patientendaten sollen nicht zu Cobas IT 1000 verschickt werden QC Daten sollen übermittelt werden. Der Fehler trat erst einige Tage nach der Inbetriebnahme auf. | Verbindung geprüft, Einstellungen der EDV Verbindung OK In der Gerätekonfiguration war die aktive Befundfreigabe aktiviert, daher wurden die Befunde erst nach aktiver Freigabe verschickt. Auf Wunsch von Herrn Kallaus die aktive Befundefreigabe deaktiviert. Anschließend wurde eine Testmessung durchgeführt, diese wurde direkt und automatisch an Cobas IT 1000 verschickt. | SWO-466369 | Setup | SW | His/Lis | His/Lis |  | The patient data should not be sent to Coba's IT 1000. QC data is to be transmitted. After the installation, this error did not occur and suddenly occurred. | Patient data should not be sent to Coba's IT 1000 QC data is to be transmitted. The error did not occur until a few days after commissioning. | Connection checked, settings of the computer connection OK In the device configuration, the active findings release was activated, so the findings were only sent after active approval. At Mr. Kallaus's request, the active approval of findings is deactivated. Subsequently, a test measurement was carried out, which was sent directly and automatically to Cobas IT 1000.
No | Visit | Complaint | RDAN | Kan ikke køre kapillær prøver | Kan ikke asp. Kap prøver | Skifter intet gasket 902-668 | SWO-464281 | Defect | HW | Inlet module | Inlet gasket |  | Cannot run capillary samples | Cannot ASP. KAP samples | Changing nothing gasket 902-668
No | Visit | Complaint | RDAN | Efter geninstallering af softwaren to gange laver ABL'en igen loader fejl efter genstart. | Undersøgt. | Datamat unit udskiftet, kundens setup installeret, slutafprøvet. | SWO-464811 | Defect | HW | DMS_PC | CPU | mail: send datamat=CPU back to RMED (Ejner) | After reinstalling the software twice, the Abls re-loads errors after restarting. | Examined. | Computer unit replaced, customer setup installed, final tested.
Yes | Visit | Complaint | RDAN | ABL har sort skærm - error in loader roma - sort skærm -Våddelen køre OK. Prøvet at tænde og slukke men uden - aftalt besøg på mandag | Undersøgt, database smadret. | Software ver. 6.18 installeret, slutafprøvet. | SWO-464229 | Re classified RMED | same day upSWO |  |  |  | ABL has black screen-error in Loader Roma-black screen-wet part run OK. Tried to turn on and off but without-agreed visit on Monday | Investigated, database smashed. | Software view. 6.18 installeret, slutafpr.vet.
No | Visit | Complaint | RFRA | Écran ne répond pas de tps en tps. | Chgt écran | Sauvegarde données et configuration. Chgt écrans nappe écran. Passage cleaning | SWO-464520 | Defect | HW | Analyzer Control Module | display |  | Screen does not respond to tps in tps. | Chgt screen | Back-up data and configuration. Chgt screens tablecloth screen. Passage cleaning
No | Visit | Complaint | RDEU | 01.12.2018 14:30, RR Pfl. Buluar 05931-1521293, medizinische Aufnahme, Fehler 2168, Fehler in Startvorgang, zuvor ist bereits eine TropT Kalibrierung (alte Charge war verbraucht, neue Charge LOT 15133) mit Fehlermeldung 1061 gescheitert, Geräteneustart wurde 3x durchgeführt ohne Erfolg, FSE Einsatz erforderlich 06.12.2018 12:15Uhr Herr Daniel Foulloin 059311524602: Das Gerät ist nicht meßbereit, es war wieder das Summen von einer Mechanik zu hören und danach stoppte das Gerät | Diverse Fehler 2168, 5652, 5658 | Gerät überprüft und gereinigt, Incubatorwheel kontrolliert, Schacht gesäubert, Incubator PCB ersetzt | SWO-464794 | Defect | HW | Incubator Module | PCB, incubator | 1061, 2168, 5652, 5658 | 01.12.2018 14:30, RR Pfl. Buluar 05931-1521293, medical recording, error 2168, error in startup process, previously a TropT calibration (old batch was used, new batch LOT 15133) failed with error message 1061, device restart was performed 3x without success, FSE use required 06.12.2018 12:15PM Mr Daniel Foulloin 059311524602: The device is not ready to measure, it was again to hear the buzzing of a mechanic and after that the device stopped | Miscellaneous Errors 2168, 5652, 5658 | Device checked and cleaned, Incubatorwheel checked, shaft cleaned, Incubator PCB replaced
No | Visit | Complaint | RKK | ｾﾝｻｰを本体に戻すときに、戻しても戻ったと認識せず測定をし続ける（下から閉めた場合）。 | 電源再起動やﾓｼﾞｭｰﾙ再取り付けするも改善せず | ﾓｼﾞｭｰﾙ交換 | SWO-466139 | [5 empty] | When you return the sensor to the console, it does not recognize that it has returned when returning it, and continues to measure it (if you close it from the bottom). | Power restart and module reinstallation does not improve | Module replacement
No | Visit | Complaint | RLTD | Jiger 01895828757 | Performed pressure tested valve vacuum failure. Noticed debri and QCs solution under the valve tubing. | Replaced valve tubing performed pressure test and ran cal QCs all okay. Analyser is fully operational. | SWO-465104 | Defect | HW | Oxi module | Pinch valve tube |  | Jiger 01895828757 | Performed pressure tested valve vacuum failure. Noticed debri and QCs solution under the valve tubing. | Replaced valve tubing performed pressure test and ran cal QCs all okay. Analyser is fully operational.
No | Visit | Complaint | RESP | Fallo de corriente de fuga en el SP | Fallo de corriente de fuga en el SP | Se detecto aguja doblada. Se coloco correctamente. Se realizo calibración de posición de la misma. Se sustituyo SP defectuoso. se calibro equipo y se comprobó correcto funcionamiento mediante QCs. Equipo apto para su uso según especificaciones del fabricante. | SWO-465936 | Assisted Replacement | SP | SP | Leak current |  | Leakage current failure in the SP | Leakage current failure in the SP | Bent needle detected. It was positioned correctly. Position calibration of the same was performed. Defective SP is replaced. equipment is cadescribed and tested for proper operation by QCs. Equipment suitable for use according to manufacturer's specifications.
No | Visit | Complaint | RDEU | 10.12.18 09:30 Herr Rein Tel.07219743285 , Einlaßmodul an der Führung für die ELD gebrochen =>Einlaßmodul defekt | Führung für Einlassdichtung am Einlaßmodul abgebrochen. | Datensicherung durchgeführt. Neues Einlassmodul installiert. ABL durchgespült und aufgelistete Teile erneuert. LP - Verriegelung leichtgängig gemacht. Einlaufverfahren durchgeführt. Fehlerhafte SC wegen pCo2 Kal.-Fehler aus Kundenbestand erneuert. Glu und Lac noch nicht messbereit. Kunde wird weitere Kal. durchführen und gegebenenfalls QC wiederholen. | SWO-464907 | Defect | HW | Auto Inlet module | Auto Inlet module | broken | 10.12.18 09:30 Mr. Rein Tel.07219743285 , insertion module broken on the guide for the ELD =>Inlet module defective | Guide for inlet seal on the inlet module aborted. | data backup. New inlet module installed. ABL rinsed through and replaced listed parts. LP - Locking made smooth. run-in procedures. Faulty SC renewed due to pCo2 cal. errors from customer inventory. Glu and Lac are not yet ready to measure. Customer will continue To Kal. and, if necessary, repeat QC.
Yes | Visit | Complaint | RDEU | Nach Aufstellen und Einschalten des Analysators trat der Fehler auf | Datenbankfehler 2 ==> Fehler trat in der Auslieferungsphase bei Einschalten des Analysators auf | Neue Datenbank installiert SC und SP installiert THb-Kalibrierung | SWO-465246 | Re classified RMED | installation report |  |  | Move to OOB trending | After setting up and switching on the analyzer, the error occurred | Database error 2 ==> Error occurred in the delivery phase when the analyzer was switched on | New database installed SC and SP installed THb calibration
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | LAC定标不过 | 更换测试卡后修复 | SWO-464780 | [5 empty] | Test card, reagent package failure | LAC calibration is not | Repair after replacing test card
No | Visit | Complaint | RCHN | 打印机故障，打出来都是黑色条码 | 打印机故障，打出来都是黑色条码 | 自行更换903-001和902-985后修复by tim | SWO-465386 | [5 empty] | Printer failure, hit out are black bar code | Printer failure, hit out are black bar code | Repair by tim after replacing 903-001 and 902-985
No | Visit | Complaint | RIND | Signal Data Board Problem | Checked and found that continous problem in Values and Calibration results. | Need to change the Signal Data Board. | SWO-465231 | [5 empty] | Signal Data Board Problem | Checked and found that continous problem in Values and Calibration results. | Need to change the Signal Data Board.
No | Visit | Complaint | RAME | Possible fluid blockage | SC check failures. Analog board failures. | Flushed analyzer. Reconnected tubings. Made Windows setting change to correct board failure errors. | SWO-465316 | [5 empty] | Possible fluid blockage | SC check failures. Analog board failures. | Flushed analyzer. Reconnected tubings. Made Windows setting change to correct board failure errors.
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | Lac定标不过 0210 | 更换测试卡后修复 | SWO-465391 | [5 empty] | Test card, reagent package failure | Lac calibration is no more than 0210 | Repair after replacing test card
No | Visit | Complaint | RCHN | 往上顶杯报错 | 往上顶杯报错 | 清理孵育盘固定轮后修复by tim | SWO-464851 | Defect | Consumables | Solution Pack | door, solution pack | cup stacking; 5257 | To the top cup to report the error | To the top cup to report the error | Repair by tim after cleaning the incubation plate fixing wheel
No | Visit | Complaint | RCHN | 进样盘试管脱落 | 进样盘试管脱落，借903-584一个 | 更换903-584后修复by tim | SWO-463970 | Defect | HW | Inlet module | Tube holder |  | The test tube of the inlet plate is off. | The test tube of the intake plate came off, borrowed 903-584 one | Repair by tim after replacement 903-584
Yes | Visit | Complaint | RCHN | 定标失败 | 定标失败，借944-212（15116）一个 | 重新定标后修复by tim | SWO-463956 | [5 empty] | Bid failed | Failed to bid, borrow 944-212 (15116) one | Fix by tim after re-calibration
No | Visit | Complaint | RCHN | 打印机故障，扫描器扫不到 | 打印机故障，扫描器扫不到，借910-275,902-822各一个 | 更换910-275,902-822后修复by tim | SWO-463974 | Defect | HW | Printer Module | Printer Module |  | Printer failure, scanner can't sweep | Printer failure, scanner can not be scanned, borrow 910-275, 902-822 each one | Repair by tim after replacing 910-275,902-822
No | Visit | Complaint | RAME | Crea out on high crea check | Crea failing on high Crea check. Inspected electrodes and wells. Found dummy electrode has blood and Crea b has oxidation. | Cleaned dummy electrodes and ordered replacement Crea A and B (w $membranes) as warranty replacement. | SWO-465241 | Cleaned | Electrode | Electrode | Crea | CAPA-790 | Crea out on high crea check | Crea failing on high Crea check. Inspected electrodes and wells. Found dummy electrode has blood and Crea b has oxidation. | Cleaned dummy electrodes and ordered replacement Crea A and B (w $membranes) as warranty replacement.
Yes | Visit | Complaint | RAME | On-going Lac issues. | Reviewed logs found lactate within range since last membrane replacement. | No action taken by FSE. Instrument calibrations and QC within manufacturer specifications. | SWO-465371 | Re classified RMED | planned SWO |  |  |  | On-going Lac issues. | Reviewed logs found lactate within range since last membrane replacement. | No action taken by FSE. Instrument calibrations and QC within manufacturer specifications.
No | Visit | Complaint | RDEU | Glukose instabil | Glukose instabil | Gerät überprüft und gereinigt, Meßkammer und Luftauslass gesäubert, diverse Elektroden membraniert, tHb Kal durchgeführt | SWO-465458 | Assisted Replacement | Electrode | Membrane | Glu |  | Glucose unstable | Glucose unstable | Device checked and cleaned, measuring chamber and air outlet cleaned, various electrodes membraneed, tHb Kal performed
Yes | Visit | Complaint | RSCH | Schulung | Schulung Membranenwartung | Ref., Glu und Lac membraniert i.o Kal. i.o System i.o | SWO-464917 | Re classified RMED | planned SWO |  |  |  | Training | Training membrane maintenance | Ref., Glu und Lac membraniert i.o Kal. i.o System i.o
No | Visit | Complaint | RDEU | 10.12.2018, 08:29 Uhr, Sr. Isabell, Tel. 06232221600: Pat. Erg. zwischen AQT und Labor bei Trop T abweichend | Gestern um 1:30 wurde eine Messung mit einem TnT - Wert von über 500 ng gemessen . Die Probe wird 2 Std. später mit dem Wert <10 ng gemessen. Messung im Labor ergibt 3 ng. | Zu diesem Vorgang wurde DCG und BVB ausgefüllt. Alle tech. relevante Punkte sind durchgeführt. Rinigung. Blankmessung. LQC und Servicedamp. | SWO-465464 | Cleaned | HW | Process Module | Process Wheel | Techline: RDEU 829699553 | 10.12.2018, 08:29, Sr. Isabell, tel. 06232221600: Pat. Erg. different between AQT and laboratory at Trop T | Yesterday at 1:30 a measurement was measured with a TnT value of over 500 ng. The sample is measured 2 hrs later with a value of 10 ng. Measurement in the laboratory yields 3 ng. | DCG and BVB have been filled in for this operation. All tech. relevant points have been carried out. Rinigung. Blank measurement. LQC and Servicedamp.
No | Visit | Complaint | RAME | Gas flow issues Phill Smith 509-897-2708 | Arrived onsite and replaced the module. Started analyzer and calibration passed without issue. Phil will interface tonight and call if experiencing any problems. | Replaced module. Analyzer operational. | SWO-465267 | [5 empty] | Gas flow issues Phill Smith 509-897-2708 | Arrived onsite and replaced the module. Started analyzer and calibration passed without issue. Phil will interface tonight and call if experiencing any problems. | Replaced module. Analyzer operational.
No | Visit | Complaint | RPOL | Wykryto skrzep w analizatorze | Wykryto skrzep w analizatorze | Przepłukanie aparatu, wykonano testy przepływów . Wymiana igły . | SWO-464857 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway | Clot | Clot detected in analyser | Clot detected in analyser | Flushing the camera, the flow tests were performed. Replacing the needle.
No | Visit | Complaint | RDEU | 11.12.2018; Fr. Roccor; Tel.: 047316903444; Na und K Kal Fehler 11.12.18 Fr Roccor Tel.047316903444 Leckage Fehler und Backlog Fehler | Kal Fehler in Elyte K und Na nach Tausch der SC wegen pCO2 Fehler in QC | 2 mal SC getauscht unterschiedliche Charge, manuell kalibriert durch kal und QC Fehler Met, Kal und QC IO Testmessungen Sensorfehler Ci und K , Qwetschventil sehr schwergängig, getauscht, Testmessung ok | SWO-466844 | Assisted Replacement | SC | SC | K/Na | Backlog | 11.12.2018; Ms. Roccor; Tel.: 047316903444; Na and K Kal Error 11.12.18 Fr Roccor Tel.047316903444 Leakage Error and Backlog Error | Kal error in Elyte K and Na after swapping SC due to pCO2 error in QC | 2 times SC exchanged different batch, manually calibrated by kal and QC error Met, Kal and QC IO test measurements sensor error Ci and K, Qwetsch valve very difficult to replace, test measurement ok
No | Visit | Complaint | RKOR | Sample clot error | . | DM 김민우씨 방문하여 해결함. | SWO-465889 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway | Clot | Sample clot error | . | DM Kim Min-woo visited and defected.
No | Visit | Complaint | RAME | Oxi hardware error | Checked the data logs and verified the problem. Tried rebooting the analyzer. I tried swapping out the spectrometer and that didn't work. The the inlet wouldn't calibrate. Put the original spectrometer in Andy the analyzer came back up like normal. Then the manifold would not calibrate. Ordering a wet section PCB and hemolyzer. | Installed the Wet Section Control PCB. | SWO-464665 | Revisit required | HW | Oxi module | Hemolyzer |  | Oxi hardware error | Checked the data logs and verified the problem. Tried rebooting the analyzer. I tried swapping out the spectrometer and that didn't work. The the inlet wouldn't calibrate. Put the original spectrometer in Andy the analyzer came back up like normal. Then the manifold would not calibrate. Ordering a wet section PCB and hemolyzer. | Installed the Wet Section Control PCB.
Yes | Visit | Complaint | RAME | Would like new equipment setup | Swap instrument from old to new | Entered in address and tested connection. Good | SWO-473377 | Re classified RMED | planned SWO |  |  |  | Would like new equipment setup | Swap instrument from old to new | Entered in address and tested connection. Good
No | Visit | Complaint | RAME | Status 2 error/Battery issue | Not retaining Wireless settings | Found that the wrong steps were taken to setup the Wireless. I spoke with David Topaloff and we were able to set it properly. | SWO-466109 | Setup | SW | RMED SW | Wireless | His/Lis | Status 2 error/Battery issue | Not retaining Wireless settings | Found that the wrong steps were taken to setup the Wireless. I spoke with David Topaloff and we were able to set it properly.
No | Visit | Complaint | RAME | Quick access is not functioning | Windows Error, Could not update configuration file: C: $ProgramData $Netop $NetOp $RemoteControl $Host $modems.nbd | Reloaded software and tested via AQURE. | SWO-465347 | Defect | SW | Windows | Windows | Reload SW | Quick access is not functioning | Windows Error, Could not update configuration file: C: /ProgramData /Netop /NetOp /RemoteControl /Host /modems.nbd | Reloaded software and tested via AQURE.
No | Visit | Complaint | RCHN | 接口板腐蚀 | 接口板腐蚀，借902-972一个 | 更换902-972后修复by Tim | SWO-465442 | [5 empty] | Interface plate corrosion | Interface plate corrosion, borrow 902-972 one | Repair by Tim after replacement 902-972
No | Visit | Complaint | RAME | Multiple SC issues | Troubleshot to a bad sensor cassette. I also found intermittent OXY Spectrum Mismatch errors, and a loud power supply fan. | Replaced sensor cassette (supplied by customer). Ordered a hemolyzer and power supply. Analyzer is completely operational. | SWO-465360 | [5 empty] | Multiple SC issues | Troubleshot to a bad sensor cassette. I also found intermittent OXY Spectrum Mismatch errors, and a loud power supply fan. | Replaced sensor cassette (supplied by customer). Ordered a hemolyzer and power supply. Analyzer is completely operational.
No | Visit | Complaint | RAME | Analog and oxi board failure messages | Customer has 11 OSM analyzers many have had intermittent board failure problems. Units will operate normally after a reboot. | Performed ABL80 WES7 serial enumerator setting change on all analyzers. Also made change to his $list settings in application file on 6 cath lab analyzers. | SWO-465322 | [5 empty] | Analog and oxi board failure messages | Customer has 11 OSM analyzers many have had intermittent board failure problems. Units will operate normally after a reboot. | Performed ABL80 WES7 serial enumerator setting change on all analyzers. Also made change to his $list settings in application file on 6 cath lab analyzers.
No | Visit | Complaint | RCHN | 定标中止 | 机器自动重启导致定标数据丢失，报错5257，借944-258（15025）、944-212（15116）各一张 | 重新定标后修复by tim | SWO-464041 | Defect | Consumables | Solution Pack | door, solution pack | cup stacking; 5257 | Bid abort | Automatic restart of the machine resulted in the loss of calibration data, error 5257, borrowing 944-258 (15025), 944-212 (15116) each one | Fix by tim after re-calibration
No | Visit | Complaint | RCHN | 血氧检测错误 | 血氧检测错误，借902-435一个 | 更换902-435后修复by tim | SWO-464190 | Defect | HW | Oxi module | Hemolyzer | 581 | Blood oxygen test error | Blood oxygen test error, borrow 902-435 one | Repair by tim after replacement 902-435
Yes | Visit | Complaint | RIND | Preventive maintenance required. | Preventive maintenance required. | Preventive maintenance carried out by cleaning the gasket,pump tubes & tubings in the liquid path,clean the dust from all the mechanical,electronics parts,Then restart the instrument & run the 2point calibration & sample ,got the desired result. | SWO-467383 | Re classified RMED | planned SWO |  |  | see mail with title "RE: SWO-467383" | Preventive maintenance required. | Preventive maintenance required. | Preventive maintenance carried out by cleaning the gasket,pump tubes & tubings in the liquid path,clean the dust from all the mechanical,electronics parts,Then restart the instrument & run the 2point calibration & sample ,got the desired result.
No | Visit | Complaint | RDEU | 11.12.2018; Fr. Niemeyer; Tel.: 0401818863360; AU nicht mit DMS verbunden | AU nicht mit DMS verbunden. Neustart durchgeführt | Fehleranalyse, nach Geräteneustart war Fehlermeldung weg, bei Versuch über FlexQ Modul Messung zu starten kommt Fehlermeldung AU nicht mit DMS verbunden wieder, Sicherungen von Analyzer Control PCB überprüft und Kontakte gereinigt, nach Geräteneustart keine Fehlermeldung mehr, QC Material sowie Einstellungen überprüft und korrigiert, Kalibrierung und Qualitätskontrolle i. O. , Probemessungen durchgeführt | SWO-465648 | Adjustment | HW | Analyzer Control Module | Anlyzer Control Module |  | 11.12.2018; Ms. Niemeyer; Tel.: 0401818863360; AU not connected to DMS | AU is not connected to DMS. Restart performed | Error analysis, after device restart error message was gone, when trying to start via FlexQ module measurement, error message AU does not reconnect to DMS, backups of Analyzer Control PCB checked and contacts cleaned, after device restart no error message more, QC material as well as settings checked and corrected, calibration and quality control in the o. , sample measurements carried out
No | Visit | Complaint | RLTD | Please call Phil McCue on 01908 995797 | Cartridge wheel jammed with a cartridge divider that has come disconnected from the base. | Stripped down cartridge wheel to remove jammed divider, divider looked ok with no visible damaged and locked back into the wheel base securely. Analyser started up with no issues and went straight into ready mode. | SWO-465652 | Adjustment | HW | Storage Module | Cartridge guide | 5784 | Please call Phil McCue on 01908 995797 | Cartridge wheel jammed with a cartridge divider that has come disconnected from the base. | Stripped down cartridge wheel to remove jammed divider, divider looked ok with no visible damaged and locked back into the wheel base securely. Analyser started up with no issues and went straight into ready mode.
No | Visit | Complaint | RKOR | CKMB Hct error | . | Wash unit clean. P4 p5 line clean. | SWO-465886 | Cleaned | HW | Aspiration Module | Needle Wash Unit | HCT error | CKMB Hct error | . | Wash unit clean. P4 p5 line clean.
No | Visit | Complaint | RDEU | ANDME wollte AQT kalibrieren, Kassetten wurden nicht eingescannt. | Scanner justage | Scanner justiert, die Kassetten wurden gescannt. | SWO-465769 | Adjustment | HW | Storage Module | Barcode reader, storage wheel |  | ANDME wanted to calibrate AQT, cassettes were not scanned. | Scanner adjustment | scanner, the cartridges have been scanned.
No | Visit | Complaint | RKK | ｾﾝｻｰを本体に戻しても戻ったと認識せず、測定し続ける | 電源再起動、ﾓｼﾞｭｰﾙ再取り付けするが改善せず | ﾓｼﾞｭｰﾙ交換 | SWO-466144 | [5 empty] | Return the sensor to the console, but do not recognize that it has returned and continue to measure | Power restart, module reinstallation, but not improved | Module replacement
No | Visit | Complaint | RAME | Sensor temp error | Temp error and customer replaced the SC and still had the error | I replaced the sensor interface PCB | SWO-467075 | Defect | HW | Basic Unit | PCB, Sensor interface | Temp | Sensor temp error | Temp error and customer replaced the SC and still had the error | I replaced the sensor interface PCB
No | Visit | Complaint | RDEU | 6.12.2018 13:30Uhr Frau Woiton 04141971495: Wegen div. Probleme bei der Bluteingabe soll eine Geräteeinweisung stattfinden. | Einweisung bedingt durch Probleme bei der Probeneingabe | Einweisung und Klärung der Fragen durchgeführt, Es sind keine weiteren Fragen offen | SWO-466390 | User Error | preanalytical | Sample handling | Sample handling |  | 6.12.2018 13:30Pm Ms Woiton 04141971495: Due to various problems with blood entry, a device briefing should take place. | Instruction due to problems with sample entry | Instruction and clarification of the questions carried out, there are no further questions unanswered
No | Visit | Complaint | RAME | Results from analyzer out put different with results come across to Meditec. It mentioned by customer | Customer complains that the results from serial output are different with the results showing on the analyzer | User mixed up the QC test results number. Checked the QC test results on both analyzer and Meditec both are completely same. No issue has seen. | SWO-466653 | No fault found | SW | External SW | External SW |  | Results from analyzer out put different with results come across to Meditec. It mentioned by customer | Customer complains that the results from serial output are different with the results showing on the analyzer | User mixed up the QC test results number. Checked the QC test results on both analyzer and Meditec both are completely same. No issue has seen.
Yes | 1 | Y | RCZE | Blank | Viacero chýb Liqud senorov, nehomogénnych roztokov |  | D55FADB46E4B09E5 | Re classified RMED | installation report |  |  | Move to OOB trending | Blank | Multiple errors Liqud Senorov, non-homogeneous solutions
No | Visit | Complaint | RCHN | 无法开机 | 无法开机加借903-111一个。 | 更换903-111一个正常。by tim | SWO-465592 | [5 empty] | Cannot power on | Cannot be turned on and borrowed 903-111 one. | Replace 903-111 one normal. by tim
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | NA质控B，C不过 报错0089 | 更换测试卡后修复 | SWO-466186 | Assisted Replacement | SC | SC | Na | 0089 | Test card, reagent package failure | NA QC B, C but error 0089 | Repair after replacing test card
No | Visit | Complaint | RCHN | 温度传感器故障 | 温度传感器故障，借944-220（15065）一个 | 重新定标正常。by Tim | SWO-465503 | Defect | HW | Process Module | Sensor, BnP | 1235 | Temperature sensor failure | Temperature sensor fault, one of the 944-220 (15065) | Recalibration is normal. by Tim
Yes | Visit | Complaint | RCHN | 定标失败，报错1024 | 借用903-364一个（使用自己库存机光单），944-394（15048）（14411缺货）944-212（14939）各一张 | 更换903-364和重新定标后修复by tim | SWO-466236 | [5 empty] | Failed to bid, error 1024 | Borrow 903-364 one (use your own stock machine light sheet), 944-394 (15048) (14411 out of stock) 944-212 (14939) each | Repair by tim after replacement 903-364 and re-calibration
No | Visit | Complaint | RCHN | 样本检测失败 | 样本检测失败，借636-542一个 | 更换636-542一个正常。by tim | SWO-465527 | Defect | HW | Storage Module | Cable, Storage |  | Sample detection failed | Sample test failed, borrow 636-542 one | Replace 636-542 one normal. by tim
No | Visit | Complaint | RCHN | TNI定标中止 1055 | TNI定标中止 1055,借944-212（15194）一张。 | 重新定标后修复 | SWO-464924 | Assisted Replacement | Consumables | Test Kit | Cartridge, cal | 1055 | TNI Calibration Suspended 1055 | TNI bid suspension 1055, borrow 944-212 (15194) one. | Repair after re-calibration
No | Visit | Complaint | RDAN | Nina klager over, at pH elektroden er utæt. | . | Pakring på pH elektrode udskiftet, preheater $probe renset, slutafprøvet. | SWO-466299 | Defect | Electrode | Electrode | Sealing ring, pH |  | Nina complains that the pH electrode is leaking. | . | Seal on pH electrode replaced, preheater $probe purified, final tested.
No | Visit | Complaint | RDAN | Nina klager over, at pH elektroden er utæt. | . | Pakring på pH elektrode udskiftet, preheater $probe renset. | SWO-466290 | Defect | Electrode | Electrode | Sealing ring, pH |  | Nina complains that the pH electrode is leaking. | . | Seal on pH electrode replaced, preheater $probe cleaned.
No | Visit | Complaint | RDAN | Nina klager over, at pH elektroden er utæt. | . | Pakring på pH elektroden udskiftet, preheater $probe renset, slutafprøvet. | SWO-466313 | Defect | Electrode | Electrode | Sealing ring, pH |  | Nina complains that the pH electrode is leaking. | . | Seal on pH electrode replaced, preheater $probe purified, final tested.
No | Visit | Complaint | RPOL | Błąd kalibracji sodu. Status poza zakresem (za niski). Dwie kasety różnych lotów z rzędu. |  |  | SWO-466567 | Assisted Replacement | SC | SC | Na |  | Sodium calibration error. Status out of range (too low). Two cartridges of different flights in a row.
No | Visit | Complaint | RPOL | Komputer wolno działa | Analizator samoczynnie się restartuje | Reinstalacja Windows . Instalacja oprogramowania 8.11 | SWO-468496 | Adjustment | SW | RMED SW | reinstall |  | Your computer is slow | The analyzer restarts automatically | Reinstalling Windows. Installing the 8.11 Software
No | Visit | Complaint | RPOL | Komputer wolno działa | Storage wheel error | Czyszczenie storage wheel | SWO-468495 | Cleaned | HW | Storage Module | storage wheel, hub |  | Your computer is slow | Storage wheel error | Czyszczenie storage wheel
No | Visit | Complaint | RAME | Turbidity too high | Follow up to SWO-462320. | Replaced units hemolyzer and performed all necessary adjustments in and tests in service mode. Once all tests passed the unit initialized with all calibrations and internal QCs passing at this time. Lab tech ran 5 patient samples with no turbidity or oxi spectrum mismatch errors detected at this time. | SWO-467202 | Defect | HW | Oxi module | Hemolyzer |  | Turbidity too high | Follow up to SWO-462320. | Replaced units hemolyzer and performed all necessary adjustments in and tests in service mode. Once all tests passed the unit initialized with all calibrations and internal QCs passing at this time. Lab tech ran 5 patient samples with no turbidity or oxi spectrum mismatch errors detected at this time.
No | Visit | Complaint | RAME | On ac power 0% number is below plug icon. Unplug 0% hangs on battery power unit shuts off within five min | Battery has 0% charge | I replaced the battery, watched the charge percentage increase noted. Yellow on battery light when off ac and noted charging green light plugged into ac | SWO-467132 | Defect | HW | Battery module | Battery module | 0% | On ac power 0% number is below plug icon. Unplug 0% hangs on battery power unit shuts off within five min | Battery has 0% charge | I replaced the battery, watched the charge percentage increase noted. Yellow on battery light when off ac and noted charging green light plugged into ac
No | Visit | Complaint | RLTD | Failing calibration with Inhomogeneous Cal 3 | Following the installation of the analyser on 27 $11 it was ok for approx 8hrs and then it errored with inhonogeneous cal 3. It had taken a while to get the analyser to startup in the first place and following conversations with a colleague we thought that the sensor bay could be at fault. | A new sensor bay was fitted and when the analyser had powered up the fluid transport service programs were all run and they all passed. When I exited the service programs there was a temperature error present and the analsyer had to be rebooted to clear this. Once all the errors had cleared the analyser was allowed to go through its start-up. Once the calibration had started I was happy that the fluid transport errors had been resolved and left the analyser to stabilise of the weekend. Keymea went in on Monday and the only errors present were calibration errors on the Glu & Lac. The customer was happy to run with this cassette while they did their validation of the analsyer. | SWO-469290 | Defect | HW | Basic Unit | Sensor Bay Unit | FT | Failing calibration with Inhomogeneous Cal 3 | Following the installation of the analyser on 27 /11 it was ok for approx 8hrs and then it errored with inhonogeneous cal 3. It had taken a while to get the analyser to startup in the first place and following conversations with a colleague we thought that the sensor bay could be at fault. | A new sensor bay was fitted and when the analyser had powered up the fluid transport service programs were all run and they all passed. When I exited the service programs there was a temperature error present and the analsyer had to be rebooted to clear this. Once all the errors had cleared the analyser was allowed to go through its start-up. Once the calibration had started I was happy that the fluid transport errors had been resolved and left the analyser to stabilise of the weekend. Keymea went in on Monday and the only errors present were calibration errors on the Glu & Lac. The customer was happy to run with this cassette while they did their validation of the analsyer.
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | PO2质控不过 | 更换测试卡后修复 | SWO-469117 | [5 empty] | Test card, reagent package failure | PO2 QC but not | Repair after replacing test card
No | Visit | Complaint | RCHN | 质控不过 | 质控不过，借902-980一个 | 更换902-980后修复by Tim | SWO-465575 | Defect | HW | Oxi module | Hemolyzer |  | But the quality control is not | But the quality control, borrow 902-980 one | Repair by Tim after replacement 902-980
No | Visit | Complaint | RCHN | 光学强度超范围，报错1031 | 光学单元超出范围，报错1031，定标底部外露，报错1024，需要借件903-364和903-236各一个 | 更换903-364后修复by Tim | SWO-465471 | Defect | HW | Process Module | Optical Unit | 1031 | Optical strength over range, error 1031 | Optical unit out of range, error 1031, calibration bottom exposed, error 1024, need to borrow 903-364 and 903-236 each one | Repair by Tim after replacing 903-364
No | Visit | Complaint | RCHN | 进样口支架破损 | 进样口支架破损，借903-584一个 老师来电抱怨工程师响应速度慢，要求工程师以后及时响应，已反馈工程师，工程师表示已在现场处理 | 更换903-584后修复by Tim | SWO-465472 | Defect | HW | Inlet module | Tube holder |  | Damaged in the inlet bracket | The inlet bracket is broken, borrow 903-584 one The teacher called to complain that the engineer was slow to respond, asked the engineer to respond in a timely manner, and the engineer has been reported back to the engineer, the engineer said it has been on site to deal with | Repair by Tim after replacing 903-584
No | Visit | Complaint | RCHN | 扫描器无法扫描 | 扫描器无法扫描 | 重新设置系统后修复by Tim | SWO-466853 | Setup | SW | RMED SW | RMED SW | Barcode reader, inlet | Scanner cannot be scanned | Scanner cannot be scanned | Fix by Tim after resetting the system
No | Visit | Complaint | RIND | analysis is not processing. | Sample Error and temperature error | Replaced inlet probe and preheater properly. then give 2 point calibration 2-3 time. after that analyser start working properly. handed over in fine working condition. | SWO-467125 | Defect | HW | Inlet module | Preheater |  | analysis is not processing. | Sample Error and temperature error | Replaced inlet probe and preheater properly. then give 2 point calibration 2-3 time. after that analyser start working properly. handed over in fine working condition.
No | Visit | Complaint | RDEU | 14.12.18 09:15 Frau Lamm Tel.068411622792 , Druckerbild zum teil unleserlich | Druckbild schlecht. | Druckereinheit ausgetauscht. Kal und QC. | SWO-467643 | Defect | HW | Printer Module | Printer Module |  | 14.12.18 09:15 Mrs. Lamb Tel.068411622792 , Printer picture partly illegible | Print image bad. | replaced by the printer unit. Kal and QC.
No | Visit | Complaint | RAME | Intermittent Hemolyzer phase regulation failed | Found intermittent Hemolyzer phase regulation failed errors on patient and QC samples. | Replaced hemolyzer. Verified operation. 12 $14- returned reviewed logs. Problem not returned since last visit. | SWO-465373 | Defect | HW | Oxi module | Hemolyzer |  | Intermittent Hemolyzer phase regulation failed | Found intermittent Hemolyzer phase regulation failed errors on patient and QC samples. | Replaced hemolyzer. Verified operation. 12 $14- returned reviewed logs. Problem not returned since last visit.
No | Visit | Complaint | RSAF | Halogen lamp intensity too high | Intensity high 3896 | Halogen lamp adjustment done. Complete oxi test ok. | SWO-466943 | Adjustment | HW | Oxi module | Lamp unit | Halogen | Halogen lamp intensity too high | Intensity high 3896 | Halogen lamp adjustment Done. Complete oxy Test OK.
No | Visit | Complaint | RKOR | network. error | . | ABL90 IP replace. ICU, NICU, ER, CCU visits => 전체 data 재전송 및 check. | SWO-467480 | Setup | SW | RMED SW | His/Lis | IP address | network. error | . | ABL90 IP replace. ICU, NICU, ER, CCU visits => 전체 data 재전송 및 check.
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 定标灵敏度超范围 | 更换测试卡后修复 | SWO-471370 | Assisted Replacement | SC | SC | SC | Cal | Test card, reagent package failure | Calibration sensitivity over range | Repair after replacing test card
Yes | Visit | Complaint | RCHN | 光学强度超范围 | 报错1031，1024,1029，借903-364一个 | 更换903-364后修复(By Ran) | SWO-464032 | [5 empty] | Optical strength over range | 1031,1024,1029, loan 903-364 one | Repair after replacement 903-364 (By Ran)
No | Visit | Complaint | RESP | Error intervención requerida | Comprobación de los registros de actividad etc... con el cliente y comprobación por soporte remoto. | Limpieza del analizador, comprobación de la entrada, cambio de la junta de conexión, junta de entrada y tubo de la válvula de oxi, Test diferentes por servicio, calibración y controles de calidad, funcionando correctamente. | SWO-467502 | Defect | HW | Basic Unit | PM Kit | FT, 1384 | Intervention error required | Checking activity logs etc... customer and remote support verification. | Analyzer cleaning, inlet check, connection joint change, inlet joint and oxy valve tube, Different test by service, calibration and quality controls, working properly.
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | CO2定标不过 | 更换测试卡后修复 | SWO-472477 | [5 empty] | Test card, reagent package failure | CO2 calibration is not the | Repair after replacing test card
No | Visit | Complaint | RCHN | 血氧单元故障，报错1018 | 血氧单元故障，报错1018，需要借件902-980 |  | SWO-467436 | Defect | HW | Oxi module | Hemolyzer |  | Blood oxygen unit failure, error 1018 | Blood oxygen unit failure, reported error 1018, need to borrow 902-980
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 定标失败 定标灵敏度超范围 | 连续更换1张测试卡和1个试剂包后修复 | SWO-471371 | Assisted Replacement | SP | SP | SP | Cal | Test card, reagent package failure | Bid failed Calibration sensitivity over range | Repair after continuous replacement of 1 test card and 1 reagent kit
No | Visit | Complaint | RCHN | HCT传感器响应错误 | HCT传感器响应错误,需要借件903-329 | 更换903-329后发现条码扫描器故障，需要加借910-304一个by tim | SWO-467350 | Defect | HW | Aspiration Module | Needle Wash Unit | HCT error | HCT sensor response error | HCT sensor response error, need to borrow 903-329 | Replacing 903-329 found a barcode scanner failure and needs to borrow 910-304 a by-tim
Yes | Visit | Complaint | RCHN | 条码扫描器故障 | 条码扫描器故障，需要借件910-304一个 |  | SWO-467446 | Re classified RMED | same day upSWO |  |  |  | Barcode scanner failure | Bar code scanner failure, need to borrow 910-304 one
No | Visit | Complaint | RCHN | 处理盘找不到位置，报错5784 | 处理盘找不到位置，报错5784，借636-533*2,903-031一个 | 更换636-533* 备件后修复by Tim | SWO-466902 | Defect | HW | Process Module | PCB, cup transport rear | 5784 | The processing disk could not find the location, the error 5784 | The processing disk could not find a location, reported error 5784, borrowed 636-533 x 2,903-031 one | Repair by Tim after replacement of 636-533-spare parts
No | Visit | Complaint | RCHN | 光学强度超范围，报错1031 | 光学单元超出范围，报错1024，需要借件903-364和944-212 14868*2 |  | SWO-467435 | Defect | HW | Process Module | Optical Unit | 1024 | Optical strength over range, error 1031 | Optical unit out of range, error 1024, need to borrow 903-364 and 944-212 14868 x2
No | Visit | Complaint | RCHN | 报错1035 | 报错1035，需要借件90-364和944-268（15132）、944-258（15185） |  | SWO-467449 | Defect | HW | Process Module | Optical Unit | 1035 | 1035 error | 1035, need to borrow 90-364 and 944-268 (15132), 944-258 (15185)
No | Visit | Complaint | RCHN | na k ca 定标失败，报错0378 | na k ca 定标失败，报错0378,需要借件945-603 |  | SWO-467453 | Assisted Replacement | Electrode | Electrode | Reference | CAPA-733 | na k ca bid failure, error 0378 | na k ca bid failure, error 0378, need to borrow 945-603
No | Visit | Complaint | RLTD | Contact Lygia D'Mello 0208 2962825 | On arrival analyser was switched off. Power lead was not fully in and battery had run flat. Re- Routed power lead to help prevent this happening again. Connected power and restarted analyser. Analyser fully working after restart. Checked data logs to see cause of issues - common error was pressure test vacuum error. After analyser had restarted I checked the pressure test values in service and they were all mid range. Analyser is working totally correctly and is ready for use. | Connected power and restarted analyser - no fault present. | SWO-468540 | No fault found | No fault found | No fault found | No fault found | Power | Contact Lygia D'Mello 0208 2962825 | On arrival analyser was switched off. Power lead was not fully in and battery had run flat. Re- Routed power lead to help prevent this happening again. Connected power and restarted analyser. Analyser fully working after restart. Checked data logs to see cause of issues - common error was pressure test vacuum error. After analyser had restarted I checked the pressure test values in service and they were all mid range. Analyser is working totally correctly and is ready for use. | Connected power and restarted analyser - no fault present.
No | Visit | Complaint | RDEU | Nach Installation keine autom. Drucken, pCO2 Fehler QC L1 | Drucker überprüft, Windowsfehlermeldung im Hintergrund offen, QC scheitert gelegentlich bei pCO2 | Fehlermeldung geschlossen und installierten Drucker Einstellungen überprüft, als Standartdrucker eingestellt, erneute Kal, QC wiederholt, i. O. | SWO-468459 | Setup | SW | RMED SW | Printer |  | Autom after installation. Printing, pCO2 Error QC L1 | Printer checks, Windows error message in background open, QC occasionally fails at pCO2 | Error message closed and installed printer settings checked, set as standard printer, re-call Kal, QC repeated, i.o.
No | Visit | Complaint | RLTD | Status code 2 error. | Site visit to repair | Arrived on site and attended analyser. Analyser found showing status code 2 error. Exited roma and attempted to replace corrupt database with blank database. Unable to do so as it was in use with sql anywhere personal server service somewhere on the analyser. Unable to find this service to stop it for repair.powered down and reinstalled windows. After reinstall windows started but hung for a very long time. Restarted and same thing happened. Unable to progress or start windows task manager. Powered down analyser and reloaded windows again. This time install was successful. Reloaded application software and rebooted. Once started, reloaded all consumables and waited for startup to complete. After startup, analyser failing cal for pco2, glucose and lactate. Rerun cal and qcs but same error occurring. Sensor cassette down to 53 tests and has been offline since Friday. Suggest sc replaced. Customer going to do this and run Thb cal. | SWO-467744 | Defect | SW | RMED SW | Database | Status code 2 | Status code 2 error. | Site visit to repair | Arrived on site and attended analyser. Analyser found showing status code 2 error. Exited roma and attempted to replace corrupt database with blank database. Unable to do so as it was in use with sql anywhere personal server service somewhere on the analyser. Unable to find this service to stop it for repair.powered down and reinstalled windows. After reinstall windows started but hung for a very long time. Restarted and same thing happened. Unable to progress or start windows task manager. Powered down analyser and reloaded windows again. This time install was successful. Reloaded application software and rebooted. Once started, reloaded all consumables and waited for startup to complete. After startup, analyser failing cal for pco2, glucose and lactate. Rerun cal and qcs but same error occurring. Sensor cassette down to 53 tests and has been offline since Friday. Suggest sc replaced. Customer going to do this and run Thb cal.
No | Visit | Complaint | RCHN | PCO2定标不过 | PCO2定标不过 借902-972一个 |  | SWO-467647 | [5 empty] | PCO2 calibration is not the | PCO2 calibration but borrowing 902-972 one
No | Visit | Complaint | RCHN | 泵轮变形，无法工作 | 泵轮变形，无法工作,需要借件902-971一个 | 更换902-971后修复(By Ran) | SWO-467375 | [5 empty] | Pump wheel deformed, not working | Pump wheel deformation, can not work, need to borrow 902-971 one | Post-replacement 902-971 repair (By Ran)
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 质控值高于统计范围 | 更换试剂包后修复 | SWO-471375 | Assisted Replacement | SP | SP | SP | QC | Test card, reagent package failure | The quality control value is higher than the statistical range | Repair after replacement of reagent kit
No | Visit | Complaint | RCHN | 光学单元超出范围，报错1031 | 光学单元超出范围，报错1031，需要借件903-364一个 | 更换903-364后修复（By Ran） | SWO-467345 | Defect | HW | Process Module | Optical Unit | 1031 | Optical unit out of range, error 1031 | Optical unit out of range, error 1031, need to borrow 903-364 one | Repair after replacement 903-364 (By Ran)
Yes | Visit | Complaint | RCHN | 光学单元超出范围，报错1031 | MYO定标不过，借944-214（14959）一张 |  | SWO-467601 | Re classified RMED | same day upSWO |  |  |  | Optical unit out of range, error 1031 | MYO calibration, borrow 944-214 (14959) a
No | Visit | Complaint | RCHN | 扫描器无法扫描 | 扫描器无法扫描 借910-339 636-666各一个 |  | SWO-467556 | Defect | HW | Inlet module | Barcode reader, Inlet | mail: send scanner to RMED | Scanner cannot be scanned | The scanner can't scan borrowing 910-339 636-666 each
No | Visit | Complaint | RCHN | 无法读取试剂包芯片信息 | 无法读取试剂包芯片信息 | 用胶水固定好原来被试剂包顶坏的配件（by chuanwei） | SWO-468339 | Defect | Consumables | Solution Pack | Solution Pack | Chip | Unable to read reagent package chip information | Unable to read reagent package chip information | Secure the original bad parts wrapped by the reagent (by chuanwei) with glue
No | Visit | Complaint | RIND | Inlet gasket leakage and issue in ph results | Inlet gasket leakage and ph results issue , | Checked the inlet gasket , run protein removal twice , cleaned the inlet gasket and gave rinse , checked the ph membrane and cleaned the ph position .Then performed calibration and run samples ,customer happy with instrument performance | SWO-469315 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | Inlet gasket leakage and issue in ph results | Inlet gasket leakage and ph results issue , | Checked the inlet gasket , run protein removal twice , cleaned the inlet gasket and gave rinse , checked the ph membrane and cleaned the ph position . Then performed calibration and run samples ,customer happy with instrument performance
No | Visit | Complaint | RFRA | Les AutoCheck ne passent plus en mode auto. La sonde d'intro ne descend plus dans la seringue lorsque l'on utilise le FlexQ | Test FlexQ en mode service. Impossible d'enregistrer les position 2 et 3 | Remplacement du FlexQ. Changement du cable Autochek. Réglage des positions du passeur en mode service, OK. | SWO-468156 | Defect | HW | FlexQ Module | FlexQ Module |  | AutoCheck no longer goes into auto mode. The intro probe no longer descends into the syringe when using FlexQ | FlexQ test in service mode. Can't save position 2 and 3 | Replacing the FlexQ. Change of the Autochek cable. Setting the smuggler's positions in service mode, OK.
No | Visit | Complaint | RDEU | 12.12.2018 Frau Freund 093727001652 meldet, einen pCO2 Fehler QC L3. 13.12.2018 Frau Freund 093727001652: pCO2-Fehler im QCLevel 3 | Empfindlichkeit der pCO2 Elektrode liegt knapp unterhalb des unteren Grenzwertes von 85%. Anwender haben mehrmals membraniert, ohne Erfolg. Im Haus unterwegs zu weiteren Service. | Einsatz einer neuen pCO2 Elektrode. Nach mehrmaliger Kalibration verändert sich die Empfindlichkeit nicht, sie bleibt bei 40%. Verwendete Elektrode Lot: BJ04, Elektrode wird reklamiert, neue bestellt. Eine Ersatz-Leihelektrode eingebaut mit Lot: RI05. Bei Lieferung der neuen pCO2 Elektrode erfolgt Einbau. Mit Ersatz Elektrode ist Qualitätskontrolle in Ordnung. | SWO-468139 | Assisted Replacement | Electrode | Electrode | pCO2 |  | 12.12.2018 Ms Friend 093727001652 reports a pCO2 error QC L3. 12/13/2018 Mrs Friend 093727001652: pCO2 error in QCLevel 3 | Sensitivity of the pCO2 electrode is just below the lower limit of 85%. Users have membraned several times, to no avail. In the house on the way to further service. | Use of a new pCO2 electrode. After several calibrations, the sensitivity does not change, it remains at 40%. Used electrode lot: BJ04, electrode is complained about, new ones are ordered. A replacement rental electrode installed with solder: RI05. When the new pCO2 electrode is delivered, installation takes place. With replacement electrode quality control is fine.
No | Visit | Complaint | RSCH | Mischerträger | Mischterträger Schlupf erkannt seit Auslieferung |  | SWO-467645 | Defect | HW | FlexQ Module | FlexQ Module | Info is in CRM: Mischerträger ausgetauscht und justiert Wellenkupplung festgezogen (war kompl lose) Referenzwerte auf Anweisung Chr. Mittl angepasst Funktionskontrolle i.o | Mischerträger | Mixed carrier slip detected since delivery
No | Visit | Complaint | RAME | During the complete OXI test, the hemolyzer failed for a Phaseregulation error. | Troubleshot to a faulty hemolyzer. | Replaced the hemolyzer, and this resolved the problem. | SWO-468230 | Defect | HW | Oxi module | Hemolyzer |  | During The complete OXI test, the Hemolyzer failed for a Phaseregulation error. | Troubleshot to a faulty hemolyzer. | Replaced the hemolyzer, and this resolved the problem.
No | Visit | Complaint | RDAN | Dorte klager over utæt pH elektrode. | Undersøgt. | Pakring for pH elektrode udskiftet, slutafprøvet. | SWO-467796 | Defect | Electrode | Electrode | Sealing ring, pH |  | Dorte complains of leaking pH electrode. | Examined. | Seal for pH electrode replaced, final tested.
No | Visit | Complaint | RDAN | Dorte klager over utæt pH elektrode samt mange fejl 500. | Undersøgt, målekammer snasket, pakring klistret fast til målekammer. | Pakring til pH elektrode udskiftet, gas $luft ventil skiftet, slutafprøvet. | SWO-467766 | Defect | Electrode | Electrode | Sealing ring, pH |  | Dorte complains of leaking pH electrode as well as many faults 500. | Checked, measuring chamber snashes, packing ring glued to measuring chamber. | Seal for pH electrode replaced, gas $luft valve switch, final tested.
No | Visit | Complaint | RDAN | Dorte klager over utæt pH elektrode. | Undersøgt, målekammer snasket, pakring sidder fast i målekammer. | Målekammer rengjort, ny pakring monteret, slutafprøvet. | SWO-467778 | Defect | Electrode | Electrode | Sealing ring, pH |  | Dorte complains of leaking pH electrode. | Checked, measuring chamber snashes, packing ring stuck in the measuring chamber. | Measuring chamber cleaned, new packing ring mounted, final tested.
Yes | Visit | Complaint | RDEU | 14.11.2018 11:00 Fr. Soll 02151363000: Probenfehler TN-T Temperatur | Probenabbrüche bei NT-proBNP. | Reparierter AQT aufgestellt, Systemüberprüfung und QC’s durchgeführt. | SWO-469253 | Re classified RMED | planned SWO |  |  |  | 14.11.2018 11:00 Fr. Should 02151363000: Sample error TN-T Temperature | Probenabbrüche bei NT-proBNP. | Repaired AQT, system verification and QC's performed.
No | Visit | Complaint | RDEU | 17.12.2018 15:30 Frau Mück 020172383539 -> Bildschirm schwarz (AQT in Betrieb und in Radiance sichtbar) | Bildschirm schwarz. | Wackelkontakt im Bildschirmkabel. Erneuerung des Bildschirmkabels. Systemüberprüfung: I.O. | SWO-468448 | Defect | HW | Cover and Chassis | Cable, display |  | 17.12.2018 15:30 Mrs. Mück 020172383539 -> Screen black (AQT in operation and visible in Radiance) | Screen black. | Wobble contact in the screen cable. Renewal of the screen cable. System check: I.O.
No | Visit | Complaint | RKOR | Cartrige door error | . | Door sensor error - Adjust. | SWO-470998 | Defect | HW | Cover and Chassis | Door, Storage |  | Cartrige door error | . | Door sensor error - Adjust.
No | Visit | Complaint | RKK | Cl キャルバックログエラー、ステータス、感度範囲外 | Cl キャルバックログエラー、ステータス、感度範囲外 | Clコンタクトプローブを未使用コンタクトプローブと入れ替えるも症状改善せず。インターフェイスPCBボード不良。代替えインターフェイスPCBへ入れ替え。正規部品入荷次第、再訪問し、代替えPCBと入れ替え予定。 | SWO-469119 | Defect | HW | Basic Unit | PCB, Sensor interface | Cl | Cl cal backlog error, status, sensitivity out of range | Cl cal backlog error, status, sensitivity out of range | Replacing the Cl contact probe with an unused contact probe does not improve symptoms. Interface PCB board bad. Replaced with an alternate interface PCB. As soon as the regular parts are in stock, we will re-visit and replace them with replacement sIPcS.
No | Visit | Complaint | RAME | Battery error | Battery error | Replace battery | SWO-468818 | Defect | HW | Battery module | Battery module | No charging | Battery error | Battery error | Replace battery
No | Visit | Complaint | RDEU | Kein Geräte upstart | Status error 2, Flüssigkeitstransportfehler | Software reinstalliert, Lösungspack ersetzt, Einlaßdichtung gereinigt | SWO-468936 | Defect | SW | RMED SW | Database | Status code 2 | No devices upstart | Status error 2, Flüssigkeitstransportfehler | Software reinstalled, solution pack replaced, inlet seal cleaned
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | LAC定标失败 | 更换测试卡后修复 | SWO-473033 | [5 empty] | Test card, reagent package failure | LAC calibration failure | Repair after replacing test card
No | Visit | Complaint | RCHN | 定标错误，气流量超范围 | 定标错误，气流量超范围 借903-111一个 | 更换903-111后修复by tim | SWO-467596 | [5 empty] | Miscalculation, air flow over range | Calibration error, air flow over range borrowed 903-111 one | Repair by tim after replacement 903-111
No | Visit | Complaint | RCHN | GLU检测结果偏高，抽吸不到样本 | GLU检测结果偏高，抽吸不到样本 | 发现科室用错了抗凝液，更换正确的抗凝液，清洁进样口后修复(by chuanwei) | SWO-468607 | User Error | preanalytical | Preanalytical | Sample handling | Glu | GLU test results are too high to draw samples | GLU test results are too high to draw samples | Found that the department used the wrong anticoagulant, replaced the correct anticoagulant, clean the sample port after repair (by chuanwei)
Yes | Visit | Complaint | RCHN | 报错1024 | 报错1024 | 自行更换903-364后修复by tim | SWO-469163 | [5 empty] | 1024 error | 1024 error | Repair by tim after replacing 903-364 yourself
No | Visit | Complaint | RCHN | 存储盘卡锁故障 | 存储盘卡锁故障，借903-378一个，测试卡断裂，断裂部位（Teeth） | 更换903-378及测试卡后修复(by chuanwei) | SWO-467718 | Defect | HW | Storage Module | Cartridge lock | 5784 | Memory card lock failure | Memory disk card lock failure, borrowing 903-378 one, test card break, broken site (Teeth) | Replace 903-378 and post-test card repair (by chuanwei)
No | Visit | Complaint | RCHN | 孵育盘报错，打印不清晰 2168 | 孵育盘报错，打印不清晰 2168,借903-138，910-286各一个。 | 更换903-138，910-286后修复（by chuanwei） | SWO-467694 | Defect | HW | Incubator Module | PCB, incubator | 2168 | Incubation plate misreported, printing is not clear 2168 | Incubate plate error, printing is not clear 2168, borrow 903-138, 910-286 each one. | Replacement 903-138,910-286 post-repair (by chuanwei)
No | Visit | Complaint | RIND | Oxi spectrum mismatch | Checked data logs and found tHb reported blank in calibration. Performed complete oximetry test, Lamp unit failed with halogen intensity low. | Performed lamp unit adjustment successfully. 2 point calibration performed successfully. | SWO-468849 | Revisit required | HW | Oxi module | Hemolyzer |  | Oxi spectrum mismatch | Checked data logs and found tHb reported blank in calibration. Performed complete oximetry test, Lamp unit failed with halogen intensity low. | Performed lamp unit adjustment successfully. 2 point calibration performed successfully.
Yes | Visit | Complaint | RKK | PHドリフト、感度エラー ドリフトはしばらく続いていた | PHドリフトが続いている 感度エラー | PH電極交換 | SWO-469054 | [5 empty] | PH drift, sensitivity error The drift had been going on for a while. | PH drift continues Sensitivity error | PH electrode replacement
No | Visit | Complaint | RDEU | Elyte Driftfehler 17:12:2018; 08:53 Uhr, Fr. Valentin, Tel. 053319344940: Kal Fehler bei Na und Ca | Kalibrations Driftgrenzen werden bei den Elektrolyten überschritten | Gerät geprüft und gereinigt, empfohlene Servicemaßnahmen durchgeführt (Reinigung BG Messkammer, Elektrodenkappe), Konfiguration im Bereich der Driftgrenzen für die Elyte angepasst, Datensicherung, Funktionskontrolle ok | SWO-468656 | Cleaned | HW | El/Met Module | measuring chamber |  | Elyte Drift Error 17:12:2018; 08:53, Ms. Valentin, tel. 053319344940: Kal error at Na and Ca | Calibration Drift limits are exceeded in electrolytes | Device tested and cleaned, recommended service measures carried out (cleaning BG measuring chamber, electrode cap), configuration in the area of drift limits adapted for the elyte, data backup, function control ok
No | Visit | Complaint | RPAC | Error 0408 Zero current error | pO2 failing calibration due to error 0408 - zero current error | Performed zero current adjustment and performed calibration. Calibration passed | SWO-469065 | Adjustment | HW | pH/BG Module | pO2 | zero current | Error 0408 Zero current error | pO2 failing calibration due to error 0408 - zero current error | Performed zero current adjustment and performed calibration. Calibration passed
No | Visit | Complaint | RFRA | Erreur de mesure oxymètrie | Module oxymètrie | Contrôle du module oxy $Constate erreur de lampe $Erreur néon position $Remplacement du module complet d'oxymètrie REF: 902-833 LOT: R2372N005 $Ajustement de la lampe $Oxy complete test ok $Calibration deux points $Calibration de la tHB en mode service $Calibration deux points $Calibration 1 point $Passage des CQ 4 niveaux $ | SWO-468551 | Defect | HW | Oxi module | Oxi module |  | Oxymeterin measurement error | Module oxymètrie | Oxy module control $Constate lamp error $Erreur neon position $Remplacement of the complete REF oxygen module: 902-833 LOT: R2372N005 $Ajustement of the lamp $Oxy complete test ok $Calibration two points $Calibration of the tHB in service mode $ Two-point calibration $Calibration 1 point $Passage of THE CQ $4 levels
Yes | Visit | Complaint | RFRA | Installation en remplacement de l’aqt90 neuf (échange sous garantie) | Installation AQT90 838R0503N007 en échange sous garantie (à la place de l’AQT90 838R0487N013) | Installation AQT90 ok. Mis sur paillasse ok. Chargement confit ok. Mis en place de la connection réseau ok. Calibration Tni ok. Passage de 3xCQ MCC3 sur Tni ok. Passage d’un patient Tni ok (test connexion concerto ok) Lancement des calibrations bhcg, Ddimer, NT-Pro, et Myo ok. Pas de cassette de Cal PCT. (Commande en cours ok) Cliente attend PCT pour inactiver ancien AQT90 838R0487N013 et remettre en caisse. Rapport à envoyer par mail à qualite@labo-barla.eu nanylutier@hotmail.com | SWO-468971 | Re classified RMED | installation report |  |  | Move to OOB trending | Installing the new aqt90 (guaranteed exchange) | Installation AQT90 838R0503N007 in exchange under warranty (in place of AQT90 838R0487N013) | Installation AQT90 ok. Put on doormat set ok. Candied loading ok. Set up the network connection ok. Tni calibration ok. Passage from 3xCQ MCC3 to Tni ok. Passing a patient Tni ok (concerto connection test ok) Launch of bhcg, Ddimer, NT-Pro, and Myo calibrations ok. No Cal PCT tape. (Order in progress ok) Customer waits for PCT to inactivate old AQT90 838R0487N013 and put back in checkout. Report to be emailed to qualite@labo-barla.eu nanylutier@hotmail.com
No | Visit | Complaint | RPOL | Wymiana taśmy od monitora | Monitor nie reaguje na dotyk | Wymiana taśmy od monitora | SWO-468636 | Defect | HW | Cover and Chassis | Cable, display |  | Replacing the tape from the monitor | Monitor does not respond to touch | Replacing the tape from the monitor
No | Visit | Complaint | RAME | QC and calibration intermittent issues with oximetry | Intermittent Thb drifting and solution errors | Installed a new Pm kit $Seems the Oximetry tubing was faulty | SWO-470693 | Defect | HW | Basic Unit | PM Kit |  | QC and calibration intermittent issues with oximetry | Intermittent Thb drifting and solution errors | Installed a new Pm kit $Seems the Oximetry tubing was faulty
No | Visit | Complaint | RDEU | 30.11.2018 09:00 Frau Holfert 03588264523 -> SP90-Wechsel-Aufforderung während Pat.-Messung 17.12.2018 13:45 Frau Holfert 03588264640 meldet einen unzuverlässigen Analysator. | Gerät meldet häufig bei BG $Oxi Kal Oxy Spectrum mismatch Fehler 581 o. Kal 3 Lösung inhomogen Gerät geht häufig in geführte Fehlersuche und will LP ersetzt haben | Austausch der Hemolyzer Unit und der Kal. internen Verschlauchung, Kal. Oxy Test i.O. Kal & QC Daten gemäß RDEU Anweisung gesichert und dokumentiert | SWO-469241 | Defect | HW | Oxi module | Hemolyzer |  | 30.11.2018 09:00 Ms. Holfert 03588264523 -> SP90 change prompt during Pat.measurement 17.12.2018 13:45 Ms. Holfert 03588264640 reports an unreliable analyzer. | Device frequently reports to BG /Oxi Kal Oxy Spectrum mismatch error 581 or Kal 3 solution inhomogeneous Device frequently goes into guided troubleshooting and wants to replace LP | Replacement of the Hemolyzer Unit and the Kal. internal hose, Kal. Oxy Test i.O. Kal & QC data backed up and documented in accordance with RDEU instructions
No | Visit | Complaint | RDEU | 18.12.2018 Herr Meinke 041617033551 meldet, dass der ABL seit zwei Wochen nicht einsatzbereit ist. | System überprüft => kein Fehler feststellbar | Einlassdichtung gereinigt, Uhrzeit richtig eingestellt, Kalibration und Qc durchgeführt | SWO-469223 | Cleaned | HW | Inlet module | Inlet gasket |  | 18.12.2018 Mr Meinke 041617033551 reports that the ABL has not been operational for two weeks. | System checked => no error detectable | Inlet seal cleaned, time set correctly, calibration and Qc performed
No | Visit | Complaint | RDEU | Oxiemetrie Hardfehfehler. | Hemolysator eingebaut, Start Phase, Fehler besteht nach wie vor. | Spectrometer bestellt. | SWO-470022 | Revisit required | HW | Oxi module | Oxi module |  | Oxiemetrie Hardfehfehler. | Hemolysator built-in, start phase, error still exists. | Spectrometer.
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 多参数定标不过 | 更换1张测试卡和1个试剂包后修复 | SWO-469764 | [5 empty] | Test card, reagent package failure | Multi-parameter calibration but | Repair after replacing 1 test card and 1 kit
No | Visit | Complaint | RCHN | 标准碳酸氢与实际碳酸氢差距过大 | 标准碳酸氢与实际碳酸氢测量结果差距过大 | 后检测系统通道号搞错导致，实际仪器无故障by tim | SWO-469198 | [5 empty] | The gap between standard bicarbonate and actual bicarbonate is too large | The gap between standard bicarbonate and actual bicarbonate measurements is too large | After the detection system channel number error led to the actual instrument fault-free by tim
No | Visit | Complaint | RCHN | 开机黑屏 | 开机黑屏，借903-381一个 | 更换903-381备件后修复（by chuanwei） | SWO-468353 | [5 empty] | Turn on the black screen | Boot black screen, borrow 903-381 one | Repair after replacement of 903-381 spare parts (by chuanwei)
No | Visit | Complaint | RCHN | LAC定标灵敏度超范围，响应错误 | LAC定标灵敏度超范围，响应错误 借903-581 一个 | 更换903-581和新测试卡后修复（by chuanwei） | SWO-467734 | Assisted Replacement | SC | SC | Lac | 1310 | LAC calibration sensitivity over range, response error | LAC calibration sensitivity over range, response error borrowed 903-581 one | Repair after replacing 903-581 and new test card (by chuanwei)
No | Visit | Complaint | RAME | Intermittent problem eating the samples | Error 374 Inhomogeneous samples | Found Ph electrode leaking replaced O-ring $Ran blood it failed first sample $Replaced reference electrode $Membrane | SWO-469565 | Revisit required | HW | pH/BG Module | Preheater |  | Intermittent problem eating the samples | Error 374 Inhomogeneous samples | Found Ph electrode leaking replaced O-ring $Ran blood it failed first sample $Replaced reference electrode $Membrane
No | Visit | Complaint | RIND | Pco2 calibration error. | Check the system. Check the calibration error. check the poc2 sensitivity. check the calibration logs. the sensitivity always on lower side. | Check the membrane & replace the membrane. new membrane also sensitivity on lower side. Replace the o-ring & Check the sensitivity. Clean chamber & Replace the gas also. still the sensitivity lower side & continuous drift for out of ranges. Cross check electrode in another system. need to check the electrode. | SWO-471050 | Assisted Replacement | Electrode | Electrode | pCO2 |  | Pco2 calibration error. | Check the system. Check the calibration error. check the poc2 sensitivity. check the calibration logs. the sensitivity always on lower side. | Check the membrane & replace the membrane. new membrane also sensitivity on lower side. Replace the o-ring & Check the sensitivity. Clean chamber & Replace the gas also. still the sensitivity lower side & continuous drift for out of ranges. Cross check electrode in another system. need to check the electrode.
No | Visit | Complaint | RDAN | Nina klager over, at pH elektroden er utæt. | . | Monteret 834-618 pakring, slutafprøvet. | SWO-469883 | Defect | Electrode | Electrode | Sealing ring, pH |  | Nina complains that the pH electrode is leaking. | . | Mounted 834-618 packing ring, final tested.
No | Visit | Complaint | RDAN | Nina klager over, at pH elektroden er utæt. | . | Monteret 834-618 pakring, slutafprøvet. | SWO-469886 | Defect | Electrode | Electrode | Sealing ring, pH |  | Nina complains that the pH electrode is leaking. | . | Mounted 834-618 packing ring, final tested.
No | Visit | Complaint | RDAN | Nina klager over pH følsomhed som svinger omkring 90%. | Undersøgt. | Kal. log. undersøgt. Der er prøvet med protein remover og elektrode cleaner. Ny monteret (BS-04). Slutafprøvet. | SWO-469880 | Assisted Replacement | Electrode | Electrode | pH |  | Nina complains of pH sensitivity which fluctuates around 90%. | Examined. | Kal. Log. Examined. There are tested with protein remover and electrode cleaner. New Mounted (BS-04). Final tested.
No | Visit | Complaint | RDEU | 19.12.2018; Sr. Laura; Tel.: 034914089030; Intervention erforderlich, Leckage erkannt 19.12.2018 Sr. Laura 034914089030 meldet, das der ABL immer wieder ein LP gewechselt haben möchte. | Intervention erforderlich Leckage erkannt Vakuumfehler seit 18.12. gegen 12 Uhr | Überprüfung Dateiarchiv => SC am 04.12. planmäßig gewechselt, LP am 06.12. wegen Kal und QC Fehler gewechselt, Einlassdichtung am 18.12. und 19.12. gewechselt => im ausgebauten Einlasshalter Dichtung nicht korrekt positioniert => Dichtung sauber in Halter gedrückt Durchführung aller Fluid Transport Programme => keine Fehler Startphase i.O. | SWO-469903 | User Error | HW | Inlet module | Inlet gasket |  | 19.12.2018; Sr. Laura; Tel.: 034914089030; Intervention required, leakage detected 19.12.2018 Sr. Laura 034914089030 reports that the ABL wants to change an LP again and again. | Intervention needed Leakage detected Vacuum fault since 18.12. around 12 o'clock | Check file archive => SC changed on 04.12. as planned, LP changed on 06.12. due to Kal and QC errors, inlet seal changed on 18.12. and 19.12.=> in the removed inlet holder seal not correctly positioned => seal cleanly pressed into holder Execution of all Fluid Transport Programs => no errors Start-up phase i.O.
No | Visit | Complaint | RDEU | Oxiemetrie Hardfehfehler. | Hemolysator ausgebaut, Spectrometer ersetzt, Startphase, Fehler besteht weiterhin. | PCB bestellt | SWO-470023 | Revisit required | HW | Oxi module | Oxi module |  | Oxiemetrie Hardfehfehler. | Hemolysator removed, spectrometer replaced, start phase, error persists. | PCB ordered
No | Visit | Complaint | RLTD | Please contact Hayley Smith Sister in charge Tel 01234 795805 or you can ask for Sarah Bailey | Analyser failing to pump solutions. | Solution pack not activated correctly, inlet gasket not cleaned. Activated solution pack and cleaned inlet gasket, | SWO-470233 | User Error | SP | SP | SP activation |  | Please contact Hayley Smith Sister in charge Tel 01234 795805 or you can ask for Sarah Bailey | Analyser failing to pump solutions. | Solution pack not activated correctly, inlet gasket not cleaned. Activated solution pack and cleaned inlet gasket,
No | Visit | Complaint | RLTD | Jerome Tel: 01480 364309 x: 4309 or Tel: 01480 830541 and ask for bleep 254 | Leak detected | Black rubber insert had become dislodged from white plastic part of inlet gasket. Refitted, now ok. | SWO-470234 | Defect | HW | Auto Inlet module | Inlet gasket | 1342 | Jerome Tel: 01480 364309 x: 4309 or Tel: 01480 830541 and ask for bleep 254 | Leak detected | Black rubber insert had become dislodged from white plastic part of inlet gasket. Refitted, now ok.
No | Visit | Complaint | RAME | Not connecting with NetOp | Netop inop for this device, Found windows firewall did not have netop program clearance | Selected netop public and private options, rebuilt windows software rame software IP address | SWO-470296 | Setup | SW | Windows | Firewall | NetOp | Not connecting with NetOp | Just inop for this device, Found Windows Firewall did not have just program clearance | Selected netop public and private options, rebuilt windows software rame software IP address
No | Visit | Complaint | RAME | Printer lid alarm not working | Missing 1298 Printer lip open | Rebooted analyzer and tested. Note: HUP conference call | SWO-470311 | Defect | SW | RMED SW | Printer lid | 1298 | Printer lid alarm not working | Missing 1298 Printer lip open | Rebooted analyzer and tested. Note: HUP conference call
No | Visit | Complaint | RDEU | 20.12.2018, 09:17 Uhr, Fr. Sellmann, Tel. 0401818828816: QC Verletzung bei pCO2 im Level A Gespräch nach Vorgabe des Gesprächsleitfadens geführt. | QC in lvl C außer Bereich, es können keine Patientenergebnisse aus dem AQURE eingesehen werden | Fehleranalyse, Einstellungen des internal remote Supports angepasst, neues LP wurde bereits gewechselt, Kalibrierung und Qualitätskontrollen i. O. | SWO-469997 | Assisted Replacement | SP | SP | pCO2 | QC A | 20.12.2018, 09:17, Ms. Sellmann, tel. 0401818828816: QC violation at pCO2 in level A conversation according to the specification of the interview guide. | QC in lvl C out of range, no patient results from AQURE can be viewed | Error analysis, internal remote support settings adjusted, new LP has already been changed, calibration and quality controls are in the way of
No | Visit | Complaint | RAME | Leaking Sp: pack completely emptied itself in <24 hours with no cycles remaining. | Leaking sp | Cleaned up dye Replaced waste manifold Only waste manifold I had was in pm kit | SWO-470277 | [5 empty] | Leaking Sp: pack completely emptied itself in <24 hours with no cycles remaining. | Leaking sp | Cleaned up dye Replaced waste manifold Only waste manifold I had was in pm kit
No | Visit | Complaint | RIND | Verification Check Failed | Verification Check Failed replace SC or SP | Check Zeorinig adjusment. Po2 value is unstable in Zeroinig. Lac value is unstable. Need to change Signal data board and sensor Board. | SWO-472499 | [5 empty] | Verification Check Failed | Verification Check Failed replace SC or SP | Check Zeorinig adjusment. Po2 value is unstable in Zeroinig. Lac value is unstable. Need to change Signal data board and sensor Board.
No | Visit | Complaint | RFRA | Fuite de l'appareil au niveau de l'évacuation. Mme DEMADE - 03.27.56.12.34 ou 03.27.56.12.30 Ouvert de 8h à 17h30 | Vérification complète de l'analyseur | Nettoyage du compartiment du pack de solution $ $ remplacement de la pompe d'évacuation des déchets et du tétons de connexion de l'évacuation des déchets $ $ relance $ $ cycle de qualité $ $ retrait du pack pour vérification : plus de fuite | SWO-470208 | [5 empty] | Leakage of the aircraft at the evacuation level. Mrs DEMADE - 03.27.56.12.34 or 03.27.56.12.30 Open 8am-5.30pm | Full verification of the analyzer | Solution pack compartment cleaning $$ replacement of waste disposal pump and waste disposal connection nipples $$ raise $$ $quality cycle $$ withdrawal of pack for verification: no more leakage
No | Visit | Complaint | RCHN | BNP温度传感器已损坏 | BNP温度传感器已损坏，借905-908一个 | 更换905-908后修复（by chuanwei） | SWO-469269 | Defect | HW | Process Module | Sensor, BnP | 1235 | BNP temperature sensor damaged | BNP temperature sensor is damaged, borrow 905-908 one | Repair after replacement 905-908 (by chuanwei)
No | Visit | Complaint | RCHN | 光学单元超出范围，报错1031 | 光学单元超出范围，报错1031,需要借件903-364、944-212（15116） 944-258（15025）缺货 | 更换903-364，重新定标后修复（by chuanwei） | SWO-469074 | Defect | HW | Process Module | Optical Unit | 1031 | Optical unit out of range, error 1031 | Optical unit out of range, error 1031, need to borrow 903-364, 944-212 (15116) 944-258 (15025) out of stock | Replace 903-364, re-calibration repair (by chuanwei)
No | Visit | Complaint | RPAC | sampling issues | Many sampling issues on patient samples. Also errors relating to inhomagenous samples and liquid arriving at liquid sensor before it is expected. | Investigate the machine and found that the LS on the ELMET lower was quite low value (20) so replaced this (Now 23) Also replaced the rinse valve on the ELMET as it was sounding hollow and not moving properly even though it would not fail the actual service test. The top chamber silo on the ELMET module was failing and coming off a little bit fromthe base. This was quite clear to see after I opened up the module. Replaced this chamber The GLU electrode was tarnished so I replaced that and its electrode cover, and the membrane itself. Reassembled then initialized then ran LS adjust and pump cal. and checked valves Ran 2 point cals till all parameters were in spec. | SWO-469756 | Defect | HW | El/Met Module | El/Met Module |  | sampling issues | Many sampling issues on patient samples. Also errors relating to inhomagenous samples and liquid arriving at liquid sensor before it is expected. | Investigate the machine and found that the LS on the ELMET lower was quite low value (20) so replaced this (Now 23) Also replaced the rinse valve on the ELMET as it was sounding hollow and not moving properly even though it would not fail the actual service test. The top chamber silo on the ELMET module was failing and coming off a little bit fromthe base. This was quite clear to see after I opened up the module. Replaced this chamber The GLU electrode was tarnished so I replaced that and its electrode cover, and the membrane itself. Reassembled then initialized then ran LS adjust and pump cal. and checked valves Ran 2 point cals till all parameters were in spec.
No | Visit | Complaint | RDEU | LS Oxifehler | LS Oxifehler | Air und Wastekanal gespült. Proteinentfernung Ls-Justage und Pumpekal durchgeführt. Cl Dummy O-Ring erneuert. | SWO-470422 | Defect | HW | Oxi module | LS |  | LS Oxifehler | LS Oxifehler | Air and waste channel. Protein removal Ls-adjustment and pump scale. Cl Dummy O-Ring renewed.
No | Visit | Complaint | RIND | Leaking from inlet gasket | Checked the analyzer found clot inlet gasket & waste was blocked | Replaced waste tube and cleaned inlet gasket | SWO-470500 | Defect | HW | Basic Unit | Waste tube | Clot | Leaking from inlet gasket | Checked the analyzer found clot inlet gasket & waste was blocked | Replaced waste tube and cleaned inlet gasket
No | Visit | Complaint | RAME | Rinse Error | Rinse Error that would not clear. | Replaced top rinse valve. | SWO-471324 | Defect | HW | pH/BG Module | Valve |  | Rinse Error | Rinse Error that would not clear. | Replaced top rinse valve.
No | Visit | Complaint | RSAF | Analyzer not switching on :burning smell from the machine | Power supply fuses ok, Control PCB fuses ok, interface cable for display unit short , screen and CPU unit PCB possibly short also. | CPU unit to be replaced. | SWO-470053 | Defect | HW | DMS_PC | CPU |  | Analyzer not switching on :burning smell from the machine | Power supply fuses ok, Control PCB fuses ok, interface cable for display unit short , screen and CPU unit PCB possibly short also. | CPU unit to be replaced.
Yes | Visit | Complaint | RSCH | Ca, PH und PCO2 ausser Bereich im QC | Ca, PH und PCO2 ausser Bereich im QC | URSST hat Gerät ausgeliefert in Liestal. Probleme mit diverse QC‘s PH, PCO2 und Ca. Urs hat Ref neu membraniert. PH Dichtung ers. Alle Parameter ausser PCO2 QC 3 ok. PCO2 Membrane ers. System i.o. Kal. + QC + Thb i.o. | SWO-470037 | Re classified RMED | installation report |  |  | Move to OOB trending | Ca, PH and PCO2 out of range in the QC | Ca, PH and PCO2 out of range in the QC | URSST has shipped device in Liestal. Problems with various QC's PH, PCO2 and Ca. Urs has re-membraneed Ref. PH seal ers. All parameters with the third time p.o. PCO2 QC 3 ok. PCO2 membrane ers. System i.o. Kal. + QC + Thb i.o.
No | Visit | Complaint | RFRA | Pendant la calibration de troponine, la cliente s'est connecté pour voir le temps restant et l'application s'est fermée. La cliente a redémarré l'appareil et s'est mis en erreur 1999. Mme SEAQUET 06.13.26.33.75 | Impossibilité de passer en mode SAV, erreur Pisa.exe Plusieurs fois idem. | Démontage AQT90 ok. Vérification mécanique de la pipette ok. Relance soft AQT90 part le bureau directement ok. Initialisation mécanique et fluidique ok, pipette init ok, plusieurs cupules sont évacuées ok. Erreur logiciel incompatible. Sauvegarde, servicedump et Re-installation du soft ok. Initialisation ok, plus d’erreur ok. Passage en mode SAV sans erreur plusieurs fois ok. Cassette de cal Tnt sortie (entamée). Contrôle système ok. Validation Radiometer ok. Passage 3x CQ MCC3 ok. Rapport à envoyer par mail à isabelle.bonnefille@labosud.fr | SWO-470255 | Defect | SW | RMED SW | RMED SW | 1999 | During troponin calibration, the client logged in to see the remaining time and the application closed. The customer restarted the device and erred in 1999. Ms. SEAQUET 06.13.26.33.75 | Unable to switch to SAV mode, Pisa.exe error Several times ditto. | Disassembly AQT90 ok. Mechanical check of the pipette ok. Relaunch soft AQT90 leaves the office directly ok. Mechanical and fluidization ok, pipette init ok, several cupules are evacuated ok. Incompatible software error. Backup, servicedump and re-installation of the app ok. Initialization ok, no more error ok. Switch to SAV mode without error several times ok. Tnt cal cassette output (started). System control ok. Radiometer validation ok. Passage 3x CQ MCC3 ok. Report to be emailed to isabelle.bonnefille@labosud.fr
No | Visit | Complaint | RKK | バッテリー チャージ不良 | バッテリー チャージ不良 | バッテリー交換 装置分解・内部点検 OXIバルブ整備実施 CAL・QC確認 | SWO-471352 | Defect | HW | Battery module | Battery module | Charging | Battery charge bad | Battery charge bad | Battery replacement Equipment decomposition and internal inspection Implemented oxi valve maintenance CAL/QC Confirmation
Yes | Visit | Complaint | RDEU | 21.12.2018/09:00 Herr Kowalski 023052945343 -> QK-fehler bei pCO2-Level A zu hoch 21.12.2018/12:00 Hr. Kowalski bittet um RÜCKRUF/SERVICE | pCO2 bei QC Level1 grenzwertig hoch | SC geliefert und installiert. | SWO-470920 | Re classified RMED | Logistical upSWO |  |  |  | 21.12.2018/09:00 Mr Kowalski 023052945343 -> QK error at pCO2 level A too high 21.12.2018/12:00 Mr. Kowalski asks for CALL/SERVICE | pCO2 at QC Level1 borderline high | SC is supplied and installed.
No | Visit | Complaint | RDEU | Clot detected , keine manuelle Spülung möglich | Verstopfung | Spülungsversuch fehlgeschlagen, mehrfach, Hämolysator ersetzt, Alle Schläuche ersetzt, Startphase, Thb kal , alles i.o. | SWO-472034 | Defect | HW | Oxi module | Hemolyzer | Clot | Clot detected , no manual rinsing possible | Constipation | Rinse attempt failed, multiple times, hemolyser replaced, all hoses replaced, start phase, thb kal, everything i.o.
No | Visit | Complaint | RDEU | Meßabbrüche | Mehrere Meßabbrüche Probenaspiration gescheitert 1253, Keine Probe währen der Ansaugung 331 | Gerät überprüft, vermutlich Handhabungsfehler, Einlaßdichtung ersetzt, Messungen o.k. | SWO-470943 | User Error | Sample handling | Sample handling | Sample handling |  | Meßabbrüche | Several measurement interruptions Sample aspiration failed 1253, No sample during suction 331 | Device checked, presumably handling error, inlet seal replaced, measurements ok
No | Visit | Complaint | RCHN | 无法识别处理包 | 无法识别处理包 | 更换新试剂包后修复 | SWO-470599 | Defect | Consumables | Solution Pack | Solution Pack | Chip | The processing package is not recognized | The processing package is not recognized | Repair after replacing a new kit
No | Visit | Complaint | RCHN | 定标底物外漏，1024 | 定标底物外漏，1024，借903-364和903-702各一个 | 更换903-364后重新定标修复，903-702未使用（by chuanwei） | SWO-469934 | Defect | HW | Process Module | Optical Unit | 1024 | Calibration substrate leakage, 1024 | Bid substrate leakage, 1024, borrow 903-364 and 903-702 each | Re-calibration repair after replacement of 903-364, 903-702 not used (by chuanwei)
No | Visit | Complaint | RCHN | 扫描头故障 | 扫描头故障，需要借件902-924 | 更换902-924后修复(by chuanwei) | SWO-469782 | Defect | HW | Cover and Chassis | Barcode reader, display |  | Scan head failure | Scan head failure, need to borrow 902-924 | Repair after replacement 902-924 (by chuanwei)
No | Visit | Complaint | RCHN | HCT值测量结果异常 | HCT值测量错误， 借903-329一个 | 更换903-329后修复（by chuanwei） | SWO-469840 | Defect | HW | Aspiration Module | Needle Wash Unit | 1267 1258 1223 | Abnormal measurement of HCT values | HCT value measurement error, borrow 903-329 one | Repair after replacement 903-329 (by chuanwei)
No | Visit | Complaint | RCHN | 做样本失败，报错1999.50.15.1008.0.-1.-1 | 做样本失败,报错1999.50.15.1008.0.-1.-1，借636-652一个 | 更换636-652备件后修复(by chuanwei) | SWO-469126 | Defect | HW | Storage Module | Cable, barcode reader |  | Failed to make samples, error 1999.50.15.1008.0.-1.1 | Sample failure, error 1999.50.15.1008.0.-1.1, borrow 636-652 one | Repair after replacement of 636-652 spare parts (by chuanwei)
No | Visit | Complaint | RCHN | 干燥前温度超出范围，报错1068 | 干燥前温度超出范围，报错1068 需要借件903-210 | 更换903-210后修复(by chuanwei) | SWO-469841 | Defect | HW | Process Module | PCB, cup dryer | 1068 | Temperature out of range before drying, error 1068 | Temperature out of range before drying, report 1068 requires a loan 903-210 | Repair after replacement 903-210 (by chuanwei)
No | Visit | Complaint | RCHN | 测试卡断裂 | 测试卡断裂，断裂部位（Teeth） | 更换测试卡后修复 by Andrew | SWO-471169 | Defect | Consumables | Test Kit | Cartridge, Broken | Teeth | Test card break | Test card break, break site (Teeth) | Fix by Andrew after replacing test card
No | Visit | Complaint | RAME | Intermittent problem eating the samples | Analyzer is intermittently presenting message 374 Inhomogeneous sample when running patient blood samples. Question marks on all parameters. | 12-20-18 Thoroughly cleaned the pH measuring Chamber. Replaced the pH electrode seal ring. Replaced the inlet rubber tubing for the inlet probe for the inlet module. Ran 16 blood samples and, no errors presented. 12-21-18 Today, Mark stated the analyzer presented errors on two blood samples. I tested the instrument with blood and, I saw the error 374. Replaced the pH $BG measuring chamber, the pH $BG lower liquid sensor and, the air and, rinse valves for the pH $BG module. Replaced the pH electrode jacket. Replaced all pump tubing from PM kit and, checked all pump rotor , found ok. Replaced the nipple kit from spares. Replaced the rubber tubing at the inlet liquid sensor. Replaced the inlet gasket. Cleaned the sample preheater with protein remover and, water. Performed the service mode liquid sensor adjust and, ran pump calibrations, all passing within specifications. I ran a blood sample and, I did not see the error again. This is an error that intermittently shows when running blood samples. I was not able to find more samples to run. After today's troubleshooting, the therapist will continue to use the analyzer and, will contact us if the problem reoccurs. | SWO-470295 | Revisit required | HW | pH/BG Module | Preheater |  | Intermittent problem eating the samples | Analyzer is intermittently presenting message 374 Inhomogeneous sample when running patient blood samples. Question marks on all parameters. | 12-20-18 Thoroughly cleaned the pH measuring Chamber. Replaced the pH electrode seal ring. Replaced the inlet rubber tubing for the inlet probe for the inlet module. Ran 16 blood samples and, no errors presented. 12-21-18 Today, Mark stated the analyzer presented errors on two blood samples. I tested the instrument with blood and, I saw the error 374. Replaced the pH $BG measuring chamber, the pH $BG lower liquid sensor and, the air and, rinse valves for the pH $BG module. Replaced the pH electrode jacket. Replaced all pump tubing from PM kit and, checked all pump rotor , found ok. Replaced the nipple kit from spares. Replaced the rubber tubing at the inlet liquid sensor. Replaced the inlet gasket. Cleaned the sample preheater with protein remover and, water. Performed the service mode liquid sensor adjust and, ran pump calibrations, all passing within specifications. I ran a blood sample and, I did not see the error again. This is an error that intermittently shows when running blood samples. I was not able to find more samples to run. After today's troubleshooting, the therapist will continue to use the analyzer and, will contact us if the problem reoccurs.
No | Visit | Complaint | RDEU | Zeitweise Elektrolytdriften, Rilibäkverletzungen pCO2 | Zeitweise Elektrolytdriften, pCO2 Rilibäkfehler | Gerät überprüft, Meßkammer gesäubert, pH Dichtungsring ersetzt, Ref. Elektrode aus PKW Bestand ersetzt, diverse Elektroden membraniert | SWO-470579 | Assisted Replacement | Electrode | Electrode | Reference | CAPA-733 | Temporary electrolyte drifts, rileb injuries pCO2 | Temporary electrolyte drifts, pCO2 rilibble error | Device checked, measuring chamber cleaned, pH sealing ring replaced, ref. Electrode replaced by passenger car stock, various electrodes membraneed
No | Visit | Complaint | RDEU | 21.12.2018 12:00 H. Kreinz 065194783788: Status rot ; keine Bedienung möglich | Sicht- und Funktionskontrolle | Gerät geöffnet und Wetsection überprüft Fehler3414 un 3412 Ftc_105, 103 Ansaugen der Solution aus Lösungspack auch manuell nicht möglich, freie Flüssigkeit im Solutionpack - Defekt. Neues Solutionpack ist unterwegs und wird morgen eingesetzt, Kunde meldet sich ggf. Solutionpack zwecks Überprüfung mitgenommen | SWO-470817 | Defect | Consumables | Solution Pack | Solution Pack, valve | 3414, 1052, 1067 | 21.12.2018 12:00 H. Kreinz 065194783788: Status red ; no operation possible | Visual and functional control | Device opened and wetsection checks error3414 un 3412 Ftc_105, 103 Suction of the solution from solution pack also manually not possible, free liquid in the solution pack - defect. New Solutionpack is on the way and will be used tomorrow, customer may contact us Solutionpack taken with you for verification
No | Visit | Complaint | RFRA | Erreur 3412 et 3415, pack de solution non dispo. Contact : Corine 04.66.59.87.10 | Erreur 3415 : FTC 106 AQT90 neuf Au bout de 2 jours erreur et sur plusieurs pack solution. | Dans SAV : erreur Pisa.exe Puis au redémarrage erreur supplémentaire logiciel incompatible. ServiceDump récupéré ok. Rechargement soft 8.11 ok. Même erreur Pisa en passage mode SAV. Rechargement soft 8.11 en clean install : erreur firmware. Relance AQT90 ok. Plus d’erreur Pisa.exe. Dans service safe position ok. Needle wash déborde, pompe 4 et 5 ne tourne pas. Échange carte wetsection et moteur P5 en DIV ok. Changement de tous les corps de pompe ok. Relance ok, needle wash ne déborde plus ok. P4 et p5 tourne ok. Test p1 $p2 ok. Test p4 $p5 NOK (liquide évacué mais valeur de hct ne fluctue pas) Changement filtre en DIV ok. Toujours même Pbm. Erreur sur test p4 $p5. Plus de pièce disponible. Client n’utilise pas le système pdt quelques mois, retour atelier pour réparation. Rapport à envoyer par mail à Beaucaire@labosud.fr | SWO-470829 | Defect | HW | Wet section | P4/P5 | 3412, 3415 | Error 3412 and 3415, solution pack not available. Contact: Corine 04.66.59.87.10 | Erreur 3415 : FTC 106 AQT90 neuf Au bout de 2 jours erreur et sur plusieurs pack solution. | In SAV: Pisa.exe error Then restart additional incompatible software error. ServiceDump recovered ok. Soft recharging 8.11 ok. Same Pisa error in passing SAV mode. Soft recharging 8.11 in clean install: firmware error. Relaunch AQT90 ok. No more Pisa.exe error. In service safe position ok. Needle wash overflows, pump 4 and 5 does not turn. Exchange wetsection card and P5 engine in DIV ok. Change all pump bodies ok. Relaunch ok, needle wash no longer overflows ok. P4 and p5 turns ok. Test p1 $p2 ok. Test p4 $p5 NOK (liquid evacuated but hct value does not fluctuate) Change filter to DIV ok. Still even Pbm. Error on test p4 $p5. More room available. Customer does not use the pdt system a few months, return workshop for repair. Report to be emailed to Beaucaire@labosud.fr
No | Visit | Complaint | RPOL | Błędy kalibracyjne | Błąd kalibracji elektrolitow. | Sprawdzenie aparatu. | SWO-471247 | Assisted Replacement | SC | SC | SC | Cal | Calibration errors | Electrolyte calibration error. | Check the camera.
No | Visit | Complaint | RAME | Analyzer down. Liquid sensor error detected. Tried to worked with FSE to replace sp/sc to no avail. Liquid sensors at sensor bay continue to fail. Service visit needed. | Confirmed issue when onsite. Tried to reset cassette bay sensors via service program but both would continuously remain beyond 254, limits of acceptance. | Performed flushing without any obstruction. Rerun liquid sensor calibration again but would remain above 254 units, beyond passing limits. Reboot analyzer to clear off system memory. Then rerun liquid cals again. Now values are below 60, well within passing limits. Reviewed datalog and found power surge on 12 $19 $2018 at 302pm when analyzer rebooted. Since then, liquid sensors started to surface with both E529 and E531, Inlet and Sensors LS failed to calibrate respectively. Will convey to Pablo Bracho, lab director importance to install Uninterrupted Power Supply with conditioner feature to level off power surge impact. | SWO-470871 | Defect | SW | RMED SW | RMED SW | LS | Analyzer down. Liquid sensor error detected. Tried to worked with FSE to replace sp/sc to no avail. Liquid sensors at sensor bay continue to fail. Service visit needed. | Confirmed issue when onsite. Tried to reset cassette bay sensors via service program but both would continuously remain beyond 254, limits of acceptance. | Performed flushing without any obstruction. Rerun liquid sensor calibration again but would remain above 254 units, beyond passing limits. Reboot analyzer to clear off system memory. Then rerun liquid cals again. Now values are below 60, well within passing limits. Reviewed datalog and found power surge on 12 $19 $2018 at 302pm when analyzer rebooted. Since then, liquid sensors started to surface with both E529 and E531, Inlet and Sensors LS failed to calibrate respectively. Will convey to Pablo Bracho, lab director importance to install Uninterrupted Power Supply with conditioner feature to level off power surge impact.
No | Visit | Complaint | RCHN | CL定标失败 | CL定标失败，借902-972一个 | 更换902-972后修复(by chuanwei) | SWO-470360 | [5 empty] | CL calibration failed | CL bid failure, borrow 902-972 one | Repair after replacement 902-972 (by chuanwei)
No | Visit | Complaint | RCHN | 重启分析仪 | 重启分析仪，借944-394（15209）一张 | 重新定标后修复（by chuanwei） | SWO-470509 | Adjustment | HW | Incubator Module | Punch up | 5297 | Restart the analyzer | Restart the analyzer, borrow 944-394 (15209) | Re-calibration repair (bywei chuan)
No | Visit | Complaint | RKOR | Sample clot error | . | Gasket module clean. Liquid flow line clean. Service mode check. Sample test. | SWO-471013 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway | Clot | Sample clot error | . | Gasket module clean. Liquid flow line clean. Service mode check. Sample test.
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | LAC,HCT,NA,K定标失败 | 更换测试卡后修复 | SWO-473027 | [5 empty] | Test card, reagent package failure | LAC, HCT, NA, K calibration failed | Repair after replacing test card
No | Visit | Complaint | RCHN | 无法开机 | 排线腐蚀，无法开机，借636-542一个 | 更换636-542后修复（by chuanwei） | SWO-470425 | Defect | HW | Wet section | Cable, wet section |  | Cannot power on | Line corrosion, can not be turned on, borrow 636-542 one | Repair after replacement 636-542 (by chuanwei)
No | Visit | Complaint | RPOL | Aparat nie przyjmuje nowego pakietu | Nie zaobserwowano uszkodzenia aparatu. Nowy pakiet uruchomił się bez problemu. | Wymiana pakietu na nowy BX-05 (seria niepewna, zostanie wysłany jeszcze jeden pakiet z biura). Zainstalowano oprogramowanie 3.4 MR1 | SWO-471019 | Assisted Replacement | SP | SP | Chip |  | The camera does not accept a new package | No damage to the camera was observed. The new package has started without a problem. | Replacing the package with the new BX-05 (Uncertain series, one more package will be sent from the office). Installed 3.4 Software MR1
Yes | Visit | Complaint | RSAF | no oximetry resuls on patient and qc results, and tHb fails on calibrations. Restarted analyzer but calibrations still fail tHb. | Complete oxi test fails, no blockages in hemolyzer | Replaced hemolyzer unit, and no solution flowing after replacememnt. New Hemolyzer blocked. Will return with another hemolyzer unit. Old Hemolyzer reinstalled. | SWO-471306 | Re classified RMED | same day upSWO |  |  |  | no oximetry resuls on patient and qc results, and tHb fails on calibrations. Restarted analyzer but calibrations still fail tHb. | Complete oxi test fails, no blockages in hemolyzer | Replaced hemolyzer unit, and no solution flowing after replacememnt. New Hemolyzer blocked. Will return with another hemolyzer unit. Old Hemolyzer reinstalled.
No | Visit | Complaint | RSAF | no oximetry resuls on patient and qc results, and tHb fails on calibrations. Restarted analyzer but calibrations still fail tHb. | Return from previous visit | Replaced hemolyzer unit. Complete oxi test still failed. posible Spectrometer or wet section control PCB fault. | SWO-471307 | Revisit required | HW | Oxi module | Oxi module |  | no oximetry resuls on patient and qc results, and tHb fails on calibrations. Restarted analyzer but calibrations still fail tHb. | Return from previous visit | Replaced hemolyzer unit. Complete oxi test still failed. posible Spectrometer or wet section control PCB fault.
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 定标状态-SC | 更换测试卡后修复 | SWO-473038 | [5 empty] | Test card, reagent package failure | Bid Status-SC | Repair after replacing test card
No | Visit | Complaint | RCHN | 测试卡验证失败 | 测试卡验证失败 | 更换库存机上的902-972后修复 by Andrew | SWO-471431 | [5 empty] | Test card validation failed | Test card validation failed | Repair by Andrew after replacing 902-972 on the stock machine
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 不均匀的定标3溶液 1539 | 更换试剂包后修复 | SWO-471546 | Assisted Replacement | SP | SP | FT | 1539 | Test card, reagent package failure | Uneven calibration 3 solution 1539 | Repair after replacement of reagent kit
No | Visit | Complaint | RCHN | 最大暗视野超范围 | 最大暗视野超范围 借903-364一个 | 更换903-364后修复 by Andrew | SWO-471189 | Defect | HW | Process Module | Optical Unit | 1035 | Maximum dark field of view over range | Maximum dark field of view over range borrowed 903-364 one | Repair by Andrew after replacing 903-364
No | Visit | Complaint | RCHN | 报错2168，定标-42 | 报错2168，定标-42 | 清洁孵育盘传感器后修复by Anderw | SWO-471503 | Cleaned | HW | Incubator Module | PCB, incubator | 2168 | 2168, bid-42 | 2168, bid-42 | Repair by Anderw after cleaning incubation disk sensor
No | Visit | Complaint | RCHN | 仓门故障 | 仓门故障，借886-691 $2个，886-638,886-639各一个 | 更换886-638,886-639，2个886-691后修复 by Andrew | SWO-470970 | Defect | HW | Cover and Chassis | Door, Inlet |  | Warehouse failure | Warehouse failure, borrow886-691 /2, 886-638, 886-639 each | Replace 886-638, 886-639, 2 886-691 after repair by Andrew
No | Visit | Complaint | RCHN | 测试卡断裂，报错5784 5689 | 测试卡断裂，报错5784 5689 断裂部位（Teeth） | 更换测试卡后修复 by Andrew | SWO-471398 | Defect | Consumables | Test Kit | Cartridge, Broken | Teeth | Test card broken, error 5784 5689 | Test card broken, error 5784 5689 Break site (Teeth) | Fix by Andrew after replacing test card
No | Visit | Complaint | RCHN | 屏幕黑屏 | 屏幕黑屏，借910-321一个 | 更换910-321后修复by Andrew | SWO-471135 | Defect | HW | DMS_PC | CPU |  | Screen black screen | Screen black screen, borrow 910-321 one | Repair by Andrew after replacement 910-321
No | Visit | Complaint | RKK | 納品して4か月くらいから印字が欠けはじめた | ﾌﾟﾘﾝﾀﾍｯﾄﾞ不良 | ﾌﾟﾘﾝﾀﾍｯﾄﾞ交換 | SWO-471531 | Defect | HW | Printer Module | Printer Head |  | After delivery, the print began to be missing from about four months. | Bad printer head | Printer head replacement
No | Visit | Complaint | RKK | GTモードが多発する | 機器状況確認 インレットガスケットが、旧タイプに 交換されていた。 | インレットガスケット再交換。 インレットプローブ洗浄及び変形チェック。IRセンサー検出部チェック。 ポンプローターチェック。インレット $排液ライン洗浄。フィールドトランスポートテストOK。OXI,pO2コンプリートチェックOK。 CALチェックOK、QCチェック(ABC)OK サンプル測定テストOK | SWO-471825 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | GT mode frequently | Equipment status check The inlet gasket is had been replaced. | Inlet gasket re-replacement. Inlet probe cleaning and deformation check. Ir sensor detection unit check. Pump rotor check. Inlet $ drainage line cleaning. Field transport test OK. OXI, pO2 Complete Check OK. CAL check OK, QC check (ABC) OK Sample measurement test OK
No | Visit | Complaint | RFRA | Impossible de changer la cassette la mâchoire reste ouverte . | Mâchoire cassette ouverte , impossible de la fermer dans le menu service. | Démontage du système d'introduction cable légèrement déconnecter, déconnection du cable . Calibration dû système d'introduction ok, ouverture fermeture mâchoire ok. Changement de cassette ok , lactate en cours de stabilisation. | SWO-471599 | Adjustment | HW | Basic Unit | Sensor door |  | Can't change the tape the jaw remains open. | Jaw cassette open, impossible to close it in the service menu. | Disassembly of the cable introduction system slightly disconnect, disconnect the cable . Calibration due introduction system ok, jaw closing aperture ok. Change of cassette ok, lactate being stabilized.
No | Visit | Complaint | RCHN | 抽吸样本失败 | 抽吸样本失败，借903-129一个 | 更换903-129备件后修复 by Andrew | SWO-471444 | Defect | HW | Oxi module | Pinch valve | 1271 | Failed to pump samples | Failed to suck samples, borrow 903-129 one | Repair by Andrew after replacement of 903-129 spare parts
No | Visit | Complaint | RCHN | 光学单元超出范围 | 光学单元超出范围。需要借件903-364一个 | 更换903-364后修复 by Andrew | SWO-471366 | Defect | HW | Process Module | Optical Unit | mail: which error code? Which type of error? | Optical unit out of range | The optical unit is out of range. Need to borrow 903-364 one | Repair by Andrew after replacing 903-364
No | Visit | Complaint | RCHN | 重复死机 | 重复死机，借944-216（15107）一个 | 重新定标后修复 by Andrew | SWO-471434 | Assisted Replacement | Consumables | Test Kit | Cartridge, cal | 5684 | Repeat the crash | Repeat the crash, borrow 944-216 (15107) one | Fix by Andrew after re-calibration
No | Visit | Complaint | RCHN | 提示样本不足 | 提示样本不足 | 清洁HCT模块正常。by Andrew | SWO-472563 | Cleaned | HW | Aspiration Module | Needle Wash Unit |  | Hint slack sample | Hint slack sample | Clean the HCT module properly. by Andrew
No | Visit | Complaint | RIND | Oxi spectrum mismatch | Performed complete Oximetry Test successfully. All parameters were passed. | Replaced Hemolyzer unit, performed 2 point calibration successfully. | SWO-472551 | Defect | HW | Oxi module | Hemolyzer |  | Oxi spectrum mismatch | Performed complete Oximetry Test successfully. All parameters were passed. | Replaced Hemolyzer unit, performed 2 point calibration successfully.
No | Visit | Complaint | RAME | Intermittent problem eating the samples | Error 374 | Replaced the preheater and preheater tubings $Ran six samples without errors | SWO-471686 | Defect | HW | pH/BG Module | Preheater |  | Intermittent problem eating the samples | Error 374 | Replaced the preheater and preheater tubings $Ran six samples without errors
No | Visit | Complaint | RIND | pO2 QC low | Checked the data logs pO2 QC low for level1 & 2 found puss in pO2 & pCO2 chambers | Cleaned the pO2, pCO2 chamber and decontamination done performed QC all the values within the range | SWO-472618 | Cleaned | HW | pH/BG Module | measuring chamber |  | pO2 QC low | Checked the data logs pO2 QC low for level1 & 2 found puss in pO2 & pCO2 chambers | Cleaned the pO2, pCO2 chamber and decontamination done performed QC all the values within the range
No | Visit | Complaint | RDEU | 27.12.2018 07:30Uhr Frau Wellmann 00401818827721 (med Technik): Die Nachricht kam per Mailbox um 0:00 an. Das Gerät würde seit der Aufstellung unzuverlässig laufen. Das Gerät würde ständig Gerinnsel erkannt melden und ungewöhnlich laut sein, wenn der LP gewechselt wird. Der Benutzer ist verunsichert, weil die Software sich beim Wechseln vom LP aufgehängt hat. | Analysator läuft seit Aufstellung nicht durch. Ständig Fehlermeldungen -Inhomogene Flüssigkeit und Gas. Fehler LS Zustand nach Probe. | Nach Rücksprache mit Frau Dr. Händel gab es bei der Aufstellung des Analysators keine passenden Clot Catcher vor Ort ca 1Woche. Aus Unwissenheit wurden dann die vom 800 er benutzt. Im Servicemenü den Bereich Fluid Transport alle Einstellungen mehrfach abgearbeitet. Mehrere Kalibrationen und QC durchgeführt. Kein Fehler festgestellt. Frau Dr.Händel sagte mir auch, dass bei der Einweisung der richtige Gerbrauch des Spülsets nicht gezeigt wurde. Bzw. das Set bei Auslieferung gar nicht vorhanden war. | SWO-471983 | User Error | Clot catcher | Clot catcher | Clot catcher |  | 27.12.2018 07:30Pm Ms Wellmann 00401818827721 (med Technik): The message arrived by mailbox at 0:00. The device would run unreliably since installation. The device would constantly report clots detected and be unusually noisy when the LP is changed. The user is unsettled because the software has hung up when switching from the LP. | Analyzer hasn't been running since installation. Constantly error messages -Inhomogeneous liquid and gas. Error LS state after sample. | After consultation with Dr. Handel, there were no suitable Clot catchers on site for about 1 week. Out of ignorance, the 800 he was then used. In the service menu, the Fluid Transport area has processed all settings several times. Several calibrations and QC performed. No error detected. Dr.Händel also told me that the correct use of the rinsing set was not shown during the briefing. The set was not available at all on delivery.
Yes | Visit | Complaint | RAME | Broken display housing. | Broken display housing. | Replaced display unit and tested. Note: Worked on installing 36 analyzers. | SWO-471680 | Re classified RMED | installation report |  |  | Move to OOB trending | Broken display housing. | Broken display housing. | Replaced display unit and tested. Note: Worked on installing 36 analyzers.
No | Visit | Complaint | RAME | Possible Blockage | No problems found upon arrival. Customer had ran QC and everything was within range. Went ahead with plan to perform PM anyway since it is due within 1 month. | Performed PM. See separate work order for PM. Ran tHb calibration after PM. Results were within range. | SWO-472266 | [5 empty] | Possible Blockage | No problems found upon arrival. Customer had ran QC and everything was within range. Went ahead with plan to perform PM anyway since it is due within 1 month. | Performed PM. See separate work order for PM. Ran tHb calibration after PM. Results were within range.
No | Visit | Complaint | RESP | Electrodo | Comprobación del funcionamiento del electrodo | Se sustituye el electrodo por uno nuevo, se le coloca una membrana nueva y se calibra. | SWO-472665 | [5 empty] | Electrode | Checking the electrode's operation | The electrode is replaced with a new one, a new membrane is placed and calibrated.
No | Visit | Complaint | RCHN | 定标失败 | 定标失败,需要借件944-216（15082） | 重新定标后修复 by Andrew | SWO-471439 | Assisted Replacement | Consumables | Test Kit | Cartridge, cal |  | Bid failed | Failed bid, 944-216 (15082) required | Fix by Andrew after re-calibration
No | 4 | Y | RCZE | Blank |  |  | D5723FB158476170 | Setup | SW | RMED SW | Database |  | Blank
No | Visit | Complaint | RKK | 印字がかすれる | 印字かすれ | プリンタユニット交換 | SWO-472381 | Defect | HW | Printer Module | Printer Module |  | The print is blurred. | Printing blurry | Printer unit replacement
Yes | Visit | Complaint | RFRA | Erreur T° ampoule QC. | cf swo | cf swo | SWO-472639 | Re classified RMED | same day upSWO |  |  |  | Error T-bulb QC. | cf swo | cf swo
No | Visit | Complaint | RFRA | Erreur T° ampoule QC. | Erreur de température dans le module autocheck | Changement de la nappe de connection entre le module autocheck et la carte électronique autocheck.parametrage des cq en manuel si bessoin.test en menu service ok | SWO-472173 | Defect | HW | Autocheck module | Cable |  | Error T-bulb QC. | Temperature error in the autocheck module | Change of the connection table between the autocheck module and the electronic card autocheck.parameting the cq in manual if bessoin.test in menu service ok
No | Visit | Complaint | RDAN | Dorte klager over utæt pH elektrode samt mange fejl 500. | Undersøgt, pH elektrodebrønd halvt fuld af væske | pH elektrode udskiftet, slutafprøvet | SWO-472054 | Assisted Replacement | Electrode | Electrode | pH |  | Dorte complains of leaking pH electrode as well as many faults 500. | Studied, pH electrode well half full of liquid | pH electrode replaced, final tested
No | Visit | Complaint | RFRA | Du liquide s'echappe du tirroir du pack. Par téléphone : Nettoyage et remplacement du pack. Attente pour voir si la fuite persiste. Nouvel appel deux jours plus tard, la fuite persiste. Intervention FSE nécessaire. Contact : 05.49.83.83.03 | Inspection des tuyaux de retour poubelle | Remise en place du tuyau de retour poubelle. Il avait sauté juste après le Y transparent. Remise en route NOK erreur de détection de cupule a l'éjection de la cartridge. En attente de pièce pour réparer. | SWO-472892 | Revisit required | HW | Incubator Module | PCB, incubator |  | Liquid escapes from the dash of the pack. By phone: Cleaning and replacing the pack. Wait to see if the leak persists. New call two days later, the leak persists. FSE intervention required. Contact: 05.49.83.83.03 | Inspection of garbage return pipes | Putting the trash return hose back in place. He jumped right after the clear Y. Restarted NOK cup detection error has the ejection of the cartridge. Waiting for room to repair.
No | Visit | Complaint | RAME | Calibration failed on tHb | Calibration failed on thb | Replaced hemolyzer assembly | SWO-473314 | Defect | HW | Oxi module | Hemolyzer |  | Calibration failed on tHb | Calibration failed on thb | Replaced hemolyzer assembly
No | Visit | Complaint | RLTD | Contact: Laura Cogwar Tel: 0203 447 0326 Please call Joseph on 07887 828 380 | Analyser screen abnormal and analyser unable to run. Issue found to be with corrupt database. Installed clean blank database, restored customer setup and restarted analyser. Installed new Sensor Cassette as requested by customer and left analyser restarting - customer to report any further issues. | Installed clean database. | SWO-472904 | Defect | SW | RMED SW | Database | corrupt | Contact: Laura Cogwar Tel: 0203 447 0326 Please call Joseph on 07887 828 380 | Analyser screen abnormal and analyser unable to run. Issue found to be with corrupt database. Installed clean blank database, restored customer setup and restarted analyser. Installed new Sensor Cassette as requested by customer and left analyser restarting - customer to report any further issues. | Installed clean database.
No | Visit | Complaint | RLTD | Contact: Laura Cogwar Tel: 0203 447 0326 Please call Joseph on 07887 828 380 | Checked hardware, wet section is working and also show as communicating to DMS in activity log. Backed up customer files. Reloaded WIN7 and analyser software. Analyser now working correctly. Restored customer data, analyser not working again. Reloaded blank database, analyser working again. Restored customer setup. Restarted analyser and verified operation, all ok. | Replaced corrupted database with clean database. | SWO-472903 | Defect | SW | RMED SW | Database | corrupt | Contact: Laura Cogwar Tel: 0203 447 0326 Please call Joseph on 07887 828 380 | Checked hardware, wet section is working and also show as communicating to DMS in activity log. Backed up customer files. Reloaded WIN7 and analyser software. Analyser now working correctly. Restored customer data, analyser not working again. Reloaded blank database, analyser working again. Restored customer setup. Restarted analyser and verified operation, all ok. | Replaced corrupted database with clean database.
No | Visit | Complaint | RLTD | Contact: Nicola O'Callahan Tel: 0203 448 4706 | Fault traced to inlet gasket with holder, not creating good seal to inlet probe and causing leak error. Customer did not have spare inlet gasket available, tried to clean inlet gasket, but unable to improve seal. Therefore carried out due service visit and fitted service kit which includes the gasket. | Replaced inlet gasket with holder. | SWO-472682 | Defect | HW | Auto Inlet module | Inlet gasket |  | Contact: Nicola O'Callahan Tel: 0203 448 4706 | Fault traced to inlet gasket with holder, not creating good seal to inlet probe and causing leak error. Customer did not have spare inlet gasket available, tried to clean inlet gasket, but unable to improve seal. Therefore carried out due service visit and fitted service kit which includes the gasket. | Replaced inlet gasket with holder.
No | Visit | Complaint | RESP | 0508 fallo de transporte de liquidos | 0508 fallo de transporte de liquidos. Se detecto aguja de entrada doblada. | Se enderezo la aguja de entrada y se sustituyo la inlet gasket. Finalmente se comprobo correcto funcionamiento mediante QC. | SWO-472707 | Defect | HW | Auto Inlet module | Probe bend | 508 | 0508 liquid transport failure | 0508 liquid transport failure. Bent input needle detected. | The inlet needle is righted and the inlet gasket is replaced. Finally, correct operation is made by QC.
No | Visit | Complaint | RAME | Going into startups regularly | Going into startups | Replaced power supply | SWO-472749 | Defect | HW | Power Supply | Power supply |  | Going into startups regularly | Going into startups | Replaced power supply
Yes | Visit | Complaint | RAME | Broken display housing. | IVW tech assit | Assist in installing 36 instruments. | SWO-473230 | Re classified RMED | installation report |  |  | Move to OOB trending | Broken display housing. | IVW tech assit | Assist in installing 36 instruments.
Yes | Visit | Complaint | RCHN | 测量仓打不开 | 测量仓打不开，借580-083,903-165（被903-644替代）各一个 | 两个备件都没有使用，需要加借903-581一个，此单先关单 by Andrew | SWO-472462 | Re classified RMED | same day upSWO |  |  |  | The measuring compartment cannot be opened | The measuring compartment cannot be opened, borrow 580-083,903-165 (replaced by 903-644) each | Both spare parts are not in use and need to be loaned 903-581 one, this single first close by Andrew
No | Visit | Complaint | RCHN | 测量仓打不开 | 机器未修复，需要加借备件 903-581一个 | 重新插拔电源线后修复 by Andrew | SWO-472520 | Defect | HW | Basic Unit | PCB, Sensor interface | Sensor Door | The measuring compartment cannot be opened | The machine has not been repaired and requires a spare part 903-581 | Fix by Andrew after replugging the power cord
No | Visit | Complaint | RCHN | 报错2168，定标-42 | 定标-42，借903-030一个 | 清洁传感器后修复 by Andrew | SWO-472552 | Cleaned | HW | Storage Module | PCB, Cartridge lock | 2168 | 2168, bid-42 | Bid-42, borrow 903-030 one | Fix by Andrew after cleaning sensor
No | Visit | Complaint | RIND | Electrolytes result variation | Checked the data logs Electrolytes result variation | Performed zeroing & level 3 QC done all the values within the range sample comparison done result was correlating | SWO-472601 | [5 empty] | Electrolytes result variation | Checked the data logs Electrolytes result variation | Performed zeroing & level 3 QC done all the values within the range sample comparison done result was correlating
No | Visit | Complaint | RIND | Calibration Error. | pO2 Calibration error. Need to replace the pO2 Membrane. | Replaced the pO2 membrane and run the instrument its working good. | SWO-472631 | Assisted Replacement | Electrode | Membrane | pO2 |  | Calibration Error. | pO2 Calibration error. Need to replace the pO2 Membrane. | Replaced the pO2 membrane and run the instrument its working good.
No | Visit | Complaint | RFRA | Du liquide s'echappe du tirroir du pack. Par téléphone : Nettoyage et remplacement du pack. Attente pour voir si la fuite persiste. Nouvel appel deux jours plus tard, la fuite persiste. Intervention FSE nécessaire. Contact : 05.49.83.83.03 | Erreur 1999 et 3553 | Remplacement de l'incubateur PCB et du capteur de détection de cupule. Remise en route et passage CQ. | SWO-472893 | Defect | HW | Incubator Module | PCB, incubator |  | Liquid escapes from the dash of the pack. By phone: Cleaning and replacing the pack. Wait to see if the leak persists. New call two days later, the leak persists. FSE intervention required. Contact: 05.49.83.83.03 | Error 1999 and 3553 | Replacement of the PCB incubator and cup detection sensor. Restarted and passage CQ.
No | Visit | Complaint | RDEU | Kalibrationsabbruch | Kalibrationsabbruch | Haltenase von Kalkassette NT-proBNP Lot 15260 abgebrochen. Defekte Kalkassette entfernt. Storagewheel ausgebaut, keine Kunststoffreste zu finden. Keine weiteren Kalkassetten vor Ort. QC TNI durchgeführt. Fehlende D-Dimer-Kalibration nachgeholt. | SWO-472700 | Defect | Consumables | Test Kit | Cartridge, Broken | Nose | Kalibrationsabbruch | Kalibrationsabbruch | Holdase of calcender NT-proBNP Lot 15260 aborted. Defective calcender removed. Storagewheel, no plastic residues to be found. No more calcins on site. QC TNI. Missing D-Dimer calibration made up for.
No | Visit | Complaint | RDEU | 28.12.2018 Frau Kresse 067219074355 meldet einen Fehler 5784 | 5784 Setsafe-22 | Storage Wheel hängt. Storage Wheel ausgebaut gereinigt. LP ausgebaut und getauscht. 5 Blankmessungen und eine QC durchgeführt. Systemüberprüfung in Ordnung | SWO-473061 | Cleaned | HW | Storage Module | storage wheel, hub | 5784; Techline? Jan | 28.12.2018 Ms Kresse 067219074355 reports an error 5784 | 5784 Setsafe-22 | Storage Wheel hangs. Storage Wheel removed cleaned. LP removed and exchanged. 5 blank measurements and one QC performed. System verification in order
No | Visit | Complaint | RDEU | 27.12.2018 Herr Görbing 0363267511303 meldet, dass sich das Storage Wheel nicht dreht. | Gerät steht im Fehler Neubooten des Analysators ohne Besserung | Überprüfung Dateiarchiv => letzte Patientenmessungen gestern Mittag abgebrochen => Fehler Pipette_424 Gerät geöffnet und Überprüfung Nadel => loses Probenröhrchen liegt im Gerät => Nadel stark verbogen => Probenröhrchen mit Stichmerkmalen auf der Unterseite => Probenröhrchen wurde verkehrt herum eingelegt Austausch Nadel => Nadelkalibrierung Zähler zurückgesetzt Systemüberprüfung und Systemreinigung i.O. Durchführung MCC Level 1 TnI | SWO-472614 | User Error | HW | Aspiration Module | Needle Assembly | Pipette_424 | 27.12.2018 Mr Görbing 0363267511303 reports that the storage wheel is not spinning. | Device is in error New boats of the analyzer without improvement | Check File Archive => last patient measurements aborted yesterday noon => Error Pipette_424 Device opened and check needle => loose sample tube lies in the device => needle strongly bent => Sample tubes with stitch characteristics on the underside => Sample tube was inserted upside down Needle Replacement => Needle Calibration Counter reset System inspection and system cleaning as the Execution MCC Level 1 TnI
No | Visit | Complaint | RDEU | Sensorbay schließt und öffnet sich kurz bevor die Sensorbay geschlossen ist | System überprüft => Fahne für die Lichtschranke an der Senorbay falsch dejustiert ,neu installierte SC defekt Schwester Fanziska hat den Service beauftragt e | Fahne richtig justiert und SC erneuert , Kalibration und Qc durchgeführt | SWO-473070 | Adjustment | HW | Basic Unit | Sensor door | Sensor Door | Sensorbay closes and opens shortly before the Sensorbay is closed | System checked => Flag for the light barrier on the Senorbay incorrectly deadjusted ,newly installed SC defective Sister Fanziska commissioned the service e | Flag correctly adjusted and SC renewed , calibration and Qc performed
No | Visit | Complaint | RCHN | 屏幕故障 | 屏幕故障，借903-239(被903-672替代）一个 | 更换903-672备件后修复 by Andrew | SWO-472534 | [5 empty] | Screen failure | Screen failure, borrow 903-239 (replaced by 903-672) one | Repair by Andrew after replacement of 903-672 spare parts
No | Visit | Complaint | RCHN | 打印机故障 | 打印机故障，借910-275一个 | 借件时报错了货号，借件好还。从库存机上拆903-001更换后正常。 by Andrew | SWO-472622 | [5 empty] | Printer failure | Printer failure, borrow 910-275 one | The loan times wrong number, the loan is good to return. Remove 903-001 from the stock machine after replacement normal. by Andrew
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | LAC定标结果偏低 | 更换测试卡后修复 | SWO-473049 | [5 empty] | Test card, reagent package failure | LAC calibration results are low | Repair after replacing test card
No | Visit | Complaint | RCHN | 试管夹断裂 | 试管夹断裂 借903-584一个 | 更换903-584一个正常。 by Andrew | SWO-472493 | Defect | HW | Aspiration Module | Tube holder |  | Test tube clip break | Test tube clip broken, borrow 903-584 one | Replace the 903-584 one normal. by Andrew
No | Visit | Complaint | RCHN | 定标扫码器故障 | 定标卡扫描器不能扫描，借902-924一个 | 更换902-924后修复 by Andrew | SWO-471956 | Defect | HW | Cover and Chassis | Barcode reader, display |  | Calibration scanner failure | The calibration card scanner cannot be scanned, borrowing 902-924 one | Repair by Andrew after replacing 902-924
No | Visit | Complaint | RCHN | 系统检查时数据丢失 | 系统检查时数据丢失,借944-214（15057），944-268（15219），944-216（15082）各一张。 | 重新定标正常。 by Andrew | SWO-472519 | Assisted Replacement | Consumables | Test Kit | Cartridge, cal |  | Data loss during system checks | Data loss during system checking, borrow 944-214 (15057), 944-268 (15219), 944-216 (15082) each. | Recalibration is normal. by Andrew
No | Visit | Complaint | RCHN | 光学强度超范围，报错1031 | 电压规格超范围 1225，借903-364,944-212（15210） $2个，944-214（14934） $2个 | 更换903-364一个，重新定标正常。by jesson | SWO-473036 | Defect | HW | Process Module | Optical Unit | 1031 | Optical strength over range, error 1031 | Voltage specification over range 1225, borrowed 903-364, 944-212 (15210) /2, 944-214 (14934) /2 | Replace the 903-364 one and re-label normally. by jesson
No | Visit | Complaint | RCHN | 测试卡验证败 | 测试卡验证败，借902-972一个 | 目测接口板腐蚀，更换902-972一个正常。 | SWO-473039 | [5 empty] | Test card validation failure | Test card validation failure, borrow 902-972 one | Visual interface plate corrosion, replacement 902-972 a normal.
Yes | Visit | Complaint | RAME | Customer called with the touch screen frozen and not responding to touch | Upon arrival, tested touchscreen and received no response. Reset the analyzer using a keyboard multiple times, but that did not correct the issue. | Replaced the CPU. Touchscreen operation restored. Normal operations restored. | SWO-473484 | [5 empty] | Customer called with the touch screen frozen and not responding to touch | Upon arrival, tested touchscreen and received no response. Reset the analyzer using a keyboard multiple times, but that did not correct the issue. | Replaced the CPU. Touchscreen operation restored. Normal operations restored.
No | Visit | Complaint | RCHN | 报错2168，定标-42 | 报错2168，定标-42，借903-030一个 | 更换903-030一个正常。by jesson | SWO-473050 | Defect | HW | Storage Module | PCB, Cartridge lock | 2168 | 2168, bid-42 | 2168, bid-42, borrow 903-030 one | Replace 903-030 one normal. by jesson
No | Visit | Complaint | RCHN | 干燥前温度超范围 1068 | 干燥前温度超范围 1068,借903-210一个。 | 科室温度太低的原因。将温度调整正常即可。备件好还。by jesson | SWO-473164 | Defect | HW | Process Module | PCB, cup dryer | 1068 | Over-range temperature before drying 1068 | The pre-drying temperature exceeds the range of 1068, borrowing 903-210 one. | The reason why the room temperature is too low. Just adjust the temperature to normal. Spare parts are good also. by jesson
No | Visit | Complaint | RAME | Cal's will not pass | The 2pt cal would not pass with the electrolytes. Sen failed | The analyzer was in stand by all weekend and when it came out of standby the sen came in on he cal | SWO-473546 | No fault found | No fault found | No fault found | No fault found |  | Cal's will not pass | The 2pt cal would not pass with the electrolytes. Sen failed | The analyzer was in stand by all weekend and when it came out of standby the sen came in on he cal
No | Visit | Complaint | RAME | Leak detected error - 0663 | Found no O-ring on PO2 electrode and loose tubings on met measuring module valve block. | Remembraned PO2 electrode with new sealing ring and properly seated tubings. | SWO-473534 | Defect | Electrode | Electrode | Sealing ring, pO2 |  | Leak detected error - 0663 | Found no O-ring on PO2 electrode and loose tubings on met measuring module valve block. | Remembraned PO2 electrode with new sealing ring and properly seated tubings.
No | Visit | Complaint | RCHN | 干燥前温度超范围 1068 | 干燥前温度超范围 1068 | 干燥单元定标，确认环境温度。by jesson | SWO-473575 | Adjustment | HW | Process Module | Cup dryer | 1068 | Over-range temperature before drying 1068 | Over-range temperature before drying 1068 | The drying unit is calibrated to confirm the ambient temperature. by jesson
No | Visit | Complaint | RDEU | 28.12.2018; Fr. Usbeck; Tel.: 071164893200; MetHb QC Messfehler B und C | tHb-Kal durchgeführt. Weiterhin QC-Fehler. OxiCompleteTest durchgeführt. Keine Fehler. | Hämolysator erneuert. Kalibrationen durchgeführt. QCs nahe Zielwert. | SWO-474021 | Defect | HW | Oxi module | Hemolyzer | QC, MET-Hb | 28.12.2018; Ms. Usbeck; Tel.: 071164893200; MetHb QC Measurement error B and C | tHb-Kal. Further QC errors. OxiCompleteTest. No errors. | Hemolyser renewed. calibrations. QCs close to target.
Yes | Visit | Complaint | RSAF | Inhomogenous cal3 solution and pump cal error | Telephonic trouble shooting done but replacement of soultion pack is required and customer has no spare solution pack avaliable. | Collect and deliver to customer site 1x new solution pack. customer to install. | SWO-474011 | Re classified RMED | Logistical upSWO |  |  |  | Inhomogenous cal3 solution and pump cal error | Telephonic trouble shooting done but replacement of soultion pack is required and customer has no spare solution pack avaliable. | Collect and deliver to customer site 1x new solution pack. customer to install.
No | Visit | Complaint | RDEU | Clot detected , keine manuelle Spülung möglich 02.01.2019; Sr. Justine; Tel.: 02191138580; Verdacht auf Gerinnsel im Oxy Modul | Gerät überprüft | Gerät komplett zerlegt und die einzelnen Bauteile manuell mit Wasser und Proteinentfernung gespült. Der Fehler besteht weiterhin. Sensorbay verstopft und muss erneuert werden | SWO-474368 | Revisit required | HW | Basic Unit | Sensor Bay Unit | Clot | Clot detected , no manual rinsing possible 02.01.2019; Sr. Justine; Tel.: 02191138580; Suspicion of clots in the Oxy module | Device checked | completely disassembled and manually rinsed the individual components with water and protein removal. The error persists. Sensorbay clogged and needs to be renewed
No | Visit | Complaint | RDEU | Gerät neu, SC neu, SP neu, pCO2-Empfindlichkeit zu gering | Alle QC pCO2 nach oben außer Bereich, mit neuer SC pCO2-Emfpkt. zu gering, aber QC in allen Leveln i.O. | SC + SP erneuert (856-41 & BY56), Empfkt. steigende Tendenz | SWO-474134 | Assisted Replacement | SC | SC | pCO2 | Sens low | Device new, SC new, SP new, pCO2 sensitivity too low | All QC pCO2 up out of range, with new SC pCO2 footprint. too low, but QC at all levels. | SC + SP renewed (856-41 & BY56), Mpfkt. rising trend
No | Visit | Complaint | RDEU | 02.01.2018; Fr. Koch; Tel.: 08216004773; Einlassdichtung muss getauscht werden | Fehlersuche vor Ort. | JW Kit verbaut wegen div Fehler Startphase abgewartet. Sensorkassette erneuert aus Kundenbestand. THb kal durchgeführt. Kal und QC bei Startphase mit durchgeführt. Danach war Analysator wieder messbereit. | SWO-474074 | Defect | HW | Basic Unit | PM Kit |  | 02.01.2018; Ms. Koch; Tel.: 08216004773; Inlet seal must be replaced | Troubleshooting on site. | JW Kit installed due to div error start-up phase. Sensor cassette renewed from customer base. THb kal. Kal and QC at the start phase. After that, the analyzer was ready for measurement again.
No | Visit | Complaint | RLTD | Claire Latham Tel: 01553 613780 or 07890 715176 - please call before 3.30pm as there will be no one there after this time today or tomorrow | Random software issues. Analyser acting strangely, when user runs a sample instead of the patient input screen, the sensor cassette replacement screen appears and tells the user someone else is logged in. Customer reported that this is doing the same on all other analysers onsite. | The analyser has been working normally over the weekend with the data cable disconnected. Ran the virus scan tool, no issues found. Reconnected data cable and ran a test sample, sample ran without issue. Reconnected the rest of the analysers and ran a test sample, all ok. | SWO-473899 | No fault found | SW | SW | SW |  | Claire Latham Tel: 01553 613780 or 07890 715176 - please call before 3.30pm as there will be no one there after this time today or tomorrow | Random software issues. Analyser acting strangely, when user runs a sample instead of the patient input screen, the sensor cassette replacement screen appears and tells the user someone else is logged in. Customer reported that this is doing the same on all other analysers onsite. | The analyser has been working normally over the weekend with the data cable disconnected. Ran the virus scan tool, no issues found. Reconnected data cable and ran a test sample, sample ran without issue. Reconnected the rest of the analysers and ran a test sample, all ok.
No | Visit | Complaint | RAME | Going into startups regularly | Replaced battery | Replaced battery and sensor. Thursday came to verify all was within specs | SWO-474154 | Defect | HW | Battery module | Battery module |  | Going into startups regularly | Replaced battery | Replaced battery and sensor. Thursday came to verify all was within specs
No | Visit | Complaint | RAME | 1063 Pressure Vacuum error | Failed vacuum pressure test | Replaced inlet probe and tested. | SWO-474241 | Defect | HW | Auto Inlet module | Probe | 1063 | 1063 Pressure Vacuum error | Failed vacuum pressure test | Replaced inlet probe and tested.
No | Visit | Complaint | RAME | Housing screws are no longer secure monitor. | Housing screws are no longer secure monitor. | Replaced monitor and created CQC. | SWO-474143 | Defect | HW | Display | Display | Screws | Housing screws are no longer secure monitor. | Housing screws are no longer secure monitor. | Replaced monitor and created CQC.
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | LAC定标值偏低 | 更换测试卡后修复 | SWO-478132 | [5 empty] | Test card, reagent package failure | LOW LAC CALIBRATION VALUE | Repair after replacing test card
No | Visit | Complaint | RCHN | 测试卡断脚 | 测试卡断脚，损坏部位：teeth | 更换测试卡后修复 | SWO-473690 | Defect | Consumables | Test Kit | Cartridge, Broken | Teeth, 5784 | Test card broken foot | Test card broken foot, damaged site: teeth | Repair after replacing test card
No | Visit | Complaint | RDEU | 02.01.2019 Frau Danner 068411622792 meldet, dass die Schläuche unter der Einlassdichtung ab gehen. | Waste Pumpenschläuche springen ab. | die Membrane vom Luftventil BG $ Module ausgetauscht. Schlauchverbindung der Vorwärmer und Messkammer undicht . Kalibration und Qualitätskontrolle. | SWO-474363 | Defect | HW | pH/BG Module | Valve, membrane |  | 02.01.2019 Ms. Danner 068411622792 reports that the hoses are going off under the inlet seal. | Waste pump hoses bounce off. | the diaphragm is replaced by the air valve BG' modules. Hose connection of the preheater and measuring chamber leaks . Calibration and quality control.
No | Visit | Complaint | RDEU | 02:01:2019; 08.36 Uhr; Fr. Franz, Tel.: 020172382700: STORAGEDOOROBSRTUCTED | Fehler trat nur einmal auf | Storagedoor im Service überprüft, Funktion OK. | SWO-473902 | No fault found | HW | Cover and Chassis | Door, Storage | 1009, 6237 | 02:01:2019; 8:36 a.m.; Ms. Franz, Tel.: 020172382700: STORAGEDOOROBSRTUCTED | Error occurred only once | Storagedoor checked in service, function OK.
No | Visit | Complaint | RDEU | 02.01.2019, 07:39 Uhr, Sr. Doreen, Tel. 061514027002: Testkassette verklemmt | Analysator ausgeschaltet Neustart Testkassetten Öffnung (Tür geöffnet) Aktivitäten Meldung 4726 Open Storage 24 | Demontage Frontabdeckung - Türmechanik gereinigt und geschmiert etc. (Tür verklemmt) - Allgemeine Überprüfung und Reinigung - Funktionsüberprüfung etc. ohne Beanstandung | SWO-473825 | Adjustment | HW | Cover and Chassis | Door, Storage | 4726 | 02.01.2019, 07:39, Sr. Doreen, tel. 061514027002: Test cassette jammed | Analyzer switched off Restart Test cassettes opening (door open) Activities Message 4726 Open Storage 24 | Disassembly front cover - door mechanism cleaned and lubricated etc. (door jammed) - General inspection and cleaning - Functional check etc. without complaint
No | Visit | Complaint | RDEU | 27.12.2018; Fr. Yaila; Tel.: 07112184060; Kein Test für Messung verfügbar. 28.12.2018 08:15Uhr Frau Yaila 071121840681: Das Gerät bootet und möchte dann wieder einen Neustart | Test NT-proBNP durch defekte Kalkassette unbrauchbar. | Durch Gerätefehler unbrauchbare Tests eingezogen. Bestände an KorLo weitergegeben. | SWO-474020 | User Error | User Error | User Error | User Error | Calibrations deleted | 27.12.2018; Ms. Yaila; Tel.: 07112184060; No test available for measurement. 28.12.2018 08:15Pm Ms Yaila 071121840681: The device boots and wants to restart again | Test NT-proBNP unusable due to defective calcanda. | Unusable tests are retracted due to device errors. stocks passed on to KorLo.
No | Visit | Complaint | RDEU | 28.12.2018 Frau Kresse 067219074355 meldet einen Fehler 5784 28.12.2018 18:40Uhr Herr Kron 067219074199: das Gerät hat die Messung abgebrochen und verlangt nach jedem Neustart wieder einen Neustart | Storage wheel Bewegung ist gestört. | Testkassetteneinlass-Führungsschienen sind abgebrochen und verhindern die Bewegung der Storage Wheel. Führungsschienen ausgebaut. Gerät von Innen gereinigt und neu gestartet. Ersatzteile werden bestellt. Die Rep. wird in einem späteren Zeitpunkt fortgesetzt. Messung ist möglich. QC durchgeführt. | SWO-473931 | Revisit required | HW | Storage Module | Cartridge guide | 5784 | 28.12.2018 Ms Kresse 067219074355 reports an error 5784 28.12.2018 18:40Mr Kron 067219074199: the device has aborted the measurement and requires a restart after each restart | Storage wheel movement is disturbed. | Test cartridge inlet guide rails are broken off and prevent the storage wheel from moving. guide rails. Device cleaned from the inside and restarted. Spare parts are ordered. The Rep. will continue at a later date. Measurement is possible. QC.
No | Visit | Complaint | RDEU | Probenaspiration gescheitert, Fehler #1271 | Probenaspiration gescheitert, Fehler #1271 Pinch Valve schwergängig | Gerät manuell gespült Pinch Valve erneuert, Jahreswartungskit verbaut Sensorkassette aus Kundenbestand erneuert, danach Na-Sensor-Impedanzfehler SCi aus Wagebestand eingebaut Gerät kalibriert und überprüft | SWO-474397 | Defect | HW | Basic Unit | PM Kit | 1271 | Sample aspiration failed, errors #1271 | Sample aspiration failed, errors #1271 Pinch Valve difficult to handle | Device manually rinsed Pinch Valve renewed, annual maintenance kit installed Sensor cartridge renewed from customer base, then Na-sensor impedance error SCi built-in from wage stock Device calibrated and checked
No | Visit | Complaint | RDEU | 02.01.2019 15:00Uhr Sr.Monika 02417501154: Trotz umfangreiche Wartungsarbeiten meldet das Gerät ständig Fehler 0.3.01.2019 08:00Uhr Herr Susen: Der Kunde hatte die falsche Gerätenummer angegeben, weswegen der Case auf das falsche Gerät geschrieben wurde 090R0943N017 - Der Bericht wurde korrigiert | S. Case - Intervention erforderlich | Fehlersuche: Mit vorhandenem SP und SC Fluidic Serviceprogramme absolviert => Lösungen werden unzuverlässig gepumpt. Einlassdichtung $-Nadel überprüft, Einlassdichtung vollständig eingerastet und auf der Nadel bewegt, Lösungen werden nun einwandfrei gefördert. Neustart, Startphase 60min => pCO2 $Lac $Glu Werte weit außerhalb des Kalibrierbereiches => SC ersetzt, erneut komplette Startphase abgewartet, Kalibrierungen sowie tHb-Kal (fällig 5.1.) und QC Messungen aller Level, ok | SWO-476548 | Adjustment | HW | Inlet module | Inlet gasket |  | 02.01.2019 15:00h Sr.Monika 02417501154: Despite extensive maintenance work, the device constantly reports errors 0.3.01.2019 08:00Am Mr Susen: The customer had given the wrong device number, so the case was written to the wrong device 090R0943N017 - The report has been corrected | S. Case - Intervention required | Troubleshooting: Completed with existing SP and SC Fluidic service programs => Solutions are pumped unreliably. Inlet seal s-needle checked, inlet seal fully snapped and moved on the needle, solutions are now properly conveyed. Restart, start phase 60min => pCO2 $Lac $Glu values far outside the calibration range => SC, again complete start phase awaited, calibrations as well as tHb-Kal (due 5.1.) and QC measurements of all levels, ok
No | Visit | Complaint | RDEU | 14.11.2018; Hr. Krasniqui; Tel.: 021513347257; pCO2 QC Fehler C 2.1.19/15:10 Herr Brack 021513347299 -> pCO2-Fehler bei QK 02.01.2019, 17.30Uhr, Herr Brack, Tel: 021513347299: weiterhin AQC Fehler bei pCO2 | pCO2 in der QC lvl A und C zu hoch. | Sensorkassette erneuert, zwei mal Kalibriert und QC’s durchgeführt. Lac muss noch einlaufen da die SC frisch aus dem Kühlschrank kam. | SWO-477487 | Assisted Replacement | SC | SC | pCO2 | QC A and C too high | 14.11.2018; Hr. Krasniqui; Tel.: 021513347257; pCO2 QC Fehler C 2.1.19/15:10 Herr Brack 021513347299 -> pCO2-Fehler bei QK 02.01.2019, 17.30Uhr, Herr Brack, Tel: 021513347299: weiterhin AQC Fehler bei pCO2 | pCO2 in der QC lvl A und C zu hoch. | Sensor cassette renewed, calibrated twice and QC's performed. Lac has yet to come in because the SC came fresh from the fridge.
No | Visit | Complaint | RAME | pCO2 QC high, Page my FSE | Pco2 qc high on all qc levels | Replaced sensor cassette, Pm instrument and replaced both sensor cassette and solution pack. | SWO-476466 | Assisted Replacement | SC | SC | pCO2 | QC all levels too high | pCO2 QC high, Page my FSE | Pco2 qc high on all qc levels | Replaced sensor cassette, Pm instrument and replaced both sensor cassette and solution pack.
No | Visit | Complaint | RDEU | Oxiemetrie Hardfehfehler. | Wet Section Control PCB ersetzt, AU Sw installiert. | Startphase. | SWO-476206 | Defect | HW | Basic Unit | PCB, wet section | Oxi | Oxiemetrie Hardfehfehler. | Wet Section Control PCB replaced, AU Sw installed. | Start.
No | Visit | Complaint | RDEU | 2.1.19/13:20 Sr. Manuela 02037332211 -> SP90 immer wieder gewechselt werden 02.01.2019, 16.45Uhr, Frau Linde, Tel: 02037330: Lösungen können nicht gepumpt werden 02.01.2019, 17.30Uhr, Tel: 02037332210: erneut Wechselschleife- LP tauschen | Inhomogene Lösung Kal 3 | Einen neuen LP vorab organisiert, Gerät durchgespült, defektes LP Lot (BY-55) mit LOT HA-02 erneuert, Gerät kalibriert und QC durchgeführt. | SWO-476109 | Assisted Replacement | SP | SP | FT | Inhomogeneous Cal 3 | 2.1.19/13:20 Sr. Manuela 02037332211 -> SP90 to be changed again and again 02.01.2019, 16.45pm, Mrs Linde, Tel: 02037330: Solutions cannot be pumped 02.01.2019, 17.30h, Tel: 02037332210: exchange again interchangeable loop LP | Inhomogeneous solution Kal 3 | A new LP pre-organized, device flushed through, defective LP lot (BY-55) renewed with LOT HA-02, device calibrated and QC performed.
No | Visit | Complaint | RLTD | Please call Lynn on 01582 497 991 | Analyser not allowing solution pack change because inlet is not functioning correctly, analyser is indicating that the inlet is being blocked. | Removed rear cover and observed inlet motor and drive shaft, checked the grub screws securing the shaft to the motor. Issue remains with inlet removed. Stripped down inlet and removed the curve wheel, the curve wheel was being held on with residual spring tension from the inlet mount. Removed wheel and reseated it, reassembled inlet and retested, solution pack changed without issue, changed two more times and tested in service mode, all operating as normal. | SWO-476188 | Adjustment | HW | Auto Inlet module | Curve Wheel |  | Please call Lynn on 01582 497 991 | Analyser not allowing solution pack change because inlet is not functioning correctly, analyser is indicating that the inlet is being blocked. | Removed rear cover and observed inlet motor and drive shaft, checked the grub screws securing the shaft to the motor. Issue remains with inlet removed. Stripped down inlet and removed the curve wheel, the curve wheel was being held on with residual spring tension from the inlet mount. Removed wheel and reseated it, reassembled inlet and retested, solution pack changed without issue, changed two more times and tested in service mode, all operating as normal.
No | Visit | Complaint | RAME | 216 General Printing error | 1) 216 General Printing error | 1) Replaced printer unit and printing cable. Tested. 2) Provided customer Dec 2018 SWO's and for 35 installs: SWO's, IVW's and VK Linearity results. | SWO-476307 | Defect | HW | Printer Module | Printer Module | 216 | 216 General Printing error | 1) 216 General Printing error | 1) Replaced printer unit and printing cable. Tested. 2) Provided customer Dec 2018 SWO's and for 35 installs: SWO's, IVW's and VK Linearity results.
No | Visit | Complaint | RIND | Parameter Failure | Found that sensor cassette is cloted. Sample is collected without hyperin. | Need to change new sensor cassette. Note given them to Use Hyperin for wash before collecting the ABG Sample | SWO-474405 | [5 empty] | Parameter Failure | Found that sensor cassette is cloted. Sample is collected without hyperin. | Need to change new sensor cassette. Note given them to Use Hyperin for wash before collecting the ABG Sample
No | Visit | Complaint | RAME | The tHb and sO2 do not compare well with the other analyzer. Consistently measures 0.5 to 1.3 g/dL lower. | The tHb measures low compared to other analyzers. QCs are close but do not compare exactly. | Replaced the hemolyzer unit. Performed tHb calibrations and ran QC. The analyzer now compares better to the other analyzers on QC but the blood samples still do not compare well. Replaced the spectrometer and had the same result, QC compares, blood does not. Reinstalled the spectrometer and replaced the CO-OX Controller Board and had the same results, QC okay, blood does not compare. Will report findings to company and request further instructions. | SWO-474212 | [5 empty] | The tHb and sO2 do not compare well with the other analyzer. Consistently measures 0.5 to 1.3 g/dL lower. | The tHb measures low compared to other analyzers. QCs are close but do not compare exactly. | Replaced the hemolyzer unit. Performed tHb calibrations and ran QC. The analyzer now compares better to the other analyzers on QC but the blood samples still do not compare well. Replaced the spectrometer and had the same result, QC compares, blood does not. Reinstalled the spectrometer and replaced the CO-OX Controller Board and had the same results, QC okay, blood does not compare. Will report findings to company and request further instructions.
No | Visit | Complaint | RCHN | 无法读取测试卡智能芯片，报错1123 | 无法读取测试卡智能芯片，下层温度超出范围，报错1123和1106，需要借件903-581一个 | 更换903-581后修复 | SWO-473666 | Defect | HW | Basic Unit | PCB, Sensor interface | Chip, temp | Failed to read test card smart chip, error 1123 | Unable to read the test card smart chip, the lower temperature out of range, reported errors 1123 and 1106, need to borrow 903-581 a | Repair after replacement 903-581
No | Visit | Complaint | RCHN | 抱怨有结果出不来 | 抱怨有结果出不来，有结果传不到系统 | 核实是超出可报告范围，调整可报告范围就好了，是对方系统的问题 | SWO-475952 | Setup | SW | RMED SW | Reportable range |  | Complaining that the results don't come out | Complaining that there are results that don't come out, that there are results that don't pass on to the system | Verification is beyond the reportable range, adjusting the reportable range is good, is the problem of the other system
No | Visit | Complaint | RCHN | 卡盘 | 卡盘，错误代码5784，借886-622一个 | 更换886-622 | SWO-473645 | Defect | HW | Storage Module | storage wheel, hub | 5784 | Chuck | Card, error code 5784, borrow 886-622 one | Replace 886-622
No | Visit | Complaint | RCHN | 干燥前温度超范围 | 干燥前温度超范围，借903-210一个 | 更换903-210后修复 by jesson | SWO-473819 | Defect | HW | Process Module | PCB, cup dryer | 1068 | Overrange temperature before drying | Over-range temperature before drying, borrow 903-210 one | Repair by jesson after replacement 903-210
Yes | Visit | Complaint | RFRA | Décalage du pH et iono | Regul stock après expédition électrode de référence | Regul stock après expédition électrode de référence | SWO-476197 | Re classified RMED | Logistical upSWO |  |  | mail: Visist: What has been done? | PH and iono shift | Regul stock after reference electrode shipping | Regul stock after reference electrode shipping
No | Visit | Complaint | RDEU | 02.01.2018, Frau reffelmann(POCT), Tel: 0210389921190: Drucker druckt nicht richtig | Weiße Linien genau in den Werten | Drucker ersetzt , Druckbild wieder gut, Kal und Qc s bestanden | SWO-476241 | Defect | HW | Printer Module | Printer Module |  | 02.01.2018, Mrs. reffelmann(POCT), Tel: 0210389921190: Printer does not print correctly | White lines exactly in the values | Printer replaced, print image again good, Kal and Qc s passed
No | 3 | Y | RCZE | Blank | vadna CPU unit |  | D579B86D1C29F3BC | Defect | HW | DMS_PC | CPU |  | Blank | Vadna CPU Unit
No | Visit | Complaint | RIND | Autocheck cover open close automatically. | Need to replace the auto check module PCB | Replace the auto check PCB & Auto check interface PCB from another old machine & it was working properly. | SWO-477107 | Defect | HW | Autocheck module | PCB |  | Autocheck cover open close automatically. | Need to replace the auto check module PCB | Replace the auto check PCB & Auto check interface PCB from another old machine & it was working properly.
No | Visit | Complaint | RIND | Instrument not switching on . | Instrument not switching ON. | Checked all the connections , restarted the instrument still the instrument is not responding | SWO-477109 | Defect | HW | DMS_PC | CPU | wait, check CRM for more info: Checked all the connections , restarted the instrument still the instrument is not responding ,changed the PC module and then restarted the instrument , upgraded the software configured it and calibrated . Run some samples ,customer satisfied with results . | Instrument not switching on . | Instrument not switching ON. | Checked all the connections , restarted the instrument still the instrument is not responding
No | Visit | Complaint | RDEU | 02.01.2019, Sr. Martina, tel. 0343218453: AQC Fehler bei pCO2, CL und Lac | Gerät meldet seit Dez.18 häufig pCO2 instabil bei QC Messung, E-Lyte Kal häufig mit Drift, QC Material auf Grund der häufigen Wiederholung fast alle Membranierung der Elektroden brachte keine Verbesserung | Gerät überprüft: E1001 ersetzt, Messkammern gereinigt, Reinigungszusatz für Reinigungslösung nicht aktiviert-> aktiviert, Dichtring an pH Elektrode ersetzt, Airanschluß im Fluidicmodul verstopft -> gereinigt Kunde veranlasst das vorziehen der nächsten automatischen ACQC Lieferung Kal & QC Daten gemäß RDEU Anweisung gesichert und dokumentiert | SWO-476222 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | 02.01.2019, Sr. Martina, tel. 0343218453: AQC error with pCO2, CL and Lac | Device since Dec.18 frequently reports pCO2 unstable in QC measurement, E-Lyte Kal frequently with drift, QC material due to frequent repetition almost all Membraneing of the electrodes did not bring any improvement | Device checked: E1001 replaced, measuring chambers cleaned, cleaning additive for cleaning solution not activated-> activated, sealing ring replaced to pH electrode, air connection clogged in fluidic module -> cleaned Customer initiates the preference of the next automatic ACQC delivery Kal & QC data backed up and documented in accordance with RDEU instructions
No | Visit | Complaint | RAME | CPU issues after power change | During power outage, analyzer shut down and is now showing Status Code 9 Error during boot up. | Exited Roma, deleted Roma.ini file that was corrupted, pasted a Roma.ini file from previous backup, rebooted analyzer, and validated. Replaced lactate and reference membrane which were due. Replaced fan filter. | SWO-476303 | Defect | SW | RMED SW | Roma.ini | Status code 9 | CPU issues after power change | During power outage, analyzer shut down and is now showing Status Code 9 Error during boot up. | Exited Roma, deleted Roma.ini file that was corrupted, pasted a Roma.ini file from previous backup, rebooted analyzer, and validated. Replaced lactate and reference membrane which were due. Replaced fan filter.
No | Visit | Complaint | RAME | PCO2 Calibration will not come in. | Calibration is failing for pCO2. Upon further investigation, it was caused by a defective electrode. | Replaced pCO2 electrode under warranty, cleaned chamber and contacts. Allowed analyzer to condition. | SWO-476305 | Assisted Replacement | Electrode | Electrode | pCO2 | Cal | PCO2 Calibration will not come in. | Calibration is failing for pCO2. Upon further investigation, it was caused by a defective electrode. | Replaced pCO2 electrode under warranty, cleaned chamber and contacts. Allowed analyzer to condition.
No | Visit | Complaint | RESP | Error 316 Neon number too low. | El test de posición de la lámpara de neón estaba por debajo del limite. | Cambiar el espectrometro del módulo OXI, calibrar y probar con controles. | SWO-476069 | Defect | HW | Oxi module | Spectrometer | 316; mail: Fiber optic between lamp unit and spectrometer , tried to change? | Error 316 Neon number too low. | The neon lamp position test was below the limit. | Change the oximeter spectrometer, calibrate and test with controls.
No | Visit | Complaint | RAME | Oximetry errors on blood samples. | The analyzer is presenting the following errors during blood sample measurements: 292 Turbidity too high 581 Oxi spectrum mismatch | Checked the patient results log and, confirmed the error messages. Performed a service Complete Oximetry test and, found the hemolyzer failing on the Cavitation absorb value at -0.09 (Range: 5.00-30.00). Replaced a new hemolyzer unit, the pinch valve and, the pinch valve tubing. Ran the Complete Oximetry test again and, now the Cavitation absorb value is within range at 13.03. Ran the ctHb calibration, passing within specifications. | SWO-477080 | Defect | HW | Oxi module | Hemolyzer | 292, 581 | Oximetry errors on blood samples. | The analyzer is presenting the following errors during blood sample measurements: 292 Turbidity too high 581 Oxi spectrum mismatch | Checked the patient results log and, confirmed the error messages. Performed a service Complete Oximetry test and, found the hemolyzer failing on the Cavitation absorb value at -0.09 (Range: 5.00-30.00). Replaced a new hemolyzer unit, the pinch valve and, the pinch valve tubing. Ran the Complete Oximetry test again and, now the Cavitation absorb value is within range at 13.03. Ran the ctHb calibration, passing within specifications.
No | Visit | Complaint | RDEU | Clot detected , keine manuelle Spülung möglich 02.01.2019; Sr. Justine; Tel.: 02191138580; Verdacht auf Gerinnsel im Oxy Modul | Verstopfung im Sensormodul | Sensormodul ersetzt, Startphase (60min), po2 und Lac Fehler , Sc erneuert, Ansprachefehler Lac -- Konditionierung , Po2 QC Level 2 zu niedrig, Lösungspack ersetzt, jetzt alle Qc Werte mittig, nach vielen Kalibrationen dann alles Grün die Messung von Proben mit Anwedern besprochen, | SWO-477502 | Defect | HW | Basic Unit | Sensor Bay Unit | Clot | Clot detected , no manual rinsing possible 02.01.2019; Sr. Justine; Tel.: 02191138580; Suspicion of clots in the Oxy module | Constipation in the sensor module | Sensor module replaced, start phase (60min), po2 and lac error, Sc renewed, speech error Lac -- conditioning, Po2 QC Level 2 too low, solution pack replaced, now all Qc values in the middle, after many calibrations then all green the measurement of samples discussed with objectors,
No | Visit | Complaint | RAME | I can’t do the precision on Cath Lab 1+ because it is asking for a new inlet gasket holder. Every time I try to go into service, it goes into a 45 minute start up. | 1) Error 1084: Inhomogeneous cal 3 solution. Elongated solution pack pump tubing. | 1) Replaced solution pack and tested. 2) Submitted Credit Claim form for SP. 3) Connected analyzer to AQURE. 4) Retrieved Service dump for Customer Quality Case RAME 831360780. | SWO-476885 | Assisted Replacement | SP | SP | Pump tube | 1084 | I can’t do the precision on Cath Lab 1+ because it is asking for a new inlet gasket holder. Every time I try to go into service, it goes into a 45 minute start up. | 1) Error 1084: Inhomogeneous cal 3 solution. Elongated solution pack pump tubing. | 1) Replaced solution pack and tested. 2) Submitted Credit Claim form for SP. 3) Connected analyzer to AQURE. 4) Retrieved Service dump for Customer Quality Case RAME 831360780.
No | Visit | Complaint | RCHN | 经常提示更换试剂包 | 经常提示更换试剂包 | 冲洗管路，重新安装废液管后修复 | SWO-476509 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | Frequent tips for replacing reagent kits | Frequent tips for replacing reagent kits | Rinse the line and repair it after reinstalling the waste pipe
No | Visit | Complaint | RCHN | 试剂包漏电流，报错0964 | 试剂包漏电流，报错0964 借903-581一个 | 更换903-581后修复 | SWO-473755 | Defect | HW | Basic Unit | PCB, Sensor interface | 964 | Reagent packet leakage current, reported error 0964 | Reagent packet leakage current, error 0964 borrowed 903-581 one | Repair after replacement 903-581
No | Visit | Complaint | RCHN | 光学强度超范围，报错1031 | 光学强度超范围，报错1031 借903-364一个 | 更换903-364后修复 by jesson | SWO-474294 | Defect | HW | Process Module | Optical Unit | 1031 | Optical strength over range, error 1031 | Optical strength over range, error 1031 borrowed 903-364 one | Repair by jesson after replacing 903-364
No | Visit | Complaint | RCHN | 定标失败 | 定标失败，借903-364,944-258（15081），944-212（15177）各一个 | 更换备件903-364和重新定标后修复 by jesson | SWO-473742 | Defect | HW | Process Module | Optical Unit | 1024 | Bid failed | Failed bid, borrow 903-364, 944-258 (15081), 944-212 (15177) each | Replacement of spare parts 903-364 and repair of by jesson after recalibration
No | Visit | Complaint | RCHN | 光学单元超范围 | 光学单元超范围，借903-364一个 | 更换903-364后修复 by jesson | SWO-474275 | Defect | HW | Process Module | Optical Unit | 1031 | Optical unit over range | Optical unit over range, borrowed 903-364 one | Repair by jesson after replacing 903-364
Yes | 2 | Y | RCZE | Blank |  |  | D57FED386A2F876C | Re classified RMED | refurbishing analyzer |  |  |  | Blank
No | Visit | Complaint | RPOL | Głośny wentylator. | Diagnoza problemu. | Wymiana uszkodzonego wentylatora w module AutoCheck. Testy serwisowe. | SWO-476633 | Defect | HW | Autocheck module | Fan |  | Loud fan. | Diagnosis of the problem. | Replacing the defective fan in the AutoCheck module. Service tests.
No | Visit | Complaint | RPOL | Neon number too low, Neon intensity too low | Testy oksymetru. | Wymian modułu OXI na nowy. Testy serwisowe ok. Kalibracja hemoglobiny. | SWO-478908 | Defect | HW | Oxi module | Oxi module |  | Neon number too low, Neon intensity too low | The test of the oksymmetry. | Replace the OXI module with a new. Service tests approx. Hemoglobin calibration.
No | Visit | Complaint | RCHN | 间歇性报5084 | 间歇性报5684 | 仪器清洁，检查排线 需要继续观察 by jesson | SWO-477170 | Cleaned | HW | Process Module | Process Wheel | 5684 | Intermittent 5084 | Intermittent 5684 | Instrument cleaning, checking the wiring need to continue to observe by jesson
No | Visit | Complaint | RCHN | TNI定标失败 | TNI定标失败 | 清洁处理盘，处理盘底下的白块，整机的除灰，但定标还是无法通过，故需要借件处理。此单先关闭，另开SWO by jesson | SWO-477111 | Revisit required | HW | Process Module | Optical Unit | 1024 | TNI calibration failed | TNI calibration failed | Clean the processing disk, the treatment plate under the white block, the whole machine to remove the ash, but the calibration is still unable to pass, so need to borrow processing. This single closes first, opens sWO by jesson
No | Visit | Complaint | RCHN | TNT定标失败，报错5257 | TNT定标失败，报错5257，借944-268（15316）一张 | 清除内部灰尘，重新定标后修复 by jesson | SWO-474384 | Cleaned | HW | Process Module | Sensor, process unit | 5257 | TNT calibration failed, 5257 error | TNT bid failed, 5257, borrow edgy 944-268 (15316) one | Remove internal dust and repair by jesson after recalibration
No | Visit | Complaint | RKOR | Interface connected error | , | Interface connection. | SWO-477277 | User Error, FSE | SW | RMED SW | IP address |  | Interface connected error | , | Interface connection.
No | Visit | Complaint | RCHN | 定标失败 | 定标失败，借903-364一个 | 更换903-364后在维修菜单中进行校准失败，需要再借一个903-364，故此单关闭，另开SWO by jesson | SWO-477112 | Defect | HW | Process Module | Optical Unit | 1024 | Bid failed | Bid failure, borrow 903-364 one | After replacing 903-364, the calibration failed in the service menu, need to borrow another 903-364, so the single closed, another SWO by jesson
No | Visit | Complaint | RCHN | TNI定标失败 | TNI定标失败，借903-364一个，944-212（14869）2张944-258（14940）一张 |  | SWO-477174 | Defect | HW | Process Module | Optical Unit | 1024 | TNI calibration failed | TNI bid failed, borrow 903-364 one, 944-212 (14869) 2 944-258 (14940) one
No | Visit | Complaint | RDEU | 04.01.2018 Herr Diebel 0662188922533 meldet, dass er bei drei unterschiedlichen SC immer einen Backlogfehler angezeigt bekommt. | Kalibrierfehler im pCo2 | Empfindlichkeit im pCo2 bei 36% Alle drei Qualitätskontrollen liegen im Bereich, trotz Kalibrierfehler. Im ganzen Haus war keine Kassette mehr zu finden, außer Intensiv 3, die aber selbst die Kassette benötigen. Kunde wartet, bis Material geliefert wird und tauscht die Kassette. In der Orthop. Klinik sind die Kassetten bereits vor Weihnachten bestellt worden und das Gerät steht still. | SWO-478203 | Assisted Replacement | SC | SC | pCO2 | Cal | 04.01.2018 Mr. Diebel 0662188922533 reports that he always gets a backlog error displayed for three different ScIs. | Calibration error in pCo2 | Sensitivity in pCo2 at 36% All three quality controls are in the range, despite calibration errors. There was no cassette left in the whole house, except Intensive 3, which even need the cassette. customer waits for material to be delivered and replaces the cassette. In the Orthop. Clinic the cassettes were ordered before Christmas and the device is standing still.
No | Visit | Complaint | RPAC | SPs x2 failured while up to provide training | tried to get SPs to work without success | FOC consumable supplied | SWO-487660 | Assisted Replacement | SP | SP | SP |  | SPs x2 failured while up to provide training | tried to get SPs to work without success | FOC consumable supplied
No | Visit | Complaint | RDEU | immer wieder werden die Pat.messungen mit ? ausgedruckt, Fehler Trübung zu hoch, Oximetriefehler, | Viele Patientenmessungen werden mit ? markiert, weil kein Hb und Derivate bestimmt werden. Fehlerangaben: Oxi Spectrum Missmatch. | Hemolyzer mit Hypochlorid gespült, Ventil gereinigt, Ventilschlauch verschlissen - ersetzt, Einlassdichtung blutig - ersetzt. Startphase durchgeführt, Kalibriert und QC’s durchgeführt. 3 Test- und 2 Patientenmessungen durchgeführt - kein Oxi Fehler mehr aufgetreten. | SWO-478221 | Cleaned | HW | Oxi module | Pinch valve |  | again and again the Pat.measurements with ? printed, error turbidity too high, oximetry errors, | Many patient measurements are carried out with ? because no Hb and derivatives are determined. Error information: Oxi Spectrum Missmatch. | Hemolyzer rinsed with hypochloride, valve cleaned, valve hose worn - replaced, inlet seal bloody - replaced. start-up phase, calibrated and QC's performed. 3 test and 2 patient measurements performed - no Oxi error occurred.
No | Visit | Complaint | RAME | Going into startups regularly | Continues to reboot $Noticed software corruption causing rebooting | Replaced CPU and reinstalled all setup $Wednesday connected wireless network | SWO-478478 | Defect | HW | DMS_PC | CPU |  | Going into startups regularly | Continues to reboot /Noticed software corruption causing rebooting | Replaced CPU and reinstalled all setup $Wednesday connected wireless network
No | Visit | Complaint | RPAC | Joanne called 6458 2753 Pump calibration failing | Pump calibration error. Ran service porgram. Passed the presure test and LS adjust. Failed on pump calibration. Calibration factor is 0 | Replaced the SP, all ok | SWO-477874 | Assisted Replacement | SP | SP | Pump cal |  | Joanne called 6458 2753 Pump calibration failing | Pump calibration error. Ran service porgram. Passed the presure test and LS adjust. Failed on pump calibration. Calibration factor is 0 | Replaced the SP, all ok
No | Visit | Complaint | RPAC | Derek called 9429 2222 / 9426 6767 Parameters red after solution pack and sensor cassette change | Customer had replaced the sensor cassette and solution pack yesterday. There are issues on the OXI parameters. Checked fluid transport in service menu and all was okay. Performed complete OXI test and it failed on the hemolyser. Flushed analyser. This did not resolve the issue. | Shutdown analyser and removed wet section. Replaced the hemolyser and then reassembled machine. Started up machine and then performed wet sections activities after reinserting the sensor cassette and solution pack. Allowed system to perform its start up. Cals and QCs all passed. Performed tHb calibration and that passed as well. | SWO-477341 | Defect | HW | Oxi module | Hemolyzer |  | Derek called 9429 2222 / 9426 6767 Parameters red after solution pack and sensor cassette change | Customer had replaced the sensor cassette and solution pack yesterday. There are issues on the OXI parameters. Checked fluid transport in service menu and all was okay. Performed complete OXI test and it failed on the hemolyser. Flushed analyser. This did not resolve the issue. | Shutdown analyser and removed wet section. Replaced the hemolyser and then reassembled machine. Started up machine and then performed wet sections activities after reinserting the sensor cassette and solution pack. Allowed system to perform its start up. Cals and QCs all passed. Performed tHb calibration and that passed as well.
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | LAC定标，质控均不过 0589,0210 | 连续更换1张测试卡和2个试剂包后修复 | SWO-481220 | Assisted Replacement | SP | SP | Lac | Cal + QC | Test card, reagent package failure | LAC calibration, quality control is not more than 0589,0210 | Repair after replacing 1 test card and 2 reagent kits in a row
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 无法抽吸溶液 1119 | 更换试剂包后修复 | SWO-485480 | [5 empty] | Test card, reagent package failure | Unable to pump solution 1119 | Repair after replacement of reagent kit
No | Visit | Complaint | RCHN | 定标失败 | 定标失败,更换903-364后出厂校准不准，需要再借一个903-364 | 并非光单原因导致故障，要求加借903-702一个by tim | SWO-477181 | Defect | HW | Process Module | Process Wheel | Very old analyzer | Bid failed | Failed to set the standard, replacement 903-364 after the factory calibration is not allowed, need to borrow another 903-364 | Not a single cause of failure, requires a 903-702 by tim
No | Visit | Complaint | RCHN | 客户抱怨仪器总是报错1068 | 客户抱怨仪器总是报错1068 | 配备暖风机，下一步计划为仪器配一个保温的罩子by tim | SWO-478187 | User Error | User Error | User Error | User Error | Placed outside temperature specifications | Customer complains that the instrument always misreported 1068 | Customer complains that the instrument always misreported 1068 | Equipped with a heater, the next step is to equip the instrument with a insulated hood by tim
No | Visit | Complaint | RCHN | hct模块无法测量 | hct模块无法测量，需要借件903-329一个 |  | SWO-477455 | Defect | HW | Aspiration Module | Needle Wash Unit |  | hct module cannot be measured | Hct module cannot be measured and requires a loan of 903-329
No | Visit | Complaint | RDEU | 07.01.2019 09:00Uhr Herr Adebar 06213833528: Es werden nicht alle QC-Ampullen erkannt. | Auto Check erkennt nicht alle Ampullen | Analysator heruntergefahren - Demontage Auto Check Modul - Reinigung - Rückmontage und Funktionsüberprüfung etc. | SWO-477766 | Cleaned | HW | Autocheck module | Autocheck module |  | 07.01.2019 09:00Am Mr Adebar 06213833528: Not all QC ampoules are detected. | Auto Check does not recognize all ampoules | Analyzer shut down - Disassembly Auto Check Module - Cleaning - Backassembly and function check etc.
No | Visit | Complaint | RKK | インレットからの液漏れ 排液ラインつまり等の異常無し。納品直後の為プローブ・ガスケット交換し様子見。 | インレットから血液漏れ | プローブ・ガスケット交換 | SWO-477486 | Defect | HW | Inlet module | Inlet gasket | send til Leif (oversættelse), OOB | Leakage from inlet No abnormalities such as drainage line, that is. Probe gasket replacement and watch immediately after delivery. | Blood leak from inlet | Probe gasket replacement
No | Visit | Complaint | RDEU | DMS not connected | System überprüft => Au Board defekt | Au Board aus schwarz bestand erneuert, Kalibration und Qc durchgeführt | SWO-477648 | Defect | HW | Analyzer Control Module | Anlyzer Control Module |  | DMS not connected | System checked => Au board defective | Au board made of black consisted renewed, calibration and Qc carried out
No | Visit | Complaint | RDEU | 03.01.2018; Hr. Zimmermann; Tel.: 0401818812527; Oxy Hardware Fehler 03.01.2019, 11:52 Uhr, H. Zimmermann, Tel. 0401818812527: Weiter Oxy Hardware Fehler | OXI Hardware Fehler | Fehleranalyse, OXI Module PCB erneuert, Softwareupdate durchgeführt, Lamp Unit adj. 3480 (3100-3700) Kalibrierung incl. tHB Kal (Fcuv 0,9564) und Qualitätskontrolle durchgeführt i.O. Aqure Verbindung hergestellt und eingerichtet, Benutzer definierte Anpassungen vorgenommen | SWO-477571 | Defect | HW | Oxi module | Oxi_PCB |  | 03.01.2018; Mr. Zimmermann; Tel.: 0401818812527; Oxy Hardware Error 03.01.2019, 11:52, H. Zimmermann, Tel. 0401818812527: Continue Oxy Hardware Error | OXI Hardware Error | Error analysis, OXI module PCB renewed, software update performed, Lamp Unit adj. 3480 (3100-3700) Calibration incl. tHB Kal (Fcuv 0.9564) and quality control carried out in cl. Aqure connection established and set up, user made defined adjustments
No | Visit | Complaint | RDEU | 05.12.2018 14:30Uhr Herr Dietz 0627183277: LQC bei Tnt außerhalb Bereich 20.12.2018; 16:10 Uhr; Hr. Dietz; Tel.: 0627183671; TnT QC Ergebnisse außerhalb der RiliBÄK | Nach Rücksprache mit Technischen Support der Firma Radiometer Überprüfung des Analysators, ziehen der Daten für Kopenhagen. Bei TnT im Level rot und gelb sind alle Messergebnisse am Oberen Hersteller Grenzwert, mit RiliBäk Grenzverletzung. Anwender haben neues TnT Testkit kalibriert und neue Qualitätskontroll Charge verwendet. | Reinigung der Pipette des Wash Blocks, Funktion in Ordnung. Reinigung der Cup Wash Unit und des Process Wheels. Alle Teile und der Analysator sind sehr sauber, da wenig in Betrieb und Analysator neu aufgestellt wurde. Systemreinigung durchgeführt. Überprüfung im Service der Pipette ob sie auf den Cupboden richtig eingestellt ist. Neue Einstellung gespeichert. Nach Service wird ersichtlich, dass die Inlet Doors nicht Richtig sich öffnen. Vorher nichts an den Inlet Doors gemacht. Mechanik der Inlet Doors nach Service Manuel mit Abstandslehre eingestellt. Datalogs gezogen und Service Dump zur Auswertung für Technischen Support von Radiometer. Blankmessungen durchgeführt, Messwerte in Ordnung. | SWO-477924 | User Error | User Error | User Error | User Error | RDEU 832293698 | 05.12.2018 14:30 Mr Dietz 0627183277: LQC at Tnt outside area 20.12.2018; 4:10 p.m.; Mr. Dietz; Tel.: 0627183671; TnT QC Results outside the RiliBÄK | After consultation with technical support of the company Radiometer review of the analyzer, pull the data for Copenhagen. At TnT in the level red and yellow all measurement results at the Upper Manufacturer limit value, with RiliBäk limit violation. Users have calibrated new TnT test kit and used new quality control batch. | Cleaning the pipette of the wash block, function fine. Cleaning the Cup Wash Unit and the Process Wheel. All parts and the analyzer are very clean, as little has been repositioned in operation and analyzer. system cleaning. Check in the service of the pipette whether it is set correctly to the cup bottom. New setting saved. After service, it becomes apparent that the Inlet Doors do not open properly. Nothing done about the Inlet Doors before. Mechanics of the Inlet Doors adjusted according to Service Manuel with distance gauge. Datalogs and Service Dump for evaluation for Technical Support of Radiometer. Blank measurements performed, measurements in order.
No | Visit | Complaint | RDEU | 07.01.2019; Fr. Reichert; Tel.: 08216004890; ABL verlangt eine Spülung | Fehlersuche vor Ort | Immer wieder hab Probleme bei QC. Kal ist i.O. Lösungspack wurde mehrmals erneuert. Hemolyzer getauscht. THb kal und kal Verifizierung durchgeführt -> i.O. Mehrere Kalibrierungen durchgeführt. QCs für alle Level durchgeführt. –> i.O. ABL wieder messbereit. Mit Frau Renner noch einige Dinge wegen den Betriebsmitteln und der Bestellung durchgesprochen. | SWO-478460 | Defect | HW | Oxi module | Hemolyzer |  | 07.01.2019; Ms. Reichert; Tel.: 08216004890; ABL requires flushing | Troubleshooting on site | Time and again I have problems with QC. Kal is i.O. Solution pack has been renewed several times. Hemolyzer. THb kal and kal verification performed -> i.O. Multiple calibrations performed. QCs for all levels. –> i.O. ABL ready to be measured again. With Mrs. Renner still some things talked through because of the equipment and the order.
No | Visit | Complaint | RSAF | SP error, then Inlet gasket needs replacement. New SP and both inlet gasket w. holder and needle replaced as reported by customer. | Fluud transport errors, Needle blocked (not new) and inlet gasket w. holder is also an old one | Replaced needle and inlet gasket with holder, and reinstalled SP. Fluid transport tests ok. Start up test failed on PCO2 and Gluc calibration. QC Errors on of parameters. tHb cal done , reran QC and all oxi parameters ok. PCO2 and Gluc still need time to condition. Customer to monitor and replace SC when necessary. | SWO-478095 | Defect | HW | Auto Inlet module | Probe |  | SP error, then Inlet gasket needs replacement. New SP and both inlet gasket w. holder and needle replaced as reported by customer. | Fluud transport errors, Needle blocked (not new) and inlet gasket w. holder is also an old one | Replaced needle and inlet gasket with holder, and reinstalled SP. Fluid transport tests ok. Start up test failed on PCO2 and Gluc calibration. QC Errors on of parameters. tHb cal done , reran QC and all oxi parameters ok. PCO2 and Gluc still need time to condition. Customer to monitor and replace SC when necessary.
No | Visit | Complaint | RLTD | Please call Abdl on 0208 869 2133 | Printer working ok. Vacuum pressure error customer has flushed system and replaced inlet gasket. | Flushed system and found pinch valve tube to be leaking. Replaced tube and run service test programs all passed ok. | SWO-478268 | Defect | HW | Oxi module | Pinch valve |  | Please call Abdl on 0208 869 2133 | Printer working ok. Vacuum pressure error customer has flushed system and replaced inlet gasket. | Flushed system and found pinch valve tube to be leaking. Replaced tube and run service test programs all passed ok.
No | Visit | Complaint | RAME | Status Code 2 error | Received call stating analyzer showing status code 2. | Reinstalled software version 3.4. Did not fix issue, reinstalled windows and software. Replaced SC and SP and analyzer is working properly. | SWO-478034 | Defect | SW | RMED SW | Windows | Status code 2 | Status Code 2 error | Received call stating analyzer showing status code 2. | Reinstalled software version 3.4. Did not fix issue, reinstalled windows and software. Replaced SC and SP and analyzer is working properly.
No | Visit | Complaint | RAME | Possible Blockage | Arrived onsite to check SC connector board as discussed from conference call regarding the escalation of errors that the customer is seeing. There was no noticeable wear or leakage on the board or inside where the sensor cassette pins go. | No replacement necessary at this time. Performed zero calibration with zeroing cassette while onsite. | SWO-478751 | [5 empty] | Possible Blockage | Arrived onsite to check SC connector board as discussed from conference call regarding the escalation of errors that the customer is seeing. There was no noticeable wear or leakage on the board or inside where the sensor cassette pins go. | No replacement necessary at this time. Performed zero calibration with zeroing cassette while onsite.
No | Visit | Complaint | RDEU | Verschiedene Parameter immerwieder fehlerhaft. | Kein Lp und SC vor Ort wird bestellt | Gerät arbeitet seit Aufstellung nicht korrekt. Fliesswege überprüft und gereinigt. QC wird nicht bestanden,immer wieder bei verschiedenen Parametern, aber überwiegend bei PCO2 und PO2. Anwender wechselt LP und misst erneut QC. Rückruf bei erneuten Problemen. | SWO-482165 | [5 empty] | Various parameters are always faulty. | No Lp and SC on site is ordered | Device has not been working correctly since installation. Flow paths checked and cleaned. QC is not passed again and again on different parameters, but predominantly with PCO2 and PO2. User changes LP and measures QC again. Recall in case of re-problems.
No | Visit | Complaint | RIND | Roller assembly broken | Roller assembly broken. | Need to change roller assembly | SWO-479655 | [5 empty] | Roller assembly broken | Roller assembly broken. | Need to change roller assembly
No | Visit | Complaint | RAME | The tHb and sO2 do not compare well with the other analyzer. Consistently measures 0.5 to 1.3 g/dL lower. | THb measures low compared to two other analyzers. | Replaced the hemolyzer and spine board and reinstalled the original CO-OX Control Board. Performed system cycle, tHb calibration and ran QC. Ran blood samples and the tHb was at least 1.0 g $dL lower than the other two analyzers. | SWO-478171 | [5 empty] | The tHb and sO2 do not compare well with the other analyzer. Consistently measures 0.5 to 1.3 g/dL lower. | THb measures low compared to two other analyzers. | Replaced the hemolyzer and spine board and reinstalled the original CO-OX Control Board. Performed system cycle, tHb calibration and ran QC. Ran blood samples and the tHb was at least 1.0 g $dL lower than the other two analyzers.
No | Visit | Complaint | RCHN | 总胆红素结果为零 | 总胆红素结果为零 | 修改截距后修复by tim | SWO-477568 | Setup | SW | RMED SW | RMED SW | Bili | Total bilirubin result is zero | Total bilirubin result is zero | Fix by tim after modifying intercept
No | Visit | Complaint | RCHN | 定标失败 | 故障并非光单原因，现在要加借备件903-236一个（被903-702替代） | 最终更换了一个903-364和一个903-702后修复by tim | SWO-477627 | Defect | HW | Process Module | Process Wheel | Process wheel 2 replaced by process wheel 3 | Bid failed | The fault is not the cause of light sheet, now add spare parts 903-236 one (replaced by 903-702) | Finally replaced a 903-364 and a 903-702 after repair by tim
No | Visit | Complaint | RIND | pO2 results are not displayed | pO2 results are not displayed | Checked the patient result log , found the problem has been since 2 days , unplugged the cassette and checked if any problem with it , cassette is good , reinstalled the cassette and gave calibration ,checked a patient sample but still the problem is not solved ,restarted the instrument and then run calibration and checked with a sample result are good , customer happy with the result. | SWO-478317 | [5 empty] | pO2 results are not displayed | pO2 results are not displayed | Checked the patient result log , found the problem has been since 2 days , unplugged the cassette and checked if any problem with it , cassette is good , reinstalled the cassette and gave calibration ,checked a patient sample but still the problem is not solved ,restarted the instrument and then run calibration and checked with a sample result are good , customer happy with the result.
No | Visit | Complaint | RIND | Cassette not detecting | Cassette not detecting | Found software is hanged , restarted the instrument , reinstalled the software , now the cassette is detected and then upgraded the key ,did calibration and run sample , customer happy with the result. | SWO-478321 | [5 empty] | Cassette not detecting | Cassette not detecting | Found software is hanged , restarted the instrument , reinstalled the software , now the cassette is detected and then upgraded the key ,did calibration and run sample , customer happy with the result.
No | Visit | Complaint | RPAC | The analyser is having drifting issues at times on teh Electrolyte parameters. The Glu and Lact parameters can also be a little unstable. There is the odd sampling error as well. | The analyser had all the electrodes checked and were replaced if they did not look correct. The Chamber of the pH $BG Module was also checked. | The analyser had the Ref, K, Na, Ca, Glu and Lact electrodes replaced. All the new electrodes had new membranes placed on them. The Reference electrode certainly was found to be not correct. The pH $BG Module Electrode Chamber was replaced, as it did look a little dirty. The Inlet Probe was replaced, as it was found to be bent. A new Inlet Gasket was also put onto the analyser, this was the Pathology Dept's gasket. The analyser was returned to normal and allowed it to stabilize via the calibrations. Once all the parameters had calibrated and stabilized, QC Levels 1 & 3 were performed, all values are within their limits. The analyser now appears to be functioning fine. Please note, the parameters that had the electrodes replaced, will most likely drift at times for the next 24 hours. | SWO-478213 | Defect | HW | Inlet module | Probe bend |  | The analyser is having drifting issues at times on teh Electrolyte parameters. The Glu and Lact parameters can also be a little unstable. There is the odd sampling error as well. | The analyser had all the electrodes checked and were replaced if they did not look correct. The Chamber of the pH /BG Module was also checked. | The analyser had the Ref, K, Na, Ca, Glu and Lact electrodes replaced. All the new electrodes had new membranes placed on them. The Reference electrode certainly was found to be not correct. The pH $BG Module Electrode Chamber was replaced, as it did look a little dirty. The Inlet Probe was replaced, as it was found to be bent. A new Inlet Gasket was also put onto the analyser, this was the Pathology Dept's gasket. The analyser was returned to normal and allowed it to stabilize via the calibrations. Once all the parameters had calibrated and stabilized, QC Levels 1 & 3 were performed, all values are within their limits. The analyser now appears to be functioning fine. Please note, the parameters that had the electrodes replaced, will most likely drift at times for the next 24 hours.
No | Visit | Complaint | RIND | PH parameter not working properly. | checked the analyzer and found PH parameter not working. | checked the calibration reports and found its showing dots on sensitivity cleaned the measuring chamber and replaced reference membrane still showing sensitivity dots .so we need to replace PH electrode. | SWO-482571 | Revisit required | Electrode | Electrode | pH |  | PH parameter not working properly. | checked the analyzer and found PH parameter not working. | checked the calibration reports and found its showing dots on sensitivity cleaned the measuring chamber and replaced reference membrane still showing sensitivity dots .so we need to replace PH electrode.
No | Visit | Complaint | RFRA | Erreur T° ampoule QC. Module Autocheck reçu par le client En attente d'intervention | Erreur température sur module autocheck | Changement du module , test SAV ok , température ok . Passage calibration contrôles ok | SWO-478903 | Defect | HW | Autocheck module | Autocheck module | temperature error | Error T-bulb QC. Autocheck module received by the customer Waiting for intervention | Temperature error on autocheck module | Module change, sAV test ok, temperature ok. Passage calibration controls ok
No | Visit | Complaint | RDEU | 8.1.2019 08:50Uhr Herr Skok 071127832585: Das Gerät bootet, bleibt aber mit einer Fehlermeldung Data upgrade fail | Error2 | Datenbank gelöscht. Datenbanksicherung vom 01.01.2019. Kalibrationen und Qualitätskontrollen in Ordnung. | SWO-478501 | Defect | SW | RMED SW | Database |  | 8.1.2019 08:50 Mr Skok 071127832585: The device boots but remains fail with an error message Data upgrade | Error2 | Database deleted. Database backup from 01.01.2019. Calibrations and quality controls in order.
No | Visit | Complaint | RIND | Storage barcode not working | Checked and found storage barcode misaligned. | Storage barcode alignment performed successfully. | SWO-480434 | Adjustment | HW | Storage Module | Barcode reader, storage wheel |  | Storage barcode not working | Checked and found storage barcode misaligned. | Storage barcode alignment performed successfully.
No | Visit | Complaint | RPOL | Błąd kalibracji pompy. | Błąd kalibracji pompy. | Sprawdzenie aparatu. Wczytanie kodów kreskowych dla operatorów . | SWO-478335 | Assisted Replacement | SP | SP | Pump cal |  | Pump calibration error. | Pump calibration error. | Check the camera. Loading bar codes for operators.
Yes | Visit | Complaint | RAME | Status Code 2 error | When finished with last visit analyzer had error in sensor cassette. Returned to verify analyzer is operating properly for CAS visit. | When arrived analyzer was green. Restarted analyzer twice to verify the status code 2 didn’t return. | SWO-479297 | Re classified RMED | planned SWO |  |  |  | Status Code 2 error | When finished with last visit analyzer had error in sensor cassette. Returned to verify analyzer is operating properly for CAS visit. | When arrived analyzer was green. Restarted analyzer twice to verify the status code 2 didn’t return.
No | Visit | Complaint | RIND | Analyzer Shutdown | Analyzer is not booting, Check power module by replacing it by new one, Power is not passing through CPU . | Found that power module is ok. Need to change CPU. | SWO-479614 | [5 empty] | Analyzer Shutdown | Analyzer is not booting, Check power module by replacing it by new one, Power is not passing through CPU . | Found that power module is ok. Need to change CPU.
No | Visit | Complaint | RCHN | 扫描头故障 | 定标卡无法扫描，借902-924一个 | 更换902-924后修复by tim | SWO-477608 | Defect | HW | Cover and Chassis | Barcode reader, display |  | Scan head failure | The bid card cannot be scanned, borrowing 902-924 one | Repair by tim after replacement 902-924
No | Visit | Complaint | RCHN | 测试卡断裂，报错5784 | 测试卡断裂，报错5784 断裂部位（Teeth） | 更换测试卡后修复 | SWO-479703 | Defect | Consumables | Test Kit | Cartridge, Broken | Teeth | Test card break, 5784 error | Test card break, 5784 error Break site (Teeth) | Repair after replacing test card
No | Visit | Complaint | RKK | キャル時やリンス時に廃液ローターから異音がする | 廃液ポンプローターが回転すると 通常とは異なる音がでる | 廃液用ポンプヘッドを交換 | SWO-487658 | [5 empty] | There is a noise from the waste water rotor at the time of the car and the rinse. | When the waste pump rotor rotates, A different sound is coming out. | Replaced pump head for waste water
No | Visit | Complaint | RCHN | 热敏打印机不能工作 | 热敏打印机不能工作 | 更新驱动后修复by tim | SWO-480234 | [5 empty] | Thermal printer does not work | Thermal printer does not work | Fix by tim after update driver
No | Visit | Complaint | RIND | Crea quire for all the samples and calibration result. | Checked the data logs crea quivery for all the samples replaced the new membrane its worked for two to three days again we are getting the frequent crea problem, Need to replace Crea and dummy electrode. | Replaced the New Crea and dummy electrodes and membranes, perform two point calibration then QC done all results are in range. | SWO-479698 | Assisted Replacement | Electrode | Electrode | Crea |  | Crea quire for all the samples and calibration result. | Checked the data logs crea quivery for all the samples replaced the new membrane its worked for two to three days again we are getting the frequent crea problem, Need to replace Crea and dummy electrode. | Replaced the New Crea and dummy electrodes and membranes, perform two point calibration then QC done all results are in range.
No | Visit | Complaint | RDEU | Elyte Driftfehler 17:12:2018; 08:53 Uhr, Fr. Valentin, Tel. 053319344940: Kal Fehler bei Na und Ca | Elyte Driftfehler | Gerät geprüft, Referenzelektrode (E1001 58-01) ausgetauscht, Funktionskontrolle ok | SWO-479601 | Assisted Replacement | Electrode | Electrode | Reference | CAPA-733 | Elyte Drift Error 17:12:2018; 08:53, Ms. Valentin, tel. 053319344940: Kal error at Na and Ca | A. M. Driftfehler | Device tested, reference electrode (E1001 58-01) replaced, function control ok
No | Visit | Complaint | RAME | Customer is getting oxi mismatch error and the error 893 when they run the samples by using 125 micro lit capillary | Customer is getting error 893 and oxi spectrum mismatch error during running the samples by using capillary | The pump solution’s tube and the modules pump tubing all replaced. 10 consequent tests performed and passed with no more errors. | SWO-479293 | Defect | HW | Met II module | Pump tube |  | Customer is getting oxi mismatch error and the error 893 when they run the samples by using 125 micro lit capillary | Customer is getting error 893 and oxi spectrum mismatch error during running the samples by using capillary | The pump solution’s tube and the modules pump tubing all replaced. 10 consequent tests performed and passed with no more errors.
Yes | Visit | Complaint | RIND | Error for sO2 and Hb results | sO2 and Hb results error | Checked the calibration log and patient log , found that question mark for many sample results for so2 and Hb , therefore checked the oximetry section ,valve and its functioning , updated the software keys and then run calibration and then run samples , customer happy with the resuts. | SWO-480870 | Re classified RMED | planned SWO SW key expired |  |  |  | Error for sO2 and Hb results | sO2 and Hb results error | Checked the calibration log and patient log , found that question mark for many sample results for so2 and Hb , therefore checked the oximetry section ,valve and its functioning , updated the software keys and then run calibration and then run samples , customer happy with the resuts.
No | Visit | Complaint | RAME | Negative bias on qc for po2. | Investigation with blood and qc inspection reveals no bias on either analyzer. Cap material is not altitude corrected and reduces po2 results. Local hospitals produce much the same results on cap material. | None necessary | SWO-479349 | No fault found | No fault found | No fault found | No fault found |  | Negative bias on qc for po2. | Investigation with blood and qc inspection reveals no bias on either analyzer. Cap material is not altitude corrected and reduces po2 results. Local hospitals produce much the same results on cap material. | None necessary
No | Visit | Complaint | RIND | Pisa application software error. | Pisa application software error occurred. | Install the application software but during the installation pisa data uploader got failed ,no data was available.Try to restore the back up files but again application software was corrupted,So install the application software again & then calibrate the tropnin I parameter,done setup & handover the instrument to the customer. | SWO-480863 | User Error | SW | RMED SW | Pisa DB |  | Pisa application software error. | Pisa application software error occurred. | Install the application software but during the installation pisa data uploader got failed ,no data was available. Try to restore the back up files but again application software was corrupted,So install the application software again & then calibrate the tropnin I parameter,done setup & handover the instrument to the customer.
No | Visit | Complaint | RKOR | Hct value cannot be correctly determined | Perform the needle wash unit $ tube #9, pump 4, pump 5 cleaning | Perform the needle wash unit $ tube #9, pump 4, pump 5 cleaning | SWO-481033 | Cleaned | HW | Aspiration Module | Needle Wash Unit | Clot | Hct value cannot be correctly determined | Perform the needle wash unit / tube #9, pump 4, pump 5 cleaning | Perform the needle wash unit $ tube #9, pump 4, pump 5 cleaning
No | Visit | Complaint | RKOR | FTC 103 error | . | Solution pack leak, p4-p5 line clean. Needle clean. Wash station clean. Sample test. | SWO-480992 | Revisit required | Consumables | Solution Pack | Solution Pack | Leak, Potentially User Error | FTC 103 error | . | Solution pack leak, p4-p5 line clean. Needle clean. Wash station clean. Sample test.
No | Visit | Complaint | RSAF | Trouble shooting done and faulty pinch valve detected as root cause. | Done investigation again and root cause is intermittent pinch valve failure | Installed new pinch valve and started up analyzer, failed pco2 and glucose and lactate parameters, installed new solution pack and calibrated but pco2 error remains, installed new sensor cassette and calibrated, no change. | SWO-479741 | Defect | HW | Oxi module | Pinch valve |  | Trouble shooting done and faulty pinch valve detected as root cause. | Done investigation again and root cause is intermittent pinch valve failure | Installed new pinch valve and started up analyzer, failed pco2 and glucose and lactate parameters, installed new solution pack and calibrated but pco2 error remains, installed new sensor cassette and calibrated, no change.
No | Visit | Complaint | RDEU | Sampler Mixer gebrochen | Sampler Mixer gebrochen | Modul getauscht, Startphase QC | SWO-479935 | Defect | HW | Basic Unit | Sample mixer |  | Sampler Mixer broken | Sampler Mixer broken | Module exchanged, start phase QC
No | Visit | Complaint | RKOR | Sample clot error | . | Liquid flow line clean. Gasket module clean. Service mode check. | SWO-480994 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway | Clot | Sample clot error | . | Liquid flow line clean. Gasket module clean. Service mode check.
No | Visit | Complaint | RESP | Han cambiado reactivos y accesorios y sigue el problema | Chequeo de circuito neumático, reactivos y accesorios |  | SWO-482058 | Defect | HW | Basic Unit | PM Kit |  | Reagents and accessories have been changed and the problem continues | Pneumatic circuit check, reagents and accessories
No | Visit | Complaint | RDEU | 08.01.2019 Herr Hagen 0401818854581 - Displayrahmen gebrochen | System überprüft => Gewinde im Displayrahmen gebrochen | Displayrahmen erneuert , Da Gerät noch in der Med.Technik steht und nicht in Betrieb ist keine Kal und QC durchgeführt | SWO-479639 | Defect | HW | Basic Unit | Display | Screws | 01/08/2019 Mr Hagen 0401818854581 - Display frame broken | System checked => Thread broken in the display frame | Display frame renewed, Since device is still in the Med.Technology and is not in operation no Kal and QC is performed
No | Visit | Complaint | RDEU | verstärkt (bis zu 50% der Messungen) Meldung Oxi-Mismatch bzw. Trübung zu hoch Case wurde in der hotline falsch zugeordnet! | Oxi-Test: Cavitation absorb zu klein (<0,5), vermutlich gab es bereits vorher sporadische Ausfälle | Hämolysator erneuert, Oxi-Test i.O. | SWO-480315 | Defect | HW | Oxi module | Hemolyzer | 581, 292 | (up to 50% of measurements) Message Oxi-Mismatch or turbidity too high Case was misassigned in the hotline! | Oxi test: Cavitation absorbs too small (-0.5), presumably there have been sporadic failures before | Hemolyser renewed, Oxi test in the o.
No | Visit | Complaint | RAME | Several combinations of solution pack and cassette glu sensitivity does not rise above 0.3 | Good cassettes and solution packs put on instrument and all fail glu calibration sensitivity low | I replaced the sensor interface board added consumables and instrument glu calibrated first on first cal of initialization | SWO-480103 | Defect | HW | Basic Unit | PCB, Sensor interface | Glu | Several combinations of solution pack and cassette glu sensitivity does not rise above 0.3 | Good cassettes and solution packs put on instrument and all fail glu calibration sensitivity low | I replaced the sensor interface board added consumables and instrument glu calibrated first on first cal of initialization
Yes | Visit | Complaint | RDEU | 09.01.2018 Herr Reineke 02083092184 meldet das keine SC (crea) geliefert wurden. | SC für Crea $ Urea nicht lieferbar. | Crea $ Urea Modus deaktiviert. LS und SC organisiert. SC und LP durch Kundeneigene Betriebsmittel installiert und Gerät kalibriert. Kunde wird sich bei der Hotline melden, sobald ein neue Crea $ Urea SC ausgeliefert wurde, um die Konfiguration wieder umzustellen. | SWO-481246 | Re classified RMED | Shipping |  |  |  | 09.01.2018 Mr. Reineke 02083092184 reports that no SC (crea) were delivered. | SC for Crea ' Urea not available. | Crea's Urea mode disabled. LS and SC. SC and LP installed and calibrated by customer-owned equipment. The customer will contact the hotline as soon as a new Crea's Urea SC has been shipped to change the configuration.
No | Visit | Complaint | RLTD | Ian Robinson Tel: 01865 221596 | The analyser was showing on Live Connect but not connecting to the customers AQURE. | I tried plugging the analyser in to a network port that was known to be working and it still didn't connect to AQURE. I took the analyser to ED Majors and moved the network cable from the ABL827 to the ABL90 and it still didn't connect. I checked the setting on the ABL827 and compared them to the ABL90 and found that the ABL827 had the server IP address in the settings where the ABL90 had the server name. I put the server IP address in the ABL90 and once that had been done it connected to AQURE. I took the analyser back to its location and connected it up then made sure Quick Access worked and that all the queued data was uploded to AQURE. | SWO-480356 | Setup | SW | RMED SW | IP address |  | Ian Robinson Tel: 01865 221596 | The analyser was showing on Live Connect but not connecting to the customers AQURE. | I tried plugging the analyser in to a network port that was known to be working and it still didn't connect to AQURE. I took the analyser to ED Majors and moved the network cable from the ABL827 to the ABL90 and it still didn't connect. I checked the setting on the ABL827 and compared them to the ABL90 and found that the ABL827 had the server IP address in the settings where the ABL90 had the server name. I put the server IP address in the ABL90 and once that had been done it connected to AQURE. I took the analyser back to its location and connected it up then made sure Quick Access worked and that all the queued data was uploded to AQURE.
No | Visit | Complaint | RAME | One of the analyzer fans is making a loud noise | Troubleshot to a bad power supply fan. I also intermittent OXY Spectrum Mismatch errors during patient samples. Troubleshot that to a faulty hemolyzer. | Replaced the power supply, and hemolyzer. | SWO-480057 | [5 empty] | One of the analyzer fans is making a loud noise | Troubleshot to a bad power supply fan. I also intermittent OXY Spectrum Mismatch errors during patient samples. Troubleshot that to a faulty hemolyzer. | Replaced the power supply, and hemolyzer.
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | CA,LAC定标失败 0378，0210 | 更换测试卡正常。 | SWO-479609 | Assisted Replacement | SC | SC | Ca, Lac | 378 | Test card, reagent package failure | CA, LAC calibration failure 0378,0210 | The replacement test card is OK.
No | Visit | Complaint | RCHN | 管夹断裂 | 管夹断裂，借903-584一个 | 更换903-584后修复by tim | SWO-478792 | Defect | HW | Aspiration Module | Tube holder |  | Pipe clip break | The tube clip is broken, borrowed 903-584 one | Repair by tim after replacement 903-584
Yes | Visit | Complaint | RCHN | 触摸屏不灵光 | 触摸屏不灵光，借636-571和910-298各一个 | 上机后发现新备件质量问题，需加借备件，此单关单by tim | SWO-479582 | [5 empty] | Touch screen is not bright | Touch screen is not bright, borrow 636-571 and 910-298 each one | New spare parts quality problems are found after the on-board, spare parts need to be borrowed, this single single by tim
Yes | Visit | Complaint | RPAC | Jan called 0240 143 057 Continuous calibration drift on sodium, calcium, etc | pCO2 QC and electrolyte drifts | electrodes already dispatched to the customer | SWO-479440 | Re classified RMED | Phone hotline |  |  |  | Jan called 0240 143 057 Continuous calibration drift on sodium, calcium, etc | pCO2 QC and electrolyte drifts | electrodes already dispatched to the customer
No | Visit | Complaint | RFRA | Err 3050 CorruptXML qui revient régulierement. | 3050 CorruptXML: 18-22 $08 $18. 17 $11 $18 (x2). 14-24 $12 $18. 06 $01 $19. laboratoire@ch-cm.fr | Sauvegarde des données. Chgt CompactFlash, installation Windows et logiciel AQT90. Configuration: réseau, sauvegarde auto. Installation configuration et sauvegarde de l'AQT90. Redémarrage routine, passage Cleaning. Vérification avec utilisatrice: code utilisateurs, remontée Identification patient. Passage LQC MCC1 (x3) | SWO-479742 | Defect | SW | RMED SW | RMED SW | 3050 | Err 3050 CorruptXML that comes back regularly. | 3050 CorruptXML: 18-22 /08 /18. /17 /11 /18 (x2). 14-24 /12 /18. 06 /01 /19. laboratoire@ch-cm.fr | Backing up the data. Chgt CompactFlash, Windows installation and AQT90 software. Configuration: network, auto backup. Installing configuration and backup of the AQT90. Routine restart, Cleaning passage. User check: user code, patient identification lift. Passage LQC MCC1 (x3)
Yes | Visit | Complaint | RFRA | Err 3050 CorruptXML qui revient régulierement. | RAS | Déplacement hôpital Moissac | SWO-480363 | Re classified RMED | same day upSWO |  |  |  | Err 3050 CorruptXML that comes back regularly. | RAS | Move Hospital Moissac
No | Visit | Complaint | RDEU | 28.12.2018 Frau Kresse 067219074355 meldet einen Fehler 5784 28.12.2018 18:40Uhr Herr Kron 067219074199: das Gerät hat die Messung abgebrochen und verlangt nach jedem Neustart wieder einen Neustart 02.01.2018 09:20Uhr Herr Cikorski 067219074444: das Gerät ist weiterhin defekt, wann kommt ein Techniker | Die bestellte Gehäuseteile sind eingetroffen und werden eingebaut. | Frontabdeckung und Kassetteneinführfenster eingebaut. Neustart. eine QC. | SWO-479925 | Defect | HW | Cover and Chassis | Door, Storage | 5784 | 28.12.2018 Ms Kresse 067219074355 reports an error 5784 28.12.2018 18:40Mr Kron 067219074199: the device has aborted the measurement and requires a restart after each restart 02.01.2018 09:20Am Mr Cikorski 067219074444: the device is still defective, when is a technician coming | The ordered housing parts have arrived and are installed. | front cover and cassette insertion windows. Restart. a QC.
No | Visit | Complaint | RPOL | Błędy przepływu płynów, błędy kalibracji. | Weryfikacja statusu analizatora | Wymiana uszczelki wlotu. Regulacja przepływu płynów. Weryfikacja kalibracji, 3 krotne wykonanie kalibracji kasety czujników - stwierdzono błąd kalibracji dla K. Kaseta czujników do reklamacji. | SWO-479846 | Defect | HW | Auto Inlet module | Inlet gasket |  | Fluid flow errors, calibration errors. | Verifying analyzer Status | Replacing the inlet gasket. Fluid flow regulation. Calibration verification, 3 times calibration of the sensor cassette-calibration error was found for K. Cassette sensor for complaints.
No | Visit | Complaint | RKOR | Pressure error | . | Liquid flow line clean. Gasket module clean. Solution pack replace. Service mode check. | SWO-481009 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | Pressure error | . | Liquid flow line clean. Gasket module clean. Solution pack replace. Service mode check.
No | Visit | Complaint | RDEU | 27.12.2018 07:30Uhr Frau Wellmann 00401818827721 (med Technik): Die Nachricht kam per Mailbox um 0:00 an. Das Gerät würde seit der Aufstellung unzuverlässig laufen. Das Gerät würde ständig Gerinnsel erkannt melden und ungewöhnlich laut sein, wenn der LP gewechselt wird. Der Benutzer ist verunsichert, weil die Software sich beim Wechseln vom LP aufgehängt hat. 03.01.2019, Frau Sellmann, Tel: 0401818828816: ABL läuft unzuverlässig, ctBil wird oft nicht gemessen obwohl der Wert nicht über der Messgrenze liegt, Werte schwanken sehr oft, ständig sollen die Consumables getauscht werden. 11.01.2019; Fr. Sellmann; Tel.: 0401818828816 (Nachricht auf dem AB); ABL nicht messbereit | Analysator nicht einsatzbereit. | Analysator mehrfach gespült in beide Richtungen. Im Servicemenü alle relevanten Parameter abgearbeitet. Mehrere Kalibrationen durchgeführt danach QC durchgeführt. Fehlermeldung Inhomogene Flüssigkeiten. Wieder mal LP Fehler. LP getauscht . Immer wieder Fehlermeldung Inhomogene Spülflüssigkeit. Analysator zerlegt Jahreswartung Kitt verbaut. Analysator wieder hochgefahren Kal und QC durchgeführt keine Verbesserung. Hemolyzer und Einlasssonde getauscht. Kal und QC durchgeführt | SWO-480412 | Defect | HW | Oxi module | Hemolyzer | Fluid transport | 27.12.2018 07:30Pm Ms Wellmann 00401818827721 (med Technik): The message arrived by mailbox at 0:00. The device would run unreliably since installation. The device would constantly report clots detected and be unusually noisy when the LP is changed. The user is unsettled because the software has hung up when switching from the LP. 03.01.2019, Mrs. Sellmann, Tel: 040181888816: ABL runs unreliably, ctBil is often not measured even though the value is not above the measuring limit, values vary very often, the consumables are constantly to be exchanged. 11.01.2019; Ms. Sellmann; Tel.: 0401818828816 (message on the AB); ABL not ready for measurement | Analyzer not ready for use. | Analyzer repeatedly rinsed in both directions. All relevant parameters are processed in the service menu. Several calibrations were then performed QC. Error message Inhomogeneous liquids. Again LP error. LP exchanged . Again and again error message Inhomogeneous rinsing liquid. Analyzer disassembled annual maintenance putty installed. Analyzer restarted restarted Kal and QC performed no improvement. Hemolyzer and inlet probe. Kal and QC carried out
No | Visit | Complaint | RCHN | 试剂包门栓断裂 | 试剂包门栓断裂，借903-363一个 | 更换903-363后修复by tim | SWO-478972 | [5 empty] | Reagent door bolt break | Reagent pack door bolt broken, borrow903-363 one | Repair by tim after replacement 903-363
No | Visit | Complaint | RKK | 画面タッチ時の動作不良 | 画面タッチ時の動作不良 | 代替え機設置、機器引き上げ。 | SWO-480398 | [5 empty] | Malfunction when touching the screen | Malfunction when touching the screen | Replacement machine installation, equipment pulling up.
No | Visit | Complaint | RCHN | 定标错误 | 定标错误，借902-778一个 | 更换902-778后修复by tim | SWO-479581 | [5 empty] | Calibration error | Miscalculation, borrow 902-778 one | Repair by tim after replacement 902-778
No | Visit | Complaint | RCHN | 处理宝无法安装 | 处理宝无法安装,需要借件903-036 | 更换903-036后修复by tim | SWO-479568 | Defect | HW | Cover and Chassis | PCB, solution pack | Potentially SP related | Processing treasure cannot be installed | Processing treasure can not be installed, need to borrow 903-036 | Repair by tim after replacement 903-036
No | Visit | Complaint | RCHN | 处理盘不转动，报错5784 | HCT 不能测定，处理盘不转动，报错5784（工程师暂不借件） | 更换903-329和580-098后修复by tim | SWO-479484 | Defect | HW | Process Module | Motor, process wheel | 5784 | The processing disc does not turn, the error 5784 | HCT can not be determined, processing disc does not turn, error 5784 (engineer temporarily borrowed) | Repair by tim after replacing 903-329 and 580-098
No | Visit | Complaint | RCHN | 仪器无法保存数据 | 仪器无法保存数据 | 维护瑞美软件后by tim | SWO-479556 | Defect | SW | External SW | His/Lis |  | The instrument could not save the data | The instrument could not save the data | Maintenance of Reme software after by tim
Yes | Visit | Complaint | RCHN | ph报告不准 | PCT报告不准 | pct扫码升级，更改报告范围到0.072后完成By tim | SWO-480841 | Re classified RMED | installation report |  |  | Move to OOB trending | Ph report not allowed | PCT reporting not permitted | pct sweep code upgrade, change report ingres to 0.072 after By tim
No | Visit | Complaint | RCHN | 接口板腐蚀 | 接口板腐蚀，需要借件902-972 |  | SWO-480883 | [5 empty] | Interface plate corrosion | Interface plate corrosion, need to borrow 902-972
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 多参数定标不过 | 更换测试卡后修复 | SWO-481886 | [5 empty] | Test card, reagent package failure | Multi-parameter calibration but | Repair after replacing test card
No | Visit | Complaint | RCHN | 做O2质控出控 | 做O2质控出控，借902-972一个 | 更换902-972后修复by tim | SWO-480296 | [5 empty] | Do O2 quality control | Do O2 QUALITY CONTROL, borrow 902-972 one | Repair by tim after replacement 902-972
No | Visit | Complaint | RCHN | 定标数据丢失 | 定标数据丢失，借944-220（15125）一张 | 重新定标后修复by tim | SWO-480258 | Assisted Replacement | Consumables | Test Kit | Cartridge, cal |  | Target data loss | Bid data loss, borrow 944-220 (15125) one | Fix by tim after re-calibration
No | Visit | Complaint | RCHN | 支撑板滑丝 | 支撑板滑丝，借886-664一个 | 更换886-664后修复by tim | SWO-480377 | Defect | HW | Incubator Module | Incubator Lid | 886-664 | Support plate filament | Support plate filament, borrow 886-664 one | Repair by tim after replacing 886-664
No | Visit | Complaint | RCHN | 歇性报错1999 | 歇性报错1999 | 现场做系统清理和维护保养后修复(by chuanwei) | SWO-481087 | Cleaned | HW | Storage Module | Wheel lock | 1999.18.28 | Rest of the year, 1999 | Rest of the year, 1999 | On-site system cleaning and maintenance repair (by chuanwei)
Yes | Visit | Complaint | RCHN | 触摸屏不灵光 | 加借910-298一个 | 更换910-298后修复by tim | SWO-480239 | [5 empty] | Touch screen is not bright | Borrow 910-298 one | Repair by tim after replacement 910-298
No | Visit | Complaint | RKOR | Solution pack error | . | Solution pack replace | SWO-481011 | Defect | Consumables | Solution Pack | Solution Pack | Leak, Potentially User Error | Solution pack error | . | Solution pack replace
No | Visit | Complaint | RESP | 1184: Fuga detectada | 1184: Fuga detectada. Aguja de entrada doblada | Se enderezo la aguja de entrada, se realizaron calibraciones de fluidica desde servicio y se comprobo correcto funcionamiento mediante QC. | SWO-485168 | Adjustment | HW | Inlet module | Probe bend | 1184 | 1184: Leak detected | 1184: Leak detected. Folded input needle | The input needle was enforced, fluidic calibrations were performed from service and proper QC operation was compensated.
No | Visit | Complaint | RDEU | 14.01.2019 10:00 H. Gieß IT 089678204758: Gerät bootet nicht mehr | Status Code 2 Data Upgrade failed | Database defekt, versucht durch erneutes installieren der Software Daten zu retten, kein Erfolg , Roma.db in Roma.db.old benannt und empty dB benutzt, Analysator Neustart , Startphase, alles i.o. | SWO-481957 | Defect | SW | RMED SW | Database | Status code 2 | 14.01.2019 10:00 H. Casting IT 089678204758: Device no longer boots | Status Code 2 Data Upgrade failed | Database broken, tried by reinstalling the software to save data, no success, Roma.db named in Roma.db.old and empty dB used, Analyzer Restart , Start Phase, everything i.o.
No | Visit | Complaint | RKOR | pO2 C Control high resulte | . | ICU pO2 QC 관련에 대해 구팀장님 미팅 진행. pO2 QC C에대한 customer letter 대처하기로함. | SWO-486137 | Assisted Replacement | SP | SP | pO2 | QC C high | pO2 C High control result | . | ICU pO2 QC related to the meeting of the former chief. Customer letter for pO2 QC C.
Yes | Visit | Complaint | RDEU | 09.01.2018 Herr Reineke 02083092184 meldet das keine SC (crea) geliefert wurden. | SC für Crea $ Urea war beim letzten Einsatz nicht lieferbar. | Gerät auf Crea $Urea umgestellt. SC und LP durch Kundeneigene Komponenten ersetzt. Referenzwerte für Crea $ Urea nach Kundenvorgabe (Frau Zölkau Zentrallabor EKO Oberhausen) eingetragen. | SWO-482185 | Re classified RMED | Shipping |  |  |  | 09.01.2018 Mr. Reineke 02083092184 reports that no SC (crea) were delivered. | SC for Crea'urea was not available at the last use. | Switch to Crea $Urea. SC and LP replaced by customer-owned components. Reference values for Crea's Urea according to customer specifications (Mrs. Zölkau Central Laboratory EKO Oberhausen).
No | Visit | Complaint | RCHN | 接口板腐蚀 | 接口板腐蚀，需要借件902-972 | 更换902-972后修复（by chuanwei） | SWO-480882 | [5 empty] | Interface plate corrosion | Interface plate corrosion, need to borrow 902-972 | Repair after replacement 902-972 (by chuanwei)
No | Visit | Complaint | RCHN | 感应板故障 | 感应板故障，借902-972一个 |  | SWO-481264 | [5 empty] | Sensing board failure | Sensor board fault, borrow 902-972 one
No | Visit | Complaint | RCHN | 气压错误，报错0679 | 气压错误，报错0679。需要借件903-165一个 | 更换903-165后修复（by chuanwei） | SWO-480976 | Defect | HW | Basic Unit | Wet section | 679 | Air pressure error, error 0679 | Air pressure error, error 0679. Need to borrow 903-165 one | Repair after replacement 903-165 (by chuanwei)
No | Visit | Complaint | RCHN | 间歇性测试卡卡住和间歇性条码扫描器无法扫描 | 间歇性测试卡卡住和间歇性条码扫描器无法扫描，需要借件910-304和903-030各一个 | 更换910-304和903-030后修复（by chuanwei） | SWO-480880 | Defect | HW | Storage Module | Barcode reader, storage wheel |  | Intermittent test card and intermittent barcode scanner scan | Intermittent test card and intermittent barcode scanners cannot be scanned and need to borrow one each for 910-304 and 903-030 | Replacement 910-304 and 903-030 post-repair (by chuanwei)
No | Visit | Complaint | RCHN | 测试卡扫码器不能扫描 | 测试卡扫码器不能扫描，借910-304一个 |  | SWO-481272 | Defect | HW | Storage Module | Barcode reader, storage wheel |  | Test card scanner cannot be scanned | Test card scanner can not be scanned, borrow 910-304 one
No | Visit | Complaint | RCHN | 开机报错 | 开机报错，借636-549一个 |  | SWO-481156 | Defect | HW | Pipette | Sensor, needle | 4794 | Boot error | Boot error, borrow 636-549 one
No | Visit | Complaint | RCHN | HCT值无法被测得 | HCT值无法被测得，借903-329,580-098各一个 |  | SWO-481268 | Defect | HW | Aspiration Module | Sensor, HCT | 1223, 5784 | HCT value cannot be measured | HCT values cannot be measured, one each of 903-329,580-098
No | Visit | Complaint | RCHN | HCT值无法被测得 | HCT值无法被测得 借903-329一个 |  | SWO-481100 | Defect | HW | Aspiration Module | Sensor, HCT | 1264 | HCT value cannot be measured | HCT value cannot be measured borrowing 903-329 one
No | Visit | Complaint | RPOL | Błąd QC : pO2 | Błąd kontroli jakości pO2 na poziomie 1,2. | Wymiana SC9 oraz SP9. | SWO-481855 | [5 empty] | Błąd QC : pO2 | Quality control error pO2 at 1.2 level. | Replacing SC9 and SP9.
No | Visit | Complaint | RCHN | 检测不到测试卡，测试卡智能芯片无法读取 | 检测不到测试卡，测试卡智能芯片无法读取 借903-623，903-612各一个 |  | SWO-481181 | [5 empty] | The test card cannot be detected, the test card smart chip cannot be read | Can not detect the test card, test card smart chip can not read borrowing 903-623, 903-612 each one
No | Visit | Complaint | RIND | Service Visit. | In oximetry complete test lamp intensity and neon position failed. | Temporary shutdown the analyzer, open the oximetry module and flush it with hypo solution, now start the analyzer and go in the service mode and adjust the lamp intensity, successfully done, do the 2 point calibration and do the sample now successfully measurement comes, analyzer handed over to staff in well working condition. | SWO-485732 | Adjustment | HW | Oxi module | Lamp unit | Neon | Service Visit. | In oximetry complete test lamp intensity and neon position failed. | Temporary shutdown the analyzer, open the oximetry module and flush it with hypo solution, now start the analyzer and go in the service mode and adjust the lamp intensity, successfully done, do the 2 point calibration and do the sample now successfully measurement comes, analyzer handed over to staff in well working condition.
No | Visit | Complaint | RAME | The analyzer is presenting status code error 13. | The analyzer is presenting status code error 13 'Low hard disk space. | 01 $14 $19 Upon arrival, found the analyzer presenting the status code error 13. Shutdown the analyzer, removed the analyzer faulty CF card and, installed a new card. Loaded windows 7 and, the analyzer software version 3.4. Loaded the analyzer setup. Loaded the customers current qualicheck qc that is in use for this analyzer. Gave the old defective CF card to Karin. Calibrations for pCO2 and, the electrolytes and, metabolites were not passing. Replaced a new sensor cassette from customer's stock. Also found sO2 not passing internal qc level B, performed the service Complete oximetry test, passing within specifications. Ran the ctHb calibration, passing within specifications. replaced a new solution pack. SO2 is still low outside of range for qc lev el B. I tested the solution pack with another analyzer and, the solution pack worked good. Determined the hemolyzer unit needs to be replaced. 01 $15 $19 Replaced a new hemolyzer unit and, the hemolyzer pinch valve tubing. Ran the service Complete Oxi test and, ran the ctHb calibration, both checks passed within specifications. | SWO-484317 | Defect | HW | DMS_PC | CF card | Status code 13 | The analyzer is presenting status code error 13. | The analyzer is presenting status code error 13 'Low hard disk space. | 01 $14 $19 Upon arrival, found the analyzer presenting the status code error 13. Shutdown the analyzer, removed the analyzer faulty CF card and, installed a new card. Loaded windows 7 and, the analyzer software version 3.4. Loaded the analyzer setup. Loaded the customers current qualicheck qc that is in use for this analyzer. Gave the old defective CF card to Karin. Calibrations for pCO2 and, the electrolytes and, metabolites were not passing. Replaced a new sensor cassette from customer's stock. Also found sO2 not passing internal qc level B, performed the service Complete oximetry test, passing within specifications. Ran the ctHb calibration, passing within specifications. replaced a new solution pack. SO2 is still low outside of range for qc lev el B. I tested the solution pack with another analyzer and, the solution pack worked good. Determined the hemolyzer unit needs to be replaced. 01 $15 $19 Replaced a new hemolyzer unit and, the hemolyzer pinch valve tubing. Ran the service Complete Oxi test and, ran the ctHb calibration, both checks passed within specifications.
No | Visit | Complaint | RDEU | Keine Verbindung zu AQURE | Keine Verbindung zu AQURE | Netzwerkverbindung zu AQURE hergestellt | SWO-482277 | Setup | SW | RMED SW | Aqure | Network | No connection to AQURE | No connection to AQURE | Network connection to AQURE established
No | Visit | Complaint | RLTD | Nessa 07534 067522 | Leak found to be due to small hole in the oxi pinch valve tube. Fitted new tube. Restarted analyser and verified that pressure test is now passed. Upgraded analyser software to v3.4 MR1 (includes fix for printer intermittently stopping printing). Verified operation analyser after restart - all parameters ok except for pCO2 which has low sensitivity on calibration. Customer to monitor pCO2 and if necessary replace the Sensor Cassette. | Replaced oxi pinch valve tube. | SWO-482312 | Defect | HW | Oxi module | Pinch valve tube |  | In this 07534 067522 | Leak found to be due to small hole in the oxi pinch valve tube. Fitted new tube. Restarted analyser and verified that pressure test is now passed. Upgraded analyser software to v3.4 MR1 (includes fix for printer intermittently stopping printing). Verified operation analyser after restart - all parameters ok except for pCO2 which has low sensitivity on calibration. Customer to monitor pCO2 and if necessary replace the Sensor Cassette. | Replaced Oxygen Pinch valve tube.
Yes | Visit | Complaint | RLTD | Software upgrade to 3.4MR1 for printer issue. | Software upgrade to 3.4MR1 for printer issue. | Software upgrade from 3.4 too 3.4MR for printer issue ok. On icu 092R0168N0065 pacu. Lab spare 092R169N004 ok | SWO-482201 | Re classified RMED | planned SWO |  |  | SW update (printer) | Software upgrade to 3.4MR1 for printer issue. | Software upgrade to 3.4MR1 for printer issue. | Software upgrade from 3.4 too 3.4MR for printer issue ok. On icu 092R0168N0065 pacu. Lab spare 092R169N004 ok
No | Visit | Complaint | RLTD | Charlene 01708 435 000 Ext 6672 or 6845 (ext 6845 is best for out of hours) |  |  | SWO-482145 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | Charlene 01708 435 000 Ext 6672 or 6845 (ext 6845 is best for out of hours)
No | Visit | Complaint | RAME | Analyzer failing to communicate to LIS | Analyzer failing to communicate with network, there was comma prior to an IT update. | Changed analyzer to static IP as requested and communication established. | SWO-482293 | Setup | SW | RMED SW | His/Lis | static IP | Analyzer failing to communicate to LIS | Analyzer failing to communicate with network, there was comma prior to an IT update. | Changed analyzer to static IP as requested and communication established.
No | Visit | Complaint | RLTD | Mori Tel: 01865221596 | Intermittent problems with the printer not printing and requiring the analyser to be rebooted to get it working again. | A software issue with the printer had been identified on Windows7 analysers and a new software version has just been released to resolve this. The analyser was upgraded to software version 3.4MR1. Also the analyser will not communicate with the Radiometer Remote Support service so a complete software re-installation was performed to see if it will resolve this. After the software had been re-installed I waited to make sure the analyser started up ok. | SWO-482624 | Defect | SW | RMED SW | RMED SW | printer | Mori Tel: 01865221596 | Intermittent problems with the printer not printing and requiring the analyser to be rebooted to get it working again. | A software issue with the printer had been identified on Windows7 analysers and a new software version has just been released to resolve this. The analyser was upgraded to software version 3.4MR1. Also the analyser will not communicate with the Radiometer Remote Support service so a complete software re-installation was performed to see if it will resolve this. After the software had been re-installed I waited to make sure the analyser started up ok.
No | Visit | Complaint | RIND | Machine is shutdown | Machine is shutdown | Periviously Check the analyzer by interchanging power module. Change the CPU ,Found that still machine is not rebooting. Change the Signal data board, sensor board, CPU . Now the analyzer is ON. | SWO-486756 | [5 empty] | Machine is shutdown | Machine is shutdown | Periviously Check the analyzer by interchanging power module. Change the CPU ,Found that still machine is not rebooting. Change the Signal data board, sensor board, CPU . Now the analyzer is ON.
No | Visit | Complaint | RCHN | 老师反映离子的水平数值有差异 | 老师反映离子的水平数值有差异 | 去现场做了校零，清洁测试卡后完成（by chuanwei） | SWO-481843 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | The teacher reported a difference in the horizontal value of the ions | The teacher reported a difference in the horizontal value of the ions | Go to the site to do school zero, clean the test card after completion (by chuanwei)
No | Visit | Complaint | RCHN | 测试卡门机械故障 | 测试卡门机械故障 | 重新调试轨道后修复（by chuanwei） | SWO-481875 | Adjustment | HW | Cover and Chassis | Door, Storage |  | Testing Carmen for mechanical failure | Testing Carmen for mechanical failure | Post-debug track repair (bywei chuan)
No | Visit | Complaint | RCHN | 测试卡锁故障 | 测试卡锁故障5785,借903-030一个。 | 更换903-030一个正常。（by chuanwei） | SWO-481358 | Defect | HW | Storage Module | PCB, Cartridge lock | 5785 | Test card lock failure | Test card lock failure 5785, borrow 903-030 one. | Replace 903-030 one normal. (by chuanwei)
No | Visit | Complaint | RIND | Rinse error occurred | Rinse error occurred. | Clean the gasket,rinse error solved but leakage from gasket continues,clean the waste path,waste pump tubes,run the pump cal & LS Adjust but still the error persist,Replace the gasket & leakage stopped. | SWO-486100 | Defect | HW | Inlet module | Inlet gasket |  | Rinse error occurred | Rinse error occurred. | Clean the gasket,rinse error solved but leakage from gasket continues,clean the waste path,waste pump tubes,run the pump cal & LS Adjust but still the error persist,Replace the gasket & leakage stopped.
No | Visit | Complaint | RSCH | Rep OXI phase regulated | OXI Hardware Fehler 269 Phase regulation Failed nur mit Blut seit Installation Komplett OXI Test Fehlerfrei | Hemolyzer ausgetauscht Alles justiert (Lampenwert nicht über 3400 einstellbar) bei Wartung evtl Spectrometer austauschen tHb Kal (1. Messung Oxi Spectrum Missmatch) Messung mit Blut i.o | SWO-481982 | Defect | HW | Oxi module | Hemolyzer |  | REP OXYGEN Phase Regulated | OXI Hardware Error 269 Phase regulation Failed Only with Blood Since Installation Complete OXI Test Error-free | Hemolyzer replaced Everything adjusted (lamp value not more than 3400 adjustable) during maintenance possibly replace spectrometer tHb Kal (1st measurement Oxi Spectrum Missmatch) Measurement with blood i.o
No | Visit | Complaint | RDAN | ABL'en taber dato når den slukkes. - den har været i kasse fra den 26.11.2018 og frem til nu - her har den tabt dato of tid. - nu når den slukkes taber den dato men ikke tiden ? Se vedhæftet billeder. | ABL'en taber dato når den slukkes. - den har været i kasse fra den 26.11.2018 og frem til nu - her har den tabt dato of tid. - nu når den slukkes taber den dato men ikke tiden ? Se vedhæftet billeder. | ABL'en taber dato når den slukkes. - den har været i kasse fra den 26.11.2018 og frem til nu - her har den tabt dato of tid. - nu når den slukkes taber den dato men ikke tiden ? Se vedhæftet billeder. | SWO-482314 | Defect | HW | DMS_PC | DMS_PC |  | The ABL loses date when it is switched off. -It has been in the box from the 26.11.2018 up to now-here has the lost date of time. -Now when it turns off loses that date but not the time? See attached pictures. | The ABL loses date when it is switched off. -It has been in the box from the 26.11.2018 up to now-here has the lost date of time. -Now when it turns off loses that date but not the time? See attached pictures. | The ABL loses date when it is switched off. -It has been in the box from the 26.11.2018 up to now-here has the lost date of time. -Now when it turns off loses that date but not the time? See attached pictures.
No | Visit | Complaint | RESP | Varios errores de calibración en múltiples parámetros | Chequeo de neumática, chequeo de electrodos y estado de equipo | Sustitución de sonda doblada, sustitución de electrodos de PH, Glucosa y lactado por deterioro de éstos | SWO-487726 | Defect | HW | Inlet module | Probe bend | 376 | Multiple calibration errors in multiple parameters | Pneumatic check, electrode check and equipment status | Replacing bent probe, replacing PH, glucose and lactating electrodes with impaired them
No | Visit | Complaint | RKOR | Solution pack leak | . | Solution pack check, Liquid flow line check. | SWO-486139 | Defect | Consumables | Solution Pack | Solution Pack | leak | Solution pack leak | . | Solution pack check, Liquid flow line check.
No | Visit | Complaint | RDEU | Spontane HCT Mess-Fehler | Fehleranalyse mit KorLo durchgeführt. Fehler 1223, 1260, 1261 | Es wird nach einer Erklärung gesucht, warum in der Pr. Dr. Scholley immer wieder der HCT Messfehler auftritt. Seit der Inbetriebnahme am 7.1.2019 ist der Fehler, wie beim alten AQT zwei mal aufgetreten. Bei einer Wiederholung der Messung ist dann alles wieder o.k. Bei QC tritt der Fehler nicht auf. ( Dies war auch beim alten AQT nicht sichtbar ) Probennehmer überprüft - o.k. Blutabnahme überprüft - o.k. Es gibt keine medikamentösen Anhaltspunkte . Dem Anwender wurde empfohlen das Reinigungsverfahren mit der Reinigungslösung nun 2 x Wöchentlich durchzuführen. Nach 4 Wochen werden dann die Messergebnisse ausgewertet. | SWO-484867 | Cleaned | HW | Aspiration Module | Sensor, HCT | send to Jan | Spontaneous HCT measurement errors | error analysis performed with KorLo. Error 1223, 1260, 1261 | An explanation is being sought as to why the HCT measurement error occurs again and again in the Pr. Dr. Scholley. Since commissioning on 7.1.2019, the error has occurred twice, as with the old AQT. When the measurement is repeated, everything is OK again. For QC, the error does not occur. ( This was not visible even with the old AQT ) Sampler checked - ok Blood collection checked - ok There is no drug evidence. The user was advised to carry out the cleaning procedure with the cleaning solution now 2 x weekly. After 4 weeks, the measurement results are evaluated.
No | Visit | Complaint | RPOL | Wykryto przeciek, wymiana uszczelki wlotu oraz pakietu nie pomogła. | Pęknięty wężyk na zaworze OXI. | Wymiana wężyka zaworowego wraz z całym zestawem przeglądowym. | SWO-481986 | Defect | HW | Oxi module | Pinch valve tube |  | Leakage detected, replacement of inlet gasket and package did not help. | Cracked tube on the OXI valve. | Replacing the valve tube with the entire inspection kit.
No | Visit | Complaint | RHUN | Oldatpakk cserét kér, gyakori folyadék áramlási hibák | Oldatpakk cserét kér, gyakori folyadék áramlási hibák | Átöblítés, Oldatpakk cseréje, Bemeneti modul ellenőrzése és megigazítása, Inicializálás, QC mérés | SWO-482744 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | Solution packs require replacement, frequent fluid flow defects | Solution packs require replacement, frequent fluid flow defects | Flush, Solution pack replacement, input module inspection and alignment, initialization, QC measurement
No | Visit | Complaint | RDEU | PH E'Lyte Empfindlichkeitsfehler, Dialysatmodus | PH E'Lyte Empfindlichkeitsfehler, Dialysatmodus | SC geliefert, SC und SP installiert, Dialysatmodus eingestellt, Einweisung durchgeführt , neues defektes SP ersetzt, Gerät kalibriert und mit QC überprüft,Ok | SWO-484855 | Assisted Replacement | SC | SC | Reference |  | PH E'Lyte Empfindlichkeitsfehler, Dialysatmodus | PH E'Lyte Empfindlichkeitsfehler, Dialysatmodus | SC supplied, SC and SP installed, dialysis mode set, instruction performed, new defective SP replaced, device calibrated and checked with QC,Ok
No | Visit | Complaint | RAME | Battery not holding charge. | Battery will not hold charge. Shows 0% on screen and analyzer shuts off immediately when unplugged. | Replaced battery. | SWO-484591 | Defect | HW | Battery module | Battery module | 0% | Battery not holding charge. | Battery will not hold charge. Shows 0% on screen and analyzer shuts off immediately when unplugged. | Replaced battery.
No | Visit | Complaint | RDEU | 15.01.2019 08:25Uhr Herr Vollmer 051218941817: Die Station meldet Div. Fehler, die Station benötigt das Gerät dringend, ein Techniker soll kommen | Hemolyzer defekt | Fehlersuche, Fehler 290, 292, 581, Oxikompletttest, cavitation absorb zu niedrig, Austausch Hemolyzer, tHb-Kal, Funktionskontrolle, Blutmehrfachmessungen | SWO-484932 | Defect | HW | Oxi module | Hemolyzer | 581 | 15.01.2019 08:25OPM Mr. Vollmer 051218941817: The station reports Div. error, the station urgently needs the device, a technician should come | Hemolyzer defective | Troubleshooting, error 290, 292, 581, oxicomplete test, cavitation absorb too low, replacement hemolyzer, tHb-Kal, function control, blood multiple measurements
Yes | 1 | Y | RCZE | Blank |  |  | D58B39C1902BF060 | Re classified RMED | planned SWO |  |  |  | Blank
No | Visit | Complaint | RAME | SC check failing | SC check failing on multiple sensor cassettes. | Replaced analog PCB. Performed oxi alignment and tHb calibration. | SWO-484677 | [5 empty] | SC check failing | SC check failing on multiple sensor cassettes. | Replaced analog PCB. Performed oxi alignment and tHb calibration.
No | Visit | Complaint | RCHN | 电池无法蓄电 | 电池无法蓄电 借431-018C | 更换431-018C一个正常。by jesson | SWO-482541 | [5 empty] | The battery can't store power. | Battery can't store power, borrow 431-018C. | Replace 431-018C one normal. by jesson
No | Visit | Complaint | RCHN | 电池无法蓄电 | 电池无法蓄电 借431-018C | 更换431-018C一个正常。by jesson | SWO-482542 | [5 empty] | The battery can't store power. | Battery can't store power, borrow 431-018C. | Replace 431-018C one normal. by jesson
No | Visit | Complaint | RCHN | 无法读取试剂包 | 无法读取试剂包，借903-158一个。 | 更换903-158后修复by tim | SWO-481304 | Defect | HW | Basic Unit | PCB, solution pack |  | Unable to read the package | The reagent package could not be read, borrowing 903-158 one. | Repair by tim after replacement 903-158
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 参比电极错误 1361 | 更换测试卡后修复 | SWO-485483 | [5 empty] | Test card, reagent package failure | Reference electrode error 1361 | Repair after replacing test card
No | Visit | Complaint | RCHN | 样本量不足 | 样本量不足 借903-329一个 | 更换903-329后修复（by chuanwei） | SWO-481825 | Defect | HW | Aspiration Module | Sensor, HCT |  | Insufficient sample size | Sample size is insufficient, borrow 903-329 one. | Repair after replacement 903-329 (by chuanwei)
No | Visit | Complaint | RCHN | 定标失败 | BNP定标失败，借944-258（14485缺货）一张 | 重新定标正常。by tim | SWO-481943 | Assisted Replacement | Consumables | Test Kit | Cal | BNP | Bid failed | BNP bid failure, borrow 944-258 (14485 out of stock) one | Recalibration is normal. by tim
No | Visit | Complaint | RCHN | 进样舱门关不住 | 进样舱门关不住，未更换备件 | 重新调整舱门齿轮后修复 by jesson | SWO-484844 | Adjustment | HW | Cover and Chassis | Door, Inlet | 4548 | The entrance door won't be closed. | The sample door cannot be closed and spare parts have not been replaced | Repair by jesson after recalibrating hatch gear
No | Visit | Complaint | RCHN | 定标失败，中止 | 定标失败，中止，借944-268（15316）一张 | 清洁液路，重新定标后修复 by jesson | SWO-481862 | Defect | Consumables | Solution Pack | Solution Pack | 3410, ftc_101 | Bid failed, aborted | Failed to bid, aborted, borrow 944-268 (15316) a | Clean the liquid road, fix by jesson after re-calibration
No | Visit | Complaint | RDEU | Hotline - Driften bei Lyte und pH | Fehlersuche, Kalarchiv durchgesehen - seit Woche gibt es immer wieder Driften bei Lyten, es wurde bereits mehrfach membraniert | Ref Elektrode ersetzt, Kal und QC | SWO-484841 | Assisted Replacement | Electrode | Electrode | Reference | CAPA-733 | Hotline - Drifting at Lyte and pH | Troubleshooting, Kalarchiv looked through - since week there are again and again drifts at Lyten, it has already been membraneed several times | Ref electrode replaced, Kal and QC
No | Visit | Complaint | RDEU | 16.01.2019, 08:35 Uhr, Fr. Glombig, Tel. 0212645645431: Kal Fehler bei Lyten | Driften bei Ca und Na | Messkammern gereinigt, Elektrodenconnectoren von Ref, Ca , und NA gereinigt, Ventile im PH Modul und El $Met Modul geprüft, Lösungsventil im ELMet Modul gereinigt und gängig gemacht , Ventilmembran ersetzt, Pumpenkal, LS Justage, mehrere Kalibrationen und Qcs bestanden , nur noch marginale Driften | SWO-484903 | Defect | HW | El/Met Module | Valve, membrane |  | 16.01.2019, 08:35, Fr. Glombig, Tel. 0212645645431: Kal Error at Lyten | Drifting at Ca and Na | Measuring chambers cleaned, electrode connectors cleaned by Ref, Ca, and NA, valves tested in PH module and El $Met module, solution valve cleaned in THE ELMet module and made commonly made, valve membrane replaced, pump scale, LS adjustment, several calibrations and qcs passed, only left marginal drifts
No | Visit | Complaint | RAME | Customer is getting oxi mismatch error and the error 893 when they run the samples by using 125 micro lit capillary | Oxi mismatch error, error 893 and lose tHb issues have seen. The sample error have seen multiple times due to flex q module line up issue. pH $BG pump won’t work properly during sample run. | Performed a Oxi complemented test and passed. Replaced Met II module’s air and rinse valves. Probe replaced. Some other tubing replaced. The flex q module re adjusted and re calibrated. 5 real samples ran as 55 micro pH+BG+oxi without issues. All tHb values were correlated. Pump Calibration in service done. | SWO-484537 | Defect | HW | Oxi module | Pump Tube | 893 | Customer is getting oxi mismatch error and the error 893 when they run the samples by using 125 micro lit capillary | Oxi mismatch error, error 893 and lose tHb issues have seen. The sample error have seen multiple times due to flex q module line up issue. pH /BG pump won’t work properly during sample run. | Performed a Oxi complemented test and passed. Replaced Met II module’s air and rinse valves. Probe replaced. Some other tubing replaced. The flex q module re adjusted and re calibrated. 5 real samples ran as 55 micro pH+BG+oxi without issues. All tHb values were correlated. Pump Calibration in service done.
No | Visit | Complaint | RPAC | Whilst on site COhb QC had failed above range. Inspected data logs to find frequent electrolyte calibration drift. | COHb QC S7835 failed above range. Frequent calibration drift on electrolytes since reference membrane replaced 10 $1 R2387. Slight leak on reference membrane. Cleaned chambers and replaced reference membrane R2387. Clean, LS, calibration. QC S7835, S7855 | Replaced reference membrane. Suspect faulty LOT R2387. Will continue to monitor and replace electrode if errors persist. FOC replacement membranes 942-058R2407for this and ED 16 $1 SO 38336 | SWO-482518 | Assisted Replacement | Electrode | Electrode | Reference | CAPA-733 | Whilst on site COhb QC had failed above range. Inspected data logs to find frequent electrolyte calibration drift. | COHb QC S7835 failed above range. Frequent calibration drift on electrolytes since reference membrane replaced 10 /1 R2387. Slight leak on reference membrane. Cleaned chambers and replaced reference membrane R2387. Clean, LS, calibration. QC S7835, S7855 | Replaced reference membrane. Suspect faulty LOT R2387. Will continue to monitor and replace electrode if errors persist. FOC replacement membranes 942-058R2407for this and ED 16 $1 SO 38336
No | Visit | Complaint | RIND | Maintenance needed | Maintenance needed | Checked the clean cycle count , reached critical limit , did cleaning then instrument came to ready condition and then run sample. | SWO-484529 | Assisted Replacement | Consumables | Cleaning solution | Cleaning solution |  | Maintenance needed | Maintenance needed | Checked the clean cycle count , reached critical limit , did cleaning then instrument came to ready condition and then run sample.
No | Visit | Complaint | RPOL | Niejednorodny roztwór cal 3 | Niejednorodny roztwór cal 3 na dwóch pakietach z tej samej serii. Aparat żąda wymiany pakietu każdego dnia. | Założono kolejny pakiet z innego lotu. Prewencyjna wymiana uszczelki wlotu. Sprawdzono igłę. | SWO-484357 | Assisted Replacement | SP | SP | FT | Cal 3 | Heterogeneous solution inch 3 | A heterogeneous solution of Cal 3 on two packages in the same series. The engine requests a package exchange each day. | Another package was assumed from another flight. Preventive replacement of inlet gasket. The needle was checked.
No | Visit | Complaint | RPOL | Błędy ciśnienia | Weryfikacja statusu analizatora | Weryfikacja drożności układu hydraulicznego. Stwierdzono brak drożności w igle aspirującej. Brak możliwości udrożnienia igły - wymiana. Kalibracja, kontrola, testowanie. | SWO-484488 | Defect | HW | Auto Inlet module | Probe | Clot | Pressure errors | Verifying analyzer Status | Verification of the hydraulic system's accuracy. There was no patency in the aspiration needle. No possibility of a needle-swapping. Calibration, inspection, testing.
No | Visit | Complaint | RAME | pCO2 failing QCs on the new SC he installed yesterday. | Pc02 still out on levels 1&3 of internal QC only. All other tests and external QC passed | Replaced SP with a different lot number since the prior 3 were the same lot. Still having the same issues with pc02. Removed wet section and SC and discovered one pin was bent on the sensor interface board. Part ordered and will return on 1 $18 $19 | SWO-485520 | Revisit required | HW | Basic Unit | PCB, Sensor interface | pCO2 | pCO2 failing QCs on the new SC he installed yesterday. | Pc02 still out on levels 1&3 of internal QC only. All other tests and external QC passed | Replaced SP with a different lot number since the prior 3 were the same lot. Still having the same issues with pc02. Removed wet section and SC and discovered one pin was bent on the sensor interface board. Part ordered and will return on 1 $18 $19
No | Visit | Complaint | RAME | pCO2 failing QCs on the new SC he installed yesterday. | Customer reported that the pc02 was out on internal QC. When I spoke with him he also said that the analyzer was locked up and wouldn’t recognize the SP. Found the analyzer stuck in the insert SP screen | Analyzer would not recognize SP and advance to close the inlet after replacing the SP. placed old SP back in and found it recognized fine and showed 0 activities. Replaced SP with customer stock and found it was recognized fine, it was the same lot number also BU-63. After getting the analyzer reinitialized the same pc02 internal QC was not passing. Also there has been a chloride sensor response issue with the past 3 SC. Ran sensor simulator test and found all tests passed. Will resume service call on 1 $17 $19 before replacing more consumables. | SWO-484802 | Revisit required | HW | Basic Unit | PCB, Sensor interface | pCO2 | pCO2 failing QCs on the new SC he installed yesterday. | Customer reported that the pc02 was out on internal QC. When I spoke with him he also said that the analyzer was locked up and wouldn’t recognize the SP. Found the analyzer stuck in the insert SP screen | Analyzer would not recognize SP and advance to close the inlet after replacing the SP. placed old SP back in and found it recognized fine and showed 0 activities. Replaced SP with customer stock and found it was recognized fine, it was the same lot number also BU-63. After getting the analyzer reinitialized the same pc02 internal QC was not passing. Also there has been a chloride sensor response issue with the past 3 SC. Ran sensor simulator test and found all tests passed. Will resume service call on 1 $17 $19 before replacing more consumables.
No | Visit | Complaint | RKOR | System check vacuume error | . | Gasket hold replace. Service check. | SWO-486144 | Defect | HW | Inlet module | Inlet gasket | Vacuume error | System check vacuume error | . | Gasket hold replace. Service check.
No | Visit | Complaint | RSCH | Startet nicht mehr, Windows beschädigt. | Startet nicht mehr, Windows beschädigt. | Flashcard ers. Windows 7 und DMS neu installiert ok. Konfiguration geladen. Ok. Diverse Einstellungen vorgenommen. Ok Gestartet ok (jedoch ohne Lac und Glu) *** Sensorkassette wird vom Kunde ers. *** Keine Kal und QC | SWO-485675 | Defect | HW | DMS_PC | CF card |  | No more starting Windows corrupted. | No more starting Windows corrupted. | Flashcard ers. Windows 7 and DMS reinstalled ok. Configuration loaded. Ok. Various settings made. Ok Started ok (but without Lac and Glu) Sensor cartridge is used by the customer. *** No Kal and QC
No | Visit | Complaint | RPAC | solution pack failured they tried 3 attempts to refit | visited site as new account SP failured | replace SP | SWO-486269 | Assisted Replacement | SP | SP | SP |  | solution pack failured they tried 3 attempts to refit | visited site as new account SP failured | replace SP
No | Visit | Complaint | RLTD | Intermittent oximetry failures on QC. | Oximetry parameters are intermittently failing QC. | Replaced Hemolyser unit. Complete oximetry test, calibration and tHb calibration. | SWO-485308 | Defect | HW | Oxi module | Hemolyzer |  | Intermittent oximetry failures on QC. | Oximetry parameters are intermittently failing QC. | Replaced Hemolyser unit. Complete oximetry test, calibration and tHb calibration.
No | Visit | Complaint | RAME | Analyzer has an oxi hardware problem 1012. | Performed a complete oxi test and it failed LED peak intensity a the cavitation absorb. | Replaced the hemolyzer. | SWO-485423 | Defect | HW | Oxi module | Hemolyzer | 1012 | Analyzer has an oxi hardware problem 1012. | Performed a complete oxi test and it failed LED peak intensity a the cavitation absorb. | Replaced the hemolyzer.
No | Visit | Complaint | RSAF | no oximetry resuls on patient and qc results, and tHb fails on calibrations. Restarted analyzer but calibrations still fail tHb. | Faulty spectrometer | Installed complete oximetry module, done zeroing of oximetry module and all passed, calibration done and ok, ran tHb cal and ok, ran Qc's A,B and C and all ok. | SWO-484911 | Defect | HW | Oxi module | Spectrometer |  | no oximetry resuls on patient and qc results, and tHb fails on calibrations. Restarted analyzer but calibrations still fail tHb. | Faulty spectrometer | Installed complete oximetry module, done zeroing of oximetry module and all passed, calibration done and ok, ran tHb cal and ok, ran Qc's A,B and C and all ok.
No | Visit | Complaint | RDEU | Der Kunde berichtet , dass er mehrfach Probenabbrüche hat, vorzugsweise am Vormittag (3-5 Messungen), Fehlermeldung: keine Probe erkannt | eventuell LS1 instabil, muss noch beobachtet werden | Kontrolle der Fließwege, Erneuerung Dichtung, Überprüfung der Flüssigkeitssensoren, Funktionskontrolle, Blutmessungen | SWO-484934 | No fault found | No fault found | No fault found | No fault found | Sample aspiration | The customer reports that he has multiple sample aborts, preferably in the morning (3-5 measurements), error message: no sample detected | possibly LS1 unstable, still needs to be observed | Control of flow paths, renewal seal, check of liquid sensors, function control, blood measurements
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 试剂包日期无法运用 1106 LAC定标不过 | 连续更换1个试剂包和1张测试卡后修复 | SWO-485628 | [5 empty] | Test card, reagent package failure | Reagent pack date cannot be used 1106 LAC calibration is not | Repair after replacing 1 reagent kit and 1 test card in a row
No | Visit | Complaint | RCHN | 系统崩溃 | 系统崩溃 （没有使用备件） | 重装系统和软件后修复 by jesson | SWO-484773 | Defect | SW | RMED SW | RMED SW | crash | System crash | System crash (No spare parts used) | Fix by jesson after reloading systems and software
No | Visit | Complaint | RCHN | 无法识别试剂包智能芯片 | 无法识别试剂包智能芯片 借903-158一个 | 更换903-158后修复 by jesson | SWO-482520 | Defect | HW | Basic Unit | PCB, solution pack |  | Unable to identify the reagent package smart chip | Unrecognizable reagent package smart chip borrowed 903-158 one | Repair by jesson after replacement 903-158
No | Visit | Complaint | RCHN | 系统温度超范围 | 系统温度超范围，借903-581一个 | 更换903-581后修复 by Andrew | SWO-482703 | Defect | HW | Basic Unit | PCB, Sensor interface |  | System temperature over range | System temperature over range, borrow 903-581 one | Repair by Andrew after replacement 903-581
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 探测到试剂包漏电流 0963 | 更换试剂包正常。 | SWO-485487 | [5 empty] | Test card, reagent package failure | Detection of reagent packet leakage current 0963 | The replacement kit is OK.
No | Visit | Complaint | RCHN | 读取条码失败 | 读取条码失败，借910-304一个 | 更换910-304后修复 by Andrew | SWO-482550 | Defect | HW | Storage Module | Barcode reader, storage wheel |  | Failed to read barcode | Failed to read barcodes, borrowing 910-304 one | Repair by Andrew after replacing 910-304
No | Visit | Complaint | RCHN | 测试卡断裂 | 测试卡断裂 断裂部位（Nose） | 检查测试卡锁，将测试卡断裂部分取出后修复 by jeeson | SWO-484832 | Defect | Consumables | Test Kit | Cartridge, Broken | Nose | Test card break | Test card break safault (Nose) | Check the test card lock and fix by jeeson after removing the broken part of the test card
No | Visit | Complaint | RCHN | 客户抱怨仪器总是报错1068 | 仪器报错1068，温度错误 | 做了一个保温罩，加了一个加热垫后修复 by Andrew | SWO-485522 | User Error | User Error | User Error | User Error | Placed outside temperature specifications | Customer complains that the instrument always misreported 1068 | Instrument error 1068, temperature error | Made a thermal shield, added a heating pad to repair by Andrew
No | Visit | Complaint | RCHN | 王老师反馈我们仪器做出的肌钙蛋白的结果与医院检验科的数据作比对，结果是有差异的 | 王老师反馈我们仪器做出的肌钙蛋白的结果与医院检验科的数据作比对，结果是有差异的 | 由于他们新装机对比，对比仪器为实验室贝克曼，所以两台仪器单纯数值上不一样。但阴阳性相符，下周AS还会过去和护士长沟通一下 by jesson | SWO-484298 | User Error | User Error | User Error | User Error | TnI harmonization misunderstanding | Mr. Wang gave feedback on the results of troponin made by our instrument satropon in comparison with the data of the hospital's testing department, and the results were different. | Mr. Wang gave feedback on the results of troponin made by our instrument satropon in comparison with the data of the hospital's testing department, and the results were different. | Because of the contrast of their new installation, the contrast instrument is the laboratory Beckman, so the two instruments are not the same in value. But the negative positive match, next week AS will go over and the head nurse to communicate by jesson
No | Visit | Complaint | RIND | sample problem error | sample problem | checked the analyser and it was giving sample problem error for all samples so cleaned the gasket and needle after that cleaned the measuring chamber and given pump calibrationa and LS adjust and calibrated the analyser to ready mode observed for few samples | SWO-486072 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway | Clot | sample problem error | sample problem | checked the analyser and it was giving sample problem error for all samples so cleaned the gasket and needle after that cleaned the measuring chamber and given pump calibrationa and LS adjust and calibrated the analyser to ready mode observed for few samples
No | Visit | Complaint | RSCH | Blank | Dérive hors gamme PH-Na-K-Ca lors de la calibration (garantie) | Remplacé l’électrode de référence d’origine e.o (garantie) Calibration e.o contrôle de qualité e.o | SWO-485161 | Assisted Replacement | Electrode | Electrode | Reference | CAPA-733 | Blank | Out-of-range drift PH-Na-K-Ca when calibrating (guaranteed) | Replaced the original e.o reference electrode (guarantee) Calibration e.o quality control e.o
No | Visit | Complaint | RDEU | 16.01.2019, 11:40 Uhr, H. Glienke, Tel. 03334691503: Kal und QC Fehler bei Elyten | Statusfehler bei pH u.d. Lyten | Fehlersuche; neue Ref.-Elektrode (alt: 98-09) membraniert u. eingesetzt; Kal. + QC i.O. | SWO-485200 | Assisted Replacement | Electrode | Electrode | Reference | CAPA-733 | 16.01.2019, 11:40 Uhr, H. Glienke, Tel. 03334691503: Kal und QC Fehler bei Elyten | Status error for pH and Lyten | Troubleshooting; new ref.-electrode (old: 98-09) membrane-infused and inserted; Kal. + QC in the first place
No | Visit | Complaint | RDEU | 14.01.2019, Frau Von Franzski, Tel. 020172382700: Bildschirm bleibt dunkel | Bildschirm wird nach Bewegung schwarz . | Erneuerung des Bildschirmes. Dabei festgestellt, dass die Kontakte des alten Bildschirms zu kurz sind und es daher bei Bewegung zu einem Wackelkontakt gekommen ist. Funktionskontrolle: I.O. | SWO-484962 | Defect | HW | Cover and Chassis | Display | Loose cable connection | 14.01.2019, Mrs. Von Franzski, Tel. 020172382700: Screen remains dark | Screen turns black after movement . | Renewal of the screen. It was determined that the contacts of the old screen are too short and therefore a wobble contact has occurred during movement. Functional control: I.O.
No | Visit | Complaint | RAME | Liquid transport Error | Found analyzer in troubleshooting needed screen, requesting a flush or reinstall SC. Customer has already run a flush as well as replacing the SP. | Performed a flush on the analyzer and cleaned the inlet connector gasket. Ran fluid transport tests in service mode but could not correct the SP failed status. Found errors 1343, unable to pump solutions and 1086 in homogeneous QC2 solution. Removed wet section and flushed hemolyzer with hypochlorite solution and installed a pm kit. | SWO-485961 | Defect | HW | Basic Unit | PM Kit | 1343 | Liquid transport Error | Found analyzer in troubleshooting needed screen, requesting a flush or reinstall SC. Customer has already run a flush as well as replacing the SP. | Performed a flush on the analyzer and cleaned the inlet connector gasket. Ran fluid transport tests in service mode but could not correct the SP failed status. Found errors 1343, unable to pump solutions and 1086 in homogeneous QC2 solution. Removed wet section and flushed hemolyzer with hypochlorite solution and installed a pm kit.
No | Visit | Complaint | RDEU | 17.01.2019 Sr. Claudia 023143423501 meldet, dass die Oximetrie nicht gemessen wird. | Gerät überprüft und gereinigt, Fließwege manuel gespült, Lösungspack erneuert, | Die Küvette des Hemolyzers ist geplatzt. Ersatzteile bei anderem Techniker abgeholt. Hemolyzer eingebaut, Gerät zusammen gebaut, neu gestartet, Startphase abgewartet, tHb Kal, Kal und QCs durchgeführt | SWO-485929 | Defect | HW | Oxi module | Hemolyzer |  | 17.01.2019 Sr. Claudia 023143423501 reports that oximetry is not measured. | Device checked and cleaned, flow paths manuel rinsed, solution pack renewed, | The hemolyzer's cuvette has burst. spare parts picked up by other technicians. Hemolyzer built-in, device assembled, restarted, start phase awaited, tHb Kal, Kal and QCs carried out
No | Visit | Complaint | RAME | Multiple errors after changing solution pack including liquid transport error and pump cal error | Found errors, tried to rerun fluid transport programs in service mode but they would not pass. | Replaced solution pack with customers supply. Submitted a credit replacement ticket for customer. Glucose and Lactate will not pass calibrations after new solution pack was installed. Need a new sensor cassette. Customer does not have one available at this time. Will install when a new one comes in. | SWO-486031 | Assisted Replacement | SP | SP | FT |  | Multiple errors after changing solution pack including liquid transport error and pump cal error | Found errors, tried to rerun fluid transport programs in service mode but they would not pass. | Replaced solution pack with customers supply. Submitted a credit replacement ticket for customer. Glucose and Lactate will not pass calibrations after new solution pack was installed. Need a new sensor cassette. Customer does not have one available at this time. Will install when a new one comes in.
No | Visit | Complaint | RAME | Battery won't hold a charge at all. | Checked battery. Showing less than 3% charge after being plugged in. | Replaced battery under warranty. Checked status and it charged correctly. | SWO-485966 | Defect | HW | Battery module | Battery module |  | Battery won't hold a charge at all. | Checked battery. Showing less than 3% charge after being plugged in. | Replaced battery under warranty. Checked status and it charged correctly.
No | Visit | Complaint | RCHN | 接口板腐蚀 | 接口板腐蚀,借902-972一个。 | 更换902-972后修复 by Andrew | SWO-484983 | [5 empty] | Interface plate corrosion | Interface plate corrosion, borrow 902-972 one. | Repair by Andrew after replacement 902-972
No | Visit | Complaint | RCHN | 定标终止 | 定标终止，借944-258（15247）一个 | 重新定标后修复(by chuanwei) | SWO-484858 | Assisted Replacement | Consumables | Test Kit | Cartridge, cal | 5257 | Bid termination | Bid termination, loan 944-258 (15247) one | Re-calibration repair (bywei chuan)
No | Visit | Complaint | RCHN | 打印机不工作 | 打印机不工作，借910-286一个 | 重新插拔线路后修复 by Andrew | SWO-484778 | Adjustment | HW | Cover and Chassis | Printer Cable |  | The printer does not work | The printer does not work, borrow 910-286 one | Fix by Andrew after replugging the line
No | Visit | Complaint | RCHN | 处理盘电极不转，报错5784 | 处理盘电机不转，处理盘杯位弹片损坏，报错5784 借903-702 580-098一个 | 更换903-702 580-098后修复（by chuanwei） | SWO-484791 | Defect | HW | Process Module | Motor, process wheel | 5784 | Processing disk electrode does not turn, reported error 5784 | The processing disk motor does not turn, the handling cup bit shrapnel damage, the error 5784 borrowed 903-702 580-098 one | Repair after replacement 903-702 580-098 (by chuanwei)
No | Visit | Complaint | RFRA | Fuites visibles des liquides en dessous pack et/ou machine Souhaite son FSE pour venir constater le problème Demande transmise tel 0660521629 | Fuite du pack solution. | Fuite venant uniquement du pack solution. Fuite sur le devant du pack solution. Nettoyage ok. Vérification de la fluidique à l’arrière de l’ABL9 ok. Vérification des connecteurs pack solution $ ABL9 ok. Changement du pack solution ok. Mise en route ok. Calibration ok. Tous les paramètres disponibles. Rapport envoyé par mail à tisi@wanadoo.fr Demande de remboursement de conso défectueux faite ok. | SWO-485850 | [5 empty] | Visible leaks of liquids below pack and/or machine Wants his ESF to come and see the problem Request forwarded tel 0660521629 | Leaked solution pack. | Leak coming only from the solution pack. Leak on the front of the solution pack. Cleaning ok. Checking the fluid at the back of the ABL9 ok. Checking connectors pack solution $ ABL9 ok. Change the solution pack ok. Start-up ok. Calibration ok. All settings available. Report emailed to tisi@wanadoo.fr Defective conso refund request made ok.
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 例子结果全部不准 | 更换测试卡后修复 | SWO-485604 | [5 empty] | Test card, reagent package failure | Example results are all not allowed | Repair after replacing test card
No | Visit | Complaint | RCRO | Often unstable pH measurement and CO2 calibration errors. Previously replaced lower sealing ring of pH electrode, but problems persited. Changed pH, Ref and CO2 electrodes and still the problem persited. 18.1. changed measuring chamber. | Učestalo nestabilno mjerenje pH i učestali povećani drift kalibracijske krivulje za CO2. Zamjenom pH, Ref i CO2 elektrode problem i dalje ostaje. | Zamjenjena mjerna komora elektroda radi sumnje u moguću mikrofrakturu. | SWO-485912 | Defect | HW | pH/BG Module | measuring chamber |  | Often unstable pH measurement and CO2 calibration errors. Previously replaced lower sealing ring of pH electrode, but problems persited. Changed pH, Ref and CO2 electrodes and still the problem persited. 18.1. changed measuring chamber. | A frequent unstable pH measurement and a frequent increased drift calibration curve for CO2. By replacing the pH, Ref and CO2 electrodes the problem still remains. | Replaced by a measuring chamber of Electros for suspicion of a possible microfracture.
No | Visit | Complaint | RAME | Status code 2 | The ABL locked up while trying to check the IP and MAC address from the command prompt. Upon reboot status code 2. | Ran Service Dump. Captured setup files and archives. Reloaded software and settings. Opened CQC. | SWO-486058 | Defect | SW | RMED SW | RMED SW | Status Code 2 | Status code 2 | The ABL locked up while trying to check the IP and MAC address from the command prompt. Upon reboot status code 2. | Ran Service Dump. Captured setup files and archives. Reloaded software and settings. Opened CQC.
No | Visit | Complaint | RDEU | 18.01.2019 08:45Uhr Frau Erika Steigerwald 08944072453: wir haben ein Blutgasanalyse Gerät, was funktioniert jetzt gerade nicht pünktlich. Hat vielmal falsche Werte angegeben. Unsére Dienstarzt kann nicht vertrauen in die Maschine, er möchte wenn jemand jetzt komm und kalibrirt die. Ich könnte Sie telefonisch nicht erreichen,deswegen melde ich mich hier per Email. | 18.01.2019 08:45Uhr Frau Erika Steigerwald 08944072453: wir haben ein Blutgasanalyse Gerät, was funktioniert jetzt gerade nicht pünktlich. Hat vielmal falsche Werte angegeben. Unsére Dienstarzt kann nicht vertrauen in die Maschine, er möchte wenn jemand jetzt komm und kalibrirt die. Ich könnte Sie telefonisch nicht erreichen,deswegen melde ich mich hier per Email. | Bei Messungen Kalibrataionsfehler vorhanden, Kd. wurde darüber informiert., Kalibrationen und QC's vor genommen. | SWO-486264 | No fault found | No fault found | No fault found | No fault found | Calibation and QC | 18.01.2019 08:45Pm Mrs. Erika Steigerwald 08944072453: we have a blood gas analysis device, what is not working on time right now. Has indicated incorrect values many times. Unsére doctor can not trust in the machine, he wants when someone comes now and calibums the. I couldn't reach you by phone, so I'll get in touch here by email. | 18.01.2019 08:45Pm Mrs. Erika Steigerwald 08944072453: we have a blood gas analysis device, what is not working on time right now. Has indicated incorrect values many times. Unsére doctor can not trust in the machine, he wants when someone comes now and calibums the. I couldn't reach you by phone, so I'll get in touch here by email. | In measurements calibratation errors present, Kd. was informed, calibrations and QC's made.
No | Visit | Complaint | RDEU | Nullstromfehler bei Kalibrierung pO2 | Nullstromfehler | Elektrodenkanäle neu kalibriert. Sprung von 25 auf 31pA. Kalibrationen und Qualitätskontrollen in Ordnung. | SWO-486304 | Adjustment | HW | pH/BG Module | pH/BG Module | zero current | Zero current error during calibration pO2 | Nullstromfehler | electrode channels. Jump from 25 pEach to 31pA. calibrations and quality controls in order.
No | Visit | Complaint | RAME | pCO2 failing QCs on the new SC he installed yesterday. | Returned to site with sensor interface board | Replaced sensor interface board and re-tested analyzer. Found all calibrations passed. Ran internal and external QC, all levels passed | SWO-486044 | Defect | HW | Basic Unit | PCB, Sensor interface | pCO2 | pCO2 failing QCs on the new SC he installed yesterday. | Returned to site with sensor interface board | Replaced sensor interface board and re-tested analyzer. Found all calibrations passed. Ran internal and external QC, all levels passed
No | Visit | Complaint | RKOR | Solution pack install error | . | User Training. (DM) | SWO-486151 | User Error | SP | SP | SP | Incorrect SP installation | Solution pack install error | . | User Training. (DM)
No | Visit | Complaint | RCHN | 仓门传感器错误 | 仓门传感器错误，借903-028一个 | 更换903-028后修复(by chuanwei) | SWO-485605 | Defect | HW | Cover and Chassis | PCB, cabinet top | 4548 | Warehouse sensor error | Warehouse sensor error, borrow 903-028 one | Repair after replacement 903-028 (by chuanwei)
No | Visit | Complaint | RCHN | 做样本报错5662,5663 | 做样本报错5662,5663 | 清洁孵育盘传感器后修复(by chuanwei) | SWO-486119 | Cleaned | HW | Incubator Module | Sensor, incubator | 5662, 5663 | Do a sample of this newspaper error 5662,5663 | Do a sample of this newspaper error 5662,5663 | Repair after cleaning incubation disk sensor (by chuanwei)
No | Visit | Complaint | RCHN | 结果打问号，HCT值不能测量 | 结果打问号，HCT值不能测量，借903-329一个 | 更换903-329后修复(by chuanwei) | SWO-485612 | Defect | HW | Aspiration Module | Needle Wash Unit |  | Results a question mark, HCT value cannot be measured | Results a question mark, HCT value can not be measured, borrow 903-329 a | Repair after replacement 903-329 (by chuanwei)
No | Visit | Complaint | RCHN | HCT值无法测得 | HCT值无法测得 借903-329一个 |  | SWO-486118 | Defect | HW | Aspiration Module | Needle Wash Unit | 1223, 1257 | HCT value cannot be measured | HCT value can not be measured borrowing 903-329 one
No | Visit | Complaint | RCHN | 进样门无法关闭报错5784，干燥前温度超范围报错1068 | 进样门无法关闭报错5784，干燥前温度超范围报错1068 借903-210,886-638,886-639各一个 |  | SWO-486110 | Defect | HW | Process Module | PCB, cup dryer | 5784, 1068 | The sample door cannot be closed to report the error 5784, and the temperature above the drying temperature is over-ranged to report the error 1068 | The sample door cannot close the report error 5784, the pre-drying temperature overrange reported error 1068 borrowed 903-210, 886-638, 886-639 each
No | Visit | Complaint | RCHN | 打印机故障 | 打印机故障 借903-001,902-985各一个 | 更换903-001后修复（by chuanwei） | SWO-486106 | [5 empty] | Printer failure | Printer failure Borrowing 903-001, 902-985 each | Repair after replacement 903-001 (by chuanwei)
No | Visit | Complaint | RCHN | CO2定标失败 | CO2定标失败 | 清洁电极仓，重新定标后修复 by Andrew | SWO-486915 | Cleaned | HW | pH/BG Module | measuring chamber | pCO2 | CO2 calibration failure | CO2 calibration failure | Clean the electrode compartment and fix by Andrew after re-calibration
Yes | Visit | Complaint | RCHN | CO2质控偏高 | CO2质控偏高 | 更换945-612后修复 by Andrew | SWO-487320 | [5 empty] | High CO2 quality control | High CO2 quality control | Repair by Andrew after replacing 945-612
No | Visit | Complaint | RDEU | 21.01.2019; Fr. Trede; Tel.: 020348800160; Das SC Fach schließt nicht mehr | Platte zur Unterbrechung der Lichtschranke verbogen | Analyse Einheit demontiert, Lichtschranke am SC Door gereinigt, Platte zur Unterbrechung der Lichtschranke verbogen, Platte und Führungsblech (in der Lichtschranke innen liegend) gerichtet. Funktion mehrfach getestet und OK. SC installiert. | SWO-487813 | Adjustment | HW | Basic Unit | Sensor Cassette Door |  | 21.01.2019; Ms. Trede; Tel.: 020348800160; The SC compartment no longer closes | Plate bent to interrupt the light barrier | Analysis unit dismantled, light barrier cleaned at SC Door, plate bent to interrupt the light barrier, plate and guide plate (inside the light barrier lying inside) directed. Function tested multiple times and OK. SC installed.
No | Visit | Complaint | RAME | Customer says analyzer stuck in loop with leak detected errors. | Customer changed solution pack and completed flush. Inlet gasket changed, all with no success. Pressure test at 0.95. | Replaced inlet probe, pressure test now 0.47 | SWO-486869 | Defect | HW | Inlet module | Probe |  | Customer says analyzer stuck in loop with leak detected errors. | Customer changed solution pack and completed flush. Inlet gasket changed, all with no success. Pressure test at 0.95. | Replaced inlet probe, pressure test now 0.47
No | Visit | Complaint | RESP | Error de no puede bombear soluciones | Error de no puede bombear soluciones Paquete de soluciones mal instalado | Limpieza de tubos y de inlet probe Reinstalacion de sp Chequeo con programas de servicio Cal, qc | SWO-486715 | User Error | SP | SP | SP activation | 1247, 1087 | Failure to pump solutions | Failure to pump solutions Poorly installed solution package | Cleaning tubes and inlet probe Reinstalling sp Checking with service programs Cal, qc
No | Visit | Complaint | RESP | Erro de calibração de pO2 | Erro de calibração de pO2 | Verificação do registro de dados. Substituição da placa de sensores da cassete. Inicialização. Teste de pO2 ok. Inicialização. Bloqueio da porta de entrada da cassete. | SWO-487988 | Defect | HW | Basic Unit | PCB, Sensor interface | pO2, 1002 | PO2 Calibration Error | PO2 Calibration Error | Checking the data record. Replacing the cassette sensor plate. Startup. PO2 Test OK. Startup. Cassette input port lock.
No | Visit | Complaint | RFRA | Pb d’écran tactile | Contrôle des connexions | Changement de l’écran complet. Essais | SWO-486607 | Defect | HW | Display | Display |  | Touch screen pb | Connection control | Change the full screen. Tests
No | Visit | Complaint | RAME | Unable to pump solution | 1) Error 1258: Liquid transport failed. Elongated pump tubing on solution pack. 2) Bent inlet probe. | 1) Replaced Solution Pack and submitted Credit Claim form. 2) Replaced Inlet module (includes probe and gasket) 3) Collected Service Dump and Data Logs to be attached to HUP Customer Quality Case. 4) Connected analyzer to AQURE. | SWO-486719 | Assisted Replacement | SP | SP | Pump Tube | 1258 | Unable to pump solution | 1) Error 1258: Liquid transport failed. Elongated pump tubing on solution pack. 2) Bent inlet probe. | 1) Replaced Solution Pack and submitted Credit Claim form. 2) Replaced Inlet module (includes probe and gasket) 3) Collected Service Dump and Data Logs to be attached to HUP Customer Quality Case. 4) Connected analyzer to AQURE.
No | 3 | Y | RCZE | Blank |  |  | D59556B2403C100A | Defect | HW | Battery module | Battery module |  | Blank
No | Visit | Complaint | RDEU | Einlass schließt nicht | Einlassdichtung verkantet | Gerät geprüft, Inletmodul demontiert und justiert, Funktionskontrolle ok | SWO-487212 | Adjustment | HW | Auto Inlet module | Auto Inlet module |  | Entrance does not close | Einlassdichtung verkantet | Device tested, inlet module dismantled and adjusted, function control ok
No | Visit | Complaint | RAME | System is having an issues connecting to Aquire | Systems could not connect to the Aquire Server. | I reset the system to use DHCP and they connected fine. It will be working on getting the IP address resolved as of now they are connecting. | SWO-489754 | [5 empty] | System is having an issues connecting to Aquire | Systems could not connect to the Aquire Server. | I reset the system to use DHCP and they connected fine. It will be working on getting the IP address resolved as of now they are connecting.
No | Visit | Complaint | RKK | ガスの減りが早い |  |  | SWO-487632 | [5 empty] | The gas is running low quickly.
No | Visit | Complaint | RPAC | The analyser is giving alot of errors on the Glu, Lact and Na parameters, regarding drift after a calibration. | The analyser has been having various issues with the Glu and Lact parameters on calibrations. The Na parameter has also been drifting alot as well. | It was decided that the pH $BG Module and El $Met Module be replaced. The Right Fluidic Module was also removed and the Solution Pump motor was replaced. Whilst the module was out of the analyser, all the eight valves were removed, checked and cleaned. It was found on the membrane of the V4 Valve, a white material, this was removed. The analyser was returned to normal and the AU Software was installed on the two new modules. A Xero Cal, LS Adjust and Pump Cal were also performed, as per the service manual. The analyser was allowed to stabilize all its parameters via calbrations, however when performing the QC Level 1, it failed the Creat parameter. The Creatinine electrodes were rememnbraned and the Cal 1 & 2 solutions replaced. Again the analyser was allowed to recalibrate and once all the parameters had settled, QC Levels 1 & 3 were performed, all parameters are now within their limits. The analyser during this time, had performed many calibrations, during this time the Glu, Lact and Na parameters were quite stable. The analyser overall, appears to now be functioning fine. The analyser was checked the following day, all fine. Nine patient samples had been through the analyser with no issues. All parameters are now stable. | SWO-487606 | Defect | HW | El/Met Module | Valve |  | The analyser is giving alot of errors on the Glu, Lact and Na parameters, regarding drift after a calibration. | The analyser has been having various issues with the Glu and Lact parameters on calibrations. The Na parameter has also been drifting alot as well. | It was decided that the pH $BG Module and El $Met Module be replaced. The Right Fluidic Module was also removed and the Solution Pump motor was replaced. Whilst the module was out of the analyser, all the eight valves were removed, checked and cleaned. It was found on the membrane of the V4 Valve, a white material, this was removed. The analyser was returned to normal and the AU Software was installed on the two new modules. A Xero Cal, LS Adjust and Pump Cal were also performed, as per the service manual. The analyser was allowed to stabilize all its parameters via calbrations, however when performing the QC Level 1, it failed the Creat parameter. The Creatinine electrodes were rememnbraned and the Cal 1 & 2 solutions replaced. Again the analyser was allowed to recalibrate and once all the parameters had settled, QC Levels 1 & 3 were performed, all parameters are now within their limits. The analyser during this time, had performed many calibrations, during this time the Glu, Lact and Na parameters were quite stable. The analyser overall, appears to now be functioning fine. The analyser was checked the following day, all fine. Nine patient samples had been through the analyser with no issues. All parameters are now stable.
No | Visit | Complaint | RIND | Crea Calibration error. | Need to replace the Crea membrane. | Replaced the Crea membrane and run the instrument its working good. | SWO-487137 | Assisted Replacement | Electrode | Membrane | Crea | Cal | Crea Calibration error. | Need to replace the Crea membrane. | Replaced the Crea membrane and run the instrument its working good.
No | Visit | Complaint | RPAC | Whilst on site COhb QC had failed above range. Inspected data logs to find frequent electrolyte calibration drift. | See previous SWO, reference membrane was replaced 16 $1 R2387 due to membrane leak and calibration drift on electrolytes. 21 $1 Electrolytes still intermittently drifting (errors 0376 $0377) and still leak on top of membrane. | Cleaned leak from reference membrane and chamber. Cleaned pH and refitted sealing ring stuck in chamber. Installed new reference electrode 58-12. Removed reference electrode 08-06 (production 3 $2018). Clean, liquid sensor, calibration x 2 and QC. | SWO-486262 | Assisted Replacement | Electrode | Electrode | Reference | CAPA-733 | Whilst on site COhb QC had failed above range. Inspected data logs to find frequent electrolyte calibration drift. | See previous SWO, reference membrane was replaced 16 /1 R2387 due to membrane leak and calibration drift on electrolytes. 21 /1 Electrolytes still intermittently drifting (errors 0376 /0377) and still leak on top of membrane. | Cleaned leak from reference membrane and chamber. Cleaned pH and refitted sealing ring stuck in chamber. Installed new reference electrode 58-12. Removed reference electrode 08-06 (production 3 $2018). Clean, liquid sensor, calibration x 2 and QC.
No | Visit | Complaint | RDAN | Den interne stregkodelæser virker ikke som den skal. | Den interne stregkodelæser virker ikke som den skal. | Den interne stregkodelæser er skiftet af Ebbe. | SWO-486339 | Defect | HW | Basic Unit | Barcode Reader |  | The internal barcode reader does not work as it should. | The internal barcode reader does not work as it should. | The internal barcode reader is switched by the EBP.
No | Visit | Complaint | RESP | Error en storage wheel Tiene algo como atascado | Chequeo de storage wheel | Ajuste de separadores de quesitos de reactivo en storage wheel | SWO-487732 | Adjustment | HW | Storage Module | Cartridge guide | 5784 | Storage wheel error He's got something like stuck. | Storage wheel check | Adjusting reagent tile separators in storage wheel
No | Visit | Complaint | RCHN | 液路有气泡和气泡有空气 | 液路有气泡和气泡有空气 | 更换903-099后修复 by Andrew | SWO-487064 | [5 empty] | There are bubbles in the liquid road and air in the air. | There are bubbles in the liquid road and air in the air. | Repair by Andrew after replacement 903-099
No | Visit | Complaint | RDEU | Telefonat mit Schw. Claudia ==> Fehlerbehebung erforderlich | Softwareupgrade wurde am 21.01.3018 durchgeführt und nun gibt es ständig Messwertausreiser bei den QC-Messungen Nullpunktkalibrierung durchgeführt, Kalibrierungen, mehrmalige QC-Messungen | Leihgerät (400237) aufgestellt und Kundengerät (400363) zur Reparatur mitgenommen | SWO-487389 | [5 empty] | Phone call with Schw. Claudia ==> Troubleshooting required | Software upgrade was carried out on 21.01.3018 and now there are constantly measuring value departures in the QC measurements zero-point calibration performed, calibrations, multiple QC measurements | Rental device (400237) and customer device (400363) taken with him for repair
No | Visit | Complaint | RCHN | 多参数定标失败 | 多参数定标失败 | 自行更换903-612后修复 by Andrew | SWO-487673 | [5 empty] | Multi-parameter calibration failed | Multi-parameter calibration failed | Fix by Andrew after replacing 903-612
No | Visit | Complaint | RAME | Crea Failing QC | QC Level 1 and 2 are failing for Crea | Removed electrodes and cleaned measuring chamber and electrode contact. Replaced pump tubing from customer stock. Ran 2 Protein Removals, Cleaning, Liquid Sensor Adjust, Pump Cal, and calibrations. | SWO-486870 | Defect | HW | El/Met Module | Pump tube | Crea QC | Crea Failing QC | QC Level 1 and 2 are failing for Crea | Removed electrodes and cleaned measuring chamber and electrode contact. Replaced pump tubing from customer stock. Ran 2 Protein Removals, Cleaning, Liquid Sensor Adjust, Pump Cal, and calibrations.
No | Visit | Complaint | RDEU | 2.01.2019, Sr. Martina, tel. 0343218453: AQC Fehler bei pCO2, CL und Lac 21.01./09:15 Uhr, Pflg. Mathias, Tel.: 0343218453: Weiterhin Probelme mit zu hohem QC Verbrauch | Gerät zeigt Fehler bei Kal bei Unterschiedlichen Parametern, bei eintreffen pH Kal Fehler seit 21.01.19 23.33Uhr-> 23.27 Uhr wurde die pCO2 Membran erneuert | Gerät überprüft: Feder in Airventil des El $Met Moduls fehlt- Ventil ersetzt, Ventile im pH $BG Ventil Block geprüft i.O., Dichtring an pH Elektrode ersetzt pH Kal trotz mehrerer Kal nicht stabil und fehlerhaft, Reinigung & Nachjustierung der Elektrodenkontakte-> keine Verbesserung, während Kal bei Ansaugung der Kal Lösung bildet sich Luftblase an pH Elektrode-> Glasbruch an Elektrodenspitze -> Elektrode ersetzt -> mehre Kal nötig bis Kal werte im Bereich Kal & QC Daten gemäß RDEU Vorgaben gesichert und dokumentiert | SWO-487181 | Assisted Replacement | Electrode | Electrode | pH |  | 2.01.2019, Sr. Martina, tel. 0343218453: AQC error with pCO2, CL and Lac 21.01./09:15 o'clock, Pflg. Mathias, Tel.: 0343218453: Continue sample lme with too high QC consumption | Device shows errors with Kal with different parameters, on arrival pH Kal error since 21.01.19 23.33h-> 23.27 o'clock the pCO2 membrane has been renewed | Device checked: Spring in air valve of the El $Met module missing valve replaced- Valves in pH $BG valve block checked i.O., sealing ring replaced at pH electrode pH Kal despite several kal not stable and faulty, cleaning & readjustment of the electrode contacts-> no improvement, while Kal forms air bubble at pH electrode-> glass breakage at electrode tip -> electrode replaces -> more kale necessary to cal values in the range Kal & QC data backed up and documented in accordance with RDEU specifications
No | Visit | Complaint | RAME | Verify if analyzer is connected to Network / Radiance | New analyzer is not communicating with Radiance | Verified all the Radiance settings were correct on the analyzer and that the connection was set to dynamic just like the other analyzers at this site. Verified that the Radiance license was also in place for the extra analyzer connection. Noticed that the IP address being obtained was in the 169. Range as opposed to the 10.0 range etc. Spoke with Myles in IT dept and he is going to get the IP correct. The analyzer should connect at this point otherwise the customer can refresh the connection | SWO-486886 | Setup | SW | RMED SW | IP address | Radiance | Verify if analyzer is connected to Network / Radiance | New analyzer is not communicating with Radiance | Verified all the Radiance settings were correct on the analyzer and that the connection was set to dynamic just like the other analyzers at this site. Verified that the Radiance license was also in place for the extra analyzer connection. Noticed that the IP address being obtained was in the 169. Range as opposed to the 10.0 range etc. Spoke with Myles in IT dept and he is going to get the IP correct. The analyzer should connect at this point otherwise the customer can refresh the connection
No | Visit | Complaint | RDAN | Dorte klager over fejl 653. | Afprøvet og undersøgt. | Gas $luft ventilen i pH $blodgas modulet udskiftet, slutafprøvet. | SWO-487162 | Defect | HW | pH/BG Module | Valve, air | 653, 664 | Dorte complains about error 653. | Tested and investigated. | Gas $luft valve in pH $blodgas module replaced, final tested.
No | Visit | Complaint | RDEU | 22.01.2019 15:45 Herr Manow 079416928122: Die Uhrzeit verstell sich ständig | Uhrzeit verstellt sich um 1h. Inhomogen Kal. 3 Lösung, Kal. Oxifehler. | Zeiteinstellung Konfiguration korrigiert. Weiterhin zeigt sich Kal. Instabilität bei Oximetrie.-Inhomogene Kal. 3 Lösung. ABL90 gespült und Teile im Umfang einer Jahreswartung erneuert. Verbogene Einlasssonde aus Kundenbestand erneuert. LP aus VSC-SW erneuert. | SWO-488445 | Setup | SW | RMED SW | Time setting |  | 22.01.2019 15:45 Mr. Manow 079416928122: The time is constantly adjusting | Time is changed at 1h. Inhomogeneous Kal. 3 Solution, Kal. Oxi error. | Time setting configuration corrected. Kal also shows up. Instability in oximetry.-Inhomogeneous Cal. 3 solution. ABL90 rinsed and parts renewed in the amount of an annual maintenance. Bent inlet probe renewed from customer base. LP from VSC-SW.
No | Visit | Complaint | RAME | Not connecting to Meditech | Customer stated no information is coming out of serial port. Also was having issues with display. Received help Jerry Kelly with Roma.ini information to setup serial port. | Recalibrate display, also added semi colon in Roma.ini file. Then saved everything and did commit command. Then connected dock light to analyzer and verified transfer. Everything transferred over. Also verified QC and patient results crossed over on Meditech with Connie. | SWO-488336 | Setup | SW | RMED SW | Roma.ini | His/Lis | Not connecting to Meditech | Customer stated no information is coming out of serial port. Also was having issues with display. Received help Jerry Kelly with Roma.ini information to setup serial port. | Recalibrate display, also added semi colon in Roma.ini file. Then saved everything and did commit command. Then connected dock light to analyzer and verified transfer. Everything transferred over. Also verified QC and patient results crossed over on Meditech with Connie.
No | Visit | Complaint | RDAN | Eva Nielsen klager over, at batteri indikatoren hele tiden viser 0% selvom ABLen kan køre 1,5 timer på batteriet. | . | Batteri udskiftet, slutafprøvet. | SWO-488104 | Defect | HW | Battery module | Battery module | 0% | Eva Nielsen complains that the battery indicator continuously shows 0% although the ABLen can run 1.5 hours on the battery. | . | Battery replaced, final tested.
No | Visit | Complaint | RSCH | Blank | Oxi-LS Error | Oxi-Liquit Sensor ersetzt. SC neu installiert. System und QA Zyklus i.o. | SWO-488033 | [5 empty] | Blank | Oxygen-LS Error | Oxi-Liquit sensor replaced. SC reinstalled. System and QA cycle i.o.
No | Visit | Complaint | RCHN | 流量选择器错误，报错1177 | 流量选择器错误，报错1177 （未更换备件） | 清洁流量选择器后修复 by Andrew | SWO-487702 | Cleaned | HW | Basic Unit | Flow selector detector | 1177 | Traffic selector error, error 1177 | Traffic selector error, error 1177 (Spare parts not replaced) | Fix by Andrew after cleaning flow selector
No | Visit | Complaint | RCHN | 扫描器无法扫描和板子断裂 | 扫描器无法扫描和板子断裂,需要借件910-304和903-209各一个 | 更换910-304和903-209后修复 by Andrew | SWO-487057 | Defect | HW | Storage Module | Barcode reader, storage wheel |  | The scanner cannot scan and the board breaks | The scanner cannot be scanned and the board is broken, and one each of the 910-304 and 903-209 loans is required | Repair by Andrew after replacing 910-304 and 903-209
No | Visit | Complaint | RCHN | TNI定标失败 | 定标失败，借903-364,903-236（被903-702替代）各一个 | 更换903-364，903-702重新定标后修复 by Andrew | SWO-486334 | Defect | HW | Process Module | Optical Unit | 1024 | TNI calibration failed | Bid failure, borrow 903-364, 903-236 (replaced by 903-702) each | Replace 903-364,903-702 After re-calibration repair by Andrew
No | Visit | Complaint | RCHN | 干燥前温度超范围 | 干燥前温度超范围，借903-118一个 | 更换903-118后修复 by Andrew | SWO-487061 | Defect | HW | Process Module | Dryer Unit | 1068 | Overrange temperature before drying | Over-range temperature before drying, borrow 903-118 | Fix by Andrew after replacing 903-118
No | Visit | Complaint | RDEU | 22.01.2019, Frau zensus, Tel: 022678897048: Patientenwerte bei pO2 zu niedrig laut den Ärzten. | Patientenwerte bei pO2 zu niedrig. | Siehe Halbjahreswartungs SWO. Oxi Complete Test. Ventile getestet. Gasflow getestet. Ok. Nichts auffälliges gefunden. Pumpen und LS Kal. Kal & QC | SWO-488491 | No fault found | No fault found | No fault found | No fault found |  | 22.01.2019, Ms zensus, Tel: 022678897048: Patient values at pO2 too low according to the doctors. | patient values at pO2 are too low. | See Half-Year Maintenance SWO. Oxi Complete Test. Valves tested. gas flow tested. Ok. Nothing conspicuous lyfound. Pumps and LS Kal. Kal & QC
No | Visit | Complaint | RCHN | 电解质数值异常 | 电解质数值异常 | 更换参比电极膜后修复 | SWO-488638 | Assisted Replacement | Electrode | Membrane | Reference |  | Electrolyte value anomaly | Electrolyte value anomaly | Repair after replacing the reference electrode membrane
No | Visit | Complaint | RDEU | 22.01.2019 17:30 Fr. Schlüter 06213835970: Probe bei der Ansaugung vorzeitig erkannt | 0500 Spüllösung von pH $BG Modul erkannt... 0531 pH $BG Upper LS Error 0653 Probe bei Ansaugung vorzeitig am pH $BG LS erkannt | Allgemeine Überprüfung - Messventile pH $BG Modul demontiert und gereinigt - Erneuerung Ventilmembran Luft im pH $BG Modul - Rückmontage und Funktionsüberprüfung etc. | SWO-487835 | Defect | HW | pH/BG Module | Valve, membrane | 653, 664 | 22.01.2019 17:30 Fr. Schlüter 06213835970: Sample detected prematurely during suction | 0500 Rinsing solution detected by pH /BG module... 0531 pH /BG Upper LS Error 0653 Sample detected prematurely at pH /BG LS in case of suction | General inspection - measuring valves pH $BG module disassembled and cleaned - Renewal valve membrane air in pH $BG module - reassembly and function verification etc.
No | Visit | Complaint | RDEU | Remote Zugriff wird im AQURE nicht korrekt dargestellt. | Remote Zugriff wird im AQURE nicht korrekt dargestellt. Farbdarstellung im AQURE nicht korrekt und zum Teil fehlerhafte Bedienung , Button reagieren zum Teil nicht korrekt. | ABL neu gestartet. NetOP deaktiviert und neu gestartet. Kein Erfolg. Info an SalVa und SasBe weitergeleitet.Netzwerkkabel Verbindung geprüft , o.k. | SWO-488446 | Revisit required | SW | RMED SW | NetOP |  | Remote access is not displayed correctly in AQURE. | Remote access is not displayed correctly in AQURE. Color display in the AQURE incorrectly and partly incorrect operation, button react partly incorrect. | ABL. NetOP disabled and restarted. No success. Info forwarded to SalVa and SasBe.Network cable connection checked, ok
No | Visit | Complaint | RIND | Service Visit. | Crea parameter is deactivate due to expired demo key. | Need to order and install the permanent key into the analyzer, do the protein removal and cleaning of the analyzer, discuss about the analyzer status with the pathologist. | SWO-488674 | Revisit required | HW | Demo key | Demo key |  | Service Visit. | Crea parameter is deactivate due to expired demo key. | Need to order and install the permanent key into the analyzer, do the protein removal and cleaning of the analyzer, discuss about the analyzer status with the pathologist.
No | Visit | Complaint | RDEU | 12.12.2018 Frau Freund 093727001652 meldet, einen pCO2 Fehler QC L3. 13.12.2018 Frau Freund 093727001652: pCO2-Fehler im QCLevel 3 | Wartungsmassnahme am Analysator, Lieferung der Elektrode. | Tausch der pCO2 Elektrode und Ausbau der gebrauchten Elektrode. Analysator einkalibriert. | SWO-488160 | Assisted Replacement | Electrode | Electrode | pCO2 | QC 3 | 12.12.2018 Ms Friend 093727001652 reports a pCO2 error QC L3. 12/13/2018 Mrs Friend 093727001652: pCO2 error in QCLevel 3 | Maintenance measure on the analyzer, delivery of the electrode. | Replacement of the pCO2 electrode and removal of the used electrode. analyzer calibrated.
No | Visit | Complaint | RAME | Printer is not feeding paper properly | The printer was no printing and jamming up the paper. | Replaced printer head. | SWO-489549 | Defect | HW | Printer Module | Printer Head | wringled paper | Printer is not feeding paper properly | The printer was no printing and jamming up the paper. | Replaced printer head.
No | Visit | Complaint | RLTD | Analyser dead, no display or lights. Phoned in direct to Ray Flint by Cecilia 07479 241036 | Fault found to be with the power supply, 12v rail has failed. Replaced faulty power supply, also fitted battery unit serial no. P0169N026. Restarted analyser, tested battery unit, working correctly. Verified operation after analyser restart was completed, all ok. | Replaced power supply. | SWO-488952 | Defect | HW | Power Supply | Power supply |  | Analyser dead, no display or lights. Phoned in direct to Ray Flint by Cecilia 07479 241036 | Fault found to be with the power supply, 12v rail has failed. Replaced faulty power supply, also fitted battery unit serial no. P0169N026. Restarted analyser, tested battery unit, working correctly. Verified operation after analyser restart was completed, all ok. | Replaced power supply.
No | Visit | Complaint | RFRA | Ecran noir | Contrôle de pc | Changement du pc complet. Reconfiguration de la machine. Reconnexion lecteur réseau. Essais | SWO-488762 | Defect | HW | DMS_PC | CPU |  | Black screen | Pc control | Change of full pc. Reconfiguration of the machine. Network drive reconnection. Tests
No | Visit | Complaint | RLTD | Kelly Mayes 01512525126 | Site visit to investigate reason for analyser unable to pump solutions. | Arrived on site and attended analyser. Analyser all in green and working fine. Entered service and ran through fluid transport programs. All fluid transport programs passing and all values well within tolerances and performing well. Inspected inside sensor bay and solution pack bay for issues with leaks or solution pump problems. Inspected inlet gasket and probe. These were both in good condition as they were changed yesterday. Ran clot detect $remove a number of times to ensure no blockages left in fluid path. Once completed, I exited to main screen, ran cal and qc to ensure analyser in spec and working correctly. These passed and analyser working well. | SWO-488710 | No fault found | No fault found | No fault found | FT |  | Kelly Mayes 01512525126 | Site visit to investigate reason for analyser unable to pump solutions. | Arrived on site and attended analyser. Analyser all in green and working fine. Entered service and ran through fluid transport programs. All fluid transport programs passing and all values well within tolerances and performing well. Inspected inside sensor bay and solution pack bay for issues with leaks or solution pump problems. Inspected inlet gasket and probe. These were both in good condition as they were changed yesterday. Ran clot detect $remove a number of times to ensure no blockages left in fluid path. Once completed, I exited to main screen, ran cal and qc to ensure analyser in spec and working correctly. These passed and analyser working well.
No | Visit | Complaint | RAME | Unable to pump solution and leak detected | Analyzer instructed customer to replace inlet gasket with holder. | Replaced inlet gasket with holder, 903-585, and tested. | SWO-488904 | Defect | HW | Auto Inlet module | Inlet gasket |  | Unable to pump solution and leak detected | Analyzer instructed customer to replace inlet gasket with holder. | Replaced inlet gasket with holder, 903-585, and tested.
Yes | Visit | Complaint | RDEU | Einrichtung der Anbindung an OSM laut Auftrag | Einrichtung der Anbindung an OSM laut Auftrag | Manuelle Geräte IP eingerichtet, | SWO-489011 | Re classified RMED | planned SWO |  |  |  | Setting up the connection to OSM according to the order | Setting up the connection to OSM according to the order | Manual devices IP set up,
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | LAC定标失败 | 更换测试卡 | SWO-489764 | [5 empty] | Test card, reagent package failure | LAC calibration failure | Replace the test card
No | Visit | Complaint | RAME | No active parameters | The analyzer was back up and running when I arrived onsite. Here is the customers steps in troubleshooting: We are in the midst of troubleshooting the ABL80 today. Here is what I have so far: 1. At 0646 this morning, we ran a capillary blood gas sample. It gave an error of possible air in flush solution. We did not report the results. 2. At ~0730, all parameters went red and were disabled. We called the hotline. They asked us to do a manual system cycle. It ran for a while and gave a blockage during system cycle error. 3. We called the hotline back. They told us to remove the sensor cassette, click the start box, shutdown the software and restart the instrument. Then they had us go to the installation screen and install the old cassette. We questioned whether we should put a new one on and were told to put the same one back on(expires on Feb 11). 4. During sensor initialization, it gave a blockage error. We called the hotline again. They told us to replace the sensor cassette. When we went into the installation screen to replace the cassette, a blockage error popped up and would not let us go further. 5. Since this happened before to me(end of Dec), I instructed the tech to click the start box and shut down the software. The instrument exited the software, but the screen still kept showing the picture of the instrument. The screen never went black. I held in on the power switch until it shut off. 6. We are currently powering back on. The next step will be to install the sensor cassette. If that works ok, we will change the Solution pack, also, because it is set to expire on 1 $26. I don’t think it is a good idea to have reagent changes on the weekend given the recent issues. 7. I will update everyone with what happens next. | The customer did the following also before I arrived: The analyzer is up and QC is within range. Here is the rest of the troubleshooting: 1. The new sensor cassette was installed. It would not pass initialization. It gave an error to investigate sources of air. 2. The tech called the hotline again, and they dispatched service. We have not heard from Trevor yet. 3. The tech shut down the instrument again and tried to install the sensor pack again. The initialization failed. 4. I installed a new solution pack. Then it told me that the sensor pack was not initialized. 5. I went through installation procedure for the same sensor pack again. 6. It passed initialization and was giving me an error saying that QC needed to be run. 7. I ran all 3 levels of QC. They were in. I feel like there may be a stability issue with the solution packs. I can’t remember 100%, but I think when this happened at the end of Dec that we might have been getting close to a solution pack expiration. Today, I know that we are 2 days from the solution pack onboard expiring. Everything started working after I changed the solution pack. However, we had also put on a new sensor cassette by this point. | SWO-488979 | [5 empty] | No active parameters | The analyzer was back up and running when I arrived onsite. Here is the customers steps in troubleshooting: We are in the midst of troubleshooting the ABL80 today. Here is what I have so far: 1. At 0646 this morning, we ran a capillary blood gas sample. It gave an error of possible air in flush solution. We did not report the results. 2. At ~0730, all parameters went red and were disabled. We called the hotline. They asked us to do a manual system cycle. It ran for a while and gave a blockage during system cycle error. 3. We called the hotline back. They told us to remove the sensor cassette, click the start box, shutdown the software and restart the instrument. Then they had us go to the installation screen and install the old cassette. We questioned whether we should put a new one on and were told to put the same one back on(expires on Feb 11). 4. During sensor initialization, it gave a blockage error. We called the hotline again. They told us to replace the sensor cassette. When we went into the installation screen to replace the cassette, a blockage error popped up and would not let us go further. 5. Since this happened before to me(end of Dec), I instructed the tech to click the start box and shut down the software. The instrument exited the software, but the screen still kept showing the picture of the instrument. The screen never went black. I held in on the power switch until it shut off. 6. We are currently powering back on. The next step will be to install the sensor cassette. If that works ok, we will change the Solution pack, also, because it is set to expire on 1 /26. I don’t think it is a good idea to have reagent changes on the weekend given the recent issues. 7. I will update everyone with what happens next. | The customer did the following also before I arrived: The analyzer is up and QC is within range. Here is the rest of the troubleshooting: 1. The new sensor cassette was installed. It would not pass initialization. It gave an error to investigate sources of air. 2. The tech called the hotline again, and they dispatched service. We have not heard from Trevor yet. 3. The tech shut down the instrument again and tried to install the sensor pack again. The initialization failed. 4. I installed a new solution pack. Then it told me that the sensor pack was not initialized. 5. I went through installation procedure for the same sensor pack again. 6. It passed initialization and was giving me an error saying that QC needed to be run. 7. I ran all 3 levels of QC. They were in. I feel like there may be a stability issue with the solution packs. I can’t remember 100%, but I think when this happened at the end of Dec that we might have been getting close to a solution pack expiration. Today, I know that we are 2 days from the solution pack onboard expiring. Everything started working after I changed the solution pack. However, we had also put on a new sensor cassette by this point.
No | Visit | Complaint | RCHN | 识别不到电极 | 识别不到电极，需要借件902-778 | 经销商提供备用仪器修复 | SWO-488621 | [5 empty] | Electrodes not recognized | Electrodes are not recognized and a loan 902-778 is required | Dealers provide spare instrument repair
No | Visit | Complaint | RCHN | 报错1068 | 报错1068，借903-210一个 | 更换903-210后修复 by Andrew | SWO-487705 | Defect | HW | Process Module | PCB, cup dryer | 1068 | 1068 error | 1068, borrow edith 903-210 | Fix by Andrew after replacing 903-210
No | 3 | Y | RCZE | Blank |  |  | D59556674045404E | Defect | HW | Oxi module | Lamp unit | Neon | Blank
No | Visit | Complaint | RIND | Service Visit. | Need to install the option key in to the analyzer. | Go in to the service mode, in option, press the new type and read it from disk, press the update type and confirm update, successfully done, after initializing wait for auto two point calibration, successfully done, analyzer handed over to staff in well working condition. | SWO-489777 | User Error, FSE | HW | Demo key | Demo key |  | Service Visit. | Need to install the option key in to the analyzer. | Go in to the service mode, in option, press the new type and read it from disk, press the update type and confirm update, successfully done, after initializing wait for auto two point calibration, successfully done, analyzer handed over to staff in well working condition.
No | Visit | Complaint | RIND | pO2 Calibration error.pO2 sensitivity error. | Found pO2 sensitivity very low and after replacing the membrane also it no increased. Found problem with pO2 electrode. | Replaced the pO2 electrode and calibrated found all results within the range. Performed the tri level QC and samples all results found within the range. | SWO-489932 | Assisted Replacement | Electrode | Electrode | pO2 | Sens | pO2 Calibration error.pO2 sensitivity error. | Found pO2 sensitivity very low and after replacing the membrane also it no increased. Found problem with pO2 electrode. | Replaced the pO2 electrode and calibrated found all results within the range. Performed the tri level QC and samples all results found within the range.
No | Visit | Complaint | RIND | Crea Electrodes failed to calibrate.Sensitivity is out of range . | Not passing Crea Electrode Calibration . Crea Sensitivity is out of range . | Replaced the electrode membranes .Still showing the error. Has to replace the electrodes . 945-662 : E8088 Crea A electrode 945-663:E8089 Crea B electrode | SWO-488972 | Assisted Replacement | Electrode | Electrode | Crea | Sens, Crea A + B | Crea Electrodes failed to calibrate. Sensitivity is out of range . | Not passing Crea Electrode Calibration . Crea Sensitivity is out of range . | Replaced the electrode membranes . Still showing the error. Has to replace the electrodes . 945-662 : E8088 Crea A electrode 945-663:E8089 Crea B electrode
No | Visit | Complaint | RAME | Customer is getting oxi mismatch error and the error 893 when they run the samples by using 125 micro lit capillary | Analyzer to be downgraded to ABL 835 From ABL 837. | The Met II module replaced with blind module. Analyzer type updated to 835. The solutions replaced. Analyzer fully tested and passed. | SWO-488890 | Defect | HW | Oxi module | Oxi module | 893 | Customer is getting oxi mismatch error and the error 893 when they run the samples by using 125 micro lit capillary | Analyzer to be downgraded to ABL 835 From ABL 837. | The Met II module replaced with blind module. Analyzer type updated to 835. The solutions replaced. Analyzer fully tested and passed.
No | Visit | Complaint | RKK | オキシハードウエアエラー（Neon number too low/ Neon Lamp intensity too low) | オキシハードウエアエラー | 代替モジュール入れ替え | SWO-489829 | Defect | HW | Oxi module | Oxi module | Neon number too low | Oxy hardware error @Neon number too low/Neon Lamp intensity too low | Oxy hardware error | Replacement of alternate modules
No | Visit | Complaint | RDEU | 24.01.2019; Fr. Lang; Tel.: 068411622792; Na Kal Fehler | Na und Hb driften sehr häufig | Ref. Elektrodemembraniert. Air- Ventil der beiden Modulen ausgebaut und gereinigt. Scharze Ablagerungen der membraneventil entfernt. Nachfüllung und LS Justage. Kal und QC in Ordnung. | SWO-489223 | Cleaned | HW | El/Met Module | Valve | Black deposits | 24.01.2019; Ms. Lang; Tel.: 068411622792; Na Kal Error | Na and Hb drift very often | Ref. electrode membranes. Air valve of the two modules removed and cleaned. Sharp deposits of the membrane valve removed. Refill and LS adjustment. Kal and QC all right.
No | Visit | Complaint | RSCH | Blank | Mischer Fehler (blockiert sporadisch) | Mischer im Serviceprogramm überprüft und getestet, alles i.o. Beim Testen blockiert der Mischer erneut nach 3 $4 Versuchen. Leihmodul installiert und SW 6.18 neu geladen (Modul war 6.16 installiert). Original Modul zur Überprüfung und Reparatur mit genommen und wird bei nächstem Service wieder eingebaut. System und Funktionskontrolle i.o. | SWO-489246 | Defect | HW | FlexQ Module | Mixer Tray |  | Blank | Mixer error (blocks sporadically) | Mixer checked and tested in the service program, all in the following detail. During testing, the mixer blocks again after 3 x 4 attempts. Rental module installed and SW 6.18 reloaded (module was installed 6.16). Original module for inspection and repair and will be reinstalled at the next service. System and function control i.o.
No | Visit | Complaint | RCRO | Pump error | Greška kalibracije pumpe i QC. Sumnja na o-ring pH elektrode | Zamjena o-ring pH elektrode Čišćenje mjernih komora Preventivna zamjena svih membrana | SWO-489421 | Defect | Electrode | Electrode | Sealing ring, pH |  | Pump error | Pump calibration error and QC. Suspicion of O-ring pH Electrodes | Replacing the O-ring pH electrode Cleaning of the measuring chambers Preventive replacement of all membranes
No | Visit | Complaint | RKK | ｽﾄﾚｰｼﾞPCB交換後、インキュベータ温度範囲外が消えなくなった。 | インキュベータPCB、FTCケーブル6芯 | インキュベータPCB、FTCケーブル6芯交換 | SWO-488437 | Defect | HW | Incubator Module | PCB, incubator |  | After replacing the storage PCB, the outside of the incubator temperature range disappeared. | Incubator PCB, 6-core FTC cable | Incubator PCB, FTC cable 6-core replacement
No | Visit | Complaint | RAME | PC02 high on all levels and failing level c | pc02 failing level c and high on other levels new pack and new sensor different lots put on no change compared analyzer trend to other units in the hospital on average it is running 1 - 3 points higher on qc and level c is in and out every 3- 4 days replaced sensor interface pcb 903-581 adjusted pins replaced solution pack with known working pack no change replaced sensor with known working sensor on floor 4 analyzer and put consumables from floor 2 on floor 4 high co2 on qc followed the cassette. | replaced sensor interface pcb and sensor cassette | SWO-489817 | Assisted Replacement | SC | SC | pCO2 | high QC | PC02 high on all levels and failing level c | pc02 failing level c and high on other levels new pack and new sensor different lots put on no change compared analyzer trend to other units in the hospital on average it is running 1 - 3 points higher on qc and level c is in and out every 3- 4 days replaced sensor interface pcb 903-581 adjusted pins replaced solution pack with known working pack no change replaced sensor with known working sensor on floor 4 analyzer and put consumables from floor 2 on floor 4 high co2 on qc followed the cassette. | replaced sensor interface pcb and sensor cassette
No | Visit | Complaint | RDEU | Sporadische Kal und QC Fehler in der Oximetrie | Sporadische Kal3 Lösung inhomogen | Gerät überprüft, Einlaßdichtung ersetzt | SWO-489375 | Defect | HW | Inlet module | Inlet gasket | 1539 | Sporadic Kal and QC errors in oximetry | Sporadic Kal3 solution inhomogeneous | Device checked, inlet seal replaced
No | Visit | Complaint | RDEU | Fehler Code 2 | Analysator startet nicht. | Software clean Installation fehlgeschlagen. Gerät bittet nicht mehr. Windows Neuinstallation fehlgeschlagen. Möglicherweise Cf Karte defekt. | SWO-490071 | Revisit required | Analyzer removed | Analyzer removed | Analyzer removed | Status code 2 | Error Code 2 | Analyzer does not start. | Software clean installation failed. Device no longer asks. Windows reinstallation failed. Possibly Cf card defective.
Yes | Visit | Complaint | RDEU | Fehler Code 2 | Deinstallation | Gerät dekontaminiert und abgebaut. | SWO-490072 | Re classified RMED | same day upSWO |  |  |  | Error Code 2 | Uninstall | decontaminated and degraded.
No | Visit | Complaint | RPAC | The dept showed disappointment with being not informed of the Expiry Date requirements for the ABL90. TheY replaced the Inlet Gasket when not needed, has been determined that there appears to be a software bug. | Cassette and Gasket issues. | L made a visit to help with some enquiries the Sleep Dept had regarding the Expiry Date system with the ABL90 and replacement dates on the ABL90. The customer indicated that they were not made aware of the expiry date requirements of the ABL90, i provided the relevant information. The Inlet Gasket they replaced much earlier than needed, this was found to be an issue with the software on the ABL90, will inform Radiometer Denmark of this issue. To help, a new cassette and Inlet Gasket was organised to be sent to them. The customer will be using these in the future. | SWO-489165 | Defect | SW | RMED SW | RMED SW | Replacement Schedule | The dept showed disappointment with being not informed of the Expiry Date requirements for the ABL90. TheY replaced the Inlet Gasket when not needed, has been determined that there appears to be a software bug. | Cassette and Gasket issues. | L made a visit to help with some enquiries the Sleep Dept had regarding the Expiry Date system with the ABL90 and replacement dates on the ABL90. The customer indicated that they were not made aware of the expiry date requirements of the ABL90, i provided the relevant information. The Inlet Gasket they replaced much earlier than needed, this was found to be an issue with the software on the ABL90, will inform Radiometer Denmark of this issue. To help, a new cassette and Inlet Gasket was organised to be sent to them. The customer will be using these in the future.
No | Visit | Complaint | RPAC | Andrew had installed new sensor 13:15PM which took a long time thermistating but still failing 2 hours later with glucose status 4497 and sens failing, Ca sens 117, repeat calibrations not improving. He reinstalled same sensor, glucose and Ca now passing but OXI hardware error 1034, 1011, 1012, and inhomogenous errors. Automatic flush and repeat calibration no improvement. OXI test gave errors on LED peak intensity = 61776 (range 80-300) and cavitation absorb = -113 (5-30). Installed new sensor cassette, LED now ok but cavitation absorb = 0. Also found 1539 inhomogeneous cal 3 errors in activity logs during January. Leo to replace hemolyser 25/01. | Hemolyzer failed | Saved setup & data logs. Removed sernsor cassette, solution pack and switch off analyzer. Replaced hemolyzer. Restarted analyzer and reinstalled sensor cassette and solution pack. Performed parameter checks and calibrations, tHb calibration, QCs. All OK. Andrew and Ash carried out blood comparison OK. New solution pack FOC provided by Suzie for troubleshooting. | SWO-489772 | Defect | HW | Oxi module | Hemolyzer | 1034, 1011, 1012 | Andrew had installed new sensor 13:15PM which took a long time thermistating but still failing 2 hours later with glucose status 4497 and sens failing, Ca sens 117, repeat calibrations not improving. He reinstalled same sensor, glucose and Ca now passing but OXI hardware error 1034, 1011, 1012, and inhomogenous errors. Automatic flush and repeat calibration no improvement. OXI test gave errors on LED peak intensity = 61776 (range 80-300) and cavitation absorb = -113 (5-30). Installed new sensor cassette, LED now ok but cavitation absorb = 0. Also found 1539 inhomogeneous cal 3 errors in activity logs during January. Leo to replace hemolyser 25/01. | Hemolyzer failed | Saved setup & data logs. Removed sernsor cassette, solution pack and switch off analyzer. Replaced hemolyzer. Restarted analyzer and reinstalled sensor cassette and solution pack. Performed parameter checks and calibrations, tHb calibration, QCs. All OK. Andrew and Ash carried out blood comparison OK. New solution pack FOC provided by Suzie for troubleshooting.
No | Visit | Complaint | RCHN | 打印报告导致分析仪重启 | 打印报告导致分析仪重启，借902-985,903-381,933-306（被933-441替代）各一个 | 更换902-985,903-381,933-441各一个正常。 by Andrew | SWO-488397 | [5 empty] | Printed report causes analyzer restart | The printed report causes the analyzer to restart, with one each of the 902-985,903-381,933-306 (replaced by 933-441) | Replace 902-985, 903-381, 933-441 each normal. by Andrew
No | Visit | Complaint | RCHN | 连续更换2张测试卡都报LAC定标失败 | 连续更换2张测试卡都报LAC定标失败，借903-581一个 | 更换903-581后修复 by Andrew | SWO-487826 | Defect | HW | Basic Unit | PCB, Sensor interface | Lac, 1310 | 2 consecutive replacement test cards report LAC calibration failure | 2 consecutive replacement test cards are reported LAC calibration failure, borrow 903-581 one | Repair by Andrew after replacement 903-581
No | Visit | Complaint | RCHN | 质控B不过 0090 | 质控B不过 0090,借902-980一个。 | 更换902-980，重新总血红蛋白定标后修复 by jesson | SWO-489788 | Defect | HW | Oxi module | Hemolyzer | 90, QC B | Quality control B but 0090 | The QC B is no more than 0090, borrowing 902-980 one. | Replace 902-980, repair by jesson after re-total hemoglobin calibration
No | Visit | Complaint | RCHN | 报错1068 | 报错1068,需要借件944-268（15367） | 重新定标后修复 by Andrew | SWO-488604 | Assisted Replacement | Consumables | Test Kit | Cartridge, cal | 1068, Potentiel User Error | 1068 error | 1068, need to borrow 944-268 (15367) | Fix by Andrew after re-calibration
No | Visit | Complaint | RCHN | 定标终止 | 定标终止，借944-212（15480）一个 | 清洁孵育盘，处理盘。重新定标正常。by Andrew | SWO-488523 | Cleaned | HW | Incubator Module | Incubator Wheel |  | Bid termination | Bid termination, loan 944-212 (15480) one | Clean the incubation tray and process the disc. Recalibration is normal. by Andrew
No | Visit | Complaint | RCHN | 条码扫描器无法扫描 | 条码扫描器无法扫描，需要借件902-924一个 | 更换902-924一个正常。by Andrew | SWO-488489 | Defect | HW | Cover and Chassis | Barcode reader, display |  | Barcode scanner cannot be scanned | Barcode scanner cannot be scanned and requires a loan of 902-924 | Replace the 902-924 one normal. by Andrew
No | Visit | Complaint | RCHN | 干燥前温度超范围，报错1068 | 干燥前温度超范围，报错1068 借903-210一个 | 更换903-210后修复 by Andrew | SWO-487668 | Defect | HW | Process Module | PCB, cup dryer | 1068 | Temperature over range before drying, reported error 1068 | The temperature exceeded the range before drying, and the error 1068 borrowed 903-210 one | Fix by Andrew after replacing 903-210
No | Visit | Complaint | RDEU | Remote Zugriff wird im AQURE nicht korrekt dargestellt. | Darstellung im $ über Remotezugriff nicht korrekt. | Datensicherung durchgeführt. OptionKey kontrolliert, o.k. Auf Win Ebene NetOP Hoste Programm deinstalliert. ABL heruntergefahren ausgeschaltet und neu gestartet. NotOP installierte sich selbstständig , Verbindungen werden erkannt NetOP gestartet. Frau ist Dill nicht erreichbar um zu kontrollieren ob das Problem behoben ist. Frau Dill vom Qualitätsmanagement wird am Mo. benachrichtigt. | SWO-490048 | Revisit required | SW | RMED SW | NetOP |  | Remote access is not displayed correctly in AQURE. | Display in the remote access is incorrect. | data backup. OptionKey controlled, ok. On Win level NetOP Hoste program uninstalled. ABL shuts down and restarts. NotOP installed itself, connections are detected NetOP started. Woman is not reachable to check if the problem is fixed. Ms. Dill from Quality Management will be at Mo. Notified.
No | Visit | Complaint | RIND | printer not working properly. | Check and found paper is not coming out or roller is not moving. Printer paper physically pulled by the user. | Open the printer found broken due to force dragging of paper. Gears of printer found broken and need to be replace with new one for working of printer. | SWO-489175 | Defect | HW | Printer Module | Printer Module |  | printer not working properly. | Check and found paper is not coming out or roller is not moving. Printer paper physically pulled by the user. | Open the printer found broken due to force dragging of paper. Gears of printer found broken and need to be replace with new one for working of printer.
No | Visit | Complaint | RDEU | 25.01.2019; Hr. Smuda; Tel.:020172382720; Pat Barcodes lassen sich nicht einscannen, Solution Pack wird nicht erkannt | Barcodescanner am Bildschirm liest weder Patientenbarcodes noch den Barcode einer Kal-Kassette ein. Kein Fehler bei Erkennung des Lösungspacks. | Gerät neu gestartet. Funktionskontrolle: I.O. (sowohl Patientenbarcod als auch Barcode der Kal-Kassette). | SWO-489671 | Defect | SW | RMED SW | RMED SW | Reboot | 25.01.2019; Mr. Smuda; Tel.:0 20172382720; Pat barcodes cannot be scanned, Solution Pack is not recognized | On-screen barcode scanner scans neither patient barcodes nor the barcode of a Kal cassette. No error in detecting the solution pack. | Device restarted. Functional control: I.O. (both patient barcode and barcode of the Kal cassette).
No | Visit | Complaint | RKOR | Interface error | . | Restart | SWO-489892 | Defect | SW | RMED SW | His/Lis |  | Interface error | . | Restart
No | Visit | Complaint | RKOR | Interface error | . | Interface PC restart | SWO-489893 | Setup | SW | RMED SW | AQURE | His/Lis | Interface error | . | Interface PC restart
No | Visit | Complaint | RCHN | 接口板腐蚀 | 接口板腐蚀,需要借件902-972一个 | 更换902-972一个正常。by Andrew | SWO-488503 | [5 empty] | Interface plate corrosion | Interface plate corrosion, need to borrow 902-972 one | Replace 902-972 one normal. by Andrew
No | Visit | Complaint | RCHN | 开机后无法启动系统 | 开机后无法启动系统（未更换备件） | 插拔CF卡后修复 by jesson | SWO-490095 | [5 empty] | The system cannot be started after power ing on | The system cannot be started after power-on (spare parts not replaced) | Fix by jesson after plugging CF card
No | Visit | Complaint | RCHN | 测试卡温度无法加热 | 测试卡温度无法加热 借902-972 903-334各一个 | 更换902-972一个正常。903-334好还。 by Andrew | SWO-487740 | [5 empty] | Test card temperature cannot be heated | Test card temperature can not be heated Borrowing 902-972 903-334 each | Replace 902-972 one normal. 903-334 good also. by Andrew
No | Visit | Complaint | RCHN | 测试卡定标失败 | 测试卡定标失败，需要借件903-030 | 更换903-030一个正常。 by Andrew | SWO-488451 | Defect | HW | Storage Module | PCB, Cartridge lock |  | Failed to test card calibration | Test card calibration failed, need to borrow 903-030 | Replace 903-030 one normal. by Andrew
No | Visit | Complaint | RCHN | po2定标失败，报错0378 | po2定标失败，报错0378 | 更换po2电极膜后修复 | SWO-490792 | Assisted Replacement | Electrode | Membrane | pO2 | 378 | po2 bid failure, error 0378 | po2 bid failure, error 0378 | Repair after replacing po2 electrode membrane
No | Visit | Complaint | RCRO | Often unstable pH measurement and CO2 calibration errors. Previously replaced lower sealing ring of pH electrode, but problems persited. Changed pH, Ref and CO2 electrodes and still the problem persited. 18.1. changed measuring chamber. | Učestali problem nestabilnosti mjerenja pH i drift na kalibraciji CO2 | Zamjena BG $El mjernog modula Zamjena pH i Ref elektroda Zamjena Ref i CO2 membrana | SWO-491414 | Defect | HW | pH/BG Module | measuring chamber |  | Often unstable pH measurement and CO2 calibration errors. Previously replaced lower sealing ring of pH electrode, but problems persited. Changed pH, Ref and CO2 electrodes and still the problem persited. 18.1. changed measuring chamber. | Frequent problem of instability of pH measurements and drift on CO2 calibration | Replacing the BG $El measuring module PH Replacement and Ref Electrda Replacement of Ref and CO2 membranes
No | Visit | Complaint | RCHN | PO2敏感性数值高 | PO2敏感性数值高,借903-334一个。 | 更换903-334后,重新校准修复 by jesson | SWO-489789 | [5 empty] | High PO2 sensitivity value | PO2 sensitivity value is high, borrowing 903-334 one. | Recalibrate repair by jesson after replacing 903-334
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 测试卡参比电极错误 1361 | 更换测试卡正常 | SWO-489841 | Assisted Replacement | SC | SC | Reference | 1361 | Test card, reagent package failure | Test card reference electrode error 1361 | Replace the test card ok
No | Visit | Complaint | RCHN | ph cl定标失败，报错0378 | ph cl定标失败，报错0378 (未使用备件) | 清洁进样支架和手动冲洗后修复 by jesson | SWO-490737 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway | pH 378 | ph cl bid failure, error 0378 | ph cl bid failure, error 0378 (Spare parts not used) | Repair by jesson after cleaning the sample holder and manual flushing
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | LAC,GLU,PH定标失败，0210 | 更换测试卡后修复 | SWO-490184 | [5 empty] | Test card, reagent package failure | LAC, GLU, PH calibration failure, 0210 | Repair after replacing test card
No | Visit | Complaint | RCHN | 定标底物分析物外露检查失败 1025，1024 | 定标底物分析物外露检查失败 1025，1024，借903-364，944-216*15331一张 | 更换903-364一个，重新定标正常。by jesson | SWO-489316 | Defect | HW | Process Module | Optical Unit | 1024 | Calibration base analysis exposure check failed 1025,1024 | Bid substrate analysis exposure check failure 1025,1024, borrow 903-364, 944-216 x 15331 | Replace the 903-364 one and re-label normally. by jesson
No | Visit | Complaint | RCHN | 定标失败 5297 | 定标失败 5297,借944-214*15174一张。 | 重新定标正常。 by Andrew | SWO-489314 | Assisted Replacement | Consumables | Test Kit | Cartridge, cal | 5297 | Bid failure 5297 | Bid failure 5297, borrow 944-214 x 15174 one. | Recalibration is normal. by Andrew
No | Visit | Complaint | RCHN | 光学强度超范围 1031 | 光学强度超范围 1031,借903-364一个，944-394*15120，944-258*15350，944-212*15388，944-220*15354各一张。 | 更换903-364，重新定标正常。 by Andrew | SWO-489310 | Defect | HW | Process Module | Optical Unit | 1031 | Optical strength over range 1031 | Optical strength super range 1031, borrow 903-364 one, 944-394 x 15120, 944-258 x 15350, 944-212 x 15388, 944-220 x 15354 each. | Replace the 903-364 and re-mark normally. by Andrew
No | Visit | Complaint | RESP | NO FUNCIONA CORRECTAMENTE. | REPARACION | EQUIPO REPARADO EN CENTRAL DE HEMOCUE. CON UN COSTE DE REPARACION DE 33€ | SWO-490264 | [5 empty] | IT'S NOT WORKING PROPERLY. | Service | EQUIPMENT REPAIRED IN HEMOCUE CENTRAL. WITH A REPAIR COST OF $33
No | Visit | Complaint | RESP | NO FUNCIONA CORRECTAMENTE. | REPARACION. | EQUIPO REPARADO EN CENTRAL HEMOCUE. CON UN COSTE DE 49€ | SWO-490265 | [5 empty] | IT'S NOT WORKING PROPERLY. | Service. | EQUIPMENT REPAIRED IN CENTRAL HEMOCUE. WITH a COST OF 49 EUROS
No | Visit | Complaint | RDEU | 25.01.2019 14:15 H. Krohn 0502192105721: Software startet nicht mehr | SW startet nicht | SW neu installiert, Konfiguration installiert und angepasst | SWO-490529 | Defect | SW | RMED SW | RMED SW |  | 25.01.2019 14:15 H. Krohn 0502192105721: Software no longer starts | SW does not start | SW reinstalled, configuration installed and customized
No | Visit | Complaint | RDEU | Glu und Lac Probleme Keine Kassette vor Ort | Glu und Lac Kal Fehler | Gerät überprüft, Lösungspack und Kassette erneuert, Konfiguration angepasst | SWO-490720 | Assisted Replacement | SC | SC | Glu/Lac | Cal | Glu and Lac Problems No cassette on site | Glu and Lac Kal Error | Device checked, solution pack and cassette renewed, configuration adapted
No | Visit | Complaint | RAME | Analyzer failing to communicate to LIS | Upon completion of last work order changes were made, I was unable to restart analyzer and commit changes, had to return at a time there was no surgeries in progress. | Committed IP change and restarted analyzer. | SWO-490251 | Setup | SW | RMED SW | IP address |  | Analyzer failing to communicate to LIS | Upon completion of last work order changes were made, I was unable to restart analyzer and commit changes, had to return at a time there was no surgeries in progress. | Committed IP change and restarted analyzer.
No | Visit | Complaint | RSCH | Pumpenkalibrierfehler + Black Screen | Pumpenkalibrierfehler + Black Screen | Allgemeine Kontrolle ok. Pumpenkalibrierfehler Servicetest calibration factror 0.76%. Gerät zerlegt alles überprüft ok. JW Kit ers ok. Pumpenrotor ers immer noch gleich. Lösungspack ers ok. Servicetest ok. Initialiesieren 2h. *** Taucht Fehler nochmal auf, werden wir ein Ersatzgerät zu verfügung stellen, Ihr Gerät wird bei RSCH überprüft ** | SWO-490277 | Assisted Replacement | SP | SP | Pump cal |  | Pumpenkalibrierfehler + Black Screen | Pumpenkalibrierfehler + Black Screen | General control ok. Pump calibration error Service test calibration factror 0.76%. Device disassembles everything checked ok. JW Kit ers ok. Pump rotor still the same. Solution pack ers ok. Service test ok. Initialies2h. If errors reappear, we will provide a replacement device, your device will be checked at RSCH **
No | Visit | Complaint | RIND | Level 2 Qc out of range for Cl- parameter. | Check the QC lot they are using & found that was Biorad Lypocheck control for immunoassay . | I have run the Biorad Liquochek control for Blood gas analysers Lot 22292,Run this twice as sample mode & QC mode & every time got the result of all the electrolytes parameter with in the range.Suggested to run the Blood gas specified QC of radiometer or other third party control. | SWO-490039 | [5 empty] | Level 2 Qc out of range for Cl- parameter. | Check the QC lot they are using & found that was Biorad Lypocheck control for immunoassay . | I have run the Biorad Liquochek control for Blood gas analysers Lot 22292,Run this twice as sample mode & QC mode & every time got the result of all the electrolytes parameter with in the range. Suggested to run the Blood gas specified QC of radiometer or other third party control.
No | Visit | Complaint | RAME | s02 not passing | Customer reported that the s02 was out on external QC | Reviewed data logs and found the s02 only out (just high) on level 2 external. Found sO2 out high on level C8704 of the system cycles intermittently as well. Customer has run the THB calibration and then the level 2 external will pass. Ran complete oxi test 3 times and found each test passed. Replaced hemolyzer and tested analyzer again. This hemolyzer was replaced less than 2 months after installation. Ran THB calibration, system cycles and manual QC, found all levels passed. | SWO-490630 | [5 empty] | s02 not passing | Customer reported that the s02 was out on external QC | Reviewed data logs and found the s02 only out (just high) on level 2 external. Found sO2 out high on level C8704 of the system cycles intermittently as well. Customer has run the THB calibration and then the level 2 external will pass. Ran complete oxi test 3 times and found each test passed. Replaced hemolyzer and tested analyzer again. This hemolyzer was replaced less than 2 months after installation. Ran THB calibration, system cycles and manual QC, found all levels passed.
No | Visit | Complaint | RCHN | 样本量不足 | 样本量不足 (未使用备件) | 清洁进样口和废液口通道 by jesson | SWO-490769 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway | Clot | Insufficient sample size | Insufficient sample size (Spare parts not used) | Clean the inlet and waste port channel by jesson
No | Visit | Complaint | RCHN | 进样仓门无法关闭 | 进样仓门无法关闭,借886-641，886-642各一个。 | 更换886-641，886-642后修复 by Andrew | SWO-489763 | Defect | HW | Cover and Chassis | Door, Storage |  | The inlet door cannot be closed | The entrance door cannot be closed, borrowing 886-641, 886-642 each. | Repair by Andrew after replacing 886-641,886-642
No | Visit | Complaint | RKK | 装置を通常使用中に繰り返し再起動が行われる | 繰り返し再起動が行われる | CPUボード交換 | SWO-492020 | Defect | HW | DMS_PC | CPU |  | Repeated restarts while the device is normally in use | Repeated restarts | CPU board replacement
No | Visit | Complaint | RIND | Gas 2 bottle empty. | Gas 2 bottle is empty. Gas 1 pressure is low. | Change gas2 bottle , run two point calibration. Calibration is ok. As pressure of gas 1 is low need to change . | SWO-491486 | Assisted Replacement | Consumables | Gas module | Gas 2 |  | Gas 2 bottle empty. | Gas 2 bottle is empty. Gas 1 pressure is low. | Change gas2 bottle , run two point calibration. Calibration is ok. As pressure of gas 1 is low need to change .
No | Visit | Complaint | RIND | Service Call | Oximetry values is not coming. Need to replace oxi module pcb. | Shut down the analyzer, replace the stand by oxi module PCB, start the analyzer, msg shows oxi module is not connected, upgrade the software, not successful because of in bootstrap analyzer is in offline mode, try 3-4 times not successful, again shut down the analyzer, replace oxi module with stand by, again problem occur, software not installed properly try old one 6.17 but not successful because of KIRA installation wizard failed again and again, install the 6.18 again, remove the pen drive when bootstrap is initializing, successfully shows online and sno of the analyzer, again insert the pen drive, successfully done the installation of the application software in to the analyzer, after initializing and 2 point calibration calibration error present in K and crea, Need to replace the membrane of K and crea in to the analyzer. | SWO-493467 | Defect | HW | Oxi module | PCB, oxi |  | Service Call | Oximetry values is not coming. Need to replace oxi module pcb. | Shut down the analyzer, replace the stand by oxi module PCB, start the analyzer, msg shows oxi module is not connected, upgrade the software, not successful because of in bootstrap analyzer is in offline mode, try 3-4 times not successful, again shut down the analyzer, replace oxi module with stand by, again problem occur, software not installed properly try old one 6.17 but not successful because of KIRA installation wizard failed again and again, install the 6.18 again, remove the pen drive when bootstrap is initializing, successfully shows online and sno of the analyzer, again insert the pen drive, successfully done the installation of the application software in to the analyzer, after initializing and 2 point calibration calibration error present in K and crea, Need to replace the membrane of K and crea in to the analyzer.
No | Visit | Complaint | RDEU | Anwender berichten, dass der Analysator häufig kalibriert, und bei K und Ca immer wieder Driften sind. | Analysator hat in unregelmäßigen Abständen immer wieder Grenzverletzungen bei Driften bei Na und Ca Parameter. Status und Empfindlichkeit aller Parameter in Ordnung. Überprüfung der Messkammern zeigt, dass eine starke Verunreinigung bei der Referenz und pH Elektrode besteht in den Messkammern und Elektrodenkontakten. | Gründliche Reinigung der Messkammern und bürsten der Elektrodenkontakte. Erneuerung des unteren Elektrodendichtrings bei der pH Elektrode, da verschmutzt. Schwankende Driften weiterhin vorhanden, Tausch der Referenz Elektrode. | SWO-490467 | Assisted Replacement | Electrode | Electrode | Reference |  | Users report that the analyzer is often calibrated, and that K and Ca are drifting again and again. | Analyzer has constant boundary violations during drifts at Na and Ca parameters at irregular intervals. Status and sensitivity of all parameters in order. Checking the measuring chambers shows that there is a strong impurity in the reference and pH electrode in the measuring chambers and electrode contacts. | Thorough cleaning of the measuring chambers and brushing of the electrode contacts. Renewal of the lower electrode sealing ring at the pH electrode, since polluted. Fluctuating drifts still present, swapping the reference electrode.
No | Visit | Complaint | RCRO | Rinse error, Inlet flow error | Rinse error, zaštopan analizator u području igle | Zamjena igle i brtve | SWO-491415 | Defect | HW | Inlet module | Probe |  | Rinse error, Inlet flow error | Rinse error, zaštopan analizator u području igle | Replacing the needle and gasket
No | Visit | Complaint | RDEU | 02336481471 Fr. Wirth Vakuum Fehler nach Anwendermaßnahme, nach Neustart lüft AU Software nicht mehr hoch | Abl90 Software startet nicht mehr. | Sicherung erstellt. Installation 3.4MR1 2. Installation 3.4MR1 Roma.db und .ini gelöscht. Software startet. Vakuum und Druckfehler. Clot removal, sol. pack Check & Pressure Test fehlgeschlagen. Wechsel des LP. Ha-07 zu HA-11. Vakuum $Druckfehler weiterhin. Schläuche und Dichtungen getestet. Ok Clotremoval durchgeführt. 2xManuellen Backflush durchgeführt. Clot Status SC value bei 1,81 SC getauscht. Hemolyzer ausgebaut, defekt, undicht | SWO-491522 | Revisit required | HW | Oxi module | Hemolyzer | 1063 | 02336481471 Fr. Wirth vacuum error after user action, after restart AU software no longer airs high | Abl90 software no longer starts. | Backup created. Installation 3.4MR1 2. Installation 3.4MR1 Roma.db and .ini deleted. Software starts. Vacuum and printing errors. Clot removal, sol. pack Check & Pressure Test failed. Changing the LP. Ha-07 to HA-11. Vacuum $Druckfehler on. hoses and seals. Ok Clotremoval. 2xManual backflush performed. Clot Status SC value at 1.81 SC. Hemolyzer removed, defective, leaky
No | Visit | Complaint | RCHN | 存储盘卡盘，测试卡断裂，报错5784，部位：wings | 存储盘卡盘，测试卡断裂，报错5784，部位：wings | 更换测试卡后修复 by jesson | SWO-491389 | Defect | Consumables | Test Kit | Cartridge, Broken | Wings | Memory disk card, test card break, error 5784, location: wings | Memory disk card, test card break, error 5784, location: wings | Fix by jesson after replacing test card
No | Visit | Complaint | RCHN | PH定标不稳定 0378 | PH定标不稳定 0378,更换PH密封圈。 | 更换PH密封圈。by jesson | SWO-491525 | Defect | Electrode | Electrode | Sealing ring, pH |  | PH calibration instability 0378 | PH calibration unstable 0378, replace the PH seal. | Replace the PH seal. by jesson
No | Visit | Complaint | RDEU | 29.01.2019 10:15Uhr Adebar 06213833528: Der bekannte Fehler wird weiterhin am Gerät angezeigt - Driftalarm bei Elektrolyten, ohne erkennbaren Grund | KAL Driften Elektrolyte 0376 Na 118 mmol $L Drift -4 Ca 4,75 mmol $L Drift -0,07 Cl 135 mmol $ L Drift 3 Bei Messung pH instabil $ gelegentlich auch Na und Ca | Allgemeine Überprüfung - Erneuerung Referenz Elektrode - Erneuerung pH Dichtring und Reinigung Elektrodenkontakte - Neustart durchgeführt -Funktionsüberprüfung etc. Service Dump und Data Logs gesichert | SWO-491145 | Revisit required | Electrode | Electrode | Reference | CAPA-733 | 29.01.2019 10:15Ade bar 06213833528: The known error is still displayed on the device - drift alarm for electrolytes, for no apparent reason | KAL Drift electrolytes 0376 Na 118 mmol /L Drift -4 Approx 4.75 mmol /L Drift -0.07 Cl 135 mmol - L Drift 3 When measuring pH unstable, occasionally Also Na and Ca | General check - Renewal reference electrode - Renewal pH sealing ring and cleaning electrode contacts - Restart performed -Function check etc. Service Dump and Data Logs Secured
No | Visit | Complaint | RAME | Oxi liquid sensor errors | OXI liquid sensor error | Flushed the air cent valve. Showed customer how to flush the air vent for future. All calibrations and 3 Levels of QC passed as per manufacturer specification | SWO-491203 | No fault found | No fault found | No fault found | No fault found |  | Oxygen Liquid Sensor Errors | OXYGEN Liquid Sensor Error | Flushed the air cent valve. Showed customer how to flush the air vent for future. All calibrations and 3 Levels of QC passed as per manufacturer specification
No | Visit | Complaint | RCHN | lac灵敏度偏低，报错0578 | lac灵敏度偏低，报错0578（未更换备件） | 更换lac膜后修复 | SWO-491407 | Assisted Replacement | Electrode | Membrane | Lac | 578 | lac sensitivity is low, error 0578 | lac sensitivity is low, error 0578 (not replaced spare parts) | Repair after replacing lac film
No | Visit | Complaint | RIND | Lac cal sensitivity low. | Lac calibration sensitivity low, need to replace the Lac membrane. | Replaced the Lac membrane and run the instrument its working good. | SWO-492320 | Assisted Replacement | Electrode | Membrane | Lac | Sens | Lac cal sensitivity low. | Lac calibration sensitivity low, need to replace the Lac membrane. | Replaced the Lac membrane and run the instrument its working good.
No | Visit | Complaint | RIND | Service Call | Need to replace the Na and crea membrane. | Replace the crea and K membrane successfully, and do all the required setting according to the user because all the data has been lost during the application software installation, user do the QC and sample, successfully done analyzer handed over to staff in well working condition. | SWO-493471 | Assisted Replacement | Electrode | Membrane | Crea |  | Service Call | Need to replace the Na and crea membrane. | Replace the crea and K membrane successfully, and do all the required setting according to the user because all the data has been lost during the application software installation, user do the QC and sample, successfully done analyzer handed over to staff in well working condition.
No | Visit | Complaint | RCHN | 扫不到质控 | 扫不到质控，借910-310一个 | 更换910-310后修复 by jesson | SWO-490692 | Defect | HW | Autocheck module | Barcode Reader |  | Sweeping no QC | Sweep not to the QC, borrow 910-310 one | Fix by jesson after replacing 910-310
No | Visit | Complaint | RSCH | Geräteaustausch und Repartur | Geräteaustausch für Rep LQC | Leihgerät 838R0268N005 in Wattwil installiert und alles manuell konfiguriert. CRP Neue Lot kalibriert Alle Parameter LQC gemacht Funktionskontrolle Vorort i.o | SWO-490807 | Defect | HW | Process Module | Process Wheel |  | Device replacement and reparture | Device replacement for Rep LQC | Rental device 838R0268N005 installed in Wattwil and all manually configured. CRP New Solder Calibrated All parameters LQC made Functional control suburb i.o
No | Visit | Complaint | RFRA | Corrupt xml | RAS, Chgt CPU | Sauvegarde des données Chgt: CompactFlash, CPU. Installation Windows et logiciel AQT. Nettoyage: process wheel, Wash Unit. Configuration Windows, installation sauvegarde de l'AQT90. Redémarrage routine ok. Passage Cleaning, LCQ. Réglage porte coulissante. | SWO-491093 | Defect | HW | Analyzer Control Module | CPU | 3050 | Corrupt xml | RAS, Chgt CPU | Backing up data Chgt: CompactFlash, CPU. Windows installation and AQT software. Cleaning: process wheel, Wash Unit. Windows configuration, installation backup of the AQT90. Reboot routine ok. Passage Cleaning, LCQ. Sliding door setting.
Yes | Visit | Complaint | RPOL | Podłączenie do sieci LIS/HIS | Podłączenie do sieci LIS $HIS | Konfiguracja aparatu. | SWO-491533 | Re classified RMED | planned SWO |  |  | See mail with title "SWO-491533" | Connection to LIS/HIS network | Connection to the LIS network /HIS | Camera setup.
No | Visit | Complaint | RPOL | Problem przepływowy. | Diagnoza błędu- problem przepływowy. | Oczyszczenie toru pomiarowego, oczyszczenie czujnika Punch down. Testy serwisowe. Uzupełnienie brakujących odczynnikow. | SWO-494973 | Cleaned | HW | Wet section | Wet section | Flow restriction | Flow Problem. | Error diagnosis-flow problem. | Cleaning the measuring track, clearing the Punch down sensor. Service tests. Replenishment of missing reagents.
No | Visit | Complaint | RKK | 検体、ＱＣ測定でスペクトルミスマッチ発生 | 状態確認 | ヘモライザ交換 調整 | SWO-491998 | Defect | HW | Oxi module | Hemolyzer | 581 | Spectral mismatch occurs in sample and QC measurements | Status Check | Hemolizer Replacement Adjust
No | Visit | Complaint | RDEU | 02336481471 Fr. Wirth Vakuum Fehler nach Anwendermaßnahme, nach Neustart lüft AU Software nicht mehr hoch | Vakuumfehler | Hemolyzer getauscht. Backup hergestellt & Konfiguration durchgeführt. Glu & Lac außerhalb der Kalibriergrenzen. Mehrere Kalibrierungen notwendig. Ok. Kal & QC. tHb Kal. | SWO-492183 | Defect | HW | Oxi module | Hemolyzer | 1063 | 02336481471 Fr. Wirth vacuum error after user action, after restart AU software no longer airs high | Vakuumfehler | Hemolyzer. Backup made & configuration performed. Glu & Lac outside the calibration limits. Multiple calibrations required. Ok. Kal & QC. tHb Kal.
No | Visit | Complaint | RAME | The back up battery needs to be replaced | The back up battery is defective | New battery installed and tested ok. The time format changed to 24 hours per customer request. The credit form setup done to have all customer information on it. | SWO-491745 | Defect | HW | Battery module | Battery module |  | The back up battery needs to be replaced | The back up battery is defective | New battery installed and tested ok. The time format changed to 24 hours per customer request. The credit form setup done to have all customer information on it.
No | Visit | Complaint | RAME | Battery error 1599 | Found that analyzer giving battery error 1599 one or more cell are not working. | Replaced battery which corrected the issue. | SWO-491865 | Defect | HW | Battery module | Battery module | 1599 | Battery error 1599 | Found that analyzer giving battery error 1599 one or more cell are not working. | Replaced battery which corrected the issue.
No | Visit | Complaint | RSAF | Marinka called to ask me if i can please call Wallace regarding the touch screen not accepting the icon pressed but rather above or below where the sisiters are pressing and can the touch screen be calibrated. | Touch screen not aligned due to the blue plastic screen casing being broken at three of the four screw mounts $fixation tabs. | New display unit to be collected from warehouse and to be installed at customers site. Display units blue casing is cracked off on three of the four screw mounts. | SWO-491737 | Revisit required | HW | Display | Display | screws | Marinka called to ask me if i can please call Wallace regarding the touch screen not accepting the icon pressed but rather above or below where the sisiters are pressing and can the touch screen be calibrated. | Touch screen not aligned due to the blue plastic screen casing being broken at three of the four screw mounts /fixation tabs. | New display unit to be collected from warehouse and to be installed at customers site. Display units blue casing is cracked off on three of the four screw mounts.
No | Visit | Complaint | RAME | Missing Operator IDs from AQURE to analyzer | Worked with Point of Care Coordinator and Radiometer SIT. | Pushed operator ID's remotely on one analyzer and manually added to other analyzer. | SWO-491909 | Setup | SW | RMED SW | Network | Aqure | Missing Operator IDs from AQURE to analyzer | Worked with Point of Care Coordinator and Radiometer SIT. | Pushed operator ID's remotely on one analyzer and manually added to other analyzer.
No | Visit | Complaint | RESP | Touch Pannel Screen | Comprobación del táctil del Display | Se sustituye el panel táctil del Display y calibramos el display por service. Comprobación final y calibración del sensor. | SWO-491932 | [5 empty] | Touch Pannel Screen | Checking the Display Touch | The Display touchpad is replaced and we calibrate the display by service. Final check and sensor calibration.
No | Visit | Complaint | RDEU | 29.01.2019 10:15Uhr Adebar 06213833528: Der bekannte Fehler wird weiterhin am Gerät angezeigt - Driftalarm bei Elektrolyten, ohne erkennbaren Grund | KAL und QC Fehler pH und Elektrolyte (Nach Erneuerung Referenz Elektrode zunächst ohne Beanstandung) Gelegentliche Kalibrier-Drift-Verletzung bei Lactat | Erneute Membranierung Referenz Elektrode (Herr Adebar) - Allgemeine Überprüfung - KAL und QC ohne Beanstandung - Funktionsüberprüfung Bemerkung: Nach Absprache mit Herr Adebar wurde zu Versuchszwecken die Kalibrierung Drifttoleranz bei Lactat von 0,2 auf 0,5 mmol $L erhöht ! Analysator-Verhalten wird beobachtet | SWO-491801 | Assisted Replacement | Electrode | Electrode | Reference | CAPA-733 | 29.01.2019 10:15Ade bar 06213833528: The known error is still displayed on the device - drift alarm for electrolytes, for no apparent reason | KAL and QC error pH and electrolytes (After renewal reference electrode initially without complaint) Occasional calibration drift injury in lactate | Re-membraneation reference electrode (Mr. Adebar) - General review - KAL and QC without complaint - Functional check Note: After consultation with Mr. Adebar, the calibration of drift tolerance for lactate was increased from 0.2 to 0.5 mmol $L for experimental purposes ! Analyzer behavior is observed
No | Visit | Complaint | RIND | Service Visit. | Display is not working. | Try to clean with spirit, not successful, corner touch of the screen totally failed, presently provide mouse for operating the analyzer. Need to replace the display unit. | SWO-491815 | Revisit required | HW | Display | Touch Screen |  | Service Visit. | Display is not working. | Try to clean with spirit, not successful, corner touch of the screen totally failed, presently provide mouse for operating the analyzer. Need to replace the display unit.
No | Visit | Complaint | RIND | Po2 values high for qc | Po2 values high for qc | checked the po2 electrode cleaned the electrode and chamber then tested the qc then also it is higher side so membrane changed and checked for latch after that tested level 1 qc now its working fine | SWO-493466 | Assisted Replacement | Electrode | Membrane | pO2 |  | Po2 values high for qc | Po2 values high for qc | checked the po2 electrode cleaned the electrode and chamber then tested the qc then also it is higher side so membrane changed and checked for latch after that tested level 1 qc now its working fine
No | Visit | Complaint | RAME | AutoCheck Carousel failure | Looked in data logs found numerous 742 error codes. Found worn 14 pole cable. | Replaced 14 pole cable. Then left and a few hours later she was having issues with the auto check doing the QC automatically. So came next day with another 14 pole cable. Found cable was crooked, reseated the connector and cable. | SWO-491283 | Defect | HW | Autocheck module | FPC cable | 742 | AutoCheck Carousel failure | Looked in data logs found numerous 742 error codes. Found worn 14 pole cable. | Replaced 14 pole cable. Then left and a few hours later she was having issues with the auto check doing the QC automatically. So came next day with another 14 pole cable. Found cable was crooked, reseated the connector and cable.
No | Visit | Complaint | RDEU | 16.01.2019, 08:35 Uhr, Fr. Glombig, Tel. 0212645645431: Kal Fehler bei Lyten | Ständig Driften bei Ca und Na | Referenz-Elektrode ersetzt, mehrere Kalibrationen, Qcs durchgeführt und bestanden | SWO-492243 | Assisted Replacement | Electrode | Electrode | Reference | CAPA-733 | 16.01.2019, 08:35, Fr. Glombig, Tel. 0212645645431: Kal Error at Lyten | Constantly drifting at Ca and Na | Reference electrode replaced, several calibrations, Qcs performed and passed
No | Visit | Complaint | RESP | Calibração de NT-proBNP abortada | Verificação do registro de erros | Limpeza dos sensores do punch cup. Inicialização. Calibração dos sensores do punch cup. Verificação do sistema. Verificação do correto funcionamento com execução de várias amostras. | SWO-491795 | Cleaned | HW | Process Module | sensor, punch out | 5769 | Aborted NT-proBNP calibration | Error log Verification | Cleaning the sensors of the Punch Cup. Startup. Calibration of the Punch Cup sensors. System check. Verification of the correct operation with the execution of several samples.
No | Visit | Complaint | RBNL | Geen Ht mogelijk | Alle metingen worden afgekeurd wegens error Geen Ht-meting mogelijk . Washblok heb gedemonteerd en doorgespoten waarna de Ht-meting weer mogelijk Was. | Washblok was verstopt. | SWO-491793 | Cleaned | HW | Process Module | Needle Wash Unit |  | No Ht possible | All measurements are rejected due to error no Ht measurement possible. Washblok have been disassembled and sprayed after which the Ht measurement Was possible again. | Washblock was clogged.
No | Visit | Complaint | RPOL | Błędy przepływowe |  |  | SWO-492148 | Defect | HW | Oxi module | Pinch valve | SWO updated with: Pump calibration error, leak detected Replacing the pinch valve | Flow errors
No | Visit | Complaint | RDEU | Verbindung mit Aqure ist unterbrochen. | Netzwerkverbindung unterbrochen. | Ip-Adresse, Gateway & DNS überprüft (Windowseinstellungen). ok. Ip-Adresse zum Aqure System korrigiert. Verbindung ok. Analysator neugestartet und getestet. ok. | SWO-492456 | Setup | SW | RMED SW | Network | Aqure, IP address | Connection to Aqure is lost. | Network connection interrupted. | Ip address, Gateway & DNS checked (Windows settings). Ok. Ip address to the Aqure system corrected. Connection ok. Analyzer restarted and tested. Ok.
No | Visit | Complaint | RLTD | Sally Thirkettle reported flow selector errors on analyser | Site visit to investigate analyser residue on sp after initial report. New pack had been fitted and a good period of time had passed for the powder to reoccur since installing. | Arrived on site. Customer was at analyser performing gasket clean. Waited for this to complete before inveatigation. Entered service and ejected sp. Found a very small amount of residue on side of sp and at same point on side of pack bay. Noticed a little of the black mark on the pack pump tube. Cleaned all of this off and removed $inspected pump roller. All rollers freely moving. Lubricated rollers as precaution. Inspected pack and found no obvious signs of damage. All gears look in good condition and sitting in correct position. Refitted all parts and ran fluid transport tests. All passed with very good results. Exited back to main menu and ran calibration and qc. Analyser running fine. Black residue does not look to be an issue and analyser running very well. Advised customer to monitor and report to radiometer if any problems arise. | SWO-492311 | No fault found | No fault found | No fault found | Flow selector detector |  | Sally Thirkettle reported flow selector errors on analyser | Site visit to investigate analyser residue on sp after initial report. New pack had been fitted and a good period of time had passed for the powder to reoccur since installing. | Arrived on site. Customer was at analyser performing gasket clean. Waited for this to complete before inveatigation. Entered service and ejected sp. Found a very small amount of residue on side of sp and at same point on side of pack bay. Noticed a little of the black mark on the pack pump tube. Cleaned all of this off and removed $inspected pump roller. All rollers freely moving. Lubricated rollers as precaution. Inspected pack and found no obvious signs of damage. All gears look in good condition and sitting in correct position. Refitted all parts and ran fluid transport tests. All passed with very good results. Exited back to main menu and ran calibration and qc. Analyser running fine. Black residue does not look to be an issue and analyser running very well. Advised customer to monitor and report to radiometer if any problems arise.
No | Visit | Complaint | RSAF | Marinka called to ask me if i can please call Wallace regarding the touch screen not accepting the icon pressed but rather above or below where the sisiters are pressing and can the touch screen be calibrated. | Collected Display unit and installed at customer. | Installed and tested Display unit, tested - ok. | SWO-492160 | Defect | HW | Display | Display | screws | Marinka called to ask me if i can please call Wallace regarding the touch screen not accepting the icon pressed but rather above or below where the sisiters are pressing and can the touch screen be calibrated. | Collected Display unit and installed at customer. | Installed and tested Display unit, tested - ok.
No | Visit | Complaint | RAME | s02 not passing | Returned too site with additional parts to address s02 issues | Analyzer was working upon arrival and the QC run today was fine internal and external. Replaced the co-ox control board, spectrometer, hemolyzer all together at once. Powered analyzer back up and it completed the Tikhonov Matrix calibration. Ran THB calibration, system cycles and manual QC found all levels passed. | SWO-492035 | [5 empty] | s02 not passing | Returned too site with additional parts to address s02 issues | Analyzer was working upon arrival and the QC run today was fine internal and external. Replaced the co-ox control board, spectrometer, hemolyzer all together at once. Powered analyzer back up and it completed the Tikhonov Matrix calibration. Ran THB calibration, system cycles and manual QC found all levels passed.
No | Visit | Complaint | RAME | s02 not passing | Returned too site to run VK kit, doing the linearity due to replacing the oximetry related parts. | Customer reported that the QC was fine this morning on s02. Reviewed the QC results before proceeding with linearity. Ran VK-R5 linearity kit and found all levels passed Picked up additional shipment from local fedex and left additional level 2 QC with customer | SWO-492847 | [5 empty] | s02 not passing | Returned too site to run VK kit, doing the linearity due to replacing the oximetry related parts. | Customer reported that the QC was fine this morning on s02. Reviewed the QC results before proceeding with linearity. Ran VK-R5 linearity kit and found all levels passed Picked up additional shipment from local fedex and left additional level 2 QC with customer
No | Visit | Complaint | RCHN | 频繁出现1068干燥前温度超范围 | 频繁出现1068干燥前温度超范围，借903-210一个 | 更换903-210后修复 by jesson | SWO-490814 | Defect | HW | Process Module | PCB, cup dryer | 1068 | Frequent 1068 pre-drying temperature overrange | Frequent 1068 pre-drying temperature over range, borrow 903-210 a | Repair by jesson after replacement 903-210
No | Visit | Complaint | RCHN | MYO定标数据丢失 | MYO定标数据丢失 借944-214（14934）一张 | 重新定标后修复 by jesson | SWO-491397 | Assisted Replacement | Consumables | Test Kit | Cartridge, cal |  | MYO calibration data loss | MYO calibration data loss Borrowed 944-214 (14934) one | Fix by jesson after re-calibration
No | Visit | Complaint | RCHN | pco2定标值偏低。报错0378 | pco2定标值偏低。报错0378 | 更换pco2电极膜后修复 | SWO-502150 | Assisted Replacement | Electrode | Membrane | pCO2 | Low Sens | The pco2 calibration value is low. Error 0378 | The pco2 calibration value is low. Error 0378 | Repair after replacing pco2 electrode membrane
No | Visit | Complaint | RDEU | Hotline - immer wieder Driften bei Lyten, QC wird dadurch verletzt, wiederholte Kal löst das Problem aber bei nächster Kal wieder Fehler | mehrfach am Tag Driften bei Lyten und pH, Membran Ref wurde schon mehrfach ersetzt | Ref Elektrode ersetzt, fällige Membran ersetzt, mehrere Kal, QC | SWO-493044 | Assisted Replacement | Electrode | Electrode | Reference | CAPA-733 | Hotline - repeatedly drifting at Lyten, QC is injured by this, repeated Kal solves the problem but again with next Kal error | drifting several times a day at lytes and pH, membrane ref has already been replaced several times | Ref electrode replaced, overdue membrane replaced, several Kal, QC
No | Visit | Complaint | RIND | Printer not working. | Checked the printer, paper jammed and also documents pending to print. | Cancel all the pending document, Clear the jam and instal a new paper roll. | SWO-492687 | Cleaned | HW | Printer Module | Printer Module | Paper jam | Printer not working. | Checked the printer, paper jammed and also documents pending to print. | Cancel all the pending document, Clear the jam and instal a new paper roll.
No | Visit | Complaint | RAME | Customer reports a broken inlet probe | Customer called to report broken probe. | Replaced probe with parts drawn from customer stock. Also replaced Inlet Tube, 841-780, for customer. Normal operations restored. | SWO-492607 | Defect | HW | Inlet module | Probe |  | Customer reports a broken inlet probe | Customer called to report broken probe. | Replaced probe with parts drawn from customer stock. Also replaced Inlet Tube, 841-780, for customer. Normal operations restored.
No | Visit | Complaint | RDEU | 02051982383705 Hr. Schick Med Technik immer wieder Na QC zu hoch gemessen | Immer wieder driften bei Na und Ca, Na QCs zu hoch, Status von Na sinkt nach einigen Kalibrationen ab, nach neuer Ref Membran Status wieder gut, | Messkammern gereinigt, Referenz Elektrode ersetzt, mehrere Kalibrationen, QC Werte jetzt deulich niedriger, lv3 vorher 130mmol $L jetzt 126mmol $L | SWO-492407 | Cleaned | Fluidic pathway | pH/BG Module | measuring chamber |  | 02051982383705 Mr. Schick Med Technique repeatedly measured Na QC too high | Again and again drift at Na and Ca, Na QCs drift too high, status of Na decreases after some calibrations, after new ref membrane status again good, | Measuring chambers cleaned, reference electrode replaced, several calibrations, QC values now significantly lower, lv3 previously 130mmol $L now 126mmol $L
No | Visit | Complaint | RESP | Error 316 Neon number too low. | Se completa la reparación colocando un módulo OXI completo. | Cambiar el módulo OXI, calibrar y probar con controles. Módulo quitado 902-833R2418N0009. | SWO-492980 | Defect | HW | Oxi module | Oxi module | Neon number too low | Error 316 Neon number too low. | The repair is completed by placing a complete OXI module. | Change the OXI module, calibrate and test with controls. Module removed 902-833R2418N0009.
No | Visit | Complaint | RFRA | L'écran tactile est très lent à répondre aux sollicitations Lorsque la cliente veut accéder au menu Consommable, l'écran devient noir et revient Erreur échantillon : tube pas assez remplis 2 ABL90 ont fréquement l'imprimante qui se deconnecte | RAS, Chgt CPU | Sauvegarde des données Chgt: CompactFlash, CPU. Installation Windows et logiciel AQT. Nettoyage: process wheel, Wash Unit, circuits vidanges. Configuration Windows, installation sauvegarde de l'AQT90. Redémarrage routine ok. Passage Cleaning. | SWO-492358 | Defect | HW | Analyzer Control Module | CPU |  | The touch screen is very slow to respond to requests When the customer wants to access the Consumable menu, the screen turns black and comes back Sample error: tube not filled enough 2 ABL90s frequently have the printer that delogs | RAS, Chgt CPU | Backing up data Chgt: CompactFlash, CPU. Windows installation and AQT software. Cleaning: process wheel, Wash Unit, drain circuits. Windows configuration, installation backup of the AQT90. Reboot routine ok. Passage Cleaning.
No | Visit | Complaint | RKK | 乾燥前温度範囲外 | 朝一の検体測定のみエラー発生 | 室内温度は常に20度以上になっている。 ドライヤーユニットと温度検出PCB交換実施。 | SWO-493645 | Defect | HW | Process Module | PCB, cup dryer |  | Outside the pre-drytemperature range | Error occurs only in the morning sample measurement | The indoor temperature is always above 20 degrees. The dryer unit and temperature detection PCB replacement were implemented.
Yes | Visit | Complaint | RDEU | 05.12.2018 14:30Uhr Herr Dietz 0627183277: LQC bei Tnt außerhalb Bereich 20.12.2018; 16:10 Uhr; Hr. Dietz; Tel.: 0627183671; TnT QC Ergebnisse außerhalb der RiliBÄK 10.01.2019; 11:00 Uhr; Fr. Drachsler; Tel.: 0627183245; TnT RiliBÄK Granzen/Abweichungen | Nach Rücksprache mit Technischen Support und Produktmanager ziehen der aktuellen Daten zur Auswertung in Kopenhagen und zur Klärung der Ursache der erhöhten Qualitätskontroll Level bei TnT. Analysator wurde von Anwender gesperrt, es werden vorab keine Patientenmessungen durchgeführt. | Ziehen der Datalogs und eines Service Dumps. Daten an Technician Support geladen. | SWO-492632 | Re classified RMED | planned SWO |  |  | RDEU 832293698 | 05.12.2018 14:30 Mr Dietz 0627183277: LQC at Tnt outside area 20.12.2018; 4:10 p.m.; Mr. Dietz; Tel.: 0627183671; TnT QC Results outside the RiliBÄK 10.01.2019; 11 a.m.; Ms. Drachsler; Tel.: 0627183245; TnT RiliBÄK Granzen/Deviations | After consultation with technical support and product managers, the current data for evaluation in Copenhagen and to clarify the cause of the increased quality control levels at TnT. Analyzer has been blocked by users, no patient measurements are carried out in advance. Performed. | Drag the data logs and a service dump. Data loaded to Technician Support.
Yes | Visit | Complaint | RPOL | Wprowadzenie kodów kreskowych dla operatorów OIOM | Wprowadzenie kodów kreskowych dla wszystkich operatorów OIT | Konfiguracja aparatu. | SWO-492975 | Re classified RMED | planned SWO |  |  | See mail with title "SWO-492975" | Introduction of barcodes for OIOM operators | Introducing barcodes for all OIT operators | Camera setup.
No | Visit | Complaint | RDEU | 01.02.2019 10:40Uhr Frau Schuller 094822026218: Das Gerät meldet Oximetriefehler bei Patientenmessung und QC | Pinch Valve fest. | Pinch Valve und Ventilschlauch erneuert. Oxi Complete Test ok. Pinch Valve schaltet korrekt. SC Fach lässt sich öffnen und schließen. Kalibration und QC’s vorgenommen. Testmessung in Ordnung. | SWO-493185 | Defect | HW | Oxi module | Pinch valve |  | 01.02.2019 10:40Am Ms Schuller 094822026218: The device reports oximetry errors in patient measurement and QC | Pinch Valve fest. | Pinch Valve and valve hose. Oxi Complete Test ok. Pinch Valve switches correctly. SC Fach can be opened and closed. calibration and QC's. Test measurement in order.
No | Visit | Complaint | RDEU | immer wieder werden die Pat.messungen mit ? ausgedruckt, Fehler Trübung zu hoch, Oximetriefehler, 01.02.2019; Hr. Dorendorf; Tel.: 021513341862; der Oxy Spectrum mismatch error ist wieder aufgetreten | Bei Patientenmessungen wird oft der Fehler gemeldet: 0292 Trübung zu hoch, 0581 OXI spectrum mismatch und 0688 ctHb $ceHb zu niedrig für Oximetrie-Berechnung. | Hemolyzer Defekt - ersetzt, Startphase durchgeführt, Kalibriert, tHb kal und QC’s durchgeführt. | SWO-493816 | Defect | HW | Oxi module | Hemolyzer | 581 | again and again the Pat.measurements with ? printed, error turbidity too high, oximetry errors, 01.02.2019; Mr. Dorendorf; Tel.: 021513341862; the Oxy Spectrum mismatch error has reappeared | Patient measurements often report the error: 0292 turbidity too high, 0581 OXI spectrum mismatch and 0688 ctHb /ceHb too low for oximetry calculation. | Hemolyzer defect - replaced, start phase performed, calibrated, tHb kal and QC's performed.
No | Visit | Complaint | RDEU | 01.02.2019 11:30Uhr Frau Trilk 03863520139: Trotz Tausch von SC und LP weiterhin QC-Fehler bei pCO2 im LevelA (zu hoch) | No | Softwareupdate 3.4mr1 | SWO-493195 | Defect | SW | RMED SW | pCO2 | QC | 01.02.2019 11:30am Ms Trilk 03863520139: Despite swapping SC and LP, QC error at pCO2 in LevelA (too high) | No | Software Update 3.4mr1
Yes | Visit | Complaint | RDEU | Patientenberichtslayout nach Vorgabe konfigurieren | Patientenberichtslayout & Benutzerverwaltung | Nach Vorgaben konfiguriert | SWO-493701 | Re classified RMED | planned SWO |  |  | See e-mail with title "AW: SWO-493701" from BORWE | Configure patient report layout by default | Patient Report Layout & User Management | Configured according to specifications
No | Visit | Complaint | RDEU | immer wieder Fehler, LP muss ersetzt werden, gerät soll gespült werden, | Fließwege müssen immer wieder gespült werden, Abf.schlauch ist verstopft, | Einl.dichtung ersetzt, LP ersetzt, Ventilschlauch ersetzt, Modul ausgebaut und Ventil gereinigt, Abf.schlauch ersetzt, Gerät kalibriert und QC durchgeführt, | SWO-493257 | Defect | HW | Basic Unit | PM Kit |  | repeated errors, LP must be replaced, device should be flushed, | Flow paths have to be rinsed again and again, drain hose is clogged, | Single seal replaced, LP replaced, valve hose replaced, module removed and valve cleaned, drain hose replaced, device calibrated and QC performed,
No | Visit | Complaint | RDEU | Gerät lässt sich im Helios Netzwerk nicht anpingen | anpingen des Gerätes im Netzwerk erlauben | Gemäß Helios Arbeitsanweisung anpingen des Gerätes inWin Firewall aktiviert, Geräteneustart: anpingen nicht möglich, Gerätekonfiguration mehrfach gemäß Anleitung wiederholt: anpingen nicht möglich -> RTP erstellt in an TS geschickt Kal & QC Daten gemäß RDEU Anweisung gesichert und dokumentiert | SWO-493137 | Setup | SW | RMED SW | His/Lis |  | Device cannot be pinged in Helios network | allow anping the device on the network | According to Helios working instruction pinging the device enabled inWin Firewall, device restart: unable to ping, device configuration repeated repeatedly according to instructions: anping not possible -> RTP created in sent to TS Kal & QC data backed up and documented in accordance with RDEU instructions
No | Visit | Complaint | RDEU | 01.02.2019 9:30 Fr. Selmann MT 0401818828816: Geräteausfall | System überprüft => PC defekt | Pc aus alt Gerät eingebaut, Kalibration und Qc durchgeführt | SWO-493106 | Defect | HW | DMS_PC | CPU |  | 01.02.2019 9:30 Fr. Selmann MT 0401818828816: Geräteausfall | System checked => PC defective | Pc built from old device, calibration and Qc performed
No | Visit | Complaint | RAME | Broken Display(s) | 1) New analyzer installation had broken display. 2) Ordered new display to replace broke display and new display was also damaged in same location. | 1) Ordered another display. 2) Submitted CQC and Installation Report. See attached. | SWO-494124 | Defect | HW | Display | Display |  | Broken Display(s) | 1) New analyzer installation had broken display. 2) Ordered new display to replace broke display and new display was also damaged in same location. | 1) Ordered another display. 2) Submitted CQC and Installation Report. See attached.
No | Visit | Complaint | RCHN | 报错1024 | 报错1024，借903-364一个 | 更换903-364后修复 by jesson | SWO-492044 | Defect | HW | Process Module | Optical Unit | 1024 | 1024 error | 1024, loan 903-364 one | Repair by jesson after replacing 903-364
No | Visit | Complaint | RCHN | 进样门无法关闭 | 进样门无法关闭 借903-028一个 | 更换903-028后修复by tim | SWO-492214 | Defect | HW | Cover and Chassis | PCB, cabinet top | 4548 | The sample door cannot be closed | The sample door can't be closed borrowing 903-028 one | Repair by tim after replacement 903-028
No | Visit | Complaint | RCHN | 液路传感错误 | 液路传感错误 | 自行更换905-908后修复by tim | SWO-493074 | Defect | HW | Process Module | Sensor, BnP | 1235 | Liquid path sensing error | Liquid path sensing error | Repair by tim after replacing 905-908 on your own
No | Visit | Complaint | RCHN | 空白杯卡住处理盘 5784，未到处理包中 | 空白杯卡住处理盘 5784，未到处理包中 | 把卡住的空白杯取出。by tim | SWO-492978 | Defect | Consumables | Solution Pack | Solution Pack | Stack | Blank cup stuck to the processing tray 5784, not in the processing package | Blank cup stuck to the processing tray 5784, not in the processing package | Remove the stuck blank cup. by tim
No | Visit | Complaint | RCHN | HCT测量结果不出值打问号 | HCT测量结果不出值打问号 借903-329一个 | 更换903-329后修复 by jesson | SWO-491363 | Defect | HW | Aspiration Module | Needle Wash Unit | 1223 | HCT measurement results do not mark the value of the question mark | HCT measurement results do not have a value to hit the question mark borrowed 903-329 one | Repair by jesson after replacing 903-329
No | Visit | Complaint | RCRO | Often unstable pH measurement and CO2 calibration errors. Previously replaced lower sealing ring of pH electrode, but problems persited. Changed pH, Ref and CO2 electrodes and still the problem persited. 18.1. changed measuring chamber. | pH mjerenje više nema problema s nestabilnošću. CO2 kalibracija i dalje ima par puta dnevno preveliki drift | Zamjena CO2 elektrode | SWO-502180 | Assisted Replacement | Electrode | Electrode | pCO2 | drift | Often unstable pH measurement and CO2 calibration errors. Previously replaced lower sealing ring of pH electrode, but problems persited. Changed pH, Ref and CO2 electrodes and still the problem persited. 18.1. changed measuring chamber. | pH measurement no longer has problems with instability. CO2 calibration still has a couple of times a day oversized drift | Replacing the CO2 electrode
No | Visit | Complaint | RCHN | 做样本时没有检测到气体断，进样针传感器错误 | 做样本时没有检测到气体段，进样针传感器错误，借902-832一个 | 更换902-832后修复 by jesson | SWO-491451 | Defect | HW | pH/BG Module | LS |  | No gas break was detected when the sample was taken, the sample needle sensor error | No gas segment detected when taking samples, error of the sample needle sensor, 902-832 one | Repair by jesson after replacement of 902-832
Yes | Visit | Complaint | RCHN | 做样本时没有检测到气体断，进样针传感器错误 | 加借902-832,902-446各一个 |  | SWO-492983 | Re classified RMED | same day upSWO |  |  |  | No gas break was detected when the sample was taken, the sample needle sensor error | Borrow 902-832, 902-446 each
No | Visit | Complaint | RIND | Oxi spectrum mismatch. | In every sample there is oxi spectrum mismatch and patient reports having ? on oximetry values. | Check the machine and run complete oximetry test, All test found passed in that. Run protein removal, 2 point calibration and tHb calibration in service program. After all this again run patient sample but still isuue is same of oxi spectrum mismatch. Than I have changed the hemolyzer in oxi module, adjust the lamp unit after this run complete oximetry test successfully. Run a 2 point calibration and patient sample successfully. | SWO-497891 | Defect | HW | Oxi module | Hemolyzer |  | Oxi spectrum mismatch. | In every sample there is oxi spectrum mismatch and patient reports having ? on oximetry values. | Check the machine and run complete oximetry test, All test found passed in that. Run protein removal, 2 point calibration and tHb calibration in service program. After all this again run patient sample but still isuue is same of oxi spectrum mismatch. Than I have changed the hemolyzer in oxi module, adjust the lamp unit after this run complete oximetry test successfully. Run a 2 point calibration and patient sample successfully.
No | 5 | Y | RCZE | Blank |  |  | D59E98917A3F7EF3 | Assisted Replacement | Electrode | Electrode | pCO2 |  | Blank
No | Visit | Complaint | RKK | オキシハードウエアエラー（Neon number too low/ Neon Lamp intensity too low) |  |  | SWO-496129 | Defect | HW | Oxi module | Pinch valve |  | Oxy hardware error @Neon number too low/Neon Lamp intensity too low
No | Visit | Complaint | RPAC | Keys need loading onto thye machine | Went in to load the keys and have a look at the locations of the ABGH machines | looked around the machines then attempted to load the keys but no success. We need to get the keys again from Denmark. | SWO-492861 | Defect | SW | RMED SW | Key | SW key | Keys need loading onto thye machine | Went in to load the keys and have a look at the locations of the ABGH machines | looked around the machines then attempted to load the keys but no success. We need to get the keys again from Denmark.
No | 3 | Y | RCZE | Blank |  |  | D5A259BBDC33AFE6 | Defect | SW | RMED SW | RMED SW | Wrong key database | Blank
No | 4 | Y | RCZE | Blank |  |  | D5AFF221EE3D8B79 | Defect | HW | Analyzer Control Module | CF card |  | Blank
No | Visit | Complaint | RAME | Hardware/Software issues | When analyzer boots up: Status Code 2 error. When loading WIN 7: Unable to create partition info. When loading Roma application software: Installation error refer to d: $install log and General installation error. | Replaced Compact Flash card and Central Processing Unit. Loaded WIN 7, Roma application software and set-up. Connected analyzer to AQURE. | SWO-493429 | Defect | HW | DMS_PC | CPU | status code 2 | Hardware/Software issues | When analyzer boots up: Status Code 2 error. When loading WIN 7: Unable to create partition info. When loading Roma application software: Installation error refer to d: /install log and General installation error. | Replaced Compact Flash card and Central Processing Unit. Loaded WIN 7, Roma application software and set-up. Connected analyzer to AQURE.
No | Visit | Complaint | RPAC | Analyser turned off during sensor cassette install. Green fuses 3 bottom and 2 top flashing, sample mixer blue light also flashing but screen completely black. See attached video. DE advised to replace power supply. RMED CK advised not seen before but try CF card or power supply. | Replaced power supply with no improvement. Replaced CF card. Analyser boots up, no longer any flashing fuses but has Offline error with 1223 message: Analyser did not connect at start-up and left hand lower 1st fuse dead. Returned original CF card: Analyser now boots up on original card with same Offline error, so the problem is not CF card related. Replaced Wetsection board because I had no stock of fuses and the customer needed the analyser. All fuses now appear correct and analyser functioning normally. The failed board reproduces the error when installed in the backup analyser Problem is presumed to be due to blown fuse,but not yet tested to confirm. | Replaced Wetsection board with board taken out of backup. | SWO-493548 | Defect | HW | Basic Unit | PCB, wet section | 1223 | Analyser turned off during sensor cassette install. Green fuses 3 bottom and 2 top flashing, sample mixer blue light also flashing but screen completely black. See attached video. DE advised to replace power supply. RMED CK advised not seen before but try CF card or power supply. | Replaced power supply with no improvement. Replaced CF card. Analyser boots up, no longer any flashing fuses but has Offline error with 1223 message: Analyser did not connect at start-up and left hand lower 1st fuse dead. Returned original CF card: Analyser now boots up on original card with same Offline error, so the problem is not CF card related. Replaced Wetsection board because I had no stock of fuses and the customer needed the analyser. All fuses now appear correct and analyser functioning normally. The failed board reproduces the error when installed in the backup analyser Problem is presumed to be due to blown fuse,but not yet tested to confirm. | Replaced Wetsection board with board taken out of backup.
No | Visit | Complaint | RCHN | 打印机一工作仪器就重启 | 打印机一工作仪器就重启，散热风扇噪音大借903-001、636-523、580-415各一个 | 更换903-001、636-523、580-415后修复 by jesson | SWO-492895 | [5 empty] | The printer restarts as soon as the instrument is working. | The printer restarts as soon as the working instrument, and the cooling fan noise borrows 903-001, 636-523, 580-415 each | Repair by jesson after replacing 903-001, 636-523, 580-415
No | Visit | Complaint | RCHN | 抽吸样本失败 | 抽吸样本失败，借903-129一个 | 更换903-129后修复 by jesson | SWO-493035 | Defect | HW | Oxi module | Pinch valve | 1271 | Failed to pump samples | Failed to suck samples, borrow 903-129 one | Repair by jesson after replacing 903-129
No | Visit | Complaint | RCHN | 光学单元上的升降马达异常 | 光学单元上的升降马达异常（未更换备件） | 调整升降马达位置后修复 by jesson | SWO-493512 | Adjustment | HW | Process Module | Motor, process unit | 5784 | Abnormal lifting motor on optical unit | Abnormal lifting motor on optical unit (spare parts not replaced) | Repair by jesson after adjusting lift motor position
No | Visit | Complaint | RCHN | 报错5784，卡杯 | 报错5784，卡杯，借944-268（15495）一张 | 重新定标后修复 by jesson | SWO-492948 | Defect | Consumables | Test Kit | Cartridge, Broken | 5784 | 5784, Card Cup | 5784, card cup, loan 944-268 (15495) one | Fix by jesson after re-calibration
No | Visit | Complaint | RCHN | 定标底物外漏 | 定标底物外漏，借903-364一个 | 更换903-364后修复 | SWO-493516 | Defect | HW | Process Module | Optical Unit | 1024 | The outer leakage of the calibration substrate | The bid substrate leak, borrow 903-364 one | Repair after replacement 903-364
No | Visit | Complaint | RCHN | 测试卡断裂，报错5784 部位:teas | 测试卡断裂，报错5784 部位:teas | 更换测试卡后修复 by jesson | SWO-493562 | Defect | Consumables | Test Kit | Cartridge, Broken | Teeth | Test card break, report error 5784 site: teas | Test card break, report error 5784 site: teas | Fix by jesson after replacing test card
No | Visit | Complaint | RDEU | 04.02.2019; Fr. Hobert; Tel.: 02811062318; Oxy Spectrum mismatch error 04.02.2019; 09:45 Uhr; Fr. Robert; Tel.: 02811062318; weiterhin Oxy Spectrum mismach error | Oxi spektrum mismatch, SHb zu hoch, keine tHb-Messung möglich, QC-Fehler, | alle Fließwege gereinigt und mit Hypochlorid gespült, Ventil im Oximodul ausgebaut und gereinigt, Gerät kalibriert und QC durchgeführt, | SWO-494018 | Cleaned | Fluidic pathway | Oxi module | Oxi module | 581 | 04.02.2019; Ms. Hobert; Tel.: 02811062318; Oxy Spectrum mismatch error 04.02.2019; 9:45 a.m.; Ms. Robert; Tel.: 02811062318; continue Oxy Spectrum mismach error | Oxi spektrum mismatch, SHb zu hoch, keine tHb-Messung möglich, QC-Fehler, | all flow paths cleaned and rinsed with hypochloride, valve removed and cleaned in the oximodul, device calibrated and QC performed,
No | Visit | Complaint | RDEU | 03916714027 Pflg. Nico tHb Messung soll stark vom Laborergebniss abweichen | nichts defekt | Mehrere Vergleichsmessungen aus unterschiedlichen Proben: maximale Abweichung 0,2 mol $L , kein Gerätedefekt feststellbar | SWO-493920 | No fault found | No fault found | No fault found | tHb |  | 03916714027 Pflg. Nico tHb Measurement should differ greatly from the laboratory result | nothing defective | Several comparative measurements from different samples: maximum deviation 0.2 mol $L, no device defect detectable
No | Visit | Complaint | RDEU | Diverse pCO2-Fehler bei QK | Diverse QK-Fehler bei pCO2 zu hoch, häufig im Level A, seltener im Level B oder C | Prüfung Serviceprogramme, Vorkonditionierte SC90 eingesetzt, Startphase, SP90 ersetzt, Kal und QK ok | SWO-494357 | Assisted Replacement | SP | SP | pCO2 | QC A too high | Diverse pCO2-Fehler bei QK | Various QK errors at pCO2 too high, often at level A, less often in level B or C | Examination service programs, pre-conditioned SC90 used, start phase, SP90 replaced, Kal and QK ok
No | Visit | Complaint | RDEU | Schraube hinten ausgerissen | Kunstoff ausgebrochen (Transportschaden?) | Rückwand Monitor erneuert | SWO-494663 | Defect | HW | Display | Display | screws | Screw ripped out at the back | Artoff erupted (transport damage?) | Backpanel monitor renewed
No | Visit | Complaint | RPAC | Calibrations failing glucose 1310 response fault. | Time: 04 $02 $19 11:39 Replacement: Sensor cassette removed due to glucose 1310: response error since 31 $1 $19. Part #: 946-010 Run #: 676-60 Expiration date: 17 $04 $19 Remaining tests: 91 Chip ID: 47378140 Max. tests: 100 First inserted: 17 $01 $19 08:32 | Replaced sensor cassette. FOC consumable supplied: Sales order 38533 946-010R0688 | SWO-493640 | Assisted Replacement | SC | SC | Glu | 1310 | Calibrations failing glucose 1310 response fault. | Time: 04 /02 /19 11:39 Replacement: Sensor cassette removed due to glucose 1310: response error since 31 /1 /19. Part #: 946-010 Run #: 676-60 Expiration date: 17 /04 /19 Remaining tests: 91 Chip ID: 47378140 Max. tests: 100 First inserted: 17 /01 /19 08:32 | Replaced sensor cassette. FOC consumable supplied: Sales order 38533 946-010R0688
No | Visit | Complaint | RKOR | Solution pack install error | . | reinstall. | SWO-496800 | User Error | SP | SP | SP activation |  | Solution pack install error | . | reinstall.
No | Visit | Complaint | RLTD | Girly Andrada 0191 569 9051 | Bent probe causing inlet to not return to home position. On arrival the customer had straigtened probe but inlet still not returing home. | Entered service menu and ran inlet in all positions but it would not calibrate. Removed inlet module and reseated in correct position. Inlet now moving freely. Analyser went into 2 hour start up. | SWO-494172 | Adjustment | HW | Auto Inlet module | Auto Inlet module | Probe bend | Girly Andrada 0191 569 9051 | Bent probe causing inlet to not return to home position. On arrival the customer had straigtened probe but inlet still not returing home. | Entered service menu and ran inlet in all positions but it would not calibrate. Removed inlet module and reseated in correct position. Inlet now moving freely. Analyser went into 2 hour start up.
No | Visit | Complaint | RDEU | 01.02.2019 Pfl. Steffen 03477171233 meldet Luftblasen in der Probe. 01.02.2019 16:40Uhr Herr Stephan 03477171233: das Gerät meldet weiterhin Luftblasen erkannt | Gerätzieht pumpt Lösung LP meldet zu wenig Lsg 2 $ wurde erst vor 3 Tagen eingesetzt | Fehlersuche Schlauch im Gerät abgesprungen, diesen wieder richtig angebracht Lösungen werden richtig durchgepumpt keine Kal mgl, da LP zu wenig Zyklen Lsg 2 keine neuer LP vor Ort zu erneuern | SWO-493939 | [5 empty] | 01.02.2019 Pfl. Steffen 03477171233 reports air bubbles in the sample. 01.02.2019 16:40Mr Stephan 03477171233: the device continues to report air bubbles detected | Device pulls solution LP reports too little Lsg 2 was used only 3 days ago | Troubleshooting hose in the device, it is properly attached again Solutions are properly pumped through no cal mgl, as LP too few cycles Lsg 2 no new LPs to renew on site
Yes | Visit | Complaint | RAME | Hours on phone and problem still there. Also won't recognize inlet flaps | No correct analyzers. Please refer to 754R2521N007 | Issue was resolved under correct analyzer 754R2521N007. | SWO-494179 | Re classified RMED | Fault in creating this SWO |  | Correct SWO: SWO-494181 |  | Hours on phone and problem still there. Also won't recognize inlet flaps | No correct analyzers. Please refer to 754R2521N007 | Issue was resolved under correct analyzer 754R2521N007.
No | Visit | Complaint | RAME | Analyzer is leaking from inlet | Analyzer is leaking from inlet gasket and pH $BG module. Waste pump damaged. | Removed all electrodes from pH $BG module and cleaned chamber. Replaced missing pO2 ring. Cleaned around inlet and lower fluidic module. Waste pump was flooded and was not functioning properly. Replaced waste pump. Flushed waste line. Removed inlet module and Autocheck module to remove clot that was by waste connector. Replaced lactate membrane and Cal 2 for customer. Replaced sealing ring on lactate electrode. Inlet tubing was corroded, replaced that as well. | SWO-494181 | Defect | HW | Fluid manifold | Waste tube | Clot | Analyzer is leaking from inlet | Analyzer is leaking from inlet gasket and pH /BG module. Waste pump damaged. | Removed all electrodes from pH $BG module and cleaned chamber. Replaced missing pO2 ring. Cleaned around inlet and lower fluidic module. Waste pump was flooded and was not functioning properly. Replaced waste pump. Flushed waste line. Removed inlet module and Autocheck module to remove clot that was by waste connector. Replaced lactate membrane and Cal 2 for customer. Replaced sealing ring on lactate electrode. Inlet tubing was corroded, replaced that as well.
No | Visit | Complaint | RDAN | Internal Barcode reader skal skiftes, da denne ikke ar OK. Kontakt Lene og aftal med hvornår det passer ind at skifte denne. 903-670 | . | Intern scanner udskiftet, slutafprøvet. | SWO-493799 | Defect | HW | Basic Unit | Barcode reader |  | Internal Barcode reader must be changed as this does not AR OK. Contact Lene and negotiate with when it will fit to change this. 903-670 | . | Internal scanner replaced, final tested.
No | Visit | Complaint | RIND | AQT90 system hang & error. | Check the system. | Check the incubation temp. pct temp error. Check the power supply exhaust . clean the fan & guide customer clean spaces for system. | SWO-494032 | Cleaned | HW | Cover and Chassis | Power supply | Fan | AQT90 system hang & error. | Check the system. | Check the incubation temp. pct temp error. Check the power supply exhaust . clean the fan & guide customer clean spaces for system.
No | Visit | Complaint | RBNL | Clip d'inlet est disparu | Clip d'inlet est disparu l'aiguile est courbé | Nouvelle Clip d'inlet & aiguile | SWO-494418 | Defect | HW | Inlet module | Probe bend |  | Inlet clip is gone | Inlet clip is gone the sharp is bent | New Inlet Clip - Sharp
No | Visit | Complaint | RKK | 代替インターフェイス入れ替え | Cl Cal バックログエラー | センサーインターフェイスPCB交換 | SWO-498343 | Defect | HW | Basic Unit | PCB, Sensor interface | Cl | Alternate interface replacement | Cl Cal Backlog Error | Sensor Interface PCB Replacement
No | Visit | Complaint | RAME | 1184 - and 1342 Leak detected | Getting errors 1184 and 1342. Pinch not working properly. | Replaced the pinch valve. Ran calibrations | SWO-495119 | Defect | HW | Oxi module | Pinch valve | 1184; 1342 | 1184 - and 1342 Leak detected | Getting errors 1184 and 1342. Pinch not working properly. | Replaced the pinch valve. Ran calibrations
No | Visit | Complaint | RDEU | Hotline - spülen des Systems erforderlich | Verdacht auf Gerinsel | ABL im Serviceprogramm überprüft und gespült, Startphase ausgeführt, Funktion OK. | SWO-494969 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway | Clot | Hotline - flushing the system required | Suspicion of Gerinsel | ABL in the service program checked and rinsed, start phase executed, function OK.
No | Visit | Complaint | RLTD | Barakat - 07557142073 | Possible clot in fluidic pathway. | Manually flushed and scrubed line, clot removed, restarted analyser all okay. | SWO-494457 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway | Clot | Barakat - 07557142073 | Possible clot in fluidic pathway. | Manually flushed and scrubed line, clot removed, restarted analyser all okay.
No | Visit | Complaint | RLTD | Jepsy 07414920186 | Analyser keeps failing due to leak detected error, pressure test failure 1063. Fault found to be small hole in the Oxi pinch valve tube. Replaced tube and ran pressure test - now passes. Restarted analyser and verified operation - all ok. | Replaced Oxi pinch valve tube. | SWO-494731 | Defect | HW | Oxi module | Pinch valve tube | 1063 | Jepsy 07414920186 | Analyser keeps failing due to leak detected error, pressure test failure 1063. Fault found to be small hole in the Oxi pinch valve tube. Replaced tube and ran pressure test - now passes. Restarted analyser and verified operation - all ok. | Replaced Oxygen Pinch valve tube.
No | Visit | Complaint | RAME | Broken Display(s) | Display screws are not secure in plastic standoffs as they are broken. | Replaced display and tested. Shipped display(s) to Westlake referencing CQC and Installation report. | SWO-494811 | Defect | HW | Display | Display | screws | Broken Display(s) | Display screws are not secure in plastic standoffs as they are broken. | Replaced display and tested. Shipped display(s) to Westlake referencing CQC and Installation report.
No | Visit | Complaint | RAME | Customer getting 0581 error during sampling | Analyzer has sample problems and error 0581 during samples. They have tried multiple tHb calibrations. | Replaced hemolyzer and cleaned pinch valve. Ran tHb calibration and verification | SWO-494794 | [5 empty] | Customer getting 0581 error during sampling | Analyzer has sample problems and error 0581 during samples. They have tried multiple tHb calibrations. | Replaced hemolyzer and cleaned pinch valve. Ran tHb calibration and verification
No | Visit | Complaint | RAME | Customer receiving 0537 Oxi Liquid Sensor Error after replacing rinse. | Analyzer is failing Level 1 QC for Crea and receiving 0537 Oxi Liquid Sensor Error | Ran 2 Protein Removal, Cleaning, Liquid Sensor Adjust, and pump calibration. Replaced reference membrane. Replaced Crea membranes. Analyzer came in for calibration and QC. After large tHb drift, ran 2 Protein Removals, 2 Cleaning and calibrations. Replaced pH sealing ring. Parameters were still drifting. Removed all electrodes, cleaned chamber with protein removal, water and dried out. Cleaned contacts and electrode caps. Went through history and noticed that issues started after inlet gasket replaced along with leakage detection. Replaced inlet gasket. | SWO-494185 | Defect | HW | Inlet module | Inlet gasket |  | Customer receiving 0537 Oxi Liquid Sensor Error after replacing rinse. | Analyzer is failing Level 1 QC for Crea and receiving 0537 Oxi Liquid Sensor Error | Ran 2 Protein Removal, Cleaning, Liquid Sensor Adjust, and pump calibration. Replaced reference membrane. Replaced Crea membranes. Analyzer came in for calibration and QC. After large tHb drift, ran 2 Protein Removals, 2 Cleaning and calibrations. Replaced pH sealing ring. Parameters were still drifting. Removed all electrodes, cleaned chamber with protein removal, water and dried out. Cleaned contacts and electrode caps. Went through history and noticed that issues started after inlet gasket replaced along with leakage detection. Replaced inlet gasket.
No | 4 | Y | RCZE | Blank | drifty na iontech |  | D5A325A8522257C6 | Assisted Replacement | Electrode | Electrode | Reference | CAPA-733 | Blank | Drifty na Iontech
No | Visit | Complaint | RFRA | tel : 01 40 94 28 00 demander le poste (7676) - Mme Ozenfant Message d'erreur depuis mercredi error : 0532 eL / Met lower liquid sensor error Cal et QC ok | Contrôle de la fluidique et des détecteurs de liquides dans le programme de service | Changement du détecteurs de liquides du bas de la chambre El $Met. Essais | SWO-494430 | Defect | HW | El/Met Module | LS | 532 | 01 40 94 28 00 apply for the position (7676) - Ms. Ozenfant Error message since Wednesday error: 0532 eL / Met lower liquid sensor error Cal and QC ok | Fluidic and liquid detector control in the service program | Changing the liquid detectors at the bottom of the El $Met chamber. Tests
No | Visit | Complaint | RPAC | Parameter not installed- 604 - fhbf, tbill | Parameter not installed Fhbf and tbill | Travelled to site. Demo oxi30 key expired. Updated option keys and restarted. Performed 2 point cal and qcs - ok | SWO-494227 | User Error, FSE | Demo key | Demo key | Demo key |  | Parameter not installed- 604 - fhbf, tbill | Parameter not installed Fhbf and tbill | Travelled to site. Demo oxi30 key expired. Updated option keys and restarted. Performed 2 point cal and qcs - ok
No | Visit | Complaint | RDEU | Zeitweise Elektrolytdriften, Rilibäkverletzungen pCO2 | Zeitweise Elektrolytdriften | Gerät überprüft, HJW durchgeführt, E1001 ersetzt | SWO-494987 | Assisted Replacement | Electrode | Electrode | Reference | CAPA-733 | Temporary electrolyte drifts, rileb injuries pCO2 | Temporary electrolyte drifts | Device checked, HJW performed, E1001 replaced
No | Visit | Complaint | RDAN | Dorte klager over nogleQC,er fra den 31/12/2018, hvor pco2 rammer helt galt. | Det viser sig | Gas $luft ventil i pH $blodgas modulet udskiftet, slutafprøvet. | SWO-494438 | Defect | HW | pH/BG Module | Valve, air | pCO2 | Dorte complains of nogleQC, is from the 31/12/2018, where pCO2 hits the wrong way. | It turns out | Gas $luft valve in pH $blodgas module replaced, final tested.
No | Visit | Complaint | RDEU | Hotline - Temp Fehler Inlet Modul 05.02.2019 12:30 H. Böckmann 05971424229: Fehler bleibt nach Restart bestehen; FSE-Einsatz | Inlet Temp Fehler | Gerät überprüft, Einlassmodul abgebaut und überprüft. Temperaturreglung im Service geprüft - ztw. 36,5 Grad , untere Grenze. Einlass mehrmals mit Wasser $ Hypochlorid gespült. Ersatzteile bestellt. | SWO-494977 | Revisit required | HW | Inlet module | Preheater |  | Hotline - Temp Error Inlet Module 05.02.2019 12:30 H. Böckmann 05971424229: Error remains after restart; FSE deployment | Inlet Temp Error | device checked, inlet module dismantled and checked. Temperature control tested in service - ztw. 36.5 degrees , lower limit. Inlet rinsed several times with water - hypochloride. spare parts ordered.
No | Visit | Complaint | RIND | Inlet door not closing properly. | Inlet door not closing properly. | Fixed door wheel C-Pin with tape,resolve the issue temporarily,Need to replace the Cartridge slider for further check. | SWO-496729 | Defect | HW | Cover and Chassis | Door, Inlet |  | Inlet door not closing properly. | Inlet door not closing properly. | Fixed door wheel C-Pin with tape,resolve the issue temporarily,Need to replace the Cartridge slider for further check.
No | Visit | Complaint | RPAC | Lun in the lab said the AQT90 had an Error message so he reboot the analyser. It now says Not Operational. I told him to leave it until it is in Ready mode. He will call back if it doesn't go to Ready. Lun called back and has reboot again and it appears that the sliding door is stuck and keeps going into Error. Can someone please follow up and call the lab on Monday morning. Customer is happy to wait until then. | Inlet doors not closing correctly causing error | Removed inlet door mechanism and replaced with spare. Also removed pc and replaced 24v b fuse. Taken back up, ran blank test and lqc - ok | SWO-494259 | Defect | HW | Cover and Chassis | Door, Inlet |  | Lun in the lab said the AQT90 had an Error message so he reboot the analyser. It now says Not Operational. I told him to leave it until it is in Ready mode. He will call back if it doesn't go to Ready. Lun called back and has reboot again and it appears that the sliding door is stuck and keeps going into Error. Can someone please follow up and call the lab on Monday morning. Customer is happy to wait until then. | Inlet doors not closing correctly causing error | Removed inlet door mechanism and replaced with spare. Also removed pc and replaced 24v b fuse. Taken back up, ran blank test and lqc - ok
No | Visit | Complaint | RPOL | Aparat porzuca próbki | Przegląd logów aparatu. | Wymiana uszczelki wlotu, odbiałczanie, regulacja LS, inletu oraz kalibracja pomp. Testy serwisowe. | SWO-497109 | Defect | HW | Inlet module | Inlet gasket | 0331 | The camera drops the sample | Overview of the camera logs. | Inlet gasket replacement, dewhiting, LS adjustment, inlet and pump calibration. Service tests.
Yes | Visit | Complaint | RDEU | Gerät fehlerhaft bei Installation | Gerät verlangt immer SP Wechsel. Startphase wird abgebrochen | SP Wechsel bringt keinen Erfolg. Fehler taucht nach 0,5 Std wieder auf. Überprüfung Flüssigkeitswege,ok. ..Überprüfung Valve ok.....Einlass Gasket ersetzt.......Fluidic Check zeigt LS und Pump Kal sowie PressureTest failt. BG LS zeigt.. 254 über GW LS ersetzt. Fluidic Check wird nun akzeptiert. Startphase abgewartet. Kalibration und QC durchgeführt und in Ordnung Probelauf. | SWO-496794 | Re classified RMED | installation report |  |  | Move to OOB trending | Device faulty during installation | Device always requires SP change. Start-up phase is canceled | SP change is not a success. Error reappears after 0.5 hrs. Checking fluid pathways,ok. .. Check Valve ok..... Inlet Gasket replaced....... Fluidic Check shows LS and Pump Kal as well as PressureTest failt. BG LS shows.. 254 via GW LS. Fluidic Check is now accepted. start-up phase. Calibration and QC carried out and in order Rehearsal.
No | Visit | Complaint | RBNL | Vele tbil onder rapporteergrens (<-2.1) | Vele tbil error 094 - onder rapporteergrens (<-2.1) | Reeds tHb cal uitgevoerd vandaag. service test hemolyzer - ok Service test algemeen - ok calibratie -ok qc - ok teststaal - ok Service dump genomen voor eventueel vervolg. | SWO-495147 | Revisit required | HW | Oxi module | Hemolyzer | Bili | Many Tbil under report limit (<-2.1) | Many Tbil error 094-under report limit (<-2.1) | Already running tHb Cal today. Service Test Hemolyzer-OK Service Test General-OK Calibration-OK QC-OK Test Steel-OK Service dump taken for any continuation.
Yes | Visit | Complaint | RDEU | Gerät fehlerhaft bei Installation | Gerät verlangt immer SP Wechsel. Startphase wird abgebrochen | SP Wechsel bringt keinen Erfolg. Fehler taucht nach 0,5 Std wieder auf. Überprüfung Flüssigkeitswege,ok. ..Überprüfung Valve ok.....Einlass Gasket ersetzt.......Fluidic Check zeigt LS und Pump Kal sowie PressureTest failt. BG LS zeigt.. 254 über GW LS ersetzt. Fluidic Check wird nun akzeptiert. Startphase abgewartet. Kalibration und QC durchgeführt und in Ordnung Probelauf. | SWO-496793 | Re classified RMED | installation report |  |  | Move to OOB trending | Device faulty during installation | Device always requires SP change. Start-up phase is canceled | SP change is not a success. Error reappears after 0.5 hrs. Checking fluid pathways,ok. .. Check Valve ok..... Inlet Gasket replaced....... Fluidic Check shows LS and Pump Kal as well as PressureTest failt. BG LS shows.. 254 via GW LS. Fluidic Check is now accepted. start-up phase. Calibration and QC carried out and in order Rehearsal.
No | Visit | Complaint | RCHN | 接口板腐蚀，导致温度不稳定， | 接口板腐蚀，导致温度不稳定， | 更换902-972后修复 by Andrew | SWO-502054 | [5 empty] | Interface plate corrosion, resulting in temperature instability, | Interface plate corrosion, resulting in temperature instability, | Repair by Andrew after replacement 902-972
No | Visit | Complaint | RPOL | Brak detekcji SP | Weryfikacja statusu analizatora, nieprawidłowy odczyt pamięci SP. Analizator zalany z powodu zsunięcia się węża przy wejściu do SC. | Wymiana płytki ze stykiem od odczytu pamięci SP (903-619). Wymiana sensora optycznego od drzwi SP(903-628). Zalanie analizatora spowodowało uszkodzenie 2 pakietów odczynnikowych SP oraz kasety sensorowej SC. Dwa pakiety odczynnikowe SP zostały dostarczone w ramach reklamacji przez inżyniera serwisu. | SWO-495095 | [5 empty] | No detection SP | Analyzer status check, incorrect SP memory reading. The analyser floated due to the hose being dried at the entrance to the SC. | Replacing the PCB with the SP memory reading contact (903-619). Optical sensor replacement from door SP (903-628). Flooding of the analyser resulted in damage to the 2 SP reagent packs and the SC sensor cassette. Two SP reagent packets were delivered as part of the complaint by the service engineer.
No | Visit | Complaint | RIND | Oximetry hardware problem. Possible to measure | Checked the data logs Oximetry hardware problem. Possible to measure found lamp unit test was getting failed | Replaced lamp unit performed complete oximetry test & 2 point calibration done processed the sample result was ok | SWO-494983 | Defect | HW | Oxi module | Lamp unit |  | Oximetry hardware problem. Possible to measure | Checked the data logs Oximetry hardware problem. Possible to measure found lamp unit test was getting failed | Replaced lamp unit performed complete oximetry test & 2 point calibration done processed the sample result was ok
No | Visit | Complaint | RFRA | Gaz1 remplacé tous les 10 jours et Gaz2 tous les mois | Aucune | Remplacement du dual valve housing et de la membrane d'air du module pH. Nettoyage des vannes de tous les module. | SWO-496762 | Defect | HW | pH/BG Module | Valve, air |  | Gas1 replaced every 10 days and Gaz2 every month | No | Replacement of the dual valve housing and the air membrane of the pH module. Cleaning the valves of all the module.
No | Visit | Complaint | RDEU | Fehler:923, 539 El-Met pumpenfehler, | Nadel mehrfach verbogen und gerichtet, El-Met Pumpenschlauch undicht und verkrustet | Im Zuge der halbjahreswartung Schläuche und vorhandene Nadel gewechselt. Und auf neu Bestellung hingewiesen. Interne Tests Kalibration und QC bestanden | SWO-497166 | Defect | HW | Inlet module | Probe bend |  | Error:923, 539 El-Met pump error, | Needle bent and directed multiple times, El-Met pump hose leaks and encrusted | In the course of the half-year maintenance hoses and existing needle changed. And pointed to new order. Internal tests calibration and QC passed
No | Visit | Complaint | RDEU | 06.02.2019 08:00Uhr Pfl Michael 0911398117630: Gerät bootet nicht | Status Error Code 10, astm.ini beschädigt | Fehlersuche, Datensicherung, Network Connection geprüft, Sicherung, astm.ini eingespielt, Reboot ok Startphase bestanden, Mehrfach Kal. (pCO2 und pO2 Kal. Error) System-und Funktionsprüfung i.O. | SWO-495105 | Defect | SW | RMED SW | astm.ini | status code 10 | 06.02.2019 08:00H Pfl Michael 0911398117630: Device does not boot | Status Error Code 10, astm.ini beschädigt | Troubleshooting, data backup, network connection checked, backup, astm.ini recorded, reboot ok Start phase passed, Multiple Kal. (pCO2 and pO2 Kal. Error) System and functional check in the end
No | Visit | Complaint | RAME | Labels on El/MET module all say blood gas parameters & are incorrectly labeled. | Inspection of the system revealed the El Met module labels only are incorrect. The hardware is functional and working well, just a labeling issue. | Put correct labels on El $Met module part# 981-625 | SWO-494908 | Defect | HW | El/Met Module | Labels | no picutre, pH/BG module labels on El/Met Module | Labels on El/MET module all say blood gas parameters & are incorrectly labeled. | Inspection of the system revealed the El Met module labels only are incorrect. The hardware is functional and working well, just a labeling issue. | Put correct labels on El $Met module part# 981-625
No | Visit | Complaint | RDEU | Hotline - Temp Fehler Inlet Modul 05.02.2019 12:30 H. Böckmann 05971424229: Fehler bleibt nach Restart bestehen; FSE-Einsatz | Inlet Temp. Fehler | Einlassmodul abgebaut und Vorwärmer erneuert. Nach Neustart diverse kal und QC gestartet. | SWO-495139 | Defect | HW | Inlet module | Preheater |  | Hotline - Temp Error Inlet Module 05.02.2019 12:30 H. Böckmann 05971424229: Error remains after restart; FSE deployment | Inlet Temp. Error | Inlet module dismantled and preheater srenewed. After rebooting various kal and QC started.
No | Visit | Complaint | RDEU | 06.02.2019; Fr. Führbach; Tel.: 020830542050; Probenabbrüche, ungenügende Probe | QC TnT nicht OK ( instabil) | Pumpe P1 aufgrund von untypischen betriebsgeräuschen erneuert, Needle Wash Unit im Umfang der HjW gereinigt, Flieswege mit HyperChlorid gereinigt, mehrer 6 Blind Messungen und QC im Umfang der HjW durchgeführt, DDimer und CRP OK. QC Ergebnisse oftmals außerhalb oberer QC Grenzen, Service Dump durchgeführt und ausgewertet, Diverse Parameter Grenzwertig, Flow ist ebenfalls Grenzwertig. Abspreche mit Anwender getroffen das über Nacht ausschließlich der AQT (R402N003) genutzt wird. | SWO-496302 | Revisit required | HW | Wet section | Wet section | Flow restriction; 1218 | 06.02.2019; Ms. Führbach; Tel.: 020830542050; Sample aborts, insufficient sample | QC TnT nicht OK ( instabil) | Pump P1 renewed due to atypical operating noises, Needle Wash Unit cleaned in the scope of HjW, flow paths cleaned with hyperchloride, several 6 blind measurements and QC performed in the scope of HjW, DDimer and CRP OK. QC results often performed outside upper QC limits, service dump performed and evaluated, various parameters limitable, flow is also limitous. Agreed with users that only the AQT (R402N003) is used overnight.
No | Visit | Complaint | RAME | Multiple replacements w/ multiple parameter failures | Instrument has experienced leak currents, pump calibration & chip read errors since start up. | Customer replaced SP but errors still intermittently are occurring. Replaced SC and instrument startup was completed without error. Calibrated and ran QC, instrument is operating normally. | SWO-495537 | Assisted Replacement | SC | SC | SC |  | Multiple replacements w/ multiple parameter failures | Instrument has experienced leak currents, pump calibration & chip read errors since start up. | Customer replaced SP but errors still intermittently are occurring. Replaced SC and instrument startup was completed without error. Calibrated and ran QC, instrument is operating normally.
No | Visit | Complaint | RBNL | Analyzer geeft sinds vannacht 0:00uur fout 1063 druk test vacuümfout aan. Al inletgasket vervangen en doorgespoeld demiwater. | Pack HA-62 gewisseld met naast staande instrument blijft hier vacuüm druk fout geven. Dit pack op andere analyzer ook vacuüm druk fout. Terugplaatsen geeft weer problemen met lezen PackID en vacuüm druk fout. Andere analyzer direct goed. | Nieuw solution pack ID pcb gemonteerd. Nieuw solution pack aangesloten BU-56. | SWO-495818 | Assisted Replacement | SP | SP | FT | 1063 | Analyzer has been giving 0:00hrs error 1063 pressure test vacuum error since last night. All inletgasket replaced and washed down demi water. | Pack HA-62 exchanged with next standing instrument, vacuum pressure is still wrong here. This pack on other analyzer also vacuum printing error. Repositioning again gives problems with read PackID and vacuum printing error. Other analyzer immediately good. | New Solution Pack ID PCB mounted. New Solution Pack connected BU-56.
No | Visit | Complaint | RCHN | 气流温度超范围 | 气流温度超范围，借847-607一个 | 更换847-607后修复by tim | SWO-495018 | Defect | HW | Process Module | Filter, cup dryer | Flow | Airflow temperature over range | Airflow temperature over range, 847-607 one | Repair by tim after replacing 847-607
No | Visit | Complaint | RCHN | 测试卡断裂（部位teeth） | 测试卡断裂（部位teeth）,未更换备件 | 更换测试卡后修复by tim | SWO-495895 | Defect | Consumables | Test Kit | Cartridge, Broken | Teeth | Test card break (site teeth) | Test card break (site teeth), spare parts not replaced | Fix by tim after replacing test card
No | 47 | Y | RCZE | Blank |  |  | D5A6E4CCB40FB00F | Defect | HW | Oxi module | Hemolyzer |  | Blank
No | Visit | Complaint | RDEU | Proben Fehler | Proben Fehler, Proben Mischerfehler | Autotray adjustiert, Probenmischer adjustiert, Kalibrationen und QC's vorgenommen | SWO-495930 | Adjustment | HW | FlexQ Module | Mixer Tray |  | Samples error | Samples Errors, Samples Mixer Errors | Autotray adjusted, sample mixer adjusted, calibrations and QC's performed
No | Visit | Complaint | RDAN | Drift på thb ved 1 punkt og 2 punkt kal. drift på 19-20. periodisk ved at gentage kalibreringen flere gange forsvinder driften. Kontroller ude på OXI parameter. (efter OK kalibrering er kontrollerne også OK) Ved komplete OXI test fejler lampe med ustabil halogen lampe. Problemet bkev ikke løst af gentagende lampe justeringer. Skiftet Lamp Unit (Joe's lager . ny bestilt FOC til JOE brugt 902-458R2984N006) Efterfølgende Lampe justering, complete OXI, Kalibrering og Kontroller OK | Drift på thb ved 1 punkt og 2 punkt kal. drift på 19-20. periodisk ved at gentage kalibreringen flere gange forsvinder driften. Kontroller ude på OXI parameter. (efter OK kalibrering er kontrollerne også OK) Ved komplete OXI test fejler lampe med ustabil halogen lampe. | Skiftet Lampe unit 902-458 Skiftet Lamp Unit (Joe's lager . ny bestilt FOC til JOE brugt 902-458R2984N006) Efterfølgende Lampe justering, complete OXI, Kalibrering og Kontroller OK | SWO-495742 | Defect | HW | Oxi module | Lamp unit |  | Operation at THB at 1 point and 2 point Kal. Operation of 19-20. Periodically by repeating the calibration several times the operation disappears. Check out on the OXI parameter. (After OK calibration, the controls are also OK) When Komplete KONTROL OXI test fails lamp with unstable halogen lamp. The problem bkev is not solved by repeating lamp adjustments. Changed Lamp Unit (Joe's Lager. New ordered FOC for JOE used 902-458R2984N006) Subsequent lamp alignment, complete OXI, calibration and controls OK | Operation at THB at 1 point and 2 point Kal. Operation of 19-20. Periodically by repeating the calibration several times the operation disappears. Check out on the OXI parameter. (After OK calibration, the controls are also OK) When Komplete KONTROL OXI test fails lamp with unstable halogen lamp. | Changed lamp Unit 902-458 switched Lamp Unit (Joe's Lager. New ordered FOC for JOE used 902-458R2984N006) Subsequent lamp alignment, complete OXI, calibration and controls OK
No | Visit | Complaint | RDAN | ABL'en taber dato når den slukkes. - den har været i kasse fra den 26.11.2018 og frem til nu - her har den tabt dato of tid. - nu når den slukkes taber den dato men ikke tiden ? Se vedhæftet billeder. | Tændte og slukkede .... abl'en huskede tid men ikke dato. | skiftede tid dato og time zone. slukkede og tændte abl'en skiftede tid til korrekt tændte og slukkede. | SWO-496242 | Defect | HW | DMS_PC | DMS_PC | Time error | The ABL loses date when it is switched off. -It has been in the box from the 26.11.2018 up to now-here has the lost date of time. -Now when it turns off loses that date but not the time? See attached pictures. | Lighted and extinguished.... The ABL remembered time but not date. | Changed time date and time zone. Switched off and lit the ABL changed time to properly lit and extinguished.
No | Visit | Complaint | RAME | Po2 electrode failure during CAS validation 3.4 sensitivity | During CAS validation Po2 was failing sensitivity (3.40) after all 2pt. Calibrations. | Swapped Po2 electrode with another analyzer to see if failure followed electrode and it did follow the electrode. This electrode is in a new install. Electrode E799 date code is BX-04. *Please approve replacement* | SWO-496020 | Assisted Replacement | Electrode | Electrode | pO2 |  | Po2 electrode failure during CAS validation 3.4 sensitivity | During CAS validation Po2 was failing sensitivity (3.40) after all 2pt. Calibrations. | Swapped Po2 electrode with another analyzer to see if failure followed electrode and it did follow the electrode. This electrode is in a new install. Electrode E799 date code is BX-04. *Please approve replacement*
No | Visit | Complaint | RDEU | 06.02.2019; Fr. Führbach; Tel.: 020830542050; Probenabbrüche, ungenügende Probe | Flieswege verschmutzt, TNT Kal Material Probleme | Jahreswartungskit verbaut, Optische Einheit ausgebaut um das Process Wheel mit Isopropoyl Alkohol zu reinigen. Mehrfach TNT QC erneuert gemessen, Werte sind in der Toleranz, jedoch nahe der oberen Herstellergrenze ( 19,9 - 39,5). TNT mit Lot: 15526 neu kalibriert, TNT QC mehrfach erneut gemessen. TNT Werte liegen liegen Nähe Sollwert. Blind Messungen und Service Dump durchgeführt. Alles OK. | SWO-496303 | Cleaned | HW | Wet section | Wet section | Flow restriction; 1218 | 06.02.2019; Ms. Führbach; Tel.: 020830542050; Sample aborts, insufficient sample | Flow ways polluted, TNT Kal material problems | Annual maintenance kit installed, optical unit removed to clean the Process Wheel with isopropoyl alcohol. Multiple TNT QC renewed measured, values are measured in tolerance, but close to the upper manufacturer limit ( 19.9 - 39.5). TNT with lot: 15526 recalibrated, TNT QC measured again several times. TNT values are close to setpoint. Blind measurements and service dump performed. All OK.
No | Visit | Complaint | RIND | Calibration adjustment background outlier check failed | Calibration failed | Needto change optical unit, Changed Optical unit , Required to calibrated all parameter but TnI Calibration is missinig . Need to change the Lot of TnI. | SWO-498570 | Defect | HW | Process Module | Optical Unit | 1024 | Calibration adjustment background outlier check failed | Calibration failed | Needto change optical unit, Changed Optical unit , Required to calibrated all parameter but TnI Calibration is missinig . Need to change the Lot of TnI.
No | Visit | Complaint | RAME | Analyzer asking for solution pack replacement. | Analyzer asking for solution pack replacement. Replaced solution pack, but then analyzer asked for second replacement. Upon arrival, analyzer was successfully using solution pack, but now failing Glucose for status and sensitivity. | Analyzer will need new sensor cassette. Cassettes on order by customer. Analyzers in process of validation. Will follow up when cassette replaced. | SWO-496596 | Assisted Replacement | SP | SP | FT |  | Analyzer asking for solution pack replacement. | Analyzer asking for solution pack replacement. Replaced solution pack, but then analyzer asked for second replacement. Upon arrival, analyzer was successfully using solution pack, but now failing Glucose for status and sensitivity. | Analyzer will need new sensor cassette. Cassettes on order by customer. Analyzers in process of validation. Will follow up when cassette replaced.
No | Visit | Complaint | RKK | オキシスペクトルが一致しません | オキシスペクトルが一致しません | ヘモライザーユニット不良の為交換。OXIバルブチューブ劣化の為、交換。OXIコンプリートテスト及び各部調整実施。作業後、Lac応答エラーの為、SC交換。 | SWO-498323 | Defect | HW | Oxi module | Hemolyzer | 581 | Oxyspectrum doesn't match | Oxyspectrum doesn't match | Replacement for malfunction of the hemolizer unit. Replacement due to OXI valve tube deterioration. OXI complete test and each part adjustment. After work, lac response error, SC exchange.
No | Visit | Complaint | RDEU | 08.02.2019, Herr Hirschel, Tel: 021513341343: Leckage erkannt | Gerät verstopft. | Gerät mehrfach durchgespült, Pumpen Kal, LS Justage durchgeführt, | SWO-497205 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway | Clot | 08.02.2019, Mr. Hirschel, Tel: 021513341343: Leakage detected | device clogged. | device repeatedly rinsed, pumps Kal, LS adjustment performed,
No | Visit | Complaint | RLTD | Contact Jepsy 07414 920186 | Fault traced to damaged inlet gasket caused by bent inlet probe. Inlewt probe had already been corrected by customer, but inlet gasket had not been replaced. Fitted new inlet gasket, analyswer now passes the pressure test and is working. | Replaced faulty inlet gasket. | SWO-496599 | Defect | HW | Auto Inlet module | Probe bend | 1063 | Contact Jepsy 07414 920186 | Fault traced to damaged inlet gasket caused by bent inlet probe. Inlewt probe had already been corrected by customer, but inlet gasket had not been replaced. Fitted new inlet gasket, analyswer now passes the pressure test and is working. | Replaced faulty inlet gasket.
No | Visit | Complaint | RBNL | Vele tbil onder rapporteergrens (<-2.1) | Ondanks controle van de hemolyzer, pinch valve en tHb cal. Is het fenomeen, in mindere mate nog steeds aanwezig. | Vervangen pinch valve & hemolyzer + thb kal. Kalibratie - ok QC - ok | SWO-496298 | Defect | HW | Oxi module | Hemolyzer | Bili | Many Tbil under report limit (<-2.1) | Despite control of the Hemolyzer, pinch valve and tHb cal. Is the phenomenon, to a lesser extent still present. | Vervangen pinch valve & hemolyzer + thb kal. Kalibratie - ok QC - ok
No | Visit | Complaint | RHUN | PM esedékes lassan, és egyben panasz is van | a waste oldat csatlakozónál szivárgás $folyás tapasztalható. Waste oldat csatlakozón a tömítés rosszul áll | készülék általános tisztítása, zéró áramok ellenőrzése, waste oldat csatlakozó tömítések cseréje, készülék init, rendszer karbantartás | SWO-496330 | [5 empty] | PM is due slowly and also has a complaint | Leakage /foly ation of the waste solution connector. The seal is poorly attached to the waste solution | General cleaning of the appliance, control of zero currents, waste solution Connector seals replacement, appliance init, system maintenance
No | Visit | Complaint | RKK | 液漏れ 発生 | 液漏れ 発生 | 内部清掃実施 排液マニホールド部Oリング確認 SPラッチ位置調整 CAL確認 | SWO-497626 | [5 empty] | Leakage | Leakage | Internal cleaning O-ring confirmation of drainage manifold SP latch position adjustment CAL confirmation
No | Visit | Complaint | RCHN | 自动关机 | 自动关机，借902-972一个 | 更换902-972后修复by tim | SWO-495016 | [5 empty] | Auto-shutdown | Automatic shutdown, borrowing 902-972 one | Repair by tim after replacement 902-972
No | Visit | Complaint | RDEU | 07.02.2019 8:00 Pfl. Michael 03473962170: Kal und QC Fehler | QC Fehler K im Level 1 | Überprüfung Dateiarchiv Kal, QC und Aktivitäten Fehler begann mit LP Wechsel am 01.02. => von Lot 267578 auf Lot 269248 gewechselt => Sprung des QC Messwertes im Level 1 an obere Grenze Am 04.02. SC Wechsel => ohne Besserung => Reinstallierung SC am 05.,06. und 07.02. => keine Besserung Zwischenzeitlich am 06.02. weiteres LP installiert => gleiches Lot 269248 wie schon am 01.02. Durchführung Zero Kal => alle Werte stabil Service Test i.O. SC reinstalliert => Kal i.O. => QC Level 1 bleibt zunächst fehlerhaft zu hoch Nach Durchführung tHb-Kal K-Wert knapp innerhalb des QC Bereiches Vermutlich LP Chargen Problem Erstellung DataLogs und Meldung an RDEU Kunde informiert dass neues LP mit anderer Charge benötigt wird falls K-Fehler wieder auftritt | SWO-496469 | [5 empty] | 07.02.2019 8:00 Pfl. Michael 03473962170: Kal und QC Fehler | QC Fehler K im Level 1 | Review File Archive Kal, QC and Activities Error started with LP change on 01.02. => changed from lot 267578 to lot 269248 => Jump of the QC reading in level 1 to upper limit On 04.02. SC change => without improvement => reinstall SC on 05.,06. and 07.02. => no improvement In the meantime on 06.02. another LP installed => same lot 269248 as already on 01.02. Execution Zero Kal => all values stable Service Test i.O. SC reinstalled => Kal i.O. => QC Level 1 remains too high at first After carrying out tHb-Kal K-value just within the QC range Probably LP Batches Problem Creation of DataLogs and notification to RDEU Customer informs that new LP with different batch is needed if K error occurs again
No | Visit | Complaint | RAME | Possible fluid blockage errors | Unit leaked internally. Cleaning no help. Hemolyzer blocked. Liquid sensor stuck to hemolyzer. | Replaced hemolyzer and liquid sensor. Cleaned internal compartment. Performed service test and tHb calibration. | SWO-496662 | [5 empty] | Possible fluid blockage errors | Unit leaked internally. Cleaning no help. Hemolyzer blocked. Liquid sensor stuck to hemolyzer. | Replaced hemolyzer and liquid sensor. Cleaned internal compartment. Performed service test and tHb calibration.
No | Visit | Complaint | RAME | Oxi errors on patient samples | Pinch valve tubing was not set correctly in between pinch valve | Once I removed the wet section found tubing outside of pinch valve. Replaced it as tubing would keep popping out of place from pinch valve. Had to parts from PM kit. Reran system cycle, thb cals and 2 Levels of QC. All calibrations and 2 Levels of QC passed as per manufacturer specification. Ran patient sample to compare between two instruments and all was well. | SWO-496652 | [5 empty] | Oxi errors on patient samples | Pinch valve tubing was not set correctly in between pinch valve | Once I removed the wet section found tubing outside of pinch valve. Replaced it as tubing would keep popping out of place from pinch valve. Had to parts from PM kit. Reran system cycle, thb cals and 2 Levels of QC. All calibrations and 2 Levels of QC passed as per manufacturer specification. Ran patient sample to compare between two instruments and all was well.
No | Visit | Complaint | RDEU | Remote Zugriff wird im AQURE nicht korrekt dargestellt. | Remote Darstellung im Radiance nicht korrekt. | Remoteeinstellungen mit SalFa über Tel.-Hotline vor Ort überprüft. Optionkey deaktiviert und wieder aktiviert. -Kein Erfolg. ABL neu gestartet. - Kein Erfolg. Datensicherung erstellt. | SWO-497707 | Revisit required | Setup | Aqure | Aqure |  | Remote access is not displayed correctly in AQURE. | Remote display in the radiance is not correct. | Remote settings checked with SalFa via local telephone hotline. Optionkey deactivated and reactivated. -No success. ABL. - No success. data backup.
No | Visit | Complaint | RDAN | Nina klager over fejl 374 selvom hun har renset med hypoklorid. | Set blodprøve med tre bobler køre ind i pH $blodgas modulet. | Inlet heather, probe og slange udskiftet, slutafprøvet. | SWO-496292 | Defect | HW | Inlet module | Inlet probe | 374 | Nina complains of error 374 although she has cleaned with hypochlorite. | Seen blood test with three bubbles run into the pH /blodgas module. | Inlet heather, probe og slange udskiftet, slutafprøvet.
No | Visit | Complaint | RPOL | Problemy z uruchomieniem analizatora. | Weryfikacja statusu analizatora | Wykonanie czynności mających na celu uruchomienie analizatora. Usunięto zanieczyszczenia w układzie hydraulicznym. Wykonano testy przepływowe. Kalibracja, kontrola. Testowanie. Upgrade oprogramowania do v.3.4 | SWO-497217 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | Problems running analyzer. | Verifying analyzer Status | Perform actions to run the analyzer. Removed the impurities in the hydraulic system. Flow tests have been performed. Calibration, control. Testing. Software Upgrade to V. 3.4
No | Visit | Complaint | RKK | リーク検出 | ケース詳細の通り | 引上げ修理 | SWO-500349 | Defect | Analyzer removed | Analyzer removed | Analyzer removed | Leak current | Leak detection | As the case details | Raise repair
No | Visit | Complaint | RAME | Recurring fluid blockage. | No apparent blockage. Unit failed on different sensor and cal pack. | Replaced analog PCB. Reinstalled failing supplies. Performed service oxi cal and tHb cal. | SWO-496703 | [5 empty] | Recurring fluid blockage. | No apparent blockage. Unit failed on different sensor and cal pack. | Replaced analog PCB. Reinstalled failing supplies. Performed service oxi cal and tHb cal.
No | Visit | Complaint | RIND | Instrument switched off and not switching on | Instrument switched off and not switching on | Checked the main supply ,found no output through supply , connected the power card to another supply board and supply is fine here , switched on the instrument and run calibration and some samples and told customer to look after the main supply , customer happy with the service . | SWO-497214 | [5 empty] | Instrument switched off and not switching on | Instrument switched off and not switching on | Checked the main supply ,found no output through supply , connected the power card to another supply board and supply is fine here , switched on the instrument and run calibration and some samples and told customer to look after the main supply , customer happy with the se rvice .
No | Visit | Complaint | RCHN | 溶血器故障 | 溶血器故障，借902-980一个 |  | SWO-496828 | Defect | HW | Oxi module | Hemolyzer |  | Hemolysis failure | Hemolytic fault, borrow 902-980 one
No | Visit | Complaint | RCHN | BNP温度传感器错误 | BNP温度传感器错误，借905-908一个 |  | SWO-496835 | Defect | HW | Process Module | Sensor, BnP |  | BNP temperature sensor error | BNP temperature sensor error, borrow 905-908 one
No | Visit | Complaint | RCHN | 进样口旁门关不起来 | 进样口旁门关不起来，借903-028一个 | 更换903-028后修复by tim | SWO-496207 | Defect | HW | Cover and Chassis | PCB, cabinet top |  | The door next to the entrance can't be closed. | The entrance door can't be closed, borrow 903-028 one | Repair by tim after replacement 903-028
Yes | Visit | Complaint | RPAC | Leak current in relation to solution pack detected. | Arrived for basic user training of labour ward midwives and noticed error 964 (Leak current in relation to solution pack detected). | FOC Solution Pack supplied. | SWO-503519 | Re classified RMED | planned SWO |  |  | The visit was planned | Leak current in relation to solution pack detected. | Arrived for basic user training of labour ward midwives and noticed error 964 (Leak current in relation to solution pack detected). | FOC Solution Pack supplied.
No | Visit | Complaint | RLTD | Beth Thomas on 0161 2220979 | Site visit to investigate $repair analyser. | Arrived on site and attended analyser. Logged on and rechecked all components checked over the phone. No issues with inlet or around pinch valve. Tried manual flush and found big restriction. No fluid transport possible. Tested from inlet to waste and found restriction inside the hemolyzer. Powered down and used fine needle to manually break down the blockage. Then managed to flush a small amount of hypochlorite into hemolyzer. Left to soak for 15 minutes before flushing multiple times with hot water to remove broken down blockage. Pink tinge to fluid flushed through. Cleared hemolyzer and refitted all parts. Ran manual flush from waste to inlet whilst full fluid path connected. Fluid flushing fine now. Refitted covers and powered up. Ran sp replacement to refit pack. Waited for system temp to come up and analyser restarted. Analyser went into 2 hour startup. Waited for this to complete. Analyser started up fine, all cals and qcs passed with no issues. Analyser running well again. Customer enquired when their next lot of consumables were due to be delivered as they now had only this fitted to the analyser left. I chased this with the office and their delivery was brought forwards to accomodate. | SWO-497730 | Cleaned | HW | Oxi module | Hemolyzer | 508 | Beth Thomas on 0161 2220979 | Site visit to investigate /repair analyser. | Arrived on site and attended analyser. Logged on and rechecked all components checked over the phone. No issues with inlet or around pinch valve. Tried manual flush and found big restriction. No fluid transport possible. Tested from inlet to waste and found restriction inside the hemolyzer. Powered down and used fine needle to manually break down the blockage. Then managed to flush a small amount of hypochlorite into hemolyzer. Left to soak for 15 minutes before flushing multiple times with hot water to remove broken down blockage. Pink tinge to fluid flushed through. Cleared hemolyzer and refitted all parts. Ran manual flush from waste to inlet whilst full fluid path connected. Fluid flushing fine now. Refitted covers and powered up. Ran sp replacement to refit pack. Waited for system temp to come up and analyser restarted. Analyser went into 2 hour startup. Waited for this to complete. Analyser started up fine, all cals and qcs passed with no issues. Analyser running well again. Customer enquired when their next lot of consumables were due to be delivered as they now had only this fitted to the analyser left. I chased this with the office and their delivery was brought forwards to accomodate.
No | Visit | Complaint | RLTD | Noreen Tel: 01708 435000 x: 6845 | Blockage in the fluid channel. | Performed intensive flush with hypochlorite and water. Ran fluid transport tests and checks. | SWO-497348 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway | Clot | Noreen Tel: 01708 435000 x: 6845 | Blockage in the fluid channel. | Performed intensive flush with hypochlorite and water. Ran fluid transport tests and checks.
No | Visit | Complaint | RBNL | Mixer werkt niet | mixer gedeconnecteerd bij laatste bezoek. | mixer verbonden - test ok. | SWO-497149 | Defect | HW | Basic Unit | Mixer |  | Mixer works niet | Mixer disconnects on last visit. | Mixer Connected-Test OK.
No | Visit | Complaint | RDEU | Oximetrie Parameter können nicht kalibriert werden | Oximetrie Parameter sind inaktiv. Kalibration bei Hb = 0,0 | Auftrag vom Hitado. Sc wurde bereits getauscht. LP ist nicht vorhanden und hat nur noch wenig Zyklen. Aus dem Grund keine tHb-Kal möglich. LP wurde bestellt. Serviceeinsatz wird auf morgen vertagt. | SWO-497246 | [5 empty] | Oximetry parameters cannot be calibrated | Oximetry parameters are inactive. Calibration at Hb = 0.0 | Order from Hitado. Sc has already been swapped. LP is not present and has few cycles left. For this reason, no tHb-Kal is possible. LP has been ordered. Service operation will be postponed until tomorrow.
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 测试卡参比电极错误 1361 | 更换测试卡正常。 | SWO-497035 | [5 empty] | Test card, reagent package failure | Test card reference electrode error 1361 | The replacement test card is OK.
No | Visit | Complaint | RCHN | 干燥前温度超范围 | 干燥前温度超范围，借903-210一个 |  | SWO-496961 | Defect | HW | Process Module | PCB, cup dryer | 1068 | Overrange temperature before drying | Over-range temperature before drying, borrow 903-210 one
No | Visit | Complaint | RCHN | 光学敏感性超范围 1031 | 光学敏感性超范围 1031 | 更换903-364一个正常。（by chuanwei） | SWO-497823 | Defect | HW | Process Module | Optical Unit | 1031 | Optical sensitivity over range 1031 | Optical sensitivity over range 1031 | Replace 903-364 one normal. (by chuanwei)
No | Visit | Complaint | RIND | Service Visit. | Need to check the analyzer. | Check the analyzer, every thing is fine in the analyzer, only the dr having the doubt in sO2 because they co-relate the value from pO2, while in our analyzer the sO2 value is directly measure from the blood in oximetry module, discussed with the doctors regarding that, successfully done, analyzer handed over to staff in well working condition. | SWO-500203 | No fault found | No fault found | No fault found | No fault found |  | Service Visit. | Need to check the analyzer. | Check the analyzer, every thing is fine in the analyzer, only the dr having the doubt in sO2 because they co-relate the value from pO2, while in our analyzer the sO2 value is directly measure from the blood in oximetry module, discussed with the doctors regarding that , successfully done, analyzer handed over to staff in well working condition.
No | Visit | Complaint | RAME | Customer reports that inlet is leaking from syringe side of inlet | Customer reports that inlet is leaking from syringe side of inlet. Upon arrival, FSE ensured all tubing was connected. Tubing was connected with no visible obstructions. | Replaced the inlet gasket with part drawn from customer inventory. Op tested by running two 2-point calibrations and all three levels of QC. Op test satisfactory. Normal operations restored. | SWO-497424 | Defect | HW | Inlet module | Inlet gasket |  | Customer reports that inlet is leaking from syringe side of inlet | Customer reports that inlet is leaking from syringe side of inlet. Upon arrival, FSE ensured all tubing was connected. Tubing was connected with no visible obstructions. | Replaced the inlet gasket with part drawn from customer inventory. Op tested by running two 2-point calibrations and all three levels of QC. Op test satisfactory. Normal operations restored.
No | Visit | Complaint | RAME | Customer reports oximetry Failures on QC. No results reporting. | Customer reports oximetry Failures on QC. No results reporting. Upon arrival, analyzer had resumed passing QCs. To ensure issue was corrected, FSE visually inspected the analyzer. A large amount of leakage was observed in the pH $BG module stemming from both the Reference electrode and pH electrode bays. | Cleaned Reference and pH electrode bays. Replaced seal ring on pH electrode. Operation test satisfactory. Normal operations restored. | SWO-497426 | Defect | Electrode | Electrode | Sealing ring, pH |  | Customer reports oximetry Failures on QC. No results reporting. | Customer reports oximetry Failures on QC. No results reporting. Upon arrival, analyzer had resumed passing QCs. To ensure issue was corrected, FSE visually inspected the analyzer. A large amount of leakage was observed in the pH /BG module stemming from both the Reference electrode and pH electrode bays. | Cleaned Reference and pH electrode bays. Replaced seal ring on pH electrode. Operation test satisfactory. Normal operations restored.
No | Visit | Complaint | RIND | Instrument not switching on . | PC module replacement | Checked the instrument , replaced the pc module , switched on the instrument , checked the initialisation , instrument initialised ,checked calibration and QC , Calibration and QC passed. | SWO-497131 | Defect | HW | DMS_PC | CPU |  | Instrument not switching on . | PC module replacement | Checked the instrument , replaced the pc module , switched on the instrument , checked the initialisation , instrument initialised ,checked calibration and QC , Calibration and QC passed.
No | Visit | Complaint | RAME | Customer called stating that PH is failing calibrations. | Found that the analyzer is getting a sensitivity of 58% on the 2 pt calibration. | Replaced the PH electrode which corrected the issue. | SWO-497475 | Assisted Replacement | Electrode | Electrode | pH |  | Customer called stating that PH is failing calibrations. | Found that the analyzer is getting a sensitivity of 58% on the 2 pt calibration. | Replaced the PH electrode which corrected the issue.
No | Visit | Complaint | RDEU | DMS not connected | pCO2 Elektrode defekt | Elektrode erneuert, Kalibrierung und Qualitätskontrolle i. O. | SWO-497062 | Assisted Replacement | Electrode | Electrode | pCO2 |  | DMS not connected | pCO2 electrode defective | Electrode renewed, calibration and quality control as a standard
No | Visit | Complaint | RAME | Po2 electrode failure during CAS validation 3.4 sensitivity | Bad pO2 electrode | Replaced electrode. Good calibration. | SWO-497441 | Assisted Replacement | Electrode | Electrode | pO2 |  | Po2 electrode failure during CAS validation 3.4 sensitivity | Bad pO2 electrode | Replaced electrode. Good calibration.
No | Visit | Complaint | RAME | Blank screen | The screen was blank | I tried a power Supply and new screen but ended up replacing the CPU to resolve the issue. The Glu was still out on cal ask the customer to replace the SC if it will not come in | SWO-497661 | Defect | HW | DMS_PC | CPU | Blank screen | Blank screen | The screen was blank | I tried a power Supply and new screen but ended up replacing the CPU to resolve the issue. The Glu was still out on cal ask the customer to replace the SC if it will not come in
No | Visit | Complaint | RAME | Oximetry hardware during CAS | Oximetry hardware during CAS validation | Replaced Hemolyzer and pinch valve. Ran THB cal | SWO-498304 | Defect | HW | Oxi module | Hemolyzer |  | Oximetry hardware during CAS | Oximetry hardware during CAS validation | Replaced Hemolyzer and pinch valve. Ran THB cal
Yes | Visit | Complaint | RDEU | Während der Durchführung der Qualitätskontrolle wird sie abgebrochen mit Fehlermeldung 0999 Undefined Error. | Analysator bricht nach zwei Minuten die Qualitätskontrolle ab und spült anschließend. Nach kurzer Zeit startet er erneut eine Qualitätskontrolle. Überprüfung im Service zeigt, dass alle Werte in Ordnung sind. Kalibrationen werden erfolgreich durchgeführt. | Tausch des Wet Section Control Platine. Einspielung der aktuellen Applikations Software, da Platine noch keine aktuelle Software. Nach Neustart des Analysators nach Software Einspielung kommt der gleiche Fehler. Rückbau des Wet Section Control Platine. Ursache anderes Bauteil, Sensor Bay bestellt. Bei Lieferung erfolgt Einbau. | SWO-498020 | Re classified RMED | installation report |  |  | Move to OOB trending; Techline "RDEU 835682385" | During the execution of the quality control, it is aborted with error message 0999 Undefined Error. | Analyzer breaks off the quality control after two minutes and then rinses. After a short time, he starts a quality control again. Checking in the service shows that all values are in order. Calibrations are successfully performed. | Replacement of the Wet Section Control board. Recording of the current application software, since board is not yet any current software. After restarting the analyzer after software recording, the same error occurs. Dismantling of the Wet Section Control board. Cause other component, Sensor Bay ordered. On delivery, installation takes place.
No | Visit | Complaint | RDEU | Oximetrie Parameter können nicht kalibriert werden | Heute ist eine Verbrachsmateriallieferung eingetroffen. Damit ist es möglich den Fehler zu untersuchen. | OXI Compl. Test hat ergeben dass die LED Lampenintensität und Kavitation schlechte Werte liefern. Lampeneinheit wird neu bestellt. OXI- Parameter vorübergehend abgeschaltet damit zumindest die Blutgase , E- Lyte und Glu. weiter gemessen werden können. | SWO-498463 | [5 empty] | Oximetry parameters cannot be calibrated | Today, a delivery of verfore-material has arrived. This makes it possible to investigate the error. | OXI Compl. Test has shown that the LED lamp intensity and cavitation provide poor values. Lamp unit is reordered. OXI parameters temporarily switched off so that at least the blood gases, E- Lyte and Glu. can be further measured.
No | Visit | Complaint | RCHN | 报错5257 | 报错5257 | 更换636-550后修复（by chuanwei） | SWO-497690 | Defect | HW | Process Module | Sensor, process unit | 5257 | 5257 error | 5257 error | Repair after replacement 636-550 (by chuanwei)
No | Visit | Complaint | RDEU | Probenfehler bei Patientenmessungen und tHb-Kal | Probenfehler bei Messungen und tHb-Kal Probe vorzeitig im Oxi Modul erkannt | Überprüfung Quetschventil => alles i.O. Überprüfung, Ausbau und Reinigung der Messmodulventile Überprüfung Nadel => Nadel gerichtet => Inlet-Justage i.O. LS-Justage und PumpenKal i.O. tHb-Kal weiterhin fehlerhaft Durchführung Proteinentfernung Austausch der Pumpenschläuche und des Lösungspumpenschlauchs Erneut PumpenKal i.O. Kal1 und erneute tHb-Kal i.O. => Fcuv = 1,0022 Nachholen der fehlerhaften QCs | SWO-497958 | Defect | HW | Fluid module | Pump tubes |  | Probenfehler bei Patientenmessungen und tHb-Kal | Sample errors in measurements and tHb-Kal Sample detected prematurely in Oxi module | Checking squeezing valve => all in the art Inspection, removal and cleaning of the measuring module valves Verification needle => Needle directed => Inlet-Adjustment i.O. LS-Adjustment and PumpsKal i.O. tHb-Kal still faulty Execution of protein removal Replacement of the pump hoses and the solution pump hose Again pumpsKal i.O. Kal1 and again tHb-Kal i.O. => Fcuv = 1,0022 Catching up with faulty QCs
No | 8 | Y | RCZE | Blank |  |  | D5AC32138C5238D4 | Assisted Replacement | Electrode | Electrode | pCO2 |  | Blank
No | Visit | Complaint | RKK | neon number too high が断続的に発生 | 装置記録で症状確認 | 代替モジュールに入れ替え | SWO-498298 | Defect | HW | Oxi module | Oxi module | 0315 | neon number too high occurs intermittently | Check symptoms by device record | Replace with alternate module
No | Visit | Complaint | RDEU | Gerät muss gespült werden, Inhomogene Lösungen, 13.02.2019 Herr Beltenin 020360083131 meldet, dass der Analysator gespült werden muss. | Gerät muß gespült werden, LP kann nicht erneut installieren werden, | Fließwege gereinigt, LP im Serv.programm installiert, Gerät kalibriert, SC erneuert, QC durchgeführt, | SWO-498756 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | device must be rinsed, inhomogeneous solutions, 13.02.2019 Mr. Beltenin 020360083131 reports that the analyzer must be rinsed. | Device must be flushed, LP cannot be reinstalled, | Flow paths cleaned, LP installed in the serv.program, device calibrated, SC renewed, QC performed,
Yes | Visit | Complaint | RDEU | Der Kunde möchte gerne Ca bis 0,1 mml/l messen | Kunde möchte gerne Low. Ca messen. | In der roma.ini. den Eintrag geändert, Kunden informiert, | SWO-498640 | Re classified RMED | planned SWO |  |  | Request | The customer would like to measure Ca up to 0.1 mml/l | Customer would like Low. Ca measure. | In the roma.ini. changed the entry, informs customers,
No | Visit | Complaint | RDEU | Fehler Code 2 | Inbetriebnahme | LP & SC eingesetzt. Pressure Test nicht bestanden. LP & Einlassdichtung stark verschmutzt. Clot im Sc Bay. Ersatzanalysator wird wieder in Betrieb genommen. Kal & QC | SWO-500574 | Defect | HW | DMS_PC | CF card | See more details in CRM in SWO-498431 | Error Code 2 | Commissioning | LP & SC. Pressure test failed. LP & inlet seal heavily polluted. Clot in Sc Bay. Replacement analyzer is put back into operation. Kal & QC
No | Visit | Complaint | RLTD | Contact POC on 0208 5 467 711 ext. 3299 after 09:00 tomorrow. | Leak detected probe appeared bent. | Straightened probe performed pressure test, all passed. | SWO-498639 | Adjustment | HW | Auto Inlet module | Probe bend | 1342 | Contact POC on 0208 5 467 711 ext. 3299 after 09:00 tomorrow. | Leak detected probe appeared bent. | Straightened probe performed pressure test, all passed.
No | Visit | Complaint | RAME | Tne oximetry results are intermittently questioned for no apparent reason on patient samples. | Performed oximeter test several times and it would fail intermittently with Spectrometer Noise = 0.00 (0.30 - 1.80); Neon Position = -10 (-0.8 - 0.8); and Hemolyzer Cavitation = 61.7 (5.00 - 30.00). | Replaced hemolyzer and spectrometer units. Performed tHb calibration and ran QC. Analyzer checks okay. | SWO-498296 | Defect | HW | Oxi module | Hemolyzer | 581 | Tne oximetry results are intermittently questioned for no apparent reason on patient samples. | Performed oximeter test several times and it would fail intermittently with Spectrometer Noise = 0.00 (0.30 - 1.80); Neon Position = -10 (-0.8 - 0.8); and Hemolyzer Cavitation = 61.7 (5.00 - 30.00). | Replaced hemolyzer and spectrometer units. Performed tHb calibration and ran QC. Analyzer checks okay.
No | Visit | Complaint | RKK | SP交換時、ローターから異音 | SP交換時、ローターから異音 | 機器引き上げ | SWO-500468 | Adjustment | HW | Basic Unit | Actuator unit |  | Noise from rotor when replacing SP | Noise from rotor when replacing SP | Equipment pulling up
No | Visit | Complaint | RSCH | Blank | Erreur transport liquide-air non homogène malgré 2X changement SP Ensuite erreur courant de fuite détecté dans l’analyseur Solution de transport non-homogène ABL installé depuis 6 jrs | Nettoyage du système fluidique e.o Divers tests à l’aide du programme de service e.o Nouveau remplacement du SP e.o Calibration e.o contrôle de qualité e.o | SWO-498692 | Assisted Replacement | SP | SP | FT |  | Blank | Non-homogeneous liquid-air transport error despite 2X SP change Then common leak error detected in the analyzer Non-homogeneous transport solution ABL installed for 6 jrs | Cleaning the fluid system e.o Various tests using the e.o service program New replacement of the SP e.o Calibration e.o quality control e.o
No | Visit | Complaint | RAME | Customer states unit Failing Oximetry and Fluid Flow Issues during Calibrations. | Upon arrival, found above errors listed. | Replaced all above parts listed. Ran all service Oximetry Alignment tests. Ran Fluid Flow Tests. Unit functioning to Specifications. | SWO-498901 | Defect | HW | Oxi module | Hemolyzer | Oxi+FT | Customer states unit Failing Oximetry and Fluid Flow Issues during Calibrations. | Upon arrival, found above errors listed. | Replaced all above parts listed. Ran all service Oximetry Alignment tests. Ran Fluid Flow Tests. Unit functioning to Specifications.
Yes | 2 | Y | RCZE | Blank |  |  | D5AFD61AEE1F7F61 | Re classified RMED | installation report |  |  | IR will be filed when root cause found (see mail with title "RE: D5AFD61AEE1F7F61 - Installation Report?" from JIRVO) | Blank
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 多参数定标不过 | 更换测试卡后修复 | SWO-498400 | [5 empty] | Test card, reagent package failure | Multi-parameter calibration but | Repair after replacing test card
No | Visit | Complaint | RCHN | 断电后仪器无法开机 | 断电后仪器无法开机 | 自行更换902-970后修复-by Hunter | SWO-498480 | [5 empty] | The instrument does not power on after power failure | The instrument does not power on after power failure | Repair-by Hunter after replacing 902-970
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | LAC,PCO2定标失败 | 更换测试卡后修复 | SWO-501224 | [5 empty] | Test card, reagent package failure | LAC, PCO2 bid failed | Repair after replacing test card
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 液体传输失败 | 更换测试卡后修复 | SWO-499715 | Assisted Replacement | SC | SC | FT |  | Test card, reagent package failure | Failed liquid transfer | Repair after replacing test card
No | Visit | Complaint | RCHN | 测试卡舱门无法关闭 | 测试卡舱门无法关闭（未更换备件） | 调整测试卡上方铝块位置后修复-by Hunter | SWO-498521 | Adjustment | HW | Basic Unit | Sensor door |  | Test card door cannot be closed | Test card door cannot be closed (spare parts not replaced) | Fix -by Hunter after adjusting the aluminum block position above the test card
No | Visit | Complaint | RCHN | 无法开机 | 无法开机，借910-342和910-285各一个 | 更换910-285后修复-by Hunter | SWO-497659 | Defect | HW | Power Supply | Power supply |  | Cannot power on | Cannot be switched on, borrow 910-342 and 910-285 each one | Repair-by Hunter after replacement 910-285
No | Visit | Complaint | RCHN | 间歇性报样本量不足 | 间歇性样本量不足 | 穿刺针到10000次，更换针，间歇性报样本量不足，重新定标针和HCT单元，观察两天，目前没再报错，BY Hunter | SWO-499642 | Defect | HW | Aspiration Module | Needle Assembly | Insufficient sample size | Intermittently reported insufficient sample size | Insufficient intermittent sample size | Piering needles to 10,000 times, replacement needles, intermittently reported insufficient sample size, re-calibration needle and HCT unit, observed for two days, no errors, BY Hunter
No | Visit | Complaint | RCHN | 进样口支架的滚轮丢失 | 进样口支架的滚轮丢失 | 自行更换903-584后修复-by Hunter | SWO-499029 | Defect | HW | Aspiration Module | Tube holder | Roller | Loss of rollers for the inlet bracket | Loss of rollers for the inlet bracket | Repair-by Hunter after replacing 903-584
No | Visit | Complaint | RCHN | 存储盘旋转异常 | 存储盘旋转异常，借886-625一个 | 更换886-625后修复-by Hunter | SWO-497004 | Defect | HW | Storage Module | Cartridge guide | 5784 | Storage disk rotation exception | Storage disk rotation exception, borrow 886-625 one | Repair-by-Hunter after replacement 886-625
No | Visit | Complaint | RIND | Lac calibration sensitivity is low. | Lac calibration sensitivity is low need to replace Lac membrane. | Replaced the Lac membrane and run the instrument its working good. | SWO-500051 | Assisted Replacement | Electrode | Membrane | Lac |  | Lac calibration sensitivity is low. | Lac calibration sensitivity is low need to replace Lac membrane. | Replaced the Lac membrane and run the instrument its working good.
No | Visit | Complaint | RAME | Customer states that the analyzer is leaking from the inlet. | Verified customer's stated issue. The leak detected was due the tubing that connects to the waste pump tubings being ripped. | Replaced the clear tubing and the waste tubings (provided by the customer) with no leaks detected. | SWO-498790 | Defect | HW | Inlet module | Tubing |  | Customer states that the analyzer is leaking from the inlet. | Verified customer's stated issue. The leak detected was due the tubing that connects to the waste pump tubings being ripped. | Replaced the clear tubing and the waste tubings (provided by the customer) with no leaks detected.
No | 4 | Y | RCZE | Blank |  |  | D5ADD0630252B02B | User Error, FSE | Demo key | Demo key | Demo key | Demo key expired | Blank
No | Visit | Complaint | RKK | K インピーダンスエラー 他 | K インピーダンスエラー 他 | センサーシミュレーターによるセンサーチェック 2回パス インピーダンス -3.555 装置分解・内部点検 回路内、路過塩素酸による洗浄実施 各部動作確認 SC交換、CAL・QC確認 | SWO-499567 | Assisted Replacement | SC | SC | K | 1079 | K impedance error, etc. | K impedance error, etc. | Sensor check by sensor simulator 2 passes Impedance -3.555 Equipment decomposition and internal inspection Cleaning in circuits and by passperchloric acid Check the operation of each part SC replacement, CAL/QC confirmation
No | Visit | Complaint | RDEU | 14.02.2019 Herr Wolfen 020360083681 meldet, dass der ABL einen schwarzen Bildschirm zeigt. | Status Error Code 2 beim Booten. | Datenbank nach Servicemanual ersetzt, Konfig. eingespielt. Funktion OK. | SWO-500919 | Defect | SW | RMED SW | Database | Status code 2 | 14.02.2019 Mr. Wolfen 020360083681 reports that the ABL shows a black screen. | Status Error Code 2 at boot. | Database replaced by service manual, Config. Recorded. Function OK.
No | Visit | Complaint | RAME | has replaced 2 sp and both of them are leaking, checked ports and they seem to be alright. But still leaking | Found door and manifold screws were loose. | Tightened screws and cleaned system | SWO-498965 | [5 empty] | has replaced 2 sp and both of them are leaking, checked ports and they seem to be alright. But still leaking | Found door and manifold screws were loose. | Tightened screws and cleaned system
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 流量选择器错误。报错1177 | 连续更换2个试剂包后修复 | SWO-504286 | Assisted Replacement | SP | SP | Flow selector | 1177 | Test card, reagent package failure | The traffic selector is wrong. 1177 error | Repair after 2 reagent kits in a row
No | Visit | Complaint | RCHN | 无法读取试剂包 | 无法读取试剂包 | 由于科室老师没有将试剂包安装到位 导致仪器无法读取试剂包 重新安装试剂包后 仪器恢复正常使用 by Andrew | SWO-500524 | User Error | Consumables | SP | SP | Installation | Unable to read the package | Unable to read the package | Because the department teacher did not install the reagent kit in place, the instrument could not read the reagent package reinstalled the refit kit, the instrument returned to normal use by Andrew
No | Visit | Complaint | RCHN | 泄露失败，血氧模块位置泄露，报错0663 | 泄露失败，血氧模块位置泄露，报错0663 | 更换902-527和842-330后修复-by Hunter | SWO-499106 | Defect | HW | Oxi module | Pinch valve | 0663 | Leak failure, blood oxygen module location leak, error 0663 | Leak failure, blood oxygen module location leak, error 0663 | Replacement 902-527 and 842-330 post-fix-by-Hunter
No | Visit | Complaint | RDEU | Aufforderung LP Wechsel | Aufforderung LP Wechsel | Hakenventil getestet, i.O., Im Service Fluid transport Programme, Lösung QC2 Fehler, aufgef. Verschleißteile ersetzt, Test i.O., Kal, QC Datensicherung Kal QC | SWO-499886 | Defect | HW | Basic Unit | PM Kit |  | Prompt LP Change | Prompt LP Change | Hook valve tested, i.O., In Service Fluid transport program, solution QC2 error, unwoven. Wear parts replaced, test i.O., Kal, QC Data backup Kal QC
No | Visit | Complaint | RBNL | Bij reguliere vervanging reagenspack wordt nieuwe pack niet vastgehouden. Maakt veel lawaai. | Bij regulier verwisselen reagenspack veel lawaai, haak blijft omhoog staan. Veel vet op as actuatormotor. Moet 15-2-2019 13:00 uur regulier sensorcassette vervangen. | As schoongemaakt. | SWO-499878 | Cleaned | HW | Basic Unit | Stepper Motor |  | With regular replacement reagent pack, new pack is not retained. Makes a lot of noise. | When regular swapping reagent pack a lot of noise, hook stays up. Lots of grease on shaft actuator motor. Must replace 15-2-2019 13:00 hours of regular sensor cassette. | As cleaned.
No | Visit | Complaint | RLTD | Mona Tel: 07503 686865 up until 5pm. After 5pm please call 0208 546 7711 bleep: 373 and speak with the nurse in charge | Reloaded software in order to recover database | Database still currupt after software reload. Renewed database and restarted, analyser operational. | SWO-499909 | Defect | SW | RMED SW | Database | corrupt | Mona Tel: 07503 686865 up until 5pm. After 5pm please call 0208 546 7711 bleep: 373 and speak with the nurse in charge | Reloaded software in order to recover database | Database still currupt after software reload. Renewed database and restarted, analyser operational.
No | Visit | Complaint | RLTD | Contact Rachel Baxter 01267 227 043 | On arrival the analyser was intervention required. The fluid transport tests were ran and the vacuum valve pressure test was failing along with clot detect in the inlet. The analyser was flushed and no clots were detected with a good flow. The oxi pinch tube and pinch valve were inspected. There appeared to be a possible leak in the Oxi pinch tube and this was replaced. On running the pressure tests once again the values improved but still failed. THe inlet gasket was inspected and slight visual damage was seen, as such it was replaced. All fluid transport tests ran and passed. | Replacement of the Oxi pinch tube and inlet gasket. The customer supplied the inlet gasket. | SWO-499837 | Defect | HW | Auto Inlet module | Inlet gasket |  | Contact Rachel Baxter 01267 227 043 | On arrival the analyser was intervention required. The fluid transport tests were ran and the vacuum valve pressure test was failing along with clot detect in the inlet. The analyser was flushed and no clots were detected with a good flow. The oxi pinch tube and pinch valve were inspected. There appeared to be a possible leak in the Oxi pinch tube and this was replaced. On running the pressure tests once again the values improved but still failed. THe inlet gasket was inspected and slight visual damage was seen, as such it was replaced. All fluid transport tests ran and passed. | Replacement of the Oxi pinch tube and inlet gasket. The customer supplied the inlet gasket.
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 温度不稳定 | 更换测试卡后修复 | SWO-500358 | [5 empty] | Test card, reagent package failure | Temperature instability | Repair after replacing test card
No | Visit | Complaint | RCHN | 蓝屏 | 蓝屏，借933-441一个 | 更换933-441和903-381后修复by Hunter | SWO-499049 | [5 empty] | Blue screen | Blue screen, borrow 933-441 one | Repair by Hunter after replacing 933-441 and 903-381
No | Visit | Complaint | RDEU | Oximetrie Parameter können nicht kalibriert werden | spectrometer defekt. | spectrometer getauscht. tHb- Kal durchgeführt. Kalibration und Qualitätskontrolle in Ordnung. | SWO-499942 | [5 empty] | Oximetry parameters cannot be calibrated | spectrometer defective. | spectrometer. tHb- Kal. Calibration and quality control in order.
Yes | Visit | Complaint | RCHN | 科室老师抱怨血氧不出值，飘红 | 科室老师抱怨血氧不出值，飘红 | 血红蛋白液重新定标，并无问题by Hunter | SWO-499663 | Re classified RMED | same day upSWO |  |  | Other in: SWO-499668 | The teacher in the department complained that the blood oxygen was not worth it, and it was red. | The teacher in the department complained that the blood oxygen was not worth it, and it was red. | Hemoglobin liquid recalibrated with out of question by Hunter
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 1361,参比电极错误 | 更换测试卡后修复 | SWO-499668 | Assisted Replacement | SC | SC | Reference | 1361 | Test card, reagent package failure | 1361, Reference Electrode Error | Repair after replacing test card
Yes | Visit | Complaint | RCHN | 报错1024，定标失败 | 报错1024，定标失败，需要借件944-212（15116） |  | SWO-499683 | [5 empty] | 1024, bid failure | 1024, failed bid, need to borrow 944-212 (15116)
Yes | Visit | Complaint | RCHN | 试管夹损坏 | 试管夹损坏，借2个903-584 |  | SWO-499685 | [5 empty] | Damage to test tube clip | Damaged tube clip, borrow 2 903-584
Yes | Visit | Complaint | RCHN | CKMB定标失败 | CKMB定标失败，借944-216（14931）一张 | 重新定标后修复by HUNTER | SWO-499001 | Re classified RMED | same day upSWO |  |  | Other in: SWO-499064 | CKMB calibration failed | CKMB bid failed, borrow 944-216 (14931) one | Fix by HUNTER after re-calibration
No | Visit | Complaint | RCHN | CKMB定标失败，AU连接不上 | 又出现了AU连接不上的故障，借903-117 903-198各一个 | 自行更换备用件903-117 903-198后修复by Hunter | SWO-499064 | Defect | HW | Cover and Chassis | PCB, Main Board | Potentially Use Error (FSE) | CKMB calibration failed, AU connection is not | There is also a failure on the AU connection, borrowing 903-117 903-198 each | Repair by Hunter after replacing spare parts 903-117 903-198
No | Visit | Complaint | RPAC | Poor print quality on ED analyzer was reported on Feb 11th, Andrew updated as breakdown on Feb 13th due to results can't be read | Printer was very streaky | Arrived and disassembled the mqchine then went ahead and replaced the printer, now its all printing well. Cals and QC all fine. | SWO-499531 | Defect | HW | Printer Module | Printer Module |  | Poor print quality on ED analyzer was reported on Feb 11th, Andrew updated as breakdown on Feb 13th due to results can't be read | Printer was very streaky | Arrived and disassembled the mqchine then went ahead and replaced the printer, now its all printing well. Cals and QC all fine.
No | Visit | Complaint | RPOL | Błędy kontroli jakości | Błędy kontroli jakości | Sprawdzenie aparatu. Wymiana pakietu odczynników. | SWO-500049 | Assisted Replacement | SP | SP | SP | QC | Quality control Errors | Quality control Errors | Check the camera. Replacement of reagent pack.
No | Visit | Complaint | RPOL | Uszkodzony moduł baterii | Niesprawny moduł baterii. | Wymiana modułu baterii. | SWO-500020 | Defect | HW | Battery module | Battery module |  | Defective Battery Module | Faulty battery module. | Replacing the battery module.
No | Visit | Complaint | RPOL | Wygięta igła. | Wygięta igła. | Wymiana igły. | SWO-499917 | Defect | HW | Auto Inlet module | Probe bend |  | Bent Needle. | Bent Needle. | Replacing the needle.
No | Visit | Complaint | RCHN | 报错2168 | 报错2168 (未使用备件) | 重新清洁传感器后修复 by Andrew | SWO-500373 | Cleaned | HW | Incubator Module | Sensor, incubator | 2168 | 2168 error | Error 2168 (unused spare parts) | Fix by Andrew after recleaning sensor
No | Visit | Complaint | RCHN | 定标分析物外漏检查失败，报错1068 | 定标分析物外漏检查失败，报错1068 借944-394（15048缺货）一张 | 重新定标后修复by Hunter | SWO-499142 | Adjustment | HW | Process Module | Cup dryer | 1068 | Bid analysis external leakage check failed, reported error 1068 | The calibration analysis leak inspection failed, reported error 1068 borrowed 944-394 (15048 out of stock) one | Fix by Hunter after re-calibration
No | Visit | Complaint | RCHN | 屏幕触摸屏失灵 | 屏幕触摸屏失灵，借910-320一个 | 更换910-320后修复by Hunter | SWO-499135 | Defect | HW | PC module | Touch Screen |  | Screen touch screen failure | Screen touch screen failure, borrow 910-320 one | Repair by Hunter after replacement 910-320
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 探测到部分漏电流0964 | 更换试剂包后修复 | SWO-500285 | [5 empty] | Test card, reagent package failure | Part leakage current 0964 detected | Repair after replacement of reagent kit
No | Visit | Complaint | RDEU | 08.02.2019, Herr Hirschel, Tel: 021513341343: Leckage erkannt | Nach einer Messung am 13.2 um 17:29 Uhr ging das Gerät erneut in Fehlerbehebung erforderlich - Leckage erkannt. | Herr Dorendorf das Benutzermanagement erklärt, Herr Dorendorf wird das Benutzermanagement einstellen um schlechte Präanalytische Proben herauszufiltern. Ampelstatus grün seit dem 13.2. Eine QC durchgeführt. | SWO-500743 | No fault found | FT | FT | Leak detected |  | 08.02.2019, Mr. Hirschel, Tel: 021513341343: Leakage detected | After a measurement on 13.2 at 17:29 the device went back to troubleshooting required - leakage detected. | Mr. Dorendorf's user management explains that Mr. Dorendorf will discontinue user management to filter out bad preanalytical samples. Traffic light status green since 13.2. A QC performed.
No | Visit | Complaint | RLTD | Philip Dixon 07341 268422 1.55pm - further call received from Phillip Dixon 07341 268422 - first fix worked for a short while, problem now reoccurred. 4.50pm - Philip Dixon called. Problem still occurring. Thinks may be an analyser fault. Tel 07341 268422. On site until around 5.15pm but can call mobile. | Found oximetry pinch valve sticking intermittently. | Strip and clean oximetry pinch valve. Continue to perform routine service. | SWO-500886 | Cleaned | HW | Oxi module | Pinch valve |  | Philip Dixon 07341 268422 1.55pm - further call received from Phillip Dixon 07341 268422 - first fix worked for a short while, problem now reoccurred. 4.50pm - Philip Dixon called. Problem still occurring. Thinks may be an analyser fault. Tel 07341 268422. On site until around 5.15pm but can call mobile. | Found oximetry pinch valve sticking intermittently. | Strip and clean oximetry pinch valve. Continue to perform routine service.
No | Visit | Complaint | RDEU | 15.02.2019 13:10 Sr. Doro 02088811292: Gerät startet nicht mehr nach dem ausschalten den den NETZSCHALTER | Gerät lässt sich nicht bedienen, Error code 2 , Datenbank ist def., | Dateb gelöscht, neue Software eingespielt, neuen roma. Ordner eingerichtet, | SWO-501734 | Defect | SW | RMED SW | Database | corrupt | 15.02.2019 13:10 Sr. Doro 02088811292: Device does not start after switching off the NETZSCHALTER | Device cannot be operated, Error code 2 , database is def., | Dateb deleted, new software added, new roma. folder set up,
No | Visit | Complaint | RAME | Analyzer was entering User Intervention due to a pump calibration error. | Troubleshot using the fluid transport program in service. Upon entering the program, I could see that the pump calibration had failed for a value of 0.79 (range = 0.80 - 1.20). After a few calibrations, I was only able to attain a calibration factor of 0.82. I replaced the solution pack, with a known good pack, and the calibration factor jumped to 0.99 on the first pump calibration. When the suspected bad solution pack was installed on another analyzer, the initial pump calibration factor was 0.79. | Replaced the solution pack, and this resolved the problem. | SWO-501103 | Assisted Replacement | SP | SP | Pump cal | 1178 | Analyzer was entering User Intervention due to a pump calibration error. | Troubleshot using the fluid transport program in service. Upon entering the program, I could see that the pump calibration had failed for a value of 0.79 (range = 0.80 - 1.20). After a few calibrations, I was only able to attain a calibration factor of 0.82. I replaced the solution pack, with a known good pack, and the calibration factor jumped to 0.99 on the first pump calibration. When the suspected bad solution pack was installed on another analyzer, the initial pump calibration factor was 0.79. | Replaced the solution pack, and this resolved the problem.
Yes | Visit | Complaint | RDEU | Während der Durchführung der Qualitätskontrolle wird sie abgebrochen mit Fehlermeldung 0999 Undefined Error. | Lieferung des notwendigen Ersatzteils. Analysator war ausgeschaltet. | Ausbau der alten Sensor Bay und Einbau der neuen Sensor Bay. Nach Neustart Phase und erster Kalibration ist der Fehler noch nicht aufgetreten. Analysator konfiguriert Sensorkassette, danach tritt Fehler auf. Kalibrationen und Qualitätskontrolle werden abgebrochen mit Fehlermeldung 0999 ERR_UNDEFINED und 1340 Fortlaufende Spülung der SC wurde unterbrochen . Analysator geht in Fehlerbehebung Verfahren und will Einlassdichtung erneuert haben. Wird dabei der Service Button betätigt, zeigt er Startphase an und nach ein paar Minuten geht er wieder in die gleiche Ansicht. Test Analysator ausschalten und Analysator ohne LP hochfahren brachte kein Erfolg. Test im Service zeigt weiterhin schwankende LS Werte. Service Dump und Datalogs gezogen. | SWO-500946 | Re classified RMED | Techline and OOB |  |  | Techline "RDEU 835682385" | During the execution of the quality control, it is aborted with error message 0999 Undefined Error. | Delivery of the necessary spare part. Analyzer was off. | Extension of the old Sensor Bay and installation of the new Sensor Bay. After rebooting phase and initial calibration, the error has not yet occurred. Analyzer configures sensor cartridge, after which errors occur. Calibrations and quality control are aborted with error message 0999 ERR_UNDEFINED and 1340 continuous flushing of the SC has been interrupted. Analyzer goes into troubleshooting procedure and wants to have inlet seal renewed. If the service button is pressed, it displays the start phase and after a few minutes it goes back into the same view. Turning off the test analyzer and booting analyser without LP did not bring any success. Test in the service continues to show fluctuating LS values. Service Dump and Datalogs.
No | Visit | Complaint | RAME | The analyzer goes in to troubleshooting mode asking to replace the solution pack. | Analyzer is going in to Troubleshooting mode asking to replace the solution pack. | I checked in the activity log and, found error 1189 FTC aborted, LS state change. Checked and, reseated the solution pack and, the sensor cassette. Back-flush the system lines. Rinse the inlet gasket holder with water. I found the Cal3 solution in the solution pack fluid transport check was failed. Performed the fluid transport checks and, now all tests are passing within specifications. | SWO-501101 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway | 1189 | The analyzer goes in to troubleshooting mode asking to replace the solution pack. | Analyzer is going in to Troubleshooting mode asking to replace the solution pack. | I checked in the activity log and, found error 1189 FTC aborted, LS state change. Checked and, reseated the solution pack and, the sensor cassette. Back-flush the system lines. Rinse the inlet gasket holder with water. I found the Cal3 solution in the solution pack fluid transport check was failed. Performed the fluid transport checks and, now all tests are passing within specifications.
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | PH,CO2，HCT,LAC定标不过 | 更换测试卡后修复 | SWO-501167 | [5 empty] | Test card, reagent package failure | PH, CO2, HCT, LAC calibration | Repair after replacing test card
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | K,CA,NA,CL,HCT定标不过 | 更换测试卡后修复 | SWO-501165 | [5 empty] | Test card, reagent package failure | K,CA,NA,CL,HCT定标不过 | Repair after replacing test card
No | Visit | Complaint | RCHN | 电极故障，温度错误 | 电极故障，温度错误，借945-377一个 |  | SWO-500496 | [5 empty] | Electrode failure, temperature error | Electrode fault, temperature error, borrow edi-945-377
No | Visit | Complaint | RCHN | 无法开机 | 无法开机，借910-285一个 |  | SWO-500378 | Defect | HW | Power Supply | Power supply |  | Cannot power on | Cannot power on, borrow 910-285 one
No | Visit | Complaint | RCHN | 定标失败，报错1024 | 定标失败，报错1024 借944-212（14715）一张 |  | SWO-500596 | Cleaned | HW | Process Module | Process Wheel | 1024 | Failed to bid, error 1024 | Bid failure, error 1024 borrow944-212 (14715) one
No | Visit | Complaint | RCHN | 样本架损坏 | 样本架损坏，借903-584 $2个 |  | SWO-500376 | Defect | HW | Aspiration Module | Tube holder |  | Damage to the sample rack | Damaged sample rack, loan 903-584 /2
No | Visit | Complaint | RCHN | 无法读取处理包芯片 | 无法读取处理包芯片 | 更换试剂包后修复 | SWO-501252 | Defect | Consumables | Solution Pack | Solution Pack | Chip | Cannot read the processing package chip | Cannot read the processing package chip | Repair after replacement of reagent kit
No | Visit | Complaint | RCHN | 数据库损坏 | 数据库损坏，借944-212（15457），944-216（15446），944-214（15428），944-258（15488），944-220（15354），944-394（15363,15623）各一个 |  | SWO-500541 | Defect | SW | RMED SW | Database | Potentially Use Error (FSE) | Database corruption | Database corruption, borrowed 944-212 (15457), 944-216 (15446), 944-214 (15428), 944-258 (15488), 944-220 (15354), 944-394 (15363, 15623) each
No | Visit | Complaint | RCHN | 扫码器灯不亮 | 扫码器灯不亮.需要借件902-924 |  | SWO-500585 | Defect | HW | Cover and Chassis | Barcode reader, display |  | The scanner light is not on | The coder light is not on.
No | Visit | Complaint | RCHN | 氧定标不稳定 | 氧定标不稳定 | 清洁电极仓 by Andrew | SWO-500540 | Cleaned | Fluidic pathway | pH/BG Module | measuring chamber |  | Oxygen calibration unstable | Oxygen calibration unstable | Cleaning electrode compartment by Andrew
No | Visit | Complaint | RDEU | 15.02.2018 Herr Vutsch 07361552250 meldet, dass der ABL unzuverlässig läuft und wünscht eine Überprüfung. | Na QC Fehler und tHb ohne Kal. Werte. Fehler 315 | Air und Wastekanal gespült. Na Elektrode membraniert und pH Dichtring erneuert. Messkammer gereinigt. Oxicompletetest durchgeführt,Halogenlampe bei 3000 und Neon Number too high. Lampenjustierung durchgeführt. Oxicomplete-Test o.k. Neon number bei 1.975 , mehrere Test durchgeführt, Wert bleibt stabil unterhalb 2.000 Grenzen. Service tHB-Kal. und mehrere 1.-2. Kal. durchgeführt. Kal und QC o.k. , keine Auffälligkeiten mehr feststellbar. | SWO-501250 | Adjustment | HW | Oxi module | Lamp unit | Neon number too high | 15.02.2018 Mr Vutsch 07361552250 reports that the ABL is running unreliably and wants a review. | Na QC Fehler und tHb ohne Kal. Werte. Fehler 315 | Air and waste channel. Na electrode membranes and pH sealing ring renewed. measuring chamber. Oxicomplete test performed,halogen lamp at 3000 and Neon Number too high. lamp adjustment. Oxicomplete test ok Neon number at 1,975 , multiple test performed, value remains stable below 2,000 limits. Service tHB-Kal. and several 1st-2nd Cal. Performed. Kal and QC ok , no more abnormalities can be detected.
No | Visit | Complaint | RPAC | Investigation by biomed Jakeer: Anlayser was not giving any audible sounds like ( close inlet flap, after sample extraction and also for selecting tabs on the screen), I had a look in service menu all setting were ok, but while I was In the service menu analyser had frozen for couple of minutes , then after analyser rebooted on its own then things came alright, user been noticing this since last Friday. Would you please depute a service Engineer or advise us to resolve the glitch. Requested staff inform us of any further issues since reboot. 11/2 Further email from Jakeer: As per the user Acoustic mute issue has reoccurred over weekend, this morning I has to do the same as last week, I selected the acoustic tab in the analyser settings then screen froze for quite long then, In the settings everything seems to be ok, to bring the acoustic signal back On has to tick the mute tab (check) and then untick (uncheck) it, for every key select analyser freezes then after unchecking the mute tab, analyser reboots on its own then things came alright. Please depute / advise to fix the issue. | Unable to reproduce the sounds loss. However found utopia~2.wav sound very small and intermittent. | Reinstall sounds. Adjusted master volume to maximum. Replaced utopia~2.wav with Windows Navigation Start.wav which are similar but louder. Tested all OK. To be monitored for any further sound issue. Due PM was done in separate SWO. | SWO-501064 | Setup | SW | RMED SW | RMED SW | Sounds | Investigation by biomed Jakeer: Anlayser was not giving any audible sounds like ( close inlet flap, after sample extraction and also for selecting tabs on the screen), I had a look in service menu all setting were ok, but while I was In the service menu analyser had frozen for couple of minutes , then after analyser rebooted on its own then things came alright, user been noticing this since last Friday. Would you please depute a service Engineer or advise us to resolve the glitch. Requested staff inform us of any further issues since reboot. 11/2 Further email from Jakeer: As per the user Acoustic mute issue has reoccurred over weekend, this morning I has to do the same as last week, I selected the acoustic tab in the analyser settings then screen froze for quite long then, In the settings everything seems to b e ok, to bring the acoustic signal back On has to tick the mute tab (check) and then untick (uncheck) it, for every key select analyser freezes then after unchecking the mute tab, analyser reboots on its own then things came alright. Please depute / advise to fix the issue. | Unable to reproduce the sounds loss. However found utopia~2.wav sound very small and intermittent. | Reinstall sounds. Adjusted master volume to maximum. Replaced utopia~2.wav with Windows Navigation Start.wav which are similar but louder. Tested all OK. To be monitored for any further sound issue. Due PM was done in separate SWO.
No | Visit | Complaint | RDEU | 18.02.2019 Herr Knorr 015254847560 meldet, dass der Analysator einen Black Screen zeigt. 18.02.2019/13:55 Herr Knorr meldet weiterhin Softwarefehler bei Neustart | Herr Knorr meldet, dass der Analysator einen Black Screen zeigt. weiterhin Softwarefehler bei Neustart | Neu Software installiert, Gerät konfiguriert, Kalibrationen und QC's vorgenommen. | SWO-500986 | Defect | SW | RMED SW | RMED SW | Black screen | 18.02.2019 Mr. Knorr 015254847560 reports that the analyzer shows a black screen. 18.02.2019/13:55 Mr Knorr continues to report software errors on reboot | Mr Knorr reports that the analyzer shows a black screen. continued software error on reboot | New software installed, device configured, calibrations and QC's made.
No | Visit | Complaint | RCHN | HCT值无法被测得 报错1223 3412 | HCT值无法被测得 报错1223 3412，借903-329一个 |  | SWO-500409 | Defect | HW | Aspiration Module | Needle Wash Unit | 1223, 3412 | HCT value cannot be measured for error 1223 3412 | HCT value can not be measured error 1223 3412, borrow903-329 one
No | Visit | Complaint | RPAC | Derek 0394292222. Unable to pump solution after changing 2 x Solution packs. Someone else tried to change the first one. They did not lift the Inlet gasket and may have damaged it. Asked Derek to replace inlet gasket and advise if still failing Pump solution. | Analyser is asking to be replaced and has a database error 1405. Put analyser into service and checked fluidic movement. All okay. Removed and checked solution pack. Solution pack had not been pushed down all the way. Reinserted solution pack after making sure its been pushed down. This has not resolved the issue. Analyser also has a OXI hardware error when trying to run the cal. LED peak intensity and cavitation absorb values are really out of range during a complete OXI test. | Replaced the solution pack with a new one. Flushed the hemolyser and shutdown analyser. Disassasembled and physcially checked hemolyser. Looks clean. Checked all cables and then reassembled. Performed fluid transport checks again and complete OXI test. Everything is now passing. Reinserted consumables and allowed analyser to perform its start up checks. The database error is no longer present and analyser passes calibrations and QCs. | SWO-501197 | User Error | SP | SP | Activation | Check where to put Activation | Derek 0394292222. Unable to pump solution after changing 2 x Solution packs. Someone else tried to change the first one. They did not lift the Inlet gasket and may have damaged it. Asked Derek to replace inlet gasket and advise if still failing Pump solution. | Analyser is asking to be replaced and has a database error 1405. Put analyser into service and checked fluidic movement. All okay. Removed and checked solution pack. Solution pack had not been pushed down all the way. Reinserted solution pack after making sure its been pushed down. This has not resolved the issue. Analyser also has a OXI hardware error when trying to run the cal. LED peak intensity and cavitation absorb values are really out of range during a complete OXI test. | Replaced the solution pack with a new one. Flushed the hemolyser and shutdown analyser. Disassasembled and physcially checked hemolyser. Looks clean. Checked all cables and then reassembled. Performed fluid transport checks again and complete OXI test. Everything is now passing. Reinserted consumables and allowed analyser to perform its start up checks. The database error is no longer present and analyser passes calibrations and QCs.
No | Visit | Complaint | RAME | Analyzer will not leave loop - unable to enter service program. | Analyzer is asking for inlet gasket replacement. When 'service' is pressed, screen goes immediately to startup screen. Unable to enter service. | Replaced SC and SP. | SWO-502015 | Defect | SW | RMED SW | RMED SW | Stuck in loop | Analyzer will not leave loop - unable to enter service program. | Analyzer is asking for inlet gasket replacement. When 'service' is pressed, screen goes immediately to startup screen. Unable to enter service. | Replaced SC and SP.
No | Visit | Complaint | RCHN | 干燥前温度超范围 1068 | 干燥前温度超范围 1068 | 更换903-210一个正常。 by Andrew | SWO-501352 | Defect | HW | Process Module | PCB, cup dryer | 1068 | Over-range temperature before drying 1068 | Over-range temperature before drying 1068 | Replace the 903-210 one normal. by Andrew
No | Visit | Complaint | RAME | Aborting calibrations | Checked data logs for errors, checked fluid flow, checked valves and liquid sensors, | Removed and replaced pH seal ring, removed and cleaned electrode chambers, removed modules and cleaned valves, replaced all pump tubing | SWO-502013 | Defect | Electrode | Electrode | Sealing ring, pH |  | Aborting calibrations | Checked data logs for errors, checked fluid flow, checked valves and liquid sensors, | Removed and replaced pH seal ring, removed and cleaned electrode chambers, removed modules and cleaned valves, replaced all pump tubing
No | Visit | Complaint | RPAC | Email from Rose Carr: Intermittent Calibration drift and control failures on calcium sodium,chloride since 20 December, ref membrane changed on 21/12/18, again on 30/12/18. Advised her to replace with reference electrode from lab (about to be decomissioned) and ordered replacement to be sent 7/1/19. | Calibration drift (error 0376) reported on 7 $1 $19, had been happening since December usually at 21:40 after first caibration 21:30. 7 $1 $19 Reference electrode 38-13 was removed and ex used biochem LOT 08-05 installed. 19 $2 $19 Installed new reference electrode as still some calibration drift. | Installed new reference electrode 945-603 LOT 58-12 Electrode removed 38-13 (production 2 $-2018) 08-05 (production 3 $2018) | SWO-501326 | Assisted Replacement | Electrode | Electrode | Reference | CAPA-733 | Email from Rose Carr: Intermittent Calibration drift and control failures on calcium sodium,chloride since 20 December, ref membrane changed on 21/12/18, again on 30/12/18. Advised her to replace with reference electrode from lab (about to be decomissioned) and ordered replacement to be sent 7/1/19. | Calibration drift (error 0376) reported on 7 /1 /19, had been happening since December usually at 21:40 after first caibration 21:30. 7 /1 /19 Reference electrode 38-13 was removed and ex used biochem LOT 08-05 installed. 19 /2 /19 Installed new reference electrode as still some calibration drift. | Installed new reference electrode 945-603 LOT 58-12 Electrode removed 38-13 (production 2 $-2018) 08-05 (production 3 $2018)
Yes | Visit | Complaint | RAME | DMS not connected to AU | Replaced power supply and tried reloading AU software and keys. Cleaned fuses with alcohol and suspect AU PCB . | Ordered new AU pcb and keys with serial number chip | SWO-502623 | Re classified RMED | same day upSWO |  |  | This one 4h prior to SWO-502622 | DMS not connected to AU | Replaced power supply and tried reloading AU software and keys. Cleaned fuses with alcohol and suspect AU PCB . | Ordered new AU pcb and keys with serial number chip
No | Visit | Complaint | RSCH | Verstopfung | Schläuche unter Druck, Verstopfung | Allgemeine Kontrolle ok. Alle Schläuche von Pumpenrotor zu Abfall ersetzt ok. Einlassdichtung ers ok. Diverse Schläuche ers ok. Gerät neu initialisiert ok. Blutprobe test i.o. Kal. I.o. | SWO-501501 | Defect | HW | Waste | Pump tubes | blocked | Constipation | Hoses under pressure, constipation | General control ok. All hoses from pump rotor to waste replaced ok. Inlet seal he ok. Various hoses he s ok. Device reinitialized ok. Blood sample test i.o. Kal. Ok.
No | Visit | Complaint | RDEU | 07.02.2019, Frau Rahn, Tel: 060329254613: QC Fehler bei CRP Level 2 | Überprüfung CRP Level2 oberhalb Referenzbereiches | Durchsicht der Aktivitäten und LQC. Flow zwischen 1300-1400 bei Patientenmessung, iO. Systemüberprüfung ohne Beanstandung. Analysator geöffnet. Reinigungsverfahren P1-P2, P4-P5 durchgeführt. P5 laute Laufgeräusche, P4 und P5 erneuert. Nadel gereinigt, neuen Collard aufgesetzt. Nadelblock gereinigt, auf Beschädigung geprüft. Nadelkolben freigängig. Needle wash block ausgebaut und gereinigt, neuen Abfallschlauch angebracht, Ablagerungen. Filter stark verschmutzt, erneuert. Cup Wash Unit ausgebaut, zerlegt und gereinigt. Processwheel und Incubatorwheel gereinigt. Gewinde gefettet. Aufwärm- und Initialisierungsphase Systemüberprüfung ohne Beanstandung. Nadelstichtiefe in Cup und Washblock kalibriert. LQC CRP Level2 3x durchgeführt. Sollwert 116mg $L oberer Grenzwert 132mg $L Ergebnisse CRP LQC (Auftauzeit 90Min.) 1. 127mg $L 2. 126mg $L 3. 128mg $L Falls LQC in den nächsten Tagen wieder Grenzwertig sein sollte, müsste CRP nochmals Kalibriert werden. LQC wurde bereits gegen neues Lot ausgetauscht. | SWO-501603 | Cleaned | HW | Wet section | Wet section | QC L2 CRP | 07.02.2019, Frau Rahn, Tel: 060329254613: QC Fehler bei CRP Level 2 | Review CRP Level2 above reference range | Review of activities and LQC. Flow between 1300-1400 for patient measurement, iO. System verification without complaint. Analyzer open. cleaning method P1-P2, P4-P5. P5 loud running noises, P4 and P5 renewed. Needle cleaned, new collard attached. Needle block cleaned, checked for damage. Needle flask free-running. Needle wash block removed and cleaned, new waste hose attached, deposits. Filter heavily polluted, renewed. Cup Wash Unit removed, disassembled and cleaned. Processwheel and Incubatorwheel cleaned. Thread greased. Warm-up and initialization phase System verification without complaint. Needle stitch depth calibrated in cup and washblock. LQC CRP Level2 3x. Setpoint 116mg $L upper limit 132mg $L Results CRP LQC (thawing time 90min.) 1. 127mg $L 2. 126mg $L 3. 128mg $L If LQC were to be limitable again in the next few days, CRP would have to be calibrated again. LQC has already been replaced with new lot.
No | Visit | Complaint | RLTD | Issue occurred during training by David Ruaux, reported by him. | On arrival found analyser to be working and no error messages present. Therefore unable to directly troubleshoot errors seen yesterday. Backed up customer data. Reloaded Windows XPe and analyser software. Restored customer data and restarted - analyser start-up failed with Database error. Installed blank database, restored customer set-up. Verified analyser operation - restart passed without error. Due to blank database calibration data for parameters has been wiped. So parameters need to be recalibrated. | Reloaded Windows XPe and AQT90 software v8.11 with blank database. | SWO-501663 | Defect | SW | RMED SW | RMED SW |  | Issue occurred during training by David Ruaux, reported by him. | On arrival found analyser to be working and no error messages present. Therefore unable to directly troubleshoot errors seen yesterday. Backed up customer data. Reloaded Windows XPe and analyser software. Restored customer data and restarted - analyser start-up failed with Database error. Installed blank database, restored customer set-up. Verified analyser operation - restart passed without error. Due to blank database calibration data for parameters has been wiped. So parameters need to be recalibrated. | Reloaded Windows XPe and AQT90 software v8.11 with blank database.
No | Visit | Complaint | RDEU | Herr Kraus meldet immer häufiger Probleme mit Lösungspacks 18.02.2019, 11:41 Uhr, Herr Asam, Tel. 0661159537: Weiter häufiger Wechsel der SC und des LP durch den ABL angefordert | In letzter Zeit verlangt das Gerät immer wieder Spülung und Tausch der Komponenten. | Ventil und Schlauch überprüft. Zusammen mit Anwendern die Einlassdichtung erneuert. Flow-Sensor gereinigt und Gerät im Service überprüft. Gerät überprüft | SWO-502537 | Defect | HW | Inlet module | Inlet gasket |  | Mr. Kraus is increasingly reporting problems with solution packs 18.02.2019, 11:41 a.m., Mr. Asam, tel. 0661159537: Further frequent change of SC and LP requested by the ABL | Lately, the device has been demanding the rinsing and replacement of the components. | valve and hose. Renewed the inlet seal together with users. Flow sensor cleaned and device checked in service. Device checked
No | Visit | Complaint | RDEU | 19.02.2019 Frau Audyova 04079006590 (-366) meldet, dass nach dem SW update kein Lac gemessen wird. | Es handelt sich nicht um den Parameter sondern das, dass der Analysator nicht nach der Kalibrationen einsatzbereit ist. Der Parameter Lac hat die Kalibration nicht bestanden. | Neue SC eingesetzt und abgewartet bis der Analysator wieder einsatzbereit ist. | SWO-502239 | Assisted Replacement | SC | SC | Lac | Cal | 19.02.2019 Mrs Audyova 04079006590 (-366) reports that no Lac is measured after the SW update. | It is not the parameter, but that the analyzer is not ready for use after calibration. The Lac parameter did not pass the calibration. | New SC inserted and waited until the analyzer is ready for use again.
No | Visit | Complaint | RDEU | Drucker sporadisch ohne Funktion | Drucker sporadisch ohne Funktion | Drucker ersetzt | SWO-502861 | Defect | HW | Printer Module | Printer Module |  | Printer sporadically without function | Printer sporadically without function | Printer replaced
No | Visit | Complaint | RDEU | 20.02.2019, 08:10 Uhr, Fr. Podstawa, Tel. 023021756565: Keine Oxy Werte; Kal und QC Fehler 20.02.2019, 08:54 Uhr, Fr. Podstawa, Tel. 023021756565: Oxy Hardware Fehler | Daten ausgewertet, Fehler 1011 Keine Nullpunkt Kal für OXI bei Kalibration war nach aus $einschalten behoben. Der Anwender hatte aber eine tHb Kal mit OXI Hardware error durchgeführt. Dadurch weiterhin Fehlermeldung | tHB Kal durchgeführt, QC wiederholt. Oximetrie OK, pCO2 bei Level C außerhalb. Software installiert, SC aus Kundenbestand erneuert. pCO2 weiterhin grenzwertig hoch. Servicedump generiert. Reklamationsausdruck an RDEU gemailt. | SWO-502423 | Adjustment | HW | Oxi module | tHb Cal | 1011, 1012 | 20.02.2019, 08:10, Ms. Podstawa, tel. 023021756565: No Oxy values; Kal and QC error 20.02.2019, 08:54, Ms. Podstawa, Tel. 023021756565: Oxy Hardware Error | Data evaluated, error 1011 No zero point cal for OXI at calibration was fixed after /einschalten. However, the user had performed a tHb Kal with OXI hardware error. As a result, error message continues | tHB Kal performed, QC repeated. Oximetry OK, pCO2 at level C outside. Software installed, SC renewed from customer base. pCO2 remains borderline high. Service dump generated. Complaint statement sent to RDEU.
No | Visit | Complaint | RAME | Inlet module keeps coming off | Inlet is not securing properly | Verified inlet positioning. Attempted to reset the inlet mechanism. Verified everything was moving correctly. Used inlet from another analyzer to verify and compare inlet locking mechanism. Will have to order a new part. | SWO-502560 | Revisit required | HW | Auto Inlet module | Auto Inlet module | module keeps falling off | Inlet module keeps coming off | Inlet is not securing properly | Verified inlet positioning. Attempted to reset the inlet mechanism. Verified everything was moving correctly. Used inlet from another analyzer to verify and compare inlet locking mechanism. Will have to order a new part.
No | Visit | Complaint | RAME | Battery error 1599 | Found that the battery is still giving the error that the one cell is dead. | replaced battery and power supply to help correct the issue. | SWO-504111 | Defect | HW | Battery module | Battery module | 1599 | Battery error 1599 | Found that the battery is still giving the error that the one cell is dead. | replaced battery and power supply to help correct the issue.
No | Visit | Complaint | RAME | Oximetry failure in the WDC report | Oximetry values failed in the WDC report | Replaced hemolyzer | SWO-502569 | Defect | HW | Oxi module | Hemolyzer |  | Oximetry failure in the WDC report | Oximetry values failed in the WDC report | Replaced hemolyzer
No | Visit | Complaint | RAME | The analyzer screen is broken | Display broken and needs to be replaced | New display installed and tested ok. | SWO-502709 | Defect | HW | Display | Display |  | The analyzer screen is broken | Display broken and needs to be replaced | New display installed and tested ok.
No | Visit | Complaint | RAME | Error 581: OXI Spectrum mismatch | 1) Error 581: OXI Spectrum mismatch and Error: 1034 OXI Hardware. 2) OXI Complete Test - Hemolyzer hardware status: Hemolyzer driver voltage, hemolyzer turningstatus and caviation absorb failures. | 1) Replaced hemolyzer and pinch valve. Tested. 2) Replaced wet section printed circuit board. Error 581: OXI Spectrum mismatch is still present. | SWO-501885 | Revisit required | HW | Oxi module | Spectrometer |  | Error 581: OXI Spectrum mismatch | 1) Error 581: OXI Spectrum mismatch and Error: 1034 OXI Hardware. 2) OXI Complete Test - Hemolyzer hardware status: Hemolyzer driver voltage, hemolyzer turningstatus and caviation absorb failures. | 1) Replaced hemolyzer and pinch valve. Tested. 2) Replaced wet section printed circuit board. Error 581: OXI Spectrum mismatch is still present.
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | K定标灵敏度超范围 | 更换测试卡后修复 | SWO-502783 | [5 empty] | Test card, reagent package failure | K calibration sensitivity over range | Repair after replacing test card
No | Visit | Complaint | RCHN | 定标通过，K离子结果不准 | 定标通过，K离子结果不准 | 做了质控，样本采样处理后无问题by tim | SWO-503639 | [5 empty] | The bid is passed, the K ion result is not allowed | The bid is passed, the K ion result is not allowed | Quality control, sample sampling after no problem by tim
No | Visit | Complaint | RCHN | 机器报温度错误 | 机器报温度错误 | 了解机器的使用情况，建议将机器移至离窗口较远处使用后修复 by Andrew | SWO-502829 | [5 empty] | Machine reported temperature error | Machine reported temperature error | To understand the use of the machine, it is recommended to repair by Andrew after moving the machine farther away from the window
No | Visit | Complaint | RAME | Blue Screen | This problem was reported within the first 24 hours of start up. Machine was getting blue screen and machine freezing. I initially thought it was a problem with machine looking to communicate with network, but even after turning off those settings customer continued to report the problems. | Replaced ABL 80 CPU, E3825. I used the original CFAST card that came in the machine from factory because the one that came with the new board was not programmed and did not work. Software with new CF card printed out: 3.18.84.1 (FLEX) software on original card was : 3.18.85.3 (CO-OX) I monitored analyzer for multiple hours and performed various activities to try and replicate the problem. It appears to be working with new CPU. I performed all necessary calibrations. | SWO-502477 | [5 empty] | Blue Screen | This problem was reported within the first 24 hours of start up. Machine was getting blue screen and machine freezing. I initially thought it was a problem with machine looking to communicate with network, but even after turning off those settings customer continued to report the problems. | Replaced ABL 80 CPU, E3825. I used the original CFAST card that came in the machine from factory because the one that came with the new board was not programmed and did not work. Software with new CF card printed out: 3.18.84.1 (FLEX) software on original card was : 3.18.85.3 (CO-OX) I monitored analyzer for multiple hours and performed various activities to try and replicate the problem. It appears to be working with new CPU. I performed all necessary calibrations.
No | Visit | Complaint | RCHN | 仪器软件损坏 | 仪器软件损坏 | 重装系统和软件，重新定标。 by Andrew | SWO-502052 | Defect | SW | RMED SW | RMED SW | corrupt | Damage to instrument software | Damage to instrument software | Reload the system and software to re-label. by Andrew
No | Visit | Complaint | RCHN | 舱门开了关不上 | 舱门开了关不上，借903-028一个 | 更换903-028后修复 by Andrew | SWO-501412 | Defect | HW | Cover and Chassis | PCB, cabinet top | 4548 | The hatch is closed. | The hatch was closed, and it was 903-028. | Fix by Andrew after replacing 903-028
No | Visit | Complaint | RCHN | 做HCT标本时打问号，报错1264 | 做HCT标本时打问号，报错1264 | 清洁HCT模块后修复 by Andrew | SWO-502050 | Cleaned | HW | Aspiration Module | Needle Wash Unit | 1264 | When doing HCT specimen, put a question mark, report the error 1264 | When doing HCT specimen, put a question mark, report the error 1264 | Repair by Andrew after cleaning HCT module
No | Visit | Complaint | RCHN | 打印机无法打印 | 打印机无法打印 需要借件910-286 636-667 636-566 | 更换910-286后修复 by Andrew | SWO-500642 | Defect | HW | Analyzer Control Module | Printer Module |  | The printer can't print | Printer can't print Need to borrow 910-286 636-667 636-566 | Repair by Andrew after replacement 910-286
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 空气检测定标失败 | 更换测试卡后修复修复 | SWO-502767 | [5 empty] | Test card, reagent package failure | Failed air detection calibration | Repair repair after replacing test card
No | Visit | Complaint | RAME | DMS not connected to AU | Checked fuses and error log. Keys now showing as demo temp keys and only during reload of au. | Tried reloading au file and keys. Unsuccessfully. | SWO-502622 | Revisit required | HW | Basic Unit | PCB, AU |  | DMS not connected to AU | Checked fuses and error log. Keys now showing as demo temp keys and only during reload of au. | Tried reloading au file and keys. Unsuccessfully.
No | Visit | Complaint | RAME | DMS not connected to AU | Previously inspected and AU Pcb suspected | Replaced power supply and au pcb. Reloaded keys and rebooted. All systems operations now within spec. | SWO-502624 | Defect | HW | Analyzer Unit | PCB, analyzer unit |  | DMS not connected to AU | Previously inspected and AU Pcb suspected | Replaced power supply and au pcb. Reloaded keys and rebooted. All systems operations now within spec.
No | Visit | Complaint | RDEU | 19.2.2019 13:45 Frau Pelkmann 05451522095 -> meldet häufige Crea-Probleme | WG Crea Probleme | Gerät überprüft, kal und QC geprüft, QC neu konfiguriert . Kunden wegen Fan und Abläufe aufgeklärt. | SWO-502870 | Setup | SW | RMED SW | QC limits | CAP-789; FAN 915-387 | Feb 19, 2019 1:45 PM Ms. Pelkmann 05451522095 -> reports common Crea problems | WG Crea Problems | Device checked, kal and QC checked, QC reconfigured. customers informed about fan and processes.
No | Visit | Complaint | RSCH | Geräteaustausch und Repartur | CRP ausser Grenzen bei LQC | Optical Unit und Prozess Wheel ausgetauscht Alle Parameter neu kalibriert Funktionskontrolle i.o | SWO-502322 | Defect | HW | Process Module | Optical Unit | QC CRP | Device replacement and reparture | CRP out of the way at LQC | Optical Unit and Process Wheel Replaced All parameters recalibrated Functional control i.o
No | Visit | Complaint | RDEU | Klappe für Kassetten schließt nicht | Klappe für Kassetten schließt nicht | Gerät geprüft, demontiert und Klappe wieder eingehängt. Funktion im Service überprüft. | SWO-502871 | Adjustment | HW | Cover and Chassis | Door, Storage |  | Flap for cassettes does not close | Flap for cassettes does not close | device is tested, dismantled and flap remounted. Function in the service checked.
No | Visit | Complaint | RPOL | Błąd aspiracji. Igła nie wysuwa się. | Analizator sprawny. | Kalibracja uszczelki wlotu. | SWO-502310 | Adjustment | HW | Auto Inlet module | Inlet gasket |  | Aspiration error. The needle does not eject. | Functional Analyzer. | Calibration of Inlet gasket.
No | Visit | Complaint | RPOL | OXI liquid sensor error. odbiałczanie i regulacja LS nie pomaga. | Zatkany króciec wlotu powietrza. | Odetkanie króćca wlotu powietrza. Testy serwisowe ok. | SWO-502224 | Cleaned | HW | Fluidic manifold | Air intake |  | OXI Liquid sensor error. Deproturization and adjustment of LS does not help. | Clogged air inlet spigot. | The air inlet nozzle is cut off. Service tests approx.
No | Visit | Complaint | RPOL | Częste występowanie komunikatu niejednorodnej próbki przy badaniach zkapilary. | Testy kontrolne. | Wymiana uszczelki wlotu. | SWO-503405 | Defect | HW | Inlet module | Inlet gasket | 374 | Frequent occurrence of the message of heterogeneous sample in the studies with capillaries. | Control tests. | Replacing the inlet gasket.
No | Visit | Complaint | RPOL | Aparat zgłasza wykrycie przecieku. Użytkownik zasugerował że igła jest nienaturalnie zgięta. Użytkownik nie posiada ani nowej igły ani uszczelki wlotu. | Pokrzywiona igła. | Wymiana igły wraz z wężykiem. Wymiana uszczelki wlotu. Testy i kalibracje serwisowe. | SWO-503406 | Defect | HW | Inlet module | Inlet gasket | Leak detected | The camera reports a leak detection. The user suggested that the needle is unnaturally bent. You do not have a new needle or gasket in the inlet. | The needle is nesty. | Replacing the needle with the tubing. Replacing the inlet gasket. Tests and service calibrations.
No | Visit | Complaint | RHUN | A készülék képernyőjén néha felvillan egy windows-os 'error' felirat, ilyenkor úgy tapasztalják, hogy nem lesz jó a minta felszívás, ami után 10 perces mosás következik. | többször hibás a minta felszívás, a készülék képernyőjén a háttérben egy windows $netop hiba üzenet látható, a készülékhez nem működik a távoli elérés | teljes win7 és analizátor újra telepítése, beállítások visszaállítása, hálózati csatlakoztatás. Kalibráció, QC mérés | SWO-503022 | Defect | SW | External SW | Windows | pop up: "remote access…" | A Windows ' error ' is sometimes displayed on the screen of the device, so it is seen that there will be no good sample absorption, followed by a 10-minute wash. | Repeatedly faulty sample absorption, the device screen in the background a Windows /netop error message is displayed, Device does not work with the remote access | Full Win7 and Analyzer reinstallation, reset settings, network connection. Calibration, QC measurement
No | Visit | Complaint | RLTD | Noreen Tel: 01708 435000 x: 6845 21.02.19 Marie also called in today to report the same issue Tel: 01708 435000 x: 6845 | The Analyser was in ready mode when I arrived. Checked data logs. Checked fluid transport in service programs, all performed well. | No further action required. | SWO-503234 | No fault found | Fluidic pathway | Fluidic pathway | Fluidic pathway | Clot | Noreen Tel: 01708 435000 x: 6845 21.02.19 Marie also called in today to report the same issue Tel: 01708 435000 x: 6845 | The Analyser was in ready mode when I arrived. Checked data logs. Checked fluid transport in service programs, all performed well. | No further action required.
No | Visit | Complaint | RCHN | 信号不稳定 | 信号不稳定 借902-972一个（未更换备件使用） | 重新指导老师正确的操作后修复 by Andrew | SWO-502157 | [5 empty] | Signal instability | Signal instability Borrowing 902-972 one (not replaced with spare parts) | Fix by Andrew after re-instructing teacher sits correctly
No | Visit | Complaint | RCHN | 每天凌晨系统自动检测到有凝块 | 每天凌晨系统自动检测到有凝块 | 清洁管路后修复 by Andrew | SWO-502798 | Cleaned | HW | Oxi module | Hemolyzer | Clot | A clot is automatically detected by the system in the early hours of the morning every day | A clot is automatically detected by the system in the early hours of the morning every day | Repair by Andrew after cleaning lines
No | Visit | Complaint | RCHN | 频繁自动卸载程序 | 频繁自动卸载程序（未使用备件） | 重装软件后修复 | SWO-504198 | Defect | SW | RMED SW | RMED SW |  | Frequent auto-uninstall programs | Frequent automatic unloading procedures (spare parts not used) | Repair after reinstalling software
No | Visit | Complaint | RCHN | HCT测量出错 | HCT测量出错，借903-329一个 | 更换903-329后修复 by Andrew | SWO-502095 | Defect | HW | Aspiration Module | Needle Wash Unit | 1223 | HCT measurement error | HCT measurement error, borrow 903-329 one | Fix by Andrew after replacing 903-329
No | Visit | Complaint | RCHN | 干燥前温度超范围 | 干燥前温度超范围，借903-210一个 | 更换903-210后修复 by Andrew | SWO-501386 | Defect | HW | Process Module | PCB, cup dryer | 1069 | Overrange temperature before drying | Over-range temperature before drying, borrow 903-210 one | Fix by Andrew after replacing 903-210
No | Visit | Complaint | RCHN | 报错5684 卡杯子 | 报错5684 卡杯子 | 处理包取出来，倒出一些杯子后修复 by Andrew | SWO-502979 | Defect | Consumables | Solution Pack | Solution Pack | Stack | 5684 card cup sin | 5684 card cup sin | The handling bag is removed, some cups are poured out to repair by Andrew
No | Visit | Complaint | RCHN | 干燥前温度超出范围，报错1068 | 干燥前温度超出范围，报错1068，需要借件903-210和905-908各一个 | 更换903-210和905-908修复by tim | SWO-502818 | Defect | HW | Process Module | PCB, cup dryer | 1068 | Temperature out of range before drying, error 1068 | The temperature exceeded the range before drying, the error 1068, the need to borrow 903-210 and 905-908 each one | Replace 903-210 and 905-908 Repair by tim
No | Visit | Complaint | RCHN | 感应器故障 | 感应器故障，借903-672一个 | 更换903-672后修复 by Andrew | SWO-502008 | [5 empty] | Sensor failure | Sensor failure, borrow 903-672 one | Repair by Andrew after replacing 903-672
Yes | Visit | Complaint | RAME | Glu parameter will not come in Cleaning done, Membrane replacement done. | Glucose not out of range when FSE arrived for maintenance. | Corrected Lactate issue that was present on separate case CAS-477895-R6J3. Confirmed issue corrected by running calibration and Quality Checks | SWO-503876 | Re classified RMED | same day upSWO |  |  |  | Glu parameter will not come in Cleaning done, Membrane replacement done. | Glucose not out of range when FSE arrived for maintenance. | Corrected Lactate issue that was present on separate case CAS-477895-R6J3. Confirmed issue corrected by running calibration and Quality Checks
No | Visit | Complaint | RAME | Lactate failing calibration due to high sensitivity | Analyzer failing calibration due to high sensitivity. | Replaced Lactate membrane. Normal operations restored. | SWO-503397 | Assisted Replacement | Electrode | Membrane | Lac | Sens too high; mail sent: lot number of failing Lac membrane | Lactate failing calibration due to high sensitivity | Analyzer failing calibration due to high sensitivity. | Replaced Lactate membrane. Normal operations restored.
No | Visit | Complaint | RDEU | 25.01.2019, Frau Lang, Tel: 08131764002: Kal und QC Fehler bei Na und pO2 | Fehlersuche vor Ort Driften bei den Elektrolyten bei der Kal (dadruch QC Fehler) | Ventile aus EleMet und PhBG Modul ausgebaut und gereingit. Luftventile waren gehangen. Referenzelektrode neu membraniert. Ph Dichtring erneuert. Pumpenkal und LS Justage durchgeführt. Mehrere Kalibrerungen durchgeführt. Keine ausergewöhnlichen Driften mehr feststellbar. QCs für alle Level durchgeführt. ABL wieder messbereit. | SWO-503218 | Cleaned | HW | pH/BG Module | Valve, air |  | 25.01.2019, Frau Lang, Tel: 08131764002: Kal und QC Fehler bei Na und pO2 | Troubleshooting on site Drifting in electrolytes at the cal (dadruch QC error) | Valves made of EleMet and PhBG module removed and gereingit. Air valves were hung. reference electrode re-membraneed. Ph Sealing ring renewed. pump scale and LS adjustment. Several calibrations performed. No more unusual drifts detectable. QCs for all levels. ABL ready to be measured again.
No | Visit | Complaint | RDEU | Fehler QC Oxi. Oxi Spektrum mismatch. | Oxiventil hakt. | Oxiventil erneuert. | SWO-503287 | Defect | HW | Oxi module | Tube, valve | 902-527 | Fehler QC Oxi. Oxi Spektrum mismatch. | Oxiventil hakt. | Oxivalve renewed.
No | Visit | Complaint | RDEU | Inbetriebnahme d. Gerätes aufgrund div. Fehler nicht möglich | Sehr häufig QC-Abbrüche Level 3, Fehler 1218 & 1065, Probe wird ca 2,5 min nach dem Einlegen abgebrochen | Bildschirm getauscht, da die Rückwandbefestigung gebrochen war, Austausch Needlewashunit und Kalibrierung, Funktionskontrolle Servicedump | SWO-503066 | Defect | HW | Analyzer Control Module | Display | screws | Commissioning of the device due to various errors not possible | Very common QC aborted level 3, error 1218 & 1065, sample is aborted about 2.5 min after insertion | screen changed because the back panel attachment was broken, needlewashunit exchange and calibration, function control Service dump
No | Visit | Complaint | RPOL | Błędy kalibracji pCO2 | Weryfikacja wyników kalibracji | Wymiana elektrody pCO2 Kalibracja, kontrola. | SWO-503089 | Assisted Replacement | Electrode | Electrode | pCO2 | Cal | PCO2 Calibration Errors | Verification of calibration Results | Replacing the pCO2 electrode Calibration, control.
No | Visit | Complaint | RPOL | Szkodzenie roma.ini podczas wyłączenia analizatora. |  |  | SWO-504127 | Defect | SW | RMED SW | roma.ini |  | Damaging the ROM. ini when switching off the analyzer.
No | Visit | Complaint | RLTD | Philip Dixon 07341 268422 1.55pm - further call received from Phillip Dixon 07341 268422 - first fix worked for a short while, problem now reoccurred. 4.50pm - Philip Dixon called. Problem still occurring. Thinks may be an analyser fault. Tel 07341 268422. On site until around 5.15pm but can call mobile. | Return visit for same problem | Fit new oximetry pinch valve | SWO-503868 | Defect | HW | Oxi module | Pinch valve | Unable to pump | Philip Dixon 07341 268422 1.55pm - further call received from Phillip Dixon 07341 268422 - first fix worked for a short while, problem now reoccurred. 4.50pm - Philip Dixon called. Problem still occurring. Thinks may be an analyser fault. Tel 07341 268422. On site until around 5.15pm but can call mobile. | Return visit for same problem | Fit new oximetry pinch valve
No | Visit | Complaint | RAME | Error 581: OXI Spectrum mismatch | Error 581: OXI Spectrum mismatch | Replaced spectrometer and tested. | SWO-503484 | Defect | HW | Oxi module | Spectrometer |  | Error 581: OXI Spectrum mismatch | Error 581: OXI Spectrum mismatch | Replaced spectrometer and tested.
No | Visit | Complaint | RCHN | PH测量结果不准，血气频繁呼碱 | PH测量结果不准，血气频繁呼碱 | 分析数据，分析前误差培训，问题解决 by Andrew | SWO-504165 | [5 empty] | PH measurement results are not allowed, blood gas frequent exhalation alkali | PH measurement results are not allowed, blood gas frequent exhalation alkali | Analyze data, pre-analyze error training, problem solving by Andrew
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | LAC定标不过 | 更换测试卡后修复 | SWO-503562 | [5 empty] | Test card, reagent package failure | LAC calibration is not | Repair after replacing test card
No | Visit | Complaint | RAME | Fluid blockage | Hemolyzer blocked. | Could not clear blockage. Will order parts and repair ASAP. | SWO-504026 | [5 empty] | Fluid blockage | Hemolyzer blocked. | Could not clear blockage. Will order parts and repair ASAP.
No | Visit | Complaint | RAME | Fluid blockage. | Hemolyzer blocked. | Replaced hemolyzer, tubings, pumps, sensor and finally solution pack. Tried analog PCB but no help. | SWO-504021 | [5 empty] | Fluid blockage. | Hemolyzer blocked. | Replaced hemolyzer, tubings, pumps, sensor and finally solution pack. Tried analog PCB but no help.
No | Visit | Complaint | RCHN | 报泄漏，报错1342 | 报泄漏，报错1342 | 更换进样针和进样口支架后修复 by Andrew | SWO-503523 | Defect | HW | Inlet module | Inlet gasket | 1342 | Reported leak, error 1342 | Reported leak, error 1342 | Repair by Andrew after replacing the sample pin and the inlet bracket
No | Visit | Complaint | RCHN | 血氧故障，报错1034 | 血氧故障，报错1034，需要借件902-980 | 更换902-980后修复by tim | SWO-502901 | Defect | HW | Oxi module | Hemolyzer | 1034 | Blood oxygen fault, error 1034 | Blood oxygen fault, error 1034, need to borrow 902-980 | Repair by tim after replacement 902-980
No | Visit | Complaint | RCHN | 干燥前温度超出范围，报错1068 | 干燥前温度超出范围，报错1068，需要借件903-210 | 更换903-210后修复by tim | SWO-503493 | Defect | HW | Process Module | PCB, cup dryer | 1068 | Temperature out of range before drying, error 1068 | Temperature out of range before drying, error 1068, need to borrow 903-210 | Repair by tim after replacement 903-210
No | Visit | Complaint | RCHN | 报错5784，卡盘 | 报错5784，卡盘 | 调整存储盘位置后修复by tim | SWO-503764 | Adjustment | HW | Storage Module | Storage wheel | 5784 | 5784, card | 5784, card | Fix by tim after adjusting the storage disk position
No | Visit | Complaint | RCHN | 马达故障和punch up单元故障，报错5784 | 马达故障和punch up单元故障，报错5784，需要借件903-371和580-098一个 | 更换580-098后修复 by Andrew | SWO-502859 | Defect | HW | Incubator Module | Motor, punch up | 5784 | Motor failure and punch up unit failure, error 5784 | Motor failure and punch up unit failure, error 5784, need to borrow 903-371 and 580-098 one | Repair by Andrew after replacement 580-098
No | Visit | Complaint | RCHN | 更换处理包后液路报错，FTC-105 | 更换处理包后液路报错，FTC-105 | 手动抽吸管路后修复by tim | SWO-503579 | Defect | Consumables | Solution Pack | Solution Pack | FTC-105 | Liquid road error after replacement of treatment package, FTC-105 | Liquid road error after replacement of treatment package, FTC-105 | Repair by tim after manual suction line
No | Visit | Complaint | RCHN | 参数无法测量 | 参数无法测量，借636-373一个 | 更换636-373后修复by tim | SWO-502952 | Defect | HW | Oxi module | Lamp unit | Cable | The parameter cannot be measured | The parameter cannot be measured, borrow 636-373 one | Repair by tim after replacement 636-373
No | Visit | Complaint | RDEU | Das Ergebnisprotokoll ist unleserlich, auch Zahlenwerte | Dots ausgefallen | Drucker bestellt | SWO-503829 | Revisit required | HW | Printer Module | Printer Head |  | The result log is illegible, even numerical values | Dots fancy | Printer ordered
No | Visit | Complaint | RDEU | Spontane HCT Mess-Fehler | Kontrollbesuch wegen HCT Fehler bei Probenmessung. | Weiterhin spontane Messabbrüche wegen HCT Fehler. Mit Stationsleitung der Dialyse ( die auch Messungen am AQT durchführt ) gesprochen wegen Probenabnahme für AQT. Es wurde mir versichert das das Patientenblut in die üblichen Sarstetdprobennehmer direkt am Patienten abgenommen wird. Datendump und Servicedump erstellt. Weiteres vorgehen muß mit KorLo besprochen werden. | SWO-504392 | No fault found | HW | Aspiration Module | Needle Wash Unit | 1223, 1260, 1261 | Spontaneous HCT measurement errors | Inspection due to HCT error in sample measurement. | Further spontaneous measurement interruptions due to HCT errors. With station management of dialysis (which also performs measurements on AQT ) spoken because of sample sampling for AQT. I have been assured that the patient's blood is taken directly into the usual sarstetd samplers directly at the patient. Data dump and service dump. Further action must be discussed with KorLo.
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | K定标不过 0377 | 更换测试卡后修复 | SWO-509565 | [5 empty] | Test card, reagent package failure | K calibration is no more than 0377 | Repair after replacing test card
Yes | Visit | Complaint | RCHN | 光学强度超范围，报错1031 | 光学强度超范围，报错1031 借903-364一个，944-212（15327）缺货 |  | SWO-504124 | [5 empty] | Optical strength over range, error 1031 | Optical strength over range, error 1031 borrowed 903-364 one, 944-212 (15327) out of stock
No | Visit | Complaint | RCHN | TNI定标失败 | TNI定标失败，借903-364,944-212（15388），944-394（15362），944-258（15247），944-220（15486）各一个 | 收到的光单有质量问题，上机后仍报错1024，要再借903-364,944-212（15388），944-220（15486）各一个，此单先关单 by Andrew | SWO-502921 | Defect | HW | Process Module | Optical Unit | 1024 | TNI calibration failed | TNI bid failure, borrow edging 903-364, 944-212 (15388), 944-394 (15362), 944-258 (15247), 944-220 (15486) each | The light order received has quality problems, still reported error 1024 after the on-board, to borrow 903-364,944-212 (15388), 944-220 (15486) each one, this single first single by Andrew
No | Visit | Complaint | RCHN | TNI定标失败 | 收到的光单有质量问题，上机后仍报错1024，要再借903-364,944-212（15388），944-220（15486）各一个 |  | SWO-504136 | Defect | HW | Process Module | Optical Unit | 1024 | TNI calibration failed | The light list received has quality problems, still reported error 1024 after the on-board, to borrow 903-364, 944-212 (15388), 944-220 (15486) each
Yes | Visit | Complaint | RCHN | 光学强度超范围 | 光学强度超范围，借903-364一个 | 更换903-364后修复 by Andrew | SWO-503668 | [5 empty] | Optical strength over range | Optical strength over range, borrowed 903-364 a | Repair by Andrew after replacing 903-364
No | Visit | Complaint | RCHN | 测试卡断裂，报错1999 | 测试卡断裂，报错1999 ，断裂部位（Wings） | 更换测试卡后修复 by jesson | SWO-505830 | Defect | Consumables | Test Kit | Cartridge, Broken | Wings | Test card break, error 1999 | Test card break, error 1999, broken site (Wings) | Fix by jesson after replacing test card
No | Visit | Complaint | RAME | Oximetry failure in the WDC report | Failing internal QC OXi Values | Replaced spectrometer | SWO-504908 | Defect | HW | Oxi module | Spectrometer | QC | Oximetry failure in the WDC report | Failing internal QC OXi Values | Replaced spectrometer
No | Visit | Complaint | RDEU | 22.02.2019 13:45Uhr Frau Wenzler 07721934110: Der neue LP zeigt Fehler - Lösungen können nicht gepumpt werden 22.02.2019 14:55 Frau Kanar 07721934130 meldet, dass der ABL nach einen LP Wechsel nicht läuft. 22.02.2019 16:25Uhr Frau Djadjo 07721-934110: wir haben gestern ein neues BGA-Gerät in unser Zentrallabor am Standort VS bekommen, einen ABL90flex PLUS mit der SN 092R0196N016. Bereits heute haben wir nur Probleme mit dem Gerät. Obwohl keine Patientenprobe gemessen wurde, war die Kalibration nicht in Ordnung: inhomogene Kal3 Lösung. Das Gerät hat den Wechsel des Lösungspack empfohlen. Meine Kollegin hat den Lösungspack gewechselt und Sie per Telefon informiert. Auch nach einem zweiten Wechsel des Lösungspack nach dem Ausschalten des Gerätes auf Ihre Empfehlung, ist das Gerät nach wie vor nicht für den Einsatz bereit. Fehler 1086 und 1085 Inhomogene QC 1 und 2 Lösung. Das Problem sollte umgehend behoben werden, es handelt sich schließlich um ein neues Gerät und das sollte nicht schon am 2. Tag funktionsunfähig sein. | Gerät bei Eintreffen messbereit. | Herr Hutt MT hatte per Fernwartung bereits Serviceprogramme durchgeschalten. Ebenfalls Servicetests durchgeführt. Alle Tests in Ordnung. | SWO-504860 | Assisted Replacement | SP | SP | FT |  | 22.02.2019 13:45pm Mrs. Wenzler 07721934110: The new LP shows errors - solutions cannot be pumped 22.02.2019 14:55 Ms Kanar 07721934130 reports that the ABL is not running after an LP change. 22.02.2019 16:25Pm Mrs Djadjo 07721-934110: yesterday we got a new BGA device in our central laboratory at the VS site, an ABL90flex PLUS with the SN 092R0196N016. Already today we only have problems with the device. Although no patient sample was measured, the calibration was not correct: inhomogeneous Kal3 solution. The device has recommended changing the solution pack. My colleague changed the solution pack and informed you by phone. Even after a second change of the solution pack after turning off the device on your recommendation, the device is still not ready for use. Errors 1086 and 1085 Inhomogeneous QC 1 and 2 solution. The problem should be fixed immediately, it is a new device and this should not be inoperable on the 2nd day. | device ready to be measured upon arrival. | Mr Hutt MT had already switched through service programs by remote maintenance. Service tests also performed. All tests fine.
No | Visit | Complaint | RCHN | 无法开机，泵轮损坏 | 无法开机，泵轮损坏，借903-663、903-381和902-971各一个 |  | SWO-504379 | [5 empty] | Unable to power on, pump wheel damaged | Unable to power on, pump wheel damage, borrow 903-663, 903-381 and 902-971 each
No | Visit | Complaint | RIND | Lac failure. | Checked the analyser, lac calibration failed. | Performed zero check procedure and observed lac value unstable. Tried no. of times the zero check but lac value not stable. | SWO-507981 | [5 empty] | Lac failure. | Checked the analyser, lac calibration failed. | Performed zero check procedure and observed lac value unstable. Tried no. of times the zero check but lac value not stable.
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 多参数定标失败 | 更换测试卡正常。 | SWO-505199 | [5 empty] | Test card, reagent package failure | Multi-parameter calibration failed | The replacement test card is OK.
No | Visit | Complaint | RCHN | 试剂包漏液、仪器重启 | 1试剂包漏液、2 仪器重启 | 1 更换试剂包 2 仪器重启因为测试卡温度错误导致，更换测试卡后解决 by jesson | SWO-505817 | [5 empty] | Reagent bag leakage, instrument restart | 1 reagent bag leak, 2 instrument restart | 1 Replacement reagent pack 2 instrument restart due to test card temperature error, replace test card after resolving by jesson
No | Visit | Complaint | RKK | 透析液測定で当該機でカリウムの値がターゲット1.9mmol/Lに対し、2.3程になる。同液で検査室ABL800で測定したところ1.9程度、かつ代替器にて同一SCにて測定したところ1.9程度であった。 |  |  | SWO-507331 | [5 empty] | The value of potassium in the machine in the dialysate measurement is about 2.3 relative to the target 1.9mmol / L. About 1.9 was measured by the laboratory ABL800 with the same liquid, and was about 1.9 was measured by the same SC in an alternative.
No | Visit | Complaint | RCHN | 重连单机LIS | 重连单机LIS，未更换备件修复 | 重新安装单机LIS后修复by tim | SWO-510247 | Setup | SW | RMED SW | His/Lis |  | Reconnected stand-alone LIS | Reconnected stand-alone LIS, no replacement parts repair | Fix by tim after reinstalling stand-alone LIS
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 泵定标失败，报错1178 | 更换试剂包后修复 | SWO-504387 | Assisted Replacement | SP | SP | Pump cal | 1178 | Test card, reagent package failure | Pump calibration failed, error 1178 | Repair after replacement of reagent kit
Yes | Visit | Complaint | RCHN | 测试卡仓门无法关闭 | 测试卡仓门无法关闭 | 拆机重新调整后修复 by jesson | SWO-505072 | [5 empty] | Test card door cannot be closed | Test card door cannot be closed | Repair by jesson after disassembly
No | Visit | Complaint | RCHN | 报错3412错误 | 报错3412错误 | 冲洗管路，维修后正常 by jesson | SWO-505711 | Cleaned | HW | Aspiration Module | Needle Wash Unit | 3412 | 3412 errors | 3412 errors | Flush the line, normal after repair by jesson
Yes | Visit | Complaint | RCHN | 试管夹断裂 | 试管夹断裂，借903-584*2个 |  | SWO-504407 | [5 empty] | Test tube clip break | Test tube clip broken, borrowed 903-584 x 2
No | Visit | Complaint | RCHN | 报错1999 | 报错1999.61.22，借836-166一个 |  | SWO-504408 | Defect | HW | Incubator Module | Rubber disk, incubator | 1999.61.22 | Error 1999 | False 1999.61.22, loan 836-166 a
No | Visit | Complaint | RCHN | 液路问题，报错3412 | 液路问题，报错3412 | 检查后发现P4，P5泵头管路变形无弹性，拆卸泵头后重新组合泵头安装使用 by jesson | SWO-504447 | Adjustment | HW | Wet section | Pumphead | P4-P5 | Liquid Road Problem, Error 3412 | Liquid Road Problem, Error 3412 | Inspection found that P4, P5 pump head line deformation is inelastic, remove the pump head after recombining the pump head installation using by jesson
No | Visit | Complaint | RCHN | NA,CL病人结果与临床不符 | NA,CL病人结果与临床不符，未更换备件修复 | 做S7440质控后正常by tim | SWO-510251 | [5 empty] | NA, CL patient results are not clinically consistent | NA, CL patient results are not clinical, no replacement parts repair | Do S7440 after QC normal by tim
No | Visit | Complaint | RAME | Cal sen was high and would not pass cal | Analyzer was green when I arrived. I saw fluid on the spillway of the Ph $BG mod. | I cleaned and checked all chambers. After the first cal Ca sen was 106%. I replaced the ref electrode and membrane then all came in the first cal. A few cal's later it went back up. I replaced the Ca membrane and it went higher. I replaced the Ph Oring and it dropped down to normal | SWO-504998 | Defect | Electrode | Electrode | Sealing ring, pH | Ca sens high | Cal sen was high and would not pass cal | Analyzer was green when I arrived. I saw fluid on the spillway of the Ph /BG mod. | I cleaned and checked all chambers. After the first cal Ca sen was 106%. I replaced the ref electrode and membrane then all came in the first cal. A few cal's later it went back up. I replaced the Ca membrane and it went higher. I replaced the Ph Oring and it dropped down to normal
No | Visit | Complaint | RDEU | 25.01.2019, Frau Lang, Tel: 08131764002: Kal und QC Fehler bei Na und pO2 | Ref Cal error | Messkammern gereinigt, Proteinentfernung und LS Justage durchgeführt, Referenz Elektrode erneuert und neu membraniert, Mehrere Kalibrationen und QC's vorgenommen. | SWO-505093 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | 25.01.2019, Frau Lang, Tel: 08131764002: Kal und QC Fehler bei Na und pO2 | Ref Cal error | Measuring chambers cleaned, protein removal and LS adjustment performed, reference electrode renewed and re-membraneed, several calibrations and QC's performed.
No | Visit | Complaint | RAME | Software issues | Status Code 9, 2, and 10 during boot up. | Used Roma.ini and roma.db backup files to restore the software. Attempted to replace roma.ini file with a new roma.ini file. Installed application software first with all data, then only with database, then only with set up. All software installations failed. Installed application software with all new data. Attempted to replace roma.ini and roma.db from original software. Both original roma.ini and roma.db showed corruption. No data was transferred. Ran tHb calibration and installed QC. Customer notified. | SWO-504944 | Defect | SW | RMED SW | roma.ini |  | Software issues | Status Code 9, 2, and 10 during boot up. | Used Roma.ini and roma.db backup files to restore the software. Attempted to replace roma.ini file with a new roma.ini file. Installed application software first with all data, then only with database, then only with set up. All software installations failed. Installed application software with all new data. Attempted to replace roma.ini and roma.db from original software. Both original roma.ini and roma.db showed corruption. No data was transferred. Ran tHb calibration and installed QC. Customer notified.
No | Visit | Complaint | RDEU | 22.02.2019, Frau Ketterer, Tel. 076512978437: FlexQ Modul liest keine Barcodes | Barcode wird nicht korrekt zugeordnet. | Barcode auf Auftrags-Id eingestellt. | SWO-504856 | Defect | HW | FlexQ Module | Barcode reader |  | 22.02.2019, Mrs Ketterer, tel. 076512978437: FlexQ module does not read barcodes | Barcode is not mapped correctly. | Barcode set to order id.
No | Visit | Complaint | RAME | Ionized Calcium Calibration Failure | Ionized Calcium Calibration Failure. Failing sensitivity on calibration. Had previously replaced just pH module and corrected issue temporarily. | Replaced pH $BG module and el $MET module to correct calcium failure. Serial number for removed pH module 902-823R2912N009. Serial number for removed el module 902-824R2766N004. Placed analyzer live and removed old analyzer. Will pack up both old analyzers and customer will ship back using RGA Y18-73439L. | SWO-504644 | Defect | HW | El/Met Module | El/Met Module | Ca | Ionized Calcium Calibration Failure | Ionized Calcium Calibration Failure. Failing sensitivity on calibration. Had previously replaced just pH module and corrected issue temporarily. | Replaced pH $BG module and el $MET module to correct calcium failure. Serial number for removed pH module 902-823R2912N009. Serial number for removed el module 902-824R2766N004. Placed analyzer live and removed old analyzer. Will pack up both old analyzers and customer will ship back using RGA Y18-73439L.
No | Visit | Complaint | RPOL | Błędnie zainstalowana kaseta kalibracyjna - kalibracja niemożliwa. | Brak wykonania kalibracji na zainstalowanej kasecie kalibracyjnej w dozwolonym czasie. | Wykasowanie plików z historią wymiany kaset, kalibracja B-hCG. Analizator sprawny. | SWO-504888 | Assisted Replacement | Consumables | Test Kit | Cartridge, cal |  | Incorrectly installed calibration cartridge-calibration not possible. | No calibration is performed on the installed calibration cartridge in the allowed time. | Erase files with cassette replacement history, B-hCG calibration. Functional Analyzer.
No | Visit | Complaint | RAME | Inlet handle described to On call as broken in pieces | arrived on site and escorted by Chad to ER. Ilet probe bent and gasket is broken into several pieces | Replaced Inlet probe. Replaced Inlet gasket (PM kit). Replaced Inlet clip. Waited for pCO2, Glucose, Lactate and Na to settle. Replaced SC and printed credit form. | SWO-505068 | Defect | HW | Inlet module | Inlet gasket |  | Inlet handle described to On call as broken in pieces | arrived on site and escorted by Chad to ER. Ilet probe bent and gasket is broken into several pieces | Replaced Inlet probe. Replaced Inlet gasket (PM kit). Replaced Inlet clip. Waited for pCO2, Glucose, Lactate and Na to settle. Replaced SC and printed credit form.
No | Visit | Complaint | RKOR | Inlet module position error | . | Service mode check. Position adjust. Inlet probe position adjust. | SWO-512467 | Adjustment | HW | Auto Inlet module | Auto Inlet module |  | Inlet module position error | . | Service mode check. Position adjust. Inlet probe position adjust.
No | Visit | Complaint | RPAC | 1559: Battery error detected. One or more failing cells in the battery detected. Analyzer will be locked if battery is not removed or replaced shortly (4 weeks = 13th March) Advised David we have 4 weeks to replaced it so one of us will be there as soon as possible as traffic light will remain yellow until fixed. | Battery error 1559 | Saved setup & logs, backed up data. Shut down analyzer and replaced battery module. Restarted & reinitialized analyzer, battery error was cleared. Calibration & QCs all OK. | SWO-505246 | Defect | HW | Battery module | Battery module | 1559 | 1559: Battery error detected. One or more failing cells in the battery detected. Analyzer will be locked if battery is not removed or replaced shortly (4 weeks = 13th March) Advised David we have 4 weeks to replaced it so one of us will be there as soon as possible as traffic light will remain yellow until fixed. | Battery error 1559 | Saved setup & logs, backed up data. Shut down analyzer and replaced battery module. Restarted & reinitialized analyzer, battery error was cleared. Calibration & QCs all OK.
No | Visit | Complaint | RAME | Battery Error | Error 1559. Battery error, replaced battery no light troubleshooting took battery control bed and swapped with previous battery pack | Replaced battery assembly | SWO-505511 | Defect | HW | Battery module | Battery module | 1559 | Battery Error | Error 1559. Battery error, replaced battery no light troubleshooting took battery control bed and swapped with previous battery pack | Replaced battery assembly
No | Visit | Complaint | RDEU | 26.02.2019 9:20 H. Haberkamp 02083092181: Interventionsmeldung Endlosschleife; kein Material mehr vor Ort | Material für Crea $Urea ABL90 nicht lieferbar. | Crea Modus deaktiviert, Sc und LP ohne Crea (durch Kundeneigene Betriebsmittel) installiert. Crea Modus wird wieder aktiviert sobald die Crea spezifischen ABL90Flex Plus Betriebsmittel wieder lieferbar sind. | SWO-507551 | Setup | SW | RMED SW | Crea/Urea |  | 26.02.2019 9:20 Haberkamp 02083092181: Intervention notification Endless loop; no more material on site | Material for Crea /Urea ABL90 not available. | Crea mode disabled, Sc and LP installed without Crea (by customer's own resources). Crea mode is reactivated as soon as the Crea specific ABL90Flex Plus equipment is available again.
No | Visit | Complaint | RLTD | New analyser not yet in use failing to pump solutions. | New analyser not yet in use failing to pump solutions. | Customer and Sale rep had both attempted to resolve. Arrived on site an ran a visual flow check, fluids flowing very sporadically, attempted to flush analyser from sensor cassette to waste, unable to get a steady flow. Seperated componants and flushed each part and reassembled, I now have a steady stream. Now passing fluid transport tests, MSS to monitor during training. | SWO-505659 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway | Clot | New analyser not yet in use failing to pump solutions. | New analyser not yet in use failing to pump solutions. | Customer and Sale rep had both attempted to resolve. Arrived on site an ran a visual flow check, fluids flowing very sporadically, attempted to flush analyser from sensor cassette to waste, unable to get a steady flow. Seperated componants and flushed each part and reassembled, I now have a steady stream. Now passing fluid transport tests, MSS to monitor during training.
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | LAC定标失败 | 更换测试卡后修复 | SWO-511784 | [5 empty] | Test card, reagent package failure | LAC calibration failure | Repair after replacing test card
No | Visit | Complaint | RAME | Fluid blockage | Sensor check failures due to possible fluid blockage. | Replaced hemolyzer, liquid sensor and installed maintenance kit. Kept getting 1218 errors. Used 2 ea 944-341 solution packs and one ea. 945-843 sensor cassette. | SWO-505589 | [5 empty] | Fluid blockage | Sensor check failures due to possible fluid blockage. | Replaced hemolyzer, liquid sensor and installed maintenance kit. Kept getting 1218 errors. Used 2 ea 944-341 solution packs and one ea. 945-843 sensor cassette.
No | Visit | Complaint | RDEU | Coox Control Board Fehler | Gerät meldet morgens weiterhin Coox Control Board Fehler Coox Board wurde bereits erneuert! | Gerät geprüft , - um Fehler einzukreisen, wurde die Original CF Karte mit Win7 wieder zurückgebaut. Remote Support ist momentan nicht möglich. | SWO-505814 | [5 empty] | Coox Control Board Error | Device continues to report Coox Control Board error in the morning Coox Board has already been renewed! | Device checked - to encircle errors, the original CF card with Win7 was restored. Remote support is not possible at this time.
No | Visit | Complaint | RCHN | 无法抽吸溶液 1343 | 无法抽吸溶液 1343,清洁进样支架和进样针恢复。未更换备件。 | 清洁进样支架和进样针恢复。未更换备件。 by jesson | SWO-505723 | Cleaned | HW | Inlet module | Inlet gasket | 1343 | Unable to pump solution 1343 | Unable to pump solution 1343, clean the sample holder and the sample needle to recover. Spare parts have not been replaced. | Clean the sample holder and the feed needle to recover. Spare parts have not been replaced. by jesson
No | Visit | Complaint | RCHN | 检测到泄漏 | 检测到泄漏 | 自行更换841-811后修复 by jesson | SWO-505212 | Revisit required | HW | Oxi module | Pinch valve tube | 1341 | A leak has been detected | A leak has been detected | Repair by jesson after replacing 841-811
No | Visit | Complaint | RCHN | 存储盘卡盘 | 存储盘卡盘 | 存储盘上有黑颜色的块掉落导致，重新安装后修复 by jesson | SWO-505087 | Adjustment | HW | Storage Module | Cartridge guide | 5784 | Memory disk card | Memory disk card | Blocks with black color on the tray drop cause, reinstall after repair by jesson
No | Visit | Complaint | RPOL | Wykryto powietrze w układzie pomiarowym. | Diagnoza problemu. | Reinstalacja pakietu odczynnikowego oraz kasety sensorycznej. Oczyszczenie toru pomiarowego, oczyszczenie kasety sensorycznej. Testy serwisowe. Analizator sprawny. W razie ponownych problemów zalecana wymiana kasety sensorowej. | SWO-505266 | [5 empty] | Air detected in the measuring system. | Diagnosis of the problem. | Reinstallation of the reagent pack and the sensory cassette. Cleaning the measuring track, cleaning the sensor cassette. Service tests. Functional Analyzer. In case of reissue, the recommended replacement of the sensor cassette.
No | Visit | Complaint | RDEU | 01.10.2018 08:30Uhr Herr Glienke 03334691503: Das Gerät meldet häufig Driftalarm bei Elektrolyte | Diverse Driften bei Elektrolyten, Referenz wurde bereits mehrfach mit verschiedenen Lot D711 membraniert | Kontrolle Lösungstransport, Kontrolle E1001 auf Risse bzw. Verschleiß, Kontrolle Ventilfunktionen im Service, starke Differenzen bei ModulpumpenschlauchKalibrierung, beide Modulpumpenschläuche ersetzt, PumpenKal ok, korrekte Kal abgewartet und QK ok | SWO-505796 | Revisit required | Electrode | Electrode | Reference | CAPA-733 | 01.10.2018 08:30Am Mr Glienke 03334691503: The device frequently reports drift alarm for electrolytes | Various drifts in electrolytes, reference has already been membraned several times with different lot D711 | Control solution transport, control E1001 for cracks or wear, control valve functions in service, strong differences in module pump hosecalibration, both module pump hoses replaced, pumpsKal ok, correct scale waited and QK ok
No | Visit | Complaint | RIND | Rinse error. | Check the system. | Clean the rinse & air valve. Calibrate the LS. | SWO-508909 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | Rinse error. | Check the system. | Clean the rinse & air valve. Calibrate the LS.
No | Visit | Complaint | RKK | 測定時Lac不安定 メンブラン交換しても頻発するため電極交換 | 測定時Lac不安定 | Lac電極交換 | SWO-505176 | Assisted Replacement | Electrode | Electrode | Lac |  | Lac instability at the time of measurement Because it occurs frequently even when the membrane is replaced, the electrode exchange | Lac instability at the time of measurement | Lac electrode replacement
No | Visit | Complaint | RAME | Flex-Q Barcode scanner issues | Unit will not scan customers barcodes but will scan safe pics barcodes. Unit was working up until yesterday. Customers second analyzer can scan barcodes. | Checked AutoCheck alignment. OK. Laser seems weak. | SWO-505678 | Defect | HW | FlexQ Module | Barcode reader |  | Flex-Q Barcode scanner issues | Unit will not scan customers barcodes but will scan safe pics barcodes. Unit was working up until yesterday. Customers second analyzer can scan barcodes. | Checked AutoCheck alignment. OK. Laser seems weak.
No | Visit | Complaint | RIND | Neon Lamp to High. | Neon Lamp too high. | Replaced Lamp unit and Oxi module PCB, but problem is remaining same, replaced the Oxi module found Oxi parameters are working instrument is functional. | SWO-511039 | Defect | HW | Oxi module | Oxi module | Neon number too high | Neon Lamp to High. | Neon Lamp too high. | Replaced Lamp unit and OXI module PCB, but problem is remaining same, replaced the Oxi module found OXI parameters are working tool is functional.
No | Visit | Complaint | RDEU | 27.02.2019, SR. Merah, Tel: 02191138582: Monitor defekt | Cover defekt, an 3 Stellen die Büchsen gerissen, keine äußerlichen Schäden erkennbar, | Cover von Kollegen geholt, altes Cover demontiert, neues Cover Verbaut Neustart, Alles grün | SWO-507225 | Defect | HW | Display | Display | screws | 27.02.2019, SR. Merah, Tel: 02191138582: Monitor defekt | Cover defective, in 3 places the cans torn, no external damage recognizable, | Cover picked up by colleagues, old cover dismantled, new cover built-in reboot, All Green
Yes | Visit | Complaint | RDEU | 27.02.2019, SR. Merah, Tel: 02191138582: Monitor defekt | Schraubenhalterung im Cover an 3 Stellen gerissen | Kein Ersatzteil auf Lager -- Erstzteil Beschaffung | SWO-507224 | Re classified RMED | same day upSWO |  |  |  | 27.02.2019, SR. Merah, Tel: 02191138582: Monitor defekt | Screw holder in the cover torn in 3 places | No spare part in stock -- First part procurement
No | Visit | Complaint | RDEU | Gehäuse am Monitor gebrochen | Monitorgehäuse erneuern, Es sind 3 Gewinde am vorderen Rahmen abgebrochen. Vermutlich sind die Schrauben zu fest angezogen worden. | Monitor ausgebaut, Gehäuse ersetzt. | SWO-507096 | Defect | HW | Display | Display | screws | Housing on the monitor broken | Renew monitor housings, There are 3 threads broken off on the front frame. Presumably the screws have been tightened too tightly. | Monitor removed, housing replaced.
No | Visit | Complaint | RLTD | Ariun 07999562191 | The customer has requested setting the Analyser up with Reference ranges. Maintenance training and Analyser settings check need to be carried out. | Settings and functions have been tested. Ran through the Analyser settings for Reference ranges. | SWO-506981 | User Error, FSE | SW | RMED SW | Demo mode | Left in Demo mode | Ariun 07999562191 | The customer has requested setting the Analyser up with Reference ranges. Maintenance training and Analyser settings check need to be carried out. | Settings and functions have been tested. Ran through the Analyser settings for Reference ranges.
No | Visit | Complaint | RDEU | 26.02.2019; Hr. Koczielka; Tel.: 03461271050; regelmäßige pCO2 QC Fehler | pCO2 QC Fehler im Lev. A | kein Fehler mehr vor Vorhanden $ letzte QC war ok Software aktualisiert pCO2 anschl. wieder knapp außerhalb Bereich im LevA SC erneuert | SWO-506919 | Defect | SW | RMED SW | pCO2 | QC A | 26.02.2019; Mr. Koczielka; Tel.: 03461271050; regular pCO2 QC error | pCO2 QC Fehler im Lev. A | no more error before existing QC was ok Software updated pCO2 followed by again just outside the area in the LevA SC renewed
No | Visit | Complaint | RKOR | Quick access error | . | OS, Software reinstall. | SWO-512470 | Defect | SW | RMED SW | RMED SW | Quick access | Quick access error | . | OS, Software reinstall.
Yes | Visit | Complaint | RIND | Verification Check failed error for SC & SP. | Verification Check failed error for SC & SP. | Found error V $C failed for 4 number of SC and SP, Diagnosed with Signal data PCB and Main distribution PCB, but the issue with Sensor connector instrument is under warranty, provided the standby sensor connector, Need to replace with new. 902-972. | SWO-507425 | [5 empty] | Verification Check failed error for SC & SP. | Verification Check failed error for SC & SP. | Found error V $C failed for 4 number of SC and SP, Diagnosed with Signal data PCB and Main distribution PCB, but the issue with Sensor connector instrument is under warranty, provided the standby sensor connector, Need to replace with new. 902-972.
No | Visit | Complaint | RCHN | 检测到泄漏 | 夹紧管漏气 | 更换841-811后修复 by jesson | SWO-508095 | Defect | HW | Oxi module | Pinch valve tube | 1341 | A leak has been detected | Clamping tube leak | Repair by jesson after replacement 841-811
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 泄漏 1342 | 更换测试卡后修复 | SWO-508101 | [5 empty] | Test card, reagent package failure | Leak 1342 | Repair after replacing test card
No | Visit | Complaint | RCHN | BNP温度传感器不存在 | BNP温度传感器不存在 | 自行更换905-908后修复 by jesson | SWO-507421 | Defect | HW | Process Module | Sensor, BNP | 1235 | BNP temperature sensor does not exist | BNP temperature sensor does not exist | Repair by jesson after replacing 905-908
No | Visit | Complaint | RIND | pH values are not coming for all patient samples. | pH values are not coming for all patient samples,clean the chamber and electrode then done the calibration,pH calibration is not stable then replaced the pH electrode and done the calibration its working good. Run the patient samples all results are good. | pH values are not coming for all patient samples,clean the chamber and electrode then done the calibration,pH calibration is not stable then replaced the pH electrode and done the calibration its working good. Run the patient samples all results are good. | SWO-508517 | Assisted Replacement | Electrode | Electrode | pH |  | pH values are not coming for all patient samples. | pH values are not coming for all patient samples,clean the chamber and electrode then done the calibration,pH calibration is not stable then replaced the pH electrode and done the calibration its working good. Run the patient samples all results are good. | pH values are not coming for all patient samples,clean the chamber and electrode then done the calibration,pH calibration is not stable then replaced the pH electrode and done the calibration its working good. Run the patient samples all results are good.
No | Visit | Complaint | RIND | Print not clear | Print not clear | Cleaned the printer head , checked the connections and checked the print , but still the print is not clear .need to replace printer head | SWO-506950 | Revisit required | HW | Printer Module | Printer Head |  | Print not clear | Print not clear | Cleaned the printer head , checked the connections and checked the print , but still the print is not clear .need to replace printer head
No | Visit | Complaint | RKK | O2モジュール不良 | チャンバーから外しても測定が開始されない | モジュール交換 | SWO-517531 | [5 empty] | O2 module bad | Measurement does not start even if removed from the chamber | Module replacement
No | Visit | Complaint | RKK | バッテリーエラーが検出されました | バッテリーエラーが検出されました | バッテリーエラー交換実施 | SWO-510198 | Defect | HW | Battery module | Battery module | 1559 | Battery error detected | Battery error detected | Battery error replacement
Yes | Visit | Complaint | RDEU | 10.01.2019, 08:43 Uhr, Sr. Sandra, Tel. 05215811351: Kd. fragt nach Einweisungstermin | Einweisung | Einweisung ABL90 | SWO-508171 | Re classified RMED | planned SWO |  |  | mail: not upSWO, agree? | 10.01.2019, 08:43, Sr. Sandra, tel. 05215811351: Kd. asks for briefing date | Instruction | Instruction ABL90
No | Visit | Complaint | RLTD | Baracat 02089343299 or 02089342052 | Software $flash card failure - failing at dmi pooling data $harddisk failure | Replaced flash card and reloaded software copied setup $database from Majors ED analyser. Named analyser accordingly. | SWO-507729 | Defect | HW | DMS_PC | CF card | frozen screen | Baracat 02089343299 or 02089342052 | Software /flash card failure - failing at dmi pooling data /harddisk failure | Replaced flash card and reloaded software copied setup $database from Majors ED analyser. Named analyser accordingly.
No | Visit | Complaint | RLTD | Barracat 0208 934 3299 or 0208 934 2052 | Replaced pinch valve tube after noticing fluid leakage. | Replaced pinch valve on reboot noticed power cable $wire unseated from connecter causing inlet module to be inoperable - reseated $ all okay | SWO-508035 | Defect | HW | Oxi module | Pinch valve tube |  | Barracat 0208 934 3299 or 0208 934 2052 | Replaced pinch valve tube after noticing fluid leakage. | Replaced pinch valve on reboot noticed power cable $wire unseated from connecter causing inlet module to be inoperable - reseated $ all okay
No | Visit | Complaint | RCHN | 屏幕不亮 | 屏幕不亮，需要借件 552-039（保外借件工程师不需要了） | 更换备用机上的552-039后修复 by jesson | SWO-505780 | [5 empty] | The screen is not lit | The screen is not bright, 552-039 is required (no need for the loan erronied agent) | Repair by jesson after replacing 552-039 on the standby
No | Visit | Complaint | RCHN | 频繁LAC定标不过 | 频繁LAC定标不过 借902-972一个 | 引脚有腐蚀，更换902-972后修复。 by jesson | SWO-505861 | [5 empty] | Frequent LAC calibration however | Frequent LAC calibration but borrow 902-972 one | Pins are corroded and repaired after replacing 902-972. by jesson
No | Visit | Complaint | RCHN | 软件不匹配 | 软件不匹配 | 重新安装系统及软件后修复 by jesson | SWO-507712 | Adjustment | SW | RMED SW | RMED SW | 0002 | Software mismatch | Software mismatch | Fix by jesson after reinstalling the system and software
No | Visit | Complaint | RCHN | TNI定标失败 | 加借903-364,944-212（15388），944-258（15247）各一个 | 更换903-364一个，重新定标正常。by jesson | SWO-505007 | Defect | HW | Process Module | Optical Unit | 1024 | TNI calibration failed | Borrow 903-364, 944-212 (15388), 944-258 (15247) each | Replace the 903-364 one and re-label normally. by jesson
No | Visit | Complaint | RCHN | 无法识别处理包，报错1999 | 无法识别处理包，报错1999,需要借件903-036 | 更换903-036后修复 by jesson | SWO-505805 | Defect | Consumables | Solution Pack | Solution Pack | PCB | The processing package was not recognized, and the error was reported in 1999 | Unrecognizable processing package, error 1999, need to borrow 903-036 | Repair by jesson after replacement 903-036
No | Visit | Complaint | RCHN | 定标卡扫描故障 排查原因。重新定标 无报错代码 | 定标卡扫描故障 排查原因。重新定标 无报错代码 | 检查定标卡，检查仪器事件日志，调整扫描器角度。重新定标后正常 by jesson | SWO-508223 | Adjustment | HW | Storage Module | Barcode reader, storage wheel |  | The bid card scan fault troubleshooting the cause. Re-calibration No error code | The bid card scan fault troubleshooting the cause. Re-calibration No error code | Check the calibration card, check the instrument event log, adjust the scanner angle. Normal by jesson after re-calibration
No | Visit | Complaint | RCHN | 进样仓门无法正常开合 | 进样仓门无法正常开合，借903-028,886-638,886-639各一个 | 更换903-028,886-638,886-639后修复 by jesson | SWO-505879 | Defect | HW | Cover and Chassis | PCB, cabinet top |  | The inlet warehouse door does not open and close properly | The entrance door does not open and close properly, borrowing 903-028, 886-638, 886-639 each | Fix by jesson after replacing 903-028, 886-638, 886-639
No | Visit | Complaint | RKK | pO2の測定値が高め。QCレベル３でpO2値が高めに範囲外 | 機器状況確認 | 装置返却作業 | SWO-508083 | [5 empty] | The measurement value of pO2 is increased. High pO2 value at QC level 3 | Equipment status check | Equipment return work
No | Visit | Complaint | RFRA | Gel de l’écran | Contrôle de la configuration du PC | Changement de la carte mère. Essais | SWO-507595 | Defect | HW | DMS_PC | CPU |  | Screen freeze | Controlling the PC configuration | Changing the motherboard. Tests
No | Visit | Complaint | RIND | Service Visit. | Blood is coming from the inlet gas kit. | Do the protein removal 5 times, hold the analyzer, clean the inlet gaskit, inlet probe and connecting tubing, successfully done, recorded no replacement, restart the analyzer and wait for auto two point calibration, successfully done. Analyzer handed over to staff in well working condition. | SWO-508512 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | Service Visit. | Blood is coming from the inlet gas kit. | Do the protein removal 5 times, hold the analyzer, clean the inlet gaskit, inlet probe and connecting tubing, successfully done, recorded no replacement, restart the analyzer and wait for auto two point calibration, successfully done. Analyzer handed over to staff in well working condition.
No | Visit | Complaint | RDEU | Demo Key's abgelaufen | Demo Key's abgelaufen | Key's angefordert und aufgespielt. Abgelaufenen QC aus Karoussel entfernt und entsorgt. Neue QC konfiguriert , diverse kal und QC gestartet. | SWO-508172 | User Error, FSE | Demo key | Demo key | Demo key |  | Demo Key's expired | Demo Key's expired | Key's requested and played. Expired QC removed from karoussel and disposed of. New QC configured, various kal and QC started.
No | Visit | Complaint | RDEU | pO2 und pCO2 Meßwerte unglaubwürdig | Zeitweise pCO2 außerhalb QC Grenzen |  | SWO-508365 | Assisted Replacement | SC | SC | pCO2 |  | pO2 and pCO2 readings implausible | Intermittent pCO2 outside QC limits
No | Visit | Complaint | RDEU | Gerät lässt sich im Helios Netzwerk nicht anpingen | Anpingen zulassen, Uhrzeit geht ein Stunde nach | Gerät überprüft: falsche Zeitzone gewählt -korrigiert, danach Uhrzeit okay, Firewall Einstellung gemäß RDEU Anleitung eingestellt, Geräteneustart, Prüfung Ping von PC im Op -> nicht möglich, Info an TS Kal & QC Daten gemäß RDEU Anleitung gesichert und dokumentiert | SWO-508333 | Setup | SW | RMED SW | RMED SW | ping | Device cannot be pinged in Helios network | Allow anping, time goes one hour after | Device checked: wrong time zone selected -corrected, then time okay, firewall setting according to RDEU instructions set, device restart, checking ping of PC in op -> not possible, info to TS Kal & QC data backed up and documented in accordance with RDEU instructions
No | Visit | Complaint | RLTD | Adam on 0208 2164934 Error message for 3 days-Leak current in relation to SP detected at midnight . | Error 0964 message for 3 days-Leak current in relation to SP detected at midnight . | Removed wet section and stripped and checked for leaks ok. Replaced pinch valve tube as worn. Pressure tested system ok. Run impedance test passed ok service programs passed ok. Run system check passed ok. For midnight ok. | SWO-508514 | Defect | HW | Oxi module | Pinch valve tube | 964 | Adam on 0208 2164934 Error message for 3 days-Leak current in relation to SP detected at midnight . | Error 0964 message for 3 days-Leak current in relation to SP detected at midnight . | Removed wet section and stripped and checked for leaks ok. Replaced pinch valve tube as worn. Pressure tested system ok. Run impedance test passed ok service programs passed ok. Run system check passed ok. For midnight ok.
No | Visit | Complaint | RBNL | Vele tbil onder rapporteergrens (<-2.1) | Probleem met tBil (patient resultaten beneden onderste rapporteergrens) is reeds lange tijd bezig | Vervanging van spectrometer, tHb calibratie , QC en patientstalen gemeten :OK | SWO-508355 | Defect | HW | Oxi module | Spectrometer | Bili | Many Tbil under report limit (<-2.1) | Problem with tBil (patient results below bottom report limit) has been busy for a long time | Replacement of spectrometer, tHb calibration, QC and Patientsteel measured: OK
No | Visit | Complaint | RKOR | ABL/DMS PC connected to wet section | . | OS Software reinstall (failed.), Demo ABL90 replace. | SWO-512477 | Revisit required | Analyzer removed | Analyzer removed | Analyzer removed | PC connection | ABL/DMS PC connected to wet section | . | OS Software reinstall (failed.), Demo ABL90 replace.
Yes | Visit | Complaint | RCHN | 接口板腐蚀 | 接口板腐蚀，需要借件902-972 | 更换902-972后修复 by jesson | SWO-505878 | [5 empty] | Interface plate corrosion | Interface plate corrosion, need to borrow 902-972 | Repair by jesson after replacement 902-972
No | Visit | Complaint | RAME | Received 581 and 292 when running samples | 292 and 581 errors on patient samples. | Ran service test. No issues. Replaced Hemolyzer and pinch valve. | SWO-508569 | [5 empty] | Received 581 and 292 when running samples | 292 and 581 errors on patient samples. | Ran service test. No issues. Replaced Hemolyzer and pinch valve.
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | NA,CL,PCO2定标失败，1319，1321，0210 | 更换测试卡后修复 | SWO-510283 | [5 empty] | Test card, reagent package failure | NA, CL, PCO2 bid failure, 1319,1321,0210 | Repair after replacing test card
No | Visit | Complaint | RCHN | 机器黑屏 | 机器黑屏，需要借件910-342 | 更换910-342后修复 by jesson | SWO-505825 | Defect | HW | DMS_PC | DMS_PC | Black screen | Machine black screen | Machine black screen, need to borrow 910-342 | Repair by jesson after replacing 910-342
No | Visit | Complaint | RCHN | 定标失败 5257 | 定标失败 5257, 借944-214*15428一张 | 重新定标后修复 by jesson | SWO-507473 | Assisted Replacement | Consumables | Test Kit | Cartridge, cal | 5257 | Bid failure 5257 | Bid failure 5257, borrowed 944-214 x 15428 | Fix by jesson after re-calibration
No | Visit | Complaint | RCHN | 3414FTC液路故障维修 | 3414FTC液路故障维修 | 拆机维修，疏通液路和HCT模块堵塞物，开机调试后正常 by jesson | SWO-508213 | Cleaned | HW | Wet section | Needle Wash Unit | 3414 | 3414FTC liquid road fault repair | 3414FTC liquid road fault repair | Demolition repair, dredging liquid road and HCT module blockage, normal after start-up commissioning by jesson
No | Visit | Complaint | RCHN | 定标失败，报错1029 | 定标失败，报错1029.需要借件903-364、944-212（15241）、944-214（15174）、944-216（15107）944-258（15185）缺货、944-220（15274） | 更换903-364重新定标后修复 by jesson | SWO-505837 | Defect | HW | Process Module | Optical Unit | 1029 | Bid failure, error 1029 | Failed bid failure, error 1029. Need to borrow 903-364, 944-212 (15241), 944-214 (15174), 944-216 (15107) 944-258 (15185) out of stock, 944-220 (15274) | Fix by jesson after replacement 903-364 recalibration
No | Visit | Complaint | RCHN | 定标失败 | 定标失败，借944-394（15624）一个 | 重新定标后修复 by jesson | SWO-507558 | Assisted Replacement | Consumables | Test Kit | Cartridge, cal | 5684 | Bid failed | Failed to bid, borrow 944-394 (15624) one | Fix by jesson after re-calibration
No | Visit | Complaint | RCHN | 机器一直重启 | 机器一直提示重启解决故障，回访老师的时候，反映机器一直频繁出现4601，3050，需要重启解决。 | 清洁进样口的皮带轮后修复 by jesson | SWO-507430 | Adjustment | HW | Cover and Chassis | Door, Inlet | 4061 | The machine has been restarting. | The machine has been prompted to restart to solve the problem, return to the teacher, reflecting that the machine has been frequently appearing 4601,3050, need to restart to solve. | Repair by jesson after cleaning the pulley at the inlet
No | Visit | Complaint | RCHN | 电机，传感器故障 | 电机，传感器故障，借636-533,580-098各一个 | 拆机后发现电机上的固定螺丝松动，拧紧后修复 by jesson | SWO-508126 | Adjustment | HW | Storage Module | Sensor, storage wheel | 5684 | Motor, sensor fault | Motor, sensor fault, one each by 636-533, 580-098 | After the removal of the machine found that the fixed screws on the motor loose, tightened after repair by jesson
No | Visit | Complaint | RCHN | 试剂包卡在机器里 | 由于用户安装试剂包，未将试剂包安装到位，造成试剂包卡在机器里 | 取出试剂包，重新安装后修复 by jesson | SWO-508300 | Assisted Replacement | Consumables | Solution Pack | Solution Pack | 5784 | The reagent bag is stuck in the machine. | Due to the user installing the reagent package, the reagent package is not installed in place, resulting in the reagent package stuck in the machine | Remove the reagent kit and reinstall the repair by jesson
No | Visit | Complaint | RLTD | Analyser memory chip not detected contact service displayed on analyser screen Dr Lewis Gough | Due to extensive wait in time for analyser ID board to arrive decided to make a site visit to check analyser over. Powered down and removed cover and located analyser ID Board, board quite loose in its mounting, removed and reinserted few times and no change on power on, analyser memory chip not detected. Removed complete board that ID board slots into and visually checked board and refitted and reconnected all cables, left | Re assembled analyser and will now chase up order for analyser ID board. | SWO-508939 | [5 empty] | Analyser memory chip not detected contact service displayed on analyser screen Dr Lewis Gough | Due to extensive wait in time for analyser ID board to arrive decided to make a site visit to check analyser over. Powered down and removed cover and located analyser ID Board, board quite loose in its mounting, removed and reinserted few times and no change on power on, analyser memory chip not detected. Removed complete board that ID board slots into and visually checked board and refitted and reconnected all cables, left | Re assembled analyser and will now chase up order for analyser ID board.
No | Visit | Complaint | RCHN | 空气检测失败 | 空气检测失败 | 重新冲洗测试卡后修复 by jesson | SWO-508121 | [5 empty] | Failed air detection | Failed air detection | Fix by jesson after re-re-flushing test card
No | Visit | Complaint | RCHN | lac电极故障 | lac电极故障.需要借件945-619一个 | 更换945-619后修复 by jesson | SWO-507378 | Assisted Replacement | Electrode | Electrode | Lac | 378 | lac electrode failure | lac electrode failure. | Repair by jesson after replacement 945-619
No | Visit | Complaint | RDEU | 01.10.2018 08:30Uhr Herr Glienke 03334691503: Das Gerät meldet häufig Driftalarm bei Elektrolyte | Wieder Sporadische Driften der Lyte bei Kal und auch gelegentlich QK-Fehler | Kontrolle und Austausch der E1001 (alt: 38-09), mehrere Kal abgewartet, Kal und QK ok | SWO-508335 | Assisted Replacement | Electrode | Electrode | Reference | CAPA-733 | 01.10.2018 08:30Am Mr Glienke 03334691503: The device frequently reports drift alarm for electrolytes | Again sporadic drifts of the Lyte at Kal and also occasional QK errors | Control and replacement of the E1001 (old: 38-09), several Kal waited, Kal and QK ok
No | Visit | Complaint | RFRA | ca trop fort sur level 3 | ca trop fort sur level 3 (baisse valeur ) | nettoyage electrode PH , changement menbrane CA , passage calibration ctrl ok0 ( valeur cq fournisseur 0,27 -0,47; valeur client 0,31 - 0 ,43) Modification des valeurs pour suivre l'évolution l'etat des membranes de ca. | SWO-508326 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | Ca too strong on level 3 | ca too strong on level 3 (lower value) | ELECTROde cleaning PH, change menbrane CA, passage calibration ctrl ok0 (supplier cq value 0.27 -0.47; customer value 0.31 - 0 .43) Changing values to track the evolution of the state of the ca membranes.
No | Visit | Complaint | RKK | 初期不良 Hemolyzer PhaseRegulationFailed ヘモライザー交換 | オキシメトリー測定エラー | ヘモライザー交換 | SWO-508321 | Defect | HW | Oxi module | Hemolyzer |  | Initial Defect Hemolyzer PhaseRegulation Failed Hemolizer Replacement | Oxymetry measurement error | Hemolyzer replacement
No | Visit | Complaint | RAME | Analyzer failing to print | Analyzer failing to print and showing error 216. | Restarted analyzer and removed and replaced printer paper. Issue cleared | SWO-508602 | Assisted Replacement | Printer | Paper | Paper |  | Analyzer failing to print | Analyzer failing to print and showing error 216. | Restarted analyzer and removed and replaced printer paper. Issue cleared
Yes | 3 | Y | RCZE | Blank |  |  | D5C45FDA4E2C827F | Re classified RMED | planned SWO |  |  |  | Blank
Yes | 3 | Y | RCZE | Blank |  |  | D5C460764E178281 | Re classified RMED | planned SWO |  |  |  | Blank
No | Visit | Complaint | RPOL | Po wymianie kasety nie kalibrują się mleczany | Testy serwisowe | Wymiana kasety sensorowej oraz uszczelki wlotu i igły. | SWO-508790 | Assisted Replacement | SC | SC | Lac |  | After replacing the cartridge, do not calibrate the lactate | Service Tests | Replacing the sensor cassette and the inlet and needle seals.
No | Visit | Complaint | RCHN | 孵育盘温度传感器线断了 | 孵育盘温度传感器线断了 | 更换排线后修复 by jesson | SWO-508749 | Defect | HW | Incubator Module | Cable, temp |  | The incubation disk temperature sensor line is broken. | The incubation disk temperature sensor line is broken. | Fix by jesson after replacing wiring
No | Visit | Complaint | RCHN | 进样舱门无法关闭 | 进样舱门无法关闭，借886-638和886-639各一个 | 更换886-638和886-639后修复 by jesson | SWO-507383 | Defect | HW | Cover and Chassis | Door, Inlet | 5784 | The sample door cannot be closed | The sample door cannot be closed, with one each of 886-638 and 886-639 | Fix by jesson after replacing 886-638 and 886-639
No | Visit | Complaint | RAME | Customer requests to go live on this system & analyzer is having Glu sensitivity issue (too low) | Inspection of the El $Met module revealed the Glucose lower sealing ring is missing. Some debris is Glu well. Connectivity testing failed, waiting for IT to resolve IP conflict. Got a new Static IP and had MAC address of analyzer added to network filter. | Replaced lower sealing ring & cleaned measuring wells. Prepped system for customer go-live. Resolved settings, new IP & had IT add MAC address to network & Verified connectivity. | SWO-508663 | Defect | HW | Electrode | pH sealing ring |  | Customer requests to go live on this system & analyzer is having Glu sensitivity issue (too low) | Inspection of the El /Met module revealed the Glucose lower sealing ring is missing. Some debris is Glu well. Connectivity testing failed, waiting for IT to resolve IP conflict. Got a new Static IP and had MAC address of analyzer added to network filter. | Replaced lower sealing ring & cleaned measuring wells. Prepped system for customer go-live. Resolved settings, new IP & had IT add MAC address to network & Verified connectivity.
No | Visit | Complaint | RIND | AQT90 sample inlet door movement error. | Check the system. check the logs. Check the error logs. | Inlet door assembly issue. Movement error cause the system hang. system fail to recovery. cross check with another inlet door assembly. need to replace the door assembly. | SWO-512429 | Defect | HW | Cover and Chassis | Door, Inlet |  | AQT90 sample inlet door movement error. | Check the system. check the logs. Check the error logs. | Inlet door assembly issue. Movement error cause the system hang. system fail to recovery. cross check with another inlet door assembly. need to replace the door assembly.
No | Visit | Complaint | RIND | Inlet door not closed properly. | Inlet door not closed properly.Automatic recovery started | Clean the inlet door sensor.Fix the washer in gear wheel with mircopore as I have identified that the gear wheel movement in the different position is responsibile for inlet door open $close. | SWO-508709 | Adjustment | HW | Cover and Chassis | Door, Inlet |  | Inlet door not closed properly. | Inlet door not closed properly. Automatic recovery started | Clean the inlet door sensor. Fix the washer in gear wheel with mircopore as I have identified that the gear wheel movement in the different position is responsibile for inlet door open $close.
No | Visit | Complaint | RKOR | ABL/DMS PC connected to wet section | . | Demo ABL90 standby | SWO-512479 | Defect | HW | DMS_PC | CPU |  | ABL/DMS PC connected to wet section | . | Demo ABL90 standby
No | Visit | Complaint | RCHN | 光学灵敏度超范围，报错1029 | 光学灵敏度超范围，报错1029 借903-364 944-212（15177）944-214（15057 15229缺货）944-216（15107）各一个 | 更换903-364重新定标后修复 by jesson | SWO-508160 | Defect | HW | Process Module | Optical Unit | 1029 | Optical sensitivity over range, error 1029 | Optical sensitivity over range, error 1029 borrowing 903-364 944-212 (15177) 944-214 (15057 15229 out of stock) 944-216 (15107) each | Fix by jesson after replacement 903-364 recalibration
No | Visit | Complaint | RCHN | 加热温度超范围 | 加热温度超范围 | 更换库存903-210后修复by tim | SWO-508851 | Defect | HW | Process Module | PCB, cup dryer |  | Overheating temperature over range | Overheating temperature over range | Repair by tim after replacing stock 903-210
No | Visit | Complaint | RHUN | Tű elgörbült, LIVE Connect | Tű elgörbült. | Felszívó tű kiegyenesítése, LIVE Connect beállítása. tHb kal. Kalibrálás, QC mérés. | SWO-509087 | Adjustment | HW | Inlet module | Probe bend |  | Needle Bent, LIVE Connect | Needle Bent. | Suction needle straightening, LIVE Connect setting. by tHb. Calibration, QC measurement.
No | Visit | Complaint | RLTD | Leak current detected |  |  | SWO-509721 | Adjustment | Analyzer restart | Analyzer restart | Leak current | CRM: Corrective Action: Analyser showing Leak current on intervention screen, but has low leak current reading in service. Restarted analyser and waited for 1 hour start up. now okay. | Leak current detected
No | Visit | Complaint | RAME | Inlet module keeps coming off | Waiting for new inlet motion mechanism. | Replaced inlet motion mechanism, inlet coupling and inlet connector gasket. Allowed analyzer to condition. | SWO-509250 | Defect | HW | Auto Inlet module | PCB, Auto inlet |  | Inlet module keeps coming off | Waiting for new inlet motion mechanism. | Replaced inlet motion mechanism, inlet coupling and inlet connector gasket. Allowed analyzer to condition.
No | Visit | Complaint | RLTD | Irene Mallett 01708435000 ext 6845 or 6194 | The customer has replaced the SC to try and rectify the issue. Checked data logs nad service programs. Clot detect $removal is showing various errors. | Performed intensive flush in servicwe mode. Ran fluid transport checks, solution pack check and clot detect $removal. All are now passing. | SWO-509665 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway | Clot | Irene Mallett 01708435000 ext 6845 or 6194 | The customer has replaced the SC to try and rectify the issue. Checked data logs nad service programs. Clot detect /removal is showing various errors. | Performed intensive flush in servicwe mode. Ran fluid transport checks, solution pack check and clot detect $removal. All are now passing.
No | Visit | Complaint | RLTD | The analyser is giving a Clot Suspected in Inlet error every night. | The analyser has been failing the midnight clot detection with a clot suspected in the inlet. The customer had been clearing the error by rebooting the analyser but when I arrived today I was able to clear the error by running the clot detect $ remove service program. | As I could not replicate the fault by re-running the clot detect $ remove service program I swapped the inlet module with one from another analyser to see if the fault will follow the inlet module. | SWO-509029 | Revisit required | HW | Basic Unit | Sensor Bay Unit | Clot detection, midnight | The analyser is giving a Clot Suspected in Inlet error every night. | The analyser has been failing the midnight clot detection with a clot suspected in the inlet. The customer had been clearing the error by rebooting the analyser but when I arrived today I was able to clear the error by running the clot detect / remove service program. | As I could not replicate the fault by re-running the clot detect $ remove service program I swapped the inlet module with one from another analyser to see if the fault will follow the inlet module.
No | Visit | Complaint | RAME | Fluid blockage | 1218 errors indicating air contamination. | Replaced manifold assembly. Got CO-OX board failure. Cleared error but it recurred another time. 1218 errors also recurring. Performed service test, tHb cal and system cycle. | SWO-509383 | [5 empty] | Fluid blockage | 1218 errors indicating air contamination. | Replaced manifold assembly. Got CO-OX board failure. Cleared error but it recurred another time. 1218 errors also recurring. Performed service test, tHb cal and system cycle.
No | Visit | Complaint | RCHN | 多参数定标不过 | 多参数定标不过，借891-051*7 | 更换891-051*7后修复by tim | SWO-508927 | Defect | HW | Basic Unit | SC contact | 378 | Multi-parameter calibration but | Multi-parameter calibration however, borrow 891-051 x 7 | Repair by tim after replacing 891-051 x 7
No | Visit | Complaint | RCHN | 定标卡扫描器的线头坏了 | 定标卡扫描器的线头坏了，借636-571一个 |  | SWO-508874 | Defect | HW | Analyzer Control Module | Cable, display |  | The line head of the calibration card scanner is broken. | The line head of the calibration card scanner is broken, borrow 636-571 one
No | Visit | Complaint | RCHN | 样本架损坏 | 进样支架断裂，借903-584一个 |  | SWO-508822 | Defect | HW | Aspiration Module | Tube holder |  | Damage to the sample rack | The sample bracket is broken, and 903-584 is a
No | Visit | Complaint | RCHN | 定标底物外露 | 定标底物外露，借944-212（15388）一个 | 重新定标后修复 | SWO-508681 | Cleaned | HW | Process Module | Optical Unit | 1024 | Calibration substrate exposure | Bid substrate exposed, borrow 944-212 (15388) one | Repair after re-calibration
No | Visit | Complaint | RCHN | 测试卡损坏导致卡盘 5784 （teeth) | 测试卡损坏导致卡盘 5784 （teeth) | 取出断角的测试卡by tim | SWO-509459 | Defect | Consumables | Test Kit | Cartridge, Broken | Teeth, 5784 | Damage to the test card caused the chuck 5784 (teeth) | Damage to the test card caused the chuck 5784 (teeth) | Take out the cut-out test card by tim
No | Visit | Complaint | RPOL | Brak parametrów raportowych, brak możliwości pomiaru. | Prawdopodobnie została zmierzona próbka QC na niewłaściwym poziomie. | Pomiar próbki kontrolnej lvl 1, po pomiarze pO2 poza granicami normy, wymiana pakietu roztworów, ponowny pomiar próbek kontrolnych lvl1 i lvl2. | SWO-509274 | [5 empty] | No reporting parameters, no measurement possible. | A QC sample at the wrong level was probably measured. | Measurement of Control Sample LVL 1, after measurement of pO2 beyond the limits of the standard, replacement of the solution package, remeasurement of control Samples LVL1 and LVL2.
No | Visit | Complaint | RCHN | 无法正常重装测试卡，休眠时间显示不正确 | 无法正常重装测试卡，休眠时间显示不正确 | 指导检查仪器设置，重启机器，重新安装测试卡后修复by tim | SWO-511735 | [5 empty] | Failed to reload the test card properly, sleep time is not displayed correctly | Failed to reload the test card properly, sleep time is not displayed correctly | Guide to check instrument settings, restart the machine, reinstall the test card to repair by tim
No | Visit | Complaint | RDEU | Das Gerät bootet nicht (AU läuft?) | Gerät bootet nicht | Computereinheit demontiert und geprüft, Ersatzteile bestellt, Ersatzgerät (Abl 90 R0840N003) installiert | SWO-509729 | Revisit required | HW | DMS_PC | DMS_PC | mail: PC module back for investigation | The device is not booting (AU is running?) | Device does not boot | Computer unit dismantled and tested, spare parts ordered, replacement unit (Abl 90 R0840N003) installed
Yes | 3 | Y | RCZE | Blank |  |  | D5C55176C4431D59 | Re classified RMED | planned SWO |  |  |  | Blank
No | 5 | Y | RCZE | Blank |  |  | D5C57AC9C4040BED | Cleaned | Fluidic pathway | Inlet module | Probe |  | Blank
No | Visit | Complaint | RCHN | 样本检测失败 | 样本检测失败，借903-371一个 |  | SWO-508897 | Defect | HW | Incubator Module | Punch up | 5298 | Sample detection failed | Sample test failed, borrow edi-903-371
No | Visit | Complaint | RIND | Service Call. | Sample measurement error, shows optical sensitivity out of range. | Run the distilled water as a sample, again and again sample measurement comes with an error optical sensitivity out of range. Need to replace optical unit. | SWO-509560 | Revisit required | HW | Process Module | Optical Unit | 1029 | Service Call. | Sample measurement error, shows optical sensitivity out of range. | Run the distilled water as a sample, again and again sample measurement comes with an error optical sensitivity out of range. Need to replace optical unit.
No | Visit | Complaint | RPOL | Błąd kalibracji LS OXi | Błędna regulacja LS | Czyszczenie odpowietrzania. Regulacja LS prawidłowa. | SWO-509498 | Adjustment | HW | Oxi module | LS |  | LS OXi Calibration Error | Incorrect adjustment of LS | Cleaning the venting. LS adjustment is correct.
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 气泡探测定标失败 | 更换测试卡后修复 | SWO-509620 | [5 empty] | Test card, reagent package failure | Bubble probe calibration failed | Repair after replacing test card
No | Visit | Complaint | RCHN | 接口板腐蚀 | 接口板腐蚀，借902-972一个 | 更换902-972及更换测试卡后修复by tim | SWO-508966 | [5 empty] | Interface plate corrosion | Interface plate corrosion, borrow 902-972 one | Repair by tim after replacing 902-972 and replacing test card
No | Visit | Complaint | RCHN | 无法抽吸溶液 | 无法抽吸溶液（未更换备件） | 冲洗液路透明管后修复by tim | SWO-509676 | [5 empty] | Unable to pump solution | Unable to pump solution (spare parts not replaced) | Repair by tim after flushing liquid road transparent tube
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 测试卡有凝血 | 更换测试卡后修复 | SWO-510248 | [5 empty] | Test card, reagent package failure | The test card has clotting. | Repair after replacing test card
No | Visit | Complaint | RCHN | 机器堵了 | 机器堵了（未更换备件） | 清洁，冲洗管路后修复by tim | SWO-509494 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | The machine's blocked. | Machine blocked (no replacement parts) | Clean, rinse the line and repair by tim
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 测试卡参比电极错误 1361 | 连续更换2张测试卡后修复 | SWO-511167 | [5 empty] | Test card, reagent package failure | Test card reference electrode error 1361 | Repair after 2 consecutive test cards
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 泄漏1184；泵定标出错1178 | 连续更换测试卡，试剂包后修复 | SWO-509600 | [5 empty] | Test card, reagent package failure | Leak 1184; pump calibration error 1178 | Continuous replacement of test cards, remediation after reagent package
Yes | Visit | Complaint | RCHN | 液路模块接头断了 | 液路模块接头断了，借924-591一个 | 更换924-591后修复 by Andrew | SWO-508751 | [5 empty] | The fluid module connector is broken. | The liquid-road module connector is broken, borrowed 924-591 one | Repair by Andrew after replacement 924-591
No | Visit | Complaint | RCHN | 仪器软件崩溃，初始化失败，指定模块不能被找到 | 仪器软件崩溃，初始化失败，指定模块不能被找到 | 重装软件后修复by tim | SWO-509624 | Defect | SW | RMED SW | Windows | Reinstall | Instrument software crashes, initialization fails, specified module cannot be found | Instrument software crashes, initialization fails, specified module cannot be found | Fix by tim after reinstalling software
No | Visit | Complaint | RFRA | Instable sur lactate | Valeur instable depuis 15 jours | Changement chambre de mesure ions, changement capteur de réponse , calibration du zéro , cal ok ctrl ok passage de 5 gaz ok à suivre . | SWO-509818 | Defect | HW | El/Met Module | measuring chamber |  | Unstable on lactate | Unstable value for 15 days | Changing chamber measurement ions, changing response sensor, calibration of zero, cal ok ctrl ok passage of 5 gas ok to follow.
No | Visit | Complaint | RFRA | Erreur calibration Ca++ | Module oxy $ | Remplacement de l'électrode de réf $Test oxy test $ Constate erreur lamp $Démontage et remplacement du module lampe $Ajustement de la lampe $Test oxy ok $Calibration de la tHB en mode service $Calibration de la tHB en mode utilisateur $Calibration deux points et passage des CQ. | SWO-509739 | Assisted Replacement | Electrode | Electrode | Reference | CAPA-733 | Ca's calibration error | Module oxy / | Ref electrode replacement $Test oxy test $ Finds error lamp $Deupting and replacement of the lamp module $Ajustement lamp $Test oxy ok $Calibration of the tHB in service mode $Calibration of the tHB in user mode $Calibration two points and passage of the CQ.
No | Visit | Complaint | RAME | The MetHb measures high on blood samples compared to two other analyzers. | The MetHb results are high on blood samples compared to two other analyzers. THb Cal has been performed but it did not affect the results. | Replaced the Oximetry Unit 902-833R2398N003. Performed the tHb calibration and complete oximetry alignment. Hemolyzer repeatedly fails the Impedance Shape test. Reinstalled original Oximetry Unit and performed the tHb calibration and complete oximetry alignment. Ran QC. Analyzer checks okay. Will return to replace the Oximeter Unit when part is available. | SWO-509392 | Revisit required | HW | Oxi module | Oxi module |  | The MetHb measures high on blood samples compared to two other analyzers. | The MetHb results are high on blood samples compared to two other analyzers. THb Cal has been performed but it did not affect the results. | Replaced the Oximetry Unit 902-833R2398N003. Performed the tHb calibration and complete oximetry alignment. Hemolyzer repeatedly fails the Impedance Shape test. Reinstalled original Oximetry Unit and performed the tHb calibration and complete oximetry alignment. Ran QC. Analyzer checks okay. Will return to replace the Oximeter Unit when part is available.
No | Visit | Complaint | RKK | neon number too high が断続的に発生 | spectrum meter が原因の為 Oxiモジュールごと交換 | oxiモジュール交換 | SWO-510190 | Defect | HW | Oxi module | Oxi module | Neon number too high | neon number too high occurs intermittently | Because of spectrum meter Replacement per Oxi module | Oxi module replacement
No | Visit | Complaint | RESP | Equipo bloqueado tras carga de reactivo No les deja hacer nada en el equipo | Chequeo de datos, comprobación de storage wheel | Se ha quedado un quesito atascado en la storage wheel causado por el papel secante que trae en la bolsita | SWO-517289 | Defect | Consumables | Test Kit | Cartridge, Broken | 5784; Teeth | Equipment locked after reagent loading He won't let them do anything on the team. | Data check, storage wheel check | It has been stuck in the storage wheel caused by the drying paper it brings in the bag
No | Visit | Complaint | RPOL | Błąd intensywności modułu optycznego | Sprawdzenie ostatnich błędów. | Czyszczenie koła procesowego i oraz punktów referencyjnych. Brak poprawy. Konieczna wymiana modułu optycznego. Prace serwisowe będą kontynuowane po otrzymaniu części zamiennych. | SWO-510101 | Revisit required | HW | Process Module | Optical Unit |  | Optical module Intensity error | Check for recent errors. | Cleaning the process wheel and reference points. No improvement. Replacement of the optical module is necessary. Maintenance will continue after the replacement parts are received.
No | Visit | Complaint | RLTD | leak detected | on arrival the analyser was showing a blue screen so a reboot was carried out. Analyser started up ok but 5 minutes into start up it asked for inlet gasket replacement. The service menu was entered and fluid transport tests carried out and pressure test failed with a vaccum value of 0.97. On inspection evidence of a leak was seen under pinch valve. | Tubing for pinch valve replaced. Fluid transport tests now passing. | SWO-510724 | Defect | HW | Oxi module | Pinch valve tube |  | leak detected | on arrival the analyser was showing a blue screen so a reboot was carried out. Analyser started up ok but 5 minutes into start up it asked for inlet gasket replacement. The service menu was entered and fluid transport tests carried out and pressure test failed with a vaccum value of 0.97. On inspection evidence of a leak was seen under pinch valve. | Tubing for pinch valve replaced. Fluid transport tests now passing.
No | Visit | Complaint | RKK | リーク検出 | リーク検出、圧力テスト吸引エラー。装置メッセージ通り溶液パック・インレットセット交換するも改善しない。2 $9引上げ。社内修理 | 原因はセンサーカセット内の詰まり。センサーカセット交換で改善。 メンテナンスキット交換、ランニングテストOK。 | SWO-510994 | Assisted Replacement | SC | SC | FT | Clot | Leak detection | Leak detection, pressure test suction error. The device message as solution pack inlet set exchange also does not improve. 2 /9 raise. In-house repairs | The cause is a jam in the sensor cassette. Improved by sensor cassette replacement. Maintenance kit replacement, running test OK.
No | Visit | Complaint | RDEU | 26.02.2019 Herr Simon 03491502099 möchte die fixen Zielwerte installiert haben. | Einrichtung fixe Zielwerte | Anpassung roma.ini Geräteneustart | SWO-510629 | Setup | SW | RMED SW | roma.ini | fixed assigned values | 26.02.2019 Mr. Simon 03491502099 wants to have the fixed target values installed. | Setting up fixed target values | Adaptation roma.ini Device restart
No | Visit | Complaint | RLTD | Besa Tel: 07534 067522 | Gasket, solution pack cassette and probe already replaced, replaced pinch valve this did not resolve problem. | Noticed gasket had a large clot buildup within it, replaced and performed pressure test - all passed. | SWO-510691 | Defect | HW | Auto Inlet module | Inlet gasket | Clot | Besa Tel: 07534 067522 | Gasket, solution pack cassette and probe already replaced, replaced pinch valve this did not resolve problem. | Noticed gasket had a large clot buildup within it, replaced and performed pressure test - all passed.
No | Visit | Complaint | RKOR | Na, K cal error, Oxi module clot, SP error | . | SC, SP replace, Liquid flow line clean. | SWO-512485 | Assisted Replacement | SC | SC | Na/K |  | Na, K cal error, Oxi module clot, SP error | . | SC, SP replace, Liquid flow line clean.
No | Visit | Complaint | RLTD | David Ryder 07580 818900 | Site visit to investigate and repair po2 issue with dark data out of range. | Attended analyser and performed test on pO2 sensor chanel in service. Found po2 failed for zpo2 dark running high at 227.08 above the 200 threshold. Cleaned the lense in sensor bay and retested. Still failing. Powered down and replaced interface pcb and realigned contact pins. Refitted and noticed screen bezel not seated correctly. Found that two of the turrets holding rear and front covers together had snapped and caused a large 3 to 4 mm gap between screen and bezel on right hand side. No sign of physical damage to external area or signs of dropping. Replaced with spare and refitted. Powered analyser back up and retested the po2 in service. Now passing with value of 85.70. Exited and run calibrations and qcs. Analyser passing all cals and qcs and running well again. | SWO-510616 | Defect | HW | Basic Unit | PCB, Sensor interface | pO2 | David Ryder 07580 818900 | Site visit to investigate and repair po2 issue with dark data out of range. | Attended analyser and performed test on pO2 sensor chanel in service. Found po2 failed for zpo2 dark running high at 227.08 above the 200 threshold. Cleaned the lense in sensor bay and retested. Still failing. Powered down and replaced interface pcb and realigned contact pins. Refitted and noticed screen bezel not seated correctly. Found that two of the turrets holding rear and front covers together had snapped and caused a large 3 to 4 mm gap between screen and bezel on right hand side. No sign of physical damage to external area or signs of dropping. Replaced with spare and refitted. Powered analyser back up and retested the po2 in service. Now passing with value of 85.70. Exited and run calibrations and qcs. Analyser passing all cals and qcs and running well again.
Yes | Visit | Complaint | RIND | SP was consumed within 10 days machine having workload of 400+ samples per month | checked calibration and patient log | checked calibration and patient log found more calibration has been taken in C $A type, the SC found faulty as after taking corrective actions the parameters are not activated so suggested to replace sp and sc. | SWO-516429 | [5 empty] | SP was consumed within 10 days machine having workload of 400+ samples per month | checked calibration and patient log | checked calibration and patient log found more calibration has been taken in C $A type, the SC found faulty as after taking corrective actions the parameters are not activated so suggested to replace sp and sc.
No | Visit | Complaint | RCHN | 提示试剂门栓打开 | 提示试剂门栓打开，借903-334一个 | 更换903-334一个正常。by tim | SWO-510267 | [5 empty] | Prompt reagent door bolt open | Prompt reagent door bolt open, borrow 903-334 one | Replace 903-334 one normal. by tim
No | Visit | Complaint | RCHN | PO2定标失败 | PO2定标失败,903-581一个，未更换备件修复 | 打开仓门，打开金属探针旁边的探头by tim | SWO-508984 | Cleaned | HW | Inlet module | Probe |  | PO2 calibration failure | PO2 calibration failure, 903-581 one, unreplaced spare parts repair | Open the door, open the probe next to the metal probe by tim
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | K定标失败1070 | 更换测试卡后修复 | SWO-511254 | [5 empty] | Test card, reagent package failure | K calibration failure 1070 | Repair after replacing test card
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 无法读取芯片 | 更换试剂包后修复 | SWO-510266 | [5 empty] | Test card, reagent package failure | The chip could not be read | Repair after replacement of reagent kit
No | Visit | Complaint | RCHN | 机器做样本失败率高 | 机器做样本失败率高 | 调整后修复by tim | SWO-510409 | Defect | HW | Process Module | PCB, cup dryer | 1068 | High failure rate of machine sample | High failure rate of machine sample | Adjusted fix by tim
No | Visit | Complaint | RCHN | 定标失败 | 定标失败 借944-216（15107），944-214（15057）各一张 | 重新定标正常。by tim | SWO-508959 | Assisted Replacement | Consumables | Test Kit | Cartridge, cal |  | Bid failed | Bid failure Borrowing 944-216 (15107), 944-214 (15057) each | Recalibration is normal. by tim
No | Visit | Complaint | RCHN | 液体传输失败 | 液体传输失败（未更换备件） | 重复执行液路传感校准by tim | SWO-511233 | Adjustment | HW | pH/BG Module | LS | 537 | Failed liquid transfer | Liquid transfer failed (spare parts not replaced) | Repeat the liquid path sensing calibration by tim
No | Visit | Complaint | RDEU | 05.03.2019, Frau Sellmann, Tel: 0401818828816: Druckbild ist schlecht lesbar | Drucker Bild ist schlecht. | Printerhead erneuert. | SWO-510368 | Defect | HW | Printer Module | Printer Head |  | 05.03.2019, Mrs. Sellmann, Tel: 0401818828816: Print image is difficult to read | Printer image is bad. | Printerhead.
No | Visit | Complaint | RAME | Flex-Q Barcode scanner issues | Unit will scan safe pico but not customer's barcodes. A second analyzer scans OK. | Replaced and aligned FlexQ module. Loaded AU engine. Had to reload AU after unit stalled. | SWO-510819 | Defect | HW | FlexQ Module | FlexQ Module |  | Flex-Q Barcode scanner issues | Unit will scan safe pico but not customer's barcodes. A second analyzer scans OK. | Replaced and aligned FlexQ module. Loaded AU engine. Had to reload AU after unit stalled.
No | Visit | Complaint | RIND | Service Call. | Start The system check, failed, solid standard deviation is more negative below -20, need to replace the optical unit. | Replace the stand by optical unit, calibrate factory reference. But previous calibration has invalid due to replacement of optical unit. Need cal cartridge of TnI and CKMB. | SWO-510812 | Revisit required | HW | Process Module | Optical Unit | 1029 | Service Call. | Start The system check, failed, solid standard deviation is more negative below -20, need to replace the optical unit. | Replace the stand by optical unit, calibrate factory reference. But previous calibration has invalid due to replacement of optical unit. Need cal cartridge of TnI and CKMB.
No | Visit | Complaint | RDEU | 07.03.2019 10:50Uhr Herr Widera tel:+492191138584: Der bereits gemeldete Fehler pCO2 im LevelA zu hoch | Pco2 Fehler Level A zu hoch ( 29,3) Bereich : 25,1-28,9 | Sp gewechselt (Lot H05 zu HD06) mehrere Kal durchgeführt, Qc Level A 7 mal durchgeführt pco2 jetzt immer 28,8 Mmhg | SWO-511629 | Assisted Replacement | SP | SP | pCO2 | QC A high | 03/07/2019 10:50Am Mr Widera tel:+492191138584: The already reported error pCO2 in LevelA too high | Pco2 Error Level A Too High ( 29.3) Range: 25.1-28.9 | Sp changed (Lot H05 to HD06) multiple scale sat, Qc Level A 7 times performed pco2 now always 28.8 mmhg
No | Visit | Complaint | RDEU | keine Verbindung zu COBAS IT, Ping zum Gerät nicht möglich | Verbindungsdaten erneut überprüft und mit anderen Geräten im Hause abgeglichen, Daten korrekt und Verbindung laut ABL auch aktiv, Ping zum Gerät nicht möglich, Firewall überprüft, | Einstellungen korrigiert und System aktualisiert, Neustart abgewartet, QC, erneuter Pingtest durch lokale IT i.O. | SWO-511343 | Setup | SW | RMED SW | His/Lis |  | no connection to COBAS IT, ping to the device not possible | Connection data re-checked and matched with other devices in the home, data correct and connection according to ABL also active, ping to the device not possible, firewall checked, | Settings corrected and system updated, restart awaited, QC, re-ping test by local IT I.O.
No | Visit | Complaint | RDEU | 07.03.2019; Fr. Krug; Tel.: 06213836285; Leckage erkannt 07.03.2019 11:10Uhr Fr.Krug 06213836285: Weiterhin Leckagefehler | Leckage erkannt Vakuum Fehler | Allgemeine Überprüfung - Einlassdichtung (bereits erneuert) und Einlassnadel demontiert und überprüft - Analysator manuell gespült - Service Tests durchgeführt (Vakuum Value erhöht) - Erneuerung SP - Startphase - Funktionsüberprüfung etc. | SWO-511424 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway | 1342 | 07.03.2019; Ms. Krug; Tel.: 06213836285; Leakage detected 03/07/2019 11:10Am Fri Krug 06213836285: Leakage errors continue | Leakage detected Vacuum error | General inspection - Inlet seal (already renewed) and inlet pin disassembled and checked - Analyzer manually rinsed - Service tests carried out (Vacuum Value increased) - Renewal SP - Start phase - Function check etc.
No | Visit | Complaint | RAME | Clot detected. | Clot detected. Performed flush, installed new SC, SP, inlet probe with no help. | Retubed instrument and replaced pinch valve. | SWO-511538 | Defect | HW | Basic Unit | PM Kit | Clot | Clot detected. | Clot detected. Performed flush, installed new SC, SP, inlet probe with no help. | Retubed instrument and replaced pinch valve.
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | pco2,glu定标失败 | 更换测试卡后修复 | SWO-514315 | [5 empty] | Test card, reagent package failure | pco2, glu bid failure | Repair after replacing test card
No | Visit | Complaint | RIND | SP was consumed within 10 days machine having workload of 400+ samples per month | replace SC and SP | Install new SC and SP,calibration done successfully machine under observation. | SWO-516431 | [5 empty] | SP was consumed within 10 days machine having workload of 400+ samples per month | replace SC and SP | Install new SC and SP,calibration done successfully machine under observation.
No | Visit | Complaint | RIND | When customer try to on instrument with power switch, instrument is not getting switched on. | Check power supply and try to on analyser but 1st and 2nd LES is red so instrument internal power supply problem | Check CPU board and power supply and remove signal data board and found that signal data board is burn. without signal data board analyser is switched on so we have to replace SINGAL DATA BOARD. | SWO-512418 | [5 empty] | When customer try to on instrument with power switch, instrument is not getting switched on. | Check power supply and try to on analyser but 1st and 2nd LES is red so instrument internal power supply problem | Check CPU board and power supply and remove signal data board and found that signal data board is burn. without signal data board analyser is switched on so we have to replace SINGAL DATA BOARD.
Yes | Visit | Complaint | RIND | ABL80 frequent calibration. | Check the calibration logs. | Take a logs. SP80 consumed in 10 days by frequent calibration. calibration interval 8 hours still calibration every 2 hours. Need to investigation. | SWO-512166 | [5 empty] | ABL80 frequent calibration. | Check the calibration logs. | Take a logs. SP80 consumed in 10 days by frequent calibration. calibration interval 8 hours still calibration every 2 hours. Need to investigation.
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 无法阅读测试卡智能芯片 1123 | 更换试剂包后修复 | SWO-512795 | Assisted Replacement | SC | SC | chip |  | Test card, reagent package failure | Can't read test card smart chip 1123 | Repair after replacement of reagent kit
No | Visit | Complaint | RCHN | TNI定标数据失效 | TNI定标数据失效 借944-212（15457）一张 | 重新定标正常。by tim | SWO-510310 | Assisted Replacement | Consumables | Test Kit | Cartridge, cal |  | TNI calibration data failure | TNI calibration data failure Borrowed 944-212 (15457) one | Recalibration is normal. by tim
No | Visit | Complaint | RCHN | Push up杯子打不上 | Push up杯子打不上 | 自行更换push up unit后修复by tim | SWO-511124 | Defect | HW | Incubator Module | Punch up |  | Push up cup can't hit | Push up cup can't hit | Fix by tim after replacing push up unit on your own
No | Visit | Complaint | RIND | Analyser status is showing intervention required | Check system massage first. and then all solution status. Cal 2 solution was empty. | installed Cal 2 solution. then give 2 point calibration 2 time. after that analyser was start working. | SWO-512384 | Assisted Replacement | Consumables | Solution | Cal 2 |  | Analyser status is showing intervention required | Check system massage first. and then all solution status. Cal 2 solution was empty. | installed Cal 2 solution. then give 2 point calibration 2 time. after that analyser was start working.
No | Visit | Complaint | RSCH | Blank | Erreur oxymétrie lors de la calibration | Ajustage lampe halogène impossible (Halogen Stability test failed) Remplacé l’unité de la lampe halogène e.o Remplacé les orings des électrodes factices e.o Test complet de l’oxymétrie e.o Calibration e.o contrôle de qualité e.o | SWO-511408 | Defect | HW | Oxi module | Lamp unit | Halogen | Blank | Oxymetry error during calibration | Halogen Stability test failed Replaced the unit of the halogen lamp e.o Replaced the orings of dummy electrodes e.o Complete oxymetry test e.o Calibration e.o quality control e.o
No | Visit | Complaint | RSAF | Ph out of range on sensitivity | Inspected for any leaks on the chamber | Replaced o-rings | SWO-511168 | Defect | Electrode | Electrode | Sealing ring, pH |  | Ph out of range on sensitivity | Inspected for any leaks on the chamber | Replaced o-rings
No | Visit | Complaint | RAME | Leaking tube. | Tubing between PH $BG and EL $MET modules leaking . | Replaced tubing. | SWO-511563 | Defect | HW | El/Met Module | Tubing |  | Leaking tube. | Tubing between PH /BG and EL /MET modules leaking . | Replaced tubing.
Yes | Visit | Complaint | RIND | Oxi module defective | Performed complete oximetery test and found lamp unit failed. | Replaced complete oximetery module. Performed 2point calibration without error. | SWO-516005 | Re classified RMED | planned SWO |  |  | Move to OOB trending | Oxy Module Defective | Performed complete oximetery test and found lamp unit failed. | Replaced complete oximetery module. Performed 2point calibration without error.
No | Visit | Complaint | RDEU | 07.03.2019 9:30 Sr. Nicole 059311523852: Gerät ohne Funktion | Fehler 5784: Setsafe_22, 2877: Cartridge_40 | Gerät überprüft, Kassetten und Punch out kontrolliert, Kassettenführung arretiert, Messung o.k, | SWO-513544 | Defect | Consumables | Test Kit | Cartridge, Broken | 5784, Teeth | 03/07/2019 9:30 Sr. Nicole 059311523852: Device without function | Error 5784: Setsafe_22, 2877: Cartridge_40 | device checked, cassettes and punch out checked, cassette guide locked, measurement ok,
No | Visit | Complaint | RIND | Service Call. | Need to clean the process wheel and check the blower of cup dryer. | Shut down the analyzer, open the front panel and back panel, remove the optical unit, clean the process wheel, connect all parts same as before, start the analyzer and go to service mode, test the cup dryer flow test, successfully done, calibrate the TnI, calibration successfully, analyzer handed over to staff in well working condition. | SWO-512394 | Defect | HW | Process Module | Optical Unit | 1029 | Service Call. | Need to clean the process wheel and check the blower of cup dryer. | Shut down the analyzer, open the front panel and back panel, remove the optical unit, clean the process wheel, connect all parts same as before, start the analyzer and go to service mode, test the cup dryer flow test, successfully done, calibrate the TnI, calibration successfully, analyzer handed over to staff in well working condition.
No | Visit | Complaint | RKK | GTモードで装置が停止。（2日間） | 機器状況確認 インレットガスケット、汚れの為 | インレットガスケット、インレットプローブ清掃 排液 $インレットライン清掃。 溶液パック再取り付け、センサーカセット交換。 CALチェックOK。QCチェック(A,B,C)OK | SWO-519566 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | The device stopped in GT mode. (2 days) | Equipment status check Inlet gaskets for dirt | Inlet gasket, inlet probe cleaning Drainage $Inlet line cleaning. Solution pack re-installed, sensor cassette replacement. CAL check OK. QC check (A, B, C) OK
No | Visit | Complaint | RAME | Battery has zero charge | Checked out analyzer. Battery not holding any charge. | Replaced battery. Adjusted settings in the control panel to remove seconds from time to try to aleviate problems with Aqure. Performed all necessary calibrations. | SWO-512306 | Defect | HW | Battery module | Battery module |  | Battery has zero charge | Checked out analyzer. Battery not holding any charge. | Replaced battery. Adjusted settings in the control panel to remove seconds from time to try to aleviate problems with Aqure. Performed all necessary calibrations.
Yes | Visit | Complaint | RKOR | Server PC error | . | PC replace. AQURE setup. | SWO-512494 | Re classified RMED | External Product |  |  | see mail with title: "RE: Installation of server PCs FW: SWO-512494" from JIMPA on 2019-03-14 12:35 | Server PC error | . | PC replace. AQURE setup.
Yes | Visit | Complaint | RKOR | AQURE server pc change | AQURE server pc change | AQURE server pc change | SWO-512606 | Re classified RMED | same day upSWO |  |  |  | AQURE server pc change | AQURE server pc change | AQURE server pc change
No | Visit | Complaint | RAME | Unresponsive screen - Intermittent issues | Responding to call for touch screen calibration. | Disabled write protection, conducted touchscreen calibration and restarted. | SWO-512226 | Adjustment | HW | Display | Touch Screen | calibration | Unresponsive screen - Intermittent issues | Responding to call for touch screen calibration. | Disabled write protection, conducted touchscreen calibration and restarted.
No | Visit | Complaint | RAME | Status Code 5-Data upgrade failed | 1) Status Code 5-Data upgrade failed 2) Go Live | 1) Reloaded OS and Roma software. 2) Go Live | SWO-512319 | Defect | SW | RMED SW | RMED SW | Status code 2 | Status Code 5-Data upgrade failed | 1) Status Code 5-Data upgrade failed 2) Go Live | 1) Reloaded OS and Roma software. 2) Go Live
No | Visit | Complaint | RKOR | Quick access error | . | OS, Software reinstall. | SWO-512491 | Defect | SW | RMED SW | RMED SW | Quick access | Quick access error | . | OS, Software reinstall.
No | Visit | Complaint | RKOR | pCO2 Calibration error | . | SC replace. | SWO-512493 | Assisted Replacement | SC | SC | pCO2 | calibration | pCO2 Calibration error | . | SC replace.
No | Visit | Complaint | RKOR | Quick access error | . | OS, Software reinstall. | SWO-512492 | Defect | SW | RMED SW | RMED SW | Quick access | Quick access error | . | OS, Software reinstall.
No | Visit | Complaint | RCHN | 仪器偶发卡死，试剂包漏液 | 仪器偶发卡死，试剂包漏液 | 检测确认问题属实，更换CF卡后修复-by Hunter | SWO-512717 | [5 empty] | Instrument occasionally stuck, reagent bag leakage | Instrument occasionally stuck, reagent bag leakage | Detection confirms the problem is true, repair after replacing CF card - by Hunter
No | Visit | Complaint | RAME | Pump making noises | Solutions pump noisy. | Reseated solutions pump. | SWO-512240 | [5 empty] | Pump making noises | Solutions pump noisy. | Reseated solutions pump.
No | Visit | Complaint | RIND | V/C failed in every SC | V $C failed in every cassette due to sensor connector is not working properly. after zeroing O2 value is not in range | Change sensor connector and zeroing done. calibration also passed. Instrument working fine now. instrument is under observation | SWO-511891 | [5 empty] | V/C failed in every SC | V /C failed in every cassette due to sensor connector is not working properly. after zeroing O2 value is not in range | Change sensor connector and zeroing done. calibration also passed. Instrument working fine now. instrument is under observation
No | Visit | Complaint | RCHN | 无法抽吸样本，进样口排线有断痕 | 无法抽吸样本，进样口排线有断痕 | 更换636-452后修复-by Hunter | SWO-512688 | Defect | HW | Inlet module | Cable |  | Cannot draw samples, the inlet line is broken | Cannot draw samples, the inlet line is broken | Repair-by-Hunter after replacement 636-452
No | Visit | Complaint | RCHN | 处理盘卡盘，干燥前温度超范围 | 处理盘卡盘，干燥前温度超范围，借580-098,903-210各一个 | 更换580-098,903-210后修复by tim | SWO-511145 | Defect | HW | Process Module | Motor, process wheel | 5784 | Processing disk, pre-drying temperature over range | Processing disk, pre-drying temperature over range, borrow 580-098, 903-210 each one | Replace 580-098,903-210 after repair by tim
No | Visit | Complaint | RKK | PO2測定 高値乖離 | PO2測定 時々 高値乖離 | センサーボードとメンテナンスキット交換で様子見 | SWO-511948 | [5 empty] | PO2 measurement high deviation | PO2 measurement sometimes high deviation | The sensor board and the maintenance kit are replaced.
No | Visit | Complaint | RCHN | 样本量不足，K定标不过，报错0210,0719 | 样本量不足，K定标不过，报错0210,0719 | 更换K电极膜，清洁测量仓后修复 | SWO-514451 | Assisted Replacement | Electrode | Membrane | K | calibration | Insufficient sample size, K calibration, error 0210, 0719 | Insufficient sample size, K calibration, error 0210, 0719 | Replace the K electrode film and clean the measuring chamber after repair
No | Visit | Complaint | RDEU | 08.03.2019, Frau Suhr, Tel: 03834865528: QC Fehler bei CRP- Frau Suhr hat bereits alles mit Herrn Ripke besprochen. | CRP QC Lvl 2 zu hoch (5x) | System überprüft - alle Key Komponenten gereinigt - keine gr. Verschmutzung gefunden SD erstellt Pumpe P2 getauscht, da Fließwege erhöht waren, nun 1300-1400 | SWO-512197 | Defect | HW | Wet section | P2 | Flow restriction | 08.03.2019, Mrs Suhr, Tel: 03834865528: QC Error at CRP- Ms. Suhr has already discussed everything with Mr. Ripke. | CRP QC Lvl 2 zu hoch (5x) | System checked - all key components cleaned - no gr. Pollution found SD created Pump P2 replaced as flow paths were increased, now 1300-1400
No | Visit | Complaint | RAME | Electrode Sensor disconnected | Electrode disconnected | Replaced module | SWO-512073 | [5 empty] | Electrode Sensor disconnected | Electrode disconnected | Replaced module
No | Visit | Complaint | RPOL | Wymiana elektrod Crea , zmiana konfiguracji Netop | Błąd kontroli jakości Crea | Wymiana elektrod i membran do kreatyniny. Sprawdzenie aparatu. Zmiana konfiguracji ustawień Netop. | SWO-512250 | Assisted Replacement | Electrode | Electrode | Crea |  | Replacing the Crea electrodes, changing the Netop configuration | Crea Quality Control Error | Replacement of electrodes and membranes for creatinine. Check the camera. Changing the configuration of Netop settings.
No | Visit | Complaint | RCHN | 科室老师反馈近期结果打问号 | 科室老师反馈近期结果打问号（未更换备件） | 工程师到现场进行质控，观察发现没有明显异常，而且测试正常，后期再和科室老师保持沟通by tim | SWO-512425 | No fault found | No fault found | No fault found | No fault found |  | Section teacher feedback recent results to question mark | Section teacher feedback recent results to hit the question mark (no replacement spare parts) | Engineers to the scene to carry out quality control, observation found no obvious anomalies, and test normal, and later and the department teacher to maintain communication by tim
Yes | Visit | Complaint | RCHN | 系统检查光学单元散光灯强度超范围 | 系统检查光学单元散光灯强度超范围，借903-364一个（技术支持田工3 $7同意借件） | 更换903-364后修复(定标报错1031）by tim | SWO-510392 | [5 empty] | System to check optical unit diffuse light intensity over range | System check optical unit diffuse light intensity over range, borrow 903-364 one (technical support field worker 3 /7 agreed to borrow) | Replacement 903-364 after repair (calibration error 1031) by tim
No | Visit | Complaint | RCHN | 所有离子的测量结果不准确 | 所有离子的测量结果不准确，借902-972和903-334各一个 |  | SWO-512527 | [5 empty] | Inaccurate measurements of all ions | The measurement sits on all ions is inaccurate, with one each of the 902-972 and 903-334
No | Visit | Complaint | RCHN | 接触探针弹性没有了 | 接触探针弹性没有了，借891-051一个 |  | SWO-512505 | Defect | HW | Basic Unit | SC contact |  | The contact probe elasticity is gone. | Contact probe elasticity is gone, borrow 891-051 a
No | Visit | Complaint | RCHN | 加热温度超范围 | 加热温度超范围，借903-210一个 |  | SWO-512503 | Defect | HW | Process Module | PCB, cup dryer | 1068, 1069 | Overheating temperature over range | Overheating temperature over range, borrowed 903-210 one
No | Visit | Complaint | RCHN | 液路堵塞 | 液路堵塞，借903-112一个（未更换备件） | 因903-112缺货，故工程师先清洗AB管，但故障仍然存在，报错FCT103，故加借903-329和902-978-by Hunter | SWO-512406 | Revisit required | HW | Wet section | PCB, wet section | 3412 | Liquid road blockage | Liquid road blockage, borrow edgy 903-112 (spare parts not replaced) | Because 903-112 is out of stock, so the engineer first clean ab tube, but the fault is still there, reported the error FCT103, so borrow 903-329 and 902-978-by Hunter
No | Visit | Complaint | RDEU | 08:00Uhr Herr Kremer tel:+4920360083681: Das Gerät schaltet sporadisch ab | Gerät schaltet sich immer wieder ab, nach Datenarchivierung fährt das Gerät herunter und macht einen Warmstart, | Ich habe die autom. Archivierung abgeschaltet um zu testen, | SWO-513218 | Adjustment | SW | RMED SW | RMED SW | sporadic breakdown | 08:00H Mr Kremer tel:+4920360083681: The device switches off sporadically | device shuts down again and again, after data archiving, the device shuts down and makes a warm start, | I have the autom. Archiving switched off to test
No | Visit | Complaint | RLTD | Yaser on 0208 869213 | Leak detected unable to pump solution clot detected. | Flushed system and found poor flow. Removed SC and flushed from inlet port out to probe and found clot in inlet connector gasket. Replaced inlet connector gasket. Flushed and run service programs and all passed ok. Analyser start up in 50 minuets. Analyser left to start up customer will monitor. | SWO-513677 | Defect | HW | Auto Inlet module | Inlet connector gasket | Clot | Yaser on 0208 869213 | Leak detected unable to pump solution clot detected. | Flushed system and found poor flow. Removed SC and flushed from inlet port out to probe and found clot in inlet connector gasket. Replaced inlet connector gasket. Flushed and run service programs and all passed ok. Analyser start up in 50 minuets. Analyser left to start up customer will monitor.
No | Visit | Complaint | RDEU | Lösungspack nicht korrekt installiert | Lösungspack nicht korrekt installiert, Chipdaten nicht lesbar seit zwei Tagen | SP entfernt, Kontakte sowie Flowselectorfenster gereinigt, Serviceprogramme ok, weiterhin SP nicht erkannt, SP90 (HC07) ersetzt, Startphase, Leckage erkannt, Einlaßdichtungshalter ersetzt, SC90 ersetzt, Startphase abgewartet, Kal und QK ok | SWO-513662 | Assisted Replacement | SP | SP | chip |  | Solution pack not installed correctly | Solution pack not installed correctly, chip data not readable for two days | SP removed, contacts and flowselector window cleaned, service programs ok, still SP not detected, SP90 (HC07) replaced, start phase, leakage detected, inlet seal holder replaced, SC90 replaced, start phase waited, Kal and QK ok
No | Visit | Complaint | RKK | キャルでｋインピーダンスエラーが突然出る | インタフェースPCB交換 | インタフェースPCB交換。 交換後、同じSC,SPでエラーなし | SWO-518837 | Defect | HW | Basic Unit | PCB, Sensor interface | K | K impedance error suddenly appears in the car | Interface PCB Replacement | Interface PCB replacement. No errors on same SC,SP after replacement
No | Visit | Complaint | RCHN | 接口板腐蚀 | 接口板腐蚀，借902-972一个 | 更换902-972及测试卡后修复-by Hunter | SWO-512528 | [5 empty] | Interface plate corrosion | Interface plate corrosion, borrow 902-972 one | Replace 902-972 and repair-by-Hunter after test card
No | Visit | Complaint | RCHN | 电池无法蓄电 | 电池无法蓄电,先给医院使用备用机。 | 先给医院使用备用机。送修维修间。 | SWO-518370 | [5 empty] | The battery can't store power. | The battery can't store power, first use a spare machine for the hospital. | Start by using a spare machine for the hospital. Send the repair room.
No | Visit | Complaint | RCHN | 报错4548，进样门偶尔打不开 | 报错4548，进样门偶尔打不开，借903-028一个 | 更换903-028后修复 | SWO-512424 | Defect | HW | Cover and Chassis | PCB, cabinet top | 4548, Door, inlet | Reported error 4548, into the sample door occasionally can not open | Reported error 4548, into the sample door occasionally can not open, borrow 903-028 one | Repair after replacement 903-028
No | Visit | Complaint | RCHN | 仓门关不上 | 仓门关不上，借886-638,886-639各一个 |  | SWO-512727 | Defect | HW | Cover and Chassis | Door, Inlet |  | The cangcang door can't be closed. | Cangmen closed, borrow 886-638, 886-639 each one
No | Visit | Complaint | RCHN | 系统崩溃 | 系统崩溃，借944-258（15488），944-212（15529缺货，15457），944-216（15446，15107），944-214（15229缺货，15057），944-268（15495），944-394（15624）各一个 |  | SWO-512754 | Defect | SW | RMED SW | Calibration data erased |  | System crash | System crashes, 944-258 (15488), 944-212 (15529 out of stock, 15457), 944-216 (15446, 15107), 944-214 (15229 out of stock, 15057), 944-268 (15495), 944-394 (15624) each
No | Visit | Complaint | RCHN | 定标失败 | 定标终止，借944-258（14986）一个 |  | SWO-512835 | Assisted Replacement | Consumables | Test Kit | Cartridge, cal |  | Bid failed | Bid termination, loan 944-258 (14986) one
No | Visit | Complaint | RCHN | 液路堵塞，FCT_103 | 液路堵塞，FCT_103，借903-329、902-978 |  | SWO-512715 | Revisit required | HW | Wet section | PCB, wet section | 3412 | Liquid Road Blocked, FCT-103 | Liquid Road Blocked, FCT-103, loan 903-329, 902-978
No | Visit | Complaint | RKK | 「重大なシステムエラー」の表示で画面が真っ赤になっていた。装置リスタートもソフト立ち上がらず。 | 立ち上がらず | OS入れ直し | SWO-516186 | [5 empty] | The screen was red with a "critical system error" display. The device restart does not stand up softly. | Don't stand up. | Os re-entry
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | NA定标不过 0375 | 更换测试卡后修复 | SWO-514485 | [5 empty] | Test card, reagent package failure | NA calibration is no more than 0375 | Repair after replacing test card
No | Visit | Complaint | RSAF | 2pt cal intermittently passes and fails on tHb. | Halogen lamp intensity too high | Lamp unit adjustment done, complete oximetry test done. 2pt cal ok. | SWO-513143 | Adjustment | HW | Oxi module | Lamp unit | halogen lamp intensity too high | 2pt cal intermittently passes and fails on tHb. | Halogen lamp intensity too high | Lamp unit adjustment done, complete oximetry test done. 2pt cal ok.
No | Visit | Complaint | RIND | Electrolyte values are coming variations for all the samples. | Found the clot in Reference membrane, then clean the chamber and membrane then run the instrument its working good. | Found the clot in Reference membrane, then clean the chamber and membrane then run the instrument its working good. | SWO-513139 | Cleaned | HW | pH/BG Module | measuring chamber | Clot | Electrolyte values are coming variations for all the samples. | Found the clot in Reference membrane, then clean the chamber and membrane then run the instrument its working good. | Found the clot in Reference membrane, then clean the chamber and membrane then run the instrument its working good.
No | Visit | Complaint | RDAN | Nina klager over fejl 374 selvom hun har renset med hypoklorid. | Der ses stadig en del prøver med prøvefejl, primært uhomogene prøver. | Aftalt forsøgsvis udskiftning af pH $BG-modulet. | SWO-513898 | Defect | HW | pH/BG Module | pH/BG Module | 374 | Nina complains of error 374 although she has cleaned with hypochlorite. | There are still a few samples of test faults, mainly unhomogeneous samples. | The pH $BG module.
No | 3 | Y | RCZE | Blank |  |  | D5CDF8CAFE1D1BC9 | Defect | SW | RMED SW | RMED SW | Database | Blank
No | Visit | Complaint | RDEU | Drucker meldet Offlinefehler | Drucker im Offlinemodus | Wackelkontakt beim Drucker - Stromversorgungskabel. Stecker war nicht vollständig aufgesteckt und löste sich . Kontakte leicht nachgebogen und Stecker bis Anschlag aufgesteckt. Drucker wieder o.k. | SWO-513551 | Adjustment | HW | Analyzer Control Module | Cable, printer |  | Printer reports offline error | Printers offline | Wobbly contact on printer - power supply cable. Plug was not fully plugged in and detached. Contacts slightly bent and plug plug plugs plugged up. Printer again OK.
No | Visit | Complaint | RDEU | Spontane HCT Mess-Fehler | Empfohlene Maßnamen von TS durchgeführt. Fehler 1223,1264 | Pumpe 3+4 , Needle Assembly und FilterUnit HCT-Version erneuert. Reinigung, Systemüberprüfung und QC durchgeführt | SWO-513552 | Defect | HW | Aspiration Module | P4 | 1223, 1260, 1261, 1264 | Spontaneous HCT measurement errors | Recommended dimension names performed by TS. Error 1223,1264 | Pump 3+4 , Needle Assembly and FilterUnit HCT version renewed. Cleaning, system verification and QC performed
No | Visit | Complaint | RPOL | Błąd kalibracji Glu, sprawdzenie aparatu. | Błąd Kalibracji Glu | Czyszczenie elektrody Glu, wymiana membrany. Ustawienie harmonogramu wymiany membran, kalibracja hemoglobiny. Szkolenie personelu technicznego z zakresu podstawowych czynności konserwacyjnych. | SWO-513043 | Assisted Replacement | Electrode | Membrane | Glu |  | Glu calibration error, check the camera. | Glu Calibration Error | Cleaning the Glu electrode, replacing the membrane. Setting of the membrane replacement schedule, hemoglobin calibration. Training of technical personnel in the field of basic maintenance.
No | Visit | Complaint | RDEU | 08.03.2019 11:30 Fr. Bernhard 0212992948027: sporadische Kal-Fehler Hb 11.03.2019 8:30 Fr. Bongartz 0212992948027: weiterhin Kal-Fehler ctHb | Sporadisch c-Thb Fehler | Im Service Oxi-complete Test durchgeführt -- bestanden, Quetschventil geprüft und gereinigt , Quetschventilschlauch verschlissen -- ersetzt, Fehler seit Lösungspack Wechsel , LP ersetzt, mehrere Kalibrationen durchgeführt und bestanden | SWO-513798 | Assisted Replacement | SP | SP | Oxi | tHb calibration, sporadic | 03/08/2019 11:30 Fr. Bernhard 0212992948027: sporadic Kal error Hb 11.03.2019 8:30 Fr. Bongartz 0212992948027: still Kal error ctHb | Sporadisch c-Thb Fehler | Performed in the Service Oxi-complete test -- passed, crushing valve tested and cleaned, squeeze valve hose worn -- replaced, error since solution pack change, LP replaced, several calibrations performed and passed
No | Visit | Complaint | RDEU | 11.03.2019, Heike, Tel: 015170824649: Disk boot Fehler | disk boot failure -> cpu def. | Fehlersuche; CF-Karte gew.; Neustart = Fehler; Versuch, Gerät v. Stick zu booten -> ohne Erfolg = CPU unit def.; CF-Karte zurückgetauscht; Abbruch d. Arbeiten; Material für den Folgetag zum Kunden bestellt | SWO-513973 | Revisit required | HW | DMS_PC | CPU | disc boot failure | 11.03.2019, Heike, Tel: 015170824649: Disk boot Error | disk boot failure -> cpu def. | Troubleshooting; CF card. Restart = Error; Attempt to boot device v. Stick -> without success = CPU unit def.; CF card exchanged; Demolition of the works; Material ordered to the customer for the following day
No | Visit | Complaint | RDEU | 07.03.2019 10:50Uhr Herr Widera tel:+492191138584: Der bereits gemeldete Fehler pCO2 im LevelA zu hoch 08.03.2019, 8.15Uhr, Herr Widera, Tel: 02191138584: erneut QC Fehler bei pCO2 level C 12.03.2019; 10:30 Uhr; Hr. Widera; Tel.: 02191138580; erneut pCO2 QC Fehler A 12.03.2019; 11:20 Uhr; Hr. Widera; Tel.: 02191138580; immer noch pCO2 QC Fehler A | Pco2 Level A zu hoch, | Komplettes Wartungskit verbaut, Goldkontakte gereinigt, Quetschventil gereinigt, Sensorkassette ersetzt, Startphase PCO2 Level A jetzt 28,4mmhg ,Pco2 Empf. 80,1% | SWO-515180 | Assisted Replacement | SC | SC | pCO2 | QC A too high and C | 07.03.2019 10:50Uhr Herr Widera tel:+492191138584: Der bereits gemeldete Fehler pCO2 im LevelA zu hoch 08.03.2019, 8.15Uhr, Herr Widera, Tel: 02191138584: erneut QC Fehler bei pCO2 level C 12.03.2019; 10:30 Uhr; Hr. Widera; Tel.: 02191138580; erneut pCO2 QC Fehler A 12.03.2019; 11:20 Uhr; Hr. Widera; Tel.: 02191138580; immer noch pCO2 QC Fehler A | Pco2 Level A too high, | Complete maintenance kit installed, gold contacts cleaned, squeezing valve cleaned, sensor cassette replaced, start phase PCO2 Level A now 28.4mmhg ,Pco2 Empf. 80,1%
No | Visit | Complaint | RDEU | Drucker sporadisch ohne Funktion | Drucker sporadisch ohne Funktion | Windows7, 3.4MR1 ,Konfiguration installiert, Gerät kalibriert und mit QCs überprüft,OK | SWO-514351 | Setup | SW | RMED SW | RMED SW | Printer | Printer sporadically without function | Printer sporadically without function | Windows7, 3.4MR1 ,Configuration installed, device calibrated and checked with QCs,OK
No | Visit | Complaint | RMAY | Oncall breakdown | -Last sample patient was clotted ( aborted) -Found blockage in waste nipple | -Check analyser data log - Perform manual flush -Run fluid transport test ( OK) | SWO-514620 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway | Clot, 1063 | Oncall breakdown | -Last sample patient was clotted ( aborted) -Found blockage in waste nipple | -Check analyser data log - Perform manual flush -Run fluid transport test ( OK)
No | Visit | Complaint | RFRA | Machine ne démarre plus suite à un redémarrage. Reste bloquée sur Data Upgrade Failed (Status 2). Essai de démarrage de l'analyseur sans le pack, meme résultat. Intervention FSE demandée. | Contrôle de la carte mère | Changement de la carte mère. Reconfiguration. Essais | SWO-514064 | Defect | HW | DMS_PC | CF card | Status code 2 | Machine no longer starts after a restart. Remains blocked on Data Upgrade Failed (Status 2). Start test of the analyzer without the pack, same result. ESF intervention requested. | Motherboard control | Changing the motherboard. Reconfiguration. Tests
No | Visit | Complaint | RAME | Liquid Sensor Error | Analyzer is showing liquid sensor error during initialization. Upon further inspection, the blood gas liquid sensor value is out of range. | Replaced sensor bay unit. | SWO-514206 | Defect | HW | Basic Unit | Sensor Bay Unit | LS | Liquid Sensor Error | Analyzer is showing liquid sensor error during initialization. Upon further inspection, the blood gas liquid sensor value is out of range. | Replaced sensor bay unit.
No | Visit | Complaint | RAME | Intervention required | Unit keeps returning to maintenance screens. In flush screen on arrival. | Flushed unit thoroughly. Installed new inlet gasket with holder. Reseated inlet gasket. Installed a new solution pack. | SWO-513916 | Defect | HW | Auto Inlet module | Inlet gasket |  | Intervention required | Unit keeps returning to maintenance screens. In flush screen on arrival. | Flushed unit thoroughly. Installed new inlet gasket with holder. Reseated inlet gasket. Installed a new solution pack.
No | Visit | Complaint | RCHN | 液路堵塞 | 液路堵塞（未更换备件） | 更换管路后修复by Hunter | SWO-514527 | Defect | HW | Wet section | Tubing |  | Liquid road blockage | Liquid road blockage (no replacement parts) | Repair by Hunter after replacing lines
No | Visit | Complaint | RCHN | 样本量不足 | 样本量不足 | 更换库存636-561,903-329后修复by Hunter | SWO-514324 | Defect | HW | Aspiration Module | Needle Wash Unit | 1218 | Insufficient sample size | Insufficient sample size | Repair by Hunter after replacing stock 636-561,903-329
No | Visit | Complaint | RCHN | 进样舱门打开已取出样本很久之后，仍然提示请取出样本 | 进样舱门打开已取出样本很久之后，仍然提示请取出样本（未更换备件） | 清洁扫试管条码器和传感器后修复(by chuanwei) | SWO-515116 | Cleaned | HW | Aspiration Module | Barcode reader, Inlet |  | After the sample door has been opened and the sample has been taken out for a long time, still prompt to remove the sample | After the intake door has been opened and the sample has been taken out for a long time, still prompt to remove the sample (spare parts not replaced) | Repair after cleaning the tube bar code and sensor (by chuanwei)
No | Visit | Complaint | RCHN | 进样口漏液 | 进样口漏液,清洁进样口，废液接口。（未更换备件。） | 清洁进样口，废液接口。（未更换备件。）by Hunter | SWO-514481 | Cleaned | HW | Basic Unit | Inlet module |  | Inlet leak | Inlet leakage, clean the sample port, waste liquid interface. (Spare parts have not been replaced.) ) | Clean the sample port, waste fluid interface. (Spare parts have not been replaced.) by Hunter
No | Visit | Complaint | RDEU | Das Gerät bootet nicht (AU läuft?) | Keine Funktion | Gerät demontiert, CPU ausgetauscht, Analysator an Aqure angeschlossen, Funktionskontrolle ok | SWO-514735 | Defect | HW | DMS_PC | DMS_PC |  | The device is not booting (AU is running?) | No function | Device dismantled, CPU replaced, analyzer connected to Aqure, function control ok
No | Visit | Complaint | RAME | The printer has streaks in the print | The printer has streaks | I replaced the printer | SWO-514777 | Defect | HW | DMS_PC | Printer Module |  | The printer has streaks in the print | The printer has streaks | I replaced the printer
No | Visit | Complaint | RAME | El/Met lower liquid sensor error | El Met lower LS error | I replaced the LS and the LS cal came in. | SWO-514776 | Defect | HW | El/Met Module | LS |  | El/Met lower liquid sensor error | El Met lower LS error | I replaced the LS and the LS cal came in.
No | Visit | Complaint | RIND | pO2 values are coming low. | Po2 values are coming low, found the clot in pO2 membrane then clean the chamber and membrane then run the instrument its working good. | Po2 values are coming low, found the clot in pO2 membrane then clean the chamber and membrane then run the instrument its working good. | SWO-516256 | Cleaned | HW | pH/BG Module | measuring chamber | pO2 QC low | pO2 values are coming low. | Po2 values are coming low, found the clot in pO2 membrane then clean the chamber and membrane then run the instrument its working good. | Po2 values are coming low, found the clot in pO2 membrane then clean the chamber and membrane then run the instrument its working good.
No | Visit | Complaint | RIND | TropninI calibration failed. | Check the calibration message & found back ground check failure. | Replace the optical & process wheel with the new one,run the tropnin I calibration & it got passed successfully. | SWO-516436 | Defect | HW | Process Module | Optical Unit | 1024 | TropninI calibration failed. | Check the calibration message & found back ground check failure. | Replace the optical & process wheel with the new one,run the tropnin I calibration & it got passed successfully.
No | Visit | Complaint | RPOL | Błąd skanera FlexQ | Brak czerwonego światła od skanera | Wymiana modułu FkexQ. Kalibracje | SWO-513828 | Defect | HW | FlexQ Module | FlexQ Module | scanner | FlexQ Scanner Error | No red light from scanner | Replacing the FkexQ module. Calibrations
No | Visit | Complaint | RDEU | 11.03.2019, Heike, Tel: 015170824649: Disk boot Fehler 12.03.2019, Sr. Carola +4917614022104. wartet noch auf Kontaktaufnahme | - disk boot failure - blue screen | Gerät aus; PC-Modul (alt: 910-342P1827N9580) gew.; SW aufgespielt (V3.4MR1); LP u. SC neu initialisiert; Startphase abgewartet; div. Windows-Einstellungen vorgenommen; wieder Startphase; Konfig. aufgespielt; IT-Anbindung aufgrund von Windows7 noch offen; an Hr. Siebert kommuniziert; Kal. + QC i.O. | SWO-514758 | Defect | HW | DMS_PC | CPU | disc boot failure | 11.03.2019, Heike, Tel: 015170824649: Disk boot Error 12.03.2019, Sr. Carola +4917614022104. is still waiting to get in touch | - disk boot failure - blue screen | device off; PC module (old: 910-342P1827N9580); SW (V3.4MR1); LP and SC reinitialized; start-up phase awaited; div. Windows settings made; start-up phase again; Config. played up; IT connection due to Windows 7 still open; communicated to Mr. Siebert; Kal. + QC in the first place
No | Visit | Complaint | RDEU | 11.03.2019, Frau Rottmann, Tel: 03834871240: Display ist lose und wackelt bei Bedienung | Rückwand des Displays an allen vier Stellen ausgebrochen | Displayrahmen gewechselt | SWO-515990 | Defect | HW | Display | Display | screws | 11.03.2019, Mrs. Rottmann, Tel: 03834871240: Display is loose and wobbles when operated | Back wall of the display erupted in all four places | Display frame changed
No | Visit | Complaint | RAME | pO2 issues | API proficiency results out on PO2. | Performed PO2 complete service test. Also performed oxi test and tHb calibration. | SWO-514873 | No fault found | No fault found | No fault found | pO2 |  | pO2 issues | API proficiency results out on PO2. | Performed PO2 complete service test. Also performed oxi test and tHb calibration.
No | Visit | Complaint | RLTD | Suspected clot in sensor cassette | I checked the fluid transport service program and found that the sensor clot detect value was quite near the upper limit, there were no other errors in the fluid transport programs. Checking the datalogs and there were a number of errors relating to clots being detected in the sensor cassette, but a new cassette has been fitted so the problem is not with that. | An old sensor cassette was fitted and the sample path was soaked with protein remover to try and break up any remnants of a clot that may have been remaining. Annual service kit fitted to replace the sample path tubings. Inlet probe replaced. Fluid transport service programs run and the clot detect value for the sensor had come back down to the middle of the range. Two capillary samples were run after the analyser had come back to ready and these both aborted with a failed to aspirate message even though the entire sample was taken in to the analyser. After this two capillary samples with water were run and these went through ok, then two blood samples obtained from members of staff went through and one went through ok and the other aborted because of air bubbles in the sample, it did aspitate it ok though. No further samples were available from the department at that time but after I had left it was reported that the analyser was still aborting samples. I said I would order futher parts and return on Friday to investigate further. | SWO-514725 | Revisit required | Fluidic pathway | Fluidic pathway | Fluidic pathway | Clot | Suspected clot in sensor cassette | I checked the fluid transport service program and found that the sensor clot detect value was quite near the upper limit, there were no other errors in the fluid transport programs. Checking the datalogs and there were a number of errors relating to clots being detected in the sensor cassette, but a new cassette has been fitted so the problem is not with that. | An old sensor cassette was fitted and the sample path was soaked with protein remover to try and break up any remnants of a clot that may have been remaining. Annual service kit fitted to replace the sample path tubings. Inlet probe replaced. Fluid transport service programs run and the clot detect value for the sensor had come back down to the middle of the range. Two capillary samples were run after the analyser had come back to ready and these both aborted with a failed to aspirate message even though the entire sample was taken in to the analyser. After this two capillary samples with water were run and these went through ok, then two blood samples obtained from members of staff went through and one went through ok and the other aborted because of air bubbles in the sample, it did aspitate it ok though. No further samples were available from the department at that time but after I had left it was reported that the analyser was still aborting samples. I said I would order futher parts and return on Friday to investigate further.
Yes | Visit | Complaint | RDEU | 11.03.2019 Frau Selmann +49401818828816 berichtet, dass ihr vor Wochen versprochen wurde, das der Analysator ausgetauscht wird. | Abholung | Analysator abgeholt. Und Kundengespräch durchgeführt. | SWO-514421 | Re classified RMED | planned SWO |  |  | Analyzer removed and changed to 092R0196N024 | 11.03.2019 Ms. Selmann +49401818828816 reports that she was promised weeks ago that the analyzer will be replaced. | Pick up | analyzer. And customer talk conducted.
No | Visit | Complaint | RCHN | CA定标失败 | CA定标失败 | 更换测试卡后修复by Hunter | SWO-514430 | [5 empty] | CA calibration failure | CA calibration failure | Fix by Hunter after replacing test card
No | Visit | Complaint | RIND | Lac failure. | Lac value not stable. | Replaced the signal data board, check the zeroing process but again lac unstable. Checked the earthing voltage its also good value. | SWO-516433 | [5 empty] | Lac failure. | Lac value not stable. | Replaced the signal data board, check the zeroing process but again lac unstable. Checked the earthing voltage its also good value.
No | Visit | Complaint | RCHN | 科室老师反馈电解质结果不准 | 科室老师反馈电解质结果不准 | 目前来看结果没有明显差异，仪器也是正常的，故工程师与科室老师针对报告进行了解读，后期会持续跟进中（by chuanwei） | SWO-515784 | [5 empty] | Department teacher feedback electrolyte results are not allowed | Department teacher feedback electrolyte results are not allowed | At present, there is no obvious difference in the results, the instrument is also normal, so the engineer and the department teacher for the report was interpreted, the latter will continue to follow up (by chuanwei)
No | Visit | Complaint | RCHN | 无法开机 | 无法开机，借910-342一个 | 更换910-342后修复by Hunter | SWO-513559 | Defect | HW | DMS_PC | CPU |  | Cannot power on | Cannot power on, borrow 910-342 one | Repair by Hunter after replacing 910-342
No | Visit | Complaint | RCHN | 进样针堵塞 | 进样针堵塞 | 更换进样针924-455后修复by Hunter | SWO-514350 | Defect | HW | Inlet module | Probe |  | The sample needle is blocked. | The sample needle is blocked. | Repair by Hunter after replacing the sample needle 924-455
No | Visit | Complaint | RCHN | 进样门仓无法关闭 | 进样门无法关闭 5784 | 调节进样仓门后修复by Hunter | SWO-514428 | Adjustment | HW | Cover and Chassis | Door, Inlet | 5784 | The sample door cannot be closed | The sample door cannot be closed 5784 | Fix by Hunter after adjusting the entrance door
No | Visit | Complaint | RCHN | 样本盘仓门卡顿 | 样本盘仓门卡顿 5689 | 调整仓门松紧度后修复by Hunter | SWO-514359 | Adjustment | HW | Cover and Chassis | Door, Inlet | 5689 | Sample warehouse door Caton | Sample Warehouse Gate Caton 5689 | Fix by Hunter after adjusting the tightness of the door
No | Visit | Complaint | RCHN | O2定标灵敏度超范围 | O2定标灵敏度超范围，借945-613一个 | 更换945-613后修复by Hunter | SWO-513671 | Assisted Replacement | Electrode | Electrode | pO2 | calibration | O2 calibration sensitivity over range | O2 calibration sensitivity over range, with 945-613 one | Repair by Hunter after replacement 945-613
No | Visit | Complaint | RIND | Thb/so2 error | Thb $so2 ? for all samples | checked the analyser thb $so2 values was not coming properly so cleaned the oxy module tubes with hypo solution and hemolyzer on and off done for few sec and then cleane dthe oxy module tubes with distill water and then complete oximetry test done then calibratead the analyser to ready mode | SWO-516414 | Cleaned | HW | Oxi module | Oxi module | thb/sO2 | Thb/so2 error | Thb /so2 ? for all samples | checked the analyser thb $so2 values was not coming properly so cleaned the oxy module tubes with hypo solution and hemolyzer on and off done for few sec and then cleane dthe oxy module tubes with distill water and then complete oximetry test done then calibratead the analyser to ready mode
No | Visit | Complaint | RAME | Leakage error code 0663 | Error code 0663, leak detected, reported. Customer had previously worked with tech support over the phone and replaced inlet gasket, pO2, Ca, and reference membranes. Upon arrival, during visual inspection, leakage was visible from the El $Met module. Inspected the electrodes and found the Ca membrane was missing the white, plastic seal ring, which was allowing a free flow of fluid to leak through the electrode bay. | Replaced the Ca membrane. Op test satisfactory. Normal operations restored. | SWO-514891 | Defect | Electrode | Electrode | Sealing ring, Ca | missing sealing ring | Leakage error code 0663 | Error code 0663, leak detected, reported. Customer had previously worked with tech support over the phone and replaced inlet gasket, pO2, Ca, and reference membranes. Upon arrival, during visual inspection, leakage was visible from the El /Met module. Inspected the electrodes and found the Ca membrane was missing the white, plastic seal ring, which was allowing a free flow of fluid to leak through the electrode bay. | Replaced the Ca membrane. Op test satisfactory. Normal operations restored.
Yes | 2 | Y | RCZE | Blank |  |  | D5D0D991EA4FD04D | Re classified RMED | planned SWO |  |  |  | Blank
No | Visit | Complaint | RDEU | Druckbild mangelhaft. | Druckbild kaum lesbar. Kalibrationsfehler mehrere Parameter. | Drucker erneuert. Mehrere Kalibrationen. | SWO-515595 | Defect | HW | DMS_PC | Printer Module | printer head | Print image deficient. | Print image hardly readable. Calibration error multiple parameters. | Printer renewed. Multiple calibrations.
No | Visit | Complaint | RDEU | 27.02.2019, Hewrr Dadendick, 0941782350, Datenverbindung verloren | Zeitweise längerer Ausfall der Netzwerkverbindung. Im Fehlerfall ist im Control Panel die Netzwerkverbindung (Settings) nicht vorhanden. CPU Unit erneuern. | CPU Unit erneuert. ABL verbindet sich mit PDMS und Conworx. Kalibrationen und QC´s vorgenommen. Testmessung durchgeführt. Daten werden lt. Hr. Reiser übertragen. | SWO-514446 | Revisit required | HW | DMS_PC | CPU | mail: send defective CPU unit back to RMED | 27.02.2019, Hewrr Dadendick, 0941782350, data connection lost | Temporarily prolonged network connection failure. In the event of an error, the network connection (settings) does not exist in the Control Panel. Cpu Unit. | CPU Unit. ABL connects with PDMS and Conworx. calibrations and QC's. test measurement. Data is transferred according to Mr. Reiser.
No | Visit | Complaint | RAME | Electrode Sensor disconnected | Module disconnected | Replaced connector pcb | SWO-515015 | [5 empty] | Electrode Sensor disconnected | Module disconnected | Replaced connector pcb
No | Visit | Complaint | RDEU | Oxymetrie und pCO2 Fehler Level A Oxymetriefehler sporadisch bei Kalibration. | Oximetrie und pCO2 Fehler Level A Oximetriefehler sporadisch bei Kalibration. 0581 bei QC Messung. Immer wieder inhomogene Lösungen. LP wurden mehrfach gewechselt (unterschiedliche Lot). Flüssigkeitsaustritt aus Hemolyzer. | Hemolyzer Unit erneuert. Oxi Complete Test - ok. tHb Kalibration, Kalibrationen und QC´s vorgenommen. 3 Testmessung durchgeführt - ok. Reparatur im Zuge der Wartung durchgeführt. | SWO-515288 | Defect | HW | Oxi module | Hemolyzer |  | Oxymetry and pCO2 Error Level A Oxymetry error sporadically during calibration. | Oximetry and pCO2 Error Level A Oximetry error sporadically during calibration. 0581 for QC measurement. Again and again inhomogeneous solutions. LP were changed several times (different lot). Liquid leakage from hemolyzer. | Hemolyzer Unit. Oxi Complete Review - ok. tHb calibration, calibrations and QC's. 3 Test measurement performed - ok. repair carried out in the course of maintenance.
No | Visit | Complaint | RAME | Inlet gasket issues | Inlet Gasket stuck in and won’t come back out after sample aspirated. Inlet probe was bent. | Replaced Inlet Probe. | SWO-515457 | Defect | HW | Inlet module | Probe bend |  | Inlet gasket issues | Inlet Gasket stuck in and won’t come back out after sample aspirated. Inlet probe was bent. | Replaced Inlet Probe.
No | Visit | Complaint | RFRA | Erreur statut code 2. | Erreur code 2 | Récupération sauvegardes sur Z. Recharge programme et sauvegarde( toutes données et configuration). Nettoyage joint d'introduction. Calibration. | SWO-515581 | Defect | SW | RMED SW | RMED SW | Status code 2 | Code 2 status error. | Code 2 error | Recovering backups on Z. Recharge program and backup (all data and configuration). Introducing joint cleaning. Calibration.
No | Visit | Complaint | RFRA | L'analyseur est lent Beug Se déconnecte de l'utilisateur | Vérification complète de l'analyseur | Sauvegarde de toutes les données et de la configuration $ $ Remplacement de la cpu $ $ Installation de windows7 $ $ configuration de windows $ $ configuration de la sauvegarde réseau (info non présente sur le CRM) $ $ installation de l'application abl90 $ $ relance $ $ chargement de toutes les données et de la configuration $ $ mise en route (1H) $ $ essais bons | SWO-515607 | Defect | HW | DMS_PC | CPU |  | The analyzer is slow Beug Disconnects from the user | Full verification of the analyzer | Backup of all data and configuration $$ Replacement of the $$ $Windows7 installation windows7 $windows configuration $$ configuration of the network backup (info not present on the CRM) $ installation of the application abl90 $raise $ $ $ loading all data and configuration $$ start (1H) $good trials
No | Visit | Complaint | RIND | When customer try to on instrument with power switch, instrument is not getting switched on. | Need to replace new signal data board | Change new signal data board and check all parameters, instrument working fine. Calibration also passed and sample results is also ok. | SWO-515473 | [5 empty] | When customer try to on instrument with power switch, instrument is not getting switched on. | Need to replace new signal data board | Change new signal data board and check all parameters, instrument working fine. Calibration also passed and sample results is also ok.
No | Visit | Complaint | RCHN | 机器堵管 | 机器堵管 | 冲洗管路后修复（by chuanwei） | SWO-515091 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway | Clot | Machine plugs | Machine plugs | Repair after flushing line (by by chuanwei)
No | Visit | Complaint | RCHN | 进样口断裂 | 进样口断裂 （未更换备件） | 更换进样口后修复（by chuanwei） | SWO-515842 | Defect | HW | Aspiration Module | Tube holder |  | The inlet is broken. | The inlet is broken (spare parts have not been replaced) | Repair after replacing the sample (by chuanwei)
No | Visit | Complaint | RCHN | 报错1257 | 报错1257 | 清洁HCT模块后修复（by chuanwei） | SWO-515840 | Cleaned | HW | Aspiration Module | Needle Wash Unit | 1257 | 1257 error | 1257 error | Repair after cleaning HCT module (by chuanwei)
No | Visit | Complaint | RCHN | 做样本报错1258.1223 | 做样本报错1258.1223 | 拆卸清理过滤网，拆卸清理HCT单元，做样本测试未出现报警，仪器修复 by Andrew | SWO-519001 | Cleaned | HW | Aspiration Module | Needle Wash Unit | 1258, 1223 | Do a sample of this newspaper error 1258.1223 | Do a sample of this newspaper error 1258.1223 | Remove the clean-up filter, remove the cleaning HCT unit, do sample test without alarm, instrument repair by Andrew
No | Visit | Complaint | RKK | 異常音 | 器械 動作時 異常音 | 排液と溶液 ポンプヘッド交換。 | SWO-515781 | [5 empty] | Abnormal tone | Instrument Operating Abnormal Sound | Drainage and solution pump head replacement.
No | Visit | Complaint | RPAC | Deb called 0260 517 688 creatinine membrane changed, qc failing | instructed customer to try another membrane change, called later all fine | instructed customer to try another membrane change, called later all fine | SWO-523385 | Assisted Replacement | Electrode | Membrane | Crea | CAPA-789 | Deb called 0260 517 688 creatinine membrane changed, qc failing | instructed customer to try another membrane change, called later all fine | instructed customer to try another membrane change, called later all fine
No | 48 | Y | RCZE | Blank |  |  | D5D30004D6553604 | Cleaned | HW | Autocheck module | Autocheck module |  | Blank
No | Visit | Complaint | RDEU | Das Gerät bootet nicht (AU läuft?) | Patientenlayout lässt sich nicht ändern (links keine Auswahlmöglichkeiten) | Gerät geprüft, Konfiguration geprüft, Software V6.18 neu installiert, Konfiguration gelöscht, Konfiguration neu eingegeben, Funktionskontrolle ok | SWO-515531 | Defect | SW | RMED SW | RMED SW | patient layout | The device is not booting (AU is running?) | Patient layout cannot be changed (no choices on the left) | Device checked, configuration checked, software V6.18 reinstalled, configuration deleted, configuration re-entered, function control ok
No | Visit | Complaint | RCHN | 液路故障 | ph垫圈故障 | 清洁垫圈后修复（by chuanwei） | SWO-515100 | Cleaned | Electrode | Electrode | Sealing ring, pH |  | Liquid road fault | ph washer fault | Post-cleaning repair (bywei chuan)
No | Visit | Complaint | RIND | Inlet Probe bend | checked and found that no sample detected during aspiration | Need to replace Inlet Probe | SWO-515931 | Defect | HW | Inlet module | Probe bend |  | Inlet Probe bend | checked and found that no sample detected during aspiration | Need to replace Inlet Probe
No | Visit | Complaint | RDAN | Stregkode læseren vil ikke læse den lille del af stregkoden på sprøjterne - med andre ord så skal sprøjten vendes med den store stregkode ind mod læseren før at denne vil læse. | Kunne ikke læse hver gang I service programmet læste den hvergang til venstre men ikke til højre | justeret stregkodelæseren nu læser den hver gang - check i både service program og i FLEXQ | SWO-515369 | Adjustment | HW | FlexQ Module | Barcode reader |  | The barcode reader will not read the small part of the barcode on the syringes-in other words, the syringe should be reversed with the large barcode against the reader before it will read. | Could not read every time In the service, it read each time to the left but not to the right | Adjusted barcode Reader Now it reads every time- Check in both service and FLEXQ
Yes | Visit | Complaint | RESP | Tras cambio de aguja da Qc bajos | Chequeo neumático, chequeo aguja | Sustitución de aguja por segunda vez, chequeo de manifiol y sustitución de bombas p2 y p4 stock cliente, limpieza de nipples , ajustes por servicio de aguja y otras partes, limpieza general, calibración de nuevo lote y carga de reactivos, pasados Qc | SWO-517295 | Re classified RMED | same day upSWO |  |  |  | After changing needle gives Qc low | Pneumatic check, needle check | Needle replacement a second time, manifiol check and replacement of p2 and p4 customer stock pumps, nipple cleaning, needle service and other adjustments, general cleaning, new batch calibration and reagent loading, past Qc
No | Visit | Complaint | RESP | Tras cambio de aguja da Qc bajos | Chequeo de neumática y partes del equipo | Sustitución de aguja, sustitución de Wash Block unit visiblemente afectada, realización de kit de tubos, chequeo y limpieza completa, calibración de lote nuevo. Realización de Qc sin éxito | SWO-517294 | Defect | HW | Aspiration Module | Needle Assembly | 588 | After changing needle gives Qc low | Pneumatic check and parts of the equipment | Needle replacement, wash block unit replacement visibly affected, tube kit, check and complete cleaning, new batch calibration. Performing Qc without success
No | Visit | Complaint | RDEU | Zur Wartungsmassnahme an einem anderen Analysator, Anwender wenden sich an mich mit der Fehlermeldung. | Start einer Wartungsmassnahme an einem anderen Analysator, Info über Defekt am Analysator. Wartungsmassnahme gestoppt. Überprüfung des Analysators. Display reagiert verzögert und deutlich verschoben. Durch Neustart keine Behebung der Situation, nur sehr kurzfristig. Analysator geht dann wieder in das bekannte Reaktionsverhalten über. | Versuch den Bildschirm neu zu Kalibrieren im Service, scheitert. Reaktion verzögert und Timeouts . Tausch des Bildschirms mit dem des funktionierenden Analysators. Es zeigt sich, dass der Bildschirm normal reagiert. Touch Screen defekt, bestellt. Wartungsmassnahme beim anderen Analysator wird gestoppt. Durchführung zu einem späteren Zeitpunkt. | SWO-515568 | Revisit required | HW | Analyzer Control Module | Display | wet cleaning | For maintenance on another analyzer, users contact me with the error message. | Start of a maintenance measure on another analyzer, info about defect on the analyzer. Maintenance measure stopped. Checking the analyzer. Display reacts delayed and significantly shifted. By restarting no fix of the situation, only very short-term. Analyzer then passes back into the known reaction behavior. | Attempting to recalibrate the screen in the service fails. Response delayed and timeouts . Swap the screen with that of the working analyzer. It turns out that the screen responds normally. Touch screen defective, ordered. Maintenance measure on the other analyzer is stopped. implementation at a later date.
No | Visit | Complaint | RLTD | Email received from Vanessa Rumens. | Checked activity log - error messages for aborted calibrations are: 4019 - Measure_23; 5375 - ReadOpticalROM_56. Analyser is intermittently losing communication with the optical unit ROM. Tested cable between Main pcb and optical unit - no fault found. Rebooted analyser and checked optical unit - now able to read optical unit ROM. Fault is with the optical unit and this needs to be replaced, part not available during this visit and will be ordered, | Optical unit is faulty and needs to be replaced. Replacement part ordered. | SWO-516109 | Revisit required | HW | Process Module | Optical Unit | 5375 | Email received from Vanessa Rumens. | Checked activity log - error messages for aborted calibrations are: 4019 - Measure_23; 5375 - ReadOpticalROM_56. Analyser is intermittently losing communication with the optical unit ROM. Tested cable between Main pcb and optical unit - no fault found. Rebooted analyser and checked optical unit - now able to read optical unit ROM. Fault is with the optical unit and this needs to be replaced, part not available during this visit and will be ordered, | Optical unit is faulty and needs to be replaced. Replacement part ordered.
No | Visit | Complaint | RDEU | Anwender berichten, dass der Scanner für die EDTA-Röhrchen nur sporadisch den Barcode erkennt. | Anwender berichten, dass der Scanner für die EDTA-Röhrchen nur sporadisch den Barcode erkennt. | Scanner im Servicemenü überprüft. Dabei wurde am Gerät kein Fehler festgestellt. Probenröhrchen der Qualitätskontrollen werden jedes Mal eingespannt. Bei Patienten-Probenröhrchen bringt das Gerät die Meldung Gescannter Strichcode hat den Strichcode des Zielfeldes überschrieben Dieser Fehler soll mit der nächsten Softwareversion behoben werden. Gerät überprüft. | SWO-515980 | Defect | SW | RMED SW | RMED SW | Barcode reader, inlet | Users report that the scanner for the EDTA tubes only recognizes the barcode sporadically. | Users report that the scanner for the EDTA tubes only recognizes the barcode sporadically. | scanner in the service menu. No error was detected on the device. Sample tubes of the quality controls are clamped every time. For patient sample tubes, the device brings the message Scanned bar code has overwritten the bar code of the target field This error should be fixed with the next software version. device checked.
No | Visit | Complaint | RSCH | Blank | MetHb. bei externer QC ungenügend (Zielwert 1.0% gemessen 2.3%). | Hemolyzer ersetzt, alles neu eingestellt und mit externer QC überprüft (gemessen 1.9%) Original Hemolyzer wieder eingebaut und Spectrometer ersetzt gemessen 1.7%). Nochmals Hemolyzer ersetzt und mit neuem Spectrometer gestartet. ABL macht eine Oximetrie Initialisierung (60min.). Kal., tHb.-Kal., QC und System i.o. Erneut mit externer QC gemessen (Resultat 0.9% bei Zielwert 1.0%) alles i.o. | SWO-516442 | Defect | HW | Oxi module | Oxi module | Hemolyzer and spectrometer | Blank | MetHb. insufficient for external QC (target value 1.0% measured 2.3%). | Hemolyzer replaced, all re-adjusted and checked with external QC (measured 1.9%) Original Hemolyzer reinstalled and Spectrometer replaced measured 1.7%). Replaced Again Hemolyzer and started with new Spectrometer. ABL makes an oximetry initialization (60min.). Cal., tHb.-Kal., QC and System i.o. Measured again with external QC (result 0.9% at target value 1.0%) everything i.o.
No | Visit | Complaint | RDEU | 25.01.2019, Herr Skog, Tel: 071127832585: Oxi Hardware Funktion | Oxihardwarefehler. | WSC-PCB erneuert. Software geladen. Wetsectioninstaller beendet Installation auch nach längerer Wartezeit nicht. Installation zweimal mit unterschiedlichen Medien wiederholt. Erfolglos. WSC zurückgebaut. Weiterhin Fehler 1010 1018 1020 1029 | SWO-516729 | Revisit required | HW | Oxi module | Spectrometer | 1010, 1018, 1020, 1029 | 25.01.2019, Mr. Skog, Tel: 071127832585: Oxi Hardware Function | Oxihardwarefehler. | WSC-PCB. Software loaded. Wetsectioninstaller does not stop installation even after a long wait. Repeated installation twice with different media. unsuccessful. WSC. Also error 1010 1018 1020 1029
No | Visit | Complaint | RKK | キャル時のKインピーダンスエラーが頻発。 | シミュレータテストでCO2がフェール | センタインタフェースPCB交換 | SWO-518834 | Defect | HW | Basic Unit | PCB, Sensor interface | K; pCO2 mail: two analyzers at same hospitals within a few days with same failure? K or pCO2? | K impedance error at the time of the car occurs frequently. | CO2 fails in simulator test | Center Interface PCB Replacement
Yes | 5 | Y | RCZE | Blank | V záznamoch chyba 1361 Prístroj si žiada výmenu SP - skrat v SP chyba 0964 Nehomogénny roztok v LS chyba 1324 |  | D5D30E0DD64F0B01 | [5 empty] | Blank | In records error 1361 device asks for Exchange SP-shorting in SP error 0964 non-homogeneous solution in LS error 1324
No | Visit | Complaint | RLTD | Suspected clot in sensor cassette | Returned to investigate why the analyser is aborting samples, when a full 70ul capillary was presented to the analyser it would apirate the whole sample in and abort with 1271 - Failed to aspirate sample . The sensor cassette had expired on arrival so a new one was installed and allowed to start up. Once the start up had completed 5 water samples were run to see if the fault could be replicated but they all went through ok. 6 blood samples were run and only 1 failed and that was due to errors obtaining the sample rather than a problem with the analyser. | No errors could be found with the analyser so it was handed back to the unit and are going to run it over the weekend and see how it goes. | SWO-516597 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway | Clot | Suspected clot in sensor cassette | Returned to investigate why the analyser is aborting samples, when a full 70ul capillary was presented to the analyser it would apirate the whole sample in and abort with 1271 - Failed to aspirate sample . The sensor cassette had expired on arrival so a new one was installed and allowed to start up. Once the start up had completed 5 water samples were run to see if the fault could be replicated but they all went through ok. 6 blood samples were run and only 1 failed and that was due to errors obtaining the sample rather than a problem with the analyser. | No errors could be found with the analyser so it was handed back to the unit and are going to run it over the weekend and see how it goes.
Yes | Visit | Complaint | RDEU | 15.03.2019 09:20Uhr Thorsten Haberkam 02083092162 2180: Der Kunde hat auf die Mailbox gesprochen und bittet um Rückruf. | Crea $Urea SC nicht lieferbar. Das Gerät muss ohne Crea Urea betrieben werden. | LP ohne Crea installiert. Kunde ist unzufrieden über die Art und Weise wie die Station einen Service kommunizieren muss. | SWO-516139 | Re classified RMED | Shipping |  |  | Crea/Urea SC not delivered | 15.03.2019 09:20pm Thorsten Haberkam 02083092162 2180: The customer has spoken to the mailbox and asks for a call back. | Crea /Urea SC not available. The device must be operated without Crea Urea. | LP installed without Crea. Customer is dissatisfied with the way the station needs to communicate a service.
No | Visit | Complaint | RCHN | 电池无法蓄电 | 电池无法蓄电先给医院使用备用机。送修维修间。 | 先给医院使用备用机。送修维修间。 | SWO-518371 | [5 empty] | The battery can't store power. | The battery cannot be stored to use a spare machine for the hospital first. Send the repair room. | Start by using a spare machine for the hospital. Send the repair room.
No | Visit | Complaint | RCHN | K,CA定标失败 | K,CA定标失败，借903-581一个 | 更换903-581后修复（by chunwei） | SWO-515071 | Defect | HW | Basic Unit | PCB, Sensor interface | K, Ca Cal | K, CA calibration failed | K, CA bid failed, borrow 903-581 one | Repair after replacement 903-581 (by chunwei)
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 无法抽洗溶液，报错1169 | 更换试剂包后修复 | SWO-518242 | [5 empty] | Test card, reagent package failure | Unable to pump the solution, report 1169 | Repair after replacement of reagent kit
No | Visit | Complaint | RCHN | 传感器故障，报错2167 | 传感器故障，报错2167，借636-539一个 | 更换636-539后修复（by chuanwei） | SWO-514362 | Defect | HW | Storage Module | sensor, cup | 2167 | Sensor fault, error 2167 | Sensor fault, error 2167, borrow636-539 one | Repair after replacement 636-539 (by chuanwei)
Yes | Visit | Complaint | RCHN | 液路堵塞，FCT_103 | 液路堵塞，FCT_103，借903-112一个 | 更换液路模块后修复（by chuanwei） | SWO-514306 | Re classified RMED | same day upSWO |  |  |  | Liquid Road Blocked, FCT-103 | Liquid Road Blocked, FCT-103, loan 903-112 one | Repair after replacing the liquid circuit module (by chuanwei)
No | Visit | Complaint | RCHN | 定标失败 | 定标失败，借944-258（15582）、944-214（15452）各一个 | 更换液路模块，重新定标后修复(by chuanwei) | SWO-515306 | Defect | HW | Wet section | PCB, wet section | 3412 | Bid failed | Failed to bid, borrow 944-258 (15582), 944-214 (15452) each | Replace the liquid circuit module and repair it after re-calibration (by chuanwei)
No | Visit | Complaint | RDEU | 04.01.2019 Herr Bebernik 071717011332 meldet, das sein ABL9 gesperrt ist. 04.01.2019 Frau Kaltenhofen 071717011887 gelingen die QCs nicht. 07.01.2019 Herr Bebernik 071717011332 Fehler bleibt bestehen. | ABL9 besteht keinen QC. Ständige QC Fehler in pCO2 und O2. | ABL9 wurde vom Kunden abgezogen und in VSC-SW überprüft. Verschlauchung des ABl9 geprüft, - keine Auffälligkeiten. Neue SC und LP installiert, Kal.. o.k. QC nach mehreren Versuchen immer Fehlerhaft. CarHi über den zustand des Gerätes informiert. | SWO-516759 | [5 empty] | 04.01.2019 Mr. Bebernik 071717011332 reports that his ABL9 is blocked. 04.01.2019 Ms. Kaltenhofen 071717011887 the QCs do not succeed. 07.01.2019 Mr. Bebernik 071717011332 Error remains. | ABL9 does not have a QC. Permanent QC errors in pCO2 and O2. | ABL9 was withdrawn from the customer and checked in VSC-SW. Hoseing of the ABl9 tested, - no abnormalities. New SC and LP installed, Kal.. o.k. QC after several attempts always faulty. CarHi informed about the condition of the device.
No | Visit | Complaint | RDEU | 21.12.2018; Fr. Schmidt; Tel.: 071717011888; pO2 QC und Messfehler 27.12.2018; Fr. Hirsch; Tel.: 071717011888; immer noch pCO2 und pO2 QC Fehler L3 03.01.2019, 09:02 Uhr, Fr. Kaltenhofer, Tel. 071717011887: QC Fehler bei Level 2 und 3 bei pCO2 und pO2 04.01.2019; Fr. Kaldenhofer; Tel.: 071717011887; weiterhin QC Fehler L2 und L3 | Fehlerhafte QC Messungen in pCO2 und pO2 | ABL9 wurde vom Kunden abgezogen und in VSC-SW überprüft. Verschlauchung des ABl9 geprüft, - keine Auffälligkeiten. Neue SC und LP installiert, Kal.. o.k. QC nach mehreren Versuchen immer Fehlerhaft. CarHi über den zustand des Gerätes informiert. | SWO-516760 | [5 empty] | 21.12.2018; Fr. Schmidt; Tel.: 071717011888; pO2 QC und Messfehler 27.12.2018; Fr. Hirsch; Tel.: 071717011888; immer noch pCO2 und pO2 QC Fehler L3 03.01.2019, 09:02 Uhr, Fr. Kaltenhofer, Tel. 071717011887: QC Fehler bei Level 2 und 3 bei pCO2 und pO2 04.01.2019; Fr. Kaldenhofer; Tel.: 071717011887; weiterhin QC Fehler L2 und L3 | Faulty QC measurements in pCO2 and pO2 | ABL9 was withdrawn from the customer and checked in VSC-SW. Hoseing of the ABl9 tested, - no abnormalities. New SC and LP installed, Kal.. o.k. QC after several attempts always faulty. CarHi informed about the condition of the device.
No | Visit | Complaint | RAME | Customer states that although the QC results appear correct on the analyzer, when they cross over to the customers LIS, the QC results logged in the LIS do not match the analyzers QC results. | Verified the customer's stated issue. Accessed the patient layout page and removed all the temperature corrected values to ensure that the LIS isn't using those QC values. After removing the temperature corrected values and sending the results to the LIS, it was discovered that the results on the customers LIS were still not matching the analyzers QC values. | Set up a trace on the analyzer to run over the weekend and will revisit the customer and will hand over the trace files to the costumer so that he may forward them to the appropriate people to get this issue resolved. | SWO-516404 | Revisit required | SW | External SW | His/Lis |  | Customer states that although the QC results appear correct on the analyzer, when they cross over to the customers LIS, the QC results logged in the LIS do not match the analyzers QC results. | Verified the customer's stated issue. Accessed the patient layout page and removed all the temperature corrected values to ensure that the LIS isn't using those QC values. After removing the temperature corrected values and sending the results to the LIS, it was discovered that the results on the customers LIS were still not matching the analyzers QC values. | Set up a trace on the analyzer to run over the weekend and will revisit the customer and will hand over the trace files to the costumer so that he may forward them to the appropriate people to get this issue resolved.
Yes | Visit | Complaint | RIND | Demo option will expire soon. |  | installed the option keys. Did calibration and samples all results found within the range .found instrument working satisfactorily. | SWO-516187 | Re classified RMED | planned SWO |  |  |  | Demo option will expire soon. |  | installed the option keys. Did calibration and samples all results found within the range .found instrument working satisfactorily.
No | Visit | Complaint | RDEU | QC Level 3 Bilirubin außer Bereich | Die letzten 4 Tage zum Teil erhöhte Driften tHb. QC Bilirubin zu hoch im Level 3. Lichtwert Oxi bei 2900. | Proteinentfernung, Kal und Service-tHb-Kal durchgeführt. Weiterhin Bili unverändert. Lampe neu justiert und erforderliche Kalibrierungen durchgeführt. Neuer Lichtwert 3400. Weiterhin Bilirubin außer Bereich. Hämolysator aus Ersatzmodul eingesetzt, justiert und kalibriert. | SWO-516730 | Defect | HW | Oxi module | Hemolyzer | QC L3 Bili | QC Level 3 Bilirubin out of range | The last 4 days partly increased drifts tHb. QC Bilirubin too high in level 3. Light value Oxi at 2900. | protein removal, kal and service tHb-Kal. Bili remains unchanged. lamp is re-adjusted and required calibrations are carried out. New light value 3400. Also bilirubin out of range. Hemolyser from replacement module, adjusted and calibrated.
No | Visit | Complaint | RDEU | Das Ergebnisprotokoll ist unleserlich, auch Zahlenwerte | No | Druckwerk erneuert | SWO-516763 | Revisit required | HW | DMS_PC | Printer Module |  | The result log is illegible, even numerical values | No | Printing unit renewed
Yes | Visit | Complaint | RFRA | CRP StDev non acceptable selon qualité demandé par le laboratoire | A la demande de Paris, vérification position aiguille. | Contrôle position aiguille dans cup et wash unit: OK. Une réunion prévue cet après-midi, Radiometer prendra contact avec le laboratoire. | SWO-516065 | Re classified RMED | planned SWO |  |  | Techline: RFRA 836351484 | CRP StDev not acceptable according to quality requested by the laboratory | At the request of Paris, check needle position. | Control needle position in cup and wash unit: OK. A meeting scheduled for this afternoon, Radiometer will contact the laboratory.
No | Visit | Complaint | RDEU | Zur Wartungsmassnahme an einem anderen Analysator, Anwender wenden sich an mich mit der Fehlermeldung. | Reaktion Bildschirm nicht verändert, weiterhin verzögert und versetzt. | Tausch des Touch Screen im Bildschirm. Nach Einbau und Neustart Funktion wie gehabt ohne Einschränkungen. Bildschirm nicht nass wischen, da sonst defekt, Info an Anwender gegeben. Defekt vermutlich durch zu feuchte Reinigung. | SWO-516101 | User Error | HW | Analyzer Control Module | Display | wet cleaning | For maintenance on another analyzer, users contact me with the error message. | Reaction screen not changed, continues to be delayed and offset. | Swap the touch screen on the screen. After installation and restart function as usual without restrictions. Do not wipe the screen wet, otherwise defective, info given to users. Defect probably due to too moist cleaning.
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | LAC定标失败 | 更换测试卡 | SWO-516652 | [5 empty] | Test card, reagent package failure | LAC calibration failure | Replace the test card
No | Visit | Complaint | RCHN | 机器无法开机，屏幕不亮 | 机器无法开机，屏幕不亮,需要借件910-285一个 | 更换910-285后修复（by chuanwei） | SWO-516425 | Defect | HW | Power Supply | Power supply |  | The machine does not power on, the screen does not light | The machine can not be turned on, the screen is not lit, need to borrow 910-285 one | Repair after replacement 910-285 (by chuanwei)
No | Visit | Complaint | RCHN | 进样口支架堵塞 | 进样口支架堵塞 | 更换进样口支架，冲洗管路后修复 by Andrew | SWO-516837 | Defect | HW | Inlet module | Inlet gasket | Clot | Blocked in the inlet bracket | Blocked in the inlet bracket | Replace the inlet bracket and repair by Andrew after flushing the line
No | Visit | Complaint | RCHN | 光学灵敏度超范围，报错1029 | 定标失败，报错1029.。需要借件 944-212（15177）、15529（缺货） 944-214（15229）*2、（缺货）944-216（15107）、944-258（15185）缺货 |  | SWO-516572 | Defect | HW | Process Module | Optical Unit | 1029 | Optical sensitivity over range, error 1029 | The bid failed, the error 1029. Loans 944-212 (15177), 15529 (out of stock) 944-214 (15229) x2, (out of stock) 944-216 (15107), 944-258 (15185) out of stock
No | Visit | Complaint | RCHN | CF卡损坏 | CF卡损坏 | 重装系统，自行更换910-373后修复 by Andrew | SWO-516720 | Defect | HW | Analyzer Control Module | CF card |  | CF card damage | CF card damage | Reinstall the system to repair by Andrew after replacing itself with 910-373
No | Visit | Complaint | RCHN | 打印机无法打印 | 打印机无法打印，需要借件910-275一个 |  | SWO-516577 | Defect | HW | DMS_PC | Printer Module |  | The printer can't print | The printer cannot print, requires a loan of 910-275
No | Visit | Complaint | RIND | Lac shows yellow | Check system message and calibration result in which Lac shows yellow and it's sensitivity is below the range while it was installed 3 days back. | Check chamber which is clean but its sensitivity was below the range very early,clean the chambers,replace Lactate membrane run cleaning and 2-point calibration which is OK. and problem has been resolved. | SWO-516614 | Assisted Replacement | Electrode | Membrane | Lac | 378 | Lac shows yellow | Check system message and calibration result in which Lac shows yellow and it's sensitivity is below the range while it was installed 3 days back. | Check chamber which is clean but its sensitivity was below the range very early,clean the chambers,replace Lactate membrane run cleaning and 2-point calibration which is OK. and problem has been resolved.
No | Visit | Complaint | RAME | SC won’t stabilize | SC won’t stabilize. NA out on Calibration. Sensitivity too high. Other parameters out due to dependencies. Ran impedance test in service. All good. Ran SC Simulator test. All good. During startup after removing SC Simulator, received Flow Selector error and analyzer wanted SP replaced. | Put SC back in analyzer after SC Simulator test. Cleaned flow Selector Sensor. Was dirty. Calibrations and QC fine now. Also ran tHb Calibration. Both analyzers had a reminder. | SWO-517185 | Cleaned | HW | Basic Unit | Flow selector | Na sens too high | SC won’t stabilize | SC won’t stabilize. NA out on Calibration. Sensitivity too high. Other parameters out due to dependencies. Ran impedance test in service. All good. Ran SC Simulator test. All good. During startup after removing SC Simulator, received Flow Selector error and analyzer wanted SP replaced. | Put SC back in analyzer after SC Simulator test. Cleaned flow Selector Sensor. Was dirty. Calibrations and QC fine now. Also ran tHb Calibration. Both analyzers had a reminder.
No | Visit | Complaint | RDEU | 25.01.2019, Herr Skog, Tel: 071127832585: Oxi Hardware Funktion | Oxifehler 1010 1018 1020 1029 | Spectrometer Unit erneuert. Kalibrierungen und tHb-Kalibrierung durchgeführt. | SWO-517158 | Defect | HW | Oxi module | Spectrometer | 1010, 1018, 1020, 1029 | 25.01.2019, Mr. Skog, Tel: 071127832585: Oxi Hardware Function | Oxifehler 1010 1018 1020 1029 | Spectrometer Unit. calibrations and tHb calibration.
No | Visit | Complaint | RMAY | Oncall breakdown | Found blockage in hemolyzer | -Check analyser data log -Perform manual flush with hypochlorite ( Failed) -Dismantle wet section module -Waiting sparepart | SWO-520515 | Revisit required | HW | Oxi module | Hemolyzer | Clot; 1352 | Oncall breakdown | Found blockage in hemolyzer | -Check analyser data log -Perform manual flush with hypochlorite ( Failed) -Dismantle wet section module -Waiting sparepart
Yes | Visit | Complaint | RAME | Customer has some difficulty to retrieve the date | Customer has difficulty to search the patient data among the archived data. Customer wants to have ability to monitor the error codes on the analyzer. | Showing to customer how to retrieve the archive data to excel file. Talking to customer and they wanted to know if they use the a cure software to be used as a data management tools to monitor data and patient results including the error codes. Customer wanted to know if the hemolysis happening on the capillary sample can cause the sensor response error. Will be returned to customer with answers to the questions shortly. | SWO-517321 | Re classified RMED | Request |  |  | See e-mail with title "RE: SWO-517321" from JGIBB on 2019-03-25 15:15 | Customer has some difficulty to retrieve the date | Customer has difficulty to search the patient data among the archived data. Customer wants to have ability to monitor the error codes on the analyzer. | Showing to customer how to retrieve the archive data to excel file. Talking to customer and they wanted to know if they use the a cure software to be used as a data management tools to monitor data and patient results including the error codes. Customer wanted to know if the hemolysis happening on the capillary sample can cause the sensor response error. Will be returned to customer with answers to the questions shortly.
No | Visit | Complaint | RDEU | 15.03.2019 08:50Uhr H. Milczarek +4922029380: Ständig ist ein defekt am Gerät und Zubehör muss getauscht werden | PCO2 Fehler | Backup erstellt. Lp gewechselt. Fehler weiterhin in Level A. SW Update auf Version 3.4 MR1. Startphase. Kal ok. QCs nicht ok. SC getauscht. Startphase. PH & Metabolit Fehler. Weitere SC eingesetzt. Kal Backlog NA, Glu & Lac außer Grenzen. Spülung + weitere Kal. Mehrfache Kal durchgeführt, wegen Glu & Lac Fehler. Kal & QC. Ok. | SWO-517014 | Assisted Replacement | SC | SC | pCO2 |  | 15.03.2019 08:50H H. Milczarek +4922029380: Constantly a defective device and accessories must be replaced | PCO2 Error | Backup created. Lp. Errors continue in Level A. SW Update to version 3.4 MR1. Start. Kal ok. QCs not ok. SC. Start. PH & metabolite error. More SC used. Kal Backlog NA, Glu & Lac out of bounds. Rinse + more Kal. Multiple Kal performed, due to Glu & Lac error. Kal & QC. Ok.
No | Visit | Complaint | RLTD | Sample pathway blocked, unable to clear with flush device. | On arrival tried flushing and was able to clear blockage in sample pathway. Refitted original inlet probe after correcting bend in it. Restarted analyser - passed fluidic tests. Customer to verify that analyser passes calibration and quality controls after full restart has completed. | Cleared blockage from sample pathway. Corrected bent inlet probe. | SWO-517313 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway | Clot | Sample pathway blocked, unable to clear with flush device. | On arrival tried flushing and was able to clear blockage in sample pathway. Refitted original inlet probe after correcting bend in it. Restarted analyser - passed fluidic tests. Customer to verify that analyser passes calibration and quality controls after full restart has completed. | Cleared blockage from sample pathway. Corrected bent inlet probe.
No | Visit | Complaint | RLTD | The analyser is giving a Clot Suspected in Inlet error every night. | Replacing the inlet module and inlet connector gasket have not resolved the problem of repeated clot detected in the inlet. The liquid sensor values for the Blood Gas Liquid Sensor and the Oxi Liquid Sensor were higher that normally seen but still within the limits. | The Wet Section Control PCB was replaced the see if this would bring the liquid sensor values back down to what would normally be seen. The liquid sensor values didn't change but I ran the Clot Detect $ Remove service program 5 times with no clots detected. The analyser will be monitored to see if the problems re-occurs. | SWO-517121 | Revisit required | HW | Basic Unit | Sensor Bay Unit | Clot detection, midnight | The analyser is giving a Clot Suspected in Inlet error every night. | Replacing the inlet module and inlet connector gasket have not resolved the problem of repeated clot detected in the inlet. The liquid sensor values for the Blood Gas Liquid Sensor and the Oxi Liquid Sensor were higher that normally seen but still within the limits. | The Wet Section Control PCB was replaced the see if this would bring the liquid sensor values back down to what would normally be seen. The liquid sensor values didn't change but I ran the Clot Detect $ Remove service program 5 times with no clots detected. The analyser will be monitored to see if the problems re-occurs.
No | Visit | Complaint | RAME | Analyzer is failing precision study validation for pO2 on level 3 QC | PO2 is not very stable during Zeroing of Sensor board | Replaced manifold unit, sensor connector board and all parts in Annual Maintenance kit. Afterwards performed Zero Check and calibrated the sensor board. Ran VK-R5 linearity kit. Afterwards performed pO2 precision on Level 3 QC for customer. | SWO-517236 | [5 empty] | Analyzer is failing precision study validation for pO2 on level 3 QC | PO2 is not very stable during Zeroing of Sensor board | Replaced manifold unit, sensor connector board and all parts in Annual Maintenance kit. Afterwards performed Zero Check and calibrated the sensor board. Ran VK-R5 linearity kit. Afterwards performed pO2 precision on Level 3 QC for customer.
No | Visit | Complaint | RCHN | NICU的老师对于GLU不出结果有疑问 | NICU的老师对于GLU不出结果有疑问 | 与老师解释情况后解决by Hunter | SWO-517411 | User Error | SC | SC | Glu | SC Replacement | NICU's teachers have questions about GLU's results | NICU's teachers have questions about GLU's results | Solve by Hunter after explaining the situation with the teacher
No | Visit | Complaint | RCHN | 储存盘卡盘，马达无法使用，报错5784 | 储存盘卡盘，马达无法使用，报错5784。需要借件580-098 | 更换580-098后修复 by Andrew | SWO-516579 | Defect | HW | Storage Module | Motor, storage wheel | 5784 | Storage disk card, motor can not be used, reported error 5784 | Storage disk card, the motor can not be used, reported error 5784. 580-098 required | Repair by Andrew after replacement 580-098
No | Visit | Complaint | RCHN | 读取不了处理包芯片 | 读取不了处理包芯片,需要借件903-036一个 | 更换903-036后修复 | SWO-516566 | Defect | HW | Cover and Chassis | PCB, solution pack |  | Can't read the processing package chip | Can not read the processing package chip, need to borrow 903-036 one | Repair after replacement 903-036
No | Visit | Complaint | RCHN | 报错5784，处理盘里的夹子松动 | 报错5784，处理盘里的夹子松动（未更换备件） | 调整处理盘里架子位置后修复by Hunter | SWO-517538 | Adjustment | HW | Storage Module | Cartridge guide | 5784 | 5784, the clip in the processing disc is loose | Error 5784, clip loose in the processing tray (spare parts not replaced) | Fix by Hunter after adjusting the shelf position in the processing tray
No | Visit | Complaint | RCHN | CO2含量有问题 | CO2含量有问题 | 重新设置单位，参考范围后解决by Hunter | SWO-517482 | Assisted Replacement | Electrode | Electrode | pCO2 |  | There is a problem with co2 content | There is a problem with co2 content | Reset the unit, reference the scope and resolve by Hunter
No | Visit | Complaint | RIND | PCO2 sensitivity out of range. | PCO2 sensitivity out of range | Check the PCO2 electrode chamber & found clot in it,also found clot in Reference electrode chamber,clean the clot,run the 2point calibration,PCO2 got passed but Po2 sensitivity was out of range,check the membrane & found it was emptied,bring the PO2 membrane,replace it & run the calibration & got passed successfully. | SWO-520259 | Cleaned | Fluidic pathway | pH/BG Module | measuring chamber | Clot; pCO2 Sens | PCO2 sensitivity out of range. | PCO2 sensitivity out of range | Check the PCO2 electrode chamber & found clot in it,also found clot in Reference electrode chamber,clean the clot,run the 2point calibration,PCO2 got passed but Po2 sensitivity was out of range,check the membrane & found it was emptied,bring the PO2 membrane, replace it & run the calibration & got passed successfully.
No | Visit | Complaint | RIND | Inlet gasket leakage and print not clear | Inlet gasket leakage and print not clear | Cleaned the instrument , checked the inlet gasket , run protein removal , still leakage not stopped , change the waste pump tubing , opened the inlet gasket and checked for clots , found clots in drain tubing , cleaned the tubing removed clots , run protein removal , checked the print , print is not proper , cleaned printer head and checked but still the print is not proper , changed the printer head , run calibration and some samples , instrument working fine , customer happy with instrument , | SWO-518331 | Defect | HW | Printer Module | Printer Head |  | Inlet gasket leakage and print not clear | Inlet gasket leakage and print not clear | Cleaned the instrument , checked the inlet gasket , run protein removal , still leakage not stopped , change the waste pump tubing , opened the inlet gasket and checked for clots , found clots in drain tubing , cleaned the tubing removed clots , run protein removal , checke d the print , print is not proper , cleaned printer head and checked but still the print is not proper , changed the printer head , run calibration and some samples , instrument working fine , customer happy with instrument ,
Yes | Visit | Complaint | RDEU | 331 Keine Probe während Ansaugung erkannt bei QC | AQC wird sporadisch abgebrochen | Ventile ünerprüft. Einlassdichtung ausgetauscht. AQC total verschmutzt. >>> gereinigt. Einlassnadel gerade. Pumpenkal und LS Justage. | SWO-516067 | [5 empty] | 331 No sample detected during suction at QC | AQC is aborted sporadically | valves. inlet seal. AQC totally polluted. >>> cleaned. Inlet pin straight. Pump scale and LS adjustment.
No | Visit | Complaint | RDEU | Hotline Einsatz Demokeys abgelaufen/ Parameter nicht installiert | Demo-Keys abgelaufen, Gerät außer Betrieb | OriginalKeys installiert, Startphase, korrekte Kal abgewartet, QK ok | SWO-516941 | User Error, FSE | Demo key | Demo key | Demo key |  | Hotline Use Demokeys expired/ Parameters not installed | Demo keys expired, device out of service | OriginalKeys installed, start phase, correct kal wait, QK ok
No | Visit | Complaint | RIND | Replacement of hemolyzer unit | Replacement of hemolyzer unit | Replaced the hemolyzer unit into the thb $so2 module and refixed the thb $so2 module to the machine and calibrated the analyser to ready mode | SWO-524035 | Defect | HW | Oxi module | Hemolyzer |  | Replacement of hemolyzer unit | Replacement of hemolyzer unit | Replaced the hemolyzer unit into the thb $so2 module and refixed the thb $so2 module to the machine and calibrated the analyser to ready mode
No | Visit | Complaint | RDEU | QC Level 3 Bilirubin außer Bereich | Hämolysator defekt. | Oximodul komplett erneuert. tHb-Kalibrierung durchgeführt. | SWO-517159 | Defect | HW | Oxi module | Oxi module | 0089 | QC Level 3 Bilirubin out of range | Hemolyser defective. | Oximodul completely renewed. tHb calibration.
No | Visit | Complaint | RIND | pH Measurement Unstable. | pH Measurement Unstable. | Cleaned the electrode with electrode cleaning solution but problem is remaining same, customer is unhappy with results, we are fours to do change of new pH electrode, calibrated found sample results are OK. It’s under observation. Defective electrode Lot: BS-03. | SWO-520263 | Assisted Replacement | Electrode | Electrode | pH |  | pH Measurement Unstable. | pH Measurement Unstable. | Cleaned the electrode with electrode cleaning solution but problem is remaining same, customer is unhappy with results, we are fours to do change of new pH electrode, calibrated found sample results are OK. It’s under observation. Defective electrode Lot: BS-03.
No | Visit | Complaint | RSCH | 0663: Leckage | 0663: Leckage | Allgemeine Kontrolle ok. Diverse Schläuche ers, Dichtungen usw. Diverse Membranen ers Ventile und Messkammer überprüft ohne Erfolg. Gerät mitgenommen und bei RSCH überprüft und rep. Komplettes Einlassmodul ers i.o. Friction detector und Einlassnadel beim Kundenmodul ers. Inlet gasket und Inlet Tubing Nippel ers. Ok. System i.o. Kal. + QC i.o. Gerät wird durch ROBRO zurückgebracht. | SWO-517191 | Defect | HW | Inlet module | Inlet module | 0663 | 0663: Leakage | 0663: Leakage | General control ok. Various hoses, seals, etc. Various membranes ers Valves and measuring chamber checked without success. device and checked with RSCH and rep. Complete inlet module ers i.o. Friction detector and inlet needle at the customer module. Inlet gasket and Inlet Tubing nipple ers. Ok. System i.o. Kal. + QC i.o. Device is returned by ROBRO.
Yes | Visit | Complaint | RPAC | Visit site and perform the Actions as per the Section 4 of the Data Collection Guide for Discrepant Results | As per Clinical Applications Specialist Request | Perform all the DATA collection, actions and Maintenace as per Section 4 of the Data Collection for Discrepant Results Procedure. | SWO-516734 | Re classified RMED | planned SWO |  |  | Techline: RPAC 837535748 | Visit site and perform the Actions as per the Section 4 of the Data Collection Guide for Discrepant Results | As per Clinical Applications Specialist Request | Perform all the DATA collection, actions and Maintenace as per Section 4 of the Data Collection for Discrepant Results Procedure.
No | Visit | Complaint | RDEU | Im Umfang der Zwischenwartung wurde festgestellt, das der AQT90 sporadisch den Fehler 2132: AUCONNECTIONTIMEOUT signalisiert. | Kabel zwischen CPU Unit und Display defekt. | Kabel zwischen CPU Unit und Display erneuert. | SWO-517526 | Defect | HW | Analyzer Control Module | Cable, display | 2132 | In the scope of intermediate maintenance, it was determined that the AQT90 sporadically signals error 2132: AUCONNECTIONTIMEOUT. | Cable between CPU unit and display defective. | Cable between CPU unit and display is renewed.
No | Visit | Complaint | RDEU | 18.03.2019 13:50Uhr Herr Zimmermann tel:+496976016010: Das Gerät meldet seit einer Woche pCO2-Fehler | Neue SW 3.4 MR1 installieren | Software installiert, Level A ist bei 29.7. Kunden gebeten das Gerät zu beobachten. | SWO-517852 | Defect | SW | RMED SW | pCO2 |  | 18.03.2019 13:50Mr Zimmermann tel:+496976016010: The device has been reporting pCO2 errors for a week | Installing new SW 3.4 MR1 | Software installed, Level A is at 29.7. Customers asked to observe the device.
No | Visit | Complaint | RDEU | 15.03.2019 08:20Uhr Frau Hägerisch 091318542993: pCO2 Fehler bei QC 18.03.2019 11:30Uhr Frau Nicole Tiliouine Station Medizin 4 D03 Tel: 09131/8542994 oder 09131/8542993 ( Fr. Hägerich) | Häufige pCO2 QC Verletzung in Level A und Level C pCO2 zu hoch in Level A und Level C - pCO2 steigt stetig an von QC zu QC Kunde hat innerhalb weniger Tage mehrfach die SC90 getauscht | ABL im Service überprüft, Einlassmodul überprüft, Einlasssonde ist gerade, Einlassmodul öffnet und schließt ohne Probleme, Quetschventil schaltet sauber nach Ansteuerung, ABL zerlegt, Schlauchsystem erneuert, Einlassdichtung und Anschlussdichtung erneuert, ABL komplettiert, Startphase ist gestartet, SC90 erwärmte sich selbst sehr langsam - kurz die Meldung: geeignete Umgebungstemp. schaffen - Temp. im Substrat außer Bereich, Startphase von 10 Minuten startete nach Erwärmung automatisch, Startphase läuft sauber runter, kalibriert, manuell alle 3 Level der QC wiederholt, ServiceDump gezogen, Blutmessung in Ordnung, System- und Funktionskontrolle in Ordnung | SWO-517859 | Assisted Replacement | SC | SC | pCO2 | QC L A and C high | 15.03.2019 08:20Uhr Frau Hägerisch 091318542993: pCO2 Fehler bei QC 18.03.2019 11:30Uhr Frau Nicole Tiliouine Station Medizin 4 D03 Tel: 09131/8542994 oder 09131/8542993 ( Fr. Hägerich) | Common pCO2 QC violation in Level A and Level C pCO2 too high in level A and level C - pCO2 is steadily increasing from QC to QC Customer has exchanged the SC90 several times within a few days | ABL checked in service, inlet module checked, inlet probe is straight, inlet module opens and closes without problems, squeeze valve switches clean after control, ABL disassembled, hose system renewed, inlet seal and connection seal renewed, ABL Complete, start phase is started, SC90 warmed itself very slowly - short the message: suitable ambient temp. create - Temp. in the substrate out of range, start phase of 10 minutes started automatically after heating, start phase runs clean down, calibrated, manually repeated all 3 levels of QC, ServiceDump pulled, blood measurement in order, system and function control in order
No | Visit | Complaint | RLTD | Sample pathway blocked, customer unable to clear with flush device - Cleared blockage with additional flush, corrected bent inlet probe. | Some oximetry parameters failing QC Level B. Cleaned hemolyser unit with hypochlorite solution. Performed tHb calibration. Ran QC Level B - all parameters now passed and are mid-range. Analyser reasdy for use. | Cleaned hemolyser. Ran tHb calibration. | SWO-518535 | Adjustment | HW | Oxi module | tHb Cal | QC B | Sample pathway blocked, customer unable to clear with flush device - Cleared blockage with additional flush, corrected bent inlet probe. | Some oximetry parameters failing QC Level B. Cleaned hemolyser unit with hypochlorite solution. Performed tHb calibration. Ran QC Level B - all parameters now passed and are mid-range. Analyser reasdy for use. | Cleaned hemolyser. Ran tHb calibration.
No | Visit | Complaint | RPAC | Machine is not passing its clot detect in the fluid transport menu. | Confirmed customers complaint. Customer has replaced the sensor cassette, and solution pack as part of troubleshooting. Checked fluid flow by running water through the system. Everything seems to be flowing well without restriction. Flushed out the fluid path a few times. Replaced inlet connector gasket and retested without success. | Checked inlet probe and found that it was bent. Replaced inlet probe and retested. Clot detect is now passing. Allowed unit to reinitialise and then performed calibrations and QCs. Everything passes and all is well with the machine. | SWO-517432 | Defect | HW | Auto Inlet module | Probe bend | Clot | Machine is not passing its clot detect in the fluid transport menu. | Confirmed customers complaint. Customer has replaced the sensor cassette, and solution pack as part of troubleshooting. Checked fluid flow by running water through the system. Everything seems to be flowing well without restriction. Flushed out the fluid path a few times. Replaced inlet connector gasket and retested without success. | Checked inlet probe and found that it was bent. Replaced inlet probe and retested. Clot detect is now passing. Allowed unit to reinitialise and then performed calibrations and QCs. Everything passes and all is well with the machine.
No | Visit | Complaint | RFRA | Reste en erreur sonde échantillon ouverte suite à une erreur de manipulation : La technicienne a voulu remettre en place le buvard en sortant le plateau durant la calibration. Le remplacement de la sonde est requise. | sonde hs | remplacemement sond calibration passage cq 5 fois | SWO-517554 | [5 empty] | Remains in error open sample probe due to a handling error: The technician wanted to put the blotter back in place by taking out the tray during calibration. Replacement of the probe is required. | hs probe | replacing probe calibration passage cq 5 times
No | Visit | Complaint | RCHN | 电池无法蓄电 | 电池无法蓄电，借431-018一个 | 更换431-018后修复 by Andrew | SWO-516896 | [5 empty] | The battery can't store power. | Battery can not store power, borrow 431-018 one | Repair by Andrew after replacement 431-018
No | Visit | Complaint | RCHN | 卡盘（断裂部位：nose） | 卡盘（断裂部位：nose） | 更换测试卡后修复 | SWO-518864 | Defect | Consumables | Test Kit | Cartridge, Broken | Nose | Card (fractured site: nose) | Card (fractured site: nose) | Repair after replacing test card
No | Visit | Complaint | RCHN | 电解质结果与临床不符和无法找到接口板，报错9011 | 电解质结果与临床不符和无法找到接口板，报错9011 未更换备件 | 零点校准和重装软件，需要观察 by Andrew | SWO-518955 | [5 empty] | Electrolyte results are not clinically consistent and no interface board can be found, error 9011 | Electrolyte results are not clinically consistent and no interface board can be found, error 9011 spare parts are not replaced | Zero-point calibration and reloading software, you need to observe by Andrew
No | Visit | Complaint | RAME | Customer states that although the QC results appear correct on the analyzer, when they cross over to the customers LIS, the QC results logged in the LIS do not match the analyzers QC results. | Set up a trace on the analyzer to run over the weekend and will revisit the customer and will hand over the trace files to the costumer so that he may forward them to the appropriate people to get this issue resolved. | Revisited customer and gave him the trace logs on a USB so that he may forward them to his LIS team. | SWO-518039 | No fault found | SW | External SW | His/Lis |  | Customer states that although the QC results appear correct on the analyzer, when they cross over to the customers LIS, the QC results logged in the LIS do not match the analyzers QC results. | Set up a trace on the analyzer to run over the weekend and will revisit the customer and will hand over the trace files to the costumer so that he may forward them to the appropriate people to get this issue resolved. | Revisited customer and gave him the trace logs on a USB so that he may forward them to his LIS team.
No | Visit | Complaint | RDEU | Pumpenkalibrierung error | Pumpenkalibrierung Fehler, leackage, Spülfehler, fehler1015 Volume cal. Failed, LS justagefehler, | Jahreswartung durchgeführt, membrandichtung Glu neuaufgesetzt, mehrmals Kalibriert und QC s durchgeführt SO2 Level 2 :95,4 (96,2-98,2) durchgefallen | SWO-519114 | Defect | Electrode | Electrode | Sealing ring, Glu | Pump cal error | Pumpenkalibrierung error | Pumpenkalibrierung Fehler, leackage, Spülfehler, fehler1015 Volume cal. Failed, LS justagefehler, | Annual maintenance performed, membrane seal Glu re-applied, several times calibrated and QC s performed SO2 Level 2:95.4 (96.2-98.2) failed
No | Visit | Complaint | RDEU | PCO2 Kal und QC Probleme, Parameter bei Fehlern sperren einstellen | Archiv überprüft, PCO2 schwankt sehr stark seit Wochen | Elektrode und Deckel überprüft neu membraniert und kalibriert. Kal. I.o. wird beobachtet und evtl. Neue Elektrode eingebaut. Parameter Unterdrückung bei Kal und QC Fehler,eingestellt. | SWO-519116 | Assisted Replacement | Electrode | Membrane | pCO2 |  | PCO2 Kal and QC problems to set parameters in case of errors lock | Archive checked, PCO2 fluctuates very strongly for weeks | Electrode and lid checked re-membraneed and calibrated. Kal. Ok. is observed and may be New electrode installed. Parameter suppression for Kal and QC error,set.
No | Visit | Complaint | RAME | Oxi hardware error codes 0251 and 0579. | Customer reported 579&251 errors after 2 point calibrations, this is an intermittent error. Ran complete oxi test and found all levels passed but the halogen lamp value is 3890. | Adjusted lamp value to 3280 and tested analyzer. Ran THB calibration, one and two point calibrations. Ran QC found all levels passed. Re-enabled oxi parameters | SWO-517399 | Adjustment | HW | Oxi module | Lamp unit | Halogen | Oxi hardware error codes 0251 and 0579. | Customer reported 579&251 errors after 2 point calibrations, this is an intermittent error. Ran complete oxi test and found all levels passed but the halogen lamp value is 3890. | Adjusted lamp value to 3280 and tested analyzer. Ran THB calibration, one and two point calibrations. Ran QC found all levels passed. Re-enabled oxi parameters
No | Visit | Complaint | RIND | pH calibration error. | Check the system. Check the calibration logs. Check the QC logs. | pH parameter unstable error intermittent in sample. protein removal & clean chamber still the error. Need to check with electrode. | SWO-520122 | Revisit required | Electrode | Electrode | pH |  | pH calibration error. | Check the system. Check the calibration logs. Check the QC logs. | pH parameter unstable error intermittent in sample. protein removal & clean chamber still the error. Need to check with electrode.
No | Visit | Complaint | RAME | Electrode Sensor disconnected | Analyzer picked up | Analyzer shipped for repairs to westlake | SWO-518696 | [5 empty] | Electrode Sensor disconnected | Analyzer picked up | Analyzer shipped for repairs to westlake
No | Visit | Complaint | RPOL | Błędne pomiary hemoglobiny występujące wraz z komunikatem: Problem z próbką (0653). | Sprawdzenie elementów układu transportu próbki. Stwierdzono wyciek na wężyku pompy pH $BG. Wykonano test oxi - wynik pozytywny. | Wymiana nieszczelnego wężyka. Kalibracje i testy serwisowe. Kalibracja hemoglobiny. Test QC - OK. | SWO-517995 | Revisit required | HW | pH/BG Module | Pump tube |  | Erroneous hemoglobin measurements occurring with the message: Problem with the sample (0653). | Inspect the sample transport system elements. A leak was found on the pH /BG pump tubing. Oxi test was performed positive. | Replacing the leaky hose. Calibrations and service tests. Hemoglobin calibration. QC Test-OK.
No | Visit | Complaint | RPOL | Wykryto prąd upływu (0963) | Wizualna kontrola elementów aparatu. Stwierdzono wyciek na zaworach V6, V7 i V8. | Wymiana zaworów V6 (rinse), V7 (air) i V8 (waste). Kalibracje i testy serwisowe. Test QC - OK. | SWO-517997 | Defect | HW | Fluid module | Valves | 963, V6, V7, V8 | Leakage current detected (0963) | Visual control of camera elements. Leaks were found on the V6, V7 and V8 valves. | Replacement of V6 (rinse), V7 (air) and V8 (waste) valves. Calibrations and service tests. QC Test-OK.
No | Visit | Complaint | RSAF | PCO2 Cal failure | PCO2 Cal error 1316 New SP installed 19 $03 $2019 New SC installed 13 $03 $2019 Cal errors started 15 $03 $2019 | Cleaned inlet gasket and probe Reinserted SC Cal acceptable after start up QC 's acceptable | SWO-518490 | Assisted Replacement | SC | SC | pCO2 | 1316 | PCO2 Cal failure | PCO2 Cal error 1316 New SP installed 19 /03 /2019 New SC installed 13 /03 /2019 Cal errors started 15 /03 /2019 | Cleaned inlet gasket and probe Reinserted SC Cal acceptable after start up QC 's acceptable
No | Visit | Complaint | RESP | Sospecha de coagulo. Equipo en intervención requerida | Sospecha de coágulo | Se procede a entrar en servicio y comprobar detección de coagulo. Se observan valores elevados. Se realizan varias limpiezas sin solventarse. Se procede a la sustitución del CE, sin aparente solución. Se procede a realizar un Kit de tubos y una limpieza de contactos, así como cambio de Inlet. Parece que los valores vuelven a estar dentro de rango pero elevados. Se realizan varias Cal y Qc’s. | SWO-519029 | Defect | HW | Basic Unit | PM Kit | Clot; 1537 | Suspicion of clot. Intervention team required | Suspected clot | Proceed to enter service and check clot detection. High values are observed. Several unsolved cleanings are performed. The EC is replaced without apparent solution. A Tube Kit and contact cleaning are performed, as well as Inlet change. It seems that the values are back in range but elevated. Several Cal and Qc's are performed.
No | Visit | Complaint | RLTD | Please call Cecilia on 07479 241 036 | Analyser failing due to leak detected 1184. Fault due to hole in the oxi pinch valve tube. Replaced Oxi pinch valve tube. Tested fluid pathway in service programs - all now ok. Restarted analyser and verified operation - all ok. | Replaced Oxi pinch valve tube. | SWO-518912 | Defect | HW | Oxi module | Pinch valve tube | 1184 | Please call Cecilia on 07479 241 036 | Analyser failing due to leak detected 1184. Fault due to hole in the oxi pinch valve tube. Replaced Oxi pinch valve tube. Tested fluid pathway in service programs - all now ok. Restarted analyser and verified operation - all ok. | Replaced Oxygen Pinch valve tube.
No | Visit | Complaint | RLTD | Irene Malik 01708 435000 ex 6845 or Anna on ex 8472 | Clot detected in OXI and SC. | Soaked Analyser fluid pathway with hypochlorite solution. Performed fluid transport checks and tests. Analyser has passed calibrations and QC's. | SWO-518513 | Cleaned | Fluidic pathway | Fluidic pathway | Oxi module | Clot | Irene Malik 01708 435000 ex 6845 or Anna on ex 8472 | Clot detected in OXI and SC. | Soaked Analyser fluid pathway with hypochlorite solution. Performed fluid transport checks and tests. Analyser has passed calibrations and QC's.
No | Visit | Complaint | RPAC | Ellen called 3349 0811 grinding noise after changing fluid pack | Inspect the SP. Broken Pump tube. | Clean Qc material from area affected by the broken pump tube. Clean and check the pump rotor. No SP on site. return to WS and install new SP. Check all functions normal. Perform calibrations and QC's. Perform tHb calibration which was overdue. Return ABL90 to site and Check Cals and QC's again. All OK. | SWO-518199 | Assisted Replacement | SP | SP | Pump tube | broken | Ellen called 3349 0811 grinding noise after changing fluid pack | Inspect the SP. Broken Pump tube. | Clean Qc material from area affected by the broken pump tube. Clean and check the pump rotor. No SP on site. return to WS and install new SP. Check all functions normal. Perform calibrations and QC's. Perform tHb calibration which was overdue. Return ABL90 to site and Check Cals and QC's again. All OK.
Yes | Visit | Complaint | RPAC | Ellen called 3349 0811 grinding noise after changing fluid pack | Returning Analyser | SP replaced in Workshop and Calibrations and QC's performed OK. Returned analyser to site and repeated Calibration and QC's. All functions normal. | SWO-518774 | Re classified RMED | same day upSWO |  |  |  | Ellen called 3349 0811 grinding noise after changing fluid pack | Returning Analyser | SP replaced in Workshop and Calibrations and QC's performed OK. Returned analyser to site and repeated Calibration and QC's. All functions normal.
Yes | Visit | Complaint | RBNL | Beeldscherm uit de scharnieren afgebroken. | Blauwe omkadering scherm is volledig losgekomen. | Display unit vervangen onder garantie. | SWO-518388 | Re classified RMED | same day upSWO |  |  |  | The display from the hinges. | Blue Bezel screen has completely come loose. | Replace the Display unit under warranty.
No | Visit | Complaint | RBNL | Battery toont 0%, analyzer functioneert nog op batterijspanning Volgens RMED moet deze dus vervangen worden | Batterij op 0% maar werkt nog wel wanneer we de spanning afzetten. | Reset van de batterij helpt niet, deze wordt vervangen. Batterij % wordt weer degelijk weergegeven. | SWO-518391 | Defect | HW | Battery module | Battery module | 0% | Battery shows 0%, analyzer still functions on battery voltage According to RMED, it must be replaced | Battery is 0% but still works when we drop the voltage. | Battery Reset does not help, it will be replaced. Battery% is displayed soundly.
No | Visit | Complaint | RKOR | QC fixed range error | . | Fixed assigned values resetting | SWO-520669 | Setup | SW | RMED SW | fixed ranges |  | QC fixed range error | . | Fixed assigned values resetting
No | Visit | Complaint | RKOR | shutdown error, Printer error | . | shutdown error, Printer error | SWO-520673 | Defect | HW | Printer Module | Printer Module |  | shutdown error, Printer error | . | shutdown error, Printer error
No | Visit | Complaint | RKOR | shutdown error | . | CPU Unit, CF card replace. Resetting | SWO-520671 | Defect | HW | DMS_PC | CPU |  | shutdown error | . | CPU Unit, CF card replace. Resetting
No | Visit | Complaint | RKOR | printer error | . | CPU unit clean. Cable reconnected. | SWO-520674 | Adjustment | HW | Printer Module | Cable |  | printer error | . | CPU unit clean. Cable reconnected.
No | Visit | Complaint | RESP | Falló en el lactato | Ajuste de cero electrónico del lactato. | Ajuste del cero electrónico de todos los parámetros. | SWO-518567 | [5 empty] | Failed in lactate | Electronic lactate zero setting. | Electronic zero adjustment of all parameters.
No | Visit | Complaint | RCHN | 电池无法蓄电 | 电池无法蓄电，借431-018一个 | 更换431-018后修复 by Andrew | SWO-516891 | [5 empty] | The battery can't store power. | Battery can not store power, borrow 431-018 one | Repair by Andrew after replacement 431-018
No | Visit | Complaint | RESP | Touch panel | Comprobación del estado de funcionamiento del equipo | Sustitución del touch panel, calibración del display, calibración del sensor, verificación final de funcionamiento. | SWO-518701 | [5 empty] | Touch panel | Checking the operating status of the equipment | Touch panel replacement, display calibration, sensor calibration, final operating check.
No | Visit | Complaint | RCHN | 仓门无法关闭 4548 | 仓门无法关闭 4548,借903-028一个。 | 更换903-028备件后修复 by Andrew | SWO-516834 | Defect | HW | Cover and Chassis | PCB, cabinet top | 4548 | Warehouse door cannot be closed 4548 | The warehouse door cannot be closed 4548, borrowing 903-028 one. | Repair by Andrew after replacement of 903-028 spare parts
No | Visit | Complaint | RLTD | Analyser memory chip not detected contact service displayed on analyser screen Dr Lewis Gough | Analyser not recognising ID chip | Fit new ID chip and analyser now booting up ok, but could not install new SP, further investigating required. | SWO-518605 | [5 empty] | Analyser memory chip not detected contact service displayed on analyser screen Dr Lewis Gough | Analyser not recognising ID chip | Fit new ID chip and analyser now booting up ok, but could not install new SP, further investigating required.
No | Visit | Complaint | RFRA | Gaz1 remplacé tous les 10 jours et Gaz2 tous les mois | Aucune | Remplacement des deux régulateurs de gaz. Remplacement du pressure transducer unit . Remplacement du joint pH. Calibration. QC hors gamme sur creat. Investigation en cours | SWO-519421 | Defect | HW | Gas module | Gas 1 |  | Gas1 replaced every 10 days and Gaz2 every month | No | Replacement of the two gas regulators. Replacing the pressure transducer unit . Replacing the pH joint. Calibration. QC off-range on creat. Investigation underway
No | Visit | Complaint | RIND | pO2 level 1 QC was getting failed | Checked the data logs pO2 level 1 QC was getting failed | Performed decontamination 2-point calibration & QC done pO2 value was within the low range | SWO-519943 | Cleaned | HW | pH/BG Module | pH/BG Module | pO2 QC L1 | pO2 level 1 QC was getting failed | Checked the data logs pO2 level 1 QC was getting failed | Performed decontamination 2-point calibration & QC done pO2 value was within the low range
No | Visit | Complaint | RAME | The printer has streaks in the print | Not printing | I replaced the printer | SWO-519130 | Defect | HW | Printer Module | Printer Module |  | The printer has streaks in the print | Not printing | I replaced the printer
No | Visit | Complaint | RDEU | Temp.fehler bei allen Modulen | Nach der Umschaltung aufm Notstrom ist Temp.fehler aufgtreten. | Neustart bringt keine Veränderung. Alle Sicherungen in WSC pcb. kontrolliert. Powersupply ausgetauscht. Neustart. | SWO-518595 | Defect | HW | Power Supply | Power supply |  | Temp.error for all modules | After switching to the emergency current, Temp.error occurs. | Restart brings no change. All backups in WSC pcb. Controlled. Powersupply. Restart.
No | Visit | Complaint | RAME | Unit is not running QC. | Upon arrival, analyzer was powered down. | Power on analyzer and let it initialize. Analyzer performed 2pt Cal automatically where PCO2 passed but observed PCO2 was on upper limit of range. Level 2 QC ran by analyzer but failed for lower limit. Performed membrane change with QC still failing after membrane change. Swapped PCO2 electrode $membrane with analyzer SN 754N2302N001. Performed L2 QC test which PCO2 stabilized on the Mean target. Performed L1 and L3 QCs which passed. | SWO-518686 | Assisted Replacement | Electrode | Electrode | pCO2 | QC | Unit is not running QC. | Upon arrival, analyzer was powered down. | Power on analyzer and let it initialize. Analyzer performed 2pt Cal automatically where PCO2 passed but observed PCO2 was on upper limit of range. Level 2 QC ran by analyzer but failed for lower limit. Performed membrane change with QC still failing after membrane change. Swapped PCO2 electrode $membrane with analyzer SN 754N2302N001. Performed L2 QC test which PCO2 stabilized on the Mean target. Performed L1 and L3 QCs which passed.
Yes | Visit | Complaint | RAME | Analyzer would not perform QC measurements. Technician performed power down/up and the analyzer would not completely startup and Status Code Error 10 occurred. | Entered the blood gas lab and the analyzer was powered down. | Powered the analyzer up and allowed for it it to run through the its standard startup protocol. Once the analyzer completed its startup, no Status Code 10 was observed. After running multiple calibrations and QC, the Status Code 10 did not return. | SWO-518685 | Re classified RMED | same day upSWO |  |  |  | Analyzer would not perform QC measurements. Technician performed power down/up and the analyzer would not completely startup and Status Code Error 10 occurred. | Entered the blood gas lab and the analyzer was powered down. | Powered the analyzer up and allowed for it it to run through the its standard startup protocol. Once the analyzer completed its startup, no Status Code 10 was observed. After running multiple calibrations and QC, the Status Code 10 did not return.
No | Visit | Complaint | RIND | Software corrrupted | Software corrrupted | Checked the supply voltages , supply is fine , restarted the instrument , but software is not loading , therefore reinstalled the software and did necessary configuration ,print setup and then run calibration and tests | SWO-519780 | Defect | SW | RMED SW | RMED SW |  | Software corrrupted | Software corrrupted | Checked the supply voltages , supply is fine , restarted the instrument , but software is not loading , therefore reinstalled the software and did necessary configuration ,print setup and then run calibration and tests
Yes | Visit | Complaint | RDEU | 18.03.2019 Pf. Konrad 03616541310 meldet einen Fehler 600. Demo Option Key wird bald ablaufen. | Demo Key läuft ab | Key eingelesen Geräteneustart aufgrund QC Fehler pCO2 und Ref. membraniert | SWO-518479 | Re classified RMED | planned SWO |  |  |  | 18.03.2019 Pf. Konrad 03616541310 reports an error 600. Demo Option Key will expire soon. | Demo Key expires | Key read in Device restart due to QC errors pCO2 and Ref. membraneed
No | Visit | Complaint | RLTD | Cartridge storage door fails to operate after restart, stepper motor chatters back and forth. | Cartridge storage door fails to operate after restart, stepper motor chatters back and forth. | Restarted and reloaded software as this was sketchy when flowing through service programs. Removed, flattened and reversed cable from main board to front PCB, issue still remained. Will have to return with new motor and front PCB. | SWO-520772 | Revisit required | HW | Analyzer Control Module | PCB, Main Board |  | Cartridge storage door fails to operate after restart, stepper motor chatters back and forth. | Cartridge storage door fails to operate after restart, stepper motor chatters back and forth. | Restarted and reloaded software as this was sketchy when flowing through service programs. Removed, flattened and reversed cable from main board to front PCB, issue still remained. Will have to return with new motor and front PCB.
No | Visit | Complaint | RFRA | Labosud NIMES connexion Réseau AQT90 | test cmd : Adresse provisoire en mémoire. | Adresse réseau perdu. Remise en place et excécusion ligne de commande ( commit) pour mise en mémoire de l'adresse IP. Test prise en main distante par AQURE : bon | SWO-518954 | Setup | SW | External SW | His/Lis | IP address | Labosud NIMES Network connection AQT90 | cmd test: Provisional address in memory. | Lost network address. Re-in-place and execution command line (commit) for memory of the IP address. AQURE remote hand-held test: good
No | Visit | Complaint | RLTD | Email received from Vanessa Rumens. | Intermittent errors causing calibrations to abort - error codes: 4019 - Measure_23; 5375 - ReadOpticalROM_56. Replaced faulty optical unit. Calibrated factory reference for new optical unit and allowed analyser to warm-up. Ran calibrations for CRP (new lot 15780), TnT (existing lot 15733) and D-Dimer (new lot 15729) - all passed. Analyser ready for use. | Replaced optical unit. Re-calibrated parameters. | SWO-518473 | Defect | HW | Process Module | Optical Unit | 4019, 5375 | Email received from Vanessa Rumens. | Intermittent errors causing calibrations to abort - error codes: 4019 - Measure_23; 5375 - ReadOpticalROM_56. Replaced faulty optical unit. Calibrated factory reference for new optical unit and allowed analyser to warm-up. Ran calibrations for CRP (new lot 15780), TnT (existing lot 15733) and D-Dimer (new lot 15729) - all passed. Analyser ready for use. | Replaced optical unit. Re-calibrated parameters.
No | Visit | Complaint | RPOL | Błąd vacume, błąd kalibracji hemoglobiny. | Weryfikacja stanu technicznego analizatora | Wykonanie testów technicznych. Stwierdzono błąd podciśnienia. Uszkodzony wężyk pod zaworem oxi. Wymiana. Wymiana uszczelki wlotu oraz uszczelki połączenia. Testowanie | SWO-519047 | Defect | HW | Oxi module | Pinch valve tube |  | Error vacume, hemoglobin calibration error. | Checking the technical condition of the analyzer | Perform technical tests. There was a vacuum error. Faulty tubing under the oxi valve. Exchange. Inlet gasket replacement and connection gaskets. Testing
No | Visit | Complaint | RPOL | Aparat porzuca próbki | Puszczenie badania próbnego - aparat porzucił próbkę. Błędy ROTATEINLETWHEEL. | Korekta montażu inlet wheel. Puszczenie próbek testowych - OK. Wykonanie procedury czyszczenia aparatu. Przeprowadzenie kontroli jakości. | SWO-518785 | Adjustment | HW | Aspiration Module | Inlet Wheel | Potentially Use Error (FSE) | The camera drops the sample | Releasing a trial test-the camera dropped the sample. ROTATEINLETWHEEL errors. | Adjustment of the inlet wheel. Releasing test samples-OK. Perform a procedure for cleaning the camera. Perform quality control.
No | Visit | Complaint | RFRA | Erreur de sensibilité Po2 sur plusieurs cassettes | Erreur calibration pO2 hors limite. techniquemc@hotmail.com | Calibration tHb en retard => calibration tHb. Chgt: carte sensor, pinch valve, aiguille. Nettoyage: à l'hypochlorite du circuit fluidique. Nettoyage et dépoussiérage. Remplacement: Joint d'introduction, Joint connexion, tubulure vanne oxi. Test service: Fluid transport, Oxi $pO2. Lors du changement du pack de solution, penser à nettoyer le capteur rectangulaire sombre au-dessus du pack côté ABL. Chgt Pack Solution, calibration et contrôle: ok. | SWO-519395 | Assisted Replacement | SP | SP | pO2 | calibration | Po2 sensitivity error on multiple cassettes | Error calibration pO2 off limits. techniquemc@hotmail.com | THb calibration late - calibration tHb. Chgt: sensor card, pinch valve, needle. Cleaning: hypochlorite of the fluid circuit. Cleaning and dusting. Replacement: Introductory joint, joint connection, tube valve oxi. Test service: Fluid transport, Oxi $pO2. When changing the solution pack, consider cleaning the dark rectangular sensor above the ABL side pack. Chgt Pack Solution, calibration and control: ok.
No | Visit | Complaint | RDEU | QC im Level A werden sporadisch zu hochgemessen. | Daten ausgelesen | Servicedump und Datalogs generiert und an TS gemailt. | SWO-519754 | Assisted Replacement | SP | SP | pCO2 | QC L A high; Techline? | QC im Level A werden sporadisch zu hochgemessen. | Data read out | Servicedump and Datalogs generated and emailed to TS.
No | Visit | Complaint | RDEU | QCs werden bei Level A sporadisch zu hoch gemessen. | Daten ausgelesen | Servicedump und Datalogs generiert und an TS gemailt. | SWO-519757 | Assisted Replacement | SP | SP | pCO2 | QC L A high; Techline? | QCs are sporadically overmeasured at Level A. | Data read out | Servicedump and Datalogs generated and emailed to TS.
No | Visit | Complaint | RDEU | QCs werden sporadisch im Level A zu hoch gemessen | Datenspeicher überprüft | Servicedump und Datalogs generiert und an TS gemailt. | SWO-519752 | Assisted Replacement | SP | SP | pCO2 | QC L A high; Techline? | QCs are measured sporadically in level A too high | Data store checked | Servicedump and Datalogs generated and emailed to TS.
No | Visit | Complaint | RDEU | QC bei Level A sporadisch zu hoch | Datenausgelesen | Servicedump und Datalogs generiert und an TS gemailt. | SWO-519758 | Assisted Replacement | SP | SP | pCO2 | QC L A high; Techline? | QC bei Level A sporadisch zu hoch | Datenausgelesen | Servicedump and Datalogs generated and emailed to TS.
No | Visit | Complaint | RDEU | QC wird bei Level A sporadisch zu hoch gemessen | Daten ausgelesen | Servicedump und Datalogs generiert und an TS gemailt. | SWO-519766 | Assisted Replacement | SP | SP | pCO2 | QC L A high; Techline? | QC is measured sporadically too high at Level A | Data read out | Servicedump and Datalogs generated and emailed to TS.
No | Visit | Complaint | RDEU | 20.03.2019 Herr König 023712128200 möchte das die Normwerte wieder eingestellt werden. | Referenzwerte bestimmen | Referenzwerte nach Kundenvorgabe eingetragen. | SWO-519714 | Setup | SW | RMED SW | Reference ranges |  | 20.03.2019 Mr. König 023712128200 wants to set the standard values again. | Determine reference values | Reference values according to customer specifications.
No | Visit | Complaint | RPAC | Monica reported long measurement time on patient samples. Checked logs, found sensor cassette was replaced on 18/3/19 and occasional Rilibak violation on pCO2 after sensor cassette replacement. | Checked logs, found sensor cassette was replaced on 18 $3 $19 and occasional Rilibak violation on pCO2 since sensor cassette replacement. | Saved logs. Advised Monia measurement may include calibration after having new sensor cassette and therefore may be longer than usual. Also continue monitoring for intermittent Rilibak violation on pCO2 until next Monday, if it does not improve consider sensor cassette replacement. | SWO-519704 | Assisted Replacement | SC | SC | pCO2 | QC Rilibäk | Monica reported long measurement time on patient samples. Checked logs, found sensor cassette was replaced on 18/3/19 and occasional Rilibak violation on pCO2 after sensor cassette replacement. | Checked logs, found sensor cassette was replaced on 18 /3 /19 and occasional Rilibak violation on pCO2 since sensor cassette replacement. | Saved logs. Advised Monia measurement may include calibration after having new sensor cassette and therefore may be longer than usual. Also continue monitoring for intermittent Rilibak violation on pCO2 until next Monday, if it does not improve consider sensor cassette replacement.
No | Visit | Complaint | RDEU | 20.03.2019, Sr. Heike, Tel: 03634521539: AQC Fehler bei pCO2 alle Level im Ergebnis zu hoch | Qualitätskontrollen in pCo2 alle zu hoch | Bei Ankunft stand Gerät auf Fehlermodus und wollte durchgespült werden. Einlassdichtung getauscht, sehr verschmutzt und verstopft. Sensorkassette wurde am 18.03.2019 getauscht und seitdem nicht wieder. Sie wurde zwar immer wieder herausgenommen und wieder eingesetzt, aber der Fehler blieb bestehen. Stattdessen wurde der Lösungspack zweimalig erneuert, was aber zu keinem Erfolg führte. Es lag kein Kalibrierfehler vor, aber die Qualitätskontrollen lagen im pCo2 weit erhöht bei allen drei Level. Sensorkassette aus Sondershausen eingebaut. Startphase abgewartet. Danach Qualitätskontrollen in pCo2 alle ok. Glu-Lac benötigten noch Zeit um in Bereich zu kommen. Anwender führen noch Kalibrierung durch und melden sich, falls Probleme auftreten sollten. Gerät überprüft. | SWO-519808 | Assisted Replacement | SC | SC | pCO2 | QC L A and B and C high | 20.03.2019, Sr. Heike, Tel: 03634521539: AQC error with pCO2 all levels too high in result | Quality controls in pCo2 all too high | On arrival, the device was in error mode and wanted to be flushed through. Inlet seal replaced, very soiled and clogged. Sensor cartridge was replaced on March 18, 2019 and has not been replaced since. It was repeatedly taken out and reinstated, but the error persisted. Instead, the solution pack was renewed twice, but this did not lead to any success. There was no calibration error, but the quality controls in the pCo2 were greatly increased at all three levels. Sensor cassette from Sondershausen installed. start-up phase. After quality control in pCo2 all ok. Glu-Lac still needed time to get into range. Users are still performing calibration and logging in if problems occur. device checked.
No | Visit | Complaint | RAME | No sound,issue with flush system errors. | Defective speaker | Replaced defective speaker. Inspected error logs, downloaded data logs to inspect for trend in flush error. Testing is done on Bovine blood and customer has also started using safe pico syringes and may have introduced air when inputting samples. | SWO-519453 | Defect | HW | DMS_PC | Speaker |  | No sound,issue with flush system errors. | Defective speaker | Replaced defective speaker. Inspected error logs, downloaded data logs to inspect for trend in flush error. Testing is done on Bovine blood and customer has also started using safe pico syringes and may have introduced air when inputting samples.
No | Visit | Complaint | RAME | analyzer is locked in flush | No investigation needed. System was in ready mode and green lights once onsite. | Discussed recent CAP proficiency changes, upcoming analyzer go live. Visited CAP website together to address questions. No correction needed. | SWO-519413 | No fault found | No fault found | No fault found | No fault found |  | analyzer is locked in flush | No investigation needed. System was in ready mode and green lights once onsite. | Discussed recent CAP proficiency changes, upcoming analyzer go live. Visited CAP website together to address questions. No correction needed.
No | Visit | Complaint | RAME | Multiple cal errors | Multiple cal errors | Replaced pinch valve and ran all service fluid tests | SWO-519515 | Defect | HW | Oxi module | Pinch valve |  | Multiple cal errors | Multiple cal errors | Replaced pinch valve and ran all service fluid tests
No | Visit | Complaint | RCHN | 电池无法蓄电 | 电池无法蓄电，借431-018一个 | 更换431-018后修复 by Andrew | SWO-517475 | [5 empty] | The battery can't store power. | Battery can not store power, borrow 431-018 one | Repair by Andrew after replacement 431-018
No | Visit | Complaint | RCHN | 条码扫描器故障 | 条码扫描器故障，借902-924一个 | 更换902-924后修复 by Andrew | SWO-517580 | Defect | HW | Display | Barcode Reader |  | Barcode scanner failure | Barcode scanner failure, borrow 902-924 one | Repair by Andrew after replacing 902-924
No | Visit | Complaint | RCHN | AU到存储盘排线磨损 | AU到存储盘排线磨损 | 更换636-544后修复 by Andrew | SWO-519776 | Defect | HW | Storage Module | Cable, cup transport front |  | AU-to-storage disk line wear | AU-to-storage disk line wear | Repair by Andrew after replacing 636-544
No | Visit | Complaint | RCHN | HCT无法测量 1267 | HCT无法测量 1267 | 清洁HCT，重新做样本。 by Andrew | SWO-518967 | Cleaned | HW | Aspiration Module | Needle Wash Unit | 1267 | HCT cannot measure 1267 | HCT cannot measure 1267 | Clean the HCT and remake the sample. by Andrew
No | Visit | Complaint | RIND | Software corrrupted | NA | Checked the calibration log , checked the patient data , activity log , no errors found , customer happy with the results . | SWO-519783 | No fault found | SW | RMED SW | RMED SW |  | Software corrrupted | In | Checked the calibration log , checked the patient data , activity log , no errors found , customer happy with the results .
No | Visit | Complaint | RDEU | Testkassette hat sich verklemmt, Fehler 5785 Setsafe35 | D-Dimere Testkassette hat sich verklemmt Fehler 5785 Setsafe35. Testkassette wurde falsch $ verdreht eingelegt. | Über Servicemenue Testkassette entriegelt und entfernt. D-Dimer Kassette erneuert. Systemtest, Reinigung und QC durchgeführt. Alles wieder o.k. | SWO-519657 | User Error | Consumables | Test Kit | Cartridge, Broken | twisted; 5785 | Test cassette got stuck, error 5785 Setsafe35 | D-Dimere test cassette has become jammed error 5785 Setsafe35. Test cartridge has been inserted incorrectly . . . . . . . . . . . . . . . . . . . . | Unlocked and removed via Servicemenue test cassette. D-Dimer cassette renewed. system test, cleaning and QC. All again OK.
No | Visit | Complaint | RLTD | Cartridge storage door fails to operate after restart, stepper motor chatters back and forth. | Cartridge storage door fails to operate after restart, stepper motor chatters back and forth. | Replaced stepper motor and disconnected everything from the front PCB, error remained. Replaced front cabinet PCB and motor, and reloaded software, issue still remained. Will have to return with a new main control PCB. | SWO-520773 | Revisit required | HW | Analyzer Control Module | PCB, Main Board |  | Cartridge storage door fails to operate after restart, stepper motor chatters back and forth. | Cartridge storage door fails to operate after restart, stepper motor chatters back and forth. | Replaced stepper motor and disconnected everything from the front PCB, error remained. Replaced front cabinet PCB and motor, and reloaded software, issue still remained. Will have to return with a new main control PCB.
No | Visit | Complaint | RPOL | Błąd Undefinded Lookup String 18.1 | Sprawdzenie przyczyna pojawienia się błędu - błąd spowodwany jest nieprawidłową pracą zamka koła kaset. | Smarowanie osi napędowej mechanizmu zamka kaset, czyszczenie czujników optycznych. | SWO-519287 | Adjustment | HW | Incubator Module | wheel lock | lubricated | Błąd Undefinded Lookup String 18.1 | Check the cause of the error-error caused by incorrect operation of the cassette wheel lock. | Lubricate the drive axle of the cassette lock mechanism, cleaning the optical sensors.
No | Visit | Complaint | RPOL | Błąd kalibracji i QCna pH | Czułość pH ok 87%, QC powyżej zakresu referencyjnego. | Testy elektrod pH oraz Ref. Sprawdzenie klapki elektrody pH. Wymiana odczynników - wadliwy odczynnik Cal1. | SWO-519274 | Assisted Replacement | Solutions | Solutions | Cal 1 |  | Calibration Error and QCna pH | PH sensitivity approx. 87%, QC above reference range. | Test electrodes pH and Ref. Check the pH electrode flap. Replacement of reagents-defective reagent Cal1.
No | Visit | Complaint | RPOL | Błędne pomiary hemoglobiny występujące wraz z komunikatem: Problem z próbką (0653). | Sprawdzenie komunikatów błędów przy badaniach pacjentów - brak komunikatów dot. problemów z próbką. Niepowodzenie kalibracji glukozy i mleczanów. Sprawdzenie stanu membran. | Kilkukrotne przeprowadzenie kalibracji - OK. Kontrola jakości - OK. Wielokrotne porównanie wyników tej samej próbki z aparatem ABL 810 oraz Sysmex - wyniki porównywalne. | SWO-519299 | Defect | HW | pH/BG Module | Pump tube |  | Erroneous hemoglobin measurements occurring with the message: Problem with the sample (0653). | Check for error messages in patient studies-no DOT messages Problems with the sample. Glucose and lactate calibration failed. Check the condition of the membrane. | Calibration carried out several times-OK. Quality Control-OK. Multiple comparison of the results of the same sample with ABL 810 and Sysmex-comparable results.
No | Visit | Complaint | RESP | Error de fuga y transporte de luquidos | Inlet gasket partido y aguja doblada | Sustitucion del inlet gasket propiedad del cliente y enderezar inlet probe Cambio de sc Cal, qc | SWO-520135 | Defect | HW | Inlet module | Inlet gasket | 1237, 1063 | Leakage error and luquidation transport | Inlet gasket split and needle bent | Substitution of customer-owned inlet gasket and inlet probe Changing sc Cal, qc
No | Visit | Complaint | RHUN | Műszer nem indul | Műszer nem indul, fekete képernyő. | Adatbázis fájl sérült, szoftver újratelepítése. | SWO-519962 | Defect | SW | RMED SW | Database | Pot. Status code 2 | Instrument does not start | Instrument does not start, black screen. | Database file corrupted, software reinstalling.
No | Visit | Complaint | RPAC | Joanne called 6458 2753 leak detected | Leak detected. Customer has no spare inlet gasket | Inlet probe a bit bended. Stright the probe. Replaced the inlet gasket. Tested, all ok | SWO-521606 | Defect | HW | Auto Inlet module | Probe bend |  | Joanne called 6458 2753 leak detected | Leak detected. Customer has no spare inlet gasket | Inlet probe a bit bended. Stright the probe. Replaced the inlet gasket. Tested, all ok
No | Visit | Complaint | RDEU | 20.03.2019 17.25Uhr Herr Reißmann tel:+493441740248: Nach LP-Tausch pCO2 Fehler im Level A und C | QC-Messfehler pCO2 in Lösung A und C ==> Fehlersuche | Einstellung Fixe Zielwerte kontrolliert Sensorkassette erneuert und defekte SC (945-009, R0876) reklamiert | SWO-519977 | Assisted Replacement | SC | SC | pCO2 | QC L A and C | 20.03.2019 17.25pm Mr Reißmann tel:+493441740248: After LP exchange pCO2 errors in level A and C | QC measurement error pCO2 in solution A and C ==> Troubleshooting | Setting Fixed target values controlled Sensor cassette renewed and defective SC (945-009, R0876) complained
No | 2 | Y | RCZE | Blank | Error in loader, Cannot load application. Error number 7, Error source: Roma, Description: Out of memory |  | D5DBD3D910434202 | Defect | SW | RMED SW | Database |  | Blank | Error in loader, Cannot load application. Error number 7, Error source: Roma, Description: Out of memory
No | Visit | Complaint | RCHN | ph值结果偏高 | ph值结果偏高(未更换备件）零点校准后修复 | 零点校准后修复 by jesson | SWO-522734 | [5 empty] | Ph-value result is high | High ph-value results (unreplaced spare parts) after zero calibration repair | Fix by jesson after zero calibration
No | Visit | Complaint | RAME | po2 out of range on External QC | Confirmed issue when onsite. Also noticed VC check error on sp. | Proceed to replace both sc $sp. system cycle passed and then ran QCs which also passed. | SWO-520173 | [5 empty] | po2 out of range on External QC | Confirmed issue when onsite. Also noticed VC check error on sp. | Proceed to replace both sc $sp. system cycle passed and then ran QCs which also passed.
No | Visit | Complaint | RIND | QC values out of range. | found customer using Bio-Rad chemistry QC for electrolytes. and user not processing QC correctly. | Entered the QC ranges and given training for QC running. explained about level selection, running QC, checking QC logs,taking print out of L-J Charts and statics. | SWO-520273 | [5 empty] | QC values out of range. | found customer using Bio-Rad chemistry QC for electrolytes. and user not processing QC correctly. | Entered the QC ranges and given training for QC running. explained about level selection, running QC, checking QC logs,taking print out of L-J Charts and statics.
No | Visit | Complaint | RCHN | 漏电流 | 漏电流，借903-581一个 | 更换903-581后修复 by Andrew | SWO-518874 | Defect | HW | Basic Unit | PCB, Sensor interface |  | Leak current | Leak current, borrow 903-581 one | Repair by Andrew after replacement 903-581
No | Visit | Complaint | RIND | Frequent SC Failure | Frequent SC Failure | Found and replace the part, Sensor Connector Board. Reinstalled the consumables , Analyser was ready And Under observation, Part No. 903-612 S $N : QS17290157 | SWO-519659 | [5 empty] | Frequent SC Failure | Frequent SC Failure | Found and replace the part, Sensor Connector Board. Reinstalled the consumables , Analyser was ready And Under observation, Part No. 903-612 S $N : QS17290157
No | Visit | Complaint | RPOL | Kontrola jakości poza zakresem dla parametrów Na+/K+ | Sprawdzenie historii kontroli jakości - serwis stwierdza przekroczony dopuszczalny zakres dla K+ $Na+ dla 4 z 5 zmierzonych kontrolek. | Wyzerowanie kanałów, zebranie danych diagnostycznych - dziennków systemu. W celu wyjaśnienia przyczyn zachowania analizatora zostanie wystosowane zapytanie do siedziby w Kopenhadze. | SWO-519925 | [5 empty] | Quality control out of range for parameters Na +/K + | Checking the quality control history-the service finds the exceeded allowable range for K + /Na + for 4 of the 5 measured controls. | Reset channels, collect diagnostic data-logs of the system. To clarify the reasons for the analyser's behaviour, a request to Copenhagen headquarters will be made.
No | Visit | Complaint | RPOL | Kontrola jakości dla parametru pO2 poza dopuszczalnym zakresem. | Serwis stwierdza przekroczenie dopuszczalnego zakresu dla QC Level 2, po niedawnej wymiane pakietu roztworów oraz kasety sensorowej. | Zerowanie kanału pO2 z poziou serwisowego, eksport dziennika kalibracji, QC oraz wymian. Zostanie wystosowane zapytanie o przyczynę pojawiania się problemu do siedziby w Kopenhadze. | SWO-519771 | [5 empty] | Quality control for the pO2 parameter outside the allowable range. | The service finds an excess of the permissible range for QC Level 2, after a recent replacement of the solution package and the sensor cassette. | Reset channel pO2 from service level, calibration log export, QC and exchanges. You will be asked about the cause of the problem to be raised in Copenhagen.
No | Visit | Complaint | RKK | SENDX指示、メンテナンスキット、センサーコネクタ―の交換。 | ケース詳細の通り | SENDX指示、メンテナンスキット、センサーコネクタ―の交換。 | SWO-522830 | [5 empty] | Replacement of SENDX instructions, maintenance kits, and sensor connectors. | As the case details | Replacement of SENDX instructions, maintenance kits, and sensor connectors.
No | Visit | Complaint | RCHN | pco2 po2质控错误 | pco2 po2质控错误 | 重新质控后修复 by Andrew | SWO-519736 | [5 empty] | pco2 po2 QC error | pco2 po2 QC error | Fix by Andrew after re-quality control
No | Visit | Complaint | RDEU | 15.03.2019 13:10 Fr. Zenses +4922678897049: Ref-Kammer-Deckel mech. defekt | Ref-Kammerdeckel defekt. | Ref-Kammerdeckel voll mit Flüssigkeit aus der Ref-Membran, auskristallisiert und somit den defekt hervorgerufen. Deckel ausgetauscht. Kal & QC | SWO-520576 | Revisit required | HW | pH/BG Module | Electrode connector |  | 15.03.2019 13:10 Fr. Zenses +4922678897049: Ref-chamber lid mech. faulty | Ref-chamber lid defective. | Ref-chamber lid full of liquid from the ref membrane, crystallized and thus caused the defective. Lid replaced. Kal & QC
No | Visit | Complaint | RIND | Ca parameter error. | Check the parameter. | Clot in chamber cause the parameter error. Guide customer for troubleshoot. | SWO-520125 | Cleaned | Fluidic pathway | El/Met Module | measuring chamber | Clot | Ca parameter error. | Check the parameter. | Clot in chamber cause the parameter error. Guide customer for troubleshoot.
No | Visit | Complaint | RAME | Inhomogenous samples | Inhomogenous samples | Straightened slightly bent probe and inspected $cleaned PH $BG chamber. Performed a new LS adjust and ran cal and QC. Customer tested with several good samples. | SWO-520068 | Cleaned | Fluidic pathway | pH/BG Module | measuring chamber |  | Inhomogenous samples | Inhomogenous samples | Straightened slightly bent probe and inspected $cleaned PH $BG chamber. Performed a new LS adjust and ran cal and QC. Customer tested with several good samples.
No | Visit | Complaint | RAME | Analyzer failing to print | Printer error. This has been an on going issue. | Replace printer unit complete. | SWO-520081 | Defect | HW | Printer Module | Printer Module | 216 | Analyzer failing to print | Printer error. This has been an on going issue. | Replace printer unit complete.
No | Visit | Complaint | RFRA | Remise en place de pièce prise pour dépanner AQT usic | Remise en place de pièce prise pour dépanner AQT usic | Remise en place du câble entre le capteur de température de l’incubateur et la carte de contrôle. Essais | SWO-519805 | Defect | HW | Incubator Module | sensor, temperature |  | Refit of part taken to troubleshoot AQT usic | Refit of part taken to troubleshoot AQT usic | Re-establishing the cable between the incubator temperature sensor and the control card. Tests
No | Visit | Complaint | RPOL | Czułość Crea 0% - QC mocno poza zakresem. | Testy Sensor Simulator - OK. | Wymian kasety sensorowej. | SWO-519887 | Assisted Replacement | SC | SC | Crea |  | Crea Sensitivity 0%-QC tight out of range. | Testy Sensor Simulator - OK. | Replace the sensor cassette.
No | Visit | Complaint | RPOL | Wymiana modułu kreatyniny, wymiana elektrody referencyjnej | Niestabilne wyniki QC Crea | Wymiana modułu kreatyniny. Wymiana elektrody referencyjnej oraz membrany. | SWO-519833 | Defect | HW | Met II module | MET II module | Crea QC unstable | Replacing the creatinine module, replacing the reference electrode | Niestabilne wyniki QC Crea | Replacement of the creatinine module. Reference electrode and membrane replacement.
No | Visit | Complaint | RAME | Keeps prompting for inlet gasket to be changed after already changed. | Inspection of the fluidic suggests a restriction in fluid flow. | Replaced all parts from a pm kit & SC for diagnostic purposes, system operational. | SWO-520196 | Defect | HW | Basic Unit | PM Kit |  | Keeps prompting for inlet gasket to be changed after already changed. | Inspection of the fluidic suggests a restriction in fluid flow. | Replaced all parts from a pm kit & SC for diagnostic purposes, system operational.
No | Visit | Complaint | RSAF | The is a leak detected in the analyzer | Inspected the analyzer for any visible crystals | Replaced inlet gasket holder | SWO-520319 | Defect | HW | Auto Inlet module | Inlet gasket holder |  | The is a leak detected in the analyzer | Inspected the analyzer for any visible crystals | Replaced inlet gasket holder
No | Visit | Complaint | RCHN | 机器黑屏 | 机器黑屏，借910-342一个 | 更换910-342后修复 by jesson | SWO-519648 | Defect | HW | DMS_PC | CPU |  | Machine black screen | Machine black screen, borrow 910-342 one | Repair by jesson after replacing 910-342
No | Visit | Complaint | RCHN | 进样门无法正常关闭，报错5784 | 进样门无法正常关闭，报错5784，借903-028（上海仓库无货，从广州发）,886-638,886-639各一个 | 更换903-028，886-638,886-639，同时加固了进样门上方的架子后修复 by Andre | SWO-518831 | Defect | HW | Cover and Chassis | PCB, cabinet top | 5784 | The sample door does not close properly, 5784 is reported | The entrance door can not be closed properly, reported error 5784, borrowed 903-028 (Shanghai warehouse is out of stock, from Guangzhou), 886-638, 886-639 each one | Replace 903-028, 886-638, 886-639, and repair by Andre after reinforcing the shelf above the sample door
No | Visit | Complaint | RCHN | K定标失败，血氧模块错误 | K定标失败，血氧模块错误，借902-833一个 | 更换902-833后修复 by jesson | SWO-519600 | Defect | HW | Oxi module | Oxi module | 379, 580 | K calibration failed, blood oxygen module error | K calibration failure, blood oxygen module error, borrow 902-833 one | Repair by jesson after replacement 902-833
No | Visit | Complaint | RCHN | 传感器电路错误 | 传感器电路错误，借902-972,903-334各一个 | 更换902-972,903-334后修复 by Andrew | SWO-519638 | [5 empty] | Sensor circuit error | Sensor circuit error, borrow 902-972, 903-334 each | Fix by Andrew after replacing 902-972, 903-334
No | Visit | Complaint | RCHN | 光学单元超出范围，报错1035 | 光学单元超出范围，报错1035，需要借件903-364 | 更换903-364后修复 | SWO-519707 | Defect | HW | Process Module | Optical Unit | 1035 | Optical unit out of range, error 1035 | Optical unit out of range, error 1035, need to borrow 903-364 | Repair after replacement 903-364
Yes | Visit | Complaint | RDEU | 22.03.2019, herr Stephan, Tel: 03477171165: Barcode Leser soll umgestellt werden- Typ Codabar | Änderung Konfiguration Barcodetyp ändern | Änderung Roma.ini Änderung Barcode auf codabar | SWO-520870 | Re classified RMED | planned SWO |  |  | Request | 22.03.2019, mr Stephan, Tel: 03477171165: Barcode reader to be changed- Type Codabar | Change configuration Change barcode type | Change Roma.ini Change barcode to codabar
No | Visit | Complaint | RAME | P02 dark data error. | Powered on analyzer and calibration failed with error 1002 P02 dark data error. | Replaced the signal data board 903-581. | SWO-521267 | Defect | HW | Basic Unit | PCB, Sensor interface | pO2, 1002 | P02 dark data error. | Powered on analyzer and calibration failed with error 1002 P02 dark data error. | Replaced the signal data board 903-581.
Yes | Visit | Complaint | RCHN | PCO2过高 | PCO2过高 | 解释原因，培训分析前误差 by jesson | SWO-522871 | [5 empty] | PCO2 too high | PCO2 too high | Explain the reason, training analysis before the error by jesson
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | PH,lac初始化失败，定标超范围 | 更换测试卡 | SWO-523680 | [5 empty] | Test card, reagent package failure | PH, lac initialization failure, calibration over range | Replace the test card
No | Visit | Complaint | RCHN | HCT值无法被测得，报错1260,1261,1264,1265 | HCT值无法被测得，报错1260,1261,1264,1265，借903-329一个 | 更换903-329后修复 by Andrew | SWO-520260 | Defect | HW | Aspiration Module | Needle Wash Unit | 1260, 1261, 1264, 1265 | HCT value cannot be measured, error 1260,1261,1264,1265 | HCT value cannot be measured, error 1260,1261,1264,1265, borrowed 903-329 one | Fix by Andrew after replacing 903-329
No | Visit | Complaint | RCHN | 气流温度超出范围，报错1070 | 气流温度超出范围，报错1070.需要借件848-037一个 |  | SWO-520486 | Defect | HW | Process Module | Compressor | 1070 | Airflow temperature out of range, error 1070 | Airflow temperature out of range, report error 1070.
No | Visit | Complaint | RLTD | Analyser memory chip not detected contact service displayed on analyser screen Dr Lewis Gough | Return visit, as on previous visit to fit analyser ID board that had arrived after fitting the ID board it had resolved the critical error 'Analyser memory chip not detected', but would not allow me to install a new SP or enter several menus. This happened very late on the day i did it, i left site to discuss and figure out what to do next to resolve the issues i had, on studying the ABL9 service manual i discovered the analyser had a factory reset option and wondered if this needed doing with the new analyser ID board fitted, and was also wondering about reinstalling the software. | Enter service and activate the factory reset option, this then allowed me to install a new SP. Reset the analyser up as per customer requirements, and chase up redline order for customers replacement SC. | SWO-520913 | [5 empty] | Analyser memory chip not detected contact service displayed on analyser screen Dr Lewis Gough | Return visit, as on previous visit to fit analyser ID board that had arrived after fitting the ID board it had resolved the critical error 'Analyser memory chip not detected', but would not allow me to install a new SP or enter several menus. This happened very late on the day i did it, i left site to discuss and figure out what to do next to resolve the issues i had, on studying the ABL9 service manual i discovered the analyser had a factory reset option and wondered if this needed doing with the new analyser ID board fitted, and was also wondering about reinstalling the software. | Enter service and activate the factory reset option, this then allowed me to install a new SP. Reset the analyser up as per customer requirements, and chase up redline order for customers replacement SC.
No | Visit | Complaint | RCHN | pco2 k na定标失败，报错0378 | pco2 k na定标失败，报错0378 | 更换pco2和k膜修复 by jesson | SWO-522726 | Assisted Replacement | Electrode | Membrane | pCO2 | 378 | pco2 k na bid failure, error 0378 | pco2 k na bid failure, error 0378 | Replace pco2 and k membrane repair by jesson
Yes | Visit | Complaint | RFRA | QC creat hors gamme | Historique remplacement et CQ | Contrôle des électrodes et de la chambre de mesure, RAS. Élimination des micros bulles dans les membranes. Nettoyage électrode fictive. Calibration, Ok, CQ, NOK. Récupération des données pour analyse au Danemark. Envoie des électrodes pour Essais au CH de Pont l'Abbé. Commande carte électronique pour une nouvelle intervention. | SWO-520968 | Re classified RMED | same day upSWO |  |  |  | QC creat off-range | Replacement history and CQ | Control of the electrodes and the measuring chamber, RAS. Elimination of micro bubbles in membranes. Fictional electrode cleaning. Calibration, Ok, CQ, NOK. Data recovery for analysis in Denmark. Send electrodes for Tests at the CH of Pont l'Abbé. Order electronic card for a new intervention.
No | Visit | Complaint | RFRA | QC creat hors gamme | Aucune | Remplacement de la chambre de mesure. Étalonnage des pompes, ajustement DL. Calibration et passage CQ | SWO-520969 | Defect | HW | Met II module | measuring chamber |  | QC creat off-range | No | Replacing the measuring chamber. Pump calibration, DL fit. Calibration and passage CQ
No | Visit | Complaint | RIND | Analyser is not working properly. | Check system massage and calibration report. observed that cal2 solution failed to calibrate. | give 2 point calibration 2 time but calibration failed. switched off analyser and on again then reinstalled Cal2 solution and give 2 point calibration 2 time. after that all parameter was active. | SWO-521752 | Assisted Replacement | Consumables | Solutions | Cal 2 |  | Analyser is not working properly. | Check system massage and calibration report. observed that cal2 solution failed to calibrate. | give 2 point calibration 2 time but calibration failed. switched off analyser and on again then reinstalled Cal2 solution and give 2 point calibration 2 time. after that all parameter was active.
No | Visit | Complaint | RAME | Lactate unstable and falling out every 2-3 days replaced membrane and | Replaced the 902-824 module | We have very erratic QC on Lactate on ABL 3. It is the newest one. As you can see the CV’s are very high and it was flagged on the WDC report. We have changed the membrane and electrode. But, it only stays good for a few days and trends back down. Replaced Electrolyte Module and customer will continue to monitor the analyzer. | SWO-521051 | Defect | HW | El/Met Module | El/Met Module |  | Lactate unstable and falling out every 2-3 days replaced membrane and | Replaced the 902-824 module | We have very erratic QC on Lactate on ABL 3. It is the newest one. As you can see the CV’s are very high and it was flagged on the WDC report. We have changed the membrane and electrode. But, it only stays good for a few days and trends back down. Replaced Electrolyte Module and customer will continue to monitor the analyzer.
No | Visit | Complaint | RDEU | 18.03.2019 Pf. Konrad 03616541310 meldet einen Fehler 600. Demo Option Key wird bald ablaufen. | nach Ablauf Demo Key keine EL mehr Meldung Key nicht installiert | Überprüfung Options Einstellung korrigiert Einlaufverfahren Lac membraniert , da QC Fehler | SWO-520872 | Setup | SW | RMED SW | RMED SW | Demo key, potentially user error | 18.03.2019 Pf. Konrad 03616541310 reports an error 600. Demo Option Key will expire soon. | after expiration Demo Key no more EL Message Key not installed | Review Options Setting corrected Entry procedure Lac membranes as QC error
No | Visit | Complaint | RSCH | Oxi Phasen Regulation Failed | 0269: Hemolyzer Phase regulation Failed | Hemolyzer Unit ausgetauscht und justiert tHb kal Funktionskontrolle i.o | SWO-520775 | Defect | HW | Oxi module | Hemolyzer | 269 | Oxi Phase Regulation Failed | 0269: Hemolyzer Phase regulation Failed | Hemolyzer Unit replaced and adjusted tHb kal Functional control i.o
No | Visit | Complaint | RLTD | Cartridge storage door fails to operate after restart, stepper motor chatters back and forth. | Cartridge storage door fails to operate after restart, stepper motor chatters back and forth. Already reloaded software, fitted new stepper and new front cabinet PCB. | Replaced main control PCB, reloaded software, reloaded service config, and tested. Stepper motor now operates correctly, tested all other funtions in service, no problems observed. | SWO-520774 | Defect | HW | Analyzer Control Module | PCB, Main Board |  | Cartridge storage door fails to operate after restart, stepper motor chatters back and forth. | Cartridge storage door fails to operate after restart, stepper motor chatters back and forth. Already reloaded software, fitted new stepper and new front cabinet PCB. | Replaced main control PCB, reloaded software, reloaded service config, and tested. Stepper motor now operates correctly, tested all other funtions in service, no problems observed.
Yes | Visit | Complaint | RDEU | 22.03.2019, Herr Stephan, Tel: 03477171165: Patientenbarcodes sollen gelesen werden- Barcode Codabar | Änderung Konfiguration Barcodetyp ändern | Änderung Barcodetyp codabar | SWO-520871 | Re classified RMED | planned SWO |  |  | Request | 22.03.2019, Mr. Stephan, Tel: 03477171165: Patient barcodes should be read- Barcode Codabar | Change configuration Change barcode type | Change barcode type codabar
Yes | Visit | Complaint | RFRA | Demande de RDV pour enlèvement de l'AQT90 de prêt et sauvegarde données et compteur Mme AGUERA et Mr BROUTIN souhaitent activer leur compte MyRadiometer | Préparation de l’aqt90 de prêt pour le transport afin de le renvoyer à Radiometer. | AQT90 de prêt du 15 $09 $2018 au 25 $03 $2019. Relevé de compteur fait ce jour 25 $03 $2019. Récupération et sauvegarde de toutes les données faites sur clé usb laissée au client (confie, toutes données, export toutes données .csv) Procédure long Time shutdown ok. Mise en caisse de l’aqt90 et demande faite à la logistique pour passage transporteur ok. Rapport envoyé par mail à uzes@bioaxiome.com vincent.broutin@bioaxiome.com | SWO-520850 | Re classified RMED | planned SWO |  |  | Request | Request for appointment for removal of the AQT90 loan and backup data and meter Ms. AGUERA and Mr. BROUTIN wish to activate their MyRadiometer account | Preparing the loan aqt90 for transport in order to send it back to Radiometer. | AQT90 loan from $15 $09 2018 to $25 $03 2019. Meter reading makes today $25 $03 2019. Recovering and backing up all data made on USB key left to the customer (confides, all data, export all data .csv) Long Time shutdown procedure ok. Cashing the aqt90 and request made to logistics for carrier pass ok. Report emailed to uzes@bioaxiome.com vincent.broutin@bioaxiome.com
No | Visit | Complaint | RIND | analyser giving startup error | mechanical error | Check system check in service mode found problem near incubator wheel , open front panel and clean dust from sensor near incubator wheel then again given startup now analyser ready. | SWO-524127 | Cleaned | HW | Incubator Module | Sensor, incubator |  | analyser giving startup error | mechanical error | Check system check in service mode found problem near incubator wheel , open front panel and clean dust from sensor near incubator wheel then again given startup now analyser ready.
No | Visit | Complaint | RAME | Electrode Sensor disconnected | Deliver the TCM | Checked everything worked ok | SWO-521192 | [5 empty] | Electrode Sensor disconnected | Deliver the TCM | Checked everything worked ok
No | Visit | Complaint | RKK | pO2 Complete test Error | pO2 Complete test Error | センサーインターフェースPCB交換 センサーシミュレーターによる動作確認 | SWO-521925 | Defect | HW | Basic Unit | PCB, Sensor interface | pO2 | pO2 Complete test Error | pO2 Complete test Error | Sensor Interface PCB Replacement Check operation by sensor simulator
No | Visit | Complaint | RAME | Time is off by one hour, getting errors on samples. | Customers Time is off by one hour, getting errors on samples. | Corrected time in control panel. | SWO-521900 | Setup | SW | RMED SW | RMED SW | time, 1h | Time is off by one hour, getting errors on samples. | Customers Time is off by one hour, getting errors on samples. | Corrected time in control panel.
No | Visit | Complaint | RDEU | Lösungen können nicht gepumpt werden/Bitte wenden sie sich an den Kundendienst | Lösungen können nicht gepumpt werden, | Einlassdichtung getauscht, Startphase, | SWO-522123 | Defect | HW | Inlet module | Inlet gasket |  | Solutions cannot be pumped/Please contact Customer Service | Solutions cannot be pumped, | Inlet seal replaced, start phase,
No | Visit | Complaint | RDEU | 22.03.2019 12:15 Fr. Holpel +49765129558: Bildschirmrahmen ist gerissen; Touch ohne Funktion | Bildschirm defekt, Touchscreen nicht arretiert, werkseitige Überspannung im Rahmen, Verbindungsstücke gebrochen | GEWÄHRLEISTUNG, Display komplett ersetzt, ,QC i.O. | SWO-521449 | Defect | HW | Display | Display | screws (defective screen) | 22.03.2019 12:15 Ms. Holpel +49765129558: Screen frame is cracked; Touch without function | Screen defective, touchscreen not locked, factory overvoltage in frame, connectors broken | WARRANTY, display completely replaced, ,QC i.O.
No | Visit | Complaint | RDEU | Bili Pat-Messungen werden nicht durchgeführt | bei einem Pat wird der Billi nicht mit ausgeben mit 713 | tHb bei diesem 24.1 g $dl und somit wird der Billi nicht ausgegeben. kurze Einweisung auf dem System gegeben | SWO-522124 | No fault found | No fault found | Oxi module | Bili | User training | Bili Pat measurements are not performed | at a Pat, the Billi will not spend with 713 | tHb at this 24.1 g $dl and thus the billi is not issued. short briefing given on the system
No | Visit | Complaint | RDAN | Eva Nielsen klager over fejl 1343. | Undersøgt; 1343+1324 | Waste studs + Slange + oxi slange og inlet pakning udskiftet, væskesensor renset og kalibreret, slutafprøvet. | SWO-522121 | Defect | HW | Basic Unit | PM Kit | 1343, 1324 | Eva Nielsen complains about error 1343. | Investigated; 1343 + 1324 | Waste Studs + Hose + oxi hose and inlet packing replaced, liquid sensor cleaned and calibrated, final tested.
No | Visit | Complaint | RLTD | Reported by Haval Ozgun 07737 457222, please contact Frederick on 07789 864100. | Cause of leak traced to failed Oximetry pinch valve tube - replaced tube. Tested fluid transport in service now passes all tests. Also noted battery charge showing as 0% incorrectly, removed battery unit and reset. Battery now showing 100% charge correctly. Restarted analyser and verified operation - pCO2 failed to calibrate. Fitted new Sensor Cassette and restarted. Calibration and quality controls passed, except QCs for Glucose. Analyser returned to ITU and Biochem spare returned to Biochem. ITU staff advised that Glucose would be ready in the next hour. Analyser ready for use on all other parameters. | Replaced Oximetry pinch valve tube. Replaced Sensor Cassette. | SWO-521778 | Defect | HW | Oxi module | Pinch valve tube | 1184 | Reported by Haval Ozgun 07737 457222, please contact Frederick on 07789 864100. | Cause of leak traced to failed Oximetry pinch valve tube - replaced tube. Tested fluid transport in service now passes all tests. Also noted battery charge showing as 0% incorrectly, removed battery unit and reset. Battery now showing 100% charge correctly. Restarted analyser and verified operation - pCO2 failed to calibrate. Fitted new Sensor Cassette and restarted. Calibration and quality controls passed, except QCs for Glucose. Analyser returned to ITU and Biochem spare returned to Biochem. ITU staff advised that Glucose would be ready in the next hour. Analyser ready for use on all other parameters. | Replaced Oximetry pinch valve tube. Replaced Sensor Cassette.
No | Visit | Complaint | RLTD | Kieran on 0141 9515917 | Analyser battery showing 0% but funtioning correct | Removed battery module and reset by removing PCB jumper switch and replacing Battery % now functioning correct | SWO-521530 | Adjustment | HW | Battery module | Battery module | 0% | Kieran on 0141 9515917 | Analyser battery showing 0% but funtioning correct | Removed battery module and reset by removing PCB jumper switch and replacing Battery % now functioning correct
No | Visit | Complaint | RSCH | Blank | Défectuosité de la fixation de l’écran tactile (garantie) | Remplacé l’unité complète de l’écran (type 910-353) Calibration e.o CQ e.o | SWO-521549 | Defect | HW | Display | Display |  | Blank | Touch screen attachment defect (guarantee) | Replaced the full screen unit (type 910-353) Calibration e.o CQ e.o
No | Visit | Complaint | RAME | Analyzer will not power on | 1) Analyzer will not power on. 2) Inlet probe bent. | 1) Replaced power supply. 2) Replaced inlet probe. | SWO-521704 | Defect | HW | Power Supply | Power supply |  | Analyzer will not power on | 1) Analyzer will not power on. 2) Inlet probe bent. | 1) Replaced power supply. 2) Replaced inlet probe.
No | Visit | Complaint | RFRA | Reste en erreur sonde échantillon ouverte suite à une erreur de manipulation : La technicienne a voulu remettre en place le buvard en sortant le plateau durant la calibration. Le remplacement de la sonde est requise. | ctrl trop bas | chagement de pipette , aligmenment preconise par constructeur calibration ctrl passage de CQ (valeur de 285 a 291 . | SWO-521447 | [5 empty] | Remains in error open sample probe due to a handling error: The technician wanted to put the blotter back in place by taking out the tray during calibration. Replacement of the probe is required. | ctrl too low | pipette, aligmently preconise by manufacturer calibration ctrl passage of CQ (value of 285 to 291 .
No | Visit | Complaint | RCHN | 接口板腐蚀 | 接口板腐蚀,借902-972一个。 | 更换902-972后修复 by jesson | SWO-521432 | [5 empty] | Interface plate corrosion | Interface plate corrosion, borrow 902-972 one. | Repair by jesson after replacement 902-972
No | Visit | Complaint | RKK | 新品のガスを取り付けてキャルを実施するとエラー発生。残量が少ないガスだと通常に稼働。 エラーが出たガスを他装置に取り付けたが通常に使用可能。 |  |  | SWO-521351 | [5 empty] | An error occurs when a new gas is installed and the car is carried out. If the remaining amount is low, it will operate normally. The gas that failed was installed in another device, but it can be used normally.
No | Visit | Complaint | RIND | Crea calibration sensitivity low. | Found the clot in chamber and air bubbles in chamber.need to clean. | Found the clot in chamber and membrane then run the instrument its working good. | SWO-521486 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway | Crea sens low | Crea calibration sensitivity low. | Found the clot in chamber and air bubbles in chamber.need to clean. | Found the clot in chamber and membrane then run the instrument its working good.
No | Visit | Complaint | RIND | Calibration error | Checked the data logs crea calibration sensitivity out of range | Performed decontamination and replaced crea membrane | SWO-523576 | Assisted Replacement | Electrode | Membrane | Crea | Sens | Calibration error | Checked the data logs crea calibration sensitivity out of range | Performed decontamination and replaced crea membrane
No | Visit | Complaint | RPAC | Thb cals failing | Analyser giving low values on the tHb calibrator, when run as a sample, after the tHb Cal. | The analyser had a Complete OXI Test, all fine. However, the Halogen Lamp was a bit high, this was adjusted. All relevant checks were made as per the service manual, but this still did not resolve the issue completely. The Hemolizer when performing one of the calibrations, was found to certainly not sound fine, it was interminttantly hemolizing. It was decided that the Hemolizer required replacement. The analyser was powered down and the Hemolizer was replaced. Again, all teh relevant adjustments and tests were performed, as per teh service manual. Once all calibrations were performed, along with the tHb calibrations, QC Levels 1 & 3 were performed. All now seems to be functioning fine. The Hemolizer that is being returned is 902-435R4803N006. | SWO-521217 | Defect | HW | Oxi module | Hemolyzer | tHB Cal | Thb cals failing | Analyser giving low values on the tHb calibrator, when run as a sample, after the tHb Cal. | The analyser had a Complete OXI Test, all fine. However, the Halogen Lamp was a bit high, this was adjusted. All relevant checks were made as per the service manual, but this still did not resolve the issue completely. The Hemolizer when performing one of the calibrations, was found to certainly not sound fine, it was interminttantly hemolizing. It was decided that the Hemolizer required replacement. The analyser was powered down and the Hemolizer was replaced. Again, all teh relevant adjustments and tests were performed, as per teh service manual. Once all calibrations were performed, along with the tHb calibrations, QC Levels 1 & 3 were performed. All now seems to be functioning fine. The Hemolizer that is being returned is 902-435R4803N006.
No | Visit | Complaint | RIND | M/c not starting up | Checked and found Windows is corrupted. | Reinstalled Windows7, reinstalled ABL800 SW v6.18 successfully. Performed setup as per customer's requirement. 2 point calibration performed successfully. | SWO-524053 | Defect | SW | External SW | Windows |  | M/c not starting up | Checked and found Windows is corrupted. | Reinstalled Windows7, reinstalled ABL800 SW v6.18 successfully. Performed setup as per customer's requirement. 2 point calibration performed successfully.
No | Visit | Complaint | RIND | Maintainance required | Maintenance required | Cleaned the instrument , checked the system messages , clean cycle reached the maximum count , did cleaning and then run samples , customer happy with results. | SWO-522093 | Assisted Replacement | Consumables | System clean | System clean | User Error | Maintainance required | Maintenance required | Cleaned the instrument , checked the system messages , clean cycle reached the maximum count , did cleaning and then run samples , customer happy with results.
No | Visit | Complaint | RDEU | 11.03.2019, Heike, Tel: 015170824649: Disk boot Fehler 12.03.2019, Sr. Carola +4917614022104. wartet noch auf Kontaktaufnahme | nach PC-Modulwechsel keine Aqure-Verbindung, WLAN-Stick nicht installiert | Datensicherung, Schreibschutz deaktiviert, Neustart, WLAN-Treiber installiert und konfiguriert, Netzwerkeinstellungen (IP: 192.168.71.12, Subnet:255.255.255.0, Gateway: 192.168.71.1)korrigiert, Schreibschutz aktiviert, Neustart, Startphase abgewartet, | SWO-522265 | Setup | SW | RMED SW | RMED SW |  | 11.03.2019, Heike, Tel: 015170824649: Disk boot Error 12.03.2019, Sr. Carola +4917614022104. is still waiting to get in touch | No Aqure connection after PC module change, Wi-Fi stick not installed | Data backup, read-only disabled, restart, Wi-Fi driver installed and configured, network settings (IP: 192.168.71.12, subnet:255.255.255.0, gateway: 192.168.71.1) corrected, read-only enabled, restart, start-up phase waited,
No | Visit | Complaint | RDEU | keine WLAN-Netzwerkverbindung | keine WLAN-Verbindung | Datensicherung, Schreibschutz deaktiviert, Neustart, WLAN-Treiber deinstalliert und wieder neu installiert, Konfiguration, Netzwerkeinrichtung, Ping aktiviert, Powermanagement deaktiviert, Schreibschutz aktiviert, Neustart, Startphase, Kal und QK ok | SWO-522266 | Setup | SW | RMED SW | RMED SW |  | No Wi-Fi network connection | No Wi-Fi connection | Data backup, read-only disabled, restart, Wi-Fi driver uninstalled and reinstalled, configuration, network setup, ping enabled, power management disabled, read-only enabled, restart, boot phase, cal and QK ok
No | Visit | Complaint | RDEU | Gerät ist in Aqure aber kein Schnellzugriff | Meldung von netop, dass ein Ordner nicht geschrieben werden kann | Softwareinstallation, komplett und nur netop, anschließend immer noch die Fehlermeldung, Kunde hat aber bestätigt, dass der Fernzugriff funktioniert | SWO-522853 | Defect | SW | RMED SW | RMED SW | Aqure quick assess | Device is in Aqure but not quick access | Message from netop that a folder cannot be written | Software installation, complete and only netop, then still the error message, but customer has confirmed that the remote access works
No | Visit | Complaint | RLTD | Cecelia - 07479 241036 | Collar at rear of inlet module has broken such that retaining clip falls off. Replaced inlet module - tested operation of inlet , all ok. | Replaced inlet module. | SWO-522303 | Defect | HW | Auto Inlet module | Auto Inlet module |  | Cecelia - 07479 241036 | Collar at rear of inlet module has broken such that retaining clip falls off. Replaced inlet module - tested operation of inlet , all ok. | Replaced inlet module.
No | Visit | Complaint | RSAF | Leaking on the hemolyzer tubing | Inspected the analyzer for any visible crystals or leaks | Replaced the hemolyzer tubing | SWO-522406 | Defect | HW | Oxi module | Pinch valve tube |  | Leaking on the hemolyzer tubing | Inspected the analyzer for any visible crystals or leaks | Replaced the hemolyzer tubing
No | Visit | Complaint | RLTD | Sarah Jordan on 0845 155 5000 ex 71602 or Annie on 0798449505 26/03 Further call received from Gavin Cooper 07930-102362 - error not resolved. Error message 1393 (inlet not mounted) | Customer has replaced inlet gasket and probe. Keeps going into intervension required error 1391 inlet is not in correct position. | Found Plastic part broken between the two location pins on the back of inlet module. fitted service spare module and test ok. SC and SP replaced as due analyser starting up. Customer will monitor analyser. | SWO-522268 | Defect | HW | Auto Inlet module | Auto Inlet module | 1393, 1391 | Sarah Jordan on 0845 155 5000 ex 71602 or Annie on 0798449505 26/03 Further call received from Gavin Cooper 07930-102362 - error not resolved. Error message 1393 (inlet not mounted) | Customer has replaced inlet gasket and probe. Keeps going into intervension required error 1391 inlet is not in correct position. | Found Plastic part broken between the two location pins on the back of inlet module. fitted service spare module and test ok. SC and SP replaced as due analyser starting up. Customer will monitor analyser.
No | Visit | Complaint | RDEU | 26.02.2019 9:20 H. Haberkamp 02083092181: Interventionsmeldung Endlosschleife; kein Material mehr vor Ort | Lieferschwierigkeiten: Sensor Kassette ABL Crea Urea Neue Sensor Kassette defekt. | Situation mit Kunden besprochen: Herr Haberkamp ist weiterhin sehr unzufrieden, weil zu den Kundenbeschwerden in der letzten Woche, nur Teilweise eine Stellungnahme auf die in der beschwerde aufgezeigten Probleme erfolgte. Herr Haberkamp und Herr Reinicke genannt hatte, nicht eingegangen worden ist, bzw. nur zum Teil eine Stellungnahme erfolgt ist. Durchgeführte Tätigkeiten: Gerät wieder auf Crea $Urea konfiguriert. SC und LP durch Kundeneigene Betriebsmittel eingesetzt. Gerät kalibriert und QC durchgeführt. Die QC Werte liegen bei Lösung C S29250 beim pCO² ausserhalb der Herstellergrenzen. Gerät mehrfach kalibriert, LP erneuert, QC Ergebniss pCO² bei Lösung C S29250 weiterhin ausserhalb der Herstellergrenzen. SC durch Kundeneigene ersetzt. | SWO-523670 | Assisted Replacement | SC | SC | pCO2 | Ki QC C | 26.02.2019 9:20 Haberkamp 02083092181: Intervention notification Endless loop; no more material on site | Delivery difficulties: Sensor cassette ABL Crea Urea New sensor cassette defective. | Situation discussed with customers: Mr Haberkamp remains very dissatisfied because the customer complaints last week were only partially commented on the problems identified in the complaint. Mr Haberkamp and Mr Reinicke have not been received, or only partially have been given an opinion. Activities carried out: device is configured back to Crea $Urea. SC and LP are used by customer-owned equipment. calibrated and QC performed. The QC values for solution C S29250 for the pCO2 are outside the manufacturer's limits. Device multiple calibrated, LP renewed, QC result pCO2 at solution C S29250 still outside the manufacturer's limits. SC replaced by customer-owned.
No | Visit | Complaint | RKOR | Software error | . | 3.4 Software reinstall. Setting | SWO-524121 | Defect | SW | RMED SW | RMED SW |  | Software error | . | 3.4 Software reinstall. Setting
No | Visit | Complaint | RAME | Unable to pump solutions error code 1169 | Unable to pump solutions error code 1169, customer changed 2 solution packs and 2 sensor cassettes in one week. Old valve 908-564R0721N0019 New valve 903-129R0521N002 | Replaced all tubing and parts from annual maintenance kit and also replaced pinch valve. Did 20 back flushes and Replaced solution pack and sensor cassette from customers stock. | SWO-522443 | Defect | HW | Oxi module | Pinch Valve | 1169 | Unable to pump solutions error code 1169 | Unable to pump solutions error code 1169, customer changed 2 solution packs and 2 sensor cassettes in one week. Old valve 908-564R0721N0019 New valve 903-129R0521N002 | Replaced all tubing and parts from annual maintenance kit and also replaced pinch valve. Did 20 back flushes and Replaced solution pack and sensor cassette from customers stock.
No | Visit | Complaint | RFRA | Reste en erreur sonde échantillon ouverte suite à une erreur de manipulation : La technicienne a voulu remettre en place le buvard en sortant le plateau durant la calibration. Le remplacement de la sonde est requise. | retour automate | remballage automate mise en place des securité | SWO-522145 | [5 empty] | Remains in error open sample probe due to a handling error: The technician wanted to put the blotter back in place by taking out the tray during calibration. Replacement of the probe is required. | automated return | automated repackaging setting up security
No | Visit | Complaint | RCHN | 定标失败 | 定标失败，借944-212（15631）一个 | 重新定标后修复 by jesson | SWO-521369 | Assisted Replacement | Consumables | Test Kit | Cartridge, cal | 5257 | Bid failed | Failed to bid, borrow 944-212 (15631) one | Fix by jesson after re-calibration
No | Visit | Complaint | RCHN | 仓门无法关闭 4548 | 仓门无法关闭 4548,借903-028一个。 | 更换903-028后修复 by jesson | SWO-521341 | Defect | HW | Cover and Chassis | PCB, cabinet top | 4548, Inlet door | Warehouse door cannot be closed 4548 | The warehouse door cannot be closed 4548, borrowing 903-028 one. | Repair by jesson after replacement 903-028
No | Visit | Complaint | RCHN | 卡盘，报错5784 | 卡盘，报错5784，未更换备件修复 | 重新安装测试卡后修复 by jesson | SWO-522028 | Defect | SW | RMED SW | RMED SW | 5784 | Card, 5784 | Card, error 5784, unreplaced spare parts repair | Fix by jesson after reinstalling test card
Yes | Visit | Complaint | RDEU | Austausch def. Elektroden | Elektrodentausch | neue Ref.- (alt: 37-05) u. pH-Elektrode (alt: EP-03) an die MT übergeben; Altelektroden zurückgenommen | SWO-523230 | Re classified RMED | Logistical |  |  |  | Exchange def. Electrodes | Elektrodentausch | new Ref.- (old: 37-05) and pH electrode (old: EP-03) handed over to mt; Old electrodes withdrawn
No | Visit | Complaint | RAME | Customer states a significant leak is coming from the side of the analyzer. | Verified the customer's stated issue. The leak was coming from the spout of the pH chamber. Removed all the electrodes from the pH measuring chamber and found the chamber to be filled with rinse solution. Found all the electrodes to have their respective sealing ring but found the pH electrode sealing ring stuck to the bottom of the chamber. Replaced the pH sealing (provided by the customer) and performed several rinse programs with a smaller leak coming from the spout of the pH chamber. | Dried the pH chamber thoroughly and replaced the waste tubings. Performed several rinse programs with no leaks detected at this time. | SWO-522355 | Defect | HW | Fluid manifold | Waste tube |  | Customer states a significant leak is coming from the side of the analyzer. | Verified the customer's stated issue. The leak was coming from the spout of the pH chamber. Removed all the electrodes from the pH measuring chamber and found the chamber to be filled with rinse solution. Found all the electrodes to have their respective sealing ring but found the pH electrode sealing ring stuck to the bottom of the chamber. Replaced the pH sealing (provided by the customer) and performed several rinse programs with a smaller leak coming from the spout of the pH chamber. | Dried the pH chamber thoroughly and replaced the waste tubings. Performed several rinse programs with no leaks detected at this time.
No | Visit | Complaint | RIND | Service Visit. | Need to replace the display of the analyzer. | Shut down the analyzer, remove the stand by display and connect the new one, start the analyzer, after successful initializing wait for auto 2 point calibration, successfully done, do the protein removal and cleaning in auxiliary program, successfully done, analyzer handed over to staff in well working condition. | SWO-522906 | Defect | HW | Display | Display |  | Service Visit. | Need to replace the display of the analyzer. | Shut down the analyzer, remove the stand by display and connect the new one, start the analyzer, after successful initializing wait for auto 2 point calibration, successfully done, do the protein removal and cleaning in auxiliary program, successfully done, analyzer handed over to staff in well working condition.
No | Visit | Complaint | RAME | Analyzer is giving pump errors | Analyzer giving pump cal errors due to asymmetry. | Replaced measuring chamber, pump rotor and pump tubing(customer stock). Analyzer is now working properly. | SWO-522486 | Defect | HW | pH/BG Module | measuring chamber |  | Analyzer is giving pump errors | Analyzer giving pump cal errors due to asymmetry. | Replaced measuring chamber, pump rotor and pump tubing(customer stock). Analyzer is now working properly.
No | Visit | Complaint | RAME | Analyzer will not Power On | Customer reports that analyzer shut down overnight as a result of a possible power surge and will not restart. A second analyzer also shut down but was able to power up. | Replaced the power supply. Analyzer restarted normally and went through the full initialization and start up process without error. Analyzer is within acceptable operating specifications. The serial # printed on the removed power supply’s pcb is 910-370R00044991 | SWO-522463 | Defect | HW | Power Supply | Power supply |  | Analyzer will not Power On | Customer reports that analyzer shut down overnight as a result of a possible power surge and will not restart. A second analyzer also shut down but was able to power up. | Replaced the power supply. Analyzer restarted normally and went through the full initialization and start up process without error. Analyzer is within acceptable operating specifications. The serial # printed on the removed power supply’s pcb is 910-370R00044991
Yes | Visit | Complaint | RIND | Routine service | NA | Checked the instrument , calibration data , activity log and system messages , no errors found , met doctor discussed about the instrument , customer happy with the instrument . | SWO-523662 | Re classified RMED | planned SWO |  |  |  | Routine service | In | Checked the instrument , calibration data , activity log and system messages , no errors found , met doctor discussed about the instrument , customer happy with the instrument .
No | Visit | Complaint | RDEU | 26.03.2019, Herr Wadwnbeck, Tel: 09417822350: Gasdruck Offset Spannung außer Bereich und Nullstrom Fehler bei pO2 | ABL Datei Archiv überprüft - Kal. $ QC $ Aktivitäten Nulstromfehler bei pO2 - Null anfänglich bei 82,.. mmHg - letzte Kal. bei 8,5 mmHg Gas-Offsetspannungsfehler - wird bei Eintreffen am ABL nicht mehr angezeigt | ABL im Service überprüft, Nullstrom Kal. für pO2, Glu und Lac durchgeführt, Gasoffsetspannung im Service überprüft - Gasflow und dazugehörige Spannungen sehr gut, LS Justage und Pumpen Kal. durchgeführt, Gas- und Spüllösungsventil im pH $BG Modul ausgebaut, gereinigt, auf Leichtgängigkeit überprüft, Membrane beider Ventile getauscht, pO2 neu membraniert, Messkammer gereinigt, Glasfläche der pO2 Elektrode vorsichtig gereinigt, Betriebsmittel (Spüllösung $Kal.1 und Kal.2 Lösung) erneuert - Füllstandsanzeige gelb, mehrmals kalibriert, System- und Funktionskontrolle in Ordnung | SWO-522204 | Defect | HW | pH/BG Module | Valves |  | 26.03.2019, Mr. Wadwnbeck, Tel: 09417822350: Gas pressure offset voltage out range and zero current error at pO2 | ABL File Archive Checked - Kal. • QC - Activities Nulcurrent error at pO2 - zero initially at 82,.. mmHg - last cal. at 8.5 mmHg Gas offset voltage error - is no longer displayed when arriving at the ABL | ABL checked in service, zero current Cal. for pO2, Glu and Lac performed, gas offset voltage checked in service - gas flow and associated voltages very good, LS adjustment and pumps Kal. gas and rinsing solution valve in the pH $BG module, cleaned, checked for smoothness, diaphragm of both valves replaced, pO2 re-membraneed, measuring chamber cleaned, glass surface of pO2 electrode carefully cleaned, operating equipment (flushing solution Kal.1 and Kal.2 solution) renewed - level indicator yellow, calibrated several times, system and function control in order
Yes | Visit | Complaint | RFRA | Faux positif TnT - Techline RFRA 838603725 - Ticket 559 Faux positif sur TnT à 41 ; deuxieme tube 3h plus tard, résultat négatif ; reprise du premier tube, résultat négatif ; reprend le deuxième, résultat négatif. | Aucune | Inspection de l'analyseur suite AU faux positif TnT du 19 $03. Remplacement de l'aiguille et de la cup wash unit. Nettoyage de la process wheel. Remise en route, passage 5 blanc et CQ, OK. Appareil qualifié conformément aux spécifications RADIOMETER. | SWO-522504 | Re classified RMED | planned SWO |  |  | Techline RFRA 838603725 | False positive TnT - Techline RFRA 838603725 - Ticket 559 False positive on TnT at 41; second tube 3 h later, negative result; resumption of the first tube, a negative result; the second, negative result. | No | Inspection of the analyzer following the false positive TnT of $19$03. Replacing the needle and cup wash unit. Cleaning the process wheel. Restarted, Passage 5 White and CQ, OK. Qualified device to RADIOMETER specifications.
No | Visit | Complaint | RLTD | Helen Hobba 07742 488150 | The analyser was giving Solution Pack not Accessable , No Tests Remaining & FTC 105 errors when trying to install a new solution pack. Another solution pack was tried but the same errors were still present. When the covers were removed I could see that one side of Pump 5 (P5) was disconnected so it was reconnected and the analyser started up with the covers off. During the start-up the top side of P5 disconnected again so it looks like there was a blockage somewhere in the system. The side of P5 that popped off was connected to the wash block so this was checked and found to be blocked. | The blockage in the wash block was remeoved and the analyser reassembled and allowed to start up without the covers on. This time P5 didn't pop off and the analyser started up ok. No tests were available so 1 cartridge of TnT, CRP & D-Dimer were installed on the analyser. D-Dimer was lock so I left it like that. | SWO-522395 | Cleaned | HW | Aspiration Module | Needle Wash Unit | FTC-105 | Helen Hobba 07742 488150 | The analyser was giving Solution Pack not Accessable , No Tests Remaining & FTC 105 errors when trying to install a new solution pack. Another solution pack was tried but the same errors were still present. When the covers were removed I could see that one side of Pump 5 (P5) was disconnected so it was reconnected and the analyser started up with the covers off. During the start-up the top side of P5 disconnected again so it looks like there was a blockage somewhere in the system. The side of P5 that popped off was connected to the wash block so this was checked and found to be blocked. | The blockage in the wash block was remeoved and the analyser reassembled and allowed to start up without the covers on. This time P5 didn't pop off and the analyser started up ok. No tests were available so 1 cartridge of TnT, CRP & D-Dimer were installed on the analyser. D-Dimer was lock so I left it like that.
No | Visit | Complaint | RAME | Disk Boot Failure | Verified setup to be identical to TCM R0347N011. Ran test on forearm and results were stable at 39mmHg. Swapped gas module and tested again with same electrode. Results were stable at 39 mmHg. | Monitor results are in normal range as electrode placed on forearm for 30 minute test. No failure is noted. | SWO-522393 | [5 empty] | Disk Boot Failure | Verified setup to be identical to TCM R0347N011. Ran test on forearm and results were stable at 39mmHg. Swapped gas module and tested again with same electrode. Results were stable at 39 mmHg. | Monitor results are in normal range as electrode placed on forearm for 30 minute test. No failure is noted.
No | Visit | Complaint | RPOL | Wygięcie sondy aspiracyjnej | Serwis stwierdza wygiętą sondę aspiracyjną. | Wymiana sondy aspiracyjnej oraz uszczelki wlotu. | SWO-522435 | Defect | HW | Auto Inlet module | Probe bend |  | Aspiration probe deflection | The service finds a curved aspiration probe. | Replacement of aspiration probe and inlet gasket.
No | Visit | Complaint | RPOL | Złamana igła | Testy serwisowe | Wymiana igły, testy serwisowe. | SWO-522370 | Defect | HW | Aspiration Module | Needle Assembly |  | Broken Needle | Service Tests | Replacing the needle, service tests.
No | Visit | Complaint | RPOL | Błąd incubator wheel | Przegląd logów, testy serwisowe. | Naprawa wyrwanego ząbka w storage wheel, testy serwisowe. | SWO-522128 | Defect | HW | Storage Module | Cartridge guide |  | Błąd incubator wheel | Log overview, service tests. | Repair of the broken tooth in the storage wheel, service tests.
No | Visit | Complaint | RAME | Inlet open errors | Inlet probe bent. | Replace inlet probe and inlet gasket with holder. Perform inlet calibration and all fluidic tests. | SWO-523180 | Defect | HW | Auto Inlet module | Probe bend |  | Inlet open errors | Inlet probe bent. | Replace inlet probe and inlet gasket with holder. Perform inlet calibration and all fluidic tests.
No | Visit | Complaint | RLTD | Ilias Haqim 0330 400 4333 Ext 6672 | Keeps returning to intervention required despite trouble shooting Vacuum error leak detected. | Run service pressure test failing vacuum test to high. Flushed system with out SC with protein remover for 5 minuets . Flushed with water and now getting a good flow of water. Refitted SC and SP and run service test again and all passed ok. Calibration QC passed ok Clot suspected. | SWO-523070 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway | Clot | Ilias Haqim 0330 400 4333 Ext 6672 | Keeps returning to intervention required despite trouble shooting Vacuum error leak detected. | Run service pressure test failing vacuum test to high. Flushed system with out SC with protein remover for 5 minuets . Flushed with water and now getting a good flow of water. Refitted SC and SP and run service test again and all passed ok. Calibration QC passed ok Clot suspected.
No | Visit | Complaint | RFRA | L'économiser d'écran à des sursauts. | Vérification configuration. | Screen saveur désactivé sur Windows 7 ok Sauvegarde des données Conseil: si écran noir : retirer clés usb puis redémarrer. | SWO-523227 | No fault found | No fault found | No fault found | No fault found | Black screen | The screen save to bursts. | Checking configuration. | Screen flavor disabled on Windows 7 ok Backing up data Tip: if black screen: remove USB keys and then restart.
No | Visit | Complaint | RLTD | The analyser is giving a Clot Suspected in Inlet error every night. | After changing the Wet Section Control PCB the analyser was still giving a clot detected in the inlet every time it did the clot check at midnight. | As the inlet module, inlet connector gasket and the wet section pcb had all been replaced with no resolution to the problem the sensor cassette bay was replaced to try and resolve the problem. The original wet section control PCB was refitted as it was not the cause of the problem. The analyser was reassembled and the fluid transport service programs run and they all passed. The clot detect $ remove program was run 6 times and it passed on each occasion with the values remaining fairly constant. The analyser was then allowed to complete it start up and it did with no errors. Will continue to monitor to see if the fault has been resolved. | SWO-523205 | Defect | HW | Basic Unit | Sensor Bay Unit | Clot detection, midnight | The analyser is giving a Clot Suspected in Inlet error every night. | After changing the Wet Section Control PCB the analyser was still giving a clot detected in the inlet every time it did the clot check at midnight. | As the inlet module, inlet connector gasket and the wet section pcb had all been replaced with no resolution to the problem the sensor cassette bay was replaced to try and resolve the problem. The original wet section control PCB was refitted as it was not the cause of the problem. The analyser was reassembled and the fluid transport service programs run and they all passed. The clot detect $ remove program was run 6 times and it passed on each occasion with the values remaining fairly constant. The analyser was then allowed to complete it start up and it did with no errors. Will continue to monitor to see if the fault has been resolved.
No | Visit | Complaint | RAME | Customer getting 0581 error during sampling | Customer getting 0581 errors on patient samples. | Performed cleaning program. Replaced Hemolyzer and pinch valve. Put parts from PM kit in analyzer. | SWO-523088 | [5 empty] | Customer getting 0581 error during sampling | Customer getting 0581 errors on patient samples. | Performed cleaning program. Replaced Hemolyzer and pinch valve. Put parts from PM kit in analyzer.
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | LAC定标超范围 | 更换测试卡 | SWO-523682 | [5 empty] | Test card, reagent package failure | LAC calibration over range | Replace the test card
No | Visit | Complaint | RCHN | 卡读不出 | 测试卡信息读不出（未借用备件） | 接触不良，拆机重装后修复 by jesson | SWO-523693 | Adjustment | HW | Basic Unit | PCB, Sensor interface | contacts | The card can't read. | Test card information cannot be read (spare parts not borrowed) | Poor contact, re-installed after re-installing by jesson
No | Visit | Complaint | RIND | PO2 results not coming for samples | PO2 results not coming for samples | Checked the calibration log ,calibration is passed , checked the results , found po2 parameter is not calculated fro samples , error showing no endpoint , did rinse , did manual calibration , run samples but the problem not solved .suspecting sensor cassette failure , Need to check with QC | SWO-523876 | [5 empty] | PO2 results not coming for samples | PO2 results not coming for samples | Checked the calibration log ,calibration is passed , checked the results , found po2 parameter is not calculated fro samples , error showing no endpoint , did rinse , did manual calibration , run samples but the problem not solved .suspecting sensor cassette failure , Need to check with QC
No | Visit | Complaint | RCHN | PCO2零点错误 | PCO2零点错误，借945-613一个 | 更换945-613后修复 by jesson | SWO-521290 | Assisted Replacement | Electrode | Electrode | pO2 | 408 | PCO2 zero error | PCO2 zero error, borrow 945-613 one | Repair by jesson after replacement 945-613
No | 4 | Y | RCZE | Blank | Obnovenie funkcnosti windows VPN |  | D5E348A3D4324EF6 | Setup | SW | RMED SW | His/Lis | VPN | Blank | Obnovenie funkcnosti windows VPN
No | Visit | Complaint | RDEU | pCO2 bei QK S7745 häufig zu niedrig, Membran wurde mehrfach erneuert | pCO2 bei QK S7745 häufig niedrig oder zu niedrig | Reinigungsarbeiten, Fehlersuche, Messkammer und Elektrode sowie deren Kontakte kontrolliert und gereinigt, pCO2-Elektrode (alt: BY-01) ersetzt, Kal und QK ok | SWO-522828 | Assisted Replacement | Electrode | Electrode | pCO2 | QC S7745 too low | pCO2 often too low at QK S7745, membrane has been renewed several times | pCO2 at QK S7745 often low or too low | Cleaning work, troubleshooting, measuring chamber and electrode as well as their contacts controlled and cleaned, pCO2 electrode (old: BY-01) replaced, Kal and QK ok
No | Visit | Complaint | RAME | No results on pleural pH samples | Confirmed issue when onsite. Discovered in syringe setup, pH is not shown in measured parameters box. | Shutdown and rebooted analyzer. Requested for software key and reloaded software key. pH now showing up in measured parameters box. Was able to run a pleural fluid sample with result successfully without any error. | SWO-523299 | Setup | SW | RMED SW | Demo key |  | No results on pleural pH samples | Confirmed issue when onsite. Discovered in syringe setup, pH is not shown in measured parameters box. | Shutdown and rebooted analyzer. Requested for software key and reloaded software key. pH now showing up in measured parameters box. Was able to run a pleural fluid sample with result successfully without any error.
No | Visit | Complaint | RPOL | Druk niemożliwy | Serwis stwierdza błąd sterownika drukarki. | Usunięcie przyczyny problemu, ponowny rozruch urządzenia, test wydruku. | SWO-522797 | Adjustment | HW | Printer Module | Printer Module |  | Printing not possible | The service finds a printer driver error. | Remove the cause of the problem, reboot the device, test the print.
No | Visit | Complaint | RDEU | pCO2 im QC häufig zu niedrig | Häufig pCO2 Fehler in der QC Messung, Level C | Software upgrade 3.4MR1 durchgeführt | SWO-524237 | Defect | SW | RMED SW | pCO2 | SW 3.4 MR1 | pCO2 in QC often too low | Common pCO2 error in QC measurement, level C | Software upgrade 3.4MR1 performed
No | Visit | Complaint | RAME | Leak detected. | Leak detected. | Blood Clot seen in analyzer. Replaced all tubing in yearly service kit. Completed all service tests. | SWO-523868 | Defect | HW | Basic Unit | PM Kit | Clot | Leak detected. | Leak detected. | Blood Clot seen in analyzer. Replaced all tubing in yearly service kit. Completed all service tests.
No | Visit | Complaint | RESP | Fuga detectada | Chequeo de datos del analizador. | Sustitución de varias piezas del analizador, limpieza del sistema observándose que está atascada la sonda de entrada de muestras, calibración y test controles de calidad funcionando correctamente. | SWO-523698 | Defect | HW | Basic Unit | PM Kit | 1184 | Leak detected | Analyzer data check. | Replacing multiple analyzer parts, cleaning the system noting that the sample input probe is jammed, calibration and test quality controls working properly.
No | Visit | Complaint | RAME | Netop failure | Netop remote failing to connect | Loaded fresh 3.4 software. Restarted Netop and verified remote access was working with customer in Aqure | SWO-523923 | Defect | SW | RMED SW | His/Lis | remote access | Just failure | Netop remote failing to connect | Loaded fresh 3.4 software. Restarted Netop and verified remote access was working with customer in Aqure
No | Visit | Complaint | RAME | Sensor temp error will not initialize | Sensor temp error Ran the diagnostic check and the Therm heater test failed | Replaced the analog board performed a annual overhaul since the analog board was replaced performed the zero check -- customer ran out of paper. Printed internal QC Unable To print external QC All QC ran Customer is only running ph turned all parameters off except PH | SWO-524012 | [5 empty] | Sensor temp error will not initialize | Sensor temp error Ran the diagnostic check and the Therm heater test failed | Replaced the analog board performed a annual overhaul since the analog board was replaced performed the zero check -- customer ran out of paper. Printed internal QC Unable To print external QC All QC ran Customer is only running ph turned all parameters off except PH
No | Visit | Complaint | RCHN | 无法抽吸样本，进样口排线有断痕 | 无法抽吸样本，借903-644一个 | 更换库存903-644后修复 by jesson | SWO-523536 | Defect | HW | Basic Unit | PCB, wet section |  | Cannot draw samples, the inlet line is broken | Unable to draw samples, borrow 903-644 one | Repair by jesson after replacing stock 903-644
No | Visit | Complaint | RCHN | 进样针堵塞 | 进样针堵塞（未借用配件） | 冲洗进样针后修复 by jesson | SWO-523694 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway | Clot | The sample needle is blocked. | Sample needle clogging (no accessories borrowed) | Fix by jesson after flushing the sample needle
No | Visit | Complaint | RCHN | 无法抽吸溶液，报错1343 | 无法抽吸溶液，报错1343 | 清洁管路后修复by tim | SWO-524338 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway | 1343 | Unable to pump solution, error 1343 | Unable to pump solution, error 1343 | Repair by tim after cleaning lines
No | Visit | Complaint | RCHN | 处理盘弹簧断裂 | 处理盘弹簧断裂，借903-702一个 | 更换903-702后修复 by jesson | SWO-523538 | Defect | HW | Process Module | Process Wheel | Spring | Treatment disc spring break | The handling disc spring breaks, borrows 903-702 one | Repair by jesson after replacement of 903-702
No | Visit | Complaint | RCHN | HCT无法测量 | HCT无法测量,借903-329一个。 | 更换903-329后修复 by jesson | SWO-522735 | Defect | HW | Aspiration Module | Needle Wash Unit | 1223 | HCT cannot be measured | HCT cannot be measured, borrowing 903-329 one. | Repair by jesson after replacing 903-329
No | Visit | Complaint | RIND | PO2 results not coming for samples | Visiting to run QC | Did rinsing , did calibration , run QC , QC failed for PO2 parameter , informed the same to Mr.Subbarao , suggested to replaced the cassette | SWO-523878 | [5 empty] | PO2 results not coming for samples | Visiting to run QC | Did rinsing , did calibration , run QC , QC failed for PO2 parameter , informed the same to Mr.Subbarao , suggested to replaced the cassette
No | Visit | Complaint | RIND | Low Po2 & tHB result. | Check the patient result log & found tHb,So2 result | Flush the SC9 with NS,run the Calibration & QC,got the desired result.Then run the sample ,got the satisfactory result.Told the customer to avoid pre analytical errors & pre nalytical training given to the customer. | SWO-524074 | [5 empty] | Low Po2 & tHB result. | Check the patient result log & found tHb,So2 result | Flush the SC9 with NS,run the Calibration & QC,got the desired result. Then run the sample ,got the satisfactory result. Told the customer to avoid pre analytical errors & pre nalytical training given to the customer.
No | Visit | Complaint | RIND | Service Call. | Need to calibrate the CKMB parameter in to the analyzer. | Due to replace the optical unit, required to calibrate the CKMB parameter, calibrate the CKMB parameter successfully, analyzer handed over to staff in well working condition. | SWO-524067 | Assisted Replacement | Consumables | Test Kit | Cartridge, cal |  | Service Call. | Need to calibrate the CKMB parameter in to the analyzer. | Due to replace the optical unit, required to calibrate the CKMB parameter, calibrate the CKMB parameter successfully, analyzer handed over to staff in well working condition.
No | Visit | Complaint | RPOL | Błędy ciśnienia. | Weryfikacja komunikatów błędów w analizatorze | Stwierdzono błędy ciśnienia. Wymiana wężyka oxi. Wymiana uszczelki wlotu. Kalibracja, kontrola, testowanie. | SWO-523756 | Defect | HW | Oxi module | Pinch valve tube |  | Pressure errors. | Verifying error messages in the analyzer | Pressure errors were found. Replacing the Oxi tube. Replacing the inlet gasket. Calibration, inspection, testing.
No | Visit | Complaint | RPOL | Błąd kalibracji pompy pH/BG i Solution (błędy 538 i 541). | Sprawdzenie stanu wężyków - nieszczelny (pęknięty) wężyk igła - inlet. | Wymiana uszkodzonego wężyka. Kalibracje i testy serwisowe. Kalibracja 2-punktowa. Test OC - OK. | SWO-523809 | Defect | HW | Inlet module | Tubing | 538, 541 | Calibration error of pH/BG and Solution pump (errors 538 and 541). | Check the condition of the tubing-leaky (cracked) needle-inlet tube. | Replacing the defective hose. Calibrations and service tests. 2-point calibration. Test OC-OK.
No | Visit | Complaint | RSAF | New SC chip could not be read | Faulty SC chip | Replaced faulty sensor cassette, startup test completed. 2pt cal and QC ok. | SWO-524008 | Assisted Replacement | SC | SC | chip | 1123 | New SC chip could not be read | Faulty SC chip | Replaced faulty sensor cassette, startup test completed. 2pt cal and QC ok.
No | Visit | Complaint | RKOR | Interface error (QC Gas item) | . | AQURE, mirth check.=> OK. Interface | SWO-524267 | Setup | SW | RMED SW | His/Lis | Aqure | Interface error (QC Gas item) | . | AQURE, mirth check.=> OK. Interface
No | Visit | Complaint | RCHN | 无法安装试剂包 | 无法安装试剂包 | 试剂包接触点接触不良，调整位置后修复by tim | SWO-525297 | [5 empty] | Reagent pack cannot be installed | Reagent pack cannot be installed | Reagent pack contact point contact poor contact, adjust by tim after adjustment of position
No | Visit | Complaint | RCHN | LAC定标不过，报错0378 | LAC定标不过，报错0378，借903-581一个（未更换备件） | 拆机后清洁板子后修复 by jesson | SWO-523513 | Cleaned | HW | Basic Unit | PCB, Sensor interface | Lac | LAC calibration but, error 0378 | LAC calibration but, error 0378, loan 903-581 one (not replaced spare parts) | Repair by jesson after cleaning the board after the removal
No | Visit | Complaint | RIND | PO2 results not coming for samples | Cassette replacement and QC | Installed the new cassette , calibration passed , run QC QC passed , run samples , customer happy with the results . instrument working fine . | SWO-525299 | [5 empty] | PO2 results not coming for samples | Cassette replacement and QC | Installed the new cassette , calibration passed , run QC QC passed , run samples , customer happy with the results . instrument working fine .
No | Visit | Complaint | RIND | Oximetry measuring error. | Oximetry measuring error. | Run complete oximetry test in service program, found hemolyzer test failed. In it impedance shape is failing regularly. Reconnect all the connection of oxi module PCB. Restart the machine run protein removal and run 2 point calibration. Than run complete oximetry test in service. But problem is still same, hemolyzer failed again. Run 2 point calibration successfully. | SWO-525227 | Cleaned | HW | Oxi module | Hemolyzer | mail: solved oxi failure? | Oximetry measuring error. | Oximetry measuring error. | Run complete oximetry test in service program, found hemolyzer test failed. In it impedance shape is failing regularly. Reconnect all the connection of oxi module PCB. Restart the machine run protein removal and run 2 point calibration. Than run complete oximetry test in service. But problem is still same, hemolyzer failed again. Run 2 point calibration successfully.
No | Visit | Complaint | RCHN | 报错5777 2168 5784 | 报错5777 2168 5784 | 电话指导老师重启无效 进服务发现safe position自检时PunchUp故障 拆卸孵育盘和处理包 清理孵育盘卡住的杯子 重新安装孵育盘和处理包 开机初始化正常通过 仪器恢复使用by tim | SWO-526945 | Adjustment | HW | Incubator Module | Incubator Wheel | 5777, 2168, 5784 | Error 5777 2168 5784 | Error 5777 2168 5784 | Phone instructor restart invalid into the service found safe position self-test when PunchUp fault remove incubation disk and processing package clean incubation disk stuck cup Reinstall incubation disk and processing package start initialization normal use by the instrument recovery by tim
Yes | Visit | Complaint | RCHN | 无法开机 | 无法开机，借910-321,910-274各一个 | 更换910-321,910-274后修复 by jesson | SWO-523617 | [5 empty] | Cannot power on | Cannot be switched on, borrow 910-321, 910-274 each | Fix by jesson after replacing 910-321, 910-274
No | Visit | Complaint | RDEU | QC Fehler pCO2 Level A und C | Fehler 1189, pCO2 Fehler Level A und C | Gerät überprüft, Überprüfung mit dem Sensorsimulator, Lösungspack, Sensorkassette und Einlaßdichtung ersetzt | SWO-525159 | Assisted Replacement | SC | SC | pCO2 | QC L A and C | QC Fehler pCO2 Level A und C | Error 1189, pCO2 Error Level A and C | Device checked, verification replaced with sensor simulator, solution pack, sensor cartridge and inlet seal
No | Visit | Complaint | RDEU | 01.04.2019, Herr Schuster, email,: Display Rahmen gebrochen | Displyramen gebrochen, | Displayramen ersetzt, | SWO-524718 | Defect | HW | Display | Display | Screws | 01.04.2019, Mr. Schuster, email,: Display frame broken | Displyramen broken, | Display rams replaced,
No | Visit | Complaint | RMAY | Oncall breakdown | Blockage from waste nipple | -Check analyser data log -Perform manual flush ( Fail) -Replace PPM Kit -Run fluid transport test ( OK) | SWO-529414 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway | Clot, 1063 | Oncall breakdown | Blockage from waste nipple | -Check analyser data log -Perform manual flush ( Fail) -Replace PPM Kit -Run fluid transport test ( OK)
No | Visit | Complaint | RLTD | Please call Jepsy on 0208 725 4450 | Found Oximetry pinch valve tube to have a hole in it. Replaced this tube. Restarted analyser and verified operation. | REplaced Oximetry pinch valve tube. | SWO-524868 | Defect | HW | Oxi module | Pinch valve tube |  | Please call Jepsy on 0208 725 4450 | Found Oximetry pinch valve tube to have a hole in it. Replaced this tube. Restarted analyser and verified operation. | REplaced Oximetry pinch valve tube.
No | Visit | Complaint | RAME | Calibrations are not passing. | The calibrations are not passing. | Jodi states she replaced the sensor cassette this morning due to pCO2 not passing. After she replaced the sensor cassette, now Potassium (K) is not passing calibrations. I performed the sensor simulator test to confirm the issues are not related to the analyzer sensor board. The passed within specifications. I removed the sensor simulator and, reinstalled the sensor cassette. Upon restart the analyzer presented a flow selector error. I removed the solution pack and, rebooted the analyzer to reset the flow selector. I reinstalled the solution pack and, I ran the fluid transport test, all passing within specifications. Now, the analyzer reinitialized without errors. | SWO-525002 | Assisted Replacement | SC | SC | K |  | Calibrations are not passing. | The calibrations are not passing. | Jodi states she replaced the sensor cassette this morning due to pCO2 not passing. After she replaced the sensor cassette, now Potassium (K) is not passing calibrations. I performed the sensor simulator test to confirm the issues are not related to the analyzer sensor board. The passed within specifications. I removed the sensor simulator and, reinstalled the sensor cassette. Upon restart the analyzer presented a flow selector error. I removed the solution pack and, rebooted the analyzer to reset the flow selector. I reinstalled the solution pack and, I ran the fluid transport test, all passing within specifications. Now, the analyzer reinitialized without errors.
No | Visit | Complaint | RIND | frequent calibration failure coming in the analyser. | Check System | Check system, done zeroing ,reinstall SP and SC, analyser now under observation. | SWO-529445 | [5 empty] | frequent calibration failure coming in the analyser. | Check System | Check system, done zeroing ,reinstall SP and SC, analyser now under observation.
No | Visit | Complaint | RSCH | Blank | Time setup ABL80Flex (failed) Décalage date et horaire après le changement horaire (hiver-été) | Erreur de configuration dans Win7 Control Panel (Date-Time-Location) e.o Demande de garantie SP - SC e.o Calibration e.o cycle de qualité e.o | SWO-524863 | [5 empty] | Blank | Time setup ABL80Flex (failed) Date and time shift after time change (winter-summer) | Configuration error in Win7 Control Panel (Date-Time-Location) e.o SP warranty request - SC e.o E.o quality cycle calibration e.o
No | Visit | Complaint | RCHN | 无法抽吸样本，进样口排线有断痕 | 频繁报错抽吸样本失败 | 更换903-199测试卡底座后修复by tim | SWO-528143 | Defect | HW | Basic Unit | Sensor Bay Unit | FT | Cannot draw samples, the inlet line is broken | Frequent lying error siphoning sample failure | Repair by tim after replacing the base of the 903-199 test card
No | Visit | Complaint | RCHN | PH,电解质定标失败 | PH,电解质定标失败，借903-581一个 | 更换903-581后修复by tim | SWO-524079 | Defect | HW | Basic Unit | PCB, Sensor interface | pH | PH, electrolyte calibration failure | PH, electrolyte calibration failure, borrow edith 903-581 | Repair by tim after replacement 903-581
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | K测量结果与临床不符 | 更换测试卡后修复 | SWO-531765 | [5 empty] | Test card, reagent package failure | K measurement sittin tinted with clinical | Repair after replacing test card
No | Visit | Complaint | RCHN | 定标失败。报错5692 | 定标失败。报错5692，借944-258（14867）一张 |  | SWO-524421 | Assisted Replacement | Consumables | Test Kit | Cartridge, cal | 5692 | The bid failed. 5692 error | The bid failed. 5692, loan 944-258 (14867) one
No | Visit | Complaint | RCHN | 打印机间歇故障，无法打印 | 打印机间歇故障，无法打印，借910-373一个 |  | SWO-524380 | Defect | HW | Analyzer Control Module | CF card | Corrective Action: Repaired by tim after replacing 910-373 | The printer is intermittently faulty and cannot print | Printer intermittent failure, unable to print, borrow 910-373 one
No | Visit | Complaint | RCHN | 处理无法识别 | 处理无法识别，借903-036一个 |  | SWO-524385 | Defect | HW | Cover and Chassis | PCB, solution pack |  | Processing is not recognized | Processing is not recognized, borrow 903-036 a
No | Visit | Complaint | RLTD | Analyser memory chip not detected contact service displayed on analyser screen Dr Lewis Gough | Return visit to complete long running service call, new SC arrived with me and due to the problems encountered on this service call i wanted to install the SC myself and ensure the job was completed. | Install SC, and perform calibration. | SWO-525257 | [5 empty] | Analyser memory chip not detected contact service displayed on analyser screen Dr Lewis Gough | Return visit to complete long running service call, new SC arrived with me and due to the problems encountered on this service call i wanted to install the SC myself and ensure the job was completed. | Install SC, and perform calibration.
No | Visit | Complaint | RKK | QCレベル１にて、pO2、ノーエンドポイント | 機器状況確認 | メンテナンスキット、センサーコネクターボード交換作業一式。 各部調整動作チェックOK。 CALチェックOK QCチェックOK | SWO-531700 | [5 empty] | At QC level 1, pO2, no endpoint | Equipment status check | Maintenance kit, sensor connector board replacement work set. Each part adjustment operation check OK. CAL check OK QC Check OK
No | Visit | Complaint | RIND | Rinse error. | Check the logs. Inlet impedance error. | Clean inlet gasket . clean inlet section & ph chamber. Run calibration multiple time. | SWO-529347 | Cleaned | HW | Inlet module | Inlet gasket | Clot | Rinse error. | Check the logs. Inlet impedance error. | Clean inlet gasket . clean inlet section & ph chamber. Run calibration multiple time.
No | Visit | Complaint | RDEU | Nach Umzug des ABL bei Neustart FehlerData upgrade failed- status code 9 | Roma ini nur noch 0 KB Inhalt | Roma ini. versucht aus Sicherung wieder einzuspielen - Fehler bleibt, Software neu instaliert mit Datenübernahme - Fehler bleibt, Software neu installiert i.O., Konfig aus dem Gerät der Intensiv eingespielt und Geräte Namen angepasst, Kal und QC durchgeführt, Verbindung zum Aqure i.O. | SWO-525154 | Defect | SW | RMED SW | Roma.ini | Status code 9 | After moving the ABL on reboot ErrorData upgrade failed- status code 9 | Roma ini only 0 KB content left | Roma ini. tries to re-play from backup - error remains, software re-installed with data transfer - error remains, software reinstalled i.O., Config from the device of intensively recorded and devices adapted names, Kal and QC performed, connection to aqure i.O.
No | 4 | Y | RCZE | Blank |  |  | D5E8B821AC5CADD9 | Revisit required | HW | Process Module | Optical Unit | QC | Blank
No | Visit | Complaint | RLTD | Aborting samples | Aborting samples. | Samples Aborting on board. FTC error issue. Checked the fluid path. Large blood clot in filter unit located at wash block. Cleaned unit and flushed lines. Cleaned the wash block. Checked needle unit for blockages or blood clots. Calibrated the needle unit. Ran 3 levels of LQC, cleaning program X 3 and 3 samples on board. Service dump and system logs taken. | SWO-524828 | Cleaned | HW | Wet section | Needle Wash Unit | Clot | Aborting samples | Aborting samples. | Samples Aborting on board. FTC error issue. Checked the fluid path. Large blood clot in filter unit located at wash block. Cleaned unit and flushed lines. Cleaned the wash block. Checked needle unit for blockages or blood clots. Calibrated the needle unit. Ran 3 levels of LQC, cleaning program X 3 and 3 samples on board. Service dump and system logs taken.
No | Visit | Complaint | RDEU | Probe abgebrochen! | Probe abgebrochen! QC abgebrochen! Fehler Meldung: 1067, 1999, 4800, 8611 Nadel fährt in allen Positionen (z.b, Cup Jet, Cup wick usw.)nicht! | Nadel überprüft, Safe Position, Startphase, Blank Messungen vorgenommen, Fehler weiterhin vorhanden, Gerät zur Reparatur an Michael Brand RDEU geschickt. | SWO-525262 | Defect | HW | Aspiration Module | PCB, pipette | Wait, not solved yet. From other SWOs: Pipette PCB adjusted. First failure is gone after that, but it reappears. Now Pipette PCB and Cable ordered. | Sample aborted! | Sample aborted! QC aborted! Error message: 1067, 1999, 4800, 8611 Needle travels in all positions (e.g., Cup Jet, Cup wick, etc.) Not! | Needle checked, safe position, start phase, blank measurements made, errors still present, device sent to Michael Brand RDEU for repair.
No | Visit | Complaint | RPOL | Restartowanie się aparatu | Serwis stwierdza uszkodzoną kartę pamięci oraz bazę danych, występowanie błędu kalibracji glukozy oraz błędu transportu płynu dla roztworu kalibracyjnego C. | Wymiana karty pamięci CF, instalacja Windows oraz oprogramowania analizatora, wymiana uszczelki wlotu i pakietu roztworów. Pakiet roztworów ze względu na błąd transportu płynu należy zgłosić do reklamacji. | SWO-524798 | Defect | HW | DMS_PC | CF card | Database | Restarting the camera | The service finds the damaged memory card and the database, the occurrence of a glucose calibration error and the liquid transport error for the calibration solution C. | Replacing the CF memory card, installing Windows and analyzer software, replacing the inlet gasket and solution package. The package of solutions due to the liquid transport error should be reported to the complaint.
No | Visit | Complaint | RPOL | Wymiana pakietu roztworów | Przegląd logów | Wymiana pakietu roztworów, testy serwisowe. | SWO-529702 | Assisted Replacement | SP | SP | SP |  | Solution Package Replacement | Log overview | Solution package replacement, service tests.
No | Visit | Complaint | RLTD | Analyser failing 2 levels of QC on several Oximetry parameters. | Analyser failing oxi parameters on 2 QC levels, could not resolve remotely. | Fit another hemolyser, perfrom oxi complete test and tHb calibration, perform calibrations and QCs. Also whilst onsite at rquest of customer extracted all the operators from all their analysers. | SWO-527242 | Defect | HW | Oxi module | Hemolyzer | QC | Analyser failing 2 levels of QC on several Oximetry parameters. | Analyser failing oxi parameters on 2 QC levels, could not resolve remotely. | Fit another hemolyser, perfrom oxi complete test and tHb calibration, perform calibrations and QCs. Also whilst onsite at rquest of customer extracted all the operators from all their analysers.
No | Visit | Complaint | RDEU | Bootfehler in Applikationssoftware V3.4 | --Garantie-- Bootfehler in der Applikationssoftware. Bootvorgang analysiert. enable File Base Write Filter = Fehler | Roma ASTM und HL7 Dateien gesichert. Applikationssoftware neu geladen. Dabei konnte der Ladevorgang bei der Wetsection nicht abgeschlossen werden. Nochmals Software installiert, und Einlassdichtung erneuert, dann alles o.k. | SWO-527102 | Defect | SW | RMED SW | RMED SW |  | Bootfehler in Applikationssoftware V3.4 | --Guarantee-- Boot error in the application software. boot ingess. enable File Base Write Filter = Error | Roma ASTM and HL7 files backed up. application software. The wetsection loading process could not be completed. Again installed software, and inlet seal renewed, then everything OK.
No | Visit | Complaint | RDEU | 01.04.2019 13:00 H. Koblich: Mischer mischt die Proben nicht | Fehlersuche vor Ort | Probenmischer überprüft. Kabel überprüft. komplette sampler mixer unit getauscht. Danach war die Funktion wieder gegeben und die Probe wurde wieder ordnungsgemäß gemischt. Kalibrierung durchgeführt.i.O. QCs für alle Levekl durchgeführt. i.O. Patientenmessung durchgeführt. i.O. ABL wieder messbereit. Noch einige Fragen von Herrn Koblich geklärt. ABL wieder messbereit. | SWO-527530 | Defect | HW | Sample mixer | Sample mixer |  | 01.04.2019 13:00 H. Koblich: Mixer does not mix the samples | Troubleshooting on site | sample mixer. Cable checked. complete sampler mixer unit. After that, the function was returned and the sample was properly mixed again. Calibration performed.i.O. QCs for all Levekl. Ok. patient measurement. Ok. ABL ready to be measured again. A number of questions have been answered by Mr Koblich. ABL ready to be measured again.
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | Lac定标不过 | 更换测试卡后修复 | SWO-539922 | [5 empty] | Test card, reagent package failure | Lac calibration is not | Repair after replacing test card
No | Visit | Complaint | RCHN | K离子不准 | K离子不准 | 做质控后确认是样本问题By tim | SWO-528139 | [5 empty] | K ion is not allowed | K ion is not allowed | After QC confirmation is sample problem By tim
No | Visit | Complaint | RCHN | LAC,K,NA,CL定标失败 | LAC,K,NA,CL定标失败，借903-581一个 | 更换903-581后修复by tim | SWO-524441 | Defect | HW | Basic Unit | PCB, Sensor interface | Lac, K, Na, Cl, 1228 | LAC, K, NA, CL calibration failed | LAC, K, NA, CL calibration failed, borrowed 903-581 one | Repair by tim after replacement 903-581
No | Visit | Complaint | RCHN | 光学单元无法升降，液路的泵没有反应 | 光学单元无法升降，液路的泵没有反应，借903-112一个 | 更换903-112后修复by tim | SWO-524530 | Defect | HW | Wet section | PCB, wet section |  | The optical unit cannot be lifted and the pump of the liquid road is not responding | Optical unit can not lift, liquid road pump does not respond, borrow 903-112 one | Repair by tim after replacement 903-112
No | Visit | Complaint | RIND | Po2 values reported low. | Checked patient results and po2 values reported low | Performed Decontamination procedure, cleaned electrode chambers, 2 point calibration performed successfully. | SWO-527051 | Cleaned | HW | pH/BG Module | pH/BG Module | pO2 | Po2 values reported low. | Checked patient results and po2 values reported low | Performed Decontamination procedure, cleaned electrode chambers, 2 point calibration performed successfully.
No | Visit | Complaint | RPAC | 18/3 Email from Andrew: When we looked at the performance validation of creatinine, we were given some data comparing Radiometer enzymatic to COBAS enzymatic creatinine. Since going live we have collected 3 x weekly data and the data we have is nowhere near as good. Every result on the Radiometer has a negative bias compared to the COBAS enzymatic creatinine … either very small (1-2%) or quite large (over 20% lower). Certainly not as good as what we were led to believe. Andrew confirmed omparisons are performed using serum with minimal time delay between ABL and lab. 1/4 Investigation whilst doing correlations for new DCC installation crea v ED (ABL) and lab confirmed ED to have low bias. QC in range but low bias on S7835 Slot 1 (with tighter range). 2/4 Replaced crea electrodes - see SWO. | Low bias on weekly serum comparisons with laboratory Cobas method. 1 $4 Checked data logs: QC in range however Slot 1 S7835 and Slot 4 S7865 low bias. Slot 11 S8377 high bias. Crea membranes last replaced 25 $3 R0835. 2 $4 Installed new Crea electrodes A HD01 and B HD03. Clean, LS and calibration. Electrodes removed: A 945-662 BR01 and B 945-663 BQ03. Also performed tHb calibration. | 2 $4 Replaced Crea electrodes A and B. QC S7835 no change, Slot 11 now closer to mean. Whole blood sample comparisons between new DCC ABL815 and Cobas now acceptable. To monitor. | SWO-525093 | Assisted Replacement | Electrode | Membrane | Crea | CAPA-789, CAPA-790 | 18/3 Email from Andrew: When we looked at the performance validation of creatinine, we were given some data comparing Radiometer enzymatic to COBAS enzymatic creatinine. Since going live we have collected 3 x weekly data and the data we have is nowhere near as good. Every result on the Radiometer has a negative bias compared to the COBAS enzymatic creatinine … either very small (1-2%) or quite large (over 20% lower). Certainly not as good as what we were led to believe. Andrew confirmed omparisons are performed using serum with minimal time delay between ABL and lab. 1/4 Investigation whilst doing correlations for new DCC installation crea v ED (ABL) and lab confirmed ED to have low bias. QC in range but low bias on S7835 Slot 1 (with tighter range). 2/4 Replaced crea electrodes - see SWO. | Low bias on weekly serum comparisons with laboratory Cobas method. 1 /4 Checked data logs: QC in range however Slot 1 S7835 and Slot 4 S7865 low bias. Slot 11 S8377 high bias. Crea membranes last replaced 25 /3 R0835. 2 /4 Installed new Crea electrodes A HD01 and B HD03. Clean, LS and calibration. Electrodes removed: A 945-662 BR01 and B 945-663 BQ03. Also performed tHb calibration. | 2 $4 Replaced Crea electrodes A and B. QC S7835 no change, Slot 11 now closer to mean. Whole blood sample comparisons between new DCC ABL815 and Cobas now acceptable. To monitor.
No | Visit | Complaint | RPAC | Investigation by biomed Jakeer: Anlayser was not giving any audible sounds like ( close inlet flap, after sample extraction and also for selecting tabs on the screen), I had a look in service menu all setting were ok, but while I was In the service menu analyser had frozen for couple of minutes , then after analyser rebooted on its own then things came alright, user been noticing this since last Friday. Would you please depute a service Engineer or advise us to resolve the glitch. Requested staff inform us of any further issues since reboot. 11/2 Further email from Jakeer: As per the user Acoustic mute issue has reoccurred over weekend, this morning I has to do the same as last week, I selected the acoustic tab in the analyser settings then screen froze for quite long then, In the settings everything seems to be ok, to bring the acoustic signal back On has to tick the mute tab (check) and then untick (uncheck) it, for every key select analyser freezes then after unchecking the mute tab, analyser reboots on its own then things came alright. Please depute / advise to fix the issue. | Sounds currently are present. Tried to mute and unmute acoustic signal from settings, analyzer responsed OK, no reboot required. | Saved set up & backed up all data. Performed complete software reinstallation as per procedure and per Denmark's advice. Reloaded setup and all data. Performed electronics checkout and adjustment, calibration & QCs all OK. AQURE connectivity OK. To be monitored for any further sound issue. | SWO-527416 | Defect | SW | RMED SW | RMED SW | Sounds | Investigation by biomed Jakeer: Anlayser was not giving any audible sounds like ( close inlet flap, after sample extraction and also for selecting tabs on the screen), I had a look in service menu all setting were ok, but while I was In the service menu analyser had frozen for couple of minutes , then after analyser rebooted on its own then things came alright, user been noticing this since last Friday. Would you please depute a service Engineer or advise us to resolve the glitch. Requested staff inform us of any further issues since reboot. 11/2 Further email from Jakeer: As per the user Acoustic mute issue has reoccurred over weekend, this morning I has to do the same as last week, I selected the acoustic tab in the analyser settings then screen froze for quite long then, In the settings everything seems to b e ok, to bring the acoustic signal back On has to tick the mute tab (check) and then untick (uncheck) it, for every key select analyser freezes then after unchecking the mute tab, analyser reboots on its own then things came alright. Please depute / advise to fix the issue. | Sounds currently are present. Tried to mute and unmute acoustic signal from settings, analyzer responsed OK, no reboot required. | Saved set up & backed up all data. Performed complete software reinstallation as per procedure and per Denmark's advice. Reloaded setup and all data. Performed electronics checkout and adjustment, calibration & QCs all OK. AQURE connectivity OK. To be monitored for any further sound issue.
No | Visit | Complaint | RCHN | 质控值设定低于范围 | LAC质控值低于质控范围 0090 | 更换LAC膜后修复 | SWO-531767 | Assisted Replacement | Membrane | Membrane | Lac |  | The control value setting is below range | LAC QC value is lower than QC range 0090 | Repair after replacing LAC film
No | Visit | Complaint | RDEU | 08.03.2019, Frau Suhr, Tel: 03834865528: QC Fehler bei CRP- Frau Suhr hat bereits alles mit Herrn Ripke besprochen. 27.03.2019, Frau Gerschau, Tel: 03834865558: QC Fehler Level 1 CRP 27.03.2019 Frau Gerschau 017616658311 wünscht einen Vorortservice. | nach Rücksprache mit der Zentrale, soll das Process Wheel und damit auch die opt. Einheit gewechselt werden | opt. Einheit gewechselt, inkl. Process Wheel CRP, NT-ProBNP und TnI kalibriert - QC i.O. Kal D-Dimer läuft aktuell QC übernimmt Fr. Cornelius Kal Cass. ßhCG und PCT nicht vorrätig und können somit nicht kalibriert werden | SWO-527169 | Defect | HW | Process Module | Optical Unit | LQC | 08.03.2019, Frau Suhr, Tel: 03834865528: QC Fehler bei CRP- Frau Suhr hat bereits alles mit Herrn Ripke besprochen. 27.03.2019, Frau Gerschau, Tel: 03834865558: QC Fehler Level 1 CRP 27.03.2019 Frau Gerschau 017616658311 wünscht einen Vorortservice. | after consultation with the control panel, the Process Wheel and thus also the opt. unit can be changed | Opt. Unit changed, incl. Process Wheel CRP, NT-ProBNP and TnI calibrated - QC in the same way Kal D-Dimer is currently QC takes over Fr. Cornelius Kal Cass. ßhCG and PCT are out of stock and therefore cannot be calibrated
No | Visit | Complaint | RDEU | 07.02.2019, Frau Rahn, Tel: 060329254613: QC Fehler bei CRP Level 2 | Überprüfung LQC CRP nach SW Update und fehlerhafter LQC LQC CRP1 Lot 15635 LQC CRP2 Lot 15119 Reagenzien CRP Lot 15598 LQC D-Dimer läuft in beiden Leveln zielwertig | Software Update V8.11.300.00 wurde durchgeführt. CRP wurde im gleichen Lot neu Kalibriert. LQC Charge wurde auf anderes Los in beiden Level gewechselt. Weiterhin LQC im CRP zu hoch. (Keine Reinigung vor Messung) LQC nach SW Update nach Auftauzeit gemessen. LQC CRP1 ist 47,8mg $L (41,8mg $L ; 36,2-47,5mg $L) zweite Messung 44,6mg $L ohne Beanstandung LQC CRP2 ist 124mg $L (117mg $L ; 101-133mg $L) ohne Beanstandung Frau Rahn beobachtet die LQC Messung und meldet sich ggf. ServiceDump und DataLogs erstellt, einsendung TS zur Abwägung des Tauschens der optischen Einheit. | SWO-527060 | Assisted Replacement | Consumables | Test Kit | Cartridge, cal | CRP QC | 07.02.2019, Frau Rahn, Tel: 060329254613: QC Fehler bei CRP Level 2 | Check LQC CRP after SW update and faulty LQC LQC CRP1 Lot 15635 LQC CRP2 Lot 15119 Reagents CRP Lot 15598 LQC D-Dimer runs in both levels | Software Update V8.11.300.00 has been performed. CRP has been recalibrated in the same lot. LQC Charge was changed to another lot in both levels. Further LQC in the CRP too high. (No cleaning before measurement) LQC measured according to SW update by thawing time. LQC CRP1 is 47.8mg $L (41.8mg $L; 36.2-47.5mg $L) second measurement 44.6mg $L without complaint LQC CRP2 is 124mg $L (117mg $L; 101-133mg $L) without complaint Ms Rahn observes the LQC measurement and, if necessary, reports. ServiceDump and DataLogs created, submit TS to weigh the exchange of the optical unit.
No | Visit | Complaint | RIND | Analyzer showing Rotate inlet wheel error. | Checked the analyzer and performed safe position and conformed inlet wheel not rotating properly . | checked the inlet wheel and dismantled and re assembled it and performed safe positions now its showing safe .and performed 10 samples are results came analyzer working fine. | SWO-529344 | Adjustment | HW | Inlet module | Inlet Wheel | inletwheel25; Potentially User Error (FSE) | Analyzer showing Rotate inlet wheel error. | Checked the analyzer and performed safe position and conformed inlet wheel not rotating properly . | checked the inlet wheel and dismantled and re assembled it and performed safe positions now its showing safe .and performed 10 samples are results came analyzer working fine.
No | Visit | Complaint | RIND | Undefined look up strings | Checked safe programs and found Inlet wheel failed. | Checked all connectors, opened and cleaned Inlet wheel. Reassembled Inlet wheel. Performed routine maintenance procedures. Calibrated new PCT lot successfully. | SWO-527502 | Adjustment | HW | Inlet module | Inlet Wheel | Potentially Use Error (FSE) | Undefined look up strings | Checked safe programs and found Inlet wheel failed. | Checked all connectors, opened and cleaned Inlet wheel. Reassembled Inlet wheel. Performed routine maintenance procedures. Calibrated new PCT lot successfully.
No | Visit | Complaint | RAME | Demographics not showing up when scanning safe pico sampler. Customer requests SW be rolled back to v 3.4 | Inspection of the system revealed there is a loss off demographics info from AQURE intermittently about once a day or two. | Rolled back SW to Version 3.4 from 3.4 MR1. | SWO-527971 | Defect | SW | RMED SW | Demographics | sent to MIGGA | Demographics not showing up when scanning safe pico sampler. Customer requests SW be rolled back to v 3.4 | Inspection of the system revealed there is a loss off demographics info from AQURE intermittently about once a day or two. | Rolled back SW to Version 3.4 from 3.4 MR1.
No | Visit | Complaint | RAME | Customer reports that a “leak detected” error has been generated. | Leak detected errors are present. Vacuum value is 0.99. Solution pack and inlet gasket replacement have not resolved the issue. | Replaced the annual PM parts and rebooted analyzer. Vacuum value is now measured at 0.47. Also replaced the sensor cassette and submitted the credit claim form on behalf of the customer due to low pCO2 sensitivity(33%). Analyzer is within acceptable operating specifications. | SWO-527301 | Revisit required | HW | Auto Inlet module | PCB, Auto inlet | Leak detected | Customer reports that a “leak detected” error has been generated. | Leak detected errors are present. Vacuum value is 0.99. Solution pack and inlet gasket replacement have not resolved the issue. | Replaced the annual PM parts and rebooted analyzer. Vacuum value is now measured at 0.47. Also replaced the sensor cassette and submitted the credit claim form on behalf of the customer due to low pCO2 sensitivity(33%). Analyzer is within acceptable operating specifications.
No | Visit | Complaint | RAME | Wireless Connection intermittently Disconnects | Suspect wireless just cuts out | I replaced the Belgian dongle with an ASUS dongle | SWO-527832 | Defect | HW | Wifi | Wifi dongle |  | Wireless Connection intermittently Disconnects | Suspect wireless just cuts out | I replaced the Belgian dongle with an ASUS dongle
No | Visit | Complaint | RAME | Broken display. | Broken display. | Replaced display. | SWO-527914 | Defect | HW | Display | Display |  | Broken display. | Broken display. | Replaced display.
No | Visit | Complaint | RAME | Unable to pump solution. Stuck in intervention mode. | Inspection of the system revealed multiple fluidic transport errors. | Replaced Sensor bay unit , sensor interface PCB, installed a complete PM kit, & reloaded SWU for diagnostic purposes. | SWO-527368 | Defect | HW | Basic Unit | Sensor Bay Unit |  | Unable to pump solution. Stuck in intervention mode. | Inspection of the system revealed multiple fluidic transport errors. | Replaced Sensor bay unit , sensor interface PCB, installed a complete PM kit, & reloaded SWU for diagnostic purposes.
No | Visit | Complaint | RCHN | 接口板腐蚀 | 接口板腐蚀，借902-972一个 | 更换902-972后修复by tim | SWO-525298 | [5 empty] | Interface plate corrosion | Interface plate corrosion, borrow 902-972 one | Repair by tim after replacement 902-972
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | LAC定标值偏低 | 更换测试卡后修复 | SWO-532542 | [5 empty] | Test card, reagent package failure | LOW LAC CALIBRATION VALUE | Repair after replacing test card
No | Visit | Complaint | RCHN | 风扇线断了 | 风扇线断了，借580-182一个 | 更换580-182后修复by tim | SWO-524473 | Defect | HW | Basic Unit | Cable, Fan |  | The fan line is broken. | Fan line broken, borrow 580-182 one | Repair by tim after replacement 580-182
No | Visit | Complaint | RCHN | 报错1031，光学单元灵敏度超范围 | 报错1031，光学强度超范围，借903-364一个 |  | SWO-527418 | Defect | HW | Process Module | Optical Unit | 1031 | Error 1031, optical unit sensitivity over range | Error 1031, optical intensity over range, borrowed 903-364 one
No | Visit | Complaint | RCHN | 孵育盘转动故障 | 孵育盘转动故障，借836-166,580-097各一个 | 更换836-166一个正常。by tim | SWO-525318 | Defect | HW | Incubator Module | Drive wheel |  | Incubation plate rotation fault | Incubation plate rotation fault, borrow 836-166, 580-097 each one | Replace 836-166 one normal. by tim
No | Visit | Complaint | RCHN | 做不了样本 | 做不了样本（未借用备件） | 清洁仪器并做了维护保养后修复by tim | SWO-528177 | User Error | Consumables | System clean | System clean | expired maintenance | Can't do the sample. | Cannot do the sample (no spare parts borrowed) | Clean the instrument and repair by tim after maintenance
No | Visit | Complaint | RAME | Waste bottle over flowed. | Customer had waste bottle overflow numerous times, | Replaced waste detector. | SWO-527945 | Defect | HW | Basic Unit | Waste detector |  | Waste bottle over flowed. | Customer had waste bottle overflow numerous times, | Replaced waste detector.
No | Visit | Complaint | RIND | pH calibration error. | Check the system. | Clean the ph & replace the ref membrane. machine under observation. | SWO-529337 | Cleaned | HW | pH/BG Module | measuring chamber | pH | pH calibration error. | Check the system. | Clean the ph & replace the ref membrane. machine under observation.
Yes | Visit | Complaint | RBNL | Verplaatsing van de AQT90 | Op verzoek van klant, AQT90 heb verplaatst naar andere ruimte vervolgens heb gecontroleerd, MCC1 LQC TNT heb bepaald, OK. | Verplaatsing AQT90 | SWO-527838 | Re classified RMED | Request |  |  | See mail with title "SWO-527838" | Relocation of the AQT90 | At customer's request, AQT90 have moved to other space then checked, MCC1 LQC TNT have determined, OK. | Displacement AQT90
No | Visit | Complaint | RPOL | Weryfikacja stanu technicznego analizatora. Zgłoszono problem z instalacją pakietu odczynnikowego. | Zgłoszono problem z instalacja pakietu roztworów. Weryfikacja kodów błędów | Analiza dziennika wymian $zadań $błędów. Wykonano techniczne testy przepływu płynów. Testy hydrauliczne prawidłowe. W przypadku ponownych problemów z instalacja pakietu zostanie wymieniona płytka odczytu pamięci. | SWO-527857 | No fault found | No fault found | No fault found | No fault found |  | Checking the technical condition of the analyzer. A problem with the installation of the reagent pack has been reported. | A problem with the installation of the solutions package has been reported. Error code Verification | Analysis of the Exchange log $zada $b Łęów. Technical fluid flow tests have been performed. Hydraulic tests correct. If you are reinstalling the package, the memory read tile will be replaced.
No | Visit | Complaint | RPAC | Analyser is intervention Required | Analyser intervention required | Replaced inlet gasket Performed Calibrations Ran QC A,B,C Tested, all ok. | SWO-528195 | Defect | HW | Inlet module | Inlet gasket |  | Analyser is intervention Required | Analyser intervention required | Replaced inlet gasket Performed Calibrations Ran QC A,B,C Tested, all ok.
No | Visit | Complaint | RPAC | Therese called 5619 0100 broken inlet gasket and bent probe | On Site inspection | Replace the inlet module individual parts. Flush the Fluidic path and decontaminate as Blood remained in system since the breakdown. Replaced the Inlet gasket to the SC. Blocked with blood. Perform Start up check. | SWO-528931 | Defect | HW | Inlet module | Inlet gasket |  | Therese called 5619 0100 broken inlet gasket and bent probe | On Site inspection | Replace the inlet module individual parts. Flush the Fluidic path and decontaminate as Blood remained in system since the breakdown. Replaced the Inlet gasket to the SC. Blocked with blood. Perform Start up check.
No | Visit | Complaint | RDEU | 04.04.2019 8:05 H. Möglich +497802801329: Gerät fällt ständig aus | Kunde hat wiederholt Störungen am Gerät | Fließwege überprüft, Fließwege durch JW-Kit erneuert | SWO-528344 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | 04.04.2019 8:05 H. Possible +497802801329: Device is constantly falling out | Customer has repeatedly malfunctioned on the device | Flow paths checked, flow paths renewed by JW kit
No | Visit | Complaint | RDEU | 29.03.2019, Herr Hart, Tel: 0401818854581: Drucker rattert total beim ausdrucken, erst druckt er gar nicht mehr, nach Neustart druckt er jetzt- rattert jedoch sehr laut. | Lautes Knattern beim Drucken. | Drucker aus dem VSC geholt und zum Kunden gefahren.Drucker defekt. Drucker getauscht. Und Kundengespräch durchgeführt. | SWO-528180 | Defect | HW | Printer Module | Printer Module |  | 29.03.2019, Mr. Hart, Tel: 0401818854581: Printer rattles completely when printing, first it doesn't print at all, after reboot he prints now - but rattles very loudly. | Loud rattling when printing. | printer from the VSC and driven to the customer. Printer defective. Printer exchanged. And customer talk conducted.
No | Visit | Complaint | RDEU | 04.04.2019, Herr Ochtmann, Tel: 020360083110: Display Rahmen ist lose- Touchsceen lässt sich verschieben | Displayramen ist lose, Schraubenhülsen sind gebrochen, | Displayramen ersetzt, Gerät alibriert, | SWO-528836 | Defect | HW | Display | Display | Screws | 04.04.2019, Mr. Ochtmann, Tel: 020360083110: Display frame is loose- Touchsceen can be moved | Display ramen is loose, screw sleeves are broken, | Display rams replaced, device alibriert,
No | Visit | Complaint | RAME | Customer reports that a “leak detected” error has been generated. | Customer reports that the inlet gasket is hitting the solution pack when it closes. Inlet position detection assembly is observed to be out of alignment and preventing normal operation of the inlet module. | Properly aligned the inlet position detection assembly and verified proper opening and closing of the inlet module. Analyzer is within acceptable operating specifications. | SWO-528604 | Adjustment | HW | Auto Inlet module | PCB, Auto inlet | Leak detected | Customer reports that a “leak detected” error has been generated. | Customer reports that the inlet gasket is hitting the solution pack when it closes. Inlet position detection assembly is observed to be out of alignment and preventing normal operation of the inlet module. | Properly aligned the inlet position detection assembly and verified proper opening and closing of the inlet module. Analyzer is within acceptable operating specifications.
No | Visit | Complaint | RDEU | 04.04.2018 Herr Reinecke meldet, dass der Scanner nicht arbeitet und die Anzeige der GFR entfernt wurde. | Barcode Typ muss konfiguriert werden. | Konfiguration anderer BGA Geräte überprüft um die korrekten Barcode Typ heraus zu bekommen. Roma.Ini konfiguriert. | SWO-528358 | Setup | SW | RMED SW | Roma.ini | Barcode reader | 04.04.2018 Mr. Reinecke reports that the scanner is not working and the display of the GFR has been removed. | Barcode type must be configured. | Configuration of other BGA devices checked to get the correct barcode type out. Roma.Ini.
No | Visit | Complaint | RLTD | Unable to pump solutions | Unable to pump solutions | Ran visual flow check, I could clearly see the analyser was having difficulty pumping solutions. Tried to flush analyser from solution pack waste connection to left side of sensor cassette port, unable to get a clear flow. Repeatedly flushed until i could get a clear stream. Analyser now passing pump cal and solution pack check, waited for analyser to finish 55min start up, now okay. Also set up remote support on new analysers. | SWO-529807 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | Unable to pump solutions | Unable to pump solutions | Ran visual flow check, I could clearly see the analyser was having difficulty pumping solutions. Tried to flush analyser from solution pack waste connection to left side of sensor cassette port, unable to get a clear flow. Repeatedly flushed until i could get a clear stream. Analyser now passing pump cal and solution pack check, waited for analyser to finish 55min start up, now okay. Also set up remote support on new analysers.
No | Visit | Complaint | RHUN | készülék folyik, waste csatlakozó rendben van, oldat gyanús | oldat kazetta waste tartálya nem fogad folyadékot, a készülék oldalon hiba nem tapasztalható | készülék ellenőrzése, folyadék műveletek tesztelése, oldat csere, rendszer karbantartás | SWO-528313 | [5 empty] | The appliance is on, waste connector is in order, a suspect solution | Solution is not receiving any liquids, there is no error on the device side. | Monitoring of apparatus, testing of liquid operations, solution replacement, system maintenance
No | Visit | Complaint | RCHN | 电源线烧坏 | 电源线烧坏，需要借件903-663一个 | 更换903-663后修复by tim | SWO-527437 | [5 empty] | Power cord burned out | Power cord burned, need to borrow 903-663 one | Repair by tim after replacement 903-663
No | Visit | Complaint | RCHN | 屏幕花屏 | 屏幕花屏，借903-239一个 | 更换903-239后修复by tim | SWO-527426 | [5 empty] | Screen flower screen | Screen flower screen, borrow 903-239 one | Repair by tim after replacement 903-239
No | Visit | Complaint | RCHN | 电池无法蓄电 | 电池无法蓄电，借431-018一个 | 更换431-018后修复by tim | SWO-527543 | [5 empty] | The battery can't store power. | Battery can not store power, borrow 431-018 one | Repair by tim after replacement 431-018
No | Visit | Complaint | RCHN | 电池无法蓄电 | 电池无法蓄电，借431-018一个 | 更换431-018后修复by tim | SWO-527549 | [5 empty] | The battery can't store power. | Battery can not store power, borrow 431-018 one | Repair by tim after replacement 431-018
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | PH,NA,K,CA定标失败，1305 | 更换测试卡后修复 | SWO-529608 | [5 empty] | Test card, reagent package failure | PH, NA, K, CA calibration failure, 1305 | Repair after replacing test card
No | Visit | Complaint | RCHN | 报错1254，HCT无法测得 | 报错1254，HCT无法测得，借903-329一个（未更换备件） | 插拔HCT模块的排线，清洁HCT模块后修复by tim | SWO-527462 | Cleaned | HW | Aspiration Module | Needle Wash Unit | 1254 | Error 1254, HCT can not be measured | Error 1254, HCT can not be measured, borrow edgy 903-329 (unreplaced spare parts) | Plug in the wiring of the HCT module and repair the by-tim after cleaning the HCT module
No | 3 | Y | RCZE | Blank |  |  | D5ED3AAE8441C470 | Defect | HW | Printer Module | Printer Module |  | Blank
No | Visit | Complaint | RMAY | Service Breakdown | External QC data shows all level positive biased . | - Check analyzer data log - Clean chamber - Flush inlet - Perform protein removal - Run Radiometer QC all level ( Pass) | SWO-533102 | Cleaned | Fluidic pathway | pH/BG Module | pH/BG Module |  | Service Breakdown | External QC data shows all level positive biased . | - Check analyzer data log - Clean chamber - Flush inlet - Perform protein removal - Run Radiometer QC all level ( Pass)
No | Visit | Complaint | RIND | Query ? showing in patient results. | Checked the ABL800 and ? reported in patient results. | performed LS adjust and pump calibration successfully. 2point calibration performed successfully. | SWO-528849 | Adjustment | HW | pH/BG Module | LS |  | Query ? showing in patient results. | Checked the ABL800 and ? reported in patient results. | performed LS adjust and pump calibration successfully. 2point calibration performed successfully.
No | Visit | Complaint | RMAY | Breakdown services | Found proteinized in po2 chamber | - Check analyzer data log - Clean chamber with hypochlorite - Flush inlet - Perform protein removal | SWO-533100 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway | pO2 | Breakdown services | Found proteinized in po2 chamber | - Check analyzer data log - Clean chamber with hypochlorite - Flush inlet - Perform protein removal
No | Visit | Complaint | RKK | QCでNa、Ca下限を切る | QCでNa、Ca下限を切る | 比較電極交換 | SWO-528760 | Assisted Replacement | Electrode | Electrode | Reference |  | Na and Ca lower in QC | Na and Ca lower in QC | Comparative electrode replacement
No | Visit | Complaint | RPAC | Occasional calibration drifts on Na Ca Cl 4/4 Replaced new Ref electrode and membrane | There were still some occasional calibration drifts on Na, Ca, and Cl after Ref and electrolytes chambers were cleaned. | Replaced new Reference electrode and membrane. Calibration, QC and blood comp OK. To be monitored for further occurrences. FOC consumable supplied. | SWO-528082 | Assisted Replacement | Electrode | Electrode | Reference | CAPA-733 | Occasional calibration drifts on Na Ca Cl 4/4 Replaced new Ref electrode and membrane | There were still some occasional calibration drifts on Na, Ca, and Cl after Ref and electrolytes chambers were cleaned. | Replaced new Reference electrode and membrane. Calibration, QC and blood comp OK. To be monitored for further occurrences. FOC consumable supplied.
No | Visit | Complaint | RDEU | Anwender haben tHb Kalibration mehrmals durchgeführt, mit Fehlermeldung; Oxi Spectrum mismatch. | Im Haus unterwegs zu weiteren Service. tHb Kalibration durch Anwender gescheitert. Analysator meldet Oximetrie Hardware Fehler. | Überprüfung der Oximetrie im Service, Oxi Complete Test durchgeführt, alle Werte in Ordnung innerhalb der Grenzwerte. Proteinentfernung durchgeführt, Analysator kalibriert und tHb Kalibration durchgeführt, Messwert wieder in Ordnung, Fehler behoben. Bei der Durchführung der Qualitätskontrolle wurde die Probe abgebrochen. Ursache krumme Einlasssonde, erneuert. In der Qualitätskontroll Historie ist ersichtlich, dass es öfters zu Abbrüchen kam. Qualitätskontroll Wert des tHb Parameters nicht zielwertig. Neues Oximetrie Modul bestellt. | SWO-528517 | Revisit required | HW | Oxi module | Oxi module |  | Users have performed tHb calibration several times, with error message; Oxi Spectrum mismatch. | In the house on the way to further service. tHb calibration by users failed. Analyzer reports oximetry hardware errors. | Checking the oximetry in the service, Oxi Complete test performed, all values in order within the limits. Protein removal performed, analyzer calibrated and tHb calibration performed, measured value back in order, bug fixed. During the quality control, the sample was aborted. Cause crooked inlet probe, renewed. In the quality control history it can be seen that there were frequent cancellations. Quality control value of the tHb parameter not targetable. New oximetry module ordered.
No | Visit | Complaint | RSAF | Analyzer keeps on wanting to restart | None | Disabled and then enabled the firewall | SWO-528589 | Adjustment | SW | RMED SW | Firewall |  | Analyzer keeps on wanting to restart | None | Disabled and then enabled the firewall
No | Visit | Complaint | RDEU | 03.04.2019, Frau Maderthaner, Tel: 0898412556: Kundin berichtet, dass bei 2 Patienten der Trop I Wert in der Praxis negativ waren- diese Ptienten wurden in ein Krankenhaus eingeliefert und hatten dort positive Troponin Werte. | Kundin berichtet, dass bei 2 Patienten der Trop I Wert in der Praxis negativ waren- diese Ptienten wurden in ein Krankenhaus eingeliefert und hatten dort positive Troponin Werte. | Fließwege mit Hypochlorid gespült. Blankmessungen vorgenommen. Servicedump erstellt und überprüft. Fließwege in Ordnung (1033 - 1102). Blankmessungen in Ordnung. Systemüberprüfung und Systemreinigung - ok. LQC MCC2 gemessen - ok. | SWO-529362 | No fault found | No fault found | No fault found | No fault found | Potentially Use Error | 03.04.2019, Ms. Maderthaner, Tel: 0898412556: Customer reports that in 2 patients the Trop I value was negative in practice - these ptients were admitted to a hospital and had positive troponin values there. | Customer reports that in 2 patients the Trop I value were negative in practice - these ptients were admitted to a hospital and had positive troponin values there. | Flow paths rinsed with hypochloride. Blank measurements. Service dump created and verified. Flow ways in order (1033 - 1102). Blank measurements in order. System check and system cleaning - ok. LQC MCC2 measured - ok.
No | Visit | Complaint | RBNL | menu service - valve pressure pas ok | pinch valve bouche pas de temps en temps | *changement de pinch valve *test service - ok *calibration et QC - ok | SWO-528905 | Defect | HW | Oxi module | Pinch valve |  | menu service - valve pressure pas ok | pinch valve mouth no time to time | Change of pinch valve Test service - ok 'calibration' and QC - ok
No | Visit | Complaint | RHUN | Inlet +SC csere | Inlet +SC csere Általános szervíz hibaelhárítás funkciók áttekintése | Szervíz tesztek, Inlet csere, Technikus betanítása, SC csere, Inicializálás, Kalibrációk, QC mérés | SWO-528893 | Defect | HW | Inlet module | Inlet module |  | Inlet +SC csere | Inlet + SC Replacement General Service Troubleshooting features overview | Service tests, Inlet replacement, technician training, SC replacement, initialization, calibrations, QC measurement
No | Visit | Complaint | RMAY | Attend on-call breakdown due to inlet gasket holder error | 1. Check analyser status and data log 2. Intervention required | 1. Enter service - calibrate inlet, flush manually, run liquid transport 2. Pull out inlet gasket, check physically 3. Swap with loan gasket holder 4. Start up machine | SWO-533505 | Defect | HW | Auto Inlet module | Inlet gasket | 1401 | Attend on-call breakdown due to inlet gasket holder error | 1. Check analyser status and data log 2. Intervention required | 1. Enter service - calibrate inlet, flush manually, run liquid transport 2. Pull out inlet gasket, check physically 3. Swap with loan gasket holder 4. Start up machine
No | Visit | Complaint | RKOR | Print paper output error | Disassemble the print module and clean the inside Cal Control data check Check print normal operation | Disassemble the print module and clean the inside Cal Control data check Check print normal operation | SWO-529475 | Adjustment | HW | Printer Module | Printer Module |  | Print paper output error | Disassemble the print module and clean the inside Cal Control data check Check print normal operation | Disassemble the print module and clean the inside Cal Control data check Check print normal operation
No | Visit | Complaint | RIND | Parameter Failure | Lac failure | Flus the cassette with normal sline. Re-install the cassette , Calibration not passed. Need to change sensor cassette. | SWO-533078 | [5 empty] | Parameter Failure | Lac failure | Flus the cassette with normal sline. Re-install the cassette, Calibration not passed. Need to change sensor cassette.
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 液体传输失败1268，在泰康仙林鼓楼医院呼吸科试用 | 更换试剂包后修复 | SWO-530281 | [5 empty] | Test card, reagent package failure | Liquid transmission failure 1268, in Taikang Xianlin Drum Tower Hospital Respiratory Department trial | Repair after replacement of reagent kit
No | Visit | Complaint | RAME | Aborted cal's | The Cal would not run just showed abort | I checked the valves in the ph $big mod and the air valve was bad. I replaced it and the rinse valve and did the same to the El $met mod | SWO-529005 | Defect | HW | pH/BG Module | Valves |  | Aborted cal's | The Cal would not run just showed abort | I checked the valves in the ph $big mod and the air valve was bad. I replaced it and the rinse valve and did the same to the El $met mod
No | Visit | Complaint | RAME | Customer states he is having intermittent Lac calibration issues with drift despite replacing membranes constantly. | Verified customer's stated issue. Looked at all the analyzers logs and found several zero calibration errors. Checked the Lac electrode to ensure it has its respective sealing ring and enough solution in its membrane. Lac electrode had its sealing ring and the membrane was recently replaced. Checked the Lac electrodes connector and it was clean with no signs of patina. Checked the El $Met chamber and ensured it was dry and verified all the other electrodes had their sealing rings. | Performed a zero calibration on the analyzer and downloaded the analyzers service dump. | SWO-529131 | Adjustment | HW | El/Met Module | zero current | Lac | Customer states he is having intermittent Lac calibration issues with drift despite replacing membranes constantly. | Verified customer's stated issue. Looked at all the analyzers logs and found several zero calibration errors. Checked the Lac electrode to ensure it has its respective sealing ring and enough solution in its membrane. Lac electrode had its sealing ring and the membrane was recently replaced. Checked the Lac electrodes connector and it was clean with no signs of patina. Checked the El /Met chamber and ensured it was dry and verified all the other electrodes had their sealing rings. | Performed a zero calibration on the analyzer and downloaded the analyzers service dump.
No | Visit | Complaint | RSAF | Patients samples are running only on QC and needle straight. | No movement breaker Status in break position. | Service checks done. Breaker not working. Stepper motor to be replaced. | SWO-528840 | Revisit required | HW | Autocheck module | Stepper Motor | Breaker | Patients samples are running only on QC and needle straight. | No movement breaker Status in break position. | Service checks done. Breaker not working. Stepper motor to be replaced.
No | Visit | Complaint | RDEU | Keine Anbindung an Aqure möglich | Es besteht keine Netzwerke Verbindung | Konfiguration geprüft, mit der Haus IT Herr T. Adam gesprochen, nach längerem Suche wurde eine fehlerhafte Einstellung in der Firewall gefunden, Funktionskontrolle ok | SWO-529173 | Setup | SW | External SW | Firewall |  | No connection to Aqure possible | There is no network connection | Configuration checked, with the house IT Mr. T. Adam talked, after a long search a faulty setting was found in the firewall, function control ok
No | Visit | Complaint | RIND | Service Visit. | Need to clean the dust from the IR sensors. | Shutdown the analyzer, open the front panel as well as back panel of the analyzer, remove the incubator wheel and clean the dust from it, clean the IR sensor of storage wheel and incubator wheel, successfully done, connect all the parts same as before, start the analyzer, analyzer is in ready mode and handed over to staff in well working condition. | SWO-529320 | Cleaned | HW | Incubator Module | Sensor, incubator | Calibrate_13 | Service Visit. | Need to clean the dust from the IR sensors. | Shutdown the analyzer, open the front panel as well as back panel of the analyzer, remove the incubator wheel and clean the dust from it, clean the IR sensor of storage wheel and incubator wheel, successfully done, connect all the parts same as before, start the analyzer , analyzer is in ready mode and handed over to staff in well working condition.
No | Visit | Complaint | RCHN | HCT测量结果打问号 | HCT测量结果打问号,借903-329一个，944-268*15733一张 |  | SWO-529376 | Defect | HW | Aspiration Module | Needle Wash Unit |  | QUESTION MARK ON HCT MEASUREMENTS | HCT measurements hit the question mark, borrow 903-329 one, 944-268 x 15733
No | Visit | Complaint | RCHN | HCT无法测量 | HCT无法测量,借903-329一个。 |  | SWO-529378 | Defect | HW | Aspiration Module | Needle Wash Unit | 1223 | HCT cannot be measured | HCT cannot be measured, borrowing 903-329 one.
No | Visit | Complaint | RIND | Oximetry measuring error. | Oximetry values is missing from patient reports. | Change oxi module, run complete oximetry test successfully. Perfom protein removal and 2 point calibration successfully. Now machine is working fine and ready to use. | SWO-531124 | Defect | HW | Oxi module | Oxi module |  | Oximetry measuring error. | Oximetry values is missing from patient reports. | Change oxi module, run complete oximetry test successfully. Perfom protein removal and 2 point calibration successfully. Now machine is working fine and ready to use.
No | Visit | Complaint | RCHN | 试剂包无法读取信息，连接片断裂 | 试剂包无法读取信息，连接片断裂 | 更换库存903-158后修复by Hunter | SWO-529576 | Defect | HW | Basic Unit | PCB, solution pack |  | The reagent package could not read the information and the connection sheet broke | The reagent package could not read the information and the connection sheet broke | Repair by Hunter after replacing stock 903-158
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 液体传输失败0508，泵定标错误1178 | 更换试剂包后修复 | SWO-530280 | [5 empty] | Test card, reagent package failure | Liquid transmission failure 0508, pump calibration error 1178 | Repair after replacement of reagent kit
No | Visit | Complaint | RCHN | 试管架断裂 | 试管架断裂, 借886-731（903-584）两个。 |  | SWO-529417 | Defect | HW | Inlet module | Tube holder |  | The test tube rack is broken. | The test tube rack is broken, borrowing 886-731 (903-584).
No | Visit | Complaint | RESP | Error de temperatura | Error temperatura | Comprobación del estado del analizador. Se procede a entrar en sistema y se observa que no hay error de temperatura, se reinicia casete dando el error. Se procede a pedir al cliente un SC nuevo, no dispone por lo que se coje de UCI pediatrica. El SC parece que funciona. Se realizan varias Cal para estabilizar y Qc’s para comprobar funcionamiento. Se indica que Glu y Lac. se esta Estabilizando tanto a supervisora como a enfermeras. No se disponen de mas casetes para su cambio. | SWO-531252 | Assisted Replacement | SC | SC | Thermistor | 1221 | Temperature error | Temperature error | Checking the status of the scanner. It proceeds to enter the system and it is observed that there is no temperature error, cassette restart giving the error. We proceed to ask the customer for a new SC, so it does not have iCU pediatric. The SC seems to be working. Several Cals are performed to stabilize and Qc's to check performance. Glu and Lac are indicated. both supervisor and nurses are stabilizing. No more cassettes are available for change.
No | Visit | Complaint | RDEU | Anhaltendes Gerinnsel erkannt | Gerinnsel erkannt | Analysator mit Spülwerkzeug gespült, Hämolysator gespült alles freigängig, Abfallschlauch ersetzt, Startphase, Sensorcassette ersetzt, Startphase | SWO-530388 | Revisit required | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | Persistent clot detected | Clots detected | Analyzer with rinsing tool rinsed, hemolyser flushed everything free-running, waste hose replaced, start phase, sensor cassette replaced, start phase
No | Visit | Complaint | RMAY | Service breakdown | Found blockage from inside of waste nipple | - Check analyzer data log - Perform manual flush from sensor bay - Run fluid transport test | SWO-533108 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | Service breakdown | Found blockage from inside of waste nipple | - Check analyzer data log - Perform manual flush from sensor bay - Run fluid transport test
No | Visit | Complaint | RLTD | ISSUES ON NIGHT SHIFT , MACHINE SAID FLUSHING REQ,NOT SURE WHAT THEY DID, BUT IT SAID SOLUTION PACK CHANGE THEN IT HAD MULTIPLE ERRORS, EVERY SOLUTION PACK THEY CHANGED SAID DATA Roshini - 0208 970 8437 | POCT had resolved the issue by the time I had arrived. They stated that they believe the user that replaced the SP initially was not trained to carry out the replacement. POCT had reset SP latch and fitted correctly. | No further action required. | SWO-532555 | No fault found | No fault found | User error | SP activation |  | ISSUES ON NIGHT SHIFT , MACHINE SAID FLUSHING REQ,NOT SURE WHAT THEY DID, BUT IT SAID SOLUTION PACK CHANGE THEN IT HAD MULTIPLE ERRORS, EVERY SOLUTION PACK THEY CHANGED SAID DATA Roshini - 0208 970 8437 | POCT had resolved the issue by the time I had arrived. They stated that they believe the user that replaced the SP initially was not trained to carry out the replacement. POCT had reset SP latch and fitted correctly. | No further action required.
No | Visit | Complaint | RFRA | Depuis une semaine, l'écran de l'appareil reste noir. L'appareil est bien sous tension, les lumières du passeur de l'agitateur de seringue sont bien allumé, et le laser du lecteur code barre s'allume également. Lorsque l'on éteint et rallume l'appareil, le BIOS démarre correctement puis l'écran bascule et reste noir. Au bout d'un certains, l'appareil se met complètement hors tension. | Y-a-t-il eu coupure secteur ? Écran noir, chgt Compact Flash. pmazaleyrat@clinique-pasteur.com | Installation Windows7 et logiciel. Configuration Windows Recharge données et configuration ABL90Plus. Chgt Pack Solution et cassette. Calibration tHB( non calibrée depuis 25 $02 $2019) Nettoyage site introduction (en attente depuis 14 $02 $2019) Lors du changement du pack de solution, penser à nettoyer le capteur rectangulaire sombre au-dessus du pack côté ABL. | SWO-529990 | Defect | HW | DMS_PC | CF card |  | For the past week, the device's screen has remained black. The device is well energized, the syringe agitator's smuggler's lights are well on, and the barcode reader's laser also lights up. When the device is turned off and re-on, the BIOS starts properly and then the screen flips over and stays black. After a few, the device goes completely off. | Was there a cut off? Black screen, chgt Compact Flash. pmazaleyrat@clinique-pasteur.com | Windows7 installation and software. Windows setup Recharge data and ABL90Plus configuration. Chgt Pack Solution and cassette. THB calibration (not calibrated since $25 $02 2019) Cleaning site introduction (pending since $14$02 2019) When changing the solution pack, consider cleaning the dark rectangular sensor above the ABL side pack.
Yes | Visit | Complaint | RCHN | 电池无法蓄电 | 电池无法蓄电，借431-018一个 | 更换431-018后修复 | SWO-528150 | [5 empty] | The battery can't store power. | Battery can not store power, borrow 431-018 one | Repair after replacement 431-018
Yes | Visit | Complaint | RCHN | 电池无法蓄电 | 电池无法蓄电，借431-018一个 | 更换431-018后修复 | SWO-528149 | [5 empty] | The battery can't store power. | Battery can not store power, borrow 431-018 one | Repair after replacement 431-018
Yes | Visit | Complaint | RCHN | 电池无法蓄电 | 电池无法蓄电，借431-018一个 | 更换431-018后修复 | SWO-528152 | [5 empty] | The battery can't store power. | Battery can not store power, borrow 431-018 one | Repair after replacement 431-018
Yes | Visit | Complaint | RCHN | 电池无法蓄电 | 电池无法蓄电，借431-018一个 | 更换431-018后修复 | SWO-528151 | [5 empty] | The battery can't store power. | Battery can not store power, borrow 431-018 one | Repair after replacement 431-018
No | Visit | Complaint | RCHN | 定标失败 | 定标失败，借944-214（15452），944-258（15247）各一个 |  | SWO-529585 | Assisted Replacement | Consumables | Test Kit | Cartridge, cal | Potentially Use Error | Bid failed | Failed to bid, borrow 944-214 (15452), 944-258 (15247) each
No | Visit | Complaint | RCHN | 存储盘卡盘导致测试卡断裂 | 存储盘卡盘导致测试卡断裂，损坏部位：teeth | 更换测试卡后修复by Hunter | SWO-529614 | Defect | Consumables | Test Kit | Cartridge, Broken | Teeth | Memory disk card causes test card to break | Memory disk card causes test card to break, damaged site: teeth | Fix by Hunter after replacing test card
No | Visit | Complaint | RKK | QCレベル１にて、pO2、ノーエンドポイント | Po2 N $E | 代替え設置 | SWO-534011 | [5 empty] | At QC level 1, pO2, no endpoint | Po2 N /E | Alternative installation
No | Visit | Complaint | RMAY | Service breakdown | Chamber clotted | - Check analyzer data log - Clean chamber - Flush inlet - Perform protein removal | SWO-533107 | Cleaned | Fluidic pathway | pH/BG Module | pH/BG Module |  | Service breakdown | Chamber clotted | - Check analyzer data log - Clean chamber - Flush inlet - Perform protein removal
No | Visit | Complaint | RAME | Customer states after cleaning a clot and troubleshooting a leak, the leak has become greater and the analyzer is leaking everywhere. | Verified customer's stated issue. Customer had incorrectly replaced the waste tubings. The uppermost tubing was stretched out further than the bottom waste tubing and not sitting correctly in the upper notch. | Replaced both waste pump tubings (provided by the customer). | SWO-530248 | Defect | HW | Waste tubings | Waste tubings |  | Customer states after cleaning a clot and troubleshooting a leak, the leak has become greater and the analyzer is leaking everywhere. | Verified customer's stated issue. Customer had incorrectly replaced the waste tubings. The uppermost tubing was stretched out further than the bottom waste tubing and not sitting correctly in the upper notch. | Replaced both waste pump tubings (provided by the customer).
No | Visit | Complaint | RDEU | Kalibrationen sporadisch abgebrochen | PH Bg LS Error | Airventil Membran durchgeschlagen , ersetzt, mehrere Kals bestanden | SWO-530387 | Revisit required | HW | pH/BG Module | Valves |  | Calibrations aborted sporadically | PH Bg LS Error | Airvalve membrane pierced, replaced, several kals passed
No | Visit | Complaint | RAME | Occasional printer slept off, not printing | Confirmed issue when onsite. Printer not responding to command. | Proceed to replace printer assembly. Print off counters successfully. Will monitor status with customer if new printer will not fall asleep, needing to reboot to wake it up. Otherwise, windows will need to be reinstalled. | SWO-530165 | Defect | HW | Printer Module | Printer Module |  | Occasional printer slept off, not printing | Confirmed issue when onsite. Printer not responding to command. | Proceed to replace printer assembly. Print off counters successfully. Will monitor status with customer if new printer will not fall asleep, needing to reboot to wake it up. Otherwise, windows will need to be reinstalled.
No | Visit | Complaint | RAME | Occasional Drift error on K, resulting in QC failures as well | Confirmed issue when onsite. Reviewed data log and noticed numerous drift errors on K. | Replaced Ref and k membrane again (customer replaced k membrane on Friday). Cal passed but if error returns, then k electrode will need to be changed. | SWO-530180 | Assisted Replacement | Electrode | Membrane | Reference |  | Occasional Drift error on K, resulting in QC failures as well | Confirmed issue when onsite. Reviewed data log and noticed numerous drift errors on K. | Replaced Ref and k membrane again (customer replaced k membrane on Friday). Cal passed but if error returns, then k electrode will need to be changed.
No | Visit | Complaint | RKK | メンブラン交換後、CO2が感度エラーが発生し何をやっても戻らず。その前からドリフトエラーがぽつりぽつりと出ていた | メンブラン交換後、CO2が感度エラーが発生し何をやっても戻らず。その前からドリフトエラーがぽつりぽつりと出ていた | CO2電極交換。 | SWO-534001 | Assisted Replacement | Electrode | Electrode | pCO2 |  | After the membrane exchange, CO2 does not return even if the sensitivity error occurs and anything is done. Drift errors were popping up from before that. | After the membrane exchange, CO2 does not return even if the sensitivity error occurs and anything is done. Drift errors were popping up from before that. | CO2 electrode replacement.
No | Visit | Complaint | RDEU | 02.04.2019 Sr Kerstin 05971421230 meldet QC Fehler bei den Elektrolyten. 04.04.2019, SR. Kerstin, Tel: 05971421230: weiterhin immer wieder Driften bei den E´lyten und QC Fehler | Kal Driften , QC ausser Bereich | Gerät geprüft , Messkammer gesäubert , Dichtring Ph Elektrode erneuert , Ref Elektrode erneuert und membraniert , O2 , Co2 und Cl membraniert. Diverse Kal und QC gestartet. | SWO-531104 | Assisted Replacement | Electrode | Electrode | reference |  | 02.04.2019 Sr Kerstin 05971421230 reports QC errors in the electrolytes. 04.04.2019, SR. Kerstin, Tel: 05971421230: continue to drift at the E'lytes and QC errors | Kal Driften , QC out of range | Device tested, measuring chamber cleaned, sealing ring Ph electrode renewed, ref electrode renewed and membraneed, O2, Co2 and Cl membraned. Various Kal and QC started.
No | Visit | Complaint | RDEU | 08.04.2019, Sr. Madeleine, Tel: 0375514700: AQT steht seit dem Wochenende im Fehler | Fehlersuche ==> SETSAFE_22 und ROTATESTORAGEWHEEL_24 | Kassetten entfernt und Storagewheel ausgebaut und gereinigt | SWO-529967 | Defect | Consumables | Test Kit | Cartridge, Broken | Mail: Teeth? | 08.04.2019, Sr. Madeleine, Tel: 0375514700: AQT has been in error since the weekend | Troubleshooting ==> SETSAFE_22 and ROTATESTORAGEWHEEL_24 | Removed cassettes and storagewheel removed and cleaned
No | Visit | Complaint | RDEU | Anhaltendes Gerinnsel erkannt 09.04.2019 8:00 H. Molitor +492191138585: Erneute Probleme nach Service-Einsatz am 08.04.2019 | Flow Selektor Fehler erkannt | Gerät ausgeschaltet, LP entfernt , Flow Selektor gereinigt , Analysator ohne LP neu gestartet und erneut installiert, | SWO-531156 | Cleaned | HW | Basic Unit | Flow selector |  | Persistent clot detected 09.04.2019 8:00 H. Molitor +492191138585: Renewed problems after service use on 08.04.2019 | Flow selector error detected | Device off, LP removed, Flow selector cleaned, analyzer without LP restarted and reinstalled,
No | Visit | Complaint | RESP | Sucessivos pedidos de substituição de consumiveis | Sucessivos pedidos de substituição de consumiveis | Verificação do registro de dados. Reinstalação da cassete de eléctrodos. Erro de calibração de K e Cl. Substituição da cassete de eléctrodos. Inicialização. Calibração e verificação do correto funcionamento com controle de qualidade. | SWO-530555 | Assisted Replacement | SC | SC | SC | K and Cl, 1167 | Successive applications for substitution of consumables | Successive applications for substitution of consumables | Checking the data record. Reinstall the electrode cassette. Calibration error of K and Cl. Replacing the electrode cassette. Startup. Calibration and verification of the correct operation with quality control.
No | Visit | Complaint | RHUN | OXI teszt hiba, sO2 pontatlan | OXI teszt hiba, sO2 pontatlan | Műszer ellenőrzése, Hemolizátor garanciális csere, Folyadék áramlási tesztek + kalibrációk, OXI tesztek, Inincializálás, Kalibrálás, tHb,QC mérés | SWO-530590 | Defect | HW | Oxi module | Hemolyzer | sO2 | OXY test failure, sO2 inaccurate | OXY test failure, sO2 inaccurate | Instrument Check, hemolator warranty replacement, liquid flow tests + calibrations, OXY tests, inincialization, calibration, tHb, QC measurement
No | Visit | Complaint | RDEU | 08.04.2019 12:20Uhr Herr Hagen tel:+49401818854581: keine Verbindung mit Netzwerk | System überprüft | Alle Einstellung bezüglich Aqure korrekt. | SWO-530512 | No fault found | No fault found | No fault found | HIS/LIS | AQURE | 04/08/2019 12:20PM Mr Hagen tel:+49401818854581: no connection to network | System checked | All setting regarding Aqure correct.
No | Visit | Complaint | RLTD | Customer removed analyser from AAU as it was asking for a flush, customer thought this was related previous problem that happened recently. | On arrival analyser was in UIR and asking for manual flush, clot suspected in inlet. On entering service the value for inlet in clot detect remove service program was out of limits. | Run clot detect remove cycle in service, all values now passing and at the mean value. Exit service and selected 'test again', analyser performed startup and calibrations and QCs. Just before leaving the analyser went into UIR again saying clot detected in inlet, runnning clot removal in servic cleared the error again. Fitted a different solution pack and analyser completed startup and all 3 QCs satisfactorily. Problem most probably caused by a user introducing a clot into the analyser. | SWO-530901 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | Customer removed analyser from AAU as it was asking for a flush, customer thought this was related previous problem that happened recently. | On arrival analyser was in UIR and asking for manual flush, clot suspected in inlet. On entering service the value for inlet in clot detect remove service program was out of limits. | Run clot detect remove cycle in service, all values now passing and at the mean value. Exit service and selected 'test again', analyser performed startup and calibrations and QCs. Just before leaving the analyser went into UIR again saying clot detected in inlet, runnning clot removal in servic cleared the error again. Fitted a different solution pack and analyser completed startup and all 3 QCs satisfactorily. Problem most probably caused by a user introducing a clot into the analyser.
No | Visit | Complaint | RFRA | Lorsque l'on procède au nettoyage du joint d'introduction la date ne s'incremente pas. Exemple : Dernier nettoyage le 05.02, prochain le 12.02. Lorsque l'on lance un nettoyage du joint d'intro la date du dernier effectué s'incremente mais pas celle du prochain : Dernier le 06.02, prochain le 12.02. Le probleme est présent sur les 3 appareils du site, ils sont tous en version 3.4. Mise à jour prévue pour tenter de corriger ce bug. | Mise à jour logiciel version 3.4MR1 | Sauvegarde toutes données Sauvegarde de la configuration Mise à jour logiciel version 3.4MR1 Relance analyseur | SWO-530553 | Defect | SW | RMED SW | replacement schedule |  | When cleaning the introductory seal, the date does not break down. Example: Last clean-up on 05.02, next on 12.02. When one launches a cleaning of the intro joint the date of the last performed is increments but not that of the next: Last on 06.02, next on 12.02. The problem is present on the 3 devices of the site, they are all in version 3.4. Update planned to try to fix this bug. | Software update version 3.4MR1 | All-data backup Save the setup Software update version 3.4MR1 Relaunch analyzer
No | Visit | Complaint | RLTD | Parameter failures | Multiple parameter failures $oximetry failure, sensor cassete and solution pack replaced however still occuring. Tried flushing and performed pressure test $oxi full test - all okay. Attempting thb calibration. | Thb cal passed - calibration and QC's all passing | SWO-531148 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway | Oxi | Parameter failures | Multiple parameter failures /oximetry failure, sensor cassete and solution pack replaced however still occuring. Tried flushing and performed pressure test /oxi full test - all okay. Attempting thb calibration. | Thb cal passed - calibration and QC's all passing
No | 4 | Y | RCZE | Blank | Cez telefon konzultovana kontrola kabelaze, restart bezuspesne. |  | D5F27C345C546FA3 | Adjustment | HW | Basic Unit | Cable |  | Blank | CEZ phone Konzultovana control Kabelaze, restart Bezuspesne.
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | LAC定标失败 | 更换测试卡后修复 | SWO-533338 | [5 empty] | Test card, reagent package failure | LAC calibration failure | Repair after replacing test card
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | K,CA定标失败 | 更换测试卡后修复 | SWO-531059 | [5 empty] | Test card, reagent package failure | K, CA calibration failed | Repair after replacing test card
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | LAC定标不过 | 更换测试卡后修复 | SWO-531012 | [5 empty] | Test card, reagent package failure | LAC calibration is not | Repair after replacing test card
No | Visit | Complaint | RCHN | 无法开机 | 无法开机,借910-285一个。 | 更换910-285后修复by HUNTER | SWO-529844 | Defect | HW | Power Supply | Power Supply |  | Cannot power on | Cannot power on, borrow 910-285 one. | Repair by HUNTER after replacement 910-285
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 无法读取测试卡智能芯片1123 | 更换测试卡后修复 | SWO-533311 | [5 empty] | Test card, reagent package failure | Cannot read test card smart chip 1123 | Repair after replacing test card
No | Visit | Complaint | RCHN | 进样仓门及存储仓门打不开 | 进样仓门及存储仓门打不开，借903-177,903-028各一个 | 更换903-177,903-028后修复by HUnter | SWO-529790 | Defect | HW | Cover and Chassis | PCB, cabinet top |  | The sample door and storage warehouse door cannot be opened | The sample door and storage warehouse door can not be opened, borrow 903-177, 903-028 each one | Repair by HUnter after replacing 903-177,903-028
No | Visit | Complaint | RCHN | 测试卡舱门开关异常，错误代码：2995 | 测试卡舱门开关异常，错误代码：2995 | 重新安装测试卡舱门弹簧后修复by Hunter | SWO-530418 | Adjustment | HW | Storage Module | Door, storage | Spring, 2995 | Test card door switch abnormal, error code: 2995 | Test card door switch abnormal, error code: 2995 | Repair by Hunter after reinstalling test card door spring
No | Visit | Complaint | RESP | Erro de calibração de pO2 | Erro da corrente zero de pO2 | Substituição do eléctrodo de pO2. Limpeza da câmara de medida. Inicialização. Calibração e verificação do correto funcionamento com controle de qualidade. | SWO-530763 | Cleaned | Fluidic pathway | pH/BG Module | pH/BG Module | 409 | PO2 Calibration Error | PO2 Zero Current Error | Replacement of the pO2 electrode. Cleaning the measuring chamber. Startup. Calibration and verification of the correct operation with quality control.
No | Visit | Complaint | RIND | Rinse error | Rinse error | checked the analyser being new rinse bottle so cleaned the gasket and needle was slightly bent made correct then checked the analyser but still its has rinse error so LS adjust done and pump calibration done still problem not solved so switch off the analyser for few min then switched on then machine came to ready mode | SWO-533039 | No fault found | Clean fluidic pathway |  |  |  | Rinse error | Rinse error | checked the analyser being new rinse bottle so cleaned the gasket and needle was slightly bent made correct then checked the analyser but still its has rinse error so LS adjust done and pump calibration done still problem not solved so switch off the analyser for few min then switched on then machine came to ready mode
No | 2 | Y | RCZE | Blank |  |  | D5F42CADD22C0CE3 | User Error | Training | Training | Training |  | Blank
No | Visit | Complaint | RKK | ミキサーエラー発生 | 装置内部増し締め及びミキサーコンプリート交換を実施しました。 | 装置内部増し締め及びミキサーコンプリート交換を実施しました。 | SWO-530277 | Defect | HW | Auto Inlet module | Mixer tray |  | Mixer error | The internal tightening of the equipment and the mixer complete replacement were carried out. | The internal tightening of the equipment and the mixer complete replacement were carried out.
No | Visit | Complaint | RFRA | La machine plante avec un écran bleu. Pas de diagnostic plus précis pour l'instant car la connexion VPN ne fonctionne pas. Analyse des logs windows à effectuer. | verification de l'historique des activitées | démontage de l'écran tactile : fixation de la coque plastique HS Fixation de la coque avec support et joint etanchéité en attente de commande d'un ecran complet | SWO-530541 | Defect | HW | Cover and Chassis | Display | Frame | The machine plants with a blue screen. No more accurate diagnosis at the moment because the VPN connection does not work. Analysis of windows logs to be performed. | verification of the history of activities | touch screen dismantling: fixing the HS plastic shell Fixing the hull with support and seal pending control of a complete screen
Yes | Visit | Complaint | RFRA | La machine plante avec un écran bleu. Pas de diagnostic plus précis pour l'instant car la connexion VPN ne fonctionne pas. Analyse des logs windows à effectuer. | TT2 pont l’abbé - Poitiers | TT2 | SWO-530918 | Re classified RMED | Same day UpSWO | Same day UpSWO | Same day UpSWO | Same day UpSWO | The machine plants with a blue screen. No more accurate diagnosis at the moment because the VPN connection does not work. Analysis of windows logs to be performed. | TT2 Abbot Bridge - Poitiers | TT2
No | Visit | Complaint | RPOL | Błąd kalibracji pompy | Sprawdzenie sondy aspiracyjnej oraz uszczelki wlotu. | Serwis stwierdza zablokowanie sondy aspiracyjnej skrzepem lub fragmentem gazy. Usunięcie ciała stałego, test transportu płynów. | SWO-530513 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | Pump Calibration Error | Check the aspiration probe and inlet gasket. | The service concludes a blockage of the aspiration probe with a clot or a fragment of gauze. Solids removal, fluid transport test.
No | Visit | Complaint | RPOL | Częste płukanie, niedziałająca bateria | Sprawdzenie działania baterii oraz diagnostyka sekcji mokrej. | Usunięcie przerwy w obwodzie ładowania baterii, test transportu płynów. Serwis nie stwierdza nieprawidłowośći w pracy sekcji mokrej. | SWO-530450 | Defect | HW | Battery module | Battery module |  | Frequent flushing, non-functioning battery | Check the battery operation and diagnostics of the wet section. | Removal of the battery charging circuit, fluid transport test. The service does not detect irregularities in the work of the wet section.
No | Visit | Complaint | RPOL | Podczas uruchamiania pojawia się komunikat Cannot load application . Problem powstał w wyniku nagłego wyłączenia aparatu (wyładowanie baterii, aparat został wcześniej podłączony do gniazdka 24V). | Błąd podczas uruchamiania aplikacji. | Przeinstalowanie aplikacji aparatu (wersja 3.4). Wczytanie ustawień. Wymiana igły oraz uszczelki wlotu. Kalibracja hemoglobiny. | SWO-530719 | Defect | SW | RMED SW | database |  | The message "Cannot load Application" appears during startup. The Problem was caused by a sudden shutdown of the camera (battery discharge, the camera was previously plugged into a 24V socket). | Error starting application. | Reinstall the Camera app (version 3.4). Load the settings. Replacing the needle and inlet gasket. Hemoglobin calibration.
No | Visit | Complaint | RDEU | Anhaltendes Gerinnsel erkannt 09.04.2019 8:00 H. Molitor +492191138585: Erneute Probleme nach Service-Einsatz am 08.04.2019 09.04.2019; 19:12 Uhr; Fr. Lehmann; Tel.: +492191138585; der Fehler ist wieder aufgetreten | Fehlermeldung Gerinnsel erkannt | Zur Überbrückung Leihgerät aus VSC West geholt und aufgebaut (SN : 090R0309N007) Leihvertrag erstellt | SWO-531384 | Revisit required | analyzer removed | Fluidic pathway | FT |  | Persistent clot detected 09.04.2019 8:00 H. Molitor +492191138585: Renewed problems after service use on 08.04.2019 09.04.2019; 7:12 p.m.; Ms. Lehmann; Tel. +492191138585; the error has occurred again | Error message clot detected | To bridge loaner from VSC West fetched and built up (SN : 090R0309N007) Loan contract created
No | Visit | Complaint | RSAF | Solution pack needed to be changed | None | Replaced the solution pack | SWO-531413 | Assisted Replacement | SP | SP | Unknown |  | Solution pack needed to be changed | None | Replaced the solution pack
No | Visit | Complaint | RLTD | 10/04/19 - Girlie Andrada on 0191 4453195 Leak current detected in relation to SP | Leak current detected in relation to SP persists after SP and SC replacements | All tubing inspected and reseated, hemolyser and sensor bay flushed and white powder cleared off flow selector sensor and chip reader. inlet gasket was cleaned and inlet probe straightened. impedance measured in the service menu and reading 84.25Mv Analyser went into 2 hour start up passing calibrations and QCs. Measured impedance again reading 394.25Mv. Measured again removing SP and swapping SC but no real change in values. On leaving analyser impedance reading 184Mv. Customer to monitor over night. | SWO-531573 | Revisit required | SC | SC | leak current |  | 10/04/19 - Girlie Andrada on 0191 4453195 Leak current detected in relation to SP | Leak current detected in relation to SP persists after SP and SC replacements | All tubing inspected and reseated, hemolyser and sensor bay flushed and white powder cleared off flow selector sensor and chip reader. inlet gasket was cleaned and inlet probe straightened. impedance measured in the service menu and reading 84.25Mv Analyser went into 2 hour start up passing calibrations and QCs. Measured impedance again reading 394.25Mv. Measured again removing SP and swapping SC but no real change in values. On leaving analyser impedance reading 184Mv. Customer to monitor over night.
No | Visit | Complaint | RDEU | 10.04.2019 10:00 Fr. Rodriges +491727222311: QC-Fehler K, Lac, Glu, pO2 10.04.2019 13:10 Fr. Rodriges +491727222311: QC-Fehler K, Lac, Glu, pO2 | Häufige Kal. und QC Fehler, besoners bei Glu und Lac | Teile im Umfang einer Jahreswartung erneuert. ABL90 gespült, Kontakte überprüft und gereinigt. SC und LP wieder installiert. Einlaufverfahren Kal. und QC Fehlerlos. Mit Fr. Rodriges vom Labor über Präanalytik gesprochen. Sie wird veranlassen das Kapillare auf jeden Fall mit Mischerstäbchen gemischt werden. z.Teil haben Anwender Probleme beim Ansetzen der Kapillare mit Gerinselfänger. JonLu wird benachrichtigt zur weiteren Abklärung. | SWO-531813 | No fault found | No fault found | No fault found | QC | Glu/Lac | 10.04.2019 10:00 Fr. Rodriges +491727222311: QC-Fehler K, Lac, Glu, pO2 10.04.2019 13:10 Fr. Rodriges +491727222311: QC-Fehler K, Lac, Glu, pO2 | Frequent Kal. and QC error, especially at Glu and Lac | Parts to the extent of an annual maintenance. ABL90 rinsed, contacts checked and cleaned. SC and LP reinstalled. Inlet method Kal. and QC errorless. Talked to Ms. Rodriges from the laboratory about preanalysis. It will cause the capillary to be mixed with mixer rods in any case. for example, users have problems attaching the capillary with a sconing catcher. JonLu will be notified for further clarification.
No | Visit | Complaint | RAME | clot in system unit giving intervention required message | clot in system customer could not flush system | customers flush device is missing , flushed analyzer to remove clot started unit up and tested operation. | SWO-532135 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | clot in system unit giving intervention required message | clot in system customer could not flush system | customers flush device is missing , flushed analyzer to remove clot started unit up and tested operation.
No | Visit | Complaint | RAME | Qualicheck adapter 925-214 fits too tight on the inlet gasket making it difficult to remove. | Qualicheck adapter 925-214 fits too tight on the inlet gasket making it difficult to remove. | Replaced 3 of the customers qualicheck adapters. | SWO-532114 | Defect | HW | Auto Inlet module | QC adaptor |  | Qualicheck adapter 925-214 fits too tight on the inlet gasket making it difficult to remove. | Qualicheck adapter 925-214 fits too tight on the inlet gasket making it difficult to remove. | Replaced 3 of the customers qualicheck adapters.
No | Visit | Complaint | RBNL | Zoals bij het toestel op spoed is de blauwe omkadering van het scherm losgekomen. (Alle 4 connectiepunten gebroken) | Blauwe omkadering van scherm is losgekomen | Vervanging van display unit en test : OK | SWO-531322 | Defect | HW | Display | Display | Screws | As with the device, the blue bezel of the screen has come loose. (All 4 connection points broken) | Blue screen framing has come loose | Replacement of display unit and test: OK
No | Visit | Complaint | RLTD | David Ryder Tel: 07580 818900 | Site visit to repair broken analyser display | Arrived on site and attended analyser. It was switched off with screen bezel taped closed to hold it into position. Investigation showed all 4 screw turrets had sheered. No signs of other damage to screen or casing. Replaced whole display for new and powered back up. Calibrated screen and rebooted. Waited for analyser to startup. Got an intervention required error. Cleaned and checked gasket and probe. Retested. 55 minute startup. Started came down to 15 minutes after 15 minutes. Waited for completion. Analyser passed cals and qcs. Tested barcode scanner which worked fine. No further action required. | SWO-531357 | Defect | HW | Display | Display | Screws | David Ryder Tel: 07580 818900 | Site visit to repair broken analyser display | Arrived on site and attended analyser. It was switched off with screen bezel taped closed to hold it into position. Investigation showed all 4 screw turrets had sheered. No signs of other damage to screen or casing. Replaced whole display for new and powered back up. Calibrated screen and rebooted. Waited for analyser to startup. Got an intervention required error. Cleaned and checked gasket and probe. Retested. 55 minute startup. Started came down to 15 minutes after 15 minutes. Waited for completion. Analyser passed cals and qcs. Tested barcode scanner which worked fine. No further action required.
No | Visit | Complaint | RAME | Set up issues/set up is not the same as the other | Instrument not sending data to Rals. Send button is there. | Inspected all settings and look correct. Light is on at the back of instrument. Working with Rals to find problem. | SWO-531642 | No fault found | Setup | RMED SW | HIS/LIS | Rals | Set up issues/set up is not the same as the other | Instrument not sending data to Rals. Send button is there. | Inspected all settings and look correct. Light is on at the back of instrument. Working with Rals to find problem.
No | Visit | Complaint | RDEU | pCO2 ist seit der Geräte-Installation im Level C an der oberen Grenze oder zu hoch | Datenbank geprüft | Beim Eintreffen hat der SP noch 16 restliche Zyklen SP erneuert --> pCO2 Level C bei 2,79 kPa Sensorkassette erneuert --> pCO2 Level A bei 3,66 kpA, Level B 8,92 kPa, Level C 2,83kPa | SWO-531187 | Assisted Replacement | SC | SC | pCO2 | QC C | pCO2 has been at the upper limit or too high since the device was installed in Level C. | Database checked | On arrival, the SP still has 16 remaining cycles SP renewed --> pCO2 Level C at 2.79 kPa Sensor cassette renewed --> pCO2 Level A at 3.66 kpA, Level B 8.92 kPa, Level C 2.83kPa
No | Visit | Complaint | RCHN | 由于近期乳酸卡连续出现问题 前往安抚客户 | 由于近期乳酸卡连续出现问题 前往安抚客户 | 由于近期乳酸卡连续出现问题 前往安抚客户(by chuanwei) | SWO-532008 | [5 empty] | Due to the recent lactic acid card continuous problems to appease customers | Due to the recent lactic acid card continuous problems to appease customers | Due to a recent series of lactic acid card problems to pacify customers (by chuanwei)
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 试剂包连接不上 | 更换试剂包后修复 | SWO-534010 | [5 empty] | Test card, reagent package failure | The reagent package is not connected | Repair after replacement of reagent kit
No | Visit | Complaint | RCHN | 无法开机 | 无法开机，借910-285一个 | 更换910-285后修复BY Hunter | SWO-530379 | Defect | HW | Power Supply | Power Supply |  | Cannot power on | Cannot power on, borrow 910-285 one | Fix BY Hunter after replacing 910-285
No | Visit | Complaint | RCHN | 报错1999.5.209 | 报错1999.5.209，借902-978一个 | 更换902-978后修复（by chuanwei） | SWO-530470 | Defect | HW | Aspiration Module | Cavro Pump |  | Error 1999.5.209 | Error 1999.5.209, loan 902-978 a | Repair after replacement 902-978 (by chuanwei)
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | NA定标不过 | 更换测试卡后修复 | SWO-532560 | [5 empty] | Test card, reagent package failure | NA calibration is not | Repair after replacing test card
No | Visit | Complaint | RCHN | 光谱不匹配和AB浓度过低，报错0688和0581 | 光谱不匹配和cthb浓度过低，报错0688和0581 | 更换902-527后修复(by chuanwei) | SWO-531726 | Defect | HW | Oxi module | Pinch valve |  | Spectrum mismatch and AB concentration is too low, error00688 and 0581 | Spectroscopy mismatch and cthb concentration is too low, error00688 and 0581 | Repair after replacement 902-527 (by chuanwei)
No | Visit | Complaint | RCHN | 质控结果不准确 | 质控结果不准确 | 重新清洁和保养后修复(by chuanwei) | SWO-531749 | Cleaned | Clean fluidic pathway | pH/BG Module | pH/BG Module | Half yearly maintenance | Inaccurate quality control results | Inaccurate quality control results | Re-cleaning and post-service repair (by chuanwei)
No | Visit | Complaint | RDEU | 10.04.2019, Herr Schick (MT), Tel: 02051982383705: keine Oximetrie Werte bei QC | Fehler 791 Parameter konnte nicht gemessen werden bei QCs und Patientenblutmessungen (betrifft: ctHb und sO2). Das erste Mal ist der Fehler heute um 2:20 Uhr aufgetreten (nach automatischer LS-Justage und Pumpenkal um 2 Uhr). | Quetschventil überprüft und gereinigt. Sonstige Ventilschaltungen optisch überprüft. Schlauchwege und Messkammern überprüft. Unteren pH-Dichtungsring erneuert. Der untere pO2-Dichtungsring fehlte - erneuert. Einlassschlauch erneuert. Lufteinlass durchgespült. LS-Justage und Pumpenkal durchgeführt. Gerät einkalibriert. Funktionskontrolle: I.O. | SWO-531345 | Defect | HW | po2 | sealing ring | 791 | 10.04.2019, Mr. Schick (MT), Tel: 02051982383705: no oximetry values at QC | Error 791 parameters could not be measured in QCs and patient blood measurements (affected: ctHb and sO2). The first time the error occurred today was at 2:20 a.m. (after automatic LS adjustment and pump scale at 2 o'clock). | Squeeze valve checked and cleaned. Other valve circuits optically checked. hoseways and measuring chambers. Lower pH sealing ring renewed. The lower pO2 sealing ring was missing - renewed. Inlet hose renewed. Air inlet rinsed through. LS-adjustment and pump scale. device calibrated. Functional control: I.O.
No | Visit | Complaint | RDEU | 10.04.2019 Herr Hex +49401818862280 möchte den ABL gerne ans Netzwerk angeschlossen haben. | Systemeinsetellungen überprüft | Alle Systemeinstellung sind korrekt, Gerät ist mit Netzwerk verbunden aber nicht mit Aqure | SWO-531270 | Setup | SW | HIS/LIS | HIS/LIS |  | 10.04.2019 Mr. Hex +49401818862280 would like to have connected the ABL to the network. | System insertion checks | All system settings are correct, device is connected to network but not connected to Aqure
No | Visit | Complaint | RIND | Lac Calibration Sensitivity Low. | Need to replace the Lac membrane. | Replaced the Lac membrane then run the instrument its working good. | SWO-531327 | Assisted Replacement | Membrane | Membrane | Lac |  | Lac Calibration Sensitivity Low. | Need to replace the Lac membrane. | Replaced the Lac membrane then run the instrument its working good.
No | Visit | Complaint | RDEU | 09.04.2019; Hr. Schmeer; Tel.: 0662188922531; Drucker druckt Streifen | Schlechtes Druckbild | Drucker auf Garantie erneuert. Startphase und Gerät überprüft | SWO-531183 | Defect | HW | Printer Module | Printer Module |  | 09.04.2019; Mr. Schmeer; Tel.: 0662188922531; Printer prints stripes | Poor print image | Printer renewed on warranty. Start-up phase and device checked
No | Visit | Complaint | RDEU | Gerät signalisiert den Fehler Einlassdurchfuss behindert. | Einlasssonde verstopft. Fliesßwege verunreinigt. | Einlasssonde und Proteinentferung durchgeführt, Gerät kalibriert und QC durchgeführt. | SWO-531480 | Cleaned | Fluidic pathway | Inlet module | Probe |  | Device signals the error inlet throughfoot obstructed. | inlet probe clogged. Flow paths contaminated. | Inlet probe and protein dispersing performed, device calibrated and QC performed.
Yes | Visit | Complaint | RIND | Pm as done | Checked the machine's activity log and system check procedures, As per periodic Schedule Pm to be done. | Preventive Maintenance as done | SWO-532571 | Re classified RMED | planned SWO |  |  | Preventive Maintenance | Pm as done | Checked the machine's activity log and system check procedures, As per periodic Schedule Pm to be done. | Preventive Maintenance as done
No | Visit | Complaint | RDEU | Vakuum Fehler | Siehe Case | Fehlersuche Einlassdichtung ersetzt, Startphase | SWO-532089 | Defect | HW | Inlet module | Inlet gasket |  | Vacuum error | See Case | Troubleshooting inlet seal replaced, start phase
No | Visit | Complaint | RLTD | Leak current detected in relation to SP | leak current detected again overnight. | SP swapped with known working SP and impedance still measuring above 200 mV. SC swapped with Known working SC and error moved to the working analyser. Impedance measuring at 77 mV on CCD analyser. With the SC swapped over both analysers went into 2 hour start up. CCD analyser has passed all calibrations and QCS. Rang morning of 12 $04 and error has not returned. | SWO-532833 | Assisted Replacement | SC | SC | leak current |  | Leak current detected in relation to SP | leak current detected again overnight. | SP swapped with known working SP and impedance still measuring above 200 mV. SC swapped with Known working SC and error moved to the working analyser. Impedance measuring at 77 mV on CCD analyser. With the SC swapped over both analysers went into 2 hour start up. CCD analyser has passed all calibrations and QCS. Rang morning of 12 $04 and error has not returned.
No | Visit | Complaint | RAME | Error code 1361 Internal reference electrode error in SC. | Customer reports multiple “1361 -internal reference electrode” errors are present. Data logs indicate the presence of this error over the past 2 months on separate sensor cassettes. | Replaced the sensor interface PCB and rebooted analyzer. Verified that sensor cassette is properly initialized. Analyzer is within acceptable operating specifications. Serial # for removed sensor interface PCB is unavailable. | SWO-532254 | Defect | HW | Basic Unit | PCB, Sensor interface | Reference | Error code 1361 Internal reference electrode error in SC. | Customer reports multiple “1361 -internal reference electrode” errors are present. Data logs indicate the presence of this error over the past 2 months on separate sensor cassettes. | Replaced the sensor interface PCB and rebooted analyzer. Verified that sensor cassette is properly initialized. Analyzer is within acceptable operating specifications. Serial # for removed sensor interface PCB is unavailable.
No | Visit | Complaint | RLTD | Please contact Tracey Lee Tel 0208 8872736 or 2495 or sam. | Bank of case screen fixings have broken. under warranty no physical Damage | Had to replace complete screen test ok. | SWO-532240 | Defect | HW | Display | Display | screws | Please contact Tracey Lee Tel 0208 8872736 or 2495 or sam. | Bank of case screen fixings have broken. under warranty no physical Damage | Had to replace complete screen test ok.
No | Visit | Complaint | RCHN | 科室老师质疑结果不准 | 科室老师质疑结果不准 | 对比结果并进行验证，仪器无问题。可能是操作不当引起的，已重新培训老师操作(by chuanwei) | SWO-532014 | [5 empty] | The teacher in the section questioned the result was not allowed. | The teacher in the section questioned the result was not allowed. | The results are compared and verified, and the instrument is free of problems. May be caused by improper operation, teacher operation has been retrained (by chuanwei)
No | Visit | Complaint | RCHN | 泵轮故障 | 泵轮故障，借902-971一个 | 更换902-971后修复(by chuanwei) | SWO-530350 | [5 empty] | Pump wheel failure | Pump wheel failure, borrow 902-971 one | Repair after replacement 902-971 (by chuanwei)
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | LAC定标失败 | 更换2张测试卡后修复 | SWO-535463 | [5 empty] | Test card, reagent package failure | LAC calibration failure | Repair after replacing 2 test cards
No | Visit | Complaint | RIND | QC Problem | Checked the machine with radiometer QC, and 2 point Calibration Results From Installation of SC. | Promoted Radiometer QC and Flushed the SC with Distilled Water and reinstalled the SC. | SWO-532584 | [5 empty] | QC Problem | Checked the machine with radiometer QC, and 2 point Calibration Results From Installation of SC. | Promoted Radiometer QC and Flushed the SC with Distilled Water and reinstalled the SC.
No | Visit | Complaint | RCHN | 漏电流 | 漏电流 | 更换废液管和夹紧阀后修复（by chuanwei） | SWO-531816 | Defect | HW | Oxi module | Pinch valve | Leak current | Leak current | Leak current | Repair after replacing waste pipes and clamping valves (by chuanwei)
No | Visit | Complaint | RCHN | 进样口垫片无法下降，发现进样针弯曲 | 进样口垫片无法下降，发现进样针弯曲 | 更换进样针后修复(by chuanwei) | SWO-531745 | Defect | HW | Inlet module | Probe bend |  | The inlet gasket could not drop and found the inlet needle bent | The inlet gasket could not drop and found the inlet needle bent | Repair after replacing the sample needle (by chuanwei)
No | Visit | Complaint | RCHN | 触屏不灵敏，显示器选项故障 | 触屏不灵敏，显示器选项故障，借910-298一个 | 更换910-298后修复(by chuanwei) | SWO-531200 | Defect | HW | Analyzer Control Module | Display |  | Insensitive touch screen, display option si | Touch screen is not sensitive, display options fail, borrow 910-298 one | Repair after replacement 910-298 (by chuanwei)
No | Visit | Complaint | RCHN | BNP传感器错误 | BNP传感器错误，借905-908一个 | 更换905-908后修复（by chuanwei） | SWO-531092 | Defect | HW | Incubator Module | Sensor, BNP | 1235 | BNP sensor error | BNP sensor error, borrow 905-908 one | Repair after replacement 905-908 (by chuanwei)
No | Visit | Complaint | RCHN | O2，CO2测量结果不准确 | O2，CO2测量结果不准确 | 做了质控，纠正老师的操作，建议老师使用雷度PICO针后解决(by chuanwei) | SWO-531729 | [5 empty] | O2, CO2 measurements are inaccurate | O2, CO2 measurements are inaccurate | Do the quality control, correct the teacher's operation, recommend the teacher to use the ray PICO needle after the solution (by chuanwei)
No | Visit | Complaint | RIND | PO2 level 1 QC getting failed | Checked the data logs PO2 level 1 QC getting failed repeatedly replaced po2 membrane its was working for some times and problem repeated again po2 level 1 qc was getting low values found problem with po2 electrode | Replaced po2 electrode and membrane performed qc values was within the range | SWO-532779 | Assisted Replacement | Electrode | Electrode | po2 |  | PO2 level 1 QC getting failed | Checked the data logs PO2 level 1 QC getting failed repeatedly replaced po2 membrane its was working for some times and problem repeated again po2 level 1 qc was getting low values found problem with po2 electrode | Replaced po2 electrode and membrane performed qc values was within the range
No | Visit | Complaint | RDEU | 11.04.2019 08:15Uhr Burger 06841162413: Das Gerät meldet ständig Kal gescheitert | pO2 und pCO2 drift. Kalibration gescheitert | Membranventil vom BG- Modul getauscht. Pumpekal. LS-Justage. Kalibration und QC . | SWO-532162 | Defect | HW | pH/BG Module | Valves |  | 11.04.2019 08:15Pm Burger 06841162413: The device constantly reports Kal failed | pO2 and pCO2 drift. Calibration failed | diaphragm valve replaced by the BG module. Pumpcal. LS-Adjustment. Calibration and QC .
No | Visit | Complaint | RDEU | 11.04.2019 Herr Virag +4993727001022 meldet einen Einlassfehler und möchte eine neue Geräte-IP-Adresse in den ABL eingehen. | Im Haus unterwegs zu weiteren Service. Analysator zeigt unter Meldungen obige Meldung. | Einlassdichtung erneuert, da alte verschmutzt. LS Justage und Pumpenkalibration durchgeführt, Fehler bleibt bestehen. Tausch des Liquid Sensors im Einlassmodul. Nach Neustart macht der Analysator eine LS Justage, danach wird der Fehler nicht mehr angezeigt. Mehrmals kalibriert bis Werte stabil. Auf Wunsch der EDV Analysator neue IP Adresse zugeteilt, auf DHCP umgestellt. Einstellung gespeichert und Datenübertragung überprüft, Funktion in Ordnung. | SWO-532200 | Defect | HW | Inlet module | LS |  | 11.04.2019 Mr. Virag +4993727001022 reports an inlet error and wants to enter a new device IP address in the ABL. | In the house on the way to further service. Analyzer displays the above message under Messages. | Inlet seal renewed as old ones dirty. LS adjustment and pump calibration, error remains. Replacement of the liquid sensor in the inlet module. After rebooting, the analyzer makes an LS adjustment, after which the error is no longer displayed. Calibrated several times to values stable. At the request of the COMPUTER analyzer new IP address assigned, switched to DHCP. Setting saved and data transfer checked, function ok.
No | Visit | Complaint | RFRA | Pb de cal pH et cq glucose et lactate | Contrôle de l’étanchéité de la machine, des vannes, des pompes | Remise en place des joints sur les électrodes de pO2, glucose. Nettoyage de la chambre de référence et des électrodes fantômes. Changement de la membrane de glucose, lactate, pO2 et créa. Essais | SWO-531904 | Cleaned | Fluidic pathway | pH/BG Module | measuring chamber |  | Pb of cal pH and cq glucose and lactate | Control of machine sealing, valves, pumps | Putting joints back on pO2 electrodes, glucose. Cleaning the reference chamber and phantom electrodes. Change of glucose membrane, lactate, pO2 and created. Tests
No | Visit | Complaint | RESP | No funciona el lector de codigos de barras | No lee los codigos de barras | Desmontaje y limpieza del lector de codigos de barras Cal, qc | SWO-532230 | Cleaned | HW | Barcode reader | Barcode Reader |  | The barcode reader doesn't work | Doesn't read the barcodes | Disassembly and cleaning of the barcode reader Cal, qc
No | Visit | Complaint | RIND | Lac sensitivity is below the range | Check system message and calibration result in which Lac shows yellow and it's sensitivity is below the range while membrane was replaced 2-3 days before. | Check Lac membrane status which is OK.Interchange Glu & Lac electrode with each other and found that Lac sensitivity increased and Glu sensitivity decreased.Replace Glu membrane,run cleaning and 2-point calibration which is OK.but Glu sensitivity started from 270 while from Glu electrode it was started from almost 900.So found the problem in Lactate electrode which need to replace to stop Lac frequent failure. | SWO-533225 | Assisted Replacement | Electrode | Electrode | Lac | 0378 | Lac sensitivity is below the range | Check system message and calibration result in which Lac shows yellow and it's sensitivity is below the range while membrane was replaced 2-3 days before. | Check Lac membrane status which is OK. Interchange Glu & Lac electrode with each other and found that Lac sensitivity increased and Glu sensitivity decreased. Replace Glu membrane,run cleaning and 2-point calibration which is OK.but Glu sensitivity started from 270 while from Glu electrode it was started from almost 900.So found the problem in Lactate electrode which need to replace to stop Lac frequent failure.
No | Visit | Complaint | RDEU | Needle Counter Grenze erreicht. | Needle und Syringe müssen aufgrund hoher Proben durchsätze erneuert werden. | Needle erneuert, Syringe an der Cavo Pumpe ersetzt, Counter Reset durchgeführt, | SWO-532566 | Adjustment | HW | Aspiration Module | Needle | Needle piercing limit reach | Needle Counter limit reached. | Needle and Syringe need to be renewed due to high samples. | Needle renewed, Syringe replaced at the Cavo pump, Counter Reset performed,
No | Visit | Complaint | RPOL | Błąd zgodności oprogramowania podczas pomiaru | W konfiguracji aparatu wystąpił błąd powodujący niezgodność wersji oprogramowania. | Przywrócenie ustawień domyślnych sekcji setup, zmiana ustawień automatycznego sprawdzania systemu, ponowny rozruch aparatu, testowanie. | SWO-532236 | Adjustment | SW | RMED SW | RMED SW |  | Software compatibility Error during measurement | There was an error in the camera configuration that caused software version mismatch. | Reset the Setup section to default settings, change the settings for automatic system check, reboot the camera, test.
No | Visit | Complaint | RPOL | Błąd podczas wykonywania procedury mieszania próbki za pomocą koła inletowego | Serwis stwierdza obluzowane mocowanie koła inletowego. | Ponowne mocowanie koła inletowego do osi napędowej, testowanie analizatora. | SWO-532235 | Adjustment | HW | Aspiration Module | Inlet Wheel | Mix wheel_25, potential FSE user error | Error when executing a sample blending procedure with an incuber wheel | The service finds a loose fitting of the inleted wheel. | Reattaching the inleted wheel to the drive axle, testing the analyser.
Yes | Visit | Complaint | RPOL | Błąd podczas wykonywania procedury mieszania próbki za pomocą koła inletowego | Serwis stwierdza obluzowane mocowanie koła inletowego. | Ponowne mocowanie koła inletowego do osi napędowej, testowanie urządzenia. | SWO-532234 | Re classified RMED | Same day UpSWO | Same day UpSWO | Same day UpSWO | Same day UpSWO | Error when executing a sample blending procedure with an incuber wheel | The service finds a loose fitting of the inleted wheel. | Reattaching the inleted wheel to the drive axle, testing the device.
No | Visit | Complaint | RDEU | Hotline CO2 ind Level C zu niedrig 11.04.2019; Fr. Sicking; Tel.: 025629157070; weiterhin pCO2 Probleme, Messzeit sehr lang | pCO2 zeitweise zu hoch Level B, C , Flowselektorfehler, Immer wieder Fehler 1189, 1249, 1084 | Gerät überprüft, Elektrodensymulation durchgeführt, Ref Fläche Flowselector kontrolliert, Einlaß gesäubert, Fließwege gepült, Geräte Neustart durchgeführt, Dichtungen geprüft, Einlaßdichtung und Lösungspack ersetzt | SWO-532736 | Assisted Replacement | SP | SP | pCO2 | QC C | Hotline CO2 ind Level C too low 11.04.2019; Ms. Sicking; Tel.: 025629157070; continue pCO2 problems, measurement time very long | pCO2 at times too high Level B, C , Flow selector error, Repeated error 1189, 1249, 1084 | Device checked, electrode symulation performed, ref surface flowselector checked, inlet cleaned, flow paths put, equipment restarted, seals tested, inlet seal and solution pack replaced
No | Visit | Complaint | RBNL | Vloeistoftransport fout | Reeds geruime tijd komt melding #1538 en #1539 (inhomogene cal 2 & 3 oplossing) voor in de ctiviteiten lijst, bijgevolg vroegtijdig vevangen van SP en inlet gasket. In service de vloeistofstroom gemonitord zonder fouten, SC en naald doorgespoten met lauw water, verbindings gasket vervangen en naald nagekeken (geen kromming vasthesteld) | Doorspuiten van naald , SC en vervanging van verbindingsgasket. | SWO-532858 | Defect | HW | Basic Unit | PM kit |  | Vloeistoftransport fout | For some time now there is mention #1538 and #1539 (inhomogeneous Cal 2 & 3 solution) in the Ctivities list, therefore early vecatching of SP and inlet gasket. In Service The liquid flow is monitored without errors, SC and needle sprayed with lukewarm water, replace connection gasket and needle checked (no curvature Vasthesteld) | By spraying needle, SC and replacing Verbindingsgasket.
No | Visit | Complaint | RFRA | mail client : Pour la première fois depuis la mise en place de l'ABL90 au SAU (le 11/03/2019), nous avons eu un cas de créatinine faussement élevée pour une patiente de 47 ans. Voici la chronologie des faits : - hier à 21h44 : dosage d'une gaz du sang artériel complet au SAU : créatininémie à 114 µmol/L (iono et Hb concordante avec résultats du labo) - hier prélèvement de 21h30 envoyé au labo (tube jaune) : créatininémie à 71 µmol/L (correspond aux antériorités et au résultat de ce matin) Nous sommes allés chercher la seringue dans la poubelle du SAU et nous l'avons redosé : - sur l'ABL90 du SAU : 73 µmol/L - et chez nous : 72 µmol/L Il y a donc eu un résultat faussement élevé hier soir. Avez vous déjà noté ce type de problème? Quelle peut en être la cause? A noter que jusque là, les résultats de créatinine obtenus sur l'ABL90 sont tout à fait bien corrélés à ceux obtenus au laboratoire. En vous remerciant par avance, RFRA 839570309 / ABL90 FLEX | Recupération servicedump ( donnée technique analyseur sans identité patient ) Transmission Servicedump maison mère via site techline | Mise à jour information Techline | SWO-532360 | Performance | Performance | Performance | Performance | Crea | customer mail: For the first time since the introduction of ABL90 at the SAU (11/03/2019), we had a case of falsely elevated creatinine for a 47-year-old patient. Here is the chronology of the facts: - yesterday at 9:44 p.m.: dosage of a complete arterial blood gas at the SAU: creatinemia at 114 'mol/L (iono and Hb consistent with lab results) - yesterday 21:30 sample sent to the lab (yellow tube): creatininemia at 71 'mol/L ( corresponds to the anteriority and result this morning) We went to get the syringe from the SAU trash can and redosed it: - on the ABL90 of the SAU: 73 'mol/L - and at home: 72 'mol/L' So there was a falsely high result last night. Have you ever noticed this type of problem? What could be the cause? Note that until then, the results of creatinine obtained on the ABL90 are quite well correlated with those obtained in the laboratory. Thanking you in advance, RFRA 839570309 / ABL90 FLEX | Servicedump recovery (technical data analyzer without patient identity) Transmission Servicedump parent company via techline site | Mise à jour Techline information
No | Visit | Complaint | RCHN | 测量结果不准 | 测量结果不准 | 老师PICO针使用不规范导致，指导老师规范操作pico针 by Andrew | SWO-533361 | [5 empty] | The measurement results are not allowed. | The measurement results are not allowed. | Teacher PICO needle use irregularity led to, guide the teacher to standardize the operation of pico needle by Andrew
No | Visit | Complaint | RIND | Lac Update | Customer Unable to print Lac in report Format | Enable report format layout | SWO-532995 | [5 empty] | Lac Update | Customer Unable to print Lac in report Format | Enable report format layout
No | Visit | Complaint | RCHN | 样本架滑轮脱落 | 样本架滑轮脱落，借903-584一个 | 更换903-584后修复(by chuanwei) | SWO-532436 | Defect | HW | Aspiration Module | Tube holder |  | Sample frame sheave sheds | Sample frame sheave sheds, borrows 903-584 one | Repair after replacement 903-584 (by chuanwei)
Yes | Visit | Complaint | RIND | Calibration error | Check system | Check calibration result, pH ,pCO2 and Cl given problem in calibrating so flush the SC ,done zeroing and again given two point calibration but still problem occur made PER. | SWO-536445 | [5 empty] | Calibration error | Check system | Check calibration result, pH ,pCO2 and Cl given problem in calibrating so flush the SC ,done zeroing and again given two point calibration but still problem occur made PER.
No | Visit | Complaint | RAME | E269 and E580 oxy hardware errors | Confirmed E269 and E580 when onsite. Performed oxy hardware error and found voltage for hemolyzer failed. | Replaced hemolyzer and rerun cal, which passed. | SWO-532965 | Defect | HW | Oxi module | Hemolyzer | E269 E580 | E269 and E580 oxy hardware errors | Confirmed E269 and E580 when onsite. Performed oxy hardware error and found voltage for hemolyzer failed. | Replaced hemolyzer and rerun cal, which passed.
No | Visit | Complaint | RKK | OXIハードウェアトラブル | OXIハードウェアトラブル | 代替モジュール交換 | SWO-532600 | Defect | HW | Oxi module | Oxi module |  | OXI Hardware Trouble | OXI Hardware Trouble | Alternative module replacement
No | Visit | Complaint | RKK | 時刻がずれる | 時刻がずれる 電源を入れなおすと2012年1月1日になる | CPU交換 | SWO-536382 | Defect | HW | DMS_PC | CPU | Time | The time is off. | The time is off. Turn it back on and it will be January 1, 2012 | CPU switch
No | Visit | Complaint | RDEU | Keine Anbindung an Aqure möglich | Kein Remote Support Key installiert | Den Remote Support Key installiert, Ca. Elektrode neu membraniert, Funktionskontrolle ok | SWO-532771 | Setup | SW | HIS/LIS | HIS/LIS | Remote key installed | No connection to Aqure possible | No Remote Support Key installed | Installing the Remote Support Key, Re-membraneing approx. electrode, function control ok
Yes | Visit | Complaint | RBNL | Er is een TnT gemeten van 1600, waarop patient naar ziekenhuis gestuurd. Labwaarde ziekenhuis is <dan. Overmeten op AQT90 geeft nu ook <17 | AQT90 FLEX data collection guide for discrepant results ingevuld en Techline aangemaakt (RBNL........). AQT90 gecontroleerd, er zijn geen afwijkingen waarneembaar. Naald wasstation, cupwash en zwart $wit filter wel schoongemaakt. Diverse checkouts uitgevoerd in het servicemenu, deze zijn allemaal oke. SD analyse uitgevoerd --> oke. 5-Blank test uitgevoerd --> oke QC gemeten --> oke Nieuwe SD analyse uitgevoerd --> oke. |  | SWO-532823 | Re classified RMED | Re-classified RMED | Re-classified RMED | Re-classified RMED | Re-classified RMED | There is a TnT measured from 1600, on which patient sent to hospital. Laboratory Value Hospital is < dan. Overmeasuring on AQT90 now also gives < 17 | AQT90 FLEX Data Collection Guide for behavior results completed and Techline created (RBNL........). AQT90 checked, there are no anomalies discernible. Needle Wash station, Cupwash and black /wit filter cleaned. Several checkouts performed in the service menu, these are all okay. SD Analysis performed--> okay. 5-Blank Test performed--> okay QC measured--> okay New SD Analysis performed--> okay.
No | Visit | Complaint | RFRA | Voir ALBCO pour historique (infos sur pdf ci-joint voir note) L'application 8.11.300 ne se charge pas (roma loader error) Réinstallation et/ou remplacement flashCard et/ou PC UNIT UIM945 En attente d'intervention - erreur apparue lors la modification runpisa.cmd par IA. | Application ne s’ouvre pas , erreur Roma | Sauvegarde ok. Mise à jour en version 8.11 (upgrade) ok. Application s’ouvre correctement ok. Plus d’erreur Roma. Sauvegarde toutes données et config ok. Vérification liveconnect ok. Pas de CQ sur place. Dernier CQ passé le 11 $03 $2019. Test vérification système ok. | SWO-532773 | Defect | SW | RMED SW | RMED SW |  | See ALBCO for history (info on pdf attached see note) The application 8.11.300 does not load (roma loader error) FlashCard and/or PC UNIT UIM945 relocation and/or replacement Waiting for intervention - error that appeared during the change runpisa.cmd by AI. | Application does not open, Roma error | Back up okay. Update da gup 8.11 (upgrade) ok. Application opens properly ok. No more Roma error. Back up all data and config ok. Check liveconnect ok. No CQ on site. Last CQ passed the $11 $03 2019. Test check system ok.
No | Visit | Complaint | RPOL | Aparat żąda wymiany uszczelki. Po wymianie uszczelki oraz podmienieniu SP problem wciąż występuje. | Zainstalowanie używanej kasety sensorowej w celu wykonania testów. Test ciśnienia pozytywny. Wymiana kasety sensorowej na nową - test ciśnienia negatywny. Szczegółowa kontrola wężyków w układzie transportu próbki - stwierdzono naciek na wężyku przy zaworze OXI. | Wymiana wężyków układu transportu próbki. Testy serwisowe - OK. | SWO-532746 | Defect | HW | Oxi module | Pinch valve tube |  | The camera requests replacement of the gasket. The problem still occurs after replacing the gasket with the SP substitution. | Install a used sensory cassette for testing. Positive pressure Test. Replacing the sensor cartridge with a new-negative pressure test. Detailed control of the tubing in the sample transport system-found infiltration on the hose at the OXI valve. | Sample transport system tubing replacement. Service Tests-OK.
No | Visit | Complaint | RPOL | Błąd bazy danych, inlet unieruchomiony w pozycji pomiarowej. | Serwis stwierdza uszkodzoną bazę danych na aparacie oraz inlet w niewłaściwej pozycji. | Oględziny, czyszczenie i zmiana pozycji wlotu, reinstalacja oprogramowania, usunięcie uszkodzonej bazy danych. | SWO-532966 | Defect | SW | RMED SW | database |  | Database error, inlet immobilized in measuring position. | The service finds the corrupted database on the camera and the inlet in the wrong position. | Visual inspection, cleaning and repositioning of the inlet, reinstallation of the software, removal of the corrupted database.
No | Visit | Complaint | RMAY | Oncall breakdown | Solution Pack not properly pressed down until locked during replacement procedure by user. | - Check analyzer data log - Clean inlet - Perform manual flush - Pressed down the solution pack - Properly reinstalled solution pack. | SWO-533113 | User Error | SP | SP | SP activation |  | Oncall breakdown | Solution Pack not properly pressed down until locked during replacement procedure by user. | - Check analyzer data log - Clean inlet - Perform manual flush - Pressed down the solution pack - Properly reinstalled solution pack.
Yes | Visit | Complaint | RIND | SC80 calibration error. | Check the logs. | SC80 failure to calibrate the system. Replace the sc80. | SWO-533249 | [5 empty] | SC80 calibration error. | Check the logs. | SC80 failure to calibrate the system. Replace the sc80.
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | CL.LAC定标不过1321,0377 | 更换测试卡后修复 | SWO-533349 | [5 empty] | Test card, reagent package failure | Cl. LAC calibration is no more than 1321,0377 | Repair after replacing test card
No | Visit | Complaint | RCHN | 检测到泄漏 | 检测到泄漏（未更换备件） | 更换进样口后修复 by jesson | SWO-539942 | Defect | HW | Inlet module | Inlet gasket | 1184 | A leak has been detected | Leak detected (spare parts not replaced) | Fix by jesson after replacing the sample
No | Visit | Complaint | RCHN | 光学强度超范围，报错1031 | 光学强度超范围，报错1031，借903-364一个，944-212（15510）944-258（15385缺货）各一张 | 更换903-364和重新定标后修复(by chuanwei) | SWO-532470 | Defect | HW | Process Module | Optical Unit | 1031 | Optical strength over range, error 1031 | Optical strength over range, error 1031, loan 903-364 one, 944-212 (15510) 944-258 (15385 out of stock) one each | Replacement 903-364 and re-calibration repair (by chuanwei)
No | Visit | Complaint | RIND | SC giving error regarding configuration while installation. | Check system | Check system, check SC configuration need high volume key to install higher configuration SC. | SWO-536450 | [5 empty] | SC giving error regarding configuration while installation. | Check system | Check system, check SC configuration need high volume key to install higher configuration SC.
No | Visit | Complaint | RIND | pH calibration error. | Check the logs. | Replace the ref membrane. calibrate the system. Replace the o-ring & clean the electrode still the sample unstable error. | SWO-533247 | Revisit required | Electrode | Electrode | ph |  | pH calibration error. | Check the logs. | Replace the ref membrane. calibrate the system. Replace the o-ring & clean the electrode still the sample unstable error.
No | Visit | Complaint | RCHN | 机器黑屏，无法操作 | 机器黑屏，无法操作，借910-342一个 | 更换910-342后故障未解决，自行更换备用件910-285后修复 by Andrew | SWO-533088 | Defect | HW | Power Supply | Power Supply |  | Machine black screen, unable to operate | Machine black screen, unable to operate, borrow 910-342 one | Replacement 910-342 after the fault is not resolved, replace the spare parts 910-285 after repair by Andrew
No | Visit | Complaint | RCHN | 最大暗视野检查失败 | 最大暗视野检查失败，借备用903-364一个（R0224N002） |  | SWO-533238 | Defect | HW | Process Module | Optical Unit | 1035 | Maximum dark field check failed | Maximum dark field check failed, with spare 903-364 one (R0224N002)
No | Visit | Complaint | RMAY | Service breakdown | Service breakdown | - Check analyzer data log - Clean chamber - Flush inlet - Perform protein removal - Reconnect AU cable from CPU - Restart analyzer | SWO-533185 | Defect | HW | Basic Unit | cable |  | Service breakdown | Service breakdown | - Check analyzer data log - Clean chamber - Flush inlet - Perform protein removal - Reconnect AU cable from CPU - Restart analyzer
No | Visit | Complaint | RCHN | ph定标失败 | ph定标失败（未更换备件） | 清洁测量仓，前置放大器 by jesson | SWO-539391 | Cleaned | HW | pH/BG Module | measuring chamber | pH, 375 | ph calibration failure | Ph calibration failure (spare parts not replaced) | Clean measuring compartment, preamplifier by jesson
No | Visit | Complaint | RMAY | Service breakdown | Qc error due to drift in calibration high | - Check analyzer data log - Clean chamber - Flush inlet - Perform protein removal - Run calibration to lower the drift. | SWO-533179 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway | pCO2 | Service breakdown | Qc error due to drift in calibration high | - Check analyzer data log - Clean chamber - Flush inlet - Perform protein removal - Run calibration to lower the drift.
No | Visit | Complaint | RDEU | Display Gehäuseschrauben gerissen | Gehäuseschrauben am Diplay gerissen | Gerät und Display kontrolliert, Display erneuert | SWO-534089 | Defect | HW | Display | Display | screws | Display housing screws ripped | Housing screws torn on diplay | Device and display checked, display renewed
No | Visit | Complaint | RAME | Customer getting leak detected on analyzer. | Found that they keep getting leak detected on analyzer. | Back flushed analyzer cleaned inlet gasket, checked inlet probe, check pinch valve, and performed sp checks in service everything is working fine now. | SWO-536705 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway | leak | Customer getting leak detected on analyzer. | Found that they keep getting leak detected on analyzer. | Back flushed analyzer cleaned inlet gasket, checked inlet probe, check pinch valve, and performed sp checks in service everything is working fine now.
No | Visit | Complaint | RAME | Customer called stating analyzer battery not charging | Found that battery is stating 0% charge. | Replaced battery on analyzer | SWO-536704 | Defect | HW | Battery module | Battery module | 0% | Customer called stating analyzer battery not charging | Found that battery is stating 0% charge. | Replaced battery on analyzer
No | Visit | Complaint | RDEU | 12.04.2019 Sr. Janette +492088811295 meldet, dass eine tHb-Kal notwendig ist. | Es sollte ein tHb Kal Fehler vorliegen, bei eintreffen Kal seit 12.04 18.00 wieder i.O ohne das ThB Kal durchgeführt wurde, | Bei Eintreffen LP Haltbarkeit überschritten seit dem Vortag, LP ersetzt Startphase 2h, Speicher überprüft es lagen keine Kal und QC Fehler am Vortag mehr an, Anwender meldet sich wenn es nach der Startphase wieder zu Problemen kommen sollte | SWO-533508 | No fault found | No fault found | No fault found | No fault found | tHb Cal | 12.04.2019 Sr. Janette +492088811295 reports that a tHb-Kal is necessary. | There should be a tHb Kal error, if Kal arrives since 12.04 18.00 again i.O without the ThB Kal was carried out, | On arrival LP shelf life exceeded since the previous day, LP replaces start phase 2h, memory checks there were no kal and QC errors the day before, user logs in if problems should occur again after the start phase
No | Visit | Complaint | RAME | SC door won't close. | On site to inspect possible faulty SC's. | Two sensor cassette's lot 1504 will install but require a firm 'push' on the right side. There is plastic that appears leftover from the molding process sticking out. Other lot numbers at customer site do not have this issue. | SWO-534059 | Assisted Replacement | SC | SC | SC | sent upSWO to CSCHO | SC door won't close. | On site to inspect possible faulty SC's. | Two sensor cassette's lot 1504 will install but require a firm 'push' on the right side. There is plastic that appears leftover from the molding process sticking out. Other lot numbers at customer site do not have this issue.
No | Visit | Complaint | RCHN | 测试卡连续上机验证不通过 | 测试卡连续上机验证不通过，借903-334和902-972各一个 | 更换测试卡后零点校准后修复 by Andrew | SWO-533079 | [5 empty] | Test card continuous on-board verification does not pass | Test card continuous on-board verification does not pass, borrow 903-334 and 902-972 each one | Fix by Andrew after zero calibration after replacing test card
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | LAC定标失败 | 更换测试卡后修复 | SWO-534760 | [5 empty] | Test card, reagent package failure | LAC calibration failure | Repair after replacing test card
No | Visit | Complaint | RCHN | 数据无法下载 | 数据无法下载，借903-381（被903-720替换），933-494各一个 |  | SWO-533353 | [5 empty] | Data cannot be downloaded | Data cannot be downloaded, borrowed 903-381 (replaced by 903-720), 933-494 each
No | Visit | Complaint | RIND | SC Replacement as done | Checked the machine, Failure in SC Temp-A, Temp-B, Temp-F Occurs when SC installed. | Need to replace the Solution Cassette for good working. | SWO-536303 | [5 empty] | SC Replacement as done | Checked the machine, Failure in SC Temp-A, Temp-B, Temp-F Occurs when SC installed. | Need to replace the Solution Cassette for good working.
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | K灵敏度偏低 | 更换测试卡正常。 | SWO-534115 | [5 empty] | Test card, reagent package failure | Low Sensitivity of K | The replacement test card is OK.
No | Visit | Complaint | RCHN | 样本抽吸错误 | 样本抽吸错误，加借903-139一个 | 更换903-129后修复 by Andrew | SWO-533081 | Defect | HW | Oxi module | pinch valve |  | Sample suction error | Sample suction error, plus 903-139 one | Repair by Andrew after replacing 903-129
No | Visit | Complaint | RCHN | 黑屏 | 黑屏，借910-342和910-285各一个 | 更换910-285后修复 by Andrew | SWO-533095 | Defect | HW | Power Supply | Power Supply |  | Black screen | Black screen, borrow 910-342 and 910-285 each | Repair by Andrew after replacement 910-285
No | Visit | Complaint | RCHN | 液路有凝块 | 液路有凝块（未更换备件） | 手动冲洗液路，重新定标后修复 by Andrew | SWO-533499 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway | Clot | There's a clot in the liquid road. | There are clots in the liquid road (no replacement parts) | Manual flushing of the liquid circuit, re-calibration and repair ingress by Andrew
No | Visit | Complaint | RCHN | 定标失败 | 定标失败，借944-268（15733）一个 |  | SWO-533357 | Assisted Replacement | Consumables | Test Kit | Cartridge, cal |  | Bid failed | Failed to bid, borrow 944-268 (15733) one
No | Visit | Complaint | RCHN | 报错4009 | 报错4009，需要借件636-538一个 |  | SWO-533316 | Defect | HW | Inlet module | Sensor, inlet lock | 4009 | 4009 error | 4009, need to borrow 636-538 one
No | Visit | Complaint | RCHN | 液路故障，报错1273 | 液路故障，报错1273。需要借件844-035一个 |  | SWO-533381 | Defect | HW | Wet section | Motor, pump | 1273 | Liquid road fault, error 1273 | Liquid road fault, error 1273. Need to borrow 844-035 one
No | Visit | Complaint | RCHN | 定标失败 | 定标失败（未更换备件） | 清洗PH模块上方冲洗阀后修复 by jesson | SWO-539421 | Cleaned | HW | pH/BG Module | pH/BG Module | Valve | Bid failed | Failed to bid (spare parts not replaced) | Repair by jesson after cleaning the flushing valve above the PH module
No | Visit | Complaint | RESP | Nos QC dão amostras não homogênea. | Verificação de uma bolha de ar no início da amostra | Substituição do material indicado e verificação do funcionamento com QC e várias amostras. | SWO-533777 | Defect | HW | Inlet module | Tubing | Preheater | In QC they give unhomogeneous samples. | Checking an air bubble at the beginning of the sample | Replacement of the indicated material and verification of operation with QC and various samples.
No | Visit | Complaint | RAME | Solution pump error | rinse not detected in Ph BG module | replaced both valve in the ph module , pCO2 not passing QC remembraned the electrode | SWO-533943 | Defect | HW | pH/BG Module | Valves | Pump error | Solution pump error | rinse not detected in Ph BG module | replaced both valve in the ph module , pCO2 not passing QC remembraned the electrode
No | 5 | Y | RCZE | Blank | 500 - pH $BG modul detekoval preplachovaci roztok ked nebol ocakavany 963 - zjisten unikajici proud |  | D5F9FD1620522FE4 | Defect | HW | pH/BG Module | Valve, rinse | 963 | Blank | 500-pH /BG module detected a flushing solution when it was not ocakavany 963-zjisten Unikajici proud
No | Visit | Complaint | RIND | Calibration failure | Calibration failure for calcium | Checked the calibration log , activity log , calcium sensitivity and drift out of range , cleaned the instrument , changed the calcium membrane , repeated the calibration ,calibration passed , done samples, customer happy with results | SWO-534113 | Assisted Replacement | Electrode | Membrane | Ca | Sens and drift | Calibration failure | Calibration failure for calcium | Checked the calibration log , activity log , calcium sensitivity and drift out of range , cleaned the instrument , changed the calcium membrane , repeated the calibration ,calibration passed , done samples, customer happy with results
No | 8 | Y | RCZE | Blank |  |  | D5FB2CC3960BFEEC | Defect | HW | Cover and Chassis | PCB, solution pack |  | Blank
No | Visit | Complaint | RLTD | Mandy Keyworth Tel: 01970 635712 | The analysers had been failing the pressure test after the problem with the inlet had been resolved but on arrival the analyser was requesting a new solution pack as the one installed had run out of activities. | I fitted a new solution pack and the analyser went through its start-up ok. I checked the operation of the inlet and it was working as it should do, it looks like the reason the analyser was failing the pressure test was due to the old solution pack | SWO-534569 | Assisted Replacement | SP | SP | Pressure Test |  | Mandy Keyworth Tel: 01970 635712 | The analysers had been failing the pressure test after the problem with the inlet had been resolved but on arrival the analyser was requesting a new solution pack as the one installed had run out of activities. | I fitted a new solution pack and the analyser went through its start-up ok. I checked the operation of the inlet and it was working as it should do, it looks like the reason the analyser was failing the pressure test was due to the old solution pack
No | Visit | Complaint | RPAC | Allison 0413835307 | Unit has a leakage current error 964 which has persisted over two solution packs with different LOT numbers. Also has a couple of inlet errors over the same period of time. Checked solution pack and it was engaged properly. Performed fluid movement activities. Performed complete OXI test and pO2 test. All okay. The error code was not present and the unit started up as per normal. The error codes were found in the activity log. Checked inlet probe and it looks very slightly bent. Shutdown unit and removed consumables. Disassembled wet section and found that the waste tubing was not put on properly on the hemolyser end. There was however no obvious signs of a leak having occured. | Replaced the waste tubing and reassembled wet section. Replaced inlet probe and gasket. Rechecked fluid movements and allowed analyser to start back up. Analyser passes all of its calibrations and QCs. The tHb calibration was due so I did that as well. All is well with the machine. | SWO-534093 | Defect | HW | Basic Unit | PM kit | 964 | Allison 0413835307 | Unit has a leakage current error 964 which has persisted over two solution packs with different LOT numbers. Also has a couple of inlet errors over the same period of time. Checked solution pack and it was engaged properly. Performed fluid movement activities. Performed complete OXI test and pO2 test. All okay. The error code was not present and the unit started up as per normal. The error codes were found in the activity log. Checked inlet probe and it looks very slightly bent. Shutdown unit and removed consumables. Disassembled wet section and found that the waste tubing was not put on properly on the hemolyser end. There was however no obvious signs of a leak having occured. | Replaced the waste tubing and reassembled wet section. Replaced inlet probe and gasket. Rechecked fluid movements and allowed analyser to start back up. Analyser passes all of its calibrations and QCs. The tHb calibration was due so I did that as well. All is well with the machine.
No | Visit | Complaint | RKOR | Qc data trouble | Inlet probe clean, check fluid line flush Gasket holder check Gasket module check calibration,control data check overflow check | Inlet probe clean, check fluid line flush Gasket holder check Gasket module check calibration,control data check overflow check | SWO-536134 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | Qc data trouble | Inlet probe clean, check fluid line flush Gasket holder check Gasket module check calibration,control data check overflow check | Inlet probe clean, check fluid line flush Gasket holder check Gasket module check calibration,control data check overflow check
No | Visit | Complaint | RAME | Screen chassis repair | Screen chassis broken | Replaced screen module | SWO-534695 | Defect | HW | Display | Display | Chassis broken | Screen chassis repair | Screen chassis broken | Replaced screen module
No | Visit | Complaint | RIND | Na is high in sample | Check Patient result Data log. Result were high. | Flus the cassette ,re-install the cassette. Calibration pass. Run a sample in pico Result is satisfactory. | SWO-536521 | [5 empty] | Na is high in sample | Check Patient result Data log. Result were high. | Flus the cassette ,re-install the cassette. Calibration pass. Run a sample in pico Result is satisfactory.
No | Visit | Complaint | RKOR | Tem.A Error | . | SC replace. Warranty DM training. | SWO-536136 | [5 empty] | Tem.A Error | . | SC replace. Warranty DM training.
No | Visit | Complaint | RIND | Analysis disabled | Checked the ABL80 and found insufficient solution pack. | Need to replace the solution pack but SP is not available on customer site. Informed to CP for arrange to supply the SP. | SWO-534465 | [5 empty] | Analysis disabled | Checked the ABL80 and found insufficient solution pack. | Need to replace the solution pack but SP is not available on customer site. Informed to CP for arrange to supply the SP.
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 参比电极错误 1361 | 更换测试卡后修复 | SWO-536108 | [5 empty] | Test card, reagent package failure | Reference electrode error 1361 | Repair after replacing test card
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | PO2质控不过 | 更换试剂包后修复 | SWO-536106 | [5 empty] | Test card, reagent package failure | PO2 QC but not | Repair after replacement of reagent kit
No | Visit | Complaint | RCHN | HCT测量不出值 | HCT测量不出值 | 清洁液路后修复 by Andrew | SWO-535572 | Cleaned | HW | Aspiration Module | Needle Wash Unit |  | HCT cannot measure value | HCT cannot measure value | Repair by Andrew after cleaning liquid road
No | Visit | Complaint | RCHN | 重启分析仪 | 重启分析仪 | 更换试剂包后修复 | SWO-534058 | Defect | Consumables | Solution Pack | Chip | 1999.35.16 | Restart the analyzer | Restart the analyzer | Repair after replacement of reagent kit
No | Visit | Complaint | RKK | PO2ノーエンドポイント頻発、PH,PCO2,Hｃｔ、K,Ca,Clノーエンドポイント頻発 | PO2ノーエンドポイント頻発、PH,PCO2,Hｃｔ、K,Ca,Clノーエンドポイント頻発 | センサーコネクターボード及びメンテナンスキット交換。 | SWO-536193 | [5 empty] | PO2 no-endpoint frequent, PH, PCO2, Hct, K, Ca, Cl no endpoint frequent | PO2 no-endpoint frequent, PH, PCO2, Hct, K, Ca, Cl no endpoint frequent | Sensor connector board and maintenance kit replacement.
No | 4 | Y | RCZE | Blank | V case 3,50 zastavena kalibrace pro neidentifikaci proplachu |  | D5FDA34E82610CFD | Defect | HW | pH/BG Module | measuring chamber | Check history | Blank | In case 3.50 The calibration is stopped for non-identification of rinse
No | Visit | Complaint | RAME | Autocheck issues | Auto-check issues. Not scanning ampoules. Return to try more cleaning and adjustment. | Ampoules are being scanned ok now. | SWO-534456 | Cleaned | HW | Autocheck module | Mirror |  | Autocheck issues | Auto-check issues. Not scanning ampoules. Return to try more cleaning and adjustment. | Ampoules are being scanned ok now.
No | Visit | Complaint | RAME | Autocheck issues | Auto-check not recognizing ampoules. Doesn’t see ampoules in service. Replaced ribbon cable on carousel. Still same. Tried alignment in service. Still bad. Cleaned mirrors. Laser engine mirror is not very shiny and has black spots on. It. Now see some ampoules once and a while. | Will return tomorrow morning to try one more thing. | SWO-533978 | Revisit required | HW | Autocheck module | Mirror |  | Autocheck issues | Auto-check not recognizing ampoules. Doesn’t see ampoules in service. Replaced ribbon cable on carousel. Still same. Tried alignment in service. Still bad. Cleaned mirrors. Laser engine mirror is not very shiny and has black spots on. It. Now see some ampoules once and a while. | Will return tomorrow morning to try one more thing.
No | Visit | Complaint | RAME | Oxi hardware error | Sample errors. Bent probe. | Followed up with customer about oximetry error from day before and they weren’t having oxi issues but were having no sample detected issues instead. Found bent probe and replaced with customer stock. | SWO-534610 | Defect | HW | Inlet module | Probe bend |  | Oxi hardware error | Sample errors. Bent probe. | Followed up with customer about oximetry error from day before and they weren’t having oxi issues but were having no sample detected issues instead. Found bent probe and replaced with customer stock.
No | Visit | Complaint | RAME | ID=95 Phone #: Ext Tel: Contact Name: Caller Nm: Yasmin Elgamal Phone #: 202-715-4252 Facility: George Washington Univ Hospital Address: 901 23rd St Nw City,State,Zip: Washington, DC, 20037 Instrument: ABL825 FLEX Serial Number: 754R1291N001 Message: THEY HAVE BRAND NEW ABLS AND THEY KEEP GETTING ERRORS SAYING QCS ARE OUT. | Lac low on level 3 QC | I checked the valves in the mods. I cleaned and remembered the lac with no change. I switched the electrodes between analyzers and the issue moved. I have a new electrode sent to them | SWO-534648 | Assisted Replacement | Electrode | Electrode | Lac | QC L3 low | ID=95 Phone #: Ext Tel: Contact Name: Caller Nm: Yasmin Elgamal Phone #: 202-715-4252 Facility: George Washington Univ Hospital Address: 901 23rd St Nw City,State,Zip: Washington, DC, 20037 Instrument: ABL825 FLEX Serial Number: 754R1291N001 Message: THEY HAVE BRAND NEW ABLS AND THEY KEEP GETTING ERRORS SAYING QCS ARE OUT. | Lac low on level 3 QC | I checked the valves in the mods. I cleaned and remembered the lac with no change. I switched the electrodes between analyzers and the issue moved. I have a new electrode sent to them
No | Visit | Complaint | RIND | SETSAFE 22 | Storage wheel not functionig | Clean the sensor . Clean the storage wheel. Power on the analyzer. | SWO-536518 | Cleaned | HW | Storage Module | Sensor, storage wheel | setsafe 22 | SETSAFE 22 | Storage wheel not functionig | Clean the sensor . Clean the storage wheel. Power on the analyzer.
No | 8 | Y | RCZE | Blank |  |  | D5FCC9650C25FEF9 | Defect | HW | Process Module | Optical Unit | QC | Blank
Yes | Visit | Complaint | RKK | 機器点検 | 機器点検 | 機器点検の為、機器引き上げ。代替え機セットアップ。 | SWO-535592 | Re classified RMED | Maintenance |  |  | See mail with title: "RE: SWO-535592" from SHIYA | Equipment inspection | Equipment inspection | For equipment inspection, equipment pullup. Alternative machine setup.
No | Visit | Complaint | RMAY | Oncall breakdown | Solution pack not properly installed | -Check analyser data log -Perform manual flush -Reinstall solution pack -Run fluid transport test | SWO-535433 | User Error | SP | SP | SP installation | 1543 | Oncall breakdown | Solution pack not properly installed | -Check analyser data log -Perform manual flush -Reinstall solution pack -Run fluid transport test
No | Visit | Complaint | RLTD | Girlie 01914041000 ext 2350 | Analyser freezing whilst entering patient demographics and needing to be rebooted | Analyser software V3.4 reinstalled | SWO-535085 | Defect | SW | RMED SW | Freeze |  | Girlie 01914041000 ext 2350 | Analyser freezing whilst entering patient demographics and needing to be rebooted | Analyser software V3.4 reinstalled
No | Visit | Complaint | RLTD | Girlie 01914041000 ext 2350 | Analyser freezing whilst entering patient demographics and needing to be rebooted | Reinstalled the analyser software V3.4 | SWO-535086 | Defect | SW | RMED SW | Freeze |  | Girlie 01914041000 ext 2350 | Analyser freezing whilst entering patient demographics and needing to be rebooted | Reinstalled the analyser software V3.4
Yes | Visit | Complaint | RSAF | Leakage | Inspected the machine for any visible leaks | Re-activated the solution pack | SWO-534951 | [5 empty] | Leakage | Inspected the machine for any visible leaks | Re-activated the solution pack
No | Visit | Complaint | RAME | Multiple Patient samples getting errors 0094, 0292- turbidity, 0581, 0688 | Inspected instrument and reviewed logs. Upon arrival Na is failing calibrations. Recalibration led to intermittent failure of Na and Lac. Replaced sensor cassette after repair. Patient samples have error codes 292 turbidity too high, 291 SHb too high, 581 OXI spectrum mismatch on all oximetry but QC and tHb calibrations pass. | Replaced hemolyzer, tubings and gaskets. Old Hemolyzer s $n 908-552R2249N0035. Printed credit forms for sensor cassettes. | SWO-535380 | Defect | HW | Oxi module | Hemolyzer | 581, 292, 688 | Multiple Patient samples getting errors 0094, 0292- turbidity, 0581, 0688 | Inspected instrument and reviewed logs. Upon arrival Na is failing calibrations. Recalibration led to intermittent failure of Na and Lac. Replaced sensor cassette after repair. Patient samples have error codes 292 turbidity too high, 291 SHb too high, 581 OXI spectrum mismatch on all oximetry but QC and tHb calibrations pass. | Replaced hemolyzer, tubings and gaskets. Old Hemolyzer s $n 908-552R2249N0035. Printed credit forms for sensor cassettes.
No | Visit | Complaint | RAME | The bezel on the back of the display is loose. | Upon arriving I also found that the glass was loose as well. | The part I received was not correct. I removed the customer’s display and adjusted and tightened the glass. I secured the bezel to the back of the display. I am ordering another display to be replaced during a future visit. | SWO-535240 | Defect | HW | Display | Display |  | The bezel on the back of the display is loose. | Upon arriving I also found that the glass was loose as well. | The part I received was not correct. I removed the customer’s display and adjusted and tightened the glass. I secured the bezel to the back of the display. I am ordering another display to be replaced during a future visit.
No | Visit | Complaint | RAME | Set up issues/set up is not the same as the other | Instrument not talking to Ralls. | Looks like the Astm.ini file might be bad. Will have another one sent to me. | SWO-535367 | Revisit required | SW | RMED SW | HIS/LIS |  | Set up issues/set up is not the same as the other | Instrument not talking to Ralls. | Looks like the Astm.ini file might be bad. Will have another one sent to me.
No | Visit | Complaint | RDEU | 11.04.2019; Hr. Reich; Tel.: 015254785507; Uhrzeit stimmt nicht | Überprüfung Uhrzeiteinstellung (WIN) Analysator steht im ´Halt´-Intervention erforderlich | Servcicetests im Menü durchgeführt. Vacuum Value zu niedrig. Drucktest nicht bestanden. Schlauch des Quetschventils erneuert. SC Defekt. Neue SC eingesetzt, Vakuumtest i.O. Defekte SC reklamiert Service Kalibrationen durchgeführt. Initialsierung Deaktivieren des Schreibschutzes Einstellungen der Uhrzeit auf Windowsebene korrigiert Aktivieren des Schreibschutzes Initialsierungsphase Kal und QC wird von OP Personal beobachtet | SWO-535203 | Assisted Replacement | SC | SC | Pressure Test |  | 11.04.2019; Mr. Reich; Tel.: 015254785507; Time is not correct | Check Time setting (WIN) Analyzer is required in the 'Halt' intervention | Servcice tests in the menu. Vacuum value too low. Pressure test failed. hose of the crushing valve. SC Defect. New SC used, vacuum test i.O. Defective SC complained Service calibrations. Initialization Disabling write protection Fixed time settings at the Windows level Activating write protection Initialization phase Kal and QC are observed by SURGICAL staff
No | Visit | Complaint | RDEU | 11.04.2019; Hr. Reich; Tel.: 015254785507; NetOp funktioniert nicht | Überprüfung NetOp | Remote ohne Funktion Überprüfung der Konfiguration des Fernzugriff; Kein Status der Software. Neustart und stoppen des NetOp. Erneuter Neustart des NetOP - Status ´running´ Fernzugriff laut Herrn Dr. Reich möglich. LP wurde aufgrund 50 verbleibende Aktivitäten und pCO2 QC Fehler C vorzeitig ausgetauscht. Initialisierung Kalibrationen durchgeführt. QC im Level A nun außer Bereich. SC aus Kühlschrank entnommen, Stationsleitung wird SC nach Aufwärmphase austauschen QC im pCO2 außer Bereich | SWO-535200 | Defect | SW | RMED SW | HIS/LIS |  | 11.04.2019; Mr. Reich; Tel.: 015254785507; NetOp is not working | Review NetOp | Remote without function Checking the configuration of remote access; No status of the software. Restart and stop the NetOp. Restarting netOP status of "running" Remote access according to Dr. Reich. LP was replaced prematurely due to 50 remaining activities and pCO2 QC error C. Initialization calibrations. QC in Level A now out of range. SC removed from refrigerator, station line will replace SC after warm-up phase QC in pCO2 out of range
No | Visit | Complaint | RDEU | 16.4.2019/12:10 Sr. Marlies 06051872888 - korrekte Uhrzeit vom Gerät stimmt nicht mit Messausdruck überein | Überprüfung der Einstellung | Uhrzeit ohne Beanstandung (Zeitserver) Uhrzeit auf Ausdruck eine -1Stunde. Analysator neu gestartet. Änderung der Zeitzone wurde geändert ohne Speichersperre zu deaktivieren. Speichersperre aufgehoben, Einstellungen auf Windowsebene aktualisiert. Speichersperre wieder aktiviert. Neustart Zeiten stimmen. QC im Bereich des pCO2 Level A Fehlerhaft Stationsleitung tauscht SC nach Aufwärmphase | SWO-535201 | Setup | SW | RMED SW | Time setting |  | 16.4.2019/12:10 Sr. Marlies 06051872888 - correct time from the device does not match measurement expression | Checking the setting | Time without complaint (time server) Time on printout one -1 hour. Analyzer restarted. Change of time zone has been changed without disabling memory lock. Memory lock, Windows-level settings updated. Memory lock is reactivated. Restart Times are right. QC in the area of pCO2 Level A Faulty Station line replaces SC after warm-up phase
No | Visit | Complaint | RAME | Calibrate touchscreen | Customer reported that the touchscreen was not reacting to touch. Found the screen to be slightly off when touching | Cleaned screen and ran screen calibration program. Found screen working much better after calibration. Ran system cycles and manual QC found all levels passed | SWO-535392 | [5 empty] | Calibrate touchscreen | Customer reported that the touchscreen was not reacting to touch. Found the screen to be slightly off when touching | Cleaned screen and ran screen calibration program. Found screen working much better after calibration. Ran system cycles and manual QC found all levels passed
Yes | Visit | Complaint | RIND | Calibration error | Check System | Check system,reset SP again Install SC and SP , Flush SC still the problem occurs in calibration suggested to replace SC. | SWO-536459 | [5 empty] | Calibration error | Check System | Check system,reset SP again Install SC and SP , Flush SC still the problem occurs in calibration suggested to replace SC.
No | Visit | Complaint | RIND | Calibration error | Check System | When installing new cassata again calibration error due to air message showing so again cleaned the tubings and tried but same problem occurs replace SP with new one still same problem happen SP having same lot no, but when these two consumables run on standby analyser working good, so waiting for new klot of SP. | SWO-536460 | [5 empty] | Calibration error | Check System | When installing new cassata again calibration error due to air message showing so again cleaned the tubings and tried but same problem occurs replace SP with new one still same problem happen SP having same lot no, but when these two consumables run on standby analyser working good, so waiting for new klot of SP.
No | Visit | Complaint | RCHN | 系统温度超范围，报错1221 | 系统温度超范围，报错1221，借903-581一个 | 更换903-581后修复 by Andrew | SWO-534891 | Defect | HW | Basic Unit | PCB, Sensor interface | 1221, temp | System temperature over range, error 1221 | System temperature over range, error 1221, borrow 903-581 one | Repair by Andrew after replacement 903-581
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 压力测试真空错误 1063 | 更换测试卡后修复 | SWO-534899 | Assisted Replacement | SC | SC | Pressure Test | 1063 | Test card, reagent package failure | Pressure test vacuum error 1063 | Repair after replacing test card
No | Visit | Complaint | RCHN | 测试卡断裂，报错5784 teeth | 测试卡断裂，报错5784 teeth | 更换测试卡后修复 by Andrew | SWO-534863 | Defect | Consumables | Test Kit | Cartridge, Broken | 5784, Teeth | Test card break, error 5784 teeth | Test card break, error 5784 teeth | Fix by Andrew after replacing test card
No | Visit | Complaint | RDEU | PCO2 Empfindlichkeitsfehler | PCO2 Empfindlichkeitsfehler | aufgef. PCO2-Elektrode ersetzt, Kal, QC Datensicherung Kal QC | SWO-535358 | Assisted Replacement | Electrode | Electrode | pCO2 | Sens | PCO2 Empfindlichkeitsfehler | PCO2 Empfindlichkeitsfehler | on the unraised. PCO2 electrode replaced, Kal, QC Data backup Kal QC
No | Visit | Complaint | RCHN | 气阀膜损坏 482 | 气阀膜损坏 482,更换834-222一个正常。 | 更换834-222一个正常。 by Andrew | SWO-534947 | Defect | HW | pH/BG Module | Valve, membrane | 482 | Damage to valve membrane 482 | The valve membrane is damaged 482, replacing 834-222 a normal. | Replace 834-222 one normal. by Andrew
No | Visit | Complaint | RKK | Na値が低い | Na値が低い（測定、QC） | Cal1溶液の汚染→Cal1.2溶液交換、比較メンブラン交換、QC測定OK | SWO-535379 | Assisted Replacement | Consumables | Solution | Cal1 | Na | Low Na value | Low Na value (measurement, QC) | Cal1 solution contamination → Cal1.2 solution exchange, comparison membrane replacement, QC measurement OK
No | Visit | Complaint | RIND | Sodium Value low | Checked the QC Value Sodium value is low in its calculated ranges, and calibration not ok Sodium value is in Drift. | Cleaned the measuring chambers and a clot in sodium chamber removed and given 2 point calibration. QC as done, Now working good | SWO-536320 | Cleaned | HW | El/Met Module | measuring chamber | Clot, Na | Sodium Value low | Checked the QC Value Sodium value is low in its calculated ranges, and calibration not ok Sodium value is in Drift. | Cleaned the measuring chambers and a clot in sodium chamber removed and given 2 point calibration. QC as done, Now working good
No | Visit | Complaint | RDEU | 16.04.2019 13:45Uhr Her Glienke tel:+493334691503 : Kal-Fehler bei Elektrolyten | häufige Driften bei Na, Ca und Cl | Ventile der Meßmodule ausgebaut und kontrolliert, da Luftblase mit Bläschen versehen war - i.O. blauer Luftauslaß Wetsection war zugesetzt da weiterhin Driften auftraten, Referenz Elektrode gewechselt Meßkammer kontrolliert - Dichtring an pCO2 war verdreht eingesetzt $zusätzlich membraniert mit anderer Charge | SWO-535102 | Assisted Replacement | Electrode | Electrode | reference | CAPA-733 | 16.04.2019 13:45H Her Glienke tel:+493334691503 : Kal fault in electrolytes | frequent drifts at Na, Ca and Cl | Valves of the measuring modules removed and controlled, since air bubble was provided with bubbles - i.O. blue air outlet Wetsection was added as drifts continued to occur, reference electrode changed Measuring chamber controlled - sealing ring on pCO2 was used twisted $zusreputably membraneed with other batch
No | Visit | Complaint | RDEU | 16.04.2019; Hr. Börste; Tel.: 0401818862280; kap. Messungen nicht möglich | Keine Kapilare Messung möglich. Flüssigkeit läuft aus der Einlassdichtung. | Analysator von Aussen gereinigt. Einlassdichtung getauscht. Kal und QC durchgeführt. Danach Probemessung mit Kapillare durchgeführt. Sytem wieder einsatzbereit. | SWO-534902 | Defect | HW | Inlet module | Inlet gasket |  | 16.04.2019; Mr. Börste; Tel.: 0401818862280; Chapter. Measurements not possible | No kapilare measurement possible. Liquid runs out of the inlet seal. | Analyzer cleaned from the outside. inlet seal is exchanged. Kal and QC. Then sample measurement performed with capillary. system is ready for use again.
No | Visit | Complaint | RDEU | 17.04.2019 09:00Uhr Frau Schulz +4923096377254 : Auf der rechten Seite unter der Flasche Reinigungslösung tropft Flüssigkeit aus. | Flüssigkeitsaustritt unterhalb des ABL im Bereich der Reinigungsflasche | Fließwege und Ventile im Fluidiq Modul überprüft, Luftkanal und Abfallwege, gespült alles trocken und Funktionsfähig, Reinigungsflasche Lot BR04 immer wieder feucht , Reinigungslösung ersetzt, keine Flüssigkeit mehr feststellbar, QC durchgeführt, Anwender beobachtet und meldet sich bei weiteren Problemen | SWO-535480 | Defect | Consumables | Solution | Cleaning | Lot BR04 | 17.04.2019 09:00Am Mrs Schulz +4923096377254 : Liquid drips out on the right under the bottle of cleaning solution. | Liquid outlet below the ABL in the area of the cleaning bottle | Flow paths and valves checked in fluidiq module, air duct and waste paths, flushed everything dry and functional, cleaning bottle lot BR04 again and again moist, cleaning solution replaced, no liquid detectable, QC performed, user observed and reports further problems
Yes | 8 | Y | RCZE | Blank |  |  | D5FCCAFF0C24FEFC | Re classified RMED | Return of analyzer |  |  |  | Blank
No | Visit | Complaint | RPOL | Zgięta igła | Wizualna ocena | Wymiana igły, testy serwisowe. | SWO-538720 | Defect | HW | Auto Inlet module | Probe bend |  | Bent Needle | Visual rating | Replacing the needle, service tests.
No | Visit | Complaint | RPOL | Wygięta igła | Serwis stwierdza wygiętą igłę pobierającą próbkę. | Wymiana igły, testowanie aparatu. | SWO-535236 | Defect | HW | Aspiration Module | Needle Assembly | Bent | Bent Needle | The service finds a curved needle that takes the sample. | Replacing the needle, testing the camera.
No | Visit | Complaint | RAME | Battery shows no charge | None | Replaced battery assembly | SWO-535448 | Defect | HW | Battery module | Battery module |  | Battery shows no charge | None | Replaced battery assembly
No | Visit | Complaint | RDEU | 16.4.2019/17:00 Frau Worchert (02191138483) - Intervention erforderlich 17.04.2019 18:00Uhr Frau Worchert 02191138483) | ABL steht in Halt, Intervention erforderlich - Leckage erkannt. | Einlassdichtung blutig und das Gummi ist aus der Einlassdichtung ist herausgedrückt. Einlassdichtung gereinigt und korrekt befestigt, Einlassnadel durchgespült, Transportwege im Service überpürft und Kalibriert, 2 std. Startphase durchgeführt - Kal ok, QC lvl C pO2 zu hoch, erneut Kalibriert und die QC wiederholt - weiterhin zu hoch, LP erneuert, Startphase durchgeführt - Kal und QC ok. | SWO-536329 | Defect | HW | Inlet module | Inlet gasket | Leak | 16.4.2019/17:00 Ms Worchert (02191138483) - Intervention required 17.04.2019 18:00pm Ms Worchert 02191138483) | ABL is in hold, intervention required - leakage detected. | Inlet seal bloody and the rubber is squeezed out of the inlet seal. Inlet seal cleaned and correctly fixed, inlet needle flushed through, transport routes in service overrun and calibrated, 2 hours. Start phase performed - Kal ok, QC lvl C pO2 too high, recalibrated and the QC repeated - continues too high, LP renewed, start phase performed - Kal and QC ok.
No | Visit | Complaint | RDEU | 08.04.2019; Hr. Weiser; Tel.: 0765129413; ABL verlängert die Messzeit 11.04.2019, Frau Guth, Tel: 0765129413: ständige Kalibrierungen ca im 30 Min. Takt 12.04.2019 08:30Uhr Frau Haas tel:+4976512978412: Das Gerät geht ständig in die Startphase | Gerät bricht Kal. ab, geht selbständig in Reboot $Initialisierungsphase | GEWÄHRLEISTUNG! CPU und Netzteil ersetzt, CF Karte übernommen, Konfiguration und Vorgaben überprüft, Vorgaben und Anbindung i.O., SC, LP und Einlasssondenhalter aus Kundenbestand erneuert, 10 min vor Ablauf der Initialisierung geht das Gerät wieder in Reboot | SWO-535906 | Revisit required | HW | DMS_PC | CPU | Reboot | 08.04.2019; Mr. Weiser; Tel.: 0765129413; ABL extends the measuring time 11.04.2019, Mrs. Guth, Tel: 0765129413: continuous calibrations approx. 30 min. 12.04.2019 08:30Pm Ms Haas tel:+4976512978412: The device is constantly going into the start-up phase | Device breaks Kal. goes off, goes into reboot /Initialisierungsphase on its own | Ensure! Replaced CPU and power supply, CF card taken over, configuration and specifications checked, specifications and connection i.O., SC, LP and inlet probe holder renewed from customer stock, 10 min before the start of the initialization the device goes back into reboot
No | Visit | Complaint | RKOR | Inhomogeneous Rinse solution | Oxi module check, Inlet check, module check repaet cal control data check | Oxi module check, Inlet check, module check repaet cal control data check | SWO-536158 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway | Clot | Inhomogeneous Rinse solution | Oxi module check, Inlet check, module check repaet cal control data check | Oxi module check, Inlet check, module check repaet cal control data check
No | Visit | Complaint | RLTD | Magda Tel. 0208 8692133 | Vacuum error pump cal error Liquid transport failed. | Flushed system and found restriction in water flow. Traced this to inlet module gasket being blocked. Removed gasket and flushed and removed clot. Refitted gasket and run all service test all passed ok. Calibration and QC run and passed ok. | SWO-535893 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway | Clot | Magda Tel. 0208 8692133 | Vacuum error pump cal error Liquid transport failed. | Flushed system and found restriction in water flow. Traced this to inlet module gasket being blocked. Removed gasket and flushed and removed clot. Refitted gasket and run all service test all passed ok. Calibration and QC run and passed ok.
No | Visit | Complaint | RLTD | Richard Barton on 0141 9515917 | Analyser falling calibration and tHb calibrations Clot suspected to have caused original issues | Flushed analyser and carried out all fluid transport calibrations in service program Restarted analyser and all ok | SWO-535774 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway | Clot | Richard Barton on 0141 9515917 | Analyser falling calibration and tHb calibrations Clot suspected to have caused original issues | Flushed analyser and carried out all fluid transport calibrations in service program Restarted analyser and all ok
No | Visit | Complaint | RAME | Sensor cassette bay will not close. | Found that the bay unit would get half way close then open again. | Found that it was catching on SC adjusted the bay unit so it could close. | SWO-536706 | Adjustment | HW | Basic Unit | Sensor Bay Unit |  | Sensor cassette bay will not close. | Found that the bay unit would get half way close then open again. | Found that it was catching on SC adjusted the bay unit so it could close.
No | Visit | Complaint | RLTD | Customer removed analyser from AAU as it was asking for a flush, customer thought this was related previous problem that happened recently. | Analyser showing clot in inlet. | Checked logs, this has been intermittent since the 25 $03 $19. Carried out a manual flush, flow not a 100%. Checked flow from SC to Inlet, which was okay, then checked flow from SC to SP, flow not great. Removed wet section and inspected fluid path, waste tube had a noticeable kink in it because the spring had slipped back and had been folded over. Liquid sensors have now improved and clot detect values have also improved. | SWO-535853 | Defect | HW | Basic Unit | Waste tubings |  | Customer removed analyser from AAU as it was asking for a flush, customer thought this was related previous problem that happened recently. | Analyser showing clot in inlet. | Checked logs, this has been intermittent since the 25 $03 $19. Carried out a manual flush, flow not a 100%. Checked flow from SC to Inlet, which was okay, then checked flow from SC to SP, flow not great. Removed wet section and inspected fluid path, waste tube had a noticeable kink in it because the spring had slipped back and had been folded over. Liquid sensors have now improved and clot detect values have also improved.
No | Visit | Complaint | RLTD | Intermittenly not getting Lactate or Calcium results | Intermittenly not getting Lactate or Calcium results. Capillary samples checked logs and patient log result shows Ca 1070 sensor response Lac 1228 correction for lactate contains errors from K+Na+ Ca+ | Advised Sister in charge that they must mix the capillary sample using the mixing wire. The sister did say they are not always doing this. Sidki Akay EBME will send out an email to the doctors regarding mixing the samples. | SWO-535891 | User Error | Sample handling | Sample handling | pre analytical | upSWO sent to MRM | Intermittenly not getting Lactate or Calcium results | Intermittenly not getting Lactate or Calcium results. Capillary samples checked logs and patient log result shows Ca 1070 sensor response Lac 1228 correction for lactate contains errors from K+Na+ Ca+ | Advised Sister in charge that they must mix the capillary sample using the mixing wire. The sister did say they are not always doing this. Sidki Akay EBME will send out an email to the doctors regarding mixing the samples.
No | Visit | Complaint | RCHN | 血红蛋白与临床不符 | 血红蛋白与临床不符（未更换备件） | 做血红蛋白校准后修复 by jesson | SWO-539927 | Adjustment | HW | Oxi module | tHb Cal |  | Hemoglobin is not clinically compatible | Hemoglobin is not clinically compatible (spare parts not replaced) | Fix by jesson after haemoglobin calibration
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 漏电流 0964 | 更换试剂包后修复 | SWO-536169 | [5 empty] | Test card, reagent package failure | Drain current 0964 | Repair after replacement of reagent kit
No | Visit | Complaint | RCHN | solution pack cannot be installed during back flush procedure | solution pack cannot be installed during back flush procedure, solution pack was tried to install several times but it seemed like a bouncing pressure push back the solution pack. | Analyzer was restarted but encountered service code 2 error, application software was reinstalled without reserved database, solution pack was then installed back and passed the service fluid transport test. | SWO-535503 | Defect | SW | RMED SW | database | Status code 2 | solution pack cannot be installed during back flush procedure | solution pack cannot be installed during back flush procedure, solution pack was tried to install several times but it seemed like a bouncing pressure push back the solution pack. | Analyzer was restarted but encountered service code 2 error, application software was reinstalled without reserved database, solution pack was then installed back and passed the service fluid transport test.
No | Visit | Complaint | RCHN | 加热温度超范围，存储盘仓门无法关闭，报错1068,2996 | 加热温度超范围，存储盘仓门无法关闭，报错1068,2996，借903-210，903-117,636-551,903-198各一个 | 更换903-210，903-198后修复 by Andrew | SWO-534294 | Defect | HW | Process Module | PCB, cup dryer | 1068 | Overheating temperature over range, storage warehouse door can not be closed, error 1068,2996 | The heating temperature is over range, the storage warehouse door cannot be closed, error 1068,2996, loan 903-210, 903-117, 636-551,903-198 each | Repair by Andrew after replacing 903-210, 903-198
No | Visit | Complaint | RCHN | 进样舱门无法正常开启 | 进样舱门无法正常开启（未更换备件） | 调节舱门位置后修复 by Andrew | SWO-536190 | Adjustment | HW | Cover and Chassis | Door, Inlet | 4548 | The intake door does not open properly | The intake door does not open properly (spare parts not replaced) | Fix by Andrew after adjusting the hatch position
Yes | Visit | Complaint | RCHN | 电源模块损坏 | 电源模块损坏借903-175一个。未更换备件修复 | 医院插座问题，换了个插座后修复 by Andrew | SWO-534907 | [5 empty] | Damagetopower module | The power supply is damaged by borrowing 903-175 one. Unreplaced spare parts repair | Hospital socket problem, replace by Andrew after changing socket
No | Visit | Complaint | RCHN | 测试卡未连接，定标失败 | 测试卡未连接，定标失败，借903-613一个 | 更换903-613后修复 by Andrew | SWO-534070 | [5 empty] | The test card is not connected and the bid failed | The test card is not connected, the bid failed, borrows 903-613 one | Fix by Andrew after replacing 903-613
Yes | Visit | Complaint | RCHN | 开机后无法进入系统 | 开机后无法进入系统，借910-332、636-660和 636-576各一个 | 更换636-660,636-576后修复 by Andrew | SWO-534223 | [5 empty] | Cannot enter the system after power-on | The system cannot be accessed after power-on, with one each of 910-332, 636-660 and 636-576 | Repair by Andrew after replacing 636-660, 636-576
No | Visit | Complaint | RIND | pH calibration error. | Check the system. | Check the calibration & electrode responses. | SWO-536454 | No fault found | No fault found | No fault found | No fault found | pH | pH calibration error. | Check the system. | Check the calibration & electrode responses.
No | Visit | Complaint | RAME | The printer has streaks in the print | Print faded on the left side | I replaced the printer for the third time. | SWO-536847 | Defect | HW | Printer Module | Printer Module |  | The printer has streaks in the print | Print faded on the left side | I replaced the printer for the third time.
No | Visit | Complaint | RAME | Printer faded | The printer is faded on the right side | I replaced the printer | SWO-536848 | Defect | HW | Printer Module | Printer Module |  | Printer faded | The printer is faded on the right side | I replaced the printer
No | Visit | Complaint | RDEU | 16.04.2019; Hr. Barrenbrügge; Tel.: 025418912191; immer wieder Na und Ca Kal Driftfehler | Zeitweise Elektrolytdriften | Gerät überprüft, E1001 erneuert, Glu und Lac membraniert | SWO-535554 | Assisted Replacement | Electrode | Electrode | reference | CAPA-733 | 16.04.2019; Mr. Barrenbrügge; Tel.: 025418912191; again and again Na and Ca Kal drift error | Temporary electrolyte drifts | Device checked, E1001 renewed, Glu and Lac membraneed
No | Visit | Complaint | RESP | Erro 1999 undefined_lookup_string | Sonda torta | Substituição da sonda de aspiração. Substituição do filtro da unidade de lavagem. Limpeza da unidade de lavagem. Inicialização. Substituição do cartucho de soluções. Limpeza e verificação do sistema. Execução de dois níveis de controle de qualidade. Realização de várias amostras. | SWO-535743 | Defect | HW | Aspiration Module | Needle Assembly |  | Erro 1999 undefined_lookup_string | Cake Probe | Replacing the suction probe. Replacing the washing unit filter. Cleaning the washing unit. Startup. Replacing the solution cartridge. Cleaning and checking the system. Execution of two levels of quality control. Conducting several samples.
No | Visit | Complaint | RFRA | Lors d'un chargement d'une cartouche de Crp la porte est restée ouverte. N'ayant plus de reponse de la porte la cliente a tenté de faire un arret temporaire de la machine, mais cela n'a pas fonctionné. Il y'a donc une coupure sauvage de l'analyseur afin de le relancer. Depuis l'appareil ne redemarre plus. Celui-ci est bloqué sur disk boot failure . | Annulation de l'intervention programmée pour intervention urgente ch gourdon. Carte cf HS. | Remplacement de la carte cf. Rechargement Windows xp. Chargement du programme Aqt 8.9. Démarrage ok. Rechargement des sauvegardes récupérées sur le serveur. Configuration adresse Ip Aqt. Chargement du run pisa.cmd Redémarrage, vérification connexion ok. Passage echantillon et cq. | SWO-535908 | Defect | HW | Analyzer Control Module | CF card |  | When loading a Crp cartridge the door remained open. With no answer from the door, the client attempted to make a temporary stop to the machine, but it did not work. So there is a wild cut from the analyzer in order to restart it. Since then the device no longer restarts. This one is blocked on disk boot failure. | Cancellation of the intervention scheduled for urgent intervention ch gourdon. Map cf HS. | Replacing the cf card. Recharging Windows xp. Loading the Aqt 8.9 program. Start ok. Reloading backups recovered from the server. Setting up Ip Aqt address. Loading run pisa.cmd Restart, login check ok. Passage sample and cq.
No | Visit | Complaint | RDEU | 18.04.2019 11:50Uhr Frau Weidenhiller tel:+498431545478: Wegen einer Fehlermeldung wurde das Gerät neu gestartet und meldet jetzt 1052 und 1051 | Wegen einer Fehlermeldung wurde das Gerät neu gestartet und meldet jetzt 1052 und 1051 FTC-106 | Systemsüberprüfung hacken raus und wieder rein, Safe Positions, Lösungspack erneuert, Systemsreinigung, LQC MCC, LQC PCT, LQC CRP gemessenen. | SWO-536335 | Defect | Consumables | Solution Pack | Valves | 1052 | 18.04.2019 11:50Pm Ms Weidenhiller tel:+498431545478: Due to an error message the device was restarted and now reports 1052 and 1051 | Due to an error message, the device was restarted and now reports 1052 and 1051 FTC-106 | Systemsüberprüfung hacken raus und wieder rein, Safe Positions, Lösungspack erneuert, Systemsreinigung, LQC MCC, LQC PCT, LQC CRP gemessenen.
No | Visit | Complaint | RMAY | Oncall breakdown | Liquid transport failed due to clot in inlet | - Check analyser data log - Perform manual flush - Run fluid transport test | SWO-540108 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway | Clot, 1343 | Oncall breakdown | Liquid transport failed due to clot in inlet | - Check analyser data log - Perform manual flush - Run fluid transport test
Yes | Visit | Complaint | RCHN | 接口板腐蚀 | 接口板腐蚀，借902-972一个 | 更换902-972备件后修复 by Andrew | SWO-535577 | [5 empty] | Interface plate corrosion | Interface plate corrosion, borrow 902-972 one | Repair by Andrew after replacement of 902-972 spare parts
No | Visit | Complaint | RCHN | 试剂泵电机不转 | 试剂泵电机不转，借580-093一个 | 需加借备件，此单关单 by Andrew | SWO-536100 | [5 empty] | Reagent pump motor does not turn | Reagent pump motor does not turn, borrow 580-093 one | Spare parts required, this single by Andrew
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | GLU,LAC定标失败 0375,0588,1310 | 连续更换2张测试卡后修复 | SWO-536167 | [5 empty] | Test card, reagent package failure | GLU, LAC calibration failure 0375,0588,1310 | Repair after 2 consecutive test cards
No | Visit | Complaint | RCHN | 漏液 | 漏液，借844-442一个 | 更换844-442后修复 by Andrew | SWO-534966 | Cleaned | HW | Aspiration Module | Needle Wash Unit | From CRM: First Cleaning and in second SWO change of 844-442 | Leak | Leak, borrow 844-442 one | Repair by Andrew after replacing 844-442
Yes | Visit | Complaint | RCHN | 漏液 | 液路故障 3415 | 清洁HCT模块后修复 by Andrew | SWO-536468 | Re classified RMED | Same day UpSWO |  |  |  | Leak | Liquid Road Fault 3415 | Repair by Andrew after cleaning HCT module
No | Visit | Complaint | RCHN | 报错1024、1025、1031 | 报错1024、1025、1031，借903-364一个 | 更换903-364后修复 by Andrew | SWO-535674 | Defect | HW | Process Module | Optical Unit | 1024, 1031 | Errors 1024, 1025, 1031 | 1024, 1025, 1031, borrow 903-364 one | Repair by Andrew after replacing 903-364
No | Visit | Complaint | RCHN | TNI结果打问号 | TNI结果打问号，借958-201,840-681各一个 | 更换958-201,840-681后修复 by Andrew | SWO-535404 | Defect | HW | Aspiration Module | Cavro, Syringe |  | TNI results hit the question mark | TNI results hit the question mark, borrowing 958-201, 840-681 each | Repair by Andrew after replacing 958-201, 840-681
No | Visit | Complaint | RBNL | Cassettes geblokkeerd | Cassette zit geblokkeerd thv punch out positie, reden is 2 afgebroken lipjes zodat cassette niet meer teruggetrokken wordt. Controle van alle cassettes, geen afwijkingen gevonden In- en uitschuiven van cassette gebeurt heel soepel. Ziet er naar uit dat de cassette met afgebroken lipjes werd geplaatst in de analyzer. | Geen | SWO-536261 | Defect | Consumables | Test Kit | Cartridge, Broken | Teeth | Cassettes blocked | Cassette is blocked THV punch out position, reason is 2 broken tabs so that cassette is no longer retracted. Checking all cassettes, no deviations found In and out of cassette is very smooth. Looks like the cassette with broken tabs was placed in the analyzer. | No
No | Visit | Complaint | RPOL | Aparat rozłącza się z serwerem | Aparat rozłącza się z serwerem codziennie o 5.00 po aktualizacji systemu. | Sprawdzenie połączenia LIS $HIS. Aktualizacja karty sieciowej na serwerze wykonana przez informatyków TK Med. | SWO-536227 | Defect | SW | External SW | HIS/LIS |  | The camera disconnects from the server | The camera disconnects from the server every day at 5.00 after the system update. | Check the LIS $HIS connection. Update of the network card on the server made by IT Professionals TK Med.
No | Visit | Complaint | RPOL | Błąd mechaniczny storage wheel | Zablokowana kaseta odczynnikowego w storage wheel | Usunięcie uszkodzonej i zablokowanej kasety z modułu storage wheel. Sprawdzenie analizatora. | SWO-537417 | Defect | Consumables | Test Kit | Cartridge, Broken |  | Mechanical Storage Wheel Error | Locked Storage Wheel | Remove the damaged and locked cassette from the storage wheel. Inspect the analyzer.
No | Visit | Complaint | RIND | QC Problem | Checked and found that in QC, All electrolytes values in low range, not in given range. | Checked with another Lot QC which I taken from Sri Ramakrishna Hospital and informed them problem in their QC and Promoted Radiometer QC. | SWO-539764 | [5 empty] | QC Problem | Checked and found that in QC, All electrolytes values in low range, not in given range. | Checked with another Lot QC which I taken from Sri Ramakrishna Hospital and informed them problem in their QC and Promoted Radiometer QC.
No | Visit | Complaint | RCHN | 进样舱门关不上 | 进样舱门关不上，借903-028一个 | 更换903-028后修复 by Andrew | SWO-536089 | Defect | HW | Cover and Chassis | PCB, cabinet top | 4548 | The entrance door doesn't close. | The entrance door is not closed, borrow 903-028 one | Fix by Andrew after replacing 903-028
No | Visit | Complaint | RMAY | Oncall breakdown | Found clot in inlet gasket | - Check analyser data log - Perform manual flush - Run fluid transport test | SWO-540127 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway | Clot, 1352 | Oncall breakdown | Found clot in inlet gasket | - Check analyser data log - Perform manual flush - Run fluid transport test
No | Visit | Complaint | RCHN | LAC响应错误 | LAC响应错误，借903-581一个 | 更换903-581后修复 by Andrew | SWO-536210 | Defect | HW | Basic Unit | PCB, Sensor interface | Lac | LAC response error | LAC response error, borrow 903-581 one | Repair by Andrew after replacement 903-581
No | Visit | Complaint | RCHN | 样本扫码器排线断裂 | 样本扫码器排线断裂,需要借件636-666 | 更换636-666后修复 by Andrew | SWO-536105 | Defect | HW | Inlet module | Cable, barcode reader |  | Sample sweeper line break | Sample scanner line break, need to borrow 636-666 | Repair by Andrew after replacement 636-666
No | Visit | Complaint | RMAY | Service breakdown | Calibration error due to clot in chamber | - Check analyser data log - Flush inlet - Clean chamber | SWO-540118 | Cleaned | HW | El/Met Module | measuring chamber | Na, Clot | Service breakdown | Calibration error due to clot in chamber | - Check analyser data log - Flush inlet - Clean chamber
No | Visit | Complaint | RMAY | Service breakdown | Calibration error due to poor membrane condition | - Check analyzer data log - Clean chamber - Flush inlet - Remembrane ( user ) | SWO-540117 | Assisted Replacement | Electrode | Membrane | Glu | 378 | Service breakdown | Calibration error due to poor membrane condition | - Check analyzer data log - Clean chamber - Flush inlet - Remembrane ( user )
No | Visit | Complaint | RKOR | Sample clot error | . | Liquid flow line clean. Gasket module clean. Inlet probe clot => Clean. Service mode check. Sample test. | SWO-539986 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway | Clot | Sample clot error | . | Liquid flow line clean. Gasket module clean. Inlet probe clot => Clean. Service mode check. Sample test.
No | Visit | Complaint | RCHN | 电解质测量结果不准 | 电解质测量结果不准 | 到现场同客户沟通，客户使用肝素抗凝配比不合适，待调整配比后进一步跟进 by jesson | SWO-539480 | [5 empty] | Electrolyte measurement results are not allowed | Electrolyte measurement results are not allowed | To the scene to communicate with customers, customers use heparin anticoagulant ratio is not appropriate, to adjust the ratio after further follow-up by jesson
No | Visit | Complaint | RAME | Analyzer is not autodetecting manual QC. | Analyzer was found to be not auto detecting QC on last visit for pCO2 QC Errors. Ordered parts and made a return trip. Upgraded customer software to version 3.18. | Upgraded the customers software and that seems to have resolved the issue. | SWO-536837 | [5 empty] | Analyzer is not autodetecting manual QC. | Analyzer was found to be not auto detecting QC on last visit for pCO2 QC Errors. Ordered parts and made a return trip. Upgraded customer software to version 3.18. | Upgraded the customers software and that seems to have resolved the issue.
No | Visit | Complaint | RCHN | 试剂泵电机不转 | 加借903-334一个 | 更换903-334备件后修复 by jesson | SWO-536386 | [5 empty] | Reagent pump motor does not turn | Borrow 903-334 one | Repair by jesson after replacement of 903-334 spare parts
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 流量选择器定标错误0978 | 更换试剂包后修复 | SWO-538524 | [5 empty] | Test card, reagent package failure | Flow Selector Calibration Error 0978 | Repair after replacement of reagent kit
No | Visit | Complaint | RCHN | TNT定标失败 | TNT定标失败，借944-268（15733）*2个 | 检查处理盘上卡杯子的弹簧有问题，处理后重新定标后修复 by jesson | SWO-536532 | Adjustment | HW | Process Module | Spring, process wheel | 3415 | TNT calibration failure | TNT calibration failure, borrowed 944-268 (15733) x 2 | Check for problems with the spring of the card cup on the processing disc and fix by jesson after re-calibration after treatment
No | Visit | Complaint | RCHN | 舱门无法打开 | 舱门无法打开， | 手动调试并加垫片后修复 by jesson | SWO-536681 | Adjustment | HW | Cover and Chassis | Door, Inlet |  | The hatch cannot be opened. | The hatch cannot be opened. | Fix by jesson after manual commissioning and gasket
No | Visit | Complaint | RCHN | 定标卡扫描器无法扫描 | 定标卡扫描器无法扫描，借902-924一个 |  | SWO-536789 | Defect | HW | Cover and Chassis | Barcode reader, display | Wait for more info in SWO: Corrective Action: Repaired after replacing 902-924 spare parts by jesson | The calibration card scanner cannot be scanned | The calibration card scanner cannot be scanned, borrowing 902-924 one
No | Visit | Complaint | RPAC | Lactate calibrations intermittently failing with 0379: calibration unstable (response fault) since 17/4, failing more frequently on 22/4 Membrane last replaced 25/3 R4073. 22/4 Installed new lactate membrane which fixed issue. Also replaced glucose, reference and crea membranes due. | Lactate calibrations failing with 0379: calibration unstable (reponse fault), intermittent errors since 17 $4. Lactate membrane R4073 installed 25 $3. Replaced lactate membrane fixed issue, calibration and QC OK. Replaced glucose, reference and crea also due. | Replaced lactate membrane. Faulty membrane removed R4073. | SWO-536780 | Assisted Replacement | Electrode | Membrane | Lac | 379, R4073 | Lactate calibrations intermittently failing with 0379: calibration unstable (response fault) since 17/4, failing more frequently on 22/4 Membrane last replaced 25/3 R4073. 22/4 Installed new lactate membrane which fixed issue. Also replaced glucose, reference and crea membranes due. | Lactate calibrations failing with 0379: calibration unstable (reponse fault), intermittent errors since 17 /4. Lactate membrane R4073 installed 25 /3. Replaced lactate membrane fixed issue, calibration and QC OK. Replaced glucose, reference and crea also due. | Replaced lactate membrane. Faulty membrane removed R4073.
No | Visit | Complaint | RIND | Po2 ? mark for all samples | PO2 ? mark for all samples | Checked the analyser po2 was giving ? mark for all samples so cleaned the po2 electrode and po2 measuring chamber then tried but its failed in samples so latch was slightly rubbed using needle and pumpa nd ls adjust calibration done then calibrated the analyser to raedy mode and analysed the sample observed for few samples working fine | SWO-539698 | Adjustment | HW | pH/BG Module | LS | pO2 | Po2 ? mark for all samples | PO2 ? mark for all samples | Checked the analyser po2 was giving ? mark for all samples so cleaned the po2 electrode and po2 measuring chamber then tried but its failed in samples so latch was slightly rubbed using needle and pumpa nd ls adjust calibration done then calibrated the analyser to raedy mode and analysed the sample observed for few samples working fine
No | Visit | Complaint | RAME | The MetHb measures high on blood samples compared to two other analyzers. | Onsite for autocheck scanning issue and had part available to help troubleshoot oximetry issue. | Replaced oximetry module for troubleshooting purposes. Ran tHb calibration. | SWO-537022 | Defect | HW | Oxi module | Oxi module |  | The MetHb measures high on blood samples compared to two other analyzers. | Onsite for autocheck scanning issue and had part available to help troubleshoot oximetry issue. | Replaced oximetry module for troubleshooting purposes. Ran tHb calibration.
No | Visit | Complaint | RAME | Autocheck does not read qc. | Checked barcode scanning is=n service mode. Adjusted the nut to change the laser. Unable to get the scanner to read all autocheck vials correctly. | Replaced laser scanner for autocheck module. Adjusted laser to read all vials. It is now reading each vial and showing missing when no vial is present. Analyzer displayed a status code 2 error when rebooting the analyzer after repair. Unable to retrieve and load customers data base from today. Customer had a thumb drive with a back up data base from 4 $18 $19. I was able to load that on the analyzer and reboot. The customer lost 4 days of data due to a corrupt data base. | SWO-537021 | Adjustment | HW | Autocheck module | Barcode reader |  | Autocheck does not read qc. | Checked barcode scanning is=n service mode. Adjusted the nut to change the laser. Unable to get the scanner to read all autocheck vials correctly. | Replaced laser scanner for autocheck module. Adjusted laser to read all vials. It is now reading each vial and showing missing when no vial is present. Analyzer displayed a status code 2 error when rebooting the analyzer after repair. Unable to retrieve and load customers data base from today. Customer had a thumb drive with a back up data base from 4 $18 $19. I was able to load that on the analyzer and reboot. The customer lost 4 days of data due to a corrupt data base.
No | Visit | Complaint | RAME | No sample detected | Found multiple pump errors as well as it was leaking from ph $BG measuring chamber front. | Found pH sealing was not on all the way. As well as very bad clot. Reseat the sealing ring and flush the clot. Ran multiple pump cals and 2 point cals after wards. All calibrations and 3 Levels of QC passed as per manufacturer specification. No more leaking from front. Cleaned out all the blood on the instrument. Customer will clean out around the analyzer. On the side note I was told by one of the tech the probe was really bent when all the issues started. | SWO-537032 | Cleaned | HW | pH/BG Module | measuring chamber | Clot, pH | No sample detected | Found multiple pump errors as well as it was leaking from ph /BG measuring chamber front. | Found pH sealing was not on all the way. As well as very bad clot. Reseat the sealing ring and flush the clot. Ran multiple pump cals and 2 point cals after wards. All calibrations and 3 Levels of QC passed as per manufacturer specification. No more leaking from front. Cleaned out all the blood on the instrument. Customer will clean out around the analyzer. On the side note I was told by one of the tech the probe was really bent when all the issues started.
No | Visit | Complaint | RDEU | 23.04.2019 11:20Uhr Herr Stallmeyer tel:+4925418947124: Leckagefehler | Vacuumfehler, Einlaßdichtung verschlissen | Gerät überprüft und gereinigt, Einlaßdichtung ersetzt, | SWO-537466 | Defect | HW | Inlet module | Inlet gasket | Vacuum error | 23.04.2019 11:20am Mr Stallmeyer tel:+4925418947124: Leakage error | Vacuum error, inlet seal worn | device checked and cleaned, inlet seal replaced,
Yes | Visit | Complaint | RDEU | 23.04.2019 11:20Uhr Herr Stallmeyer tel:+4925418947124: Leckagefehler | Einlaßdichtung defekt, pCO2 Empf. Fehler (Gerät 4 Tage nicht kalibriert) | Gerät überprüft, Einlaßdichtung ersetzt, Meßkassette erneuert, Kal und QC o.k. | SWO-537510 | Re classified RMED | Same day UpSWO |  |  |  | 23.04.2019 11:20am Mr Stallmeyer tel:+4925418947124: Leakage error | Inlet seal defective, pCO2 empf. Error (device not calibrated for 4 days) | Device checked, inlet seal replaced, measuring cassette renewed, Kal and QC ok.
No | Visit | Complaint | RDEU | 23.04.2019 11:40Uhr Sr.Heike tel:+4920360083011 : Zwei Gerät zeigen den Fehler pCO2 QC-Verletzung bei LevelA | Gerät zeigt einen CO2 QC Fehler. | LP ersetzt und Gerät kalibriert, QC durchgeführt, | SWO-537875 | Assisted Replacement | SP | SP | pCO2 | QC A high | Apr 23, 2019 11:40am Sr.Heike tel:+4920360083011 : Two devices show the error pCO2 QC violation at LevelA | Device shows a CO2 QC error. | LP replaced and device calibrated, QC performed,
No | Visit | Complaint | RPAC | Stacey called 0448 887 058 Connectivity issue, turned machine off. Turned back on, message replace inlet gasket holder | Phonecall | Perform Fluidic checks and back flushes. Pressure test failed. Check the pinch valve for leaks. Non found. Remove inlet gasket and clean. Repeat Fluidics check and Pressure test OK. Repeat pressure test several times. All OK. Start up Cal and QC's all good. | SWO-537161 | Cleaned | HW | Auto Inlet module | Inlet gasket |  | Stacey called 0448 887 058 Connectivity issue, turned machine off. Turned back on, message replace inlet gasket holder | Phonecall | Perform Fluidic checks and back flushes. Pressure test failed. Check the pinch valve for leaks. Non found. Remove inlet gasket and clean. Repeat Fluidics check and Pressure test OK. Repeat pressure test several times. All OK. Start up Cal and QC's all good.
No | Visit | Complaint | RKOR | Vacumme error | . | Liquid flow line clean. Gasket module clean. Service mode check. | SWO-539987 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway | Clot | Vacumme error | . | Liquid flow line clean. Gasket module clean. Service mode check.
No | Visit | Complaint | RAME | pO2 issues | Solution pack was out of activities because was having PO2 calibration errors. | Replaced solution pack from N003 and all Cals and internal and external QC passed. | SWO-537773 | Assisted Replacement | SP | SP | pO2 | Cal | pO2 issues | Solution pack was out of activities because was having PO2 calibration errors. | Replaced solution pack from N003 and all Cals and internal and external QC passed.
No | Visit | Complaint | RAME | Set up issues/set up is not the same as the other | Install Astm.ini file. | Install new Astm.ini file so instrument will communicate with Ralls. | SWO-537720 | Setup | SW | RMED SW | HIS/LIS |  | Set up issues/set up is not the same as the other | Install Astm.ini file. | Install new Astm.ini file so instrument will communicate with Ralls.
No | Visit | Complaint | RIND | In Level 3 QC Na and pCO2 going out continuously. | Check System | Check System, check calibration, check QC data log, done zeroing, install new cassate , still same problem occur in QC level3 ,system need tobe clean using cleaning SP and SC. | SWO-538245 | [5 empty] | In Level 3 QC Na and pCO2 going out continuously. | Check System | Check System, check calibration, check QC data log, done zeroing, install new cassate , still same problem occur in QC level3 ,system need tobe clean using cleaning SP and SC.
No | Visit | Complaint | RCHN | 机器无法开机 | 机器无法开机，需要借件910-342一个 |  | SWO-537233 | Defect | HW | DMS_PC | CPU |  | The machine does not power on | The machine cannot be switched on, need to borrow 910-342 one
No | Visit | Complaint | RCHN | 进样口支架断裂 | 进样口支架断裂 | 更换库存进样口支架，清洁管路后修复 by jesson | SWO-537151 | Defect | HW | Inlet module | Inlet gasket |  | The inlet bracket is broken | The inlet bracket is broken | Replace the stock inlet bracket and repair by jesson after cleaning the line
No | Visit | Complaint | RCHN | 无法读取试剂包信息 | 无法读取试剂包信息,借903-158一个。 |  | SWO-537247 | Defect | HW | Basic Unit | PCB, solution pack |  | Unable to read package information | The package information could not be read, borrowing 903-158 one.
No | Visit | Complaint | RCHN | 条码扫描器无法扫描 | 条码扫描器无法扫描,需要借件 910-304 | 更换910-304后修复 by jesson | SWO-536676 | Defect | HW | Storage Module | Barcode reader, storage wheel |  | Barcode scanner cannot be scanned | Barcode scanner cannot be scanned and needs to be borrowed 910-304 | Repair by jesson after replacing 910-304
No | Visit | Complaint | RCHN | 条码扫不出 | 条码扫不出，借910-304一个 |  | SWO-537113 | Defect | HW | Storage Module | Barcode reader, storage wheel | Wait for more info in SWO: Corrective Action: Repair after replacing 910-304 by jesson | Bar code can't be scanned | Bar code can't sweep, borrow 910-304 a
No | Visit | Complaint | RCHN | TNT值偏高 | TNT值偏高，借958-201一个 |  | SWO-537174 | Defect | HW | Aspiration Module | Cavro, Syringe | Wait for more info in SWO: Corrective Action: Replacement 958-201 is normal. By jesson | TNT value is high | TNT value is high, borrow 958-201 one
No | Visit | Complaint | RCHN | 试管夹断裂 | 试管夹断裂，借903-584 $2个 |  | SWO-537122 | Defect | HW | Inlet module | Tube holder |  | Test tube clip break | Test tube clip broken, borrowed 903-584 /2
No | Visit | Complaint | RCHN | 定标失败 | 定标失败，借944-258（15732）一个 |  | SWO-537150 | Cleaned | HW | Process Module | Process Wheel | Wait for more info in SWO: Corrective Action: Clean the processing disk and fix it after recalibration by jesson | Bid failed | Failed to bid, borrow 944-258 (15732) one
Yes | Visit | Complaint | RCHN | 定标失败 | 定标过程中报错1025 | 清洁处理盘后修复 by jesson | SWO-539267 | Re classified RMED | Same day UpSWO |  |  |  | Bid failed | Error 1025 during the bid calibration process | Repair by jesson after cleaning the processing disc
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | lac定标失败 | 更换测试卡后修复 | SWO-537165 | [5 empty] | Test card, reagent package failure | lac calibration failed | Repair after replacing test card
No | Visit | Complaint | RAME | The creatinine value not match with LIS | The creatinine value not matched between LIS and analyzer print out | The analyzer output tested by using Docklight software and the output were completely matched with the print out and screen For the sample. With meaning the QC analyzer is sending two different values for the temperature corrected parameters and LIS should pick up what customer needs to be reported. The analyzer works properly as it designed. Customer and IT Department adviced to work on LIS configurations. | SWO-537550 | No fault found | SW | RMED SW | HIS/LIS |  | The creatinine value not match with LIS | The creatinine value not matched between LIS and analyzer print out | The analyzer output tested by using Docklight software and the output were completely matched with the print out and screen For the sample. With meaning the QC analyzer is sending two different values for the temperature corrected parameters and LIS should pick up what customer needs to be reported. The analyzer works properly as it designed. Customer and IT Department adviced to work on LIS configurations.
No | Visit | Complaint | RPOL | Aparat zawiesza się podczas wymiany pakietu roztworów. | Serwis stwierdza zawieszenie aparatu w ostatnim kroku wymiany pakietu roztworów. | Wyłączenie aparatu, wyjęcie pakietu roztworów przy pomocy awaryjnego wysuwania, ponowny rozruch i instalacja pakietu roztworów oraz kasety sensorowej. | SWO-537385 | Assisted Replacement | SP | SP | SP |  | The camera freezes when replacing the solution package. | The service finds the camera suspension in the last step of replacing the solution package. | Shutdown of the apparatus, removal of the solution package by means of emergency ejecting, rebooting and installation of the solution package and the sensor cassette.
No | Visit | Complaint | RESP | Erro de vácuo | Erro de vácuo com valor de 92 | Substituição da gasket sem sucesso, substituição do pack sem sucesso, substituição da entrada com o equipamento ao lado sem sucesso, teste de vácuo com a SC antiga e resolveu, substituição de um SC nova e verificação de problema resolvido, estabilização e passagem do QC. | SWO-538394 | Assisted Replacement | SC | SC | Vaccuum error | 1063 | Vacuum error | Vacuum error with 92 value | Replacing the gasket without success, replacing the pack without success, replacing the input with the equipment beside the unsuccessful, vacuum test with the old SC and solved, replacing a new SC and verifying problem solved, stabilization and passage of the Qc.
No | Visit | Complaint | RPAC | Suzanne called 5722 5201 snapped off where sample is placed into port | Inlet Gaslet broken and Inlet Probe bent. | Replaced the Inlet Gasket and Inlet Probe. Returned the analyser back to normal and allowed to re-initialise. All is now functioning. | SWO-537846 | Defect | HW | Inlet module | Inlet gasket |  | Suzanne called 5722 5201 snapped off where sample is placed into port | Inlet Gaslet broken and Inlet Probe bent. | Replaced the Inlet Gasket and Inlet Probe. Returned the analyser back to normal and allowed to re-initialise. All is now functioning.
No | Visit | Complaint | RDEU | 08.04.2019; Hr. Weiser; Tel.: 0765129413; ABL verlängert die Messzeit 11.04.2019, Frau Guth, Tel: 0765129413: ständige Kalibrierungen ca im 30 Min. Takt 12.04.2019 08:30Uhr Frau Haas tel:+4976512978412: Das Gerät geht ständig in die Startphase | Aufstellung Ersatzgerät. Reboot bei fehlerhaftem Gerät findet regelmässig nach automatischer Archivierung auf. Archivierung abgewählt. Festplattenspeicher kleiner 15%, Festplatte von 4GB temporärer Explorerdateien bereinigt. | Konfiguration und CCO-Dateien übertragen. Netzwerkanbindung durchgeführt. Gerät angebunden, Verbindung zu aqure überprüft, Daten auf CF-Karte gesichert. Automatische Archivierung abgeschaltet. Gerät in der VSC in Betrieb genommen. | SWO-538610 | Revisit required | analyzer removed | Analyzer removed | Analyzer removed | Reboot | 08.04.2019; Mr. Weiser; Tel.: 0765129413; ABL extends the measuring time 11.04.2019, Mrs. Guth, Tel: 0765129413: continuous calibrations approx. 30 min. 12.04.2019 08:30Pm Ms Haas tel:+4976512978412: The device is constantly going into the start-up phase | Installation of replacement device. Reboot in case of faulty device is regularly resonated after automatic archiving. Archiving has been deselected. Disk space less than 15%, hard drive cleaned of 4GB of temporary explorer files. | configuration and CCO files. network connection. Device connected, connection to aqure checked, data backed up to CF card. Automatic archiving switched off. device in the VSC.
Yes | Visit | Complaint | RHUN | Oldat pumpálási hibák | Oldat pumpálási hibák(Cal2, Cal3 véletlenszerüen) Felkészítés üzembehelyezésre | Alaktrészek cseréje, Szerviz-tesztek,10x Solution Pack check, Inicializálás, QC mérés | SWO-538588 | [5 empty] | Solution for pumping defects | Solution for pumping defects (Cal2, Cal3 randomization) Preparation for commissioning | Replace shape, service tests, 10x Solution Pack Check, initialization, QC measurement
No | Visit | Complaint | RLTD | Contacted direct by customer Haval Ozgun - leak dected even after troubleshooting. | Confirmed that cause of leak detected was due to Oximetry pinch valve tube having a hole in it. Replaced tube as part of annual preventative maintenance which was due soon. Tested in fluid transport after repair - leak test now passed. | Replaced Oximetry pinch valve tube. | SWO-538120 | Defect | HW | Oxi module | Pinch valve tube | Leak | Contacted direct by customer Haval Ozgun - leak dected even after troubleshooting. | Confirmed that cause of leak detected was due to Oximetry pinch valve tube having a hole in it. Replaced tube as part of annual preventative maintenance which was due soon. Tested in fluid transport after repair - leak test now passed. | Replaced Oximetry pinch valve tube.
No | Visit | Complaint | RAME | Clot stuck in analyzer. |  |  | SWO-540241 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway | Clot | Clot stuck in analyzer.
No | Visit | Complaint | RAME | Clot detected in inlet. | Looked in data logs, found clot detected on inlet gasket.. | Did 20 back flushes with warm water, Replaced solution pack from customers stock and inlet gasket from maintenance kit. | SWO-537776 | Revisit required | HW | Basic Unit | PM kit | Clot | Clot detected in inlet. | Looked in data logs, found clot detected on inlet gasket.. | Did 20 back flushes with warm water, Replaced solution pack from customers stock and inlet gasket from maintenance kit.
No | Visit | Complaint | RDEU | 23.04.2019, 16:45 Uhr ==> Anrufer Hebamme Elisabeth, Tel.: 073150052631 ==>Unglaubwürdige pH-Messwerte | bei MBU anderer pH Wert wie später mit Nabelschnurblut gemessen. | Untersuchung Dateien. Verdacht auf Gerinnsel in Messkammer,das später herausgespült wurde. Weitere pH Werte unauffällig. Vergleichsmessungen mit Blut bei FU Intensiv und Gyn OP Geräten über 3 Testproben ohne Beanstandung. Gerät nach QC wieder freigegeben. | SWO-538210 | No fault found | No fault found | No fault found | No fault found | Clot, pH | 23.04.2019, 16:45 ==> Caller Midwife Elisabeth, Tel.: 073150052631 ==>Unbelievable pH readings | at MBU other pH value as later measured with umbilical cord blood. | Investigation files. Suspicion of clots in measuring chamber, which was later rinsed out. Further pH values inconspicuous. Comparative measurements with blood in FU Intensive and Gyn OP devices over 3 test samples without complaint. Device after QC.
No | Visit | Complaint | RAME | LS errors unable to pump solutions | unable to pump solutions and LS errors user changed pack multiple times and sensor once with no resolve user also flushed the analyzer multiple times checked service menu LS errors on oxi and sensor - ran service menu checks cleared the erros re installed solution pack within 30 minutes got the errors to come back. flushed the analyzer - seems to be fluid flow problems between waste and oxi module - removed waste line and replaced re calibrated LS ran a visual flow didnt look right tried flushing and it was still not flushing ok. Took the unit apart flushed waste line- ok flushed hemolyzer - ok flushed probe ok flushed bay unit from hemolyzer side blocked | replaced waste line inlet gasket holder gasket and oxi tube re calibrated LS replaced sensor due to clot still in sensor replaced bay unit due to blockage that I was unable to remove | SWO-538230 | Defect | HW | Basic Unit | Sensor Bay Unit | Clot | LS errors unable to pump solutions | unable to pump solutions and LS errors user changed pack multiple times and sensor once with no resolve user also flushed the analyzer multiple times checked service menu LS errors on oxi and sensor - ran service menu checks cleared the erros re installed solution pack within 30 minutes got the errors to come back. flushed the analyzer - seems to be fluid flow problems between waste and oxi module - removed waste line and replaced re calibrated LS ran a visual flow didnt look right tried flushing and it was still not flushing ok. Took the unit apart flushed waste line- ok flushed hemolyzer - ok flushed probe ok flushed bay unit from hemolyzer side blocked | replaced waste line inlet gasket holder gasket and oxi tube re calibrated LS replaced sensor due to clot still in sensor replaced bay unit due to blockage that I was unable to remove
No | Visit | Complaint | RCHN | 检测到泄露，报错1342 | 检测到泄露，报错1342（未更换备件） | 清洗进样口支架后修复 by jesson | SWO-537872 | Cleaned | HW | Inlet module | Inlet gasket | 1342, leak | A leak was detected, with an error of 1342 | A leak was detected and error 1342 was reported (spare parts not replaced) | Repair by jesson after cleaning the inlet bracket
No | Visit | Complaint | RCHN | HCT值测量出错，不能被正确测量 | HCT值测量出错，不能被正确测量，经销商工程师在调换AU的时候上传config文件时可能出错，HCT界面里的K值变成-10了。 | 清洁HCT，晾干后重新做了cell constant定标。 | SWO-537999 | Cleaned | HW | Aspiration Module | Needle Wash Unit | 1258, 1259, 1265 | HCT value measurement error, can not be measured correctly | HCT value measurement error, can not be measured correctly, dealer engineers in the transfer of AU when uploading config files may be error, HCT interface K value sturned to -10. | Clean the HCT, dry and re-do the cell constant calibration.
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 无法读取测试卡智能芯片 | 更换测试卡后修复 | SWO-539825 | [5 empty] | Test card, reagent package failure | Failed to read test card smart chip | Repair after replacing test card
No | Visit | Complaint | RCHN | CTHB定标结果没有显示 | CTHB定标结果没有显示，借902-458一个 | 更换902-458备件后修复 by jesson | SWO-536765 | Defect | HW | Oxi module | Lamp unit | Old lamp unit | CTHB calibration results are not displayed | CTHB calibration results are not shown, borrowing 902-458 a | Repair by jesson after replacement of 902-458 spare parts
No | Visit | Complaint | RDEU | 23.04.2019, 14:30 Uhr ==> Anrufer Herr Mercan, Tel.: 030450675376 ==> AC-House hebt sich bei Messung nicht (Fehlermeldung 0741) | Ersatzteil beschafft | AC Modul erneut ausgebaut und Motor $Getriebeeinheit ersetzt, alles zusammengesetzt, Neustart, mehrere Kal, Flüssigkeit tritt am Einlaß aus - Einlaßdichtung ersetzt, QC | SWO-538587 | Defect | HW | Autocheck module | Stepper Motor |  | 23.04.2019, 14:30 ==> Caller Mr. Mercan, Tel.: 030450675376 ==> AC-House does not stand out when measured (error message 0741) | Spare part procured | AC module re-removed and engine $Getriebeeinheit replaced, all assembled, restart, several cal, liquid exits at inlet - inlet seal replaced, QC
Yes | Visit | Complaint | RDEU | 23.04.2019, 14:30 Uhr ==> Anrufer Herr Mercan, Tel.: 030450675376 ==> AC-House hebt sich bei Messung nicht (Fehlermeldung 0741) | Fehlersuche, AC Modul im Service getestet - keine Bewegung des inneren Moduls, | Gerät abgeschalten, AC Modul ausgebaut, Mechanik überprüft und Schrittmotor für Hub ausgebaut - Getriebe zeigt kaum Wiederstand - möglicherweise defekt - muß beschafft werden - Gerät zusammengesetzt und manuell QC durchgeführt, erneuter Einsatz erforderlich | SWO-538586 | Re classified RMED | Same day UpSWO |  |  |  | 23.04.2019, 14:30 ==> Caller Mr. Mercan, Tel.: 030450675376 ==> AC-House does not stand out when measured (error message 0741) | Troubleshooting, AC module tested in service - no movement of the inner module, | Switch off the device, AC module removed, mechanics checked and stepper motor removed for stroke - gearbox shows hardly any rectairreadiness - possibly defective - must be procured - device assembled and manually QC performed, re-use required
No | Visit | Complaint | RDEU | Kalium Instabil bei Kal und QC | Ca und Na sind sporadisch aus der Rilibäk raus. | Ref. und Na. Elektroden sind bereits ausgetauscht. Membraniert wurden die Elektroden mehrfach. Airventil von beiden Modulen ausgebaut und deren Sitz gereinigt. ( Ablagerungen der Membranventil waren ersichtlich). Nachfüllung. Pumpenkal. LS- Justage. Kalibration 2x. QC ind Ordnung. | SWO-538290 | User expectation | Specifications | Electrode | QC | Na and Ca | Potassium Unstable in Kal and QC | Ca and Na are sporadic out of the Rilibäk. | Ref. and, well, Electrodes have already been replaced. The electrodes were membrane-finished several times. Air valve removed from both modules and their seat cleaned. (Deposits of the diaphragm valve were visible). Refill. Pump scale. LS adjustment. Calibration 2x. QC ind order.
No | Visit | Complaint | RDEU | Na und Lac. Instabil | Natrium bei QC Level1 und 3 instabil. Lac Instabil bei der Kal und QC | Dichtungsring der Lac. Membran eingesetzt. Airvalves der beiden Modulen ausgebaut und den gereinigt. Membranvalve der PH $ BG Module ausgetauscht. Gasflaschen erneuert. Nachfüllung, Pumpenkal und LS- Justage. | SWO-538292 | Defect | HW | pH/BG Module | Valve, membrane |  | Na and Lac. Unstable | Sodium at QC Level1 and 3 unstable. Lac Unstable in Kal and QC | Sealing ring of the Lac. membrane. Airvalves of the two modules and cleaned the. Diaphragm valve of the PH BG modules. gas cylinders renewed. Refilling, pump scale and LS adjustment.
No | Visit | Complaint | RCHN | 仪器界面频繁跳跃到试剂包 | 仪器界面频繁跳跃到试剂包.需要借件902-502和902-503各一个 | 更换902-503后修复 by jesson | SWO-537953 | Defect | HW | Basic Unit | PCB, reagent illumination |  | Instrument interface jumps frequently to reagent package | Instrument interface frequent jumps to reagent package. | Repair by jesson after replacement 902-503
No | Visit | Complaint | RDAN | Mogens Hald siger, at printeren skriver dårligt på midten. Han har prøvet, at rense den uden held. | Set. | Printer udskiftet, slutafprøvet. | SWO-537978 | Defect | HW | Printer Module | Printer Module |  | Mogens Hald says the printer is writing badly in the middle. He has tried to clean it without success. | Set. | Printer replaced, final tested.
No | Visit | Complaint | RDEU | 24.04.2019 08:00Uhr Herr Nather tel:+493036501330 : : Die Pufferkassette wird nicht angenommen, Das Display hat einen Wackelkontakt, Die Software will ständig neu booten | - Bildschirm flackert bei Bewegung - Reagent Pack wird nicht eingelesen | Fehlersuche; Gerät aus; Bildschirmkabel gew.; div. Serviceprogr. durchgef. -> Reagent Pack PCB def.; Abbruch d. Arbeiten; Material bestellt | SWO-538098 | Defect | HW | Analyzer Control Module | Cable, Display |  | 24.04.2019 08:00OPm Mr Nather tel:+493036501330 : : The buffer cartridge is not accepted, The display has a wobbly contact, the software wants to reboot constantly | - Screen flickers when moving - Reagent Pack will not be read in | Troubleshooting; device off; Screen cable. div. Serviceprogr. through it. -> Reagent Pack PCB def.; Demolition of the works; Material ordered
No | Visit | Complaint | RFRA | Configuration. | Pas de connexion Aqure | Modification des adresses ip. Essais connexion ok. Vérification fonctionnement sauvegarde sur réseau ok. | SWO-538550 | Setup | SW | RMED SW | IP address |  | Configuration. | No Aqure connection | Changing ip addresses. Try connection ok. Checking running backup on network ok.
No | Visit | Complaint | RPOL | Błąd QC na parametrach OXI | Przegląd logów aparatu. | Odbiałczanie modułu OXI, kalibracja tHb, testy serwisowe. | SWO-538722 | Revisit required | HW | Oxi module | Hemolyzer |  | QC error on OXI parameters | Overview of the camera logs. | Deleukation of the OXI module, tHb Calibration, service tests.
No | Visit | Complaint | RPOL | Parametry nie zostały zainstalowane. | Przegląd logów | Aparat stracił klucze. Ponowne wgranie nie dało powiodło się. Wstawiono pastylkę z kluczami demo. Wymieniono membrany ref oraz crea i elektrody pH. Kalibracje. | SWO-538723 | User Error, FSE | Demo key | Demo key | Demo key |  | The parameters have not been installed. | Log overview | The camera lost its keys. Re-uploading failed. Inserted a tablet with demo keys. Ref and Crea membranes and pH electrodes were listed. Calibrations.
No | Visit | Complaint | RPOL | Błąd próbki | Przegląd logów. Ocena wizualna. | Wymiana zgiętej igły, kalibracja wlotu, testy. | SWO-538724 | Defect | HW | Inlet module | Probe bend |  | Sample Error | Log overview. Visual evaluation. | Replacing the bent needle, calibration of the inlet, tests.
No | Visit | Complaint | RKOR | pressure vacuum error | Gasket holder replace, line flush, cal control data check. | Gasket holder replace, line flush, cal control data check. | SWO-539889 | Defect | HW | Inlet module | Inlet gasket | Vacuum error | pressure vacuum error | Gasket holder replace, line flush, cal control data check. | Gasket holder replace, line flush, cal control data check.
No | Visit | Complaint | RDEU | 25.04.2019 08:10Uhr SrHeike tel:+4920360083011 : Das Gerät meldet Gerinnsel erkannt, die Nachschicht hat bereits alle nötigen Maßnahmen gestartet, aber erfolglos. | Gerät steht im Haltmodus, Fließwege sollen gespült werden, | Fließwege gespült, alle Serv.programme durchgespielt, Gerät kalibriert und QC durchgeführt, | SWO-539311 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway | Clot | 25.04.2019 08:10H SrHeike tel:+4920360083011 : The device detects clots, the post-layer has already started all necessary measures, but to no avail. | device is in holding mode, flow paths should be flushed, | Flow paths rinsed, all Serv.programs played through, device calibrated and QC performed,
No | Visit | Complaint | RSAF | No more activities left in the solution pack | None | Replaced solution park | SWO-539170 | Assisted Replacement | SP | SP | SP | Potentially Use Error (Lack of training) | No more activities left in the solution pack | None | Replaced solution park
No | Visit | Complaint | RIND | Service Visit. | Need to check the analyzer. | After observing the analyzer, there is air in sample that is why *** is coming. Report to Dr. Gupta and Dr, Vivek Joshi. Analyzer handed over to staff in well working condition. | SWO-539852 | [5 empty] | Service Visit. | Need to check the analyzer. | After observing the analyzer, there is air in sample that is why *** is coming. Report to Dr. Gupta and Dr, Vivek Joshi. Analyzer handed over to staff in well working condition.
No | Visit | Complaint | RCHN | 条码扫描器无法扫描 | 条码扫描器无法扫描，需要借件910-304一个 | 更换910-304一个正常。by jesson | SWO-538003 | Defect | HW | Storage Module | Barcode reader, storage wheel |  | Barcode scanner cannot be scanned | Barcode scanner cannot be scanned and requires a loan of 910-304 | Replace 910-304 one normal. by jesson
No | 7 | Y | RCZE | Blank | Za posledny mesiac cca 3x chyba 963, po restarte respektive rekalibracii vsetko ok, pri kontrole v isolation programe toto v norme. Respektive pred tyzdnom kontrola ventilov na pH $BG a El $Met pri chybe 500. Dnes o 18,00 hlasena opatovne tato chyba. |  | D6071A04BC625075 | Defect | HW | Fluidic manifold | Valve | V8 | Blank | For the last month approx. 3x error 963, after Restarte recalibration all OK, when checking in isolation headphones program this in the standard. Respektive before the rod control of the valves to pH /BG and El /Met at Error 500. Today about 18.00 The voice of the custody of the mistake.
No | 3 | Y | RCZE | Blank |  |  | D6071562BC6260C4 | Assisted Replacement | Electrode | Electrode | reference | CAPA-733 | Blank
No | Visit | Complaint | RIND | Printer not workinig | Printer is not giving print on paper. Found D lock was open. | Lock the D lock with roller. Power on the instrument. Take a print. Print is correct. | SWO-539071 | User Error | HW | Printer Module | Printer Module | Paper | Printer not workinig | Printer is not giving print on paper. Found D lock was open. | Lock the D lock with roller. Power on the instrument. Take a print. Print is correct.
No | Visit | Complaint | RDEU | 24.04.2019 08:00Uhr Herr Nather tel:+493036501330 : : Die Pufferkassette wird nicht angenommen, Das Display hat einen Wackelkontakt, Die Software will ständig neu booten | Reagent Pack wird nicht gelesen | Gerät aus + zerlegt; Reagent Pack PCB (alt: R2002N346) gew.; Gerät ein; Reagentpack installiert; Startphase abgewartet; keine QC, da ZI im Anschluß durchgef. (siehe folgender SWO) | SWO-538676 | Defect | HW | Cover and Chassis | PCB, solution pack |  | 24.04.2019 08:00OPm Mr Nather tel:+493036501330 : : The buffer cartridge is not accepted, The display has a wobbly contact, the software wants to reboot constantly | Reagent Pack is not read | Device disassembled from +; Reagent Pack PCB (old: R2002N346); device; Reagentpack installed; start-up phase awaited; no QC, as ZI is subsequently (see the following SWO)
No | Visit | Complaint | RIND | NT Pro BNP value outside the reportable range. | NT Pro BNP value outside the reportable range for lfew samples of last few days. | Run the Biorad QC & got the desired result for NT Pro BNP,also run the same sample 3times in two different AQT90 machine & got the satisfactory result. | SWO-539853 | No fault found | No fault found | No fault found | No fault found |  | NT Pro BNP value outside the reportable range. | NT Pro BNP value outside the reportable range for lfew samples of last few days. | Run the Biorad QC & got the desired result for NT Pro BNP,also run the same sample 3times in two different AQT90 machine & got the satisfactory result.
No | Visit | Complaint | RPOL | Błąd QC na parametrach OXI | Dokończenie pracy z dnia wczorajszego. | Wymiana hemolizatora, complete OXI test, kalibracja tHb, testy serwisowe. | SWO-538725 | Defect | HW | Oxi module | Hemolyzer |  | QC error on OXI parameters | Completing the work of yesterday. | Replacement of Hemolyser, complete OXI test, tHb Calibration, service tests.
Yes | Visit | Complaint | RPOL | Upgrade oprogramowania | Sprawdzenie wersji oprogramowania | Upgrade oprogramowania do wersji 3.4MR1 | SWO-538726 | Re classified RMED | Same day UpSWO |  |  | SW update | Software Upgrade | Check the software version | Software Upgrade to version 3.4 MR1
No | Visit | Complaint | RPOL | Aparat nie przeczytał dwóch pakietów z rzędu | Przegląd logów i ocena wizualna. | Stwierdzono mechaniczne uszkodzenie płytki z chipem odczytującym pakiet. Wymiana tej płytki. Testy serwisowe. | SWO-539428 | Defect | HW | Cover and Chassis | PCB, solution pack |  | The camera has not read two packages in a row | Log review and visual assessment. | There was mechanical damage to the plate with the packet reading chip. Replacing this tile. Service tests.
No | Visit | Complaint | RLTD | The sensor cassette was replaced but shortly after it asked for the cassette to be replaced again and when it was taken out it was very hot to the touch. | System temperature error and overheating SC | Service dump and data logs collected for further investigation | SWO-539683 | Assisted Replacement | SC | SC | Temp | mail: send to MRM | The sensor cassette was replaced but shortly after it asked for the cassette to be replaced again and when it was taken out it was very hot to the touch. | System temperature error and overheating SC | Service dump and data logs collected for further investigation
No | Visit | Complaint | RHUN | Oldat pumpálási hibák, 04.26-án a készülék hibákat jelez, tisztításat kér, a hiba a folyadék érzékelőből fakad | - | készülék elszállítása Radiometer telephelyére szerviz és tesztelés céljából | SWO-539508 | Revisit required | analyzer removed | Analyzer removed | FT | LS | Solution for pumping defects, On 04.26, the device indicates defects, it asks for cleaning, the defect is from the liquid sensor | - | Equipment to Radiometer's site for service and testing purposes
No | Visit | Complaint | RSAF | All parameters were red on top | Backflash | Replaced sensor cassete and deleted all manual qc | SWO-539761 | Assisted Replacement | SC | SC | SC |  | All parameters were red on top | Backflash | Replaced sensor cassete and deleted all manual QC
No | Visit | Complaint | RAME | Clot detected in inlet. | Analyzer has clot detected in inlet, error code 1350. | Flushed entire system, replaced all tubing and parts from maintenance kit. Also replaced solution pack and sensor cassette from FSE stock. | SWO-539157 | Defect | HW | Basic Unit | PM kit | Clot | Clot detected in inlet. | Analyzer has clot detected in inlet, error code 1350. | Flushed entire system, replaced all tubing and parts from maintenance kit. Also replaced solution pack and sensor cassette from FSE stock.
No | Visit | Complaint | RCHN | lac定标失败 | 多张测试卡lac定标失败，发现接口板腐蚀。需要借件902-972一个 | 更换902-972备件后修复 by jesson | SWO-538729 | [5 empty] | lac calibration failed | Several test cards failed to calibrate and the interface plate was found to be corroded. Need to borrow 902-972 one | Repair by jesson after replacement of 902-972 spare parts
No | Visit | Complaint | RIND | Service Visit. | Need to change the settings according to the user. | Change the setting of the analyzer, add fiO2 in the sample information, successfully done, give the training and meet to Dr. Shah, successfully done, analyzer handed over to staff in well working condition. | SWO-539864 | [5 empty] | Service Visit. | Need to change the settings according to the user. | Change the setting of the analyzer, add fiO2 in the sample information, successfully done, give the training and meet to Dr. Shah, successfully done, analyzer handed over to staff in well working condition.
No | Visit | Complaint | RSCH | Rep. pH/pCO2 | Seit 23.04.2109 Probleme nach SC und LP Wechsel Kunde bereits 2x SC und 1x LP ersetzt Kontaminierung mit Luft pH $pCO2 $Hct Kunde hat am 25.04.2019 angerufen Am Abend um 19Uhr alles Parameter grün | => Nach Wartung separater Rapport Parameter reaktiviert, ohne Erfolg Neustart und SC nochmals initialisiert, ohne Erfolg LP ersetzt und SC mit Wasser gespült QA Zyklus i.o 2ter QA Zyklus i.o 3ter QA Zyklus i.o System i.o 2x SC und 2x LP werden gutgeschrieben Ausdruck und Gutschrift durch Techniker Eingesetztes LP wird verrechnet Kunde geschult wie Reklamationsdatei ausgefüllt werden muss | SWO-539527 | [5 empty] | Rep. pH/pCO2 | Since 23.04.2109 Problems after SC and LP change Customer already replaced 2x SC and 1x LP Contamination with air pH /pCO2 /Hct Customer called on 25.04.2019 In the evening at 7 pm all parameters green
No | Visit | Complaint | RIND | In Level 3 QC Na and pCO2 going out continuously. | Check System | Check system ,check calibration, done valve cleaning using cleaning solution internally, clean all the tubing internally ,done zeroing, reset SP , re-install SP and SC, The problem comes while installing SC after 4-5 try the SC installed, run QC level 3 with Different lot still problem persist. | SWO-539855 | [5 empty] | In Level 3 QC Na and pCO2 going out continuously. | Check System
No | Visit | Complaint | RCHN | 试剂包探测到漏电流，报错0964 | 试剂包探测到漏电流，报错0964，需要借件903-581一个 | 更换903-581后修复 by jesson | SWO-538601 | Defect | HW | Basic Unit | PCB, Sensor interface | 964 | Reagent pack detects leakage current, error 0964 | Reagent package detected leakage current, reported error 0964, need to borrow 903-581 a
No | Visit | Complaint | RCHN | 样本条码无法扫描 | 样本条码无法扫描 | 更换902-924后修复 | SWO-539582 | Defect | HW | Cover and Chassis | Barcode reader, display |  | Sample barcodes cannot be scanned | Sample barcodes cannot be scanned
No | Visit | Complaint | RDEU | 25.04.2019, Herr Kleine, Tel: 021714092175: schlechtes Druckbild | Druckbild nicht lesbar | Druckerkopf getauscht. ok. | SWO-539477 | Defect | HW | Printer Module | Printer Head |  | 25.04.2019, Mr. Kleine, Tel: 021714092175: bad print image | Print image not readable
No | Visit | Complaint | RDAN | Mogens Hald melder, at ABLen slår net sikringerne både sine egne og KBAs!! | Undersøgt. | Strømforsyning udskiftet, slutafprøvet. | SWO-539368 | Defect | HW | Power Supply | Power Supply |  | Mogens Hald reports that the ABLen beats the net fuses both his own and the KBAs!! | Examined.
No | Visit | Complaint | RAME | Analyzer not detecting any gas 2 pressure. | Gas 2 is always empty. | Exchanged the Gas 2 regulator and the error was not resolved. Need to replace the entire gas module. Will order replacement part and schedule a visit when the part arrives. | SWO-539658 | Revisit required | HW | Gas module | Valve |  | Analyzer not detecting any gas 2 pressure. | Gas 2 is always empty.
No | Visit | Complaint | RKOR | CORRUPTXML | Cleaning the neddle All position safe check system check, Optical measurement check | Cleaning the neddle All position safe check system check, Optical measurement check | SWO-540187 | No fault found | No fault found | No fault found | No fault found |  | CORRUPTXML | Cleaning the neddle All position safe check system check, Optical measurement check
No | Visit | Complaint | RLTD | Error codes 1258, 1259, and 1264. | Error codes 1258, 1259, and 1264. Customer had already ran system cleans and replaced solution pack. | Replaced needle wash unit, carried out needle calibration, carried out HCT cell calibration. High level CRP QC and ran a blank test | SWO-539722 | Defect | HW | Aspiration Module | Needle Wash Unit | 1258, 1259, 1264 | Error codes 1258, 1259, and 1264. | Error codes 1258, 1259, and 1264. Customer had already ran system cleans and replaced solution pack.
No | Visit | Complaint | RPOL | Zmiana zakresów raportowalnych | Zmiana zakresów raportowalnych | Analizator sprawny technicznie | SWO-539533 | Setup | SW | RMED SW | Reportable range |  | Changing the reportable Ranges | Changing the reportable Ranges
No | Visit | Complaint | RPOL | Zbyt wysoki zmętnienie widma w czasie pomiarów parametrów oksymetrii | Przegląd logów aparatu | Odbiałczanie, kalibracja tHb od serwisu, testy serwisowe. | SWO-539429 | Adjustment | HW | Oxi module | tHb Cal |  | Spectral opacity too high during measurements of the symmetry parameters | Overview of Camera Logs
No | Visit | Complaint | RAME | Analyzer is unable to pump solutions | Troubleshot to a possible solution pack or sensor cassette problem. After replacing both the sensor cassette and the solution pack, but there is no change. Using the fluid transport program the Solution Pack Check fails on every test, and the Clot Detect fails for a low Sensor Cassette Value and a High OXI module value. I flushed the analyzer several times, and saw no problems with the fluid flow, but the analyzer finally came in. | Replaced the sensor cassette, solution pack and flushed the analyzer. | SWO-540706 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway | Unable to Pump | Analyzer is unable to pump solutions | Troubleshot to a possible solution pack or sensor cassette problem. After replacing both the sensor cassette and the solution pack, but there is no change. Using the fluid transport program the Solution Pack Check fails on every test, and the Clot Detect fails for a low Sensor Cassette Value and a High OXI module value. I flushed the analyzer several times, and saw no problems with the fluid flow, but the analyzer finally came in.
No | Visit | Complaint | RKOR | Solutio pack detected error | . | Solution pack smart chip position adjust. | SWO-540020 | [5 empty] | Solutio pack detected error | .
Yes | Visit | Complaint | RIND | In Level 3 QC Na and pCO2 going out continuously. | Check system | Keep one stand by unit and tried QC Level3 but same problem occurred ,done zeroing and run QC also but problem as it is, this time problem comes only in Na. | SWO-542961 | Re classified RMED | Same day UpSWO |  |  | See SWO-542998 | In Level 3 QC Na and pCO2 going out continuously. | Check system
No | Visit | Complaint | RIND | In Level 3 QC Na and pCO2 going out continuously. | Check system | Check system and tried with a new lot cassate from Biosys still the problem is not resolved tried on standby analyser and ABG analyser ,problem faced in Level 3 Na | SWO-542998 | [5 empty] | In Level 3 QC Na and pCO2 going out continuously. | Check system
No | Visit | Complaint | RCHN | HCT值测量错误 | HCT值测量错误，借903-329一个 | 更换903-329后修复 by jesson | SWO-539369 | Defect | HW | Aspiration Module | Needle Wash Unit | 1223, 1258, 1259 | HCT value measurement error | HCT value measurement error, borrow 903-329 one
Yes | Visit | Complaint | RMAY | Attend breakdown service after work hour | 1. Check analyser status and data log 2. 0376 : Calibration drift out of range (Ca++) 3. 0379 : Calibration error (Cl-) | 1. Replace with pH electrode loan from other analyser (S $N:754R0959N0005) 2. Run 2 points calibration until stable 3. Cl value failed - 0379:Calibration error 4. Perform cleaning 5. Run calibration 6. Run QC | SWO-550324 | [5 empty] | Attend breakdown service after work hour | 1. Check analyser status and data log 2. 0376 : Calibration drift out of range (Ca++) 3. 0379 : Calibration error (Cl-)
No | Visit | Complaint | RKOR | Solutio pack replace error | . | Motor module, Sensor board check. | SWO-540019 | Revisit required | Consumables | Solution Pack | Solution Pack |  | Solutio pack replace error | .
No | Visit | Complaint | RCHN | 试剂包针脚断裂 | 试剂包针脚断裂。需要借件903-158和903-183各一个 |  | SWO-540065 | Defect | HW | Basic Unit | PCB, solution pack | CRM Corrective Action: reagent package pin break, replace 903-158 repair; parameter calibration has problems, replace 903-183 repair. By tim | Reagent pack pin break | The reagent pack pins are broken. One each of 903-158 and 903-183 is required
No | Visit | Complaint | RCHN | 进样口抬起时，机器重启 | 进样口抬起时，机器重启，借903-644,903-581,902-918各一个 | 更换903-644,903-581,902-918各一个备件后修复 | SWO-539873 | Defect | HW | Basic Unit | PCB, wet section |  | When the sample port is lifted, the machine restarts | When the sample is lifted, the machine restarts, borrowing 903-644, 903-581, 902-918 each
No | Visit | Complaint | RCHN | 检测到泄漏 | 检测到泄漏（未更换备件） | 更换进样扣支架后修复 by jesson | SWO-553450 | Defect | HW | Inlet module | Inlet gasket | 1184 | A leak has been detected | Leak detected (spare parts not replaced)
No | Visit | Complaint | RCHN | 提示重启分析仪，报错3264 | 提示重启分析仪，报错3264，借903-198一个 |  | SWO-540091 | Defect | HW | Analyzer Control Module | PCB, Main Board | CRM Corrective Action: Repaired by tim after replacing 903-198 | Prompt to restart analyzer, error 3264 | Prompt to restart analyzer, report error 3264, borrow 903-198 one
No | Visit | Complaint | RCHN | 试管夹断裂 | 试管夹断裂 | 更换903-584*2后修复by tim | SWO-540215 | Defect | HW | Inlet module | Tube holder |  | Test tube clip break | Test tube clip break
No | Visit | Complaint | RCHN | 无法识别处理包 | 无法识别处理包,借903-036一个。 | 更换903-036后修复 by jesson | SWO-539375 | Defect | HW | Cover and Chassis | PCB, solution pack | 1052 | The processing package is not recognized | The processing package is not recognized and 903-036 is borrowed.
No | Visit | Complaint | RCHN | 定标中断 | 用户在定标过程中，点击屏幕的按键时，仪器自动退出到window 桌面。借944-394（15754）一个 | 重启机器，重新定标后修复 by jesson | SWO-539307 | Defect | SW | RMED SW | RMED SW |  | Calibration interruption | During the calibration process, the instrument automatically exits to the window desktop when the button of the screen is clicked. Borrow 944-394 (15754) one
No | Visit | Complaint | RSAF | Restarting, disconnecting UPS and connecting direct to wall socket and still not switching on. | Checked UPS and power cable Checked power supply and CPU unit, fault suspected on Main board | Replaced faulty main board, software upgrade and uploaded configuration settings. LQC on all parameters done | SWO-540152 | Defect | HW | Analyzer Control Module | PCB, Main Board |  | Restarting, disconnecting UPS and connecting direct to wall socket and still not switching on. | Checked UPS and power cable Checked power supply and CPU unit, fault suspected on Main board
No | Visit | Complaint | RDEU | 24.04.2019, 12:19 Uhr ==> Anrufer Herr Connemann, Tel.: 015209322790 ==> Analysator hängt in der Fehlerbehebungsroutine 26.04.2019, 10:15 Uhr ==> Anrufer Herr Connemann, Tel.: 015209322790 ==> Fehler tritt häufig auf, SP muss oft erneuert werden, jetzt u.a. Fehler o964: Ableitstrom im Anschluss SP erkannt | Sichtkontrolle Ableitstrom erkannt | Div Test im Servicemenü nach bearbeiten des schwergängigen Quetsch-Ventils einwandfrei Neustart, SP und Dichtung wird gefordert pressure Test nicht bestanden, Ableitstrom... SP ersetzt, Gerät zerlegt, Hemolyzer Sichtkontrolle, Wartungskit verbaut, Test im Service wieder OK, Neustart, Kalibration, QC-Kontrolle | SWO-540456 | Defect | HW | Basic Unit | PM kit | Leak current | 24.04.2019, 12:19 o'clock ==> Caller Mr. Connemann, Tel.: 015209322790 ==> Analyzer hangs in the troubleshooting routine 26.04.2019, 10:15 a.m. ==> Caller Mr. Connemann, Tel.: 015209322790 ==> Error occurs frequently, SP often needs to be renewed, now among others error o964: discharge current detected in the following SP | Visual inspection Discharge current detected
Yes | Visit | Complaint | RESP | Cliente indica que no puede cambiar correctamente tipo de muestra. | Pantalla parece no responder correctamente | Comprobación del analizador. Se procede a entrar en el sistema y se procede a re calibrar pantalla. Se realizan pruebas y parece que funciona correctamente. | SWO-541095 | [5 empty] | Client indicates that it cannot correctly change sample type. | Screen seems not to respond correctly
No | Visit | Complaint | RESP | Não aceita SP e com erro de selector de fluxo | Erro de selector de fluxo após colocação do SP | Substituição do pack, ajuste do sensor de fluxo, estabilização e passagem do QC. | SWO-540573 | Assisted Replacement | SP | SP | Flow selector | 1177 | Does not accept SP and with Flow selector error | Flow selector error after SP placement
No | Visit | Complaint | RESP | O cabo de alimentação faz mal contacto e a máquina desliga-se e verificar a máquina porque deu problemas no fim de semana. | Verificação de mal contacto do cabo de alimentação com o equipamento. | Substituição do cabo de alimentação, estabilização, verificação do funcionamento e passagem do QC. | SWO-540646 | Defect | HW | Power Supply | Power Cable |  | The power cord makes contact poorly and the machine turns off and checks the machine because it has been in trouble over the weekend. | Check the power cord's contact with the equipment.
Yes | Visit | Complaint | RMAY | Oncall breakdown | Printer print 1 month patient result due to user wrongly pressed print icon. | -Check analyser data log - Delete pending printing document -Tested OK | SWO-544798 | [5 empty] | Oncall breakdown | Printer print 1 month patient result due to user wrongly pressed print icon.
No | Visit | Complaint | RKOR | Pressure Test Vacuum error | Gasket holder replace, line flush, cal control data check. | Gasket holder replace, line flush, cal control data check. | SWO-543280 | Defect | HW | Auto Inlet module | Inlet gasket | Vacuum error | Pressure Test Vacuum error | Gasket holder replace, line flush, cal control data check.
No | Visit | Complaint | RAME | Status code 2 | Status code 2 | Reloaded the software with blank database. Supplies were expired so I could not do any more. Bonnie will be here tomorrow and put supplies on | SWO-540893 | Defect | SW | RMED SW | database | Status code 2 | Status code 2 | Status code 2
No | Visit | Complaint | RKOR | Pressure Test Vacuum error | Gasket holder replace, line flush, cal control data check. | Gasket holder replace, line flush, cal control data check. | SWO-543279 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway | CRM: Not same hospital as the other on the same day. | Pressure Test Vacuum error | Gasket holder replace, line flush, cal control data check.
No | Visit | Complaint | RCHN | 外壳故障 | 外壳故障，借888-725一个 | 更换888-725后修复by tim | SWO-526933 | [5 empty] | Housing failure | Housing failure, borrow 888-725 one
No | Visit | Complaint | RCHN | PO2病人测量结果异常 | PO2病人测量结果异常，借903-334一个 |  | SWO-540254 | [5 empty] | Abnormal measurements in PO2 patients | PO2 patient measurement situ abnormal, borrow903-334 one
No | Visit | Complaint | RCHN | 仪器频繁自动关机 | 仪器频繁自动关机，借903-663一个 | 更换903-663后修复by tim | SWO-540044 | [5 empty] | The instrument shuts down automatically frequently | The instrument automatically shuts down frequently, borrows 903-663 one
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 无法读取试剂包芯片信息 | 更换试剂包后修复 | SWO-541071 | [5 empty] | Test card, reagent package failure | Unable to read reagent package chip information
No | Visit | Complaint | RCHN | AU板烧毁 | AU板烧毁，借903-198一个 |  | SWO-540239 | Defect | HW | Analyzer Control Module | PCB, Main Board |  | AU board burned | AU board burned, borrowed 903-198 one
No | Visit | Complaint | RMAY | Service breakdown | Qc error due to poor membrane condition | - Check analyzer data log - Clean chamber - Flush inlet - Perform protein removal - Remembrane Na ( user ) | SWO-543325 | User Error, FSE | Electrode | Membrane | Lifetime exceeded | RMAY: planned visit 1-3 times / week. Membrane exceeded lifetime (see e-mail from Azlan AHZAW "RE: SWO-543325 - comments on ABL800 upSWOs from RMAY") | Service breakdown | Qc error due to poor membrane condition
No | Visit | Complaint | RIND | Intervention required | Intervention required | Checked calibration log , system messages , activity log , found system keys expired , updated system keys , run calibration , calibration passed , run samples . | SWO-540431 | User Error, FSE | Demo key | Demo key | Demo key |  | Intervention required | Intervention required
No | Visit | Complaint | RAME | MetHb out high L1 and L2 | MetHb out high L1 and L2 QC. Lamp unit adjustment would not go over 2062. | Replaced OXI Module. Perform complete OXI alignment. Also added a Cal Verification report and Sample button to setup. | SWO-540687 | Defect | HW | Oxi module | Oxi module | JGIBB informed about Service Tip: 780902523 | MetHb out high L1 and L2 | MetHb out high L1 and L2 QC. Lamp unit adjustment would not go over 2062.
No | Visit | Complaint | RDEU | 25.04.2019 8:15 H. Funcke +496063501472: Signaltöne fehlen | Bei Service festgestellt, dass Signaltöne nicht erzeugt werden, trotz richtiger Einstellung. Analysator hat sich aufgehängt, musste über Tatstatur ein Neustart durchgeführt werden. Sicherung aller Daten und der Konfiguration. Im Haus unterwegs zu weiteren Service. | Erneute Einspielung des Betriebssystems, Windows 7, dadurch Festplatte formatiert. Anschließend Applikations Software eingespielt. Danach die Daten und Konfiguration mit Netzwerkverbindung eingespielt und neu konfiguriert auch die Konfiguration für das Land. Kalibrationen und Qualitätskontrolle in Ordnung. Anwender beobachten nun Verhalten des Analysators. | SWO-540678 | Defect | SW | RMED SW | Loudspeaker |  | 25.04.2019 8:15 H. Funcke +496063501472: Beeps are missing | When service determines that beeps are not generated, despite correct setting. Analyzer has hung up, had to perform a reboot via Tatstatur. Backing up all data and configuration. In the house on the way to further service.
No | Visit | Complaint | RDEU | 18.04.2019 11:50Uhr Frau Weidenhiller tel:+498431545478: Wegen einer Fehlermeldung wurde das Gerät neu gestartet und meldet jetzt 1052 und 1051 29.04.2019 10:15 Uhr Herr Mikolajek Medizintechnik tel:+498431545470: Fehler 1052 und 1051 ist wieder aufgetreten | erneut kein Zugriff auf Pufferabfallkassette - Fehler 3414 FTC105 | Solution Pack im Service ausgeworfen Safe Positions angefahren - ok mitgebrachten Solution Pack installiert Systemreinigung und -überprüfung durchgeführt LQC gemessen | SWO-540518 | Defect | Consumables | Solution Pack | Solution Pack | FTC 105, 1052 | 18.04.2019 11:50Pm Ms Weidenhiller tel:+498431545478: Due to an error message the device was restarted and now reports 1052 and 1051 29.04.2019 10:15 Mr Mikolajek Medical Technology tel:+498431545470: Error 1052 and 1051 has reappeared | again no access to buffer waste cartridge - Error 3414 FTC105
No | Visit | Complaint | RDEU | 29.04.2019; Fr. Jenner; Tel.: 071819399921; FTC_103 und gemessene Flüssigkeitsmenge oberhalb Warngrenze | Fehler 3412 FTC_103 Fehler 1067, 1223,1264,1265 HCT-Fehler | Zur weiteren Fehleranalyse Servicedump erstellt. P2 und P4 erneuert. Systemreinigung durchgeführt. 5 Blanktests durchgeführt-o.k. Vorhandene QC Tests durchgeführt 2.Servicedump erstellt. | SWO-541792 | Defect | HW | Wet section | Pump 2 & 4 | 3412, FTC 103 | 29.04.2019; Ms. Jenner; Tel.: 071819399921; FTC_103 and measured liquid quantity above warning limit | Fehler 3412 FTC_103 Fehler 1067, 1223,1264,1265 HCT-Fehler
No | Visit | Complaint | RPOL | Brak możliwości odczytu kodu kreskowego z butelki z odczynnikiem | Weryfikacja statusu analizatora oraz funkcjonowania czytnika kodów | Wymiana czytnika, konfiguracja kontroli, testowanie. Wykonanie badania kontrolnego. | SWO-540836 | Defect | HW | Basic Unit | Barcode Reader |  | Unable to read barcode from reagent bottle | Verification of Analyzer status and code reader functioning
No | Visit | Complaint | RDEU | 30.04.2019; 9:30 Uhr; Frau Ridacco; Tel.: 089476036; ABL pO2 und tHb Fehler | Kundin stand nicht am Gerät. Sie sagte nur dass wieder pO2 und dieses mal auch die tHb Werte in der Paramterleiste rot sind. Der pO2 Fehler war auch schon am 26.4. Kundin hat SC und LP bereits erneuert. Kein Erfolg Sonst konnte sie keine Auskünfte geben. | Kundin wünscht Serviceeinsatz. Kein Material mehr vorhanden!!!! | SWO-541125 | Revisit required | Fluidic pathway | Fluidic pathway | Fluidic pathway | pO2/tHb Cal | 30.04.2019; 9:30 Uhr; Frau Ridacco; Tel.: 089476036; ABL pO2 und tHb Fehler | Customer was not on the device. She just said that again pO2 and this time the tHb values in the paramter bar are red. The pO2 error was also already on 26.4. Kundin has already renewed SC and LP. No success Otherwise, it could not provide any information.
No | Visit | Complaint | RAME | Display Damaged Contact Lori Black. | Ordered touchscreen unit for overnight. Picked up touchscreen in Roseburg and touchscreen was not in box. Arrived onsite and found the touchscreen housing broken. Will order an new touchscreen unit and replace. Analyzer working. | Order a new touchscreen unit. | SWO-541536 | Defect | HW | Display | Touch Screen |  | Display Damaged Contact Lori Black. | Ordered touchscreen unit for overnight. Picked up touchscreen in Roseburg and touchscreen was not in box. Arrived onsite and found the touchscreen housing broken. Will order an new touchscreen unit and replace. Analyzer working.
No | Visit | Complaint | RLTD | Battery error | Battery error | Reset jumper connection all okay | SWO-541435 | Adjustment | HW | Battery module | Battery module | 0% | Battery error | Battery error
No | Visit | Complaint | RDEU | pCO2 ist seit der Geräte-Installation im Level C an der oberen Grenze oder zu hoch | siehe SWO | neues Verbrauchsmaterial installiert --> Neustart ( SP ( HJ-01 ) $SC (176 ) Dateiexport und Servicedump | SWO-541355 | Assisted Replacement | SC | SC | pCO2 | QC C high | pCO2 has been at the upper limit or too high since the device was installed in Level C. | see SWO
No | Visit | Complaint | RFRA | Impossible d'ajouter des utilisateurs |  | Récupération de la base de données et envoi par mail chez Advanced. | SWO-541488 | [5 empty] | Can't add users
No | Visit | Complaint | RCHN | po2,pco2结果有异常 | po2,pco2结果有异常，借903-334一个 | 更换903-334后修复by tim | SWO-540319 | [5 empty] | po2, pco2 results are abnormal | po2, pco2 results have an exception, borrow 903-334 one
No | Visit | Complaint | RIND | In samples electrolytes results are showing very high | Check analyzer status in which calibration is OK,but in patient result electrolytes results are showing * or very high values which are not acceptable. | Check and clean waste bin and waste line.Check zeroing which is not OK,reset zero,upgrade the software,reinstall SP & SC and check for calibration which is OK.Run RL which have satisfactory result for electrolytes and problem has been resolved but under observation. | SWO-543294 | [5 empty] | In samples electrolytes results are showing very high | Check analyzer status in which calibration is OK,but in patient result electrolytes results are showing * or very high values which are not acceptable.
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | pco2定标失败1316 | 更换测试卡后修复 | SWO-543191 | Assisted Replacement | SC | SC | pCO2 | 1316 | Test card, reagent package failure | pco2 calibration failure 1316
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 流量选择器错误 0978 | 更换测试卡后修复 | SWO-541034 | [5 empty] | Test card, reagent package failure | Traffic Selector Error 0978
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | NA,LAC,GLU定标不过0378,0377,0375,1319 | 更换一张测试卡和一个试剂包后修复 | SWO-543500 | [5 empty] | Test card, reagent package failure | NA, LAC, GLU calibration but 0378,0377,0375,1319
No | Visit | Complaint | RCHN | 仪器不断重启，报错1999.12 | 仪器不断重启，报错1999.12,需要借件 903-030 | 更换903-030后修复by tim | SWO-540330 | Defect | HW | Storage Module | PCB, cartridge lock |  | The instrument kept restarting, and the error was 1999.12. | The instrument is constantly restarted, error 1999.12, need to borrow 903-030
No | Visit | Complaint | RDEU | Na und Ca instabil | Na und Ca sind sporadisch bei der Kalibration instabil | Fehlersuche. Ref. Elektrode getauscht | SWO-541398 | Assisted Replacement | Electrode | Electrode | reference | CAPA-733 | Na and Ca unstable | Na and Ca are sporadic unstable during calibration
No | Visit | Complaint | RDEU | Ausdruck unleserlich | Ausdruck unleserlich | Druckereinheit ausgetauscht. Neustart. Kal und QC | SWO-541397 | Defect | HW | Printer Module | Printer Module |  | Expression illegible | Expression illegible
No | Visit | Complaint | RSCH | Rep CREA und Na | CREA immer Grenzwertig und ausser Bereich auch mit Lot R0859 Na immer hoher Drift | Referenz Elektrode und Membrane ausgetauscht Met 2 PCB ausgetauscht Met 2 Messkammer ausgetauscht Funktionskontrolle danach i.o | SWO-541302 | Assisted Replacement | Electrode | Electrode | reference | CAPA-733; failing Reference 58-05 | Rep CREA and Na | CREA always limitable and out range also with lot R0859 Na always high drift
No | Visit | Complaint | RAME | Barcode issue - not reading the ampules | Ampoules not being scanned. Laser engine is suspect. Mirror in Laser Engine is not very shinny. | Replaced Laser Engine. Performed alignment in service. | SWO-541404 | Defect | HW | Autocheck module | Barcode reader |  | Barcode issue - not reading the ampules | Ampoules not being scanned. Laser engine is suspect. Mirror in Laser Engine is not very shinny.
No | Visit | Complaint | RAME | Autocheck issues | Although cleaning worked from previous visit, a new Laser Engine needs to be installed. I didn’t have one when I was here last visit. Before running QC, getting inlet flap removed. | Replaced Laser Engine. Performed alignment. Turns out inlet cable that plugs into Inlet PC was almost unplugged. Pushed cable all the way in. | SWO-541440 | Defect | HW | Autocheck module | Barcode reader |  | Autocheck issues | Although cleaning worked from previous visit, a new Laser Engine needs to be installed. I didn’t have one when I was here last visit. Before running QC, getting inlet flap removed.
No | Visit | Complaint | RIND | Service Visit. QC analysis is coming regularly out of range. | Need to check the analyzer. | Check the calibration and patient result found ok, QC out regularly, need to replace the reference and sodium membrane, clear the doubt regarding the non-radiometer QC Set up. | SWO-542932 | Revisit required | Electrode | Membrane | Na | Non-radiometer QC | Service Visit. QC analysis is coming regularly out of range. | Need to check the analyzer.
No | Visit | Complaint | RAME | Customer called stating that analyzer is not running the qc automatically | Found that the analyzer would open the autocheck it would go vertical and then just rattle. | Check all the connections on the autocheck to try and get it to work needs a new 14 pole cable it was to late to order it today will order tomorrow morning. | SWO-541865 | Revisit required | HW | Autocheck module | FPC Cable |  | Customer called stating that analyzer is not running the qc automatically | Found that the analyzer would open the autocheck it would go vertical and then just rattle.
No | Visit | Complaint | RDAN | Efter udskiftning af strømforsyning vil ABLen ikke printe. | Afprøvet og undersøgt. | Sikringsholder til printer +24v loddet fast, slutafprøvet. | SWO-541191 | Defect | HW | Basic Unit | Fuse box |  | After replacing the power supply, the ABLen will not print. | Tested and investigated.
No | Visit | Complaint | RAME | Oxi hardware error | 269 hemolyzer phase regulation error | Previous call checked connection and powered off analyzer and error went away for a couple weeks. Today I replaced the Hemolyzer when problem came back. Performed lamp adjust, complete oximetry test, and service THB. All passed except level 3 QC resulted in 581 errors with new Hemolyzer. Ordering new hemolyzer and optic cable. | SWO-540818 | Defect | HW | Oxi module | Hemolyzer | 269 | Oxi hardware error | 269 hemolyzer phase regulation error
No | Visit | Complaint | RAME | Analyzer keeps having a 663 leak error. | Checked the data logs and verified the problem. | Performed a couple of zero current cals to solve the problem. | SWO-541454 | Adjustment | HW | pH/BG Module | zero current | 663 leak error | Analyzer keeps having a 663 leak error. | Checked the data logs and verified the problem.
No | Visit | Complaint | RAME | Leak detected after replacing solution a and all electrodes. | Checked the data logs and found 508 rinse error. | Ran a rinse cycle to solve the problem. | SWO-541507 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway | 508 | Leak detected after replacing solution a and all electrodes. | Checked the data logs and found 508 rinse error.
No | Visit | Complaint | RSAF | HCT measurements errors on patient samples | Buffer solution overflows in Needle wash block Errors 1258 1259 1264 and 1265 | Replaced pump4, cleaned filter, and replaced tubing from wash block to Solution Pack. Patient samples ran but hct errors not resolved. Needle wash block to be replaced. | SWO-540898 | Revisit required | HW | Aspiration Module | Needle Wash Unit | Blocked | HCT measurements errors on patient samples | Buffer solution overflows in Needle wash block Errors 1258 1259 1264 and 1265
No | Visit | Complaint | RKOR | Solution pack lock error | . | Service mode check. | SWO-543344 | Defect | Consumables | Solution Pack | Solution Pack |  | Solution pack lock error | .
No | Visit | Complaint | RPOL | Skrzywiona igła | Uszkodzenie igły | Wymiana igły oraz uszczelki wlotu | SWO-541369 | Defect | HW | Auto Inlet module | Probe bend |  | Winted Needle | Needle damage
No | Visit | Complaint | RPOL | Wykryto prąd upływu w relacji do pakietu roztworów. Wymiana pakietu oraz kasety sensorowej nie rozwiązała probemu. | Testy serwisowe. Sprawdzenie analizatora pod kątem wycieków. | Wymiana wszsytkich wężyków. Wymiana płyty Sensor Interface PCB. Testy serwisowe ok. Wartoś prądu upływu w normie. | SWO-546454 | Defect | HW | Basic Unit | PM kit | 964 | Leakage current in relation to the solution package has been detected. Replacing the package and the sensor cassette did not solve the probem. | Service tests. Check the analyser for leaks.
Yes | Visit | Complaint | RMAY | Oncall breakdown | Problem occurred due to clot in inlet probe . Inlet probe cannot be flushed therefore is faulty . | -Check analyser data log -Flush from left sensor bay to waste (OK) -Flush from right sensor bay to gasket (FAIL) -Dismantle inlet module -Flush without inlet module(OK) -Swap inlet probe -Assemble inlet module -Run fluid transport test ( OK) | SWO-544803 | [5 empty] | Oncall breakdown | Problem occurred due to clot in inlet probe . Inlet probe cannot be flushed therefore is faulty .
No | Visit | Complaint | RAME | Analyzer is not starting up, showing a Roma error. | Upon arrival, I found the analyzer software is not loading and, the screeen is showing a Roma loader error #7. | Restarted the analyzer in windows explorer and, saved all the roma, setup files and, database files. Then I reinstalled windows 7 and, the analyzer software. I reloaded the analyzer roma, setup and, the database files but, upon restarting the analyzer the software did not startup again, presenting the same roma loader error #7. I determined that the files were corrupt and, they are the root cause that prevented the analyzer software from starting up. I reinstalled the analyzer software without the corrupt files and, now the analyzer software loaded properly at startup. I copied the setup files from a nearby ABL90FlexPlus and, loaded it on this analyzer. | SWO-541620 | Defect | SW | RMED SW | Roma.ini | corrupt files | Analyzer is not starting up, showing a Roma error. | Upon arrival, I found the analyzer software is not loading and, the screeen is showing a Roma loader error #7.
No | Visit | Complaint | RAME | LS errors unable to pump solutions | 1343 unable to pump solutions system check shows solution pack errors - replaced solution pack no change in state replaced inlet gasket no change ran service menu check on solution pack all passing but when the system goes into a full check it fails back to the unable to pump solutions flushed the analyzer no issues analyzer has good fluid flow replaced the sensor cassette with a known working one from another unit - analyzer came up with no issues - old sensor was put in the other analzyer and no isses reported on the other unit went through the change solution pack tasks again to varify analyer comes up without error | replaced sensor cassette | SWO-542429 | Assisted Replacement | SC | SC | FT | Unable to Pump | LS errors unable to pump solutions | 1343 unable to pump solutions system check shows solution pack errors - replaced solution pack no change in state replaced inlet gasket no change ran service menu check on solution pack all passing but when the system goes into a full check it fails back to the unable to pump solutions flushed the analyzer no issues analyzer has good fluid flow replaced the sensor cassette with a known working one from another unit - analyzer came up with no issues - old sensor was put in the other analzyer and no isses reported on the other unit went through the change solution pack tasks again to varify analyer comes up without error
No | Visit | Complaint | RCHN | 机器断电，定标数据丢失 | 机器断电，定标数据丢失，借903-198,910-373各一个 | 更换903-198后修复by tim | SWO-541029 | Defect | HW | Analyzer Control Module | PCB, Main Board |  | Machine power down, calibration data loss | Machine power loss, calibration data loss, borrowing 903-198, 910-373 each one
No | Visit | Complaint | RMAY | 0510:EL/MET Liquid sensor did not receive Rinse from Above. | Check analyser status, Check valve and pump, Check tubing and pump. | Troubleshoot the analyser. Check and restart the analyzer. Run calibration - Passed Run QC -Passed. | SWO-554823 | Adjustment | HW | El/Met Module | LS | 510 | 0510:EL/MET Liquid sensor did not receive Rinse from Above. | Check analyser status, Check valve and pump, Check tubing and pump.
No | Visit | Complaint | RCHN | 定标失败 | PO2定标错误 0476 | 更换氧电极膜后修复 | SWO-552772 | Assisted Replacement | Electrode | Membrane | pO2 | 509; 476 | Bid failed | PO2 calibration error 0476
No | Visit | Complaint | RAME | Oxi hardware error | Return trip for 581 errors on level 3 QC with new Hemolyzer replacement | Replaced Hemolyzer again. Ran complete oximetry test, cals, and service THB. All cals and QC’s passed. | SWO-541894 | Defect | HW | Oxi module | Hemolyzer | 581: communicated: if error happens again: exchange the whole Oxi module. | Oxi hardware error | Return trip for 581 errors on level 3 QC with new Hemolyzer replacement
No | Visit | Complaint | RDEU | 30.04.2019; 9:30 Uhr; Frau Ridacco; Tel.: 089476036; ABL pO2 und tHb Fehler | ABL pO2 und tHb Fehler | BGA manuell gespült, Enlassdichtung erneuert, Lösungspack und Sensor Kassette erneuert, Kalibrationen und QC's vorgenommen. | SWO-542633 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway | pO2/tHb Cal | 30.04.2019; 9:30 Uhr; Frau Ridacco; Tel.: 089476036; ABL pO2 und tHb Fehler | ABL pO2 und tHb Fehler
No | Visit | Complaint | RDEU | PCo2 QC Fehler Level A leicht oberhalb der Grenze zeitweise | PCo2 QC Level A zeitweise oberhalb der Grenze, nach Wiederholung der QC wieder i.O. | Fehler am 30.04 aufgetreten, nach SC Wechsel am 1.05. QC innerhalb des Bereichs, Daten gesichert und Service dump gezogen, QC A und QC B von der Duchführungszeit getauscht, SC ersetzt siehe SWO- 542637 Software update | SWO-542638 | Defect | SW | RMED SW | pCO2 | Proposal for this one: Defect - SW - RMED SW - pCO2 (as the SW is installed this day and trigger is pCO2 failure.) CRM SW SWO-542637: Alle Daten gesichert, Software V3.4MR1 installiert, Grund der Maßnahme mit Anwender durchgesprochen, Kal und QC durchgeführt, Fehler bei Ca Empf nach update zu niedrig, Einlassdichtung und Enlassdichtung mit Halter ersetzt wegen aufgetretenem Leckage Fehler, Kal wiederholt- Ca Emp bleibt zu niedrig - SC ersetzt, Kal und QC durchgeführt, Ca i.O. QC und Kal La und Gluc noch fehlerbehaftet - SC muss noch einlaufen Anwender informiert, bei bleibenden Problemen werden sie die SC erneut ersetzen | PCo2 QC Error Level A slightly above the limit intermittently | PCo2 QC Level A temporarily above the limit, after repetition of the QC again i.O.
No | Visit | Complaint | RAME | Noise from back of analyzer | Ticking noise from speaker | Reseated speaker and noise disappeared | SWO-542291 | Adjustment | HW | Basic Unit | Speaker |  | Noise from back of analyzer | Ticking noise from speaker
Yes | Visit | Complaint | RAME | Netop failure 04/29/2019 JMG- | Retrieve QC results from previous service call | Attached QC results so case can be resolved | SWO-542304 | Re classified RMED | planned SWO |  |  | This SWO: only get QC results. SWO with failure: SWO-523923 | Netop failure 04/29/2019 JMG- | Retrieve QC results from previous service call
Yes | Visit | Complaint | RDEU | pCO2 ist seit der Geräte-Installation im Level C an der oberen Grenze oder zu hoch | QC-Daten geprüft | ABL90 beim Eintreffen fehlerfrei QC-Daten geprüft --> i.O alle QC-Level noch einmal durchgeführt --> pCO2-Werte sind stabil ABL90 wird bis zum 06.05. 2019 weiter getestet | SWO-542074 | Re classified RMED | Courtesy |  |  |  | pCO2 has been at the upper limit or too high since the device was installed in Level C. | QC data checked
No | Visit | Complaint | RIND | In Level 3 QC Na and pCO2 going out continuously. | Check system | Received new control from Ahmedabad having same lot no 104 , run on 3 analyser out of 3 two are ABG and one is lyte the control run on lytes gives result but on ABG it goes out, in this complaint tried all different lot of SC, SP, Analysers, QCs bt still problem remains as it is. | SWO-543008 | [5 empty] | In Level 3 QC Na and pCO2 going out continuously. | Check system
Yes | 3 | Y | RCZE | Blank |  |  | D6108C516C39DEDA | Re classified RMED | planned SWO |  |  |  | Blank
No | Visit | Complaint | RESP | Las medidas comparadas con el otro TCM difieren mucho. | La configuración de los TCM son distintas. | Poner la configuración con los mismos parámetros y probar. | SWO-542671 | [5 empty] | The measures compared to the other TCM differ greatly. | The configuration of MTCs are different.
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 多参数定标失败 | 更换测试卡后修复 | SWO-552782 | Assisted Replacement | SC | SC | Multi |  | Test card, reagent package failure | Multi-parameter calibration failed
No | Visit | Complaint | RCHN | 测试卡条码阅读器识别不了测试卡 | 测试卡条码阅读器识别不了测试卡，借910-304一个（未更换备件） | 调整条码扫码器角度后修复by tim | SWO-541697 | Adjustment | HW | Storage Module | Barcode reader, storage wheel |  | Test card barcode reader does not recognize the test card | Test card barcode reader does not recognize the test card, borrow 910-304 one (spare parts not replaced)
No | Visit | Complaint | RIND | SP memory chip corrupt. | SP memory chip corrupt | Clean the Solution pack chip ,try to install it but still it is showing corrupt,need to replace solution pack. | SWO-543138 | [5 empty] | SP memory chip corrupt. | SP memory chip corrupt
Yes | Visit | Complaint | RMAY | Breakdown services | Result patient for po2 unstable due to protein in chamber | - Check analyzer data log - Clean chamber with hypochlorite - Flush inlet - Perform protein removal | SWO-544810 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway | pO2, 476; check age (mail) | Breakdown services | Result patient for po2 unstable due to protein in chamber
No | Visit | Complaint | RMAY | Services breakdown | QC error due to clot in chamber | - Check analyzer data log - Clean chamber - Flush inlet - Perform protein removal | SWO-544817 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway | Clot, Na | Services breakdown | QC error due to clot in chamber
No | Visit | Complaint | RIND | Na calibration sensitivity is low. | Need to replace the Na membrane . | Replaced the Na membrane and run the instrument its working good. | SWO-542297 | User Error, FSE | Electrode | Membrane | Lifetime exceeded | see mail from PRARA with title "RE: RIND - ABL8XX - upSWO questions" | Na calibration sensitivity is low. | Need to replace the Na membrane .
No | Visit | Complaint | RAME | Customer called stating that analyzer is not running the qc automatically | Found that the qc does not spin when the housing goes vertically. | Replaced the 14 pole cable from board to housing to correct the issue. | SWO-546301 | Defect | HW | Autocheck module | FPC Cable |  | Customer called stating that analyzer is not running the qc automatically | Found that the qc does not spin when the housing goes vertically.
No | Visit | Complaint | RIND | Crea Calibration is not stable. | Found the clot in chamber and air bubbles in membrane then clean the chamber and removed the air bubbles then run the instrument its working good. | Found the clot in chamber and air bubbles in membrane then clean the chamber and removed the air bubbles then run the instrument its working good. | SWO-542292 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway | Clot | Crea Calibration is not stable. | Found the clot in chamber and air bubbles in membrane then clean the chamber and removed the air bubbles then run the instrument its working good.
No | Visit | Complaint | RIND | Startup failure | Check system | During startup analyser giving error so switch of the analyser clean the dust internally and from incubation wheel then again given startup now analyser working ok. | SWO-546995 | Cleaned | HW | Incubator Module | Incubator Wheel |  | Startup failure | Check system
No | Visit | Complaint | RDEU | 02.05.2019 11:10 Fr. Erbe +496661812394: QC-Fehler PCT beide Level | PCT-QC in beiden Leveln am unteren Bereich. | PCT Level1 LOT 15292 und Level 2 LOT 15015 am unteren Bereich. Neues LOT wurde letzte Woche zm ersten mal gemessen, davor waren die Werte mittig.Anwender bestellen sich neues Material und melden sich ggf. Gerät überprüft. | SWO-542765 | Defect | Consumables | QC | QC | Low PCT | 02.05.2019 11:10 Fr. Erbe +496661812394: QC-Fehler PCT beide Level | PCT-QC in both levels at the bottom.
No | Visit | Complaint | RPOL | W kontroli jakośc oraz na próbkach pacjentów pojawia się komunkat: Niezgodność widma OXI (0581) . Błędy przy pomiarach ctHb, Bil i Crea/Urea. | Testy modułu OXI. Błędne wyniki testu Spectrometer noise (37 mV). Diagnostyka z użyciem modułu OXI oraz płyty Wet Section PCB. Stwierdzono problemy z płytą Wet Section. | Instalacja zastępczej płyty Wet Section PCB. Testy i kalibracje serwisowe. | SWO-543306 | Defect | HW | Wet section | PCB, wet section | 581 | In quality control and patient samples, there is a komunkat: the Spectrum mismatch OXI (0581). Errors when measuring ctHb, Bil and Crea/Urea. | Tests of the OXI module. Incorrect Spectrometer noise (37 mV) test results. Diagnosis using OXI module and Wet Section PCB. There were problems with the Wet Section plate.
No | Visit | Complaint | RESP | AJUSTE INLET. | OBSTACULO EN INLET | NO CIERRA INLET PORQUE DETECTA UN OBSTACULO, SIENDO NECESARIO DESMONTAJE Y LIMPIEZA DE LA MISMA ASI COMO AJUSTE POR SERVICIO, UNA VEZ FINALIZADO DE EFECTUAN PRUEBAS DE FUNCIONAMIENTO Y CONTROLES DE CALIDAD. | SWO-542780 | Adjustment | HW | Auto Inlet module | Auto Inlet module | Obstacle | ADJUSTMENT INLET. | OBSTASC IN INLET
No | Visit | Complaint | RAME | Customer called with calibration and internal QC errors. 1002 PO2 dark data is out of range | Calibration and QC errors - 1002 PO2 dark data is out of range Customer replaced the SC for troubleshooting with same results. Performed the PO2 test in service to find the Dark value at 436.59 (10.00 - 200.00) | I will order parts and return when available. | SWO-543029 | Revisit required | HW | Basic Unit | PCB, Sensor interface | pO2 dark data | Customer called with calibration and internal QC errors. 1002 PO2 dark data is out of range | Calibration and QC errors - 1002 PO2 dark data is out of range Customer replaced the SC for troubleshooting with same results. Performed the PO2 test in service to find the Dark value at 436.59 (10.00 - 200.00)
No | Visit | Complaint | RIND | pO2 is not calibrated properly | Check system | Check system, check calibration data, remove SC , flush SC with Dwater, Done zeroing of the analyser, again re-install SC., Done calibration | SWO-546996 | [5 empty] | pO2 is not calibrated properly | Check system
No | Visit | Complaint | RKOR | Temperature error, cal fail | sensor cassette check,pin status check, flush check, line clean, solution pack check, sensor value check, operation check. | sensor cassette check,pin status check, flush check, line clean, solution pack check, sensor value check, operation check. | SWO-543288 | [5 empty] | Temperature error, cal fail | sensor cassette check,pin status check, flush check, line clean, solution pack check, sensor value check, operation check.
No | Visit | Complaint | RKOR | Calibration error | . | Liquid flow line clean. SC replace. | SWO-543346 | [5 empty] | Calibration error | .
Yes | Visit | Complaint | RCHN | 定标数据丢失 | 定标数据丢失，借944-212（15457）、944-214（15428）、944-216（15427）缺货、944-258（15260）、944-220（15486）各一个 | 重装系统，重新定标后D-二聚体依然定标失败，加借一张定标卡by tim | SWO-541998 | [5 empty] | Target data loss | Bid data loss, borrowing 944-212 (15457), 944-214 (15428), 944-216 (15427) out of stock, 944-258 (15260), 944-220 (15486) each
Yes | Visit | Complaint | RCHN | 定标数据丢失 | 定标数据丢失，借944-220（15486）一张 | 重新定标后修复by tim | SWO-543134 | [5 empty] | Target data loss | Bid data loss, borrow 944-220 (15486) one
No | Visit | Complaint | RIND | SP memory chip corrupt. | Solution pack need to replace. | Replace the solution pack with the new one,Air detection calibration got failed,pull the solutions from each chamber from the solution pack,it was coming properly.Then gain try to install & air detection calibration failed again,then flush the waste tube & solution tube,then again install the solution pack & this time it is working properly. | SWO-543139 | [5 empty] | SP memory chip corrupt. | Solution pack need to replace.
Yes | Visit | Complaint | RIND | Printer is not working | Check printer which is creating noise when commands for print but not printing paper. | Check printer head in which one gear is damaged and need to replace printer head.Replace new printer head and check for printing,now printer is working fine.Clean al the chambers,run protein removal and 2-point calibration which is OK and instrument working fine. | SWO-543176 | Defect | HW | Printer Module | Printer Head |  | Printer is not working | Check printer which is creating noise when commands for print but not printing paper.
No | Visit | Complaint | RIND | Chloride range is too low in QC | Checked and found that Chloride level is low in QC and Lactate membrane replacement Schedule as arrived. | Replaced Chloride and Lactate membrane and done an Successful 2 point Calibration, now machine works good. | SWO-543057 | User Error, FSE | Electrode | Membrane | Lifetime exceeded | see mail from PRARA with title "RE: RIND - ABL8XX - upSWO questions" | Chloride range is too low in QC | Checked and found that Chloride level is low in QC and Lactate membrane replacement Schedule as arrived.
No | Visit | Complaint | RIND | Intervention required | Intervention required error | Checked the analyser it was giving intervention required error so checke d for solution in that cal 2 was empty changed it after that cleaned measuring chamber then given decontamination then calibrated the analyser to ready mode | SWO-543119 | Assisted Replacement | Consumables | Solution | Cal2 | Empty | Intervention required | Intervention required error
No | Visit | Complaint | RIND | Service Call. | Need to clean the vision barcode. | Open the front panel as well as back panel of the analyzer, remove and clean the dust from its vision barcode reader, successfully done, analyzer handed over to staff in well working condition. | SWO-543002 | Cleaned | HW | Storage Module | Barcode reader, storage wheel |  | Service Call. | Need to clean the vision barcode.
No | Visit | Complaint | RKOR | pressure test error, Oxi value error | gasket holder replace, line flush, connter gaske replace, probe clean adjust, solution pack, Oxi, inlet value check | gasket holder replace, line flush, connter gaske replace, probe clean adjust, solution pack, Oxi, inlet value check | SWO-543292 | Defect | HW | Auto Inlet module | Inlet gasket |  | pressure test error, Oxi value error | gasket holder replace, line flush, connter gaske replace, probe clean adjust, solution pack, Oxi, inlet value check
Yes | Visit | Complaint | RMAY | Breakdown | Problem occurred due to clot in inlet | -Check analyser data log - Perform manual flush -Run fluid transport test ( OK) | SWO-544819 | [5 empty] | Breakdown | Problem occurred due to clot in inlet
Yes | Visit | Complaint | RCHN | po2定标失败，报错0588 0375 0210 | po2定标失败，报错0588 0375 0210.需要借件903-581一个 |  | SWO-543220 | [5 empty] | po2 bid failure, error 0588 0375 0210 | po2 bid failure, error 0588 0375 0210. Need to borrow 903-581 one
No | Visit | Complaint | RCHN | TNI、MYO质控低于统计范围的值 | TNI、MYO质控低于统计范围的值，借903-364、903-702、944-216（15446）、944-214（15229）缺货、944-258（15732）缺货、944-212（15677、15828）、944-268（15802）944-394（15753）各一张 | 更换903-364和重新定标后发现新光单报错1024，需要重新借件，固此单先关单by Hunter | SWO-542743 | Revisit required | Consumables | QC | QC | stored improperly (see e-mail…) | TNI, MYO QUALITY control is below the statistical range of the value | TNI, MYO QC is lower than the value of the statistical range, borrowed 903-364, 903-702, 944-216 (15446), 944-214 (15229) out of stock, 944-258 (15732), 944-212 (15677, 15828), 944-268 (15802)944-394 (15753) each
No | Visit | Complaint | RCHN | 屏幕花屏 | 屏幕花屏，借636-571,636-539各一个 |  | SWO-543270 | Defect | HW | Incubator Module | sensor, cup | CRM Corrective Action: Replace the 636-571 screen 636-539 after repairing the tray cup | Screen flower screen | Screen flower screen, borrow 636-571, 636-539 each
No | Visit | Complaint | RCHN | 进样门关不上 | 进样门关不上,需要借件886-638和886-639各一个 |  | SWO-543239 | Adjustment | HW | Cover and Chassis | Door, Inlet | CRM Corrective Action: Use ab glue to bond the black track to the worn part, then fix it by Hunter | The entrance door doesn't close. | The sample door is not closed, and it is necessary to borrow one each of 886-638 and 886-639
No | Visit | Complaint | RCHN | 定标分析物检测外漏 | 定标分析物检测外漏，借944-394（15887*2） |  | SWO-543201 | Assisted Replacement | Consumables | Test Kit | Cartridge, cal | 1024; CRM Corrective Action: Repair by Hunter after recalibration | Calibration analyte detection of leakage | Calibration anagency detects leakage, borrowed 944-394 (15887 x2)
No | Visit | Complaint | RCHN | 老师反映CTHB结果不 | 准老师反映CTHB结果不准 | 重新比对结果后修复-by Hunter | SWO-543356 | [5 empty] | Teacher reflects CTHB results not | The teacher-to-be reflected that CTHB results were not allowed
No | Visit | Complaint | RCHN | 无法扫描试剂 | 无法扫描试剂，借902-822一个 |  | SWO-543267 | Defect | HW | Basic Unit | Barcode Reader |  | Unable to scan reagents | Unable to scan reagents, borrow 902-822 one
No | Visit | Complaint | RCRO | Inlet gasket | Greška brtve | Zamjena brtve | SWO-548038 | Defect | HW | Inlet module | Inlet gasket |  | Inlet gasket | GASKET Error
Yes | Visit | Complaint | RMAY | Services breakdown | Clot found in the inner layer of inlet gasket | - Check analyzer data log - Perform manual flush - Check reagent - Swap inlet gasket (used) - Run fluid transport test ( OK) - Run calibration | SWO-544836 | [5 empty] | Services breakdown | Clot found in the inner layer of inlet gasket
No | Visit | Complaint | RAME | Customer called with calibration and internal QC errors. 1002 PO2 dark data is out of range | PO2 failing calibrations - PO2 dark data out of range Sensor Interface Serial# old-903-213R2008N0466 Monitor Serial# - Old - 910-353P1843N9968 | Replaced the Sensor Interface PCB for PO2 dark data error. PO2 dark value dropped within range from 436 to 75. Replaced the Monitor for defects Calibrations and QC to verify performance | SWO-543848 | Defect | HW | Basic Unit | PCB, Sensor interface | pO2 dark data | Customer called with calibration and internal QC errors. 1002 PO2 dark data is out of range | PO2 failing calibrations - PO2 dark data out of range Sensor Interface Serial# old-903-213R2008N0466 Monitor Serial# - Old - 910-353P1843N9968
Yes | Visit | Complaint | RDEU | pCO2 ist seit der Geräte-Installation im Level C an der oberen Grenze oder zu hoch | siehe SWO | pCO2-Werte stabil im Bereich Station informiert MedTech wird am 07.05.2019 informiert Daten gesichert | SWO-543792 | Re classified RMED | Courtesy |  |  |  | pCO2 has been at the upper limit or too high since the device was installed in Level C. | see SWO
No | Visit | Complaint | RIND | In Level 3 QC Na and pCO2 going out continuously. | QC out of range for Na. | Performed the zeroing found zeroing values are within the range, reinstalled the SC found installation is OK all the parameters are OK, performed the QC which I have carried different Lot 102, found Na values are below the range, need to replace with new analyzer and check. | SWO-546690 | [5 empty] | In Level 3 QC Na and pCO2 going out continuously. | QC out of range for Na.
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | GLU定标失败 1310,0210 | 更换测试卡后修复 | SWO-544101 | [5 empty] | Test card, reagent package failure | GLU calibration failure 1310,0210
No | Visit | Complaint | RCHN | po2定标失败，质控AB不过，报错 0375和0210、0588 | po2定标失败，质控AB不过，报错 0375和0210、0588，需要借件903-183一个 |  | SWO-543458 | Assisted Replacement | SP | SP | pO2 | 375, 588; wait for more info in CRM; check dates and revisit RAW DATA vs CRM | po2 calibration failure, QC AB but, error 0375 and 0210, 0588 | po2 calibration failure, QC AB however, misreported 0375 and 0210, 0588, need to borrow 903-183 one
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | po2定标失败，质控AB不过，报错 0588、0375、0210 | 更换测试卡后修复 | SWO-543460 | [5 empty] | Test card, reagent package failure | po2 calibration failure, QC AB but, error 0588, 0375, 0210
No | Visit | Complaint | RCHN | TNI、MYO质控低于统计范围的值 | TNI、MYO质控低于统计范围的值，错误代码：0588、0090 借903-364、944-214(15517） 944-212(15828) 944-216（15446）944-268（15802） 944-258（15732）缺货 |  | SWO-543374 | Revisit required | Consumables | QC | QC | stored improperly (see e-mail…) | TNI, MYO QUALITY control is below the statistical range of the value | TNI, MYO QC is lower than the statistical range of the value, error code: 0588, 0090 borrowing 903-364, 944-214 (15517) 944-212 (15828) 944-216 (15446) 944-268 (15802) 944-258 (15732) out of stock
No | Visit | Complaint | RCHN | 进样门打不开 | 进样门打不开，借886-638、886-639各一个 | 更换886-638、886-639后修复by Hunter | SWO-543156 | Defect | HW | Cover and Chassis | Door, Inlet | 5784 | The entrance door won't open. | The entrance door cannot be opened, borrow 886-638, 886-639 each one
No | Visit | Complaint | RAME | Customer called stating that they have been failing calibrations on analyzer for random errors. | Found that they were getting aborted calibrations for no rinse detected from above. Notice that the plug on the crea module was not seated all the way on dual valve housing. | Reseated the plug on the dual valve housing of crea module and flushed air valve and gave user some more basic troubleshooting advice about the analyzer. | SWO-543949 | Adjustment | HW | Met II module | Plug, Valve |  | Customer called stating that they have been failing calibrations on analyzer for random errors. | Found that they were getting aborted calibrations for no rinse detected from above. Notice that the plug on the crea module was not seated all the way on dual valve housing.
No | Visit | Complaint | RSCH | Rep Storage Door | Storage Door bleibt sporadisch offen. Analysator muss dann neu gestartet werden Ticket 013121 | Cabinet PCB ausgetauscht und justiert Funktionskontrolle i.o | SWO-543532 | Defect | HW | Cover and Chassis | PCB, cabinet | Storage Door | Rep Storage Door | Storage Door remains open sporadically. Analyzer must then be restarted Ticket 013121
No | Visit | Complaint | RESP | Impresora imprime hasta la mitad y se para. | Fallo de la impresora térmica, (imprime hasta la mitad y se para) | Comprobación del fallo, se procede a la sustitución de la misma. Revisión de funcionamiento y pasar QC´s. | SWO-544405 | Defect | HW | Printer Module | Printer Module |  | Printer prints halfway and stops. | Thermal printer failure, (prints halfway and stops)
Yes | Visit | Complaint | RFRA | Inlet reste en l’air | Après maintenance annuelle du matin, 2 capillaires ok, puis les capillaires ne passent plus. (Erreur échantillon) Les 2 premiers capillaires ok, puis résultat sans valeur puis erreur échantillon avec module Inlet reste ouvert. | Hotline fait des manipulations avec technicienne référente en Neonat. Passage capillaire, aspiration échantillon, inlet ne redescend pas, erreur échantillon puis l’écran se fige. Relance système ok, inlet se baisse ok. À mon arrivée à 17h30, je vérifie mécaniquement le module Inlet ok. Changement de la sonde ok (consommable du client, légèrement tordue). Meilleur coulissement du joint le long de la sonde ok. Vérification aspiration ok. Relance système ok. Validation Radiometer ok. Calibration ok, passage de tous les CQ ok. Tous les paramètres disponibles ok. Passage de 3 capillaires montés par technicienne ok. Passage d’1 capillaire par infirmière ok. Passage d’1 autre capillaire par infirmière prélevé d’un bébé ok. Pas d’autre capillaire disponibles. Plus d’erreur sur les 5 capillaires passés depuis changement de la sonde tordue (consommable). NB : le passage de capillaire augmente le risque de tordre la sonde. Pensez à humidifier le joint d’introduction blanc pour obtenir un meilleur coulissement le long d’une sonde non tordue. | SWO-545404 | Re classified RMED | Same day UpSWO |  |  | Inlet stays in the air Hotline does some manipulations with a reference technician in Neonat. Capillary flow, sample aspiration, inlet does not come back down, sample error then the screen freezes. System restart ok, inlet gets off ok. When I arrive at 17:30, I mechanically check the Inlet module ok. Probe change ok (customer's consumable, slightly twisted). Better sliding of the seal along the probe ok. Suction check ok. System restart ok. Radiometer validation ok. Calibration ok, move all CQ ok. All available settings ok. Passage of 3 capillaries mounted by technician ok. Passage of 1 capillary by nurse ok. Passage of 1 other capillary by nurse taken from a baby ok. No other capillaries available. More error on the 5 past capillaries since change of the twisted probe (consumable). NB: the capillary passage increases the risk of twisting the probe. Remember to moisten the white introducer seal for better sliding along a non-twisted probe. | Inlet stays in the air | After annual maintenance in the morning, 2 capillaries ok, then the capillaries no longer pass. (Sample error) The first 2 capillaries ok, then worthless result then sample error with Inlet module remains open.
No | Visit | Complaint | RFRA | Erreur de mesure du Vacuum value. Essai de débouchage du circuit, sans résultats. Vérifcation des différents éléments remplacés, ras. Un trace rose est présente sous le tuyau de la valve d'oxy. Nettoyage de la zone et relance de divers tests fluidiques. Du liquide est réapparut sous le tuyau, il y'a une fuite. Intervention ingénieur terrain nécessaire. | Tuyau vanne oxy percé | Changement tuyau ok. Changement cassette car cassette à séchée (faire demande de dédommagement) Maintenance annuelle ok. Test SAV ok. Vacuum value ok. Relance système ok, et attente remise en route ok. Validation Radiometer ok. Passage calibration ok. Passage de tous les CQ ok. Tous les paramètres disponibles ok. Pas de capillaire disponible à passer. NB : attention pas de sauvegarde automatique de paramètré sur abl90 de Neonat. (Ni sur clé usb, ni sur réseau) | SWO-544338 | Defect | HW | Oxi module | Pinch valve tube |  | Vacuum value measurement error. Test to unblock the circuit, without results. Verifcation of the various replaced elements, ras. A pink trace is present under the hose of the oxy valve. Cleaning the area and relaunching various fluid tests. Liquid reappeared under the pipe, there's a leak. Field engineer intervention required. | Oxy valve pipe pierced
No | Visit | Complaint | RAME | Customer called stating that Multiple parameters are failing calibrations. | Found that the analyzer was green when I got in. How ever based on log file they were getting drift out of range, and sensor impedance error on analyzer. | Went through service check out for impedance all values were in range. I then explain some more troubleshooting thing they can do when they get certain errors. Analyzer is all green now. | SWO-544483 | No fault found | No fault found | No fault found | No fault found |  | Customer called stating that Multiple parameters are failing calibrations. | Found that the analyzer was green when I got in. How ever based on log file they were getting drift out of range, and sensor impedance error on analyzer.
No | Visit | Complaint | RAME | Customer called stating that LAC and other parameters keep failing calibrations. | When I got there the analyzer was all green. But looking through the logs found that the customer kept having issue with LAC not passing was getting sensor response error and calibration error for sensitivity out of range. | Informed the customer that the analyzer was all green when I got here but went through so service test make sure everything was ok. All gave him some more basic troubleshooting information about the analyzers and what to do. | SWO-549759 | No fault found | No fault found | No fault found | No fault found | Lac | Customer called stating that LAC and other parameters keep failing calibrations. | When I got there the analyzer was all green. But looking through the logs found that the customer kept having issue with LAC not passing was getting sensor response error and calibration error for sensitivity out of range.
No | Visit | Complaint | RAME | CUSTOMER RAN CLOT IN ANALYZER. PROBE CLOTTED | CUSTOMER RAN CLOT IN ANALYZER. PROBE CLOTTED. BATTERY SHOWING 0 %. I TRIED MY NEW BATTERY AND IT GAVE A 1559:BATTERY ERROR DETECTED. RESET ORIGINAL BATTERY AND NOW SHOWS 100%. | FLUSHED ANALYZER. REPLACED INLET PROBE WITH CUSTOMER STOCK. REPLACED INLET CONNECTOR GASKET. RESET BATTERY. | SWO-544515 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway | Clot; probe, inlet gasket and connector gasket replaced | CUSTOMER RAN CLOT IN ANALYZER. PROBE CLOTTED | CUSTOMER RAN CLOT IN ANALYZER. PROBE CLOTTED. BATTERY SHOWING 0 %. I TRIED MY NEW BATTERY AND IT GAVE A 1559:BATTERY ERROR DETECTED. RESET ORIGINAL BATTERY AND NOW SHOWS 100%.
No | Visit | Complaint | RLTD | Wont calibrate liquid transport error. 07/05/19 09:33am - Analyser still not working. Please call POC. | Checked data logs and service programs. Clots suspected, Pressure test and vaious fluid transport errors. Checked Oxi tubing and valve operation, all ok. Checked gasket, connector gasket and inlet probe, all looked ok. Performed manual flush and found a solid blockage. Traced the blockage to the inlet probe and gasket. | Replaced the inlet probe and gasket. Replaced the SC as after an hour of calibrations pCO2, Glucose and Lactate sensitivties are all very low. | SWO-544392 | Defect | HW | Auto Inlet module | Probe | Clot | Wont calibrate liquid transport error. 07/05/19 09:33am - Analyser still not working. Please call POC. | Checked data logs and service programs. Clots suspected, Pressure test and vaious fluid transport errors. Checked Oxi tubing and valve operation, all ok. Checked gasket, connector gasket and inlet probe, all looked ok. Performed manual flush and found a solid blockage. Traced the blockage to the inlet probe and gasket.
No | Visit | Complaint | RAME | On-going issues carrying over from last week. Keeps asking for the SP to be replaced. | Reoccurrence of fluidic errors. Errors 1352 - clot suspected in Oxi module, 1351 - clot suspected in sensor cassette. Flushed entire fluidics, no clot $clog found or observed. Ran service diagnostics, fluid flow for solution pack failed all areas except QC3, failed clot detection, sensor cassette value 0.01 (.8 - 1.2) and failed Oxi module value 3.18 (.8 - 1.2). Found peristaltic pump spring mounting block not moveable, appears out of alignment. Suspect problem with solution pack not pumping solutions correctly. | Replaced Actuator stepper motor, actuator rod, fork & Pump rotor. Reassembled peristaltic pump spring block assembly confirming alignment and movement is smooth and consistent. Reassembled ABL90FP, performed service diagnostics, no failures. All tests now passing. | SWO-544223 | Defect | HW | Basic Unit | Stepper Motor | 1352, 1351 | On-going issues carrying over from last week. Keeps asking for the SP to be replaced. | Reoccurrence of fluidic errors. Errors 1352 - clot suspected in Oxi module, 1351 - clot suspected in sensor cassette. Flushed entire fluidics, no clot /clog found or observed. Ran service diagnostics, fluid flow for solution pack failed all areas except QC3, failed clot detection, sensor cassette value 0.01 (.8 - 1.2) and failed Oxi module value 3.18 (.8 - 1.2). Found peristaltic pump spring mounting block not moveable, appears out of alignment. Suspect problem with solution pack not pumping solutions correctly.
No | Visit | Complaint | RAME | Failed proficiency testing for pCO2 | Analyzer failed proficiency testing for pCO2. Customer ran a Linearity kit and then failed pO2 during linearity testing. | Replaced main distribution board, signal data board, sensor board and sensor board cable. Ran zeroing cassette and zeroed all channels. Ran calibrations and analyzer operating properly. Ran linearity with no issues. | SWO-544597 | [5 empty] | Failed proficiency testing for pCO2 | Analyzer failed proficiency testing for pCO2. Customer ran a Linearity kit and then failed pO2 during linearity testing.
Yes | Visit | Complaint | RCHN | 屏幕无法点动 | 屏幕无法点动，借903-720,933-494，各一个 | 更换903-720,933-494后修复BY Hunter | SWO-543564 | [5 empty] | The screen cannot be clicked | The screen cannot be clicked, borrow 903-720, 933-494, one each
Yes | Visit | Complaint | RCHN | TNI、MYO质控低于统计范围的值 | TNI、MYO质控低于统计范围的值,需要借件944-258（15711缺货）944-394（15784） | 更换903-364，处理盘，重新定标后修复by Hunter | SWO-544857 | Re classified RMED | Same day UpSWO |  |  |  | TNI, MYO QUALITY control is below the statistical range of the value | TNI, MYO quality control is lower than the statistical range of the value, need to borrow 944-258 (15711 out of stock) 944-394 (15784)
No | Visit | Complaint | RCHN | TNI、MYO质控低于统计范围的值 | TNI、MYO质控低于统计范围的值，样本的结果与科室的另一台仪器有差异，现在更换光单后，结果不准的问题已解决，但是质控还是有点偏差 | 对仪器的质控和结果进行对比，均没有发现问题，工程师最后怀疑是医院自行购买的质控品有问题，需要AS申请质控，后续再去现场验证by Hunter | SWO-545031 | Revisit required | Consumables | QC | QC | stored improperly (see e-mail…) | TNI, MYO QUALITY control is below the statistical range of the value | TNI, MYO quality control is lower than the value of the statistical range, the sample results and the department of another instrument is different, now replace the light sheet, the result is not allowed to solve the problem, but the quality control is still a bit biased
No | Visit | Complaint | RHUN | Blank Data | - | pH alsó tömítés csere, mérőkamra tisztítás, kalibráció, QC mérés *** GLU és LAC membrán csere esedékes, LAC kalibrációs hiba van | SWO-544291 | Defect | Electrode | Electrode | Sealing ring, pH |  | Blank Data | -
No | Visit | Complaint | RMAY | Services breakdown | Qc error due to calibration drift status high. | - Check analyzer data log - Clean chamber - Flush inlet - Perform protein removal | SWO-544843 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway | pCO2 | Services breakdown | Qc error due to calibration drift status high.
No | Visit | Complaint | RAME | Customer still getting sample errors after troubleshooting with tech support. Sample error 0722 Inhomogenous samples. | Looked in data logs and found numerous inhomogeneous samples on QC. | Replaced tubing and verified Auto check alignment. found cracked inlet nipples, replaced inlet nipples. Completed PM and wrote up on separate SWO. | SWO-544695 | Defect | HW | Inlet module | Nipple Kit | 722 | Customer still getting sample errors after troubleshooting with tech support. Sample error 0722 Inhomogenous samples. | Looked in data logs and found numerous inhomogeneous samples on QC.
No | Visit | Complaint | RIND | Calibration error | Checked data logs found pH calibration sensitivity out of range | Replaced reference membrane | SWO-546674 | User Error, FSE | Electrode | Membrane | Lifetime exceeded | see mail from PRARA with title "RE: RIND - ABL8XX - upSWO questions" | Calibration error | Checked data logs found pH calibration sensitivity out of range
No | Visit | Complaint | RIND | Calibration error | Checked data logs found pH calibration sensitivity out of range | Replaced reference membrane | SWO-546694 | User Error, FSE | Electrode | Membrane | Lifetime exceeded | see mail from PRARA with title "RE: RIND - ABL8XX - upSWO questions" | Calibration error | Checked data logs found pH calibration sensitivity out of range
No | Visit | Complaint | RSCH | Notfall HD | Probleme mit HD | Hard-Disk ersetzt ok Windows 7 und DMS installiert ok Membranen ersetzt am Crea A+B ok Alle Settings kontrolliert ok Kal und QC durchgeführt ok | SWO-546341 | Defect | HW | DMS_PC | SSD |  | Emergency HD | Problems with HD
No | Visit | Complaint | RDAN | From: Trine Rennebod Larsen <Trine.Rennebod.Larsen@rsyd.dk> Sent: 1. maj 2019 13:07 To: Heerfordt, Niels NHE <niels.heerfordt@radiometer.dk> Cc: Christina Myrén <Christina.Myren@rsyd.dk>; Birgit Christensen <Birgit.Christensen@rsyd.dk> Subject: uhomogene prøver Hej Niels. Vi har af og til nogle prøver, hvor ABL’en melder om en uhomogen prøve. Kan dette forekomme i andre tilfælde end hvis prøven ret faktisk ikke er blandet ordentligt? F.eks. oplever vi, at tipcap’en ikke altid bliver sat godt nok på og ikke altid fyldes. Kan dette give ophav til en melding om uhomogen prøve? Det opleves, at meldingen også kommer for prøver, som bliver bragt ned af en person, som rent faktisk går og vender prøven. Mange hilsner, Trine Venlig hilsen Trine Rennebod Larsen Biokemiker, Ph.D. Trine.Rennebod.Larsen@rsyd.dk - www.ouh.dk Tlf. 51779101 Klinisk Biokemisk Afdeling, Svendborg Svendborg Sygehus Baagøes Alle 15, 5700 Svendborg From: Trine Rennebod Larsen <Trine.Rennebod.Larsen@rsyd.dk> Sent: 8. maj 2019 11:20 To: Søeborg, Susanne SUSOE <susanne.soeborg@radiometer.dk> Cc: Christina Myrén <Christina.Myren@rsyd.dk>; Birgit Christensen <Birgit.Christensen@rsyd.dk>; Heerfordt, Niels NHE <niels.heerfordt@radiometer.dk> Subject: SV: Uhomogene prøver Kære Susanne. Tak for det. Vi får sat noget sammen, som vi kan sende ud til afdelingerne. Og sætte op ved apparaterne. Mange hilsner, Trine Venlig hilsen Trine Rennebod Larsen Biokemiker, Ph.D. Trine.Rennebod.Larsen@rsyd.dk - www.ouh.dk Tlf. 51779101 Klinisk Biokemisk Afdeling, Svendborg Svendborg Sygehus Baagøes Alle 15, 5700 Svendborg Fra: Søeborg, Susanne SUSOE [mailto:susanne.soeborg@radiometer.dk] Sendt: 8. maj 2019 09:18 Til: Trine Rennebod Larsen Cc: Christina Myrén; Birgit Christensen; Heerfordt, Niels NHE Emne: RE: Uhomogene prøver Hej Trine Tak for sidst. Jeg har vedhæftet et dokument hvor jeg har beskrevet vores observationer fra i går. Jeg har også vedhæftet insert til sprøjterne, hvor | Kiggede i spanden og der var 2-4 ud af 10 sprøjter hvor TipCap ikke var fyldt med blod. Og der var en del af sprøjterne TipCappen hvor den ikke sad rigtigt på (sad løst) | Vi konkluderede at dette var skyld i fejl koderne. Aftalen blev at kontakt bioana. ville undervise brugerne i af monterer TipCap rigtigt og fylde denne korrekt. | SWO-545873 | User Error | Sample handling | Sample handling | Sample handling | User error, Tip Cap was not mounted correctly; Error Code (Related Case) : 722, 374 Corrective Action : **** Primary Fault Corrective Action (Related Case) : Vi konkluderede at dette var skyld i fejl koderne. Aftalen blev at kontakt bioana. ville undervise brugerne i af monterer TipCap rigtigt og fylde denne korrekt. | From: Trine Rennebod Larsen <Trine.Rennebod.Larsen@rsyd.dk> Late: May 1, 2019 13:07 To: Heerfordt, Niels NHE <niels.heerfordt@radiometer.dk> Cc: Christina Myrén <Christina.Myren@rsyd.dk>; Birgit Christensen <Birgit.Christensen@rsyd.dk> Subject: Unhomogeneous samples Hi Niels. We sometimes have some tests in which the ABL reports an unhomogeneous sample. Can this occur in other cases than if the sample right is not actually mixed properly? For example, we find that the tipcap is not always put on well enough and not always filled. Can this give the creator of a message of unhomogeneous test? It is experienced that the announcement also comes to samples that are brought down by a person who actually goes and turns the test. Many greetings, Trine Regards Trine Rennebod Larsen Biochemist, Ph.D. Trine.Rennebod.Larsen@rsyd.dk-www.ouh.dk Tel. 51779101 Clinical Biochemical Department, Svendborg Svendborg Hospital Baagøes Alle 15, 5700 svendborg From: Trine Rennebod Larsen <Trine.Rennebod.Larsen@rsyd.dk> Late: May 8, 2019 11:20 Two: Søeborg, Susanne SUSOE <susanne.soeborg@radiometer.dk> Cc: Christina Myrén <Christina.Myren@rsyd.dk>; Birgit Christensen <Birgit.Christensen@rsyd.dk>; Heerfordt, Niels NHE <niels.heerfordt@radiometer.dk> Subject: A: unhomogeneous samples Dear Susanne. Thank you for that. We get something together that we can send out to the departments. And set up by the appliances. Many greetings, Trine Regards Trine Rennebod Larsen Biochemist, Ph.D. Trine.Rennebod.Larsen@rsyd.dk-www.ouh.dk Tel. 51779101 Clinical Biochemical Department, Svendborg Svendborg Hospital Baagøes Alle 15, 5700 svendborg From: Søeborg, Susanne SUSOE [mailto:susanne.soeborg@radiometer.dk] Posted: May 8, 2019 09:18 To: Trine Rennebod Larsen Cc: Christina Myrén; Birgit Christensen; Heerfordt, Niels NHE Subject: RE: unhomogeneous samples Hi Trine Thank you for the last time. I have attached a document where I have described our observations from yesterday. I also attached the insert to the syringes, where | Looked in the bucket and there were 2-4 out of 10 syringes where TipCap was not filled with blood. And there was a part of the sprayed Tipcapp where it was not seated properly (sat loose)
No | Visit | Complaint | RDAN | Hej Niels. Vi har af og til nogle prøver, hvor ABL’en melder om en uhomogen prøve. Kan dette forekomme i andre tilfælde end hvis prøven ret faktisk ikke er blandet ordentligt? F.eks. oplever vi, at tipcap’en ikke altid bliver sat godt nok på og ikke altid fyldes. Kan dette give ophav til en melding om uhomogen prøve? Det opleves, at meldingen også kommer for prøver, som bliver bragt ned af en person, som rent faktisk går og vender prøven. Mange hilsner, Trine Venlig hilsen Trine Rennebod Larsen Biokemiker, Ph.D. Trine.Rennebod.Larsen@rsyd.dk - www.ouh.dk Tlf. 51779101 Klinisk Biokemisk Afdeling, Svendborg Svendborg Sygehus Baagøes Alle 15, 5700 Svendborg Hej Trine Tak for sidst. Jeg har vedhæftet et dokument hvor jeg har beskrevet vores observationer fra i går. Jeg har også vedhæftet insert til sprøjterne, hvor der er en illustration der viser brugen af sprøjterne. Mvh Susanne Søeborg Applikationsspecialist Radiometer Danmark Direkte + 45 3827 2828 Mobil +45 20465127 e-mail Susanne.soeborg@radiometer.dk From: Trine Rennebod Larsen <Trine.Rennebod.Larsen@rsyd.dk> Sent: 2. maj 2019 14:28 To: Heerfordt, Niels NHE <niels.heerfordt@radiometer.dk> Cc: Christina Myrén <Christina.Myren@rsyd.dk>; Birgit Christensen <Birgit.Christensen@rsyd.dk>; Søeborg, Susanne SUSOE <susanne.soeborg@radiometer.dk> Subject: SV: Uhomogene prøver Hej Niels. Det lyder godt. Birgit, vil du informere FAM og intensiv? Mange hilsner, Trine Venlig hilsen Trine Rennebod Larsen Biokemiker, Ph.D. Trine.Rennebod.Larsen@rsyd.dk - www.ouh.dk Tlf. 51779101 Klinisk Biokemisk Afdeling, Svendborg Svendborg Sygehus Baagøes Alle 15, 5700 Svendborg Fra: Heerfordt, Niels NHE [mailto:niels.heerfordt@radiometer.dk] | Kiggede i spanden og der var 2-4 ud af 10 sprøjter hvor TipCap ikke var fyldt med blod. Og der var en del af sprøjterne TipCappen hvor den ikke sad rigtigt på (sad løst) | Vi konkluderede at dette var skyld i fejl koderne. Aftalen blev at kontakt bioana. ville undervise brugerne i af monterer TipCap rigtigt og fylde denne | SWO-545875 | User Error | Sample handling | Sample handling | Sample handling | User error, Tip Cap was not mounted correctly; Error Code (Related Case) : 722, 374 Corrective Action : **** Primary Fault Corrective Action (Related Case) : Vi konkluderede at dette var skyld i fejl koderne. Aftalen blev at kontakt bioana. ville undervise brugerne i af monterer TipCap rigtigt og fylde denne korrekt. | Hi Niels. We sometimes have some tests in which the ABL reports an unhomogeneous sample. Can this occur in other cases than if the sample right is not actually mixed properly? For example, we find that the tipcap is not always put on well enough and not always filled. Can this give the creator of a message of unhomogeneous test? It is experienced that the announcement also comes to samples that are brought down by a person who actually goes and turns the test. Many greetings, Trine Regards Trine Rennebod Larsen Biochemist, Ph.D. Trine.Rennebod.Larsen@rsyd.dk-www.ouh.dk Tel. 51779101 Clinical Biochemical Department, Svendborg Svendborg Hospital Baagøes Alle 15, 5700 svendborg Hi Trine Thank you for the last time. I have attached a document where I have described our observations from yesterday. I also attached the insert to the syringes, where there is an illustration showing the use of the syringes. Mvh Susanne Søeborg Application Specialist Radiometer Denmark Direct + 45 3827 2828 Mobile + 45 20465127 E-mail Susanne.soeborg@radiometer.dk From: Trine Rennebod Larsen <Trine.Rennebod.Larsen@rsyd.dk> Late: May 2, 2019 14:28 To: Heerfordt, Niels NHE <niels.heerfordt@radiometer.dk> Cc: Christina Myrén <Christina.Myren@rsyd.dk>; Birgit Christensen <Birgit.Christensen@rsyd.dk>; Susan SUSOE <susanne.soeborg@radiometer.dk> Subject: A: unhomogeneous samples Hi Niels. That sounds good. Birgit, will you inform FAM and intensive? Many greetings, Trine Regards Trine Rennebod Larsen Biochemist, Ph.D. Trine.Rennebod.Larsen@rsyd.dk-www.ouh.dk Tel. 51779101 Clinical Biochemical Department, Svendborg Svendborg Hospital Baagøes Alle 15, 5700 svendborg From: Heerfordt, Niels NHE [mailto:niels.heerfordt@radiometer.dk] | Looked in the bucket and there were 2-4 out of 10 syringes where TipCap was not filled with blood. And there was a part of the sprayed Tipcapp where it was not seated properly (sat loose)
No | Visit | Complaint | RAME | Analyzer not detecting any gas 2 pressure. | Gas 1 Empty at all times. | Installed loaner gas assembly and still have no gas 1 pressure. Found fault to be the control board. New board on order. Will return when part arrives. Analyzer available for ICa only. | SWO-544650 | Revisit required | HW | Analyzer Control Module | Fuse |  | Analyzer not detecting any gas 2 pressure. | Gas 1 Empty at all times.
No | 4 | Y | RCZE | Blank |  |  | D618B6B430476189 | User Error | Electrode | Membrane | Replacement |  | Blank
No | Visit | Complaint | RKK | ミキサーエラーが発生してしまう（1日1回程度） | 再度ミキサーモジュール交換実施。 またモジュールのみICU-LAB入れ替え実施、様子見。 | 再度ミキサーモジュール交換実施。 またモジュールのみICU-LAB入れ替え実施、様子見。 | SWO-547075 | Defect | HW | FlexQ Module | Mixer module |  | Mixer error occurs (about once a day) | The mixer module replacement was carried out again. In addition, only the module is performed, and the appearance is done.
Yes | Visit | Complaint | RPOL | Wyciek krwi pod aparatem | Ocena wizualna | Poprawiono wężyk pompowy od pompy nr 4, który spadł. Testy serwisowe. | SWO-544526 | [5 empty] | Blood leakage under the camera | Visual assessment
No | Visit | Complaint | RDEU | Druckbild besteht nur aus Streifen - Papier wurde ersetzt | Fehlersuche, Druckbild besteht nur aus Streifen - mechanisch keine Fehler, evtl Treiber abgestürzt | Update fällig (siehe weiterer Bericht)- nach Update und Systemneustart die Startphase abgewartet, Drucker getestet, Kal und QC | SWO-545127 | Defect | SW | RMED SW | Printer |  | Print image consists only of strips - paper has been replaced | Troubleshooting, print image consists only of stripes - mechanically no errors, possibly driver crashed
No | Visit | Complaint | RLTD | Please call Besa on 07534 067 522 | Waste tubing and or pinch valve tubing causing pressure errors, as yearly maintenance is due next month - fitted maintenance kit | Fitted maintenance kit - pressure error resolved. | SWO-545645 | Defect | HW | Basic Unit | PM kit | Pump cal error, Pressure error | Please call Besa on 07534 067 522 | Waste tubing and or pinch valve tubing causing pressure errors, as yearly maintenance is due next month - fitted maintenance kit
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 传感器阻抗错误，和定标失败 | 更换试剂包后修复 | SWO-544826 | [5 empty] | Test card, reagent package failure | Sensor impedance error, and calibration failed
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | K,CA定标失败 1318 阻抗错误 1076 | 连续更换2张测试卡后修复 | SWO-544939 | Assisted Replacement | SC | SC | K + Ca | 1318, 1076 | Test card, reagent package failure | K, CA calibration failure 1318 Impedance Error 1076
No | Visit | Complaint | RCHN | 定标底物外漏，报错1024 | 定标底物外漏，报错1024，借903-364,903-702各一个 | 更换903-364,903-702后修复by Hunter | SWO-544221 | Defect | HW | Process Module | Optical Unit | 1024 | The bid substrate leak, reported error 1024 | The bid substrate leak, reported error 1024, borrowed 903-364, 903-702 each one
No | Visit | Complaint | RCHN | 试管夹断裂 | 试管夹断裂，借903-584 $2个 | 更换903-584 $2个后修复by Hunter | SWO-543620 | Defect | HW | Aspiration Module | Tube holder |  | Test tube clip break | Test tube clip broken, borrowed 903-584 /2
No | Visit | Complaint | RCHN | 存储盘卡盘 | 存储盘卡盘，借903-378一个 | 更换903-378备件后修复by Hunter | SWO-544282 | Defect | HW | Storage Module | Cartridge lock | 5784 | Memory disk card | Memory disk card, borrow 903-378 one
No | Visit | Complaint | RCHN | 界面直接从应用程序退到windows系统 | 界面直接从应用程序退到windows系统，借910-343一个（未更换配件） | 间歇性故障，重启后解决，现科室不愿换备件，暂时可正常使用，此单关单by Hunter | SWO-544220 | Adjustment | SW | RMED SW | Windows |  | Interface back to windows system directly from application | Interface directly from the application to the windows system, borrow 910-343 one (unreplaced accessories)
No | Visit | Complaint | RCHN | 故障代码4548 | 故障代码4548 | 调节进样门支架位置后修复（by chuanwei） | SWO-544923 | Adjustment | HW | Cover and Chassis | Door, Inlet | 4548 | Fault code 4548 | Fault code 4548
No | Visit | Complaint | RDEU | 08.05.2019; Hr. Adebar; Tel.: +496213833528; Oxy LS Fehler 08.05.2019; 08:10; Hr. Adebar; Tel.: +496213833528; weiterhin Oxy LS Fehler | Oxi LS Error | Allgemeine Überprüfung - Luftstutzen gereinigt und gespült - Funktionsüberprüfung etc. | SWO-545106 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway | Oxi, 537 | 08.05.2019; Mr. Adebar; Tel.: +496213833528; Oxy LS Error 08.05.2019; 08:10; Mr. Adebar; Tel.: +496213833528; continue Oxy LS error | Oxygen LS Error
No | Visit | Complaint | RSCH | Crea QC Level 1 | Crea QC Level out of Range | Kunden Crea Elektroden wieder eingebaut Neu membraniert mit Lot R0868 Neue Kal Lösungen eingesetzt Kontrolle aktuel i.o | SWO-545115 | Assisted Replacement | Electrode | Membrane | Crea | QC L1 | Crea QC Level 1 | Crea QC Level out of Range
No | Visit | Complaint | RCHN | 血氧定标无法通过 | 血氧定标无法通过，借636-373一个 | 更换636-373备件后修复by Hunter | SWO-543607 | Defect | HW | Oxi module | Lamp unit | 258, 316 | Blood oxygen calibration can't pass | Blood oxygen calibration can not pass, borrow 636-373 one
No | Visit | Complaint | RDEU | Häufige Driften Elektrolyte. | Häufige Driften Elektrolyte trotz mehrfachem membranieren der Referenzmembran. | Meßkammern gereinigt, keine Verschmutzungen oder Undichtigkeiten feststellbar. Referenzelektrode erneuert. | SWO-545241 | Assisted Replacement | Electrode | Electrode | reference | CAPA-733 | Frequent drifting electrolytes. | Frequent drifting electrolytes despite multiple membraneing of the reference membrane.
No | Visit | Complaint | RDEU | unregelmäßige Kal.- u. QC-Fehler bei pH u.d. Lyten | Kal.-Fehler d. Lyte | Fehlersuche; neue Ref.-Elektrode (alt: 58-10) bezogen; div. Kal.; Kal. + QC i.O. | SWO-545290 | Assisted Replacement | Electrode | Membrane | reference |  | irregular cal and QC errors in pH and lytes | Kal.-Fehler d. Lyte
No | Visit | Complaint | RDEU | 06.05.2019 07:30Uhr Stein 0611433260: Faxauftrag - Einlassnadel Halter verbogen | Nadelhalter etwas verbogen. | Nadelhalter ausgetauscht. Funktionsprüfung i.O | SWO-545411 | Defect | HW | Inlet module | Probe bend |  | 06.05.2019 07:30h Stone 0611433260: Fax order - inlet pin holder bent | Needle holder slightly bent.
No | Visit | Complaint | RIND | Analyzer not working | Checked analyzer and found that software 6.18 was crashed. | Installed 6.18 software and restarted analyzer.Also performed clean up of c and d drive.They done all necessary setup. Checked with samples.Unit is working fine. | SWO-546260 | Defect | SW | RMED SW | RMED SW |  | Analyzer not working | Checked analyzer and found that software 6.18 was crashed.
No | Visit | Complaint | RSAF | HCT measurements errors on patient samples | Return visit | Replaced needle assembly and needle wash block. Hct calibration done. MCC 1 and 2 AND CRP done. Patient samples ok. | SWO-545367 | Defect | HW | Aspiration Module | Needle Wash Unit | 1264 | HCT measurements errors on patient samples | Return visit
No | Visit | Complaint | RDEU | 18.04.2019 11:50Uhr Frau Weidenhiller tel:+498431545478: Wegen einer Fehlermeldung wurde das Gerät neu gestartet und meldet jetzt 1052 und 1051 29.04.2019 10:15 Uhr Herr Mikolajek Medizintechnik tel:+498431545470: Fehler 1052 und 1051 ist wieder aufgetreten | Auf Pufferkassette kann nicht zugegriffen werden. Probe wurde abgbrochen. 3414: FTC_105 3415: FTC_106 | Pufferkassette im Service entfernt. Neue Pufferkassette eingelegt. AQT Messbereit. LQC TnI durchgeführt - ok. Servicedump und Datalogs gezogen. Vorsorglich Ersatzteile bestellt. | SWO-545348 | Defect | Consumables | Solution Pack | Valves | 1052; 3414; 3415 | 18.04.2019 11:50Pm Ms Weidenhiller tel:+498431545478: Due to an error message the device was restarted and now reports 1052 and 1051 29.04.2019 10:15 Mr Mikolajek Medical Technology tel:+498431545470: Error 1052 and 1051 has reappeared | Auf Pufferkassette kann nicht zugegriffen werden. Probe wurde abgbrochen. 3414: FTC_105 3415: FTC_106
No | Visit | Complaint | RPOL | Wykryto przeciek, 1184, 0508, 1063 | Uszkodzona uszczelka wlotu. | Wymiana uszczelki, czyszczenie selektora, uruchomienie prawidłowe. | SWO-548651 | Defect | HW | Auto Inlet module | Inlet gasket | 1184, 508, 1063 | Leak detected, 1184, 0508, 1063 | Damaged inlet gasket.
No | Visit | Complaint | RDEU | Analysatorverbindung zum CobasIT zeitweise gestört Anpassung der Netzwerkeinstellung zum Test | Netzwerkverbindung zum CobasIT zeitweise gestört | Mit Hr. Zukowski und Mitarbeiter der EDV Netzwerkverbindung überprüft Zur Testung Gerät auf DHCP umgestellt und ping freigegeben Testmessungen werden odnungsgemäß ans Cobis IT 1000 übertragen | SWO-545805 | Defect | SW | HIS/LIS | HIS/LIS |  | Analyzer connection to CobasIT temporarily disturbed Adjusting the network setting to the test | Network connection to CobasIT temporarily disrupted
No | Visit | Complaint | RLTD | Pressure errors - Vacuum valve error | On arrival the analyser was in operator intervention required. The service programs were entered and the fluid transport tests ran, vacuum valve pressure failed along with LS sensors and the solution pack check. The solution pack was removed and it was ensured that it was fully activated and that there were no clots in the analyser. The inlet gasket, inlet connector gasket and Oxi pinch tube were replaced, (part of a service kit, the service was completed in conjuncture with the repair and a separate report produced for this). The fluid transport tests were ran and passed. | A combination of ensuring that solution pack was activated and fitting the service kit has resolved errors. It was not possible to identify the precise failure. | SWO-546009 | Defect | HW | Basic Unit | PM kit | Pressure and vacuum error | Pressure errors - Vacuum valve error | On arrival the analyser was in operator intervention required. The service programs were entered and the fluid transport tests ran, vacuum valve pressure failed along with LS sensors and the solution pack check. The solution pack was removed and it was ensured that it was fully activated and that there were no clots in the analyser. The inlet gasket, inlet connector gasket and Oxi pinch tube were replaced, (part of a service kit, the service was completed in conjuncture with the repair and a separate report produced for this). The fluid transport tests were ran and passed.
No | Visit | Complaint | RAME | Analyzer no communicating with Viper. | Arrived onsite and found that the patient layout did not match from old analyzer to new. Changed patient layout to exactly match. Checked connection and found rs232 adapter in back of old analyzer. Migrated adapter on to new analyzer and sent results across. Viper received results. Analyzer working. | Connected adapter to com1 port. | SWO-546655 | Setup | HW | HIS/LIS | HIS/LIS | rs232 adapter | Analyzer no communicating with Viper. | Arrived onsite and found that the patient layout did not match from old analyzer to new. Changed patient layout to exactly match. Checked connection and found rs232 adapter in back of old analyzer. Migrated adapter on to new analyzer and sent results across. Viper received results. Analyzer working.
No | Visit | Complaint | RLTD | 1223 : Analyser did not connect at start-up | The analyser was powered up with Offline displayed in the top left hand corner. It looks like the wet section had not connected to the DMS when the analyser started up. | I powered the analyser off and removed and reconnected the USB cable between the Wet Section and the DMS. I powered the analyser back up and this time the Wet Section did connect to the DMS. I waited for the start up to complete and it went through without any problems. Several parameters have failed the calibration & QC most likely because the analyser has been out of action for so long. They may come back with time but if not the sensor cassette will need replacing, there wasn't a Crea sensor cassette available while I was on site. I also noticed that the battery icon was missing from the screen and the battery wasn't keeping the analyser powered up. I removed the battery and reset its charging circuit. The battery was re-installed and the analyser powered up and now the icon was present and after the battery had charged it kept the analyser powered up. Also when the analyser booted up there were problems with the Wet Section and DMS connecting. | SWO-545758 | Adjustment | HW | DMS_PC | Cable | 1223 | 1223 : Analyser did not connect at start-up | The analyser was powered up with Offline displayed in the top left hand corner. It looks like the wet section had not connected to the DMS when the analyser started up.
No | Visit | Complaint | RCHN | 报错9098，无法识别清洁管 | 报错9098，无法识别清洁管，借905-953一个 | 更换905-953后修复（by chuanwei） | SWO-544121 | Defect | HW | Inlet module | Barcode reader, Inlet | 9098 | Error 9098, unrecognizable cleaning tube | Error 9098, unrecognizable cleaning tube, borrowed 905-953 one
Yes | Visit | Complaint | RCHN | 液体流量选择器过高，报错 1278 | 液体流量选择器过高，报错 1278，需要借件903-119一个（未更换备件） | 疏通管路，清洁杯子冲洗单元（by chuanwei） | SWO-544935 | [5 empty] | Liquid flow selector too high, error 1278 | Liquid flow selector is too high, error 1278, need to borrow 903-119 one (not replaced spare parts)
No | Visit | Complaint | RCHN | BNP定标失败 | BNP定标失败,需要借件 944-258 15895 | 重新定标后修复（by chuanwei） | SWO-544906 | Assisted Replacement | Consumables | Test Kit | Cartridge, cal |  | BNP bid failed | BNP bid failed, need to borrow 944-258 15895
No | Visit | Complaint | RCHN | 仓门无法关闭 4548 | 仓门无法关闭 4548借903-028一个。 | 更换903-028后修复（by chuanwei) | SWO-544909 | Defect | HW | Cover and Chassis | PCB, cabinet top | 4548 | Warehouse door cannot be closed 4548 | The warehouse door cannot be closed 4548 borrowed 903-028 one.
No | Visit | Complaint | RCHN | 报错 4548 | 报错 4548 | 更换库存903-717后修复（by chuanwei） | SWO-545009 | Defect | HW | Cover and Chassis | Door, Inlet | 4548; top bracket | Error 4548 | Error 4548
No | Visit | Complaint | RCHN | 存储盘卡盘 | 存储盘卡盘 | 更换测试卡后修复 | SWO-547206 | Defect | Consumables | Test Kit | Cartridge, Broken | 5784 | Memory disk card | Memory disk card
Yes | Visit | Complaint | RMAY | Service breakdown | Calibration sensitivity and drift error due to ref membrane contaminated . | - Check analyzer data log - Clean chamber - Replace ref membrane ( user ) - Perform decontamination 2x - Run calibration | SWO-546896 | Assisted Replacement | Electrode | Membrane | reference |  | Service breakdown | Calibration sensitivity and drift error due to ref membrane contaminated .
No | Visit | Complaint | RCHN | CO2电极故障 | CO2电极故障，借945-612U | 更换945-612后修复（by chuanwei） | SWO-544142 | Assisted Replacement | Electrode | Electrode | pCO2 |  | CO2 electrode failure | CO2 electrode failure, 945-612U
No | Visit | Complaint | RAME | Unable to Print Cal Ver Results | Verified customer's stated issue. The CAL VER results appear under the patient results log but when trying to retrieve the results the screen just blinks and no results are displayed. This is only happening on the CAL VER results. When accessing the proficiency test results ran on the same day and within minutes of the last CAL VER the results appear on the screen and the customer is able to print them. | Re-downloaded sw keeping the customers data base and settings with no results displaying for CAL VER. Re-downloaded sw deleting the customers data base and settings and then importing them back in, with no results displaying for CAL VER. The customers current data base is corrupted. Suggested to the customer to start with a fresh data base when these units go live. | SWO-546108 | Defect | SW | RMED SW | database | corrupt files | Unable to Print Cal Ver Results | Verified customer's stated issue. The CAL VER results appear under the patient results log but when trying to retrieve the results the screen just blinks and no results are displayed. This is only happening on the CAL VER results. When accessing the proficiency test results ran on the same day and within minutes of the last CAL VER the results appear on the screen and the customer is able to print them.
No | Visit | Complaint | RESP | Keys Oxy10 | Comprobación de licencias instaladas de manera definitiva | Se instala paquete nuevo de licencias Oxi. Verificación de la correcta instalación. | SWO-546712 | Defect | SW | Demo key | Demo key | Potentially User Error, FSE | Keys Oxy10 | Checking for permanently installed licenses
No | Visit | Complaint | RCHN | 打印机走纸跑偏 | 打印机走纸跑偏，需要借件910-275一个 | 更换910-275后修复（by chuanwei） | SWO-544871 | Defect | HW | Printer Module | Printer Module |  | Printer paper run-off | Printer paper run off, need to borrow 910-275 one
No | Visit | Complaint | RESP | Keys | Comprobación de licencias instaladas | Instalación de paquete nuevo de licencias Oxy Comprobación de las licencias instaladas. | SWO-546713 | Defect | SW | Demo key | Demo key | Potentially User Error, FSE | Keys | Checking installed licenses
No | Visit | Complaint | RDEU | 07.05.2019; Fr. Bernauer; Tel.: 089529921; Probenabbrüche 08.05.2019 Sr. Martini +4989529921 meldet das weiterhin Probenabbrüche bestehen. | Probenabbrüche, Cartidge_25 | Storage Wheel ausgebaut, überprüft und adjustiert, 5x blankmessungen vorgenommen, Systemsreinigung durchgeführt, LQC MCC2 vorgenommen, Probe gemessen, ok. | SWO-545916 | Defect | Consumables | Test Kit | Cartridge, Broken | Cartridge_25 | 07.05.2019; Ms. Bernauer; Tel.: 089529921; Sample aborts 08.05.2019 Sr. Martini +4989529921 reports that there are still sample cancellations. | Probenabbrüche, Cartidge_25
No | Visit | Complaint | RPOL | Wykryto przeciek. | Wykryto przeciek. | Wymiana uszczelki, dociśnięcie pakietu, testy sekcji mokrej w normie. | SWO-548655 | Defect | HW | Auto Inlet module | Inlet gasket |  | A leak has been detected. | A leak has been detected.
No | Visit | Complaint | RPOL | Brak połączenia sieci LIS/HIS | Błąd transmisji danych pacjenta | Zamiana ustawień aparatu, zmiana sterownika przez firmę Marcel. | SWO-545876 | Setup | SW | HIS/LIS | HIS/LIS |  | No LIS/HIS network connection | Patient Data transmission Error
No | Visit | Complaint | RKK | プローブを変形させてしまい、正常な動作を実行できなくなってしまった。 |  |  | SWO-547818 | Defect | HW | Inlet module | Probe bend |  | The probe has been deformed, and it is no longer possible to perform normal operation.
No | Visit | Complaint | RLTD | Jamie Keegan Tel: 0191 445 5587 | Not getting full results on samples, showing insuffcient sample | The pinch valve was stripped, cleaned and re lubricated. Inlet connector gasket reseated correctly. Ran several test samples showing insufficient sample. On further inspection it was seen that the analyser had been setup for short probe. Once this was changed again the analyser was now accepting samples. | SWO-547371 | Setup | SW | RMED SW | Sample mode | Short probe | Jamie Keegan Tel: 0191 445 5587 | Not getting full results on samples, showing insuffcient sample
Yes | Visit | Complaint | RLTD | Inlet stuck open | Inlet stuck open | Talked user through removing the inlet, functioning the inlet without the inlet module on, then refitting the inlet, now ok | SWO-546525 | Re classified RMED | Phone support | Auto Inlet module | Auto Inlet module | See mail from JOHND with title "RE: SWO-546525 - Phone support or Visit?" | Inlet stuck open | Inlet stuck open
No | Visit | Complaint | RAME | Display Damaged Contact Lori Black. | Arrived onsite with part and replaced monitor unit 910-298. Powered on analyzer and calibrated touchscreen. Analyzer operational. | Replaced monitor. | SWO-546813 | Defect | HW | Display | Display |  | Display Damaged Contact Lori Black. | Arrived onsite with part and replaced monitor unit 910-298. Powered on analyzer and calibrated touchscreen. Analyzer operational.
No | Visit | Complaint | RHUN | újabb folyás, látszólag nincs oka | a folyás oka nem egyértelmű, minden funkció működik, oldat rendben, készülék oldal rendben | csere készülék beüzemelése, beállítások átköltöztetése, kalibráció, rendszer karbantartás, ***** eredeti készülék elszállítása tesztelés céljából | SWO-546441 | [5 empty] | Another discharge, apparently no reason | The cause of the discharge is unclear, all functions work, solution is OK, device page is OK
No | Visit | Complaint | RCHN | 老师投诉质控偏低 | 老师投诉质控偏低 | 做了质控瓶发觉质控正常，跟老师解释后认可 by Andrew | SWO-547294 | User Error | Consumables | QC | QC | stored improperly (see e-mail…) | Teacher complains about low quality control | Teacher complains about low quality control
No | Visit | Complaint | RCHN | 试管夹断裂 | 试管夹断裂，借903-584*2个 | 更换903-584*2个后修复（by chuanwei） | SWO-545577 | Defect | HW | Inlet module | Tube holder |  | Test tube clip break | Test tube clip broken, borrowed 903-584 x 2
No | Visit | Complaint | RCHN | 进样门开关故障 报错5784 | 进样门开关故障 报错5784 | 调节进样门支架位置后修复（by chuanwei） | SWO-546409 | Adjustment | HW | Cover and Chassis | Door, Inlet | 5784 | The sample door switch fault reported error 5784 | The sample door switch fault reported error 5784
No | 6 | Y | RCZE | Blank | S obsluhou najskor riesene cez telefon, obsluha nevidela nic podozrive, vymenila vstupne tesnenie a bez napravy... |  | D619B384A6576428 | Defect | Electrode | Membrane | MemPack missing | Ca | Blank | With the service of the probably ECT through the telephone, the operator did not see anything suspicious, replaced the entrance seal and without the execution...
No | Visit | Complaint | RDEU | 08.05.2019, Herr Weber MT: immer wieder Driften bei Elektrolyten | sporadisch Elektrolytdriften, pH ok | Reinigungsarbeiten, Fehlersuche, Ref.-elektrode (alt:08-17) ersetzt, Kal und QK ok | SWO-546377 | Assisted Replacement | Electrode | Electrode | reference | CAPA-733 | 08.05.2019, Mr. Weber MT: drifting in electrolytes again and again | sporadisch Elektrolytdriften, pH ok
No | Visit | Complaint | RAME | Analyzer not detecting any gas 2 pressure. | Checked data logs, found ICA not reading status or sensitivity. Checked electrodes in the service mode. Unable to determine if ICA is responding appropriately. Checked electrodes, cleaned chambers. Removed back cover and found the lower fuse to not be working, the light was off. | Swapped fuses and now both are working. Restarted analyzer, still not passing calibrations. Made multiple changes, finally tried an older calcium electrode and immediately got status and sensitivity numbers. Status is 10.38 and sensitivity is 100.0 | SWO-546868 | Defect | HW | Analyzer Control Module | Fuse |  | Analyzer not detecting any gas 2 pressure. | Checked data logs, found ICA not reading status or sensitivity. Checked electrodes in the service mode. Unable to determine if ICA is responding appropriately. Checked electrodes, cleaned chambers. Removed back cover and found the lower fuse to not be working, the light was off.
No | Visit | Complaint | RSAF | Neon intensity too low | None | Replaced neon lamp | SWO-547992 | Defect | HW | Oxi module | Lamp unit | Neon lamp | Neon intensity too low | None
No | Visit | Complaint | RDEU | häufige sporadische Elektrolytdriften | häufige sporadische Elektrolytdriften, pH ok | Reinigungsarbeiten, Fehlersuche, defekte Ref.-Elektrode (alt: 58-10) ersetzt | SWO-546512 | Assisted Replacement | Electrode | Electrode | reference | CAPA-733 | frequent sporadic electrolyte drifts | frequent sporadic electrolyte drifts, pH ok
No | Visit | Complaint | RDEU | 18.04.2019 11:50Uhr Frau Weidenhiller tel:+498431545478: Wegen einer Fehlermeldung wurde das Gerät neu gestartet und meldet jetzt 1052 und 1051 29.04.2019 10:15 Uhr Herr Mikolajek Medizintechnik tel:+498431545470: Fehler 1052 und 1051 ist wieder aufgetreten | Sporadisch kann auf Pufferkassette nicht zugegriffen werden. Fehler FTC_105 und FTC_106 | P1 bis P5 erneuert. Needle Wash Unit erneuert. Needle Kalibration (Washblock) druchgeführt. Needle Wash Unit kalibriert. Systemüberprüfung durchgeführt - ok. 1x LQC Tni, 1x LQC d-dimere, 1x LQC CRP, 1x LQC PCT und 1x LQC ßHCG vorgenommen. | SWO-547213 | Defect | HW | Aspiration Module | Needle Wash Unit |  | 18.04.2019 11:50Pm Ms Weidenhiller tel:+498431545478: Due to an error message the device was restarted and now reports 1052 and 1051 29.04.2019 10:15 Mr Mikolajek Medical Technology tel:+498431545470: Error 1052 and 1051 has reappeared | Sporadisch kann auf Pufferkassette nicht zugegriffen werden. Fehler FTC_105 und FTC_106
No | Visit | Complaint | RDEU | 02.05.2019 11:10 Fr. Erbe +496661812394: QC-Fehler PCT beide Level | QC im PCT außer Bereich | Neu gelieferte PCT-Kontrolle in beiden Bereichen mittig gemessen. Fehlerhafte Kontrolle wurde vom Anwender entsorgt. | SWO-546539 | Assisted Replacement | Consumables | LQC | LQC | Techline? Mail | 02.05.2019 11:10 Fr. Erbe +496661812394: QC-Fehler PCT beide Level | QC in PCT out of range
No | Visit | Complaint | RPOL | Wymiana modułu FLEXQ | Przegląd logów | Wymiana modułu FLEXQ na nowy z powodu błędu mieszalnika. Mieszalnik nie działał wcale i aparat zgłaszał ciągły błąd próbki. Testy serwisowe. Podmiana pastylki serwisowej na nową. | SWO-546663 | Defect | HW | FlexQ Module | FlexQ Module |  | Replacing the FLEXQ module | Log overview
Yes | Visit | Complaint | RPOL | Przeciek | Wizualna ocena | Wymiana pompy ściekowej, testy serwisowe. | SWO-546664 | [5 empty] | Leak | Visual rating
Yes | Visit | Complaint | RMAY | Oncall breakdown | Sample cannot aspirate due to clot in inlet gasket | - Check analyser data log -Perform manual flush -Run fluid transport test | SWO-547494 | [5 empty] | Oncall breakdown | Sample cannot aspirate due to clot in inlet gasket
No | Visit | Complaint | RCHN | 血氧光谱不匹配。报错 0581 | 血氧光谱不匹配。报错 0581。需要借件902-980一个 | 更换902-980备件后修复（by chuanwei） | SWO-546405 | Defect | HW | Oxi module | Hemolyzer | 581 | The blood oxygen spectrum does not match. Error 0581 | The blood oxygen spectrum does not match. Error 0581. Need to borrow 902-980 one
No | Visit | Complaint | RCHN | 干燥前温度超范围，报错1068 | 干燥前温度超范围，报错1068，借903-118,636-561各一个 | 更换903-118,636-561备件后修复（by chuanwei） | SWO-546247 | Defect | HW | Process Module | Cup dryer | 1068 | Temperature over range before drying, reported error 1068 | Over-range temperature before drying, reported error 1068, borrowed 903-118, 636-561 each
No | Visit | Complaint | RCHN | 机器一直重启，间歇性报错5784 | 机器一直重启，间歇性报错5784。需要借件 903-031 636-545 580-098 |  | SWO-546878 | Defect | HW | Process Module | Motor, process wheel | 5784 | The machine has been restarting, intermittently reporting error 5784 | The machine has been restarting, intermittently reporting an error of 5784. Borrowed 903-031 636-545 580-098
No | Visit | Complaint | RCHN | 医院断电之后，机器无法启动 | 医院断电之后，机器无法启动。需要借件910-343 | 更换910-343备件后修复(by chuanwei) | SWO-544075 | Defect | HW | Analyzer Control Module | CPU |  | The machine failed to start after the hospital lost power | The machine does not start after the hospital loses power. Need to borrow 910-343
No | Visit | Complaint | RCHN | 测试卡定标失败 | 测试卡定标失败，需要借件903-612一个 |  | SWO-546932 | [5 empty] | Failed to test card calibration | Test card calibration failed, need to borrow 903-612 one
No | Visit | Complaint | RIND | Printer not working | Printer not workinig | Check the liver , found broken. Check D lock ,found broken . | SWO-547898 | Defect | HW | Printer Module | Printer Head |  | Printer not working | Printer not workinig
No | Visit | Complaint | RIND | Printer Head problem | Checked and found that when printing given, end of the paper is dimmed and some end are lightened, Quality of Printer head is not good. | Need to replace the printer head for good printer working. | SWO-547133 | Defect | HW | Printer Module | Printer Head |  | Printer Head problem | Checked and found that when printing given, end of the paper is dimmed and some end are lightened, Quality of Printer head is not good.
No | Visit | Complaint | RKK | プリンタ印字がかすれているためプリンタヘッドを交換した。 | プリンタの印字かすれ | プリンタヘッド交換 | SWO-548835 | Defect | HW | Printer Module | Printer Head |  | The printer head was replaced because the printer print is blurry. | Printer print ing
Yes | Visit | Complaint | RMAY | Service breakdown | Leak detected due to clot | - Check analyser data log - Perform manual flush - Run fluid transport test ( OK) | SWO-547500 | [5 empty] | Service breakdown | Leak detected due to clot
Yes | Visit | Complaint | RCHN | TNI定标失败 | TNI定标失败，需要借件903-364 系统需要升级至4G，需要借910-373 | 更换903-364和910-373后修复 by Andrew | SWO-546911 | [5 empty] | TNI calibration failed | TNI bid fails, need s903-364 The system needs to be upgraded to 4G and needs to be borrowed 910-373
No | Visit | Complaint | RCHN | BNP定标失败 | BNP定标失败,需要借件944-258（15733）缺货 |  | SWO-546966 | Assisted Replacement | Consumables | Test Kit | Cartridge, cal |  | BNP bid failed | BNP bid failure, need to borrow 944-258 (15733) out of stock
No | Visit | Complaint | RCHN | 光学单元超出范围，报错1031 | 光学单元超出范围，报错1031,需要借件944-212 15957 |  | SWO-547004 | Revisit required | HW | Process Module | Optical Unit | 1031; revisit required | Optical unit out of range, error 1031 | Optical unit out of range, error 1031, need to borrow 944-212 15957
Yes | Visit | Complaint | RCHN | 光学单元超出范围，报错1031 | 光学单元超出范围，报错1031,需要借件903-364（需要西安办自提） |  | SWO-547005 | Re classified RMED | Same day UpSWO |  |  |  | Optical unit out of range, error 1031 | Optical unit out of range, error 1031, need to borrow 903-364 (need to Xi'an to do self-pick)
No | Visit | Complaint | RIND | Printer failure | Printer failure | Replace printer head, customer handheld damage, | SWO-548198 | Defect | HW | Printer Module | Printer Head |  | Printer failure | Printer failure
No | Visit | Complaint | RDEU | Intervention erforderlich - Leckage erkannt | Leckage erkannt | Reinigungsarbeiten, Fehlersuche, Gerät gespült, Serviceprogramme ok, Neustart ok, Kal & QK ok | SWO-547648 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway | Leak | Intervention required - leakage detected | Leakage detected
No | Visit | Complaint | RBNL | Oximetrie fouten bij stalen (ong 10%) | Bij ongeveer 10% van de stalen komen volgende foutmeldingen voor : 0292 (te hoge troebelheid) en 0581 (OXI spectrum mismatch) Calibraties en QC zijn OK | Vervanging van hemolyser, opstart, cal tHb en QC : OK Stalen gemeten : OK | SWO-547450 | Defect | HW | Oxi module | Hemolyzer | 292, 581 | Oximetry errors in samples (approx. 10%) | In approximately 10% of the samples, the following error messages occur: 0292 (too high turbidity) and 0581 (OXI spectrum mismatch) Calibrations and QC are OK
No | Visit | Complaint | RFRA | Site d’introduction ouvert | Contrôle du site d’introduction | Changement de la sonde d’aspiration. Remise en place du site d’introduction. Essais | SWO-547531 | Defect | HW | Auto Inlet module | Probe |  | Open introductory site | Control of the introductory site
No | Visit | Complaint | RPAC | Danny called 0428 759 737 liquid sensor error | Arrived to find machine not passine initializations. In service mode I was having a variety of problems with the liquid sensors and the pressure tests. Flushed it etc but could not shake the errors. | In the end I went ahead and replaced the pinch valve and did a full annual service on it. I was able to let it initialise. The machine has given a 2 hour startup so I wil check in the morning how the machine has come up. | SWO-547857 | Defect | HW | Oxi module | Pinch Valve | LS and pressue test error | Danny called 0428 759 737 liquid sensor error | Arrived to find machine not passine initializations. In service mode I was having a variety of problems with the liquid sensors and the pressure tests. Flushed it etc but could not shake the errors.
No | Visit | Complaint | RKK | 安定化中でキャリブレーションがとれない | 安定化中でキャリブレーションがとれない | 代替電極設置 | SWO-548583 | [5 empty] | Unable to calibrate during stabilization | Unable to calibrate during stabilization
No | Visit | Complaint | RCHN | 样本量不足，报错9098 | 样本量不足，报错9098，需要借件905-953（缺货）一个 | 使用备用机上的905-953备件后修复 by Andrew | SWO-546888 | Defect | HW | Inlet module | Barcode reader, Inlet | 9098 | Insufficient sample size, error 9098 | Insufficient sample size, error 9098, need to borrow 905-953 (out of stock) one
No | Visit | Complaint | RCHN | 液体量过高，报错1278 | 液体量过高，报错1278 | 清洁管路后修复 by Andrew | SWO-547204 | Cleaned | HW | Wet section | Fluidic pathway | 1278 | The amount of liquid is too high, the error 1278 | The amount of liquid is too high, the error 1278
No | Visit | Complaint | RCHN | 频繁要求重启仪器，处理盘卡盘 | 频繁要求重启仪器，处理盘卡盘，借636-533*2 |  | SWO-547102 | Cleaned | HW | Process Module | Sensor, process wheel | Corrective Action: Repair after re-cleaning the sensor by Andrew | Frequent requests for restart of instruments, processing disk | Frequent requests for restart of instruments, processing disk, borrowing 636-533 x 2
No | Visit | Complaint | RCHN | 测试卡门栓无法关闭 | 测试卡门栓无法关闭,需要借件903-627一个 | 更换903-627备件后修复 by Andrew | SWO-546930 | [5 empty] | Test Carmen bolt can't be closed | Test Carmen bolt can not be closed, need to borrow 903-627 one
Yes | Visit | Complaint | RMAY | Service breakdown | Cleaning required | - Check analyzer data log - Clean chamber - Flush inlet - Perform decontamination | SWO-547509 | [5 empty] | Service breakdown | Cleaning required
No | 5 | Y | RCZE | Blank |  |  | D61E6CE37E511B9A | Setup | SW | RMED SW | RMED SW | Hbf not shown | Blank
No | Visit | Complaint | RPAC | Clear waste tube popped off and spilled blood sample during weekend | Clear waste tube popped off and spilled blood sample during weekend, the tube has been reconnected. | Replaced clear waste tube and circular junction. Cleaned dried spill, waste pump rotor and tubings. Calibration & QC passed. Chris performed blood comparison, Ca varied, redid calibrations & blood comp, checked & replaced pH sealing O ring, re-calibrated & blood comp OK. | SWO-547053 | Adjustment | HW | Basic Unit | Waste tube |  | Clear waste tube popped off and spilled blood sample during weekend | Clear waste tube popped off and spilled blood sample during weekend, the tube has been reconnected.
No | Visit | Complaint | RIND | Sample aborted. | Check the patient result log & found last few samples got aborted | Check the gasket condition & found broken part of capillary glass tube was stuck into the gasket,remove this by using forcep,then flush gasket few times,run few sample,got the desired result & handover the instrument to the customer. | SWO-550906 | Cleaned | HW | Inlet module | Inlet gasket |  | Sample aborted. | Check the patient result log & found last few samples got aborted
No | Visit | Complaint | RFRA | Une fuite est présente au niveau du pack de réactif. Malgré le remplacement du pack la fuite persiste. | Tubulure évacuation débranchée avant le T. (Dérivation) | Remplacement des tubulures évacuation + T et des têtes de pompes du circuit de vidange. Nettoyage préventif du téton de raccordement au pack solution. Test amorçage et vidange x 7 ok. Pas d'action sur la mesure, uniquement évacuation déchet. | SWO-547667 | Defect | HW | Wet section | Waste tube |  | There is a leak in the reagent pack. Despite the replacement of the pack the leak persists. | Evacuation pipe unplugged before T. (Derivation)
Yes | Visit | Complaint | RDEU | Gerät quietscht | Gerät überprüft, | Potentialausgleichsfeder justiert, Gerät neu gestartet. | SWO-547355 | [5 empty] | Device squeaks | Device checked,
No | Visit | Complaint | RPOL | Drukarka termiczna nie drukuje na całej szerokości papieru. | Serwis stwierdza uszkodzoną głowicę drukarki termicznej. | Wymiana głowicy drukarki termicznej. | SWO-547414 | Defect | HW | Printer Module | Printer Head |  | The thermal printer does not print across the entire paper width. | The service finds a damaged thermal printer head.
No | Visit | Complaint | RLTD | The sensor cassette was replaced but shortly after it asked for the cassette to be replaced again and when it was taken out it was very hot to the touch. 10.05.2019 - SC is still overheating, door is hot, liquid is bubbling - Lee Murray Tel: 07388 044242 | sensor cassette overheating | replaced sensor bay and sensor interface PCB | SWO-548414 | Defect | SC | SC | Temp | Overheating | The sensor cassette was replaced but shortly after it asked for the cassette to be replaced again and when it was taken out it was very hot to the touch. 10.05.2019 - SC is still overheating, door is hot, liquid is bubbling - Lee Murray Tel: 07388 044242 | sensor cassette overheating
Yes | Visit | Complaint | RAME | Intervention Required and Leak Detected message | Error Leak detected message | Replaced Inlet Gasket Holder. Replaced Solutions Pack for troubleshooting purposes. | SWO-548639 | [5 empty] | Intervention Required and Leak Detected message | Error Leak detected message
Yes | Visit | Complaint | RLTD | Analyser failing with Oxi hardware issues | Failing Oxi copmplete test on hemolyser | Replaced Hemolyser, ran a complete oxi test (now passing), ran calibration and thb calibration then all QC's (all passed). | SWO-549647 | [5 empty] | Analyser failing with Oxi hardware issues | Failing Oxi copmplete test on hemolyser
No | Visit | Complaint | RDEU | Hotline - Intervention erforderlich | Intervention erforderlich , Gerät in Warteschleife Lösungspack verbraucht - 0 Zyklen ztw. Fehler - Gerät steht nich waagerecht | Gerät überprüft, Fliesswege im Service gespült , Lösungspack und Sensorcassette installiert, diverse Kal und QC gestartet. | SWO-548679 | Assisted Replacement | SP | SP | Empty |  | Hotline - Intervention required | Intervention required, device on hold Solution pack consumed - 0 cycles ztw. Error - device is not horizontal
No | Visit | Complaint | RAME | Broken monitor | Monitor case-multiple cracks at screw standoffs. | Replaced monitor and tested. | SWO-547790 | Defect | HW | Display | Display | Screws | Broken monitor | Monitor case-multiple cracks at screw standoffs.
No | Visit | Complaint | RFRA | Pbm sur imprimante, imprime que sur un côté. Nous vous confirmons la bonne prise en compte de votre demande sous la référence Ticket N° : 934 pour votre analyseur ABL90 FLEX PLUS 092R0196N019. Vous nous avez signalé le(s) problème(s) suivant : - L'imprimante de l'ABL90 de Kdy Mater (I393-092R0196N0019 ) n'imprime que sur la partie gauche - Merci d'intervenir - Nous avons également une K7 pour ABL80 qui périme fin mai que nous n'utiliserons pas - - Action menée: - La relance de l'analyseur a déjà été effectuée sans amélioration - Demande de remplacement de l'imprimante transmise à l'ingénieur SAV - Nous vous remercions de nous avoir contacté et vous confirmons que Guillaume Cavalier, votre ingénieur SAV terrain vous contactera dans les meilleurs délais. Votre demande a été traitée par Dominic Ritondale | Tête impression mécaniquement cassée, Impossibilité d’imprimer. | Changement imprimante Complète ok. Relance système ok. Essai impression ok. Validation Radiometer ok. Passage calibration ok. Passage de tous les CQ ok. Tous les paramètres disponibles ok. Rapport à envoyer par mail à bruno.poirey@labosud.fr lydie.saque@labosud.fr | SWO-548462 | Defect | HW | Printer Module | Printer Module |  | Pbm on printer, prints only on one side. We confirm that your application is properly considered under Ticket No. 934 for your ABL90 FLEX PLUS 092R0196N019 analyzer. You have reported the following problems: - Kdy Mater's ABL90 printer (I393-092R0196N0019) prints only on the left side - Thank you for intervening - We also have a K7 for ABL80 that expires at the end of May that we will not use - - Action: - The relaunch of the analyzer has already been carried out without improvement - Request to replace the printer sent to the SAV engineer - We thank you for contacting us and confirm that Guillaume Cavalier, your sAV field engineer, will contact you as soon as possible. Your application has been processed Dominic Ritondale | Mechanically broken print head, Unable to print.
No | Visit | Complaint | RAME | Analyzer keeps asking to have the SC and The SP replaced. | Analyzer keeps going into troubleshooting cycle wanting SP or inlet gasket to be replaced errors 1084, 1189, 1087 (LS and inhomogeneous errors on mor than one solution. When the old SP #HD59, was removed the winding was pink indicating a defective pump tube. | Replaced all tubing, inlet gasket, pinch valve tube, and connector gasket $ pm kit $ found a possible pinch in the waste tube at the waste fitting, Replaced SP and SC with new stock from James White. Ran diag - all passed and copied logs and ran service dump. | SWO-547850 | Defect | HW | Basic Unit | Waste tube |  | Analyzer keeps asking to have the SC and The SP replaced. | Analyzer keeps going into troubleshooting cycle wanting SP or inlet gasket to be replaced errors 1084, 1189, 1087 (LS and inhomogeneous errors on mor than one solution. When the old SP #HD59, was removed the winding was pink indicating a defective pump tube.
Yes | Visit | Complaint | RAME | Analyzer keeps asking to have the SC and The SP replaced. | Analyzer keeps going into troubleshooting cycle wanting SP or inlet gasket to be replaced errors 1084, 1189, 1087 (LS and inhomogeneous errors on mor than one solution. When the old SP #HD59, was removed the winding was pink indicating a defective pump tube. | Replaced all tubing, inlet gasket, pinch valve tube, and connector gasket $ pm kit $ found a possible pinch in the waste tube at the waste fitting, Replaced SP and SC with new stock from James White. Ran diag - all passed and copied logs and ran service dump. | SWO-547765 | Re classified RMED | Same day UpSWO |  |  |  | Analyzer keeps asking to have the SC and The SP replaced. | Analyzer keeps going into troubleshooting cycle wanting SP or inlet gasket to be replaced errors 1084, 1189, 1087 (LS and inhomogeneous errors on mor than one solution. When the old SP #HD59, was removed the winding was pink indicating a defective pump tube.
Yes | Visit | Complaint | RIND | Calibration error in O2 | Check system | Check system, reinstall SP and SC, recalibrate the analyzer, point 2 calibration success | SWO-551527 | [5 empty] | Calibration error in O2 | Check system
No | Visit | Complaint | RCHN | 做样本时有气泡 | 做样本时有气泡，需要借件903-099 | 更换903-099后修复 by Adrew | SWO-547859 | [5 empty] | There are bubbles when you take the sample | There are bubbles when taking samples, and a loan is required 903-099
No | Visit | Complaint | RCHN | 液体传输失败 | 液体传输失败，借902-527一个 | 更换902-527后修复 by Andrew | SWO-547290 | Defect | HW | Oxi module | Pinch Valve |  | Failed liquid transfer | Liquid transfer failed, borrowed 902-527 one
No | Visit | Complaint | RCHN | 科室反应机器频繁冲洗，进样口支架更换频繁 | 科室反应机器频繁冲洗，进样口支架更换频繁 | 注意样本抗凝，增加进样口支架备用 by jesson | SWO-551098 | User Error | Sample handling | Sample handling | pre analytical |  | Department reaction machine frequent flushing, inlet bracket replacement frequently | Department reaction machine frequent flushing, inlet bracket replacement frequently
No | Visit | Complaint | RCHN | 科室反应机器频繁冲洗，进样口支架更换频繁 | 科室反应机器频繁冲洗，进样口支架更换频繁 | 注意样本抗凝，增加进样口支架备用 by jesson | SWO-551099 | User Error | Sample handling | Sample handling | pre analytical |  | Department reaction machine frequent flushing, inlet bracket replacement frequently | Department reaction machine frequent flushing, inlet bracket replacement frequently
Yes | Visit | Complaint | RCHN | 处理包漏液 | 处理包漏液 | 更换试剂包后修复 | SWO-547864 | [5 empty] | Handling package leakage | Handling package leakage
No | Visit | Complaint | RCHN | 仓门无法关闭 | 仓门无法关闭,借903-717一个。 | 更换903-717一个正常。 by Andrew | SWO-547338 | Defect | HW | Cover and Chassis | Door, Inlet | Top bracket | The warehouse door cannot be closed | The warehouse door cannot be closed, borrowing 903-717 one.
No | Visit | Complaint | RIND | PCO2,PO2 calibration failed. | Check the calibration log & few parameters like pH,PCO2,PO2 got failed. | Flush the sensor cassette with the normal salaine,then run the calibration few times & it got passed successfully. | SWO-550909 | [5 empty] | PCO2,PO2 calibration failed. | Check the calibration log & few parameters like pH,PCO2,PO2 got failed.
No | Visit | Complaint | RDEU | 02.05.2019; Fr. Zenses; Tel.: 022678897049; lautes Pfeifen aus dem ABL | Lautes Pfeifen | Lüfter im AQC Modul getauscht. | SWO-548677 | Defect | HW | Autocheck module | Fan |  | 02.05.2019; Ms. Zenses; Tel.: 022678897049; loud whistling from the ABL | Loud whistling
No | Visit | Complaint | RIND | Crea Calibration sensitivity low. | Need to replace the crea membrane. | Replaced the Crea membrane then run the instrument its working good. | SWO-548837 | Assisted Replacement | Electrode | Membrane | Crea | Sens low | Crea Calibration sensitivity low. | Need to replace the crea membrane.
No | Visit | Complaint | RPAC | QC Ampoules intermittantly not being broken. | QC Ampoules intermittantly not being broken. | The analyser was placed into the Service mode and the Autocheck functions were checked. The Ampoule Breaker was the main focus of attention, however there were no issues found. The Ampoule Breaker Arm movement was checked many time, all fine. The analyser, once the QC Ampoules had their time warming up, Levels 1, 2 and 3 were performed, all fine. At present no issues could be found. However, if this issue does continue, then the module will be replaced. | SWO-547816 | No fault found | HW | Autocheck module | Autocheck module |  | QC Ampoules intermittantly not being broken. | QC Ampoules intermittantly not being broken.
No | Visit | Complaint | RAME | Analyzer giving rinse error | Inspected instrument and reviewed logs. Found rinse errors 751, 657, 510, and 509 $508. Piston on pH $bg module air valve showing in fluid. Lactate failing calibrations. | Replaced air valve and sealing ring on Lactate. | SWO-548463 | Defect | HW | pH/BG Module | Valve, air |  | Analyzer giving rinse error | Inspected instrument and reviewed logs. Found rinse errors 751, 657, 510, and 509 /508. Piston on pH /bg module air valve showing in fluid. Lactate failing calibrations.
No | Visit | Complaint | RAME | Qc stopped running. | Before arriving on-site I contacted Aqure to check their end of things out. Robert Newland dialed in and ran reports and confirmed everything seemed to be fine on their end. Once on site I went through data logs and found that machine is overlapping things in the schedule causing it to lock up. Daily cleaning and 2pt calibrations was programmed to start at 23:30 (which takes about 30 minutes to complete both) and Qc was schedule to start at 23:45. In addition the setting was on in corrective actions to lock analyzer immediately if qc is overdue. | I adjusted the daily cleaning to start at 22:00. This gives the machine plenty of time to complete its cleaning and the 2 pt calibration that follows it immediately. This should alleviate a lot of the problems customer has experienced with this problem. | SWO-547825 | Setup | SW | RMED SW | QC schedule |  | Qc stopped running. | Before arriving on-site I contacted Aqure to check their end of things out. Robert Newland dialed in and ran reports and confirmed everything seemed to be fine on their end. Once on site I went through data logs and found that machine is overlapping things in the schedule causing it to lock up. Daily cleaning and 2pt calibrations was programmed to start at 23:30 (which takes about 30 minutes to complete both) and Qc was schedule to start at 23:45. In addition the setting was on in corrective actions to lock analyzer immediately if qc is overdue.
No | Visit | Complaint | RAME | El/Met Module issues | Analyzer has been having drift intermittently on lytes and metabolites. This has been going on for a few months but recently reported to Radiometer. Tech support dispatched the call and recommended replacement of the El $Met module. | Reviewed data logs and found the drift is never more than + $-3 and on parameters CL, Ca2, lactate and Glu sometimes. Downloaded data logs and sent to Matt Rutrough for review. Customer has replaced all the membranes more than once as well as the calibration solutions have been changed.The customer has run several protein removals and decontamination programs on their regular schedule yet the drift continues. Went ahead and cleaned both measuring chambers and re membrane the Ca,Cl and Na electrodes. Also replaced the electrode pump tubings and the solution pump tubing. Replaced the cleaning solution bottle as well. Will let analyzer run overnight before replacing additional parts. I do have a El $Met module on hand and can replace based on Matt Rutrough’s findings. | SWO-547794 | Revisit required | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | El/Met Module issues | Analyzer has been having drift intermittently on lytes and metabolites. This has been going on for a few months but recently reported to Radiometer. Tech support dispatched the call and recommended replacement of the El /Met module.
No | Visit | Complaint | RAME | El/Met Module issues | Returned too site to check on calibration progress | Found only lactate has been drifting since the service done on 5 $13 $19. All other lytes are calibrating fine now. Customer has already re-membraned the lactate. Cleaned lactate chamber out again and ran more calibrations. Lactate is now passing calibration as well as all other parameters. Customer will monitor the calibrations and report to me if any of the drift errors return. Ran calibrations and QC found all levels passed. | SWO-548526 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | El/Met Module issues | Returned too site to check on calibration progress
No | Visit | Complaint | RIND | Windows corrupted. | Check the analyzer, Roma loader error present. | Install window software successfully. Install ABL800 software application successfully. Create new setup and settings. Performed 2point calibration without any error. | SWO-549104 | Defect | SW | RMED SW | Windows |  | Windows corrupted. | Check the analyzer, Roma loader error present.
No | Visit | Complaint | RCHN | ph状态不稳定，报错0375 | ph状态不稳定，报错0375，需要借件945-614U | 更换945-614U后发现为电极故障，现在需要借件945-614新电极一个 by Andrew | SWO-547295 | Revisit required | Electrode | Electrode | pH | 375 | Ph state is unstable, error 0375 | Ph state is unstable, error 0375, need to borrow 945-614U
Yes | Visit | Complaint | RPOL | Wyciek ścieków | Przyciśnięty wężyk od pompy nr 4 do trójnika | Wymiana wężyka | SWO-548317 | [5 empty] | Effluent leakage | Pressed hose from Pump No. 4 to Tee
No | Visit | Complaint | RKK | カセット不良が続くための対応策 |  |  | SWO-548804 | Defect | HW | Basic Unit | PCB, Sensor interface |  | What to do to keep the cassette defective
No | Visit | Complaint | RESP | Valores diferentes de hb entre gasimetria e hemograma | Verificação dos registos | Calibracao da hb e passagem do QC Recomenda-se a passagem do hemograma no aparelho de gasimetria para comparar | SWO-549188 | No fault found | No fault found | Oxi module | tHb |  | Different values of HB between Gasimetry and Hemogram | Verification of Records
No | Visit | Complaint | RLTD | Battery showing 0% | Battery showing 0% | reset battery jumper as per Radiometer guidelines. On start up analyser detecting battery error. New battery ordered and will return tomorrow to replace. | SWO-549071 | Revisit required | HW | Battery module | Battery module | 0% | Battery showing 0% | Battery showing 0%
No | 4 | Y | RCZE | Blank |  |  | D625DA41424F5F01 | Defect | HW | Display | Display | Frame | Blank
No | Visit | Complaint | RCHN | 温度过高 | 温度过高.需要借件903-334和902-972各一个 | 更换903-334和902-972后修复 by Andrew | SWO-547941 | [5 empty] | Too high temperature | The temperature is too high. One each for borrowings 903-334 and 902-972
No | Visit | Complaint | RIND | SC calibration failure | Check system | Check calibration, flush SC, again reinstall , given calibration ,calibration for few paraters then given active parameters same problem persist replace SC,calibration passed. | SWO-550879 | [5 empty] | SC calibration failure | Check system
Yes | Visit | Complaint | RCHN | 试管夹断裂 | 试管夹断裂，借903-584*2 | 更换903-584*2后修复 by jesson | SWO-547142 | [5 empty] | Test tube clip break | Test tube clip broken, borrowed 903-584 x 2
No | Visit | Complaint | RCHN | 光学强度超范围 1031 | 光学强度超范围 1031,借903-364一个。 | 更换903-364后修复 by Andrew | SWO-547356 | Defect | HW | Process Module | Optical Unit | 1031 | Optical strength over range 1031 | Optical strength over range 1031, borrowed 903-364 one.
No | Visit | Complaint | RCHN | 报错 4548 | 报错 4548，借903-717一个 | 更换903-717一个正常。 by Andrew | SWO-547387 | Defect | HW | Cover and Chassis | Door, Inlet | Top bracket | Error 4548 | Error 4548, loan 903-717 one
No | Visit | Complaint | RCHN | 光学单元超出范围，报错1031 | 加借903-364一个 | 更换903-364后修复 by Andrew | SWO-547833 | Defect | HW | Process Module | Optical Unit | 1031 | Optical unit out of range, error 1031 | Borrow 903-364 one
No | Visit | Complaint | RCHN | 温度无法被读取，排线破损，报错1020 | 温度无法被读取，排线破损，报错1020,需要借件636-452 | 更换636-452后修复 by Andrew | SWO-547999 | Defect | HW | Incubator Module | Cable, temp sensor | 1020 | The temperature cannot be read, the line is broken, the error 1020 is reported | The temperature cannot be read, the line is broken, the error 1020 is reported, the loan is required 636-452
No | Visit | Complaint | RIND | Multiple parameters failure in auto calibration. | Check the calibration results, pH, Pco2 and Cl- parameters are showing blank. | Active all parameters and performed 2point calibration. Performed zeroing successfully. | SWO-549166 | [5 empty] | Multiple parameters failure in auto calibration. | Check the calibration results, pH, Pco2 and Cl- parameters are showing blank.
No | Visit | Complaint | RKK | PCO2がQCで加減下回る。PO2が突然６以上ドリフトする | PCO2がQCで加減下回る。 PO2が６以上ドリフトする。 | PO2電極交換 PCO2のメンブランをロット違うものに交換。 ポンプローターをPHモジュールと溶液用を入れ替え。 | SWO-551085 | Assisted Replacement | Electrode | Electrode | pO2 | drift | PCO2 is lower than qc. PO2 suddenly drifts more than 6 | PCO2 is lower than qc. PO2 drifts more than 6.
No | Visit | Complaint | RKK | チャンバー汚れ、流路内の汚れが |  |  | SWO-549366 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | Chamber dirt, dirt in the flow path
No | Visit | Complaint | RKK | チャンバ内汚れ、流路の汚れの可能性 |  |  | SWO-549364 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | Possibility of dirt in the chamber and dirt in the flow path
No | Visit | Complaint | RSAF | Neon number too low | Performed complete oxi test,thb cal | An engineer will go back to site at a later stage | SWO-549272 | Revisit required | HW | Oxi module | Spectrometer | Neon number too low | Neon number too low | Performed complete oxi test,thb cal
Yes | 2 | Y | RCZE | Blank |  |  | D61FD12BF407417E | Re classified RMED | planned SWO |  |  | Installation | Blank
No | Visit | Complaint | RDEU | Meßunterschiede zum Labor, Jahreswartung vorziehen | Meßunterschiede zum Labor | Jahreswartung vorgezogen | SWO-549459 | Revisit required | HW | Process Module | Cup wash unit | discr. Results | Measurement differences to the laboratory, annual maintenance prefer | Measurement differences to the laboratory
No | Visit | Complaint | RPOL | Wystąpił błąd nieprawidłowego odporwadzenia ścieków podczas pracy analizatora. | Serwis stwierdza nieprawidłowe odprowadzenie ścieków. | Wykonanie 3x czyszczenia systemu, testowanie aparatu, błąd nie występuje po wykonanych czynnościach. | SWO-549195 | Cleaned | HW | Wet section | Fluidic pathway |  | An incorrect fault resistance occurred during the operation of the analyser. | The service finds incorrect drainage of wastewater.
No | Visit | Complaint | RAME | Clot Suspected | 1) Error 589: QC level C p02 above statistical range. | 1) Replaced Solution Pack. 2) Reviewed troubleshooting procedures with tech. 3) Provided customer Instruction For Use manual. 4) Customer has ordered spare inlet gaskets and inlet probes. | SWO-549896 | Assisted Replacement | SP | SP | pO2 |  | Clot Suspected | 1) Error 589: QC level C p02 above statistical range.
No | Visit | Complaint | RLTD | Battery showing 0% | Battery error detected | Battery replaced and seen to be charging correctly | SWO-550386 | Defect | HW | Battery module | Battery module | 0% | Battery showing 0% | Battery error detected
No | Visit | Complaint | RDEU | 20.03.2019 9:30 Fr. Anüsch 08912183306: SC-Thermistorfehler nach SC-Wechsel | Thermistorfehler (20.03.2019) | Kein Fehler vorhanden Gerät kalibriert und überprüft | SWO-549740 | No fault found | SC | SC | Thermistor |  | 20.03.2019 9:30 Fr. Anüsch 08912183306: SC thermistor error after SC change | Thermistorfehler (20.03.2019)
No | Visit | Complaint | RAME | Unable to Pump solution | Unable to pump solutions, found errors 1324,1352 & 508. In homogeneous solution errors as well as liquid transport errors. Customer had replaced SP on 5 $14 but this did not correct the problems. Upon arriving I found the analyzer was up and running fine. | The customer has had to replace the SP twice within the last two weeks and this analyzer was installed on 4 $22 $19. Due to the intermittent issues I have installed a pm kit in the analyzer. Tested analyzer once it re-initialized and found no issues with the analyzer. Ran calibrations, internal and external QC and found all levels passed, verbally signed off on by Scott Perry | SWO-549315 | Defect | HW | Basic Unit | PM kit | 1324, 1352, 508 | Unable to Pump solution | Unable to pump solutions, found errors 1324,1352 & 508. In homogeneous solution errors as well as liquid transport errors. Customer had replaced SP on 5 /14 but this did not correct the problems. Upon arriving I found the analyzer was up and running fine.
No | Visit | Complaint | RCHN | 触摸屏无反应 | 触摸屏无反应，借552-039一个 | 更换552-039后修复 by Andrew | SWO-548706 | [5 empty] | Touch screen does not respond | Touch screen non-reactive, borrow 552-039 a
No | Visit | Complaint | RIND | Calibration error | Check system | Check standby analyser , install new lot SP and SC on the analyser it is working ok, calibration pass and analyser handover to the concerning staff | SWO-550882 | [5 empty] | Calibration error | Check system
No | Visit | Complaint | RIND | Po2 not calibrating | Po2 not calibrating | Removed the cassette and reinstalled but its failed in the calibration so again reinstalled and while calibration time pressed the solution pack at that time it came to ready mode and | SWO-550761 | [5 empty] | Po2 not calibrating | Po2 not calibrating
No | Visit | Complaint | RCHN | 高铁血红蛋白质控高于统计范围，报错0089 | 高铁血红蛋白质控高于统计范围，报错0089，借902-980一个 | 更换902-980一个正常。 by Andrew | SWO-548744 | Defect | HW | Oxi module | Hemolyzer | 89 | High-speed blood red protein control higher than the statistical range, reported error 0089 | High-speed blood red protein control higher than the statistical range, reported error 0089, borrowed 902-980 one
No | Visit | Complaint | RCHN | 定标底物外漏检查失败 | 定标底物外漏检查失败，借903-364一个 | 更换903-364一个正常。 by Andrew | SWO-548710 | Defect | HW | Process Module | Optical Unit | 1024 | Failed to check for external leakage of the calibration substrate | The bid substrate leakage inspection failed, borrow 903-364 one
No | Visit | Complaint | RCHN | 定标失败，报错1067 | 定标失败，报错1067,需要借件944-220 15646 | 重新定标后修复 by Andrew | SWO-548630 | Assisted Replacement | Consumables | Test Kit | Cartridge, cal | 1067 | Failed to bid, error 1067 | Failed bid, error 1067, need to borrow 944-220 15646
No | Visit | Complaint | RCHN | 报错1030，光学单元超出范围 | 报错1031，光学单元超出范围，需要借件903-364 | 科室有两台AQT，这一台停用不修了。备件已好还。 by Andrew | SWO-549521 | Revisit required | HW | Process Module | Optical Unit | 1031 | Error 1030, optical unit out of range | Error 1031, optical unit out of range, need to borrow 903-364
No | Visit | Complaint | RCHN | 界面直接从应用程序退到windows系统 | 界面直接从应用程序退到windows系统,借910-343一个。 | 更换910-343一个正常。 by Andrew | SWO-549566 | Defect | HW | Analyzer Control Module | CPU |  | Interface back to windows system directly from application | The interface goes straight from the application to the windows system, borrowing 910-343 one.
No | Visit | Complaint | RKK | 溶液パックが突然認識できなくなる | 溶液パックが突然認識できなくなる | SPセンサー位置調整 SP取り付け確認 各部調整・点検 CAL確認 | SWO-550129 | [5 empty] | The solution pack suddenly becomes unrecognizable | The solution pack suddenly becomes unrecognizable
No | Visit | Complaint | RCHN | pco2定标失败 | pco2定标失败（未更换备件） | 清洁测量仓后修复 by jesson | SWO-553473 | Cleaned | Fluidic pathway | pH/BG Module | pH/BG Module | pCO2, 378 | pco2 calibration failure | pco2 calibration failure (spare parts not replaced)
No | Visit | Complaint | RAME | 722 & 374 in homogeneous sample error messages on qc. | Inhomogeneous samples on QC and blood samples | Replaced pH measuring chamber, inlet liquid sensor on inlet module and on PH $BG module, replaced all 3 inlet nipples and replaced all tubing. | SWO-549307 | Defect | HW | pH/BG Module | measuring chamber | 722, 374 | 722 & 374 in homogeneous sample error messages on qc. | Inhomogeneous samples on QC and blood samples
No | Visit | Complaint | RIND | Service Visit. | Need to check the analyzer status, Menu, Analyzer Status, Data Log, Tutorials hyperlink is not showing on the screen. | Connect the key board, start task manager, end the task of ROMA.ini and AU Engine in task manager, restart the application and platform software, successfully done, replace the rinse and waste container, analyzer handed over to staff in well working condition. | SWO-550189 | Defect | SW | RMED SW | RMED SW |  | Service Visit. | Need to check the analyzer status, Menu, Analyzer Status, Data Log, Tutorials hyperlink is not showing on the screen.
No | Visit | Complaint | RIND | Calibration drift 1 and 2 out of range | Checked the data logs crea Calibration drift 1 and 2 out of range for all 3 new crea membranes and its not getting stable | Need to replace crea electrode | SWO-550486 | Assisted Replacement | Electrode | Electrode | Crea |  | Calibration drift 1 and 2 out of range | Checked the data logs crea Calibration drift 1 and 2 out of range for all 3 new crea membranes and its not getting stable
No | Visit | Complaint | RPAC | Andrew said AED printouts showed a non-printed area 8/5/19 Checked AED printouts, there is a vertical non-printed thin line but results can still be read OK. Chris checked with AED charge nurse and they preferred not to work on machine today as results are still readable. 16/5/19 Replaced printer head | Poor printouts quality | Replaced printer head. Reinitialized analyzer, tested printout quality OK. Calibration & QCs passed. | SWO-549748 | Defect | HW | Printer Module | Printer Head |  | Andrew said AED printouts showed a non-printed area 8/5/19 Checked AED printouts, there is a vertical non-printed thin line but results can still be read OK. Chris checked with AED charge nurse and they preferred not to work on machine today as results are still readable. 16/5/19 Replaced printer head | Poor printouts quality
No | Visit | Complaint | RCHN | ph定标失败 | ph定标失败 ，需要借件945-614一个 | 更换945-614后修复 by Andrew | SWO-548567 | Assisted Replacement | Electrode | Electrode | pH | cal | ph calibration failure | Ph bid failure, need to borrow 945-614 one
No | Visit | Complaint | RIND | Replacement of pO2 | pO2 sensitivity is low . Training regarding consumable replacement. | Change the pO2 Membrane as per procedure. Run two point calibration . Calibration pass | SWO-554613 | Assisted Replacement | Electrode | Membrane | pO2 | Sens low | Replacement of pO2 | pO2 sensitivity is low . Training regarding consumable replacement.
No | Visit | Complaint | RCHN | PH测量不稳定 | PH测量不稳定 未更换备件 | 清洁测量室后修复 by jesson | SWO-553485 | Cleaned | Fluidic pathway | pH/BG Module | pH/BG Module | pH; 476 | PH measurement instability | PH measurement unstable Unreplaced spare parts
No | Visit | Complaint | RBNL | Bloedstalen worden afgebroken | Wanneer een meting start en een cup in de incubator geduwd wordt, loopt het blijkbaar fout. | Na grondige reiniging van de analyzer (onderhoud uitgevoerd) start de analyzer op zonder fouten en kunnen opnieuw stalen gemeten worden zonder fouten. | SWO-550290 | Cleaned | HW | Incubator Module | Incubator Module |  | Blood languages are broken down | When a measurement starts and a cup is pushed into the incubator, it apparently runs foul.
No | Visit | Complaint | RDEU | Meßunterschiede zum Labor, Jahreswartung vorziehen | Cup Wash Unit defekt | Cup wash Unit ersetzt, Blank Tests durchgeführt, optische Einheit gesäubert, Blank Tests wiederholt, QC o.k. 1 Ausreisser beim Blanktest | SWO-550186 | Defect | HW | Process Module | Cup wash unit | discr. Results | Measurement differences to the laboratory, annual maintenance prefer | Cup Wash Unit defective
No | Visit | Complaint | RDEU | 14.05.2019 15:45Uhr +493331271100 : Das Gerät verlangt einen Neustart nach dem Einschalten 1505.2019 07:30Uhr Fehler besteht weiterhin Ein Vorort Einsatz wird gewünscht. | Code: 5784 (Setsafe22) | Reinigung, Fehlersuche, defekte Testkassette (CRP Lot 15598) entfernt, Montage, Startphase, QK ok | SWO-549777 | Defect | Consumables | Test Kit | Cartridge, Broken | 5784 | 14.05.2019 15:45h +493331271100 : The device requires a restart after power on 1505.2019 07:30h Error persists A suburban deployment is desired. | Code: 5784 (Setsafe22)
Yes | Visit | Complaint | RIND | Service Visit. | Inlet door is not closing properly, analyzer shows error set safe_22. | Shut down the analyzer, open the front panel of the analyzer, check the inlet door, tighten the screw of the gear mechanism, successfully done, close the panel and start the analyzer, analyzer handed over to staff in well working condition. | SWO-549913 | Re classified RMED | Same day UpSWO |  |  |  | Service Visit. | Inlet door is not closing properly, analyzer shows error set safe_22.
No | Visit | Complaint | RIND | Service Visit. | Again the issue is with inlet door is not closing properly, error is same. | Shutdown the analyzer, open the panel, open the screw of both left and right gear, tie the Teflon tape on the supporter of the gear for tighten the mechanism, check it with the service program, successfully done, close the panel, start the analyzer and do the system cleaning successfully, analyzer handed over to staff in well working condition. | SWO-549917 | Adjustment | HW | Cover and Chassis | Door, Inlet |  | Service Visit. | Again the issue is with inlet door is not closing properly, error is same.
No | Visit | Complaint | RESP | Equipamento não liga | Verificação de que o equipamento não liga, desmontagem e verificação das ligações, teste da fonte ok, montagem e inicializacao.correcta, estabilização e passagem dos QCs. | Não foi detectada a razão da anomalia no equipamento. | SWO-550926 | No fault found | No fault found | No fault found | Power On |  | Equipment does not turn on | Verification that the equipment does not turn on, disassembly and verification of the connections, test the source OK, assembly and initialization. Stabilisation and passage of the QCs.
No | Visit | Complaint | RDEU | 16.05.2019 15:30 H. Bexter +491727657693: Aqureanbindung gestört | Keine Verbindung zu aqure. Keine korrekte EMail zugewiesen. | Konfigurationsfehler. Fehler trat nach Neustart auf. Daten auf CF gesichert. | SWO-550908 | Setup | SW | RMED SW | AQURE |  | 16.05.2019 15:30 H. Bexter +491727657693: Aqure connection disturbed | No connection to aqure. No correct email assigned.
No | Visit | Complaint | RKK | ヘモライザ初期不良によるリーク発生 |  |  | SWO-551053 | Defect | HW | Oxi module | Hemolyzer |  | Leak age due to initial defect of the hemolizer
No | Visit | Complaint | RDEU | Nach Sturz des Geätes ist das Gehäuse defekt | Gehäuse defekt nach Sturz | Tausch Monitorgehäuse, Funktionskontrolle | SWO-550522 | User Error | User Error | User Error | Display |  | After the fall of the saddle, the housing is defective | Housing defective after fall
No | Visit | Complaint | RAME | Battery - reading 0 percent | Battery issue. | Reset battery jumper during PM. Issue resolved. | SWO-550681 | Adjustment | HW | Battery module | Battery module | 0% | Battery - reading 0 percent | Battery issue.
No | Visit | Complaint | RLTD | Internal fixings broken | Internal damage on display unit case fixings | Replaced display unit | SWO-552319 | Defect | HW | Display | Display | Screws | Internal fixings broken | Internal damage on display unit case fixings
Yes | Visit | Complaint | RLTD | Change analyser from Crea/Urea to standard due to no consumables. Ref Andrew Masters and Sidki Akay | Change analyser from Crea $Urea to standard due to no consumables available. | Change analyser from Crea $Urea using service buttom to standard due to no consumables. In stalled SC900 and SPXL Start up 50 minuets. Cassette failing with impedance and response errors part 946-009 run 918-70. New case from NICU fitted and analyser 50 minuets start up calibration QCok. | SWO-550427 | Re classified RMED | Shipping |  |  | Crea/Urea consumabels not available | Change analyser from Crea/Urea to standard due to no consumables. Ref Andrew Masters and Sidki Akay | Change analyser from Crea /Urea to standard due to no consumables available.
Yes | Visit | Complaint | RLTD | Change analyser from Crea/Urea to standard due to no consumables available. Ref Andrew Masters and Sidki Akay. | Change analyser from Crea $Urea to standard due to no consumables available. | Analyser down graded using service button to standard and new consumables fitted SC 900,SP90XL Analyser 50 minutes start up. Calibration ok and QC ok apart from lactate sensitivity low SC to settle down. | SWO-550426 | Re classified RMED | Shipping |  |  | Crea/Urea consumabels not available | Change analyser from Crea/Urea to standard due to no consumables available. Ref Andrew Masters and Sidki Akay. | Change analyser from Crea /Urea to standard due to no consumables available.
No | Visit | Complaint | RPOL | Uszkodzona bateria UPS-a | Serwis stwierdza uszkodzoną baterię w zasilaczu UPS. | Wymiana baterii zasilacza awaryjnego. | SWO-550505 | [5 empty] | Defective UPS Battery | The service finds a defective battery in the UPS.
No | Visit | Complaint | RAME | Analyzer is giving a fluid blockage error. | Check the data logs. Tried to reinstall the sensor cassette and there was no fluid moving through it. | The solutions tube had come off the manifold and was put back. | SWO-550674 | [5 empty] | Analyzer is giving a fluid blockage error. | Check the data logs. Tried to reinstall the sensor cassette and there was no fluid moving through it.
No | Visit | Complaint | RAME | Possible fluid blockage | Customer reported that her existing SP was leaking therefore needed replacing. Upon replacing the new SP she said it says it’s not connected to the analyzer. | Upon arriving I found that the old leaking SP was in the analyzer with a fluid blockage message, cleaned up all the QC from the leaking SP. Started analyzer back up and found that the SP was installed on 4 $10 $19 therefore the customer did not install the SP correctly and that’s why it said pack not connected. Installed her SP and found it works fine, system cycles ran with no issues. I did replace the door latch as it was stained from the leaking SP. Ran system cycles and manual QC found all levels passed. Customer was not available upon completion, verbally acknowledgment from Margaret Jones | SWO-550751 | [5 empty] | Possible fluid blockage | Customer reported that her existing SP was leaking therefore needed replacing. Upon replacing the new SP she said it says it’s not connected to the analyzer.
No | Visit | Complaint | RCHN | 机器抽吸样本失败 | 机器抽吸样本失败 | 机器进样针弯了，更换进样针后，使用正常 by jesson | SWO-551100 | Defect | HW | Inlet module | Probe bend | 1271 | Failed to pump samples by the machine | Failed to pump samples by the machine
No | Visit | Complaint | RCHN | 进样支架断裂 | 进样支架断裂，借903-584一个 | 更换903-584一个 by Andrew | SWO-549490 | Defect | HW | Inlet module | Tube holder |  | The sample bracket breaks | The sample bracket is broken, and 903-584 is a
No | Visit | Complaint | RDEU | Anwender (Labor) sehr unzufrieden über Stabilität des Analysators. | Analysator läuft instabil. Glucose driftet und es liegt eine Qualitätskontrolle Grenzverletzung vor. Elektrolyt fehlt in der Membran, Anwender Fehler. Oximetrie Grenzverletzung bei tHb, alle Level werden zu hoch gemessen an der oberen Herstellergrenze mit häufiger Grenzverletzung. Anwender haben vor ein paar Tagen tHb Kalibration durchgeführt. Es kommt immer wieder zu Driften in den Elektrolyten und pH, besonders bei Ca, Elektrode nicht stabil. Analysator zeigt unter Meldungen: 0529 Inlet liquide Sensor error an. | Glucose Elektrode gereinigt und neu membraniert. Alle Messkammern gereinigt und Elektroden überprüft. pH Elektrode gereinigt. Einbau eines neuen Oximetrie Moduls. Durchführung der tHb Kalibration im Service und im Anwender Bereich. Tausch der Referenz Elektrode notwendig, keine auf dem Fahrzeug (Lieferengpass), gebrauchte Referenz Elektrode eingesetzt (6811). Bei Lieferung erfolgt Einbau. Nach Neustart des Analysators macht er gleich eine LS Justage, Fehler behoben. Überprüfung der Messwerte im Service zeigt, dass die Werte innerhalb der Grenzwerte sind. | SWO-550621 | Assisted Replacement | Electrode | Electrode | reference | CAPA-733 | User (laboratory) very dissatisfied about the stability of the analyzer. | Analyzer runs unstable. Glucose drifts and there is a quality control limit violation. Electrolyte is missing in the membrane, user error. Oximetry boundary violation at tHb, all levels are too high measured at the upper manufacturer's limit with frequent boundary violation. Users performed tHb calibration a few days ago. Drifting in the electrolytes and pH, especially with Ca, electrode is not stable again and again. Analyzer displays under messages: 0529 Inlet liquid sensor error.
No | 5 | Y | RCZE | Blank |  |  | D625DC5A42045F0D | Assisted Replacement | Electrode | Electrode | reference | 815; | Blank
No | Visit | Complaint | RSAF | Neon number too low | None | Replaced spectrometer with a loan unit | SWO-550576 | Defect | HW | Oxi module | Spectrometer | Neon number too low | Neon number too low | None
No | Visit | Complaint | RIND | pCO2 Sensitivity is low | Sensitivity is low in calibration. Training regarding consumable replacement is to be given. | Check the calibration log , found Sensitivity is low . Run protein removal , calibration given. All ok in calibration. Training regarding replacement of consumables is given to staff. | SWO-554614 | Cleaned | Fluidic pathway | pH/BG Module | pH/BG Module | pCO2; Sens low | pCO2 Sensitivity is low | Sensitivity is low in calibration. Training regarding consumable replacement is to be given.
No | Visit | Complaint | RCHN | ph状态不稳定，报错0375 | ph状态不稳定，报错0375，需要借件945-614一个 | 更换945-614一个正常。 by Andrew | SWO-548553 | Assisted Replacement | Electrode | Electrode | pH | 375 | Ph state is unstable, error 0375 | Ph state is unstable, misreported 0375, need to borrow 945-614 one
No | Visit | Complaint | RDEU | Meßunterschiede zum Labor, Jahreswartung vorziehen | Meßunterschiede zum Labor, optische Einheit defekt | Gerät überprüft, optische Einheit erneuert, Parameter kalibriert, Blank Tests durchgeführt weiterhin zeitweise Meßausreisser | SWO-550518 | Defect | HW | Process Module | Optical Unit | discr. Results | Measurement differences to the laboratory, annual maintenance prefer | Measurement differences to the laboratory, optical unit defective
No | Visit | Complaint | RPOL | Pozostałość krwi w łączeniach między złączami elektrod. | Serwis stwierdza obecność zakrzepłej krwi w łączeniach pomiędzy złączami elektrod. | Czyszczenie modułów za pomocą podchlorynu, przepłukanie wodą, ponowna instalacja elektrod, kalibracja. Pozostawiono 2 brakujące elektrody ślepe. | SWO-550503 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway | Clot | Blood residue in joints between electrode joints. | The service finds the presence of a blood clot in the joints between the electrode connectors.
No | Visit | Complaint | RPOL | Wyciek z uszczelki wlotu podczas płukania. | Serwis stwierdza obecność skrzepu w wężyku ściekowym. | Usunięcie skrzepu, odbiałczanie, testowanie aparatu. | SWO-550504 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway | Clot | Leakage from the inlet gasket during flushing. | The service finds the presence of a clot in the sewage hose.
No | Visit | Complaint | RCHN | 孵育盘温度错误 | 孵育盘温度错误，借903-138和902-478各一个 | 更换903-138一个正常。 by Andrew | SWO-549641 | Defect | HW | Incubator Module | PCB, incubator | 3564 | Incubation plate temperature error | Incubation plate temperature error, borrow 903-138 and 902-478 each one
No | Visit | Complaint | RCHN | 样本的舱门无法关闭，报错5784 | 样本的舱门无法关闭，报错5784，需要借件903-717 | 调整齿条螺丝正常。 by Andrew | SWO-550259 | Adjustment | HW | Cover and Chassis | Door, Inlet | 5784 | The sample door could not be closed, 5784 was reported | The sample door could not be closed, 5784 was reported, and a loan of 903-717 was required.
No | Visit | Complaint | RCHN | 进样门打不开 | 进样门打不开，借903-198一个 | 更换903-198一个正常。 by Andrew | SWO-549570 | Defect | HW | Cover and Chassis | Door, Inlet | Potentially Use Error, FSE | The entrance door won't open. | The entrance door won't open, borrow 903-198 one.
No | Visit | Complaint | RCHN | 进样门无法关闭 5784 | 进样门无法关闭 5784,借903-224，903-717各一个。 | 更换903-224，903-717各一个正常。 by Andrew | SWO-548814 | Defect | HW | Cover and Chassis | Door, Inlet | Top bracket | The sample door cannot be closed 5784 | The inlet door cannot be closed 5784, borrowing 903-224, 903-717 each.
No | Visit | Complaint | RCHN | 舱门无法打开 | 舱门无法打开，借903-717一个 | 更换903-717一个正常 by Andrew | SWO-549407 | Defect | HW | Cover and Chassis | Door, Inlet | Top bracket | The hatch cannot be opened. | The hatch can't be opened, borrow 903-717.
No | Visit | Complaint | RCHN | 试剂泵不对称错误，吸样过程中液体传感器过早检测到样本 | 试剂泵不对称错误，吸样过程中液体传感器过早检测到样本。调整液体传感器，泵定标。关机重启正常。未更换备件。 | 调整液体传感器，泵定标。关机重启正常。未更换备件。 by Andrew | SWO-550961 | Adjustment | HW | Fluidic manifold | Pumps |  | Reagent pump asymmetry error, liquid sensor detects sample sits prematurely during suction | The reagent pump is asymmetrically incorrect and the liquid sensor detects the sample prematurely during the suction. Adjust the liquid sensor and pump calibration. The shutdown restarts normally. Spare parts have not been replaced.
No | Visit | Complaint | RIND | Service Visit. | Qc is not coming in the range, membrane replace already but sensitivity is not coming. After observation inlet not filled properly during the bidirectional pump calibration, inlet not filled properly during rinse, air is detected during aspiration. These are the issue having in the analyzer. | Check the calibration status, replace the membrane of Na, successfully done, do the inlet calibration, LS adjust and pump calibration all are successful, but in EI $Met chamber observe that the bubble has generated, check the electrode, by cross verification one electrode doesn't create the vacuum in to the channel, need to keep the dummy electrode in all the in to the channel and try with Na electrode because already we replace the 2 membranes and 15 days before reference has been replaced by the user. | SWO-551046 | Assisted Replacement | Electrode | Membrane | Na |  | Service Visit. | Qc is not coming in the range, membrane replace already but sensitivity is not coming. After observation inlet not filled properly during the bidirectional pump calibration, inlet not filled properly during rinse, air is detected during aspiration. These are the issue having in the analyzer.
No | Visit | Complaint | RSAF | Ph failing callibration | Replaced solution pack | Replaced sensor cassete | SWO-550998 | Assisted Replacement | SC | SC | pH | Cal | Ph failing callibration | Replaced solution pack
No | Visit | Complaint | RCHN | BNP定标失败 1999 | BNP定标失败 1999,借94-4258*15895一 张 |  | SWO-550958 | Cleaned | HW | Process Module | Process Wheel | CRM: Corrective Action: Clear internal dust and recalibrate normally by jesson | BNP calibration failure 1999 | BNP bid failure 1999, borrow edding 94-4258 x 15895
No | Visit | Complaint | RCHN | 开机报错 2132，0002系统软件不匹配 | 开机报错 2132，0002系统软件不匹配借903-198一个。 |  | SWO-550968 | Defect | HW | Analyzer Control Module | PCB, Main Board | 2132; 0002; CRM: Corrective Action: Repaired after replacing 903-198 by jesson | Boot error 2132,0002 system software mismatch | The boot error 2132,0002 system software does not match to borrow 903-198 one.
No | Visit | Complaint | RCHN | 处理单元升降故障 5257 | 处理单元升降故障 5257,借944-220*15599，944-258*15611 | 重新定标正常。 by Andrew | SWO-550350 | Assisted Replacement | Consumables | Test Kit | Cartridge, cal | 5257 | Treatment unit lift fault 5257 | Treatment unit lift fault 5257, borrowed 944-220 x 15599, 944-258 x 15611
No | Visit | Complaint | RCHN | 存储盘卡盘 定标失败 5784 5689 | 存储盘卡盘 定标失败 5784 5689。借944-394*15886一张 |  | SWO-550972 | Defect | Consumables | Test Kit | Cartridge, Broken | 5784; 5689; CRM: Corrective Action: Repair after recalibration by jesson | Memory disk card bid failure 5784 5689 | The memory disk card disk bid failed 5784 5689. Borrow 944-394 x 15886
No | Visit | Complaint | RCHN | 触摸屏没反映 | 触摸屏没反映，借910-320一个 | 更换910-320后故障仍存在。加借910-320一个。by Andrew | SWO-550128 | Defect | HW | Display | Display |  | The touch screen doesn't reflect. | Touch screen does not reflect, borrow 910-320 a
No | Visit | Complaint | RCHN | 触摸屏没反映 | 更换910-320后故障仍存在。加借910-320一个 |  | SWO-550956 | Defect | HW | Display | Display | upSWO Date changed to 2019-05-21 | The touch screen doesn't reflect. | The fault still exists after replacing the 910-320. Borrow 910-320 one
No | Visit | Complaint | RDEU | 09.05.2019, Herr Manhoff, tel:+4979416928122: Glu Werte zu hoch im Vergleich zu einem anderen BZ Gerät. | Fehlerhafte Glu und diferse Elytefehler bei Kal., QC, und Probenmessung. Negativer Statuswert bei Glu. | Kal. und QC Werte analysiert. Teile im Umfang einer Jahreswartung erneuert. ( Jahreswartung vorgezogen, siehe SWO Jahreswartung). Kontakte der SC-PCB überprüft und gereinigt. SC und LP aus VSC-SW-Bestand erneuert. Flachbandkabel AU-PCB->SC-PCB erneuert. Einlaufverfahren durchgeführt. Alle Kal.-Werte sofort im Normbereich, QC nach mehreren Kal. o.k. Vergleichsmessungen und 10xQC durchgeführt, keine Auffälligkeiten mehr feststellbar. | SWO-552690 | Cleaned | HW | Basic Unit | PCB, Sensor interface |  | 05/09/2019, Mr. Manhoff, tel:+4979416928122: Glu values too high compared to another BZ device. | Faulty Glu and diferse Elyte errors in Cal., QC, and sample measurement. Negative status value for Glu.
No | Visit | Complaint | RDEU | 08.04.2019; Hr. Weiser; Tel.: 0765129413; ABL verlängert die Messzeit 11.04.2019, Frau Guth, Tel: 0765129413: ständige Kalibrierungen ca im 30 Min. Takt 12.04.2019 08:30Uhr Frau Haas tel:+4976512978412: Das Gerät geht ständig in die Startphase | 08.04.2019; Hr. Weiser; Tel.: 0765129413; ABL verlängert die Messzeit 11.04.2019, Frau Guth, Tel: 0765129413: ständige Kalibrierungen ca im 30 Min. Takt 12.04.2019 08:30Uhr Frau Haas tel:+4976512978412: Das Gerät geht ständig in die Startphase | Gerät beim Kunden aufgestellt und einkalibriert, Kal. und QC i.O., Probleme mit Netzwerkanbindung, Anbindung mit POC IT abgestimmt | SWO-552311 | Defect | HW | DMS_PC | CF card | Repair in workshop SWO-552286 | 08.04.2019; Mr. Weiser; Tel.: 0765129413; ABL extends the measuring time 11.04.2019, Mrs. Guth, Tel: 0765129413: continuous calibrations approx. 30 min. 12.04.2019 08:30Pm Ms Haas tel:+4976512978412: The device is constantly going into the start-up phase | 08.04.2019; Mr. Weiser; Tel.: 0765129413; ABL extends the measuring time 11.04.2019, Mrs. Guth, Tel: 0765129413: continuous calibrations approx. 30 min. 12.04.2019 08:30Pm Ms Haas tel:+4976512978412: The device is constantly going into the start-up phase
No | Visit | Complaint | RAME | Boot failed | Arrived onsite and found analyzer not booting into system software. Reloaded software without Roma and databases. Analyzer powered up to green. Loaded setup from duplicate analyzer. Named NICU. Ran QC. Placed analyzer and tested connection. Analyzer is operational. | Reloaded software | SWO-551696 | Defect | SW | RMED SW | RMED SW |  | Boot failed | Arrived onsite and found analyzer not booting into system software. Reloaded software without Roma and databases. Analyzer powered up to green. Loaded setup from duplicate analyzer. Named NICU. Ran QC. Placed analyzer and tested connection. Analyzer is operational.
Yes | Visit | Complaint | RAME | Received message fatal error occurred when shutdown analyzer will not turn on anymore. | Machine trys to boot, but goes to black screen. | Replaced CF card. Reinstalled solution pack and sensor cassette. Replaced customer settings to the best of my ability with Bob's assistance. | SWO-551699 | [5 empty] | Received message fatal error occurred when shutdown analyzer will not turn on anymore. | Machine trys to boot, but goes to black screen.
No | Visit | Complaint | RCHN | glu llac定标失败 | glu llac定标失败，需要借件903-581一个 |  | SWO-551151 | Defect | HW | Basic Unit | PCB, Sensor interface | Glu/Lac | glu llac calibration failure | glu llac calibration failed, need to borrow 903-581 one
Yes | Visit | Complaint | RCHN | 间歇性显示屏一直在闪烁 | 间歇性显示屏一直在闪烁，借910-342和910-285各一个 |  | SWO-551079 | [5 empty] | Intermittent display is flashing all the time | Intermittent displays are flashing all the time, with one each of the 910-342 and 910-285
No | Visit | Complaint | RCHN | CO2定标不过 | CO2定标不过，借891-051*5 |  | SWO-551162 | Defect | HW | Basic Unit | SC contact |  | CO2 calibration is not the | CO2 calibration, however, borrows 891-051 x 5
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 质控po2失败，0589 | 更换试剂包后修复 | SWO-551840 | [5 empty] | Test card, reagent package failure | QC po2 failed, 0589
No | Visit | Complaint | RCHN | 处理包无法识别 | 处理包无法识别，借903-036一个 |  | SWO-551159 | Defect | Consumables | Solution Pack | Chip |  | Processing package is not recognized | Processing package is not recognized, borrow 903-036 one
No | Visit | Complaint | RLTD | Just restarted analyser but not working, saying calibration required. | Advised the customer how to carry out membrane replacements and general Analyser check. | Analyser has passed calibrations and QC's. | SWO-551402 | User Error | Training | Training | Training |  | Just restarted analyser but not working, saying calibration required. | Advised the customer how to carry out membrane replacements and general Analyser check.
No | Visit | Complaint | RAME | OXI error 0688 tHb too low for OXI calculation and error 0581 OXI spectrum mismatch. | OXO error 0688 tHb too low for OXI calculation and error 0581 spectrum mismatch when running autocheck 3+. | Cleaned hemolyzer and replaced the OXI module pinch valve tubing. | SWO-552275 | Defect | HW | Oxi module | Pinch valve tube | 688; 581 | OXI error 0688 tHb too low for OXI calculation and error 0581 OXI spectrum mismatch. | OXO error 0688 tHb too low for OXI calculation and error 0581 spectrum mismatch when running autocheck 3+.
No | Visit | Complaint | RIND | Clot in Chamber | Checked and found that clot in K chamber and PH Chamber. | Manually Cleaned the measuring chamber and cleaning process as done and given an Successful 2 Point Calibration. | SWO-554508 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway | Clot | Clot in Chamber | Checked and found that clot in K chamber and PH Chamber.
No | Visit | Complaint | RIND | Service Visit. | Calibration found OK, already replace the membrane of Na and reference. Need to run the radiometer QC and observe. | Run the radiometer QC of Lot No. 0114 having expiry 04 $2020, the measurement of QC lie in the range and some parameter hit the target the value, QC measurement successful, for user satisfaction same lot run on the other department same model radiometer analyzer (CTVS & Main ICU), measurement comes successfully for electrolytes. Analyzer handed over to staff in well working condition. | SWO-554472 | No fault found | No fault found | No fault found | No fault found | QC | Service Visit. | Calibration found OK, already replace the membrane of Na and reference. Need to run the radiometer QC and observe.
No | Visit | Complaint | RDEU | Analysator ohne Funktion | Fehler: Analysator ohne Funktion (schwarzer Bildschirm) -->Anwender hat selbständig schon die Sicherungen des Netzteils überprüft, eine andere Steckdose, wie auch ein anderes Stromkabel ausprobiert, jedoch ohne Erfolg Fehlersuche --> Netzteil, Sicherungen, Hauptplatine, CPU-Unit überprüft --> Netzteil getauscht --> Fehler weiterhin vorhanden --> Netzteil wieder zurückgebaut --> Netzteil funktioniert , aber eine LED auf der Hauptlatine leuchtet nicht --> Sicherung getauscht, LED auf der Hauptlatine leuchtet immer noch nicht --> vermutlich Hauptplatine oder CPU-Unit defekt | vermutlich Hauptlatine oder CPU-Unit defekt --> beide Ersatzteile nicht im Bestand des Technikers --> Ersatzteile müssen bestellt werden! | SWO-551314 | Revisit required | HW | DMS_PC | CPU | Black screen; mail: Ejner: want failing screen/CPU back | Analyzer without function | Error: Analyzer without function (black screen) -->User has already checked the fuses of the power supply on its own, tried another socket, as well as another power cable, but to no avail Troubleshooting --> Power supply, fuses, motherboard, CPU unit checked --> Power supply replaced --> Error still present --> Power supply restored --> Power supply works, but an LED on the main latin does not light --> fuse exchanged, LED on the Main latin still does not glow --> probably motherboard or CPU unit defective
No | Visit | Complaint | RDEU | Remote Support Error, Analysator verlangt Neuinstallation von Win7 | Remote Support Error, der Analysator verlangt eine Neuinstallation von Win7 | Gerät geprüft, Software Win7 und 6.18MR1 neu installiert, Konfig. eingespielt, Permanent Keys installiert, Funktionskontrolle ok | SWO-552063 | Defect | SW | Windows | Windows | Remote Support Error | Remote Support Error, Analyzer Requests Reinstallation of Win7 | Remote Support Error, the analyzer requires a new installation of Win7
No | Visit | Complaint | RPOL | Błąd odczytu kasety sensorowej | Błąd odczytu kasety odczynnikowej | Sprawdzenie aparatu. | SWO-551311 | Adjustment | HW | Storage Module | Barcode reader, storage wheel |  | Sensor Cassette Read error | Reagent Cassette read error
No | Visit | Complaint | RLTD | Please call John on 0208 934 2052 | Blockage found to be in inlet connector rub worn out. | Replaced inlet connector tested pressure all okay | SWO-552370 | Defect | HW | Auto Inlet module | connector gasket | 1342; leak detected | Please call John on 0208 934 2052 | Blockage found to be in inlet connector rub worn out.
No | Visit | Complaint | RLTD | The pO2 channel is failing both calibrations and QCs. Georgina 07884273427 | Over the weekend the 1200 test cassette which was on the ED analyser was put on the spare analyser and it worked on that one. A 300 test cassette was put on the ED analyser yesterday and other than the errors you would expect to see while a new cassette was stabilising this also was ok when I arrived. | I ran the complete pO2 test on both analysers and with the 1200 test cassette in the Spare analyser and the 300test cassette in the ED analyser they both passed, the 1200 test cassette was near the lower limit for the zpO2_Energy parameter however. I swapped the cassettes over and ran the complete pO2 test again and this time the 1200 test cassette failed the zpO2_Energy parameter when it was on the ED analyser and the 300 test cassette gave similar values on both analysers. It looks to be a problem with the sensor cassette rather than the analyser. | SWO-552747 | Assisted Replacement | SC | SC | pO2 |  | The pO2 channel is failing both calibrations and QCs. Georgina 07884273427 | Over the weekend the 1200 test cassette which was on the ED analyser was put on the spare analyser and it worked on that one. A 300 test cassette was put on the ED analyser yesterday and other than the errors you would expect to see while a new cassette was stabilising this also was ok when I arrived.
No | Visit | Complaint | RDEU | Anwender berichtet, dass der Bildschirm flackert un dnicht ruhig in seiner Darstellung ist. | Nur der Bildschirmschoner flackert | Anwendern erklärt das es keine Fehlfunktion ist, Softwareupdate durchgeführt | SWO-552699 | Defect | SW | RMED SW | RMED SW | screen saver flickers | User reports that the screen flickers and is not quiet in its display. | Only the screen saver flickers
No | Visit | Complaint | RLTD | Keeps asking to change SP and SC Sidki AKay | Keeps asking to change SP and SC | Reflushed system found restriction in flow path flushed out clot. Run service programs failed vacuum pressure test. Replaced inlet gasket customer part now passing test. all other test ok. Calibration and QC ok. | SWO-552724 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway | Clot | Keeps asking to change SP and SC Sidki AKay | Keeps asking to change SP and SC
No | Visit | Complaint | RLTD | Keeps asking to replace SP and SC Sidki Akay 07913 642189 | Clot detected in oxi module and inlet module. pump calibration error. | Flushed system found restriction in flow path cleared clot. Run service test programs failing pump calibration replaced SP and run test several times passing ok. All otertest ok. Run calibation and QC passed ok. | SWO-552725 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway | Clot | Keeps asking to replace SP and SC Sidki Akay 07913 642189 | Clot detected in oxi module and inlet module. pump calibration error.
No | Visit | Complaint | RCHN | PO2定标偏高，接口板腐蚀 | PO2定标偏高，接口板腐蚀 | 更换备用的902-972,903-334后修复 by jesson | SWO-551359 | [5 empty] | PO2 calibration is high, interface plate corrosion | PO2 calibration is high, interface plate corrosion
No | Visit | Complaint | RCHN | 试剂包门栓断裂 | 试剂包门栓断裂，借902-975一个 | 更换902-975后修复 by jesson | SWO-551230 | [5 empty] | Reagent door bolt break | Reagent pack door bolt fracture, borrow 902-975 one
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | hct值无法被测量 | 更换测试卡后修复 | SWO-552602 | [5 empty] | Test card, reagent package failure | hct value cannot be measured
Yes | Visit | Complaint | RCHN | 夹紧阀管子破裂 | 夹紧阀管子破裂（未更换备件） | 更换夹紧阀管子后修复 by jesson | SWO-552049 | [5 empty] | Clamping valve pipe crack | Clamping valve pipe rupture (spare parts not replaced)
No | Visit | Complaint | RCHN | 测试卡的针角断裂 | 测试卡的针角断裂，断裂部位：teethn 未更换备件 | 更换测试卡后修复 by jesson | SWO-551868 | Defect | Consumables | Test Kit | Cartridge, Broken | Teeth | The needle angle of the test card is broken | Needle angle of test card break, fracture site: teethn spare parts not replaced
No | Visit | Complaint | RCHN | 内置芯片错误 | 内置芯片错误，借903-624一个 |  | SWO-551861 | [5 empty] | Built-in chip errors | Built-in chip error, borrow 903-624 one
No | Visit | Complaint | RDEU | Anwender (Labor) sehr unzufrieden über Stabilität des Analysators. | KAL-Driften Elektrolyte 0529 Inlet Liquid Sensor Error | Allgemeine Überprüfung - Spülung Lufteinlass im Fluidic Modul - Messkammern und Kontakte gereinigt - Erneuerung pH Dichtring (fest geklebt) - Erneuerung Referenz Elektrode - Membranierungen Elektrolyte (27.5. fällig) - LS Justage, Pumpenkalibration und Funktionsüberprüfung etc. | SWO-552277 | Assisted Replacement | Electrode | Electrode | reference |  | User (laboratory) very dissatisfied about the stability of the analyzer. | KAL Driftelectrodes Electrolytes 0529 Inlet Liquid Sensor Error
No | Visit | Complaint | RIND | Leakage from gasket. | Found the clot in gasket as well as waste tubing then clean the gasket and waste tubing then run the instrument its working good. | Found the clot in gasket as well as waste tubing then clean the gasket and waste tubing then run the instrument its working good | SWO-552170 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway | Clot | Leakage from gasket. | Found the clot in gasket as well as waste tubing then clean the gasket and waste tubing then run the instrument its working good.
No | Visit | Complaint | RDEU | 20.05.2019 Herr Rudolph meldet ein falsch zu hohes TnT (AQT 90: 0,39 vs 0,016 Kliniklabor), | TnT-Messwert bei einem Patienten gegenüber Messwert aus Labor einmalig nach oben abweichend Probe für Messung im Labor wurde laut Aussage vom Notarzt vor Transport in die Klinik abgenommen | Daten abgefragt Servicedump erstellt mehrmalige Systemreinigung Arbeiten am Analysator laut Checkliste durchgeführt ==> u.a. Reinigung mit Staubsauger, Wash unit, Nadelwascheinheit und Probennadel gereinigt Blankmessungen durchgeführt und ausgewertet ==> i.O. | SWO-552480 | Cleaned | HW | Wet section | Wet section | discr. Results | 20.05.2019 Mr Rudolph reports a falsely too high TnT (AQT 90: 0.39 vs 0.016 clinic laboratory), | TnT reading in one patient differs from laboratory reading once upwards Sample for measurement in the laboratory was taken by the emergency doctor before transport to the clinic, according to the statement
No | Visit | Complaint | RFRA | La machine plante avec un écran bleu. Pas de diagnostic plus précis pour l'instant car la connexion VPN ne fonctionne pas. Analyse des logs windows à effectuer. | Arret de l'instrument | remplacement ecran + tactile dans le cadre de la garantie calibration ecran tactile Contact FSE Modification paramétrage du controle système relance analyseur test lecture code barre ok essai Ok | SWO-552073 | Defect | HW | Analyzer Control Module | Display |  | The machine plants with a blue screen. No more accurate diagnosis at the moment because the VPN connection does not work. Analysis of windows logs to be performed. | Stopping the instrument
No | Visit | Complaint | RDEU | Spontane HCT Mess-Fehler | Kontrollbesuch wegen HCT Fehler | Seit 10.5.2019 ist der HCT Fehler 1260, 1223,1264 nicht mehr aufgetereten. Backupgerät AQT $0137N004 läuft ohne Fehler. Alle Funktionen und QC o.k,. Dialyse misst jetzt nur noch auf dem Backupgerät AQT $0137N004. Servicedump und Data Log erstellt. | SWO-552840 | No fault found | No fault found | No fault found | No fault found | HCT | Spontaneous HCT measurement errors | Verification visit due to HCT error
No | Visit | Complaint | RDEU | 07.05.2019; Fr. Bernauer; Tel.: 089529921; Probenabbrüche 08.05.2019 Sr. Martini +4989529921 meldet das weiterhin Probenabbrüche bestehen. 16.05.2019, 17.10Uhr, Frau Joana, Tel: 089524145: weiterhin Probenfehler | Das Gerät bricht öfter die Proben mit Fehler 1263 + 1264 ab. | Die HCT Zelle mit Hypochlorid gespült, Pumpenkopf 2 und 4 erneuert, Nadel ausgebaut und durchgespült, Schwarz weißen Abfallfilter hinter der HCT Zelle erneuert, Jahreswartung Kit verbaut, 6xBlankmessungen vorgenommen, Systemsreinigung und 3xLQC MCC2 und 3xLQC CRP2 vorgenommen, AQT90 ok. | SWO-552509 | Cleaned | HW | Aspiration Module | Needle Wash Unit | 1263; 1264; | 07.05.2019; Ms. Bernauer; Tel.: 089529921; Sample aborts 08.05.2019 Sr. Martini +4989529921 reports that there are still sample cancellations. 16.05.2019, 17.10h, Mrs. Joana, Tel: 089524145: continue sample error | The device cancels the samples more often with error 1263 + 1264.
Yes | Visit | Complaint | RPOL | Niedziałający skaner kodów kreskowych, wcześniejsza wymiana samego skanera nie przyniosła rezultatów. | Uszkodzony skaner kodów kreskowych, wymiana samego skanera nie przyniosła rezultatów. | Wymiana kompletnego ekranu dotykowego wraz ze skanerem kodów kreskowych, kalibracja ekranu dotykowego, testowanie aparatu. | SWO-552024 | [5 empty] | A non-working barcode scanner, an earlier replacement of the scanner itself did not produce results. | A damaged barcode scanner, replacing the scanner itself did not produce results.
No | Visit | Complaint | RKOR | PCO2 CAL DATA ERROR | replace sensor cassette , flush, fluid line clean, cal data check, gasket module check | replace sensor cassette , flush, fluid line clean, cal data check, gasket module check | SWO-554631 | Assisted Replacement | SC | SC | pCO2 | cal | PCO2 CAL DATA ERROR | replace sensor cassette , flush, fluid line clean, cal data check, gasket module check
No | Visit | Complaint | RLTD | Katy 01914453195 | Sample mixer not working | On inspection the cable connection was found to be loose. Cable re seated and analyser powered on. Sample mixer now working. | SWO-552937 | Adjustment | HW | Sample mixer | Cable |  | Katy 01914453195 | Sample mixer not working
No | Visit | Complaint | RLTD | Omed 02089675480 | Inlet module damaged misuse locking pin broken. | New inlet module fitted and calibrated and test operation ok. Auto Inlet module 903-338 Cost £376+vat PO LNW198339. New SC fitted as the one in analyser has dried out due to the broken inlet module. Calibration QC ok. | SWO-553218 | User Error | HW | Auto Inlet module | Auto Inlet module |  | Omed 02089675480 | Inlet module damaged misuse locking pin broken.
No | Visit | Complaint | RLTD | Dr Ellie Dawkins on 0208 8872000 Bleep 184 Calcium readings to low | Calcium readings to low on some samples compared to the Lab results EG 0.7 | Run a arterial adult syring sample through the ICU ABL90 R0183N0035 2 times and got these results. ICU TEST1 1.11 Ca TEST2 1.09 Ca Run same sample through SCBU ABL90 R0184N0011 3 time and got the same results. SCBU TEST1 1.11ca TEST2 1.10 Ca TEST 3 1.10 Ca Results checked out ok. Also at the time me and Sidki Akay Clinical medical manager watched the staff run 4 capillary sample and they got Ca results of around the same 1.11 Ca. analysers seems to be working ok. Calibration QC ok. | SWO-553219 | No fault found | Sample handling | Sample handling | pre analytical | Ca; Clot? | Dr Ellie Dawkins on 0208 8872000 Bleep 184 Calcium readings to low | Calcium readings to low on some samples compared to the Lab results EG 0.7
No | Visit | Complaint | RIND | Calibration failure during 2 point calibration | Check System | Check System, check calibration data, flush the SC and clean the waste line , found clot in the Sc which was removed by flushing and again given 2 point calibration , all parameters now green. | SWO-554695 | [5 empty] | Calibration failure during 2 point calibration | Check System
No | Visit | Complaint | RCHN | 样本定标检查终止 | 样本定标检查终止，在维修菜单检查液路，发现V2检测不通过 借902-977一个 | 更换902-977后修复 by jesson | SWO-551360 | Defect | HW | Wet section | Valves | 3412-FTC-103 | Sample calibration check terminated | Sample calibration check terminated, in the repair menu check the liquid road, found that V2 detection did not pass the loan 902-977 one
No | 49 | Y | RCZE | Blank |  |  | D629D629A438BB5A | Defect | HW | Autocheck module | Autocheck module |  | Blank
No | Visit | Complaint | RIND | Barcode scanner not working | Found barcode Scanner is not scanning the consumables barcode. | Open the barcode unit . Clean the barcode unit. Fix it. Turn on the machine. | SWO-554625 | Cleaned | HW | Basic Unit | Barcode reader |  | Barcode scanner not working | Found barcode Scanner is not scanning the consumables barcode.
No | Visit | Complaint | RIND | Barcode reader not detecting barcode | Barcode reader not detecting barcode | Checked the barcode reader , it is not detecting the barcode , turned off the instrument , checked the connections , cleaned it and fixed back , checked it again still it is not detecting the barcode , need to replace the barcode reader . | SWO-553490 | Revisit required | HW | Basic Unit | Barcode reader |  | Barcode reader not detecting barcode | Barcode reader not detecting barcode
No | Visit | Complaint | RDEU | pCO2 Kal fehler | Elektrode wurde mehrfach membraniert. Elektrodenkörper gebrochen. | Elektrode ausgetauscht. Kal und QC in Ordnung Handhabungsfehler. | SWO-554239 | User Error | Electrode | Electrode | pCO2 |  | pCO2 Cal error | Electrode has been membraneed several times. electrode body broken.
No | Visit | Complaint | RCHN | 定标飘移出界 | 定标飘移出界 | 自行更换942-064后修复 by jesson ' | SWO-553550 | Assisted Replacement | Electrode | Membrane | pO2 | 376 | Marked drift out of bounds | Marked drift out of bounds
No | Visit | Complaint | RCHN | 定标失败，报错 1024 | 定标失败，报错 1024,需要借件903-364一个 | 更换903-364后修复 by jesson | SWO-551855 | Defect | HW | Process Module | Optical Unit | 1024 | Bid failure, error 1024 | Failed bid, error 1024, need to borrow 903-364 one
No | Visit | Complaint | RPOL | Błędy kalibracji i kontroli jakości dla pCO2. | Serwis stwierdza problem z elektrodą pCO2. | Wymiana elektrody i membrany pCO2, kalibracja, pomiar QC. | SWO-552939 | Assisted Replacement | Electrode | Electrode | pCO2 | QC | Calibration and quality control errors for pCO2. | The service finds a problem with the pCO2 electrode.
No | Visit | Complaint | RDEU | Anruf Hr. Wiedig LP Erkennung scheitert => Reinstallation erfolgreich Kontaktfeder verbogen | LP Chip wird zeitweise nicht erkannt | Austausch Solution Pack ID PCB Geräteneustart i.O. | SWO-554049 | Defect | HW | Basic Unit | PCB, solution pack | 1134 | Call Mr. Wiedig LP Detection fails => Reinstallation successful Contact spring bent | LP chip is temporarily not recognized
No | Visit | Complaint | RDEU | 21.05.2015 15:30 H. Buckner +4921514527214: QC-Fehler pCO2 23.05.2019 08:44Uhr Herr Buckner tel:+4921513341772 : Der Fehler wird wieder angezeigt, ein Techniker soll kommen | Seit der Erneuerung des LP am 20.05 wird pCO2 in der QC Level A zu hoch gemessen. | Mitgebrachtes LP installiert, Startphase durchgeführt, QC's durchgeführt - ok. Defekte Betriebsmittel Reklamiert. | SWO-553760 | Assisted Replacement | SP | SP | pCO2 | QC A too high | 21.05.2015 15:30 H. Buckner +4921514527214: QC error pCO2 23.05.2019 08:44OPM Mr Buckner tel:+4921513341772 : The error is displayed again, a technician should come | Since the renewal of the LP on 20.05, pCO2 has been measured too high in QC Level A.
No | Visit | Complaint | RESP | Fallo de software (Data upgrade failed . status code 2) | Fallo de software, pese al reinicio del equipo. | Se entra en el archivo del equipo, se limpia la aplicación y se reinstala de nuevo la v3.4MR1. Se configura el equipo, calibra y pasan los QC´s. | SWO-553672 | Defect | SW | RMED SW | RMED SW | Status code 2 | Software failure (Data upgrade failed . status code 2) | Software failure, despite the computer restart.
No | Visit | Complaint | RDEU | 23.05.2019 Herr Ochtmann +4920360083110 meldet eine mögliche Verstopfung. 23.05.2019, Pfl. Marc, Tel: 020360083110: weiterhin wird keine Probe eingesaugt | Proben werden nicht richtig angesaugt, Proben werden abgebrochen, Einl.nadel ist verbogen, | Einl.nadel ersetzt, Probenmessungen durchgeführt, Gerät kalibriert, Fließwege gereinigt, | SWO-554064 | Defect | HW | Inlet module | Probe bend |  | 23.05.2019 Mr. Ochtmann +4920360083110 reports a possible blockage. 23.05.2019, Pfl. Marc, Tel: 020360083110: no sample is sucked in | Samples are not sucked in properly, samples are aborted, einl.needle is bent,
No | Visit | Complaint | RDEU | ramona.scholz@krankenhaus-angermuende.de | Datum steht auf 24.5.2019 | Datum korrigiert und Kunden die Anmeldung zur Korrektur in der Konfiguration erläutert Kunde bestellt neue SC80, da die installierte Kassette nach der Datumskorrektur verfallen ist | SWO-553782 | [5 empty] | ramona.scholz@krankenhaus-angermuende.de | Date is on May 24, 2019
No | Visit | Complaint | RCHN | 模拟板故障 | 模拟板故障 | 更换备用903-334后修复by tim | SWO-554760 | [5 empty] | Analog board failure | Analog board failure
No | Visit | Complaint | RCHN | 仪器运行时自动停止，显示运行超出时间 | 仪器运行时自动停止，屏幕显示运行超出时间，借910-343一个（缺货，经销商先用备用件去维修） | 更换备用机上的备件910-343后修复 by jesson | SWO-552594 | Defect | HW | Analyzer Control Module | CPU |  | The instrument stops automatically while running, showing that it is running out of time | The instrument is automatically stopped when running, the screen shows that the operation is out of time, borrows 910-343 one (out of stock, the dealer first uses spare parts to repair)
No | Visit | Complaint | RCHN | 显示器不显示 | 显示器不显示，借910-343各一个 | 更换备用件910-343后修复 by jesson | SWO-552708 | Defect | HW | Analyzer Control Module | CPU |  | The monitor does not display | The monitor does not display, borrow 910-343 each one
No | Visit | Complaint | RCHN | 报错 5684.处理盘卡盘 | 报错 5684.处理盘卡盘，需要借件944-394 15783 | 重新定标后修复 by jesson | SWO-551891 | Defect | Consumables | Test Kit | Cartridge, Broken | 5684 | False 5684.Processing Disk | False 5684.Processing disk, need to borrow 944-394 15783
No | Visit | Complaint | RCHN | 软件故障，无法进入系统 | 软件故障，无法进入系统，未更换备件 | 重新安装系统，重新定标后修复 by jesson | SWO-553981 | Defect | SW | RMED SW | RMED SW |  | Software failure, no access to the system | Software failure, no access to the system, no replacement parts
No | Visit | Complaint | RCHN | ph测量不稳定，报错0476 | ph测量不稳定，报错0476,需要借件945-614一个 | 更换945-614后修复 by jesson | SWO-552732 | Assisted Replacement | Electrode | Electrode | pH | 476 | Ph measurement instability, error 0476 | Ph measurement unstable, misreported 0476, need to borrow 945-614 one
No | Visit | Complaint | RIND | Barcode reader not detecting barcode | To replace barcode reader and to train new staff | Turned off the instrument , replaced the barcode reader with standby , checked al the connections , turned on the instrument , observed the working of barcode reader , it is working fine , trained the staff about loading QC , sample processing , loading consumables , customer happy with instrument | SWO-554586 | Defect | HW | Basic Unit | Barcode reader |  | Barcode reader not detecting barcode | To replace barcode reader and to train new staff
No | Visit | Complaint | RIND | Barcode Scanner is not workinig | Barcode Scanner is not working. Found Sensitivity of Crea is low | Open the barcode unit clean it. Turn on the analyzer . Now barcode is working fine. Change the membrane of crea , run two point calibration. Calibration pass. | SWO-554634 | Cleaned | HW | Basic Unit | Barcode reader |  | Barcode Scanner is not workinig | Barcode Scanner is not working. Found Sensitivity of Crea is low
No | Visit | Complaint | RAME | pCO2 Fails 2-point calibration for high sensitivity | Troubleshot and found that a new membrane was required. I also found that both membrane valves on the pH $BG module and El $Met module required replacement. | Replaced the pCO2 membrane, and the valves on the pH $BG module and El $Met module. | SWO-553293 | Revisit required | HW | Gas module | Flow regulator |  | pCO2 Fails 2-point calibration for high sensitivity | Troubleshot and found that a new membrane was required. I also found that both membrane valves on the pH /BG module and El /Met module required replacement.
No | Visit | Complaint | RDEU | QC häufig ausser Bereich. | Driften pH und Elektrolyte. | Referenzelektrode durch Testelektrode 81-12 ersetzt. Kalibrationen und Qualitätskontrollen in Ordnung. | SWO-554021 | Assisted Replacement | Electrode | Electrode | reference |  | QC is often out of the mix. | Drifting pH and electrolytes.
No | Visit | Complaint | RIND | Neon Lamp voltage too high. | Neon Lamp voltage too high. | Recently replaced the Oxi module, but problem is remaining same, replaced the Power supply unit and Oxi module and restarted analyzer performed the 2p calibration found Oximetry parameters are OK, instrument is under observation. | SWO-554310 | Defect | HW | Power Supply | Power Supply | Neon lamp voltage too high | Neon Lamp voltage too high. | Neon Lamp voltage too high.
No | Visit | Complaint | RDEU | Kunde wünscht eine sofortige Inbetriebnahme. Das Leihgerät ist aufgrund eines Problem mit dem FlexQ Modul defekt. | Inbetriebnahme erfoderlich. | ABL vom Magazin Dezernat Technik zum ZOM2 AWR transportiert. Gerät aufgestellt, konfiguriert nach Kundenstandart, Gerät mehrfach kalibriert, und an das Radiance System angebunden. Leihgerät in die Medizintechnik Dezernat Technik gebracht. | SWO-553617 | Revisit required | analyzer removed | Analyzer removed | Analyzer removed | FlexQ Module | Customer wants immediate commissioning. The rental device is defective due to a problem with the FlexQ module. | commissioning is essential.
No | Visit | Complaint | RPOL | W kontroli jakości oraz na próbkach pacjentów pojawia się komunikat: Niezgodność widma OXI (0581) . Błędy przy pomiarach ctHb, Bil i Crea/Urea. | Weryfikacja dzienników danych. | Wymiana płyty Wet Section Control PCB. Kalibracje i testy serwisowe. | SWO-553590 | Defect | HW | Wet section | PCB, wet section | 581 | In quality control and patient samples, a message appears: Spectrum mismatch OXI (0581). Errors when measuring ctHb, Bil and Crea/Urea. | Verification of data logs.
No | Visit | Complaint | RPOL | Zgłoszono usuwanie kalibracji CRP lot 15598 w ciągu kilku dni od skalibrowania. | Zgłoszono usuwanie aktywnej kalibracji CRP w ciągu kiku dni od skalibrowania. | Sprawdzenie historii kalibracji, ponowna kalibracja CRP LOT 15598, ważność kalibracji do 30.11.2019. | SWO-553540 | User Error | Consumables | Test Kit | Cartridge, cal |  | Removal of the CRP calibration flight 15598 has been reported within a few days of calibrating. | It was reported to remove active CRP calibration within the Kiku days after calibrating.
No | Visit | Complaint | RPOL | Aparat nie czytającej troponiny | Sprawdzenie logów, testy serwisowe | Aparat nie czyta kasetek troponiny, podczas gdy czyta d-dimmery oraz CRP. Regulacja ani wymiana skanera nie zmieniły sytuacji. W logach widnieje komunikat: ,,aktywna kalibracja TnI została usunięta”. | SWO-554092 | User Error | Consumables | Test Kit | Cartridge, cal |  | Apparatus not reading Troponin | Log check, service tests
No | Visit | Complaint | RAME | Printer very loud and customer wants replaced in warranty period | Checked for paper james or other issue. Printer just loud and seems like it's failing | Replaced printer. Check printout, good. Not loud | SWO-553925 | Defect | HW | Printer Module | Printer Module |  | Printer very loud and customer wants replaced in warranty period | Checked for paper james or other issue. Printer just loud and seems like it's failing
Yes | Visit | Complaint | RFRA | Erreur 331, pas d'aspiration d'échantillon. Malgré tout, l'appareil aspire bien les seringues. Cassette remplacé le 17 mai et pack le 20 mai, le phénomène se produisait déjà avant les remplacement des consommables. | Contrôle de la fluidique | Contrôle de la fluidique $Contrôle fichier activités $Constate erreurs echantillon abandonné et DL $Démontage et remplacement de la bay unit $Remplacement du kit de maintenance annuelle $ $Remplacement du tuyau de vidange, joint connexion, joint d'intro et tuyaux hémolyseur $sensor cassette. Ajustement DL et calibration de la pompe. Test du circuit fluidique en mode service. Test PO2 & OXI en mode service. Calibration et passage CQ. | SWO-554113 | [5 empty] | Error 331, no sample suction. Nevertheless, the device sucks the syringes well. Cassette replaced on May 17 and pack on May 20, the phenomenon was already occurring before the replacement of consumables. | Fluid control
Yes | Visit | Complaint | RSCH | Rep. Bildschirm | Bildschirm bei Auslieferung durch Verkäufer zerkratzt | Display unit ersetzt i.o | SWO-554260 | Re classified RMED | installation report |  |  | Move to OOB trending | Rep. Screen | Screen on delivery scratched by seller
No | Visit | Complaint | RCHN | 频繁定标不过 | 频繁定标不过，借902-972一个 | 故障未修复，需加借备件903-334一个，此单先关单处理 by jesson | SWO-553536 | [5 empty] | Frequent calibration is not | Frequent calibration, however, borrow 902-972 a
No | Visit | Complaint | RSCH | Notfall ABL 80 Altendorf | Luftblasenerkennungssystem gescheitert | JW Kit Ersetzt ok Zeroing durchgeführt ok Kal+ QA Zyklus durchgeführt ok Sensorkassette ersetzt ok End Kontrolle durchgeführt ok | SWO-554405 | [5 empty] | Emergency ABL 80 Altendorf | Air bubble detection system failed
No | Visit | Complaint | RCHN | 接口板腐蚀 | 接口板腐蚀，借902-972一个 | 更换902-972后修复 by jesson | SWO-552572 | [5 empty] | Interface plate corrosion | Interface plate corrosion, borrow 902-972 one
No | Visit | Complaint | RCHN | 定标灵敏度错误 | 定标灵敏度错误，借945-737一个 | 更换945-737后修复 by jesson | SWO-553527 | [5 empty] | Calibration sensitivity error | Calibration sensitivity error, borrow 945-737 one
No | Visit | Complaint | RCHN | 无法抽洗样本，报错1235 | 无法抽洗样本，报错1235，需要借件903-129一个 | 更换903-129后修复 by jesson | SWO-552623 | Defect | HW | Oxi module | Pinch valve | 1235 | Unable to wash samples, error 1235 | Unable to pump samples, reported error 1235, need to borrow 903-129 one
No | Visit | Complaint | RCHN | 溶血器温度超出范围，报错1069 | 溶血器温度超出范围，报错1069,需要借件903-644一个 | 更换903-644一个正常。by jesson | SWO-553514 | Defect | HW | Wet section | PCB, wet section | 1069 | Hemolysisr temperature out of range, error 1069 | Hemolytic temperature out of range, reported error 1069, need to borrow 903-644 one
No | Visit | Complaint | RCHN | 定标数据丢失 | 定标数据丢失,需要借件944-258 15940 | 重新定标正常。by jesson | SWO-553512 | Assisted Replacement | Consumables | Test Kit | Cartridge, cal |  | Target data loss | Bid data loss, need to borrow 944-258 15940
No | Visit | Complaint | RCHN | 进样门关不上，报错5784 | 进样门关不上，报错5784，未更换备件 | 调节进样门位置后修复 by jesson | SWO-554028 | Adjustment | HW | Cover and Chassis | Door, Inlet | 5784 | The entrance door is not closed, the wrong 5784 | The sample door is not closed, 5784 error is reported, spare parts are not replaced
No | Visit | Complaint | RIND | During sample run Parameters not showing values | Check System | Check calibration, run the sample , Value of pO2 and pH is not given when run same sample again pH given bt pO2 not this problem coming frequently so change the SC also bt same problem persisit again replace analyser also but the problem is not resolved tried all the possible ways to resolve the issue. | SWO-554698 | [5 empty] | During sample run Parameters not showing values | Check System
Yes | Visit | Complaint | RIND | Need to replace the membrane. | pO2 Parameter is showing yellow in the parameter bar, Need to replace the pO2 membrane. | Do the protein removal and cleaning of the analyzer, hold the analyzer, replace the pO2 membrane successfully, recorded replacement, restart the analyzer and wait for auto 2-point calibration, successfully done, analyzer handed over to staff in well working condition. | SWO-554570 | [5 empty] | Need to replace the membrane. | pO2 Parameter is showing yellow in the parameter bar, Need to replace the pO2 membrane.
No | Visit | Complaint | RAME | Analyzer intermittently fails calibration for high pCO2 sensitivity. | Ran the gas module check in service and found that the Gas 1 flow regulator was very unstable. | Placed the Gas 1 Flow Regulator on order. | SWO-554408 | Defect | HW | Gas module | Flow regulator |  | Analyzer intermittently fails calibration for high pCO2 sensitivity. | Ran the gas module check in service and found that the Gas 1 flow regulator was very unstable.
No | Visit | Complaint | RIND | N + reading higher side | N + reading higher side | Done Protein removal and decontamination, guide customer how to cleaned analyser | SWO-554585 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway | Clot; Na | N + reading higher side | N + reading higher side
No | Visit | Complaint | RDEU | 23.05.2019 8:00 Fr. Schupfner +499471705456: Geräteausfall wg. falsch eingesetzter Probenkassette | Testkassette wurde falsch herum in Gerät eingelegt. Kundenverschulden Auftrag kam von Herrn Brill | Gerät geöffnet. Testkassette aus Gerät entfernt. Systemüberprüfung ok. LQC TnT und CKMB durchgeführt. | SWO-554863 | User Error | Consumables | Test Kit | Cartridge | Inserted upside down | 23.05.2019 8:00 Fr. Schupfner +499471705456: Device failure wg. incorrectly inserted sample cassette | Test cartridge was inserted incorrectly around in device. Customer debt Order came from Mr. Brill
No | Visit | Complaint | RPOL | Wyięta sonda aspracyjna. | Serwis stwierdza wygięcie sondy aspiracyjnej. | Wymiana sondy aspiracyjnej, testowanie aparatu. | SWO-554200 | Defect | HW | Auto Inlet module | Probe bend |  | Wyi'ta probe aspracyjna. | The service concludes the bending of the aspiration probe.
Yes | Visit | Complaint | RDEU | Krumme Nadel | Krumme Nadel , Thb Level C nicht bestanden , Thb nicht bestanden | Nadel ersetzt, Startphase bestanden, Kal und QC bestanden | SWO-554849 | [5 empty] | Krumme Needle | Krumme Needle, Thb Level C failed, Thb failed
No | Visit | Complaint | RCHN | 频繁定标不过 | 故障未修复，加借备件903-334一个 | 更换903-334，902-972各一个正常。by jesson | SWO-554106 | [5 empty] | Frequent calibration is not | Failure not repaired, spare parts 903-334 one
No | Visit | Complaint | RCHN | TNI定标中止，报的5257 | TNI定标中止，报错5257，借944-212（15883）一张 | 重新定标正常。by jesson | SWO-553979 | Cleaned | HW | Process Module | Process Unit | 5257 | TNI calibration suspended, reported 5257 | TNI bid suspension, error 5257, loan 944-212 (15883) one
No | Visit | Complaint | RCHN | 数据无法传输到电脑 | 数据无法传输到电脑 | 软件升级。 | SWO-554953 | [5 empty] | Data cannot be transferred to your computer | Data cannot be transferred to your computer
No | Visit | Complaint | RCHN | 接口板氧化，传感器电路错误 | 接口板氧化，传感器电路错误,借902-972一个。 |  | SWO-554678 | [5 empty] | Interface board oxidation, sensor circuit error | Interface board oxidation, sensor circuit error, borrow 902-972 one.
No | Visit | Complaint | RCHN | 电源损坏 | 电源损坏,借903-663一个。 |  | SWO-554708 | [5 empty] | Power damage | Power supply damage, borrow 903-663 one.
No | Visit | Complaint | RCHN | 假阳性故障判断 | 清洁处理盘 | 机器做空白实验，排查机器故障by tim | SWO-557384 | Cleaned | HW | Process Module | Process Wheel | DR | False positive fault judgment | Cleaning the handling disc
No | Visit | Complaint | RDEU | Anruf Hr. Wiedig Gerät meldet Leckage LP wurde gestern gewechselt, dabei aber nicht richtig aktiviert Hr. Wiedig hat LP nochmal neu eingesetzt aber ohne Besserung | wieder Leckage Fehlermeldung 1342 QC Fehler Oximetrie | Überprüfung Quetschventil und Verschlauchung => Lösungsaustritt am Quetschventilschlauch, Ventil schaltet schwergängig => Austausch Ventil, Schlauch neu angebracht Durchführung aller Fluid Transport Programme Startphase i.O. | SWO-555371 | Defect | HW | Oxi module | Pinch valve | Leak detected; 1342 | Call Mr. Wiedig Device reports leakage LP was changed yesterday, but not activated correctly Mr. Wiedig has re-installed LP but without improvement | again leakage error message 1342 QC Error Oximetry
No | Visit | Complaint | RDEU | der Analsator startet nicht, Status Code 2 | Status Code 2 | Gerät geprüft, Roma gesichert und überschrieben, Win7 und Analysator SW neu installiert, Ersatzteile bestellt | SWO-555609 | Revisit required | HW | DMS_PC | CF card | Status code 2 | the analsator does not start, status code 2 | Status Code 2
Yes | 9 | Y | RCZE | Blank |  |  | D633ADD1543E25C2 | Re classified RMED | Request |  |  | See mail from JIRVO with title "RE: D633ADD1543E25C2" | Blank
No | Visit | Complaint | RCHN | 接口板腐蚀 | 接口板腐蚀，借902-972一个 |  | SWO-554852 | [5 empty] | Interface plate corrosion | Interface plate corrosion, borrow 902-972 one
No | Visit | Complaint | RIND | Service Visit. | Value of the Na not satisfactory, need to install the new SC. | Install the new SC, not successful, bubbles is generating automatically during calibration, SC has failed. | SWO-556999 | [5 empty] | Service Visit. | Value of the Na not satisfactory, need to install the new SC.
No | Visit | Complaint | RCHN | 测试卡下层温度超范围，报错1065 | 测试卡下层温度超范围，报错1065，借903-581一个 |  | SWO-554794 | Defect | HW | Basic Unit | PCB, Sensor interface | 1065 | The test card lower temperature exceeds the range, the error 1065 | The test card lower temperature exceeds the range, the error 1065, borrows 903-581 one
No | Visit | Complaint | RCHN | 提示更换试剂包 | 提示更换试剂包（未更换备件） | 管路有凝块，重新管路 by jesson | SWO-566504 | Defect | HW | Basic Unit | PM kit | Clot | Prompt to replace the reagent kit | Prompt replacement kit (no replacement parts)
Yes | Visit | Complaint | RCHN | 血氧单元硬件错误，报错1034 | 血氧单元硬件错误，报错1034，借903-197一个 |  | SWO-554765 | [5 empty] | Blood oxygen unit hardware error, report error 1034 | Blood oxygen unit hardware error, report error 1034, borrow 903-197 one
No | Visit | Complaint | RCHN | 黑屏 | 黑屏 | 重装系统by tim | SWO-554998 | Defect | SW | Windows | Windows | Black screen | Black screen | Black screen
No | Visit | Complaint | RCHN | 显示屏有黑点 | 显示屏有黑点.需要借件910-298一个 |  | SWO-554770 | Defect | HW | Analyzer Control Module | Display | Potential User Error | The display has a black dot | The display has a black dot.
No | Visit | Complaint | RIND | PH & pCO2 Calibration not passed | SC Cleaning is required | Clean SC and Reinstall, Calibration passed but PH sensitivity is to low, Analyser is under observation and need to replace SP and SC. | SWO-557370 | [5 empty] | PH & pCO2 Calibration not passed | SC Cleaning is required
No | Visit | Complaint | RPOL | Podczas pomiaru pojawiają się błędy: zbyt niska temperatura oraz wykryto powietrze w próbce. | Przegląd historii analizatora, serwis stwierdza pojawianie się błędu: wykryto powietrze w próbce podczas pomiaru krwi oraz błędu zbyt niska temperatura podczas pomiaru próbki QC. | Czyszczenie styków sprężystych kasety sensorowej, ponowna instalacja kasety, pomiar QC 3 poziomu, 2x pomiar na próbce krwi. | SWO-555128 | [5 empty] | During the measurement, errors occur: temperature too low and air detected in the sample. | Overview of the Analyzer's history, the service determines the occurrence of the error: Air was detected in the sample during blood measurement and error temperature too low during the measurement of the QC sample.
No | Visit | Complaint | RIND | K Result variation | Check System | Check system, check calibration ask to rerun the sample , run sample by me ,problem found in sample handling and running process, given training for sample handling | SWO-557819 | [5 empty] | K Result variation | Check System
No | Visit | Complaint | RDEU | 06.05.2019; Hr. Adebar; +496213833528; E'Lyte Kal Drift Fehler | Referenz Elektrode eingetroffen KAL - Driften Elektrolyte | Allgemeine Überprüfung - Erneuerung Referenz Elektrode - Funktionsüberprüfung etc. ohne Beanstandung | SWO-555051 | Assisted Replacement | Electrode | Electrode | reference | CAPA-733 | 06.05.2019; Hr. Adebar; +496213833528; E'Lyte Kal Drift Fehler | Reference electrode arrived KAL - Drifting Electrolytes
No | Visit | Complaint | RKK | インレットフロー妨害と液漏れ | インレットフロー妨害 液漏れ対応 Cl QC範囲外 | 32mmリンスラインチューブ交換 排液ラインチューブ交換と排液ニップル清掃 インレットモジュール入れ替え Cal 1溶液交換 | SWO-557369 | Defect | HW | Inlet module | Inlet module |  | Inlet flow interference and leakage | Inlet flow interference Liquid leakage response Cl QC out of range
No | Visit | Complaint | RDEU | 28.05.2019 Herr Prengel +493817022771 meldet, das er immer wieder das LP wechseln soll. | zwei Lösungspacks (HL02 und HJ02) haben keine Rinse-Lösung, alle anderen Lösungen werden fehlerfrei geördert, Test mit verfallenem SP zeigte keine Probleme | Kunde bekommt aus dem Hauptlabor neues SP keine Reparatur erforderlich | SWO-555615 | Assisted Replacement | SP | SP | FT |  | 28.05.2019 Mr. Prengel +493817022771 announces that he is to change the LP again and again. | two solution packs (HL02 and HJ02) have no rinse solution, all other solutions are error-free, Test with dilapidated SP showed no problems
No | Visit | Complaint | RAME | Barcode reader intermittently giving wrong information | Barcode reader misinterpreting barcodes. | Replaced barcode reader and edited .ini file per RMED change order. See notes. Original barcode reader serial number 902-924R2009N0756. | SWO-555819 | Defect | HW | Basic Unit | Barcode reader |  | Barcode reader intermittently giving wrong information | Barcode reader misinterpreting barcodes.
No | Visit | Complaint | RLTD | Ilyaas Hakem on 0330 4004333 ex 6672 | Performed manual flush. Soaked the fluid pathway with Hypochlorite solution for twenty minutes and flushed with water. No flow restrictions or clots were found. Pressure drawn value is at 1.25 or very close when repeated. Replaced all tubings and gaskets. Pressure drawn value is now closer to 1, but pump calibration is failing as it is too low. | Replaced SP. Ran a successful test sample. Short sample mode has been removed from the syringe mode and is now the standard 65uL mode. Analyser is ready for use. | SWO-555727 | Revisit required | HW | Basic Unit | Sensor Bay Unit | Clot | Ilyaas Hakem on 0330 4004333 ex 6672 | Performed manual flush. Soaked the fluid pathway with Hypochlorite solution for twenty minutes and flushed with water. No flow restrictions or clots were found. Pressure drawn value is at 1.25 or very close when repeated. Replaced all tubings and gaskets. Pressure drawn value is now closer to 1, but pump calibration is failing as it is too low.
No | Visit | Complaint | RDEU | der Analsator startet nicht, Status Code 2 | Status Code 2 | CF Card ausgetauscht, Software installiert, CPU ersetzt, Win7 und Analysator SW neu installiert, Gerät neu konfiguriert, anschluß ans Hausnetzwerk, SC 90 erneuert, Funktionskontrolle ok | SWO-555610 | Defect | HW | DMS_PC | CF card | Status code 2 | the analsator does not start, status code 2 | Status Code 2
No | Visit | Complaint | RCHN | 废液口堵塞 | 废液口堵塞 | 清洁废液口修复by tim | SWO-555462 | [5 empty] | Waste mouth blockage | Waste mouth blockage
No | Visit | Complaint | RCHN | 显示屏后盖螺丝老化 | 显示屏后盖螺丝老化，借910-298一个 | 收到备件有质量问题，需重新加借，此单关单by tim | SWO-555352 | Revisit required | HW | Analyzer Control Module | Frame, Display |  | Screen screw aging | Display back cover screw aging, borrow 910-298 one
No | Visit | Complaint | RCHN | 进样舱门损坏 | 进样舱门损坏，借886-638一个 | 更换886-638后发现故障依旧，需要重新借件by tim | SWO-554989 | Revisit required | HW | Cover and Chassis | Door, Inlet | Top bracket | Damage to the hatch | Damage to the hatch, borrow 886-638
No | Visit | Complaint | RCHN | 提示气瓶空，气压不足 | 提示气瓶空，气压不足 | 检查各气瓶后，更换气瓶后解决by tim | SWO-555436 | Assisted Replacement | Consumables | Gas | Cal |  | Suggestthat gas cylinder empty, insufficient air pressure | Suggestthat gas cylinder empty, insufficient air pressure
No | Visit | Complaint | RIND | QC vials are not detected in Autocheck | QC vials not detected in Auto check module | Customer loaded the QC ampules in auto check module but instrument is not detecting them , checked the auto check in service mode , did barcode reading and observed that barcode is not detecting the vials , opened the auto check module cleaned the laser scan unit , restarted the instrument and checked the alignments , alignments at are fine but still the instrument is not detecting the auto check vials , need to replace the laser scan unit . | SWO-557529 | Revisit required | HW | Autocheck module | Barcode reader |  | QC vials are not detected in Autocheck | QC vials not detected in Auto check module
No | Visit | Complaint | RPAC | Steve called 6237 1211 software corrupted after power outage | Machine will not start up - Software issue. | Attend site and examine machine. Discussed with customer and machine failed safter a scheduled power outage as the machine was not shutdown. Windows software is ok, but ABL Software and database are corrupt. Tried to replace database without luck. Renamed ABL700 folder to ABL700 NSP. Reloaded ABL Software with version 6.18 MR1. Restart machine and reload setup from backup. Restart agaian and allow to stratup. Configure machine and setup QC schedule and reload QC's. Perform muliple 1 and 2 point calibrations, perform Protien Clean and also tHb calibration. Run 2 levels of QC, OK. | SWO-556716 | Defect | SW | RMED SW | RMED SW | Power outage | Steve called 6237 1211 software corrupted after power outage | Machine will not start up - Software issue.
No | Visit | Complaint | RIND | Service Visit. | Check the status of the calibration, need to replace the reference and pCO2 membrane. | Hold the analyzer, remove the electrode of reference, pH and pCO2, clean the chamber, replace the new membrane in reference and pCO2 as per protocol, recorded replacement, successfully done, restart the analyzer and wait for auto 2 point calibration, successfully done, analyzer handed over to staff in well working condition. | SWO-557397 | User Error, FSE | Electrode | Membrane | Lifetime exceeded |  | Service Visit. | Check the status of the calibration, need to replace the reference and pCO2 membrane.
No | Visit | Complaint | RAME | Analyzer intermittently fails calibration for high pCO2 sensitivity. | Troubleshot to a bad gas 1 flow regulator. | Replaced the gas 1 flow regulator | SWO-555970 | Defect | HW | Gas module | Gas valve |  | Analyzer intermittently fails calibration for high pCO2 sensitivity. | Troubleshot to a bad gas 1 flow regulator.
No | Visit | Complaint | RKOR | pCO2 QC error | . | Electrode replace. Liquid flow line clean. Membrane clean. | SWO-557972 | Cleaned | Fluidic pathway | pH/BG Module | pCO2 |  | pCO2 QC error | .
No | Visit | Complaint | RKOR | Lac QC low result | . | Electrode clean. Liquid flow line clean. Membrane clean. | SWO-557973 | Cleaned | Fluidic pathway | El/Met Module | Lac |  | Lac QC low result | .
No | Visit | Complaint | RPOL | Częste problemy z analizatorem, ciagle płukania. Częste interwencje wymagane. | Testy serwisowe. Sprawdzenie uszczelki wlotu. | Wymiana uszczelki wlotu. Przepłukanie manualne podchlorynem. Zamontowanie wspornika pakietu odczynnikowego. Testy serwisowe ok. | SWO-555549 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | Common problems with the analysing, continuous rinse. Frequent interventions required. | Service tests. Check Inlet gasket.
No | Visit | Complaint | RPOL | Aparat nie czytającej troponiny | Dokończenie prac | Reinstalacja systemu Windows XP oraz oprogramowania w wersji 8.11.300.0 na czysto. Kalibracja parametrów, testy serwisowe, konfiguracja aparatu. Konfiguracja połączenia do systemu AMMS. | SWO-556191 | Defect | SW | RMED SW | RMED SW | Barcode, Cartridge | Apparatus not reading Troponin | Completing the work
No | Visit | Complaint | RDEU | 27.05.2019 Frau Lohmann +493391394636 meldet, dass der ABL einen Vakuum Fehler zeigt. 28.05.2019/11:00 Herr Götz (MT) 03391394473 meldet den gleichen Fehler mit der Bitte um Serviceeinsatz | immer wieder Vakuum Fehler bei 0,95 Oxi Schlauch defekt | Oxi Schlauch gewechselt Service Programme nun i.O. SC gewechselt und Startup abgewartet pCO2 Empf. bei 22% -> SC erneut gewechselt | SWO-556288 | Defect | HW | Oxi module | Pinch valve tube | 1063 | 27.05.2019 Ms. Lohmann +493391394636 reports that the ABL shows a vacuum error. 28.05.2019/11:00 Mr. Götz (MT) 03391394473 reports the same error with the request for service | repeated vacuum error at 0.95 Oxi hose defective
No | Visit | Complaint | RESP | Fuga detectada en el Paq. Soluc (964). | Fuga detectada en el Paq. Soluc (964) | Se revisa el estado de registros y se extraee informe de error; así como gestión de la devolución del SP. Se revisan los gaskets y tubo de waste (sustitución del mismo). Se comprueban los diferentes programas de servicio. Se reinstala la aplicación 3.4MR1 ya que durante el proceso de arranque se corrompe durante (data upgrade failed - Status code 2). Se realiza carga de la configuración, calibración y pasar los QC´s. | SWO-556560 | Revisit required | HW | Basic Unit | PM kit | 964 | Leak detected in the Paq. Soluc (964). | Leak detected in the Paq. Soluc (964)
No | Visit | Complaint | RDEU | 29.05.2019, 11:16 Uhr ==> Anrufer Fr. Schuler, Tel.: 07315000 ==> Intervention erforderlich, Fehler 1343: Lösungen können nicht gefördert werden |  | Vacuum Error bei Eintreffen QC2 Lös nicht erkannt. Fluidic Wege überprüft. Inlet ok. SP entnommen und reinstalliert. Startphase 2Std. Kal und QC ok. | SWO-556405 | Adjustment | SP | SP | FT | 1343; SP reinstalled | 29.05.2019, 11:16 a.m. ==> Caller S.A. Ms. Schuler, tel.: 07315000 ==> Intervention required, error 1343: Solutions cannot be supported
No | Visit | Complaint | RDEU | der Analsator startet nicht, Status Code 2 29.05.2019 8:00 Fr. Nirentz +4951218943613: Nach Reparatur funktioniert der Patientendatentransfer nicht mehr | Pat. Daten werden nicht gelesen | Konfig. geprüft und angepasst, Funktionskontrolle ok | SWO-556518 | User Error, FSE | SW | HIS/LIS | HIS/LIS | Setup | the analsator does not start, status code 2 29.05.2019 8:00 Ms. Nirentz +4951218943613: After repair, patient data transfer no longer works | Pat. Data is not read
No | Visit | Complaint | RAME | Battery error | Battery Error 1559 displayed, battery cell error. Says analyzer will shutdown. | Replaced battery assembly and found percentages displayed correctly. Ran calibrations and manual QC found all levels passed. After battery showed 25% charge I disconnected the AC power and found the battery symbol displayed correctly. | SWO-555927 | Defect | HW | Battery module | Battery module | 1559 | Battery error | Battery Error 1559 displayed, battery cell error. Says analyzer will shutdown.
No | Visit | Complaint | RAME | Customer dissatisfied with amount of failures on analzyer from sensor cassettes to solution packs. | Analyzer was failing pCO2 high on internal QC A & C. Ran sensor simulator and all test passed within specifications. | Replaced sensor cassette from customers stock to correct pCO2 internal QC failure. | SWO-556507 | Assisted Replacement | SC | SC | pCO2 | QC A&C high | Customer dissatisfied with amount of failures on analzyer from sensor cassettes to solution packs. | Analyzer was failing pCO2 high on internal QC A & C. Ran sensor simulator and all test passed within specifications.
No | Visit | Complaint | RCHN | 接口板腐蚀 | 接口板腐蚀,，需要借件902-972一个 | 更换902-972一个正常。by tim | SWO-555460 | [5 empty] | Interface plate corrosion | Interface plate corrosion, need to borrow 902-972 one
No | Visit | Complaint | RDEU | Patientenbarcode wird nicht erkannt. Info kam von Hr. Maximilian Neumann. | Patientenbarcode wird nicht erkannt. | Scanner konfiguriert. Barcode wird nun gescannt. Automatisches Drucken von 2 Punkt Kalibration und QC´s deaktiviert. Kalibration und QC´s vorgenommen. | SWO-556229 | [5 empty] | Patient bar code is not recognized. Info came from Mr. Maximilian Neumann. | Patient bar code is not recognized.
No | Visit | Complaint | RCHN | Punch up故障，报错5297 | Punch up故障，报错5297，借903-031一个 | 自行更换903-031后修复by tim | SWO-556647 | Defect | HW | Storage Module | PCB, Cup transport rear | 5297 | Punch up fault, error 5297 | Punch up fault, error 5297, loan 903-031 one
No | Visit | Complaint | RCHN | 舱门破裂 | 舱门破裂，需要借件886-638、886-639各一个 |  | SWO-556072 | Adjustment | HW | Cover and Chassis | Door, Inlet | Corrective Action: Repaired by HUnter on both sides of the door with a cotton swab | The hatch is broken. | The hatch is broken and requires a loan of 886-638, 886-639 each
No | Visit | Complaint | RCHN | 加样单元运动错误，显示器右侧螺丝位破裂开缝,报错：1999 37.522 | 加样单元运动错误，显示器右侧螺丝位破裂开缝,报错：1999 37.522 | 清洁调整加样单元，胶布固定显示器后修复by tim | SWO-556144 | Cleaned | HW | Aspiration Module | Needle Bushing |  | Sample unit motion error, right screw position crack edgy crack, error: 1999 37.522 | Sample unit motion error, right screw position crack edgy crack, error: 1999 37.522
No | Visit | Complaint | RCHN | 进样舱门损坏 | 进样舱门损坏,需要重新借件886-638 | 更换886-638后发现故障依旧，需要重新借件by tim | SWO-555425 | Revisit required | HW | Cover and Chassis | Door, Inlet | Top bracket | Damage to the hatch | The sample door is damaged and needs to be re-borrowed 886-638
No | Visit | Complaint | RCHN | 进样门无法关闭 | 进样门无法关闭，借903-717一个 |  | SWO-556103 | Defect | HW | Cover and Chassis | Door, Inlet | Corrective Action: Repair by Hunter after replacing 903-717; Top bracket | The sample door cannot be closed | The sample door cannot be closed, borrow 903-717 one
No | Visit | Complaint | RCHN | 舱门关闭不好 | 舱门关闭不好，借903-717一个 |  | SWO-555966 | Defect | HW | Cover and Chassis | Door, Inlet | Corrective Action: Repair by Hunter after replacing 903-717; Top bracket | The door is not closed. | The hatch closed badly, borrow edges 903-717 one
No | Visit | Complaint | RKK | SC取付時に、内部のチューブが外れ液漏れした(SCと本体のニップルが詰まっていた) それに伴い、液漏れが発生し、モーターの異音、コネクタボードが腐食した。 また、SP取付不良が続発したため、ラッチ交換(オプトセンサーチェック済み、モルッドアチェック済み) ディスク容量がいっぱい、の表示 | 詳細の通り | 以下交換 | SWO-556743 | [5 empty] | When installing the SC, the internal tube leaked out of the liquid (the nipple of the SC and the main body was jammed) Along with it, leakage occurs, the noise of the motor, the connector board has corroded. In addition, because of a succession of sp-mounting failures, latch replacement (optosensor checked, mold-a-checked) Full disk space, view | As for the details
No | Visit | Complaint | RCHN | 故障代码1035，pO2无端点 | 故障代码1035，pO2无端点 | 重新安装接口板后修复by tim | SWO-556184 | [5 empty] | Fault code 1035, pO2 unprovoked point | Fault code 1035, pO2 unprovoked point
No | Visit | Complaint | RIND | Unstable sodium level in patient sample | Unstable sodium level in patient sample | Done Protein removal and decontamination | SWO-556835 | Cleaned | Fluidic pathway | pH/BG Module | reference |  | Unstable sodium level in patient sample | Unstable sodium level in patient sample
No | Visit | Complaint | RIND | Oximetery Value Missing | Check the result | Run LS adjust , Pump Calibration then two point calibration. Run a sample found Oximetery value Missing. Open the oximetery module, Check the Pintch Valve , found tube is loose between valve . Adjust the tube . | SWO-558184 | Adjustment | HW | Oxi module | Pinch valve tube |  | Oximetery Value Missing | Check the result
No | Visit | Complaint | RAME | AutoCheck is not reading several of the QC vials. | AutoCheck not reading QC vial barcodes. Found laser not reading properly. | Replaced laser unit on AutoCheck module. AutoCheck now reads QC vial barcodes as it should. Analyzer checks okay. | SWO-556529 | Defect | HW | Auto Check | Scanner |  | AutoCheck is not reading several of the QC vials. | AutoCheck not reading QC vial barcodes. Found laser not reading properly.
No | Visit | Complaint | RDEU | Analysator ohne Funktion | Fehler: Analysator ohne Funktion Fehlersuche --> Netzteil, Hauptplatine, CPU-Unit überprüft --> LED an Hauptplatine an bei F1 leuchtet nicht --> Sicherung F1 erneuert --> ohne Erfolg --> zusätzlich Netzteil erneuert --> Sicherungen von Netzteil geschossen -->Sicherungen am Netzteil erneuert --> Gerät macht nun überhaupt keine Reaktion --> vermutlich neues Netzteil defekt! --> altes Netzteil und Hauptlatine wieder eingebaut --> gleiches Bild wie vor der Reparatur!!!! | Netzteil muss bestellt werden, da keines mehr im Technikerbestand, da dies beim Reparaturversuch defekt wurde! | SWO-556276 | Revisit required | HW | Power Supply | Power Supply |  | Analyzer without function | Error: Analyzer without function Troubleshooting --> Power supply, motherboard, CPU unit checked --> LED on motherboard on at F1 does not light up --> Fuse F1 renewed --> without success --> additional power supply renewed --> Fuses shot from power supply -->Fuses on the power supply renewed --> Device now makes no reaction at all --> probably new power supply defective! --> old power supply and main latin reinstalled --> same image as before the repair!!!!
No | Visit | Complaint | RKOR | Liquid sensor error | . | pH Membrane clean. pH shilling replace. | SWO-558002 | Defect | Electrode | Electrode | Sealing ring, pH |  | Liquid sensor error | .
No | Visit | Complaint | RSCH | 0731: Oberer MET II Flüssigkeitsfühler hat keine Spüllösung von ... 0508: Spülfehler | 0731: Oberer MET II Flüssigkeitsfühler hat keine Spüllösung von ... 0508: Spülfehler | Allgemeine Kontrolle ok. Alle Ventile ausgebaut und kontrolliert ok. Jahreswartung durchgeführt ok. (Flüssigkeittransport) Tubing Nippel Einlassdichtung ers ok. System i.o. Kal. + QC i.o. | SWO-556130 | Cleaned | Fluidic pathway | Met II module | Crea |  | 0731: Upper MET II liquid sensor has no rinsing solution of ... 0508: Flushing error | 0731: Upper MET II liquid sensor has no rinsing solution of ... 0508: Flushing error
No | Visit | Complaint | RIND | Touch screen not working. | Checked the ABL800 and found touch screen not working. | Restart the analyzer but touch is still not working. cross checked the display unit by exchanging with another analyzer found working properly. Crosschecked the CPU unit with other analyzer found CPU unit is defective. | SWO-556090 | Defect | HW | CPU | CPU |  | Touch screen not working. | Checked the ABL800 and found touch screen not working.
No | Visit | Complaint | RDEU | 28.05.2019 14:50 Fr. Kratzfeld +4923052945236: QC-Probleme CRP | CRP bei QC sporadisch zu hoch | Daten ausgewertet, anstieg erfolgte nach Chargenwechsel QC. Servicedump generiert, Anwender informiert neue QC zu bestellen. | SWO-557273 | Assisted Replacement | Consumables | LQC | LQC | CRP too high | 28.05.2019 14:50 Fr. Kratzfeld +4923052945236: QC-Probleme CRP | CRP bei QC sporadisch zu hoch
No | Visit | Complaint | RAME | Analyzer is not communicating with the LIS. Analyzer is setup like all of the other ABL90 analyzers and it is the only one not working. | IT unable to establish communication between the ABL and the LIS. | Checked all settings and could find no errors. Will obtain the needed parts to troubleshoot the analyzer's TCPIP output. | SWO-559224 | No fault found | SW | HIS/LIS | HIS/LIS |  | Analyzer is not communicating with the LIS. Analyzer is setup like all of the other ABL90 analyzers and it is the only one not working. | IT unable to establish communication between the ABL and the LIS.
No | Visit | Complaint | RAME | Calibrations failing due to THB drift and inhomogeneous cal 3 QC failing with 581 and 291 errors | Calibrations failing THB drift and QC failing with 581 ( OXI Spectrum Mismatch) and 291 (SHB too high) Logs show several calibration and QC failures | I was able to return the analyzer to the ready state by repeating the calibrations to remove the drift error. QC passed all three levels with no 581 or 291 errors after the calibrations passed. I am removing the analyzer for to monitor for THB drift to ensure a seamless transition at time of go live. The analyzer is not live at this time and there is no go live date scheduled at this time. Pam Smaldino released the analyzer to me for monitoring. A complete linearity will be performed upon return of the analyzer. | SWO-557035 | Adjustment | HW | Oxi module | tHb Cal | drift; 581; 291 | Calibrations failing due to THB drift and inhomogeneous cal 3 QC failing with 581 and 291 errors | Calibrations failing THB drift and QC failing with 581 ( OXI Spectrum Mismatch) and 291 (SHB too high) Logs show several calibration and QC failures
No | Visit | Complaint | RAME | THB failing QC | 1011 Oxi has no blank cal error | Tried swapping SP and SC without luck. Rebooted analyzer and received status code 2 error. Tried reloading 3.4 software but kept getting status code 2 error. Reloaded win7 and 3.4 software. Ran successful complete oxi test. Calibrations and QC passed. | SWO-556577 | Adjustment | HW | Oxi module | tHb Cal | 1011 | THB failing QC | 1011 Oxi has no blank lime error
No | Visit | Complaint | RAME | Analyzer is in the replace the gasket loop | Analyzer keeps asking to replace the gasket. | I replace the gasket with a new on and flushed the system all is working as expected | SWO-556927 | Defect | HW | Auto Inlet module | Inlet gasket |  | Analyzer is in the replace the gasket loop | Analyzer keeps asking to replace the gasket.
No | Visit | Complaint | RAME | LS errors unable to pump solutions | Inspection of the unit revealed the Sensor Interface PCB has failed (903-213R2009N0213). Inspection of the pinch valve & tubing tubing shows a small amount of white debris deposit on it. | Replaced the Sensor interface PCB, pinch valve & pinch tubing. | SWO-556596 | Defect | HW | Basic Unit | PCB, Sensor interface | unable to pump solutions | LS errors unable to pump solutions | Inspection of the unit revealed the Sensor Interface PCB has failed (903-213R2009N0213). Inspection of the pinch valve & tubing tubing shows a small amount of white debris deposit on it.
No | Visit | Complaint | RCHN | HCT无法测得，报错1267 | HCT无法测得，报错1267，借903-329一个 | 更换903-329后修复by tim | SWO-555459 | Defect | HW | Aspiration Module | Needle Wash Unit | 1267 | HCT can't be measured, error 1267 | HCT undetectable, error 1267, loan 903-329 one
No | Visit | Complaint | RCHN | 系统数据丢失 | 系统数据丢失,需要借件944-212 15696（缺货）944-214（15452）944-216（15557）944-258（15582） | 重新定标后修复by tim | SWO-555497 | Assisted Replacement | Consumables | Test Kit | Cartridge, cal |  | System data loss | System data loss, need to borrow 944-212 15696 (out of stock) 944-214 (15452)944-216 (15557) 944-258 (15582)
No | Visit | Complaint | RCHN | HCT值无法测量 | HCT值无法测量，借903-329,636-549各一个 | 更换636-549后修复by tim | SWO-556054 | Defect | HW | Aspiration Module | Cable, needle wash unit | 1223, 1258 | HCT value cannot be measured | HCT values cannot be measured, one for each of 903-329,636-549
No | Visit | Complaint | RCHN | 显示屏后盖螺丝老化 | 加借910-298一个 | 更换910-298备件后修复by tim | SWO-555954 | Defect | HW | Analyzer Control Module | Frame, Display |  | Screen screw aging | Borrow 910-298 one
No | Visit | Complaint | RCHN | 进样舱门损坏 | 加借886-639一个 | 更换886-638后故障依旧，需要加借903-717一个by tim | SWO-556043 | Revisit required | HW | Cover and Chassis | Door, Inlet | Top bracket | Damage to the hatch | Borrow 886-639 one
No | Visit | Complaint | RCHN | 进样舱门损坏 | 进样舱门损坏，需要加借903-717 |  | SWO-556798 | Defect | HW | Cover and Chassis | Door, Inlet | Top bracket | Damage to the hatch | The sample door is damaged and needs to be borrowed 903-717
No | Visit | Complaint | RKK | サンプル移動エラー/インレットフロー妨害が発生。流体回路を確認したが、漏れや異常はなかった。液センサーを中古のものに交換したところ改善。後日新品の日交換。（2019/5/30対応) | 液センサー交換後、異状なく測定できていたため、原因が液センサーであることが分かった。 | 中古品から、新品パーツに交換。 | SWO-556858 | Defect | HW | pH/BG Module | Liquid sensor |  | Sample move error/inlet flow interference occurred. The fluid circuit was confirmed, but there were no leaks or abnormalities. Improvement when the liquid sensor was replaced with a used one. A new day exchange at a later date. (For May 30, 2019) | After the liquid sensor was replaced, it was possible to measure without any problem, so it was found that the cause was a liquid sensor.
No | Visit | Complaint | RIND | Printing Problem | Checked and found , when printing some letters are not printing properly, and get blank image, found that printer head quality is not good. | Need to replace the Printer Head for good print. | SWO-558169 | Revisit required | HW | Printer Module | Printer Head |  | Printing Problem | Checked and found , when printing some letters are not printing properly, and get blank image, found that printer head quality is not good.
No | Visit | Complaint | RAME | Analyzer on hold - Leak from waste line | Instrument on hold with blockage in waste line. | Replaced waste line | SWO-557120 | Defect | HW | Waste tubings | Waste tubings |  | Analyzer on hold - Leak from waste line | Instrument on hold with blockage in waste line.
No | Visit | Complaint | RKOR | Pump calibration failed | . | Liquid flow line clean. Waste line clean. Oxi flow line clean. | SWO-558014 | Cleaned | Fluidic pathway | Waste tubings | Waste tubings |  | Pump calibration failed | .
No | Visit | Complaint | RPOL | Powtarzający się komunikat: Zmierzony przepływ cieczy powyżej zakresu ostrzegawczego. | Serwis stwierdza częśćiowe zablokowanie sekcji mokrej aparatu. | Odbiałczanie sekcji mokrej za pomocą podchlorynu, czyszczenie filtra, czyszczenie przy użyciu próbki czyszczącej, pomiar LQC. | SWO-556890 | Cleaned | HW | Wet section | Wet section |  | Repeated message: The measured flow of liquids above the warning range. | The service finds a partial blockage of the wet section of the camera.
No | Visit | Complaint | RPOL | Aparat nie zaczytuje wszystkich kodów kasetek. | Serwis stwierdza nieprawidłowe pozycjonowanie wiązki lasera czytnika kodów kreskowych. | Czyszczenie czytnika kodów kreskowych, justowanie pozycji wiązki lasera, testowanie czytnika. | SWO-556891 | Cleaned | HW | Storage Module | Barcode reader, storage wheel |  | The camera does not read all the cassette codes. | The service finds incorrect positioning of the laser beam barcode reader.
No | Visit | Complaint | RAME | Calibration errors and, OIR messages. | Calibration errors and, OIR messages. | Checked the data logs and, found the analyzer presenting several fluid transport errors, 1184 Leak detected, 1063 Pressure test vacuum error, 0508 Liquid transport error during rinse and, 1084 Inhomogeneous cal 3 solution. I checked the fluid transport values and, found the Pressure test failing for the vacuum value running high outside of range 84 (0.25-0.75), indicating a leak in the fluid path. I replaced the inlet gasket holder , the inlet connector gasket , the hemolyzer pinch valve tube (Found with pin hole), the waste tubing and, the waste connector. To check for the calibrations errors, I ran the sensor simulator test and, confirmed the sensor interface board was properly working, the test is passing within specifications thus indicating no eerrros with the sensor interface board. I determined that the cause of the calibration errors was related to the fluid transport errors the instrument was presenting. | SWO-557803 | Defect | HW | Basic Unit | PM kit | 1184, 1063, 508, 1084 | Calibration errors and, OIR messages. | Calibration errors and, OIR messages.
No | Visit | Complaint | RDEU | 29.05.2019, 15:08 Uhr ==> Anrufer Herr Brill, Tel.: 09431888277 ==> Keine Verbindung zu Radiance | 29.05.2019, 15:08 Uhr ==> Anrufer Herr Brill, Tel.: 09431888277 ==> Keine Verbindung zu Radiance | Zugang bei Herrn Prehn angefragt Laut MICKO ist dies ein bekanntes Problem beim Kunden. Der Kunde prüft intern sein Netzwerk | SWO-557573 | No fault found | SW | HIS/LIS | HIS/LIS |  | 29.05.2019, 15:08 ==> Caller Mr. Brill, Tel.: 09431888277 ==> No connection to Radiance | 29.05.2019, 15:08 ==> Caller Mr. Brill, Tel.: 09431888277 ==> No connection to Radiance
No | Visit | Complaint | RESP | Fuga detectada en el Paq. Soluc (964). | El cliente llama por no haber recibido el SP solicitado (festivo el Dinamarca). Solicita visita técnica ya que ante el fin de semana está intranquilo. | Se visita al cliente, se le realiza un nuevo kit de tubos y cambio de a la aguja de entrada de muestras; así como solicitud de un SP urgente a C. Fernández el cual debe recibirse el lunes. Se realiza calibraciones y QC´s. | SWO-558056 | Defect | HW | Basic Unit | PM kit | 964 | Leak detected in the Paq. Soluc (964). | The customer calls for not receiving the requested SP (Danish holiday). Request a technical visit as it is uneasy at the weekend.
No | Visit | Complaint | RAME | Inlet module stuck in mid position. | Inspection of the system revealed error 1391, rebooting didn't help. | Re-Assembled inlet gear in correct position & Calibrated inlet in service screen. Still intermittently erroring, ordered inlet Assy. | SWO-557760 | Revisit required | HW | Auto Inlet module | Auto Inlet module |  | Inlet module stuck in mid position. | Inspection of the system revealed error 1391, rebooting didn't help.
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | LAC定标不过 | 更换测试卡后修复 | SWO-558800 | [5 empty] | Test card, reagent package failure | LAC calibration is not
Yes | Visit | Complaint | RIND | Qc Level going out | Check System | Check system, check calibration and QC log, run QC by own mix properly and run still QC going out replace SC, QC coming Ok. | SWO-561582 | [5 empty] | Qc Level going out | Check System
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | CA,K定标失败 | 更换测试卡后修复 | SWO-562280 | [5 empty] | Test card, reagent package failure | CA, K bid failed
No | Visit | Complaint | RCHN | 质控不过 | 质控不过，借903-099,933-494各一个 | 更换903-099,933-494后修复by tim | SWO-556758 | [5 empty] | But the quality control is not | However, one each is 903-099,933-494
Yes | Visit | Complaint | RCHN | 触摸屏失灵 | 触摸屏失灵, 借910-324一个。 | 更换910-324后修复by tim | SWO-556160 | [5 empty] | Touch screen failure | The touch screen is not working, borrowing 910-324 one.
No | Visit | Complaint | RCHN | 显示屏排线故障 | 显示屏排线故障,需要借件636-571一个 | 更换636-571后修复 by tim | SWO-556120 | Defect | HW | Analyzer Control Module | Cable, Display |  | Display line failure | Display line failure, need to borrow 636-571 one
Yes | Visit | Complaint | RCHN | 定标失败 1024 | 定标失败 1024,借903-364一个。 | 更换903-364备件后故障依旧，需要重新借903-364一个和TNT定标卡一张by tim | SWO-556839 | [5 empty] | Bid Failure 1024 | Bid failure 1024, borrowing 903-364 one.
No | Visit | Complaint | RKK | 機器より異音 | 機器より異音 | CPUファン不良の為交換。 | SWO-558094 | [5 empty] | Noise from equipment | Noise from equipment
No | Visit | Complaint | RIND | PH & pCO2 Calibration not passed | PH sensitivity is low and calibration failed | Reverse Sensor connector and zeroing done also reinstall software and check, same problem PH Sensitivity is still low now need to check with interface board and new sensor connector. Waiting for Spare | SWO-557471 | [5 empty] | PH & pCO2 Calibration not passed | PH sensitivity is low and calibration failed
No | Visit | Complaint | RMAY | Oncall breakdown | Calibration failed due to clot in ph chamber | -Check analyzer data log -Clean chamber -Run calibration | SWO-563824 | Cleaned | HW | pH/BG Module | measuring chamber | 379 | Oncall breakdown | Calibration failed due to clot in ph chamber
No | Visit | Complaint | RDEU | Oxi Fehler bei Pat Messung und QC ´Hemolyzer phase regulation` | Hemolyzer phase regulation failed, Oxi Hardware Fehler pH instabil Installationsdatum 24.09.2018 | Oxi Complette Test ohne Beanstandung Hemolysator ausgetauscht. pH Elektrode und Messkammer gereinigt, neuen Dichtring aufgesetzt. Aufwärm- und Initialisierungsphase Lamp Unit auf 3370 justiert. Oxi Complette Test i.O Kalibrationen durchgeführt tHb Kal QC in allen 3 Leveln ohne Beanstandung Patientenmessung ohne Beanstandung | SWO-557510 | Defect | HW | Oxi module | Hemolyzer |  | Oxi error in Pat measurement and QC 'Hemolyzer phase regulation' | Hemolyzer phase regulation failed, Oxi hardware error pH unstable Installation date 24.09.2018
No | Visit | Complaint | RAME | Waste Detector Error | Waste Detector Error. Found various sampling errors on patient samples. Inlet gasket was dirty, inlet probe was bent, pH electrode missing seal ring. | Replaced the waste detector. Replaced inlet gasket, inlet probe and pH seal ring from customers stock. | SWO-557596 | Defect | HW | Basic Unit | Waste detector | 618 | Waste Detector Error | Waste Detector Error. Found various sampling errors on patient samples. Inlet gasket was dirty, inlet probe was bent, pH electrode missing seal ring.
No | Visit | Complaint | RAME | Connect to the interface | Unit indicates connection made but is not transmitting. | Was not able to make work with current configuration. Reconfigured with serial settings and cable. | SWO-557651 | Setup | SW | HIS/LIS | HIS/LIS |  | Connect to the interface | Unit indicates connection made but is not transmitting.
No | Visit | Complaint | RCHN | PH值定标灵敏度低 | PH值定标灵敏度低 | 自行更换945-614后修复by tim | SWO-557478 | Assisted Replacement | Electrode | Electrode | pH | 378 | Low P-value calibration sensitivity | Low P-value calibration sensitivity
No | Visit | Complaint | RIND | Touch screen not working. | Checked the ABL800 and found touch screen not working | Connected Keyboard and Mouse for interface with defective touch screen. | SWO-559054 | Revisit required | HW | Display | Display |  | Touch screen not working. | Checked the ABL800 and found touch screen not working
No | Visit | Complaint | RPOL | Wykryto prądy upływu | Przegląd logów, testy serwisowe. | Wymiana uszkodzonego zaworu V8, testy serwisowe, kalibracje. | SWO-558054 | Revisit required | HW | Fluidic manifold | Fluidic manifold | Leak current | Leakage Currents detected | Log overview, service tests.
Yes | Visit | Complaint | RCHN | 废液口接口破损 | 废液口接口破损，需要借件903-205一个 | 更换903-205后修复by tim | SWO-557393 | [5 empty] | The waste port interface is broken | The waste port interface is broken and requires a loan of 903-205
No | Visit | Complaint | RCHN | 进样仓门关不上 | 进样仓门关不住 5784（未更换备件） | 拿了纸垫了下后修复by HUNTER | SWO-558939 | Adjustment | HW | Cover and Chassis | Door, Inlet | 5784 | The door of the warehouse can't be closed. | The entrance door does not hold 5784 (spare parts not replaced)
No | Visit | Complaint | RCHN | 测试卡卡盘 | 测试卡卡盘 | 更换测试卡后修复 | SWO-558935 | Defect | Consumables | Test Kit | Cartridge, Broken | 5784 | Testing the kaka plate | Testing the kaka plate
No | Visit | Complaint | RCHN | 数据库损坏，定标数据丢失 | 数据库损坏，定标数据丢失，借944-212（15836缺货,16030）944-214（15653），944-216（15755），944-258（16028）各一张 | 重装系统，重新定标后修复by tim | SWO-556674 | Defect | SW | RMED SW | RMED SW |  | Database corruption, bid data loss | Database corruption, data loss of bid, borrow 944-212 (15836 out of stock, 16030) 944-214 (15653), 944-216 (15755), 944-258 (16028) each
Yes | Visit | Complaint | RCHN | 定标失败 1024（在常州市妇幼保健院装机发生故障） | 定标失败 1024 ，需要借件903-364、944-268（15132） | 更换903-364后故障未排除，报错1024，加借903-236一个by Hunter | SWO-557250 | [5 empty] | Bid failure 1024 (failure of installation at Changzhou Maternal and Child Health Hospital) | Bid failure 1024, need to borrow 903-364, 944-268 (15132)
No | Visit | Complaint | RIND | Printing Problem | Checked and found that when printing, some letters are not printing properly, and get blank image. found that printer head quality is not good. | Need to replace the Printer Head for good print | SWO-558197 | Defect | HW | Printer Module | Printer Head |  | Printing Problem | Checked and found that when printing, some letters are not printing properly, and get blank image. found that printer head quality is not good.
No | Visit | Complaint | RIND | Unstable sodium level in patient sample | pH reading outside the acceptable range | Cleaned pH electrode and chamber, replace Ref. Membrane, analyser working fine | SWO-557892 | Assisted Replacement | Electrode | Membrane | reference |  | Unstable sodium level in patient sample | pH reading outside the acceptable range
No | Visit | Complaint | RAME | Leaking out the front and the MET is low | Leaking out the front and the MET is low | Ran service THB system also had a. Lot in the tubing and had to retube the inlet. | SWO-557772 | Defect | HW | Inlet module | Waste tubings |  | Leaking out the front and the MET is low | Leaking out the front and the MET is low
No | Visit | Complaint | RMAY | Error occur due to LS2 detected missing liquid during pump 5 tubing fill. | Error occur due to blockage at tubing line No.7 | Flush tubing line No.7. Perform system clean. | SWO-557831 | Cleaned | HW | Wet section | Wet section |  | Error occur due to LS2 detected missing liquid during pump 5 tubing fill. | Error occur due to blockage at tubing line No.7
No | Visit | Complaint | RDEU | 31.5.2019/11:00 Frau Adam 05115321458 -> Ableitstrom im Anschluß SP erkannt | Ableitstrom erkannt, pH Impedanzfehler | Gerät geprüft, Schlauchsystem ersetzt, Funktionskontrolle nicht ok | SWO-558352 | Revisit required | SC | SC | Leak current |  | 31.5.2019/11:00 Mrs. Adam 05115321458 -> Discharge current detected in connection SP | Discharge current detected, pH impedance error
No | Visit | Complaint | RDEU | 03.06.2019 Sr. Mebela +492191139003 meldet, das der Analysator in der Nacht ausgeschaltet wurde. Sie wünscht einen Vorortservice. | Lp Motor macht laute Geräusche, Motorspindel bis Anschlag ausgefahren, | Motor ausgebaut und von Hand zurückgeschraubt, Lösungspack zum Test mehrfach eingesetzt und herausgenommen, Startphase 2 std. Backlog Fehler Sensorcassette , Sensorkassette ersetzt, Startphase 50min,Kal und Qc Leistungstest GLU und LAC, mehrfach Kalibriert, | SWO-559033 | Adjustment | HW | Basic Unit | Stepper Motor |  | 06/03/2019 Sr. Mebela +492191139003 reports that the analyzer was turned off at night. She wants a suburban service. | Lp motor makes loud noises, motor spindle extended to stop,
No | Visit | Complaint | RESP | El equipo está parado | El equipo da fallo de coágulos y cal de bombas | Se realiza un flush y se intercambia tanto la entrada como el pack y el cassette con el otro equipo, se observa que no parece haber un coágulo y al cambiar tanto el pack como el cassette se arregla el fallo. Se procede a pedir sustitución del fungible en mal estado | SWO-558924 | Assisted Replacement | SC | SC | FT |  | The team is standing still. | The team gives clot failure and bomb lime
No | Visit | Complaint | RLTD | error message- SC blocked even after try various things 29.05.19 10:33 - Clot issue | Clot detected in SC again. Suspected issue with sensor bay and $or Liquid sensors. | Replaced Sensor bay unit under Warranty. Aligned contact pins and ran fluid transport checks. All ok. Analyser has passed calibrations and QC's. | SWO-558433 | Defect | HW | Basic Unit | Sensor Bay Unit | Clot | error message- SC blocked even after try various things 29.05.19 10:33 - Clot issue | Clot detected in SC again. Suspected issue with sensor bay and /or Liquid sensors.
No | Visit | Complaint | RAME | Safepico mixer making clicking noise | Safepico mixer making clicking noise. | Found part of plastic chipped and in the mixer. Removed plastic pease. | SWO-558663 | Cleaned | HW | Sample mixer | Sample Mixer | plastic part | Safepico mixer making clicking noise | Safepico mixer making clicking noise.
No | Visit | Complaint | RIND | Pco2 value higher side | Pco2 values higher side for some patient | Checked the analyser removed the sc aand reinstalled it and told the respiratory technican to run the sample so they were not mixing the sample properly so given some training on that and then tested the sample all values are ok | SWO-561574 | [5 empty] | Pco2 value higher side | Pco2 values higher side for some patient
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 参比电极错误 1361 | 更换测试卡后修复 | SWO-558905 | [5 empty] | Test card, reagent package failure | Reference electrode error 1361
No | Visit | Complaint | RCHN | 试管架断裂 | 试管架断裂,需要借件903-584*2个 | 更换903-584*2后修复by HUnter | SWO-558107 | Defect | HW | Aspiration Module | Tube holder |  | The test tube rack is broken. | The test tube rack is broken and requires a loan of 903-584 x 2
No | Visit | Complaint | RCHN | 测试卡卡盘，报错5784，卡盘部位： teethn | 测试卡卡盘，报错5784，卡盘部位： teethn | 更换测试卡后修复by Hunter | SWO-558118 | Defect | Consumables | Test Kit | Cartridge, Broken | Teeth | Test card plate, error 5784, card plate location: teethn | Test card plate, error 5784, card plate location: teethn
Yes | Visit | Complaint | RCHN | 定标失败 1024（在常州市妇幼保健院装机发生故障） | 定标失败 1024（在常州市妇幼保健院装机发生故障）,更换903-364后故障未排除，报错1024，加借903-236一个 | 更换903-236后修复by Hunter | SWO-557968 | [5 empty] | Bid failure 1024 (failure of installation at Changzhou Maternal and Child Health Hospital) | Bid failure 1024 (in Changzhou City Maternal and Child Health Hospital installation failure), replacement 903-364 after the fault was not corrected, reported error 1024, plus 903-236 one
No | Visit | Complaint | RIND | pCO2 QC outside the range. | pCO2 QC outside the range. | Guided the Biomedical Eng for Calibration reports, found pCO2 sensitivity is very low, Visited the customer check the stock of pCO2 membranes replaced the same calibrated analyzer found pCO2 sensitivity OK performed the QC found results are OK. | SWO-561264 | Assisted Replacement | Electrode | Membrane | pCO2 | Sens low; QC outside range | pCO2 QC outside the range. | pCO2 QC outside the range.
Yes | Visit | Complaint | RPAC | Move analyser into Theatre and enter TCP/IP Addresses. | Move analyser into Theatre and enter TCP $IP Addresses. | Move analyser into Theatre for Dorevitch Pathology and power On. Enter the provided TCP $IP Addresses into the analyser via Windows Control Panel. | SWO-558165 | Re classified RMED | planned SWO |  |  | Setup | Move analyser into Theatre and enter TCP/IP Addresses. | Move analyser into Theatre and enter TCP /IP Addresses.
No | Visit | Complaint | RKK | 電極交換 廃液検出器l交換作業 |  |  | SWO-560329 | Defect | HW | Basic Unit | Waste detector |  | Electrode replacement Waste water detector l replacement work
No | Visit | Complaint | RIND | Ph unstable in patient results. | Checked the patient result log, Ph reported unstable. | Cleaned all electrode chambers. performed 2point calibration. Performed sample measurement but pH is still reported unstable. Cross checked the pH electrode with other analyzer and performed sample measurement without error. | SWO-559162 | Assisted Replacement | Electrode | Electrode | pH |  | Ph unstable in patient results. | Checked the patient result log, Ph reported unstable.
No | Visit | Complaint | RPOL | Aparat nie przyjmuje pakietów odczynnikowych z błędem : błąd odczytu układu scalone | Błąd odczytu układu scalone przy każdym pakiecie odczynnikowym | Czyszczenie styków , sprawdzenie aparatu. Zainstalowano nowy SP90. | SWO-559904 | Cleaned | HW | Basic Unit | PCB, solution pack |  | The engine does not accept reagent packets with an error: The read error of the merged circuit | Merged circuit read error at each reagent package
No | Visit | Complaint | RPOL | Wykryto prądy upływu. | Przegląd logów aparatu, testy serwisowe. | Wymiana modułu płynów, testy serwisowe, kalibracje. | SWO-559974 | Defect | HW | Fluidic manifold | Fluidic manifold | Leak current | Leakage currents detected. | Overview of camera logs, service tests.
No | Visit | Complaint | RDEU | 03.06.2019; Pfl. Christian; Tel.: +49232349892701; Lac Kal Fehler und Messabweichungen 03.06.2019; 14:30 Uhr; Sr. Kathleen; Tel.:+49232349892701; weiterhin Lac Kal Fehler | Nach SC Wechsel Lac Probleme und pCO2 bei LevelA zu hoch | SC 946-009 und LP 944-197 geliefert und installiert . Reklamationsausdrucke an RDEU gemailt. Nach Startphase noch Lactat Status zu hoch und pCO2 bei Level1 zu hoch. LP erneuert, Bei Installation Lsg3 inhomogen. Funktionen im Service überprüft, OK. Startphase wiederholt. | SWO-558993 | Assisted Replacement | SC | SC | Lac | Cal | 03.06.2019; Pfl. Christian; Tel: +49232349892701; Lac Kal Errors and Measurement Deviations 03.06.2019; 2:30 p.m.; Sr. Kathleen; Phone:+49232349892701; Continue Lac Kal Error | After SC change Lac problems and pCO2 at LevelA too high
No | Visit | Complaint | RLTD | Peter Williams Tel: 01243 788122 x: 33591 Further call rec'd at 16.15 from Peter Williams. Analyser still has problems. Has now moved it to the lab should you visit. Could you call him please on above number. | Pinch valve tube at fault causing leak error. | Pinch valve tube replaced - ran pressure test in services all okay. | SWO-559393 | Defect | HW | Oxi module | Pinch valve tube | 1063 | Peter Williams Tel: 01243 788122 x: 33591 Further call rec'd at 16.15 from Peter Williams. Analyser still has problems. Has now moved it to the lab should you visit. Could you call him please on above number. | Pinch valve tube at fault causing leak error.
No | Visit | Complaint | RLTD | David Ryder Tel: 07580 818900 03.06.19 13:37pm - David called back to also log that the battery says 0% but is plugged in | Casing for display unit cracked. Battery showing 0% | Display unit replaced with new. Battery jumper reset and battery now charging correctly | SWO-559513 | Defect | HW | Display | Display | Screws | David Ryder Tel: 07580 818900 03.06.19 13:37pm - David called back to also log that the battery says 0% but is plugged in | Casing for display unit cracked. Battery showing 0%
No | Visit | Complaint | RLTD | 03/06/19 - Now failing on bilirubin on QC's please phone Debbie Donovan on 02381 206721 Jose on 02381 206721 | Oxi mismatch errors - hemolyser blockage. | Soaked hemolyser manually and flushed. Ran thb calibration QC's all now passing. | SWO-559392 | Revisit required | HW | Oxi module | Hemolyzer | 581, Clot | 03/06/19 - Now failing on bilirubin on QC's please phone Debbie Donovan on 02381 206721 Jose on 02381 206721 | Oxi mismatch errors - hemolyser blockage.
No | Visit | Complaint | RKOR | Replace gasket holder | replace gasket holder , Inlet probe adjust, pressure test ok, cal control data check. | replace gasket holder , Inlet probe adjust, pressure test ok, cal control data check. | SWO-561000 | Defect | HW | Auto Inlet module | Inlet gasket |  | Replace gasket holder | replace gasket holder , Inlet probe adjust, pressure test ok, cal control data check.
No | Visit | Complaint | RFRA | Défaut lors de l’installation dès pack de solution Problème rencontré sur les trois analyseurs | Vérification complète de l’analyseur | Vérification et tests en mode service $ $ tests dès pack de solution défaillants -> en erreur factor calibration pump <0.80 $ $ installation d’un kit de maintenance $ $ relance $ $ mise en route $ $ calibration $ $essais bons Lot défaillant de pack de solution HJ-69 | SWO-559280 | Defect | HW | Basic Unit | PM kit |  | Failure to install as soon as solution pack Problem encountered on the three analyzers | Full verification of the analyzer
No | Visit | Complaint | RAME | Inlet module stuck in mid position. | Inspection of the system revealed the inlet module intermittently would get out of position. | Replaced auto inlet assembly. | SWO-558770 | Defect | HW | Auto Inlet module | Auto Inlet module |  | Inlet module stuck in mid position. | Inspection of the system revealed the inlet module intermittently would get out of position.
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 多参数定标失败 | 更换测试卡后修复 | SWO-561795 | [5 empty] | Test card, reagent package failure | Multi-parameter calibration failed
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | LAC定标失败 | 更换测试卡后修复 | SWO-559114 | [5 empty] | Test card, reagent package failure | LAC calibration failure
No | Visit | Complaint | RCHN | 黑屏 | 黑屏，借903-175一个（因910-343缺货所以暂时不借） |  | SWO-559140 | Defect | HW | Analyzer Control Module | CPU | CRM Corrective Action: Repaired by replacing 910-343 by Andrew | Black screen | Black screen, borrow 903-175 one (because 910-343 is out of stock so temporarily not to borrow)
No | Visit | Complaint | RDEU | sporadische Driften bei Lyte/Metabolite | sporadische Driften bei Lyte $Metabolite, resultierende Meßfehlerbei QK und Blut | Reinigungsarbeiten, verschlissene (gerissene) E1001 (alt: 48-02) ersetzt, alle Membranventile demontiert und gereinigt sowie wieder montiert, beide Modulpumpenschläuche ersetzt, PumpenKal, Nachfüllung, mehrere korrekte Kal abgewartet, QK ok | SWO-559244 | Cleaned | HW | Fluidic pathway | Fluidic pathway |  | sporadic drifts in lyte/metabolites | sporadic drifts in Lyte /Metabolite, resulting measurement errors in QK and blood
No | Visit | Complaint | RDAN | From: Trine Rennebod Larsen <Trine.Rennebod.Larsen@rsyd.dk> Sent: 28. maj 2019 14:15 To: Heerfordt, Niels NHE <niels.heerfordt@radiometer.dk> Cc: Birgit Christensen <Birgit.Christensen@rsyd.dk>; Gunhild Hansen <Gunhild.Hansen1@rsyd.dk> Subject: fejl på Intensiv Kære Niels. Vi har oplevet, at der er en del fejl på Intensivs ABL. Afdelingen har også selv henvendt sig. Jeg har lavet et udtræk over opgaver for maj måned. De sidste 14 dage har der været mange fejlkalibreringer, især på glucose og laktat. Vil du kigge på det? Mange hilsner, Trine Venlig hilsen Trine Rennebod Larsen Biokemiker, Ph.D. Trine.Rennebod.Larsen@rsyd.dk - www.ouh.dk Tlf. 51779101 Klinisk Biokemisk Afdeling, Svendborg Svendborg Sygehus Baagøes Alle 15, 5700 Svendborg | Afprøvet og undersøgt, luft i lactat elektrode. | Luft i lactat elektrode frigjort, personale instrueret i membranering og vigtigheden i, at elektrolytten er tæt på 37 grader. Nulpunktjusteret, luftventil checket og slutafprøvet. | SWO-559077 | Cleaned | HW | El/Met Module | Lac | Air bubble | From: Trine Rennebod Larsen <Trine.Rennebod.Larsen@rsyd.dk> Late: 28. May 2019 14:15 To: Heerfordt, Niels NHE <niels.heerfordt@radiometer.dk> Cc: Birgit Christensen <Birgit.Christensen@rsyd.dk>; Gunhild Hansen <Gunhild.Hansen1@rsyd.dk> Subject: Failure of intensive Dear Niels. We have seen that there are a lot of mistakes at the Intensivs ABL. The department has also approached itself. I have made an extraction over tasks for the month of May. The last 14 days there have been many error calibrations, especially on glucose and lactate. You want to look at it? Many greetings, Trine Regards Trine Rennebod Larsen Biochemist, Ph.D. Trine.Rennebod.Larsen@rsyd.dk-www.ouh.dk Tel. 51779101 Clinical Biochemical Department, Svendborg Svendborg Hospital Baagøes Alle 15, 5700 svendborg | Tested and investigated, air in lactate electrode.
No | Visit | Complaint | RDEU | 03.06.2019, 14:28 Uhr ==> Anrufer Herr Brech, Tel.: 05971424042 ==> pH-Messung instabil 03.06.2019, 15:24 Uhr ==> Anrufer Herr Brech, Tel.: 05971424042 ==> Die durchgeführten Maßnahmen haben nicht zum Erfolg geführt, Fehler besteht weiterhin | Ph bei Messung instabil | Elektrode und Messkammer gereinigt und Dichtring Ph Elektrode erneuert. | SWO-559686 | Defect | Electrode | Electrode | Sealing ring, pH |  | 03.06.2019, 14:28 ==> Caller Mr. Brech, Tel.: 05971424042 ==> pH measurement unstable 03.06.2019, 15:24 ==> Caller Mr. Brech, tel.: 05971424042 ==> The measures carried out have not been successful, errors still exist | Ph unstable when measured
No | Visit | Complaint | RDEU | Analysator ohne Funktion | Fehler: Analysator ohne Funktion --> LED neben der Sicherung S1 auf der Hauptplatine leuchtet nicht auf --> vermutliche Spannungsverlust $abfall!!! siehe auch SWO-556276 Fehlersuche --> Hauptplatine (Analyzer Control PCB), die einzelnen Module (AC-Modul, pH $BG-, El $Met-, Einlass-, Blind- und Oxy-Modul, FlexQ-Modul) , welche an der Hauptplatine angesteckt, sowie die Steckverbindungen (Kabel) überprüft. Alle Module der Hauptlatine abgesteckt und dann eines nach dem anderen angesteckt, um zu sehen, von welchem Modul der Spannungsabfall kommt! --> Nach Lokalisierung des Moduls, welches den Fehler verursacht, auf weitere und intensivere Fehlersuche gegangen! --> Kabel vom AC-Modul zur Verteilerplatine auf der Rückseite des Analysators defekt (eine Ader des Kabels verletzt, welches dann die Spannung zusammenbrechen lässt) --> FlexQ-Modul ungenutzt, deswegen ist ein Flachbandkabel übrig! -->Elektrodenwartung fällig | Defektes Flachbandkabel vom AC-Modul ausgebaut und das übrige Flachbandkabel vom nicht genutzten FlexQ-Modul verwendet, um das AC-Modul zu betreiben. Tausch des Flachbandkabels war erfolgreich und die Spannung bricht nicht mehr zusammen. Erforderliche Elektrodenwartung durchgeführt, Pumpenkal., LS-Justage, Startphase, Kal., QC, System und Funktion i.O. | SWO-559123 | Defect | HW | Basic Unit | Flat cable |  | Analyzer without function | Error: Analyzer without function --> LED next to the fuse S1 on the motherboard does not light up --> probable voltage loss /abfall!!! see also SWO-556276 Troubleshooting --> Motherboard (Analyzer Control PCB), the individual modules (AC module, pH /BG, El /Met, inlet, blind and oxy module, FlexQ module) which is plugged into the motherboard, as well as the connectors (cables) checked. Plugged all modules of the main latin and then plugged in one by one to see which module the voltage drop comes from! --> After localizing the module that causes the error, went to further and more intensive troubleshooting! --> Cable from the AC module to the distribution board on the back of the analyzer defective (one wire of the cable is injured, which then causes the voltage to collapse) --> FlexQ module unused, so a ribbon cable is left! -->Electrode maintenance due
No | Visit | Complaint | RKOR | pCO2 QC out of range | . | Membrane replace. Electrode replace. Cal QC check. Service mode check. | SWO-561781 | Assisted Replacement | Electrode | Electrode | pCO2 | QC out of range | pCO2 QC out of range | .
No | Visit | Complaint | RLTD | When I came in to do the Half-Yearly service I found that the analyser was dead with the exception of the lights around the inlet & storage doors. | I had previously tried powering the analyser off & on but that didn't help. I checked all of the fuses and they were all ok so I have returned with a power supply and a Main board. I swapped the main board with a new one and then powered the analyser back up but it was still the same. I replaced the power supply unit and powered it up and this time it powered up normally. | The power supply unit was replaced. Once the analyser had powered back up and was ready I completed the half yearly service and no additional problems were found. | SWO-559733 | Defect | HW | Cover and Chassis | Power Supply |  | When I came in to do the Half-Yearly service I found that the analyser was dead with the exception of the lights around the inlet & storage doors. | I had previously tried powering the analyser off & on but that didn't help. I checked all of the fuses and they were all ok so I have returned with a power supply and a Main board. I swapped the main board with a new one and then powered the analyser back up but it was still the same. I replaced the power supply unit and powered it up and this time it powered up normally.
No | Visit | Complaint | RKOR | Monitor defect | Monitor defect | Replace display unit | SWO-561687 | Defect | HW | Display | Display |  | Monitor defect | Monitor defect
No | Visit | Complaint | RLTD | 03/06/19 - Now failing on bilirubin on QC's please phone Debbie Donovan on 02381 206721 Jose on 02381 206721 04.06.2019 15:38pm lost oxsymmetry values on patient results Deborah Donovan 02381 206721 | After flushing oxi issue persisting possibly a hemolyser fault | Replaced hemolyser | SWO-559917 | Defect | HW | Oxi module | Hemolyzer | 581, Clot | 03/06/19 - Now failing on bilirubin on QC's please phone Debbie Donovan on 02381 206721 Jose on 02381 206721 04.06.2019 15:38pm lost oxsymmetry values on patient results Deborah Donovan 02381 206721 | After flushing oxi issue persisting possibly a hemolyser fault
No | Visit | Complaint | RAME | Touch screen issue. monitor frame is loose | 1) Display unit cracked around the edges. | 1) Replaced display. 2) Replenished HUP inventory. 3) Updated Credit Claim form. | SWO-560161 | Defect | HW | Display | Display |  | Touch screen issue. monitor frame is loose | 1) Display unit cracked around the edges.
No | Visit | Complaint | RCHN | 泵轮故障 | 泵轮故障，借902-971一个 | 更换902-971后修复by Hunter | SWO-558119 | [5 empty] | Pump wheel failure | Pump wheel failure, borrow 902-971 one
Yes | Visit | Complaint | RCHN | 机器无法开机 | 机器无法开机,需要借件910-342一个 | 更换910-342后修复(by chuanwei) | SWO-557210 | [5 empty] | The machine does not power on | The machine cannot be switched on, need to borrow 910-342 one
No | Visit | Complaint | RCHN | 系统报警连接错误，然后自动关机 | 系统报警连接错误，然后自动关机，借910-379一个 | 更换910-379后修复by hUnter | SWO-558863 | Defect | HW | DMS_PC | CF card |  | System alarm connection error, then automatically shut down | System alarm connection error, and then automatically shut down, borrow 910-379 one
No | Visit | Complaint | RCHN | 溶血器不工作 | 溶血器不工作，借902-980、903-644、636-518各一个 | 更换902-980、903-644备件后修复(by chuanwei) | SWO-559062 | Defect | HW | Oxi module | Hemolyzer |  | The hemolytic does not work. | Hemolysis does not work, borrowing 902-980, 903-644, 636-518 each
No | Visit | Complaint | RCHN | TNT质控连续失控 | TNT质控连续失控，借903-702一个 | 机器未修复，加借580-096,924-653,903-364,903-119各一个，此单先关单(by chuanwei) | SWO-559006 | Defect | analyzer removed | Analyzer removed | Analyzer removed |  | TNT QC continuously out of control | TNT QC continuously out of control, borrowing 903-702 one
Yes | Visit | Complaint | RCHN | TNT质控连续失控 | 加借580-096,924-653,903-364,903-119各一个 |  | SWO-559832 | Re classified RMED | Same day UpSWO |  |  |  | TNT QC continuously out of control | Borrow 580-096, 924-653, 903-364, 903-119 each
No | Visit | Complaint | RCHN | 定标步骤失败 | 定标步骤失败，借903-364一个 | 更换903-364后修复by Hunter | SWO-558945 | Defect | HW | Process Module | Optical Unit | 1059 | The bid step failed | The bid step failed, borrowing 903-364 one
No | Visit | Complaint | RCHN | 试管夹断裂 | 试管夹断裂，借903-584一个 | 更换903-584后修复(by chuanwei) | SWO-558101 | Defect | HW | Aspiration Module | Tube holder |  | Test tube clip break | Test tube clip broken, borrow 903-584 one
No | Visit | Complaint | RIND | New 300/60 sensor cassette installation in 24th May early morning 12.13 AM but its showing re installed in 4/10/2019. | checked the analyzer system logs and sc status and conformed its showing wrong installation date. | taken sc installation report its showing re installed .sc installed in 24th midnight 12.13 am but its showing already installed in 4 $10 $2019.now its showing 5 days reaming for expiration date. | SWO-560442 | [5 empty] | New 300/60 sensor cassette installation in 24th May early morning 12.13 AM but its showing re installed in 4/10/2019. | checked the analyzer system logs and sc status and conformed its showing wrong installation date.
No | Visit | Complaint | RIND | In samples K+ result some times shows high some times very low | Check analyzer status and calibration result which is OK. Check patient results in which for some samples K result is not satisfactory. | Check sensor cassette status,zeroing which are OK ,Run QC which is also OK and found no problem in the instrument.Give training for preanalytical,and also provide PICO sampler to use for sampling.Because found the problem in heparin use .Run a sample and found satisfactory result. | SWO-568382 | [5 empty] | In samples K+ result some times shows high some times very low | Check analyzer status and calibration result which is OK. Check patient results in which for some samples K result is not satisfactory.
No | Visit | Complaint | RPOL | W dziennikach systemu pojawiają się okresowo powiadomienia otwarte wejście SC . | Weryfikacja danych systemu. | Wymiana Opto Sensor oraz Interface Board. Testy. | SWO-559764 | [5 empty] | In the system logs, open SC input notifications periodically appear. | System data verification.
No | Visit | Complaint | RKK | Neon Lamp intensity too low | Neon Lamp intensity too low QCで電解質が日内変動 | ランプユニット、比較電極交換 | SWO-561722 | Defect | HW | Oxi module | Lamp unit | 258; old version | Neon Lamp intensity too low | Neon Lamp intensity too low Electrolytes fluctuate in the day in QC
No | Visit | Complaint | RDEU | 15.05.2019 10:00 H. Teufel +49911398118263: Kal-Fehler pO2 27.5.2019/09:15 HerrTeufel 0911398118263 - pO2-Nullwertfehler | pO2 Kal. und QC Fehler Nullstromfehler | Proteinentfernung durchgeführt, Lufteinlass gespült, ABL Datei Archiv überprüft - Kal. und QC, ABL im Service überprüft, Tausch der pO2 Elektrode E799, Messkammer gereinigt, Glasfläche der neuen Elektrode gereinigt, mehrfach kalibriert, Kal. und QC in Ordnung, System- und Funktionskontrolle in Ordnung | SWO-560636 | Assisted Replacement | Electrode | Electrode | pO2 | Zero current error; 408 | 15.05.2019 10:00 H. Teufel +49911398118263: Kal error pO2 27.5.2019/09:15 MrTeufel 0911398118263 - pO2-Nullwertfehler | pO2 Kal. und QC Fehler Nullstromfehler
No | Visit | Complaint | RIND | Error SETSAFE 22 | Setsafe22 | Run Safe position in service program. Shutdown the Analyzer . Clean the Storage wheel. Clean the sensor. Restart the analyzer but error remain same | SWO-565006 | Revisit required | HW | Process Module | Motor, punch out | setsafe22 | Error SETSAFE 22 | Setsafe22
No | Visit | Complaint | RIND | Hct Filter Clot | Checked and found that HCT filter clot, | Cleaned the HCT filter and new lot Calibration as done. | SWO-561570 | Cleaned | HW | Wet section | Wet section | Waste filter | Hct Filter Clot | Checked and found that HCT filter clot,
No | Visit | Complaint | RPOL | Analizator przerywa badanie w trakcie suszenia z powodu błędu temperatury. | Weryfikacja komunikatów błędów. Stwierdzono dodatkowo mechaniczne uszkodzenie obudowy monitora. Ekran działa prawidłowo, dotyk również. | Ustawienie analizatora na przygotowanych podporach w celu poprawy wentylacji w obszarze pracy analizatora. Mechaniczne uszkodzenie monitora. Zamówiono monitor. | SWO-561659 | User Error | Temperature | Temperature | Temperature | Temp. In environment too high | The analyser interrupts the test during drying due to a temperature error. | Error message validation. Additionally, mechanical damage to the monitor casing was observed. The screen works properly, Touch also.
No | Visit | Complaint | RPOL | Aparat nie aspiruje kontrolki QC z AutoCheck | Weryfikacja dzienników danych, sprawdzenie igły - igła krzywa, nie trafia w ampułkę QC. | Wymiana igły. Wymiana uszczelki wlotu. Kalibracje. Test AutoCheck - kontrolka oznaczona. | SWO-559915 | Defect | HW | Inlet module | Probe bend |  | Camera does not aspirate QC control with AutoCheck | Verifying the data logs, checking the needle-needle curve, does not hit the QC Amposette.
No | Visit | Complaint | RDEU | 31.5.2019/11:00 Frau Adam 05115321458 -> Ableitstrom im Anschluß SP erkannt | pH Impedanz Fehler, Ableitstromfehler | Gerät geprüft, SC 90 ersetzt, Funktionskontrolle ok | SWO-560553 | Assisted Replacement | SC | SC | Leak current |  | 31.5.2019/11:00 Mrs. Adam 05115321458 -> Discharge current detected in connection SP | pH Impedance Error, Leakage Current Error
No | Visit | Complaint | RLTD | MACHINE SAYING A PUMP CALIBRATION WAS DETECTED 1178 INSERT SALUTION PACK DOES NOT KNOW WHERE SALUTION PACK HAS TO GO Sister office 02087 251 934 or bleep 6993 | Tested fluid transport is service programs - failing pump calibration, value too low <0.80, also failing clot detection with low values. Tested solution pack on another ABL90 in Champney Ward that was working fine. Pack failed with same errors. Tried 6 other another solution packs of same lot HL-53 - all doing the same fault, pump calibration error. Removed Solution pack from Champney Ward and fitted to Neonatal Unit ABL90 to get this working. Passing fluid transport test now, pump calibration within range at 0.86 (range is 0.80 to 1.20, value when on Champney was 0.88). Verified operation of Neonatal Analyser - all now working and ready for use. | Replaced solution pack with one of a different lot i.e. not lot HL-53. Replacement solution packs for faulty lot HL-53 to be supplied if possible tomorrow Friday 7th June. | SWO-560872 | Assisted Replacement | SP | SP | Pump Cal | 1178 | MACHINE SAYING A PUMP CALIBRATION WAS DETECTED 1178 INSERT SALUTION PACK DOES NOT KNOW WHERE SALUTION PACK HAS TO GO Sister office 02087 251 934 or bleep 6993 | Tested fluid transport is service programs - failing pump calibration, value too low <0.80, also failing clot detection with low values. Tested solution pack on another ABL90 in Champney Ward that was working fine. Pack failed with same errors. Tried 6 other another solution packs of same lot HL-53 - all doing the same fault, pump calibration error. Removed Solution pack from Champney Ward and fitted to Neonatal Unit ABL90 to get this working. Passing fluid transport test now, pump calibration within range at 0.86 (range is 0.80 to 1.20, value when on Champney was 0.88). Verified operation of Neonatal Analyser - all now working and ready for use.
No | Visit | Complaint | RLTD | Richard Barton Tel: 0141 951 5917 | Analyser failing leakage check and various fluidic errors Already changed SC, SP, Gasket and flushed analyser | Replaced inlet probe and connector gasket but still same Replaced pinch valve tuning, now ok | SWO-560653 | Defect | HW | Oxi module | Pinch valve tube |  | Richard Barton Tel: 0141 951 5917 | Analyser failing leakage check and various fluidic errors Already changed SC, SP, Gasket and flushed analyser
No | Visit | Complaint | RLTD | 03/06/19 - Now failing on bilirubin on QC's please phone Debbie Donovan on 02381 206721 Jose on 02381 206721 04.06.2019 15:38pm lost oxsymmetry values on patient results Deborah Donovan 02381 206721 | Spectrometer noise failure - replaced spectrometer and PSU however noise still persisting. | Arranged for wet section board to be replaced. | SWO-560745 | Revisit required | HW | Basic Unit | PCB, wet section |  | 03/06/19 - Now failing on bilirubin on QC's please phone Debbie Donovan on 02381 206721 Jose on 02381 206721 04.06.2019 15:38pm lost oxsymmetry values on patient results Deborah Donovan 02381 206721 | Spectrometer noise failure - replaced spectrometer and PSU however noise still persisting.
No | Visit | Complaint | RSCH | Einlass Undicht | Einlass undicht | Einlass Nadel ausgetauscht Einlassdichtung ausgetauscht Funktionskontrolle i.o | SWO-560566 | Defect | HW | Auto Inlet module | Inlet gasket |  | Entrance leak | Entrance leaks
No | Visit | Complaint | RDAN | Jane kan ikke tilgå ABL90 (NY) via remote. Den melder fejl i windows, og der kommer hvid skærm. Det er forsøgt at frakoble Proxy Host samt genstart af apparat. | . | CF kort formateret, win 7 geninstalleret, ver. 3.4MR1 geninstalleret, opsætning geninstalleret, afprøvet, hurtigadgang fra Aqure virker! | SWO-561088 | Defect | SW | RMED SW | RMED SW | Aqure quick access | Jane cannot access ABL90 (NEW) via remote. It reports errors in Windows, and a white screen appears. It is attempted to disconnect the Proxy Host as well as restarting the device. | .
No | Visit | Complaint | RCHN | 无法开机 | 无法开机，借902-970一个 | 更换902-970一个正常。(by chuanwei) | SWO-559723 | [5 empty] | Cannot power on | Cannot power on, borrow 902-970 one
No | Visit | Complaint | RCHN | lac值质控偏高 | lac值质控偏高 | 重新做质控后修复(by chuanwei) | SWO-560501 | [5 empty] | Lac-value quality control is high | Lac-value quality control is high
No | Visit | Complaint | RAME | Solution pack not recognized by analyzer, Blockage error | Analyzer was locked, system error message 2106 indicating blockage had occurred. performed visual inspection of analyzer, did not observe leaks. Verified SP was replaced day before. | Performed manual system cycle test which failed. Performed diagnostics test - LS calibration failed. Re-initialized analyzer and performed diagnostics. LS test passed however Hemolyzer test failed cavitation absorb. Unassembled wet section, noting fluidics tubing between manifold block and hemolyzer pump separated. Dried QC material observed in wet section area. Removed dried residue materials and installed replaced all tubings and pumps using PM kit 905-896. Initialized analyzer and performed manual system cycle tests and external QC's. All tests passed. Analyzer is performing to specifications | SWO-560871 | [5 empty] | Solution pack not recognized by analyzer, Blockage error | Analyzer was locked, system error message 2106 indicating blockage had occurred. performed visual inspection of analyzer, did not observe leaks. Verified SP was replaced day before.
Yes | Visit | Complaint | RAME | Solution pack not recognized by analyzer, Blockage error | Manual System Cycle and QC Data for June 6 service visit | Manual System Cycle and QC Data for June 6 service visit. *Having connectivity issues with network so having difficulty uploading large files... | SWO-560910 | Re classified RMED | Same day UpSWO |  |  |  | Solution pack not recognized by analyzer, Blockage error | Manual System Cycle and QC Data for June 6 service visit
No | Visit | Complaint | RAME | All oximetry parameters failing system cycles. | All OXI parameters are failing System Cycle. Customer has replaced both the SC and SP with same results. Unable to perform THB calibrations or blank Calibrations. Possible clot or air leak in the tubing. | Verified all actions performed by the customer. Installed complete Annual Maintenance Kit for possible air contamination. Flushed the Hemolyzer for possible clot. Visual and operational inspection of pinch valve. Started the analyzer with a previously failed SC with passing results. System Cycle, THB calibration and QC to verify performance. | SWO-560835 | [5 empty] | All oximetry parameters failing system cycles. | All OXI parameters are failing System Cycle. Customer has replaced both the SC and SP with same results. Unable to perform THB calibrations or blank Calibrations. Possible clot or air leak in the tubing.
No | Visit | Complaint | RDEU | Qualitätskontrollen fehlerhaft in pCO2 und pH. | Qualitätskontrollen fehlerhaft in pCO2 und pH. Von 3 kurz hintereinander eingelegten QCs des gleichen Levels sind 2 fehlerhaft, einer gut. | Gerät mit Nullpunktkassette überprüft und justiert. Alle Werte bereits nach kurzer Zeit stabil. Lösungspack erneuert. Weiterhin identisches Verhalten. Ersatzteile bestellt. | SWO-560853 | [5 empty] | Quality controls faulty in pCO2 and pH. | Quality controls faulty in pCO2 and pH. Of 3 QCs of the same level inserted in quick succession, 2 are faulty, one is good.
No | Visit | Complaint | RCHN | pco2质控A运行不过 | pco2质控A运行不过（未更换备件） | 多做几个质控就通过了(by chuanwei) | SWO-560431 | Adjustment | SC | SC | pCO2 | QC A | pco2 QC A does not run | pco2 QUALITY CONTROL A DOES NOT OPERATE (NO REPLACEMENT SPARE PARTS)
No | Visit | Complaint | RCHN | 光谱不匹配.报错：0581 | 光谱不匹配.报错：0581，需要借件902-980一个 | 更换902-980后修复(by chuanwei) | SWO-559028 | Defect | HW | Oxi module | Hemolyzer | 581 | Spectrum mismatch.Report error: 0581 | Spectroscopic mismatch.False: 0581, need to borrow 902-980 one
No | Visit | Complaint | RCHN | 屏幕提示手动冲洗管路 | 屏幕提示手动冲洗管路（未更换备件 ） | 冲洗管路，重新定标、质控(by chuanwei) | SWO-560434 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | Screen prompt scoffing lines manually | Screen prompt scoffing the line manually (spare parts not replaced)
No | Visit | Complaint | RCHN | 气流温度超范围 | 气流温度超范围，借848-037一个 | 更换848-037后修复(by chuanwei) | SWO-559133 | Defect | HW | Process Module | Membrane Pump |  | Airflow temperature over range | Airflow temperature over range, 848-037 one
No | Visit | Complaint | RCHN | 处理包无法识别 | 处理包无法识别，需要借件903-036一个 | 更换903-036备件后修复(by chuanwei) | SWO-558886 | Defect | Consumables | Solution Pack | Chip | 1052 | Processing package is not recognized | Processing package is not recognized, need to borrow 903-036 one
No | Visit | Complaint | RCHN | 仓门无法开启 | 仓门无法开启，借903-717一个 | 更换903-717备件后修复(by chuanwei) | SWO-559083 | Defect | HW | Cover and Chassis | Door, Inlet | Top bracket | The door cannot be opened | The door cannot be opened, borrow 903-717 a
No | Visit | Complaint | RCHN | 仓门无法关闭 | 仓门无法关闭，借903-717一个 | 更换903-717备件后修复(by chuanwei) | SWO-559728 | Defect | HW | Cover and Chassis | Door, Inlet | Top bracket | The warehouse door cannot be closed | The door cannot be closed, borrow 903-717 one
Yes | Visit | Complaint | RIND | Service Visit. | No air detected during measurement, need to replace the sensor cassette. | Replace the sensor cassette, successfully done, user do the sample measurement in analyzer, successfully measurement comes. Analyzer handed over to staff in well working condition. | SWO-561678 | [5 empty] | Service Visit. | No air detected during measurement, need to replace the sensor cassette.
No | Visit | Complaint | RCHN | PO2无端点 | PO2无端点，借903-612一个 | 更换903-612后修复(by chuanwei) | SWO-558958 | [5 empty] | PO2 Unprovoked Point | PO2 No.
No | Visit | Complaint | RIND | Printing Problem | Checked and found that Problem in printer head and Gasket to be cleaned. | Replaced the Printer head assembly, and cleaned the waste tubes and gasket, Now machine works good. | SWO-561571 | Defect | HW | Printer Module | Printer Head |  | Printing Problem | Checked and found that Problem in printer head and Gasket to be cleaned.
No | Visit | Complaint | RCHN | 总血红蛋白定标漂移，报错0376 | 总血红蛋白定标漂移，报错0376，未更改备件 | 重新定标后修复(by chuanwei) | SWO-560339 | Adjustment | HW | Oxi module | tHb Cal | 376 | Total hemoglobin calibration drift, error 0376 | Total hemoglobin calibration drift, error 0376, unchanged spare parts
No | Visit | Complaint | RFRA | Pas de mesure oxy | Contrôle module oxy | Contrôle du module oxymétrie,constate blocage de la pinch valve $ $Démontage ,nettoyage et lubrification de la pinch valve $ $Test pinch valve ok $ $Calibration de la tHB $ $Contrôle de la tHB ok $ $Calibration deux points et passage des 4 niveaux de CQ . | SWO-561171 | Cleaned | HW | Oxi module | Pinch valve |  | No oxy measurement | Oxy module control
Yes | Visit | Complaint | RDAN | Oxi parameter uden for reportable range, blev reporteret med .... og pile. Ønske at istedet blev reporteret med værdi >< og værdi på grændsen. Romai.ini rettet / ShowValuesOutOfReportableRange=1 / Dette blev testet. | Rettet Roma INI. | ShowValuesOutOfReportableRange=1 | SWO-561263 | Re classified RMED | Request |  |  |  | Oxi parameter outside the reportable range, was reported with.... and arrows. Wish that instead was reported with value > < and value on the grout. Romai. ini corrected/ShowValuesOutOfReportableRange = 1/This was tested. | Save Roma INI.
Yes | Visit | Complaint | RDAN | Oxi parameter uden for reportable range, blev reporteret med .... og pile. Ønske at istedet blev reporteret med værdi >< og værdi på grændsen. Romai.ini rettet / ShowValuesOutOfReportableRange=1 / Dette blev testet. | Oxi parameter uden for reportable range, blev reporteret med .... og pile. Ønske at istedet blev reporteret med værdi >< og værdi på grændsen. Romai.ini rettet $ ShowValuesOutOfReportableRange=1 $ Dette blev testet. | Oxi parameter uden for reportable range, blev reporteret med .... og pile. Ønske at istedet blev reporteret med værdi >< og værdi på grændsen. Romai.ini rettet $ ShowValuesOutOfReportableRange=1 $ Dette blev testet. | SWO-561257 | Re classified RMED | Request |  |  |  | Oxi parameter outside the reportable range, was reported with.... and arrows. Wish that instead was reported with value > < and value on the grout. Romai. ini corrected/ShowValuesOutOfReportableRange = 1/This was tested. | Oxi parameter outside the reportable range, was reported with.... and arrows. Wish that instead was reported with value > < and value on the grout. Romai. ini corrected / ShowValuesOutOfReportableRange = 1 / this was tested.
No | Visit | Complaint | RDEU | 06.06.2019 9:00 Fr. Seitz +49911398114183: Gerät startet nicht mehr | ABL bootet nicht weiter Fehlermeldung... roma.ini failed | Software neu aufgespielt - clean install, Konfiguration von einem anderen Gerät aufgespielt, Druckereinstellungen überprüft, alle Elektroden membraniert, tHb Kal. durchgeführt, Konfig. überprüft, Init.Phase abgewartet, mehrfach kalibriert, QC Chargen eingelesen, manuell alle 3 Level der QC wiederholt, System- und Funktionskontrolle in Ordnung | SWO-561227 | Defect | SW | RMED SW | RMED SW |  | 06.06.2019 9:00 Fr. Seitz +49911398114183: Device no longer starts | ABL no longer boots Error message... roma.ini failed
No | Visit | Complaint | RFRA | Le QC lvl1 de NT-Pro est hors gamme bas. Il est conseillé de passer 4 tubes de nettoyages et de refaire un ajustement cal sur le lot du paramètre. | $ | Réalisation de la VPS, voir le SWO sur case Halfyearly. | SWO-561454 | Adjustment | HW | Aspiration Module | Needle | also see SWO-561455 | The NT-Pro lvl1 QC is off-range low. It is advisable to pass 4 cleaning tubes and redo a cal adjustment on the lot of the setting. | /
Yes | Visit | Complaint | RFRA | Les QC NCC1 et Bhcg ne passent pas correctement. Demande de passage de cleaning x5. Attente resultats. Nouvel appel : Le probleme persiste. Lots utlisés : Bhcg 15459 Niv 1 et Ntpno 09171 Niv 1 | Réalisation de la VPS en avance. | Voir SWO sur case VPS. Le pb de QC NTpro venait très certainement du vieux lot de test NTpro. | SWO-562616 | Re classified RMED | Same day UpSWO |  |  |  | The NCC1 and Bhcg QC do not pass properly. Request for cleaning x5. Waiting results. New call: The problem persists. Used Lots: Bhcg 15459 Niv 1 and Ntpno 09171 Niv 1 | Achieving the VPS in advance.
No | Visit | Complaint | RIND | Service Call and Visit. | Need to replace the gear mechanism of the inlet door. | Replace the gear mechanism of the analyzer, successfully done, analyzer handed over to staff in well working condition. | SWO-561672 | Defect | HW | Cover and Chassis | Door, Inlet |  | Service Call and Visit. | Need to replace the gear mechanism of the inlet door.
No | Visit | Complaint | RPOL | Błąd transportu płynów, błąd kalibracji pompy. | Błąd transportu płynów, błąd kalibracji pompy. | Przepłukanie toru pomiarowego, wymiana uszczelki wlotu. Zainstalowano nowa SC90 lot R0909N173. | SWO-560674 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | Fluid transport error, pump calibration error. | Fluid transport error, pump calibration error.
No | Visit | Complaint | RDEU | Anwender hat, die SC wegn O2 und CL Kalfehlern ausgetauscht, Lösungspack wurde 2 mal erneuert, Gerät meldet immer wieder Leckage erkannt. FSE Einsatz erwünscht. | Gerät überprüft, Servicetests, Leckage im System | Einlassdichtung erneuert, Servicetests i.O, Startphase i.O. Funktionsprüfung i.O | SWO-561960 | Defect | HW | Inlet module | Inlet gasket | Leak | User has replaced the SC away O2 and CL cal errors, solution pack has been renewed 2 times, device reports repeatedly leakage detected. FSE use desired. | Device checked, service tests, leakage in the system
No | Visit | Complaint | RAME | Costumer states that the analyzer is giving a boot up error | Verified customer's stated issue. | Replaced analyzers CF card and installed the relative software. | SWO-561501 | Defect | HW | DMS_PC | CF card |  | Costumer states that the analyzer is giving a boot up error | Verified customer's stated issue.
Yes | Visit | Complaint | RDEU | 06.06.2019; Hr. Meister; Tel.: +4962138371427; Fehler Ref. Elektrode in SC | SC defekt | Lieferung Sensor-Kassette 946-005 (LOT R1552) Defekte SC von Kunde reklamiert | SWO-561121 | Re classified RMED | Logistical SWO |  |  |  | 06.06.2019; Mr. Meister; Tel. +4962138371427; Error Ref. Electrode in SC | SC defective
Yes | Visit | Complaint | RLTD | MACHINE SAYING A PUMP CALIBRATION WAS DETECTED 1178 INSERT SALUTION PACK DOES NOT KNOW WHERE SALUTION PACK HAS TO GO Sister office 02087 251 934 or bleep 6993 | Delivered direct to POC 11 x 944-457 lot HL-51 as replacements for the 11 faulty 944-457 of lot HL-53. | Delivered replacement Solution Packs. | SWO-561442 | Re classified RMED | Logistical SWO |  |  |  | MACHINE SAYING A PUMP CALIBRATION WAS DETECTED 1178 INSERT SALUTION PACK DOES NOT KNOW WHERE SALUTION PACK HAS TO GO Sister office 02087 251 934 or bleep 6993 | Delivered direct to POC 11 x 944-457 lot HL-51 as replacements for the 11 faulty 944-457 of lot HL-53.
No | Visit | Complaint | RDEU | 06.06.2016, 09:32 Uhr, Fr. Butter, Tel. +4973915865635: Chipdaten nicht lesbar 07.06.2019; 12:50 Uhr, Fr. Butter; Tel.: +4973915865635; weiterhin Chipdaten nicht lesbar | . | Beide Sol.Pack check--Chip Data nicht lesbar. SP ID PCB ersetzt. Neuer Versuch mit SP aus VSC Startphase abgewartet. Kal ok QC Level C --PCO2 0,5 mmHg zu tief. Kein Zubehör mehr vorhanden. Kunde unterrichtet. | SWO-563133 | Defect | HW | Basic Unit | PCB, solution pack | Chip | 06.06.2016, 09:32, Fr. Butter, Tel. +4973915865635: Chip data unreadable 07.06.2019; 12:50 p.m., Fr. Butter; Tel.: +4973915865635; chip data still unreadable | .
No | Visit | Complaint | RAME | Remote connection will not work on analyzer. | Quick access through netops not working. Tried reinstalling Roma, 3.4 SW. Reinstalled Win7. Reinstalled 3.4. Restored setup and backup. | Reinstalled all software on analyzer. | SWO-561435 | Defect | SW | RMED SW | RMED SW | Aqure quick access | Remote connection will not work on analyzer. | Quick access through netops not working. Tried reinstalling Roma, 3.4 SW. Reinstalled Win7. Reinstalled 3.4. Restored setup and backup.
No | Visit | Complaint | RAME | Unable to Quick Access | Unable to quick access the analyzer from AQURE. Reinstalled win 7 software, 3.4 and reloaded setup and backup. | Reinstalled windows 7. | SWO-561516 | Defect | SW | RMED SW | RMED SW | Aqure quick access | Unable to Quick Access | Unable to quick access the analyzer from AQURE. Reinstalled win 7 software, 3.4 and reloaded setup and backup.
No | Visit | Complaint | RLTD | 03/06/19 - Now failing on bilirubin on QC's please phone Debbie Donovan on 02381 206721 Jose on 02381 206721 04.06.2019 15:38pm lost oxsymmetry values on patient results Deborah Donovan 02381 206721 | Spectrum Mismatch errors on Oxi parameters and the complete Oxi test failed due to spectrometer noise. The hemolyzer, spectrometer and power supply have already been replaced but they did not resolve the problem. | I replaced the wet section control PCB and updated the software to v3.4MR1. When the analsyer had started up again I ran the complete Oxi test and it passed all parameters, the noise value was 0.72mV (0.30-1.80 is the range). A calibration was then run and passed, a tHb calibration was performed and that passed and all 3 levels of QC run and passed. | SWO-562594 | Defect | HW | Basic Unit | PCB, wet section |  | 03/06/19 - Now failing on bilirubin on QC's please phone Debbie Donovan on 02381 206721 Jose on 02381 206721 04.06.2019 15:38pm lost oxsymmetry values on patient results Deborah Donovan 02381 206721 | Spectrum Mismatch errors on Oxi parameters and the complete Oxi test failed due to spectrometer noise. The hemolyzer, spectrometer and power supply have already been replaced but they did not resolve the problem.
No | Visit | Complaint | RCHN | 泵定标出错，测试卡安装失败，报错1167 | 泵定标出错，测试卡安装失败，报错1167，借903-129一个 | 更换903-129后修复(by chuanwei) | SWO-560612 | Defect | HW | Oxi module | Pinch Valve | 1167 | Pump calibration error, test card installation failure, error 1167 | Pump calibration error, test card installation failure, error 1167, borrow 903-129 one
No | Visit | Complaint | RCHN | 进样口支架密封不严，报泄漏 | 进样口支架密封不严，报泄漏 | 更换进样口支架后修复(by chuanwei) | SWO-561006 | Defect | HW | Inlet module | Inlet gasket |  | The inlet bracket seal is not strict, the leak is reported | The inlet bracket seal is not strict, the leak is reported
No | Visit | Complaint | RCHN | 无法开机，系统软件不匹配 | 无法开机，系统软件不匹配，借910-343和903-198各一个 | 更换910-343和903-198后修复(by chuanwei) | SWO-559730 | Defect | HW | Analyzer Control Module | CPU | 0002 | Cannot power on, system software does not match | Cannot power on, system software does not match, borrow 910-343 and 903-198 each one
No | Visit | Complaint | RCHN | 进样门无法关闭 | 进样门无法关闭，借903-717一个 | 更换903-717后修复(by chuanwei) | SWO-559773 | Defect | HW | Cover and Chassis | Door, Inlet | Top bracket | The sample door cannot be closed | The sample door cannot be closed, borrow 903-717 one
No | Visit | Complaint | RIND | Patient data log full . | Patient data log full . | Checked data logs , exported and deleted the patient data , trained the technician and biomedical staff about exporting the data . Instrument working fine , customer happy with instrument . | SWO-561426 | [5 empty] | Patient data log full . | Patient data log full .
No | Visit | Complaint | RIND | Smoke coming out from backside of the analyser. | Check system | check analyser open back cover burning smell coming from power supply , open power supply unit found it gets damaged, need tobe replaced. | SWO-561589 | [5 empty] | Smoke coming out from backside of the analyser. | Check system
No | Visit | Complaint | RKK | QCで電解室が日内変動 | QCで電解室が日内変動 | 比較電極交換 | SWO-561726 | User expectation | Electrode | Electrode | reference | QC values fluctuated within RMED specifications | The electrolysis chamber fluctuates in the day in QC | The electrolysis chamber fluctuates in the day in QC
No | Visit | Complaint | RIND | Na and K status out of range. | Checked the calibration results, Na and K status reported out of range. | Cleaned all electrode chambers. Cross checked reference electrode. performed 2point calibration. | SWO-561908 | Cleaned | HW | El/Met Module | measuring chamber |  | Na and K status out of range. | Checked the calibration results, Na and K status reported out of range.
Yes | Visit | Complaint | RDEU | pH-Empf.-keit bei 81% | Lieferung pH-Elektrode, Rücknahme defekte Elektrode | neue pH-Elektrode geliefert, alte $defekte pH-Elektrode zurückgenommen (BU-01) | SWO-562337 | Re classified RMED | Logistical SWO |  |  |  | pH absorption at 81% | Delivery of pH electrode, return of defective electrode
No | Visit | Complaint | RIND | Error SETSAFE 22 | SETSAFE22 | Run Safe position and shutdown the analyzer. Change the stepper motor punchout complete. Restart the analyzer. Machine is ready. | SWO-565013 | Defect | HW | Process Module | Motor, punch out | setsafe22 | Error SETSAFE 22 | SETSAFE22
No | Visit | Complaint | RIND | Lac value not calculated in patient report. | Checked the patient report, lac was not calculated in report due to low Po2 value. | Checked the calibration report, Calibration was passed. Runed a fresh sample and get the result of lac. | SWO-564522 | [5 empty] | Lac value not calculated in patient report. | Checked the patient report, lac was not calculated in report due to low Po2 value.
No | Visit | Complaint | RCHN | K定标不过 | K定标不过，借903-581一个 | 更换903-581后修复(by chuanwei) | SWO-560347 | Defect | HW | Basic Unit | PCB, Sensor interface | K | K calibration but | K calibration, however, borrow 903-581 a
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 参比电极错误 1361 | 更换测试卡后修复 | SWO-567767 | [5 empty] | Test card, reagent package failure | Reference electrode error 1361
Yes | Visit | Complaint | RSCH | Pikett Crea | Kunde hat am 07.06.2019 Crea Membranen Express bestellt Am 08.06.2019 aber falsche Lieferung erhalten Kunde wünscht die sofortige Lieferung der Membranen, ist aber nicht bereit Wochenendpauschale zu bezahlen | Membranen Crea in Thalwil besorgt Membranen ausgeliefert | SWO-561600 | Re classified RMED | Shipping |  |  |  | Pikett Crea | Customer announced on 07.06.2019 Crea Membranen Express ordered On 08.06.2019 but received incorrect delivery Customer wishes the immediate delivery of the membranes, but is not willing to pay weekend lump sum
No | Visit | Complaint | RCHN | 孵育盘定标失败，报错2168 | 孵育盘定标失败，报错2168，借636-536,903-138,903-215各一个 | 更换903-138,903-215后修复 by Andrew | SWO-561101 | Defect | HW | Incubator Module | PCB, incubator | 2168 | Incubation plate calibration failed, 2168 error | Incubation plate calibration failed, reported error 2168, borrowed 636-536, 903-138, 903-215 each
No | Visit | Complaint | RKK | 液漏れ | 液漏れ | 溶液パック初期不良 | SWO-564777 | Assisted Replacement | SP | SP | Leak |  | Leakage | Leakage
Yes | Visit | Complaint | RAME | Calibrations failing due to THB drift and inhomogeneous cal 3 QC failing with 581 and 291 errors | THB Drift errors. | Returning the analyzer after monitoring for THB drift since 5-30-2019. The analyzer has not failed any calibrations since the original failure. Analyzer will remain idle until the Go Live date. Calibrations and QC to verify performance at start up. | SWO-562095 | Re classified RMED | Moved analyzer |  |  | see SWO-557035 | Calibrations failing due to THB drift and inhomogeneous cal 3 QC failing with 581 and 291 errors | THB Drift errors.
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 气泡探测定标失败 | 更换测试卡后修复 | SWO-564002 | [5 empty] | Test card, reagent package failure | Bubble probe calibration failed
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | K传感器阻抗错误1079 | 更换测试卡后修复 | SWO-567765 | [5 empty] | Test card, reagent package failure | K sensor impedance error 1079
No | Visit | Complaint | RCHN | 测试卡无法扫描 | 测试卡无法扫描，借910-304一个 | 更换910-304一个正常。by Andrew | SWO-561640 | Defect | HW | Storage Module | Barcode reader, storage wheel | BR001 | Test card can't be scanned | Test card can not be scanned, borrow 910-304 one
No | Visit | Complaint | RCHN | 处理包智能芯片无法识别 | 处理包智能芯片无法识别 | 老师操作问题导致，讲解后解决 by jesson | SWO-565976 | User Error | Consumables | Solution Pack | Solution Pack |  | Processing package smart chip is not recognized | Processing package smart chip is not recognized
No | Visit | Complaint | RCHN | 无法读取测试卡 | 无法读取测试卡,借905-960一个。 |  | SWO-561904 | Defect | HW | Storage Module | Barcode reader, storage wheel | CRM Corrective Action: Replacement 905-960 is normal. By Andrew | The test card could not be read | The test card could not be read, and 905-960 was borrowed.
No | Visit | Complaint | RCHN | PCT定标失败 1031 | PCT定标失败 1031,借944-394*15754一张 |  | SWO-561905 | Assisted Replacement | Consumables | Test Kit | Cartridge, cal | 1031; CRM Corrective Action: The recalibration is normal. By Andrew | PCT calibration failure 1031 | PCT bid failure 1031, loan 944-394 x 15754
No | Visit | Complaint | RCHN | 标本放进去报错5784 | 标本放进去报错5784（未更换备件） | 清洁传感器灰尘 | SWO-561894 | Cleaned | HW | Storage Module | Sensor, storage wheel | 5784 | The specimen was put in and reported the error 5784. | Specimens put in to report error 5784 (unreplaced spare parts)
No | Visit | Complaint | RCHN | 科室老师反映样本无法识别 | 科室老师反映样本无法识别（未更换备件） | 工程师到现场排查原因，仪器并无故障，再次培训演示样本操作 by Andrew | SWO-565000 | User Error | HW | Inlet module | Tube holder |  | The teacher of the section reflected that the sample was not recognized | The section teacher reflected that the sample was not identified (spare parts were not replaced)
No | Visit | Complaint | RCHN | PO2定标失败 | PO2定标失败 | 更换O2膜后修复 | SWO-561807 | Assisted Replacement | Electrode | Membrane | pO2 | 378 | PO2 calibration failure | PO2 calibration failure
No | Visit | Complaint | RKK | Oxi項目が表示されない | Oxi項目が表示されない ミキサーエラー | ヘモライザー交換 FlexQモジュール分解点検 | SWO-562257 | Defect | HW | Oxi module | Hemolyzer |  | Oxi items are missing | Oxi items are missing Mixer error
No | Visit | Complaint | RAME | Received Cannot find required Map name now analyzer will not boot-up | During boot up analyzer gives error cannot find required map name. Checked connections on CPU. Tried to reinstall win 7 software but found I needed a different version. | Ordered parts. | SWO-562180 | Defect | HW | DMS_PC | CPU | mail: send CPU back to Ejner | Received Cannot find required Map name now analyzer will not boot-up | During boot up analyzer gives error cannot find required map name. Checked connections on CPU. Tried to reinstall win 7 software but found I needed a different version.
No | Visit | Complaint | RPAC | Unable to pump solution | Unable to pump solution. Pump calibration error. | Replaced the inlet gasket with holder. Tested, all ok | SWO-562760 | Defect | HW | Inlet module | Inlet gasket | Unable to pump | Unable to pump solution | Unable to pump solution. Pump calibration error.
No | Visit | Complaint | RDEU | 29.05.2019, 11:56 Uhr ==> Anrufer Fr. Betz, Tel.: 08362500196 ==> Fehler 0581: Oxi spectrum mismatch und Fehler 0292: Trübung zu hoch 07.06.2019, 14:26 Uhr ==> Anrufer Fr.Betz, Tel.: 08362500196 ==> Fehler 0581: Oxi spectrum mismatch bei Patientenmessungen besteht weiterhin | Fehler #1324, Inhomogene Spüllösung sporadisch Oxi Spectrum mismatch, Fehler #0581+ 0292 Hemolyzer macht laute Pfeiffgeräusche kein SP mehr vor Ort | Gerät manuell gespült Pinchvalve im Service überprüft - ok Hemolyzer erneuert, Jahreswartungskit verbaut SP von Intensivstation geholt und eingesetzt tHb kalibriert, Gerät kalibriert und überprüft Blutmessung durchgeführt - ok | SWO-562443 | Defect | HW | Oxi module | Hemolyzer | 581; 292; 1324 | 29.05.2019, 11:56 a.m. ==> Caller Ms. Betz, Tel.: 08362500196 ==> Error 0581: Oxi spectrum mismatch and error 0292: turbidity too high 06/07/2019, 14:26 ==> Caller Fr.Betz, Tel.: 08362500196 ==> Error 0581: Oxi spectrum mismatch in patient measurements still exists | Errors #1324, Inhomogeneous Rinsing Solution sporadically Oxi Spectrum mismatch, error #0581+ 0292 Hemolyzer makes loud whistles no more SP on site
No | Visit | Complaint | RDEU | 04.06.2019 13:00 H. Simon +493491502099: Gerät verbraucht zu schnell Zyklen aus LP 07.06.2019, 15:03 Uhr ==> Anrufer Herr Simon, Te.: 03491502099 ==> Solutionpack ist sehr schnell leer und die Kalibration wird automatisch sehr häufig durchgeführt | ABL verbraucht zu viel Lösungspack Aktivitäten Fehler 1406 Analysator steht nicht waagerecht | SC wurde am 13.05. wurde am gewechselt => seitdem Fehler 1406 => ABL führte nach jeder Meldung zusätzliche Aktivität durch => 3 LPs innerhalb jeweils einer Woche leer nach SC Wechsel am 05.06. keine Fehlermeldung mehr und normales Lösungsverbrauch | SWO-563151 | No fault found | SC | SC | Na | 1406; Analyzer is not horizontal | 06/04/2019 13:00 H. Simon +493491502099: Device consumes cycles from LP too fast 06/07/2019, 15:03 ==> Caller Mr. Simon, Te.: 03491502099 ==> Solutionpack is empty very quickly and the calibration is carried out very frequently automatically | ABL consumes too much solution pack activities Error 1406 Analyzer is not horizontal
Yes | Visit | Complaint | RDEU | 11.06.2019 8:15 H. Schölling +4933196822241: Probleme in den Fließwegen | Fehler im Fluid transport Programm Solution Pack Check | Fluid transport Programm durchgeführt, aufgef. Verschleißteile ersetzt, gleicher Fehler, Abbruch der Arbeiten, Gerät mitgenommen zur Reparatur im VSC NO | SWO-562985 | Re classified RMED | analyzer removed |  |  |  | 11.06.2019 8:15 H. Schölling +4933196822241: Problems in the flow paths | Error in Fluid transport Program Solution Pack Check
No | Visit | Complaint | RAME | The display housing is broken allowing the touch screen to slide around. | Broken housing on touchscreen. | Replaced touchscreen assembly | SWO-563019 | Defect | HW | Display | Display |  | The display housing is broken allowing the touch screen to slide around. | Broken housing on touchscreen.
No | Visit | Complaint | RAME | Broken display housing. | Broken housing on the touchscreen. | Replaced housing assembly. | SWO-563020 | Defect | HW | Display | Display |  | Broken display housing. | Broken housing on the touchscreen.
No | Visit | Complaint | RAME | The METHB is not passing their comparisons. | Ask the customer to explain the problem. The customer showed me the EP Evaluator report on the METHB. | Checked the thb calibration. Performed a complete oxi test and a thb cal. Ran blood samples between the ABL90 and the ABL825's. The ABL90 was almost an exact match with analyzer number 2. | SWO-562969 | Adjustment | HW | Oxi module | tHb Cal | MetHb | The METHB is not passing their comparisons. | Ask the customer to explain the problem. The customer showed me the EP Evaluator report on the METHB.
No | Visit | Complaint | RKK | 透析液測定でNa値が高い(145～153meq/L)。SC交換後も変化なし。SP交換後142meq/Lまで下がる。 | 透析液測定でNa値が高い(145～153meq $L)。 | SC交換後も変化なし。SP交換後142meq $Lまで下がる。 | SWO-568031 | [5 empty] | Na value is high in dialysate measurement (145~153meq/L). No change even after SC exchange. Drops to 142meq/L after SP replacement. | Na value is high in dialysate measurement (145~153meq/L).
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | po2质控值偏高 | 更换测试卡后修复 | SWO-562343 | [5 empty] | Test card, reagent package failure | Po2 QC value is high
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | po2质控值偏高 | 更换测试卡后修复 | SWO-562342 | [5 empty] | Test card, reagent package failure | Po2 QC value is high
No | Visit | Complaint | RDEU | Qualitätskontrollen fehlerhaft in pCO2 und pH. | Üerbrückungsgerät zur Verfügung gestellt. | ABL80 Basic 304528 aufgestellt und eingerichtet. Alte Daten gelöscht. Bei ERsatzgerät alle Werte innerhalb Referenzbereich. Bei 317085 unreproduzierbare Sprünge pCO2. | SWO-562665 | [5 empty] | Quality controls faulty in pCO2 and pH. | bridging device.
No | Visit | Complaint | RCHN | 测试卡断裂 | 测试卡断裂，断裂部位TEETH（未更换备件） | 更换测试卡后修复 by Andrew | SWO-564055 | Defect | Consumables | Test Kit | Cartridge, Broken | Teeth | Test card break | TEST CARD BREAK, FAULT SITE TEETH (NOT REPLACED SPARE PARTS)
No | Visit | Complaint | RCHN | 进样门关不上 | 进样门关不上（未更换备件） | 清洁传感器 by Andrew | SWO-562447 | Cleaned | HW | Cover and Chassis | Door, Inlet |  | The entrance door doesn't close. | The sample door is not closed (no spare parts replaced)
No | Visit | Complaint | RCHN | 处理包卡盘 | 处理包卡盘 | 重启分析仪后修复 by Andrew | SWO-562306 | Defect | Consumables | Solution Pack | Cup Stack |  | Handling the package chuck | Handling the package chuck
No | Visit | Complaint | RDEU | 11.06.2019 Frau Scholz +4933312713146 meldet, dass der Bildschirm sich nicht bedienen lässt. | Anschluß am Touchscreen korrodiert | Touchscreen gewechselt - Test i.O. Einsatz ist kostenpflichtig, da Feuchtigkeit im Monitor | SWO-562788 | User Error | HW | Analyzer Control Module | Display | Moisture | 11.06.2019 Ms. Scholz +4933312713146 reports that the screen cannot be operated. | Connection to the touchscreen corroded
No | Visit | Complaint | RIND | In result message optical Unit sensitivity error | Find Error message , Optical unit senstivity | Run Safe position in service program . Shutdown the analyzer. Clean the Process wheel and clean the optical unit bottom part. Restart the analyzer. | SWO-564779 | Cleaned | HW | Process Module | Process Wheel | Optical Unit sensitivity | In result message optical Unit sensitivity error | Find Error message , Optical unit senstivity
No | Visit | Complaint | RDEU | 07.06.2019 Frau Möst +49898412556 berichtet, das ein Patient falsch Positiv gemessen wurde. | Patient falsch Positiv gemessen wurde. | Jahreswartung durchgeführt, Blank Messung Photo counts 650-700, Optical Unit erneuert, Alle Parameter kalibriert, Blankmessungen vorgenommen. Servicedump erstellt und überprüft. Fließwege in Ordnung (1172 - 1321). Blankmessungen in Ordnung. Systemüberprüfung und Systemreinigung - ok. 3xLQC MCC2, 1x LQC CRP2 und 1x LQC D-dimer gemessen - ok. | SWO-562896 | Defect | HW | Process Module | Optical Unit | Tn F+ | 07.06.2019 Ms. Möst +49898412556 reported that a patient was wrongly measured positively. | patient was wrongly measured positively.
No | Visit | Complaint | RPOL | Błąd kalibracji pompy | Błąd kalibracji pompy po wymianie SP90 | Ponowna instalacja pakietu roztworów, wymiana uszczelki wlotu. Wykonanie testów przepływów w menu serwisowym | SWO-562639 | Adjustment | SP | SP | Pump Cal | SP Re-installed | Pump Calibration Error | Pump calibration error after replacing SP90
No | Visit | Complaint | RSCH | Blank | Neustart des Analysators nach 10 Tagen Ausfallzeit aufgrung eines SP (Solutions Pack) Lieferproblems nicht möglich Christalisierung im Fluidteil | Reinigung des Fluidteils i.o Maintenance kit ausgetauscht i.o 1X SP HL-55 und 1X SC R0751-79 ins Garantie i.o Kalibration i.o QC i.o | SWO-563295 | Defect | HW | Basic Unit | PM kit | Potential User Error | Blank | Unable to restart the analyzer after 10 days of downtime to raise an SP (Solutions Pack) delivery problem Christalization in the Fluid Part
No | Visit | Complaint | RESP | Erro de calibracao de bombas | Verificação de valor de sonda torta e entrada suja. | Substituição da sonda e da entrada, limpeza do circuito e verificação da bomba sem sucesso, valor de calibracao por volta dos 0,79. Substituição do Pack de soluções , testes ok, estabilização e passagem do QC | SWO-563632 | Assisted Replacement | SP | SP | Pump Cal | 1178 | Pump Calibration Error | Check for pie probe value and dirty entry.
No | Visit | Complaint | RDEU | 11.06.2019 8:15 H. Schölling +4933196822241: Probleme in den Fließwegen | Drucktestfehler | aufgef. Verriegelungsmotor ersetzt, Analysator wieder vor Ort installiert,Kal, QC Datensicherung Kal QC | SWO-563696 | Defect | HW | Basic Unit | Actuator unit | 1350, 1351, 1352 (Clot) | 11.06.2019 8:15 H. Schölling +4933196822241: Problems in the flow paths | Drucktestfehler
No | Visit | Complaint | RAME | Turbidity too high errors on analyzer | Error 292 turbidity too high errors on analyzer. Investigated and found that the oxi complete test failed cavitation absorb range. Cycled the pinch valve and stretched out the pinch tubing. Re-ran the complete oxi test and all passed. Ordered pinch valve for Thursday. | Analyzer down. | SWO-563089 | Revisit required | HW | Oxi module | Hemolyzer | Turbidity too high | Turbidity too high errors on analyzer | Error 292 turbidity too high errors on analyzer. Investigated and found that the oxi complete test failed cavitation absorb range. Cycled the pinch valve and stretched out the pinch tubing. Re-ran the complete oxi test and all passed. Ordered pinch valve for Thursday.
No | Visit | Complaint | RKOR | Sample clot error | , | Sample clot error | SWO-568681 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway | Clot | Sample clot error | ,
No | Visit | Complaint | RDEU | Qualitätskontrollen fehlerhaft in pCO2 und pH. | QC pCO2 häufig fehlerhaft. | Sensor board erneuert. | SWO-563440 | [5 empty] | Quality controls faulty in pCO2 and pH. | QC pCO2 often faulty.
No | Visit | Complaint | RCHN | BNP定标失败，报错1025 | BNP定标失败，报错1025需要借944-258 15802 缺货 | 使用其他批号的定标卡定标后修复 by Andrew | SWO-562275 | Assisted Replacement | Consumables | Test Kit | Cartridge, cal | 1025 | BNP bid failure, error 1025 | BNP bid failure, misquoted 1025 need to borrow 944-258 15802 out of stock
No | Visit | Complaint | RCHN | 卡试剂包的传感器损坏 | 卡试剂包的传感器损坏,借903-627一个。 | 收到备件后故障依旧，需要加借903-629和903-628各一个 by Andrew | SWO-563367 | [5 empty] | Damage to the sensor of the card kit | The sensor of the card reagent pack is damaged, borrowing 903-627 one.
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 测试卡漏液 | 现场同客户沟通解释。更换测试卡 | SWO-563438 | [5 empty] | Test card, reagent package failure | Test card leakage
No | Visit | Complaint | RCHN | O2电极密封圈老化 | O2电极密封圈老化 | 更换密封圈后修复 by Andrew | SWO-563993 | Defect | Electrode | Electrode | Upper sealing ring, pO2 |  | O2 electrode seal ageing | O2 electrode seal ageing
No | Visit | Complaint | RIND | Windows Corrupt | Windows corrupt. | Install new window ,restart the analyzer. Install software 6.18 and restart the analyzer. Run two point calibration. | SWO-564765 | Defect | SW | Windows | Windows | Old CPU | Windows Corrupt | Windows corrupt.
No | 5 | Y | RCZE | Blank |  |  | D642E913DC53CE4C | Defect | HW | Inlet module | Probe bend |  | Blank
No | Visit | Complaint | RAME | Oxi hardware errors | 1) Error 251: Halogen lamp intensity too high. 2) Error 379: unstable calibration. | 1) Adjusted lamp within range. 2) Remembraned electrode. | SWO-563088 | Adjustment | HW | Oxi module | Lamp unit | 251; halogen lamp intensity too high | Oxy Hardware Errors | 1) Error 251: Halogen lamp intensity too high. 2) Error 379: unstable calibration.
No | Visit | Complaint | RIND | Na+ blood sample reading high | found Na+ blood sample reading high | Cleaned inlet gasket,chamber, wash electrode,replace Na+membrane,done 2 point calibration , run control, control found whiten range. | SWO-564403 | Assisted Replacement | Electrode | Membrane | reference | Na+ | Na+ blood sample reading high | found Na+ blood sample reading high
No | Visit | Complaint | RDEU | 06.06.2019 9:00 Fr. Seitz +49911398114183: Gerät startet nicht mehr 11.06.2019, 08:34 Uhr, Fr. Seitz, Tel. +49911398114183: ABL bootet nicht mehr | ABL bootet nicht mehr Fehlermeldung wie letzte Woche roma.ini Fehler nach Neustart | Tausch des SSD Speichers, versucht Windows7 aufzuspielen - ohne Erfolg, mehrere Versuche gestartet, alte SSD Speicherkarte eingesteckt und versucht diese neu aufzusetzen, ebefalls ohne Erfolg - USB Stick scheint defekt zu sein, anderen USB Stick versucht - ohne Erfolge - Datei boot.mgr würde fehlen, Kontakt mit TS aufegnommen, boot.mgr auf Stick gespielt und mit alter SSD Speicherkarte Win7 versucht zu installieren - funktioniert, alte SSD Karte wieder getauscht gegen neue SSD Speicherkarte, Win7 aufgespielt, Applikations Software aufgespielt, Konfig eingespielt und überprüft, Ausdrucklayout angepasst und mit Kunden besprochen, QC Chargen eingelesen, mehrmals kalibriert, QC in Ordnung, System- und Funktionskontrolle in Ordnung | SWO-563957 | Defect | HW | DMS_PC | SSD | mail: send SSD card to Ejner | 06.06.2019 9:00 Fr. Seitz +49911398114183: Device no longer starts 11.06.2019, 08:34, Fr. Seitz, Tel. +49911398114183: ABL no longer boots | ABL no longer boots Error message like last week roma.ini Error after restart
No | Visit | Complaint | RIND | Leakage from inlet casket is observed on ABL800 | Need to check the inlet Gasket for any clots, also need to check the pH chamber for any clots trapped | Cleaned the inlet casket using sodium hypo-chloride solution, clot is trapped in the pH chamber so cleaned the pH electrode, done protein removal, after 2 point calibration machine is working fine. | SWO-564399 | Cleaned | HW | pH/BG Module | measuring chamber | Clot | Leakage from inlet casket is observed on ABL800 | Need to check the inlet Gasket for any clots, also need to check the pH chamber for any clots trapped
No | Visit | Complaint | RPOL | Błąd kalibracji pompy, błędy przepływów | Błędy przepływowe | Wymiana uszczelki wlotu | SWO-563710 | Defect | HW | Auto Inlet module | Inlet gasket | Pump cal | Pump calibration error, flow errors | Flow errors
No | Visit | Complaint | RPOL | Wygięta sonda aspiracyjna | Stwierdzono wygięcie sondy aspiracyjnej. | Wymiana sondy aspiracyjnej, testowanie aparatu, sprawdzenie logów komunikacyjnych połączenia LIS $HIS, testowanie połączenia. | SWO-563494 | Defect | HW | Auto Inlet module | Probe bend |  | Bent aspiration probe | The aspiration probe was bent.
No | Visit | Complaint | RPOL | Dostarczenie nowej SC90 | Błąd kalibracji pH oraz elektrolitów | Dostarczenie oraz wymiana nowej SC90 | SWO-563712 | Assisted Replacement | SC | SC | reference |  | Delivery of the new SC90 | PH and Electrolyte calibration error
No | Visit | Complaint | RDEU | 12.06.2019 13:15 Fr. Drunkemühle +496151278870: pO2 QC-Fehler Lev C 12.6.2019/14:50 Sr. Claudia Tel.: 02811062310 -> weiterhin pO2-Fehler im Level C | PO2 im Level C zu hoch 39mmHg, und im Level B zu niedrig 287 mmHg | Einlassnadel verbogen, Nadel ersetzt, Einlassdichtung ersetzt, LP ersetzt, keien Veränderung, SC wurde nur reiniitialisiert und nicht ersetzt, SC aus Kundenbestand ersetzt, Kal und QC durchgeführt i.O. einen LP als Ersatz vor Ort gelassen | SWO-564781 | Assisted Replacement | SC | SC | pO2 | QC C too high; QC B too low | 12.06.2019 13:15 Fr. Drunkemühle +496151278870: pO2 QC Error Lev C 12.6.2019/14:50 Sr. Claudia Tel.: 02811062310 -> continue pO2 errors in level C | PO2 at level C too high 39mmHg, and in level B too low 287 mmHg
Yes | Visit | Complaint | RDEU | 06.06.2019 14:00Uhr Herr Krüger tel:+491607479479 : Die Gerät werden an ein neues Netzwerk angebunden. Die Konfiguration von allen Geräten soll angepasst werden. | Anbindung aller 4 neuen Geräte im Haus an Kühndata. | IP Einstellungen nach Kundenvorgabe konfiguriert, IP Adresse ist nicht anpingbar - Herr Öz von der EDV benötigt die Einstellung, dass das Gerät gepingt werden kann. Auf Kundenwunsch die Pingfreigabe in der Firewall akitiviert. Gerät kann gepingt werden. LIS $KIS Verbindung konfiguriert und Daten gesendet. Kalibriert und eine QC durchgeführt. | SWO-564298 | Re classified RMED | Request |  |  |  | 06.06.2019 14:00H Mr. Krüger tel:+491607479479 : The device will be connected to a new network. The configuration of all devices should be adjusted. | Connection of all 4 new devices in the house to Kühndata.
No | Visit | Complaint | RAME | Leak Detected | Leak detected. | Tried replacing another SC and SP for troubleshooting purposes. Still same issue. Disassemble analyzer. Pinch valve tubing had a hole in it. Clean and lubricate pinch valve. Put in parts from PM kit. | SWO-564379 | Defect | HW | Oxi module | Pinch valve tube | Leak | Leak Detected | Leak detected.
No | Visit | Complaint | RLTD | Doha on 0208 2164934 | Customer called back still having issue seems to be leaking from under ware the sp goes | Removed wet section and found the pinch valve tube to have a small hole in. Cleaned area to remove dried liquid. Fitted new valve tube and refitted wet section. Run service test programs all passed ok. Run several calibration and Qc all passed ok. | SWO-564816 | Defect | HW | Oxi module | Pinch valve tube | Leak | Doha on 0208 2164934 | Customer called back still having issue seems to be leaking from under ware the sp goes
No | Visit | Complaint | RSCH | Nach Sensor tausch Urea Fehler Kal | Urea Kalibrierung fehlerhaft nach Sensoraustausch | Kal Urea Fehlerhaft. Keine Reparatur möglich da Kunde keine Sensorkassetten mehr hat aber bestellt wurden und keine Lieferung erfolgt. Analysatorstatus nicht i.o | SWO-564222 | Assisted Replacement | SC | SC | Urea | Cal | After sensor exchange Urea error Kal | Urea calibration faulty after sensor replacement
No | Visit | Complaint | RAME | cracked screen case | cracked screen casing looks like it was pulled apart. damage was not pressent on install or durring super user trainings. database status code 2 - need to replace database file roma file and remove log file | ordered screen will be here 6 $18 to replace re set database files analyzer restart will need to re run linearity on 6 $18 with screen replacement | SWO-565708 | Revisit required | HW | Display | Touch Screen |  | cracked screen case | cracked screen casing looks like it was pulled apart. damage was not pressent on install or durring super user trainings. database status code 2 - need to replace database file roma file and remove log file
No | Visit | Complaint | RCHN | HCT,CA定标不过 | HCT,CA定标不过，未更换备件 | 冲洗测试卡后重新安装修复 by Andrew | SWO-564003 | [5 empty] | HCT, CA calibration but | HCT, CA calibration, no replacement of spare parts
No | Visit | Complaint | RCHN | 病人结果与临床不符 | 病人结果与临床不符（未更换备件） | 做质控校准后修复 by Andrew | SWO-563896 | [5 empty] | Patient results are not clinically consistent | Patient results are not clinically consistent (spare parts not replaced)
No | Visit | Complaint | RIND | Proactive Visit. | Need to check the analyzer. | Check the status of SC and SP, check the status of calibration, found OK, analyzer handed over to staff in well working condition. | SWO-564983 | [5 empty] | Proactive Visit. | Need to check the analyzer.
No | Visit | Complaint | RCHN | 接口板腐蚀 | 接口板腐蚀，借902-972一个 | 更换902-972后修复 by Andrew | SWO-562508 | [5 empty] | Interface plate corrosion | Interface plate corrosion, borrow 902-972 one
No | Visit | Complaint | RCHN | 参比液泄漏，导致针脚腐蚀 | 参比液泄漏，导致针脚腐蚀，借891-051*32个 |  | SWO-563998 | Defect | HW | Basic Unit | PCB, Sensor interface | Leak from Reference electrode | Reference fluid leakage, resulting in pin corrosion | Reference fluid leakage, resulting in pin corrosion, borrowed 891-051 x 32
No | Visit | Complaint | RCHN | 不均匀的冲洗溶液 | 不均匀的冲洗溶液（未更换备件） | 自行更换924-816后修复 by Andrew | SWO-564124 | Defect | HW | Inlet module | Inlet gasket | 1324 | Uneven flushing solution | Uneven flushing solution (no replacement parts)
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 测试卡底层温度超范围 | 更换测试卡后修复 | SWO-564690 | [5 empty] | Test card, reagent package failure | Test card underlying temperature over range
No | Visit | Complaint | RCHN | 进样门开关故障 报错5784 | 进样门开关故障 报错5784，需要借件903-717 | 更换903-717备件后修复 by Andrew | SWO-563267 | Defect | HW | Cover and Chassis | Door, Inlet | 5784; Top bracket | The sample door switch fault reported error 5784 | Sample door switch fault reported error 5784, need to borrow 903-717
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | pco2定标失败 | 更换测试卡后修复 | SWO-563958 | [5 empty] | Test card, reagent package failure | pco2 calibration failure
No | Visit | Complaint | RIND | pH pCO2 Lac not calibrating | Check system | Check system, flush SC with D $water, again reinstall same problem occurs try new solution pack but still problem not resolved replace SC and make PER. | SWO-564494 | [5 empty] | pH pCO2 Lac not calibrating | Check system
No | Visit | Complaint | RCHN | 废液口管道连接器断裂 | 废液口管道连接器断裂,借905-847一个。 |  | SWO-564114 | Defect | HW | Inlet module | Nipple Kit |  | Waste port pipe connector breaks | The waste mouth pipe connector breaks, borrowing 905-847 one.
No | Visit | Complaint | RAME | Received Cannot find required Map name now analyzer will not boot-up | ABL will not boot up. Error cannot find required map name during boot up. After correcting the boot up issue I had error 0316 neon number too low. Performed a complete OXI test. Lamp test failed for neon position number too low. | Replaced the CPU and hard drive. Reinstalled software, customers setup and database. Replaced the lamp unit for the neon number too low error. Performed a complete OXI test with no errors. Performed a service tHb calibration. | SWO-564338 | Defect | HW | Oxi module | Lamp unit | Neon number too low | Received Cannot find required Map name now analyzer will not boot-up | ABL will not boot up. Error cannot find required map name during boot up. After correcting the boot up issue I had error 0316 neon number too low. Performed a complete OXI test. Lamp test failed for neon position number too low.
No | Visit | Complaint | RAME | Oximetry errors | 581 and 269 oximetry errors during patient samples | Replaced Hemolyzer S $N R4936N007 and ran service THB cal. Complete oximetry test passed | SWO-564601 | Defect | HW | Oxi module | Hemolyzer | 581; 269 | Oximetry errors | 581 and 269 oximetry errors during patient samples
No | Visit | Complaint | RAME | Failing the high Crea Check. He has spoken to his FSE who indicated that he would need to order in parts. The FSE asked him to call in for a job ticket. | High Creatine control is high |  | SWO-564607 | Revisit required | Electrode | Electrode | Crea | S8377 too high | Failing the high Crea Check. He has spoken to his FSE who indicated that he would need to order in parts. The FSE asked him to call in for a job ticket. | High Creatine control is high
No | Visit | Complaint | RAME | Inlet nipple connector is damaged | Checked the data logs. | Replaced the inlet nipples. | SWO-563827 | Defect | HW | Inlet module | Nipple Kit |  | Inlet nipple connector is damaged | Checked the data logs.
No | Visit | Complaint | RESP | Han ido a cargar un quesito y no lo lee | Chequeo lector códigos de barra quesito | Limpieza lector, limpieza externa e interna de equipo observando gran cantidad de polvo, carga de reactivo, calibración de nuevo lote de NTproBNP, pasados Qc. | SWO-564577 | Cleaned | HW | Storage Module | Barcode reader, storage wheel |  | They've gone to carry a body and he doesn't read it. | Check reader cheesy barcodes
No | Visit | Complaint | RESP | Han ido a cargar un quesito y no lo lee | Chequeo lector códigos de barra quesito | Limpieza lector, limpieza externa e interna de equipo observando gran cantidad de polvo, carga de reactivo, calibración de nuevo lote de NTproBNP, pasados Qc. | SWO-564577 | Re classified RMED | Duplicate |  |  |  | They've gone to carry a body and he doesn't read it. | Check reader cheesy barcodes
No | Visit | Complaint | RDEU | Falsch Positiv gemessener Tni | Nach Anleitung von RDEU den AQT überprüft, Service Dump vor und nach der Überprüfung erstellt, zur Analyse nach Radiometer Kopenhagen über Radiometer Service zugesendet. | Jahreswartung durchgeführt, Schläuche, Filter und Dichtungen ersetzt. Alle Scanner gereinigt. Geräten innen und außen desinfiziert $ Gereinigt. Allgemeine Sicht und Funktionskontrolle. Fünf „Blindmessungen gemessen und verglichen>> I.O. Drei LQC’s gemessen. | SWO-564182 | Cleaned | HW | Wet section | Wet section | Tni F+; yearly maintenance | False positively measured Tni | According to the guidance of RDEU the AQT reviewed, service dump created before and after the review, sent to Radiometer Copenhagen for analysis via Radiometer Service.
No | Visit | Complaint | RDEU | 14.06.2019 Herr Telbert +496976014231 meldet, dass der ABL immer wieder ein neues LP verlangt | Überprüfung im Servicemenü | LP Check im Bereich Spülung durchgefüallen, Drucktest zu hoch. Quetschventil hängt zeitweise, Ventil gangbar gemacht. Pumpenkopf kontrolliert, i.O. Spülung durchgeführt. Servicetest erfolgreich bestanden Initialisierungsphase Kal und QC ohne Beanstandung | SWO-565059 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | 14.06.2019 Mr Telbert +496976014231 reports that the ABL is constantly demanding a new LP | Check in the service menu
No | Visit | Complaint | RDEU | ein Fernzugriff (Schnell-) aus Aqure auf den Analysator ist nicht möglich | Win7 Fehlermeldung scheint die Ursache zu sein | Gerät komplett neu aufgesetzt, alte Daten aufgespielt, Zugriff i.O. | SWO-565072 | Defect | SW | RMED SW | Windows | Aqure quick access | remote access (quick) from Aqure to the analyzer is not possible | Win7 error message appears to be the cause
No | Visit | Complaint | RLTD | Claire Reynolds on 07546 728582 17.06.19 David Ryder 07580 818900 - also failing po2 on qc | SP not locking in correctly. Pump calibration errors found to be faulty SP | stripped and cleaned actuator motor. SP now locking in correctly. | SWO-569800 | Adjustment | HW | Basic Unit | Stepper Motor |  | Claire Reynolds on 07546 728582 17.06.19 David Ryder 07580 818900 - also failing po2 on qc | SP not locking in correctly. Pump calibration errors found to be faulty SP
No | Visit | Complaint | RAME | Turbidity too high errors on analyzer | Arrived onsite and replaced the pinch valve and tubing. Ran complete oxi test in service and passed. Ran a calibration and passed. Ran three levels of QC all passed until level 3 lac. Reran QC and lac passed. Reran oxi text and failed. Problem still exists. Will speak with specialist in the morning and order parts for Monday. Informed Mario on status of analyzer. | Ordered parts for Monday. Analyzer not operational. | SWO-564606 | Revisit required | HW | Oxi module | Hemolyzer | Turbidity too high | Turbidity too high errors on analyzer | Arrived onsite and replaced the pinch valve and tubing. Ran complete oxi test in service and passed. Ran a calibration and passed. Ran three levels of QC all passed until level 3 lac. Reran QC and lac passed. Reran oxi text and failed. Problem still exists. Will speak with specialist in the morning and order parts for Monday. Informed Mario on status of analyzer.
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | CO2，CL定标不过0378 | 更换测试卡后修复 | SWO-569427 | [5 empty] | Test card, reagent package failure | CO2, CL is not only 0378
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | CO2,CL定标不过0378 | 更换测试卡后修复 | SWO-570396 | [5 empty] | Test card, reagent package failure | CO2, CL is not only 0378
No | Visit | Complaint | RCHN | 血氧光谱不匹配 | 血氧光谱不匹配，借902-980,902-979各一个 |  | SWO-564717 | Revisit required | HW | Basic Unit | PCB, wet section | 581 | Blood oxygen spectroscopy mismatch | Blood oxygen spectroscopy mismatch, borrow 902-980, 902-979 each
No | Visit | Complaint | RCHN | 定标失败 | 定标失败，借 944-268（15897）一张 |  | SWO-564907 | Assisted Replacement | Consumables | Test Kit | Cartridge, cal | 3412 | Bid failed | Failed bid, borrow 944-268 (15897)
Yes | Visit | Complaint | RCHN | 处理盘卡盘，报错5784 | 处理盘卡盘，报错5784,需要借件903-031和636-452一个 |  | SWO-564719 | Re classified RMED | No Visit |  |  | see e-mail from ROCAO with title: 答复: SWO-564719 - why has this SWO been removed? | Processing disk, 5784 error | Processing disk, 5784, need to borrow 903-031 and 636-452 one
No | Visit | Complaint | RCHN | 进样门关不上，报错5784 | 进样门关闭不上 5784（未更换备件） | 调整位置后修复 by Andrew | SWO-564758 | Adjustment | HW | Cover and Chassis | Door, Inlet | 5784 | The entrance door is not closed, the wrong 5784 | The sample door is closed not to 5784 (spare parts not replaced)
No | Visit | Complaint | RCHN | TNT定标失败 | TNT定标失败，借944-268（15781）*2 |  | SWO-564786 | Cleaned | HW | Process Module | Process Wheel | 1024; CRM Corrective Action: Clean internal dust and recalibrate after repair by jesson | TNT calibration failure | TNT calibration failure, borrowed 944-268 (15781) x2
No | Visit | Complaint | RCHN | 频繁自动启动 | 频繁自动启动，借910-343,910-373各一个 | 更换910-343,910-373后修复(by chuanwei) | SWO-563141 | Defect | HW | Analyzer Control Module | CPU |  | Frequent auto-start | Frequent auto-start, borrow 910-343, 910-373 each
No | Visit | Complaint | RCHN | 定标失败 1024（在常州市妇幼保健院装机发生故障） | 定标失败，借944-214（15958）一张 |  | SWO-564825 | Defect | HW | Process Module | Process Wheel | 1024 | Bid failure 1024 (failure of installation at Changzhou Maternal and Child Health Hospital) | Failed to bid, borrow 944-214 (15958)
No | Visit | Complaint | RDAN | Lene har problemer med at apparatet ikke kan detektere, at inlet er lukket. De har slukket og tændt apparatet og derefter kommer en fejl på tipenhed. De har også oplevet problemer med skærmen, som ikke har responderet som den skulle. | Sluk $tænd | Væskepakke skiftet u.b. | SWO-567533 | [5 empty] | Lene has problems with the appliance not being able to detect that inlet is closed. They have switched the appliance off and on and then there is an error on the Tipene. They have also experienced problems with the screen, which have not responded as it should. | Turn off /t
No | Visit | Complaint | RCHN | 卡试剂包的传感器损坏 | 卡试剂包的传感器损坏,需要借件 903-629和903-628各一个 |  | SWO-564929 | [5 empty] | Damage to the sensor of the card kit | The sensor of the card kit is damaged and requires one each of the 903-629 and 903-628
No | Visit | Complaint | RIND | Service Call. | Clotted sample has aspirated by the user which was tangled in the waste line, need to remove it. | Hold the analyzer, flush the inlet gaskit, inlet probe and its tubing, check the waste line and flush it with hypo, successfully done, recorded no replacement and restart the analyzer, successfully done, analyzer handed over to staff in well working condition. | SWO-564992 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway | Clot | Service Call. | Clotted sample has aspirated by the user which was tangled in the waste line, need to remove it.
No | Visit | Complaint | RCHN | k测量的结果不准确 | k测量的结果不准确(未更换备件） | 更换k和lac还有参比电极膜后修复 | SWO-564919 | Assisted Replacement | Electrode | Membrane | K |  | k Inaccurate measurement results | k Inaccurate measurement results (spare parts not replaced)
No | Visit | Complaint | RDEU | Häufige Driften pH und Elektrolyte. | Häufige Driften Elektrolyte und pH. | Referenzelektrode erneuert. | SWO-565104 | Assisted Replacement | Electrode | Electrode | reference | CAP-733; failing reference electrode: from before CM | Frequent drifts pH and electrolytes. | Frequent drifting electrolytes and pH.
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | lac定标不过 | 更换测试卡后修复 | SWO-570343 | [5 empty] | Test card, reagent package failure | lac calibration is not
No | Visit | Complaint | RCHN | PH,PCO2，K,NA,CA,CL定标失败 | PH,PCO2，K,NA,CA,CL定标失败（未更换备件） | 重新安装测试卡，零点校准 by jesson | SWO-565468 | [5 empty] | PH,PCO2，K,NA,CA,CL定标失败 | PH, PCO2, K, NA, CA, CL calibration failure (spare parts not replaced)
No | Visit | Complaint | RCHN | 进样口支架断裂 | 进样口支架断裂（未更换备件） | 更换进样口支架后修复 by jesson | SWO-565466 | Defect | HW | Inlet module | Inlet gasket |  | The inlet bracket is broken | The inlet bracket is broken (spare parts have not been replaced)
No | Visit | Complaint | RDEU | 17.06.2019 12:40 H. Meier +492191138380: Durchgangsdichtung muß gewechselt werden; Einlaßsonde ist verbogen | Gerät in Messbereit, Einlassanschlussdichtung wird angefordert, Nadel etwas krumm, | mit Anwender zusammen die Dichtung und die Nadel aus Kundenbestand ersetzt. | SWO-565981 | Defect | HW | Inlet module | Probe bend |  | 17.06.2019 12:40 H. Meier +492191138380: Through seal must be changed; Inlet probe is bent | Device in ready to measure, inlet connection seal is requested, needle slightly crooked,
No | Visit | Complaint | RESP | Vários erros com a instalação do cartucho e posteriormente erro de sw | Status erro 2 | Instalação do sw , configuração e verificação do funcionamento. | SWO-565383 | Defect | SW | RMED SW | RMED SW | Status code 2 | Multiple errors with cartridge installation and subsequently SW error | Status Error 2
No | Visit | Complaint | RESP | Entrada partida ( gasket) | Verificação de gasket partido e sonda torta | Substituição do material indicado e passagem do QC | SWO-565304 | Defect | HW | Inlet module | Inlet gasket | 1367 | Starter input (GASKET) | Check of party gasket and PIE probe
No | Visit | Complaint | RLTD | Katie Surtees 0191-445 3195 or 07875 335682 |  |  | SWO-566203 | Defect | HW | Oxi module | Pinch valve tube | CRM Corrective Action: Found OXI pinch valve tubing to be leaking. Replaced tubing. Analyser started up ok and passed all fluid transport tests. | Katie Surtees 0191-445 3195 or 07875 335682
No | Visit | Complaint | RDEU | 17.06.2019; Hebamme Conny; Tel.: +493991772516; Leckage erkannt | Einlaßhalter gewechselt | Service Programme durchlaufen lassen - alles i.O. SC erneuert, da Fehler seit Samstag besteht Startup läuft noch | SWO-565530 | Defect | HW | Auto Inlet module | Inlet gasket | Leak | 17.06.2019; Midwife Conny; Tel. +493991772516; Leakage detected | Inlet holder changed
No | Visit | Complaint | RKOR | Out of range calibration(status, sens) | Out of range calibration(stauts, sens) | Re-perform calibration, QC | SWO-568572 | No fault found | No fault found | No fault found | No fault found | Cal status / sens | Out of range calibration(status, sens) | Out of range calibration(stauts, sens)
No | Visit | Complaint | RKOR | Value below range(Glue) | Value below range(Glue) | Analyzer status check Calibration, QC perform Data logs check and backup Service dump backup | SWO-568569 | No fault found | No fault found | No fault found | Glu | Cal | Value below range(Glue) | Value below range(Glue)
No | Visit | Complaint | RAME | Analyzer failed a system cycle due to error 2106 | Troubleshot and found that there was no fluid flow. I did find evidence of clots in the waste tubing. | Replaced the hemolyzer and all the tubing. This resolved the problem. | SWO-565564 | [5 empty] | Analyzer failed a system cycle due to error 2106 | Troubleshot and found that there was no fluid flow. I did find evidence of clots in the waste tubing.
No | Visit | Complaint | RAME | pO2 external QC out of range high on multiple consumables. | Tested channels with cassette. Replaced pump and waste manifold. Ran level 1 QC. PO2 still really high. Reviewed logs and lot changes. SP changed before problem started. | Removed SP lot 276816 and replaced with SP lot 277155. pO2 is still really high. Replaced the sensor interconnect board. | SWO-565641 | [5 empty] | pO2 external QC out of range high on multiple consumables. | Tested channels with cassette. Replaced pump and waste manifold. Ran level 1 QC. PO2 still really high. Reviewed logs and lot changes. SP changed before problem started.
No | Visit | Complaint | RIND | Pco2 and Hct values problem | Customer complaining pco2 and hct values are lower side | Checked the analyser reinstalled the cassette and suggested the customer to mix the sample nicely then tried all values came in the range working fine | SWO-568379 | [5 empty] | Pco2 and Hct values problem | Customer complaining pco2 and hct values are lower side
No | Visit | Complaint | RCHN | 检测到泄漏 1184 | 检测到泄漏 1184 更换夹紧阀管路正常。（未更换备件） | 更换夹紧阀管路正常。（未更换备件） | SWO-565738 | Adjustment | HW | Oxi module | Pinch valve tube | 1184 | Leak detected 1184 | Leak 1184 Replacement clamping valve lines are detected as correct. (Spare parts not replaced)
No | Visit | Complaint | RCHN | 进样口支架断裂，试管条码器故障 | 进样口支架断裂，试管条码器故障 | 自行更换903-584、910-339后修复 by jesson | SWO-565294 | Defect | HW | Inlet module | Tube holder |  | The inlet bracket is broken, the test tube barcode renter is faulty | The inlet bracket is broken, the test tube barcode renter is faulty
No | Visit | Complaint | RCHN | 存储盘卡盘 | 存储盘卡盘，测试卡卡住，重新取出后安装 | 测试卡卡住，重新取出后安装 by jesson | SWO-565385 | Defect | Consumables | Test Kit | Cartridge, Broken |  | Memory disk card | Memory disk card, test card hold, re-remove after installation
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | PO2没有端点1035 | 更换测试卡后修复 | SWO-570899 | [5 empty] | Test card, reagent package failure | PO2 has no endpoint 1035
No | Visit | Complaint | RCHN | 科室老师怀疑离子结果不准 | 科室老师怀疑离子结果不准（未更换备件） | 给仪器做质控，给客户讲解分析前误差 by jesson | SWO-565889 | [5 empty] | The teacher in the department suspected that the ion results were not allowed. | The teacher suspected that the ion results were not allowed (no spare parts were replaced)
Yes | Visit | Complaint | RDEU | 13.06.2019, email: Demo Options keys laufen bald ab | Demo Option Key läuft aus | Final Key eingespielt; Startphase abgewartet; Kal. + QC i.O. | SWO-565039 | Re classified RMED | Request |  |  |  | 13.06.2019, email: Demo Options keys will expire soon | Demo Option Key expires
No | Visit | Complaint | RDEU | Bei Wechseln der Einlassnadel ist Verriegelung gebrochen | Verriegelung Einlassnadel gebrochen. Kundenverschulden. Telefonische Reparaturfreigabe von Hr. Florian Ziaja (Medizintechnik Synlab) | Probe Lock erneuert. Klibrationen und QC´s vorgenommen. | SWO-565190 | User Error | HW | Inlet module | Probe lock |  | When changing the inlet needle, the locking is broken | Locking inlet needle broken. Customer debt. Telephone repair clearance by Mr. Florian Ziaja (Medical Technology Synlab)
No | Visit | Complaint | RESP | El rodillo de la impresora se rompe imposibilitando así la impresión de resultados | La impresora no funciona debido a que falta una pieza del rodillo de impresión. | Se sustituye la impresora y se comprueba el correcto funcionamiento de la misma | SWO-569683 | Defect | HW | Printer Module | Printer Module |  | The printer roller breaks, making it impossible to print results | The printer does not work due to missing a part of the print roller.
No | Visit | Complaint | RDEU | 17.06.2019, 09:54 Uhr, Fr. Seiler/Fr. Schahle, Tel. 08341424735: Schwarzer Bildschirm | Fehlersuche vor Ort | CPU Unit getauscht. Leider Keine Besserung. CF Karte defekt. Frau Gebele über die Situation informiert. CF Karte bestellt. Wie ursprünglich schon geplant wird am Freitag ein Techniker kommen und die CF Karte einbauen. ABL nicht messbereit. CPU Unit für Freitag für Techniker vor Ort an Gerät gelassen. | SWO-566225 | Revisit required | HW | DMS_PC | CF card |  | 17.06.2019, 09:54, Fr. Seiler/Fr. Schahle, tel. 08341424735: Black screen | Troubleshooting on site
No | Visit | Complaint | RDEU | 17.06 10:15 Frau Huber TouchScreen ohne Funktion | Ersatzteil zu AchBe nach Marbach a. N. | Ersatzteillieferung wegen Geräteausfall an AchBe | SWO-567882 | Revisit required | HW | Display | Display | Screws | 17.06 10:15 Ms Huber TouchScreen without function | Spare part to AchBe to Marbach a. N.
No | Visit | Complaint | RDEU | 17.06 10:15 Frau Huber TouchScreen ohne Funktion | Touchscreen ohne Funktion. Schraubengewinde Monitorrückseite ausgebrochen. Monitor sitzt schief. | Touchscreen aus Ersatzgerät eingesetzt. Monitor sitzt trotz Neujustierung schief. Ersatzteile bestellt. | SWO-566165 | Defect | HW | Display | Display | Screws | 17.06 10:15 Ms Huber TouchScreen without function | Touch screen without function. Screw thread monitor back broken out. Monitor sits crooked.
No | Visit | Complaint | RESP | El equipo se queda bloqueado con error de código 2 | Debido a una mala conexión del cable de la red eléctrica pudo provocar el problema. No se ha podido recuperar los datos almacenados. | Poner en la aplicación ficheros roma vacíos y e iniciar el equipo. | SWO-566449 | Adjustment | HW | DMS_PC | Cable | Status code 2 | Your computer gets stuck with code error 2 | Due to a poor connection of the mains cable could cause the problem. The stored data could not be retrieved.
No | Visit | Complaint | RLTD | Laura Anderson Tel: 0141 951 5917 | Analyser battery working but showing 0% on icon | Shutdown analyser, reset jumper on battery module, no showing proper battery % | SWO-566139 | Adjustment | HW | Battery module | Battery module | 0% | Laura Anderson Tel: 0141 951 5917 | Analyser battery working but showing 0% on icon
No | Visit | Complaint | RAME | Turbidity too high errors on analyzer | Arrived onsite with new hemolyzer. Replaced the hemolyzer. Ran complete oxi test. Passed. Ran calibration. Went back into service and ran complete oxi test. All passed. Ran calibration and tHb calibration. All results within radiometer specifications. | Replaced the hemolyzer. (Old SN R2251N007) | SWO-566403 | Defect | HW | Oxi module | Hemolyzer | Turbidity too high | Turbidity too high errors on analyzer | Arrived onsite with new hemolyzer. Replaced the hemolyzer. Ran complete oxi test. Passed. Ran calibration. Went back into service and ran complete oxi test. All passed. Ran calibration and tHb calibration. All results within radiometer specifications.
No | Visit | Complaint | RAME | pO2 issues | Confirmed issue when onsite. PO2 was having E1002, dark data out of range. tHb cal last performed on 6 $10 $2019 was 100.4%. | Replaced wet section and sensor interface pcb per manual to resolve E1002. After first cal, error was resolved. | SWO-566357 | Defect | HW | Basic Unit | PCB, wet section | 1002 | pO2 issues | Confirmed issue when onsite. PO2 was having E1002, dark data out of range. tHb cal last performed on 6 /10 /2019 was 100.4%.
No | Visit | Complaint | RAME | cracked screen case | replacement of screen due to cracked shell | replaced screen | SWO-567555 | Defect | HW | Display | Touch Screen |  | cracked screen case | replacement of screen due to cracked shell
No | Visit | Complaint | RAME | unable to pump solutions cycle of change solution pack | unable to pump solutions change solution pack screen and flush screen replaced solution pack flushed sensor and lines no resolve gets to 5 seconds and goes back to flush screen checked the inlet gasket connector gasket and fluid flow no issues with flow however inlet gasket looks like it is going to crack and connector gasket looks off shape and pushed in to far replaced inlet gasket and connector gasket | replaced solution pack connector gasket inlet gasket | SWO-567580 | Defect | HW | Basic Unit | PM kit | 1343 | unable to pump solutions cycle of change solution pack | unable to pump solutions change solution pack screen and flush screen replaced solution pack flushed sensor and lines no resolve gets to 5 seconds and goes back to flush screen checked the inlet gasket connector gasket and fluid flow no issues with flow however inlet gasket looks like it is going to crack and connector gasket looks off shape and pushed in to far replaced inlet gasket and connector gasket
No | Visit | Complaint | RCHN | po2结果偏高，打印机故障 | po2结果偏高，打印机故障，借903-001、902-972、903-334各一个 | 校准零点，更换903-001、902-972后校准零点，修复 by jesson | SWO-565513 | [5 empty] | Po2 results are high, printer failure | Po2 results are high, printer failure, borrow 903-001, 902-972, 903-334 each one
No | Visit | Complaint | RDEU | Qualitätskontrollen fehlerhaft in pCO2 und pH. | Geräteinbetriebnahme beim Kunden. | Einlaufphase. Gerät mit QCs überprüft. Kalibrationen und Qualitätskontrollen in Ordnung. | SWO-566166 | [5 empty] | Quality controls faulty in pCO2 and pH. | Device commissioning at the customer's station.
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 定标失败 | 更换测试卡后修复 | SWO-568713 | Assisted Replacement | SC | SC | SC | Cal | Test card, reagent package failure | Bid failed
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | CO2定标失败 0378 | 更换测试卡后修复 | SWO-567230 | Assisted Replacement | SC | SC | pCO2 | 378 | Test card, reagent package failure | CO2 calibration failure 0378
No | Visit | Complaint | RCHN | 间歇性报错1999 | 间歇性报错1999 | 清洁灰尘和清洗管路后修复 by jesson | SWO-566687 | Cleaned | HW | Wet section | Wet section |  | Intermittent error 1999 | Intermittent error 1999
No | Visit | Complaint | RCHN | TNI定标失败 | TNI定标失败，借944-212（16080）一张 | 重新定标后修复 by jesson | SWO-565539 | Assisted Replacement | Consumables | Test Kit | Cartridge, cal |  | TNI calibration failed | TNI bid failed, borrow 944-212 (16080) one
No | Visit | Complaint | RBNL | Spoelfouten | Sinds enkele dagen veel Spoelfouten Klant heeft reeds reagentia , naald , gasketel ed vervangen zonder verbetering | Ter plaatse alle kleppen en tubings nagekeken: alles OK Sinds de laatste vervanging gisteren van de spoelvloeistof zijn er geen problemen meer geweest! Waarschijnlijk was er iets fout met deze spoelfles | SWO-565951 | Revisit required | HW | Wet section | Tubing |  | Spoelfouten | Since a few days many flushing errors Customer has already replaced reagents, needle, gas boiler ed without improvement
No | Visit | Complaint | RESP | Equipamento apresenta desvios sucessivos, principalmente no Na | Verificação dos registros, descontaminacao, limpeza das válvulas e circuito dos calibradores, limpeza das câmaras de medida, substituição das membranas dos eléctrodos, estabilização e passagem dos QCs | Limpeza do circuito e substituição das membranas | SWO-566144 | Assisted Replacement | Electrode | Membrane | Na | 376 | Equipment presents successive deviations, mainly in the Na | Checking of records, decontamination, cleaning of valves and circuit of calibrators, cleaning of measurement chambers, replacement of membranes of electrodes, stabilization and passage of QCs
No | Visit | Complaint | RKK | pO2パラメーターバーが黄色と徳永技師より連絡あり |  |  | SWO-565988 | Unknow | Unknown | Unknown | pO2 |  | PO2 parameter bar is contacted by Yellow and Tokunaga Engineer
No | Visit | Complaint | RAME | Printout illegible. | Columns of printout missing. Printer unit recently replaced. | Replaced printer unit. | SWO-566333 | Defect | HW | Printer Module | Printer Module |  | Printout illegible. | Columns of printout missing. Printer unit recently replaced.
No | Visit | Complaint | RESP | Erros vários de pedidos de limpeza e calibracoes | Nenhum erro detectado, verificação de cabo de rede desligado . | Colocação do cabo de rede, substituição da gasket, testes e passagem do QC. | SWO-567025 | Defect | HW | Inlet module | Inlet gasket | 1571 | Multiple errors of cleaning requests and calibracoes | No error detected, network cable check turned off.
No | Visit | Complaint | RDEU | 19.06.2019. Herr Junus, Tel: 02842708422: Nadel verbogen und Einlassdichtung ist abgebrochen | Intervention erforderlich Einlasdichtung abgebrochen, Einlassnadel verbogen | Einlassdichtung mit Halter abgebrochen und die Einlassnadel ist stark verbogen ( Anwenderverschulden ), - Dichtung und Nadel ersetzt, ABL in Startphase 55min, Kal durchgeführt PCo2 Empf Fehler bei Empfindlichkeit 32,5%,QC durchgeführt, Werte innerhalb der Grenzen Kal Fehler wird markiert bei PCO2 und dadurch bei CL, Anwender lässt die SC wieder einlaufen und wird die Kal und dann die QC wiederholen, sollte die PCo 2 Empf. nicht steigen wird er die SC ersetzen, bei weiteren Problemen wird er sich wieder melden | SWO-567045 | Defect | HW | Inlet module | Probe bend |  | 19.06.2019. Mr. Junus, Tel: 02842708422: Needle bent and inlet seal is broken off | Intervention required lascivation aborted, inlet needle bent
No | Visit | Complaint | RDEU | Gerät urplötzlich ohne Benutzung in Intervention erforderlich 1541 | Spülung durchgeführt, keine Verstopfung ACHTUNG: Das Gerät hatte bereits sehr viele Störungen und es wurde bereits vor Wochen eine Auslieferung in einem anderen Haus gestoppt, ein echter Fehler wurde nicht gefunden... | keine Neustart und QC i.O. | SWO-567387 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway | 1541 | Device suddenly required without use in intervention 1541 | Rinse performed, no constipation ATTENTION: The device already had a lot of faults and a delivery was stopped weeks ago in another house, a real error was not found...
No | Visit | Complaint | RESP | Kan partido el gasket por la mitad | Han partido el gasket por la mitad al pasar una muestra | Se cambia gasket sonda y se cambian reactivos que se había estropeado por el tiempo que el equipo ha estado en error, se comprueba funcionamiento con test de servicio calibración y qc. | SWO-566850 | Defect | HW | Inlet module | Inlet gasket | 1249 | Khan split the gasket in half | They split the gasket in half by passing a sample
No | Visit | Complaint | RFRA | Capot de l'écran cassé. | Support cadre écran cassé aux niveau des 4 vis de maintient. Entretoise en plastique fendues et cassées à l’intérieur de l’écran sans doute dû à un mauvais montage (serrerage trop fort) Tactile fendu en bas à gauche lors de la réparation temporaire du cadre de l’écran avec du scotch. | Changement de l’écran en garantie. (Défaut de montage) Relance abl90 ok. Validation Radiometer ok. Passage calibration ok. Passage de tous les CQ ok. Tous les paramètres disponibles ok. Rapport à envoyer par mail à oualid.semouma@ch-arles.fr Romain.NURY@ch-arles.fr Btissame.QARBAA@ch-arles.fr martine.daumas@ch-arles.fr Valerie.FERREAU@ch-arles.fr luc.demaimay@ch-arles.fr | SWO-567095 | Defect | HW | Display | Display | Screws | Broken screen cover. | Broken screen frame support at the 4 hold screws. Cracked and broken plastic stowed inside the screen probably due to a bad mount (tightening too hard) Tactile split at the bottom left when temporarily repairing the screen frame with scotch.
No | Visit | Complaint | RDAN | ABL90 viser 0% i batteri. | ABL90 viser 0% i batteri. | Batteri er skiftet. Alt er ok igen. | SWO-566708 | Defect | HW | Battery module | Battery module | 0% | ABL90 displays 0% in battery. | ABL90 displays 0% in battery.
No | Visit | Complaint | RCHN | K结果偏低 | K结果偏低（未更换备件） | 零点校准，质控，重新定标 by jesson | SWO-567438 | [5 empty] | Low K results | Low K results (spare parts not replaced)
No | Visit | Complaint | RCHN | 电解质结果不准 | 电解质结果不准（未更换备件） | 质控，给客户讲解分析前误差， by jesson | SWO-567440 | [5 empty] | Electrolyte results are not allowed | Electrolyte results are not allowed (spare parts not replaced)
No | Visit | Complaint | RCHN | 上层温度超范围 | 上层温度超范围，借903-581一个 | 更换903-581后修复by HUNTER | SWO-565195 | Defect | HW | Basic Unit | PCB, Sensor interface | 1064, 1066 | Upper temperature over range | Upper temperature over range, borrow 903-581 one
No | Visit | Complaint | RCHN | 提示更换试剂包 | 提示更换试剂包（未更换备件） | 重新安装试剂包后修复 by jesson | SWO-567297 | User Error | SP | SP | SP |  | Prompt to replace the reagent kit | Prompt replacement kit (no replacement parts)
No | Visit | Complaint | RCHN | 仪器堵管 | 仪器堵管,冲洗管路。（未更换备件） | 冲洗管路。（未更换备件）by Andrew | SWO-567478 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | Instrument blocking pipe | The instrument plugs the pipe and flushes the pipe. (Spare parts not replaced)
No | Visit | Complaint | RCHN | 科室老师质疑结果不准 | 科室老师质疑结果不准（未更换备件） | 质控、分析前误差BY Hunter | SWO-566926 | User Error | Preanalytical | Preanalytical | Preanalytical |  | The teacher in the section questioned the result was not allowed. | The teacher of the section questioned the result of the indoubt (no replacement of spare parts)
No | Visit | Complaint | RCHN | 孵育盘卡盘 | 孵育盘卡盘（未更换备件） | 清洁孵育盘，重启仪器后修复 by jesson | SWO-566723 | Cleaned | HW | Incubator Module | Incubator Wheel |  | Incubating plate chuck | Incubating plate chuck (spare parts not replaced)
Yes | Visit | Complaint | RCHN | 显示样本量不足 | 显示样本量不足（未更换备件） | 清洁过滤网，HCT单元、液潞by hunter | SWO-566854 | Re classified RMED | Same day UpSWO |  |  |  | Show insufficient sample size | Show insufficient sample size (spare parts not replaced)
No | Visit | Complaint | RCHN | 试剂包漏液 | 试剂包漏液 | 更换试剂包后修复 | SWO-566855 | Defect | Consumables | Solution Pack | Solution Pack, Leak |  | Reagent bag leak | Reagent bag leak
No | Visit | Complaint | RCHN | 存储盘卡盘 | 存储盘卡盘（未更换备件） | 除尘by HUnter | SWO-566915 | Cleaned | HW | Storage Module | Storage wheel | 5784 | Memory disk card | Memory tray chuck (spare parts not replaced)
No | Visit | Complaint | RCHN | 存储盘卡盘 | 存储盘卡盘，借944-214（15653）一个 | 由于存储盘上黑色的测试卡导轨，没有正确的卡在存储盘上。将它正确安装后，重新定标后修复 by jesson | SWO-565736 | Adjustment | HW | Storage Module | Cartridge guide | 5784 | Memory disk card | Memory disk card, borrow 944-214 (15653) one
No | Visit | Complaint | RFRA | Panne | ABL9 bloqué sur paramètres pH et pco2 en rouge. Suite à un CQ niveau 1 hors gamme (vieux et mauvais CQ 4+ périmé et passé) | Problème de configuration ABL9. Dans Analyse $suppression des résultats $résultats de CQ manuel hors gamme désactivé. Cela débloque les paramètres pH et pco2 en vert mais avec avertissement que le dernier CQ niveau 1 en erreur. Modification du CQ niveau 1 et passage d’un CQ niveau 1 ok, message disparaît ok. Tous les paramètres disponibles ok. Vérification absence de plan de CQ $ absence de verrouillage d’analyse ok. Validation Radiometer ok. Calibration ok. Mme Bascou voit avec cliente concernant les conso changées (1pack et 1cassette). Rapport à envoyer par mail à loueratducros@orange.fr | SWO-567093 | [5 empty] | breakdown | ABL9 blocked on pH and pco2 settings in red. Following an off-range level 1 QC (old and bad CQ 4- outdated and past)
No | 5 | Y | RCZE | Blank |  |  | D64E559D0250D28B | Defect | HW | Oxi module | Hemolyzer |  | Blank
No | Visit | Complaint | RDEU | 18.06.2019, Herr Thiel, Tel: 0581836051: Einlassnippel abgebrochen | System überprüft = > Anschlußnippel bei wechseln an der ELD hängen geblieben , Nippel defekt | Anschlußnippel erneuert , Kalibration und QC durchgeführt | SWO-566858 | Defect | HW | Inlet module | Nipple Kit | Potential User Error | 18.06.2019, Mr. Thiel, Tel: 0581836051: Inlet nipple aborted | System checked = > Connection nipple stuck on the ELD when changing, nipple defective
No | Visit | Complaint | RDEU | 19.06.2019 11:45Uhr Sr.Daniela +49343218452: Das Gerät meldet bei der Messung ständig Probenfehler | Flüssigkeitsstrecke und Ventile geprüft --> Ventilmembran ( Air-Ventil pH $BG-Modul gerissen ) | Ventilmembran defekt--> erneuert LS-Justage und Pumpenkal. durchgeführt Kal und QC i.O Testmessungen --> pH instabil --> Messkammer bei pH gereinigt ( pH-Elektrode und Messkammer mit Elektrolytlösung verschmutzt ) | SWO-566983 | Defect | HW | pH/BG Module | Valve, membrane |  | 19.06.2019 11:45am Sr.Daniela +49343218452: The device constantly reports sample errors during the measurement | Liquid section and valves tested --> Valve membrane ( Air valve pH /BG module torn )
No | Visit | Complaint | RBNL | Spoelfouten | In de logging zijn de foutmeldingen 340 inlet liquid sensor not empty within time , 1014 Volume meting is mislukt van pH naar oxisensor en 1015 Volume meting is mislukt van EL $MET naar oxisensor . ABL is wel meetklaar, nadat reeds de spoelvloeistof werd vervangen. In het servicemenu liquidsensor kalibratie uitgevoerd, deze is goed. Hierna pompkalibratie uitgevoerd, deze mislukt op de vloeistofpomp (in logging verschijnen weer dezelfde foutmeldingen. Pompslang lijkt erg uitgerekt, daarom deze vervangen en pomprotor schoongemaakt. Hierna pompkalibratie meerdere keren uitgevoerd, deze zijn alle goed. ABL gekalibreerd en QC gemeten --> OK. ABL meetklaar. | Pompslang van de vloeistofpomp vervangen en pomprotor schoongemaakt. | SWO-566732 | Defect | HW | Wet section | Tubing |  | Spoelfouten | In the logging, the error messages 340 inlet liquid sensor not empty within time, 1014 Volume measurement failed from pH to oxisensor and 1015 Volume measurement failed from EL /MET to Oxisensor. ABL is now measuring ready, after the rinse fluid has already been replaced. In the Service menu liquid sensor calibration performed, this is good. After that pump calibration is performed, it fails on the liquid pump (in logging again the same error messages appear. Pump hose seems very stretched, therefore this replaced and pump rotor cleaned. After that, pump calibration performed several times, these are all good. ABL calibrated and QC measured--> OK. ABL measuring ready.
No | Visit | Complaint | RESP | Actividade anormal da ventoinha do auto check | Desmontagem e limpeza do interior do autocheck que se encontrava com bastante pó, o filtro do equipamento já tinha sido limpo e posteriormente relatado que os períodos de funcionamento da ventoinha do autocheck haviam diminuído. | Não foi detetada a situação relatada durante a visita | SWO-567137 | No fault found | HW | Autocheck module | Fan |  | Abnormal auto Check fan activity | Disassembly and cleaning of the interior of the autocheck that was very dusty, the filter of the equipment had already been cleaned and subsequently reported that the periods of operation of the fan of the AutoCheck had diminished.
No | Visit | Complaint | RSCH | Cardridge im Analysator verklemmt | Test Cardridge lassen sich nicht mehr entfernen. Gerät muss jedesmall neu gestartet werden. Analysator extrem langsam beim umschalten in den Modus Messbereit Kassetten Einlass blockiert. Mögliche Behinderung entfernen 2995 Closestorage_12 | Keine Behinderung festgestellt. Cabinett PCB bereits ausgetauscht. Software zweimal bereis installiert ohne Erfolg Im Serviceprogramm überprüft. Kein Fehler festgestellt. Bleibt sehr lange auf Startphase (Bildschirm hängt siehe Videos) Aktuell Funktionskontrolle i.o | SWO-566838 | Revisit required | HW | Cover and Chassis | Door, Storage | 2995 | Cardridge stuck in the analyzer | Test Cardridge can no longer be removed. Device must be restarted every small. Analyzer extremely slow when switching to ready-to-measure mode Cassettes inlet blocked. Remove possible disability 2995 Closestorage_12
No | Visit | Complaint | RKK | 測定時にカバーオープンの際に片方が開かない為機器フリーズ。再起動⇒フリーズを数回繰り返す。 |  |  | SWO-567420 | Adjustment | HW | Cover and Chassis | Door, Inlet | 5784 | Equipment freezes because one side does not open when the cover is opened at the time of measurement. Restart ⇒ Freeze several times.
No | Visit | Complaint | RPOL | Błąd kalibracji elektrody Cl | Weryfikacja statusu analizatora | Wymiana elektrody Cl wraz z membraną. Kalibracja HGB. Kalibracja, Kontrola, testowanie. | SWO-567680 | Assisted Replacement | Electrode | Electrode | Cl | Cal | Cl Electrode Calibration Error | Verifying analyzer Status
No | Visit | Complaint | RESP | El equipo presenta algún coagulo | El equipo esta en pantalla de intervención técnica por coagulo | Se realizan varios enjuagues, se pasan los test y se comprueba el correcto funcionamiento del equipo | SWO-571125 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway | Clot | The team has some clot | The team is on a surgical intervention screen
No | Visit | Complaint | RAME | pc02 errors qc high lvl 1 2 and 3 | pc02 high on all levels of qc changed pack no resolve changed sensor resolved issue | changed sensor cassette | SWO-567556 | Assisted Replacement | SC | SC | pCO2 | qc abc high | pc02 errors qc high lvl 1 2 and 3 | pc02 high on all levels of qc changed pack no resolve changed sensor resolved issue
No | Visit | Complaint | RAME | unable to pump solutions | unable to pump solutions analyzer asking for a flush checked service fluid flow saw clot detected in inlet area flushed unit had to take unit apart to flush hemolyzer clot was pushed out of hemolyzer fluid flow back to normal turned unit on to a status code 2 error data base error had to get rid of the old database ini file and log file analyzer start up as new analyzer reloaded setup will have to follow up with a linearity | removed clot re loaded data on a new database | SWO-567557 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | unable to pump solutions | unable to pump solutions analyzer asking for a flush checked service fluid flow saw clot detected in inlet area flushed unit had to take unit apart to flush hemolyzer clot was pushed out of hemolyzer fluid flow back to normal turned unit on to a status code 2 error data base error had to get rid of the old database ini file and log file analyzer start up as new analyzer reloaded setup will have to follow up with a linearity
No | Visit | Complaint | RCHN | pO2定标失败 | pO2定标失败，借902-972一个 | 更换902-972后修复 by jesson | SWO-566669 | [5 empty] | pO2 calibration failure | pO2 calibration failure, borrow 902-972 one
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | po2结果不准 | 针对仪器做质控，更换测试卡后修复 | SWO-567319 | [5 empty] | Test card, reagent package failure | Po2 results are not allowed
No | Visit | Complaint | RCHN | 无法备份数据库压缩文件 | 无法备份数据库压缩文件（未更换备件） | 取消设置，对仪器进行质控 by Andrew | SWO-567804 | [5 empty] | Database compression files could not be backed up | Database compression files could not be backed up (spare parts not replaced)
No | Visit | Complaint | RCHN | 管路堵塞 | 管路堵塞（未更换备件） | 清洁管路后 by jesson | SWO-567376 | [5 empty] | Blocked lines | Line blockage (spare parts not replaced)
No | Visit | Complaint | RCHN | PO2定标失败 | PO2定标失败。做零点校准。(未更换备件） | 做零点校准。(未更换备件）by Andrew | SWO-567470 | [5 empty] | PO2 calibration failure | THE PO2 CALIBRATION FAILED. Do zero calibration. (Spare parts not replaced)
No | Visit | Complaint | RCHN | 无法阅读试剂包智能芯片 | 无法阅读试剂包智能芯片（未更换备件） | 重新安装试剂包后修复 by jesson | SWO-567370 | Adjustment | SP | SP | Chip |  | Can't read reagent package smart chip | Unable to read reagent package smart chip (spare parts not replaced)
No | Visit | Complaint | RCHN | 管路堵塞 | 管路堵塞（未更换备件） | 清洗进样垫片，进样针后修复 by jesson | SWO-567402 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | Blocked lines | Line blockage (spare parts not replaced)
No | Visit | Complaint | RCHN | 抽吸样本失败，报错0331 | 抽吸样本失败，报错0331，（未更换备件） | 清洁进样口垫片后修复 by jesson | SWO-567286 | Cleaned | HW | Inlet module | Inlet gasket | 331 | Failed to pump samples, error 0331 | Failed to pump samples, error 0331, (spare parts not replaced)
Yes | Visit | Complaint | RCHN | 定标信息消失 | 定标信息消失，借910-373一个 | 更换910-373后重新安装系统，回复数据，修复 by jesson | SWO-566083 | [5 empty] | Bid information disappears | Bid information disappears, borrow 910-373 one
No | Visit | Complaint | RCHN | 扫码器无光 | 扫码器无光 未更换备件 | 重启仪器后修复 by jesson | SWO-567374 | Defect | SW | RMED SW | RMED SW | Scanner | The scanner has no light | No light for the scanner, no replacement spare parts
No | Visit | Complaint | RCHN | 做不了样本，提示重启分析仪 | 做不了样本，提示重启分析仪（未更换备件） | 重启分析仪 恢复 by jesson | SWO-567454 | Defect | SW | RMED SW | RMED SW | 5784 | Can't do the sample, prompt to restart the analyzer | Cannot do the sample, prompt to restart the analyzer (spare parts not replaced)
No | Visit | Complaint | RCHN | 测试卡齿条断裂 | 测试卡齿条断裂，借886-619一个 |  | SWO-567899 | Revisit required | HW | Storage Module | Cartridge slider |  | Test card strip break | Test card tooth strip break, borrow 886-619 one
No | Visit | Complaint | RCHN | 进样门卡住 | 进样门卡住（未更换备件） | 调整进样门位置 by jesson | SWO-567959 | Adjustment | HW | Cover and Chassis | Door, Inlet | 5784 | The entrance door is stuck. | The sample door is stuck (no spare parts replaced)
No | Visit | Complaint | RCHN | 仓门关不紧 | 仓门关不紧，借903-717一个 | 更换903-717后修复 by jesson | SWO-565941 | Defect | HW | Cover and Chassis | Door, Inlet | Top bracket; 4548 | The door is not tight. | The door is not tight, borrow 903-717 one
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | ph k na定标失败 | 更换测试卡后修复 | SWO-567335 | [5 empty] | Test card, reagent package failure | ph k na bid failed
No | Visit | Complaint | RCHN | 机器开机蓝屏 | 机器开机蓝屏，借910-321一个 | 更换910-321后修复 by jesson | SWO-565940 | Defect | HW | DMS_PC | CPU | blue screen | Machine power-on blue screen | Machine boot blue screen, borrow 910-321 one
No | Visit | Complaint | RDEU | JörPr infomiert, dass Fernzugriff zur Gerätenamenänderung nicht möglich ist, bitte Vor Ort service | Änderung des Gerätenames & prüfung Remote Zugriff | Gerätenamen entsprechend RDEU POC IT info geändert, exteren Remote Support Proxy Network Host gestartet | SWO-567372 | User Error, FSE | SW | RMED SW | RMED SW | Remote Support Setup | JörPr indents that remote access to device name change is not possible, please on site service | Change of device name & check remote access
No | Visit | Complaint | RAME | Customer would like radiometer to look at the analyzer | Customer cleaned waste lines and measuring chamber Instrument started to work but still wants us to look at it | Inspection check, nothing out of spec accept pH sealing oring Replaced pH oring | SWO-569230 | Defect | Electrode | Electrode | Sealing ring, pH |  | Customer would like radiometer to look at the analyzer | Customer cleaned waste lines and measuring chamber Instrument started to work but still wants us to look at it
No | Visit | Complaint | RPAC | Pressure vacuum test error, Unbale to pump calibration solution 2 | Pressure vacuum test error | Inlet probe was a bit bended. Replaced the inlet probe and inlet gasket. Tested, all ok | SWO-570937 | Defect | HW | Inlet module | Probe bend | Pressure test error | Pressure vacuum test error, Unbale to pump calibration solution 2 | Pressure vacuum test error
No | Visit | Complaint | RDEU | 17.06.2019, 09:54 Uhr, Fr. Seiler/Fr. Schahle, Tel. 08341424735: Schwarzer Bildschirm | Schwarzer Bildschirm | CF Card erneuert, Windows 7 Complete Software installiert, ABL90 Flex Software Version 3.4 MR1 installiert, Touch Screen kalibriert, Gerät Inbetriebnahme durchgeführt, BGA Konfiguriert, Data Upgrade failed, Clean Installation durchgeführt, Software reinstalliert, LP und SC erneuert, Startphase, Glu und Lac besteht Kalibrationen und QC's nicht, Mehrere Kalibrationen Und QC's vorgenommen. | SWO-568533 | Defect | HW | DMS_PC | CF card |  | 17.06.2019, 09:54, Fr. Seiler/Fr. Schahle, tel. 08341424735: Black screen | Black screen
No | Visit | Complaint | RDEU | 19.06.2019 18:00Uhr Sebastian Hodisch tel:+4915201575272 : Der Touchscreen reagiert stark versetzt | Flüssigkeit im Display | Anwenderfehler; Gerät aus; touch screen gew.; Gerät ein; Kal. + QC i.O. | SWO-568176 | Defect | HW | Display | Touch Screen | Liquid in display | 19.06.2019 18:00hrSebastian Hodisch tel:+4915201575272 : The touchscreen reacts strongly offset | Liquid in the display
No | Visit | Complaint | RAME | 581 and 688 error messages on patient samples. | Checked the data logs and verified the problem. | Replaced the hemolyzer and the pinch valve. | SWO-568225 | Defect | HW | Oxi module | Hemolyzer | 581, 688 | 581 and 688 error messages on patient samples. | Checked the data logs and verified the problem.
No | Visit | Complaint | RPAC | Marisa. Following Installation, there was an problem with samples and fluid transport. | Test in Service | Minor lubrication was need to enable smooth operation of the Valve. Perform a Full Fluidic check and a valve test. Calibrations and QC's all passed. | SWO-571681 | Adjustment | HW | Oxi module | Pinch valve |  | Marisa. Following Installation, there was an problem with samples and fluid transport. | Test in Service
No | Visit | Complaint | RFRA | Défaut oximetrie Test complet oxi en défaut sur spectrometer noise | Vérification complète de l’analyseur | Essais de remplacement de: hemoliseur, spectromètre, alimentation, cassette de capteur, pack de solution -> probleme toujours présent Remplacement de la wet section pcb $ $ réinstallation de la 3.4MR1 $ $ mise en route $ $ attente initialisation oxi $ $ calibration thb $ $ calibration $ $ passage des contrôles $ $ essais bons 903-644R0017N010 wet section pcb à sortir | SWO-568120 | Defect | HW | Basic Unit | PCB, wet section |  | Defect oximetrie Complete test oxi in defect on spectrometer noise | Full verification of the analyzer
No | Visit | Complaint | RDEU | 21.06.19 08:35 Monika Halli Tel.021621049826 , Kunde beschwert sich über neuen ABL da dauernd ausfällt bzw. intervention erforderlich ist ; Kunde bekommt Gerät immer mit gewünschten Aktionen wieder zum laufen aber beschwert sich über den Zeitaufwand | Dauernd Intervention, FTC Fehler bei einer LS Zustandsänderung | Einlassdichtung gereinigt, Flow Selektor gereinigt, im Service Fluid flow pumpenkal usw. durchgeführt, Kal und QC bestanden | SWO-568437 | No fault found | No fault found | No fault found | FT |  | 21.06.19 08:35 Monika Halli Tel.021621049826 , customer complains about new ABL as permanent or intervention is required ; Customer always gets device to run again with desired actions but complains about the time required | Continuous intervention, FTC error in LS state change
No | Visit | Complaint | RCHN | 液路堵塞 | 液路堵塞（未更换备件） | 清洁管路 by jesson | SWO-567987 | [5 empty] | Liquid road blockage | Liquid road blockage (no replacement parts)
No | Visit | Complaint | RKK | キャル液流れずキャルエラー 内部のチューブ抜け。 | キャル液流れずキャルエラー | 内部のチューブ抜け。 メンテナンスキット交換 | SWO-570882 | [5 empty] | Cal error without cal liquid flow Tube vents inside. | Cal error without cal liquid flow
No | Visit | Complaint | RSCH | Anbindung CGM | CGM Anbindung mit Herr Kropf | 21.06.2019 08:30 - 09:30 Vor Ort ASTM auf ASTM6xx umgestellt Port und IP Lis Einstellungen vorgenommen Seite CGM Daten kommen an müssen aber richtig interpretiert werden An Labor Situation geschildert Rapport bleibt offen | SWO-568114 | [5 empty] | Connection CGM | CGM connection with Mr. Kropf
No | Visit | Complaint | RCHN | 液路堵塞 | 液路堵塞 未更换备件 | 清洗管路后修复 by jesson | SWO-567873 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway |  | Liquid road blockage | Liquid road blockage Unreplaced spare parts
No | Visit | Complaint | RCHN | 总血红蛋白质控值低 | 总血红蛋白质控值低（未更换备件） | 重新定标总血红蛋白 后正常 by jesson | SWO-567933 | Adjustment | HW | Oxi module | tHb Cal |  | Low total red protein control | Low total red protein control (no replacement parts)
No | Visit | Complaint | RCHN | 压力测试真空错误，检测到泄漏 | 压力测试真空错误，检测到泄漏，借903-129一个（未更换备件） | 夹紧阀接触不良，重新插拔后，机器正常 by jesson | SWO-567357 | Adjustment | HW | Oxi module | Pinch valve | 1063 | Pressure test vacuum error, leak detected | Pressure test vacuum error, leak detected, borrow edgy 903-129 (spare parts not replaced)
No | Visit | Complaint | RCHN | 抽吸样本失败 | 抽吸样本失败（未更换备件） | 清洁进样口垫片后修复 by jesson | SWO-567862 | Cleaned | HW | Inlet module | Inlet gasket | 1271 | Failed to pump samples | Failed to suck samples (spare parts not replaced)
No | Visit | Complaint | RCHN | 冲洗时液体传输错误 | 冲洗时液体传输错误（未更换备件） | 清洁液路，检查进样口支架、采样针，重新安装试剂包 by jesson | SWO-567877 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway | 508 | Liquid transfer error during flushing | Liquid transfer error during flushing (spare parts not replaced)
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 泵定标错误，1178 | 更换试剂包后修复 | SWO-568034 | [5 empty] | Test card, reagent package failure | Pump calibration error, 1178
No | Visit | Complaint | RCHN | 进样针有凝块 | 进样针有凝块（未更换备件） | 冲洗进样针后修复 by Andrew | SWO-567836 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway | Clot | The sample needle has a clot. | The sample needle has a clot (no replacement spare parts)
No | Visit | Complaint | RCHN | 死机，做样本中断 | 死机，做样本中断，借910-343,910-373各一个 |  | SWO-568431 | Defect | HW | Analyzer Control Module | CF card | CRM Corrective Action: Repair tim after replacing 910-373 and reinstalling the system and cleaning the data | Crash, do sample interruption | Crash, do sample interruption, borrow 910-343, 910-373 each one
No | Visit | Complaint | RCHN | 处理盘卡盘，报错5784 | 处理盘卡盘，报错5784，借636-452,903-031各一个 | 更换636-452,903-031后修复 by Andrew | SWO-566549 | Defect | HW | Process Module | PCB, Cup transport rear | 5784 | Processing disk, 5784 error | Processing disk, 5784, borrowing 636-452, 903-031 each
No | Visit | Complaint | RCHN | 老师反映质控有问题 | 老师反映质控有问题（未更换备件） | 拿质控液去做测试后得到客户认可 by jesson | SWO-568404 | User Error | Consumables | LQC | LQC |  | The teacher reported that there was a problem with the quality control | The teacher reported a problem with the quality control (spare parts were not replaced)
No | Visit | Complaint | RCHN | hct值大于62% | hct值大于62% 未更换备件 | 将标本进行离心处理后，检测正常 by jesson | SWO-567934 | User Error | Preanalytical | Preanalytical | Preanalytical | hct 62% | hct value greater than 62% | hct value greater than 62% Unreplaced spare parts
No | Visit | Complaint | RCHN | 仓门无法关闭 | 仓门无法关闭，借903-028一个 | 更换903-028后修复 by jesson | SWO-567302 | Defect | HW | Cover and Chassis | PCB, cabinet top |  | The warehouse door cannot be closed | The warehouse door cannot be closed, borrow 903-028 a
No | Visit | Complaint | RIND | PO2 Parameter become inactive. | Check calibration report to observe error. it is found that P02 parameter fail to calibrate. | First of all go to setting option and enable po2 parameter and give 2 point calibration but PO2 again fail to calibrate. remove sensor cassette hydrate for 30 min. and then again install. all parameter become active. handed over in fine working condition | SWO-568296 | [5 empty] | PO2 Parameter become inactive. | Check calibration report to observe error. it is found that P02 parameter fail to calibrate.
No | Visit | Complaint | RCHN | ph pco2定标失败 | ph pco2定标失败，未更换备件 | 重新手动定标后修复 by jesson | SWO-567937 | [5 empty] | phpco2 calibration failure | Phpco2 calibration failed, spare parts not replaced
Yes | 2 | Y | RCZE | Blank | Instalacia na II detskej klinike DFN BB. |  | D64E4A1902292656 | Re classified RMED | planned SWO |  |  |  | Blank | Instalacia on the II child Clinic DFN BB.
No | Visit | Complaint | RIND | Temperature error in sample result | Find temperature error in result | Shutdown the analyzer , Clean the moisture on pre-heater. Turn on the Analyzer , Run a sample but error remain same. Change the preheater , Run a sample result is satisfactory. | SWO-571165 | Defect | HW | Inlet module | Preheater |  | Temperature error in sample result | Find temperature error in result
No | Visit | Complaint | RSCH | Cardridge im Analysator verklemmt | Immer Fehler 2995 CLOSESTORAGE_12 | Mehrere Massnahmen ohne Erfolg Komplette Front Cabinett leihweise ausgetauscht (Kunden front wird bei RSCH getestet und repariert LQC i.o Funktionskontrolle i.o | SWO-567988 | Defect | HW | Cover and Chassis | Door, Storage | 2995 | Cardridge stuck in the analyzer | Always Error 2995 CLOSESTORAGE_12
No | Visit | Complaint | RAME | Ini.file correction | status code 2 it looks like someone copied the set up from the other machine and on reboot error source was from the Roma | loaded software, QC insert sheet was missing had to loaded different lots than analyzer 189N004 due to insert sheets missing. sensor cassette error code life time of sensor expired came up--replaced sensor cassette --Lac failed for sensor response error and sensitivity not passing -- replaced sensor cassette emailed in the replacement form and copied Karren -- internal QC is passing external is failing Glu on level 2 , calibrated some more and all parameters came in -- data lost due to corrupt file, | SWO-568395 | Defect | SW | RMED SW | RMED SW | Status code 2 | Ini.file correction | status code 2 it looks like someone copied the set up from the other machine and on reboot error source was from the Roma
No | Visit | Complaint | RCHN | CA,NA报错N/E | CA,NA报错N $E，借903-334一个 | 更换903-334后修复 by jesson | SWO-568023 | [5 empty] | CA, NA报错N/E | CA, NA error N /E, borrowing 903-334 one
No | Visit | Complaint | RCHN | 血氧光谱不匹配 | 血氧光谱不匹配，0581 | 更换备用件903-644后修复by tim | SWO-568808 | Defect | HW | Basic Unit | PCB, wet section | 581 | Blood oxygen spectroscopy mismatch | Blood oxygen spectrometry mismatch, 0581
No | Visit | Complaint | RCHN | 测试卡齿条断裂 | 加借886-686一个 |  | SWO-568865 | Defect | HW | Storage Module | Cartridge slider |  | Test card strip break | Borrow 886-686 one
No | Visit | Complaint | RIND | Inlet Gasket leakage | Waste pump tube leakge | Need to replace waste pump tubing, due to leakage solution is coming out from inlet gasket. | SWO-568882 | Revisit required | HW | Basic Unit | Tubing | Waste | Inlet Gasket leakage | Waste pump tube leakge
No | Visit | Complaint | RIND | Analysis Disabled | Analysis Disabled | Found that analyser lockout, sample analysis is currently locked, logon to service found that analyser lockout, remove lockout tick from menu, analyser working fine | SWO-568999 | [5 empty] | Analysis Disabled | Analysis Disabled
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | PH定标失败 | 更换测试卡后修复 | SWO-568504 | Assisted Replacement | SC | SC | pH |  | Test card, reagent package failure | PH calibration failure
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 泵定标失败 1178 | 更换试剂包后修复 | SWO-568492 | Assisted Replacement | SP | SP | Pump Cal | 1178 | Test card, reagent package failure | Pump calibration failure 1178
No | Visit | Complaint | RCHN | 处理盘卡杯 | 处理盘卡杯，借903-702一个 检查时发现卡杯子的弹簧卡不好 | 更换903-702后修复 by jesson | SWO-568048 | Defect | HW | Process Module | Process Wheel | Spring; 5297 | Handling the disc cup | Handling the disc cup, borrow 903-702 a check found that the card cup's spring card is not good
No | Visit | Complaint | RCHN | 孵育盘卡盘 | 孵育盘卡盘（未更换备件） | 情节内不会成，重新运行安全位后修复 by jesson | SWO-568497 | Cleaned | HW | Incubator Module | Incubator Wheel |  | Incubating plate chuck | Incubating plate chuck (spare parts not replaced)
No | Visit | Complaint | RCHN | 进样门无法关闭 | 进样门无法关闭,需要借件903-717 | 更换903-717一个正常。by tim | SWO-568866 | Defect | HW | Cover and Chassis | Door, Inlet | Top bracket | The sample door cannot be closed | The sample door cannot be closed and requires a loan of 903-717
No | Visit | Complaint | RCHN | 仓门关不上 | 仓门关不上，借903-717,886-639,886-638各一个 | 更换903-717,886-639,886-638后修复 by jesson | SWO-567835 | Defect | HW | Cover and Chassis | Door, Inlet | Top bracket | The cangcang door can't be closed. | Cangmen closed, borrow 903-717, 886-639, 886-638 each
No | Visit | Complaint | RDEU | 12.06.2019, schriftlicher Auftrag MT, : Blende am Display gebrochen | Fehlersuche vor Ort | Nasen der Schrauben von Display waren gebrochen. Display Gehäuse getauscht. Kalibrierung durchgeführt. QCs für alle Level durchgeführt. Mit Herrn Häger, Frau Kitzinger und Herrn Rothbauer noch einige organisatorische Dinge und Fragen geklärt da dies der erste Technikerbesuch war. ABL wieder messbereit. | SWO-569728 | Defect | HW | Display | Display | Screws | 12.06.2019, written order MT, : aperture broken on the display | Troubleshooting on site
No | Visit | Complaint | RDEU | Frau Laska 02241-187104 meldet dass LP getauscht wurde, danach musste ABL90 neu gestartet werden. Analysator bleibt im Startbildschirm hängen ´Program can not continue´ | Beim Eintreffen nach Neustart Status code 2 Fehler | Mehrere Neustarts aufgerufen, immer mit o.g. Fehler. V3.4MR2 aufgespielt OHNE Datenbank-Sicherung; relevante CCO Dateien gesichert und nach SW Update zurück gespielt. Nach Neustart kontrolliert, ok. SP und SC abgelaufen => beide ersetzt, Startphase, Kal und QC Messungen aller Level, mehrfach nach kalibriert und QC gemessen, jetzt OK | SWO-569685 | Defect | SW | RMED SW | Database | Status code 2 | Ms. Laska 02241-187104 reports that LP was exchanged, after which ABL90 had to be restarted. Analyzer gets stuck in the home screen 'Program can't continue' | When arriving after restart status code 2 error
No | Visit | Complaint | RPAC | Failing Pressure test, leak, liquid transport Sennie 0864582753 Pump is making loud noise ?not pumping changed solution pack and flushed numerous times Manager did not approve a call out, need visit Monday morning I think this is the right analyser, she said Delivery | Failing Pressure test, leak, liquid transport | Failed the pressure test on service program. Replaced the inlet gasket with holder. Pressure test passed. Ran Calibration and QCs. Tested, all ok | SWO-570943 | Defect | HW | Auto Inlet module | Inlet gasket |  | Failing Pressure test, leak, liquid transport Sennie 0864582753 Pump is making loud noise ?not pumping changed solution pack and flushed numerous times Manager did not approve a call out, need visit Monday morning I think this is the right analyser, she said Delivery | Failing Pressure test, leak, liquid transport
No | Visit | Complaint | RCHN | 离子病人结果与临床不符 | 离子病人结果与临床不符（未更换备件） | 零点校准，质控液水平2检测后正常by tim | SWO-568926 | [5 empty] | Ion patient results are not clinically consistent | Ion patient results are not clinically consistent (spare parts not replaced)
No | Visit | Complaint | RCHN | 液路堵塞 | 液路堵塞未更换备件 | 清洁管路后修复 by tim | SWO-568894 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway | Clot | Liquid road blockage | Unreplaced spare parts with liquid road blockage
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 血氧参数定标失败 | 更换测试卡后修复 | SWO-568703 | Assisted Replacement | SC | SC | pO2 | Cal | Test card, reagent package failure | Failed to calibrate blood oxygen parameters
No | Visit | Complaint | RCHN | 泄漏 1342 | 泄漏 1342,冲洗进样针凝块（未更换备件） | 冲洗进样针凝块by tim | SWO-568906 | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway | Clot | Leak 1342 | Leak 1342, rinse into sample needle clot (spare parts not replaced)
No | Visit | Complaint | RCHN | 测试卡，试剂包故障 | CO2定标失败 0378 | 更换测试卡后修复 | SWO-568825 | Assisted Replacement | SC | SC | pCO2 | Cal | Test card, reagent package failure | CO2 calibration failure 0378
No | Visit | Complaint | RPOL | Przekroczona wartość kontroli jakości dla parametru pO2. | Serwis stwierdza przekroczenie wartości kontroli jakości dla parametru pO2 lub brak wyniku dla pO2 z komunikatem błędu 1035: Brak punktów końcowych. | Aktualizacja oprogramowania do wersji 1.1.3, czyszczenie styków sprężystych gniazda kasety sensorowej, zerowanie kanałów pomiarowych, kalibracja, pomiar QC poziomu 2. | SWO-569056 | [5 empty] | Quality control value exceeded for pO2 parameter. | The service determines that the quality control value for the pO2 parameter is exceeded, or no result for pO2 with error message 1035: no endpoints.
No | Visit | Complaint | RPOL | Błąd kalibracji | Błąd kalibracji. Wynikł z niego również błąd QC. | Wymiana kasety i ponowne qc. | SWO-569816 | [5 empty] | Calibration Error | Calibration error. It also resulted in a QC error.
No | Visit | Complaint | RDEU | 24.06.2019 Frau Thobuu tel:+49401818828817: Das Gerät meldet Leckagefehler 24.06.2019, 10.14 Uhr, Frau Hobuu, Tel: 0401818827721: erneut Leckage Fehler | Analysator geht nach Membranierung immer wieder in den Leckage fehler. | Fehlersuche am Analysator durchgeführt. Fehler bei CA. Weisser Dichtring war nicht unten an der Membrane vorhanden. | SWO-568968 | Defect | Electrode | Membrane | Ca | Mempack missing | 24.06.2019 Mrs. Thobuu tel:+49401818828817: The device reports leakage error 24.06.2019, 10.14 a.m., Mrs. Hobuu, Tel: 0401818827721: again leakage error | Analyzer repeatedly goes into the leakage error after membraneization.
No | Visit | Complaint | RIND | Service Call. | The analyzer is not switching on. | Check the fuse of EMI filter, found OK, check the fuse of 5, 1.5, 0.25, 5 Amp fuse, fuse has lost their continuity, need to replace the fuse, check the CPU and power supply board, by the cross verification, need to replace the power supply board of the analyzer. Presently provide the standby fuse for analyzer control board and power supply board to PC unit, Gas 1 and Gas 2 empty , need to replace both. | SWO-571968 | Defect | HW | DMS_PC | Power Supply |  | Service Call. | The analyzer is not switching on.
No | Visit | Complaint | RPAC | Lisa.....Error Inhomogenous Cal2 and unable to pump solution error 1169 | Check Data Logs in Activity Log | Perform Pinch Valve test. Check Pump Rotor. Perform Fluidic Checks. All passed. Perform Calibration. Error 1079 Impedance error K channel appeared. Repeat calibration. Same error. Replace the Sensor Cassette. | SWO-569475 | Assisted Replacement | SC | SC | K | 1079 | Lisa..... Error Inhomogenous Cal2 and unable to pump solution error 1169 | Check Data Logs in Activity Log
No | Visit | Complaint | RDEU | Display fällt auseinander | Display Montage CPH fehlerhaft, Befestigungsschrauben ausgebrochen. | Display ersetzt-Scanner ummontiert-Touch Screen Kal durchgeführt. Startphase abgewartet. Kal und QC ok. | SWO-569973 | Defect | HW | Display | Display | Screws | Display falls apart | Display mounting CPH faulty, fastening screws broken out.
No | Visit | Complaint | RLTD | Please call David Ryder on 07580 818 900 | Site visit to investigate and repair analyser screen issue. 1 x 910-324 abl90 touch screen @ £109.00 +VAT | Arrived on site and attended analyser. Found obvious signs of liquid ingress damage and confirmed touch screen not functioning. Powered down and replaced touch screen panel, refitted all parts and powered up. Analyser went through 55 min startup. Recalibrated screen and then Ran cals and qcs and these passed. Analyser now back in working order again. No further action required. | SWO-569970 | Defect | HW | Display | Touch Screen | Liquid in display | Please call David Ryder on 07580 818 900 | Site visit to investigate and repair analyser screen issue. 1 x 910-324 abl90 touch screen @ £109.00 +VAT
No | Visit | Complaint | RAME | Leak detected during system cycles and measurements | Leak detected during system cycles and the measurements. Customer tried different Solution packs and sensor Cassette. Also the gaskets changed and didn’t resolve the issue | The Oxi tube was punctured, replaced and fixed the issue. Customer will apply to get credit for 1 SP and 1 SC. | SWO-569996 | Defect | HW | Oxi module | Pinch valve tube | Leak | Leak detected during system cycles and measurements | Leak detected during system cycles and the measurements. Customer tried different Solution packs and sensor Cassette. Also the gaskets changed and didn’t resolve the issue
Yes | Visit | Complaint | RBNL | Na vervangen van solution pack en recente inlet gasket. | Lek gedetecteerd. | Service menu: - controle solution pack ok - controle pinch valve en sensor cassette ok - inlet controle: vervangen van de connector gasket (beschadigt) en de inlet gasket. Probe is ok. - pressure test ok - andere service testen ok - kalibratie en qc (3 niveaus) ok | SWO-569615 | Re classified RMED | Same day UpSWO |  |  |  | After replacing the solution pack and recent inlet gasket. | Leak detected.
No | Visit | Complaint | RBNL | Batterij 0% | Batterij 0% | Batterij test ok Reset van de batterij verandert de waarde naar 100% | SWO-569623 | Adjustment | HW | Battery module | Battery module | 0% | Battery 0% | Battery 0%
No | Visit | Complaint | RLTD | Please call Sarah Mills on 02381 206 721 | Restarted analyser after fully shutting down PC $Windows, a minor error is present where windows asks the user to restart the analyser | Restarted analyser and clicked ignore to restart windows - printer error on start up reseated printer connection. Installed MR1 software to correct minor windows error. | SWO-569809 | Revisit required | HW | DMS_PC | CPU |  | Please call Sarah Mills on 02381 206 721 | Restarted analyser after fully shutting down PC /Windows, a minor error is present where windows asks the user to restart the analyser
Yes | Visit | Complaint | RLTD | 03/06/19 - Now failing on bilirubin on QC's please phone Debbie Donovan on 02381 206721 Jose on 02381 206721 04.06.2019 15:38pm lost oxsymmetry values on patient results Deborah Donovan 02381 206721 | Removed spectrometer and PSU as were not faulty | Removed non-faulty parts. | SWO-569760 | Re classified RMED | planned SWO |  |  | For failures see e.g. SWO-562594 | 03/06/19 - Now failing on bilirubin on QC's please phone Debbie Donovan on 02381 206721 Jose on 02381 206721 04.06.2019 15:38pm lost oxsymmetry values on patient results Deborah Donovan 02381 206721 | Removed spectrometer and PSU as were not faulty
No | Visit | Complaint | RFRA | ERR ETALONNAGE POMPE | CONTROLE DU CIRCUIT FLUIDIQUE $ | Contrôle menu service $Fluid transport , constate erreur étalonnage pompe $Remplacement des tuyaux évacuation et hémolyseur sensor cassettte $Remplacement du joint de connexion et introduction $Remplacement de la pinch valve $ Remise en route et test menu service fluid transport $Erreur pompe persiste toujours $ $Remplacement du pack et de la cassette $ Erreur de pompage des solutions car err pack et err pompe calibration car err cassette $ Test fluid transport ok $ Calibration et passage des CQ. | SWO-569954 | Assisted Replacement | SP | SP | Pump Cal |  | ERR ETALONNAGE POMPE | CONTROLOf THE FLUIDIC CIRCUIT /
No | Visit | Complaint | RCHN | glu定标失败 0210 | glu定标失败 0210,借903-581一个。（在北京市朝阳区医院 急诊检验中心新装机） | 更换903-581后修复by tim | SWO-568885 | Defect | HW | Basic Unit | PCB, Sensor interface | Glu | glu bid failure 0210 | Glu bid failure 0210, borrow 903-581 one. (Newly installed in the Emergency Inspection Center, Chaoyang District Hospital, Beijing)
No | Visit | Complaint | RCHN | 做标本打问号 | 做标本打问号，借903-210一个 | 更换903-210后修复by tim | SWO-568750 | Defect | HW | Process Module | PCB, cup dryer | 1068 | Make a question mark for a specimen. | Do the specimen question mark, borrow 903-210 one
No | Visit | Complaint | RCHN | 测试卡无法扫描 | 测试卡无法扫描 未更换备件 | 清洁测试卡和调整位置后修复by tim | SWO-569564 | Adjustment | HW | Storage Module | Barcode reader, storage wheel |  | Test card can't be scanned | Test card can't scan Unreplaced spare parts
No | Visit | Complaint | RCHN | 进样仓门故障，报错4548 | 进样仓门故障，报错4548（未更换备件） | 重新校准进样门后修复by tim | SWO-570494 | Adjustment | HW | Cover and Chassis | Door, Inlet | 4548 | Sample door failure, error 4548 | Sample door failure, error 4548 (not replaced spare parts)
No | Visit | Complaint | RCHN | 舱门无法正常打开 | 舱门无法正常打开，借903-717一个 | 更换903-717后修复by tim | SWO-568947 | Defect | HW | Cover and Chassis | Door, Inlet | Top bracket | The hatch does not open properly. | The hatch does not open properly, borrow 903-717 one
No | Visit | Complaint | RIND | Service Call. | Need to connect the gas 1 and gas 2. | Connect the Gas 1 and Gas 2 in to the analyzer, successfully done, rinse barcode not recognized by the analyzer, upgrade the software from build 799 to 877, successfully done, reboot the analyzer, after initializing wait for auto 2point calibration, successfully done, analyzer handed over to staff in well working condition. | SWO-571988 | Defect | SW | RMED SW | RMED SW | SW Upgrade; rinse barcode | Service Call. | Need to connect the gas 1 and gas 2.
No | Visit | Complaint | RAME | Analyzer not recognizing proper solution percentage | Confirmed issue when onsite. Alpha coefficient in consumables was set at 0.68. User also wanted to disable Cl- replacement notification as Cl- was not being used. | Reset consumable alpha to default value of 1.0 and changed Cl- replacement schedule to “Never”. Provided training to better use of analyzer. | SWO-570019 | Setup | SW | RMED SW | RMED SW |  | Analyzer not recognizing proper solution percentage | Confirmed issue when onsite. Alpha coefficient in consumables was set at 0.68. User also wanted to disable Cl- replacement notification as Cl- was not being used.
No | Visit | Complaint | RDEU | 24.6.2019/09:45 Frau Kramer (Telefonat mit Joachim Hein RDEU, keinKundentelefon bekannt) -> Reagenzienpack wird nicht erkannt | Fehlersuche vor Ort | Reagenzienpack wird nicht erkannt. Reagent pack PCB getauscht. alte PCB (903-036R2002N0524)ausgebaut und neue PCB (903-036R2002N983) eingebaut. danach wurde Lösungspack wieder erkannt. Systemüberprüfung druchgeführt. i.O. Systemreinigung durchgeführt. i.O. QC Level 2 Multicheck 2 mal durchgeführt. i.O. AQT wieder messbereit. | SWO-569796 | Defect | HW | Cover and Chassis | PCB, solution pack |  | 24.6.2019/09:45 Mrs. Kramer (telephone call with Joachim Hein RDEU, no customer phone known) -> Reagent pack is not recognized | Troubleshooting on site
No | Visit | Complaint | RDEU | 26.06.2019 09:00Uhr Frau Subei tel:+49308811955 :Das Gerät meldet ständig einen Fehler und läuft unzuverlässig. | wiederholt Aufforderung zum Spülen des Systems | komplettes System gespült, Fluid programm im Service durchgeführt, i.O., Kal, QC Datensicherung Kal QC | SWO-570588 | No fault found | No fault found | No fault found | FT |  | 26.06.2019 09:00Am Ms Subei tel:+49308811955 :Das device constantly reports an error and runs unreliably. | repeated request to rinse the system
No | Visit | Complaint | RLTD | Please call Stephanie on 07415 52 49 49 | Cleaned hemolyser with hypochlorite unable to improve QC result, result showing signs of hemolyser failure. Replaced hemolyser with non stock spare hemolyser. | Replaced hemolyser, ran thb cal and full calibration, QC's passing. | SWO-570720 | Defect | HW | Oxi module | Hemolyzer | sO2 | Please call Stephanie on 07415 52 49 49 | Cleaned hemolyser with hypochlorite unable to improve QC result, result showing signs of hemolyser failure. Replaced hemolyser with non stock spare hemolyser.
No | Visit | Complaint | RAME | Battery issue | Battery does not charge | Replaced battery pack. And now it is changing. Ran cals and a QC. All calibrations and 3 Levels of QC passed as per manufacturer specification | SWO-570813 | Defect | HW | Battery module | Battery module |  | Battery issue | Battery does not charge
No | Visit | Complaint | RBNL | Na vervangen RP krijgen we pompkalibratie fouten. | RP pompkalibratie niet ok. clot gedeteerd in inlet. | service menu: *controle van de service testen - ok, waaronder pompkalibratie; *5x pompkalibratie ok *clot removal test ok Opstart toestel ok Kalibratie en QC ok Nog lucht in de leidingen na service menu testen gisteren? Meer liquid sensor calibration moeten uitvoeren. | SWO-570585 | Assisted Replacement | SP | SP | Pump Cal |  | After replacing RP we get pump calibration errors. | RP pompkalibratie niet ok. clot gedeteerd in inlet.
No | Visit | Complaint | RBNL | Beeldscherm losgekomen | beeldscherm losgekomen | Ziet er uit als een voorgaand probleem waarbij de schermen te hard werden aangespannen bij productie. Er wordt een nieuw scherm geinstalleerd onder garantie. | SWO-570597 | Defect | HW | Display | Display | Screws | Display unreleased | Display unreleased
No | Visit | Complaint | RIND | K+ parameter failed in Calibration. | Check the Calibration & found that K+ sensitivity was out of response range. | Flush the sensor cassette with the NS,then kept it on the fridge for 10mins,then again reinstall the sensor cassette but still K+ got failed,need to replace the sensor cassette for further check. | SWO-572095 | [5 empty] | K+ parameter failed in Calibration. | Check the Calibration & found that K+ sensitivity was out of response range.
No | Visit | Complaint | RCHN | 压力测试真空错误，报错1063 | 压力测试真空错误，报错1063 | 更换进样口支架后修复by tim | SWO-570481 | Defect | HW | Inlet module | Inlet gasket | 1063 | Pressure test vacuum error, error 1063 | Pressure test vacuum error, error 1063
Yes | Visit | Complaint | RCHN | 测试卡，试剂包故障 | 无法读取测试卡智能芯片 | 更换测试卡后修复 | SWO-571029 | [5 empty] | Test card, reagent package failure | Failed to read test card smart chip
No | Visit | Complaint | RCHN | GLU结果不准 | GLU结果不准（未更换备件） | 做样本比对，与客户接受后认可by tim | SWO-570476 | No fault found | No fault found | No fault found | Glu |  | GLU results are not allowed. | GLU results are not allowed (spare parts not replaced)
No | Visit | Complaint | RCHN | 处理盘卡盘，报错5784 | 处理盘卡盘，报错5784（未更换备件） | 重新安装测试卡后修复by tim | SWO-570315 | Defect | Consumables | Test Kit | Cartridge, Broken | 5784 | Processing disk, 5784 error | Processing disk, error 5784 (spare parts not replaced)
No | Visit | Complaint | RCHN | 无法进入服务菜单 | 无法进入服务菜单（未更换备件） | 重启仪器by tim | SWO-571509 | Adjustment | SW | RMED SW | RMED SW | Analyzer restarted | Unable to access the service menu | No access to the service menu (spare parts not replaced)
No | Visit | Complaint | RCHN | 进样口支架断裂，试管条码器故障 | 扫描器无法扫描 | 更换备用件636-666后修复by tim | SWO-570257 | Defect | HW | Inlet module | Cable, barcode reader |  | The inlet bracket is broken, the test tube barcode renter is faulty | Scanner cannot be scanned
No | Visit | Complaint | RCHN | 处理单元移动，报错5784 | 处理单元移动，报错5784（未更换备件） | 清洁传感器与马达后修复by tim | SWO-571130 | Cleaned | HW | Process Module | Sensor, process unit | 5784 | Processing unit move, error 5784 | Processing unit movement, error 5784 (spare parts not replaced)
No | Visit | Complaint | RCHN | 舱门破裂 | 仓门破裂，886-638,886-639各一个 | 更换886-638,886-639后修复by tim | SWO-568696 | Defect | HW | Cover and Chassis | Door, Inlet |  | The hatch is broken. | Barn door crack, 886-638, 886-639 each
No | Visit | Complaint | RCHN | BNP结果不准确 | BNP结果不准确，（未更换备件） | 重新做质控后修复by tim | SWO-570306 | Assisted Replacement | Consumables | LQC | LQC | Potential User Error | BNP results are inaccurate | BNP results are inaccurate (spare parts not replaced)
No | Visit | Complaint | RCHN | 进样门无法关闭 | 进样门无法关闭，借903-717一个 | 更换903-717后修复by tim | SWO-569677 | Defect | HW | Cover and Chassis | Door, Inlet | Top bracket | The sample door cannot be closed | The sample door cannot be closed, borrow 903-717 one
No | Visit | Complaint | RIND | Frequent SC Failure | De-installations of old ABL9 | Unplug ABL9, check conditions, Pack with all accessories | SWO-570288 | [5 empty] | Frequent SC Failure | De-installations of old ABL9
No | Visit | Complaint | RKK | SC不良。院内在庫で交換するも気泡検知エラー発生で訪問。 装置内中のチューブが抜けていた。この現象は、在庫のSCを交換してから。 チューブつなぎ直しで改善。SCは院内在庫をのそのまま使用する。 | ケース詳細の通り | SC不良。院内在庫で交換するも気泡検知エラー発生で訪問。 装置内中のチューブが抜けていた。この現象は、在庫のSCを交換してから。 チューブつなぎ直しで改善。SCは院内在庫をのそのまま使用する。 | SWO-570876 | [5 empty] | SC bad. It visits by the bubble detection error occurrence though it exchanges it with the hospital stock. The tube in the device was missing. This behavior is from replacing the SC in stock. Improved by reconnecting the tube. SC uses the in-house stock as it is. | As the case details
No | 22 | Y | RCZE | Blank | Tlaciaren netlaci 1.7. sa ma spustit prevadzka labaku a oni riesia zakladne veci ako spotrebak a LIS pripojenie |  | D6549441503F313C | Setup | SW | RMED SW | RMED SW | mail: complaint? | Blank | Tlaciaren Netlaci 1.7. The Labaku is running out of the way and they diluted the things as a consumer and a LIS connection
No | 3 | Y | RCZE | Blank |  |  | D656E30B3C05D2B6 | Assisted Replacement | Electrode | Electrode | reference | CAPA-733; mail: lot number for failing and newly inserted | Blank
No | Visit | Complaint | RDEU | 21.06.2019 19:45Uhr Frau Stickel tel:+4971127836211 : Der Drucker hat nicht gedruckt, danach wurde der Bildschirm dunkel und das Gerät lässt sich nicht mehr einschalten | Gerät lässt sich nicht einschalten. | CPU aus Ersatzgerät eingebaut. Gerät startet nicht. Powersupply eingebaut, Gerät startet. Software auf Version 6.18MR1 aktualisiert. Nach Neustart druckt Drucker nicht. Nach erneutem Neustart startet BIOS nicht. Erneut CPU aus Ersatzgerät eingebaut und Powersupply zurückgetauscht. Gerät startet, Drucker druckt. Nach Neustart Laktat instabil, Laktatelektrode membraniert. | SWO-570632 | Defect | HW | DMS_PC | CPU |  | 21.06.2019 19:45Pm Ms Stickel tel:+4971127836211 : The printer did not print, after that the screen went dark and the device can no longer be turned on | Device cannot be turned on.
No | Visit | Complaint | RDEU | 24.06.2019, email: Lautsprecher ist defekt | Lautsprecher ohne Funktion | Fehlersuche; Gerät aus; Lautsprecher gew.; Gerät ein; Kal. + QC i.O. | SWO-570615 | Defect | HW | DMS_PC | Loudspeaker |  | 24.06.2019, email: loudspeaker is defective | Loudspeaker without function
No | Visit | Complaint | RESP | Errores en varios parámetros Electrodos enviados por agencia | Errores en PH, lactanto y calcio | Sustitución de electrodos de Ref, láct y Ca | SWO-572376 | Assisted Replacement | Electrode | Electrode | reference | CAPA-733; mail: lot number for failing and newly inserted | Errors in various parameters Electrodes sent by agency | Errors in PH, lactanto and calcium
No | Visit | Complaint | RIND | Inlet Gasket leakage | Checked waste tubing and found that it was broken | Replaced waste tube and cleaned analyzer. Then given 2 point calibration and checked with samples. Unit is working fine. | SWO-571986 | Defect | HW | Basic Unit | Tubing | Waste | Inlet Gasket leakage | Checked waste tubing and found that it was broken
No | Visit | Complaint | RFRA | Er 1068 température hors gamme avant séchage, lors du 1er passage d'une série. | Vérification Contrôle système ok | En mode service, vérification débit air et température chauffage air: ok. Re-vérification Contrôle Système: ok. | SWO-570674 | Defect | HW | Process Module | Cup dryer | 1068 | Er 1068 temperature out of range before drying, during the first passage of a series. | Check System Control ok
Yes | Visit | Complaint | RKOR | Solution pack remaining test | . | Sample count and Cal QC check. | SWO-571141 | Re classified RMED | No fault found |  |  | Mail: why complaint case? | Solution pack remaining test | .
No | Visit | Complaint | RSCH | NT-proBNP bei LQC ausser Bereich | NT-proBNP bei LQC ausserhalb Messbereich Level 2 und CRP sporadisch ausserhalb Messbereich | Neue Lot 16197 kalibriert und LQC erneut durchgeführt(MCC 2 u. 3) Biorad Level 2 konfiguriert und gemessen i.o Cleaning durchgeführt Funktionskontrolle i.o | SWO-570389 | Defect | Consumables | Test Kit | Test Kit |  | NT-proBNP at LQC out of range | NT-proBNP for LQC outside measuring range level 2 and CRP sporadically outside measuring range
No | Visit | Complaint | RPOL | Błąd kalibracji pompy po wymianie pakietu roztworów. | Serwis stwierdza błędy kalibraci pompy oraz detekcji skrzepu na wlocie. | Wymiana wężyków, uszczelki wlotu, zaworka zaciskowego modułu OXI, kasetki sensorowej. Po przetestowaniu 3 pakietów roztworów na innym aparacie stwierdza się wadę 2 pakietów o numerze LOT HL-57 oraz 1 pakietu o numerze LOT HJ-57. Aparat uruchomił się na kolejnym pakiecie roztworów o numerze LOT HJ-57. | SWO-570783 | Assisted Replacement | SP | SP | Pump Cal |  | Pump calibration error after replacing the solution package. | The service finds errors in the pump calibrators and the detection of the clot at the inlet.
No | Visit | Complaint | RPOL | Błąd płukania, pojawia się komunikat brak płynu płuczącego w górnym czujniku phBG . | Aparat nie zaciąga Rinse do modułu pH $BG. Błędy 0509, 0657. Widoczne uszkodzenie membrany zaworu gazowego w module pH $BG. | Wymiana membrany w zaworze gazowym modułu pH $BG. Wymiana wężyka łączącego Inlet z modułem pH $BG. Kalibracje i testy serwisowe. Test QC-OK. | SWO-570594 | Defect | HW | pH/BG Module | Valve, membrane | 509, 657; Valve gas | Rinse error, no flushing fluid appears in the upper phBG sensor. | The camera does not incur Rinse to the pH module /BG. 0509 errors, 0657. Visible damage to the gas valve membrane in the pH module /BG.
No | Visit | Complaint | RPOL | Nieczytelny wydruk raportów pacjenta. | Nieczytelny wydruk raportów pacjenta. | Wymiana głowicy w drukarce. Testy. | SWO-570671 | Defect | HW | Printer Module | Printer Head |  | Unreadable printout of patient reports. | Unreadable printout of patient reports.
No | Visit | Complaint | RLTD | Please call Sarah Mills on 02381 206 721 | CPU has not fixed the issue PSU and CPU at fault potientially. Still restarting occuring reloaded windows and analyser software. Restarted again, have reverted to a black database and will leave over-night to see if anything changes, Both parts are warrenty parts to return to Denmark | Replaced CPU and PSU, reloaded windows software. Installed blank database. | SWO-570027 | Defect | HW | DMS_PC | CPU |  | Please call Sarah Mills on 02381 206 721 | CPU has not fixed the issue PSU and CPU at fault potientially. Still restarting occuring reloaded windows and analyser software. Restarted again, have reverted to a black database and will leave over-night to see if anything changes, Both parts are warrenty parts to return to Denmark
No | Visit | Complaint | RESP | Fallo en TC PO2 Electrode en garantía. | Se le entrega el Sensor nuevo y se instala en el equipo. Se realiza calibración correctamente. | Entregado al cliente el SC por parte del comercial. | SWO-571372 | [5 empty] | Failure in TC PO2 Electrode under warranty. | The new Sensor is delivered to you and installed on your computer. Calibration is performed correctly.
No | Visit | Complaint | RCHN | 数据无法正常传输 | 数据无法正常传输，借910-342一个 | 更换910-342后修复by tim | SWO-569716 | Defect | HW | DMS_PC | CPU |  | Data does not transfer properly | Data cannot be transferred properly, borrowing 910-342 one
No | Visit | Complaint | RCHN | 吸样失败 | 吸样失败(未更换备件) | 更换测试卡后修复 | SWO-571729 | Assisted Replacement | SC | SC | SC |  | Suction failed | Suction failed (spare parts not replaced)
No | Visit | Complaint | RCHN | 检测到泄漏，报错1184 | 检测到泄漏，报错1184（未更换备件） | 更换进样口支架后修复by tim | SWO-571164 | Defect | HW | Inlet module | Inlet gasket | 1184 | Leak detected, error 1184 | Leak detected, error 1184 (spare parts not replaced)
No | Visit | Complaint | RCHN | 液体传输失败 | 液体传输失败（未更换备件） | 冲洗分析仪，更换试剂包后修复by tim | SWO-571314 | Assisted Replacement | SP | SP | FT |  | Failed liquid transfer | Liquid transfer failed (spare parts not replaced)
No | Visit | Complaint | RCHN | 存储盘的门开关不上 | 存储盘的门开关不上（未更换备件） | 调整位置by tim | SWO-571504 | Adjustment | HW | Cover and Chassis | Door, Storage | 2995 | The door switch of the storage disk is not | The door switch of the storage tray is not on (spare parts are not replaced)
No | Visit | Complaint | RCHN | 开机报错5297 | 开机报错5297，借903-031一个 | 更换903-031后修复by tim | SWO-570447 | Defect | HW | Process Module | PCB, Cup transport rear | 5297 | Boot error 5297 | Boot error 5297, borrow 903-031 one
No | Visit | Complaint | RCHN | 显示器故障 | 显示器故障,借636-571，910-298各一个。 | 更换636-571一个正常。by tim | SWO-568892 | Defect | HW | Analyzer Control Module | Cable, Display |  | Display failure | Display failure, borrow 636-571, 910-298 each.
No | Visit | Complaint | RDEU | sporadisch Probenabbruch - Nabelschnurblut! | sporadisch Probenabbruch - Nabelschnurblut! 1 x Fehler 0328 (kein vorderes Luftsegment im Einlass LS Fühler) am 27.06 - 4 Messungen ok 5 Messungen am 26.06. ok am Aufstellungstag 25.06 15 Messungen insgesamt, davon 5 abgebrochen bei Eintreffen QC Level 1 + 2 ok, Ampelstatus grün - kein Fehler | Proteinentfernung durchgeführt, Luftweg im Fluidic Modul gespült pH Dichtring erneuert, Inlet Gasket aus Kundenbestand erneuert Abfallschlauchweg gespült, Membranen erneuert Gerät kalibriert und überprüft, kein Fehler feststellbar | SWO-571396 | No fault found | Fluidic pathway | Fluidic pathway | Fluidic pathway | 328 | sporadically sample abort - umbilical cord blood! | sporadically sample abort - umbilical cord blood! 1 x error 0328 (no front air segment in the inlet LS sensor) on 27.06 - 4 measurements ok 5 Measurements on 26.06. ok on the installation day 25.06 15 measurements in total, of which 5 aborted on arrival QC level 1 + 2 ok, traffic light status green - no error
No | Visit | Complaint | RAME | Neon number to low/ Intensity | Inspection of the system revealed multiple oxi errors and failures of complete oxi test. Neon lamp low, halogen intensity, and 581 oxi spectrum errors. Also customer having connectivity issues. | Replaced oxi module 902-833R2424N042. Troubleshot to port isn’t live, customer putting in a ticket with his IT dept. | SWO-570865 | Defect | HW | Oxi module | Oxi module | 581 | Neon number to low/ Intensity | Inspection of the system revealed multiple oxi errors and failures of complete oxi test. Neon lamp low, halogen intensity, and 581 oxi spectrum errors. Also customer having connectivity issues.
No | 3 | Y | RCZE | Blank |  |  | D655E095C64739A3 | User Error, FSE | Demo key | Demo key | Demo key | mail: complaint? | Blank
No | Visit | Complaint | RSCH | Undicht | Einlassdichtung Kapillar undicht Schlauch nicht richtig auf Nadel gesetzt Ablass-Schlauch nach Nippel Einlass Kapillare verstopft | Schlauch Nadel ersetzt (BS) i.o Ablass-Schlauch Kapillare gekürzt i.o Nachfüllung, LS Justage und Pumpen Kal. i.o System i.o | SWO-571286 | Defect | HW | Inlet module | Inlet probe |  | Leaking | Inlet seal capillary leaking Hose not set correctly on needle Drain hose after nipple inlet capillary clogged
No | Visit | Complaint | RDEU | 26.06.2019 13:50Uhr Sr.Daniela tel:+492414142491: Das Gerät meldet jeden morgen: LP wechseln, nach einem Neustart funktioniert das Gerät wieder für 24Sttd 27.06.2019 8:30 Sr.Daniela +492414142491: Problem bleibt bestehen und kommt auch nach LP-Wechsel zurück |  |  | SWO-571895 | Defect | HW | Basic Unit | PM kit |  | 26.06.2019 13:50h Sr.Daniela tel:+492414142491: The device reports every morning: change LP, after a restart the device works again for 24hd 27.06.2019 8:30 Sr.Daniela +492414142491: Problem remains and comes back even after LP change
No | Visit | Complaint | RDEU | 25.06.2019 Frau Neitzsch +4928411072476 meldet Verbindungsprobleme. | QC und Patientendaten werden auf ganze Zahlen gerundet. Patientendatenabfrage über das Netzwerk meldet den Fehler Zeitüberschreitung. | Die QC und Patientendaten werden vom Gerät korrekt mit allen Nachkommastellen gesendet. Mit Herrn Wenting von Kühndata gesprochen - Nachkommastellen werden momentan vom System gerundet - Herr Wenting hat dies bei pH auf drei Nachkommastellen und bei allen anderen Parametern auf eine Nachkommastelle geändert. Daten werden jetzt korrekt übertragen. Die Patientendatenabfrage wird momentan von der Firma Kühndata programmiert. Frau Neitzsch bleibt mit Herrn Wenting in Kontakt. | SWO-571892 | No fault found | SW | External SW | HIS/LIS |  | 25.06.2019 Ms. Neitzsch +4928411072476 reports connection problems. | QC and patient data are rounded to integers. Patient data retrieval over the network reports the error timed out.
No | Visit | Complaint | RLTD | Please call Jepsy on 07414 920 186 | Leak detected, tried changing SP and SC, flushed analyser. OXI tubing has split. | Replaced OXI tubing and tested. Analyser has passed calibrations, QC's and is now ready for use. | SWO-571793 | Defect | HW | Oxi module | Pinch valve tube | Leak | Please call Jepsy on 07414 920 186 | Leak detected, tried changing SP and SC, flushed analyser. OXI tubing has split.
No | Visit | Complaint | RAME | Turbidity too high and spectrum mismatch errors on multiple patient samples | Customer is still getting 0581, 0291, 0292 Oxi spectrum mismatch with different blood samples. | Replaced hemolyzer. replaced all tubing from annual maintenance kit. Did Complete oximetry test and THB calibration. Also sending customer VK kit. Old hemolyzer is under warranty. Serial 908-552R2347N0010 | SWO-571611 | Defect | HW | Oxi module | Hemolyzer | 581, 291, 292 | Turbidity too high and spectrum mismatch errors on multiple patient samples | Customer is still getting 0581, 0291, 0292 Oxi spectrum mismatch with different blood samples.
No | Visit | Complaint | RCHN | 接口板腐蚀 | 接口板腐蚀,需要借件902-972一个 | 更换902-972后修复by tim | SWO-570933 | [5 empty] | Interface plate corrosion | Interface plate corrosion, need to borrow 902-972 one
No | Visit | Complaint | RCHN | 接口板腐蚀，蠕动泵头断裂 | 接口板腐蚀，蠕动泵头断裂，借902-971，902-972各一个 | 更换902-971，902-972后修复by tim | SWO-571061 | [5 empty] | Interface plate corrosion, peristaltic pump head fracture | Interface plate corrosion, peristaltic pump head fracture, borrow edi-902-972 each
No | Visit | Complaint | RCHN | 检测到泄露.报错1342 | 检测到泄露.报错1342 未更换备件 | 清洁进样口支架后修复by tim | SWO-571752 | Cleaned | HW | Inlet module | Inlet gasket | 1342 | A leak was detected.Reporting error 1342 | A leak was detected.The error 1342 did not replace the spare parts
No | Visit | Complaint | RCHN | 血氧报错1034 | 血氧报错1034，借902-980一个 | 更换902-980后修复by tim | SWO-570958 | Defect | HW | Oxi module | Hemolyzer | 1034 | Blood oxygen report error 1034 | Blood oxygen report error 1034, borrow 902-980 one
No | Visit | Complaint | RCHN | 传感器错误，报错4794 | 传感器错误，报错4794，借636-549一个 | 更换636-549后修复by tim | SWO-571691 | Defect | HW | Aspiration Module | Sensor, needle | 4794 | Sensor error, error 4794 | Sensor error, error 4794, borrow636-549 one
No | Visit | Complaint | RCHN | 暗视野测量超范围，报错1024 | 暗视野测量超范围，报错1024，已更换光单，现加借944-216（15557）、944-220（15822）、944-394（15754）、944-214（15517）、944-212（15986）缺货各一个 |  | SWO-572173 | Revisit required | HW | Process Module | Optical Unit | 1024 | Dark field measurement over range, error 1024 | Dark field measurement over range, reported error 1024, replaced light sheet, now added 944-216 (15557), 944-220 (15822), 944-394 (15754), 944-214 (15517), 944-212 (15986) missing one
No | Visit | Complaint | RAME | The printout is very difficult to read. The printer needs to be replaced. | Verified that printout was difficult to read. It was light and many columns of print were missing. | Replaced printer. Performed test printout. Printing is now acceptable. | SWO-572179 | Defect | HW | Printer Module | Printer Module |  | The printout is very difficult to read. The printer needs to be replaced. | Verified that printout was difficult to read. It was light and many columns of print were missing.
No | Visit | Complaint | RAME | Printer not working Properly | Inspection of printer revealed printout has a column of missing printed area. | Replaced print head. | SWO-571631 | Defect | HW | Printer Module | Printer Head |  | Printer not working Properly | Inspection of printer revealed printout has a column of missing printed area.
No | Visit | Complaint | RAME | Barcode not working properly | Inspection of the bar code reader revealed it has a cracked lens on left top corner and bar code reader is slow and intermittently doesn’t scan bar codes. | Replaced bar code reader. | SWO-571633 | Defect | HW | Basic Unit | Barcode reader | cracked lens | Barcode not working properly | Inspection of the bar code reader revealed it has a cracked lens on left top corner and bar code reader is slow and intermittently doesn’t scan bar codes.
No | Visit | Complaint | RDEU | Drucker ist offline, kein Ausdruck möglich. Feher wurde bei Auslieferung bemerkt. | Drucker ist offline, kein Ausdruck möglich. Fehler trat bei Auslieferung auf. | Printer Unit erneuert. Ausdruck ok. LQC wurden noch nicht geliefert. AQT nach Absprache mit Hr. Badendick gesperrt. LQC wird von FST bei anderem Kunden geholt. | SWO-572603 | Defect | HW | Printer Module | Printer Module |  | Printer is offline, no printing possible. Feher was noticed on delivery. | Printer is offline, no printing possible. Errors occurred on delivery.
Yes | Visit | Complaint | RDEU | Drucker ist offline, kein Ausdruck möglich. Feher wurde bei Auslieferung bemerkt. | AQT entsperren und LQC´s durchführen. | Sperrung AQT aufgehoben. Systemüberprüfung durchgeführt - ok. LQC TnI und NT-proBNP vorgenommen. | SWO-572604 | Re classified RMED | Same day UpSWO |  |  |  | Printer is offline, no printing possible. Feher was noticed on delivery. | AQT and perform LQC's.
No | Visit | Complaint | RPOL | Wygięta sonda aspiracyjna. | Serwis stwierdza uszkodzoną sondę aspiracyjną. | Wymiana sondy aspiracyjnej, testowanie aparatu. | SWO-572139 | Defect | HW | Inlet module | Probe bend |  | Bent aspiration probe. | The service finds a defective aspiration probe.
No | Visit | Complaint | RCHN | 样本无法测量 | 样本无法测量,需要借件902-980一个 | 更换902-980后修复by tim | SWO-571694 | [5 empty] | Samples cannot be measured | Samples cannot be measured, require a loan of 902-980
No | Visit | Complaint | RCHN | TNI定标失败 | TNI定标失败，借944-212（16160）一张 | 清洁，重新定标by tim | SWO-571804 | [5 empty] | TNI calibration failed | TNI bid failed, borrow 944-212 (16160)
No | Visit | Complaint | RCHN | 暗视野测量超范围，报错1024 | 暗视野测量超范围，报错1024,需要借件903-364一个 | 更换903-364后需要重新定标，加借定标卡，故此单关闭by tim | SWO-571021 | [5 empty] | Dark field measurement over range, error 1024 | Dark field measurement over range, error 1024, need to borrow 903-364 one
No | Visit | Complaint | RCHN | 软件自动退出或死机 | 软件自动退出或死机，借910-343、636-564、 903-198各一个 |  | SWO-572252 | [5 empty] | Software automatically exits or crashes | Software automatically exit sepsis or crash, borrow 910-343, 636-564, 903-198 each
No | Visit | Complaint | RIND | Cl showing very high values in every samples | Check analyzer status and calibration result in system log which is OK.Check patient result in which Chloride values shows very high in samples. | Check complete data log in which found that Cl shows high values in the current running SC while in previous Sc result were OK.Replace SP and check but still same problem with Cl result only.Replace sensor cassette and check calibration which is OK.Run QC which is also OK,then run a sample and found satisfactory result of Cl and all parameters. | SWO-572180 | [5 empty] | Cl showing very high values in every samples | Check analyzer status and calibration result in system log which is OK. Check patient result in which Chloride values shows very high in samples.
No | Visit | Complaint | RCHN | 不均匀的样本 | 不均匀的样本（未更换备件） | 清洁，更换进样针的管by tim | SWO-572143 | [5 empty] | Uneven samples | Uneven samples (no replacement parts)
No | Visit | Complaint | RCHN | 接口板腐蚀 | 接口板腐蚀，借902-972一个 | 更换902-972后修复by tim | SWO-571790 | [5 empty] | Interface plate corrosion | Interface plate corrosion, borrow 902-972 one
No | Visit | Complaint | RCHN | 条码扫码器无法扫描 | 条码扫码器无法扫描,需要借件910-304一个 | 更换910-304后修复by tim | SWO-571805 | [5 empty] | Barcode scanner cannot be scanned | Barcode scanner cannot be scanned, need slot 910-304 one
No | Visit | Complaint | RCHN | 漏液，试管架损坏 | 漏液，试管架损坏，借944-212（16030）、903-584一个 | 更换903-584，重新定标by tim | SWO-571845 | [5 empty] | Leak, damage to test tube rack | Leak, damage to test tube rack, borrow edgy 944-212 (16030), 903-584 one
No | Visit | Complaint | RCHN | 测试卡装反，报错5784 | 测试卡装反，报错5784 未更换备件 | 更换测试卡后修复 by Andrew | SWO-572433 | [5 empty] | Test card installed back, reported error 5784 | Test card installed reverse, error 5784 did not replace spare parts
No | Visit | Complaint | RCHN | 定标终止 | 定标终止，借944-394（15623）一个 |  | SWO-572445 | [5 empty] | Bid termination | Bid termination, loan 944-394 (15623) one
No | Visit | Complaint | RMAY | On Call : 0722 : Sample Error | On Call : 0722 : Sample Error | 1, Check inlet gasket - clogged 2. Flush inlet gasket 3. Run 2 point cal 4. Run Qc | SWO-572291 | [5 empty] | On Call : 0722 : Sample Error | On Call : 0722 : Sample Error
No | Visit | Complaint | RCHN | 液路堵塞 | 液路堵塞（未更换备件） | 清洁液路后修复 by Andrew | SWO-572561 | [5 empty] | Liquid road blockage | Liquid road blockage (no replacement parts)
No | Visit | Complaint | RCHN | 抽吸样本失败 | 抽吸样本失败 未更换备件 | 更换进样口支架后修复 | SWO-572441 | [5 empty] | Failed to pump samples | Sample pumping failed No replacement spare parts
No | Visit | Complaint | RCHN | BNP定标失败 1024 | BNP定标失败 1024,借944-258*16042一张。 | 重新定标后修复 by Andrew | SWO-571131 | [5 empty] | BNP calibration failure 1024 | BNP bid failure 1024, borrow 944-258 x 16042 one.
No | Visit | Complaint | RPAC | Lauren called 4454 9160 QC out of range QC 7 days old, and left out of fridge, ask customer to test new QC, ,she will come back to me called 26.06 - forgot to run called 27.06 - new QC passed 28.06 - QC now failing again on level 2 and 3 - will go through on Monday and have a look | Noticed that the filter at the front was blocked, cavro pump struggling to aspirate buffer (line 5 blocked), needle slightly bent, cup wash unit very dirty | Do a protein clean from pump 2, replace filter, replaced needle (new non stock) and cup wash unit, replace line 5 | SWO-572472 | [5 empty] | Lauren called 4454 9160 QC out of range QC 7 days old, and left out of fridge, ask customer to test new QC, ,she will come back to me called 26.06 - forgot to run called 27.06 - new QC passed 28.06 - QC now failing again on level 2 and 3 - will go through on Monday and have a look | Noticed that the filter at the front was blocked, cavro pump struggling to aspirate buffer (line 5 blocked), needle slightly bent, cup wash unit very dirty