## Sheet2
ANALYZER_TYPE | INVESTIGATION | CORRECTION | SWO Number | Filtering (Credit/Age/Class) | INESTIGATION (ENG) | CORRECTION(ENG) | RMED FaultCode L1(New) | RMED FaultCode L2(New) | RMED FaultCode L3(New) | RMED FaultCode L4(New) | Level 5
ABL5 | Dummy | Dummy | Dummy1 | Yes | Dummy | Dummy
ABL90 |  | SP90 Lot LI-04 ersetzt | CC80B80BFA3B0999 | No |  | SP90 Lot LI-04 replaced | Defect | SP | Pump tubing | broken | x
ABL90 | im Servicebereich Tests OK, neues Lösungspack eingesetzt, wieder Fehler, im Servicebereich alles OK, Lösungen werden sichtbar gut gepumpt |  | CC80A854FA57E622 | No | in the service area tests OK, new solution pack used, again errors, in the service area everything OK, solutions are visibly well pumped |  | Defect | SP | Pump tubing | pulled out | x
ABL90 | Revisión de funcionamiento con test de servicio. Sustitución de casete de electrodos. Limpieza de circuíto neumático. Revisión de sistema mecánico y neumático Calibración y verificación de funcionamiento. Correcto según pruebas. |  | CC8097C2FA562435 | No | Operation review with service test. Replacing electrode cassette. Cleaning of pneumatic ring. Mechanical and pneumatic system review Calibration and operation verification. Correct according to evidence. |  | Defect | SC | Chip error |  | x
ABL825 FLEX |  |  | CC8127FFFA0BB543 | No |  |  | Defect | Electrode | Membrane | pCO2
ABL820 FLEX | Error code 476 -measurement unstable and 328 air segment in sample | Replaced inlet probe. replaced pH BG tubing. | CC81083BFA4B8089 | No | Error code 476 -measurement unstable and 328 air segment in sample | Replaced inlet probe. replaced pH BG tubing. | Defect | Inlet module | Probe
AQT90 FLEX |  |  | CC80D253FA144C27 | No |  |  | Cleaned | Barcode scanner | Dust - Barcode Scanner
ABL90 | Customer recently replaced solution pack and sensor cassette. Checked data logs, pO2 is failing high out of range on internal and external QC level 3. | Re-run internal level 3 QC multiple times with pass results. Ran a total of 5 waste bloods. Cleaned QC bottle adapters. | CC82489D70433B18 | No | Customer recently replaced solution pack and sensor cassette. Checked data logs, pO2 is failing high out of range on internal and external QC level 3. | Re-run internal level 3 QC multiple times with pass results. Ran a total of 5 waste bloods. Cleaned QC bottle adapters. | User Error | waste blood | ext QC pO2 level 3 |  | x
ABL90 | On arrival on site tried back flushes (X9) LS adjusts and all fluid transport service tests (which passed). Thought problem could be related to SP, installed old SP from other analyzer and it seemed to work but only had 7 tests remaining. Went to install what was supposed to be a new SP but found it was previously activated and would not install properly (Credit Form to be sent by K. McKenzie). Installed new Sp and still had inhomogeneous rinse problem. Replaced tubing valve and inlet connector along with installing new inlet gasket and probe (new style). | Could not correct problem. | CC827B2F700C1262 | No | On arrival on site tried back flushes (X9) LS adjusts and all fluid transport service tests (which passed). Thought problem could be related to SP, installed old SP from other analyzer and it seemed to work but only had 7 tests remaining. Went to install what was supposed to be a new SP but found it was previously activated and would not install properly (Credit Form to be sent by K. McKenzie). Installed new Sp and still had inhomogeneous rinse problem. Replaced tubing valve and inlet connector along with installing new inlet gasket and probe (new style). | Could not correct problem. | Unknown | Revisit required | pO2 int QC Lev A low | New SP, pinch valve tube, connector gasket, probe & gasket (New) | x
ABL90 | On arrival on site verified error. Tried a calibration and rerunning level 1 QC, it still failed. Did back flush X5, went through all fluid transport and pO2 service checks they passed but error still present. | Replaced SP (on 39 test left on original SP) and installed new style inlet probe and inlet gasket. Replaced missing inlet clip. | CC8278AE7011125E | No | On arrival on site verified error. Tried a calibration and rerunning level 1 QC, it still failed. Did back flush X5, went through all fluid transport and pO2 service checks they passed but error still present. | Replaced SP (on 39 test left on original SP) and installed new style inlet probe and inlet gasket. Replaced missing inlet clip. | Defect | SP | pO2 int QC Lev A low | & new probe/inlet + clip | x
ABL90 | Found Unable to pump solutions and issues with Lac, K+ and Po2. Also found Flow Selector errors. | Performed Software upgrade to 3.0 MR2 Replaced SP and SC due to Flow selector issues and pO2 problems Performed Waste Blood aspiration and could not get the analyzer to stablize on pO2 and Lac. Will need to return Monday to complete second analyzer software upgrade and follow up with this analyzer as I could not bring it within range. | CC826D397034DA72 | No | Found Unable to pump solutions and issues with Lac, K+ and Po2. Also found Flow Selector errors. | Performed Software upgrade to 3.0 MR2 Replaced SP and SC due to Flow selector issues and pO2 problems Performed Waste Blood aspiration and could not get the analyzer to stablize on pO2 and Lac. Will need to return Monday to complete second analyzer software upgrade and follow up with this analyzer as I could not bring it within range. | Unknown | Revisit required | Lac, K, pO2 out on int ext QC | Replaced SC/SP, installed SW, waste blood | x
ABL90 |  |  | CC821EA77050A19B | No |  |  | Defect | HW | Waste tube |  | x
ABL825 FLEX |  | Austausch 945-612 | CC81EAE4703EB6E4 | No |  | Replacement 945-612 | Defect | Electrode | Electrode | pCO2
ABL825 FLEX | Previously spoke with Tech Support on bad CPU received. They shipped replacement. Installed replacement. | Installed new CPU assy. | CC81FDF7701C6941 | No | Previously spoke with Tech Support on bad CPU received. They shipped replacement. Installed replacement. | Installed new CPU assy. | Defect | DMS PC | CPU
ABL825 FLEX | Tested the sensor in service, the system indicated a calibration error. Tested bar with new generated codes, could not duplicate error. | Perform a flush of the air and waste valves in the fluid module. then ran calibration program to resolve error message. futher investigation required, customer is snding the bar code information. Contact is checking barcode print settings. | CC8211DF7019188D | No | Tested the sensor in service, the system indicated a calibration error. Tested bar with new generated codes, could not duplicate error. | Perform a flush of the air and waste valves in the fluid module. then ran calibration program to resolve error message. futher investigation required, customer is snding the bar code information. Contact is checking barcode print settings. | Cleaned | Fluidic Pathway | -
ABL835 FLEX |  |  | CC8241F8705A0344 | No |  |  | User Error | Replacement process | double o-ring | Reference
ABL825 FLEX |  | Einlassdichtung erneuert,div. Maschinenprogramme | CC81EE267045A249 | No |  | Inlet seal renewed,div. machine programs | Defect | Inlet module | Gasket
AQT90 FLEX |  |  | CC8223737016A9E3 | No |  |  | Adjustment | Cartridge Guide | Aborted calibration | 2877 Cardrige_4
AQT90 FLEX | Scanner überprüft->defekt, Funktionsprüfung, Kal.- LQC- und Aktivitätendateien überprüft | Softwareupdate auf v8.5.70, Sampler Modifikation Kit eingebaut | CC81DC2B703F82D6 | No | Scanner checked->defective, function check, Cal.- LQC and activity files checked | Software update built on v8.5.70, Sampler Modification Kit | Exchange of Item | Barcode scanner | Inlet Barcode
ABL80 CO-OX | The scanner has a very short scan (the red line is about 1 inch long at a distance of 10 inches from the analyzer) and the scan line moves around and changes shapes. | Setup analyzer, but unable to complete the installation. | CC845B915C07F8A3 | No | The scanner has a very short scan (the red line is about 1 inch long at a distance of 10 inches from the analyzer) and the scan line moves around and changes shapes. | Setup analyzer, but unable to complete the installation.
ABL90 |  |  | CC85EEFBD2477185 | Yes |  |  | Re classified RMED | Error occured during planned SW update |  | CAA/MRM | x
ABL90 |  |  | CC85EC44D261717D | No |  |  | Defect | HW | SP PCB |  | x
ABL90 | Instrument still in UIR on arrival. As per instructions from J. Gibbons and Tech Support, I tried new software, replacing tubings, flushing lines with bleach solution and doing all service program checks. Still have error code 1324 (in homogeneous rinse solution) Software upgrade as per FAN. Additionally changed time to 24 hr and set for seasonal time change. | Backed up all data and set-up. Installed Version 3.0 MR2 software and set to military time and automatic seasonal time change. No change in problem error message 1324 still present.. | CC85DDBFD240911E | No | Instrument still in UIR on arrival. As per instructions from J. Gibbons and Tech Support, I tried new software, replacing tubings, flushing lines with bleach solution and doing all service program checks. Still have error code 1324 (in homogeneous rinse solution) Software upgrade as per FAN. Additionally changed time to 24 hr and set for seasonal time change. | Backed up all data and set-up. Installed Version 3.0 MR2 software and set to military time and automatic seasonal time change. No change in problem error message 1324 still present.. | Unknown | Revisit required | See previous next upSWO same analyzer | Backflush, new tubes, SW upgrade | x
ABL90 |  |  | CC85DE4BD23AA1A4 | No |  |  | Defect | SC | pCO2 | Sens | x
ABL810 FLEX |  |  | CC85CFAAD22FADB6 | No |  |  | Cleaned | Gas module | Sirai
ABL820 FLEX | No power to unit. Power source checks OK. | Replaced line fuses. | CC85E8E4D23374E8 | No | No power to unit. Power source checks OK. | Replaced line fuses. | Defect | Power Supply | Fuses blown
ABL90 |  |  | CC87092D4805D794 | Yes |  |  | Re classified RMED |  |  |  | x
ABL90 | Checked battery status in service program. Found cycle count and percent remaining at zero | Replaced battery pack. | CC87736848199A19 | No | Checked battery status in service program. Found cycle count and percent remaining at zero | Replaced battery pack. | Defect | HW | Battery module | Shuts down | x
ABL90 | Disassembled analyzer and found leakage around both in out ports of the hemolyzer. Cleaned and changed all tubings. Installed new waste tubing 842-345. All good. Reassembled analyzer. All good. Tried several solution packs as new solution packs had not arrived yet.Managed to get one to work. Advised customer that it would be good practice to install new solution pack when it arrives tomorrow. | Replaced waste tubing. | CC873AAD485AAD97 | No | Disassembled analyzer and found leakage around both in out ports of the hemolyzer. Cleaned and changed all tubings. Installed new waste tubing 842-345. All good. Reassembled analyzer. All good. Tried several solution packs as new solution packs had not arrived yet. Managed to get one to work. Advised customer that it would be good practice to install new solution pack when it arrives tomorrow. | Replaced waste tubing. | Defect | HW | Waste tube | See 2 previous UpSWO same month | x
ABL90 |  |  | CC86E3A548058FD7 | Yes |  |  | Re classified RMED | Logistic upSWO |  |  | x
ABL90 |  | SC erneuert; Konditionierung abgewartet Einlass gereinigt Kal & QC i.O. | CC8737D048639E75 | No |  | SC renewed; Conditioning awaited inlet cleaned Kal & QC i.O. | Defect | SC | Lac QC |  | x
ABL825 FLEX | Trend over the the last 2 months was on the high border of the 2sd range. even after changing electrodes. | Replaced electrode. QC back close to the mean. | CC86F81648300043 | No | Trend over the the last 2 months was on the high border of the 2sd range. even after changing electrodes. | Replaced electrode. QC back close to the mean. | Defect | Electrode | Electrode | pCO2
ABL80 CO-OX | EL CLIENTE TIENE DUDAS SOBRE ALGUNOS PARAMETROS DEL GASOMETRO CON RESPECTO AL LABORATORIO, SE EFECTUAN PRUEBAS DE REPRODUCIBILIDAD CON LOS DOS SISTEMAS Y CON CONTROLES DE CALIDAD, PRUEBAS CORRECTAS. CALIBRADO Y VERIFICADO MEDIANTE Q.C. |  | CC885B25BE416698 | No | THE CUSTOMER HAS DOUBTS ABOUT SOME PARAMETERS OF THE GASOMETRO WITH RESPECT TO THE LABORATORY, REPRODUCIBILITY TESTS ARE EFFECTED WITH THE TWO SYSTEMS AND WITH QUALITY CONTROLS, CORRECT TESTS. CALIBRATED AND VERIFIED BY Q.C.
ABL80 CO-OX | I flushed out the line and that did not help. I inspected all tubings and found the oxymeter tubing off. | I reinstalled the tubing. | CC882FA1BE5436EA | No | I flushed out the line and that did not help. I inspected all tubings and found the oxymeter tubing off. | I reinstalled the tubing.
ABL80 CO-OX | Factory tubing leak from Hemolyzer to Peri pump, flowed to Analog board. | Removed and replaced tubing and Analog board. | CC887E2BBE0B385C | No | Factory tubing leak from Hemolyzer to Peri pump, flowed to Analog board. | Removed and replaced tubing and Analog board.
ABL90 | Performed back flush no resolve. Manual recommends replacing hemolyzer. After hemolyzer was replaced the cals would not come in. Impeadance errors. | Replaced hemolyzer. Replaced SC. tHB calibration. Customer will monitor as bloods comes in. | CC883734BE4C1128 | No | Performed back flush no resolve. Manual recommends replacing hemolyzer. After hemolyzer was replaced the cals would not come in. Impeadance errors. | Replaced hemolyzer. Replaced SC. tHB calibration. Customer will monitor as bloods comes in. | Defect | HW | Hemolyzer | Error code 581 | x
ABL90 | Print out at METHB becomes a line intermittently and software upgrade is required. | Replaced printer and upgraded software to V-3.0 MR2. Calibrated and ran QC. Instrument operates normally. | CC888B5BBE4C438E | No | Print out at METHB becomes a line intermittently and software upgrade is required. | Replaced printer and upgraded software to V-3.0 MR2. Calibrated and ran QC. Instrument operates normally. | Defect | HW | Printer | & Installed SW - Potentially the same problem as Supressed MetHb result in TL line 13 of print report | x
ABL90 |  |  | CC8872B2BE56C848 | No |  |  | Setup | Fixed assigned values |  |  | x
ABL825 FLEX | Customer saved the bar code and tested several times. Scanned barcode would intermittently capture incorrect information. | Duplicated the issue, captured data and service dump for Quality Assurance case. | CC8869B5BE589731 | No | Customer saved the bar code and tested several times. Scanned barcode would intermittently capture incorrect information. | Duplicated the issue, captured data and service dump for Quality Assurance case. | Unknown | Basic unit | Barcode Reader
ABL827 FLEX | 1-7-14 I worked with Kevin and techsupport over the phone yesterday. 1-8-14 On site. I wanted to come in today and take a look at the analyzer and speak with the customer. I am sending them the procedure for changeing the pCO2 membrane until there is a fix for this issue. There are techline cases opended for the pCO2 running low on level 1 QC. I spoke with Jeremy and explained the letter and the techline case. | Checked data logs. spoke with Jeremy and sent the customer letters explaining the pCO2 issue. Kevin McKenzie is also sending a new lot of pCO2 membranes. I gave Darlene and Jeremy my cell number ans asked that they keep me informed on the pCO2. I will let them know when there has been a fix from RMED. | CC882154BE3FE0E8 | No | 1-7-14 I worked with Kevin and techsupport over the phone yesterday. 1-8-14 On site. I wanted to come in today and take a look at the analyzer and speak with the customer. I am sending them the procedure for changeing the pCO2 membrane until there is a fix for this issue. There are techline cases opended for the pCO2 running low on level 1 QC. I spoke with Jeremy and explained the letter and the techline case. | Checked data logs. spoke with Jeremy and sent the customer letters explaining the pCO2 issue. Kevin McKenzie is also sending a new lot of pCO2 membranes. I gave Darlene and Jeremy my cell number ans asked that they keep me informed on the pCO2. I will let them know when there has been a fix from RMED. | Defect | Electrode | Electrode | pCO2
ABL825 FLEX |  |  | CC886312BE077644 | No |  |  | Defect | Electrode | Electrode | Lac
ABL835 FLEX |  |  | CC888155BE5D6DEF | No |  |  | Defect | ph/BG Module | Pump tubing
ABL825 FLEX | Lamp unit test failed under oxi complete test. | Lamp module not adjusted, Installed loaner OXI module. | CC992AF9BC185471 | No | Lamp unit test failed under Oxi complete test. | Lamp module not adjusted, Installed loaner OXI module. | Defect | HW | Oxi Module | OXI PCB
ABL835 FLEX | DMS disconnect intermittant 2x air valve short on module BG et Met module ,Au PCB et BG pcb et Met Pcb defective |  | CC8887FBBE00140C | No | DMS disconnect intermittant 2x air valve short on module BG et Met module ,Au PCB et BG pcb et Met Pcb defective |  | Unknown | PCB?
ABL80 CO-OX |  |  | CC89950A345F15EB | No
ABL80 CO-OX |  |  | CC89B8E1340F28CA | No
ABL80 FLEX Basic version |  |  | CC8EBE7D0C3292A5 | No
ABL90 | Confirmed issues. Software update needed. Found analyzer IP to be dynamic. Found analyzer area's ambient temp to be at 28.5 degrees C. Analyzers Max room operating temp 30 degrees C. This could impact pO2 calibrations as well as internal and external QC per Operators Manual. Found external controls to be around 25 degrees C. Input temp for manual run has been Ambient Room Temp. | Worked with IS to drop a new Scope of IP addresses to use for the analyzers. Installed software version 3.0 MR2. Scope given and configured on this analyzer. Power management has been disabled so the NIC will not sleep upon no activity. Installed new SP and SC as well as flushed all lines. Tested NetOp remote support from Radiance with success. Performed waste blood to condition SC. | CC89A45D34408B47 | No | Confirmed issues. Software update needed. Found analyzer IP to be dynamic. Found analyzer area's ambient temp to be at 28.5 degrees C. Analyzers Max room operating temp 30 degrees C. This could impact pO2 calibrations as well as internal and external QC per Operators Manual. Found external controls to be around 25 degrees C. Input temp for manual run has been Ambient Room Temp. | Worked with IS to drop a new Scope of IP addresses to use for the analyzers. Installed software version 3.0 MR2. Scope given and configured on this analyzer. Power management has been disabled so the NIC will not sleep upon no activity. Installed new SP and SC as well as flushed all lines. Tested NetOp remote support from Radiance with success. Performed waste blood to condition SC. | User Error | Ext QC temp correction | new SW + Setup LIS |  | x
ABL90 | Icon needed to be pressed several times before task can be performed | Performed screen calibration | CC89949E3406CEB4 | No | Icon needed to be pressed several times before task can be performed | Performed screen calibration | Adjustment | Screen calibration |  |  | x
ABL90 |  |  | CC8AA890AA4D9D48 | No |  |  | Defect | SP | Fluid transport |  | x
ABL90 |  |  | CC899D6C343834D6 | No |  |  | Setup | Roma.ini | Barcode Reader |  | x
ABL90 |  |  | CC89AB8C3451A913 | No |  |  | Replaced | Inlet gasket | Bili + inhomo sample |  | x
ABL90 | Limpeza do circuito de amostra. Substituição da sonda e porta de entrada. Ajuste da entrada, ajuste do SL, calibração da bomba, execução do teste de pressão, execução do teste de detecção remoção de coágulos. Inicialização. Comprovação do histórico de calibrações, QC, medidas e erros do sistema. Verificação do correcto funcionamento do equipamento. |  | CC8980FC3453FE6D | No | Cleaning the sample circuit. Replacement of probe and inlet port. Inlet adjustment, SL adjustment, pump calibration, pressure test run, test run of clotting removal detection. Startup. Proof of the history of calibrations, QC, Measurements and errors of the system. Checking the correct operation of the equipment. |  | Cleaned | Fluidic Pathway | replaced probe/gasket |  | x
ABL825 FLEX |  |  | CC896F373426D2B0 | No |  |  | Defect | Electrode | Electrode | pCO2
ABL825 FLEX | PATIENT DAT LOGS SHOWED THAT THE CUSTOMER HAD RUN THE CAP SURVEY IN 95U MICRO MODE DATA LOGS SHOW THAT THE LEVEL 1 QC WAS RUNNING AT 63.7.INSPECTION OF CHAMBER SHOWS RINSE IN PH CHAMBER DURING WHEN DRAINEF | REPLACED PH O RING ANF REFERENCE MEMBRANE.RAN CALIBRATIONS AND LEVEL 1 QC AT 65.6 | CC897F233420D96F | No | PATIENT DAT LOGS SHOWED THAT THE CUSTOMER HAD RUN THE CAP SURVEY IN 95U MICRO MODE DATA LOGS SHOW THAT THE LEVEL 1 QC WAS RUNNING AT 63.7.INSPECTION OF CHAMBER SHOWS RINSE IN PH CHAMBER DURING WHEN DRAINEF | REPLACED PH O RING ANF REFERENCE MEMBRANE. RAN CALIBRATIONS AND LEVEL 1 QC AT 65.6 | Defect | Electrode | Gasket | pH
ABL825 FLEX | CPU is damaged and possible damage to other circuit boards. Water and ceiling materialr was everywhere inside the casing. | Dissasembled and let dry over night. Will return in the morning | CC89C42F3431A3A0 | Yes | CPU is damaged and possible damage to other circuit boards. Water and ceiling materialr was everywhere inside the casing. | Dissasembled and let dry over night. Will return in the morning | Re classified RMED | Waterline burst
ABL835 FLEX | pH measurement unstable for patient samples, pH QC results excessively bouncing around on all levels, pH drift during calibrations. Inspected sample path, tested and inspected all valves. Air valves for pH BG module and for El Met module stuck in open position. Replaced jammed air valves. Successfully performed calibrations and first level of QC. Second level of QC failed, indicating measurement unstable for pH, and pH results were far outside the lower tolerance limit. | Unable to repair at this time. Started CQC. | CC89D8CD343ED5DF | No | pH measurement unstable for patient samples, pH QC results excessively bouncing around on all levels, pH drift during calibrations. Inspected sample path, tested and inspected all valves. Air valves for pH BG module and for El Met module stuck in open position. Replaced jammed air valves. Successfully performed calibrations and first level of QC. Second level of QC failed, indicating measurement unstable for pH, and pH results were far outside the lower tolerance limit. | Unable to repair at this time. Started CQC. | Defect | ph/BG Module | Valve, air
ABL815 FLEX |  |  | CC8986FD34046910 | No |  |  | Defect | Printer module | Printer head
AQT90 FLEX | Sustitución de la unidad óptica. Limpieza interna (exceso de polvo), | Sustitución de la unidad óptica. Limpieza interna (exceso de polvo) | CC89BC583439DCD5 | No | Replacing the optical drive. Internal cleaning (excess dust), | Replacing the optical drive. Internal cleaning (excess dust)
TCM-COMBIM | Tried rebooting the analyzer but the sensor was not read by the analyzer. old CombiM module 903-111R0161N012 | Replaced CombiM module. | CC89B7BE345531BD | No | Tried rebooting the analyzer but the sensor was not read by the analyzer. old CombiM module 903-111R0161N012 | Replaced CombiM module. | Defect | HW | Combi M Module
ABL90 | Flüssigkeitsstrecke gespült, Dichtung und Einlassnadel geprüft | SP90 erneuert, Hart ausgeschaltet | CC8AAF98AA3D933F | No | liquid section rinsed; Seal and inlet needle tested | SP90 renewed, Hard off | Defect | SP | Pump cal error | Potential pump tube | x
ABL90 |  |  | CC8B1C24AA0139F0 | No |  |  | Performance | Repeatability pO2 | Potential wetability problem | SW installed | x
ABL90 |  |  | CC8B1AD9AA3DE7DD | Yes |  |  | Re classified RMED |  |  |  | X
ABL90 | Comprobación general del analizador chequeo de funcionamiento, cambio del casette de electrodos. |  | CC8ACD92AA2E7F80 | Yes | General check of the operating check analyzer, change of the electrode cassette. |  | Re classified RMED | User training | Perhaps Defect SC |  | x
ABL800Basic |  |  | CC8AFAF3AA2F6140 | No |  |  | Defect | Electrode | Electrode | Reference
ABL825 FLEX |  |  | CC8ACAAFAA209344 | No |  |  | Defect | Electrode | Electrode | pH
ABL800Basic |  |  | CC8B146EAA464CF9 | No |  |  | Defect | Inlet module | Inlet Gasket
ABL835 FLEX | Analyzer generates LIS HIS: Failed to send packet when demographics are being requested from it. Patient sample information is received by LIS successfully. Checked all connections and settings, rebooted term server. Customer using Quadra Med LIS. | Asked customer for trace log, communicated with RAME Tech support. | CC8AC734AA5FD5E2 | No | Analyzer generates LIS HIS: Failed to send packet when demographics are being requested from it. Patient sample information is received by LIS successfully. Checked all connections and settings, rebooted term server. Customer using Quadra Med LIS. | Asked customer for trace log, communicated with RAME Tech support. | Setup | His/Lis Setup, not upSWO
ABL825 FLEX | checked hemolyzer and lamp unit, checked power supply, performed complete oxi test, | replaced Oxi board and no change, removed and cleaned fuses on AU board, replaced power supply, performed complete oxi test, performed tHb cal | CC8B177EAA473D8E | No | checked hemolyzer and lamp unit, checked power supply, performed complete oxi test, | replaced Oxi board and no change, removed and cleaned fuses on AU board, replaced power supply, performed complete oxi test, performed tHb cal | Defect | Power Supply | Missing Trigger
ABL90 | siehe ausgeführte Arbeiten | siehe ausgeführte Arbeiten | CC8E6E6C0C2D2B5F | No | see work carried out | see work carried out | Defect | SP | Fluid transport | Perhaps pump tube | x
ABL90 |  | Hakenventil überpüft ->i.O. Spectrometer getauscht -> 1h initialisieren lassen; Oxi Complete Test durchgeführt neuen LP eingesetzt, da leer -> gl. Fehler Hemolyzer getauscht -> Oxi Complete Test, Kal, tHb Kal -> Fehler weg | CC8EAAF20C47F770 | No |  | Hook valve over ->i.O. Spectrometer exchanged -> 1h initialize; Oxi Complete Test performed new LP used because empty -> gl. Error Hemolyzer exchanged -> Oxi Complete Test, Kal, tHb Kal -> Error away | Defect | HW | Hemolyzer | Spectrum mismatch | x
ABL835 FLEX |  |  | CC9391F2E421E4FE | No |  |  | Defect | Electrode | Membrane | Na+
ABL825 FLEX | Create case and submit information. | Created case, notified customer and submit service dump. | CC8EC4A20C1FF1E3 | No | Create case and submit information. | Created case, notified customer and submit service dump. | Setup | Barcode scanner | -
ABL825 FLEX | REVIEWED DATA LOGS.SENSITIVITY AT 90% LEVEL 1 QC AT 63.8 | REPLACED PH BG MEASURING CHAMBER,HEAT DISTRIBUTION PLATE AND SEALING RINGS.REPLACED PCO2 MEMBRANE WITH A NEW LOT.REPLACED ALL PUMP,INLET AND ASSOCIATED TUBINGS.PERFORMED MULTIPLE 2 POINT CALIBRATIONS.PCO2 AT 97% SENSITIVITY AND LEVEL 1 QC AT 65.8 | CC8EBA4F0C185442 | No | REVIEWED DATA LOGS. SENSITIVITY AT 90% LEVEL 1 QC AT 63.8 | REPLACED PH BG MEASURING CHAMBER,HEAT DISTRIBUTION PLATE AND SEALING RINGS. REPLACED PCO2 MEMBRANE WITH A NEW LOT. REPLACED ALL PUMP,INLET AND ASSOCIATED TUBINGS. PERFORMED MULTIPLE 2 POINT CALIBRATIONS. PCO2 AT 97% SENSITIVITY AND LEVEL 1 QC AT 65.8 | Defect | Electrode | Electrode | pCO2
ABL835 FLEX | pH measurement unstable, pH QC and calibration issues. Both measuring module air valves jammed in open position. Ran isolation test successfully. While re-installing the AU engine software on the new AU board, error encountered at appbase.out stage and software install was terminated. Re-tried the install but this would not start, even after rebooting the analyzer. Attempted to re-install Windows, but process is very sluggish from the start. Process stalls when checking HDD during install. Rebooted and tried again, but keyboard was now unresponsive. Tried new power supply despite power supply check being successful. | Replaced analyzer control pcb temporarily (see SWO B92M4ZVYO) | CC8EE75E0C14D5E7 | No | pH measurement unstable, pH QC and calibration issues. Both measuring module air valves jammed in open position. Ran isolation test successfully. While re-installing the AU engine software on the new AU board, error encountered at appbase.out stage and software install was terminated. Re-tried the install but this would not start, even after rebooting the analyzer. Attempted to re-install Windows, but process is very sluggish from the start. Process stalls when checking HDD during install. Rebooted and tried again, but keyboard was now unresponsive. Tried new power supply despite power supply check being successful. | Replaced analyzer control pcb temporarily (see SWO B92M4ZVYO) | Unknown | Basic unit | Analyzer Control PCB
ABL825 FLEX | The halogen lamp intensity is 2500. It needs to be adjusted. | Adjusted the halogen lamp intensity to the manufacturer's specifications. There was some maintainence that was due. I performed that as well. | CC8E9A5F0C2C6CCE | No | The halogen lamp intensity is 2500. It needs to be adjusted. | Adjusted the halogen lamp intensity to the manufacturer's specifications. There was some maintainence that was due. I performed that as well. | Adjustment | HW | Oxi Module | Lamp unit
AQT90 FLEX |  |  | CC8EAA750C442634 | Yes |  |  | Re classified RMED | Training new user
TCM-COMBIM |  |  | CC96AE22D036E3D8 | No |  |  | Defect | HW | PC Module | Sound Corrupted
ABL80 CO-OX | Ablagerungen am Sensor un am Hemolyzer | Hemol. & Sensor neu | CC8FB1068244576D | No | Deposits on the sensor and at the Hemolyzer | Hemol. & Sensor new
ABL90 |  |  | CC901436820401ED | Yes |  |  | Re classified RMED | Logistic upSWO |  |  | x
ABL825 FLEX |  |  | CC90177B820CBA95 | No |  |  | Defect | Inlet module | Probe bend
ABL835 FLEX | Tried new HDD, eventually able to successfully install OS, but unable to install AU engine software despite seeking RAME Tech Support's assistance. Tried old Analyzer Control PCB, was able to reload Windows, Roma and the AU Engine software, but the problem with pH still exists. | Replaced Hard Disk Drive | CC9009388210D5ED | No | Tried new HDD, eventually able to successfully install OS, but unable to install AU engine software despite seeking RAME Tech Support's assistance. Tried old Analyzer Control PCB, was able to reload Windows, Roma and the AU Engine software, but the problem with pH still exists. | Replaced Hard Disk Drive | Defect | DMS PC | HDD
ABL800Basic |  |  | CC9001FF821C5192 | No |  |  | Defect | Electrode | Membrane | Lac
TCM-COMBIM | Inspection of analyzer revealed all connections are correct. Power supply is functioning as expected with specifications. Suspect CPU. | The CPU was replaced to resolve the reported Sound H W corruption issue. | CC8FE76F82596F2F | No | Inspection of analyzer revealed all connections are correct. Power supply is functioning as expected with specifications. Suspect CPU. | The CPU was replaced to resolve the reported Sound H W corruption issue. | Defect | HW | PC Module | Sound Corrupted
ABL80 FLEX Basic version | Verificar el estado del Pack de Soluciones. Borrado del mismo por haber sido instalado previamente, verificar funcionamiento correcto del cassette de electrodos y del Pack de soluciones. Revisión y comprobación del equipo con QC´s. | Borrado de los datos del Cal Pack y verificar correcto funcionamiento del equipo. | CC9136BFF84238A0 | No | Check the status of the Solution Pack. Erase it because it has been previously installed, check the correct operation of the electrode cassette and the Solution Pack. Review and check the equipment with QC's. | Erase the Cal Pack data and verify the correct operation of the equipment.
ABL90 | Checked CF card and CPU type | removed and replaced CF card, entered Bios and changed boot sequence, attempted to install windows install failed with the error 'Unable to detect a valid disk' tried another CD Rom readed with the same error result, | CC917F03F8633DC0 | No | Checked CF card and CPU type | removed and replaced CF card, entered Bios and changed boot sequence, attempted to install windows install failed with the error 'Unable to detect a valid disk' tried another CD Rom readed with the same error result, | Unknown | Revisit required | Disk boot failure | New UpSWO next day - CF card replace | x
ABL90 |  |  | CC9769C9460CC9A9 | No |  |  | Defect | SP | pO2 int QC | Rilibæk??? | x
ABL827 FLEX | I spoke with another FSE and he said replacing the measuring chamber fixed his pCO2 problem so I wanted to try it. | Replaced the pH BG measuring chamber. Remembraned pCO2 (lot # R0504) electrode with customer stock. | CC90F1D5F80CB6CB | No | I spoke with another FSE and he said replacing the measuring chamber fixed his pCO2 problem so I wanted to try it. | Replaced the pH BG measuring chamber. Remembraned pCO2 (lot # R0504) electrode with customer stock. | Defect | Electrode | Electrode | pCO2
ABL825 FLEX |  | Flow regulator left erneuert | CC9129BCF83E7510 | No |  | Flow regulator left renewed | Defect | Gas module | Gas valve, left
ABL835 FLEX | Inspected measuring module valves and noticed that both air valves were stuck again. Replaced the AU board with a new one, and successfully installed the AU Engine software. Ran several cals, but eventually pH became unstable again. Changed pH electrode connector, upon restarting, leak test failed. Inspected measuring module valves, both air valves had stuck again. Changed power supply and used the old AU board. Unable to calibrate inlet. Used new AU in conjunction with new power supply. Performed several cals, pH sens dropped and pH was drifting. Tried new pH electrode. Replaced pH BG board with a used one for troubleshooting purposes. | Replaced power supply, Analyzer control PCB, pH electrode connector, pH BG board (with used one for temporary troubleshooting purposes). | CC9108E6F85DD5F2 | No | Inspected measuring module valves and noticed that both air valves were stuck again. Replaced the AU board with a new one, and successfully installed the AU Engine software. Ran several cals, but eventually pH became unstable again. Changed pH electrode connector, upon restarting, leak test failed. Inspected measuring module valves, both air valves had stuck again. Changed power supply and used the old AU board. Unable to calibrate inlet. Used new AU in conjunction with new power supply. Performed several cals, pH sens dropped and pH was drifting. Tried new pH electrode. Replaced pH BG board with a used one for troubleshooting purposes. | Replaced power supply, Analyzer control PCB, pH electrode connector, pH BG board (with used one for temporary troubleshooting purposes). | Defect | Basic unit | Aanalyzer Control PCB
ABL800Basic |  |  | CC911EA1F839F22D | No |  |  | Defect | Electrode | Electrode | pCO2
ABL820 FLEX | Looking at activity logs, there were cleaning and cals aborted.Must be air leak in system | Removed alll electrodes and cleaned chamber and reseated elctrodes. Ran 10 water samples all OK. | CC91259CF84A818D | No | Looking at activity logs, there were cleaning and cals aborted. Must be air leak in system | Removed alll electrodes and cleaned chamber and reseated elctrodes. Ran 10 water samples all OK. | Cleaned | Fluidic Pathway
ABL820 FLEX | pH passing Level 2 and 3 level one failing range is 6.793-6.833, coming in between 6.819 and 6.889, customer stated she does not want to train until the QC materials come in, will return when customer has all supplies | installed ABL820; pH high on level one, cleaned chamber and rinsed electrode reran for pH only; reran Calibration; no help ; set up time; set up calibration schedule; set up reports; option keys installed | CC914383F852B909 | No | pH passing Level 2 and 3 level one failing range is 6.793-6.833; coming in between 6.819 and 6.889; customer stated she does not want to train until the QC materials come in; will return when customer has all supplies | installed ABL820; pH high on level one, cleaned chamber and rinsed electrode reran for pH only; reran Calibration; no help ; set up time; set up calibration schedule; set up reports; option keys installed | Unknown | Installation
AQT90 FLEX |  |  | CC914D22F81B7038 | No
AQT90 FLEX |  |  | CC9E13019404AD05 | No
AQT90 FLEX |  |  | CC9E1A45940DAD06 | No
Tosca-500 | Fehler wie oben auch mit Kundensensor am Werkstattgerät aufgetreten, Kundengerät mit Werkstattsen-sor o.k. | Bitte Gerät ohne Befund zum Kunden zurückschicken. | CC912326F829D9BA | No | Error as above also occurred with customer sensor on the workshop device, customer device with workshop sor o.k. | Please return the device to the customer without any findings. | Defect | Electrode |  |  | TechLine 161
ABL90 | Unit starts to boot but stops with a disk boot error. Invalid disk. The same error occurs with a new flash card. | Replaced CPU unit. Installed original flash card and installed version 3.0 MR2 software. | CC9279796E17CE56 | No | Unit starts to boot but stops with a disk boot error. Invalid disk. The same error occurs with a new flash card. | Replaced CPU unit. Installed original flash card and installed version 3.0 MR2 software. | Defect | HW | CPU | Disk boot failure - previous UpSWO CF card replaced | x
ABL90 |  |  | CC9278C36E4172F1 | No |  |  | Defect | SC | Leak detected | SP & probe + cleaned | x
ABL825 FLEX | Remembraned lactate. cleaned contacts. All other parameters OK. Remembraned reference electrode. There is a shift in level 4 QC on ph and all electrolytes . looked at ranges and discovered ranges needed to be adjusted. | Replaced Lactate Electrode. Remembraned reference electrode. Customer will monitor. | CC92548E6E11541D | No | Remembraned lactate. cleaned contacts. All other parameters OK. Remembraned reference electrode. There is a shift in level 4 QC on ph and all electrolytes . looked at ranges and discovered ranges needed to be adjusted. | Replaced Lactate Electrode. Remembraned reference electrode. Customer will monitor. | Defect | Electrode | Electrode | Lac
ABL80 | Comprobación del estado del analizador. Se comprueban registros y se verifica funcionamiento, se revisan canales de fluídica y bomba peristática, se realiza un a comprobación del canal de zero (valores dentro de márgenes). Limpieza del conector (pins) del equipo hacia la cassette de electrodos. Se sustituye el Manifold unit así como casssette de electrodos y se realizan ciclos de calidad para verificar estado de los electrodos. Revisión y verificación del equipo. | Se sustituye el Manifold unit y se realizan ciclos de calidad para verificar estado de los electrodos. Revisión y verificación del equipo. | CC93AFAAE427FA36 | No | Checking the status of the scanner. Records are checked and operation is checked, fluidic channels and peristhetic pump are checked, a zero channel check (values within margins) is performed. Cleaning the connector (pins) of the equipment towards the electrode cassette. The Manifold unit is replaced as well as electrode casssette and quality cycles are performed to check electrode status. Equipment review and verification. | The Manifold unit is replaced and quality cycles are performed to check the status of the electrodes. Equipment review and verification.
ABL90 |  |  | CC9387F6E450F112 | No |  |  | Cleaned | Fluidic Pathway | replaced SC + SP |  | x
ABL90 |  |  | CC99C33332579751 | Yes |  |  | Re classified RMED | Telephone support |  |  | LEH
ABL90 |  |  | CC99C1C63242974F | Yes |  |  | Re classified RMED | Telephone support |  |  | LEH
ABL90 | Flüssigkeitsstrecke freigespült, alle Tests innerhalb Service i.O., Kal, QC |  | CC93E2FAE40C58E2 | No | Liquid track rinsed free, all tests within service i.O., Kal, QC |  | Cleaned | Fluidic Pathway |  |  | x
AQT90 FLEX | Retirada de tubo que bloqueaba el posicionamiento de los casete. realizacion de limpieza y muestras | Retirada de tubo que bloqueaba el posicionamiento de los casete. | CC9557555A59FB5F | No | Removing tube that blocked cassette positioning. cleaning and samples | Removing tube that blocked cassette positioning.
ABL835 FLEX |  |  | CC9375BBE44339B3 | No |  |  | Cleaned | Auto Check Module | Dust - Barcode Scanner
ABL825 FLEX | On arrival on site checked data logs and found 26 2 point cals that have failed for drift in the last 30 days and over 40 QC (all levels) failing with pCO2 values high out of insert ranges. Visual inspection found that the pO2 electrode was missing its lower sealing ring and the pCO2 electrode was missing the lower plastic electrode cover and seal ring. | Replaced pO2 seal ring. Ordered new pCO2 electrode (current electrode under warranty). | CC93E63CE40B98E1 | No | On arrival on site checked data logs and found 26 2 point cals that have failed for drift in the last 30 days and over 40 QC (all levels) failing with pCO2 values high out of insert ranges. Visual inspection found that the pO2 electrode was missing its lower sealing ring and the pCO2 electrode was missing the lower plastic electrode cover and seal ring. | Replaced pO2 seal ring. Ordered new pCO2 electrode (current electrode under warranty). | User Error | Replacement process | Electrode | Gasket missing
ABL837 FLEX |  |  | CC965716D045CBF3 | No |  |  | Defect | HW | Oxi Module | OXI PCB
ABL820 FLEX | On arrival on-site checked data logs and found that on 1 16 the 7AM QC did not run due to a calibration error could not find anything in logs about today. Did hear noise coming from the Auto Check fan. Visual inspection of the Auto Check module found that the plastic protective cover had fallen off, this possibly could have affected the scheduled QC run. | Replaced the Auto Check fan. Removed protective cover. | CC93C357E44C5BC7 | No | On arrival on-site checked data logs and found that on 1 16 the 7AM QC did not run due to a calibration error could not find anything in logs about today. Did hear noise coming from the Auto Check fan. Visual inspection of the Auto Check module found that the plastic protective cover had fallen off, this possibly could have affected the scheduled QC run. | Replaced the Auto Check fan. Removed protective cover. | Defect | Auto Check Module | Leak guard
ABL825 FLEX |  |  | CC93838FE42C89E5 | No |  |  | Defect | Power Supply | -
ABL825 FLEX |  |  | CC962340D03954F9 | No |  |  | Defect | HW | Oxi Module | OXI PCB
ABL80 FLEX Basic version |  |  | CC986E25BC635501 | No
ABL80 OSM | Poor precision for tHb on ABL80OSM s n 308478. A variance on the tHb of 8.0. Precision run of 24 patient samples between AVOX and another ABL80OSM s n 308477. ABL80OSM s n308477 precision results were good. tHb calibration had not been performed since the installation on October 23, 2013. | Replaced Hemolyzer and associated fluidic tubing. Executed Cleaning Procedure. | CC975FF64600F681 | No | Poor precision for tHb on ABL80OSM s n 308478. A variance on the tHb of 8.0. Precision run of 24 patient samples between AVOX and another ABL80OSM s n 308477. ABL80OSM s n308477 precision results were good. tHb calibration had not been performed since the installation on October 23, 2013. | Replaced Hemolyzer and associated fluidic tubing. Executed Cleaning Procedure.
ABL90 |  |  | CC9793DE4640B0EB | No |  |  | Cleaned | Fluidic Pathway | replaced SP |  | x
ABL90 |  |  | CC9789FC46305AF5 | No |  |  | Defect | HW | Touch screen |  | x
ABL5 | Ausdruck zu schmal, Mechanik Fehler. | 910-206 Thermo Printer erneuert. pO² Kal. Fehler beseitigt. | CC97C818465E0AF2 | No | Expression too narrow, mechanics error. | 910-206 Thermo Printer renewed. pO2 Kal. Fixed a bug.
ABL837 FLEX | Performed a complete oximetry test. Hemolyzer tuning frequency, neon lamp intensity and neon lamp voltage failed. | Replaced the oximetry module PB board. Neon lamp checked okay but still had the tuning frequency error. Replaced the hemolyzer unit. | CC977E3D464357CF | No | Performed a complete oximetry test. Hemolyzer tuning frequency, neon lamp intensity and neon lamp voltage failed. | Replaced the oximetry module PB board. Neon lamp checked okay but still had the tuning frequency error. Replaced the hemolyzer unit. | Defect | HW | Oxi Module | OXI PCB
AQT90 FLEX | Limpieza de la unidad otpitca pruebas del sistema por servicio tecnico, sustitución de la process wheell y de la unidad optica, ajuste de calibración de referencia, test de los diferentes dispositivos, calibración de parametros del analizador y introdución de muestra de prueba, funcionando correctamente. | Cambio de la unidad optica y de la process wheell. | CC99EACE324C200B | No | Cleaning the otpitca unit tests the system by technical service, replacement of the process wheell and the optical unit, reference calibration adjustment, testing of the different devices, calibration of the analyzer parameters and sample introduction of test, working properly. | Change of the optical unit and process wheell.
ABL90 | Error status 7- DLL not found. Unit halts during data up grade | Install 3.0 MR2 software. | CC991673BC48667D | No | Error status 7- DLL not found. Unit halts during data up grade | Install 3.0 MR2 software. | Defect | SW | During SW upgrade | Error status 2 | x
ABL90 | Error status 7. Dll not found. Unit halts during data up grade | Installed 3.0 MR2 | CC991826BC376681 | No | Error status 7. Dll not found. Unit halts during data up grade | Installed 3.0 MR2 | Defect | SW | During SW upgrade | Error status 2 | x
ABL90 | Picked up a Solution pack and a Flow Selector motor from RAME Will install. | Installed Flow Selector stepper motor Installed Solution Pack Flow selector errors have diminished Flushed analyzer free of debris Removed Waste inlet to check for obstructions. Found QC buildup on outside of tubing and manifold. Cleaned free of debris and installed. | CC98AD64BC51EEAD | No | Picked up a Solution pack and a Flow Selector motor from RAME Will install. | Installed Flow Selector stepper motor Installed Solution Pack Flow selector errors have diminished Flushed analyzer free of debris Removed Waste inlet to check for obstructions. Found QC buildup on outside of tubing and manifold. Cleaned free of debris and installed. | Defect | HW | Waste tube |  | x
ABL90 | PCO2 would not cal, Sen was low | I tryed back flushing and still the PCO2 sen was low I wll have them replace the SC. I upgraded the software to 3.0 MR2 | CC98B072BC168451 | No | PCO2 would not cal, Sen was low | I tryed back flushing and still the PCO2 sen was low I wll have them replace the SC. I upgraded the software to 3.0 MR2 | Defect | SC | pCO2 | sens low | x
ABL837 FLEX | Found blockage in air vent line | Flushed out air vent line. Errior is gone | CC98D2A1BC20969C | No | Found blockage in air vent line | Flushed out air vent line. Errior is gone | Cleaned | Fluidic Pathway | Air vent
ABL90 | Einlasssonde verbogen. | Analysator gespült, Einlasssonde aus Kundenbestand erneuert. | CC99CF7E321BAED4 | No | Einlasssonde verbogen. | Analyzer rinsed, inlet probe renewed from customer stock. | Replaced | probe | Bend |  | x
ABL825 FLEX |  |  | CC99CB5D322D93AE | No |  |  | Cleaned | Fluidic Pathway | Unknown
ABL800Basic |  |  | CC99C76D32208A6D | No |  |  | Defect | Gas module | Gas 2 Valve
ABL825 FLEX |  |  | CC9A01EC3202BEC5 | No |  |  | Unknown | Power down | Oxi errors
AQT90 FLEX |  |  | CC9A08DF32251EF8 | No
TCM-COMBIM | Instrument not recognizing Combi module. Tried different module, re-configured unit. Tried different electrode. | Replaced interface board and re-configured analyzer. | CC9A1414320DD62D | No | Instrument not recognizing Combi module. Tried different module, re-configured unit. Tried different electrode. | Replaced interface board and re-configured analyzer. | Defect | HW | PCB
ABL80 CO-OX | Hämolysator Schlauch undicht | Austausch 905-897 | CC9B1DDFA804B9CB | No | Hemolyser hose leaking | Replacement 905-897
ABL90 | Unit has same problem as R0297N004 unit. Pump tubing has pulled loose. |  | CC9B8AB4A8253EA9 | Yes | Unit has same problem as R0297N004 unit. Pump tubing has pulled loose. |  | Re classified RMED | Double upSWO | CC9B7716A82E3EA6 |  | LEH
ABL90 | Gasket change program not correcting problem. Pump noisy when activated. 1343, 1342, 1063, 0508 and other errors present. Pump tubing loose on lot # LI-13 solution pack. | Replaced solution pack. Will drop off solution pack for second unit in AM. | CC9B7716A82E3EA6 | No | Gasket change program not correcting problem. Pump noisy when activated. 1343, 1342, 1063, 0508 and other errors present. Pump tubing loose on lot # LI-13 solution pack. | Replaced solution pack. Will drop off solution pack for second unit in AM. | Defect | SP | Pump tube | loose | x
ABL90 | Pump tubing pulled out of solution pack lot# LI-13. Several errors due to pump tubing failure. | I have a solution pack I can use to replace customers pack. Returned 1 24 to replace solution pack. | CC9B94B3A800289E | No | Pump tubing pulled out of solution pack lot# LI-13. Several errors due to pump tubing failure. | I have a solution pack I can use to replace customers pack. Returned 1 24 to replace solution pack. | Defect | SP | Pump tube | loose | x
ABL90 | The analyzer is presenting error 1169 Unable to pump solutions and it is linked to the Clean inlet error. | Cleaned the inlet gasket and ran the fluid transport programs in service mode, all passing ok. | CC9B67E0A831FAC8 | No | The analyzer is presenting error 1169 Unable to pump solutions and it is linked to the Clean inlet error. | Cleaned the inlet gasket and ran the fluid transport programs in service mode, all passing ok. | Cleaned | inlet gasket |  |  | x
ABL90 | No cHCO3 when pH or pO2 out of reportable range. Patient results show either pH below reportable range or pO2 above reportable range when this occurs. All calibrations and quality controls are fine. Also last Linerarity verification results were normal | Dowloaded data logs (to include patient results) and opening techline case to investigate. | CC9B9532A81E0067 | No | No cHCO3 when pH or pO2 out of reportable range. Patient results show either pH below reportable range or pO2 above reportable range when this occurs. All calibrations and quality controls are fine. Also last Linerarity verification results were normal | Dowloaded data logs (to include patient results) and opening techline case to investigate. | No fault found | HCO3 not reported when pH or pO2 outside reportable range |  |  | x
ABL90 |  |  | CC9B47FDA8497D9B | No |  |  | Replaced | probe | Bend | & maintenance | x
ABL90 | LP (LK-05) defekt | Austausch des LP | CC9AEBB3A839CF8D | No | LP (LK-05) defective | Replacement of the LP | Defect | SP | pump tube |  | x
ABL90 |  |  | CC9B0F5AA8402446 | No |  |  | Cleaned | Fluidic Pathway | inlet gasket + SP replace + probe + gasket + Hemolyzator with Hypochlorit | fluid transport at midnight | x
ABL825 FLEX |  |  | CC9B2156A84C0F2B | No |  |  | Defect | Printer module | Printer module
ABL837 FLEX |  |  | CC9E1EDB9446AD0F | Yes |  |  | Re classified RMED | Pre maintenance
ABL800Basic |  |  | CC9B3AAEA835DC5A | No |  |  | Adjustment | HW | Oxi Module | Lamp unit
ABL820 FLEX | Replaced blue inlet tubing with clear tubing and ran sample- No air bublles detected. Problem seems to be intermittent. | Replaced inlet probe tubing and tube between inlet and PH BG module. | CC9B1AB8A8039B1E | No | Replaced blue inlet tubing with clear tubing and ran sample- No air bublles detected. Problem seems to be intermittent. | Replaced inlet probe tubing and tube between inlet and PH BG module. | Defect | Inlet module | Tubings blue
AQT90 FLEX | Comprobación del estado del analizador. Se entra en servicio y se realizan los test pertinentes,l equipo funciona correctamente. Se introduce una muestra y se observa que la muestra no llega a ser perforada por el equipo, aborta cuando se selecciona pocillo. Se procede a comprobar funcionamiento storage wheel, se observa rozamiento, quesito con punch. Se desmonta punch down y se intenta dejar un poco de espacio se procede a la limpieza de sensores. Se realizan varias muestras así como Qc´s para comprobar funcionamiento del equipo. Funcionamiento correcto. | Error producido por quesito, se intenta dejar mas holgura entre punch down y quesito. En próxima visita se realizará Servicetip 60623210 | CC9B3989A81EC45B | No | Checking the status of the scanner. You enter service and the relevant tests are performed,the equipment works properly. A sample is entered and it is observed that the sample is not pierced by the equipment, aborts when well is selected. It proceeds to check storage wheel operation, friction is observed, quesito con punch. It is disassembled punch down and tries to leave some space proceed to the cleaning of sensors. Several samples are made as well as Qc's to check operation of the equipment. Correct operation. | Error produced by quesito, you try to leave more slack between punch down and quesito. Servicetip 60623210 will be held next visit
TCM400 | Hier werden beide Elekrodenmodule repro-duzierbar nicht erkannt, es liegt eine lose Schraube (Mutter M3 vom Lüfter, 2. Lüfter-schraube ebenfalls lose!) im Gerät! Main CPU PCB, Interface Board und Power Sup-ply als Defektteile ausgeschlossen, Commu-nication PCB als Defektteil erkannt, | Communication PCB erneuert, | CC9AFC3FA8434616 | No | Here both elekroden modules are not recognized repro-coded, there is a loose screw (nut M3 from the fan, 2nd fan screw also loose!) in the device! Main CPU PCB; Interface board and power sup-ply excluded as defective parts, Commu-nication PCB detected as defective part, | Communication PCB renewed, | Defect | HW | PCB
ABL90 | siehe ausgeführte Arbeiten | siehe ausgeführte Arbeiten | CC9C9C831E427E43 | No | see work carried out | see work carried out | Defect | SC | Glu/Lac Sens/status |  | x
ABL825 FLEX | Troubleshot using service program. Gain, Offset circuit adjustment, residual error, Forward reference Position, Backward reference position, and Gasket life all failed. Tried replacing Inlet Gasket - same. | Removed Inlet Module and reseated all connectors. Replaced Inlet Probe. | CC9CCDDC1E2DE437 | No | Troubleshot using service program. Gain, Offset circuit adjustment, residual error, Forward reference Position, Backward reference position, and Gasket life all failed. Tried replacing Inlet Gasket - same. | Removed Inlet Module and reseated all connectors. Replaced Inlet Probe. | Defect | Inlet module | Probe
ABL90 |  | Reinigung | CC9F01B00A24F012 | Yes |  | Cleaning | Re classified RMED | Courtesy |  | CAA | x
ABL90 |  |  | CCBF8FD8065FDFBB | No
ABL80 | Cassette Temperature failure. | Replaced the sensor PCB, performed the Zero calibration and Temperature check. Brougth system to ready. | CCA075818029EFAB | No | Cassette Temperature failure. | Replaced the sensor PCB, performed the Zero calibration and Temperature check. Brougth system to ready.
ABL90 |  | Leihgerät aufgestellt. | CCA02823803B5CD5 | No |  | rental device. | Unknown | Unresolved | Analyzer replaced | Leak detected | x
ABL90 | SP had been changed on 1-7 to lot LI-04 when QC errors occurred. SP was changed again on 1-20 with another lot LI-04 and the same QC errors occurred. | Changed SP to lot LI-26. Calibration, internal QC and external QC are all in control. Instrument operates normally. | CCA05809805938D6 | No | SP had been changed on 1-7 to lot LI-04 when QC errors occurred. SP was changed again on 1-20 with another lot LI-04 and the same QC errors occurred. | Changed SP to lot LI-26. Calibration, internal QC and external QC are all in control. Instrument operates normally. | Adjustment | tHb cal | tHb QC C | Elio contacted | x
ABL90 |  | SC90 ersetzt | CCA41310E21226CE | Yes |  | SC90 replaced | Re classified RMED | Courtesy |  | CAA | x
ABL90 | Limpiar el circuito neumático y cambiar junta de entrada. | Cambiar la junta de entrada. | CCA038C4803D5870 | No | Clean the pneumatic circuit and change inlet joint. | Change the input joint. | Replaced | Inlet gasket |  |  | X
ABL90 | Comprobar el paso de líquidos por el SC, cambiar el SP, calibrar y probar con muestras y controles. | Cambiar el SP. | CCA035DA8024586D | No | Check the passage of liquids through the SC, change the SP, calibrate and test with samples and controls. | Change the SP. | Defect | SP | pO2 QC C |  | X
ABL835 FLEX |  |  | CCA02FCD8060ED8F | No |  |  | User Error | Capillary tube | Broken | Inlet gasket
ABL825 FLEX | Inspected QC logs. found Lactate has just recently had a shift on the high end of the level 3QC only within the last couple days after membrane replacement. Lactate chamber is clotted up. Sealing ring on lactate electrode is deteriorating. No fix. K+ out of calibration range and will not come in. | Cleaned measuring chamber with hypochlorite solution. Remembraned electrode with new lot of membranes and replaced sealing ring. Ran protien removal and cleaning program. Replaced K+ and reference membrane. Suggest reastablishment of ranges. Customer will monitor QC. | CCA01EB380330A4E | No | Inspected QC logs. found Lactate has just recently had a shift on the high end of the level 3QC only within the last couple days after membrane replacement. Lactate chamber is clotted up. Sealing ring on lactate electrode is deteriorating. No fix. K+ out of calibration range and will not come in. | Cleaned measuring chamber with hypochlorite solution. Remembraned electrode with new lot of membranes and replaced sealing ring. Ran protien removal and cleaning program. Replaced K+ and reference membrane. Suggest reastablishment of ranges. Customer will monitor QC. | Defect | Electrode | Electrode | Gasket | Lac
ABL800Basic |  |  | CCA08ADD8036AA1C | No |  |  | Adjustment | Auto Check Module | Breaker arm | loose
ABL80 CO-OX |  | Verriegelung Lösungspack erneuert | CCA1718BF6239DB7 | No |  | Locking solution pack renewed
ABL80 FLEX Basic version | Sustitucion de la bateria Realizacion de shutdown |  | CCA14BB2F6329893 | No | Replacement of the battery Realization of shutdown
ABL820 FLEX | observed air bubbles entering the chloride chamber at end of every rinse cycle. Replaced both rinse and air valve membranes. No fix. | Retubbed analyzer and replace inlet gasket and probe. Customer will monitor as bloods become available. | CCA19969F61AED37 | No | observed air bubbles entering the chloride chamber at end of every rinse cycle. Replaced both rinse and air valve membranes. No fix. | Retubbed analyzer and replace inlet gasket and probe. Customer will monitor as bloods become available. | Defect | Inlet module | Probe
ABL825 FLEX |  |  | CCA14ABEF64AF8FF | No |  |  | Defect | Electrode | Electrode | Reference
TCM-COMBIM |  |  | CCA1AAD0F603502F | No |  |  | Defect | Consumables | Electrode | Temperature
Tosca-500 | Nach 5 Tagen ist die Gas-Anzeige um ca. 8 mm abgesunken (im Messbereit-Modus)! Das Problem mit der Datumseinstellung ist hier nicht aufgetreten. | TechLine-Nr: RDEU 675 909 133 erstellt. | CCA14179F6051CAF | No | After 5 days, the gas indicator has dropped by approx. 8 mm (in ready to measure)! The date setting issue did not occur here. | TechLine No: RDEU 675 909 133. | Defect | Consumables | Gas
ABL80 FLEX Basic version |  |  | CCA30B6E6C3D25D2 | No
ABL80 OSM | Fan is a little louder than normal, but it is not screaming as before according to the techs. Blew air through the fan blades and it rumbled some. Fan should be replaced. | Replaced the cooling fan. | CCA2A3286C030B5C | No | Fan is a little louder than normal, but it is not screaming as before according to the techs. Blew air through the fan blades and it rumbled some. Fan should be replaced. | Replaced the cooling fan.
ABL90 |  |  | CCA299536C19AA80 | No |  |  | Cleaned | Fluidic Pathway | Clot |  | x
ABL90 |  |  | CCA2BE996C30D09C | No |  |  | Defect | SP | Fluid transport |  | x
ABL90 | Performed fluidic checks thru service diagnostics and failed all. Tried different solution pack and sensor cassette with no resolve. Could hear excess groaning from peri pump on both solution pack removal and installation indicating the pump might not be engaging and pumping solutions. | Adjusted peri pump head and passed fluidic checks through service diagnostics. Verified instrument operation. | CCA2E7B86C4FCDD9 | No | Performed fluidic checks thru service diagnostics and failed all. Tried different solution pack and sensor cassette with no resolve. Could hear excess groaning from peri pump on both solution pack removal and installation indicating the pump might not be engaging and pumping solutions. | Adjusted peri pump head and passed fluidic checks through service diagnostics. Verified instrument operation. | Adjustment | Peri pump | Keeps saying replace inlet gasket | LEH > Elio > Log files > MTA | X
ABL80 | Desmontaje de equipo. Revisión de tensiones, contactos, etc. Se prueba con powe supply. Se solicitan piezas para proseguir con la reparación. |  | CCA2DF4A6C20B3AF | No | Equipment disassembly. Review of voltages, contacts, etc. It is tested with powe supply. Parts are requested to continue the repair.
ABL800Basic | Inllet Modul PCB defekt | Inlet Modul auf Grund von Ersatzteilmangel ersetzt | CCA2CC596C421B66 | No | Inllet Module PCB defective | Inlet module replaced due to lack of spare parts | Defect | Inlet module | Inlet PCB
ABL835 FLEX |  |  | CCAB7FCFA63F0281 | No |  |  | User Error | Sample process | Clot
ABL825 FLEX | Returned with lactate electrode. Inspected WDC reports and discovered the electrolyte ranges are fixed and need to be manually adjusted to reestablish the appropriate 2SD range | Replaced lactate electrode. Set up meeting with product specialist to discuss ranges on Friday. Customer will monitor over the next 30 days. | CCA2A2866C24059C | No | Returned with lactate electrode. Inspected WDC reports and discovered the electrolyte ranges are fixed and need to be manually adjusted to reestablish the appropriate 2SD range | Replaced lactate electrode. Set up meeting with product specialist to discuss ranges on Friday. Customer will monitor over the next 30 days. | Defect | Electrode | Electrode | Lac
ABL825 FLEX |  |  | CCA297CA6C1A4E2A | No |  |  | Defect | Inlet module | Probe bend
ABL800Basic |  | s. Arb. | CCA2CD506C0FE55F | No |  | s. Arb. | Defect | Basic unit | Gas module | unknown
AQT90 FLEX |  |  | CCA520BA58204C93 | No
AQT90 FLEX |  |  | CCA2D3676C13BBF2 | No
AQT90 FLEX |  |  | CCA29BE06C3AF6FE | No
AQT90 FLEX | Sustitución de la Tapa completa frontal de analizador, comprobación de funcionamiento general. | Sustitución de la tapa frontal del AQT. | CCA2F6436C59CB4D | No | Replacing the Full Front Analyzer Cover, General Operation Check. | Replacing the front cover of the AQT.
ABL90 |  |  | CCA41939E20C89E2 | No |  |  | Cleaned | Fluidic Pathway | Used detect remove in service |  | x
ABL5 |  |  | CCA42512E230AAD3 | No
ABL815 FLEX |  |  | CCA3DB8BE214DE95 | No |  |  | Defect | Electrode | Electrode | pCO2
ABL825 FLEX | Removed Inlet Gasket and found leak coming from connector for syringe. Flushed tubing and made sure they were not being pinched by AutoCheck Module. Tubing connectors had blood in them. | Removed and replaced Inlet tubing connectors. | CCA415C7E22E4B9E | No | Removed Inlet Gasket and found leak coming from connector for syringe. Flushed tubing and made sure they were not being pinched by AutoCheck Module. Tubing connectors had blood in them. | Removed and replaced Inlet tubing connectors. | Defect | Inlet module | Nipple kit
ABL90 |  |  | CCA54FFF5863823F | No |  |  | Defect | SP | Fluid transport | Guided troubleshooting --> Clean inlet | x
ABL90 | The analyzer was in a inlet clean loop but you could here that the pump tube was not working correctly | I went into service, because i could not by pass the inlet clean, and replaced the SP. I had to replace the SC due to Lac issues. I left them an extra SC because they have been dis satisfied with the reliably of the supplies | CCA53D08580B5D65 | No | The analyzer was in a inlet clean loop but you could here that the pump tube was not working correctly | I went into service, because i could not by pass the inlet clean, and replaced the SP. I had to replace the SC due to Lac issues. I left them an extra SC because they have been dis satisfied with the reliably of the supplies | Defect | SP | MR2 were Installed at date of upSWO occurence | Guided troubleshooting --> Clean inlet | x
ABL90 | Limpieza de contactos de la tarjeta y limpieza del cristal optico, sustitcuión de la tarjeta electronica ajuste de la misma, chequeo por servicio tecnico, cambio de junta de entrada, test controles de calidad y calibración general funcionando corretcamente. | Sustitución de la sensor interface. | CCA51AC5581C4C27 | No | Cleaning of card contacts and cleaning of the optical glass, replacement of the electronic card adjustment of the same, check by technical service, change of input joint, test quality controls and general calibration working correctly. | Replacing the interface sensor. | Defect | HW | SC PCB | LEH spørge om SW version | x
ABL815 FLEX |  |  | CCA57739584873AC | No |  |  | Defect | Basic unit | Display
ABL837 FLEX | Checked the patient and qc logs, confirmed that the send ket is missing. | Checked the LIS HIS communication settings and they are ok. CHecked the analyzer software keys, found the HIS LIS Communication keys were installed as 'Demo keys' and they are now expired. This is causing the 'Send button' not to be present on patient and qc results. | CCA52EFE5820C805 | No | Checked the patient and qc logs, confirmed that the send ket is missing. | Checked the LIS HIS communication settings and they are ok. CHecked the analyzer software keys, found the HIS LIS Communication keys were installed as 'Demo keys' and they are now expired. This is causing the 'Send button' not to be present on patient and qc results. | Setup | HIS/LIS | Demo Key Exp
Tosca-500 |  |  | CCA501B8584667F6 | No |  |  | Defect | HW | PC Module
ABL90 |  |  | CCA91135BA01CBA1 | No |  |  | Defect | HW | SP PCB | + flush and tubes | x
ABL90 |  |  | CCA9133DBA1CBDF8 | No |  |  | Cleaned | Fluidic Pathway | QC C - Parameter unknown |  | x
ABL90 |  |  | CCBFCD75062C7DF1 | No
ABL725 |  |  | CCA8FD0CBA46727C | Yes |  |  | Re classified RMED | User training
ABL827 FLEX | There is a techline case opened for the pCO2 running low for level 1 QC. This analyzer will run fine for a couple days after the pCO2 is membraned then it will start trending down. I wanted to try a pCO2 electrode to see if it would make a difference. RMED is still looking into the issue to determine a fix. This analyzer just seems to be worse than some of my other analyzers. Per techsupport we needed to try a pCO2 electrode. | Replaced the pCO2 electrode. The electrode that is in the analyzer is under warranty and the lot # is 945-612JP02. Will end back to Westlake under warranty. | CCA8D020BA325EDA | No | There is a techline case opened for the pCO2 running low for level 1 QC. This analyzer will run fine for a couple days after the pCO2 is membraned then it will start trending down. I wanted to try a pCO2 electrode to see if it would make a difference. RMED is still looking into the issue to determine a fix. This analyzer just seems to be worse than some of my other analyzers. Per techsupport we needed to try a pCO2 electrode. | Replaced the pCO2 electrode. The electrode that is in the analyzer is under warranty and the lot # is 945-612JP02. Will end back to Westlake under warranty. | Defect | Electrode | Electrode | pCO2
ABL825 FLEX |  | Druckerkopf getauscht. | CCA8FCDABA59D453 | No |  | printer head. | Defect | Printer module | Printer head
ABL90 |  |  | CCAA28D6300F5FAE | No |  |  | Defect | HW | PS |  | x
ABL90 | Inspected SP and SC and Service checks | Re-inserted the SP. Performed Service Checks | CCB4EEA9565993C9 | No | Inspected SP and SC and Service checks | Re-inserted the SP. Performed Service Checks | Adjustment | Reinstall SP | Service programs
ABL80 CO-OX | level 1 and level 4 low results, found manifold crusted | i replaced the manifold ,zeroed instrument, tested unsuccessful po2 low, replaced the sensor cassette | CCAA29F53012F5D1 | No | level 1 and level 4 low results, found manifold crusted | i replaced the manifold ,zeroed instrument, tested unsuccessful po2 low, replaced the sensor cassette
ABL837 FLEX | When loading a carousel with new qc ampoules the autocheck unit is not detecting the correct amount of ampoules. | Replaced the FPC cable. Cleaned the autocheck laser scanner and the laser mirror. The problem is not resolved. Reference membrane is due for replacement, replaced it for the customer. cCrea sensitivity 3 low on calibrations, performed a service Zero calibration. | CCAA4D1330435755 | No | When loading a carousel with new qc ampoules the autocheck unit is not detecting the correct amount of ampoules. | Replaced the FPC cable. Cleaned the autocheck laser scanner and the laser mirror. The problem is not resolved. Reference membrane is due for replacement, replaced it for the customer. cCrea sensitivity 3 low on calibrations, performed a service Zero calibration. | Defect | Auto Check Module | FPC Cable | Before new cable install
ABL837 FLEX | Revisit to reinstall HIS LIS software keys. As of my last service visit, I discovered they were expired because they installed as 'Demo keys'. | Installed the Communication HIS LIS keys. Checked and confirmed that all purchased software keys have been installed as permanent keys. | CCAA24F630116B95 | No | Revisit to reinstall HIS LIS software keys. As of my last service visit, I discovered they were expired because they installed as 'Demo keys'. | Installed the Communication HIS LIS keys. Checked and confirmed that all purchased software keys have been installed as permanent keys. | Setup | Demo key exp | Demo Key Exp | Demo key expired
ABL825 FLEX | Data logs indicate 720 insufficient sample, 688 oxi calculation, 504 nosample in El Met and QC out of range errors. Customer remembraning off schedule fairly often. Both flat membrane valves in measuring modules are stuck in activated position. | Freed and cleaned flat membrane valves. Will order replacements and install ASAP. | CCAA315B300C2997 | No | Data logs indicate 720 insufficient sample, 688 oxi calculation, 504 nosample in El Met and QC out of range errors. Customer remembraning off schedule fairly often. Both flat membrane valves in measuring modules are stuck in activated position. | Freed and cleaned flat membrane valves. Will order replacements and install ASAP. | Defect | El/Met Module | Valve, air | Valve
ABL825 FLEX | Data logs indicate inhommogenous sample, QC out of range, unstable, 474 El Met sample and 719 insufficient sample errors in data logs. | Replaced El Mey module flat membrane valve. Freed and cleaned stuck flat membrane valve in pH BG module.Cleaned pCO2 contacts. | CCAA1C5D301A2993 | No | Data logs indicate inhommogenous sample, QC out of range, unstable, 474 El Met sample and 719 insufficient sample errors in data logs. | Replaced El Mey module flat membrane valve. Freed and cleaned stuck flat membrane valve in pH BG module. Cleaned pCO2 contacts.
ABL835 FLEX |  |  | CCAA49AA30474513 | No |  |  | Adjustment | El/Met Module | Lac Zero
AQT90 FLEX |  | Aus-Einschalten. Staub entfernt. | CCAA1CE430183AE4 | No |  | Off-turn. Dust removed.
AQT90 FLEX | siehe SLR | siehe SLR | CCAA311630316156 | No | see SLR | see SLR
ABL80 | TECHLINE SUGGESTS REPLACING SIGNAL DATA PCB | REPLACED THE SIGNAL DATA PCB AND COMPLETED ZERO CURRENT AND REPLACED LOST SUPPLIES DUE TO REPAIR | CCAB99BBA61DFCF3 | No | TECHLINE SUGGESTS REPLACING SIGNAL DATA PCB | REPLACED THE SIGNAL DATA PCB AND COMPLETED ZERO CURRENT AND REPLACED LOST SUPPLIES DUE TO REPAIR
ABL90 |  |  | CCABAF18A65CEFA3 | No |  |  | Defect | SP | Potential pump tube |  | x
ABL90 |  |  | CCABA499A64566DF | No |  |  | Cleaned | Fluidic Pathway | tHb cal error |  | x
ABL835 FLEX |  |  | CCABCADDA633E6A0 | No |  |  | Defect | Inlet module | Probe bend
ABL820 FLEX | Check system, fluid path and Electrode channel. | Remove the electrode from the chamber and check the electrode. Clean top of electrode with alcohol pad. Check electrode channel according to service manual, result within specifications. Check and adjusted the Zero point in service. Clean measuring chamber with cleaning solution and dry. | CCAB858AA61733EF | No | Check system, fluid path and Electrode channel. | Remove the electrode from the chamber and check the electrode. Clean top of electrode with alcohol pad. Check electrode channel according to service manual, result within specifications. Check and adjusted the Zero point in service. Clean measuring chamber with cleaning solution and dry. | Adjustment | ph/BG Module | pO2 zero | Cleaning
ABL835 FLEX |  |  | CCABE58DA65C8213 | No |  |  | Defect | Electrode | Membrane | Ca++
AQT90 FLEX |  |  | CCAB6208A663F647 | No
ABL80 | Fans on analyzer do sound louder then on the other analyzer. | Replaced fans , test ok. | CCACB0481C2A7785 | No | Fans on analyzer do sound louder then on the other analyzer. | Replaced fans , test ok.
ABL90 | Needle a little bent and gasket with holder had cam apart. | Flushed and replaced gasket and connector gasket lit. | CCAD052F1C3A1667 | No | Needle a little bent and gasket with holder had cam apart. | Flushed and replaced gasket and connector gasket lit. | Defect | HW | Inlet Gasket |  | X
ABL90 | PRUEBAS DE FUNCIONAMIENTO Y TEST DE SERVICIO, DANDO PROBLEMAS GENERALMENTE DE FUGAS Y PRESION. SE VAN CAMBIANDO ACCESORIO Y HACIENDO PRUEBAS NO SIENDO EFECTIVO. DESMONTAJE PARCIAL Y DEL EQUIPO Y CAMBIO DE TUBOS INTERIORES. SE REALIZAN TODOS LOS TEST DE SERVICIO TECNICO SIENDO ESTOS CORRECTOS. CALIBRADO Y VERIFICADO CON Q.C. | TOBOS | CCACB9AB1C283AD0 | No | OPERATION TESTS AND SERVICE TEST, GIVING GENERAL PROBLEMS OF LEAKS AND PRESS. THEY ARE CHANGING ACCESSORY AND MAKING TESTS NOT BEING EFFECTIVE. PARTIAL DISASSEMBLY AND EQUIPMENT AND CHANGE OF INTERIOR TUBES. ALL TECHNICAL SERVICE TESTS ARE PERFORMED BY BEING THESE CORRECT. CALIBRATED AND VERIFIED WITH Q.C. | TOBOS | Unknown | Revisit required | Leak | UpSWO 4 days later | x
ABL837 FLEX |  |  | CCADF96F925BDDE0 | No |  |  | Defect | Inlet module | Probe bend
ABL825 FLEX | Inspected logs and instrument and nothing was found. | No correction needed. Will monitor instrument | CCACBB431C1EC542 | No | Inspected logs and instrument and nothing was found. | No correction needed. Will monitor instrument | Unknown | Electrode | Electrode | pCO2
ABL827 FLEX | I ran the liquid sensor in the service mode and it failed the OXI. I replaced the OXI liquid sensor and it passed. After that the hemolyzer would not turn off. When I took the hemolyzer out it was so hot I could barely grab it. I replaced the hemolyzer just in case the heat damaged it. I also replaced the OXI control PCB and the hemolyzer turned off. | Replaced the OXI Control PCB, hemolyzer and OXI liquid sensor. Ran liquid sensor calibration and complete OXI test. All passed. | CCACD0C81C4D711C | No | I ran the liquid sensor in the service mode and it failed the OXI. I replaced the OXI liquid sensor and it passed. After that the hemolyzer would not turn off. When I took the hemolyzer out it was so hot I could barely grab it. I replaced the hemolyzer just in case the heat damaged it. I also replaced the OXI control PCB and the hemolyzer turned off. | Replaced the OXYGEN Control PCB, Hemolyzer and OXYGEN liquid sensor. Ran Liquid sensor Calibration and complete OXYGEN test. All passed. | Defect | HW | Oxi Module | OXI PCB | Thermostating
ABL825 FLEX | Client changed membranes twice on Ca++ and once on REF. The glue around the base of the PH and REF looks dry routed and the edges of the inside of chamber can not be cleaned. CA++ sens is to high over 100, REF is in range. | Replaced measuring chamber and checked tubing.Replaced REF membrane with client stock. No help ordered REF electrode. | CCAD13B11C3D166A | No | Client changed membranes twice on Ca++ and once on REF. The glue around the base of the PH and REF looks dry routed and the edges of the inside of chamber can not be cleaned. CA++ sens is to high over 100, REF is in range. | Replaced measuring chamber and checked tubing. Replaced REF membrane with client stock. No help ordered REF electrode. | Defect | Electrode | Electrode | Reference
ABL825 FLEX | UPON ARRIVAL PCO2 LEVEL 1 AT 64.4 PULLED SERVICE DUMP,DATA LOGS AND SETUP. | FOUND PH BG RINSE VALVE STUCK BUT NO ERRORS.REPLACED PH BG RINSE AND AIR VALVES REPLACED GAS 1 AND REFERENCE MEMBRANE.CHECKED BAROMETER .PERFORMED CALIBRATIONS.AM ESCALATING THE CALL AND OPENING UP A TECH LINE CASE AS THE ANALYZER PERFORMS OKAY BUT QC DROPS TWO WEEKS LATERLEVEL 1 QC NOW 65.4 | CCACCE001C311CE4 | No | UPON ARRIVAL PCO2 LEVEL 1 AT 64.4 PULLED SERVICE DUMP,DATA LOGS AND SETUP. | FOUND PH BG RINSE VALVE STUCK BUT NO ERRORS. REPLACED PH BG RINSE AND AIR VALVES REPLACED GAS 1 AND REFERENCE MEMBRANE. CHECKED BAROMETER . PERFORMED CALIBRATIONS.AM ESCALATING THE CALL AND OPENING UP A TECH LINE CASE AS THE ANALYZER PERFORMS OKAY BUT QC DROPS TWO WEEKS LATERLEVEL 1 QC NOW 65.4 | Defect | ph/BG Module | Valve | Both
ABL825 FLEX | Both flat membrane valves were stuck in activated position two days after being freed up. | Replaced flat membrane valves in measuring modules. | CCACF68E1C127045 | No | Both flat membrane valves were stuck in activated position two days after being freed up. | Replaced flat membrane valves in measuring modules. | Defect | ph/BG Module | Valve, air | Valve
ABL825 FLEX | Flat membrane valve in pH BG module is stuck in activated position. | Replaced flat membrane valve in pH BG module. | CCACE7941C337041 | No | Flat membrane valve in pH BG module is stuck in activated position. | Replaced flat membrane valve in pH BG module. | Defect | ph/BG Module | Valve, air | Valve
ABL835 FLEX | I worked with Jeff over the phone and determined that the intensity on his lamp is to high. I came on site and verified it was reading (MAX) for value and that is outside the 2900-3700 range. | I adjusted the lamp to ~3250 and it resolved the problem. | CCACD8661C5BC890 | No | I worked with Jeff over the phone and determined that the intensity on his lamp is to high. I came on site and verified it was reading (MAX) for value and that is outside the 2900-3700 range. | I adjusted the lamp to ~3250 and it resolved the problem. | Adjustment | HW | Oxi Module | Lamp unit
ABL90 |  |  | CCAE1A6F9211BF2C | Yes |  |  | Re classified RMED | Telephone support |  |  | x
ABL90 |  |  | CCADFDFE922E1EB5 | No |  |  | Defect | SC | pH impedance |  | x
ABL90 | The PCO2 would not pass cal. The sen would not come in range. | I replaced the SC and ran cal's. The Lac and Glu took a few cal's before comming in. | CCADFAA1921A7E03 | No | The PCO2 would not pass cal. The sen would not come in range. | I replaced the SC and ran cal's. The Lac and Glu took a few cal's before comming in. | Defect | SC | Sens pCO2 |  | x
ABL90 | The PCO2 Sen would not come in it would not go above 56.6% | I replaced the SC and ran set up. the SC i ordered were only 100 test so i gave them two. This analyzer took a bunch of cal's for Lac to come in and then it failed B QC. I replaced the SP at this time and all passed. | CCADF10192617DFF | No | The PCO2 Sen would not come in it would not go above 56.6% | I replaced the SC and ran set up. the SC i ordered were only 100 test so i gave them two. This analyzer took a bunch of cal's for Lac to come in and then it failed B QC. I replaced the SP at this time and all passed. | Defect | SC | Sens pCO2 |  | x
ABL90 | Se reinicia el equipo previa desconexión de cable de red y pendrive, renicio programado y encendido de equipo sin cable de red para descarga, el equipo no reinica de forma normal apareciendo durante 1 segundo un mensaje en letras azules imposible de leer. renicio de equipo de 50 min y chequeo de todas sus partes. Revisión general. | Se reinicia el equipo previa desconexión de cable de red y pendrive, renicio programado y encendido de equipo sin cable de red para descarga | CCAE094E9201920A | No | The previous computer restarts network cable and pendrive disconnection, restarts scheduled restart and power on of computer without network cable for download, the computer does not resound normally by appearing for 1 second a message in blue letters that is not readable. reindenial of 50 min equipment and check of all its parts. General review. | Pre-equipment restarts network cable and pendrive disconnection, restarts scheduled computer and power on computer without network cable for download | Adjustment | Power off/on | take cable out |  | x
ABL835 FLEX |  | Tausch der pCO2-Elektrode. | CCADE91C920C312A | No |  | Replacement of the pCO2 electrode. | Defect | Electrode | Electrode | pCO2
ABL835 FLEX | System experienced tHb drift on 2pt Cal sufficient to cause error on tHb. Initial result was 526, error cal was 495, next cal was back to 427. All cal's since are 426. Customer iconcerned that this is related to an earlier occurrence of Neon Lamp low volts. Also, this system experienced Lac sense decline sufficient to trigger drift errors on 1pt Cal. We have seen this before and it is membrane related. Customer indicated (NY) that they may seek to replace this analyzer. I collected data logs on Feb 5 and found no other occurences of Neon Lamp or tHb errors . I ran Oxi Complete Test - all pass. Follow up phone call yesterday with Kevin Chow - I explained the isolated nature of these errors but he indicated they will seek replacement | No correction, documentation purposes only. | CCAE00E792240C31 | No | System experienced tHb drift on 2pt Cal sufficient to cause error on tHb. Initial result was 526, error cal was 495, next cal was back to 427. All cal's since are 426. Customer iconcerned that this is related to an earlier occurrence of Neon Lamp low volts. Also, this system experienced Lac sense decline sufficient to trigger drift errors on 1pt Cal. We have seen this before and it is membrane related. Customer indicated (NY) that they may seek to replace this analyzer. I collected data logs on Feb 5 and found no other occurences of Neon Lamp or tHb errors . I ran Oxi Complete Test - all pass. Follow up phone call yesterday with Kevin Chow - I explained the isolated nature of these errors but he indicated they will seek replacement | No correction, documentation purposes only. | Unknown | HW | Oxi Module | tHb drift | Neon lamp
ABL825 FLEX | Checked for other errors. Checked Inlet using service program. Tried to calibrate. Failed same tests as on 1-24-14. Noticed a jerky motion on the probe as it was moving forward and back. | Removed Inlet Module and disassembled the motor assembly. Reassembled and turned unit back on. Probe movement was very smooth. | CCAE5EDC925C09CF | No | Checked for other errors. Checked Inlet using service program. Tried to calibrate. Failed same tests as on 1-24-14. Noticed a jerky motion on the probe as it was moving forward and back. | Removed Inlet Module and disassembled the motor assembly. Reassembled and turned unit back on. Probe movement was very smooth. | Adjustment | Inlet module | Motor assembly
ABL835 FLEX | HEMOLYZER NEEDS ADJUSTING AND CLEANING | CLEANED AND RECALIBRATED OXI MODULE. | CCAF337F085011A3 | No | HEMOLYZER NEEDS ADJUSTING AND CLEANING | CLEANED AND RECALIBRATED OXI MODULE. | Cleaned | HW | Oxi Module | Fluid transport
AQT90 FLEX |  |  | CCAF30EF081128E7 | No
ABL90 |  |  | CCB46346E023E9BC | No |  |  | Defect | SC | Lac |  | x
ABL90 |  |  | CCB58F675604E3FE | No |  |  | No fault found | SP not detected | Installed SP - OK | Cleaned fluidic + inlet gasket | x
ABL90 | Checked data logs and this level has been failing since 2 5 14. Tried clot detection program and several calibrations with no positive results. Tried re-installing Sensor Cassette and it still fails. Replaced Sensor Cassette (Lot 99-7) that was installed on 1 27 14. New SC (Lot 0092) is from Customer Stock. | Sensitivity is still below 100% but has climbed steadily from 0.9 to 61.4%. | CCB1CC6FF426E50B | No | Checked data logs and this level has been failing since 2 5 14. Tried clot detection program and several calibrations with no positive results. Tried re-installing Sensor Cassette and it still fails. Replaced Sensor Cassette (Lot 99-7) that was installed on 1 27 14. New SC (Lot 0092) is from Customer Stock. | Sensitivity is still below 100% but has climbed steadily from 0.9 to 61.4%. | Defect | SC | Lac | QC B Low | x
ABL90 |  |  | CCB1AEBCF443AA6C | No |  |  | Replaced | inlet gasket | + pinch valve tube |  | x
ABL90 | TRAS VISITAS ANTERIORES DONDE EL EQUIPO QUEDABA FUNCIONANDO DURANTE UN TIEMPO Y VOLVIENDOSE A REPRODUCIRSE EL MISMO PERFIL DE AVERIA SE SUSTITUYE BAHIA DE SENSOR ELECTRODOS COMPLETA CON SL INCLUIDOS. SE EFECTUAN PRUEBAS DE INNICIALIZACION Y POR SERV. TEC. SIENDO ESTAS CORRECTAS. SE PASAN DIFERENTES MUESTRAS DE SANGRE SIENDO ESTAS CORRECTAS. CALIBRADO Y VERIFICADO CON Q.C. | SUSTITUCION BAHIA | CCB1E365F410C83A | No | AFTER PREVIOUS VISITS WHERE THE EQUIPMENT IS WORKING FOR A TIME AND RETURN TO PLAY THE SAME AVERIA PROFILE IS SUBSTITUTE FULL ELECTRODE SENSOR WITH SL INCLUDED. INICIALIZATION AND SERV TESTS ARE EFFECTED. Tec. BEING THESE CORRECT. DIFFERENT BLOOD SAMPLES ARE PASSED BEING THESE CORRECT. CALIBRATED AND VERIFIED WITH Q.C. | BAHIA SUBSTITUTION | Defect | HW | Sensor Cassette retainer |  | x
ABL825 FLEX | Examined pH BG module. | Replaced meas light PCB on the pH BG module. Verified operation of the ABL825. | CCB1F9F2F40D8AB0 | No | Examined pH BG module. | Replaced meas light PCB on the pH BG module. Verified operation of the ABL825. | Defect | ph/BG Module | Illumination PCB
ABL825 FLEX | Verificação das válvulas, substituição da membrana da válvula do rinse de PH BG, estabilização e passagem do QC |  | CCB22E9FF401EE42 | No | Check valves, PH BG rinse valve membrane replacement, QC stabilization and passage |  | Defect | ph/BG Module | Valve, wet
ABL835 FLEX | PATIENT SAMPLES STILL RESULT INTERMITTENTLY WITH ERROR | REPLACED HEMOLYZER AND COMPLETED ALIGNMENTS | CCB1C5D4F4001967 | No | PATIENT SAMPLES STILL RESULT INTERMITTENTLY WITH ERROR | REPLACED HEMOLYZER AND COMPLETED ALIGNMENTS | Defect | HW | Oxi Module | Hemolyzer
ABL835 FLEX | siehe arbeiten | siehe arbeiten | CCB1EDC0F431A61E | No | see working | see working | Adjustment | HW | Oxi Module | Lamp unit | LEH > why not during install
AQT90 FLEX |  |  | CCB1BF3EF460C85D | No
TCM4 |  |  | CCB1A741F41042C7 | No |  |  | Defect | Consumables | Gas
ABL90 | On arrival on-site checked calibration and QC data logs and found the analyzer finally passed on the lactate parameter. The calibration came in at 0217 hrs this morning and the final level of internal QC came in at 0223 hrs. | Installation on new SC Lot 92 SN 92-86 yesterday at 1226 hrs took 14 hrs before lactate parameter was ready for use. Information passed on to Tech Support. | CCB2EDAD6A32F610 | No | On arrival on-site checked calibration and QC data logs and found the analyzer finally passed on the lactate parameter. The calibration came in at 0217 hrs this morning and the final level of internal QC came in at 0223 hrs. | Installation on new SC Lot 92 SN 92-86 yesterday at 1226 hrs took 14 hrs before lactate parameter was ready for use. Information passed on to Tech Support. | Defect | SC | Lac | slow startup | x
ABL90 | Customer tried multiple scenarios to rectify problem. Verified Customer complaint. Screen had locked customer out and required service password to continue diagnosing problems. Pump calibration detection Cal back log errors, liquid sensor errors, chip solution pack cannot be read. | Multiple fllushingss, cleanings, change out of cusromer sensor cassette and solution pack. Replaced inlet probe and gaskets. ABL DMS PC | CCB548B85605F41E | No | Customer tried multiple scenarios to rectify problem. Verified Customer complaint. Screen had locked customer out and required service password to continue diagnosing problems. Pump calibration detection Cal back log errors, liquid sensor errors, chip solution pack cannot be read. | Multiple fllushingss, cleanings, change out of cusromer sensor cassette and solution pack. Replaced inlet probe and gaskets. ABL DMS PC | Unknown | unclear Root cause | Could be SP, activation
ABL90 | TOOK BRIAN THROUGH ALL SERVICE ADJUSTMENTS AND FOUND LIQ SENSOR CALIBRATION NOT VARING WITH TESTS | DECIDED TO POWER CYCLE THE SYSTEM AND ON POWER UP LIQUID SENSOR TEST PASSED | CCB391026A1FB60D | No | TOOK BRIAN THROUGH ALL SERVICE ADJUSTMENTS AND FOUND LIQ SENSOR CALIBRATION NOT VARING WITH TESTS | DECIDED TO POWER CYCLE THE SYSTEM AND ON POWER UP LIQUID SENSOR TEST PASSED
ABL800Basic | siehe arbeiten | siehe arbeiten | CCB326626A0A9617 | No | see working | see working | Defect | Electrode | Electrode | Reference
ABL825 FLEX | Check operation of the probe, noted the calibration failed. | Replaced the probe and tube to ensure a proper seal. The inlet calibration passed and allowed for calibrations and QC to run. | CCB35CFD6A50D0F3 | No | Check operation of the probe, noted the calibration failed. | Replaced the probe and tube to ensure a proper seal. The inlet calibration passed and allowed for calibrations and QC to run. | Defect | Inlet module | Probe
ABL800Basic |  |  | CCB5AECD564307B0 | No |  |  | Cleaned | ph/BG Module | Measuring chamber
ABL805 FLEX | The analyzer is presenting message 0675 Inlet error, position calculation error. | Upon arrival, I inspected the inlet probe and found bent. replaced a new inlet probe. Flushed system lines. | CCB31CFC6A2D5B79 | No | The analyzer is presenting message 0675 Inlet error, position calculation error. | Upon arrival, I inspected the inlet probe and found bent. replaced a new inlet probe. Flushed system lines. | Defect | Inlet module | Probe bend
ABL80 CO-OX | siehe SLR | siehe SLR | CCB487BEE0460395 | No | see SLR | see SLR
ABL80 FLEX Basic version | Limpieza de circuíto hidráulico. Sustitución de casete de electrodos. Revisión de sistema mecánico y neumático Calibración y verificación de funcionamiento. Correcto según pruebas. |  | CCB41ABEE04CE8D4 | No | Hydraulic fluid cleaning. Replacing electrode cassette. Mechanical and pneumatic system review Calibration and operation verification. Correct according to evidence.
ABL90 |  |  | CCB480ADE0024993 | No |  |  | Defect | HW | SP PCB | chip not read | x
ABL90 | Ran fluidics service diagnostics to verify problem. | Flushed system several times with water using the flush device. Cleaned and adjusted bay home flag. PCO2 and PO2 will not calibrate. | CCB5A9D6560E4727 | No | Ran fluidics service diagnostics to verify problem. | Flushed system several times with water using the flush device. Cleaned and adjusted bay home flag. PCO2 and PO2 will not calibrate. | Cleaned | Fluidic pathway | Clot
ABL90 |  | LP erenuert | CCB47C39E02C8461 | No |  | LP erenuert | Defect | SP | Leak detected |  | x
AQT90 FLEX | Comprobación de process wheel, se ve algo de exceso de liquido en copa vacia, se comprueba que desague bien mediante bombas P1 P2. No obstante se cambia wash unir al tener tubo metálico (al tubo 4) algo doblado. Comprobación unidad óptica en principio bien. Cal TnI , Ok pero con dos atipicos de blancos y valores que se acercan a 700. Se cambia unidad óptica, cal factor referencia de la misma, calibración TnI OK (blancos entre 200-300) calibración resto parámetros (PCR, Dímero D y CKMB) realización blancos | Comprobación de process wheel, se ve algo de exceso de liquido en copa vacia, se comprueba que desague bien mediante bombas P1 P2. No obstante se cambia wash unir al tener tubo metálico (al tubo 4) algo doblado. Comprobación unidad óptica en principio bien. Cal TnI , Ok pero con dos atipicos de blancos y valores que se acercan a 700. Se cambia unidad óptica, | CCB4A6AEE0113F32 | No | Checking process wheel, you see some excess liquid in empty glass, it is checked that it drains well with P1 P2 pumps. However, wash attaches when having a somewhat bent metal tube (to tube 4). Optical drive check in principle well. Cal TnI , Ok but with two atypical whites and values approaching 700. It changes optical unit, cal reference factor of the same, calibration TnI OK (whites between 200-300) calibration rest parameters (PCR, Dim D and CKMB) white realization | Checking process wheel, you see some excess liquid in empty glass, it is checked that it drains well with P1 P2 pumps. However, wash attaches when having a somewhat bent metal tube (to tube 4). Optical drive check in principle well. Cal TnI , Ok but with two atypical whites and values approaching 700. Optical drive is changed,
ABL825 FLEX |  |  | CCB457ABE02F8802 | No |  |  | Defect | Gas module | Valve, Sirai | Right
ABL825 FLEX |  |  | CCB44E6DE034BC21 | No |  |  | Unknown | Module | Valve ?
ABL825 FLEX | Review data for the barcode misreading, no sample codes available. Review the data from the analyzer in Radiance. | Capture data from Radiance for review, noted the QC that is failing has a tighter range than package inserts, analyzer results were within the insert range but outside the customer range. Analyzer performing as expected at the time, did not duplicate any errors. Contact is going to ask the staff to save the suspect barcode. | CCB4D2FCE04BC788 | No | Review data for the barcode misreading, no sample codes available. Review the data from the analyzer in Radiance. | Capture data from Radiance for review, noted the QC that is failing has a tighter range than package inserts, analyzer results were within the insert range but outside the customer range. Analyzer performing as expected at the time, did not duplicate any errors. Contact is going to ask the staff to save the suspect barcode. | Unknown | QC Range
ABL837 FLEX | Performed a complete OXI service test with no errors. Still had the same problem. | Replaced the hemolyzer. | CCB45C57E05B44BA | No | Performed a complete OXI service test with no errors. Still had the same problem. | Replaced the hemolyzer. | Defect | HW | Oxi Module | Hemolyzer | Squeling noise
ABL837 FLEX | Revisit to replace the autocheck carousel spin motor and or the qc ampoule barcode laser scanner. | Replaced the carousel spin motor and, the problem was not resolved. Tried a new qc ampoule barcode laser scanner but it was defective from stock. Reinstalled the original barcode scanner and adjusted the height. Found the instrument presenting error 643 DMS disconnected from AU. replaced the F1 and F4 fuses for the AU board. Glucose is presenting unstable measurement errors on QC Level 1. Cleaned the electrode and the connector. Cleaned the measuring chambers. Glucose is still failing qc level 1. The electrode lot #JR-02 is still under warranty. | CCB45B6BE05E8B6C | No | Revisit to replace the autocheck carousel spin motor and or the qc ampoule barcode laser scanner. | Replaced the carousel spin motor and, the problem was not resolved. Tried a new qc ampoule barcode laser scanner but it was defective from stock. Reinstalled the original barcode scanner and adjusted the height. Found the instrument presenting error 643 DMS disconnected from AU. replaced the F1 and F4 fuses for the AU board. Glucose is presenting unstable measurement errors on QC Level 1. Cleaned the electrode and the connector. Cleaned the measuring chambers. Glucose is still failing qc level 1. The electrode lot #JR-02 is still under warranty. | Adjustment | Auto Check Module | Barcode Reader
ABL825 FLEX |  |  | CCBD52191A1BCC39 | No |  |  | Defect | Electrode | Electrode | Reference
ABL825 FLEX | Removed the AU PCB fuses, cleaned and reseated. Checked voltages on the power supply and observed that the 12V reference is low. | I need to order a power supply. | CCB4BD02E0351885 | No | Removed the AU PCB fuses, cleaned and reseated. Checked voltages on the power supply and observed that the 12V reference is low. | I need to order a power supply. | Defect | PS
ABL800Basic |  |  | CCB49832E04C4A24 | No |  |  | Defect | HW | Oxi Module | Hemolyzer
TCM/Tosca |  |  | CCB42E85E0106A92 | No |  |  | Defect | HW | Mechanics | Breakage after fall
ABL90 |  |  | CCB59EBE5662F148 | Yes |  |  | Re classified RMED | User training |  |  | LEH
ABL90 |  |  | CCB5C0095607F993 | No |  |  | Defect | HW | Hemolyzer | or cleaned fluidic pathway | x
ABL90 | Tried back flush and new SP-No change Inlet gasket with holder not lining up with SP manifold properly | Removed and positioned inlet module Replaced inlet probe, connector and gasket with holder Lactate out on external QC;Replaced SC | CCB593AA561DF6CF | No | Tried back flush and new SP-No change Inlet gasket with holder not lining up with SP manifold properly | Removed and positioned inlet module Replaced inlet probe, connector and gasket with holder Lactate out on external QC; Replaced SC | Replaced | Probe | Clot
ABL825 FLEX | Ran complete oximetry test with no errors. Inlet clear tubing not positioned correctly. | Replaced hemolyzer unit for faults isolation. Adjusted halogen lamp. Ran tHB calibrations in service mode. Replace and reposition clear inlet tubing. | CCB5AC4C56336B5A | No | Ran complete oximetry test with no errors. Inlet clear tubing not positioned correctly. | Replaced hemolyzer unit for faults isolation. Adjusted halogen lamp. Ran tHB calibrations in service mode. Replace and reposition clear inlet tubing. | Defect | Basic unit | Waste tube
ABL825 FLEX | SE LE EFECTUA LIMPIEZA DEL SISTEMA Y SUSTITUCION DE SONDA, ASI COMO ENSEÑAR AL CLIENTE LA FORMA DE REALIZAR EL CAMBIO EN CASO DE NECESIDAD. SE LE DEJA SONDA DE RECAMBIO. CALIBRADO Y VERIFICADO MEDIANTE Q.C. |  | CCB5CD0156344C16 | No | SYSTEM CLEANING AND SOUND REPLACEMENT IS EFFECTED, HOW TO TEACH THE CUSTOMER HOW TO MAKE THE CHANGE IN THE EVENT OF NEED. YOU'RE LEFT A REPLACEMENT PROBE. CALIBRATED AND VERIFIED BY Q.C. |  | Defect | Inlet module | Probe bend
ABL835 FLEX |  |  | CCB5C0A8562FF14D | No |  |  | Defect | Electrode | Membrane | Reference
ABL837 FLEX | PO2 missing sealing ring - still generates the same errors.Replaced inlet tubing. Same problem. Suspect inlet nipples. | replaced sealing ring. Replaced inlet tubing. Replaced inlet nipples. Tested OK. | CCB5778C56077C67 | No | PO2 missing sealing ring - still generates the same errors. Replaced inlet tubing. Same problem. Suspect inlet nipples. | replaced sealing ring. Replaced inlet tubing. Replaced inlet nipples. Tested OK. | Defect | Inlet module | Nipple kit
ABL825 FLEX |  |  | CCB5BEB6563DA08F | No |  |  | User Error | Replacement process | Gas cylinder
AQT90 FLEX |  |  | CCB577B4560B3614 | No
ABL80 FLEX Basic version |  |  | CCB6F3C9CC00A8FF | No
ABL80 FLEX Basic version |  |  | CCB6EE88CC3D8FD9 | No
ABL90 |  |  | CCB6C8AECC43C376 | No |  |  | Defect | SC | Lac - slow startup |  | x
ABL90 | tHb in Lvl. A falsch hoch, verdacht auf Verschmutzung des Hemolyzers. | Analysator mehrfach gespült und mit Proteinentferner gereinigt. Kal. und tHb Kal. durchgeführt, tHb Kal. Fehler aufgetreten, OXI Spectrum mismatch. Reinigung, Kal., und tHb Kal. mit S7770 anderer Lot. wiederholt. tHb in Lvl. A i.O., FO²Hb und FMetHb QC Fehler in Lvl. B und C aufgetreten. Analysator neu gestartet. | CCB6E584CC41E93A | No | tHb in Lvl. A falsely high, suspects contamination of the hemolyzer. | Analyzer rinsed several times and cleaned with protein remover. Kal. and tHb Kal. tHb Kal. Error occurred, OXI Spectrum mismatch. Cleaning, Kal., and tHb Kal. with S7770 other lot. repeated. tHb in Lvl. A i.O., FO2Hb and FMetHb QC errors occurred in Lvl. B and C. Analyzer restarted. | Cleaned | Fluidic pathway |  | + protein remover | x
ABL90 | siehe ausgeführte Arbeiten | siehe ausgeführte Arbeiten | CCBAA7D72E15B840 | Yes | see work carried out | see work carried out | Re classified RMED | Demo install. |  |  | x
ABL835 FLEX |  |  | CCB6D240CC3C2A37 | No |  |  | Defect | Printer module | Printer head
ABL825 FLEX |  |  | CCB6F565CC05D80D | No |  |  | Defect | Electrode | Electrode | pCO2
ABL825 FLEX |  |  | CCB6E481CC0A1C77 | No |  |  | Cleaned | Fluidic pathway | Air vent
ABL800Basic |  |  | CCB6BEA7CC61480F | No |  |  | Defect | El/Met Module | Valve, air | Membrane
ABL800Basic |  |  | CCB6B625CC32DD47 | No |  |  | Defect | Electrode | Membrane | Glu
ABL825 FLEX | Troubleshooting documented on previous work order. The power supply needs to be replaced. | Replaced the power supply. | CCB857EE42548977 | No | Troubleshooting documented on previous work order. The power supply needs to be replaced. | Replaced the power supply. | Defect | Power Supply | -
ABL80 CO-OX |  |  | CCBAB19C2E40E64E | No
ABL90 | On arrival on-site checked NUVC, NUAC and NACP report lay outs. Found that only the NUAC layout was set up incorrectly to print out the COOX results. | Edited the NUAC report to exclude COOX results after talking with Tawanda Burris (NICU Lab Manager). | CCBA98032E02EA3C | No | On arrival on-site checked NUVC, NUAC and NACP report lay outs. Found that only the NUAC layout was set up incorrectly to print out the COOX results. | Edited the NUAC report to exclude COOX results after talking with Tawanda Burris (NICU Lab Manager). | Setup | Report layout
ABL90 | Glu. in Lvl. C falsch niedrig. 841-797 Tubing for valve undicht. | Analysator gespült, 841-797 Tubing for valve erneuert, 946-053 Sensor R0181N053 erneuert. | CCBAA0952E38C9D7 | No | Glu. in Lvl. C wrongly low. 841-797 Tubing for valve leak. | Analyzer rinsed, 841-797 tubing for valve renewed, 946-053 sensor R0181N053 renewed.
ABL90 | The cal would not pass the PCO2 sen | I came to replace the SC and the SP needed changed. The customer didn't have any. There was some ordered and i will return | CCBB071D2E51D55F | No | The cal would not pass the PCO2 sen | I came to replace the SC and the SP needed changed. The customer didn't have any. There was some ordered and i will return | Unknown | revisit required | No SC at customer |  | x
ABL825 FLEX | Comprobación del estado del analizador. Se recoge información proporcionada por el cliente, se comprueba camara de medida, pre-amplificadores y estado de los electrodos. Se procede a sustituir membranas de ref. y Ca. Y se realiza una descontaminación del analizador, tambien se verifica estado de los reactivos del equipo y se procede a programar recambio de membranas. Se procede a realizar varias Cal. y Qc´s para comprobar funcionamiento. | sustitución de membranas y comprobación del estado del analizador. | CCBA952D2E435227 | No | Checking the status of the scanner. Customer-provided information is collected, measuring camera, pre-amps and electrode status are checked. Ref membranes are replaced. and Ca. And a decontamination of the analyzer is performed, the condition of the reagents of the equipment is also checked and membrane replacement is scheduled. Several Cal. and Qc's are performed to check operation. | membrane replacement and analyzer status check. | Defect | Electrode | Membrane | Reference
ABL825 FLEX | Desmontaje de modulo entrada. Revisión de contactos y ajuste mod. entrada. Revisión de sistema mecánico y neumático Calibración y verificación de funcionamiento. Correcto según pruebas. |  | CCBA89622E4395C8 | No | Dismantling of the input module. Contact review and mod setting. Input. Mechanical and pneumatic system review Calibration and operation verification. Correct according to evidence. |  | Adjustment | Inlet module | unknown
ABL825 FLEX |  |  | CCBD54B81A40CC3E | No |  |  | Defect | Electrode | Electrode | Reference
ABL825 FLEX | Checked out Inlet in service program. Test passed however, I also heard the grinding and noticed the erratic movement of the probe. | Removed and replaced the stepping motor for the Inlet. | CCBAF2772E1350DA | No | Checked out Inlet in service program. Test passed however, I also heard the grinding and noticed the erratic movement of the probe. | Removed and replaced the stepping motor for the Inlet. | Defect | Inlet module | Stepper motor
ABL825 FLEX | METHb reading lower when compared to Resp. | Performed lamp adjustment, THb cal service cal | CCBAD24D2E3E13B1 | No | METHb reading lower when compared to Resp. | Performed lamp adjustment, THb cal service cal | Adjustment | HW | Oxi Module | Lamp unit
ABL820 FLEX | Ran bloods between both 820 analyzers. Found newer 820 reading .4 to .5 higher. | Performed lamp adjustment. complete oxi test | CCBAA2BB2E21FA94 | No | Ran bloods between both 820 analyzers. Found newer 820 reading .4 to .5 higher. | Performed lamp adjustment. Complete oxy Test | Adjustment | HW | Oxi Module | Lamp unit
ABL800Basic |  |  | CCBAC6352E2BC58A | No |  |  | Defect | Electrode | Electrode | pCO2
TCM-COMBIM | Flüssigkeit ist in den Touch Screen eingedrun-gen, keine Funktion, | Touch Screen erneuert, kalibriert, | CCBA6B8A2E2B91CC | No | liquid is enclosed in the touch screen, no function, | Touch Screen renewed, calibrated, | Defect | HW | Touch screen
ABL80 CO-OX |  |  | CCBBBC43A43E9465 | No
ABL90 | Realización de la FAN 915-320 rev2 explicación de cambios y entrega de la carta explicativa de dicha FAN, instalación del cartucho de soluciones, explicación de los extractores de sangre y su uso, calibración y test controles de calidad, funcionando correctamente. |  | CCBC0848A425CFAE | No | Realization of the FAN 915-320 rev2 explanation of changes and delivery of the explanatory letter of said FAN, installation of the solution cartridge, explanation of the blood extractors and their use, calibration and test quality controls, working properly.
ABL820 FLEX |  |  | CCBBF278A42642CF | No |  |  | Defect | Electrode | Electrode | pCO2
ABL825 FLEX |  | Tausch Gas-Ventil2. | CCBC019AA4562A09 | No |  | Exchange gas valve2. | Defect | Electrode | Electrode | pCO2
ABL810 FLEX |  |  | CCBBF64AA409BE0E | No |  |  | Defect | Electrode | Electrode | pCO2
TCM-COMBIM | Replaced COMBI module, CF card , guide for optional modules, all did not help | Ordered a 903-100 CPU board.Tested four Combi electrodes on another TCM. Three are bad and one is good | CCBBF57CA43462BF | No | Replaced COMBI module, CF card , guide for optional modules, all did not help | Ordered a 903-100 CPU board. Tested four Combi electrodes on another TCM. Three are bad and one is good | Defect | HW | CF card
TCM-COMBIM |  |  | CCC3428A68090524 | No |  |  | Defect | HW | Power Supply | Sound corrupted
ABL80 CO-OX | Fan on power supply no turning and noise reported. | Replaced power supply. | CCBD44421A48ED92 | No | Fan on power supply no turning and noise reported. | Replaced power supply.
ABL90 | siehe ausgeführte Arbeiten | siehe ausgeführte Arbeiten | CCBD34341A0964C4 | No | see work carried out | see work carried out | Replaced | Inlet gasket | Cleaned |  | x
ABL90 | Error 1324 , 1084, Analyzer saying intervention required. thb cal due. Multiple issues noticed since installation. Data logs and Service dump obtained if techline case needed. | Back flushed the system, Reinserted the SP, Cleaned the Inlet gasket, Cleaned the connection gasket. | CCBD2AC11A271463 | No | Error 1324 , 1084, Analyzer saying intervention required. thb cal due. Multiple issues noticed since installation. Data logs and Service dump obtained if techline case needed. | Back flushed the system, Reinserted the SP, Cleaned the Inlet gasket, Cleaned the connection gasket.
ABL90 | QC failing on the Na and Lac. Cals not accepted as well. | Back Flushed the analyzer, Re-inserted the SC , Conditioning done, Ran wholed blood couple times. | CCBD35201A3D1466 | No | QC failing on the Na and Lac. Cals not accepted as well. | Back Flushed the analyzer, Re-inserted the SC , Conditioning done, Ran wholed blood couple times.
ABL90 | The PCO2 sen would not come up into range | I replaced the SC and all came in | CCBD0AE91A5C1A31 | No | The PCO2 sen would not come up into range | I replaced the SC and all came in | Defect | SC | pCO2 sens |  | x
ABL90 | The PCO2 Sen would not come up into range | I replaced the SC and all came in | CCBD0D0C1A31BCDA | No | The PCO2 Sen would not come up into range | I replaced the SC and all came in | Defect | SC | pCO2 sens |  | x
ABL90 | The PCO2 Sen would not come up into range on cal's | I replaced the SC and all came in | CCBD0E041A4BBCDC | No | The PCO2 Sen would not come up into range on cal's | I replaced the SC and all came in | Defect | SC | pCO2 sens |  | x
ABL90 | Reailización de FAN 915-320 Rev2 y entrega de documentación al cliente. Procedimiento de limpieza de circuíto neumático. Revisión de sistema mecánico y neumático. Calibración y verificación de funcionamiento. Correcto según pruebas. |  | CCBCE7D31A5E95D9 | No | Reailization of FAN 915-320 Rev2 and delivery of documentation to the customer. Pneumatic circus cleaning procedure. Mechanical and pneumatic system revision. Calibration and operation verification. Correct according to evidence. |  | Cleaned | Fluidic pathway |  |  | x
ABL80 CO-OX | the numbers are all low | tested instrument changed cassette and ran vk kit | CCBCEBE01A1DA605 | No | the numbers are all low | tested instrument changed cassette and ran vk kit
ABL800Basic |  |  | CCBD13C41A58F1F8 | No |  |  | Defect | Printer module | Printer head
ABL835 FLEX |  |  | CCBD51841A243003 | No |  |  | Defect | Inlet module | Probe bend
ABL820 FLEX | Found that the QC was above package insert range for PCO2 and that sensitivity was above 100% | Cleaned measuring chamber,repl ref membrane and PCO2 membrane ,cleaned PCO2 electrode and replaced the o ring, replaced gas1 and gas2 , ran numberous 2 poit cals and also rinsed the PCO2 electrode before running the 2 point cals and qc.Pulled all reagents and reinstalled and did a refill. | CCBD431D1A045128 | No | Found that the QC was above package insert range for PCO2 and that sensitivity was above 100% | Cleaned measuring chamber,repl ref membrane and PCO2 membrane ,cleaned PCO2 electrode and replaced the o ring, replaced gas1 and gas2 , ran numberous 2 poit cals and also rinsed the PCO2 electrode before running the 2 point cals and qc. Pulled all reagents and reinstalled and did a refill. | Unknown | pCO2 sensitivity
ABL825 FLEX | calibrations data was off chart the oxi parameters were yellow, i noted the halogen lamp was at 3000. | i back flushed and adjusted lamp ending value 3500 good | CCBD42EA1A17806E | No | calibrations data was off chart the oxi parameters were yellow, i noted the halogen lamp was at 3000. | i back flushed and adjusted lamp ending value 3500 good | Adjustment | HW | Oxi Module | Lamp unit
ABL800Basic |  |  | CCBD428F1A54B365 | No |  |  | Defect | ph/BG Module | Valve, air
AQT90 FLEX |  |  | CCBD482F1A43E97F | No
TCM-COMBIM | Have replaced module and CF card on previous call, Will try new CPU PCB | New CPU PCB does not solve problem. The mother board is on the way from RMED. RAME will ship PCB on Thursday and will install it this Friday | CCBD24211A4CFCBB | No | Have replaced module and CF card on previous call, Will try new CPU PCB | New CPU PCB does not solve problem. The mother board is on the way from RMED. RAME will ship PCB on Thursday and will install it this Friday | Revisit required | Problem re occur | PC Module
TCM400 | Was able to perform two calibrations after which alarm sounded and unit locked up. Re booted unit and unit stayed in the inialization mode. Re booted again when the unit returned to main screen, unit indiciated that the modules were missing. | Returning unit to RAME on RGA WUJT. | CCBD8B861A3DE954 | No | Was able to perform two calibrations after which alarm sounded and unit locked up. Re booted unit and unit stayed in the inialization mode. Re booted again when the unit returned to main screen, unit indiciated that the modules were missing. | Returning unit to RAME on RGA WUJT. | Basic setup | Analyzer | Module missing
ABL90 | On arrival on site found analyzer still in OIR. Service programs reveal that the SP is constantly failing on various solutions. The visual check in the Fluid Transport service program shows fluids flowing with out problem, Possible SC problem. | Problem not corrected due to lack of replacement SC. | CCBE957C9061A254 | No | On arrival on site found analyzer still in OIR. Service programs reveal that the SP is constantly failing on various solutions. The visual check in the Fluid Transport service program shows fluids flowing with out problem, Possible SC problem. | Problem not corrected due to lack of replacement SC. | Unknown | Revisit required | FT - SP check OK | No SC at customer | x
ABL90 |  |  | CCBE68949060E7F7 | No |  |  | Defect | HW | Barcode Reader |  | x
ABL90 |  |  | CCBE20DF902EB431 | No |  |  | Defect | SP | FT
ABL825 FLEX |  |  | CCBE65C5901686FB | No |  |  | Defect | Electrode | Electrode | pCO2
ABL820 FLEX | gender changer is missing. Wrong adapter is being used. Cannot find the original cat5 to DB9 connector. | Will order parts and return. | CCBE45F9905DE09A | Yes | gender changer is missing. Wrong adapter is being used. Cannot find the original cat5 to DB9 connector. | Will order parts and return. | Re classified RMED | Setup
TCM-COMBIM |  |  | CCBE66D2903BDD8D | No |  |  | User Error | Setup | Sensor handling
ABL80 CO-OX | CPU EMI Enclosure Cooling fan out | Cleaned fan but barely turning fan on order for 2 24 2014 and SC. | CCBFD9D0064E0535 | No | CPU EMI Enclosure Cooling fan out | Cleaned fan but barely turning fan on order for 2 24 2014 and SC.
ABL80 CO-OX | On arrival on site verified error codes 292 (turbidity too high) and 581 (Oximeter spectrum mismatch) on last 4 samples run. Visually inspected hemolyzer for connections and tubings problems and none noted. Performed Oximetery test in diagnostic program where the hemolyzer cavitation absorbance failed with a value of -0.35 instead of being in range (5-30). | No correction can be made part must be ordered. Did install remote support option key. Downloaded data logs for CQC case. | CCBF7C4106276ADD | No | On arrival on site verified error codes 292 (turbidity too high) and 581 (Oximeter spectrum mismatch) on last 4 samples run. Visually inspected hemolyzer for connections and tubings problems and none noted. Performed Oximetery test in diagnostic program where the hemolyzer cavitation absorbance failed with a value of -0.35 instead of being in range (5-30). | No correction can be made part must be ordered. Did install remote support option key. Downloaded data logs for CQC case.
ABL90 | Checked unit and did notice a smell around power supply area. Upon power up no smoke was noticed, but unit would not power off and on. Removed unit, will order parts and work on it over weekend. Removed battery pack and unit was allowed to run without incident. | 2 22 Replaced battery pack and unit was allowed to run without incident. Also performed PM | CCC0128306604A81 | No | Checked unit and did notice a smell around power supply area. Upon power up no smoke was noticed, but unit would not power off and on. Removed unit, will order parts and work on it over weekend. Removed battery pack and unit was allowed to run without incident. | 2 22 Replaced battery pack and unit was allowed to run without incident. Also performed PM
ABL90 | Arrived on-site to find analyzer in same OIR screen with Clean Inlet Gasket message. Waited for new SC to be delivered. Verified same error messages 1084-Inhomogeneous cal3 solution, 1189-Unable to pump solutions and 1343-Unable to pump solutions are still present. | Replaced SC lot R0103 with Lot R0113 (customer stock). | CCBFC11B064E0901 | No | Arrived on-site to find analyzer in same OIR screen with Clean Inlet Gasket message. Waited for new SC to be delivered. Verified same error messages 1084-Inhomogeneous cal3 solution, 1189-Unable to pump solutions and 1343-Unable to pump solutions are still present. | Replaced SC lot R0103 with Lot R0113 (customer stock). | Defect | SC | FT problems |  | x
ABL90 | The screen went black all of a sudden and rebooting the system did not help as well.When I rebooted the system the screen flickered couple times and then the analyzer went quiet. Checked the Power supply, Checked the battery module. Checked connections. | Replaced the battery module. After installation of the battery module the battery sign did not show up. In the service mode the Battery option showed Roma.ini battery check not enabled. Went into D:ODIN Roma.INI and changed the value under feature settings for battery check from 0 to1. Charge sign came up and battery module active. | CCBF9F410661146F | No | The screen went black all of a sudden and rebooting the system did not help as well. When I rebooted the system the screen flickered couple times and then the analyzer went quiet. Checked the Power supply, Checked the battery module. Checked connections. | Replaced the battery module. After installation of the battery module the battery sign did not show up. In the service mode the Battery option showed Roma.ini battery check not enabled. Went into D:ODIN Roma.INI and changed the value under feature settings for battery check from 0 to1. Charge sign came up and battery module active.
ABL825 FLEX |  |  | CCBFC048062EC9E0 | No |  |  | Cleaned | Fluidic Module | Air vent
ABL825 FLEX |  |  | CCBFF252061A6700 | No |  |  | Instruction | QC Range | Process
ABL835 FLEX |  |  | CCBF896206409746 | No |  |  | Defect | Electrode | Membrane | pCO2
ABL825 FLEX | Troubleshot. Looked at data logs. Noticed downward trend on pCO2. Tubing between Inlet and pH BG Modules appears to have a loose seal on connector. | Removed pCO2 electrode and cleaned. Replaced tube between Inlet and pH BG Modules. | CCBFB76A06000FD1 | No | Troubleshot. Looked at data logs. Noticed downward trend on pCO2. Tubing between Inlet and pH BG Modules appears to have a loose seal on connector. | Removed pCO2 electrode and cleaned. Replaced tube between Inlet and pH BG Modules. | Defect | Basic unit | Pre heater tube
AQT90 FLEX |  |  | CCBFC92106021784 | No
AQT90 FLEX |  |  | CCBF961A0615D2D7 | No
AQT90 FLEX |  |  | CCBF8E8306513FE2 | No
ABL90 |  |  | CCC1472A7C35B8D4 | Yes |  |  | Re classified RMED | Logistic upSWO | and user training |  | x
ABL80 CO-OX | EMI fan barly turning and stopping, computer overheated | replaced computer assy and EMI fan | CCC3B07F682E5B41 | No | EMI fan barly turning and stopping, computer overheated | replaced computer assy and EMI fan
ABL80 OSM | ASTM file to be modified. | Made necesary changes to ASTM file for Comm. | CCCEFE4E8E0BF216 | No | ASTM file to be modified. | Made necesary changes to ASTM file for Comm.
ABL80 CO-OX | On arrival on site verified that the hemolyzer was still failing cavitation absorbance (value of -0.35 with a range of 5-30) during the COOX test. | Replaced hemolyzer assembly. Performed COOX test and value is now 17.91 with a range of 5-30. | CCC36A38680402CB | No | On arrival on site verified that the hemolyzer was still failing cavitation absorbance (value of -0.35 with a range of 5-30) during the COOX test. | Replaced hemolyzer assembly. Performed COOX test and value is now 17.91 with a range of 5-30.
ABL90 |  |  | CCC34640682A7F88 | No |  |  | Defect | SP | Leak current |  | x
ABL90 | Comprobar junta de entrada, tubos y SP, cambiar la junta de entrada. Limpiar el circuito neumatico, quitar el polvo del interior, comprobar el registro de errores, guardar y borrar datos, calibrar y probar con controles. | Cambiar la junta de entrada de muestras. | CCC38CEC6826B224 | No | Check inlet joint, tubes and SP, change the inlet joint. Clean the tire circuit, remove dust from the inside, check the error log, save and erase data, calibrate and test with controls. | Change the sample input joint. | Replaced | Inlet Gasket |  |  | x
ABL90 | REALIZAR PROCESO DE LIMPIEZA POR OBTRUCCION EN CIRCUITO HIDRAULICO. PRUEBAS DE FUNCIONAMIENTO. CALIBRADO Y VERIFICADO CON Q.C. | LIMPIEZA | CCC3AF0D6832E1A3 | No | PERFORM CLEANING PROCESS BY HYDRAULIC CIRCUIT OBSTRUCTION. OPERATION TESTS. CALIBRATED AND VERIFIED WITH Q.C. | Cleaning | Cleaned | Fluidic pathway | Clot |  | x
ABL90 | D |  | CCC3A7E168542581 | No | D |  | Unknown | Insufficient information |  | Alex email:is an error or an incomplete SWO. TBC. | x
AQT90 FLEX | Se libera cartucho de TNI con una de las pestañas rotas ausentes que impedia el movimiento de la storagewheel. Finalmente se comprobó correcto funcionamiento del equipo mediante QC. |  | CCD0FF5A7A42AEE7 | No | TNI cartridge is released with one of the missing broken lashes that prevented the movement of the storagewheel. The equipment was finally found to be functioning correctly using QC.
ABL835 FLEX |  |  | CCC369EE68059914 | No |  |  | Unknown | QC process
ABL825 FLEX | checked gaskets and valves | Replaced waste tubing. | CCE5D04FDA38F89F | No | checked gaskets and valves | Replaced waste tubing. | Defect | Fluidic Module | Waste pump tube
ABL820 FLEX | Keith had already been on site and wanted me to change the parts that were ordered. there were also oxi issues that keith diagnosed and he wanted me to replace the hemolyzer. | I replaced the parts and performed a carousel alingment, and I set the Lamp to ~3250 | CCCC5904A20DB856 | No | Keith had already been on site and wanted me to change the parts that were ordered. there were also oxi issues that keith diagnosed and he wanted me to replace the hemolyzer. | I replaced the parts and performed a carousel alingment, and I set the Lamp to ~3250 | Adjustment | Auto Check Module | Alignment | Carroussel
ABL825 FLEX | Sample not detected or 328 no leading air segment on patients. | Flushed sample path lines and air vent. Checked valve and pump operation. OK. | CCC3C77A685ADF54 | No | Sample not detected or 328 no leading air segment on patients. | Flushed sample path lines and air vent. Checked valve and pump operation. OK. | Cleaned | Fluidic pathway | Waste line
ABL837 FLEX | Performed complete oximetry test. Lamp unit failed. Checked lamp unit status. Lamp intensity too High. Chloride out on quality control. | Performed lamp adjustment. Replaced chloride membrane. | CCC33216685F3F05 | No | Performed complete oximetry test. Lamp unit failed. Checked lamp unit status. Lamp intensity too High. Chloride out on quality control. | Performed lamp adjustment. Replaced chloride membrane. | Adjustment | HW | Oxi Module | Lamp unit
TCM-COMBIM |  |  | CCC3752468166E27 | No |  |  | Defect | Electrode | Temperature
ABL90 |  |  | CCC49824DE3A75E3 | No |  |  | Cleaned | Inlet gasket |  |  | x
ABL90 | Level1 and Level 2 QC indicated no po2. Had to reboot the analyzer due to screen freeze. | Performed manual back flush. Analyzer took 50 mins to warm up after back flush. | CCC7AE56CA42ED43 | No | Level1 and Level 2 QC indicated no po2. Had to reboot the analyzer due to screen freeze. | Performed manual back flush. Analyzer took 50 mins to warm up after back flush. | Cleaned | Fluidic pathway |  |  | x
ABL800Basic |  |  | CCCC2C61A24450E4 | No |  |  | Defect | Inlet module | Probe
ABL825 FLEX |  |  | CCC5C7985457EBAD | No |  |  | Defect | Gas module | Flow regulator, right | and left
ABL825 FLEX |  |  | CCC48FF1DE4BC8AC | No |  |  | Defect | Electrode | Membrane | Reference
ABL810 FLEX |  |  | CCC4DD14DE5DB097 | No |  |  | Defect | Electrode | Electrode | pCO2
AQT90 FLEX | Sustitución y ajuste de la puerta de entrada de los quesitos hacia la storage wheel, chequeo de funcionamiento. | Sustitución de la entrada de storage wheel | CCC4D8F2DE6047B6 | Yes | Replacing and adjusting the door of the containers towards the storage wheel, operation check. | Replacing the storage wheel input | Re classified RMED | Analzer well functioning since last upSWO
ABL90 | Until this week, analyzer has been without oxi error since last visit (1 8 2014). Performed complete oxi test. No problems found. During investigation it was found that the same sample questioned for oximetry on analyzer, was measured accurately on other analyzer. Manual suggests to replace the spectrometer and or wet section PCB. After replacement, The sensor cassette had impeadance and back log errors that would not clear. Replaced sensor cassette. Solution pack ran out of activities. Replaced solution pack. | Replaced pinch valve and spectrometer. Had to replace sensor cassette and solution pack after repair. | CCC5CA36545512C7 | No | Until this week, analyzer has been without oxi error since last visit (1 8 2014). Performed complete oxi test. No problems found. During investigation it was found that the same sample questioned for oximetry on analyzer, was measured accurately on other analyzer. Manual suggests to replace the spectrometer and or wet section PCB. After replacement, The sensor cassette had impeadance and back log errors that would not clear. Replaced sensor cassette. Solution pack ran out of activities. Replaced solution pack. | Replaced pinch valve and spectrometer. Had to replace sensor cassette and solution pack after repair. | Defect | HW | Spectrometer | tHb too high | x
ABL90 |  |  | CCC620DD5400E1F9 | No |  |  | Cleaned | Fluidic pathway | Hypo cl. |  | x
ABL90 |  |  | CCC6057C5416BEA5 | No |  |  | Defect | SP | Leak |  | x
ABL815 FLEX |  |  | CCC73F11CA5DD8B4 | No |  |  | Setup | Radiance
ABL800Basic |  | Tausch Gas Ventil2 + pH BG Modul. | CCC5DD56540AF3D5 | No |  | Exchange gas valve2 + pH BG module. | Defect | Electrode | Electrode | pCO2
AQT90 FLEX |  |  | CCC60CAF5462CFC5 | No
ABL90 |  |  | CCC74D73CA047477 | No |  |  | Adjustment | SC Door |  |  | x
ABL90 |  |  | CCC74CCDCA3981AD | No |  |  | User Error | SP expired |  |  | x
ABL90 |  |  | CCC779E1CA5F0D4B | No |  |  | Unknown | Revisit required | CPU | See CCD12E5E7A598395 | x
ABL90 |  |  | CCC74488CA18E212 | No |  |  | Defect | SC | pO2 |  | x
ABL90 | Testes de fluidica no modulo de serviço técnico, verificação de erro de QC1 e 2, limpeza do circuito, substituição do tubo de válvula, tubo de entrada e sonda de entrada. Colocação de um novo SP e verificação do funcionamento sem problemas. Tudo OK. |  | CCC7571BCA106FFF | No | Fluid testing in the Technical Service module, QC1 and 2 error checking, circuit cleaning, valve tube replacement, inlet tube and inlet probe. Placement of a new SP and verification of operation smoothly. All OK. |  | Replaced | Inlet gasket | - |  | x
ABL825 FLEX | siehe ausgeführte Arbeiten | siehe ausgeführte Arbeiten | CCC76ADCCA5AF2E7 | No | see work carried out | see work carried out | Defect | Electrode | Electrode | pCO2
ABL820 FLEX | I looked through the data logs and it only missed a single QC and it said that there was a calibration error. | I told the customer to keep track but after talking to David in Tech support we agree that there was not a problem. Customer will keep an eye on it. | CCCC56CAA217B852 | Yes | I looked through the data logs and it only missed a single QC and it said that there was a calibration error. | I told the customer to keep track but after talking to David in Tech support we agree that there was not a problem. Customer will keep an eye on it. | Re classified RMED | Courtesy
ABL825 FLEX | Checked for other error messages. Checked all tubing. Customer has membraned numerous times and even installed a new electrode. While trying to remove electrodes, found electrode connector for pH very hard to open. | Replaced measuring chamber and electrode connector for pH electrode. Removed Rinse and Gas Valves for pH BG Module and cleaned. | CCC72303CA19ED40 | No | Checked for other error messages. Checked all tubing. Customer has membraned numerous times and even installed a new electrode. While trying to remove electrodes, found electrode connector for pH very hard to open. | Replaced measuring chamber and electrode connector for pH electrode. Removed Rinse and Gas Valves for pH BG Module and cleaned. | Defect | Electrode | Electrode | pCO2
TCM/Tosca |  |  | CCC76086CA21AF59 | No |  |  | Defect | Hardware | PC Module
ABL80 FLEX Basic version |  |  | CCC862DC40369D1A | No
ABL90 |  |  | CCC87EDC4004CCC8 | No |  |  | Cleaned | Fluidic pathway | Hypo cl. |  | x
ABL837 FLEX | Checked activity logs and this has been occuring for over a month sometimes multiple times per day. Inspected cables. no damage evident. | Replaced Power supply. Customer will monitor. | CCC866EF400453D5 | No | Checked activity logs and this has been occuring for over a month sometimes multiple times per day. Inspected cables. no damage evident. | Replaced Power supply. Customer will monitor. | Defect | Power Supply
ABL800Basic |  |  | CCC8896340027672 | No |  |  | Defect | El/Met Module | Valve, air
AQT90 FLEX |  |  | CCC898054052D153 | No
TCM/Tosca |  |  | CCC86089400FAF5C | No |  |  | Defect | Hardware | Communication PCB
TCM400 | 2-27-14 I worked with Susan over the phone. I asked her to swap the #1 electrode with the #2 electrode and it still did not calibrate the electrode in the #1 slot. She swapped the #1 module with the #2 module and the electrode in the #1 slot is now calibrating. The #2 slot is now not calibrating. It followed the module. 2-28-14 Julie said they put the new module in it all electrodes are now calibrating. She will mail the bad electrode back to me so I can turn it in for warranty. | Shipped customer an O2 module and Susan installed it. All is working now. | CCC897FF403E7308 | No | 2-27-14 I worked with Susan over the phone. I asked her to swap the #1 electrode with the #2 electrode and it still did not calibrate the electrode in the #1 slot. She swapped the #1 module with the #2 module and the electrode in the #1 slot is now calibrating. The #2 slot is now not calibrating. It followed the module. 2-28-14 Julie said they put the new module in it all electrodes are now calibrating. She will mail the bad electrode back to me so I can turn it in for warranty. | Shipped customer an O2 module and Susan installed it. All is working now. | Defect | Hardware | TC Module
TCM-COMBIM |  |  | CCC8AAA140268F04 | No |  |  | Defect | Hardware | Display
ABL825 FLEX | Checked data logs for other errors. Verified the unstable errors. Found 2 O rings on reference electrode. | Removed electrodes from pH BG Module and cleaned chamber with cleaning solution. Membraned reference electrode. Will monitor. | CCCA14D2B6298973 | No | Checked data logs for other errors. Verified the unstable errors. Found 2 O rings on reference electrode. | Removed electrodes from pH BG Module and cleaned chamber with cleaning solution. Membraned reference electrode. Will monitor. | User Error | Replacement process | O ring left in chamber | Reference
ABL825 FLEX | Analyzer still acts like the power supply needs to be replaced. It was replaced on the last visit, but I am going to replace it again. | Replaced the power supply. | CCC9D62CB66001C6 | No | Analyzer still acts like the power supply needs to be replaced. It was replaced on the last visit, but I am going to replace it again. | Replaced the power supply. | Unknown | PS | Replaced | Temp errors
ABL825 FLEX |  |  | CCC9893AB6498701 | No |  |  | Defect | HW | Oxi Module | OXI PCB
ABL90 | Kunde hat in den letzten zwei Tagenb jeweils ein LP Lot LM-09 eingesetzt->fehler blieb | Austausch des LP gegen anders Lot SC defekt durch lange Standphase-<SC ersetzt | CCCAEECF2C5AF9C2 | No | Customer has used one LP Lot LM-09 in the last two days each->error remained | Replacement of the LP with other Lot SC defective replaced by long stand-up phase-SC | Defect | SP | FT |  | x
ABL90 | After running a pleural sample, a clot was introduced to the analyzer. Used clear tubing and flush device to track down the clot. Clot was found in inlet probe. | Removed inlet probe and flushed thoroughly to remove clot. | CCCC625CA2313D4C | No | After running a pleural sample, a clot was introduced to the analyzer. Used clear tubing and flush device to track down the clot. Clot was found in inlet probe. | Removed inlet probe and flushed thoroughly to remove clot. | Cleaned | Fluidic pathway | Clot |  | x
ABL90 |  |  | CCCC7EBFA23BAFDB | No |  |  | Cleaned | Fluidic pathway | LS error |  | x
ABL825 FLEX | Examined OXi Module. Checked data logs. | Replaced Hemolyzer. Performed lamp Adj. Did complete Oxi Test. Did service tHb cal. Verified operation of ABL825. | CCCC4020A247E4E6 | No | Examined OXi Module. Checked data logs. | Replaced Hemolyzer. Performed lamp Adj. Did complete Oxi Test. Did service tHb cal. Verified operation of ABL825. | Defect | HW | Oxi Module | Hemolyzer | MetHb high
ABL835 FLEX |  |  | CCCC1610A240AF79 | No |  |  | Adjustment | HW | Oxi Module | tHb | Fcuv(,9639)
ABL837 FLEX | Reviewed los, sensitivity on cals are low. | Remembraned Glucose. Recommended to store the membranes in refig. instead of at room temp. | CCCC3550A22794D5 | No | Reviewed los, sensitivity on cals are low. | Remembraned Glucose. Recommended to store the membranes in refig. instead of at room temp. | User Error | Membrane storage | - | Glu
ABL835 FLEX | Intermittent measurement unstable errors on Patient samples and QC. Customer has printouts from Liquid Sensor cal that shows an interesting phenomenon on this analyzer and several other analyzers too. Also spent time reviewing some old membranes that appear to have damage to membrane and did customer training. | Currently the system is passing all cal's and QC. Customer will change Lac membrane tonight. | CCCC5A95A2174391 | Yes | Intermittent measurement unstable errors on Patient samples and QC. Customer has printouts from Liquid Sensor cal that shows an interesting phenomenon on this analyzer and several other analyzers too. Also spent time reviewing some old membranes that appear to have damage to membrane and did customer training. | Currently the system is passing all cal's and QC. Customer will change Lac membrane tonight. | Re classified RMED | Customer training
ABL835 FLEX | COllected all data logs for Techline. Customer noted that when a cal or QC fails, if he goes into service LS program then the Upper Met LS reports empty. This happens on several systems. Also performed user training and reviewed data logs with customer. One suspicion is that silicone is getting inot fluidics but this system run very few patient samples. | No correction, system is currently operating within spec. Will monitor performance while I submit data to Tachline for review | CCCC71D1A20B4392 | No | COllected all data logs for Techline. Customer noted that when a cal or QC fails, if he goes into service LS program then the Upper Met LS reports empty. This happens on several systems. Also performed user training and reviewed data logs with customer. One suspicion is that silicone is getting inot fluidics but this system run very few patient samples. | No correction, system is currently operating within spec. Will monitor performance while I submit data to Tachline for review | Unknown | Electrode | Electrode | GLu
ABL825 FLEX | Comprobación del estado de funcionamiento del equipo, test completo del módulo Oxi por servicio, se ajusta el rango de la lámpara hallógena por estar éste fuera de rango, se calibra el equipo en varias ocasiones y verificamos el correcto funcionamiento con cal de la tHb por servicio-usuario y QC. |  | CCCC7D9FA234226F | No | Checking the operating status of the equipment, full test of the Oxi module per service, adjusting the range of the hallogenic lamp because it is out of range, the equipment is calibrated several times and we verify the correct operation with lime of the tHb by service-user and QC. |  | Adjustment | HW | Oxi Module | Lamp unit
ABL80 CO-OX |  |  | CCCD8E3118588CD7 | No
ABL80 CO-OX | error code 581 on intermittent samples | Performed cleaning program - service. | CCCDB9CE1851606F | No | error code 581 on intermittent samples | Performed cleaning program - service.
ABL80 CO-OX |  |  | CCCD4BF31823F3A9 | No
ABL90 |  |  | CCCD6F5C185ECEBB | No |  |  | Replaced | Inlet Gasket |  |  | x
ABL90 |  |  | CCCD7EE91832B027 | No |  |  | Replaced | Probe | Bend |  | x
ABL90 |  |  | CCCD8B8D18028321 | No |  |  | Setup | QC Limits | B - C |  | x
ABL90 | instrument does not pass calibrations, the solution pack fails in service test under qc1. | i replaced the solution pack | CCCE7CB28E040A87 | No | instrument does not pass calibrations, the solution pack fails in service test under qc1. | i replaced the solution pack | Defect | SP | QC1 |  | x
ABL810 FLEX |  |  | CCCD64AC185970BA | No |  |  | Defect | Electrode | Electrode | pCO2
ABL825 FLEX | Unit was previously aborting samples and QC but is not doing that now. No clear reason for variations in QC results which are less dramatic than what was seen recently. | Checked valves and pumps again. OK. Remembraned reference electrode to correct possible carry over problems. Cleaned reference and Ca electrode contacts. | CCCD847C181DDB40 | No | Unit was previously aborting samples and QC but is not doing that now. No clear reason for variations in QC results which are less dramatic than what was seen recently. | Checked valves and pumps again. OK. Remembraned reference electrode to correct possible carry over problems. Cleaned reference and Ca electrode contacts. | Defect | Electrode | Membrane | Reference
ABL825 FLEX | The most recent QC's run had spiked upward. Before that they were OK. | Checked valves and pumps. OK. Flushed inlet lines. Remembraned reference electrode | CCCD9E3D1806DB43 | No | The most recent QC's run had spiked upward. Before that they were OK. | Checked valves and pumps. OK. Flushed inlet lines. Remembraned reference electrode | Defect | Electrode | Membrane | Reference
ABL820 FLEX | This customer is trying to connect through VAMC in Gainesville. We got the settings from Gainesville but they are still unable to connect. Working with Larry Logar in Tech Support we tried several different combinations of settings but we still had no luck. I talked with Joedy Smith at VAMC Gainesville and found that they were going through a Radiance server for QC and Data Log transmission. Larry pointed out that the VA in Jacksonville did not have a Radiance License on file, but Joedy said that one was purchased at the time the analyzer was installed. I looked up the PO and there was a Raidenance License purchased. Larry advised to get in touch with Derrick Barkalow to ensure all installation steps were completed from the Raidance side of the house. | Sent an email to Derrick Barkalow (Radiometer SIT) and informed him of the situation. Further action is needed when I receive his response. | CCCDA492180BFAD7 | Yes | This customer is trying to connect through VAMC in Gainesville. We got the settings from Gainesville but they are still unable to connect. Working with Larry Logar in Tech Support we tried several different combinations of settings but we still had no luck. I talked with Joedy Smith at VAMC Gainesville and found that they were going through a Radiance server for QC and Data Log transmission. Larry pointed out that the VA in Jacksonville did not have a Radiance License on file, but Joedy said that one was purchased at the time the analyzer was installed. I looked up the PO and there was a Raidenance License purchased. Larry advised to get in touch with Derrick Barkalow to ensure all installation steps were completed from the Raidance side of the house. | Sent an email to Derrick Barkalow (Radiometer SIT) and informed him of the situation. Further action is needed when I receive his response. | Re classified RMED | Logistic upSWO | Radiance connection | Key
ABL825 FLEX | NA | Skanner skiftet | CCCD9C2C1861452F | No | In | Scanning the switch | Defect | Basic unit | Barcode Reader
ABL80 CO-OX | Se localiza equipo con obstrucción interna. Se desmonta modulo fluídica y se limpia sistema. Revisión de funcionamiento y limpieza de manifold unit. Revisión de sistema mecánico y neumático Calibración y verificación de funcionamiento. Correcto según pruebas. |  | CCCEA6DF8E4EBA80 | No | Equipment with internal obstruction is located. It disassembles fluid module and cleans system. Operation review and cleaning of manifold unit. Mechanical and pneumatic system review Calibration and operation verification. Correct according to evidence.
ABL90 |  |  | CCCEA6CB8E603BE3 | No |  |  | Cleaned | Fluidic pathway | Hypo cl. |  | x
ABL825 FLEX | Temp errors from multiple modules lead me to believe that it could be the Analyzer Control PCB. | Replaced the Analyzer Control PCB. | CCCF0B528E610629 | No | Temp errors from multiple modules lead me to believe that it could be the Analyzer Control PCB. | Replaced the Analyzer Control PCB. | Defect | Basic unit | Analyzer Control PCB | Temp errors
ABL800Basic |  |  | CCCEA0618E5E1F4E | No |  |  | Defect | HW | Oxi Module | OXI PCB | Missing trigger
ABL835 FLEX |  |  | CCCEE2748E589976 | No |  |  | Defect | Electrode | Electrode | pO2
TCM-COMBIM | Pco2 Module not connecting to the Computer Mike had tried the Connector Computer PCB, but problem persisted. | Replaced the Communication PCB. Brought system to ready. Module connected properly. | CCCEC5208E3418FA | No | Pco2 Module not connecting to the Computer Mike had tried the Connector Computer PCB, but problem persisted. | Replaced the Communication PCB. Brought system to ready. Module connected properly. | Defect | Hardware | Communication PCB
ABL90 | client replaced 3 SC and 2 SP no help Sens too low. I tried SC different lot still failed.Ran | SC and SP on order for 3 7 2014 | CCD046EC04385FA3 | No | client replaced 3 SC and 2 SP no help Sens too low. I tried SC different lot still failed. Ran | SC and SP on order for 3 7 2014 | Unknown | Revisit required | SC pCO2 sens |  | x
ABL90 | Sometimes samples are being run before the labels are printed and the analyzer locks out user from manually entering the accession number. | Enabled the confirm patient demographics option. and advised that if possible operators wait until labels are printed prior to running samples. | CCCFDD560402C565 | No | Sometimes samples are being run before the labels are printed and the analyzer locks out user from manually entering the accession number. | Enabled the confirm patient demographics option. and advised that if possible operators wait until labels are printed prior to running samples. | User Error | Sample process |  |  | x
ABL90 |  |  | CCCFEA500452FAA4 | No |  |  | Defect | HW | Scanner |  | x
ABL90 |  |  | CCCFE58204111124 | No |  |  | Cleaned | Fluidic pathway | Oxi mismatch |  | x
ABL90 |  |  | CCD02D8B043D7AAA | No |  |  | No fault found | pCO2 patient result |  |  | x
AQT90 FLEX |  |  | CCD042B0046259CF | No
ABL800Basic |  | e788 neu | CCCFEF2804047D00 | No |  | e788 new | Defect | Electrode | Electrode | pCO2
ABL835 FLEX |  |  | CCD032B904556177 | No |  |  | Defect | Basic unit | Radiance connection | Reboot
ABL825 FLEX | pco2 wdc is low , last 2 months | i replaced the pco2 electrode.old electrodejr-03 new LM-01 | CCCFFACE041D2158 | No | pco2 wdc is low , last 2 months | i replaced the pco2 electrode.old electrodejr-03 new LM-01 | Defect | Electrode | Electrode | pCO2
ABL835 FLEX |  |  | CCD028CA043AC25B | No |  |  | Defect | Power Supply | A/D converter
ABL820 FLEX | Ongoing WDC issue. The Data is being written to the USB drive, but when the information is uploaded to QA portal she does not recieve any information back. Tech support has been involved in trying to resolve. Today she recieved Feb data but still no Jan data. | I emailed WDC data to support box. Tech support is working on a solution. | CCD008C3040F47F0 | Yes | Ongoing WDC issue. The Data is being written to the USB drive, but when the information is uploaded to QA portal she does not recieve any information back. Tech support has been involved in trying to resolve. Today she recieved Feb data but still no Jan data. | I emailed WDC data to support box. Tech support is working on a solution. | Re classified RMED | WDC upload issue
ABL825 FLEX | checked and found f4 blown on au pcb | replaced fuse | CCCFDB7D0447F7C8 | No | checked and found f4 blown on au pcb | replaced fuse | Defect | Basic unit | Analyzer Control PCB | Fuse F4 blown
ABL825 FLEX | New Hospital will go live by March 17. Extended travel to pick up UPS' at FedEx. Install UPS, install ABL825 solutions and electrodes. Lamp check, Oxi Complete test, service tHb Cal, Install option keys and update type. Will allow system to run overnight and return tomorrow to run linearity. | Begin installation of ABL825. | CCD0232A044738C0 | Yes | New Hospital will go live by March 17. Extended travel to pick up UPS' at FedEx. Install UPS, install ABL825 solutions and electrodes. Lamp check, Oxi Complete test, service tHb Cal, Install option keys and update type. Will allow system to run overnight and return tomorrow to run linearity. | Begin installation of ABL825. | Re classified RMED | Installation
AQT90 FLEX |  |  | CCD01764044B978A | No
ABL80 OSM |  |  | CCD144D67A074575 | No
ABL80 OSM | Ran service test on diagnostics and failed neon lamp position and neon lamp intensity. Suspect is malfunctioning spectrometer. | Replaced spectrometer. Performed thb cal. Verified instrument operation. | CCD12FE47A2D8F09 | No | Ran service test on diagnostics and failed neon lamp position and neon lamp intensity. Suspect is malfunctioning spectrometer. | Replaced spectrometer. Performed thb cal. Verified instrument operation.
ABL90 |  |  | CCD15ADA7A465805 | No |  |  | Cleaned | Fluidic pathway | QC3 |  | x
ABL90 | will change SC | replaced SC with client stock | CCD13C4B7A5E5FA9 | No | will change SC | replaced SC with client stock | Defect | SC | pCO2 sens |  | x
ABL90 |  |  | CCD12E5E7A598395 | No |  |  | Defect | HW | CPU |  | x
ABL90 |  | der Daten in der Konfiguration, Warmstart durchgeführt. | CCD1610A7A218389 | No |  | data in the configuration, warm-start performed. | Adjustment | Reboot analyzer | His/lis |  | x
ABL90 |  |  | CCD152BF7A5B6095 | No |  |  | Setup | Roma.ini | His/lis | Eckhard: In the first SWO there was a connecting error. The FSE switch off the access to the customer network and checked the data configuration, but there was not found a mistake. After that he plugged in the cable to get access to the network and it works. We doesn’t know what was the reason. It could be a cable or connecting problem. | x
ABL90 |  | Erneuerung der SC. | CCD172A07A5A7861 | No |  | Renewal of sc. | Defect | SC | pCO2 |  | x
TCM-COMBIM | Confirmed the touch screen issues also noted Sound HW corruption errors. | Replacing the touch screen resolved that issue and replacing the CombiM module resolved the sound HW errors. | CCD15CBD7A16578D | No | Confirmed the touch screen issues also noted Sound HW corruption errors. | Replacing the touch screen resolved that issue and replacing the CombiM module resolved the sound HW errors. | Defect | Hardware | Touch Screen
ABL80 | REMOVE BACK PANEL AND FOUND DRIED CAL SOLUTION | ORDERED DOSING UNIT FOR FUTURE REPAIR | CCD561A5DC5C66F5 | No | REMOVE BACK PANEL AND FOUND DRIED CAL SOLUTION | ORDERED DOSING UNIT FOR FUTURE REPAIR
ABL90 |  | Aus-Einschalten | CCD51BA6DC0FC9F4 | No |  | Switching off | Cleaned | Fluidic pathway |  |  | x
ABL90 |  |  | CCD4ED88DC31FF7A | No |  |  | Defect | HW | Pinch valve |  | x
ABL90 | Unit in request flush intervention on arrival. Unit will not pass pump test. Solution pack test fails everything exceppt air test. | Flushed analyzer and replaced solution pack. No help. Replaced inlet gasket with holder. | CCD4ED8FDC59E323 | No | Unit in request flush intervention on arrival. Unit will not pass pump test. Solution pack test fails everything exceppt air test. | Flushed analyzer and replaced solution pack. No help. Replaced inlet gasket with holder. | Unknown | Revisit required | Techline |  | x
ABL825 FLEX | siehe SLR | siehe SLR | CCD4E083DC1C51D4 | No | see SLR | see SLR | Cleaned | Fluidic Module | Air vent
ABL825 FLEX | SWO marked as complaint type per Andy S.). Inspection of system revealed the QC has gradually ellevated up out of insert range. Lamp will not adjust correctly. Suspect hemolyzer. | The hemolyzer was replaced top resolve the reported issue. MetHb is now passing near the mean of each level. System is operational. | CCD518F1DC19AF2D | No | SWO marked as complaint type per Andy S.). Inspection of system revealed the QC has gradually ellevated up out of insert range. Lamp will not adjust correctly. Suspect hemolyzer. | The hemolyzer was replaced top resolve the reported issue. MetHb is now passing near the mean of each level. System is operational. | Defect | HW | Oxi Module | Hemolyzer | MetHb
ABL820 FLEX | The end of the pO2 electrode is missing. Ai valve is stuck in actuated position. | Replaced pO2 electrode and membrane kit. Replaced stuck air valve in pH.BG module. | CCD54CA6DC32AE2B | No | The end of the pO2 electrode is missing. Ai valve is stuck in actuated position. | Replaced pO2 electrode and membrane kit. Replaced stuck air valve in pH.BG module. | Defect | Electrode | Electrode | pO2 | Missing end
ABL805 FLEX | The analyzer is presenting system message 0675 Inlet error: Position calculation error. | Replaced the inlet friction detection unit. Replaced the inlet rubber tubing. Cleaned the inlet probe. | CCD5189ADC22D1F7 | No | The analyzer is presenting system message 0675 Inlet error: Position calculation error. | Replaced the inlet friction detection unit. Replaced the inlet rubber tubing. Cleaned the inlet probe. | Defect | Inlet module | Friction detector unit
ABL825 FLEX | Breaker arm moving erretically due to failure of motor mounting plate. | Replaced ampoule opener motor. | CCD50DA4DC1DAE27 | No | Breaker arm moving erretically due to failure of motor mounting plate. | Replaced ampoule opener motor. | Defect | Auto Check Module | Stepper motor | Breaker arm
TCM/Tosca |  |  | CCD53BFADC0DF219 | No |  |  | Defect | Hardware | Communication PCB
ABL90 | The printed METHb result is intermittently double printed or not advanced while printing which makes the METHb result unreadable. | Replaced the printer unit. Calibrated and ran QC. Instrument operates normally. | CCD659735250E992 | No | The printed METHb result is intermittently double printed or not advanced while printing which makes the METHb result unreadable. | Replaced the printer unit. Calibrated and ran QC. Instrument operates normally. | Defect | SW | SW driver Printer module | TEP MetHb printed double | x
ABL90 |  |  | CCD69BF35232D402 | No |  |  | Defect | HW | Touch screen |  | x
ABL90 | 508, 1063, 1178, 1081, 1177 errors in log. No eveidence of leaks. | Reflushed several times. Replaced oxi pinch tube and waste line. Replaced sensor and cal pack.Replaced hemolyzer. Replaced inlet tube, No help. | CCD65DFC5220BB29 | No | 508, 1063, 1178, 1081, 1177 errors in log. No eveidence of leaks. | Reflushed several times. Replaced oxi pinch tube and waste line. Replaced sensor and cal pack. Replaced hemolyzer. Replaced inlet tube, No help. | Unknown | Revisit rquired | Multi | Matt: We spoke with the local FSE and he stated the instrument is now fully operable for the last 1.5 weeks. The last visit they replaced the following: Sensor PCB, the bay unit, and the inlet module. Testing showed that SP replacement would work for ~24 Hrs, but testing in the service identified blockage errors. | x
ABL800Basic |  |  | CCD665BD522D7FF6 | No |  |  | Cleaned | ph/BG Module | Measuring chamber
ABL835 FLEX | Archiv, Service | Elektrodenfläche gereinigt, Nullstromkalibrierung | CCD64C785228986D | No | Archive, Service | Electrode surface cleaned, zero current calibration | Adjustment | El/Met Module | Zero current
ABL825 FLEX | Unit operating normally on arrival. | Reloaded permanent keys. | CCD68728521521C2 | No | Unit operating normally on arrival. | Reloaded permanent keys. | Setup | Parameter keys
ABL825 FLEX |  | Replaced OXI Module PCB. | CCDF058C8C448C5E | No |  | Replaced OXYGEN Module PCB. | Defect | HW | Oxi Module | OXI PCB | unknown
ABL825 FLEX |  |  | CCD7A0FCC82FE344 | No |  |  | Cleaned | Electrode | Electrode | pH
ABL80 CO-OX | Realización de TU0114. ABL80 y entrega de documentación al cliente. Desmontaje de mod. fluídica. Sustitución de waste pump head y waste manifold. Sustitución de tubing set y manifold unit. Revisión de sistema mecánico y neumático Calibración y verificación de funcionamiento. Correcto según pruebas. |  | CCD8BCAB3E62D106 | No | Realization of TU0114. ABL80 and delivery of documentation to the customer. Mod disassembly. Fluidic. Substitution of waste pump head and waste manifold. Substitution of tubing set and manifold unit. Mechanical and pneumatic system review Calibration and operation verification. Correct according to evidence.
ABL90 |  |  | CCE72C0C50318A87 | No |  |  | Defect | SC | pCO2 QC |  | x
ABL90 |  |  | CCD8CB943E593E5C | No |  |  | Defect | SC | FT |  | x
ABL90 | checked data logs for errors, checked sample path for clots and flow of solutions, checked all connections, | flushed sample path, replaced inlet and pinch tube, replaced inlet probe and inlet holder, performed pump, pressure, liquid sensor, solution pack and clot detection test multiple times, | CCD91E033E469660 | No | checked data logs for errors, checked sample path for clots and flow of solutions, checked all connections, | flushed sample path, replaced inlet and pinch tube, replaced inlet probe and inlet holder, performed pump, pressure, liquid sensor, solution pack and clot detection test multiple times, | Replaced | Probe | Revisit |  | x
ABL90 | Took the inlet module off and inspected. Didn't observe any problems. However, I did notice that the inlet gasket doesn't always seat properly on the solution pack. | I swapped the inlet module with the 422N010 analyzer. | CCD8E1673E4340E2 | No | Took the inlet module off and inspected. Didn't observe any problems. However, I did notice that the inlet gasket doesn't always seat properly on the solution pack. | I swapped the inlet module with the 422N010 analyzer. | Adjustment | Inlet module | Swobbed with other analyzer |  | x
ABL810 FLEX |  |  | CCD8DDA63E5D6788 | No |  |  | Defect | Fluidic Module | Valve, air | Membrane
ABL825 FLEX | Comprobación del estado de funcionamiento del equipo, se realizan programas de test por servicio, se sustituye la PCB del Oxi, se desmonta el módulo Inlet, se comprueba el estado, se engrasa el vástago del eje, se comprueba las puertas de entrada, tubos, Inlet, se carga de nuevo el software del equipo, se instalan de nuevo las llaves de remoto, se reinicia el módulo PC, se calibra por servicio la tHb, se calibra el equipo a 2 puntos. Comprobación final con QC. |  | CCD8CB063E317221 | No | Checking the operating status of the equipment, testing programs are carried out by service, replacing the PCB of the Oxi, dismanting the Inlet module, checking the status, greasethe shaft stem, checking the entrance doors, tubes, Inlet, loading new computer software, remote keys are reinstalled, PC module is restarted, tHb is calibrated by service, the computer is calibrated to 2 points. Final check with QC. |  | Defect | HW | Oxi Module | OXI PCB
ABL825 FLEX | Went into control panel for correct time zones | Corrected time zone and checked day lite savings. Tested QC and now runs at correct time. | CCD8CEC43E5C2687 | No | Went into control panel for correct time zones | Corrected time zone and checked day lite savings. Tested QC and now runs at correct time. | Setup | QC interval | Timing | Daylight savings
ABL825 FLEX | This changed when day lite savings took place | Corrected time zone and checked daylit savings. QCs now run on time. | CCD8D0983E18268A | No | This changed when day lite savings took place | Corrected time zone and checked daylit savings. QCs now run on time.
ABL825 FLEX | Checked time zone in control panel settings are off. | Reset correct Zone and checked daylite savings. QC now runs at correct times | CCD8D1C23E36268D | No | Checked time zone in control panel settings are off. | Reset correct Zone and checked daylite savings. QC now runs at correct times | Setup | QC interval | Timing | Daylight savings
ABL825 FLEX | This happened after the time change | I control panel I reset time zone and checked daylite savings. QCs now run on time. | CCD8D2E73E52268F | No | This happened after the time change | I control panel I reset time zone and checked daylite savings. QCs now run on time. | Setup | QC interval | Timing | Daylight savings
ABL835 FLEX | Notified while on site for another system. This analyzer and R2 have more frequent errors than other systems in Lab. Customer has already tried different lot of membranes, decontamination and tubing replacements. Request that I change upper tubing from Met Valve to LS. I also verified zero current was in spec. Did not adjust. | Replace tubing to upper Met Module (customer stock) | CCD8D8CC3E5738EE | No | Notified while on site for another system. This analyzer and R2 have more frequent errors than other systems in Lab. Customer has already tried different lot of membranes, decontamination and tubing replacements. Request that I change upper tubing from Met Valve to LS. I also verified zero current was in spec. Did not adjust. | Replace tubing to upper Met Module (customer stock) | Defect | El/Met Module | pump tube | Lac | Unstable
ABL820 FLEX | Examined inlet module | Replaced friction detector unit and inlet probe. Verified operations of ABL820. | CCD8C3D63E272234 | No | Examined inlet module | Replaced friction detector unit and inlet probe. Verified operations of ABL820. | Defect | Inlet module | Friction detector unit
ABL835 FLEX |  |  | CCD8ED843E25F6D8 | No |  |  | Defect | Electrode | Electrode | pCO2
ABL90 | Ca not passing calibrations. | Ran 5 waste bloods then ran the calibration and then ran QC. | CCDA201FB424DA1F | No | Ca not passing calibrations. | Ran 5 waste bloods then ran the calibration and then ran QC. | User Error | waste blood
ABL90 | siehe ausgeführte Arbeiten | siehe ausgeführte Arbeiten | CCD9E68DB408E7A7 | No | see work carried out | see work carried out | Adjustment | HW | Fan noise |  | x
ABL90 | System geprüft | Spectrometer und Hemolyzer getauscht | CCDA178FB420417B | No | System checked | Spectrometer and Hemolyzer exchanged | Defect | HW | Sepctrometer | OOB | x
ABL90 | Actualizar el software a la versión 3.0MR2. ES, equipo pendiente de ser instalado. |  | CCDA11CDB46356CF | Yes | Update the software to version 3.0MR2. ES, equipment pending to be installed. |  | Re classified RMED | FAN / TU
ABL800Basic |  |  | CCDA3FFFB4123ECE | No |  |  | Defect | Inlet module | Probe bend
ABL835 FLEX |  |  | CCDA3AD8B450080A | No |  |  | Defect | Electrode | Membrane | pCO2
ABL825 FLEX |  |  | CCDA094DB42F00AA | No |  |  | Defect | Power Supply
AQT90 FLEX | Calibración sensor de copas con mismo resultado (error 13). apertura equipo. Presenta gran cantidad de polvo por todas las partes del equipo, especialmente en la delantera. Aspirado y soplado con aspirador. Desmontaje motor punch out y limpieza del sensor de la copa con pincel (se veia opacado). Retirada de pelusas de los ejes de los motores de la process unit, cartridge lock y punch out y de desbloqueo del cartucho de soluciones. Engrase de los mismos. Limpieza tubos 7 y 8. chequeos de la wet section. Cal del sensor de copa OK varias veces. Chequeos de la wet section, safe y puesta en marcha. Blancos de TnI | Calibración sensor de copas con mismo resultado (error 13). apertura equipo. Presenta gran cantidad de polvo por todas las partes del equipo, especialmente en la delantera. Aspirado y soplado con aspirador. Desmontaje motor punch out y limpieza del sensor de la copa con pincel (se veia opacado). Retirada de pelusas de los ejes de los motores de la process unit, cartridge lock y punch out y de desbloqueo del cartucho de soluciones. Engrase de los mismos. Limpieza tubos 7 y 8. | CCDA2093B4247755 | No | Calibration cup sensor with same result (error 13). team opening. It has a lot of dust all over the equipment, especially in the front. Vacuumed and blown with vacuum cleaner. Disassembly motor punch out and cleaning the cup sensor with brush (it looks opaque). Removing lint from the shafts of the process unit, cartridge lock and punch out motors and unlocking the solution cartridge. Grease them. Cleaning tubes 7 and 8. wet section checks. OK cup sensor cal several times. Wet section checks, safe and commissioning. Whites of TnI | Calibration cup sensor with same result (error 13). team opening. It has a lot of dust all over the equipment, especially in the front. Vacuumed and blown with vacuum cleaner. Disassembly motor punch out and cleaning the cup sensor with brush (it looks opaque). Removing lint from the shafts of the process unit, cartridge lock and punch out motors and unlocking the solution cartridge. Grease them. Cleaning tubes 7 and 8.
TCM-COMBIM |  |  | CCDA21F6B4556DE0 | No |  |  | Defect | Electrode | Regeneration
ABL825 FLEX | SE PRUEBA A INTERCAMBIAR MODULO DE PH GS NO SIENDO ESTE EL PROBLEMA LA MUESTA SE QUEDA A 3 4 PARTE DEL LLENADO DE LA CAMARA DE MEDIDA.SE EFECTUAN DIFERNTES PRUEBAS PARA DETERMINAR LA ENTRADA DE AIRE QUE IMPIDE LA PERFECTA TRACCION DE LA MUESTRAS, SE SUSTITUYEN TUBOS SE INTERCAMBIAN MODULOS NO ENCONTRANDOSE EL PROBLEMA. SE DESMONTA MODULO DE FLUIDICA Y SE TESTEAN LAS VALVULAS ESTANDO TODO EN PERFECTO ESTADO. TRAS EL MONTAJE DE DICHO MODULO DE FLUIDICA Y MONTAJE DE BOMBA DE FLUIDICA DEL MODULO NUEVA, SE REINICIA TODO EL SISTEMA SE CALIBRAM BOMBAS Y SL, SE EFECTUAN NUMEROSAS PRUEBAS CON MUESTRAS REALES SIENDO ESTAS CORRECTAS Y NO DETACTANDOSE PRESENCIA DE AIRE EN EL SISTEMA. CALIBRADO Y VERIFICADO CON Q.C. | LIMPIEZA MODULO FLUIDICA | CCDB7BD22A5A9A74 | No | IT IS TESTED TO EXCHANGE PH GS MODULE NOT BEING THIS THE PROBLEM THE MUESTA IS 3 4 PART OF THE HOUSE FILLING OF MEDIDA.SE EFFECT DIFERNTES TESTS TO DETERMINE THE AIR INPUT THAT PREVENTS THE PERFECT TRACTION OF THE SAMPLES , TUBES ARE SUBSTITUTED, MODULOS ARE NOT FOUND. FLUIDIC MODULE IS DISMONTATED AND THE VALVE IS DESTED BEING EVERYTHING IN PERFECT STATE. AFTER THE ASSEMBLY OF SUCH FLUIDIC MODULE AND NEW MODULE FLUIDIAL PUMP ASSEMBLY, ALL SYSTEM IS RESTARTED PUMPS AND SL, NUMEROUS TESTS ARE EFFECTED WITH REAL SAMPLES BEING THESE CORRECT AND NOT DETACING AIR PRESENCE IN THE SYSTEM. CALIBRATED AND VERIFIED WITH Q.C. | CLEANING FLUID MODULE | Defect | Fluidic Module | Pump tubing | Oxi
ABL80 | PREVIOUS SWO FOUND PROBABLE CAUSE TO BE MANIFOLD, HOWEVER FURTHER CLEANING FOUND PUMP HAD SMALL HOLE | REPLACED PUMP, WASTE MANIFOLD AND CONNECTORS. CLEANED SYSTEM COMPLETELY FOR VISUAL FOLLOWUP | CCDDCAF31629A278 | No | PREVIOUS SWO FOUND PROBABLE CAUSE TO BE MANIFOLD, HOWEVER FURTHER CLEANING FOUND PUMP HAD SMALL HOLE | REPLACED PUMP, WASTE MANIFOLD AND CONNECTORS. CLEANED SYSTEM COMPLETELY FOR VISUAL FOLLOWUP
ABL90 |  |  | CCDDD452164460CE | Yes |  |  | Re classified RMED | Telephone support | Hotline
ABL90 | Unit working on arrival except for pO2 ststus error. Unit not passing all fluid transport tests. | Replaced and aligned inlet module, sensor bay and sensor interface. | CCDDF6581644389E | No | Unit working on arrival except for pO2 ststus error. Unit not passing all fluid transport tests. | Replaced and aligned inlet module, sensor bay and sensor interface. | Defect | HW | multi (Macomb) | Matt: We spoke with the local FSE and he stated the instrument is now fully operable for the last 1.5 weeks. The last visit they replaced the following: Sensor PCB, the bay unit, and the inlet module. Testing showed that SP replacement would work for ~24 Hrs, but testing in the service identified blockage errors. | x
ABL825 FLEX | Sustitucion del display untit Cal, qc. | Sustitucion del display untit | CCE22C80780B8A2D | No | Replacement of the display untit Cal, qc. | Untit display replacement | Defect | Basic unit | Display | Very old
ABL825 FLEX | Tested system, checked settings, checked and tested windows settings. | Shut analyzer down and replace speaker. | CCDDFC541612376D | No | Tested system, checked settings, checked and tested windows settings. | Shut analyzer down and replace speaker. | Defect | DMS PC | Loud speaker
ABL820 FLEX | Troubleshoot the error, testing valves and fluid path in service. Noted the air line was partial blocked with dust. | Flush air line and the ran a LS adjustment. the test passed and analyzer was operational. | CCDDFD89165E3772 | No | Troubleshoot the error, testing valves and fluid path in service. Noted the air line was partial blocked with dust. | Flush air line and the ran a LS adjustment. the test passed and analyzer was operational. | Cleaned | Fluidic module | Air vent
ABL825 FLEX |  |  | CCDDDD29161AD496 | No |  |  | Defect | Electrode | Electrode | Glu
ABL820 FLEX |  |  | CCDE27C6160D3C31 | No |  |  | Defect | Connection cable | His/lis
ABL805 FLEX | The analyzer is intermittently presenting system message 675 Inlet position calculation error. | Replaced the inlet module unit. Replaced the inlet controller PCB. Performed a service inlet calibration. Worked with the customer's IT technician in setting up all ABL800s to connect to their network system using a static IP address. The SEND button kept disapearing. Found the instrument setup to use a DHCP IP address and that is not recommended for our analyzers. | CCDDC3B51635C9F5 | No | The analyzer is intermittently presenting system message 675 Inlet position calculation error. | Replaced the inlet module unit. Replaced the inlet controller PCB. Performed a service inlet calibration. Worked with the customer's IT technician in setting up all ABL800s to connect to their network system using a static IP address. The SEND button kept disapearing. Found the instrument setup to use a DHCP IP address and that is not recommended for our analyzers. | Defect | Inlet module | Controller PCB
ABL827 FLEX | 0581 errors on blood. I ran a complete oxy test which passed. | I replaced the hemo;yzer unit. | CCDDC1F8164F39F1 | No | 0581 errors on blood. I ran a complete oxy test which passed. | I replaced the hemo;yzer unit. | Defect | HW | Oxi Module | Hemolyzer
ABL835 FLEX |  |  | CCDDF2F916351962 | No |  |  | Defect | Electrode | Electrode | pCO2 | although ph/BG module
AQT90 FLEX |  |  | CCDE26A516143C2E | No
TCM-COMBIM | Bei Start des Gerätes im betriebswarmen Zustand erscheint die Meldung: Sound HW corrupted, call service . | Da es sich um ein bereits modifiziertes Modul (1. Modifikation) handelt, wurde das Main CPU PCB erneuert, | CCDDB185163E79B1 | No | When the device is started in a warm state, the message appears: Sound HW corrupted, call service . | Since it is an already modified module (1st modification), the Main CPU PCB has been renewed, | Defect | Hardware | PC Module
ABL80 CO-OX |  | Gerät runtergefahren und demontiert, Undichtigkeit im hinteren Teil des Manifolds und hinten unten gereinigt (stark ausgelaufen), Schläuche gespült, Lösungspack ersetzt, da fast leer, Kal, QC | CCDEFFAE8C14F2A1 | No |  | Device shut down and dismantled, leakage in the back of the manifold and cleaned at the back down (strongly leaked), hoses flushed, solution pack replaced as almost empty, Kal, QC
ABL80 | Po2 is out on all three external controls. | I performed the Zero calibration It was off. I performed a couple of calibrations and brought system to ready. | CCDF375E8C5ECF3B | No | Po2 is out on all three external controls. | I performed the Zero calibration It was off. I performed a couple of calibrations and brought system to ready.
ABL90 |  |  | CCDEEE068C371305 | No |  |  | Defect | HW | Oxi Module | Pinch Valve tube | x
ABL90 | Unit operating on arrival. No QC out of limits, no cal errors since unit was lsat seen. | Np fault found. | CCDF84278C23F066 | No | Unit operating on arrival. No QC out of limits, no cal errors since unit was lsat seen. | Np fault found. | No fault found | pH out |  |  | x
ABL90 |  |  | CCE029A8020B552E | No |  |  | Defect | HW | Gasket Inlet |  | x
ABL825 FLEX | INTENTO DE AJUSTES MEDIANTE SERVICIO. REVISION DE SONDA ENTRADA Y DEMAS COMPONENTES. SE DESMONTA MODULO COMPLETAMENTE Y SE REVISAN CONEXIONES, CABLES PLANOS Y PARTES MECANICAS. CARGA DE SW AU. AL NO PODER SER REPARADO SE PIDEN PIEZAS PARA SU REPARACION. |  | CCDF3E248C4F27CD | No | ATTEMPTING TO ADJUST BY SERVICE. REVISION OF INPUT SOUND AND OTHER COMPONENTS. COMPLETELY DISASSEMBLED AND FLAT CONNECTIONS, FLAT CABLES AND MECHANICAL PARTS ARE REVIEWED. SW AU LOAD. WHEN IT CANNOT BE REPAIRED, PARTS ARE ORDERED FOR REPAIR. |  | Defect | Inlet module | unknown | revisit required
ABL825 FLEX |  | Tausch der Glucoe Elektrode. | CCDF29428C302B90 | No |  | Replacement of the Glucoe electrode. | Defect | Electrode | Electrode | Glu
ABL825 FLEX | The MetHb measures about 30 percent higher on blood samples than all of the other analyzers. QC checks okay. Symptoms indicate a hemolyzer issue. | Replaced the hemolyzer unit and performed alignments as required. | CCDF07158C55CCF7 | No | The MetHb measures about 30 percent higher on blood samples than all of the other analyzers. QC checks okay. Symptoms indicate a hemolyzer issue. | Replaced the hemolyzer unit and performed alignments as required. | Defect | HW | Oxi Module | Hemolyzer | MetHb
ABL800Basic |  |  | CCDF76468C5BF312 | No |  |  | Cleaned | ph/BG Module | Measuring chamber | Clots
ABL820 FLEX | Have replaced electrode, remembraned ,cleaned measuring chamber and replaced blue tubing but they are still getting intermittant PCO2 cal failures. | Replaced the air valve on the PH BG module. Suspected that it may have been sticking. | CCDF0BEB8C4DDBE7 | No | Have replaced electrode, remembraned ,cleaned measuring chamber and replaced blue tubing but they are still getting intermittant PCO2 cal failures. | Replaced the air valve on the PH BG module. Suspected that it may have been sticking. | Defect | Electrode | Electrode | pCO2
ABL800Basic |  | Tausch des barcodesscanners. | CCDEFB6E8C06A392 | No |  | Replacement of the barcode scanner. | Defect | Basic unit | Barcode Reader
ABL90 |  |  | CCE02FC0024A0F17 | No |  |  | User Error | Incorrect SP activation |  | Eckhard: We had installed the new needle and gasket last year and we had place one set at customer site. During the last three month the customer has bend both two needle so that he works with the old version now. The problem was the activation of the solution pack. And this bends the needles. So the problem with clots in the device was not with the new version. | x
ABL90 | Liquid sensor error present on arrival. Found signs of QC leakage at hemolyzer. Replacement of pinch valve tubing did not satisfy LS error. Found pin hole leak in waste line tubing. | Replaced all tubings and cleaned hemolyzer. calibrated and ran QC. Instrument operates normally. | CCE0737C0214C08A | No | Liquid sensor error present on arrival. Found signs of QC leakage at hemolyzer. Replacement of pinch valve tubing did not satisfy LS error. Found pin hole leak in waste line tubing. | Replaced all tubings and cleaned hemolyzer. calibrated and ran QC. Instrument operates normally. | Defect | HW | Waste Tube |  | x
ABL80 | se procedio a sustituir placa interfase. tras esto se realiza ajuste del cero según manual. se solventa problema inicial de la placa. se observan fallos al calibrar casete de electrodos. se procederá a desmontar equipo para limpiar neumatica al completo. Se realizo la FAN TU114 (actualización a la versión de SW 3.11) |  | CD0F0D7C104A62C5 | No | proceeded to replace the interface plate. after this zero adjustment is made according to manual. solves initial board problem. faults are observed when calibrating electrode cassette. equipment to clean the entire tire will be carried out. THE TU114 FAN was performed (upgrade to SW version 3.11)
ABL825 FLEX | SE DESMONTA MODULO INLET Y SE SUSTITUYE INLET PCB, AL EFECTUAR LAS PRUEBAS DE AJUSTE SIGUE DANDO LOS MISMOS SINTOMAS Y SE PROCEDE A LA SUSTITUCION DE LA UNIDAD DETECTORA DE FRICCION Y A COLOCAR LA INLET PCB SUSTITUIDA ANTERIORMENTE. SE MONTA TODO EL SISTEMA SE EFECTUAN PRUEBAS DE AJUSTE Y FUNCIONAMIENTO, SINDO ESTAS SATISFACTORIAS. CALIBRADO Y VERIFICADO CONQ.C. | SUSTITUCION FRICCION DETECTOR UNIT | CCE0241302047EE8 | No | THE INLET MODULE IS DISMONTATED AND SUBSTITUTE INLET PCB, BY MAKING THE ADJUSTMENT TESTS STILL GIVE THE SAME SYMPTOMS AND THE REPLACEMENT OF THE FRICCION DETECTOR UNIT IS PROVIDED AND THE SUBSTITUTE OF THE PROVIDED PCB INLET. ALL SYSTEM IS MOUNTED, ADJUSTMENT AND OPERATION TESTS ARE EFFECTED, WITHOUT THESE SATISFACTORY. CALIBRATED AND VERIFIED WITH. C. | REPLACEMENT FRICCITION DETECTOR UNIT | Defect | Inlet module | Friction detector unit
ABL825 FLEX | Checked electrode connector. Found nothing wrong with it. | Opened connector and removed electrode. Cleaned chamber. | CCE0C1CC0206A25F | No | Checked electrode connector. Found nothing wrong with it. | Opened connector and removed electrode. Cleaned chamber. | User Error | Training new user | Electrode replacement
ABL825 FLEX | checked flow of solutions, checked valves, checked hemolyzer and lamp unit, | performed lamp adjustment, performed complete oxi test, | CCE0DA69024598C9 | No | checked flow of solutions, checked valves, checked hemolyzer and lamp unit, | Performed lamp adjustment, performed complete Oxi test, | Adjustment | HW | Oxi Module | Lamp unit | New install
AQT90 FLEX |  |  | CCE02E130257184E | No
ABL90 | Limpiar el circuito neumático, desobstruir y enderezar la aguja , calibrar y probar con controles. | Desobstruir y enderezar la aguja de entrada de muestras. | CCE1CFC6783856EA | No | Clean the pneumatic circuit, unblock and straighten the needle, calibrate and test with controls. | Unblock and straighten the sample entry needle. | Cleaned | Fluidic pathway |  |  | x
ABL835 FLEX | siehe aushegführte Arbeiten | siehe aushegführte Arbeiten | CCE1B7307832F86C | No | see excerpted work | see excerpted work | Cleaned | Auto Check Module | Barcode Reader
ABL825 FLEX | torn Air valve membrane on BG Module top side. | Replaced membrane on air valve.After unit was back up the pc02 was reading to high on level 3 QC, changed membrane with client stock. | CCE1EA7C785020C1 | No | torn Air valve membrane on BG Module top side. | Replaced membrane on air valve. After unit was back up the pc02 was reading to high on level 3 QC, changed membrane with client stock. | Defect | ph/BG Module | Valve, air | Membrane
ABL837 FLEX | Fails the OXI tests in service. | Replaced defective PCB.Load AU software. Now passes all OXI tests. Complete OXI alingment. tHb cal in service | CCE1B8717849500A | No | Fails the OXI tests in service. | Replaced defective PCB. Load AU software. Now passes all OXI tests. Complete OXI alingment. tHb cal in service | Defect | HW | Oxi Module | OXI PCB
ABL820 FLEX | DISCUSSED WITH JEFF ROBISON AND THE DEPT. LEADS ALL THE ISSUES WITH THIER NEW ANALYZERS. MOST SEEM TO BE SETUP ISSUES WHICH WILL BE FOLLOWED UP BY JEFF. I ALSO VERIFIED THE REBOOT ISSUE. | TECHLINE CASE OPENED AND DATA LOGS WITH SERVICE DUMP ATTACHED. RAME WILL INVESTIGATE. | CCE1CD887848864E | No | DISCUSSED WITH JEFF ROBISON AND THE DEPT. LEADS ALL THE ISSUES WITH THIER NEW ANALYZERS. MOST SEEM TO BE SETUP ISSUES WHICH WILL BE FOLLOWED UP BY JEFF. I ALSO VERIFIED THE REBOOT ISSUE. | TECHLINE CASE OPENED AND DATA LOGS WITH SERVICE DUMP ATTACHED. RAME WILL INVESTIGATE. | Unknown | DMS or AU or PS
ABL825 FLEX | Error 742. Autocheck carousel failure. Found FPC cable incorrectly mounted.Cable causing undue stress on housing. | Replaced cable. | CCE1A67678482D61 | No | Error 742. Autocheck carousel failure. Found FPC cable incorrectly mounted. Cable causing undue stress on housing. | Replaced cable. | Adjustment | Auto Check Module | FPC Cable | Misaligned
ABL825 FLEX |  |  | CCE1CDA578182A47 | No |  |  | Defect | DMS PC | HDD
TCM-COMBIM | Checked power supply, checked fuses on mother board. | Replaced CPU PCB. | CCE1F8507840F4BB | No | Checked power supply, checked fuses on mother board. | Replaced CPU PCB. | Defect | Hardware | PC Module
ABL80 | Entrenar al usuario en el cambio de los consumibles. | Entrenamiento a los usuarios. | CCE2DC1AEE0456F5 | No | Train the user in changing consumables. | Training users.
ABL90 |  |  | CCE2F1B9EE55006D | No |  |  | Adjustment | tHb Cal | Oxi complete test |  | x
ABL90 |  |  | CCE680045003D745 | No |  |  | Cleaned | Fluidic pathway |  |  | x
ABL90 | On arrival on site to change roma.ini, customer informed me of the lactate issue with the Level 3 external QC (Lot 118). Found that problem occurred after SP was changed on 3 19. Changed SP but it did not change problem. Tried running blood samples, that did not work. Changed SC and now pCO2 cal sensitivity is ? low (value 41.6 low limit is 60). flushed analyzer and still no improvement. | Changed roma.ini file to show values out of reportable range. Installed 2 new SC. Customer Credit forms sent for 1 SP and 2 SC. | CCE2FE9DEE63F4EB | Yes | On arrival on site to change roma.ini, customer informed me of the lactate issue with the Level 3 external QC (Lot 118). Found that problem occurred after SP was changed on 3 19. Changed SP but it did not change problem. Tried running blood samples, that did not work. Changed SC and now pCO2 cal sensitivity is ? low (value 41.6 low limit is 60). flushed analyzer and still no improvement. | Changed roma.ini file to show values out of reportable range. Installed 2 new SC. Customer Credit forms sent for 1 SP and 2 SC. | Re classified RMED | by RAME
ABL90 |  |  | CCE2CDBFEE2AC628 | No |  |  | Defect | HW | SP PCB |  | x
ABL90 |  | Geräteneustart, Spülung der Fließwege, diverse Serviceprogramme durchgeführt | CCE30416EE4F5F0F | No |  | Device restart, flushing of flow paths, various service programs carried out | Cleaned | Fluidic pathway | Lac Cal + tHb spectrum mismatch |  | x
ABL90 |  |  | CCE2DFA5EE10667B | No |  |  | Cleaned | Fluidic pathway | Unable to pump |  | x
ABL90 | Colocar correctamente la goma de unión de la aguja, cambiar junta de entradad de muestras. Limpiar el circuito neumático, quitar el polvo del interior, comprobar el registro de errores, guardar y borrar datos, calibrar y probar con controles. | Colocar correctamente la goma de unión de la aguja, cambiar junta de entradad de muestras. | CCE312DFEE1156F8 | No | Place the needle attachment rubber correctly, change sample inlet joint. Clean the pneumatic circuit, remove dust from the inside, check the error log, save and erase data, calibrate and test with controls. | Place the needle attachment rubber correctly, change sample inlet joint. | Replaced | Inlet gasket |  | The black inlet gasket was twisted over the SP output… So the FSE solved the problem (Error 1342 leak detected) and performed a preventive maintenance scheduled for April | x
ABL815 FLEX |  |  | CCE2F384EE070DF9 | No |  |  | Defect | Electrode | Electrode | K+
ABL825 FLEX | Customer performed troubleshooting outlined in the OM. removed the inlet gasket and found the syringe nipple partially clogged. found the tubing organizer waste connector cracked and the waste tube loose on y connector. | Replace waste nipples, install new tubing organizer and install anew waste tube from customer stock. Straightened bent inlet probe. | CCE314B7EE58431C | No | Customer performed troubleshooting outlined in the OM. removed the inlet gasket and found the syringe nipple partially clogged. found the tubing organizer waste connector cracked and the waste tube loose on y connector. | Replace waste nipples, install new tubing organizer and install anew waste tube from customer stock. Straightened bent inlet probe. | Defect | Inlet module | Nipple kit | Leaking
ABL837 FLEX | Found blockage in waste lines | Flushed out waste lines, changed inlet nipples. Removed all electrodes and cleaned measuring chambers, reseated electodes. No leakage at inlet. | CCE2D06BEE2BCDA3 | No | Found blockage in waste lines | Flushed out waste lines, changed inlet nipples. Removed all electrodes and cleaned measuring chambers, reseated electodes. No leakage at inlet. | Defect | Inlet module | Nipple kit
ABL825 FLEX |  | Ventilmembran vo Ventil V7 und V8 ersetzt, Ventile gereinigt, im Complete Oxymetrietest Position Neon Position Failed Bereich 0,1 bis 2 --> 0,099 zu niedrig, Service tHb Kal 2x fehlerhaft, probehalber Tausch der Oxy Modul PCB --> gleicher Fehler, Wechsel des Oxy-Moduls (Modul ausgebaut 902-833R1546N007), Complete Oxymetrie Test i.O. -->Wert jetzt 0,487 i.O., Kal, QC | CCE2B75FEE44D21D | No |  | Valve diaphragm vo valve V7 and V8 replaced, valves cleaned, in Complete Oxymetry Test position Neon Position Failed range 0.1 to 2 --> 0.099 too low, Service tHb Kal 2x faulty, test replacement of the Oxy Module PCB --> same error, change of the Oxy module (module extended 902-833R1546N007), Complete Oxymetry Test i.O. -->Value now 0.487 i.O., Kal, QC | Defect | HW | Oxi Module | 316 and 579
ABL820 FLEX | Tried new pCO2 membrane | Replaced pH BG air valve | CCE310D2EE12E585 | No | Tried new pCO2 membrane | Replaced pH BG air valve | Defect | ph/BG Module | Valve, air
AQT90 FLEX |  |  | CCEBA5F4280E1A1E | No
ABL80 FLEX Basic version | Verificação do histórico, substituição SP porque estava a terminar sem sucesso, substituição da SC, establização e passagem do QC |  | CCE6C5E1503C43B2 | No | Checking the history, replacing SP because it was terminating without success, replacing the SC, establishing and passing the QC
ABL90 |  |  | CCE686875002DD06 | No |  |  | Cleaned | Fluidic pathway | Email Here we have an “old” analyzer. The first installation was January 213 and it was for demo. | In this case the customer has problems with clots in fluid path and the customer calls for FSE. The technician does a yearly maintenance and cleans the fluid path. The problem was fixed. The maintenance was in the correct time difference to the last maintenance. | x
ABL90 |  |  | CCE730C050368A8B | Yes |  |  | Re classified RMED | Phone support
ABL90 |  |  | CCE6B7BE50274E2F | No |  |  | Defect | SP | FT |  | x
ABL90 |  |  | CCE7AEC1C6512A30 | No |  |  | Replaced | Probe |  |  | x
ABL825 FLEX |  |  | CCE6F133503BE991 | No |  |  | Cleaned | Gas module | Tubings | ??
ABL820 FLEX |  |  | CCE68EE5501A9E2A | Yes |  |  | Re classified RMED | Courtesy | Replaced Inlet Probe
ABL810 FLEX | Inspected pH BG air valve, it was jammed in open position. | Tried used pH BG module for troubleshooting. | CCE69CA3502AE592 | No | Inspected pH BG air valve, it was jammed in open position. | Tried used pH BG module for troubleshooting. | Defect | ph/BG Module | Valve, air | Stuck in open
AQT90 FLEX |  |  | CCE6A182500B2588 | No
ABL90 |  | Tausch der gesamten Verschlauchung, Neuinitialisierung. | CCE7DA2FC61D9AD8 | No |  | Replacement of the entire hose, reinitialization. | Adjustment | power off/on | Flow selector cal | ECKHI mail: The problem was fixed with the reinitialization of the analyzer. | x
ABL90 | Comprobación del estado de funcionamiento del equipo, se desmonta el equipo para la comprobación de la conexión y del estado del cable del display, se desconecta y se vuelve a conectar dicho cable, se reinicia el equipo y verificamos el correcto funcionamiento.. | Desconexión y conexión del cable del display | CCE7DAAEC6226889 | No | Checking the operating status of the computer, dismantling the equipment for the connection check and the status of the display cable, disconnecting and reconnecting the cable, restarting the computer and verifying the correct operation.. | Disconnecting and connecting the display cable | Adjustment | Cable | re attached display |  | x
ABL800Basic |  |  | CCE83AB0C6297E44 | No |  |  | Cleaned | Auto Check Module | Barcode Reader | dust
ABL837 FLEX | Upon arrival, I found the instrument leaking at the inlet gasket port. | Upon visual inspection, I found a waste line disconnected. reconnected the waste line and performed a full flush of all waste lines. Performed a liquid sensor adjust and a pump calibration. | CCE80A93C641E587 | No | Upon arrival, I found the instrument leaking at the inlet gasket port. | Upon visual inspection, I found a waste line disconnected. reconnected the waste line and performed a full flush of all waste lines. Performed a liquid sensor adjust and a pump calibration. | Defect | Inlet module | Waste tube disconnected
ABL820 FLEX | Reading on 2 point cal is above 100% for PCO2. After doing a couple more it comes in to the range. Happens every couple of weeks. | Replaced the regulators , replaced all o-rings on the electrodes, remembraned the PCO2 electrode. | CCE7E346C64202DA | No | Reading on 2 point cal is above 100% for PCO2. After doing a couple more it comes in to the range. Happens every couple of weeks. | Replaced the regulators , replaced all o-rings on the electrodes, remembraned the PCO2 electrode. | Defect | Gas module | Left flow regulator | according to EKH
ABL835 FLEX | Customer previously verified all cables, connections and net ports are good. I reset netwok adapter - no help. Set static IP address and was able to ping local and AU PCB but not Rad server or peer analyzers. At suggestion of SIT I removed an older analyzer from network, set R8 network for DNS and DHCP, rebooted and system now connects to Radiance. Older system, when reconnected to network reports license not available. System R8 then transferred some 15,000 record entries to Rad Server | Remove a 6th analyzer from network. I notofied Sales and they will work with SIT to sort out the number of licenses available. | CCE7E6A2C63942A0 | No | Customer previously verified all cables, connections and net ports are good. I reset netwok adapter - no help. Set static IP address and was able to ping local and AU PCB but not Rad server or peer analyzers. At suggestion of SIT I removed an older analyzer from network, set R8 network for DNS and DHCP, rebooted and system now connects to Radiance. Older system, when reconnected to network reports license not available. System R8 then transferred some 15,000 record entries to Rad Server | Remove a 6th analyzer from network. I notofied Sales and they will work with SIT to sort out the number of licenses available. | Setup | HIS/LIS | License
ABL80 OSM | Printer will not print approximately one hour delay. Replacing printer and printer control board did not resolve the problem. Possible issue with CPU. Request that the analyzer be removed from the lab to complete service work off site. | Ordered CPU. | CCE944C03C327B91 | No | Printer will not print approximately one hour delay. Replacing printer and printer control board did not resolve the problem. Possible issue with CPU. Request that the analyzer be removed from the lab to complete service work off site. | Ordered CPU.
ABL90 |  |  | CCE91EF13C394092 | No |  |  | Defect | HW | Display | Locking mechanism for cable is defect | x
ABL90 |  | Afprøvet og undersøgt, kommunikation afprøvet med ASTM program og serielt X kabel, virker fint, sender alt hvad den skal!! | CCE959663C0077D2 | No |  | Tested and investigated, communication tested with ASTM program and serial X cable, works fine, sends everything it needs!! | No fault found | HIS/LIS | ABL9 works as intended |  | x
ABL90 | On arrival on-site verified that charge was 0% on status screen. | Replaced Battery option. Tightened power connection. | CCE915063C32D8FC | No | On arrival on-site verified that charge was 0% on status screen. | Replaced Battery option. Tightened power connection. | Defect | HW | Battery module |  | x
ABL90 | Battery not charging. Unit % charge at zero | Replaced battery. | CCE96AB03C1984C3 | No | Battery not charging. Unit % charge at zero | Replaced battery. | Defect | HW | Battery module |  | x
ABL800Basic |  |  | CCE975833C0D63F0 | No |  |  | Defect | Electrode | Membrane | Glu
ABL835 FLEX | This SWO is to document the warranty replacement of Lac and Glu electrodes. Lac sense decline rapidaly following mebrane change. Glu has measurement unstable on numerous patient samples. Have changed membranes multile times. Have performed zero adjust, decontamintion, rinse elctrode, clean electrode well. No help. | Replace Lac and Glu electrodes and one box membrane each. | CCE8D4B93C0242AB | No | This SWO is to document the warranty replacement of Lac and Glu electrodes. Lac sense decline rapidaly following mebrane change. Glu has measurement unstable on numerous patient samples. Have changed membranes multile times. Have performed zero adjust, decontamintion, rinse elctrode, clean electrode well. No help. | Replace Lac and Glu electrodes and one box membrane each. | Defect | Electrode | Membrane | Lac
AQT90 FLEX |  | Klammer entfernt | CCE905AE3C44D292 | No |  | Bracket removed
ABL80 CO-OX |  | Austausch des rechten Stand off für Frontcover | CCEA3A83B216B3F8 | No |  | Replacement of the right stand off for front cover
ABL90 |  |  | CCEA99D2B228748F | No |  |  | Defect | HW | SC PCB |  | x
ABL90 | Kontrolleret lækstrømsfejl på væskepakken. Kontrolleret impedansen for alle parameterne. | Udskiftet sensorkassetten. | CCEBC9F028384D6A | No | Leak on the liquid package. Controlled impedance for all the parameters. | Replaced the sensor cassette.
ABL835 FLEX | Error occurs during mixed venous sampling. Causing pO2 not to be displayed. | Disable air detection in the ini file. Sample were not avialable inorder to check. Customer will monitor and inform. | CCEA6917B226661E | No | Error occurs during mixed venous sampling. Causing pO2 not to be displayed. | Disable air detection in the ini file. Sample were not avialable inorder to check. Customer will monitor and inform. | Unknown | Air detected | Disable air detection
ABL825 FLEX | Found inlet nipple broken | Replaced broken inlet nipple, left spare and instructed how to replace.Leaking has stopped. | CCEA4298B214D2AD | No | Found inlet nipple broken | Replaced broken inlet nipple, left spare and instructed how to replace. Leaking has stopped. | Defect | Inlet module | Nipple kit | Broken
ABL80 | Level three is out too low. Calibration is failing. I found the Zero point too high again. | Replaced the Signal Data Pcb and performed the Zero point calibration. Replaced the solution pack for a different lot number. Old lot #203361 New# 203416. Also, replaced the cassette, but the PO2 woud not come in. I did noticed air in the line, cleared the line by tapping on the pump tubing. I noticed the Zero would not be stable. I checked with the techline and they recomment to replace the distribution PCB Ordered for Monday. | CCEB99A7286025B7 | No | Level three is out too low. Calibration is failing. I found the Zero point too high again. | Replaced the Signal Data Pcb and performed the Zero point calibration. Replaced the solution pack for a different lot number. Old lot #203361 New# 203416. Also, replaced the cassette, but the PO2 woud not come in. I did noticed air in the line, cleared the line by tapping on the pump tubing. I noticed the Zero would not be stable. I checked with the techline and they recomment to replace the distribution PCB Ordered for Monday.
ABL90 | After customer replaced an expiring SC, the new SC would not calibrate due to the errors listed above. Customer replaced SC twice and SP per tech support, but the same errors remain. Tested sensor impedance in service and PH and electrolytes are all very high. Dissassembled analyzer and inspected for bad connections damage. Found leak at waste manifold and replaced waste manifold and tubing. Rebooted - Problem remains. Analyzer had a microsoft C++ error and shut down by itself. When it rebooted, it passed all calibrations and QC. while running waste bloods, analyzer jumped into an internal (S9040) QC and failed PH PCO2, and all OXi with 581 errors. reran QC same problem. | Ran a calibration and an external QC with similar ranges for PH and PCO2. Tested OK. Reran S9040 tested OK. | CCEB9D5D284EAA69 | No | After customer replaced an expiring SC, the new SC would not calibrate due to the errors listed above. Customer replaced SC twice and SP per tech support, but the same errors remain. Tested sensor impedance in service and PH and electrolytes are all very high. Dissassembled analyzer and inspected for bad connections damage. Found leak at waste manifold and replaced waste manifold and tubing. Rebooted - Problem remains. Analyzer had a microsoft C++ error and shut down by itself. When it rebooted, it passed all calibrations and QC. while running waste bloods, analyzer jumped into an internal (S9040) QC and failed PH PCO2, and all OXi with 581 errors. reran QC same problem. | Ran a calibration and an external QC with similar ranges for PH and PCO2. Tested OK. Reran S9040 tested OK. | Defect | HW | Waste connector unit |  | x
ABL90 | Ramesh was running qc and he noticed the analyzer would not aspirate the qc solution and it would abort the qc sample. | Upon arrival, I noticed the sensor cassette life is expired. I advised the customer but they do not have any additional SC. I back-flush the analyzer so that instrument is clean and ready for the new SC. I showed Ramesh how to remove and install the inlet clip when running qc. This could be the cause for the qc being aborted. | CCEB9E6B281DA048 | No | Ramesh was running qc and he noticed the analyzer would not aspirate the qc solution and it would abort the qc sample. | Upon arrival, I noticed the sensor cassette life is expired. I advised the customer but they do not have any additional SC. I back-flush the analyzer so that instrument is clean and ready for the new SC. I showed Ramesh how to remove and install the inlet clip when running qc. This could be the cause for the qc being aborted. | User Error | Clip not removed - Ext QC | Email matt, no reply | Maybe Courtesy | x
ABL90 |  |  | CCEEA3021411DCD4 | No |  |  | Defect | SC | Lac | Ask for log files | x
ABL90 | Limpieza inversa del circuito. Realizacion de test de servicio. Reinstalacion de SP. Cal y qc. | Limpieza inversa del circuito. Realizacion de test de servicio. Reinstalacion de SP. Cal y qc. | CCEBE6EE284D7652 | No | Reverse cleaning of the circuit. Performing service tests. REinstalling SP. Cal and qc. | Reverse cleaning of the circuit. Performing service tests. REinstalling SP. Cal and qc. | Cleaned | Fluidic pathway |  |  | x
ABL800Basic |  |  | CCEBA7602842D350 | No |  |  | Defect | HW | Oxi Module | Pinch Valve
ABL825 FLEX | Examined Oxi Module. and Data Logs. | Replaced hemolyzer and lamp unit. Performed lamp Adj and complete Oxi test. Did service tHb cal. Verified operation of ABL825. | CCEBA670282114FB | No | Examined Oxi Module. and Data Logs. | Replaced hemolyzer and lamp unit. Performed lamp Adj and complete Oxi test. Did service tHb cal. Verified operation of ABL825. | Defect | HW | Oxi Module | Hemolyzer | tHb
ABL825 FLEX |  |  | CCEBE3262811EE73 | No |  |  | Defect | Inlet module | Inlet Gasket
ABL835 FLEX |  |  | CCEBE694282CB26F | Yes |  |  | Re classified RMED | not upSWO | Cancelled
ABL835 FLEX |  |  | CCEBE39A2803B26C | No |  |  | Setup | HIS/LIS | -
ABL825 FLEX |  |  | CCEBC37A283DD573 | No |  |  | Defect | ph/BG Module | Electrode connector | pCO2
ABL837 FLEX | OXI test failes, massive spillage of QC from hemolyzer | Removed modules to clean chassis of spillage. Replaced defective hemolyzer, complete OXI alignmenet, tHb cal in service. Multi 2pt Cals were required. | CCEB950E284D2C71 | No | OXI test failes, massive spillage of QC from hemolyzer | Removed modules to clean chassis of spillage. Replaced defective hemolyzer, complete OXI alignmenet, tHb cal in service. Multi 2pt Cals were required. | Defect | HW | Oxi Module | Hemolyzer | Leakage
ABL835 FLEX | GLU ELECTRODE REPLACEMENT REDUCED THE SYSTEM UNSTABLE ERROR. HOWEVER OTHER SYSTEMS TRANSFERED IN DID NOT REPRODUCE THE UNSTABLE ISSUE. | REPLACED THE EL MET PCB AND COMPLETED THE ZERO CALIBRATION. | CCEB9EC9282952D6 | No | GLU ELECTRODE REPLACEMENT REDUCED THE SYSTEM UNSTABLE ERROR. HOWEVER OTHER SYSTEMS TRANSFERED IN DID NOT REPRODUCE THE UNSTABLE ISSUE. | REPLACED THE EL MET PCB AND COMPLETED THE ZERO CALIBRATION. | Defect | El/Met Module | PCB
ABL80 | Level three of the internal control failing on the low end. | Replaced the Disctribution PCB Performed Zero calibration. The inlet was not detected Made sure all connection were in ok. Brought system to ready again-Had to replace the cassette to calibrate properly. | CCEF95498A49697A | No | Level three of the internal control failing on the low end. | Replaced the Disctribution PCB Performed Zero calibration. The inlet was not detected Made sure all connection were in ok. Brought system to ready again-Had to replace the cassette to calibrate properly.
ABL825 FLEX | REVIEWED DATA LOGS THAT SHOW IT WAS IN FACT OUT OF PEER GROUP UNTIL MARCH 3RD WHEN BOB ZULLO CHANGED THE MEMBRANE.AFTER THE MEMBRANE CHANGE UNTIL NOW THE PO2 HAS BEEN AT THEIR HISTORICAL NUMBER AND BACK ON PEER GROUP MEAN | WILL REVISIT THIS WITH PAULA BAPTISTA AFTER THE MARCH WDC COMES IN | CCEF60E88A2EE567 | Yes | REVIEWED DATA LOGS THAT SHOW IT WAS IN FACT OUT OF PEER GROUP UNTIL MARCH 3RD WHEN BOB ZULLO CHANGED THE MEMBRANE. AFTER THE MEMBRANE CHANGE UNTIL NOW THE PO2 HAS BEEN AT THEIR HISTORICAL NUMBER AND BACK ON PEER GROUP MEAN | WILL REVISIT THIS WITH PAULA BAPTISTA AFTER THE MARCH WDC COMES IN | Re classified RMED | Telephone support
ABL810 FLEX | pCO2 membranes have arrived. Put analyzer's orriginal pH BG module back in the analyzer. Rinse errors upon startup. Inspected pH BG air valve, it was jammed in open position. Encountered inlet calibration failure despite several attempts and service calibration. | Changed pCO2 membrane, replaced pH BG air valve and inlet probe. | CCEF6AE88A101982 | No | pCO2 membranes have arrived. Put analyzer's orriginal pH BG module back in the analyzer. Rinse errors upon startup. Inspected pH BG air valve, it was jammed in open position. Encountered inlet calibration failure despite several attempts and service calibration. | Changed pCO2 membrane, replaced pH BG air valve and inlet probe. | Defect | ph/BG Module | Valve, air
ABL817 FLEX |  |  | CCF1C353765BBBE8 | No |  |  | Defect | Basic unit | Bottle detection PCB
AQT90 FLEX |  |  | CCEFCAEF8A53C5BA | No
TCM400 | Modules 2 and 4 or not recognized. Swapped module with good units - problem does not transfer. | Ordered new PCB's | CCEF94A48A54FCE9 | No | Modules 2 and 4 or not recognized. Swapped module with good units - problem does not transfer. | Ordered new PCB's | Defect | HW | PC Module
TCM-COMBIM | customer had remambraned the sensor but culd not get the error message to clear | as per ASM nad sales i was instructed to replace the sensor. when unit was powered up the message communication error appeared. i re installed sensor modual but could not clear the error. unit is being shipped into London office for repair under RGA # Y14-44890E | CCEF73828A5F1A94 | No | customer had remambraned the sensor but culd not get the error message to clear | as per ASM nad sales i was instructed to replace the sensor. when unit was powered up the message communication error appeared. i re installed sensor modual but could not clear the error. unit is being shipped into London office for repair under RGA # Y14-44890E | Defect | HW | Combi M Module
TCM-COMBIM | customer had remambraned the sensor and could not clear message | As per ASM and sales i was instructed to install a new sensor, replaced sensor and tested | CCEF6E858A301A90 | No | customer had remambraned the sensor and could not clear message | As per ASM and sales i was instructed to install a new sensor, replaced sensor and tested | Defect | HW | Electrode
ABL90 | Returned with sensor interface PCB and CF card. | Replaced CF card and reloaded software. Impedance errors returned upon restart. Replaced Interface PCB. | CCF0B68E00222635 | Yes | Returned with sensor interface PCB and CF card. | Replaced CF card and reloaded software. Impedance errors returned upon restart. Replaced Interface PCB. | Re classified RMED | Double upSWO | See CCFDFDD4325F2
ABL90 | Jake Donald has performed numerous service calls with limited results. Myself and Jim Gibbons are assisting on this service call | Jake reloaded software and replaced the sensor interface PCB | CCF0DFDD004325F2 | No | Jake Donald has performed numerous service calls with limited results. Myself and Jim Gibbons are assisting on this service call | Jake reloaded software and replaced the sensor interface PCB | Defect | HW | SC PCB
ABL90 |  |  | CCF3664CEC19B0CE | No |  |  | Defect | SP | pO2 QC C | LEH: Get datalogs | x
ABL90 | I checked the log and it was getting LS error 1324. and 1343 | It asked to have everything replaced in its time but the issue was still there. I went in and ran all the fluid test and they all passed. at this time the analyzer passed all test and came up. | CCF0C72E000D0D08 | No | I checked the log and it was getting LS error 1324. and 1343 | It asked to have everything replaced in its time but the issue was still there. I went in and ran all the fluid test and they all passed. at this time the analyzer passed all test and came up. | No fault found | FT
ABL800Basic |  |  | CCF1038B005BD1E2 | No |  |  | Defect | DMS PC | SW ctw2k
ABL825 FLEX |  |  | CCF0FF19004099FA | No |  |  | Setup | HIS/LIS | Radiance connection
ABL827 FLEX | Checked voltages on power supply and they look good. Suspect power supply. | I replaced the power supply. | CCF084DA00034DD7 | No | Checked voltages on power supply and they look good. Suspect power supply. | I replaced the power supply. | Defect | Power Supply | Spectrum mismatch | Not Human blood
AQT90 FLEX |  |  | CCF0FB120021558F | No
AQT90 FLEX |  |  | CCF0D913004004B3 | No
TCM400 | Replaced interface PCB with no success. Ordered main computer PCB. | Ordered new PCB's | CCF1242A00393A9A | No | Replaced interface PCB with no success. Ordered main computer PCB. | Ordered new PCB's | Defect | HW | PCB
ABL90 |  | Austausch 905-885, 842-345, 841-797 | CCF1ED457635A0C0 | No |  | Replacement 905-885, 842-345, 841-797 | Defect | HW | Hemolyzer | spectrum mismatch | x
ABL90 | Revisit to install a new sensor cassette the customer purchased from Radiometer and check for qc being aborted by the analyzer. | Replaced the sensor cassette. After calibrations and internal qc passed, I ran external qualichekc qc and none were aborted. After speaking with the laboratory technologist, i found that he did not remove the inlet clip. Explained the customer that he has to remove the clip when running qc. Once he is done he needs to reinstall the clip on the inlet gasket. Showed the customer how to remove and install the inlet clip. Replaced a new inlet gasket and a connector gasket from PM kit. | CCF225CD761C6569 | Yes | Revisit to install a new sensor cassette the customer purchased from Radiometer and check for qc being aborted by the analyzer. | Replaced the sensor cassette. After calibrations and internal qc passed, I ran external qualichekc qc and none were aborted. After speaking with the laboratory technologist, i found that he did not remove the inlet clip. Explained the customer that he has to remove the clip when running qc. Once he is done he needs to reinstall the clip on the inlet gasket. Showed the customer how to remove and install the inlet clip. Replaced a new inlet gasket and a connector gasket from PM kit. | Re classified RMED | Courtesy
ABL90 |  |  | CCF203DD765A8E96 | No |  |  | Cleaned | Fluidic pathway | sO2 QC-B |  | x
ABL90 |  |  | CCF1F94C7632ACB1 | Yes |  |  | Re classified RMED | To collect data
ABL80 CO-OX | Setup and tested. | Ran cleaning pack; installed upgrades including v 1.60 software; tested oximetry system. Calibrated tHb. | CCF1CA977630DCB8 | No | Setup and tested. | Ran cleaning pack; installed upgrades including v 1.60 software; tested oximetry system. Calibrated tHb.
ABL800Basic |  | und Reinigung der Ventile in den Messmodulen; Austausch defektes Airventil im pH BG Modul; Elektrodenwartung durchgeführt; Autocheckmodul gereinigt; Kal1- und StHb Kal durchgeführt; Kal2- und QC Level 1,2 | CCF1BE917648D3BD | No |  | and cleaning of the valves in the measuring modules; Replacement of defective air valve in pH BG module; electrode maintenance performed; Autocheck module cleaned; Kal1 and StHb Kal; Kal2 and QC Level 1.2 | Defect | pH/BG Module | Valve, air
ABL825 FLEX |  |  | CCF4B82D623CDC24 | No |  |  | Defect | HW | Oxi Module | Lamp unit | Neon lamp
ABL825 FLEX | Sensitivity is 84.2%. When user went to change membrane the white plastic with the sealing ring stayed in the membrane jacket. | Replaced defective electrode.Multiple 2pt Cals, all OK. | CCF1EBFB7623D6C7 | No | Sensitivity is 84.2%. When user went to change membrane the white plastic with the sealing ring stayed in the membrane jacket. | Replaced defective electrode. Multiple 2pt Cals, all OK. | Defect | Electrode | Electrode | pCO2
TCM400 | Replaced computer PCB. Modules still not working. Additional PCB's to be ordered. | Ordered new PCB's | CCF26F9976122FA0 | No | Replaced computer PCB. Modules still not working. Additional PCB's to be ordered. | Ordered new PCB's | Defect | HW | PCB
TCM-COMBIM |  |  | CCF27958765E052C | No |  |  | Defect | HW | Display
ABL80 CO-OX |  |  | CCF35ACCEC1AB7F0 | No
ABL90 |  |  | CCF3205CEC42CD08 | No |  |  | Replaced | Inlet gasket |  |  | x
ABL90 | Linearity test requested by ASM | Ran linearity test | CCF32125EC3D19CD | Yes | Linearity test requested by ASM | Ran linearity test | Re classified RMED | Courtesy
ABL825 FLEX |  |  | CCF8018CC446A5CA | No |  |  | Defect | Electrode | Membrane | pCO2
ABL837 FLEX | Readjusted OXI and tHb cal in service, did not help. tBILwent out when Hemplyzer was changed Friday 3-28-14.results are well within package insert ranges. | Set all three ABLs to run with package insert for tBIL ranges for a month to establish new 2SDs | CCF32158EC1A9254 | No | Readjusted OXI and tHb cal in service, did not help. tBILwent out when Hemplyzer was changed Friday 3-28-14.results are well within package insert ranges. | Set all three ABLs to run with package insert for tBIL ranges for a month to establish new 2SDs | Adjustment | QC Limits
ABL825 FLEX | Checked inlet module | Replaced Tubing nipple kit for inlet module. Pump rotor and performed 6 month PM | CCF851E4C4241493 | No | Checked inlet module | Replaced Tubing nipple kit for inlet module. Pump rotor and performed 6 month PM | Defect | Inlet module | Nipple kit | Leaking
ABL825 FLEX | RINSE WAS PERFORMED AND THE WASTE TUBING WAS FED NTO A BEAKER. THERE WAS STILL LEAKING FROM THE INLET GASKET. | THE LEFT NIB FOR THE WASTE TUBING IN THE INLET MODULE WAS REPLACED, THE WASTE ROTOR WAS REPLACED AND A PM WAS PERFORMED. THIS WORK WAS DONE IN CONJUNCTION WITH CLAUDE L. | CCF34DC2EC150DE0 | No | RINSE WAS PERFORMED AND THE WASTE TUBING WAS FED NTO A BEAKER. THERE WAS STILL LEAKING FROM THE INLET GASKET. | THE LEFT NIB FOR THE WASTE TUBING IN THE INLET MODULE WAS REPLACED, THE WASTE ROTOR WAS REPLACED AND A PM WAS PERFORMED. THIS WORK WAS DONE IN CONJUNCTION WITH CLAUDE L. | Defect | Inlet module | Nipple kit
ABL820 FLEX | VOLTAGE CHECKS SHOW 24 VOLTS LOW INTERMITTENTLY AND PH BG AIR VALVES FAILS TO CYCLE EVEN AFTER REPLACEMENT. TECHLINE ADVISES REPLACEMENT OF DMS & PH BG MODULE. | REPLACED DMS, AU PCB & PH BG MODULE. AU PCB FAILS TO LOAD (ORIGINAL REINSTALLED) PH MODULE NOW NORMAL. SYSTEM LOCKUP AFTER CAL 1 BOTTLE REMOVED AND REPLACED FLUIDIC MODULE. | CCF3797FEC2DA0A1 | No | VOLTAGE CHECKS SHOW 24 VOLTS LOW INTERMITTENTLY AND PH BG AIR VALVES FAILS TO CYCLE EVEN AFTER REPLACEMENT. TECHLINE ADVISES REPLACEMENT OF DMS & PH BG MODULE. | REPLACED DMS, AU PCB & PH BG MODULE. AU PCB FAILS TO LOAD (ORIGINAL REINSTALLED) PH MODULE NOW NORMAL. SYSTEM LOCKUP AFTER CAL 1 BOTTLE REMOVED AND REPLACED FLUIDIC MODULE. | Unknown | Techline
ABL825 FLEX |  |  | CCF9C1BC3A15C00B | No |  |  | Defect | Electrode | Electrode | pCO2
ABL80 FLEX Basic version |  |  | CCF4912862096386 | No
ABL90 |  |  | CCF468D86250ACEE | No |  |  | Defect | SC | pH, pCO2 | HCO3 | x
ABL90 |  | durch den Akku läßt sich das Gerät nicht mehr Ausschalten, auch nicht durch Gedrückthalten des PC Knopfes Gerät aufgeschraubt und PC Spannungsversorgung abgezogen danach bootete das Gerät wieder SC gewechselt, da pCO2 nur noch 40% | CCF466BA623FC44D | No |  | due to the battery, the device can no longer be switched off, also not screwed up by holding down the PC button device and PC power supply disconnected after then the device booted again SC changed, since pCO2 only 40% | Adjustment | Disconnect power cable | Reboot analyzer |  | x
ABL90 |  | tausch der SC. | CCF45C0E621C9213 | No |  | exchange the SC. | Defect | SC | Sensitivity |  | x
ABL90 |  |  | CCF4966B62441814 | No |  |  | Defect | HW | CF card |  | x
ABL90 |  |  | CCF4649962498917 | Yes |  |  | Re classified RMED | Customer setup
ABL820 FLEX | FOUND ANALYZER CONTINUALLY RUNNING BOTH THE FLUIDIC PUMPS (RUNNING THE RINSE INTO THE WASTE AND ON THE COUNTER) AND DMS NOT CONECTING TO AU | REPLACED THE DMS TO ORIGINAL AND REPLACED INLET MODULE AND ONE OF THE BLANK MODULES. | CCF4A3316228A0A5 | No | FOUND ANALYZER CONTINUALLY RUNNING BOTH THE FLUIDIC PUMPS (RUNNING THE RINSE INTO THE WASTE AND ON THE COUNTER) AND DMS NOT CONECTING TO AU | REPLACED THE DMS TO ORIGINAL AND REPLACED INLET MODULE AND ONE OF THE BLANK MODULES. | Defect | DMS PC | revisit
ABL800Basic |  |  | CCF4AE4662195810 | No |  |  | User Error | Replacement process | Membrane | Reference
ABL810 FLEX | pCO2 has been constantly running far above the insert mean on level 3 QC. Was awaiting more rinse in shop. Membranes from previous call did not resolve issue. | Replaced rinse bottle and pCO2 electrode | CCF4537D625BB106 | No | pCO2 has been constantly running far above the insert mean on level 3 QC. Was awaiting more rinse in shop. Membranes from previous call did not resolve issue. | Replaced rinse bottle and pCO2 electrode | Defect | Electrode | Electrode | pCO2
ABL825 FLEX | New system, installed 30 days ago. pO2 electrode (Lot LB02) fails sense at 4.7. QC is within spec but error msg cal error present Cleaned electrode well, elec. tip - no help. installed a used electrode from customer and sense immediately passed at 11.1. Cleaned electrode tip with brush and tried again. Sense declined to 3.9 Reinstalled used electrode so customer can complete their validations. Peer analyzer seems okay (Lot LH03) sense=9.5 | Will order warranty replacement electrode | CCF46E14621452A2 | No | New system, installed 30 days ago. pO2 electrode (Lot LB02) fails sense at 4.7. QC is within spec but error msg cal error present Cleaned electrode well, elec. tip - no help. installed a used electrode from customer and sense immediately passed at 11.1. Cleaned electrode tip with brush and tried again. Sense declined to 3.9 Reinstalled used electrode so customer can complete their validations. Peer analyzer seems okay (Lot LH03) sense=9.5 | Will order warranty replacement electrode | Defect | Electrode | Electrode | pO2
AQT90 FLEX |  |  | CCF46C0A6202B07E | No
ABL820 FLEX | SAT MORNING FINDS INSTRUMENT GREEN BUT CHECK OF DATA LOGS FIND SEVERAL REBOOTS | REPLACED AU PCB AND LOADED ALL KEYS. | CCF5B1FCD81814E5 | No | SAT MORNING FINDS INSTRUMENT GREEN BUT CHECK OF DATA LOGS FIND SEVERAL REBOOTS | REPLACED AU PCB AND LOADED ALL KEYS. | Defect | Basic unit | Analyzer Control PCB
ABL80 CO-OX | Tried SP and ran all diag, everything is passing | Replaced SP with Client stock, Ordered new SP and SC for ASAP | CCF862C6C43531B0 | No | Tried SP and ran all diag, everything is passing | Replaced SP with Client stock, Ordered new SP and SC for ASAP
ABL90 | Finger for optical sensor binding against spring. | Realighned finger | CCF845F5C421D73D | No | Finger for optical sensor binding against spring. | Realighned finger | Adjustment | Sensor cassette door
ABL90 |  |  | CCF866A7C461DF22 | No |  |  | Defect | SW | Connection lost to Wet.sec |  | x
AQT90 FLEX | Se traslada equipo a Taller RIBE (se instala equipo de sustitución). De acuerdo a recomendaciones de Technical Support se sustituye unidad óptica. Calibración factor de referencia de la misma. Puesta en marcha. Calibración PCR y realización LQC PCR OK. Cal TnI Multicheck OK. Se pide calibradores Dimero y CKMB. 07 04 14. Calibración CKMB y Dimero-d OK. Se lleva equipo a cliente, puesta en marcha, instalación cartuchos restantes y realización LQC PCR, dimero y MC | Se traslada equipo a Taller RIBE (se instala equipo de sustitución). De acuerdo a recomendaciones de Technical Support se sustituye unidad óptica. | CCF84054C42D3243 | No | Equipment is moved to TALLER RIBE (replacement equipment is installed). According to Technical Support recommendations optical drive is replaced. Calibration reference factor of the same. Start-up. PCR calibration and LQC PCR OK performance. Cal TnI Multicheck OK. Dimero and CKMB calibrators are requested. 07 04 14. CALIBRATION CKMB and Dimero-d OK. Equipment is brought to the customer, commissioning, installation of remaining cartridges and realization LQC PCR, dimer and MC | Equipment is moved to TALLER RIBE (replacement equipment is installed). According to Technical Support recommendations optical drive is replaced.
ABL800Basic |  |  | CCF82DFDC449F0F2 | No |  |  | Defect | Electrode | Electrode | Lower Sealing ring
ABL815 FLEX |  |  | CCF883C5C43AA14B | No |  |  | Unknown | RiliBäk limits | Electrode | pCO2
ABL835 FLEX | ANalyzer is 10 days old. Initial pO2 sensitivity at installwas 11.0 Sense has since declined to 5.7 Customer has already cleaned electrode and electrode well (no blood has been run on this system) No help. | Exchanged new elctrode for the original. Warranty replacement. Lot LP-01 for LI-03 | CCF841FBC44452B7 | No | ANalyzer is 10 days old. Initial pO2 sensitivity at installwas 11.0 Sense has since declined to 5.7 Customer has already cleaned electrode and electrode well (no blood has been run on this system) No help. | Exchanged new elctrode for the original. Warranty replacement. Lot LP-01 for LI-03 | Defect | Electrode | Electrode | pO2
AQT90 FLEX | Blockade am Storage Wheel | beschädigte Testkassette TnI (LOT8852) erneuert | CCF86725C4036D21 | No | Blockade am Storage Wheel | damaged test cassette TnI (LOT8852) renewed
TCM400 | Modules 2 and 3 not functioning. | Replaced Comm PCB - All checks OK. Replaced original PCB's in all other locations. | CCF83A6AC4024781 | No | Modules 2 and 3 not functioning. | Replaced Comm PCB - All checks OK. Replaced original PCB's in all other locations. | Defect | HW | Communication PCB
TCM-COMBIM | Fehler wie oben reproduzierbar! Bei Quer-tausch der div. Module zum Werkstattgerät wanderte der Fehler mit dem Main CPU PCB mit! Main CPU PCB defekt | Main CPU PCB erneuert | CCF81343C4601A9C | No | Errors reproducible as above! In the case of cross-exchange of the various modules to the workshop device, the error with the Main CPU PCB moved along! Main CPU PCB defective | Main CPU PCB erneuert | Defect | HW | PCB
ABL80 CO-OX |  |  | CCF981353A23A33C | No
ABL80 CO-OX | Suspect bad SC | Replaced SC, and SP Issue resolved. | CCF9853B3A0231B5 | No | Suspect bad SC | Replaced SC, and SP Issue resolved.
ABL90 |  |  | CCF9928A3A38136E | No |  |  | Defect | HW | Oxi Module | Pinch Valve tube | Leakage
ABL90 | FOUND DOORS ON CLAMSHELL MISALIGNED SLIGHTLY AND WET SECTION BASE BROKEN AT RIGHTHAND ATTACH POINT. WASHER AT ATTACHING POINT SLIPPED AND CAUSED THE WET SECTION TO RISE WHEN SOLUTION PACK IS INSTALLED. | ORDERED NEW WETSECTION BASE. | CCF9DA593A2846E3 | No | FOUND DOORS ON CLAMSHELL MISALIGNED SLIGHTLY AND WET SECTION BASE BROKEN AT RIGHTHAND ATTACH POINT. WASHER AT ATTACHING POINT SLIPPED AND CAUSED THE WET SECTION TO RISE WHEN SOLUTION PACK IS INSTALLED. | ORDERED NEW WETSECTION BASE. | Unknown | Revisit Required | wetsection base broken | See later upswo | x
ABL90 |  |  | CCF944873A02904B | No |  |  | Defect | SP | Leak current |  | x
ABL90 |  |  | CCF98F693A1D447C | No |  |  | Defect | SP | pO2 QC |  | x
ABL90 | Liquid transport errors and inhommogenous errors in logs. | Tried flushing analyzer and replacing inlet gasket with holder. No help. Replaced solution pack. | CCF9C51D3A4B687A | No | Liquid transport errors and inhommogenous errors in logs. | Tried flushing analyzer and replacing inlet gasket with holder. No help. Replaced solution pack. | Defect | SP | Fluid transport |  | x
ABL90 |  |  | CCF989643A36BD66 | No |  |  | Defect | SC | Fluid transport |  | x
ABL90 |  |  | CCF9B0573A3AEA02 | No |  |  | Defect | SP | Pump cal error |  | x
ABL825 FLEX | SUSTITUCION DE PCB OXI POR JMC. AJUSTE COOXIMETRO Y PRUEBAS DE FUNCIONAMIENTO. |  | CCF9D8643A4EEBDD | No | REPLACEMENT OF OXI PCB BY JMC. COOXIMETRO SETTING AND OPERATION TESTS. |  | Defect | HW | Oxi Module | OXI PCB | Neon lamp
ABL825 FLEX | Comprobación de configuración del equipo, configuración de las derivas de calibración por estar estas mal configuradas, se calibra el equipo a 2 puntos, se verifica el funcionamiento final con Qc | Configuración de derivas de calibración | CCF9B7BF3A183E5F | No | Equipment configuration check, calibration drift configuration because these are misconfigured, 2-point equipment is calibrated, final operation is checked with Qc | Setting up calibration drifts | unknown | Setup | Cal limits | pH | LEH mail
ABL825 FLEX | Checked activity log for other errors - none. Checked QC log and did a trend report. QC for Ca++ has been going in and out of range since 4-1-14. The customer's range for level 1 is 1.01 to 1.05 mmol L. Removed electrodes and inspected both measuring chambers - both good. | Cleaned both chamber with cleaning solution. Membraned both the reference electrode and the Ca++ electrode. Ca++ came in at 1.02. | CCF992FE3A5B4596 | No | Checked activity log for other errors - none. Checked QC log and did a trend report. QC for Ca++ has been going in and out of range since 4-1-14. The customer's range for level 1 is 1.01 to 1.05 mmol L. Removed electrodes and inspected both measuring chambers - both good. | Cleaned both chamber with cleaning solution. Membraned both the reference electrode and the Ca++ electrode. Ca++ came in at 1.02. | Cleaned | El/Met Module | Measuring chamber | Ca++
ABL825 FLEX | Checked data logs for errors, checked valves and liquid sensors, checked electrodes and electrode chamber, | removed electrodes and chamber housing, straightened and repositioned bent grounding strap electrode housing, | CCF9E21B3A2920DF | No | Checked data logs for errors, checked valves and liquid sensors, checked electrodes and electrode chamber, | removed electrodes and chamber housing, straightened and repositioned bent grounding strap electrode housing, | Defect | Electrode | Electrode | pCO2
ABL835 FLEX |  |  | CCF9A7FC3A0C2E18 | No |  |  | Defect | Electrode | Electrode | pO2
TCM-COMBIM | Fehler am Anfang aufgetreten, mit Werk-stattsensor nicht! Danach wieder mit Kun-densensor, sowie Meldung: Sensor neu bespannt? bei Messung am anderen Ge-rät, Funktionstest Gerät mit Werkstattsensor o.k., | bitte Gerät zurück zum Kunden schicken, | CCF98F853A634973 | No | Error occurred at the beginning, with factory-instead sensor not! Then again with kun-densensor, as well as message: Sensor reloaded? when measuring on the other device, function test device with workshop sensor ok, | please send the device back to the customer, | Defect | Consumables | Electrode
ABL90 |  |  | CCFBD11F2614EC26 | No |  |  | Defect | HW | Display | Touch dont react | x
ABL90 | On arrival on site found that SC had been replaced at 0026 hrs and internal QC (all 3 levels) were failing lactate low and level 3 was also failing pO2 high. 2 levels of external QC were also failing lactate low. Calibrations did not help. Re-installing the SC did not help. | Replaced SP from customer stock. CCF faxed for credit on unused portion of SP. Did quick training for lab tech on troubleshooting. | CCFACE7CB03D2826 | No | On arrival on site found that SC had been replaced at 0026 hrs and internal QC (all 3 levels) were failing lactate low and level 3 was also failing pO2 high. 2 levels of external QC were also failing lactate low. Calibrations did not help. Re-installing the SC did not help. | Replaced SP from customer stock. CCF faxed for credit on unused portion of SP. Did quick training for lab tech on troubleshooting. | Defect | SP | pO2 QC C high | & Lac Int + ext C low | x
ABL90 |  |  | CCFAC912B0037B4D | No |  |  | Defect | SC | Lac
ABL90 |  |  | CCFABE45B035E238 | No |  |  | Unknown | Revisit Required | HbF QC | fault went away - see later upSWO | x
ABL80 OSM | I found the barcode scanner not scanning every time. | Replaced the barcode scanner Verified it was scanning. I packaged and shipped analyzer to Mike Smith. | CCFA8311B02B2107 | No | I found the barcode scanner not scanning every time. | Replaced the barcode scanner Verified it was scanning. I packaged and shipped analyzer to Mike Smith.
ABL825 FLEX | Confirmed pCO2 control values are outside peer groups reading. Noted several control reading failing Evaluation: Possible measuring chamber per tech support. Tests: 2 point calibration x 3 are passing. Level 3 is failing pCO2 to high. Replaced electrode. 2pt x5 passing. New controls had to be added. S7755 from last lot is empty. A.Abramov created this SWO for Bob Lancellotti. | Replaced Measuring chamber, pCO2 electrode. | CD014F56FE513F59 | No | Confirmed pCO2 control values are outside peer groups reading. Noted several control reading failing Evaluation: Possible measuring chamber per tech support. Tests: 2 point calibration x 3 are passing. Level 3 is failing pCO2 to high. Replaced electrode. 2pt x5 passing. New controls had to be added. S7755 from last lot is empty. A.Abramov created this SWO for Bob Lancellotti. | Replaced Measuring chamber, pCO2 electrode. | Defect | Electrode | Electrode | pCO2
ABL835 FLEX | The issues began after replacing the membrane on March 31. The QC and calibrations are okay, but about 25% of the patient samples have unstable Lactate results. User replaced the membrane again on March 7 and the number of errors has decreased, but still occurring about 10% of the time. | Unable to recreate the error. Since replacing the membrane decreased the number of errors, and a similar issue is occurring on another analyzer using the same lot number of membranes (Lot 2233), I suspect the issue may be the membranes. Membranes from a different lot number are due in later today and will be installed into the analyzer. | CCFAA927B04D1E5C | No | The issues began after replacing the membrane on March 31. The QC and calibrations are okay, but about 25% of the patient samples have unstable Lactate results. User replaced the membrane again on March 7 and the number of errors has decreased, but still occurring about 10% of the time. | Unable to recreate the error. Since replacing the membrane decreased the number of errors, and a similar issue is occurring on another analyzer using the same lot number of membranes (Lot 2233), I suspect the issue may be the membranes. Membranes from a different lot number are due in later today and will be installed into the analyzer. | Defect | Electrode | Membrane | Lac
ABL835 FLEX | The electrode often takes four hours or more before the sensitivity reaches a high enough level to pass calibrations. After the four to six hour startup period, the Lac works okay. | The temperatures appear to be okay. Performed a zero current calibration. The zero current value changed by nearly 36pA. User will monitor the analyzer and if the electrode continues to be slow to respond, then recommend replacing the electrode. | CCFAC1D7B054D6FB | No | The electrode often takes four hours or more before the sensitivity reaches a high enough level to pass calibrations. After the four to six hour startup period, the Lac works okay. | The temperatures appear to be okay. Performed a zero current calibration. The zero current value changed by nearly 36pA. User will monitor the analyzer and if the electrode continues to be slow to respond, then recommend replacing the electrode. | Defect | Electrode | Membrane | Lac | Startup
ABL827 FLEX |  |  | CCFAE41BB00807EE | No |  |  | Defect | Electrode | Electrode | Dummy | Crea
ABL820 FLEX | Unit will not boot up, gets a error after windows loads. Kira is not working. Dill files are corrupt | will need to reload SW but had to leave building due to fire. will return 4 10 | CCFC217C261CF726 | No | Unit will not boot up, gets a error after windows loads. Kira is not working. Dill files are corrupt | will need to reload SW but had to leave building due to fire. will return 4 10 | Defect | DMS PC | ??
ABL80 OSM | Analyzer was installed in October odf 2013. Analyzer is not being used to report patient results. Customer has not completed linearity and or comparison studies. Printer failed. Installing new printer assemblie did not resolve issue. Corrupt CF Card, Computer failure. | Replaced computer with new computerand base board assembly, 945 kit. Replaced CF Card. | CD1520A05E28162F | No | Analyzer was installed in October odf 2013. Analyzer is not being used to report patient results. Customer has not completed linearity and or comparison studies. Printer failed. Installing new printer assemblie did not resolve issue. Corrupt CF Card, Computer failure. | Replaced computer with new computerand base board assembly, 945 kit. Replaced CF Card.
ABL90 | NEW WET SECTION CHASSIS NOT APPLICABLE | NEW WET SECTION ON ORDER FOR FRIDAY APRIL 11TH | CCFC719926597DAD | Yes | NEW WET SECTION CHASSIS NOT APPLICABLE | NEW WET SECTION ON ORDER FOR FRIDAY APRIL 11TH | Re classified RMED | Same day/issue
ABL90 | RECIEVED ORDER FOR INSTALL AND PART NOT APPLICABLE | REORDERED WET SECTION COMPLETE FOR INSTALLATION ON 04 11 2014 | CD003BC6883461DD | Yes | RECIEVED ORDER FOR INSTALL AND PART NOT APPLICABLE | REORDERED WET SECTION COMPLETE FOR INSTALLATION ON 04 11 2014 | Re classified RMED | Same day/issue
ABL90 |  |  | CCFC2223263073FA | No |  |  | Adjustment | run clot detect remove in service |  | tHb + sO2 QC | x
ABL90 | Comprobar SP y SC limpiar el circuito neumático y módulo OXI, calibrar y probar con controles. | Limpiar el circuito neumático. | CCFBCF72262B5C54 | No | Check SP and SC clean pneumatic circuit and OXI module, calibrate and test with controls. | Clean the pneumatic circuit. | Cleaned | Fluidic pathway | pH/Oxi QC |  | x
ABL80 OSM |  |  | CCFC0266260AB473 | Yes |  |  | Re classified RMED | FAN / TU
ABL825 FLEX |  |  | CCFC48BF264038D5 | No |  |  | Defect | Electrode | Membrane | Lac
ABL827 FLEX | When booting up it is getting to the Starting Up screen and stopping. It is not loading Windows. I tried a power supply and it still stops at the Starting Up screen. | Power supply did not solve the issue. Analyzer is still down. | CCFBCCB22650BD67 | No | When booting up it is getting to the Starting Up screen and stopping. It is not loading Windows. I tried a power supply and it still stops at the Starting Up screen. | Power supply did not solve the issue. Analyzer is still down. | Unknown | Revisit Required
ABL800Basic |  | LP Schlauch erneuert | CCFBD1212644046B | No |  | LP hose renewed | Defect | Fluidic module | Solution pump tubing
ABL825 FLEX | Troubleshot using data logs. QC falls low on all levels but especially on level 1. Problem has been going on since April 1, 2014. | Replaced Calcium and Reference electrodes. Gave Customer box of Reference membranes as we had used over a box in trying to troubleshoot. Will monitor and open a Customer Quality Case. Replaced pump tubing for electrodes and solutions. | CCFC478E260123F4 | No | Troubleshot using data logs. QC falls low on all levels but especially on level 1. Problem has been going on since April 1, 2014. | Replaced Calcium and Reference electrodes. Gave Customer box of Reference membranes as we had used over a box in trying to troubleshoot. Will monitor and open a Customer Quality Case. Replaced pump tubing for electrodes and solutions. | Defect | Fluidic Module | Pump tubing
ABL820 FLEX | WILL RELOAD sW COMPLETE | reloaded windows and app SW. | CCFC28E02623F729 | Yes | WILL RELOAD sW COMPLETE | reloaded windows and app SW. | Re classified RMED | See CCFC217C261CF726
ABL810 FLEX |  |  | CCFC05592611134D | No |  |  | Defect | Inlet module | Probe bend
ABL90 | NONE NECESSARY | REPLACED WET SECTION AND COMPLETED THB CAL AND OVERHAULED UNIT. | CD003CE7885D61E0 | No | NONE NECESSARY | REPLACED WET SECTION AND COMPLETED THB CAL AND OVERHAULED UNIT. | Defect | HW | wetsection base | broken | x
ABL90 | THE FLUID SYSTEM WAS FLUSHED AND A RINSE WAS PERFORMED. ERRORS WERE STILL OCCURING. THE OXI UNIT WAS REMOVED AND THERE WAS QC BUILT UP ON TH4 WHITE DISK | A HEMOLYZER HAS BEEN ORDERED | CD0B9E7BAE4A2619 | No | THE FLUID SYSTEM WAS FLUSHED AND A RINSE WAS PERFORMED. ERRORS WERE STILL OCCURING. THE OXI UNIT WAS REMOVED AND THERE WAS QC BUILT UP ON TH4 WHITE DISK | A HEMOLYZER HAS BEEN ORDERED | Unknown | Revisit required | ??
ABL90 |  |  | CCFD26F79C2212A2 | Yes |  |  | Re classified RMED | SWU
ABL90 |  |  | CCFD1B5F9C587FB4 | No |  |  | Defect | HW | Spectrometer | HbF QC | x
ABL825 FLEX | Realización del mantenimiento preventivo. Limpieza de los tubos de desechos. Limpieza de cámara de medida. Chequeo de circuitos electrónicos y neumáticos. Chequeo de sensores de líquidos y calibración de bomba. Test de oxímetro. Limpieza del equipo. Calibración y control de calidad. |  | CCFD54489C126910 | Yes | Performing preventive maintenance. Cleaning the waste pipes. Cleaning of measuring chamber. Checking electronic circuits and tires. Liquid sensor check and pump calibration. Oximeter test. Cleaning the equipment. Calibration and quality control. |  | Re classified RMED | Preventive maintenance
ABL827 FLEX | I tried a power supply yesterday but did not solve the issue. After replacing the HD I got a boot disk failure. I found the 12DC was at 7 volts in the BIOS so I will replace the power supply also. Leaving HD with customer since I can't scramble the data. I broke the wire to the speaker so I will order one and return. | Replaced the SATA HD and power supply. Copied customer's logs from old HD to new one. | CCFD25FE9C02BD6A | No | I tried a power supply yesterday but did not solve the issue. After replacing the HD I got a boot disk failure. I found the 12DC was at 7 volts in the BIOS so I will replace the power supply also. Leaving HD with customer since I can't scramble the data. I broke the wire to the speaker so I will order one and return. | Replaced the SATA HD and power supply. Copied customer's logs from old HD to new one. | Defect | DMS PC | HDD
ABL800Basic |  |  | CCFD22979C39BBDC | No |  |  | Defect | Printer module | Printer head
ABL800Basic |  |  | CCFD1B5B9C515F43 | No |  |  | Defect | El/Met Module | Solution pump tubing
ABL90 |  |  | CD0467FC6038E102 | Yes |  |  | Re classified RMED | No info
ABL90 |  |  | CD03B04FEA258641 | No |  |  | Defect | HW | Oxi Module | Pinch Valve tube | High pO2 values
ABL90 | This is the second time in two days thay had called for this. | I tryed cleaning it and runnig the fluid transport programs and still it came back with clean the inlet. Unable to pump solutions and inhomogeneous rinse. I replaced the SP and it came in | CD0102ECFE34975E | No | This is the second time in two days thay had called for this. | I tryed cleaning it and runnig the fluid transport programs and still it came back with clean the inlet. Unable to pump solutions and inhomogeneous rinse. I replaced the SP and it came in | Defect | SP | Fluid transport | guided troubleshotting say clean inlet | x
ABL90 |  |  | CD0112D8FE4AA290 | Yes |  |  | Re classified RMED | not upSWO |  |  | LEH
TCM4 | Comprobar fuente de alimentación y batería, cambiar la batería, calibrar y probar. | Cambiar la batería. | CD01033AFE41ECD3 | No | Check power supply and battery, change battery, calibrate and test. | Change the battery. | Defect | HW | Battery module
ABL90 | Conexion con Openlab Configuracion como cliente Apertura de socket. Envio tramas. |  | CD01830AFE0A919E | Yes | Connection with Openlab Configuration as Socket Opening client. He sent plots. |  | Re classified RMED | Planned | Alex: Confirmed: It is not an upSWO. It was an error from the FSE.
ABL825 FLEX | replaced inlet gasket twice | Service dump obtained. | CD1B534CAC0AAE7D | No | replaced inlet gasket twice | Service dump obtained.
AQT90 FLEX | Comprobar la sección húmeda, cambiar SP, realizar una limpieza y probar con control. | Cambiar el SP. | CD015A12FE25ECDC | No | Check the wet section, change SP, perform a cleaning and test with control. | Change the SP.
AQT90 FLEX |  |  | CD0233087452E654 | No
ABL80 CO-OX | DESMONTAJE NPARCIAL DE CIRCUITO NEUMATICO Y DE DESECHOS, SE LIMPIA POR OBSTRUCCION Y SE PASA AGUA DESTILADA POR TODOS NLOS NCIRCUITOS COMPROBANDO SU FUNCIONAMIENTO. SE MONTAN NUEVOS REACTIVOS SE CALIBRA Y SE EFECTUAN PRUEBAS DE FUNCIONAMIENTO, SIENTO ESTAS CORRECTAS. | LIMPIEZA | CD02377574069A22 | No | NPARCIAL DISASSEMBLY OF NEUMATIC CIRCUIT AND DISCLOSURES, IS CLEANED BY OBSTRUCTION AND WATER IS PASSED DISTILLED BY ALL NCIRCUITOS CHECKING ITS OPERATION. NEW REACTIVES ARE MOUNTED ARE CALIBRATED AND OPERATION TESTS ARE EFFECTED, I FEEL THESE CORRECT. | Cleaning
ABL90 |  | Austausch 924-455 | CD02184B74398D10 | No |  | Replacement 924-455 | Replaced | Probe | Bend |  | x
ABL90 | The PCO2 sen would not come in after replacement | I replaced the SC and ran several cal's and then PCO2 came in. | CD020D2F745CEB7D | No | The PCO2 sen would not come in after replacement | I replaced the SC and ran several cal's and then PCO2 came in. | Defect | SC | pCO2 sens |  | x
ABL835 FLEX |  |  | CD024B7D741B2716 | No |  |  | Defect | Electrode | Electrode | pCO2
ABL825 FLEX | Customer has replaced membranes early to allow analyzer to pass calibrations and QC. Tested the electronics in service, all systems were within specification. | Replace electrode, failed to pass calibrations after re-installation with new membrane under warranty less than 1 year from install. Used membranes to troubleshoot the issue with calibrations. Set up NetOp remote support through the Radiance server. | CD026FFE74064811 | No | Customer has replaced membranes early to allow analyzer to pass calibrations and QC. Tested the electronics in service, all systems were within specification. | Replace electrode, failed to pass calibrations after re-installation with new membrane under warranty less than 1 year from install. Used membranes to troubleshoot the issue with calibrations. Set up NetOp remote support through the Radiance server. | Defect | Electrode | Electrode | Lac
AQT90 FLEX |  |  | CD03AA19EA2C8635 | No
Tosca-500 |  |  | CD03A671EA080872 | No |  |  | Defect | Consumables | Electrode
ABL80 CO-OX |  |  | CD04AF1E6001205A | No
ABL80 | Level 1 2 external QC are failing on the low end. Also, blood sample results are lower than other analyzer. | Replaced the manifold and the Sensor Cassette pcb. Also, performed the zero point calibration for Po2. | CD03F525EA591381 | No | Level 1 2 external QC are failing on the low end. Also, blood sample results are lower than other analyzer. | Replaced the manifold and the Sensor Cassette pcb. Also, performed the zero point calibration for Po2.
ABL90 |  |  | CD038F71EA3324F8 | Yes |  |  | Re classified RMED | PM
ABL90 |  |  | CD04D3086029513F | No |  |  | Defect | SC | Lac
ABL90 |  |  | CD14FAEB5E46168A | No
ABL825 FLEX | Examined Oxi module. | Replaced hemolyzer. Adj lamp unit and did complete Oxi Test. Did tHb calibration. Membrane the Glu electrode. Verified operation of ABL825. | CD039460EA1D667E | No | Examined Oxi module. | Replaced hemolyzer. Adj lamp unit and did complete Oxi Test. Did tHb calibration. Membrane the Glu electrode. Verified operation of ABL825. | Defect | HW | Oxi Module | Hemolyzer | phase reg
ABL837 FLEX | Performed a complete oximetry service test. Hemolyzer tuning voltage failed. Voltage was 44.0 and the limits are 10 - 40V. | Replaced the hemolyzer unit. QC failed glucose reading high. Remembraned the glucose electrode. | CD03A3BCEA01553E | No | Performed a complete oximetry service test. Hemolyzer tuning voltage failed. Voltage was 44.0 and the limits are 10 - 40V. | Replaced the hemolyzer unit. QC failed glucose reading high. Remembraned the glucose electrode. | Defect | HW | Oxi Module | Hemolyzer | phase reg
ABL825 FLEX | This SWO is to document the warranty repalcement of pO2 Electrode (Lot LB-02) with new electrode (Lot LP-01) This electrode failed within 30 days of new system installation. Previous work order B9503V3PW documents the initial call. | Customer replaced pO2 electrode | CD03B7D9EA166C47 | Yes | This SWO is to document the warranty repalcement of pO2 Electrode (Lot LB-02) with new electrode (Lot LP-01) This electrode failed within 30 days of new system installation. Previous work order B9503V3PW documents the initial call. | Customer replaced pO2 electrode | Re classified RMED | Warranty issue
ABL825 FLEX | This SWO documents warranty replacement of pO2 electrode to correct low sense. This analyzer was installed on 3-24-14. Began failing Cal's for sense on 3-30-14. Warranty electrode (Lot LP-01) to replace lot LI-01 | Replaced pO2 electrode | CD03B144EA216C40 | Yes | This SWO documents warranty replacement of pO2 electrode to correct low sense. This analyzer was installed on 3-24-14. Began failing Cal's for sense on 3-30-14. Warranty electrode (Lot LP-01) to replace lot LI-01 | Replaced pO2 electrode | Re classified RMED | Warranty issue
ABL825 FLEX | This SWO documents warranty replacement of pO2 electrode to correct low sense. This analyzer was installed on 3-24-14. Began failing Cal's for sense on 3-30-14. Warranty electrode (Lot LP-01) to replace lot LI-01 | Replaced pO2 electrode | CD03B3E8EA016C43 | Yes | This SWO documents warranty replacement of pO2 electrode to correct low sense. This analyzer was installed on 3-24-14. Began failing Cal's for sense on 3-30-14. Warranty electrode (Lot LP-01) to replace lot LI-01 | Replaced pO2 electrode | Re classified RMED | Warranty issue
ABL825 FLEX | Oxi Complete test passes but lamp is at 3900. Checked lamp, adjustment lock down screw is tight. Adjusted lamp to spec and tested okay. pCO2 is failing Cal's for sense. Tried rinsing electrode and cleaning well - no help. Replaced membrane. Also replaced failed pO2 Electrode on two systems (separate SWO) | Adjusted Lamp, multiple Oxi Complete tests, Service tHb Cal. Replaced pCO2 Membrane. Cleaned elec. well | CD038BF5EA486C3C | No | Oxi Complete test passes but lamp is at 3900. Checked lamp; adjustment lock down screw is tight. Adjusted lamp to spec and tested okay. pCO2 is failing Cal's for sense. Tried rinsing electrode and cleaning well - no help. Replaced membrane. Also replaced failed pO2 Electrode on two systems (separate SWO) | Adjusted Lamp, multiple Oxi Complete tests, Service tHb Cal. Replaced pCO2 Membrane. Cleaned elec. well | Adjustment | HW | Oxi Module | Lamp unit
TCM/Tosca |  |  | CD03BF59EA570881 | No |  |  | Defect | HW | Power Supply
ABL90 |  |  | CD050EEA600457F4 | No |  |  | Defect | SC | Glu/Lac | Cal/QC | x
ABL90 | Battery not charging. % charge at zero | Replaced battery. | CD04D865603919AE | No | Battery not charging. % charge at zero | Replaced battery. | Defect | HW | Battery module
ABL90 | Instrument screen goes dark and hot electrical smell is present when battery is in use. | Replaced battery and power supply. Calibrated and ran QC. Tested battery function multiple times over several hours without issue. Instrument operates normally. | CD04E5EE6060E78B | No | Instrument screen goes dark and hot electrical smell is present when battery is in use. | Replaced battery and power supply. Calibrated and ran QC. Tested battery function multiple times over several hours without issue. Instrument operates normally. | Defect | HW | Battery module | strong smell | x
ABL90 |  |  | CD04B6936006CD0F | No |  |  | Defect | HW | Hemolyzer | Leaking | x
ABL90 | THE SENSOR CASSETTE WAS REMOVED AND INSTALLED SEVERAL TIMES AS WELL AS TH EINSTRUMENT BEING SHUT DOWN | A WET CONTROL SECTION PCB AND ANALYZER ID PCB WERE ORDERED. THE WET SECTION BOARD HAD COME IN BUT THE ANALYZER ID PCB WAS BACK ORDERED FROM COPENHAGEN. THE WET SECTION PCB WAS REPLACED AND DURING THE SOFTWARE DOWNLOAD IT POSTED AN ERROR. THE OLD BOARD WAS REINSTALLED AND THE ERROR CLEARED ITSELF (1311). ALSO A LOANER INSTRUMENT WAS BROUGHT IN BUT GURJIT DID NOT WANT TO RUN CROSS MATCHED BECAUSE OF THE TIME REQUIRED. | CD1B6659AC2C1916 | No | THE SENSOR CASSETTE WAS REMOVED AND INSTALLED SEVERAL TIMES AS WELL AS TH EINSTRUMENT BEING SHUT DOWN | A WET CONTROL SECTION PCB AND ANALYZER ID PCB WERE ORDERED. THE WET SECTION BOARD HAD COME IN BUT THE ANALYZER ID PCB WAS BACK ORDERED FROM COPENHAGEN. THE WET SECTION PCB WAS REPLACED AND DURING THE SOFTWARE DOWNLOAD IT POSTED AN ERROR. THE OLD BOARD WAS REINSTALLED AND THE ERROR CLEARED ITSELF (1311). ALSO A LOANER INSTRUMENT WAS BROUGHT IN BUT GURJIT DID NOT WANT TO RUN CROSS MATCHED BECAUSE OF THE TIME REQUIRED.
ABL90 |  |  | CD04B484600781CD | No |  |  | Defect | SP | pouch pen bend |  | x
ABL90 |  |  | CD04B92D6028F0C4 | No |  |  | No fault found | leak current detected |  |  | x
ABL5 | Found customer was failing multiple parameters, but fluidics were functioning properly. Suspect is malfunctioning electrodes. | Remembraned Ref, PCO2 and PO2 electrodes. Verified instrument operation. | CD04D36E602CFFAD | No | Found customer was failing multiple parameters, but fluidics were functioning properly. Suspect is malfunctioning electrodes. | Remembraned Ref, PCO2 and PO2 electrodes. Verified instrument operation.
ABL800Basic |  |  | CD051023601C4CDF | No |  |  | Defect | Electrode | Membrane | Lac
ABL825 FLEX | Speaking with Larry, he stated the pCO2 and MetHb both were out on QC. They found the pCO2 electrode dry with no electrolyte in membrane. He stated every time there is an issue he needs to have one of his techs in the core lab go up to the OR and spend hours trying to resolve issues, therefore leaving him short staffed in the lab. Going over all the SWO's, numerous parts have been replaced on this analyzer. They are very upset with this analyzer. | As of this afternoon the analyzer was back up and running. They remembraned the pCO2 and ran a tHb cal for the MetHb. I will monitor the situation and look for a resolution that will return confidence to the customer. | CD0538E56009BEBD | No | Speaking with Larry, he stated the pCO2 and MetHb both were out on QC. They found the pCO2 electrode dry with no electrolyte in membrane. He stated every time there is an issue he needs to have one of his techs in the core lab go up to the OR and spend hours trying to resolve issues, therefore leaving him short staffed in the lab. Going over all the SWO's, numerous parts have been replaced on this analyzer. They are very upset with this analyzer. | As of this afternoon the analyzer was back up and running. They remembraned the pCO2 and ran a tHb cal for the MetHb. I will monitor the situation and look for a resolution that will return confidence to the customer. | Defect | Electrode | Membrane | pCO2
ABL825 FLEX |  |  | CD073BDB4C4747E0 | No |  |  | Defect | pH/BG Module | Valve, air
ABL817 FLEX |  |  | CD073DCB4C50DC18 | No |  |  | Defect | Solution | Crea
ABL835 FLEX | hologen lamp at 4000 | adjust lamp at 3200 full oxi test and thb call | CD08B2ACC213CD0A | No | hologen lamp at 4000 | adjust lamp at 3200 full oxi test and thb call | Adjustment | HW | Oxi Module | Lamp Unit
ABL80 | The fan is making a loud noise. | Replaced the Fan Performed a system cycle. Noise is gone. | CD0A14F138553992 | No | The fan is making a loud noise. | Replaced the Fan Performed a system cycle. Noise is gone.
ABL827 FLEX | I replaced the HD and broke the speaker wire. They alos said the autocheck fan was roaring sometimes. I did the annual also. See SWO B95H2YSGJ for annual. | Replaced the speaker for CPU and autocheck fan. | CD09A5BE3817B8D3 | No | I replaced the HD and broke the speaker wire. They alos said the autocheck fan was roaring sometimes. I did the annual also. See SWO B95H2YSGJ for annual. | Replaced the speaker for CPU and autocheck fan. | Defect | DMS PC | speaker | Wire broke at last upSWO
ABL825 FLEX | On arrival on-site verified error code message. Tried using the bak files as per the service manual and now error code is 5. Tried reloading windows software x 4 and it will not load gets to Preparing the disk and stops. | No correction can be made. Possible hard drive failure. | CD0A277C38106561 | No | On arrival on-site verified error code message. Tried using the bak files as per the service manual and now error code is 5. Tried reloading windows software x 4 and it will not load gets to Preparing the disk and stops. | No correction can be made. Possible hard drive failure. | Unknown | DMS PC | HDD
ABL825 FLEX | Inspected the tubing organizer and observed that the bottom of it is broken. | Replaced the tubing organize. Spent some time training on maintenance. | CD0A5F78380F9E33 | No | Inspected the tubing organizer and observed that the bottom of it is broken. | Replaced the tubing organize. Spent some time training on maintenance. | Defect | Basic unit | Y piece
ABL825 FLEX | Tried laser alignment in service-No change | Replaced auto check flex cable Alignment in service | CD09C11B38471C29 | No | Tried laser alignment in service-No change | Replaced auto check flex cable Alignment in service | Defect | Auto Check Module | FPC Cable
ABL80 CO-OX |  | Gerät aus ein, LP ersetzt daten für TS erhoben | CD0B14BDAE0B01BF | No |  | Device from on, LP replaces data collected for TS
ABL80 FLEX Basic version | Instalación del equipo, instalación del reactivo, instalación del papel de la impresora, configuración del equipo, formación de manejo y funcionamiento del equipo a usuarios, verificación final del correcto funcionamiento. |  | CD0B9BF3AE5B6A6C | No | Installation of the equipment, installation of the reagent, installation of the printer paper, configuration of the equipment, training of handling and operation of the equipment to users, final verification of proper operation.
ABL90 |  |  | CD0B247CAE1C079D | No |  |  | Cleaned | Fluidic pathway
ABL90 | THE HEMOLYZER WAS REPLACED | THE HEMOLYZER WAS REPLACED AND LEAK ERRORS WERE STILL OCCURRING. THE SOFTWARE WAS ASKING FOR THE SOLUTION PACK, AND GASKETS TO BE CLEANED. WHILE CLEANING THE GASKETS CINDY CAME BY AND STATED THAT THEY WOULD LIKE A NEW ANALYZER BECAUSE THIS ONE WAS BRAND NEW AND THERE WERE ISSUES WITH IT ALREADY . DARCY AND ANAND WERE CONTACTED AND INFORMED OF THE CONVERSATION. NEW GASKETS HAVE BEEN ORDERED BECAUSE THE CUSTOMER HAS NO SUPPLIES. | CD0B9FEEAE01261C | No | THE HEMOLYZER WAS REPLACED | THE HEMOLYZER WAS REPLACED AND LEAK ERRORS WERE STILL OCCURRING. THE SOFTWARE WAS ASKING FOR THE SOLUTION PACK, AND GASKETS TO BE CLEANED. WHILE CLEANING THE GASKETS CINDY CAME BY AND STATED THAT THEY WOULD LIKE A NEW ANALYZER BECAUSE THIS ONE WAS BRAND NEW AND THERE WERE ISSUES WITH IT ALREADY . DARCY AND ANAND WERE CONTACTED AND INFORMED OF THE CONVERSATION. NEW GASKETS HAVE BEEN ORDERED BECAUSE THE CUSTOMER HAS NO SUPPLIES. | Unknown | Revisit required | ??
ABL90 | 521 inhomigeneous sample and when geven a value it is aroude 100 mmHg's less then the other analyzer | I tryed back flushing and i checked all parts and still the same. I replaced the SC | CD0AFDEBAE2E3C5F | No | 521 inhomigeneous sample and when geven a value it is aroude 100 mmHg's less then the other analyzer | I tryed back flushing and i checked all parts and still the same. I replaced the SC | Unknown | Revisit Required | Cleaned Fluidic pathway
ABL90 | The pCO2 and K would not pass the cal | I tried flushing and nothing would help until I replaced the SC | CD0B157FAE2C3C62 | No | The pCO2 and K would not pass the cal | I tried flushing and nothing would help until I replaced the SC | Defect | SC | pCO2 and K+
ABL90 |  |  | CD0F00B21052F7D4 | No
ABL90 |  |  | CD0C348824572212 | No |  |  | Defect | SP | cannot pump
ABL90 | Checked for damage of solution pack. | Reseated solution pack | CD0AE37CAE5D81F2 | No | Checked for damage of solution pack. | Reseated solution pack | No fault found | Reseated SP
ABL90 | Limpieza del circuito. Limpieza de junta. Reinstalacion de SP Varias cal, qc. | Limpieza del circuito. Limpieza de junta. Reinstalacion de SP Varias cal, qc. | CD0B095FAE1DD2F3 | No | Cleaning the circuit. Gasket cleaning. Reinstalling SP Various lime, qc. | Cleaning the circuit. Gasket cleaning. Reinstalling SP Various lime, qc. | Cleaned | Inlet gasket |  |  | x
ABL80 CO-OX |  | Austausch der O-Ringe am Abfallstutzen; Gerät gereinigt; Austausch des Lösungspacks | CD0C702924190AF9 | No |  | Replacement of the O-rings on the waste nozzle; Device cleaned; Replacing the solution pack
ABL80 CO-OX | pO2 question on some blood samples | Ran zero calibration. Set refererence ranges to match other ABL80 on site | CD0CB2B2242EB2AE | No | pO2 question on some blood samples | Ran zero calibration. Set refererence ranges to match other ABL80 on site
ABL80 FLEX Basic version | Actualización de software a la versión 3.11.ES y entregado manual al cliente. Mantenimiento preventivo. Descontaminación según procedimiento técnico, revisión de histórico, limpieza de circuitos de deshecho y soluciones, puerta de entrada de muestras, offset del cero electrónico, estado de la batería. Calibración y ciclo de calidad. |  | CD0C45A82401FBC5 | Yes | Software update to version 3.11.ES and delivered manually to the customer. Preventive maintenance. Decontamination according to technical procedure, historical review, cleaning of waste circuits and solutions, sample entry door, electronic zero offset, battery status. Calibration and quality cycle. |  | Re classified RMED | FAN / TU
ABL80 FLEX Basic version | Se sustituyo signal data y se ajusto el equipo segun protocolo del fabricante. se comprobo correcto funcionamiento mediante QC. |  | CD12768D724D2B65 | No | Signal data is replaced and the equipment is adjusted according to the manufacturer's protocol. it is properly compensated for QC operation.
ABL90 |  |  | CD1C794D225CE1A7 | No
ABL90 | The analyzer kept asking for the inlet gsk to be cleaned or replaced. | I talked with the customer and it was stuck in the loop and would not go passed. I was on my way and they called back and it cleared. This is a complaint because this customer has called for this issue many times with both sites they have. I have been there many times and had to replace SC's and SP's. I did not go on site today. | CD0C4BA224178C25 | Yes | The analyzer kept asking for the inlet gsk to be cleaned or replaced. | I talked with the customer and it was stuck in the loop and would not go passed. I was on my way and they called back and it cleared. This is a complaint because this customer has called for this issue many times with both sites they have. I have been there many times and had to replace SC's and SP's. I did not go on site today. | Re classified RMED | No visit |  |  | LEH
ABL90 |  |  | CD0C2834242A4CCC | No |  |  | No fault found | Lac
ABL90 |  |  | CD0C70E92443221A | No |  |  | User Error | SP install
ABL825 FLEX | CUSTOMER IS NOT HAPPY WITH THE ANLYZER SINCE ITS INSTALLATION IN AUGUST OF LAST YEAR.PULLED DATA LOGS AND SERVICE DUMP.ANALYZER IS RUNNING AT 13.6 0N LEVEL3 QC S7755 LOT 507 PEER1 11.6 TO 13.7 AND PEER 2 11.4 TO 13.5 | ESCALATE THE CALL TO RMED AND OPEN A TECHLINE CASE FOR FURTHER INVESTIGATION.REPLACED THE DUMMY ELECTRODE,PH SEALING RING AND PCO2 SEALING RING.REPLACED THE PCO2 MEMBRANE WITH A NEW LOT A NEW LOT.QC BACK DOWN TO 12.6 | CD0C5458245EF0BE | No | CUSTOMER IS NOT HAPPY WITH THE ANLYZER SINCE ITS INSTALLATION IN AUGUST OF LAST YEAR. PULLED DATA LOGS AND SERVICE DUMP. ANALYZER IS RUNNING AT 13.6 0N LEVEL3 QC S7755 LOT 507 PEER1 11.6 TO 13.7 AND PEER 2 11.4 TO 13.5 | ESCALATE THE CALL TO RMED AND OPEN A TECHLINE CASE FOR FURTHER INVESTIGATION. REPLACED THE DUMMY ELECTRODE,PH SEALING RING AND PCO2 SEALING RING. REPLACED THE PCO2 MEMBRANE WITH A NEW LOT A NEW LOT. QC BACK DOWN TO 12.6 | Defect | Electrode | Electrode | pCO2
ABL837 FLEX | The analyzer screen is black and it will not start up. | Replaced a new power supply but the analyzer is still not powering up. suspect a faulty CPU unit. | CD0C8700241624E3 | No | The analyzer screen is black and it will not start up. | Replaced a new power supply but the analyzer is still not powering up. suspect a faulty CPU unit. | Defect | DMS PC | CPU | revisit required
ABL817 FLEX |  |  | CD0C362D240B2216 | No |  |  | Defect | HW | Inlet module | Leaking
ABL825 FLEX | On arrival verified that analyzer was still in the windows software load screen stuck at preparing disk. | Replaced SATA hard drive, reloaded Windows and ABl software. Re-installed customer set-up. Re-installed QC solutions and schedule. Reinstalled analyzer Radiance information. | CD0C57F5245ACDD1 | No | On arrival verified that analyzer was still in the windows software load screen stuck at preparing disk. | Replaced SATA hard drive, reloaded Windows and ABl software. Re-installed customer set-up. Re-installed QC solutions and schedule. Reinstalled analyzer Radiance information. | Defect | DMS PC | HDD
ABL827 FLEX | Inspected all tubings and lines. | I flushed out the waste line | CD0C2437242C2A7F | No | Inspected all tubings and lines. | I flushed out the waste line | Cleaned | Fluidic pathway | leaking from inlet
ABL810 FLEX | QC within insert limits but either much lower than insert mean or just above. Performed gas module check and inspected sample path. | Techline started. | CD0C28ED2421829B | No | QC within insert limits but either much lower than insert mean or just above. Performed gas module check and inspected sample path. | Techline started. | Defect | Electrode | Electrode | pCO2
ABL825 FLEX | Comprobación del funcionamiento del autocheck, realización de varios QC´s, comprobación por servicio, se ajusta el posicionamiento del carrusel del QC por servicio, se comprueba el correcto funcionamiento, se sustituye la membrana de Ref. | Ajuste del autocheck | CD0C488A241D0614 | No | Checking the operation of the autocheck, performing several QC's, checking by service, adjusting the positioning of the QC carousel by service, checking the correct operation, replacing the ref membrane. | Adjusting the autocheck | Adjustment | Auto Check Module | ??
AQT90 FLEX |  |  | CD0C896E244EB155 | No
ABL90 |  |  | CD0DFA379A3750A8 | Yes |  |  | Re classified RMED | Logistic upSWO | SP
ABL90 | SYSTEM ON POLE WHICH MAKES SOLUTION PACK MORE INSTALLATION MORE DIFFICULT | NONE NECESSARY | CD0EA0B3101823BB | No | SYSTEM ON POLE WHICH MAKES SOLUTION PACK MORE INSTALLATION MORE DIFFICULT | NONE NECESSARY | User Error | SP install
ABL90 |  |  | CD0ED0C5103B2B44 | No |  |  | Replaced | Probe | Bend
ABL90 |  |  | CD0EE0BA10470B94 | No |  |  | Defect | HW | Wet sec PCB
ABL90 |  |  | CD0ECACC1007044C | Yes |  |  | Re classified RMED | See CDEEBA147B94
ABL90 |  |  | CD0ECD7A100408CB | Yes |  |  | Re classified RMED | See CDEEBA147B94
ABL90 |  |  | CD4B16723042BDEB | No
ABL90 | The calibrations and adjustments wee run in Service - Fluid Transport section. There were errors posted in all sections. | The inlet gasket with holder and the inlet connector gasket were replaced because the they were suspect in the transport of fluid. The calibrations were performed in the service mode for the liquid sensors, pump cal, pressure tests also a OXI test was performed with no issues to report. Spare parts have been left behind as the customer does not have any. | CD0D776F9A3DB954 | No | The calibrations and adjustments wee run in Service - Fluid Transport section. There were errors posted in all sections. | The inlet gasket with holder and the inlet connector gasket were replaced because the they were suspect in the transport of fluid. The calibrations were performed in the service mode for the liquid sensors, pump cal, pressure tests also a OXI test was performed with no issues to report. Spare parts have been left behind as the customer does not have any. | Replaced | Inlet gasket | #NAME?
ABL90 | none | Replaced inlet connector, gasket holder ,SC and SP | CD1BA825AC299A51 | No | none | Replaced inlet connector, gasket holder ,SC and SP
ABL90 |  |  | CD0DAD809A2B56E4 | No |  |  | Cleaned | Fluidic pathway
ABL90 | The Po2 was showing about 100 mmgs less then the others. | I flushed the analyzer many times and then ran cal's and QC. The blood samples are runnig closer now. | CD0D93B19A4F8C27 | No | The Po2 was showing about 100 mmgs less then the others. | I flushed the analyzer many times and then ran cal's and QC. The blood samples are runnig closer now. | Cleaned | Fluidic pathway | pO2
ABL90 | The analyzer keeps asking to have the linlet cleaned or replaced | I cleaned the inlet and checked all other tubes. I first replaced the SP and still the error then I replaced the SC and now it is gone. I wated until the cal's came in | CD0DAB7E9A448C2A | No | The analyzer keeps asking to have the linlet cleaned or replaced | I cleaned the inlet and checked all other tubes. I first replaced the SP and still the error then I replaced the SC and now it is gone. I wated until the cal's came in | Defect | SC | Fluid transport
ABL90 | I inspected the inlet gasket and probe for damage and none was found. | I replaced the inlet probe. | CD0D79F69A5AC638 | No | I inspected the inlet gasket and probe for damage and none was found. | I replaced the inlet probe. | Defect | HW | Probe
ABL835 FLEX |  |  | CD0EDBF4103C584E | Yes |  |  | Re classified RMED | Courtesy | User Training
ABL837 FLEX | Revisit to replace the CPU unit due to the instrument not starting up and presenting a black screen. | Replaced the CPU unit. Found lactate presenting error 379 Calibration unstable-response fault. Cleaned the lactate electrode and replaced a new membrane. cCrea membranes and the cleaning solution bottle due for replacement, replaced for the customer. | CD0DE59B9A374BAE | No | Revisit to replace the CPU unit due to the instrument not starting up and presenting a black screen. | Replaced the CPU unit. Found lactate presenting error 379 Calibration unstable-response fault. Cleaned the lactate electrode and replaced a new membrane. cCrea membranes and the cleaning solution bottle due for replacement, replaced for the customer. | Defect | DMS PC | CPU
ABL825 FLEX | Call from last night. System does not recognize Oxi PCB. Tried reboot, reload AU firmware - no help. Update type was no help and then system booted up as an ABL805. Checked cables, fueses etc. No help. Unable to run Oxi test as system does not recognize Oximetry. Also, when done I ran 3 blood checks to compare MetHb between peers. Sample 1 was out on analyzer 1 (-1.9) This system #2 compared favorably on all 3 samples. | Replaced Oxi PCB, update type (2X) | CD0DA6919A5644A3 | No | Call from last night. System does not recognize Oxi PCB. Tried reboot, reload AU firmware - no help. Update type was no help and then system booted up as an ABL805. Checked cables, fueses etc. No help. Unable to run Oxi test as system does not recognize Oximetry. Also, when done I ran 3 blood checks to compare MetHb between peers. Sample 1 was out on analyzer 1 (-1.9) This system #2 compared favorably on all 3 samples. | Replaced oxygen PCB, update type (2X) | Defect | HW | Oxi Module | OXI PCB
ABL80 FLEX Basic version | Revisión del equipo. Chequeo del estado de la cassette de electrodos. Limpieza de la tubería. Realización del varias calibraciones pero pese a todo no entra el Lac. por lo que se procede a la sustitución de la cassette de electrodos. Revisión y comprobación del equipo con QC´s. | Cambio de la cassette de electrodos. | CD0EDE021063A5B9 | No | Equipment review. Checking the status of the electrode cassette. Cleaning the pipe. Performing the various calibrations but nevertheless does not enter the Lac. so the electrode cassette is replaced. Review and check the equipment with QC's. | Changing the electrode cassette.
ABL810 FLEX |  |  | CD0EFDFE10491140 | No |  |  | Defect | pH/BG Module | Valve, air
ABL820 FLEX | Checked data logs. Several mismatch errors present. | Checked valves. Cleaned analyzer. Performed oxi alignment. | CD0EECF81035D525 | No | Checked data logs. Several mismatch errors present. | Checked valves. Cleaned analyzer. Performed oxi alignment. | Adjustment | HW | Oxi Module | Alignment
ABL820 FLEX | Athena reported to me that she suspects the Gas 2 is leaking out. The analyzer is presenting the message that the Gas 2 has a low pressure and it show as empty on the screen. | Inspect the gas flow in service mode. Confirmed Gas 2 is leaking. Replaced the Gas 2 flow regulator. Replaced the Pressure Transducer with Chassis s unit. Replaced the air valve in the pH BG module. | CD0ED279101F18AE | No | Athena reported to me that she suspects the Gas 2 is leaking out. The analyzer is presenting the message that the Gas 2 has a low pressure and it show as empty on the screen. | Inspect the gas flow in service mode. Confirmed Gas 2 is leaking. Replaced the Gas 2 flow regulator. Replaced the Pressure Transducer with Chassis s unit. Replaced the air valve in the pH BG module. | Defect | Gas module | Pressure transducer
ABL825 FLEX | 1) This has been an ongoing issue that appeared to be resolved about two months ago and now the analyzer has started running additional QC for no apparent reason. 2) Remote support key had been installed, but the analyzer needs to have the software modified in order for the program to function properly. | 1) Gathered data and performed a service dump to submit to the engineers. 2) Modified the setup to enable remote support to function. | CD0F2F1D101D0841 | No | 1) This has been an ongoing issue that appeared to be resolved about two months ago and now the analyzer has started running additional QC for no apparent reason. 2) Remote support key had been installed, but the analyzer needs to have the software modified in order for the program to function properly. | 1) Gathered data and performed a service dump to submit to the engineers. 2) Modified the setup to enable remote support to function. | Setup | QC schedule | Double QC performed
ABL825 FLEX | 1) This has been an ongoing issue that appeared to be resolved about two months ago and now the analyzer has started running additional QC for no apparent reason. 2) Remote support key had been installed, but the analyzer needs to have the software modified in order for the program to function properly. 3) Analyzer connected properly and shows connection, but the analyzer is missing in Radiance. | 1) Gathered data and performed a service dump to submit to the engineers. 2) Modified the setup to enable remote support to function. 3) Disconnected and reconnected Radiance communication and the analyzer now shows up in Radiance. | CD0F323110520845 | No | 1) This has been an ongoing issue that appeared to be resolved about two months ago and now the analyzer has started running additional QC for no apparent reason. 2) Remote support key had been installed, but the analyzer needs to have the software modified in order for the program to function properly. 3) Analyzer connected properly and shows connection, but the analyzer is missing in Radiance. | 1) Gathered data and performed a service dump to submit to the engineers. 2) Modified the setup to enable remote support to function. 3) Disconnected and reconnected Radiance communication and the analyzer now shows up in Radiance. | Setup | QC schedule | Double QC performed
ABL825 FLEX |  |  | CD0EB3551033B62D | No
ABL80 CO-OX | siehe ausgeführte Arbeiten | siehe ausgeführte Arbeiten | CD117D3BFC41C01D | No | see work carried out | see work carried out
AQT90 FLEX |  |  | CD1277E272639415 | No
ABL90 |  |  | CD12D75E724B3B32 | Yes |  |  | Re classified RMED | Logistic upSWO | SC backlog
ABL90 | the analyzer in hold due to unable to pump solution. I have hearded about or been here about this analyzer. | I replaced all tubings and put on the new style inlet probe and GSK. and still given errors so I replaced the SP. Analyzer is still asking for inlet GSK I will order a bay unit and return. | CD1286367220207C | No | the analyzer in hold due to unable to pump solution. I have hearded about or been here about this analyzer. | I replaced all tubings and put on the new style inlet probe and GSK. and still given errors so I replaced the SP. Analyzer is still asking for inlet GSK I will order a bay unit and return. | Unknown | Revisit required
ABL90 |  |  | CD1276ED72469411 | No |  |  | Replaced | Probe | Clot
ABL90 |  |  | CD12D40972418C16 | No |  |  | Unknown | Revisit required
ABL825 FLEX | Confirmed printer inoperative. | Removed and replaced printer. | CD131CF6723BE562 | No | Confirmed printer inoperative. | Removed and replaced printer. | Setup | QC schedule | Double QC performed
ABL835 FLEX | ACTIVITY LOG SHOWS 675 INLET POSITION CALCULATION ERRORS.SERVIVE PROGRAM FOR INLET CALIBRATION FAILS. | REPLACED BENT PROBE.REPLACED FRICTION DETECTOR UNIT.SUCCESSFULLY PERFORMED INLET CALIBRATION.PERFORMED 2 POINT CALIBRATION. | CD12A15A721AC204 | No | ACTIVITY LOG SHOWS 675 INLET POSITION CALCULATION ERRORS. SERVIVE PROGRAM FOR INLET CALIBRATION FAILS. | REPLACED BENT PROBE. REPLACED FRICTION DETECTOR UNIT. SUCCESSFULLY PERFORMED INLET CALIBRATION. PERFORMED 2 POINT CALIBRATION. | Setup | QC schedule | Double QC performed
TCM-COMBIM | Schwenkfuß ist stark verbogen und hat das Gehäusevorderteil verkratzt, | Schwenkfuß und Gehäusevorderteil erneuert, | CD12B5B7721A0759 | No | Swivel foot is strongly bent and has scratched the front part of the housing, | Swivel foot and front part of the housing renewed, | User Error | HW | Broken
ABL90 | Software. | Install software upgrade version 3.0 MR2. | CD146B1FE81AB959 | No | Software. | Install software upgrade version 3.0 MR2.
ABL90 |  | Reinigung, Austausch 842-345 | CD13CD6BE84FD90A | No |  | Cleaning, replacement 842-345 | Defect | HW | Waste tube
ABL90 | The analyzer would keep telling them to replace the inlet GSK | I replaced the Sensor bay and Sensor | CD13C9FCE8240C07 | No | The analyzer would keep telling them to replace the inlet GSK | I replaced the Sensor bay and Sensor | Defect | HW | Sensor Cassette retainer
ABL90 |  |  | CD13B1A3E8118C1A | No |  |  | Defect | HW | CPU | USB
ABL90 | Checked the inlet gasket for clots | reseated solution pak. | CD13B6CFE85173FD | No | Checked the inlet gasket for clots | reseated solution pak. | No fault found | Solution activation?
ABL825 FLEX |  |  | CD158B595E10EAE1 | No |  |  | Adjustment | Auto Check Module | Lid | Stepper motor
ABL837 FLEX | In service, zero current was incorrect. | Zeroed all amplifiers. performed multiple cals, all ok. | CD1422F0E8082025 | No | In service, zero current was incorrect. | Zeroed all amplifiers. performed multiple cals, all ok. | Adjustment | pH/BG Module | Zero current | pO2
AQT90 FLEX |  |  | CD144275E80A741D | No
Tosca-500 |  |  | CD153E615E1FDC85 | No |  |  | Defect | HW | Gas Module
ABL90 |  |  | CD157B5A5E59B18F | No |  |  | Defect | SC
ABL90 |  | ID PCb getauscht. | CD154A485E55A09F | No |  | ID PCb getauscht. | Defect | SC
ABL90 |  |  | CD150EA95E2DFD55 | No |  |  | Defect | SW
ABL815 FLEX | AC Modul | Ampullenbrecher steht zu weit ins AutoCheck-Karussell => Modul gereinigt, Brecher manuell zuruck gesetzt | CD1537A05E17B240 | No | AC Module | Ampoule crusher is too far into the AutoCheck carousel => module cleaned, crusher manually set | Setup | QC schedule | Double QC performed
ABL825 FLEX |  |  | CD163672D4089F4D | No |  |  | Defect | Inlet module | Probe | Unknown
ABL815 FLEX |  | Tausch der Referenz Elektrode. | CD1515265E0F8F40 | No |  | Replacement of the reference electrode. | Defect | Electrode | Electrode | Reference
TCM-COMBIM |  |  | CD155B2C5E11DC93 | No |  |  | Defect | HW | PC Module
TCM-COMBIM |  |  | CD1532485E4FDC82 | No |  |  | Defect | HW | Gas Module
ABL90 | Instrument was having intermittent IP address conflict and getting kicked out of Radiance. Upon reboot, the analyzer would print out the IP settings. Suspect is corruption in windows software or incorrect IP settings. Spoke with in house IT department and they verified IP address, but will need help from Radiometer SIT to verify settings and change IP address if necessary. | Installed new cf card with windows xp and system software version 3.0 MR2. performed service dump to capture data. Verified instrument operation. | CD164F87D44CFFCA | No | Instrument was having intermittent IP address conflict and getting kicked out of Radiance. Upon reboot, the analyzer would print out the IP settings. Suspect is corruption in windows software or incorrect IP settings. Spoke with in house IT department and they verified IP address, but will need help from Radiometer SIT to verify settings and change IP address if necessary. | Installed new cf card with windows xp and system software version 3.0 MR2. performed service dump to capture data. Verified instrument operation. | Setup | Radiance
ABL800Basic |  |  | CD24B5FAE60267D9 | No |  |  | Defect | Electrode | Electrode | Reference
ABL835 FLEX | Prior investigation revealed poor sense on new electrode on recently installed analyzer. Had warranty replacement pO2 electrode sent to customer. This SWO is to document receipt of old electrode for return to Depot | Replace pO2 Electrode Lot LI-03 with new electrode Lot LP-01 | CD167B42D450F2C7 | No | Prior investigation revealed poor sense on new electrode on recently installed analyzer. Had warranty replacement pO2 electrode sent to customer. This SWO is to document receipt of old electrode for return to Depot | Replace pO2 Electrode Lot LI-03 with new electrode Lot LP-01 | Defect | Electrode | Electrode | pO2
ABL80 CO-OX | SC-Smart Chip Fehler | Austausch Klappe für Lösungspack, da defekt Sensorkassette erneuert SW aktuallisiert | CD1783F44A0B2665 | No | SC-Smart Chip Error | Replacement flap for solution pack, as defective sensor cartridge renewed SW is updated
ABL80 OSM |  |  | CD27EE904808989D | No
ABL90 | 1343 UNABLE TO PUMP SOLUTIONS |  | CD17C4934A3B7340 | No | 1343 UNABLE TO PUMP SOLUTIONS
ABL90 |  | LP ersetzt | CD1C832222514291 | No |  | LP replaced | Defect | SP | Lac QC
ABL90 | Customer had flushed & cleaned the inlet gasket, but no change. Pressure vacuum test in service mode is too high. Found the inlet gasket had a puncture in sealing surface causing the vacuum test to fail. | Replaced inlet gasket with non-inventory inlet gasket and all test passed. | CD17CBBF4A4F6560 | No | Customer had flushed & cleaned the inlet gasket, but no change. Pressure vacuum test in service mode is too high. Found the inlet gasket had a puncture in sealing surface causing the vacuum test to fail. | Replaced inlet gasket with non-inventory inlet gasket and all test passed. | Defect | Hw | Inlet Gasket
ABL825 FLEX |  | Tausch der Air Ventil Membran im pH Bg Modul. | CD1796E24A2FD9D9 | No |  | Replacement of the Air Valve membrane in the pH Bg module. | Defect | pH/BG Module | Valve, air | Membrane
ABL810 FLEX |  |  | CD17B5D34A2F1FD3 | No |  |  | Defect | pH/BG Module | Valve, air | Membrane
ABL815 FLEX |  |  | CD17B1314A0A290C | No |  |  | Defect | Electrode | Membrane | pCO2
ABL825 FLEX | Support recommends deleting the setup and re-entering the setup manually. | Deleted the setup and re-entered manually. | CD1799884A0CAA2D | No | Support recommends deleting the setup and re-entering the setup manually. | Deleted the setup and re-entered manually. | Defect | Basic unit | AutoCheck Module | SW Bug QC Schedule
ABL825 FLEX | checked logs and membranes | try new pco2 electrode- OPEN TECHLINE CASE | CD1783694A39AAF5 | No | checked logs and membranes | try new pco2 electrode- OPEN TECHLINE CASE | Defect | Electrode | Electrode | pCO2
ABL825 FLEX | Analyzer is running the wrong level of QC. | Checked that laser scanner is securley mounted. Aligned laser scanner. Checked forproper scan. | CD1784334A2E9351 | No | Analyzer is running the wrong level of QC. | Checked that laser scanner is securley mounted. Aligned laser scanner. Checked forproper scan. | Adjustment | Auto Check Module | Laser alignment
ABL90 |  |  | CD18EBBAC03E550D | No |  |  | User Error | Incorrect Probe install
ABL90 | THE DEMO UNIT WAS DROPPED OFF AT CLAUDE LEBLOND'S HOUSE |  | CD1B69E7AC301919 | No | THE DEMO UNIT WAS DROPPED OFF AT CLAUDE LEBLOND'S HOUSE
ABL90 | Kal.- QC-Ergebnisse, Flüssigkeitsweg | Waste-Tubing erneuert; gespült | CD1A20F1361CD4B3 | No | Kal.- QC-Ergebnisse; Flüssigkeitsweg | Waste tubing renewed; Rinsed | Defect | HW | Waste Tube
ABL80 OSM |  |  | CD1B5D3BAC48C86B | No
ABL80 | pCO2 is out on quality control level 1 (71 , upper limit is 69). Also several solutions pack and two sensor cassettes had been replaced due to pCO2 being out on system cycle. | Opened and checked internal tubings. Perform zero calibration. Downloaded data logs. | CD1B7A98AC14D41D | No | pCO2 is out on quality control level 1 (71 , upper limit is 69). Also several solutions pack and two sensor cassettes had been replaced due to pCO2 being out on system cycle. | Opened and checked internal tubings. Perform zero calibration. Downloaded data logs.
ABL90 |  |  | CD1DB9C69805594A | No |  |  | Defect | HW | Wet sec PCB
ABL90 |  |  | CD295F19BE48CE5B | Yes |  |  | Re classified RMED | In Workshop
ABL90 |  |  | CD2960F2BE5ECE5D | Yes |  |  | Re classified RMED | In Workshop
ABL90 |  |  | CD29634DBE2FCE60 | Yes |  |  | Re classified RMED | In Workshop
ABL90 | THE WET SECTION BOARD WAS REPLACED AND IT HAD NO EFFECT ON THE ERROR SO A ANALYZER ID PCB WAS ORDERED | THERE WAS A SOFTWARE ERROR WHEN THE NEW WET SECTION BOARD WAS INSTALLED. THE OLD WET SECTION BOARD WAS REINSTALLED WITH NO ERRORS TO REPORT EXCEPT FOR THE 1311 ERROR CODE. ONCE THE NEW ANALYZER ID PCB WAS RECEIVED IT WAS REPLACED AND THE 1311 ERROR CODE WENT AWAY. HOWEVER THE SENSOR CASSETTE HAD TO BE REPLACED BECAUSE IT HAD NOT BEEN USED IN SOMETIME AND THE ANALYZER WAS POSTING GLU ERRORS. A NEW SENSOR CASSETTE WAS REPALCED AND THIS ONE WAS FAILING PCO2 INTERNAL QC. A NEW SENSOR CASSETTES WAS INSTALLED AND THERE ARE STILL ISSUES WITH PCO2 FAILING THIRD LEVEL INTERNAL QC. | CD1B98AFAC366121 | No | THE WET SECTION BOARD WAS REPLACED AND IT HAD NO EFFECT ON THE ERROR SO A ANALYZER ID PCB WAS ORDERED | THERE WAS A SOFTWARE ERROR WHEN THE NEW WET SECTION BOARD WAS INSTALLED. THE OLD WET SECTION BOARD WAS REINSTALLED WITH NO ERRORS TO REPORT EXCEPT FOR THE 1311 ERROR CODE. ONCE THE NEW ANALYZER ID PCB WAS RECEIVED IT WAS REPLACED AND THE 1311 ERROR CODE WENT AWAY. HOWEVER THE SENSOR CASSETTE HAD TO BE REPLACED BECAUSE IT HAD NOT BEEN USED IN SOMETIME AND THE ANALYZER WAS POSTING GLU ERRORS. A NEW SENSOR CASSETTE WAS REPALCED AND THIS ONE WAS FAILING PCO2 INTERNAL QC. A NEW SENSOR CASSETTES WAS INSTALLED AND THERE ARE STILL ISSUES WITH PCO2 FAILING THIRD LEVEL INTERNAL QC. | Defect | HW | Analyzer ID PCB
ABL90 | Mantem-se erro de Glu no QC nivel 3. Nova substituição do cartucho de soluções. Mantem-se o erro de Glu no QC nivel 1 e 3. Substituição da cassete de electrodos. Erro de calibração de pH. Nova substituição da cassete de electrodos. Verificação do correcto funcionamento. |  | CD1B8CD6AC465403 | No | It keeps Glu error at QC level 3. New cartridge replacement solutions. It keeps the error of Glu in QC Level 1 and 3. Replacing the electro-cassette. PH calibration error. New replacement of the electro-cassette. Verification of proper operation. |  | Defect | SC | GLU, and pH on second SC
ABL800Basic |  |  | CD1B6D63AC62EFE6 | No |  |  | User Error | Inlet module | Waste tubing | Disconnected during cleaning
ABL825 FLEX | Performed service liquid sensor test with no errors. Checked fluidic module valves and pumps. | Flushed fluidic module air intake line. | CD1B65CFAC1F7CC9 | No | Performed service liquid sensor test with no errors. Checked fluidic module valves and pumps. | Flushed fluidic module air intake line. | Cleaned | Fluidic pathway | Air vent
ABL825 FLEX |  | Tausch Einlassmodul, da Ersatzteillieferung zu späterem Zeitpunkt. | CD1B71C0AC4EDF5F | No |  | Exchange inlet module, as spare parts delivery at a later time. | Defect | Inlet module | -
ABL825 FLEX |  | Tausch waste Sensor. | CD1B881FAC262005 | No |  | Exchange waste sensor. | Defect | Basic unit | Waste detector
ABL835 FLEX | Std Dev Dark value failed on spectrometer in oxi compete test. Glucose not giving any Sens, status. | Replaced power supply. Customer had replaced new Glucose membranes many times. Replaced electrode. | CD1B9845AC444C1C | No | Std Dev Dark value failed on spectrometer in oxi compete test. Glucose not giving any Sens, status. | Replaced power supply. Customer had replaced new Glucose membranes many times. Replaced electrode. | Defect | Power Supply | Dark value
ABL825 FLEX | Inlet probe looks ok. Reboot analyzer. Still same error. After reseating inlet pickup board cable with reboot - still same error. Worked for 1.5 hours and errors came back. | Remove Inlet Module and reseat pickup board cable. Replaced Inlet Probe. Replaced Inlet Gasket from customer stock. | CD1B3D85AC22B201 | No | Inlet probe looks ok. Reboot analyzer. Still same error. After reseating inlet pickup board cable with reboot - still same error. Worked for 1.5 hours and errors came back. | Remove Inlet Module and reseat pickup board cable. Replaced Inlet Probe. Replaced Inlet Gasket from customer stock. | Defect | Inlet module | Gasket Inlet
ABL825 FLEX |  |  | CD1B6F02AC5E5487 | No |  |  | Defect | Inlet module | Nipple kit
AQT90 FLEX | Se abre la pantalla y se apreta conector de la misma se ajustan los tornillos de sujección se limpia la pantalla y se realiza calibración de la misma se guarda la configuración y se comprueba la misma realizando muestra, se activa el codigo de barras externo para leer codebar y se comprueba el funcionamiento general. |  | CD1B8561AC5D59AC | No | The screen is opened and the connector is tightened the fastening screws are adjusted and the screen is calibrated and the configuration is calibrated and checked by displaying, the external barcode is activated to read codebar and s and check the overall performance.
AQT90 FLEX |  |  | CD2C112CAA2EA7FB | No
Tosca-500 | Trotz langem Probelauf ist der Fehler hier nicht aufgetreten, während Telefonat bestätigt der Kunde obige Fehlerangabe! | TechLine-Nr.: RDEU 683 942 829 erstellt. | CD1B2D9AAC215EB5 | No | Despite a long trial run, the error did not occur here, while phone call confirms the above error indication! | TechLine No.: RDEU 683 942 829. | Defect | Basic unit
Tosca-500 | Sehr harter Sturzschaden, sicherheitshalber nicht an Netz angeschlossen! | TechLine-Nr.: RDEU 683 763 331 erstellt. | CD1B2BB8AC495EB3 | No | Very hard fall damage, not connected to the grid for safety reasons! | TechLine No.: RDEU 683 763 331. | User Error | HW | Broken
ABL90 |  |  | CD1CDBBF22448B32 | No |  |  | Cleaned | Fluidic pathway | and cleaned Inlet
ABL90 |  |  | CD1CC8C3222EEEB1 | No |  |  | Defect | SW | Re install Database | Status code 2
ABL90 | The Cal's would not come in for NA and Lac and QC was out for PCO2 | I flushed the analyzer and found that there was a leak in the waste tube. I replaced the waste tube and ran cal's. The Na would still not come in. I replaced the SC and ran cal's | CD1C903A22465822 | No | The Cal's would not come in for NA and Lac and QC was out for PCO2 | I flushed the analyzer and found that there was a leak in the waste tube. I replaced the waste tube and ran cal's. The Na would still not come in. I replaced the SC and ran cal's | Defect | HW | Waste Tube
ABL90 | The analyzer is green but will give a error when running a blood sample | I flushed the analyzer and then the cassette and still will not run. I replaced the Sc and all is well | CD1C9E2F221F5825 | No | The analyzer is green but will give a error when running a blood sample | I flushed the analyzer and then the cassette and still will not run. I replaced the Sc and all is well | Cleaned | Fluidic pathway | and Replaced SC | Potential Silicone issue
ABL90 |  |  | CD1C9053221CC097 | No |  |  | Cleaned | Fluidic pathway | and tHb Cal
ABL90 | INTERNAL QC WAS RUN SEVERAL TIMES AND CALIBRATIONS AS WELL. 5 SAMPLES OF BLOOD WERE RUN TO CONDITION THE CASSETTE | THE INSTRUMENT WAS LEFT OVERNIGHT SO THAT THE CASSETTE COULD BE CONDITIONED. IN THE MORNING ALL LEVELS OF INTERNAL AND EXTERNAL QC PASSED WITH NO ISSUES | CD1C80F4223BBD0F | Yes | INTERNAL QC WAS RUN SEVERAL TIMES AND CALIBRATIONS AS WELL. 5 SAMPLES OF BLOOD WERE RUN TO CONDITION THE CASSETTE | THE INSTRUMENT WAS LEFT OVERNIGHT SO THAT THE CASSETTE COULD BE CONDITIONED. IN THE MORNING ALL LEVELS OF INTERNAL AND EXTERNAL QC PASSED WITH NO ISSUES | Re classified RMED | Follow up to CD1B98AFAC366121
ABL90 | Flüssigkeitswege geprüft | SP und SC erneuert | CD1CE3892224185D | No | Liquid pathways tested | SP and SC renewed | Defect | SP | FT
ABL80 CO-OX | Suspect Hemolyzer failing. | Ran cleaning procedure, recal tHb. | CD2D52D920623EC2 | No | Suspect Hemolyzer failing. | Ran cleaning procedure, recal tHb.
ABL800Basic |  |  | CD1CFAFE220EE1AA | No |  |  | Defect | Electrode | Membrane | Lac
ABL825 FLEX |  |  | CD1CCF2622546BD1 | No |  |  | Adjustment | HW | Oxi Module | tHb cal
ABL825 FLEX | A RINSE WAS PERFORMED AND INLET GASKET WAS LEAKING | THE WASTE TUBING WAS FLUSHED AND A NEW INLET GASKET WAS INSTALLED. SEVERAL RINSES WERE PERFORMEDWITH NO ERRORS TO REPORT. AS WELL THE CUSTOMER WAS RETRAINED ON HOW TO REPLACE THE GASKET | CD1CC018220F6815 | No | A RINSE WAS PERFORMED AND INLET GASKET WAS LEAKING | THE WASTE TUBING WAS FLUSHED AND A NEW INLET GASKET WAS INSTALLED. SEVERAL RINSES WERE PERFORMEDWITH NO ERRORS TO REPORT. AS WELL THE CUSTOMER WAS RETRAINED ON HOW TO REPLACE THE GASKET | User Error | Maintenance | Gasket Inlet
AQT90 FLEX |  |  | CD1DC3A698073F76 | No
ABL90 | siehe ausgeführte Arbeiten | siehe ausgeführte Arbeiten | CD1E0C199810E516 | No | see work carried out | see work carried out | Cleaned | Fluidic pathway | tHb cal
ABL90 |  |  | CD1DF5FC9818BC49 | No |  |  | Cleaned | Fluidic pathway
ABL90 | I checked the inlet for proper seating andit was not seated properly. | I reseated the inlet gasket and ran all test in the service mode. | CD1DCA789832E970 | No | I checked the inlet for proper seating andit was not seated properly. | I reseated the inlet gasket and ran all test in the service mode. | Adjustment | Inlet Gasket
ABL80 CO-OX | Ran Patient samples between 2 analyzers. this unit is still high on blood comparisons. | Replaced suspect Hemolyzer. | CD2D544D20363EC5 | No | Ran Patient samples between 2 analyzers. this unit is still high on blood comparisons. | Replaced suspect Hemolyzer.
ABL80 CO-OX | Suspect Hemolyzer failing on other ABL unit.No errors indicated on this unit.Ran cleaning procedure, then had customer run samples between both analyzers.Samples are OK on this unit. | No problem found with this unit. MetHb values within range. | CD2D566220333EC9 | No | Suspect Hemolyzer failing on other ABL unit. No errors indicated on this unit. Ran cleaning procedure, then had customer run samples between both analyzers. Samples are OK on this unit. | No problem found with this unit. MetHb values within range.
ABL825 FLEX |  |  | CD1DE45598030E59 | No |  |  | Defect | pH/BG Module | Valve, air
ABL825 FLEX | WALL OUTLET MEASURES 117 VOLTS | REPLACED POWER SUPPLY,PERFORM STARTUP AND CALIBRATIONS | CD1DE33F98314B58 | No | WALL OUTLET MEASURES 117 VOLTS | REPLACED POWER SUPPLY,PERFORM STARTUP AND CALIBRATIONS | Defect | Power Supply | -
ABL825 FLEX |  |  | CD1DE06D98044292 | No |  |  | Cleaned | Fluidic pathway | -
ABL825 FLEX | all levels are below pee,r i swapped tubing and cleaned valves, examined gas lines, swapped dummy electrode in cl and checked ph sealing | repalced the measuring chamber | CD1DE8A5980D1C96 | No | all levels are below pee,r i swapped tubing and cleaned valves, examined gas lines, swapped dummy electrode in cl and checked ph sealing | repalced the measuring chamber | Defect | Electrode | Electrode | pCO2
ABL815 FLEX |  |  | CD1DDF04981D152E | No |  |  | Defect | Electrode | Electrode | pCO2
ABL820 FLEX | VERIFIED DATA | SEVERAL LOTS OF MEMBRANES AND TECHLINE RECOMENDS WARR REPLACEMENT OF ELECTRODE | CD1DED109819DECE | No | VERIFIED DATA | SEVERAL LOTS OF MEMBRANES AND TECHLINE RECOMENDS WARR REPLACEMENT OF ELECTRODE | Defect | Electrode | Electrode | pCO2
AQT90 FLEX | NA | Der er udført årseftersyn på AQT90. Pumpe P1 er smurt da den gav mislyd fra sig. Nål er udskiftet. Den optiske unit er skiftet. | CD1E0A579800E954 | No | In | Annual surveys have been conducted on AQT90. Pump P1 is lubricated as it caused the missound from it. The optical unit is switched.
TCM-COMBIM |  |  | CD1E09319801306F | No |  |  | Defect | HW | Communication PCB
TCM-COMBIM |  |  | CD1E07D19815306D | No |  |  | No fault found
ABL80 CO-OX |  |  | CD1F45360E256D92 | No
ABL80 | Went to Fedex to pick up part. The pO2 was out on the system cycle when I arrived. Giving possible air in cal solution. Fan is roaring. | Replaced Computer unit cooling fan. Flushed all solution lines with DI water. Replaced SC and SP with customer stock. Gave customer the cosumable credit form. | CD1F19AD0E2E6F27 | No | Went to Fedex to pick up part. The pO2 was out on the system cycle when I arrived. Giving possible air in cal solution. Fan is roaring. | Replaced Computer unit cooling fan. Flushed all solution lines with DI water. Replaced SC and SP with customer stock. Gave customer the cosumable credit form.
ABL90 | Austasuch LP brachte keine Verbesserung | erneuter Austasuch des LP | CD1F38570E450A64 | Yes | Austasuch LP brought no improvement | re-examination of the LP | Re classified RMED | Telephone support
ABL90 |  |  | CD254DDC5C37FE5D | Yes |  |  | Re classified RMED | Analyzer replaced | Repair in =R= Workshop
ABL90 | instrument shuts off when ac is unplugged | i replacced battery | CD1F9F240E445E82 | No | instrument shuts off when ac is unplugged | i replacced battery | Defect | HW | Battery module |  | x
ABL90 | Battery not charging | Replaced Battery install new software 3.1 | CD1F16560E109690 | No | Battery not charging | Replaced Battery install new software 3.1 | Defect | HW | Battery module |  | x
ABL90 | Inservice on ABL90. |  | CD20A69E8415DCC2 | No | Inservice on ABL90.
ABL820 FLEX | Recurring problem. 688, 581, 588 and093 errors on QC results. | Replaced and aligned hemolyzer unit. | CD1F44880E0A3CDA | No | Recurring problem. 688, 581, 588 and093 errors on QC results. | Replaced and aligned hemolyzer unit. | Defect | HW | Oxi Module | Hemolyzer
ABL825 FLEX | Customer had already performed protei removal | Replaced OXI PCB. adjusted lamp. Ran complete OXI test. Performed tHB cal in service and customer mode | CD1F3FBC0E0B4388 | No | Customer had already performed protei removal | Replaced OXI PCB. adjusted lamp. Ran complete OXI test. Performed tHB cal in service and customer mode | Defect | HW | Oxi Module | OXI PCB
ABL805 FLEX | COMPLETE OXIMETRY TEST SHOWS HALOGEN LAMP AT 2100 | REPLACED HEMOLYZER AND LAMP UNIT.ADJUST LAMP TO 3350 PERFORM THB CAL IN SERVICE AND CUSTOMER MODES.PERFORM CALIBRATIONS | CD1F4D790E3DDB2B | No | COMPLETE OXIMETRY TEST SHOWS HALOGEN LAMP AT 2100 | REPLACED HEMOLYZER AND LAMP UNIT. ADJUST LAMP TO 3350 PERFORM THB CAL IN SERVICE AND CUSTOMER MODES. PERFORM CALIBRATIONS | Adjustment | HW | Oxi Module | Lamp Unit
ABL837 FLEX | The sample was ran again and measured OK. The sample in question was edited and shows an audit trail. | Ran several comparisons and all tested OK. Ran tHB calibration in service mode. Customer will continue to compare and monitor samples for other critical discrepancies. | CD1F41960E40EF04 | No | The sample was ran again and measured OK. The sample in question was edited and shows an audit trail. | Ran several comparisons and all tested OK. Ran tHB calibration in service mode. Customer will continue to compare and monitor samples for other critical discrepancies. | Unknown | HW | Oxi Module | tBil + crea
AQT90 FLEX | Sustitución de la pantalla completa ya que despues de calibrarla pierde al tiempo el punto de señalización, se comprueba y ajusta realizando una calibración, se realiza una muestra, funcionando correctamente. |  | CD1F3D180E5A37FD | No | Replacing the full screen since after calibrating it loses at the time the signaling point, is checked and adjusted by performing a calibration, a sample is performed, working correctly.
ABL90 | QCs out of limits. | Perform back flush procedure. Perform service alignments. | CD240BB3E65689CC | No | QCs out of limits. | Perform back flush procedure. Perform service alignments. | Cleaned | Fluidic pathway |  |  | x
ABL90 | the analyzer keep given response errors on cal's | I replaced the sensor interface PCB and still had the error. I replaced the SC and it came right in. | CD2066A88440AA75 | No | the analyzer keep given response errors on cal's | I replaced the sensor interface PCB and still had the error. I replaced the SC and it came right in. | Defect | SC | ?
TCM4 |  |  | CD27F488480E95EC | Yes |  |  | Re classified RMED | Not upSWO
ABL90 | Comprobar posibles obstrucciones en el circuito neumático, sp y sc, limpiar restos de líquidos derramados, cambiar el tubo de la válvula OXI, calibrar y probar con muestras y controles. | Cambiar tubo de la válvula OXI. | CD2060188400ED8C | No | Check for possible obstructions in the pneumatic circuit, sp and sc, clean debris from spilled liquids, change the OXI valve tube, calibrate and test with samples and controls. | Change OXI valve tube. | Defect | HW | Pinch Valve
ABL815 FLEX | pH und Ref | unterer Dichtungsring bei pH lose => Ersatz | CD20B1A7840697A3 | No | pH and Ref | lower sealing ring at pH loose => Replacement | Defect | Electrode | Gasket | pH
ABL825 FLEX |  |  | CD205289842A3FD4 | No |  |  | Defect | pH/BG Module | Valve, air | Membrane
ABL825 FLEX | RINSE WAS PERFORMED AND THERE WAS LEAKING AT THE INLET GASKET | THE OLD INLET CALIBRATIONS WERE PRINTED. AN INLET CALIBRATION WAS PERFORMED AND THE THE ACTUAL STEP VALUE CHANGED FROM 98 TO 163 AND THE RESIDUAL ERROR CHANGED FROM 47 TO 22. A NEW INLET GASKET WAS INSTALLED. SEVERAL RINSES WERE PERFORMED AND A 1 PT CALIBRATION AS WELL WITH NO ERRORS TO REPORT. TECH SUPPORT WAS ALSO CALLED TO QUERY THE CHANGE OF STEPS. | CD2059FE844C03C8 | No | RINSE WAS PERFORMED AND THERE WAS LEAKING AT THE INLET GASKET | THE OLD INLET CALIBRATIONS WERE PRINTED. AN INLET CALIBRATION WAS PERFORMED AND THE THE ACTUAL STEP VALUE CHANGED FROM 98 TO 163 AND THE RESIDUAL ERROR CHANGED FROM 47 TO 22. A NEW INLET GASKET WAS INSTALLED. SEVERAL RINSES WERE PERFORMED AND A 1 PT CALIBRATION AS WELL WITH NO ERRORS TO REPORT. TECH SUPPORT WAS ALSO CALLED TO QUERY THE CHANGE OF STEPS. | Defect | Inlet module | Inlet Gasket
ABL825 FLEX |  |  | CD209A958431059B | No |  |  | Defect | Printer module | Printer head
ABL835 FLEX | AU fan is making grinding noise. I pulled back cover and see no obvious visual obstruction but fan is quite noisy. I tightened lock down scewss on fan cage - no help. Will order replacement fan ASAP | Clean and inspect. New fan on order. | CD207A0D84555181 | No | AU fan is making grinding noise. I pulled back cover and see no obvious visual obstruction but fan is quite noisy. I tightened lock down scewss on fan cage - no help. Will order replacement fan ASAP | Clean and inspect. New fan on order. | Defect | Basic unit | Fan | Noisy
ABL835 FLEX | Cal and QC seem normal. Intermittent but more frequent. Occurs on 25 sample in the last week. I verified Lac Zero Current in spec. Cleaned electrode well. Customer already replaced membrane recently. Will monitor | Clean electrode and well. Will open Techline case for this 2 month old analyzer | CD20858D84515184 | No | Cal and QC seem normal. Intermittent but more frequent. Occurs on 25 sample in the last week. I verified Lac Zero Current in spec. Cleaned electrode well. Customer already replaced membrane recently. Will monitor | Clean electrode and well. Will open Techline case for this 2 month old analyzer | Cleaned | Fluidic pathway | El/met Measuring chamber
ABL90 | ABL geht nicht aus Fehlerbehebungsprozedur raus Drucküberprüfung Vakuumfehler z.t. Flüssigkeitstransportfehler | Fehlersuche Einlaß gereinigt div Verschleißteile erneuert Einlaßnadel durchgespült im Service Programm durchgetestet | CD21AB15FA361139 | No | ABL does not exit from troubleshooting procedure Pressure checking Vacuum fault, in some part liquid transport error | Troubleshooting inlet cleaned div wear parts renewed inlet needle flushed through in the service program thoroughly tested | Cleaned | Inlet gasket | fluidic pathway | & replaced parts | x
ABL90 |  |  | CD22F433702CDA74 | No |  |  | Defect | HW | CPU |  | x
ABL80 CO-OX | Fließwege | LP zurück gesetzt Installierung neuer SC, da alte nicht zu reinstallieren ging | CD246575E60252A0 | No | Fließwege | LP reset installation of new SC as old one did not go to reinstall
ABL80 FLEX Basic version |  |  | CD245DF5E61EBF30 | No
ABL90 | Gerät aus ein, LP ersetzt brachte keine Verbesserung | Gerät aus, Solution Pack PCB ersetzt, Gerät ohne pack neu gestartet, LP intstalliert, | CD246628E653396A | No | Device off, LP replaced brought no improvement | device off, Solution Pack PCB replaced, device without pack restarted, LP intstalled, | Defect | HW | SP PCB |  | x
ABL90 |  |  | CD24388EE6557113 | No |  |  | Cleaned | Fluidic pathway | Clot
ABL90 |  |  | CD241FB6E6306AE8 | No |  |  | Setup | HIS/LIS | Network |  | x
ABL800Basic |  |  | CD246178E6450FB3 | No |  |  | Defect | Auto Check Module | FPC Cable | Old Analyzer
ABL825 FLEX | checked data logs for errors, checked for clots, checked flow of samples, checked electrode and membrane | removed electrode and replaced membrane, removed dual housing valves and cleaned valves, replaced ripped El Met air valve membrane. | CD247BFEE6117F30 | No | checked data logs for errors, checked for clots, checked flow of samples, checked electrode and membrane | removed electrode and replaced membrane, removed dual housing valves and cleaned valves, replaced ripped El Met air valve membrane. | Defect | El/Met Module | Valve, air | Membrane
ABL825 FLEX | Sensitivity on electrode reading below 5 | Replaced electrode | CD2499EDE60B2215 | No | Sensitivity on electrode reading below 5 | Replaced electrode | Defect | Electrode | Electrode | pO2
ABL825 FLEX | Found all the tHb values to be below 1 g dl Oxi pinch vave seems sticky | Cleaned and lubricated oxi pinch valve Test in service-OK | CD24236FE631BEA6 | No | Found all the tHb values to be below 1 g dl Oxi pinch vave seems sticky | Cleaned and lubricated oxi pinch valve Test in service-OK | Defect | HW | Oxi Module | Pinch Valve
ABL835 FLEX |  |  | CD245D8FE655241F | Yes |  |  | Re classified RMED | Local Analyzer setup
TCM-COMBIM | Communication error. | I tried a new Electrode Module, but the communication error persisted. Replaced the Commnucation PCB. Problem solved. | CD245170E614A711 | No | Communication error. | I tried a new Electrode Module, but the communication error persisted. Replaced the Commnucation PCB. Problem solved. | Defect | HW | Communication PCB
TCM-COMBIM |  |  | CD245CF3E61A8E64 | No |  |  | Defect | HW | Display
ABL90 |  |  | CD2565965C3B3166 | No |  |  | Replaced | Inlet Gasket
ABL90 |  |  | CD258DE85C0F1B25 | No |  |  | Cleaned | Fluidic pathway | Yearly Maintenance
ABL90 | Customer has been getting this error at least 1 time a day | Replaced solution pack ID PCB. Loaded software version 3.1. Installed new inlet probe and inlet gasket holder | CD2598255C3505A8 | No | Customer has been getting this error at least 1 time a day | Replaced solution pack ID PCB. Loaded software version 3.1. Installed new inlet probe and inlet gasket holder | Defect | HW | Solution pack PCB id
ABL825 FLEX | Einlassnadel geprüft | Einlassnadel erneuert | CD2537885C30F98A | No | Inlet needle checked | Inlet needle renewed | Defect | Inlet module | Probe bend
ABL825 FLEX | THE INLET GASKET WAS LEAKING | HE INLET PROBE WAS REPLACED AND THE WASTE TUBING FROM THE WASTYE CONNECTOR TO THE Y JUNCTION WAS REPLACED. ALSO THE WASTE TUBING FROM THE INLET GASKET WAS REPLACED. 30 RINSES WERE RUN WITH NO LEAKING TO REPORT. | CD256B625C2E158A | No | THE INLET GASKET WAS LEAKING | HE INLET PROBE WAS REPLACED AND THE WASTE TUBING FROM THE WASTYE CONNECTOR TO THE Y JUNCTION WAS REPLACED. ALSO THE WASTE TUBING FROM THE INLET GASKET WAS REPLACED. 30 RINSES WERE RUN WITH NO LEAKING TO REPORT. | Defect | Basic unit | Waste pump
ABL810 FLEX | The AutoCheck lid failed to open at the time scheduled for the QC to run. | I tested but could not duplicate the problem. | CD2611FE5C165F57 | No | The AutoCheck lid failed to open at the time scheduled for the QC to run. | I tested but could not duplicate the problem. | Unknown | Auto Check Module | Lid
AQT90 FLEX |  |  | CD25924D5C1592EC | No
ABL80 CO-OX |  |  | CD3E6DDE9449D69D | No
ABL80 CO-OX | Upon arrival, found unit exhibiting No Endpoint errors on PO2 when running VK-R5 kit. Spoke to Technical support on issue. Recommendations include replacing Spine Board and Battery on unit. | Replaced Spine Board and Battery on unit. Performed spine board setup and Zero calibration alignment on unit. | CD26B050D2100AF9 | No | Upon arrival, found unit exhibiting No Endpoint errors on PO2 when running VK-R5 kit. Spoke to Technical support on issue. Recommendations include replacing Spine Board and Battery on unit. | Replaced Spine Board and Battery on unit. Performed spine board setup and Zero calibration alignment on unit.
ABL90 | Inlet not closing properly all the time over the Inlet probe. Analyzer is running cals and QC without problems. | Replaced the Inlet Module. | CD26CADAD23083DC | No | Inlet not closing properly all the time over the Inlet probe. Analyzer is running cals and QC without problems. | Replaced the Inlet Module. | Defect | HW | Inlet module
ABL90 |  |  | CD26A1BDD25032E3 | No |  |  | Defect | SP | pO2 QC
ABL90 |  |  | CD26A25FD220CDDB | No |  |  | Defect | HW | Hemolyzer
ABL90 |  |  | CD26AB9BD2331316 | No |  |  | Replaced | Inlet probe | + Gasket
TCM/Tosca | Comprobar el equipo, cambiarles las tapas frontal y trasera, reparar conector de la placa base, calibrar y probar. | Cambiar tapas y reparar conector de la placa base. | CD26BF36D252DF9D | No | Check the equipment, change the front and back covers, repair motherboard connector, calibrate and test. | Replace caps and repair motherboard connector. | Defect | HW | PC Module
ABL825 FLEX | Limpieza de camara de medida. Sustitucion de la membrana de ref y lactato. Limpeza salida de desperdicios y de aire. Ajuste sensores de liquidos. Cal, qc y muestas. | Limpieza de camara de medida. Sustitucion de la membrana de ref y lactato. Limpeza salida de desperdicios y de aire. Ajuste sensores de liquidos. Cal, qc y muestas. | CD26CF59D245B301 | No | Measuring camera cleaning. Replacement of the ref and lactate membrane. Cleaning waste and air output. Adjust liquid sensors. Cal, qc and you. | Measuring camera cleaning. Replacement of the ref and lactate membrane. Cleaning waste and air output. Adjust liquid sensors. Cal, qc and you. | Cleaned | Fluidic pathway | Air vent
ABL820 FLEX | Limpieza de cámara de medida y electrodo. Se comparan resultados con otro gasómetro. Se sustituye electrodo de pH. Se realizan calibraciones y Qc. Se prueba el equipo con usuarios |  | CD36FACBD0590FF7 | No | Cleaning of measuring chamber and electrode. Results are compared with another gasometer. PH electrode is replaced. Calibrations and Qc are performed. The computer is tested with users
ABL837 FLEX | Calibrations are ok. Customer has ran several cleaning and protein removals. The electrode has been remembraned. Took out electrode for inspection and found a clot unerneath electrode. | Cleaned measuring chamber and electrode with cleaning solution and DI water. | CD26BF89D2231BB6 | No | Calibrations are ok. Customer has ran several cleaning and protein removals. The electrode has been remembraned. Took out electrode for inspection and found a clot unerneath electrode. | Cleaned measuring chamber and electrode with cleaning solution and DI water. | Cleaned | Fluidic pathway | El/met Measuring chamber | Ca
ABL825 FLEX | Troubleshot using Complete Oximetry test. Hemolyzer failed for Tuning Voltage. When I removed Hemolyzer I noticed leakage at the inlet side of the Hemolyzer. | Replaced Hemolyzer and valve tube due to loose fitting after removing from old Hemolyzer. Aligned Lamp assembly. Performed tHb calibration. | CD26F69ED234916F | No | Troubleshot using Complete Oximetry test. Hemolyzer failed for Tuning Voltage. When I removed Hemolyzer I noticed leakage at the inlet side of the Hemolyzer. | Replaced Hemolyzer and valve tube due to loose fitting after removing from old Hemolyzer. Aligned Lamp assembly. Performed tHb calibration. | Defect | HW | Oxi Module | Hemolyzer | Leakage
ABL800Basic |  |  | CD26CCD0D248AF83 | Yes |  |  | Re classified RMED | Not upSWO | Preventive Maintenance
ABL800Basic |  |  | CD26E568D239F2F6 | No |  |  | Unknown | Radiance
ABL820 FLEX | nalyzer is normal. Data logs show prior Oxi liquid sensor errors. | Checked fluidic module valves Flush fluidic module air vent. Ran liquid sensor test (service). Also checked solutions bottle for forecasting of levels. | CD2710F2D21305A4 | No | nalyzer is normal. Data logs show prior Oxi liquid sensor errors. | Checked fluidic module valves Flush fluidic module air vent. Ran liquid sensor test (service). Also checked solutions bottle for forecasting of levels. | Cleaned | Fluidic pathway | Air vent
ABL820 FLEX | Analyzer is normal now. Data logs show prior liquid sensor errors. | Flush air vent in fluidic modules. Check fluidic module valves. | CD270EC0D23E05A2 | No | Analyzer is normal now. Data logs show prior liquid sensor errors. | Flush air vent in fluidic modules. Check fluidic module valves. | Cleaned | Fluidic pathway | Air vent
ABL800Basic | V8im Fluidicmodul defekt | V8 ersetzt | CD2684F0D225603B | No | V8im Fluidicmodul defective | V8 replaced | Defect | Fluidic Module | Waste valve | Iso
ABL825 FLEX | Did complete oxi test in service, Failed phase regulaation Checked oxi tubes for leaks-OK | Replaced hemolyzer Alignment in service Did a complete oxi test-Passed | CD26EC35D24C9CB8 | No | Did complete oxi test in service; Failed phase regulaation Checked oxi tubes for leaks-OK | Replaced hemolyzer Alignment in service Did a complete oxi test-Passed | Defect | HW | Oxi Module | Hemolyzer | phase reg
AQT90 FLEX | NA. | Jeg har udskiftet den optiske unit. Jeg har udført et årsservice. | CD2711DAD24392AC | No | In. | I have replaced the optical unit. I've done an annual service.
ABL80 FLEX Basic version |  | Sensor Board getauscht. | CD27F4A44839317C | No |  | Sensor board.
ABL80 OSM |  |  | CD27EAEF4820F9A2 | No
ABL90 |  |  | CD280AF448003609 | Yes |  |  | Re classified RMED | Not upSWO
ABL805 FLEX | se observa suciedad en la cámara de medidas, se sustituye membrana de po2. se limpia junta del transductor de flujo. se calibra equipo y se comprueba correcto funcionamiento mediante qc. se realizan muestras en los 2 equipos sin observarse diferencias apreciables en los resultados. equipo ok. |  | CD36FD32D0450FFB | No | dirt is observed in the measuring chamber, po2 membrane is replaced. flow transducer gasket is cleaned. equipment is calibrated and properly functioned by qc. samples are made on the 2 teams without appreciable differences in results. team ok.
ABL820 FLEX |  | Tausch des Oximetrie-Moduls. | CD27DEAB481836CE | No |  | Replacement of the oximetry module. | Defect | HW | Oxi Module | Hemolyzer | phase reg
ABL825 FLEX |  |  | CD2CE9552053843B | No |  |  | Cleaned | Fluidic pathway
ABL835 FLEX | AU fan makes excessive noise. Possible bearing. I replaced AU fan as per service manual. | Replaced Fan | CD2822934829CEDE | No | AU fan makes excessive noise. Possible bearing. I replaced AU fan as per service manual. | Replaced Fan | Defect | Basic unit | Fan | Noisy
TCM-COMBIM | Alarmlautstärke war auf 1, aber beide Alar-me (pCo2 und pO2) waren ausgeschaltet, Akku war kpl. entladen, | Akku geladen, Alarme aktiviert, | CD2805A94848AD40 | No | Alarm volume was on 1, but both Alar-me (pCo2 and pO2) were off, battery was cpl. Unloading | Battery charged, alarms activated, | Defect | HW | Power Supply
ABL80 CO-OX | Meldung über HOTLINE ABL erkennt Luftblasen im System Kunde hat bereits LP und SC erneuert | Fehlersuche Überprüfung Fließsystem, Reinigung div. Schläuche SC80 überprüft, durchgespült tHb Kal | CD291A58BE2F84D3 | No | Message via HOTLINE ABL detects air bubbles in the system customer has already renewed LP and SC | Troubleshooting inspection flow system, cleaning various hoses SC80 checked, rinsed tHb Kal
ABL90 |  |  | CD2938E0BE08BB71 | No |  |  | Defect | HW | Hemolyzer
ABL90 |  |  | CD2938C9BE5B05FA | No |  |  | Defect | SP | FT
ABL90 | siehe SLR | siehe SLR | CD2940EFBE348D23 | No | see SLR | see SLR | Cleaned | Fluidic pathway | although guided
ABL90 |  |  | CD2929D7BE2CF84F | No |  |  | Defect | HW | Pinch Valve | Tube
ABL835 FLEX | Limpieza salida de aire y de desperdicios. Ajuste de sensores de liquidos. Desmonstaje y limpieza de las valvulas de ambos modulos. Calibracion de bombas. Apagado y encendido. Cal, qc y varias muestras. | Limpieza salida de aire y de desperdicios. Ajuste de sensores de liquidos. Desmonstaje y limpieza de las valvulas de ambos modulos. Calibracion de bombas. Apagado y encendido. Cal, qc y varias muestras. | CD29D16ABE03B740 | No | Cleaning air and waste outlet. Adjustment of liquid sensors. Demonist and cleaning of the valves of both modules. Pump calibration. Off and on. Cal, qc and several samples. | Cleaning air and waste outlet. Adjustment of liquid sensors. Demonist and cleaning of the valves of both modules. Pump calibration. Off and on. Cal, qc and several samples. | Cleaned | Fluidic pathway | Air vent
ABL825 FLEX | Tested the valves and inspected the tubing. Didn't observe any failures or leaks. Suspect a valve seal somewhere. | Removed the dual valves, cleaned and reseated. | CD2AC623343F0E4B | No | Tested the valves and inspected the tubing. Didn't observe any failures or leaks. Suspect a valve seal somewhere. | Removed the dual valves, cleaned and reseated. | No fault found | Fluidic pathway
ABL810 FLEX | Intermittent but increasingly more frequent. Oxi spectrum mismatch on QC. Oxi complete test passes. | Replaced Hemolyzer as per service manual. Adjust lamp, service tHb Cal | CD2BEC3DAA10CEEB | No | Intermittent but increasingly more frequent. Oxi spectrum mismatch on QC. Oxi complete test passes. | Replaced Hemolyzer as per service manual. Adjust lamp, service tHb Cal | Defect | HW | Oxi Module | Hemolyzer
ABL80 | I found the pO2 was high on all levels of external QC. The analyzer is also giving outside the response range on the system cycle. I saw where it gave a Temp F error and then the system cycle and external QC starting failing. The room is at about 59 degrees. The QC is at 15C and they are running it at 25C. I told customer this will cause the pO2 to run high on external QC. Spoke with Ryan at tech support and he will email Sendx to verify the response error. Joaquin said all 3 of the 80s are giving the Temp F error. | I told the customer that they will have to change the environment to a warmer temperature. | CD2D043320206612 | No | I found the pO2 was high on all levels of external QC. The analyzer is also giving outside the response range on the system cycle. I saw where it gave a Temp F error and then the system cycle and external QC starting failing. The room is at about 59 degrees. The QC is at 15C and they are running it at 25C. I told customer this will cause the pO2 to run high on external QC. Spoke with Ryan at tech support and he will email Sendx to verify the response error. Joaquin said all 3 of the 80s are giving the Temp F error. | I told the customer that they will have to change the environment to a warmer temperature.
ABL90 |  |  | CD33E0E46E0DE1CE | No |  |  | Defect | SC | Reference
ABL90 |  |  | CD2D02E020117647 | No |  |  | Setup | Language | Keyboard
TCM4 |  |  | CD43D8356C43DD4F | Yes |  |  | Re classified RMED | Logistic upSWO
ABL825 FLEX | checked electrode and electrode membrane, checked valves and liquid sensors, checked electrode chamber and sample path, | removed electrodes and cleaned electrode chamber, removed modules and cleaned dual housing valves, | CD2D6B0F2045F97B | No | checked electrode and electrode membrane, checked valves and liquid sensors, checked electrode chamber and sample path, | removed electrodes and cleaned electrode chamber, removed modules and cleaned dual housing valves, | Cleaned | El/Met Module | Measuring chamber | Lac
ABL825 FLEX | checked electrode and electrode membrane, checked valves and liquid sensors, checked electrode chamber and sample path, | removed electrodes and cleaned electrode chamber, removed modules and cleaned dual housing valves, removed and replaced lactate membrane, | CD2D6BFC2014F97E | No | checked electrode and electrode membrane, checked valves and liquid sensors, checked electrode chamber and sample path, | removed electrodes and cleaned electrode chamber, removed modules and cleaned dual housing valves, removed and replaced lactate membrane, | Cleaned | El/Met Module | Measuring chamber | Lac
ABL825 FLEX | checked electrode and electrode membrane, checked valves and liquid sensors, checked electrode chamber and sample path, | removed electrodes and cleaned electrode chamber, removed modules and cleaned dual housing valves, | CD2D6C362022F980 | No | checked electrode and electrode membrane, checked valves and liquid sensors, checked electrode chamber and sample path, | removed electrodes and cleaned electrode chamber, removed modules and cleaned dual housing valves, | Cleaned | El/Met Module | Measuring chamber | Lac
AQT90 FLEX |  |  | CD2D244F205B59DF | No
AQT90 FLEX |  |  | CD2D0E162022784E | No
ABL90 | Comprovação do histórico de calibrações, QC, medidas e erros do sistema. Detecção da anomalia (Clip da entrada mal montado e erros de Ca e Lac em QC nivel B. Substituição da sonda de asperação e entrada, posicionamento correcto do clip de entrada. Passagem de QC nivel B. Verificação do correcto funcionamento do equipamento. |  | CD2E28669606ED1B | No | Proof of the history of calibrations, QC, Measurements and errors of the system. Anomaly detection (poorly assembled inlet Clip and Ca and LAC errors in QC level B. Replacing the sprinkler and inlet probe, correct positioning of the input Clip. QC Pass level B. Verification of the correct operation of the equipment. |  | User Error | Inlet Clip | In correct install
ABL825 FLEX | I found the error 0500-Rinse detected in pH BG module when not expected. This is usually the tip broken off of the air valve. The tip was broken. The Lac would not come in so I membraned it with customer stock. | Replaced the pH BG air valve. Membraned Lac with customer stock. | CD2E434296006619 | No | I found the error 0500-Rinse detected in pH BG module when not expected. This is usually the tip broken off of the air valve. The tip was broken. The Lac would not come in so I membraned it with customer stock. | Replaced the pH BG air valve. Membraned Lac with customer stock. | Defect | pH/BG Module | Valve, air | Valve
ABL800Basic |  | Austausch des DMS Covers | CD2E435C96472EB9 | No |  | Replacing the DMS Cover | User Error | Cover Broken
ABL835 FLEX |  |  | CD2E6D7B964D352E | No |  |  | Adjustment | HW | Oxi Module | Lamp Unit
AQT90 FLEX |  | Reinigung | CD2E3D93964031A0 | No |  | Cleaning
ABL80 CO-OX |  |  | CD2F62380C4C84ED | No
ABL837 FLEX |  |  | CD2F96C20C1FA5B3 | No |  |  | Defect | Electrode | Membrane | Reference
ABL825 FLEX | Tried cleaning hemolyzer and tHb calibration in service-No change | Replaced hemolyzer Alignment and tHb calibration in service Complete oxi test-OK | CD2F5E360C4C5725 | No | Tried cleaning hemolyzer and tHb calibration in service-No change | Replaced Hemolyzer Alignment and tHb calibration in service Complete oxygen test-OK | Defect | HW | Oxi Module | Hemolyzer | spectrum mismatch
ABL800Basic |  |  | CD30B941821325B8 | No |  |  | Defect | Electrode | Electrode | Reference
ABL90 |  |  | CD3E788E9454C993 | No
ABL90 |  |  | CD35F7F55A123428 | No |  |  | Defect | HW | Barcode Reader
ABL90 | Lösungspack nicht richtig aktiviert. Tubing for valve undicht. | LP richtig aktiviert. | CD3606995A4D264D | No | Solution pack not activated correctly. Tubing for valve leaks. | LP is activated correctly. | Defect | HW | Oxi Module | Pinch Valve tube
TCM-COMBIM | Sustitución del modulo según service tip 629941548, el equipo se queda con el mensaje constantemente en la pantalla, se comprueba con el nuevo modulo funcionando correctamente. | Cambio del modulo combi. | CD35FDF85A2FECE6 | No | Replacing the module according to service tip 629941548, the computer keeps the message constantly on the screen, it is checked with the new module working properly. | Change of combi module.
ABL825 FLEX | SE DETECTAN COAGULOS QUE OBSTRUYEN TUBOS DE DESECHOS PROVOCANDO SALIDA DE DESECHOS POR ENTRADA, ASI MISMO SE DETECTA TUBO PERFORADO POR MALA MANIPULACION AL INTENTAR SOLUCIONAR EL PROBLEMA. SE RESTITUYE TUBO Y SE EFECTUA LIMPIEZA DEL CIRCUITO DE DESECHOS Y SE DOTA DE ENTRADA NUEVA. PRUEBAS DE FUNCIONAMIENTO CON MUESTRAS Y CONTROLES DE CALIDAD. | LIMPIEZA | CD362B655A26C352 | No | COAGULOS ARE DETECTED THAT OBSTRUCT THE DESTY TUBES CAUSED OUT OF INPUT, THIS IS DETECTED TUBE PERFORATED BY MALA MANIPULATION WHEN TRYING to SOLVE THE PROBLEM. TUBE RESTITUDE AND CLEANING OF THE CIRCUIT OF DISCLOSURES IS PERFORMED AND NEW INPUT IS AVAILABLE. OPERATION TESTS WITH SAMPLES AND QUALITY CONTROLS. | Cleaning | Cleaned | Fluidic pathway | Waste
ABL825 FLEX |  |  | CD35CBE65A5F3424 | No |  |  | Defect | pH/BG Module | Valve, air | Membrane
ABL90 |  |  | CD37570BD0597A38 | No |  |  | Defect | SP
ABL90 |  |  | CD37559BD0467A34 | Yes |  |  | Re classified RMED | Courtesy
ABL90 |  |  | CD529F9DF444B99E | No |  |  | Setup | HIS/LIS
ABL90 |  |  | CD371181D049679C | Yes |  |  | Re classified RMED | SWU
ABL90 | pO² in Lvl. C Falsch hoch. Waste tubing undicht. | Waste tubing erneuert. | CD3750C7D05B265D | No | pO2 in Lvl. C Falsehigh. Waste tubing leaks. | Waste tubing erneuert. | Defect | Hw | Waste tube
ABL90 | Installed wet-section PCB. No further hardware errors. Problems with fully bringing internal QC levels to within spec. | Cleaned, flushed, replaced SC, ran multiple blood samples, Ultimately brought QC into spec. | CD378299D0537165 | No | Installed wet-section PCB. No further hardware errors. Problems with fully bringing internal QC levels to within spec. | Cleaned, flushed, replaced SC, ran multiple blood samples, Ultimately brought QC into spec. | Defect | HW | Wet sec PCB
ABL90 |  |  | CD3732CED0441DA4 | No |  |  | Defect | HW | Inlet Gasket
ABL837 FLEX | Fails inlet calibration in service. Triweed new inlet gasket, inlet probe all does not help. | Ordered Inlet PCB and friction PCB. Will be back in AM | CD370015D03F5619 | No | Fails inlet calibration in service. Triweed new inlet gasket, inlet probe all does not help. | Ordered Inlet PCB and friction PCB. Will be back in AM | Unknown | Inlet module | Calibration
ABL835 FLEX | siehe arbeiten | siehe arbeiten | CD3707DFD0012C49 | No | see working | see working | Defect | Electrode | Membrane | Lac
ABL825 FLEX | checked data logs for errors, checked samples path, checked electrode and membrane, | performed lactate electrode chamber decontamination, performed zero current cal, replaced electrode membrane, changed sample size from 95ml to 195ml, | CD379089D032F9A0 | No | checked data logs for errors, checked samples path, checked electrode and membrane, | performed lactate electrode chamber decontamination, performed zero current cal, replaced electrode membrane, changed sample size from 95ml to 195ml, | Setup | Sample Size | Unstable | Lac
ABL825 FLEX | checked data logs for errors, checked samples path, checked electrode and membrane, | performed lactate electrode chamber decontamination, performed zero current cal, replaced electrode membrane, changed sample size from 95ml to 195ml, | CD378E7ED04EF99C | No | checked data logs for errors, checked samples path, checked electrode and membrane, | performed lactate electrode chamber decontamination, performed zero current cal, replaced electrode membrane, changed sample size from 95ml to 195ml, | Setup | Sample Size | Unstable | Lac
AQT90 FLEX |  |  | CD3708DCD0351580 | No
ABL80 |  |  | CD3842AD4607F97F | No
ABL90 | Instrument battery fails to power when unplugged. | Replaced power supply and battery. Fully charged and tested battery operation. Calibrated and ran QC. Instrument operates normally. | CD3874A9465690F8 | No | Instrument battery fails to power when unplugged. | Replaced power supply and battery. Fully charged and tested battery operation. Calibrated and ran QC. Instrument operates normally. | Defect | HW | Battery module
ABL90 | INSPECTED ALL FLUID PATHS FOR CLOT | REMOVED WET SECTION AND INSTALLED ALL TUBINGS AND SEALS. FLUSHED INLET SENSOR AND RECALIBERATED. | CD38CABB461D4477 | No | INSPECTED ALL FLUID PATHS FOR CLOT | REMOVED WET SECTION AND INSTALLED ALL TUBINGS AND SEALS. FLUSHED INLET SENSOR AND RECALIBERATED. | Cleaned | Fluidic pathway | LS error
ABL90 |  |  | CD3874AD46570844 | No |  |  | Cleaned | Fluidic pathway | and several other actions
ABL90 | ne peut-aspir solution du pack de solution, inlet module n'est pas aligné avec le cone du pack de solution | Replaced inlet module | CD386F6B46298C40 | No | can't-aspir solution pack solution, inlet module is not aligned with the solution pack cone | Replaced inlet module | Defect | HW | Inlet module
ABL90 | Previously changed wet-section PCB. Allowed analyzer to set overnight to confirm Cal and QC in AM. | Confirmed fix. | CD387FCA4634163F | Yes | Previously changed wet-section PCB. Allowed analyzer to set overnight to confirm Cal and QC in AM. | Confirmed fix. | Re classified RMED | upSWO to confirm previous upSWO
ABL800Basic |  |  | CD3847C6461E1B89 | No |  |  | Cleaned | pH/BG Module | Measuring chamber
ABL837 FLEX | Fails inlet calibration in service program | Replaced defective friction detector. Now passes inlet calibration test. Replaced Creatine membranes, multiple 2 pt cals for the creatine to come in. | CD385D3B4627F70C | No | Fails inlet calibration in service program | Replaced defective friction detector. Now passes inlet calibration test. Replaced Creatine membranes, multiple 2 pt cals for the creatine to come in. | Defect | Inlet module | Friction detector unit
AQT90 FLEX |  |  | CD384452464A7471 | No
ABL90 | Battery percent sign at 0. Reset the capacitor on the battery but nothing changed. Old Battery S N P0031N033. LIS not propperly communicating with analyzer. Checked all settings and connections. Checked patient lookup setup and sample pre-registration setup. | Replaced Battery module. LIS receives data from analyzer correctly but does not send back demographics to the Analyzer back. Meditech is investigating the programming of the interface. | CD397762BC197758 | No | Battery percent sign at 0. Reset the capacitor on the battery but nothing changed. Old Battery S N P0031N033. LIS not propperly communicating with analyzer. Checked all settings and connections. Checked patient lookup setup and sample pre-registration setup. | Replaced Battery module. LIS receives data from analyzer correctly but does not send back demographics to the Analyzer back. Meditech is investigating the programming of the interface. | Defect | HW | Battery module
ABL90 | Detecção da anomalia, coagulo na sonda entrada. Remoção do coágulo. Teste de detecção e remoção de coágulos. Ajuste da entrada. Calibração da bomba. Ajuste dos sensores de líquidos. Teste de pressão. Inicialização. Verificação do correcto funcionamento. |  | CD3985B4BC4445BB | No | Anomaly detection, clot in the inlet probe. Clot removal. Test for detection and removal of clots. Input adjustment. Pump calibration. Adjustment of liquid sensors. Pressure test. Startup. Verification of proper operation. |  | Cleaned | Fluidic pathway | Clot
ABL825 FLEX |  |  | CD39D284BC337229 | Yes |  |  | Re classified RMED | PM
AQT90 FLEX |  |  | CD3987FABC59DE78 | No
ABL80 OSM | Problem was found by the product specialist's during correlation. A tHb calibration would temporaraily resolve the problem. Troubleshot to a faulty hemolyzer. | Replaced the hemolyzer and this resolved the problem. | CD3B0FD1324AC376 | No | Problem was found by the product specialist's during correlation. A tHb calibration would temporaraily resolve the problem. Troubleshot to a faulty hemolyzer. | Replaced the hemolyzer and this resolved the problem.
ABL90 |  |  | CD3AC8BC32471711 | No |  |  | Defect | SC | Clot
ABL90 | Comprobar SC y SP, cambiar SC, calibrar y pasar controles. | Cambiar SC. | CD3ADE4E322FAA24 | No | Check SC and SP, change SC, calibrate and pass controls. | Change SC. | Defect | SC | pO2
ABL825 FLEX |  |  | CD3AC8283248ACFE | Yes |  |  | Re classified RMED | PM
ABL825 FLEX | Getting a 0258 Neon lamp intensity too low and 0316 number too low. This is usually the OXI control PCB. | Replaced the OXI control PCB. Adjusted lamp and ran complete OXI test aqnd tHb cals. | CD3ACDC1325F51C0 | No | Getting a 0258 Neon lamp intensity too low and 0316 number too low. This is usually the OXI control PCB. | Replaced the OXI control PCB. Adjusted lamp and ran complete OXI test aqnd tHb cals. | Defect | HW | Oxi Module | OXI PCB | Lamp voltage
ABL820 FLEX | Upon arrival, found unit leaking at inlet. | Flushed out waste lines on unit under inlet gasket. | CD3B0B25320FA61F | No | Upon arrival, found unit leaking at inlet. | Flushed out waste lines on unit under inlet gasket. | Cleaned | Inlet module | Nipple kit
ABL800Basic |  |  | CD3AE4B732171717 | No |  |  | Adjustment | HW | Oxi Module | Lamp Unit
AQT90 FLEX |  |  | CD3AD8EF32155703 | No
ABL90 |  |  | CD3E94199430B08A | No |  |  | Replaced | Probe | Leak
ABL825 FLEX |  |  | CD3EBFDD945BBC84 | Yes |  |  | Re classified RMED | PM
ABL825 FLEX |  |  | CD3EC1DC940CD723 | No |  |  | Adjustment | El/Met Module | Zero calibration
ABL90 | Everytime the customer would run a blood gas it would want to run a cal to optimize the result. Found that pinch valve tubing was | Replaced | CD413DCC800C4FBC | No | Everytime the customer would run a blood gas it would want to run a cal to optimize the result. Found that pinch valve tubing was | Replaced | Defect | HW | Oxi Module | Pinch Valve tube
ABL90 | In service required to do all the fluidic programs. this seemed to clear all the issues. made sure inlet was clean. All good. Checked data logs for errors. All good. | Ran fluidic programs in service mode. | CD3FE7F50A2B4FB7 | No | In service required to do all the fluidic programs. this seemed to clear all the issues. made sure inlet was clean. All good. Checked data logs for errors. All good. | Ran fluidic programs in service mode. | Unknown | Revisit required
ABL90 | Issues with fluidic errors and replacing multiple solution packs to resolve the fluidic errors. Fluidic errors seem to have escalated in the past 4 weeks. Checked fluidic integrity. Possible issue with solution packs not being pushed down enough when activating. | Replaced inlet connecor, inlet gasket with holder, inlet probe and tubing for valve. Instructed lab personnel on the proper method to activatesolution pack. reviewed the importance of using the inlet clip on all syring samples until new probe is available. Discussed fluidic errors and trouble shooting tips. | CD3FEE760A4DC25B | No | Issues with fluidic errors and replacing multiple solution packs to resolve the fluidic errors. Fluidic errors seem to have escalated in the past 4 weeks. Checked fluidic integrity. Possible issue with solution packs not being pushed down enough when activating. | Replaced inlet connecor, inlet gasket with holder, inlet probe and tubing for valve. Instructed lab personnel on the proper method to activatesolution pack. reviewed the importance of using the inlet clip on all syring samples until new probe is available. Discussed fluidic errors and trouble shooting tips. | Defect | HW | Oxi Module | Pinch Valve tube
ABL90 |  |  | CD5156457E3DB1C6 | No |  |  | Replaced | Inlet probe
ABL90 |  |  | CD3FA60C0A43D511 | No |  |  | Defect | SC | pCO2
ABL90 | Repuestos enviados por agencia | Repuestos enviados por agencia | CD3FEF500A265296 | Yes | Spare parts sent by agency | Spare parts sent by agency | Re classified RMED | Logistic upSWO
ABL835 FLEX | Sustitución de membranas de referencia y sodio. Calibración y comprobación de funcionamiento mediante Qc. |  | CD506158080B9B82 | No | Replacement of reference membranes and sodium. Calibration and operation check using Qc. |  | Defect | Electrode | Membrane | Reference
ABL825 FLEX |  |  | CD4034D20A1BBA0E | No |  |  | Defect | El/Met Module | PCB
ABL825 FLEX | Elektroden und Ventile | Austausch der Airventile; Reinigung Ref und Neumembranierung Ref; CO2 neu in Membran eingesetzt | CD3FE0210A08D96F | No | Electrodes and valves | Replacement of the air valves; Cleaning Ref and Remembraneation Ref; CO2 newly inserted in membrane | Defect | pH/BG Module | Valve, air | Valve
ABL835 FLEX | Pco2 sensibilit 45% et maintenance Bi-annuel | Effecué la maintenance bi-annuel selon les normes remplacé les tubes de pompe nettoyé pompes et flush v7.v8full oxi test et calibration service THB calibration nettoyé les pompes vérifié electronique calibration zero po2 ,lactate . nettoyé autocheck laser. et remplacé electrode pco2 | CD3FD6C90A2324D8 | No | Pco2 sensitivity 45% and maintenance Bi-annual | Effecué bi-annual maintenance according to standards replaced pump tubes cleaned pumps and flush v7.v8full oxi test and calibration service THB calibration cleaned pumps verified electronic calibration zero po2, lactate. cleaned laser autocheck. and replaced electrode pco2 | Defect | Electrode | Electrode | pCO2
TCM-COMBIM | Initially, analyzer apparently was requesting membrane changes frequently. When the monitor arrived, it was not recognizing the CombiM module. Tried different module and sensor. Tried new interface board and ran instrument overnight. Tried new interface board with new module and sensor. | Changed CF card. | CD400A210A3B43CD | No | Initially, analyzer apparently was requesting membrane changes frequently. When the monitor arrived, it was not recognizing the CombiM module. Tried different module and sensor. Tried new interface board and ran instrument overnight. Tried new interface board with new module and sensor. | Changed CF card. | Defect | HW | CF card
ABL90 |  |  | CD4225AFF60931CE | No |  |  | Defect | HW | Inlet Gasket | Locking mechanism
ABL90 | Comprobar SP y SC, cambiar la junta de entrada, calibrar y probar con controles. Limpiar el circuito neumático, quitar el polvo del interior, comprobar el registro de errores, guardar y borrar datos, calibrar. | Cambiar la entrada de muestras. | CD414ACD800D9012 | No | Check SP and SC, change the input joint, calibrate and test with controls. Clean the pneumatic circuit, remove dust from the inside, check the error log, save and erase data, calibrate. | Change the sample entry. | Replaced | Inlet gasket w holder
ABL825 FLEX | 6-3-14 I worked with Jim Halbleib and Vince Sigmund yesterday concerning the flash drives. The analyzer would no recognize the 2 flashdrives sent by Sendx. Becky said the display was freezing up and flickering. When I arrived I put the 2 flashdrives from Sendx on the analyzer and down loaded the data logs onto ezch. They worked fine. I believe they just needed time to load the drivers. I replaced the display unit with one the Brian Vazquez. This is qa used one and not on inventory. Will not claim as part. Worked with customer on trouble shooting before they start the testing. | Replaced the complete display unit and both flashdrives are working. | CD414FAB805F51D5 | No | 6-3-14 I worked with Jim Halbleib and Vince Sigmund yesterday concerning the flash drives. The analyzer would no recognize the 2 flashdrives sent by Sendx. Becky said the display was freezing up and flickering. When I arrived I put the 2 flashdrives from Sendx on the analyzer and down loaded the data logs onto ezch. They worked fine. I believe they just needed time to load the drivers. I replaced the display unit with one the Brian Vazquez. This is qa used one and not on inventory. Will not claim as part. Worked with customer on trouble shooting before they start the testing. | Replaced the complete display unit and both flashdrives are working. | Defect | DMS PC | Display
ABL825 FLEX | Elektroden und Klappen | Reinigung der Messkammern Ersatz oberer Dichtungsring bei Glu | CD40F817800BD707 | No | Electrodes and flaps | Cleaning of the measuring chambers Replacement of upper sealing ring at Glu | Cleaned | Inlet module | Used analyzer
ABL825 FLEX |  |  | CD4114B6803BBE5B | No |  |  | Defect | Auto Check Module | Breaker arm | Motor
ABL800Basic |  |  | CD4107CC8009371D | No |  |  | Defect | Electrode | Membrane | Reference
ABL800Basic |  |  | CD4106EB802C3718 | No |  |  | Cleaned | pH/BG Module | Measuring chamber
ABL800Basic |  |  | CD4105F4804D3712 | No |  |  | Cleaned | pH/BG Module | Measuring chamber
ABL820 FLEX | I got called by the product specialist and there was heavy build up all over the machine. I came on site and it was all over every module and inside the fluidic block. I found a small hole in the waste tube going from the y piece and the waste bottle right next to the y piece. | I had to remove every module and scrub the machine down using sani wipes. I cut the tube down 1 8th inch and it resolved the problem. | CD4248C0F60C7B73 | No | I got called by the product specialist and there was heavy build up all over the machine. I came on site and it was all over every module and inside the fluidic block. I found a small hole in the waste tube going from the y piece and the waste bottle right next to the y piece. | I had to remove every module and scrub the machine down using sani wipes. I cut the tube down 1 8th inch and it resolved the problem. | Defect | Basic unit | Waste tube | Fluidic module
ABL825 FLEX | sO2 and HBCO both reading on the high end of package insert | Performed lamp adjustment | CD41431780490943 | No | sO2 and HBCO both reading on the high end of package insert | Performed lamp adjustment | Adjustment | HW | Oxi Module | Lamp Unit
ABL825 FLEX |  |  | CD42BBC0F637CCB2 | No |  |  | Adjustment | HW | Oxi Module | Lamp Unit
ABL825 FLEX | B) The analyzer Hb would not pass calibration. Found lamp intensity over 3900, Adjusted to within limits. Had error 0316 Neon number too low. Replaced lamp assembly and had same error. Replaced the controller board and still had same error. The spectrometer needs to be replaced. | A) Setup, programmed, aligned, calibrated and checked operation of new analyzer. B) Replaced the spectrometer which resolved the error. User is familiar with the analyzer so no training required. | CD41735D800515C5 | No | B) The analyzer Hb would not pass calibration. Found lamp intensity over 3900, Adjusted to within limits. Had error 0316 Neon number too low. Replaced lamp assembly and had same error. Replaced the controller board and still had same error. The spectrometer needs to be replaced. | A) Setup, programmed, aligned, calibrated and checked operation of new analyzer. B) Replaced the spectrometer which resolved the error. User is familiar with the analyzer so no training required. | Defect | HW | Oxi module | Spectrometer
AQT90 FLEX |  |  | CD415D0980042173 | No
AQT90 FLEX | Separador de quesitos en la storagewheel mal posicionado. Se recoloca y se comprobueba funcionamiento. Pasados QC Multichech lvl 1 y muestras de TcT de pacientes. Todo ok |  | CD410FE1805C44F0 | No | Quesitos separator on the poorly positioned storagewheel. It is repositioned and performance-compropate. Passed QC Multichech lvl 1 and TcT samples from patients. All ok
ABL90 |  |  | CD425822F62F4835 | No |  |  | Defect | HW | Waste tube
ABL825 FLEX |  |  | CD428A2DF663580F | Yes |  |  | Re classified RMED | Logistic upSWO
ABL825 FLEX |  |  | CD4282F0F62DA792 | No |  |  | Defect | DMS PC | HDD
ABL825 FLEX | Investigate the issue, monitor electronic result and test channel in service. current results for service program are within expected range. | Perform a complete decontamination per service manual. Replace membranes and solutions. Install new module pump tubes and inlet gasket from customer supplies. | CD427777F660B3FA | No | Investigate the issue, monitor electronic result and test channel in service. current results for service program are within expected range. | Perform a complete decontamination per service manual. Replace membranes and solutions. Install new module pump tubes and inlet gasket from customer supplies. | Defect | Electrode | Membrane | Lac
ABL815 FLEX |  |  | CD428CE1F64D8A8B | No |  |  | Defect | Inlet module | Inlet Tube (Blue)
ABL827 FLEX | checked data logs for errors and replacements, checked valves and liquid sensors, checked electrodes and membranes, checked electrode chambers, checked pump tubing and rotors, | removed and replaced all pump tubing, replaced all electrode membranes, freshened connector tubing, removed all dual housing and cleaned all valves, replaced PH met air valve membranes, adjusted QC ranges to better reflect peer group range, | CD42DC59F6359F26 | No | checked data logs for errors and replacements, checked valves and liquid sensors, checked electrodes and membranes, checked electrode chambers, checked pump tubing and rotors, | removed and replaced all pump tubing, replaced all electrode membranes, freshened connector tubing, removed all dual housing and cleaned all valves, replaced PH met air valve membranes, adjusted QC ranges to better reflect peer group range, | Defect | pH/BG Module | Valve, air
ABL835 FLEX | Inlet calibration error intermittant et inlet coule | Remplac inlet probe tube inlet et de la probe | CD425CD8F658035F | No | Inlet calibration error intertermittant and inlet flows | Remplac inlet probe tube inlet et de la probe | Defect | Inlet module | Probe | Calibration
ABL835 FLEX | Inlet calibration error intermitly and inlet leaking | Replaced friction detector no help inlet pcb and stepper motor resolve issue for now. | CD5019C0082C6907 | No | Inlet calibration error intermitly and inlet leaking | Replaced friction detector no help inlet pcb and stepper motor resolve issue for now. | Defect | Inlet module | Inlet PCB
Tosca-500 | Datum stand auf 01.01.2001, ließ sich z. Zt. aber auf 02.06.2014 einstellen! Funktions-test Sensor 2:09 min, o.k., Gerätefehler hier nicht feststellbar, trotzdem defekt. | TechLine.: RDEU 686 882 915 erstellt, Sensor bitte zurück zum Kunden senden. | CD4221BAF6350154 | No | Date was on 01.01.2001, but could be set to 02.06.2014 at the time! Function test sensor 2:09 min, ok, device error here not detectable, nevertheless defective. | TechLine.: RDEU 686 882 915 created, sensor please send back to customer. | Defect | HW | PCB
ABL80 OSM | During my first visit I replaced the CF card, but this only worked for a short period. Troubleshot to a faulty computer board. | Replaced the 945 computer board, and this resolved the problem. | CD4409B76C4C44B9 | No | During my first visit I replaced the CF card, but this only worked for a short period. Troubleshot to a faulty computer board. | Replaced the 945 computer board, and this resolved the problem.
ABL90 |  |  | CD4413076C2F12B9 | No |  |  | Cleaned | Fluidic pathway
ABL90 |  |  | CD4A17CFBA0E35F8 | No |  |  | Unknown | Display | Settings?
ABL90 |  |  | CD43A25F6C1D3C70 | No |  |  | Defect | hw | Battery module
ABL800Basic |  |  | CD4392F06C4E1C04 | No |  |  | Defect | HW | Oxi Module | OXI PCB
ABL835 FLEX | THE OXI MODULE THAT CLAUDE HAD REMOVED FROM THE INSTRUMENT TO RESOLVE A PREVIOUS PROBLEM WAS REINSTALLED. | REPLACEMENT OF LOANER OXI MODULE | CD437E026C3F1471 | Yes | THE OXI MODULE THAT CLAUDE HAD REMOVED FROM THE INSTRUMENT TO RESOLVE A PREVIOUS PROBLEM WAS REINSTALLED. | REPLACEMENT OF LOANER OXI MODULE | Re classified RMED | Revisit to replace loaner
ABL825 FLEX | Also found this analyzer has had MetHb disabled due to previously failing to compare to peer on blood comp's. I was unaware of the MetHb issue so I worked on this as well. I enabled the MetHb parameter and ran Oxi Complete test - pass. Ran service tHb Cal 2X - all values exactly on mean. I ran same on peer. Blood checks now compare favorably but I have a very small sample size. On the Hct results I see that all samples were run in 95uL mode. As per CAP document this analyzer should have proficiencies run at largest available sample size. This was the only issue | Oxi complete test, service tHb Cal | CD43ABF76C5924F9 | No | Also found this analyzer has had MetHb disabled due to previously failing to compare to peer on blood comp's. I was unaware of the MetHb issue so I worked on this as well. I enabled the MetHb parameter and ran Oxi Complete test - pass. Ran service tHb Cal 2X - all values exactly on mean. I ran same on peer. Blood checks now compare favorably but I have a very small sample size. On the Hct results I see that all samples were run in 95uL mode. As per CAP document this analyzer should have proficiencies run at largest available sample size. This was the only issue | Oxi complete test, service tHb Cal | User Error | Setup | Sample Size | CAP Survey
ABL825 FLEX | On the original Hct results, run on May 2, I see that all samples were run in 95uL mode. As per CAP documentation this analyzer should have proficiencies run at largest available sample size which is 195 uL. This was the only issue I could find at this time | Ran Oxi complete test and service tHb Cal, all pass. | CD43CA4F6C2324FC | No | On the original Hct results, run on May 2, I see that all samples were run in 95uL mode. As per CAP documentation this analyzer should have proficiencies run at largest available sample size which is 195 uL. This was the only issue I could find at this time | Ran Oxygen Complete test and service tHb Cal, all pass. | User Error | Setup | Sample Size | CAP Survey
ABL825 FLEX |  |  | CD4B40EF305A07ED | Yes |  |  | Re classified RMED | Not upSWO
ABL827 FLEX | Performed complete oximetry test and found hemolyzer phase regulation error. Need to replace the hemolyzer. | Replaced the hemolyzer unit. Performed complete oximetry test. Analyzer now passes all tests. | CD44CCFCE2107D19 | No | Performed complete oximetry test and found hemolyzer phase regulation error. Need to replace the hemolyzer. | Replaced the hemolyzer unit. Performed complete oximetry test. Analyzer now passes all tests. | Defect | HW | Oxi Module | Hemolyzer | outside Peer group
AQT90 FLEX |  |  | CD4643D15846BCA7 | No
ABL90 | Detecting leak and pump not able to pump solution. Found waste connector partially blocked. | Install version 3.1 software upgrade. Replace waste connector. Perform back flush procedure. Perform annual maintenance and performance check. | CD479208CE14416D | No | Detecting leak and pump not able to pump solution. Found waste connector partially blocked. | Install version 3.1 software upgrade. Replace waste connector. Perform back flush procedure. Perform annual maintenance and performance check. | Cleaned | Fluidic pathway | Waste connector blocked
ABL90 | Error code 1343 -Solutions not pumping | Flush analyzer using flush kit. | CD47B8E7CE23E176 | No | Error code 1343 -Solutions not pumping | Flush analyzer using flush kit. | Cleaned | Fluidic pathway
ABL825 FLEX | Checked AutoCheck barcode scanner. Several ampoules did not read. | Clean AutoCheck barcode scanner mirror assemble. Retest scanner. | CD4742AACE54DBE6 | No | Checked AutoCheck barcode scanner. Several ampoules did not read. | Clean AutoCheck barcode scanner mirror assemble. Retest scanner. | Cleaned | Auto Check Module | Barcode Reader | Mirror
ABL800Basic |  |  | CD47C9A7CE23C61F | No |  |  | Defect | Electrode | Membrane | Reference
ABL827 FLEX |  |  | CD48CC9C4403438F | No |  |  | Defect | Electrode | Electrode | Na+
ABL825 FLEX |  |  | CD476A22CE0B00C7 | No |  |  | Adjustment | pH/BG Module | Electrode connector | Contact bent
ABL80 CO-OX |  |  | CD4AFEE03051C56C | No
ABL80 OSM | The CF card and 945 Computer Board have already been changed with no resolution to the problem. After consulting with my area specialist, we decided that the screen or the power supply must be causing the problem. | Replaced the touch screen and power supply. This resolved the problem. | CD48BB7144637F95 | No | The CF card and 945 Computer Board have already been changed with no resolution to the problem. After consulting with my area specialist, we decided that the screen or the power supply must be causing the problem. | Replaced the touch screen and power supply. This resolved the problem.
ABL90 |  |  | CD48CA754431A6E2 | No |  |  | Defect | SC | Glu
ABL90 | Followup to previous workorder # B96X5PXYB -After flushing, analyzer came up but failed on several parameters during calibrations. | Replaced sensor cassette after it arrived. Ran blood samples. | CD48E32A444EC459 | No | Followup to previous workorder # B96X5PXYB -After flushing, analyzer came up but failed on several parameters during calibrations. | Replaced sensor cassette after it arrived. Ran blood samples. | Defect | SC | Previous upSWO should have replaced
ABL90 |  |  | CD489FA9440C1E50 | No |  |  | Defect | SP | FT
ABL835 FLEX |  | pH Modul ausgebaut , flache Ventilmembran im Air-Ventil war ausgestanzt, ersetzt, Kal, QC | CD48A4D1444AC7D5 | No |  | pH module removed, flat valve membrane in the air valve was punched out, replaced, Kal, QC | Defect | pH/BG Module | Valve, air | Membrane
ABL820 FLEX |  |  | CD48DF234403CB16 | No |  |  | Defect | HW | Oxi Module
ABL835 FLEX | Sustitucion de la aguja. Cal, qc y muestras | Sustitucion de la aguja. Cal, qc y muestras | CD48C3EE44209255 | No | Needle replacement. Cal, qc and samples | Needle replacement. Cal, qc and samples | Defect | Inlet module | Probe
ABL827 FLEX | Rinse flow is not correct through the module valves. Took pH BG valve apart and found defective air valve membrane. Need to replace the membrane. | Replaced the air valve membrane and the solutions now flow as they should. | CD48D67244265323 | No | Rinse flow is not correct through the module valves. Took pH BG valve apart and found defective air valve membrane. Need to replace the membrane. | Replaced the air valve membrane and the solutions now flow as they should. | Defect | pH/BG Module | Valve, air | Membrane
ABL820 FLEX | Checked halogen lamp intensity and found voltage outside the acceptable limits. | Adjusted the halogen lamp voltage to specifications. | CD48C385445037C3 | No | Checked halogen lamp intensity and found voltage outside the acceptable limits. | Adjusted the halogen lamp voltage to specifications. | Adjustment | HW | Oxi Module | Lamp Unit
ABL825 FLEX | Reviewed data logs. tHb status on cals are low:445pA | Complete OXI alignment, passes all OXI tests in service. tHB cal done in service | CD48B5CC4422DF05 | No | Reviewed data logs. tHb status on cals are low:445pA | Complete OXYGEN Alignment, passes all OXYGEN tests in service. tHB Cal done in service | Adjustment | HW | Oxi Module | Complete Oxi
ABL80 OSM | Larry Logar asked that i come out and down load the data logs from the analyzer and open a techline case. I down loaded the WDC report and the analyzer did not include the manual QC logs even though they were present on the analyzer. The manual QC logs were from April through May. | Gathered data logs and opened techline case. Sent data logs to Larry logar and support email. | CD4A1B7EBA020FA5 | No | Larry Logar asked that i come out and down load the data logs from the analyzer and open a techline case. I down loaded the WDC report and the analyzer did not include the manual QC logs even though they were present on the analyzer. The manual QC logs were from April through May. | Gathered data logs and opened techline case. Sent data logs to Larry logar and support email.
ABL90 |  |  | CD4A05E6BA2EDBCB | No |  |  | Replaced | Inlet gasket
ABL90 |  |  | CD49C1DBBA0126AD | Yes |  |  | Re classified RMED | Hotline call
ABL90 | Unit had 1066 reference substrate temperture errors. Also 1361 internal reference electrode. | Customer replaced sensor and unit was OK on arrival. | CD49D2CBBA1AAB27 | Yes | Unit had 1066 reference substrate temperture errors. Also 1361 internal reference electrode. | Customer replaced sensor and unit was OK on arrival. | Re classified RMED | Courtesy
ABL835 FLEX |  |  | CD4A0F5ABA1E9C53 | No |  |  | Defect | HW | Inlet module | Leaking
ABL820 FLEX |  |  | CD49C690BA1BC3C2 | No |  |  | Defect | HW | Oxi Module | - | phase reg
ABL815 FLEX |  |  | CD4A15DCBA57DA81 | No |  |  | Defect | DMS PC | Printer head
ABL825 FLEX |  |  | CD4C6F08A65ACFC3 | No |  |  | Cleaned | Basic unit | Barcode Reader
ABL820 FLEX | When looking at various menus, the monitor would show lines and flashes across the screen | Replace display unit. Also decontaminated,descrambled patient records, packed and shipped old analyzers (with David's help) which had been placed out of service or inactive, but hadn't been shipped. Tracking # 7701 5886 5209 and 7701 5912 2486 | CD4A6724BA30BC74 | No | When looking at various menus, the monitor would show lines and flashes across the screen | Replace display unit. Also decontaminated,descrambled patient records, packed and shipped old analyzers (with David's help) which had been placed out of service or inactive, but hadn't been shipped. Tracking # 7701 5886 5209 and 7701 5912 2486 | Defect | DMS PC | Display
AQT90 FLEX |  |  | CD4A2F73BA1C99CF | No
ABL90 | Troubleshot using service programs. Pump,Liquid Sensors, Pressure Test, and Clot Detection all failed. Back flushed system - same. Removed sensor and hooked up piece of tubing to run Hypochlorite solution. Followed with flush of water - same. Replaced solution pack with one from customer. Will send in credit form. | Removed and replaced sensor cassette from customer's supply. Powered down Analyzer and turn back on. Will send in credit form for customer. | CD4B7E1430345265 | No | Troubleshot using service programs. Pump,Liquid Sensors, Pressure Test, and Clot Detection all failed. Back flushed system - same. Removed sensor and hooked up piece of tubing to run Hypochlorite solution. Followed with flush of water - same. Replaced solution pack with one from customer. Will send in credit form. | Removed and replaced sensor cassette from customer's supply. Powered down Analyzer and turn back on. Will send in credit form for customer. | Defect | SP | FT
ABL90 | The analyzer showed OXI issues and then the QC would keep going out. the cal's looked good but QC was failing Po2 and Na. | It looked like a clot but there was no samples ran around the time of the first OXI QC error. I flushed the analyzer with the SC in and with it out. I flushed the SC by itself and then it would not pass cal due to Na sensor response errors. I replaced the SC and all came in but Lac and Glu status. I swaped the SP with another analyzer and it did nothing. The number was moving so I continued to cal and it came in. The QC would not come in wo I replaced the SP. | CD4B4FFE301460F6 | No | The analyzer showed OXI issues and then the QC would keep going out. the cal's looked good but QC was failing Po2 and Na. | It looked like a clot but there was no samples ran around the time of the first OXI QC error. I flushed the analyzer with the SC in and with it out. I flushed the SC by itself and then it would not pass cal due to Na sensor response errors. I replaced the SC and all came in but Lac and Glu status. I swaped the SP with another analyzer and it did nothing. The number was moving so I continued to cal and it came in. The QC would not come in wo I replaced the SP. | Defect | SC | Na+
ABL80 CO-OX | Set up analyzer and found that the CF Card did not survive the transport. Customer decided to purchase a new analyzer | Ran SVC Cleaning Pack; replaced the hemolyzer and installed PM Kit; replaced the CF Card and calibrated the touchscreen. | CD4C4840A62F8AAE | No | Set up analyzer and found that the CF Card did not survive the transport. Customer decided to purchase a new analyzer | Ran SVC Cleaning Pack; replaced the hemolyzer and installed PM Kit; replaced the CF Card and calibrated the touchscreen.
ABL90 | Comprovação do histórico de calibrações, QC, medidas e erros do sistema. Cópia de segurança de dados. O equipamento perde a conexão com o chip do cartucho daí pedir para o substituir. Limpeza dos contactos equipamento cartucho. Comprovação do correcto funcionamento com um cartucho substituído ontem. Limpeza do circuito da amostra. Substituição da sonda e porta de entrada. Inicialização. |  | CD4B24B2301A127E | No | Proof of the history of calibrations, QC, Measurements and errors of the system. Data backup. The equipment loses the connection to the cartridge chip and asks to replace it. Cleaning the contacts cartridge equipment. Proof of proper operation with a cartridge replaced yesterday. Cleaning the sample circuit. Replacement of probe and inlet port. Startup. |  | Cleaned | Fluidic pathway
ABL800Basic |  |  | CD4B1964302D5856 | No |  |  | Defect | Electrode | Electrode | pCO2
ABL835 FLEX | DATA LOGS SHOW 678 HEATER ERROR AND 356 EL MET MODULE TEMP ERROR.SERVICE PROGRAM SHOWS EL MET MODULE AT 26 DEGREES | REPLACED MEASURING LIGHT BOARD IN EL MET MODULE.PERFORMED MULTIPLE CALIBRATIONS | CD4B1F743018C8E4 | No | DATA LOGS SHOW 678 HEATER ERROR AND 356 EL MET MODULE TEMP ERROR. SERVICE PROGRAM SHOWS EL MET MODULE AT 26 DEGREES | REPLACED MEASURING LIGHT BOARD IN EL MET MODULE. PERFORMED MULTIPLE CALIBRATIONS | Defect | El/Met Module | Illumination PCB
ABL820 FLEX | Customer has been entering in barcodes with key pad. Cleaned reseated connections. No Fix. | Replaced internal barcode reader. Tested OK. | CD4B33A8302F312F | No | Customer has been entering in barcodes with key pad. Cleaned reseated connections. No Fix. | Replaced internal barcode reader. Tested OK. | Defect | Basic unit | Barcode Reader
ABL825 FLEX | Inlet nipple bent. | Replace inlet nipples. | CD4BB3BF3038BC4B | No | Inlet nipple bent. | Replace inlet nipples. | Defect | Inlet module | Nipple kit
ABL80 OSM | VK-R5 kit |  | CD4C74F2A65C52D7 | No | VK-R5 kit
ABL80 OSM | Waste not being pumped out of sensor. No errors in logs. | Opened unit checked tubing positions. Reseated connectors. Also set up both new analyzers for cath lab site samples. | CD4C6AA2A6363678 | No | Waste not being pumped out of sensor. No errors in logs. | Opened unit checked tubing positions. Reseated connectors. Also set up both new analyzers for cath lab site samples.
ABL90 |  |  | CD4C8C5CA652FE04 | No |  |  | Defect | HW | Waste tube
ABL90 |  |  | CD4C6E68A6037DCC | No |  |  | Cleaned | Fluidic pathway
ABL90 |  |  | CD4C6DEBA620EB3C | Yes |  |  | Re classified RMED | Logistic upSWO
ABL90 | SE ENVIAN PIEZAS POR AGENCIA DIURECTAMENTE A CLIENTE | SE ENVIAN PIEZAS POR AGENCIA DIURECTAMENTE A CLIENTE | CD55A7C9E0517659 | Yes | PARTS ARE SENT BY AGENCY WILL BE SENT TO CUSTOMER | PARTS ARE SENT BY AGENCY WILL BE SENT TO CUSTOMER | Re classified RMED | Logistic upSWO
ABL825 FLEX |  |  | CD4F5C93924D585C | No |  |  | Cleaned | El/Met Module | Measuring chamber
ABL837 FLEX | The analyzer is presenting error 580 Oximetry hardware problem and 259 Hemolyzer phase error on calibrations. | Visited to replace a new hemolyzer unit. Adjusted the lamp unit. Performed a Complete Oxi test, passing ok. Performed the customer and the service ctHb calibration, passing ok. | CD4C9250A6543A72 | No | The analyzer is presenting error 580 Oximetry hardware problem and 259 Hemolyzer phase error on calibrations. | Visited to replace a new hemolyzer unit. Adjusted the lamp unit. Performed a Complete Oxi test, passing ok. Performed the customer and the service ctHb calibration, passing ok. | Defect | HW | Oxi Module | Hemolyzer | phase reg
ABL820 FLEX | pCO2 ?? on random cals and QCs. Drifting, Status, Sens errors on cals. Low results on some QC. Found leaking in Reference electrode well. Electrolyte in reference membrane is pink in color. | Replaced reference electrode membrane. Lot 4113. Discussed new analyzers with Biomed. | CD4C554CA62CD527 | No | pCO2 ?? on random cals and QCs. Drifting, Status, Sens errors on cals. Low results on some QC. Found leaking in Reference electrode well. Electrolyte in reference membrane is pink in color. | Replaced reference electrode membrane. Lot 4113. Discussed new analyzers with Biomed. | Defect | Electrode | Membrane | Reference
ABL827 FLEX | Instrument has been overflowing measuring chamber well do to electrode membrane not being seated properly. Instrument needs all modules pulled to clean up waste. | Replaced Ca++ membrane and cleaned instrument. Calibrated and ran QC. Instrument operates normally. | CD4C8636A6042EA3 | No | Instrument has been overflowing measuring chamber well do to electrode membrane not being seated properly. Instrument needs all modules pulled to clean up waste. | Replaced Ca++ membrane and cleaned instrument. Calibrated and ran QC. Instrument operates normally. | User Error | Replacement process | Membrane | Overflow measuring chamber
ABL90 |  |  | CD501F9A08050B49 | No |  |  | User Error | Inlet probe not aligned w gasket
ABL90 |  |  | CD501DC90812C890 | No |  |  | Defect | HW | Battery module
ABL90 | Configuracion de la ip. Guardar en memoria con commit Chequeo del equipo por programa de servicio. Cal, qc |  | CD5058700861064D | Yes | Ip configuration. Save to memory with commit Computer check by service program. Cal, qc |  | Re classified RMED | Setup | IP
ABL825 FLEX |  |  | CD504D4108257E03 | No |  |  | Defect | Inlet module | Nipple kit
ABL837 FLEX | They use to get a month's life from membrane, but now only a week's life. | Will order a new Lactate electrode and starting a TechLine case. Downloaded data logs for case. | CD50342208395D84 | No | They use to get a month's life from membrane, but now only a week's life. | Will order a new Lactate electrode and starting a TechLine case. Downloaded data logs for case. | Defect | Electrode | Electrode | Lac
ABL837 FLEX | In service the OXI calibration fails. See crystals in air vent line | Flushed out air vent line. Now OXI LS calibrates. | CD501FD1081755B6 | No | In service the OXI calibration fails. See crystals in air vent line | Flushed out air vent line. Now OXI LS calibrates. | Cleaned | Fluidic Module | Air vent
AQT90 FLEX |  |  | CD504C0E08420B50 | No
ABL90 |  |  | CD5174437E377140 | No |  |  | Defect | HW | Waste Tube
ABL90 | Test in service clot detection program Fails the inlet check | Back flush , change listed parts and the SP Checked clot detection program in service-OK | CD51685C7E58B672 | No | Test in service clot detection program Fails the inlet check | Back flush , change listed parts and the SP Checked clot detection program in service-OK | Replaced | Probe
ABL90 |  |  | CD518E9B7E501992 | No |  |  | Adjustment | Reboot analyzer | Radiance
ABL90 |  |  | CD517DFE7E43E1FC | No |  |  | Replaced | Inlet Gasket
ABL90 |  |  | CD51A4207E203646 | No |  |  | Cleaned | Fluidic pathway | Clot
ABL820 FLEX |  | Ventile überprüft und gereinigt; LS Justage Pumpenkal | CD5196D17E3D8892 | No |  | Valves checked and cleaned; LS Adjustment Pump Scale | Defect | Auto Check Module | Carroussel
ABL800Basic | kein Fehlerfeststellbar bei Eintreffen Gerät kpl. grün |  | CD5142A57E217B0A | No | no error-detectable on arrival device kpl. green |  | No fault found | PM
ABL825 FLEX | Did a trend report on Ca++. QC comes in after membrane change but bottoms out after a couple of hours. | Removed all electrodes in both chambers. Cleaned with protein remover solution and cleaning solution. Membraned both the reference electrode and Ca++ electrode. Gave customer a box of Ca++ membranes. | CD518D827E2BF32F | No | Did a trend report on Ca++. QC comes in after membrane change but bottoms out after a couple of hours. | Removed all electrodes in both chambers. Cleaned with protein remover solution and cleaning solution. Membraned both the reference electrode and Ca++ electrode. Gave customer a box of Ca++ membranes. | Defect | Electrode | Membrane | Ca++
ABL835 FLEX | THE REPAIRED OXI MODULE WAS GIVING 269 ERROR CODES ON LEVEL 3 QC. | THE OXI MODULE WAS REPLACED | CD515BBD7E62504C | No | THE REPAIRED OXI MODULE WAS GIVING 269 ERROR CODES ON LEVEL 3 QC. | THE OXI MODULE WAS REPLACED | Defect | HW | Oxi Module | all
ABL835 FLEX | Display softkey freeze often | replaced display | CD517A6A7E347624 | No | Display softkey freeze often | replaced display | Defect | DMS PC | Display | Touch screen
ABL825 FLEX | Customer replaced inlet gasket but same error. | Put on a new Inlet Probe and calibrate. | CD515ECA7E08ACB5 | No | Customer replaced inlet gasket but same error. | Put on a new Inlet Probe and calibrate. | Defect | Inlet module | Probe
TCM-COMBIM |  |  | CD5A7CD2B8611442 | No |  |  | Defect | HW | Touch screen
ABL837 FLEX | No red light from lazar engine | Repositioned ribbon cable now have red light.. Ran scan in service, picks up all ampoules | CD52E3F2F443DA73 | No | No red light from lazar engine | Repositioned ribbon cable now have red light.. Ran scan in service, picks up all ampoules | Adjustment | Auto Check Module | Cable seating
ABL825 FLEX | Tried new membrane-No change Performed zero calibration in service-No change | Replaced pO2 electrode pO2 electrode defective LI-03 | CD52CD5EF4199587 | No | Tried new membrane-No change Performed zero calibration in service-No change | Replaced pO2 electrode pO2 electrode defective LI-03 | Defect | Electrode | Electrode | pO2
ABL825 FLEX | pCO2 sensitivity fluctuating between 73 and 85%. Status fluctuating between 65 and 118. | Cleaned measuring chamber. Cleaned pCO2 electrode tip. | CD52B7F9F44AF74E | No | pCO2 sensitivity fluctuating between 73 and 85%. Status fluctuating between 65 and 118. | Cleaned measuring chamber. Cleaned pCO2 electrode tip. | Cleaned | Electrode | Electrode | pCO2
ABL80 CO-OX | Upon arrival found unit failing Oximetry calibrations. | Replaced Hemolyzer unit. Performed complete service Oximetry testing.Realigned Hemolyzer. | CD53E9876A36FEAE | No | Upon arrival found unit failing Oximetry calibrations. | Replaced Hemolyzer unit. Performed complete service Oximetry testing. Realigned Hemolyzer.
ABL90 |  |  | CD53E24A6A5684B3 | No |  |  | Replaced | Probe | Clot
ABL90 |  | Solution pack ID PCB ausgetauscht | CD5418AB6A3B88BA | No |  | Solution pack ID PCB ausgetauscht | Defect | HW | SP PCB
ABL90 |  |  | CD540D506A498445 | No |  |  | Cleaned | Fluidic pathway | Clot
ABL90 |  |  | CD53F1DF6A4C843F | No |  |  | Defect | HW | Barcode Reader
ABL827 FLEX |  |  | CD5539BEE0584BDA | No |  |  | Defect | DMS PC | Printer head
ABL810 FLEX |  |  | CD557B79E012798F | No |  |  | Defect | Basic unit | Analyzer Control PCB
ABL90 | siehe SLR | siehe SLR | CD556060E057D219 | No | see SLR | see SLR | Setup | HIS/LIS
ABL835 FLEX |  | Austausch Ref Elektrode | CD551E94E04E88C3 | No |  | Replacement Ref Electrode | Defect | Electrode | Electrode | Reference
ABL800Basic |  |  | CD55726FE00B20A5 | No |  |  | Cleaned | pH/BG Module | Valves
ABL825 FLEX | Data logs show numerous pCO2 status drift and sense errors. Sense fails both high and low. I performed Gas reset procedure which appears to have corrected status drift. Still failing pCO2 sensitivity and QC. All 3 levels out high (above insert range) Customer has already changed membranes numerous times, no help. I changed to different lot code of membrane - no help. Customer had a used Electrode, lot ER-04. I installed the used electrode which is performing withing spec. Failing electrode is lot LR-01. | Gas Reset procedure and replaced electrode with used, but working electrode. Will order warranty replacement on Monday. | CD554C4DE0262556 | No | Data logs show numerous pCO2 status drift and sense errors. Sense fails both high and low. I performed Gas reset procedure which appears to have corrected status drift. Still failing pCO2 sensitivity and QC. All 3 levels out high (above insert range) Customer has already changed membranes numerous times, no help. I changed to different lot code of membrane - no help. Customer had a used Electrode, lot ER-04. I installed the used electrode which is performing withing spec. Failing electrode is lot LR-01. | Gas Reset procedure and replaced electrode with used, but working electrode. Will order warranty replacement on Monday. | Defect | Electrode | Electrode | pCO2
ABL825 FLEX | Protein, cleaning, pump cal and LS adjust-No change Replaced pH sealing ring-No change pH seal does not seem right | Moved pH electrode form ABL725 to ABL825. pH seal is fine now | CD557F74E05CD69A | No | Protein, cleaning, pump cal and LS adjust-No change Replaced pH sealing ring-No change pH seal does not seem right | Moved pH electrode form ABL725 to ABL825. pH seal is fine now | Defect | Electrode | Electrode | pH
ABL800Basic | siehe ausgeführte Arbeiten | siehe ausgeführte Arbeiten | CD551488E02E464B | No | see work carried out | see work carried out | Defect | Electrode | Membrane | pCO2
ABL825 FLEX | Inspected hemolyzer for leaks or damage and none was found.Ran lamp unit test and found reading at 3842 which high end is 3700. | Adjusted the lamp to 3554. | CD55402AE061F605 | No | Inspected hemolyzer for leaks or damage and none was found. Ran lamp unit test and found reading at 3842 which high end is 3700. | Adjusted the lamp to 3554. | Adjustment | HW | Oxi Module | Lamp Unit
ABL80 FLEX Basic version |  |  | CD5B544C2E4CF772 | Yes |  |  | Re classified RMED | SWU
ABL80 OSM | Checked flow of solutions, checked hemolyzer and liquid sensors, checked analyzer level on analyzer stand | analyzer tipped backwards, leveled analyzer, performed oximetry test, ran tHb cal, adjusted liquid sensors, ran 5-10 QC samples and 15-20 blood samples withou any errors, sent picture of tipped analyzer to tech support for evaluation | CD592FBA421D31B9 | No | Checked flow of solutions, checked hemolyzer and liquid sensors, checked analyzer level on analyzer stand | analyzer tipped backwards, leveled analyzer, performed oximetry test, ran tHb cal, adjusted liquid sensors, ran 5-10 QC samples and 15-20 blood samples withou any errors, sent picture of tipped analyzer to tech support for evaluation
ABL90 | Bay door tries to close, but them reverses it self and remains open. | Replaced lid closed indicator and upper lid. | CD590874423B0990 | No | Bay door tries to close, but them reverses it self and remains open. | Replaced lid closed indicator and upper lid. | Defect | HW | Detector Plate
ABL90 |  |  | CD58D324425CB5D2 | Yes |  |  | Re classified RMED | Hotline call | for Credit
ABL820 FLEX |  | Reinigung Airventil pH BG Modul und Spülung Airkanal | CD590C3442450C65 | No |  | Cleaning air valve pH BG module and flushing air duct | Cleaned | Fluidic pathway | - | pCO2
ABL810 FLEX |  | Reinigung AC-Modul | CD5903DF424D0C60 | No |  | Cleaning AC module | Cleaned | Auto Check Module | Barcode Reader
ABL837 FLEX |  |  | CD5928A342381763 | No |  |  | Defect | Electrode | Gasket | pO2
ABL825 FLEX | Autocheck cable damaged | replaced FPC 14-Pole cable | CD58FE03425A1A34 | No | Autocheck cable damaged | replaced FPC 14-Pole cable | Defect | Auto Check Module | FPC Cable
ABL825 FLEX | Tested the halogen lamp intensity and observed that it was 2700. | Adjusted the halogen lamp intensity to 3400. Performed an extended protein cleaning of the hemolyzer. | CD597FC74254BB16 | No | Tested the halogen lamp intensity and observed that it was 2700. | Adjusted the halogen lamp intensity to 3400. Performed an extended protein cleaning of the hemolyzer. | Adjustment | HW | Oxi Module | Lamp Unit
ABL800Basic |  |  | CD58DA11422833C4 | No |  |  | Unknown | Radiance | Restart analyzer
ABL800Basic | siehe ausgeführte Arbeiten | siehe ausgeführte Arbeiten | CD58FF9E421EA119 | No | see work carried out | see work carried out | Defect | Electrode | Electrode | pCO2
TCM-COMBIM | Läuft einwandfrei, Kunde kann keine genauere Fehlerangabe machen, nach ca. einer Woche Probelauf: Funktionstest o.k., | ohne Befund zurück! | CD58CF0E423C44EB | No | Runs flawlessly, customer cannot make a more precise error, after about a week of trial run: functional test ok, | without findings! | No fault found
ABL90 |  | Kassette getauscht. Gerät gereinigt. | CD5A5332B856D18F | No |  | Cassette is exchanged. device cleaned. | Defect | SC | Unknown
ABL90 | They issue started with Lac failing QC. The customer replaced the SC and tried backflushing. Lac came in but Pco2 started failing. Then I came in checked the inlet gasket, connector gasket and probe. Checked activity logs. Checked calibration trends for pco2. Senstivity was at 16-19 %. Tried flushing the analyzer couple times. Sensitivity upto 21%. Tried running 5 waste blood samples sensitivity did not go past 26 %. Tried multiple cals and QC. Replaced SC. After initializing sensitivity still at 26%. Opened the sensor bay to check for bend contact probe. Used the sensor simulator to check the electronics. The simulator test passed. replaced the SP. No change in sensitivity still at 26%. Opened the analyzer checked leaks and replace the Sensor interface PCB as advised by tech support. | Found large deposits on QC and blood on the hemolyzer, between the hemolyzer and sensor bay unit. Replaced Hemolyzer unit.Sensitivity still too low. Replaced SC again. | CD5A80DDB83B9799 | No | They issue started with Lac failing QC. The customer replaced the SC and tried backflushing. Lac came in but Pco2 started failing. Then I came in checked the inlet gasket, connector gasket and probe. Checked activity logs. Checked calibration trends for pco2. Senstivity was at 16-19 %. Tried flushing the analyzer couple times. Sensitivity upto 21%. Tried running 5 waste blood samples sensitivity did not go past 26 %. Tried multiple cals and QC. Replaced SC. After initializing sensitivity still at 26%. Opened the sensor bay to check for bend contact probe. Used the sensor simulator to check the electronics. The simulator test passed. replaced the SP. No change in sensitivity still at 26%. Opened the analyzer checked leaks and replace the Sensor interface PCB as advised by tech support. | Found large deposits on QC and blood on the hemolyzer, between the hemolyzer and sensor bay unit. Replaced Hemolyzer unit. Sensitivity still too low. Replaced SC again. | Defect | SC | Lac
ABL90 |  |  | CD5A18EDB85B75D7 | No |  |  | Adjustment | SP PCB
ABL90 |  |  | CD5A35E7B834DD51 | No |  |  | Cleaned | Fluidic pathway
ABL90 |  |  | CD5A5605B8310616 | No |  |  | Defect | SC | pCO2
ABL90 |  |  | CD5A5153B83FAAF9 | No |  |  | Replaced | Probe | Bend
ABL90 | Unit stalls during re boot. Unit not loading windows. Halts at the BIOS setup screen. | Replaced CPU. | CD5A5181B806ACC6 | No | Unit stalls during re boot. Unit not loading windows. Halts at the BIOS setup screen. | Replaced CPU. | Defect | HW | CPU
TCM400 | Set up the instrument and it showed no failure. Dismantled to check interior, verified that the battery is holding charge. | Found a wire in the display area with insulation worn off; may be grounding out. Insulated the wire; calibrated the touchscreen. | CD5A3729B84EE4C6 | No | Set up the instrument and it showed no failure. Dismantled to check interior; verified that the battery is holding charge. | Found a wire in the display area with insulation worn off; may be grounding out. Insulated the wire; calibrated the touchscreen. | Defect | HW | Broken
ABL815 FLEX |  |  | CD5A57B6B838945C | No |  |  | Defect | pH/BG Module | Valve, air | Membrane
ABL825 FLEX |  |  | CD5A55BBB8541204 | No |  |  | Defect | HW | Oxi Module | Hemolyzer
ABL825 FLEX |  |  | CD5A61CEB8126DE9 | No |  |  | User Error | tHb cal exp
ABL825 FLEX | Can see left over fluid in front of Reference Electrode. | Remembrane Reference Electrode and clean Measuring Chamber. Same issue. Replaced PCO2 Electrode and Calibrate. | CD5A563BB8548711 | No | Can see left over fluid in front of Reference Electrode. | Remembrane Reference Electrode and clean Measuring Chamber. Same issue. Replaced PCO2 Electrode and Calibrate. | Defect | Electrode | Electrode | pCO2
ABL825 FLEX | Analyzer was in hold when arrived because of a leak. Leak was caused by no gasket on end of Ca Electrode membrane. | Replaced Ca Membrane to resolve leak. Also replaced Inlet Probe. | CD5A2D01B84DF3A5 | No | Analyzer was in hold when arrived because of a leak. Leak was caused by no gasket on end of Ca Electrode membrane. | Replaced Ca Membrane to resolve leak. Also replaced Inlet Probe. | Defect | Electrode | Membrane | Ca++
ABL815 FLEX |  |  | CD5A6B5DB8348495 | No |  |  | Defect | FlexQ Module | Not functioning
ABL80 CO-OX |  |  | CD5B8D972E5FAEEA | No
ABL90 |  |  | CD5B6E812E5E72BB | Yes |  |  | Re classified RMED | PM
ABL825 FLEX | System seeing large PH drift and is upgrading 1 point PH BG calibration to a full 1 point calibration. | Pulled down PH BG module, changed measuring chamber, PH.BG PCB. Found foil cap still on cleaning bottle. Bottle not activated. Installed new bottle after activating. Advised customer on protocol. | CD5D18A4A45AB67A | No | System seeing large PH drift and is upgrading 1 point PH BG calibration to a full 1 point calibration. | Pulled down PH BG module, changed measuring chamber, PH.BG PCB. Found foil cap still on cleaning bottle. Bottle not activated. Installed new bottle after activating. Advised customer on protocol. | User Error | Replacement process | Solution bottle
ABL835 FLEX | error 581 oxi spectrum mismatch on qc and samples | ran complete oxi test issue with dvdark parameter, replaced power supply. unit was then moved from the 4th floor to the main lab and has been put into active use | CD5B8F7E2E1687DB | No | error 581 oxi spectrum mismatch on qc and samples | ran complete oxi test issue with dvdark parameter, replaced power supply. unit was then moved from the 4th floor to the main lab and has been put into active use | Defect | Power Supply | Oxi Dark Value
ABL810 FLEX |  |  | CD5B66C42E461B2A | No |  |  | Cleaned | pH/BG Module | Electrode connector | pCO2
ABL825 FLEX |  |  | CD5BA6F32E2119C8 | Yes |  |  | Re classified RMED | Radiance | Setup
ABL825 FLEX | When arrived today, took out new electrode that was left in overnight to settle in - still same issue. Tried replacing measuring Chamber - same error. Replaced Inlet Probe, pH sealing ring and Inlet Nipple kit. Put measuring chamber back in analyzer. | Remove electrode that I put in yesterday. Tried replacing measuring Chamber. Replaced Inlet Probe, pH sealing ring and Inlet Nipple kit. My new electrode was defective from stock. Customer had a new PCO2 Electrode. Soak electrode and install. John will stop by next week with a new electrode. | CD5B6DA52E0BEDB5 | No | When arrived today, took out new electrode that was left in overnight to settle in - still same issue. Tried replacing measuring Chamber - same error. Replaced Inlet Probe, pH sealing ring and Inlet Nipple kit. Put measuring chamber back in analyzer. | Remove electrode that I put in yesterday. Tried replacing measuring Chamber. Replaced Inlet Probe, pH sealing ring and Inlet Nipple kit. My new electrode was defective from stock. Customer had a new PCO2 Electrode. Soak electrode and install. John will stop by next week with a new electrode. | Defect | Electrode | Electrode | pCO2
AQT90 FLEX |  |  | CD5B860B2E01C769 | No
ABL90 |  |  | CD5CFBD6A4042CE6 | No
ABL90 |  |  | CD5E083C1A103575 | No
ABL90 |  |  | CD5CCC91A42BA77A | Yes |  |  | Re classified RMED | Logistic upSWO
ABL820 FLEX |  | Austausch Airventil im pH BG Modul | CD5C9F8AA448482A | No |  | Replacement air valve in pH BG module | Defect | pH/BG Module | Valve, air | pCO2
ABL825 FLEX | Replaced pH electrode sealing rind and then pH electrode on last visit. | Replaced inlet probe and inlet tube Tested of inhomogeneous sample on QC-OK | CD5CFCFFA4050E49 | No | Replaced pH electrode sealing rind and then pH electrode on last visit. | Replaced inlet probe and inlet tube Tested of inhomogeneous sample on QC-OK | Defect | Inlet module | Probe
TCM400 | Powered on the unit was not able to reproduced the issue he was seeing. Called the user confirmed that it was an issue with the setup. the report print out was not setup right after adjusting it everything looked like the customer wanted. tested the timer function it worked fine. | I change the setup for the customer so it looks more like they want. | CD5CCF68A4598941 | No | Powered on the unit was not able to reproduced the issue he was seeing. Called the user confirmed that it was an issue with the setup. the report print out was not setup right after adjusting it everything looked like the customer wanted. tested the timer function it worked fine. | I change the setup for the customer so it looks more like they want. | User Error | Setup
ABL90 |  |  | CD5DEE581A58172F | No
ABL90 |  |  | CD5DE50F1A56DF7E | No
AQT90 FLEX |  | Nadel erneuert | CD5DFE161A04D1A5 | No |  | Needle renewed
AQT90 FLEX |  |  | CD5E04B81A352F6F | No
ABL80 FLEX Basic version |  |  | CD61C7AA7C3BF9B4 | Yes |  |  | Re classified RMED | Preventive maintenance
ABL80 CO-OX |  |  | CD61C6BA7C39D469 | No
ABL90 | Two SP's, both Lot LO-05, causing inconsistent software error. I retrieved customers database to backup drive. I then installed new SP, Lot LT-08, which immediately corrected the inconsistent SW error. I setup auto data backup. I ran Cal's and QC. Internal QC all pass except Cl on level three. Upper limit is 99 and system is getting 99 (rounding) External QC all pass exccept Level 3 fails Na and K. Both are near upper limit. Multiple Cal's did not improve. I tried to run a blood sample to help condition SC but no bloods available. WIll call customer after system stabilizes to get passing QC results | New SP, backup all data, setup auto archive. Customer is due a warranty replacement for one of these Solution Packs. I replaced the other with a service replacement. | CD61DD1B7C1E9F6C | No | Two SP's, both Lot LO-05, causing inconsistent software error. I retrieved customers database to backup drive. I then installed new SP, Lot LT-08, which immediately corrected the inconsistent SW error. I setup auto data backup. I ran Cal's and QC. Internal QC all pass except Cl on level three. Upper limit is 99 and system is getting 99 (rounding) External QC all pass exccept Level 3 fails Na and K. Both are near upper limit. Multiple Cal's did not improve. I tried to run a blood sample to help condition SC but no bloods available. WIll call customer after system stabilizes to get passing QC results | New SP, backup all data, setup auto archive. Customer is due a warranty replacement for one of these Solution Packs. I replaced the other with a service replacement. | Defect | SP | Inconsistante SW after change |  | x
ABL90 |  |  | CD63BCF2684D94E5 | No |  |  | Replaced | inlet gasket | to new version updated SW | Fluid transport | x
ABL90 |  |  | CD61F46B7C52D4CC | No |  |  | Defect | SC | pCO2 Cal - sens | No new SC available | x
ABL825 FLEX |  |  | CD632279F22D3F95 | No |  |  | Defect | pH/BG Module | PCB
ABL827 FLEX | data falls to 11.7 outs side of 2 sd range. 1 changed the membrane | i replaced the sealing on the ph electrode | CD61C0207C60E7C7 | No | data falls to 11.7 outs side of 2 sd range. 1 changed the membrane | i replaced the sealing on the ph electrode | Defect | Electrode | Membrane | pCO2 | Sealing ring
ABL820 FLEX | PO2 has been failing 2 point calibrations for zero current. | Performed zero current calibration in service program. ran several calibrations. Error has cleared. | CD61E1EC7C49A026 | No | PO2 has been failing 2 point calibrations for zero current. | Performed zero current calibration in service program. ran several calibrations. Error has cleared. | Adjustment | pH/BG Module | Zero calibration | pO2
ABL835 FLEX | Issue occurs on patients and QC. Inspected sample path, tested all valves. Inspected measuring module valves and found both air valves jammed in opened position | Replaced pH BG air valve and El Met air valve Performed thb cal (due) Obtained inhomogeneous sample error on thb cal verification. Noticed that newly installed pH BG air valve had also stuck in open position. Replaced pH BG air valve with regular- styled membrane valve. Encountered Ca drift issues, replaced calcium membrane then reference membrane (both were due in about a week). Inhomogeneous sample issue came back, pH BG air valve had jammed again, replaced pH BG PCB (and tried newly replaced pH BG air valve again). Put the 902-700 valve back in since the 902-701 was stuck | CD61EDCD7C09B8E8 | No | Issue occurs on patients and QC. Inspected sample path, tested all valves. Inspected measuring module valves and found both air valves jammed in opened position | Replaced pH BG air valve and El Met air valve Performed thb cal (due) Obtained inhomogeneous sample error on thb cal verification. Noticed that newly installed pH BG air valve had also stuck in open position. Replaced pH BG air valve with regular- styled membrane valve. Encountered Ca drift issues, replaced calcium membrane then reference membrane (both were due in about a week). Inhomogeneous sample issue came back, pH BG air valve had jammed again, replaced pH BG PCB (and tried newly replaced pH BG air valve again). Put the 902-700 valve back in since the 902-701 was stuck | Unknown | Fluidic pathway | Inhomogenous sample
ABL825 FLEX |  |  | CD61DF307C5F964D | No |  |  | Defect | Fluidic pathway | Waste Tube
ABL815 FLEX |  |  | CD61C6E07C5FCFD2 | No |  |  | Adjustment | HW | Oxi Module | Lamp Unit
ABL825 FLEX | inspected logs and found the instrument to be operating within spec. | power off on and instrumect | CD61D03E7C2464A8 | No | inspected logs and found the instrument to be operating within spec. | power off on and instrumect | Unknown | Power off/on
ABL815 FLEX |  |  | CD6208E77C063306 | No |  |  | Adjustment | HW | Oxi Module | Lamp Unit
ABL80 CO-OX | Reboot of analyzer no help. Returns to 2106 error. Will not allow sensor change. | Flushed waste lines and hemolyzer with protien remover. Checked for leaks. Liquid sensor calibration OK. | CD62FDB9F25715DB | No | Reboot of analyzer no help. Returns to 2106 error. Will not allow sensor change. | Flushed waste lines and hemolyzer with protien remover. Checked for leaks. Liquid sensor calibration OK.
ABL90 |  |  | CD630241F2297DC6 | No |  |  | Cleaned | Fluidic pathway | Inlet Gasket |  | x
ABL90 |  |  | CD65CAEEDE60B318 | No |  |  | Defect | SP | Fluid transport |  | x
ABL90 |  |  | CD62EFC9F21C9D57 | No
ABL90 |  |  | CD633FA9F225553A | No |  |  | Defect | SC | pCO2 - QC |  | x
ABL90 |  |  | CD633CE8F2359D61 | No
ABL90 |  |  | CD632FE7F21D4C4D | No |  |  | Cleaned | fluidic pathway |  |  | x
ABL810 FLEX |  |  | CD62E657F22AB2FB | No |  |  | Cleaned | Fluidic pathway | Air vent
ABL835 FLEX | Called customer first thing this morning to follow- up. After reviewing the data on her instrument, Stéphanie mentionned that the issue had come back overnight. | Went onsite and grabbed data logs. Started Techline | CD632404F231B8F8 | No | Called customer first thing this morning to follow- up. After reviewing the data on her instrument, Stéphanie mentionned that the issue had come back overnight. | Went onsite and grabbed data logs. Started Techline
ABL835 FLEX | Performed power supply check Checked Roma.ini file for any signs of corruption. Display previously changed, display plugs into CPU | Replaced CPU. | CD62FE5EF22FB8F3 | No | Performed power supply check Checked Roma.ini file for any signs of corruption. Display previously changed, display plugs into CPU | Replaced CPU.
ABL800Basic |  |  | CD632D3CF253EFBC | No |  |  | Defect | Electrode | Electrode | pO2
ABL835 FLEX |  |  | CD62E020F21D694F | No |  |  | Defect | Electrode | Electrode | pCO2
ABL815 FLEX | siehe SLR | siehe SLR | CD62F2F4F24733FA | No | see SLR | see SLR | Defect | HW | Oxi module | Spectrometer
AQT90 FLEX |  |  | CD62FFFEF2244F9C | No
ABL90 |  |  | CD65A1E7DE46439A | No |  |  | Unknown | RC & CM unclear |  |  | x
ABL90 |  |  | CD649060683EB022 | No
ABL5 |  |  | CD6439C4681B3E3E | No
ABL800Basic |  |  | CD6453FA6852B012 | No |  |  | Defect | Inlet module | Probe bend
ABL827 FLEX |  |  | CD6443BE68070E90 | No |  |  | Defect | FlexQ Module
ABL825 FLEX | Previous visit isolated error to the electrode.) Back on site with warranty electrode to exchange for failed pCO2 Electrode. Also, tHb value is low on Cal's at 460 (This # was 500 in April) I adjusted Lamp from 2270 to 3350. New tHb # is now 550 which closely resembles peer analyzer which is ~ 560. Oxi complete test and passing service tHb Cal | Exchange pCO2 Electrode with warranty replacement. Adjust Lamp, Oxi Complete and tHb Cal | CD644AC46805E398 | No | Previous visit isolated error to the electrode.) Back on site with warranty electrode to exchange for failed pCO2 Electrode. Also, tHb value is low on Cal's at 460 (This # was 500 in April) I adjusted Lamp from 2270 to 3350. New tHb # is now 550 which closely resembles peer analyzer which is ~ 560. Oxi complete test and passing service tHb Cal | Exchange pCO2 Electrode with warranty replacement. Adjust Lamp, Oxi Complete and tHb Cal | Adjustment | HW | Oxi Module | Lamp Unit
ABL835 FLEX |  |  | CD64875B684B67C1 | No |  |  | Setup | HIS/LIS
ABL825 FLEX | 3 of 4 Syringe buttons had different Layouts assigned. 1-R-default format, 2 others were Cal Verification format assigned. pH Pleaural fluid was assigned Cal. Verification. | Assigned correct Layouts to corresponding Syringe button. | CD648CD16850A79F | No | 3 of 4 Syringe buttons had different Layouts assigned. 1-R-default format, 2 others were Cal Verification format assigned. pH Pleaural fluid was assigned Cal. Verification. | Assigned correct Layouts to corresponding Syringe button. | Setup | Sample types
ABL825 FLEX |  |  | CD6D06F0A21EDD98 | No |  |  | Defect | Inlet module | Probe bend
ABL835 FLEX | Customer experienced above issues at separate times. Replaced inlet gasket. Screen froze (three separate occasions today) and customer was unable to run a rinse or a sample. Attended service visit, analyzer was working ok upon arrival Probe nice and straight. Saw inlet gasket leaking when performing QC. | Performed inlet calibration. Attempted to recreate the issue, performed calibrations and QC. Re- formated instrument and attempted to re-install Windows and software without success since the Windows install kept stalling. Tried new HDD, power supply and different CPU. Ordered new CPU and HDD, also asked peer to send their external CD drive and Windows install disk. Advised customer of situation and ordered parts. | CD647EB4684CB8FD | No | Customer experienced above issues at separate times. Replaced inlet gasket. Screen froze (three separate occasions today) and customer was unable to run a rinse or a sample. Attended service visit, analyzer was working ok upon arrival Probe nice and straight. Saw inlet gasket leaking when performing QC. | Performed inlet calibration. Attempted to recreate the issue, performed calibrations and QC. Re- formated instrument and attempted to re-install Windows and software without success since the Windows install kept stalling. Tried new HDD, power supply and different CPU. Ordered new CPU and HDD, also asked peer to send their external CD drive and Windows install disk. Advised customer of situation and ordered parts. | Unknown | DMS PC
ABL835 FLEX |  |  | CD648902685167C4 | Yes |  |  | Re classified RMED
ABL835 FLEX |  |  | CD64798E682AB019 | No |  |  | Defect | Electrode | Electrode | pCO2
ABL825 FLEX | checked inlet module. | Replaced ref membrane. Replaced inlet tube | CD6E667D1839FDF0 | No | checked inlet module. | Replaced ref membrane. Replaced inlet tube | Defect | Inlet module | Inlet Tube (Blue)
ABL825 FLEX | Analyzer was locked up. PO2 would not pass cal. | Had to perform a reboot. Replaced membrane from customer stock. | CD645DB568449708 | No | Analyzer was locked up. PO2 would not pass cal. | Had to perform a reboot. Replaced membrane from customer stock. | Defect | Electrode | Electrode | pO2
AQT90 FLEX |  |  | CD64B6C76843A1E1 | No
ABL80 CO-OX |  | Austausch CF-Card; Softwareupdate V3.11 durchgeführt; Gerät konfiguriert; tHb Kal durchgeführt | CD65BCF2DE3AF612 | No |  | Replacement CF-Card; Software update V3.11 performed; Device configured; tHb Kal performed
ABL90 |  |  | CD6A837DB65B3C4B | No |  |  | Defect | SC | Ca - Cal - sens |  | x
ABL90 |  | CF-Card getauscht. | CD658532DE43D1E6 | No |  | CF-Card. | Defect | HW | CF card | will not reboot | x
ABL90 |  |  | CD65B5BDDE50971A | No
ABL90 |  |  | CD65A19BDE4E7663 | Yes |  |  | Re classified RMED | planned SW upgrade |  |  | x
ABL90 | Analyzer will not connect to Radiance without disconnecting on of the Anesthesia analyzers. Analyzer will connect to Test system but not Live system. Work with Steve McClain as public wifi doesn't allow me to remote to server. Demo key is installed and says 50 connections possible. | Work with IT person to test wall port for connectivity. I can ping Live server from analyzer. Steve McClain loaded new analyzer connection key and all analyzers are connected to Live system. | CD658954DE1C2287 | No | Analyzer will not connect to Radiance without disconnecting on of the Anesthesia analyzers. Analyzer will connect to Test system but not Live system. Work with Steve McClain as public wifi doesn't allow me to remote to server. Demo key is installed and says 50 connections possible. | Work with IT person to test wall port for connectivity. I can ping Live server from analyzer. Steve McClain loaded new analyzer connection key and all analyzers are connected to Live system. | Setup | Radiance |  |  | x
ABL90 | analyzers were shipped to FSE home due to the long weekend cusstomes requesting install complete on Monday | unpacked analyzer and started up ABL90 Battery is not working; Causes analyzer to reboot constantly checked all connections and cables; swapped batteries and confirmed battery is bad; called tech support and ordered a new battery for Saturday | CD657686DE0B922A | Yes | analyzers were shipped to FSE home due to the long weekend cusstomes requesting install complete on Monday | unpacked analyzer and started up ABL90 Battery is not working; Causes analyzer to reboot constantly checked all connections and cables; swapped batteries and confirmed battery is bad; called tech support and ordered a new battery for Saturday | Re classified RMED | OOB
ABL90 |  |  | CD65D333DE5F2AB6 | No |  |  | Defect | SP | Fluid transport |  | x
ABL825 FLEX |  | Austausch Ref-Elektrode | CD65A90ADE6198E9 | No |  | Replacement ref electrode | Defect | Electrode | Electrode | Reference
ABL800Basic |  |  | CD65B1CADE429717 | No |  |  | Defect | Electrode | Electrode | Reference
ABL815 FLEX |  | Austausch Drucker | CD65717DDE3298D7 | No |  | Replacement printer | Defect | Printer module | Printer head
ABL825 FLEX |  |  | CD65662DDE2BF14A | No |  |  | Defect | Inlet module | Inlet Gasket
ABL835 FLEX | Analyzerand UPS was shutdown . | Turned both Ups and Analyzer | CD6E6748181CFDF3 | No | Analyzerand UPS was shutdown . | Turned both Ups and Analyzer | Unknown | Power off/on
ABL825 FLEX | Tried and new pH o-ring and a new pH electrode-No change Steve Powers did service call lastweek-No change | Replaced all parts and tubes from the pH BG air valve to the waste tubes. | CD6580A8DE54DF46 | No | Tried and new pH o-ring and a new pH electrode-No change Steve Powers did service call lastweek-No change | Replaced all parts and tubes from the pH BG air valve to the waste tubes. | Defect | pH/BG Module | Measuring chamber
ABL820 FLEX |  |  | CD65BFF5DE5A6207 | No |  |  | Defect | El/Met Module | Measuring chamber
ABL810 FLEX |  |  | CD6611CFDE4D82F8 | Yes |  |  | Re classified RMED | Setup
ABL825 FLEX | Tried replacing membranes for pcor. Gas module seems faulty. Replaced flow regulators. After power on Disk boot failure came up. Replaced power supply and that did not help. Replaced Hard drive , no help. Ordered another hard rive still no help. The XP emebedded installation starts on the new hard drives but its stuck at Cheking Hard disk and Preparing hard disk. | Consulted Waqas Shah and he consulted RAME. Rame advised to order new power supply. | CD76B0BD524110BF | No | Tried replacing membranes for pcor. Gas module seems faulty. Replaced flow regulators. After power on Disk boot failure came up. Replaced power supply and that did not help. Replaced Hard drive , no help. Ordered another hard rive still no help. The XP emebedded installation starts on the new hard drives but its stuck at Cheking Hard disk and Preparing hard disk. | Consulted Waqas Shah and he consulted RAME. Rame advised to order new power supply. | Unknown | PS
ABL825 FLEX | Tested the valves and inspected the tubing. Reviewed the datalogs and observed that the drift was 0 but questioned. | Changed the drift tolerances back to the factory standard. | CD65C81BDE596295 | No | Tested the valves and inspected the tubing. Reviewed the datalogs and observed that the drift was 0 but questioned. | Changed the drift tolerances back to the factory standard. | User Error | Calibration limits | K C++
AQT90 FLEX | SE LE HA ROTO EL SENSOR DE ENCIMA DE LA WASH UNIT Y POR ESTE MOTIVO SE HA DOBLADO LA AGUJA, SE SUSTITUYEN AMBAS PIEZAS, DE MANERA PREVENTIVA SE SUSTITUYEN BOMBAS 2 Y 4 Y FLITROS DE LIQUIDOS, LA WASH UNIT DA ERROR AL MEDIR EL HCT, SE LIMPIA Y SIGUE SIN FUNCIONAR, SE SUSTITUYE Y SE CALIBRA, SE HACE COMPROBACION CON MUESTRAS DE BLANCOS Y QC MULTICHECK NIVEL 2 | SE LE HA ROTO EL SENSOR DE ENCIMA DE LA WASH UNIT Y POR ESTE MOTIVO SE HA DOBLADO LA AGUJA, SE SUSTITUYEN AMBAS PIEZAS, DE MANERA PREVENTIVA SE SUSTITUYEN BOMBAS 2 Y 4 Y FLITROS DE LIQUIDOS, LA WASH UNIT DA ERROR AL MEDIR EL HCT, SE LIMPIA Y SIGUE SIN FUNCIONAR, SE SUSTITUYE Y SE CALIBRA, SE HACE COMPROBACION CON MUESTRAS DE BLANCOS Y QC MULTICHECK NIVEL 2 | CD658208DE268C7E | No | THE WASH UNIT'S SENSOR HAS BEEN BROKEN AND FOR THIS REASON THE NEEDLE HAS BEEN BENT, BOTH PARTS ARE SUBSTITUTED , IS CLEANED AND STILL NOT WORKING, SUBSTITUTE AND CALIBRATED, CHECKISED WITH WHITE AND QC MULTICHECK LEVEL 2 SAMPLES | THE WASH UNIT'S SENSOR HAS BEEN BROKEN AND FOR THIS REASON THE NEEDLE HAS BEEN BENT, BOTH PARTS ARE SUBSTITUTED , IS CLEANED AND STILL NOT WORKING, SUBSTITUTE AND CALIBRATED, CHECKISED WITH WHITE AND QC MULTICHECK LEVEL 2 SAMPLES
AQT90 FLEX |  |  | CD66EA7E544CCBD0 | No
ABL80 CO-OX |  |  | CD66E461545A79CC | No
ABL80 FLEX Basic version |  | Lösungspack erneuert. | CD66BFC25462D1F8 | No |  | Solution pack renewed.
ABL90 |  |  | CD66E579542A42E7 | No |  |  | Cleaned | fluidic pathway | SP replace + replaced gasket/probe |  | x
ABL90 |  |  | CD66B66E54155C97 | No |  |  | Defect | SP | Fluid transport | & Cleaned fluidic pathway | x
ABL90 |  |  | CD66E4215429DD82 | No |  |  | Defect | HW | CPU | Blue screen | x
ABL90 |  |  | CD670DCF542933E3 | No
ABL825 FLEX |  |  | CD66F92E545B33DF | No
ABL835 FLEX |  |  | CD66D6F954496831 | No |  |  | Defect | Electrode | Electrode | Lac
ABL835 FLEX | Attended service call July 2 believed that cause of issues was due to a software problem, but was unable to re-install software or resolve the issue with a new HDD, CPU or Power supply | Tried yet another new HDD and CPU, as well as trying different external CD drive and Windows disk. Tried accessing the first two HDD via laptop without success, despite new hardware found and ready for use popping up. Started Techline Discussed issue with SIT (to get software hardware advice) and Jim Gibbons, who told me about an issue with the hard drives. Returning the Drives to Jim. | CD6BB6F12C60B915 | No | Attended service call July 2 believed that cause of issues was due to a software problem, but was unable to re-install software or resolve the issue with a new HDD, CPU or Power supply | Tried yet another new HDD and CPU, as well as trying different external CD drive and Windows disk. Tried accessing the first two HDD via laptop without success, despite new hardware found and ready for use popping up. Started Techline Discussed issue with SIT (to get software hardware advice) and Jim Gibbons, who told me about an issue with the hard drives. Returning the Drives to Jim.
ABL825 FLEX |  | tHb Kalibration durchgeführt | CD66D94D545FF61F | No |  | tHb calibration performed | User Error | tHb cal exp
ABL825 FLEX | Ventile und Elektroden | Proteinentferung Messkammer gesäubert Austausch Cl-Elektrode | CD66D9A2545DBACA | No | Valves and electrodes | Protein draining measuring chamber cleaned replacement Cl electrode | Defect | Electrode | Electrode | Cl
ABL90 | Na status and sensitivity errors. Customer replaced solution pack. Recalibration no help. | Replaced sensor cassette. New sensor had pCO2 sensitivity errors. Sensitivity dropped further out of range with each calibration. Replaced sensor again. Vacuum pressure test failing at upper limit of 0.75. Replaced oxi valve tubing and inlet gasket w holder (installed 5 21 14). | CD6AD96DB6382C4E | No | Na status and sensitivity errors. Customer replaced solution pack. Recalibration no help. | Replaced sensor cassette. New sensor had pCO2 sensitivity errors. Sensitivity dropped further out of range with each calibration. Replaced sensor again. Vacuum pressure test failing at upper limit of 0.75. Replaced oxi valve tubing and inlet gasket w holder (installed 5 21 14). | Defect | SC | Na Cal | & new SC pCO2 | x
ABL90 | Archiv | SC erneuert | CD6A948EB6352A6B | No | Archive | SC renewed | Defect | SC | Lac Cal |  | x
ABL90 | Datenbank geprüft | Flüssigkeitsstrecke gespült--> Neustart | CD6ACAFBB60F025E | No | Database checked | Liquid track rinsed--> Restart | Cleaned | Fluidic pathway |  |  | x
ABL90 | when battery is installed power reboots and will not stay on with switch in the on possition | replaced the battery | CD6A5F79B610E694 | No | when battery is installed power reboots and will not stay on with switch in the on possition | replaced the battery | Defect | HW | Battery module |  | x
ABL815 FLEX |  |  | CD6A849EB63D5DAC | No |  |  | Cleaned | Basic unit | Barcode Reader | Dust
ABL800Basic |  |  | CD6AB3D8B647B0AE | No |  |  | Defect | DMS PC | CPU
ABL835 FLEX | Instrument freeze or impossible loading software Know issue according Jim Giibbons | installed a loner DMS load software reinstalled data and setup waiting for indication from Jim fix original DMS with new hard drive | CD6A94A6B616F1DD | No | Instrument freeze or impossible loading software Know issue according Jim Giibbons | installed a loner DMS load software reinstalled data and setup waiting for indication from Jim fix original DMS with new hard drive | Defect | DMS PC | HDD
ABL827 FLEX |  |  | CD6AE7B1B6436863 | No |  |  | Unknown | Power off/on
ABL815 FLEX |  |  | CD6AB6CDB658581F | No |  |  | Defect | Auto Check Module | Fan
TCM-COMBIM |  |  | CD6A82B8B6572A9D | No |  |  | Defect | Electrode | tc Sensor 84, pCO2/pO2
ABL90 |  |  | CD6C13262C25D3EB | No |  |  | Defect | SP | Fluid transport | & SC pCO2 | x
ABL90 | siehe ausgeführte Arbeiten | siehe ausgeführte Arbeiten | CD6C0F562C48513A | No | see work carried out | see work carried out | Cleaned | fluidic pathway | tHb cal |  | x
ABL90 |  |  | CD6BF9232C554718 | No |  |  | Defect | SC | pCO2 |  | x
ABL90 | Reiniciar equipamento. Limpeza do circuito de amostra. Inicialização. Substituição da sonda e entrada de amostras. Verificação do correcto funcionamento do equipamento. |  | CD6C0ED12C4C3D73 | No | Restart equipment. Cleaning the sample circuit. Startup. Replacing the probe and entering samples. Checking the correct operation of the equipment. |  | Cleaned | Fluidic pathway | Replaced probe/gasket | & power off/on | x
ABL815 FLEX |  |  | CD6BC62D2C56E43F | No |  |  | Cleaned | Auto Check Module | Barcode Reader
ABL835 FLEX |  |  | CD6BE84A2C240576 | No |  |  | Defect | El/Met Module | Pump tube
ABL825 FLEX | Checked valves ,gas valves and electrode. Pco2 reporting low on QC | Replaced Flow regulator right and left | CD6E6B57181DFDFE | No | Checked valves ,gas valves and electrode. Pco2 reporting low on QC | Replaced Flow regulator right and left | Defect | Gas module | Right and Left Regulator
ABL825 FLEX | Ran complete oxy test which passed. | I replaced the OXY module pcb. | CD6BE6EE2C257283 | No | Ran complete oxy test which passed. | I replaced the OXY module pcb. | Defect | HW | Oxi Module | PCB
ABL825 FLEX | Inspected entire oxy modul for leaks, damage burns and none was found | replaced the oxy pcb | CD6C0F7B2C2B897A | No | Inspected entire oxy modul for leaks, damage burns and none was found | replaced the oxy pcb
TCM/Tosca |  |  | CD6BF39D2C213491 | No |  |  | Defect | HW | Communication PCB
ABL80 FLEX Basic version | Comprobación del estado del analizador. Se comprueban registros y se verifica funcionamiento, se revisan canales de fluídica y bomba peristática, se realiza un a comprobación del canal de zero (valores dentro de márgenes). Limpieza del conector (pins) del equipo hacia la cassette de electrodos. Se realizan ciclos de calidad para verificar estado de los electrodos. Revisión y verificación del equipo. | Revisión preventiva según plan visitas. | CD6D435AA249D9D3 | Yes | Checking the status of the scanner. Records are checked and operation is checked, fluidic channels and peristhetic pump are checked, a zero channel check (values within margins) is performed. Cleaning the connector (pins) of the equipment towards the electrode cassette. Quality cycles are performed to check the status of the electrodes. Equipment review and verification. | Preventive review according to plan visits. | Re classified RMED | Preventive maintenance
ABL90 |  |  | CD74C6256660D2AD | No |  |  | Defect | SC | pCO2 Cal - sens |  | x
ABL90 |  |  | CD6D3AA5A22867A2 | No
ABL90 |  |  | CD6D5CA8A23A2F16 | No |  |  | Replaced | gasket | Install SW 3.1 + yearly maintenance kit |  | x
ABL90 |  |  | CD6D26B8A219960E | Yes |  |  | Re classified RMED | ECKHI: I can tell you that you are right, this is a pSWO. The FSE install a customer wish. | ROMA.Ini - Extra Decimals |  | x
ABL90 | siehe ausgeführte Arbeiten | siehe ausgeführte Arbeiten | CD6D3B44A225B8D5 | No | see work carried out | see work carried out | Unknown | Revisit required | Power supply replaced but new UpSWO later | Analyzer turns off/on sometimes several times a day | x
ABL825 FLEX | Checked qc and calibration logs. | Removed measuring chamber and checked ok. Inspected rinse and air valve-replaced air valve membrane. replaced pump tubing ph module.i nspected all electrodes. performed gas reset Gas 1 532 gas 2 498. after reset 532,532. | CD6D2E10A2223E49 | No | Checked qc and calibration logs. | Removed measuring chamber and checked ok. Inspected rinse and air valve-replaced air valve membrane. replaced pump tubing ph module.i nspected all electrodes. performed gas reset Gas 1 532 gas 2 498. after reset 532,532. | Defect | pH/BG Module | Valve air | Membrane
ABL825 FLEX | Error had cleared upon arrival. | Flushed Fluidic lines, checked valves, tubings. | CD6D50A8A2280E78 | No | Error had cleared upon arrival. | Flushed Fluidic lines, checked valves, tubings. | Cleaned | Fluidic Module | Air vent
ABL825 FLEX |  |  | CD6E3D9118260F2A | No |  |  | Defect | Inlet module | Probe bend
ABL825 FLEX | Intermittent oxi errors | Readjusted lamp unit. | CD6E6CFC1855FE04 | No | Intermittent Oxi Errors | Readjusted lamp unit. | Adjustment | HW | Oxi Module | Lamp Unit
ABL80 FLEX Basic version | Revision trimestral Limpieza de conductos Cero electronico |  | CD6EACB11856968E | No | Quarterly Review Zero Electronic Duct Cleaning
ABL90 |  |  | CD6F616E8E3CC521 | No |  |  | Defect | SC | temperature |  | x
ABL90 |  |  | CD6E83FC1811E34E | No |  |  | Defect | SC | pCO2 Cal - sens |  | x
ABL835 FLEX | Customer experiencing intermittent inhomogeneous sample errors on QC and patients samples Tested and inspected all valves, noticed pO2 o-ring was missing. Inspected all tubing, probe, and sample path Tried new pH BG module for troubleshooting purposes, and ran 10 waste bloods without error coming back | Replaced inlet preheater, and ran an additional 8 waste bloods without the inhomogeneous sample error coming back. | CD6E5E371853A532 | No | Customer experiencing intermittent inhomogeneous sample errors on QC and patients samples Tested and inspected all valves, noticed pO2 o-ring was missing. Inspected all tubing, probe, and sample path Tried new pH BG module for troubleshooting purposes, and ran 10 waste bloods without error coming back | Replaced inlet preheater, and ran an additional 8 waste bloods without the inhomogeneous sample error coming back. | Defect | Inlet module | Preheater
ABL825 FLEX | checked flow of samples, checked valves and liquid sensors, checked sample probe | removed and replaced bent inlet probe, | CD6EDF1C182A8327 | No | checked flow of samples, checked valves and liquid sensors, checked sample probe | removed and replaced bent inlet probe, | Defect | Inlet module | Probe bend
ABL835 FLEX |  |  | CD6E8D4E18241E47 | No |  |  | Cleaned | pH/BG Module | Measuring chamber | pCO2
ABL835 FLEX | Customer complaint ABL froze very almost every 2-3 days | replaced CPU | CD6E462F183F14A2 | No | Customer complaint ABL froze very almost every 2-3 days | replaced CPU | Defect | DMS PC | CPU
ABL827 FLEX |  |  | CD6E4AFD184D0302 | Yes |  |  | Re classified RMED | Setup
ABL820 FLEX |  |  | CD6E75AF1834A8CB | No |  |  | Adjustment | HW | Oxi Module | Lamp Unit
ABL825 FLEX | Checked remote support settings and regional settings. Qc on pco2 low | Replaced Ref Membrane with diff lot. | CD6E626F1820FDE9 | No | Checked remote support settings and regional settings. Qc on pco2 low | Replaced Ref Membrane with diff lot. | Defect | Electrode | Electrode | pCO2
AQT90 FLEX |  |  | CD6E7C95180FBA7B | No
ABL80 CO-OX | Performed a zero check and determined that the analog board was faulty. | Replaced the analog board, and this resolved the problem. | CD6FD10D8E32918B | No | Performed a zero check and determined that the analog board was faulty. | Replaced the analog board, and this resolved the problem.
ABL90 |  |  | CD6FC48C8E2A7AB4 | No |  |  | Defect | HW | Battery module |  | x
ABL90 | Detecção da anomalia. Sonda de entrada mal montada, não estava não estava a passar pela entrada. Substituição da sonda e entrada de amostras. Inicialização. Verificação do correcto funcionamento. Reposição do equipamento na urgência que tinha sido substituído pelo do laboratório. |  | CD6F82648E41EB0E | No | Anomaly detection. I wasn't going through the driveway. Replacing the probe and entering samples. Startup. Verification of proper operation. Replacement of the equipment in the urgency that had been replaced by the laboratory. |  | Replaced | probe | Clot |  | x
ABL825 FLEX | Erreur pompe Met module et | Liquide dans les chambres des electrode CA++, K++, NA+,nettoyé lles chambre vérifier les valves rotorles o'ring demarré pompe ajustement et liquid sensor | CD6F94998E4A2FCD | No | Error pump Met module and | Liquid in the chambers of electrode CA, K, NA, cleaned the chamber check the rotor valves o'ring demarré pump fit and liquid sensor | Cleaned | El/Met Module | Measuring chamber
ABL827 FLEX | . data was averaging around 12mmhg. mean for peer is 12.8 | i replaced the measuring chamber | CD6FCB918E1009B2 | No | . data was averaging around 12mmhg. mean for peer is 12.8 | i replaced the measuring chamber | Defect | pH/BG Module | Measuring chamber | pCO2
ABL820 FLEX | Lamp intensity is reading 3900 and needs to be adjusted down. | Adjusted halogen lamp intensity down to 3400. | CD6FF0CB8E1575C0 | No | Lamp intensity is reading 3900 and needs to be adjusted down. | Adjusted halogen lamp intensity down to 3400. | Adjustment | HW | Oxi Module | Lamp Unit
ABL835 FLEX | Noticed odd bubbles in measuring module rinse valves. Inspected measuring module valves and noticed both air valves were stuck in the open position. | Replaced pH BG and El Met air valves. Replaced pH BG pump tubing with spare. | CD7018628E61A53E | No | Noticed odd bubbles in measuring module rinse valves. Inspected measuring module valves and noticed both air valves were stuck in the open position. | Replaced pH BG and El Met air valves. Replaced pH BG pump tubing with spare. | Defect | pH/BG Module | Valve air
ABL80 | Customer said they had replaced the cassette but then external Qc would fail with no end point . I found the cassette was faulty. | Replaced cassette, but it would take a very long time to initialize. I replaced the connector PCB Signal data PCB. Performed Zero calibration and checked the temperature. All passed Brought system to ready. | CD748FA86618F97F | No | Customer said they had replaced the cassette but then external Qc would fail with no end point . I found the cassette was faulty. | Replaced cassette, but it would take a very long time to initialize. I replaced the connector PCB Signal data PCB. Performed Zero calibration and checked the temperature. All passed Brought system to ready.
ABL80 CO-OX | Upon arriavl, found unit printer cover switch broken. | Replaced Printer unit. | CD7385CAF01DB26D | No | Upon arriavl, found unit printer cover switch broken. | Replaced Printer unit.
ABL90 | Customer had replaced multiple SCs and SPs. Replaced connector gasket and holder multiple times. Customer had seen pco2 sensitivity errors. I investigated data logs. Customer had seen intermittent issues with the analyzer. Looked at tubings of the analyzer. Pinch valve seems faulty. | Replaced all tubings and Pinch valve. Replaced SP. To observe the performance of the analyzer over next few days. SW upgraded to 3.1 as well. | CD754C91DC57DE3E | No | Customer had replaced multiple SCs and SPs. Replaced connector gasket and holder multiple times. Customer had seen pco2 sensitivity errors. I investigated data logs. Customer had seen intermittent issues with the analyzer. Looked at tubings of the analyzer. Pinch valve seems faulty. | Replaced all tubings and Pinch valve. Replaced SP. To observe the performance of the analyzer over next few days. SW upgraded to 3.1 as well. | Defect | HW | Oxi Module | Pinch Valve tube | x
ABL90 | Retested the analyzer and got a Inhomogeneous QC1 solution error, and an Inhomogeneous Cal 3 solution error. I ran the sensor checks in service and got a solution pack error. Using the service manual, the inlet gasket, inlet connector gasket and sensor bay unit are the likely cause. | Replaced the inlet gasket holder, and the inlet connector gasket, but this did not resolve the problem. I performed the back flush and this resolved the SP failure, but not the inhomogeneous problem. Placed a sensor bay unit on order for overnight delivery. | CD739E82F04AB3EE | No | Retested the analyzer and got a Inhomogeneous QC1 solution error, and an Inhomogeneous Cal 3 solution error. I ran the sensor checks in service and got a solution pack error. Using the service manual, the inlet gasket, inlet connector gasket and sensor bay unit are the likely cause. | Replaced the inlet gasket holder, and the inlet connector gasket, but this did not resolve the problem. I performed the back flush and this resolved the SP failure, but not the inhomogeneous problem. Placed a sensor bay unit on order for overnight delivery. | Unknown | Revisit required | Cleaned Fluidic pathway |  | x
ABL90 | siehe ausgeführte Arbeiten | siehe ausgeführte Arbeiten | CD7388F9F0135111 | No | see work carried out | see work carried out | Defect | SP | Fluid transport |  | x
ABL90 | siehe ausgeführte Arbeiten | siehe ausgeführte Arbeiten | CD736353F020510D | No | see work carried out | see work carried out | Defect | HW | CPU | See previous UpSWO | x
ABL90 | Lmpieza del sistema con agua. Desmontaje y limpieza de la aguja cal, qc y mustras | Lmpieza del sistema con agua. Desmontaje y limpieza de la aguja cal, qc y mustras | CD736F2FF05DA76A | No | It's a water-containing system. Disassembly and cleaning of the lime needle, qc and mustras | It's a water-containing system. Disassembly and cleaning of the lime needle, qc and mustras | Cleaned | Fluidic pathway | Clot |  | x
ABL825 FLEX | Cleaned measuring chamber and waste lines-No change | Replace listed parts and inlet gasket tHb calibration in service mode X2 | CD7395DCF00B1BF9 | No | Cleaned measuring chamber and waste lines-No change | Replace listed parts and inlet gasket tHb calibration in service mode X2 | Defect | Inlet module | Probe
ABL825 FLEX | DATA LOGS SHOW INTERMITTENT LIQUID DETECTED PREMATURELY DURING GAS CALIBRATION AND 653 SAMPLE IN PH BG LIQUID SENSOR DURING QC MEASUREMENT | REPLACED PH BG AND EL MET AIR VALVES AND PUMP TUBINGS.REPLACED SOLUTIONS AND WASTE TUBINGS.REPLACED LEAKING REFERENCE MEMBRANES.PERFORMED CALIBRATIONS | CD736B71F0075484 | No | DATA LOGS SHOW INTERMITTENT LIQUID DETECTED PREMATURELY DURING GAS CALIBRATION AND 653 SAMPLE IN PH BG LIQUID SENSOR DURING QC MEASUREMENT | REPLACED PH BG AND EL MET AIR VALVES AND PUMP TUBINGS. REPLACED SOLUTIONS AND WASTE TUBINGS. REPLACED LEAKING REFERENCE MEMBRANES. PERFORMED CALIBRATIONS | Defect | pH/BG Module | Valve air
AQT90 FLEX |  |  | CD738E21F0564B17 | Yes |  |  | Re classified RMED | Preventive maintenance | and SW upgrade
AQT90 FLEX | Verbindungsstatus | servicedump, RemoteSupport Key | CD73B2C2F03B267D | No | Verbindungsstatus | servicedump, RemoteSupport Key
AQT90 FLEX | Verbindungsstatus | servicedump; RemoteSupport Key installiert | CD73AFC2F0172678 | No | Verbindungsstatus | servicedump; RemoteSupport Key installed
AQT90 FLEX |  |  | CD788004C816800F | No
ABL80 FLEX Basic version |  |  | CD74ADC26652B49A | No
ABL90 |  |  | CD74C6CC66346CDC | Yes |  |  | Re classified RMED | planned SW upgrade |  |  | x
ABL90 |  |  | CD74853C662F69A6 | No |  |  | Defect | HW | Flow selector motor |  | x
ABL90 | Troubleshot to a faulty Sensor Cassette Bay Unit | Replaced the sensor cassette bay unit, but this did not resolve the problem. The analyzer requested a new solution pack so I changed the pack and got error code 1343 (Unable to pump solutions). I installed a new inlet probe and performed a back flush. This produced a steady stream. Retested the analyzer and I am still getting error code 1343. Replaced the sensor cassette, and this resolved the problem. | CD74C2A1663635BB | No | Troubleshot to a faulty Sensor Cassette Bay Unit | Replaced the sensor cassette bay unit, but this did not resolve the problem. The analyzer requested a new solution pack so I changed the pack and got error code 1343 (Unable to pump solutions). I installed a new inlet probe and performed a back flush. This produced a steady stream. Retested the analyzer and I am still getting error code 1343. Replaced the sensor cassette, and this resolved the problem. | Defect | SC | Fluid transport |  | x
ABL827 FLEX |  |  | CD74F3D5662879FC | No |  |  | LEH - Unknown error
ABL825 FLEX |  |  | CD748C9B6619164B | No |  |  | Cleaned | pH/BG Module | Measuring chamber
ABL815 FLEX | Verificação dos registros, descontaminação, limpeza das camaras de medida com solução de limpeza e água destilada, limpeza, inspecção dos vedantes e troca das membranas dos electrodos, inspecção dos tubos e roletes de bombas, execução do teste completo da oximetria, teste das valvulas, limpeza do ciruito da valvula do ar, calibração do zero da corrente, ajuste dos SLs, calibração das bombas, inspecção e ajuste da sonda, troca da entrada, limpeza do filtro, estabilização e passagem dos QCs. |  | CD74D40F662F90C0 | No | Checking of records, decontamination, cleaning of the measuring chambers with cleaning solution and distilled water, cleaning, inspecting the seals and replacing the membranes of the electrodes, inspection of the tubes and roller pumps, execution of the complete test of the Oximetry, valve test, air valve circuit cleaning, current zero calibration, SLs adjustment, pump calibration, probe inspection and adjustment, inlet exchange, filter cleaning, QCs stabilization and passage. |  | Defect | Electrode | Membrane | Lac
ABL825 FLEX |  |  | CD75E270DC42D3D5 | No |  |  | Defect | Electrode | Membrane | Glu
ABL825 FLEX |  |  | CD74EB626625355A | No |  |  | Defect | FlexQ Module | Mixer misaligned
AQT90 FLEX | Desmontaje de la rueda de incubacion y retirata del pocillo. Realizacion de limpieza. Comprobacion con qc. | Desmontaje de la rueda de incubacion y retirata del pocillo. Realizacion de limpieza. Comprobacion con qc. | CD74DAB8661F25D6 | No | Disassembly of the incubation wheel and removal from the well. Cleaning. Check with qc. | Disassembly of the incubation wheel and removal from the well. Cleaning. Check with qc.
ABL90 |  |  | CD75D836DC5273BA | No |  |  | Defect | HW | Touch screen |  | x
ABL90 |  |  | CD75F562DC4187BE | No |  |  | Defect | SC | pCO2 Cal - sens |  | x
ABL90 | tHb Kal. i.O. 842-345 Waste tubing undicht. | System mit Proteinentferner gespült, 842-345 Waste tubing erneuert. | CD761FC7DC261968 | No | tHb Kal. i.O. 842-345 Waste tubing undicht. | System rinsed with protein remover, 842-345 waste tubing renewed. | Defect | HW | waster tube | & Cleaned fluidic pathway | x
ABL90 |  |  | CD75D8C3DC4864BE | Yes |  |  | Re classified RMED | TU116
ABL90 | Battery pack smoke | replaced battery pack | CD75D69ADC39AFD6 | No | Battery pack smoke | replaced battery pack | Defect | HW | Battery module | smoke | x
ABL90 | None | Unpacked analyzer. Installed reagents. Ran system cycles. Error code 1070 and 1079. Ordered and installed new sensor cassette but same error codes. Ran several calibrations. Ordering sensor interface PCB. Performed sensor module alignment with alignment tool. Replaced Sensor Interface board. | CD75E02DDC5BEFAE | Yes | None | Unpacked analyzer. Installed reagents. Ran system cycles. Error code 1070 and 1079. Ordered and installed new sensor cassette but same error codes. Ran several calibrations. Ordering sensor interface PCB. Performed sensor module alignment with alignment tool. Replaced Sensor Interface board. | Re classified RMED | OOB
ABL90 | siehe ausgeführte Arbeiten | siehe ausgeführte Arbeiten | CD7600DCDC152BC5 | No | see work carried out | see work carried out | Defect | SP | Fluid transport |  | x
ABL825 FLEX | pH und Ref | Reinigung Austausch pH Elektrode | CD761346DC5B11A9 | No | pH and Ref | Cleaning replacement pH electrode | Defect | Electrode | Electrode | pH
ABL800Basic |  |  | CD75E263DC40B7B8 | No |  |  | Defect | Electrode | Membrane | Lac
ABL835 FLEX | Oximetry Complete Test Funktion vom Quetschventil | Austausch Quetschventil | CD76072CDC2811A3 | No | Oximetry Complete Test Function from the crushing valve | Replacement of squeezing valve | Defect | HW | Oxi Module | Pinch Valve
ABL827 FLEX |  |  | CD75DE9CDC1DD3C0 | No |  |  | Defect | Electrode | Electrode | Reference
ABL815 FLEX |  |  | CD75E4E0DC2071F4 | No |  |  | Defect | Inlet module | Gasket
ABL825 FLEX | Reviewed data logs, put clear tubings to observe for air bubbles. | Replaced Measuring chamber, inlet preheater. Multiple 2pt Cals fpor all parameters to come in. | CD7615C1DC0002EF | No | Reviewed data logs, put clear tubings to observe for air bubbles. | Replaced Measuring chamber, inlet preheater. Multiple 2pt Cals fpor all parameters to come in. | LEH - Unknown error
ABL825 FLEX | OXI alignment verified. | Perfor lamp alignment. Perform Complete OXI Test. Perform service tHb calibration. | CD760900DC3DD201 | No | OXI alignment verified. | Perfor lamp alignment. Perform Complete OXI Test. Perform service tHb calibration. | Adjustment | HW | Oxi Module | Lamp Unit
ABL825 FLEX | OXI Module alignment verified. | Perform alignment of lamp unit. Perform Complete OXI Test. Perform service tHb calibration. | CD76059BDC25D1FE | No | OXI Module alignment verified. | Perform alignment of lamp unit. Perform Complete OXI Test. Perform service tHb calibration. | Adjustment | HW | Oxi Module | Lamp Unit
AQT90 FLEX |  |  | CD75E779DC016D75 | Yes |  |  | Re classified RMED | Planned Setup
ABL80 FLEX Basic version |  |  | CD7710D4525BEEAA | No
ABL825 FLEX | Inspection of the print head revealed all connections are correct and print head is not dirty. | Replaced print head to resolve the reported issue. | CD773BA052056937 | No | Inspection of the print head revealed all connections are correct and print head is not dirty. | Replaced print head to resolve the reported issue. | Defect | Printer module | Printer head
ABL827 FLEX | checked flow of solutions, checked pump tubing and waste pump rotor, checked inlet gasket | removed and replaced broken inlet nipples, replaced clear waste tube | CD77ACD3520B835D | No | checked flow of solutions, checked pump tubing and waste pump rotor, checked inlet gasket | removed and replaced broken inlet nipples, replaced clear waste tube | Defect | Inlet module | Nipple kit
AQT90 FLEX |  |  | CD7726265241270F | No
ABL80 CO-OX | PO2 failed on level 2 of the linearity kit. | Did a zero calibration. Ran aa VK-R5 kit. Douwnloaded the data logs. Set the baud rate so the unit will work with their LIS system. | CD787EBAC84848B8 | No | PO2 failed on level 2 of the linearity kit. | Did a zero calibration. Ran aa VK-R5 kit. Douwnloaded the data logs. Set the baud rate so the unit will work with their LIS system.
ABL825 FLEX |  | Waste Detector getauscht | CD788889C834EBE9 | No |  | Waste Detector exchanged | Defect | Basic unit | Waste detector
TCM/Tosca |  |  | CD785C58C80D9AC2 | No |  |  | Defect | SW | Compact Flash
ABL90 |  |  | CD7C46DA2A6325CA | Yes |  |  | Re classified RMED | TU116
ABL90 | 1324, 1343,1086 error codes. Primary issue is inhomogeneous rinse errors. Customer changed consumables, backflushed analyzer and changed SP several times. Error code is intermittent during cal and QC. Troubleshot looking for air introduction into system, intstalled new tubings, probe, gasket, and cleaned and trimmed waste line to fitting.Replaced SP. Ran diagnostics and troubleshooting procedures in service manual appendex. Note: Customer has several other ABL 90s in same area. All show instances of 1324 inhomogeneous rinse errros. This analyzer, however, cannot be totally cleared of the error. Suspect either SP consumable transport storage issue or analyzer issue with LS sensors in bay unit. Diagnostics good. | No correction at present. | CD7DB0A4A05C2614 | No | 1324, 1343,1086 error codes. Primary issue is inhomogeneous rinse errors. Customer changed consumables, backflushed analyzer and changed SP several times. Error code is intermittent during cal and QC. Troubleshot looking for air introduction into system, intstalled new tubings, probe, gasket, and cleaned and trimmed waste line to fitting. Replaced SP. Ran diagnostics and troubleshooting procedures in service manual appendex. Note: Customer has several other ABL 90s in same area. All show instances of 1324 inhomogeneous rinse errros. This analyzer, however, cannot be totally cleared of the error. Suspect either SP consumable transport storage issue or analyzer issue with LS sensors in bay unit. Diagnostics good. | No correction at present.
ABL90 | 1324, 1343 and 1086 errors. Customer changed consumables, backflushed and changed SP several times. States analyzer has started asking for a SP change at varying intervals for no apparent reason. Troubleshot for suspected flow or sensing related problem. Installed new tubings to address any air being introduced into system. Installed probe, inlet gasket, trimmed waste line to fitting. Replaced SP. Ran diagnostics and troubleshooting procedures in service manual appendix. | No correction at present. | CD8153ED0231A4CC | No | 1324, 1343 and 1086 errors. Customer changed consumables, backflushed and changed SP several times. States analyzer has started asking for a SP change at varying intervals for no apparent reason. Troubleshot for suspected flow or sensing related problem. Installed new tubings to address any air being introduced into system. Installed probe, inlet gasket, trimmed waste line to fitting. Replaced SP. Ran diagnostics and troubleshooting procedures in service manual appendix. | No correction at present. | Unknown | Revisit required | Fluid transport | see next UpSWO | x
ABL825 FLEX | All oxi vaalues are failing cals and QC. Found that it had a 579 error message. Ran the comple oxi test per the instructions.ffound the the Dark standard deviation failed. | Ordering a power supply, hemolyze, and an oxi module pcb. Will return on 7-22-14 to repair. | CD7C41EF2A5E9309 | No | All oxi vaalues are failing cals and QC. Found that it had a 579 error message. Ran the comple oxi test per the instructions.ffound the the Dark standard deviation failed. | Ordering a power supply, hemolyze, and an oxi module pcb. Will return on 7-22-14 to repair. | Defect | HW | Oxi Module | Next upSWO CD7DC713A4B7FA
ABL835 FLEX | Measurement unstable for all parameters, inlet leaking and it was also discovered that inlet spits fluid. Instrument occasionally failing inlet calibration. Inspected all lines, no pinched lines or clogged waste tubing found. Flushed sample probe from inlet pre- heater and flushed inlet pre- heater from inlet probe nipple. Inspected measuring module valves and found both air valves jammed in open position. Encountered issues with inlet flaps not being detected (tried different inlet PCB and flap detector PCB) | Replaced measuring module valves. Replaced Friction detector board and inlet flap detector pcb with spare (will return with new one) | CD7C39062A410BCD | No | Measurement unstable for all parameters, inlet leaking and it was also discovered that inlet spits fluid. Instrument occasionally failing inlet calibration. Inspected all lines, no pinched lines or clogged waste tubing found. Flushed sample probe from inlet pre- heater and flushed inlet pre- heater from inlet probe nipple. Inspected measuring module valves and found both air valves jammed in open position. Encountered issues with inlet flaps not being detected (tried different inlet PCB and flap detector PCB) | Replaced measuring module valves. Replaced Friction detector board and inlet flap detector pcb with spare (will return with new one) | Defect | Inlet module | Friction detector unit | Caused by leak
ABL80 CO-OX | Kal.- IQC- Aktivitätendateien, Schlauchwege | Schlauchwege erneuert und Analysator gereinigt | CD7D80D6A014D74E | No | Kal.- IQC- Aktivitätendateien, Schlauchwege | Hoseways renewed and analyzer cleaned
ABL80 CO-OX | 1011, 1034 CO-OX errors in event log. | Flushed tubings. Cleaned front of analyzer at sensor waste connector. | CD7D9ED6A00A79DA | No | 1011, 1034 CO-OX errors in event log. | Flushed tubings. Cleaned front of analyzer at sensor waste connector.
ABL90 |  |  | CD7D986AA051CE1E | No |  |  | Defect | HW | Inlet Gasket | twisted - old version replaced with new version | x
ABL90 |  |  | CD7D9FEBA0296548 | No |  |  | Cleaned | Fluidic pathway |  |  | x
ABL815 FLEX |  |  | CD7D87B1A031A271 | No
ABL820 FLEX |  |  | CD7DA477A0506DFA | No |  |  | Defect | Electrode | Electrode | pO2
ABL827 FLEX | Customer shut down analyzer due to leak out front of inlet gasket. Brought system back up and troubleshot to blockage in one inlet fitting below inlet gasket. Problems with analyzer equilibrating due to being shut down for extended period of time. Mets, Lac, Glu, and Crea calibration and QC took extended time due to analyzer instability. | Cleared blockage, allowed extended time for analyzer equilibration. Ultimately ran good calibration and QC with no further problems. Recommended to customer to allow analyzer to settle for another 2 hours before use. | CD7DBED5A03B2619 | No | Customer shut down analyzer due to leak out front of inlet gasket. Brought system back up and troubleshot to blockage in one inlet fitting below inlet gasket. Problems with analyzer equilibrating due to being shut down for extended period of time. Mets, Lac, Glu, and Crea calibration and QC took extended time due to analyzer instability. | Cleared blockage, allowed extended time for analyzer equilibration. Ultimately ran good calibration and QC with no further problems. Recommended to customer to allow analyzer to settle for another 2 hours before use. | Cleaned | Fluidic pathway
ABL825 FLEX | PO2 and all oxi values will not pas cal. | Replaced the PO2 membrane from the customer stock. Adjusted the lamp on the hemolyzer. | CD7D8779A05E6ECA | No | PO2 and all oxi values will not pas cal. | Replaced the PO2 membrane from the customer stock. Adjusted the lamp on the hemolyzer. | Adjustment | HW | Oxi Module | Lamp Unit
ABL825 FLEX | All oxi values fail calibrations. | Replaced the power suppy and performed a thb calibration. | CD7DC713A04B70FA | No | All oxi values fail calibrations. | Replaced the power suppy and performed a thb calibration. | Defect | HW | Oxi Module | Revisit
ABL835 FLEX | Issue was identified yesterday as a faulty inlet flap detector PCB, possibly caused by water damage during the leaking inlet issues. | Replaced Flap detector PCB | CD7D6DB3A0410BDD | No | Issue was identified yesterday as a faulty inlet flap detector PCB, possibly caused by water damage during the leaking inlet issues. | Replaced Flap detector PCB | Defect | Inlet module | Flap detector unit | Caused by leak
ABL835 FLEX | On site to attend construction saftey training, get into UC - Mission Bay - badging system. Install option keys, setup NetOp Host and test. | Training, NetOp Host and option keys | CD7D3F7DA03D9A19 | Yes | On site to attend construction saftey training, get into UC - Mission Bay - badging system. Install option keys, setup NetOp Host and test. | Training, NetOp Host and option keys | Re classified RMED | Installation
TCM-COMBIM | Turned monitor on and observed colors that should be red are orange. Tried resetting cable - same. | Removed and replaced cable between the display and the main CPU. | CD7D942FA062E774 | No | Turned monitor on and observed colors that should be red are orange. Tried resetting cable - same. | Removed and replaced cable between the display and the main CPU. | Defect | HW | Cable | Cable
ABL90 | On site to install previously ordered parts. | Installed wet section PCB and bay assy. No further error codes related to original issue. Upon initialization, found neon lamp errors. Suspect module jarred from install of bay assy. | CD81589D025BA4D1 | No | On site to install previously ordered parts. | Installed wet section PCB and bay assy. No further error codes related to original issue. Upon initialization, found neon lamp errors. Suspect module jarred from install of bay assy. | Defect | HW | SC retainor | See previous UpSWO | x
ABL90 |  |  | CD8022588C50CA1D | No
ABL90 |  |  | CD7EC3FE16281EDD | No |  |  | Defect | HW | Inlet Gasket | old version replaced with new version | x
ABL90 |  |  | CD7ED4F5162A03DE | No
ABL835 FLEX | COllected data logs, met with customer to discuss this and other issues at Lab. | No correction at this time, investigation ongoing. | CD7F252B162A9A2A | No | COllected data logs, met with customer to discuss this and other issues at Lab. | No correction at this time, investigation ongoing. | Cleaned | Fluidic pathway | Membrane | Lac
ABL825 FLEX |  |  | CD7E9D3B160C23A3 | No |  |  | Defect | Electrode | Electrode | Reference
ABL820 FLEX | On site for the other analyzer when I noticed this system has been failing Cal's with Oxi error 579. Ran 3X Protein removal - no help. Oxi complete test fails all tested function. Recycled power on system and oximetry recovered normally | Cycle power and run Oxi complete test. | CD7E9FD8164E9A23 | No | On site for the other analyzer when I noticed this system has been failing Cal's with Oxi error 579. Ran 3X Protein removal - no help. Oxi complete test fails all tested function. Recycled power on system and oximetry recovered normally | Cycle power and run Oxi complete test. | Unknown | Power off/on
ABL820 FLEX | Replaced all tubings, inlet gasket and inlet probe. Had to open a PM Kit for parts. | Replaced parts from PM Kit | CD7E9EFD16129A22 | No | Replaced all tubings, inlet gasket and inlet probe. Had to open a PM Kit for parts. | Replaced parts from PM Kit | Defect | Basic unit | Tubings
ABL835 FLEX | DATA LOGS SHOW TEMP ERROR 356 EL MET MODULE TEMP ERROR | REPLACED EL MET LIGHT PCB | CD7ECEED16109DA6 | No | DATA LOGS SHOW TEMP ERROR 356 EL MET MODULE TEMP ERROR | REPLACED EL MET LIGHT PCB | Defect | El/Met Module | PCB | Meas. Chamber light
ABL835 FLEX | Went through each user screen and found that all reagent bottles were in yellow. Scanned bottles back in and hit restart. Instrument went through initialisation procedure and performed initial 2pt cal. Was able to run a second 2pt cal. Inspected all tubing carefully for pinholes. Tried new inlet pre-heater. Noticed small pool of liquid at base of pH BG measuring chamber, also inspected pH BG and El Met air valves. Replaced inlet gasket, probe tubing, inlet pre-heater tube, inlet to pH BG tube, pH o-ring, reference membrane. | Scanned in all bottles Replaced jammed (in open position) measuring module air valves (pH BG and El Met) and pH BG measuring chamber. Encountered issues with waste tubings (under inlet) falling off. Flushed both lines, but one end of syringe waste tubing kept falling off when I connected the other end. Replaced left inlet nipple and changed tubing. Changed clear tubing to black due to persistant issues of tubing falling off. | CD7EA19D16230BE7 | No | Went through each user screen and found that all reagent bottles were in yellow. Scanned bottles back in and hit restart. Instrument went through initialisation procedure and performed initial 2pt cal. Was able to run a second 2pt cal. Inspected all tubing carefully for pinholes. Tried new inlet pre-heater. Noticed small pool of liquid at base of pH BG measuring chamber, also inspected pH BG and El Met air valves. Replaced inlet gasket, probe tubing, inlet pre-heater tube, inlet to pH BG tube, pH o-ring, reference membrane. | Scanned in all bottles Replaced jammed (in open position) measuring module air valves (pH BG and El Met) and pH BG measuring chamber. Encountered issues with waste tubings (under inlet) falling off. Flushed both lines, but one end of syringe waste tubing kept falling off when I connected the other end. Replaced left inlet nipple and changed tubing. Changed clear tubing to black due to persistant issues of tubing falling off. | Defect | Inlet module | Nipple kit
ABL825 FLEX | Customer has tried 2 different lots of membranes - same issue. | Replaced PCO2 Electrode. | CD7EE15E165C5D1B | No | Customer has tried 2 different lots of membranes - same issue. | Replaced PCO2 Electrode. | Defect | Electrode | Electrode | pCO2
ABL825 FLEX | tHb reads 489 PA on Calibration. Hemolyzer tuning voltage failed in service test. | Replaced Hemolyzer. Performed alignment. tHb reads 569 PA on Calibrations. | CD7EBF9B161A7A8E | No | tHb reads 489 PA on Calibration. Hemolyzer tuning voltage failed in service test. | Replaced Hemolyzer. Performed alignment. tHb reads 569 PA on Calibrations. | Defect | HW | Oxi Module | Hemolyzer
ABL825 FLEX |  |  | CD7EB60A163A0A6D | No |  |  | Defect | HW | Oxi Module | Unknown
ABL80 CO-OX |  |  | CD8013148C202E04 | No
ABL90 |  |  | CD803F2C8C011B8E | No
ABL827 FLEX |  |  | CD8135D402083B54 | No |  |  | Defect | Electrode | Electrode | Crea
ABL835 FLEX | Performed protein removal, inspected sample path and cleaned measuring chambers. Replaced pH o-ring since an an air bubble was forming at the tip and fluid crystalisation on outside of pH electrode assembly. Noticed Lactate sens was too low and had been dropping over the past few days. Lactate membrane due to be changed on July 25. | Adjusted lamp unit intensity and performed complete oxi test. Replaced pH o-ring and Lactate membrane. | CD8017928C0F0C07 | No | Performed protein removal, inspected sample path and cleaned measuring chambers. Replaced pH o-ring since an an air bubble was forming at the tip and fluid crystalisation on outside of pH electrode assembly. Noticed Lactate sens was too low and had been dropping over the past few days. Lactate membrane due to be changed on July 25. | Adjusted lamp unit intensity and performed complete oxi test. Replaced pH o-ring and Lactate membrane. | Adjustment | HW | Oxi Module | Lamp Unit
ABL835 FLEX | Customer becoming quite unhappy of new analyzer. | Escalated issue and notified local SWAT team of the issues pertaining to SWOs B98739KF7, B9883LUS7, B97L3E46N, B9753WX84, B9554RZ2M, and B96U3RAD7 | CD8021328C0A0C0B | Yes | Customer becoming quite unhappy of new analyzer. | Escalated issue and notified local SWAT team of the issues pertaining to SWOs B98739KF7, B9883LUS7, B97L3E46N, B9753WX84, B9554RZ2M, and B96U3RAD7 | Re classified RMED | upSWO only to provide info.
ABL835 FLEX | Inspected sample path and noticed that there was a bubble forming under pH electrode. | Replaced pH electrode. Replaced QC due to heavy useage for troubleshooting purposes yesterday. | CD7FF23A8C170C02 | No | Inspected sample path and noticed that there was a bubble forming under pH electrode. | Replaced pH electrode. Replaced QC due to heavy useage for troubleshooting purposes yesterday. | Defect | Electrode | Electrode | gasket
ABL90 | Battery not functioning correctly. | Replace battery Module. | CD813C0F02437E60 | No | Battery not functioning correctly. | Replace battery Module.
ABL90 | Initial problem fixed on last related visit. Module was jarred in changing bay assy. Neon lamp failed. Ordered new hemolyzer and on site to install. | Installed new hemolyzer. | CD815B850200A4D7 | No | Initial problem fixed on last related visit. Module was jarred in changing bay assy. Neon lamp failed. Ordered new hemolyzer and on site to install. | Installed new hemolyzer. | Defect | HW | Hemolyzer | Perhaps damaged by FSE during last visit | x
ABL90 |  |  | CD81720202416E13 | No
ABL90 | Error code 1343 (Unable to pump solutions.Error 1083, 1085 and 1087 (inhomogenous Cal 2 solution, QC 1 and 2). | Replaced solutions pack. Replaced inlet gasket and probe. Reinstalled original inlet gasket and probe. | CD814079024BC5B6 | No | Error code 1343 (Unable to pump solutions. Error 1083, 1085 and 1087 (inhomogenous Cal 2 solution, QC 1 and 2). | Replaced solutions pack. Replaced inlet gasket and probe. Reinstalled original inlet gasket and probe. | Defect | SP | Fluid transport |  | x
ABL825 FLEX | CHEQUEO DE CAMARA DE MEDIDA Y LIMPIEZA, CHEQUEO DE ELECTRODO DE PH Y LIMPIEZA SE SUSTITUYE ELECTRODO DE PH Y AUNQUE LOS QC EN EL VALOR DE PH SALE DENTRO DE RANGOS, EL ESTADO EN CALIBRACIONES NO SUBE DE 6.9 SUSTITUCIÓN DE ELECTRODO DE REF Y CALIBRACIONES VARIAS, PASADO QC, CUMPLIMENTACIÓN DE PARTE DE TRABAJO | SUSTITUCIÓN DE ELECTRODOS DE PH Y REF | CD841A73EE51DA83 | No | MEASURE AND CLEANING CAMERA CHECK, PH ELECTRODE CHECK AND CLEANING IS SUBSTITUTE PH ELECTRODE AND EVEN THE QC IN THE VALUE OF PH SALE WITHIN RANGES, THE STATUS IN CALIBRATIONS DOES NOT RISE OF 6.9 REPLACEMENT OF REF ELECTRODO , PAST QC, COMPLETION OF PART OF WORK | REPLACEMENT OF PH AND REF ELECTRODES
ABL835 FLEX | Inspected sample path. Discovered that the calibration of electrolytes was failing intermittently due to high drift, pO2 calibrations were also failing from time to time due to drift being outside of tolerence limits, and QC results were elevated for oximetry values (according to customeer's ranges). Reference membrane changed on July 11, near the onset of issues. Reference membrane lot R1700 Changed reference membrane and pH o-ring. Tried different pH electrode. Successfully ran patient sample, so tried customer's electrode again, still running sampled ok for 4 samples, but issues returned. | Replaced inlet pre-heater tube, inlet - pH BG tube and inlet probe tubing. | CD810B70020CF4E7 | No | Inspected sample path. Discovered that the calibration of electrolytes was failing intermittently due to high drift, pO2 calibrations were also failing from time to time due to drift being outside of tolerence limits, and QC results were elevated for oximetry values (according to customeer's ranges). Reference membrane changed on July 11, near the onset of issues. Reference membrane lot R1700 Changed reference membrane and pH o-ring. Tried different pH electrode. Successfully ran patient sample, so tried customer's electrode again, still running sampled ok for 4 samples, but issues returned. | Replaced inlet pre-heater tube, inlet - pH BG tube and inlet probe tubing.
ABL80 OSM | Using the Zero Check the Therm check was within range (18-30), but higher than normal at 27.7 degrees. | Replaced the Sensor Connector Board and Analog Board. The I performed the Zero Check a second time and the Therm Check was now at 24.7 degrees. The customer is notify me if the analyzer shutsdown again. | CD852CA7641847D8 | No | Using the Zero Check the Therm check was within range (18-30), but higher than normal at 27.7 degrees. | Replaced the Sensor Connector Board and Analog Board. The I performed the Zero Check a second time and the Therm Check was now at 24.7 degrees. The customer is notify me if the analyzer shutsdown again.
ABL90 |  |  | CD85156E642AF3D5 | No |  |  | Cleaned | Fluidic pathway
ABL90 |  |  | CD851801641F9A21 | No |  |  | User Error | Incorrect SP activation | Cleaned Fluidic pathway |  | x
ABL90 |  |  | CD84DE3B642BF2E9 | No
ABL90 |  |  | CD84ECB264475CDB | No |  |  | Defect | SC | FT |  | x
ABL800Basic |  |  | CD865542DA1A8E6E | No |  |  | Defect | Electrode | Membrane | Lac
ABL800Basic |  | Ventile überprüft--> beide Airventile in den Messkammer fest--> gereinigt; Lac neu membraniert | CD84FC8C645A500A | No |  | Valves checked--> both air valves in the measuring chamber fixed-->cleaned; Lac re-membraneed | Defect | Electrode | Membrane | Lac
ABL800Basic |  | Massekontakt gerichtet | CD8513F36417500F | No |  | Mass contact directed | Defect | El/Met Module | Electrode connector
ABL800Basic |  | Nullpunktkalibrierung durchgeführt; pH Elektrode gereinigt | CD84E5D864325005 | No |  | zero-point calibration performed; pH electrode cleaned | Defect | pH/BG Module | Electrode | pO2
ABL80 OSM |  |  | CD86649EDA2A9A14 | Yes |  |  | Re classified RMED | Preventive maintenance
ABL80 CO-OX | Tried another, same results | Replace I Button and installed new solution pack. Now recognizes new pack. | CD863C85DA3C3369 | No | Tried another, same results | Replace I Button and installed new solution pack. Now recognizes new pack.
ABL90 | Checked activity log for errors. Found Inhomogeneous rinse and QC errors. Also found inlet gasket with holder not seated correctly. Software at version 3.0 MR2. Needs upgraded to 3.1 Customer using Marquest syringes 50% of the time. | Reseated inlet components. Flushed system multiple times. After flushing replaced parts from maintenance kit as needed. Also replaced inlet probe connector gasket. Replaced sensor cassette to clear remaining errors because of possible silicone contamination. Upgraded software to revision 3.1. Glucose and Lactate parameters would not pass internal QC. Ran multiple calibrations to help bring sensitivity values up. | CD868717DA4452BB | No | Checked activity log for errors. Found Inhomogeneous rinse and QC errors. Also found inlet gasket with holder not seated correctly. Software at version 3.0 MR2. Needs upgraded to 3.1 Customer using Marquest syringes 50% of the time. | Reseated inlet components. Flushed system multiple times. After flushing replaced parts from maintenance kit as needed. Also replaced inlet probe connector gasket. Replaced sensor cassette to clear remaining errors because of possible silicone contamination. Upgraded software to revision 3.1. Glucose and Lactate parameters would not pass internal QC. Ran multiple calibrations to help bring sensitivity values up. | Cleaned | Fluidic pathway | Potential silicone
ABL800Basic |  | Trotz Reinigung des Scanners keine Besserung--> Austausch des Laserscanners | CD8650E7DA3D501D | No |  | Despite cleaning the scanner no improvement--> replacement of the laser scanner | Defect | Auto Check Module | Barcode Reader
ABL835 FLEX | This problem is occurring on several analyzers but it more pronounced on this system. QC fails, out of range low, on all levels necessitating a membrane change after 30 days vice the usual 90 days. Affected lot codes are R0486 and R0458. | Will open Techline case. Have sent customer a box of membranes to replenish those used in troubleshooting | CD86062EDA2DA924 | No | This problem is occurring on several analyzers but it more pronounced on this system. QC fails, out of range low, on all levels necessitating a membrane change after 30 days vice the usual 90 days. Affected lot codes are R0486 and R0458. | Will open Techline case. Have sent customer a box of membranes to replenish those used in troubleshooting | Defect | Electrode | Membrane | Ca+
ABL825 FLEX |  |  | CD864FEADA5823CE | Yes |  |  | Re classified RMED | Setup
ABL825 FLEX | Oximetry messuring error. Error message 269. | Performed a complete Oxi test. All test passed. Performed a service THB test. | CD86A627DA4C00FA | No | Oximetry messuring error. Error message 269. | Performed a complete Oxi test. All test passed. Performed a service THB test. | Adjustment | HW | Oxi Module | tHb cal
AQT90 FLEX | Recarga de software fallida. Se piden piezas |  | CD864DB5DA58927E | Yes | Failed software reload. Parts are ordered |  | Re classified RMED | SWU
ABL80 OSM |  |  | CD877C2C5038101D | No
ABL80 CO-OX | Found CPU cooling fan making noise. | Replaced CPU cooling fan. | CD87B968501DAB18 | No | Found CPU cooling fan making noise. | Replaced CPU cooling fan.
ABL827 FLEX | found the demo keys had expired and instrument not sending data to lis | ordered and installed lis his permanent key | CD87A6A5503F9141 | Yes | found the demo keys had expired and instrument not sending data to lis | ordered and installed lis his permanent key | Re classified RMED | Setup | Exp demo
ABL825 FLEX |  |  | CD8783215023402B | No |  |  | Cleaned | Basic unit | Barcode Reader
ABL80 FLEX Basic version |  |  | CD88FC04C656AF0A | No
ABL80 CO-OX | ran service test and fails hemolyzer test. | tried to run thb calibration | CD88B4E8C62C28E4 | No | ran service test and fails hemolyzer test. | tried to run thb calibration
ABL90 | Troubleshot looking through data logs for other errors. Did Fluidic Check in Service - All passed. Did Complete Oximetry Test - all passed. | Removed and replaced Solution Pack. | CD89000FC662365F | No | Troubleshot looking through data logs for other errors. Did Fluidic Check in Service - All passed. Did Complete Oximetry Test - all passed. | Removed and replaced Solution Pack. | Adjustment | tHb cal | replaced SP?
ABL90 | Comprobación del estado del analizador. Equipo no dispone de más reactivo en pack de soluciones. Se comenta con cliente la incidencia, piden material al almacen, se espera para recibir material, cuando se obtiene se procede a la sustitución del pack de soluciones y se explica como se debe cambiar así como casete de electrodos, que también se cambia. No disponen de Cal. de Hb, se realizan varias Cal y Qc´s para estabilizar el analizador, se observa fallo en la Oximetría durante los Qc´s posibles restos en el hemolizador, se realizan test, todos OK y varias limpiezas internas, finalmente se consigue que los Qc´s esten OK. Funcionamiento correcto. | Limpieza interna del analizador. | CD88D28FC6566B13 | No | Checking the status of the scanner. Equipment does not have more reagent in solution pack. The incident is discussed with the customer, they ask for material at the warehouse, it is expected to receive material, when it is obtained proceeds to the replacement of the solution pack and explains how it should be changed as well as electrode cassette, which is also changed. They do not have Hb Cal., several Cal and Qc's are performed to stabilize the analyzer, oxymetry failure is observed during the possible residues in the hemolizer, tests are performed, all OK and several internal cleanings, finally the Qc's are OK. Correct operation. | Internal cleaning of the analyzer. | Cleaned | Fluidic pathway | Oxi
ABL825 FLEX | A RINSE WAS PERFORMED AND THE INLET GASKET WAS LEAKING. cUSTOMER DID NOT HAVE A NEW INLET GASKET TO USE | THE INLET GASKET WAS REPLACED AND OPERATION WAS CHECKED OVER SEVERAL HOURS WITH NO LEAKS TO REPORT | CD92BA137627F660 | No | A RINSE WAS PERFORMED AND THE INLET GASKET WAS LEAKING. cUSTOMER DID NOT HAVE A NEW INLET GASKET TO USE | THE INLET GASKET WAS REPLACED AND OPERATION WAS CHECKED OVER SEVERAL HOURS WITH NO LEAKS TO REPORT | User Error | Replacement process | Inlet Gasket
ABL800Basic |  | Lac-Elektroden spitze gereinigt und Elektrode neu membraniert | CD893B2FC65C89CE | No |  | Lac electrodes pointedly cleaned and electrode re-membraneed | Defect | Electrode | Membrane | Lac
ABL800Basic | Nullpunktkalibrierung, 2Pkt.-Kalibrierungen durchgeführt | Lac-Elektrode membraniert; pH BG-Modul ausgetauscht | CD9007C68A5D90D2 | No | zero calibration; 2pc calibrations performed | Lac electrode membranes; pH BG module replaced | Defect | pH/BG Module | PCB
ABL80 CO-OX |  |  | CD89DBF53C017C6F | No
ABL90 |  |  | CD8A5AD93C308611 | No |  |  | Replaced | Inlet gasket | Perhaps old version |  | x
ABL90 |  |  | CD8A35E63C607217 | No |  |  | Replaced | Probe | Bend | & SP not correct activated | x
ABL90 |  |  | CD8A28BF3C0B6FCF | Yes |  |  | Re classified RMED | Preventive maintenance
ABL825 FLEX | Reviewed logs. Checked for re-occurring problems or unexplainable errors. | Tested all the valves, pumps and liquid sensors. Replaced the tubing, fan filter and inlet gasket using Yearly Maintenance Service Kit. Flushed the air valve in the fluidic block. Cleaned and reseated the dual valves. Cleaned analyzer. | CD8A38153C30FAAE | Yes | Reviewed logs. Checked for re-occurring problems or unexplainable errors. | Tested all the valves, pumps and liquid sensors. Replaced the tubing, fan filter and inlet gasket using Yearly Maintenance Service Kit. Flushed the air valve in the fluidic block. Cleaned and reseated the dual valves. Cleaned analyzer. | Re classified RMED | Preventive maintenance
ABL825 FLEX | Reviewed logs. Checked for re-occurring problems or unexplainable errors. | Tested all the valves, pumps and liquid sensors. Replaced the tubing, fan filter and inlet gasket using Yearly Maintenance Service Kit. Flushed the air valve in the fluidic block. Cleaned and reseated the dual valves. Cleaned analyzer. | CD8A32833C37FAAA | Yes | Reviewed logs. Checked for re-occurring problems or unexplainable errors. | Tested all the valves, pumps and liquid sensors. Replaced the tubing, fan filter and inlet gasket using Yearly Maintenance Service Kit. Flushed the air valve in the fluidic block. Cleaned and reseated the dual valves. Cleaned analyzer. | Re classified RMED | Preventive maintenance
ABL825 FLEX | When running blood unit reading lower when compared to Resp. and other ABL825. In regards to unit not displaying HBCO. The low end of the reportable range is set to .9% and any readings below .9 will not be displayed. | Checked lamp adjustment on all three analyzer. Performed service THb cal. | CD8A273A3C08C87F | No | When running blood unit reading lower when compared to Resp. and other ABL825. In regards to unit not displaying HBCO. The low end of the reportable range is set to .9% and any readings below .9 will not be displayed. | Checked lamp adjustment on all three analyzer. Performed service THb cal. | Adjustment | HW | Oxi Module | tHb cal
ABL825 FLEX |  |  | CDA5CE5C604637A4 | No |  |  | Defect | Inlet module | Inlet Gasket
AQT90 FLEX |  |  | CD8A50793C00AB0B | No
TCM-COMBIM | THERE WAS NO COMMUNICATION BETWEEN THE SENSOR MODULE AND THE MONITOR | THE CFG FILE WAS DELETED AND THE TCM 4 AND COMBI M SOFTWARE WAS REINSTALLED | CD89E69E3C575C2E | No | THERE WAS NO COMMUNICATION BETWEEN THE SENSOR MODULE AND THE MONITOR | THE CFG FILE WAS DELETED AND THE TCM 4 AND COMBI M SOFTWARE WAS REINSTALLED | Defect | SW | SW upgrade
ABL90 | Wasteschlauch undicht | Austausch 842-345 | CD8B8CC2B235E118 | No | Waste hose leaks | Replacement 842-345 | Defect | HW | Waste tube | leak current detected | x
ABL80 CO-OX |  | Patiententestmessungen im pO2 sehr instabil--> Set Zero Kalibration durchgeführt; Kal- und QC Level 1,2; Daten gezogen | CD8DE6ED9E595E57 | No |  | Patient test measurements performed in pO2 very unstable--> Set Zero calibration; Cal and QC Level 1.2; Data drawn
ABL80 CO-OX |  | Patiententestmessungen im pO2 sehr instabil--> Set Zero Kalibration durchgeführt; Kal- und QC Level 1,2; Daten gezogen | CD8DE1049E635E52 | No |  | Patient test measurements performed in pO2 very unstable--> Set Zero calibration; Cal and QC Level 1.2; Data drawn
ABL80 CO-OX | Leak coming from around wadte connecter. | Cleaned analyzer interior, exterior and counter area affected by leak. Checked tubings valves and pumps. OK. Performed all petinent diagnostic tests with no failures. | CD8DD7CA9E353A8F | No | Leak coming from around wadte connect. | Cleaned analyzer interior, exterior and counter area affected by leak. Checked tubings valves and pumps. OK. Performed all petinent diagnostic tests with no failures.
ABL80 CO-OX | problem started when a clot caused leak into ABL.Sensor bd corroded. | Replaced defective parts.Calibrated Unit. | CDA4B9D9EA4472E8 | No | problem started when a clot caused leak into ABL. Sensor bd corroded. | Replaced defective parts. Calibrated Unit.
ABL825 FLEX | Problem appears to be intermittent. | Performed Zero Current calibration. New Lot of membranes will be used at next replacement. | CD8E032F9E0543BA | No | Problem appears to be intermittent. | Performed Zero Current calibration. New Lot of membranes will be used at next replacement. | Adjustment | El/Met Module | Zero calibration | Glu
AQT90 FLEX |  |  | CD8F16E4145CCD89 | No
AQT90 FLEX |  |  | CD8E1C4A9E3309D7 | No
ABL80 CO-OX | ran system test and it fails hemolyzer test. | I replaced the hemolyzer. | CD8EE53314269CD3 | No | ran system test and it fails hemolyzer test. | I replaced the hemolyzer.
ABL90 |  |  | CD8ED5FE14006FEA | No |  |  | Defect | HW | SC PCB | pO2 dark current | x
ABL90 |  |  | CD8F13461444F655 | No |  |  | No fault found | Leak detected | User couldn't solve by GT | In service no problems detected | x
ABL90 |  |  | CD8F4299145B4152 | No |  |  | Defect | HW | SC PCB | Temp problem with multi SC | x
ABL90 | The analyzer gave unable to pump sol and asked for a SP. | I ran all test and until i removed the SC the SP program failed. I replaced the SC and waited for Lac to come in. | CD8F5C5D14382067 | No | The analyzer gave unable to pump sol and asked for a SP. | I ran all test and until i removed the SC the SP program failed. I replaced the SC and waited for Lac to come in. | Defect | SC | Fluid transport
ABL90 |  |  | CD8F2945144E6FF5 | No |  |  | Defect | HW | Battery module |  | x
ABL90 | Comprobación del estado del analizador. A mi llegada se observa pack de soluciones mal instalado, se procede a re-intalar y explicar la activación del pack de soluciones. Se entra en servicio y se realizan test de servicio, se comprueba que todo es correcto. Se realizan Cal y Qc´s para comprobar funcionamiento. Funcionamiento correcto. | Activación pack de soluciones y explicación al cliente de como realizarlo. | CD8F0454141FA3CD | Yes | Checking the status of the scanner. On my arrival we notice poorly installed solution pack, proceed to re-install and explain the activation of the solution pack. You go into service and perform service tests, you check that everything is correct. Cal and Qc's are performed to check operation. Correct operation. | Activation pack of solutions and explanation to the client how to perform it. | Re classified RMED | Telephone support
ABL825 FLEX |  |  | CD8F022A142A5A87 | No |  |  | Cleaned | pH/BG Module | Measuring chamber | pO2
ABL825 FLEX |  |  | CD8F37A0145F10D5 | No |  |  | Defect | Electrode | Electrode | pCO2
ABL835 FLEX |  |  | CD8EE570143E0D3B | No |  |  | Defect | Gas 2 | -
ABL80 CO-OX |  | Austausch des CPU Lüfters | CD9031A28A5A8CC9 | No |  | Replacing the CPU fan
ABL90 |  |  | CD904F558A25670D | Yes |  |  | Re classified RMED | Logistic upSWO
ABL90 |  |  | CD9050328A09670F | Yes |  |  | Re classified RMED | Logistic upSWO
ABL820 FLEX | linker Abfallschlauch war lose (Gerinnsel im Stutzen unter der Einlassdichtung), gesäubert, Schlauch befestigt, Kal, QC, Reinigungslösung war nicht ordnungsgemäß aktiviert (Streptokinaseadditiv), aktiviert |  | CD9081788A5F73F2 | No | left waste hose was loose (clot in the nozzle under the inlet seal), cleaned, hose attached, Kal, QC, cleaning solution was not properly activated (Streptokinase additive), activated |  | Defect | Inlet module | Tubings | Loose
ABL825 FLEX | checked logs marc poirier was there on August 1 to do service call. | none. Dropped off lactate membranes as they have to be changed every 2 weeks. Changed cal1 frequency to 1hr. | CD9082028A1DF782 | Yes | checked logs marc poirier was there on August 1 to do service call. | none. Dropped off lactate membranes as they have to be changed every 2 weeks. Changed cal1 frequency to 1hr. | Re classified RMED | Logistic upSWO
ABL835 FLEX |  |  | CD90763F8A24C006 | No |  |  | Defect | Electrode | Electrode | pCO2
ABL825 FLEX |  |  | CD9047EA8A3B5AC9 | No |  |  | Defect | pH/BG Module | Pump tubing
ABL815 FLEX |  | Membrane R0576 abgezogen | CD9078568A493C33 | No |  | Membrane R0576 subtracted | Defect | Electrode | Membrane | pCO2
ABL825 FLEX | EL MET pump error. | Entered the service program, ran a pump calibration. | CD90314B8A322666 | No | EL MET pump error. | Entered the service program, ran a pump calibration. | Adjustment | El/Met Module | Pump cal error
TCM-COMBIM |  |  | CDA00020120361CD | No |  |  | Defect | HW | Cleaned
ABL90 | Inspection of the analyzer revealed no apparent external cold or warm source that would adversely affect the proper operating temperature of the analyzer. Service menu temperature check revealed substrate SC temp was out of specs and needs to be calibrated. | Performed service menu temperature calibration to resolve all reported issues. | CD92A6E476211786 | No | Inspection of the analyzer revealed no apparent external cold or warm source that would adversely affect the proper operating temperature of the analyzer. Service menu temperature check revealed substrate SC temp was out of specs and needs to be calibrated. | Performed service menu temperature calibration to resolve all reported issues. | Defect | SC | Termistor Cal.
ABL90 | keine Undichtigkeit im schlauchsystem festgestellt | Austausch LP, danach sinkt Leak Voltage | CD91ABD100568CEB | No | no leakage detected in the hose system | Replacement LP, after which Leak Voltage sinks | Defect | SP | Leak current |  | x
ABL90 | Found signs of leakage around Hemolyzer. | Replaced Hemolyzer. Ran blood sample Spectrum mismatch still present. Returning to change out Spectrometer. Replaced spectrometer. | CD91BED6005DFC19 | No | Found signs of leakage around Hemolyzer. | Replaced Hemolyzer. Ran blood sample Spectrum mismatch still present. Returning to change out Spectrometer. Replaced spectrometer. | Defect | HW | Spectrometer | Follow Up ?
ABL90 | Comprobar SP y circuito neumático, desobstruir cubeta de medida de la hemoglobina, calibrar y probar con controles. |  | CD9187D9004B65AD | No | Check SP and pneumatic circuit, unblock hemoglobin measurement bucket, calibrate and test with controls. |  | Cleaned | Fluidic pathway | Oxi | -124
ABL815 FLEX |  |  | CD916797005E3783 | No |  |  | Cleaned | El/Met Module | Measuring chamber | Lac
ABL800Basic |  | Reinigung und Neumembranierung | CD919A9B004431FA | No |  | Cleaning and re-membraneing | Defect | Electrode | Membrane | Lac
ABL825 FLEX | Reviewed WDC report, and investigated QC and calibration logs. Replaced PH, PO2, and PCO2 sealing ring - No fix. Cleaned and inspected measuring chamber for damage - no damage evident. Replaced membrane from a different lot - No fix. | Replaced PCO2 electrode calibrated and QC was right on the peer 2 mean. Customer will continue to monitor. | CD917BC30001D7C9 | No | Reviewed WDC report, and investigated QC and calibration logs. Replaced PH, PO2, and PCO2 sealing ring - No fix. Cleaned and inspected measuring chamber for damage - no damage evident. Replaced membrane from a different lot - No fix. | Replaced PCO2 electrode calibrated and QC was right on the peer 2 mean. Customer will continue to monitor. | Defect | Electrode | Electrode | pCO2
ABL835 FLEX | Leaking from inlet gasket and from inlet probe tubing. Back pressure is occurring. I replaced Inlet Gasket with new gsaket but no help. I then replaced blue inlet tubing and waste pump tubings (to correct back pressure) No help. I replaced inlet gasket with a second gasket which corrected leaks and rinse errors. | Replaced waste pump tubings, inlet tubing and two inlet gaskets. Had to open a PM Kit for some of these parts | CD91891F00464D3F | No | Leaking from inlet gasket and from inlet probe tubing. Back pressure is occurring. I replaced Inlet Gasket with new gsaket but no help. I then replaced blue inlet tubing and waste pump tubings (to correct back pressure) No help. I replaced inlet gasket with a second gasket which corrected leaks and rinse errors. | Replaced waste pump tubings, inlet tubing and two inlet gaskets. Had to open a PM Kit for some of these parts | Defect | Inlet module | Waste tubing
ABL825 FLEX | VERIFIED BLOOD COMPARISONS | REALIGNED THE LAMP,THB CAL AND OPTICAL FIBER | CD92C654763F43B9 | No | VERIFIED BLOOD COMPARISONS | REALIGNED THE LAMP,THB CAL AND OPTICAL FIBER | Adjustment | HW | Oxi Module | Lamp Unit
ABL815 FLEX | siehe SLR | siehe SLR | CD9178CC000E41AC | Yes | see SLR | see SLR | Re classified RMED | SWU
AQT90 FLEX |  |  | CD91B27E0028F8A0 | No
AQT90 FLEX |  |  | CD917D44001B68A5 | No
ABL80 CO-OX | Flushed out cassette with water. | Air detected disapeared, but PCO2 on solution 2 reads too high. Took 4 initialzations for PCO2 to come in. | CD92F2D3762689E3 | No | Flushed out cassette with water. | Air detected disapeared, but PCO2 on solution 2 reads too high. Took 4 initialzations for PCO2 to come in.
ABL80 CO-OX | Radiance support suggested replacing cross over cable. That didn't work. Troubleshot with Radiance support via phone | Changed network setting in analyzer to match server. | CD92F8E57637CEFA | No | Radiance support suggested replacing cross over cable. That didn't work. Troubleshot with Radiance support via phone | Changed network setting in analyzer to match server.
ABL90 |  |  | CD92EC7B765378B5 | No |  |  | Cleaned | Fluidic pathway | Clot |  | & replaced gasket (Old)
ABL90 | siehe SLR | siehe SLR | CD930963763A270C | No | see SLR | see SLR | Defect | HW | waste tube |  | x
ABL90 |  |  | CD9670F2D835BADE | No |  |  | Unknown | Unclear UpSWO description |  |  | x
ABL90 |  |  | CD97BAA14E390693 | No |  |  | Cleaned | Fluidic pathway | Oxi
ABL90 |  |  | CD97DC524E28069F | No |  |  | Cleaned | Fluidic pathway | Oxi
ABL90 |  |  | CD970ED2D85F854F | No |  |  | Defect | HW | Spring for Waste ventil |  | x
ABL827 FLEX | i touched the screen to see if the screen was not working properly and some of the touch screen keyboard keys are intermittent. | I replaced the display | CD92BF2876639A45 | No | i touched the screen to see if the screen was not working properly and some of the touch screen keyboard keys are intermittent. | I replaced the display | Defect | DMS PC | Display
ABL815 FLEX |  | Lac Meßkammer gereinigt, Lac Elektrode überprüft, Kal, QC | CD92DFDB76511BA9 | No |  | Lac measuring chamber cleaned, Lac electrode checked, Kal, QC | Cleaned | El/Met Module | Measuring chamber
ABL825 FLEX |  |  | CD92B8227651C0BA | No |  |  | Defect | pH/BG Module | Valve, air | Membrane
ABL825 FLEX | Multiple inlet errors. | Inlet module or replacement analyzer not available. Removed analyzer prepared for return to Westlake office. | CD95851D62285BEE | No | Multiple inlet errors. | Inlet module or replacement analyzer not available. Removed analyzer prepared for return to Westlake office. | Defect | Inlet module | Unknown
TCM4 | Battery voltage will not rise above 3VDC. | Replaced battery. Calibrated & tested battery function. Monitor operates normally. | CD92F385763AAA54 | No | Battery voltage will not rise above 3VDC. | Replaced battery. Calibrated & tested battery function. Monitor operates normally. | Defect | HW | Battery module
TCM/Tosca |  |  | CD92BB727629F42E | No |  |  | Defect | Consumables | Electrode
TCM-COMBIM | Discussed issue with client. TCM4 indicating no gas with new bottle. Mera WO # 303130 SWO written up by FSE P. Franklin | Replaced CombiM module. Gas errror is gone. Calibration started and failed. System error is now membrane replacement needed. Reset software on membrane and attempted calibration again. Calibration still failing. Swapped Electrode with another one from clients stock and calibration is passing after two attempts. New electrode will need to be sent to client. System operational at this time. Start time 11:45 Finish time is 2:00pm Ckd I O with Jerry Kelly Part need to go back fedex trk # 770818446388 | CDA434C3EA5F368E | No | Discussed issue with client. TCM4 indicating no gas with new bottle. Mera WO # 303130 SWO written up by FSE P. Franklin | Replaced CombiM module. Gas errror is gone. Calibration started and failed. System error is now membrane replacement needed. Reset software on membrane and attempted calibration again. Calibration still failing. Swapped Electrode with another one from clients stock and calibration is passing after two attempts. New electrode will need to be sent to client. System operational at this time. Start time 11:45 Finish time is 2:00pm Ckd I O with Jerry Kelly Part need to go back fedex trk # 770818446388 | Defect | HW | Combi M Module | Gas error
ABL80 CO-OX |  |  | CD95C3E762020C55 | No
ABL80 CO-OX | Found fiberous material in tubings. Replaced all tubings. All good. Verfied liquid sensors by running program in service. All good. Also used SC during trouble shooting. | Replaced tubings. | CD96AC10D836618C | No | Found fiberous material in tubings. Replaced all tubings. All good. Verfied liquid sensors by running program in service. All good. Also used SC during trouble shooting. | Replaced tubings.
ABL90 |  |  | CD96C41CD8065B3D | No |  |  | Defect | HW | Touch screen |  | x
ABL90 |  |  | CD9AA4843A00B375 | No |  |  | No fault found | Unable to install SP | Phone support did not help | FSE found no problem to install SP | x
TCM4 |  |  | CD96A26CD821B4AC | No |  |  | Defect | HW | Noise on signal
ABL815 FLEX |  |  | CD96B108D813F2ED | No |  |  | Defect | HW | Oxi Module | PCB
ABL825 FLEX | checked logs . | replaced gas1 and gas2 regulators | CD96B47CD829292F | No | checked logs . | replaced gas1 and gas2 regulators | Defect | Basic unit | Gas Module | Gas1 and Gas2
ABL800Basic | Fehler bestnd beriets am 14.05.14 und wurde da behoben | im Fulidicmodul V6,7,8 undicht->ersetzt | CD9675EAD83689E2 | No | Bug bestst advised on 14.05.14 and has been fixed since | in the Fulidic module V6,7,8 leak->replaced | Defect | Fluidic Module | V8-V7-V6
ABL825 FLEX | Checked in service, Oxi liquid sensor failed | Replaced inlet probe and tube Flushed air vent line with DI water Test liquid sensors in service-OK | CD9681D4D8460C1D | No | Checked in service; Oxi liquid sensor failed | Replaced inlet probe and tube Flushed air vent line with DI water Test liquid sensors in service-OK | Defect | Inlet module | Probe
ABL800Basic | Lac Empfindlichkeit außer Bereich | Reinigung Kontakte gewartet, Meßkammer gereinigt Lac membraniert | CD96A1EAD8344815 | No | Lac Sensitivity Out Of Range | Cleaning contacts maintained, measuring chamber cleaned Lac membraneed | Defect | Electrode | Membrane | Lac
ABL820 FLEX | checked data logs for errors, checked flow of solutions, checked valves and liquid sensors, | removed and cleaned dual housing valves, removed electrodes and cleaned chambers, replaced all electrode membranes, | CD970F40D8607882 | No | checked data logs for errors, checked flow of solutions, checked valves and liquid sensors, | removed and cleaned dual housing valves, removed electrodes and cleaned chambers, replaced all electrode membranes, | Defect | Electrode | Membrane | pCO2
AQT90 FLEX |  |  | CD968CD6D80B6BC3 | No
AQT90 FLEX |  |  | CD96A7BAD853F902 | No
ABL90 | On site to invesigate. While driving to site the customer told me that they received their order of July 28. I asked him to replace SP which corrected the error 1169. Customer called back to say they are now failing QC all levels for pCO2. (Immediately following init) WHen I arrived i found that connector gasket replacement is past due but do not have a connector gasket. I ran another Cal and then all QC passed. | Customer corrected fluidics problem with SP replacement. I will order a connector gasket. I also left an inlet probe and holder with customer just in case. I also installed inlet clip. | CD98D0BCC437EDBB | No | On site to invesigate. While driving to site the customer told me that they received their order of July 28. I asked him to replace SP which corrected the error 1169. Customer called back to say they are now failing QC all levels for pCO2. (Immediately following init) WHen I arrived i found that connector gasket replacement is past due but do not have a connector gasket. I ran another Cal and then all QC passed. | Customer corrected fluidics problem with SP replacement. I will order a connector gasket. I also left an inlet probe and holder with customer just in case. I also installed inlet clip. | Defect | SP | Fluid transport
ABL90 | Analyzer asking to replaced SP repeatedly. 3 SP used by customer . No success. Back flushed the analyer multiple times. Checked inlet probe and holder. Checked connector gasket.Checked data logs. Checked waste tubing. Checked oxi valve tubing. Performed all service checks. Consulted with Tech support @ RAME. They suggested replacing the inlet probe. Replacing probe did not help entirely. Opened the analyzer to check for leaks on oxi module and waste. Leakage found around hemolyzer and found traces of blood and QC deposits. SC gave response error on K and Na. | Replaced SP out of customer stock. Replaced inlet probe. Replaced Inlet gasket with holder. Cleaned connector gasket. Replaced Hemolyzer. Replaced SC out of customer stock. Let analyzer Run overnight just to make sure everything is running fine. Ran waste blood samples. | CD9812324E58AA55 | No | Analyzer asking to replaced SP repeatedly. 3 SP used by customer . No success. Back flushed the analyer multiple times. Checked inlet probe and holder. Checked connector gasket. Checked data logs. Checked waste tubing. Checked oxi valve tubing. Performed all service checks. Consulted with Tech support @ RAME. They suggested replacing the inlet probe. Replacing probe did not help entirely. Opened the analyzer to check for leaks on oxi module and waste. Leakage found around hemolyzer and found traces of blood and QC deposits. SC gave response error on K and Na. | Replaced SP out of customer stock. Replaced inlet probe. Replaced Inlet gasket with holder. Cleaned connector gasket. Replaced Hemolyzer. Replaced SC out of customer stock. Let analyzer Run overnight just to make sure everything is running fine. Ran waste blood samples. | Unknown | Revisit required
TCM4 | Após erro de calibração o cliente substituiu a membrana e voltou a calibrar o equipamento, este calibrou 2 a 3 vezes correctamente. Depois de desligar e voltar a ligar o equipamento este não voltou a dar erro de calibração. Substituí a membrana e fiz vários testes sempre com calibração ok. |  | CD97FABC4E211397 | No | After the calibration error the client replaced the membrane and recalibrated the equipment, this calibrated 2 to 3 times correctly. After turning the equipment off and on again, it did not give the calibration error again. I replaced the membrane and did several tests always with calibration OK. |  | Defect | Consumables | Electrode | Calibration
ABL90 | Comprobar y limpiar el circuito neumático, desmontar y probar motor de bomba, sensores y mecanismos, cambiar tubos, probar con hasta 4 SP del lote LW-13, sin éxito. Probar el equipo con otro sp de lote LV, el problema queda resuelt,. el equipo calibra y queda operativo. | Poner un SP de lote LV-. | CD9823C74E1ECFA3 | No | Check and clean the pneumatic circuit, disassemble and test pump engine, sensors and mechanisms, change tubes, test with up to 4 SP of the LW-13 lot, without success. Testing the equipment with another LV batch sp, the problem remains resolved,. the equipment calibrates and is operational. | Put a batch SP LV-. | Defect | SP | Pump cal error | LW-13
ABL825 FLEX |  |  | CD97EEF04E499FFC | No |  |  | Cleaned | Auto Check Module | Barcode Reader
ABL800Basic | Fehlermeldung 356 Temp. Fehler im EL-Modul tritt seit ca 1Woche gehäuft auf, nach Neustart zeitweise behoben | Austausch EL-Module PCB Neustart Einlaufverfahren Reinigen Lüftungsfilter | CD97B9C94E58481C | No | Error message 356 Temp. Errors in the EL module have been accumulating for about 1week; Fixed intermittently after reboot | Replacement EL-Module PCB Restart Start-up Procedure Cleaning Ventilation Filter | Defect | El/Met Module | PCB
AQT90 FLEX |  |  | CD980C954E22C92A | No
ABL80 OSM |  |  | CD993A85C443400C | No
ABL90 | Kal.- QC- Aktivitätendateien | Spülungen durchgeführt, Schlauchwege erneuert, Dichtungen neu, LP und SC neu | CD9919A1C42104F0 | No | Kal.- QC- Aktivitätendateien | Rinses performed, hose ways renewed, seals new, LP and SC new | Defect | HW | Oxi Module | Pinch Valve tube | & Cleaned fluidic + installed new inlet + SC/SP
ABL90 |  |  | CD98F972C4217806 | No |  |  | User Error | Incorrect SP activation | old style inlet gasket dismantled and incorrectly assembled |  | x
ABL90 | Analyzer is not used for patient results. Validation study is currently being conducted. Error 1350, Clot suspected in inlet. Error occurred on 8-05-14. Clot Detection Inlet Value failed,value out high. | Back flush ABL90. Performed Fluidic check in Diagnostics. Replaced Inlet Probe. | CD98F4FDC45EDBFD | No | Analyzer is not used for patient results. Validation study is currently being conducted. Error 1350, Clot suspected in inlet. Error occurred on 8-05-14. Clot Detection Inlet Value failed,value out high. | Back flush ABL90. Performed Fluidic check in Diagnostics. Replaced Inlet Probe. | Replaced | Probe | Clot Probe
ABL800Basic | Fluidic channel unit undicht |  | CD991116C4170DB4 | No | Fluidic channel unit undicht |  | Defect | Fluidic Module | V8
ABL835 FLEX | DATA LOG SHOWS EL MET TEMP ERROR 356.SERVICE PROGRAM FOR THERMOSTATTING SHOWS 17 DEGREES..EL MET MODULE IS ILLUMINATED AND NO SIGNS OF LEAKAGE FROM MEASURING CHAMBER. | AM ESCALATING THE CALL DUE TO 3RD REPLACEMENT OF EL MET MODULE ILLUMINATION BOARD.ON THE OTHER TWO REPLACEMENTS OF THE EL MET ILLUMINATION BOARD THE PROBLEM WAS RESOLVED FOR 3 OR 4 WEEKS.THIS TIME I CHANGED THE EL MET MEASURING CHAMBER AND HEAT DISTRIBUTION PLATE AS WELL AS THE EL MET MODULE CONTROL BOARD AND ILLUMINATION BOARD | CD9912D6C45E9F8A | No | DATA LOG SHOWS EL MET TEMP ERROR 356.SERVICE PROGRAM FOR THERMOSTATTING SHOWS 17 DEGREES.. EL MET MODULE IS ILLUMINATED AND NO SIGNS OF LEAKAGE FROM MEASURING CHAMBER. | AM ESCALATING THE CALL DUE TO 3RD REPLACEMENT OF EL MET MODULE ILLUMINATION BOARD. ON THE OTHER TWO REPLACEMENTS OF THE EL MET ILLUMINATION BOARD THE PROBLEM WAS RESOLVED FOR 3 OR 4 WEEKS. THIS TIME I CHANGED THE EL MET MEASURING CHAMBER AND HEAT DISTRIBUTION PLATE AS WELL AS THE EL MET MODULE CONTROL BOARD AND ILLUMINATION BOARD | Defect | El/Met Module | PCB
ABL800Basic |  | Austausch Lac -Elektrode | CD993A4EC404483D | No |  | Replacement Lac electrode | Defect | Electrode | Electrode | Lac
ABL835 FLEX |  |  | CDA2FA7474320B26 | No |  |  | Adjustment | HW | Oxi Module | Lamp Unit
ABL820 FLEX | Upon returning discovered pH started fail 2pt. cal during evening. Unstable Response, Drift, Sensitivity,and Status errors were present. | Cleaned Decontaminated Chambers Cleaned Electrode and Connectors contacts. | CD9BD990B0175643 | No | Upon returning discovered pH started fail 2pt. cal during evening. Unstable Response, Drift, Sensitivity,and Status errors were present. | Cleaned Decontaminated Chambers Cleaned Electrode and Connectors contacts. | Cleaned | Fluidic pathway | -
AQT90 FLEX |  |  | CD991AFDC40B65B8 | No
AQT90 FLEX | Se revisa equipo a petición del cliente. Se realizó mantenimiento preventivo según protocolo del fabricante |  | CD9A571C3A57E9F2 | Yes | Equipment is reviewed at the customer's request. Preventive maintenance was performed according to manufacturer's protocol |  | Re classified RMED |  | Alejandro Gonzales (214-8-25) 838R33N2 Megalab - Las Américas 14-8-214. It is a preventive revision ask for the customer.
ABL80 CO-OX | When customer gets these errors, they either reboot analyzer or replace SC + SP and issue goes away for a day or two and comes back. | Take analyzer apart and flush out fittings on front panel, no blockages. Flush out hemolyzer, no blockages. Replace liquid sensor. Could see red material where liquid sensor port meets hemolyzer. Flush hemolyzer again with bleach till red went away. Also put in a PM kit to be safe. Perform tHb Calibration. | CD9A4D233A0A99F7 | No | When customer gets these errors, they either reboot analyzer or replace SC + SP and issue goes away for a day or two and comes back. | Take analyzer apart and flush out fittings on front panel, no blockages. Flush out hemolyzer, no blockages. Replace liquid sensor. Could see red material where liquid sensor port meets hemolyzer. Flush hemolyzer again with bleach till red went away. Also put in a PM kit to be safe. Perform tHb Calibration.
ABL80 CO-OX | Error 581 spectrum error on blood samples. | Replaced the hemolyzer Installed a new cassette. | CD9AC57A3A24697D | No | Error 581 spectrum error on blood samples. | Replaced the hemolyzer Installed a new cassette.
ABL90 | Errors 1343 coming up again. SP replacement repeatedly asked by analyzer. | Flushed the system multiple times. Flushed the Sensor cassette bay unit. Checked probe and gasket holder. | CD9BAC41B056550E | No | Errors 1343 coming up again. SP replacement repeatedly asked by analyzer. | Flushed the system multiple times. Flushed the Sensor cassette bay unit. Checked probe and gasket holder. | Cleaned | Fluidic pathway | Techline: RAME 69323344
ABL800Basic |  |  | CD9A96803A287F07 | No |  |  | Cleaned | Auto Check Module | Barcode Reader
ABL825 FLEX | SE CAMBIA ELECTRODO DE REF Y SE COMPRUEBA FUNCIONAMIENTO CONN TEST DE SERVICIO, CAL Y QC | SE CAMBIA ELECTRODO DE REF Y SE COMPRUEBA FUNCIONAMIENTO CONN TEST DE SERVICIO, CAL Y QC | CD9A510D3A55C856 | No | REF ELECTRODE IS CHANGED AND OPERATION IS CHECKED WITH SERVICE TEST, CAL AND QC | REF ELECTRODE IS CHANGED AND OPERATION IS CHECKED WITH SERVICE TEST, CAL AND QC | Defect | Electrode | Electrode | Reference
ABL825 FLEX |  | Drucker erneuert | CD9A7C4C3A34F66B | No |  | Printer renewed | Defect | Printer module | Printer head
ABL835 FLEX | Erreur 793 asymetrie BG v chambre de mesure OK | flush V&, V8 et remplacé tube de pompe BG | CD9A446E3A09FFE3 | No | Error 793 asymetrie BG v OK measuring chamber | flush V, V8 and replaced PUMP tube BG | Defect | pH/BG Module | Pump tube
ABL817 FLEX |  |  | CD9A8BDC3A55E671 | No |  |  | Adjustment | HW | Oxi Module | tHb cal
ABL835 FLEX |  |  | CDA30DF7743F0B3F | Yes |  |  | Re classified RMED | HIS/LIS
AQT90 FLEX |  |  | CD9A72133A23B08D | No
ABL90 |  |  | CD9F48B6120501FB | No |  |  | Unknown | Parts ordered - but unknown which one | Hemolyzer or tubes or SP |  | x
ABL90 | Comprobar juntas de entrada de muestras, cambiar la junta de entrada, calibrar y probar con muestras y controles. | Cambiar la junta de entrada. | CD9BE7C4B035CFBC | No | Check sample input joints, change the input joint, calibrate and test with samples and controls. | Change the input joint.
ABL80 FLEX Basic version |  |  | CDA094058840E4F4 | No
ABL80 FLEX Basic version | Revisión de la cassette en cuestión (limpieza con agua y contactos con alcohol). Se revisa la parte electrónica y chequean los diferentes programas de servicio. Se sustituye la cassette por otra nueva, la cual funciona correctamente. Calibración y pasar los QC´s. | Revisión de la cassette en cuestión (limpieza con agua y contactos con alcohol). Se revisa la parte electrónica y chequean los diferentes programas de servicio. Se sustituye la cassette en cuestión. | CD9F5BF81263657E | No | Review of the cassette in question (cleaning with water and contacts with alcohol). The electronic part is reviewed and the different service programs checked. The cassette is replaced with a new one, which works properly. Calibration and pass the QC's. | Review of the cassette in question (cleaning with water and contacts with alcohol). The electronic part is reviewed and the different service programs checked. The cassette in question is replaced.
ABL90 |  |  | CD9F89C2125B19F0 | No |  |  | Cleaned | Inlet gasket | New type gasket
ABL90 | SP geprüft, Einlassdichtung | SP korrekt aktiviert; Einlassdichtung erneuert | CD9F5B22122A6D1C | No | SP checked; Inlet seal | SP correctly activated; Inlet seal renewed | User Error | Incorrect SP activation | defect Inlet gasket NEW | LEH gasket returned | x
ABL90 |  |  | CDA0AAFB880F75AC | Yes |  |  | Re classified RMED | Telephone support | SP send
ABL800Basic | blauen Stutzen mit Aqua gespült, dazu mit Schlauch und Spritze Druck auf Stutzen gegeben, da sonst nichts gezogen wird, Kal, QC |  | CD9F775812159008 | No | blue nozzles rinsed with aqua, with hose and syringe pressure on nozzles, otherwise nothing is pulled, Kal, QC |  | Cleaned | Fluidic pathway | -
ABL835 FLEX | customer runs paatient comparrisons between there 3 analyzers, they have noticed that this unit reports lower values thatnthe other two units. The electrode has been repalced twice within the last 5 months | inspected the measuring chamber, glu electrode connector and found no issues. qc and cals appear to be normal but the sensitivity for Glu is lower thant the other two units. they are running the same membranes. i will contact my ASM and discuss the situation. i have done a service dump and exported all the data logs | CD9F69171236CDC2 | No | customer runs paatient comparrisons between there 3 analyzers, they have noticed that this unit reports lower values thatnthe other two units. The electrode has been repalced twice within the last 5 months | inspected the measuring chamber, glu electrode connector and found no issues. qc and cals appear to be normal but the sensitivity for Glu is lower thant the other two units. they are running the same membranes. i will contact my ASM and discuss the situation. i have done a service dump and exported all the data logs | Unknown | Comparison study
ABL825 FLEX |  |  | CD9F7F5D120FF5BC | No |  |  | Cleaned | Basic unit | Fan
ABL835 FLEX | I cleaned the Lac electrode well and verified Lac Zero adjustment (ne correction needed) I recalibrated and Lac sense improved from ~ 306 to ~400. I ran 2 samples of Cal Ver Level 1 and results were higher than last week but still in range. Range is 189 - 297 todays results are 270 and 277. Last weeks results were 260, 264, 262, 266. Well within range but apparently exceed the slope. | I changed the Lac membrane and will allow system to recove overnight before Tish re runs the Cal Ver again. | CD9FA625120EEDD7 | No | I cleaned the Lac electrode well and verified Lac Zero adjustment (ne correction needed) I recalibrated and Lac sense improved from ~ 306 to ~400. I ran 2 samples of Cal Ver Level 1 and results were higher than last week but still in range. Range is 189 - 297 todays results are 270 and 277. Last weeks results were 260, 264, 262, 266. Well within range but apparently exceed the slope. | I changed the Lac membrane and will allow system to recove overnight before Tish re runs the Cal Ver again. | Defect | Electrode | Membrane | Lac
ABL835 FLEX | During previous service call it was observed that there has been leaking from, or around, the pre heater. Pre heater is tarnished and shows signs of leak. | Ran all tests IAW Manuf. Spec. | CD9FAD561216EDDC | No | During previous service call it was observed that there has been leaking from, or around, the pre heater. Pre heater is tarnished and shows signs of leak. | Ran all tests IAW Manuf. Spec. | Defect | Inlet module | Preheater | Leaking
ABL800Basic |  |  | CDA09BA98835E7C8 | No |  |  | User Error | Electrode | Broken
ABL800Basic | Lac Empfindlichkeitsfehler Fehler tritt gehäuft auf, trotz mebran. und Elektrodenwechsel (siehe auch letzte SLR´s) | Reinigen Meßkammer, Kontakte Ventile am EL-Modul demontiert und gewartet Lac membraniert | CDA084B1885BD388 | No | Lac sensitivity error error occurs more frequently, despite mebran. and electrode changes (see also last SLR's) | Cleaning measuring chamber, contacts valves on the EL module disassembled and maintained lac membraneed | Defect | Electrode | Membrane | Lac
ABL815 FLEX | siehe ausgeführte Arbeiten | siehe ausgeführte Arbeiten | CDA0D9C4882959AC | No | see work carried out | see work carried out | Cleaned | Fluidic pathway | Measuring chamber | pH
ABL800Basic |  |  | CDA07842884131DC | No |  |  | Defect | Electrode | Membrane | Lac
ABL815 FLEX |  | Austausch 945-614 | CDA0AC458800FA2C | No |  | Replacement 945-614 | Defect | Electrode | Electrode | pH
ABL825 FLEX | Reference electrode electrolyte had migrated into measuring chamber and pH electrode seal had leaked. | Cleaned measuring chamber and re-seated pH O-ring. Calibrated and ran QC. Instrument operates normally. | CDA0C89D8815DAFF | No | Reference electrode electrolyte had migrated into measuring chamber and pH electrode seal had leaked. | Cleaned measuring chamber and re-seated pH O-ring. Calibrated and ran QC. Instrument operates normally. | Cleaned | pH/BG Module | Measuring chamber
ABL820 FLEX | 251 and 579 oximetry errors due to halogen lamp voltage high. Unit has also run out of cal1 solution. | Replaced cal 1 solution. Realigned halogen lamp and oximeter X2r. Rechecked for stability. | CDA097658859274C | No | 251 and 579 oximetry errors due to halogen lamp voltage high. Unit has also run out of cal1 solution. | Replaced cal 1 solution. Realigned halogen lamp and oximeter X2r. Rechecked for stability. | Adjustment | HW | Oxi Module | Lamp Unit | tHb
ABL800Basic |  | Austausch des DMS covers | CDA1EB3AFE14DA9B | No |  | Replacement of the DMS cover | User Error | DMS PC | Printer cover
TCM-COMBIM |  |  | CDA2180DFE10B555 | No |  |  | Defect | Consumables | Electrode | Regeneration
ABL80 OSM |  |  | CDA35C2074223A86 | No
ABL90 |  | Schlauch gekürzt | CDA328B57414627A | No |  | Hose shortened | Defect | HW | Waste tube |  | x
ABL90 |  |  | CDA3355F7418C34D | No
ABL827 FLEX | Upojn arrival verified Lactate errors. | Found Lactate, Glucose and PO2 electrodes ALL missing sealing O-rings. Also found probe bent. Re O-ringed all electrodes and replaced probe. Reviewed with customer need to make sure sample is completed before removing syringe to avoid probe bending. Also reviewed Electrode needs to havr O-ring in place when remembraning to avoid instablility on parameters. | CDA3196374416DC6 | No | Upojn arrival verified Lactate errors. | Found Lactate, Glucose and PO2 electrodes ALL missing sealing O-rings. Also found probe bent. Re O-ringed all electrodes and replaced probe. Reviewed with customer need to make sure sample is completed before removing syringe to avoid probe bending. Also reviewed Electrode needs to havr O-ring in place when remembraning to avoid instablility on parameters. | User Error | Electrode | Sealing ring
ABL825 FLEX | Tried membranes, different lot of membranes, cleaning electrode and cleaning contacts-No change | Replaced pCO2 electrode pCO2 electrode code LN03 defective, under warranty X2 Both electrodes showed the same problem This is not an instrument problem, it is an electrode problem and I need 2 electrodes to replace the defective ones | CDA309B47448D8AE | No | Tried membranes, different lot of membranes, cleaning electrode and cleaning contacts-No change | Replaced pCO2 electrode pCO2 electrode code LN03 defective, under warranty X2 Both electrodes showed the same problem This is not an instrument problem, it is an electrode problem and I need 2 electrodes to replace the defective ones | Defect | Electrode | Electrode | pCO2
ABL90 |  | Ventilschlauch erneuert, war undicht. | CDA45FADEA146286 | No |  | valve hose was leaking. | Defect | HW | Oxi Module | Pinch Valve tube | x
ABL90 |  |  | CDA4730EEA0FB748 | Yes |  |  | Re classified RMED | Logistic PER
ABL90 |  |  | CDA451C3EA506360 | No |  |  | Unknown | Black screen | Power off/on |  | x
ABL90 |  |  | CDA44D70EA42635A | No |  |  | Replaced | Inlet gasket | Old or new ?? Unknown | LEH gasket returned | x
ABL90 |  | Anwender über Aktivierung des Lösungspacks eingewiesen, aktuelle Einlassdichtung und Nadel eingesetzt, Abfallweg gespült, graue Dichtung ersetzt, Kal, QC | CDA4A0E0EA2410FC | No |  | User instructed via activation of the solution pack, current inlet seal and needle inserted, waste path rinsed, grey seal replaced, Kal, QC
ABL825 FLEX |  |  | CDA455D1EA5C9B77 | No |  |  | Defect | HW | Oxi Module | PCB | # 58
ABL820 FLEX | Replaced electrode connector - No fix. | Replaced PCO2 electrode. Tested OK. | CDA48AEDEA2EAEB4 | No | Replaced electrode connector - No fix. | Replaced PCO2 electrode. Tested OK. | Defect | Electrode | Electrode | pCO2
ABL90 |  |  | CDA78047D621C35D | No |  |  | Defect | SP
ABL815 FLEX |  |  | CDA72146D627F25C | No |  |  | Defect | HW | Oxi Module | PCB
ABL80 | CASSETTE LEEKING AND WAS REPLACED | RERAN SAMPLES | CDA8329E4C56E03F | No | CASSETTE LEEKING AND WAS REPLACED | RERAN SAMPLES
ABL90 |  |  | CDA962D0C2264B59 | No |  |  | Defect | HW | Waste tube
ABL90 |  |  | CDA84ED54C12C5F8 | No |  |  | User Error | Insufficent SP activation
ABL90 | SO2 and cohb QC errors | Ran Oxi test and po2 test. Checked activity logs. | CDB1673786049D0A | No | SO2 and cohb QC errors | Ran Oxygen test and PO2 test. Checked activity logs. | Unknown | Oxi Module | Oxi test | Techline: RAME 69323344
ABL90 | Fließwege SC | Spülung der Fließwege und der SC(verstopft) | CDA8619F4C45018A | No | Fließwege SC | Flushing of the flow paths and the SC (clogged) | Cleaned | Fluidic pathway | SC
ABL90 |  |  | CDA82BBD4C5D9A41 | No |  |  | Cleaned | Fluidic pathway
ABL90 | Sensor cassette needs changing, error 971. Customer does not have another sensor cassette. | Swapped the sensor cassette with a good sensor cassette from anothe abl 90 and the unit came up fine. Ordering a new sensor cassette. | CDA825E44C341474 | No | Sensor cassette needs changing, error 971. Customer does not have another sensor cassette. | Swapped the sensor cassette with a good sensor cassette from anothe abl 90 and the unit came up fine. Ordering a new sensor cassette. | Defect | SC | Expired | Fse because no SC available
ABL800Basic |  |  | CDA8142C4C43A9C3 | No |  |  | Defect | Electrode | Electrode | Reference
ABL800Basic |  |  | CDA8601F4C36D19C | Yes |  |  | Re classified RMED | Preventive maintenance
ABL827 FLEX | Inlet nipple had broken at measuring chamber. | Replaced measuring chamber. During startup instrument had A D to convertor error. OXI module PCB has failed. Ordering part for replacement 8-27. | CDA83D4B4C0ADE74 | No | Inlet nipple had broken at measuring chamber. | Replaced measuring chamber. During startup instrument had A D to convertor error. OXI module PCB has failed. Ordering part for replacement 8-27. | Defect | pH/BG Module | Measuring chamber | Nipple broken
ABL825 FLEX | Tried changing membranes, cleaning contacts and measuring chamber-Problem comes back | Replaced pCO2 electrode pCO2 electrode lot LN-05 defective and in warranty | CDA82ED74C093DE4 | No | Tried changing membranes, cleaning contacts and measuring chamber-Problem comes back | Replaced pCO2 electrode pCO2 electrode lot LN-05 defective and in warranty | Defect | Electrode | Electrode | pCO2
ABL90 |  |  | CDA9A32AC2336A23 | No |  |  | Defect | SC | Temperature
ABL90 | Gerät und Spülung der Fließwege | Austausch LP aus Kundenbestand | CDA9B064C22B745B | No | Device and flushing of the flow paths | Replacement LP from customer base | Defect | SP | -
ABL90 |  |  | CDA967C6C206CE87 | No |  |  | Cleaned | Fluidic pathway | If continued we need Log files | Bacteria??
ABL90 | Error 1343, Error 1169. Customer replaced SP. Initialization wont complete. Checked Datalogs and activity logs. Flushed the analyzer multiple times. | Flushed the SC bay. Re inserted SC and SP. Cleaned inlet gasket and flushed probe. Ran Service Checks and fluid transport tests. | CDA9FF74C2577522 | Yes | Error 1343, Error 1169. Customer replaced SP. Initialization wont complete. Checked Datalogs and activity logs. Flushed the analyzer multiple times. | Flushed the SC bay. Re inserted SC and SP. Cleaned inlet gasket and flushed probe. Ran Service Checks and fluid transport tests. | Re classified RMED | Cleaned Fluidic pathway | Err 1343 and 1169 | From: Huyben, Darcy DHUYB Sent: Friday, August 29, 214 1:51 PM The call in Lethbridge could have potentially been taken care of over the phone. The FSE was nearby doing a TC installation and went onsite as the customer is new to the ABL9 line and had to troubleshoot interface issues as well. This SWO can be changed to a courtesy call as the main fix was to flush the analyzer.
ABL90 |  |  | CDA961D6C2404B53 | No |  |  | Defect | SP | Cal3
ABL90 | data logs show all qc level 3s failed on lactate | Replaced sensor cassette. | CDA9CBEFC2008380 | No | data logs show all qc level 3s failed on lactate | Replaced sensor cassette. | Defect | SC | Lac | QC 3
ABL90 | The ABL90 says to change the sensor cassette. Rebooted the ABL90 and that did not solve the problem. | Replaced the sensor cassette. Ran several cals trying to get LAC to cal. The Sensitivity was on going up by 1 or 2 per cal cycle. Left letting the ABL 90 run over night to see if let the LAC settle and pass on its on. | CDA9D435C22EE0AC | Yes | The ABL90 says to change the sensor cassette. Rebooted the ABL90 and that did not solve the problem. | Replaced the sensor cassette. Ran several cals trying to get LAC to cal. The Sensitivity was on going up by 1 or 2 per cal cycle. Left letting the ABL 90 run over night to see if let the LAC settle and pass on its on. | Re classified RMED | Courtesy
ABL827 FLEX |  |  | CDA98174C208E896 | No |  |  | Defect | HW | Oxi Module | PCB
ABL835 FLEX | Leak detected. | Installed new Tubing Nipple kit. | CDA97920C2567EED | No | Leak detected. | Installed new Tubing Nipple kit.
AQT90 FLEX |  |  | CDA9839AC2469481 | No
AQT90 FLEX |  |  | CDA9A408C22EAAFA | No
TCM/Tosca |  |  | CDA9451DC2407D5D | No |  |  | Defect | HW | PC module | Sound corrupted
ABL90 |  |  | CDAAA3E238610FE7 | Yes |  |  | Re classified RMED | TU 116
ABL90 |  |  | CDABB1BDAE61AC8E | No |  |  | User Error | SP activate
ABL90 |  | Completed credit form for 2nd sensor cassette that had been replaced. | CDAAB4B138571135 | Yes |  | Completed credit form for 2nd sensor cassette that had been replaced. | Re classified RMED | CCN
ABL827 FLEX | OXI module main PCB has failed causing multiple OXI errors. | Replaced OXI PCB per Radiometer specifications. Calibrated and ran QC. Instrument operates normally. | CDAAECA0384CE5D1 | No | OXYGEN Module main PCB has failed causing multiple OXYGEN errors. | Replaced OXI PCB per Radiometer specifications. Calibrated and ran QC. Instrument operates normally.
ABL825 FLEX | Value in Lamp Adjustment was too high - 3.641. | Perform complete tHb alignment. | CDAA90AE385BD1EE | No | Value in Lamp Adjustment was too high - 3.641. | Perform complete tHb alignment. | Adjustment | HW | Oxi Module | Lamp Unit
ABL827 FLEX |  |  | CDBC40D2AC0913C0 | No
ABL827 FLEX |  |  | CDAAD77B3857119A | No
ABL825 FLEX | Autocheck error may be caused by damaged house cable, LS and sample errors are caused by inlet gasket being damaged internally. | Replaced Inlet gasket and autocheck house cable. Cleaned air valves and sample pathway. Tested AC operation and sample aspiration. Calibrated and ran QC. Instrument operates normally. | CDAB3A36384BBBD8 | No | Autocheck error may be caused by damaged house cable, LS and sample errors are caused by inlet gasket being damaged internally. | Replaced Inlet gasket and autocheck house cable. Cleaned air valves and sample pathway. Tested AC operation and sample aspiration. Calibrated and ran QC. Instrument operates normally.
ABL90 |  |  | CDAC38CCAE29AFB9 | No |  |  | Defect | SP | Cal3
ABL90 |  |  | CDAC0579AE20608C | Yes |  |  | Re classified RMED | Telephone support
ABL90 |  |  | CDAC00CBAE3175D3 | No |  |  | Cleaned | Fluidic pathway
ABL800 BASIC | Cl elektróda újramembránozása. Kalibrálás, QC mérés. | András Hegedüs | SWO-990 | No | Cl electrode regrate. Calibration, QC measurement. | Madhu A
ABL90 |  |  | CDAD2E482436D60C | No |  |  | Defect | SP | Leak current
ABL90 | siehe SLR | siehe SLR | CDAD3BAE241445A7 | No | see SLR | see SLR | Defect | SP | Inhomogenous sample
ABL5 |  | Justage | CDAD2CB1241D4B62 | No |  | Adjustment
ABL835 FLEX | Se realiza desproteinización de la cámara de medida y lumpieza de sulfitos. Comprobación de funcionamiento mediante QC |  | CDCCCB9AAA3DCEA1 | No | Deproteinization of the measuring chamber and lumpieza of sulphites. Operational check via QC
ABL835 FLEX |  |  | CDB175A88639EB4F | No
ABL800Basic | ABL bricht Kalib. ab | Kal.lösung 2 erneuert K membraniert u. Meßkammer gereinigt | CDAD47BF2403E4A3 | No | ABL breaks Kalib. Off | Cal.solution 2 renews K membraneized and measuring chamber cleaned | Cleaned | El/Met Module | Measuring chamber | K+
ABL800Basic |  |  | CDAD544A241966B1 | No
AQT90 FLEX |  |  | CDB388E4723B91AB | No
TCM/Tosca |  |  | CDAD1A65240A72E0 | No |  |  | No fault found
ABL90 |  |  | CDB0EE4E862213A3 | No |  |  | Defect | SP
ABL800Basic |  |  | CDB27981FC30B38B | No
AQT90 FLEX |  |  | CDB1470186299548 | No
ABL80 CO-OX |  |  | CDB22C2DFC48ADD5 | No
ABL90 |  |  | CDB26589FC05272F | No |  |  | Defect | HW | Waste tube
ABL90 | clot detected flush ABL replaced tubing not able to removed clot stuck on the left side of bay unit | not able to fix instrument instrument will be send out to london office to be repair | CDB24EE6FC4ACCA6 | No | clot detected flush ABL replaced tubing not able to removed clot stuck on the left side of bay unit | not able to fix instrument instrument will be send out to london office to be repair | Unknown | Clot issue
ABL90 | Pump calibration error | replaced all tubinsg inlet gasket ,found hemolyser leaking a bit replaced hemolyser fluh analyser | CDB250A7FC14CCA9 | No | Pump calibration error | replaced all tubinsg inlet gasket ,found hemolyser leaking a bit replaced hemolyser fluh analyser | Defect | HW | Hemolyzer
ABL815 FLEX |  |  | CDB22C32FC35C4E5 | No
ABL815 FLEX |  |  | CDB26ACEFC62C5ED | No
ABL825 FLEX |  |  | CDB24189FC2320CC | No
ABL825 FLEX |  |  | CDB24ADDFC0F08F0 | No
ABL825 FLEX |  |  | CDB24821FC03E815 | No
ABL835 FLEX | Analzyer leaking constantly. | Investigated all tubing. Found tubing that runs up to pH BG module (20A) was an inch too long. Cut tubing to proper length, installed and checked that there were no more leaks. Completed servcie tests. | CDB26712FC0BF1DD | No | Analzyer leaking constantly. | Investigated all tubing. Found tubing that runs up to pH BG module (20A) was an inch too long. Cut tubing to proper length, installed and checked that there were no more leaks. Completed servcie tests.
AQT90 FLEX | Se sustituyen varios componentes del analizador como bombas tubos de conexiones, unidades de lavado, unidad de process wheel, unidad optica, filtros, etc... chequeo general de señales del equipo, limpiezas de los diferentes modulos, calibraciones e introdución de muestras, funcionando correctamente. | Sustitución de la unidad optica. | CDB23D01FC002908 | No | Several analyzer components are replaced such as connection tube pumps, washing units, process wheel unit, optical unit, filters, etc... General equipment signal check, cleaning of different modules, calibrations and introduction of samples, working properly. | Replacing the optical unit.
ABL90 FLEX | Jeg har tager et service dump fra ABL90 så vores support i Radiometer kan se hvad der er gået galt. Jeg har også tager datalogs fra ABL90 Jeg har taget den defekte væskepakke med hjem. Nanna har sat en helt ny SP på ABL90 og nu er der ca. 600 prøver at gøre godt med. | Rasmus Pedersen | SWO-1028 | No | I have take a service dump from ABL90 so our support in Radiometer can see what has gone wrong. I also take data logs from ABL90 I have taken the defective liquid package home. Nanna has put a brand new SP on ABL90 and now there are about 600 tries to do well. | Rasmus Pedersen
ABL90 | CHECK LOGS. SP HAD BEEN CHANGED 2X. IN SVC, PUMP CAL, CLOT DETECT, LS FAILED REPEATEDLY. PUMP CAL FINALLY PASSED AT 1.19. BY AM PUMP FAILED | INSTALLED NEW SP. ALL TESTS PASSED. SC FAILED PCO2 & LAC, RUNNING OVER NIGHT. NEW SC ACCEPTED. TRIED PREVIOUSLY FAILED SP, ( 1 + 2) FAILURES CONSISTANT. NEW SP | CDC6CC265C0B8669 | No | CHECK LOGS. SP HAD BEEN CHANGED 2X. IN SVC, PUMP CAL, CLOT DETECT, LS FAILED REPEATEDLY. PUMP CAL FINALLY PASSED AT 1.19. BY AM PUMP FAILED | INSTALLED NEW SP. ALL TESTS PASSED. SC FAILED PCO2 & LAC, RUNNING OVER NIGHT. NEW SC ACCEPTED. TRIED PREVIOUSLY FAILED SP, ( 1 + 2) FAILURES CONSISTANT. NEW SP | Defect | SP | Pump cal error | Ji?í Volf
ABL90 | Lactate low results on quality control 3 - started 9 2 2014 | Replaced solution pack for troubleshooting purposes. Replace sensor cassette with other onsite analyzer's cassette (basically swap cassettes). After cassette replacement, quality controls came in. Upgraded software to version 3.1 MR1. | CDB3C7F7725864D9 | No | Lactate low results on quality control 3 - started 9 2 2014 | Replaced solution pack for troubleshooting purposes. Replace sensor cassette with other onsite analyzer's cassette (basically swap cassettes). After cassette replacement, quality controls came in. Upgraded software to version 3.1 MR1. | Defect | SC | Lactate
ABL90 | Analyzer was in Demo mode when I arrived. I changed it to regular mode and calibrated the system. The Ph backlog error would not clear out. | Mike Smith brought a New Cassette Replaced the Cassette after Mike delivered it to the Hospital. Brought system to ready. | CDB3929A72343149 | Yes | Analyzer was in Demo mode when I arrived. I changed it to regular mode and calibrated the system. The Ph backlog error would not clear out. | Mike Smith brought a New Cassette Replaced the Cassette after Mike delivered it to the Hospital. Brought system to ready. | Re classified RMED | Analyzer in Demo
TCM4 | Após erro de calibração o cliente substituiu a membrana e voltou a calibrar o equipamento. No dia seguinte, quando voltou a ligar o equipamento, este pede para calibrar e dá sempre erro de calibração. Substituí a membrana, activei a função smartcal e iniciei a calibração. O equipamento durante uma hora calibrou 3 vezes, sempre correctamente. |  | CDB3A7D57207DA2E | No | After calibration error The client replaced the membrane and recalibrated the equipment. The next day, when you switched the equipment back on, it prompts you to calibrate and always gives a calibration error. I replaced the membrane, activated the smartcal function and started the calibration. The equipment for 1 hours has been calibrated 3 times, always correctly. |  | No fault found
ABL820 FLEX | siehe ausgeführte Arbeiten | siehe ausgeführte Arbeiten | CDB3791E72448BA9 | No | see work carried out | see work carried out
ABL825 FLEX |  |  | CDB4DE60E83C7037 | No
ABL825 FLEX |  |  | CDB4DF93E851703A | No
ABL835 FLEX | siehe ausgeführte arbeiten | siehe ausgeführte arbeiten | CDB3A1A37212EDAE | No | see executed work | see executed work
ABL827 FLEX | pCO2 membrane was changed per schedule on 9-1 and calibrations failed for status, drift & sensitivity. Second remembraning did not help errors. Found a portion of the Ag AgCl reference band to be missing. | Replaced electrode. Calibrated and ran QC. Instrument operates normally. | CDB36DE67249906D | No | pCO2 membrane was changed per schedule on 9-1 and calibrations failed for status, drift & sensitivity. Second remembraning did not help errors. Found a portion of the Ag AgCl reference band to be missing. | Replaced electrode. Calibrated and ran QC. Instrument operates normally.
ABL827 FLEX | Inspected the electrodes, measuring chamber and tubing. | Cleaned the contacts with alcohol. | CDB3BE2C7206C757 | No | Inspected the electrodes, measuring chamber and tubing. | Cleaned the contacts with alcohol.
ABL835 FLEX | Unable to remote support into Mission Bay analyzers from Parnassus campus. Customer has analyzer control (can call a Cal or Rinse) but no remote desktop. These new systems are Win XP 5.1 | Temporarily deactivated Windows firewall to restore full connectivity | CDB38EA7723BD304 | No | Unable to remote support into Mission Bay analyzers from Parnassus campus. Customer has analyzer control (can call a Cal or Rinse) but no remote desktop. These new systems are Win XP 5.1 | Temporarily deactivated Windows firewall to restore full connectivity
ABL800Basic |  |  | CDB48D9BE8584D54 | No
ABL80 FLEX Basic version |  |  | CDB49779E842BE81 | No
ABL90 |  |  | CDB4A464E85FD3BE | No |  |  | Defect | HW | Barcode Reader
ABL90 | not able to pump solution sensor bay plug | replaced sensor bay but the new senso bayr from my stock was plug also I will returned that sensor bay to rame | CDB4E989E8143193 | No | not able to pump solution sensor bay plug | replaced sensor bay but the new senso bayr from my stock was plug also I will returned that sensor bay to rame | Defect | HW | Cassette retainer
ABL90 |  |  | CDB4E970E84E3EAE | Yes |  |  | Re classified RMED | FAN / TU
ABL815 FLEX |  | Austausch Positionskalibrierung | CDB4A885E83D99E8 | No |  | Replacement position calibration
ABL825 FLEX |  |  | CDB4CF5FE81FE757 | No
ABL825 FLEX |  |  | CDB4B69FE8536A6F | No
ABL825 FLEX |  | Ventile überprüft und gereinigt; LS-Justage und Pumpenkal | CDB498A3E8157AAD | No |  | Valves checked and cleaned; LS-Adjustment and Pump Scale
ABL800Basic |  |  | CDB4BF65E85C50C9 | No
ABL800Basic |  | Scanner Austausch Handscanner | CDB4CC48E80AEE09 | No |  | Scanner Exchange Hand Scanner
AQT90 FLEX |  |  | CDBEC46C9803BC9B | No
TCM/Tosca | Inspection of touch screen revealed it is non- operational & will not calibrate. | Replaced touch screen. | CDB4E0C6E82FB4DE | No | Inspection of touch screen revealed it is non- operational & will not calibrate. | Replaced touch screen. | Defect | HW | Touch screen
ABL80 | Explicación del uso del analizador ajuste de configuración del sistema de deteción de las posibles burbujas de aire, explicación cambios y chequeos del analizador funcionando correctamente. |  | CDB614B05E3C336F | Yes | Explanation of the use of the analyzer setting the system to detect possible air bubbles, explain changes and checks of the analyzer working properly. |  | Re classified RMED | Setup and training
ABL90 | The analyzer was in the replace the inlet GSK and was on software 3.0 MR2 | I upgraded the software and the analyzer came up with out any errors. | CDB5E9AC5E4D0A80 | No | The analyzer was in the replace the inlet GSK and was on software 3.0 MR2 | I upgraded the software and the analyzer came up with out any errors. | Defect | SW | 3. MR2
ABL90 | Analyzer has a fluid transport problem. Entered the service program and found that the SP & clot detection programs had failed. | Ran the clot detection program & SP program again, but it still failed. Removed the SP & performed the back flush until a steady stream was achieved, but this did not resolve the problem. Replaced the inlet gasket, probe, and inlet gasket connector. I also pushed down on the top of the SP to ensure it was activated, & a snap was heard, indicating improper activation, and this resolved the problem. | CDB628285E15E812 | No | Analyzer has a fluid transport problem. Entered the service program and found that the SP & clot detection programs had failed. | Ran the clot detection program & SP program again, but it still failed. Removed the SP & performed the back flush until a steady stream was achieved, but this did not resolve the problem. Replaced the inlet gasket, probe, and inlet gasket connector. I also pushed down on the top of the SP to ensure it was activated, & a snap was heard, indicating improper activation, and this resolved the problem. | User Error | SP activate | Incorrect activation
ABL825 FLEX |  |  | CDBA32A1C0171D99 | No
ABL825 FLEX | inspected air line for blockage and it was slightly blocked. | I flushed out the air line. | CDB5EC165E3619C1 | No | inspected air line for blockage and it was slightly blocked. | I flushed out the air line.
ABL827 FLEX |  |  | CDBC47EDAC2D3DEC | No
ABL837 FLEX | Tried new Crea electrodes, did not help. Remembraned Crea electrodes on 754R1848N003, the sensitivity went from 8.0pA to 15.2pA. | The membranes came from lot R0533 & R0534. Put lot#R0536 and the sensitivities dropped to 10.3 and to 11.0 on R1848N003. Put their electrodes in and cals and QC are good. | CDB60BB05E0A4F94 | No | Tried new Crea electrodes, did not help. Remembraned Crea electrodes on 754R1848N003, the sensitivity went from 8.0pA to 15.2pA. | The membranes came from lot R0533 & R0534. Put lot#R0536 and the sensitivities dropped to 10.3 and to 11.0 on R1848N003. Put their electrodes in and cals and QC are good.
ABL835 FLEX | No FO2Hb,FCOHb, FMetHb results for qc and samples. Checked the result messages and found error 604 'Parameter not installed. | Found the Oxi 20 keys installed as 'Demo' and found expired. Installed the 'Permanent Oxi 20' option keys. | CDB74975D436365C | No | No FO2Hb,FCOHb, FMetHb results for qc and samples. Checked the result messages and found error 604 'Parameter not installed. | Found the Oxi 20 keys installed as 'Demo' and found expired. Installed the 'Permanent Oxi 20' option keys.
ABL835 FLEX | No FO2Hb,FCOHb, FMetHb results for qc and samples. Checked the result messages and found error 604 'Parameter not installed. | Found the Oxi 20 keys installed as 'Demo' and found expired. Installed the 'Permanent Oxi 20' option keys. | CDB74C65D445365F | No | No FO2Hb,FCOHb, FMetHb results for qc and samples. Checked the result messages and found error 604 'Parameter not installed. | Found the Oxi 20 keys installed as 'Demo' and found expired. Installed the 'Permanent Oxi 20' option keys.
ABL90 | Leckage erkannt Meldung über HOTLINE ABL nicht meßbereit Kunde hat bereits LP 2x getauscht | Fehlersuche Fließwege gespült Jahreswartung durchgeführt (Tausch aller Schläuche, Dichtungen) Anderen Lösungspack getestet kein neuer LP vorhanden,aus ein neuen aber lt. Datum verfallenen LP eingesetzt (Datum angepaßt)- Kunden informiert, sollte nach Lieferung LP diesen erneuern und Dartum korrigieren | CDB9AC60C05008F0 | No | Leakage detected message via HOTLINE ABL not ready to measure Customer has already replaced LP 2x | Troubleshooting flow paths flushed Annual maintenance carried out (exchange of all hoses, seals) Other solution pack tested no new LP available, from a new but according to date forfeited LP used (date adapted)- customers should inform after delivery LP this renew and correct dartum | Defect | SP | FT
ABL90 |  |  | CDB9BD23C03C3461 | No |  |  | Adjustment | Oxi Module | Reboot | Error 134
ABL90 |  | Abfallschlauch am Anschlussstutzen des Hemolyzers perforiert--> System gereinigt und Schlauch nachgesetzt | CDB9A54AC05FE4CD | No |  | Waste hose at the connection socket of the hemolyzer perforated--> system cleaned and hose added | Defect | HW | Waste tube
ABL90 |  |  | CDB9E7F6C0056212 | Yes |  |  | Re classified RMED | TU DWU
ABL810 FLEX |  |  | CDBA0226C025B574 | No
ABL825 FLEX | Tried a new lot of rinse solution-No change | Replaced the pH electrode; lot LN03 under warranty | CDB9C114C00743BF | No | Tried a new lot of rinse solution-No change | Replaced the pH electrode; lot LN03 under warranty
ABL820 FLEX | Inlet tubing was stretched out on both ends of tubing making for leakage.Also discovered tubing at lower chamber Liquid sensor was overlly streched out. | Replaced tubings, flushed fluid pathing, inspected pHBG and Fluidic valves. | CDB9E4F4C03AFAA0 | No | Inlet tubing was stretched out on both ends of tubing making for leakage. Also discovered tubing at lower chamber Liquid sensor was overlly streched out. | Replaced tubings, flushed fluid pathing, inspected pHBG and Fluidic valves.
ABL80 CO-OX | Ran the service test, and the hemolyzer failed for a low cabitation absorb value. | Removed and replaced the hemolyzer, and this resolved the problem. | CDBB1D6836468738 | No | Ran the service test, and the hemolyzer failed for a low cabitation absorb value. | Removed and replaced the hemolyzer, and this resolved the problem.
ABL90 | the instrument does not power up , i found the fuses blown in line cord housing and the line cord missing a ground | na | CDBB52AA3635C8E5 | No | the instrument does not power up , i found the fuses blown in line cord housing and the line cord missing a ground | In | User Error | Replacement process | SC | Life time exp
ABL90 |  |  | CDBB3D33360AF274 | No |  |  | Defect | HW | CPU
ABL825 FLEX |  |  | CDBAFCB43615A8FF | No
ABL825 FLEX |  |  | CDBC2303AC5AAEE8 | No
TCM-COMBIM | THE SENSOR AND THE MODULE WERE INSERTED INTO ANOTHER TCM. THERE WERE NO GENERAL ERRORS TO REPORT | THE MAIN CPU OR INTERFACE BOARD WILL HAVE TO BE REPLACED. | CDBB293236259945 | No | THE SENSOR AND THE MODULE WERE INSERTED INTO ANOTHER TCM. THERE WERE NO GENERAL ERRORS TO REPORT | THE MAIN CPU OR INTERFACE BOARD WILL HAVE TO BE REPLACED. | Defect | HW | PCB, Main CPU | PCB, Main CPU
ABL80 FLEX BASIC | Adapter cseréje szükséges. | András Hegedüs | SWO-1106 | No | Adapter replacement required. | Madhu A |  |  |  | PCB, Main CPU
ABL90 |  |  | CDBC7847AC2FC7BC | No |  |  | Replaced | Inlet gasket |  | PCB, Main CPU
ABL90 | The customer is trying to download the WDC data using a USB without success. | I tried my USB drive and the analyzers were able to detect it properly. However, the analyzers cannot detect the customer's USB drive due to a security software loaded on the USB drive. I contacted technical support and made Rame aware of the customer's concern. Dino Garcia from Radiometer explained the our operating system does not support USB drive that have security software. | CDBC2D69AC557BE4 | No | The customer is trying to download the WDC data using a USB without success. | I tried my USB drive and the analyzers were able to detect it properly. However, the analyzers cannot detect the customer's USB drive due to a security software loaded on the USB drive. I contacted technical support and made Rame aware of the customer's concern. Dino Garcia from Radiometer explained the our operating system does not support USB drive that have security software. | No fault found | Cannot upload to USP memory | Security SW | PCB, Main CPU
ABL90 | looked into data log transfer, have offered to help with next months files if problem continues, | came on site and downloaded WDC report to USB; opened files in notebook and all looks good; sent files from customers USB to radiometer and files are not corrupted; believe customers computer or web email is causing the problem; no problem with the analyzer; cusotmer had questions about QC answered all questions; downloaded Aug WDC and sent May-Aug files to westlake | CDBC4971AC5633E0 | No | looked into data log transfer; have offered to help with next months files if problem continues; | came on site and downloaded WDC report to USB; opened files in notebook and all looks good; sent files from customers USB to radiometer and files are not corrupted; believe customers computer or web email is causing the problem; no problem with the analyzer; cusotmer had questions about QC answered all questions; downloaded Aug WDC and sent May-Aug files to westlake | No fault found | Cannot upload to WDC | Security SW | PCB, Main CPU
ABL90 | Inlet cover not fit propely touching inlet handle | I believe, i will need to replaced front panel or magnet not hold hard enough the inlet cover, i will dicuss that issue with my ASM | CDBD5BCE223E0625 | No | Inlet cover not fit propely touching inlet handle | I believe, i will need to replaced front panel or magnet not hold hard enough the inlet cover, i will dicuss that issue with my ASM | Defect | HW | Inlet Cover | PCB, Main CPU
ABL90 | Battery option needs installed in this unit. | First installed battery option failed. Started to produce a electrical burn smell. Reported issue to technical support. Installed second battery option with any issues. | CDBCCCF5AC4C5E83 | No | Battery option needs installed in this unit. | First installed battery option failed. Started to produce a electrical burn smell. Reported issue to technical support. Installed second battery option with any issues. | Defect | HW | Battery module | PCB, Main CPU
ABL825 FLEX |  |  | CDBCA457AC3C880B | No | [5 empty] | PCB, Main CPU
ABL800Basic |  |  | CDBC2936AC627A70 | No | [5 empty] | PCB, Main CPU
ABL825 FLEX |  |  | CDBC73B9AC28FC84 | No | [5 empty] | PCB, Main CPU
ABL835 FLEX |  |  | CDBCA845AC23880E | No | [5 empty] | PCB, Main CPU
ABL90 |  |  | CDBD913A223FDE35 | Yes |  |  | Re classified RMED | Preventive maintenance | Additional customer training | PCB, Main CPU
ABL90 |  |  | CDBD87A022113C1E | Yes |  |  | Re classified RMED | Preventive maintenance | Could not finalize on the 11/9-14
ABL5 |  |  | CDBD89C52248DE90 | No
ABL825 FLEX | PCO2 élevée sur controles bas sur niveaux 3 vérifier gas module OK et nettoyé contact | Remplacé électrode de PCO2 | CDBD7C132248CB87 | No | High PCO2 on low controls on levels 3 check gas module OK and cleaned contact | Replaced PCO2 electrode
ABL825 FLEX |  | alle Ventile überprüft und gereinigt; Jahreswartung durchgeführt | CDBD7792223CBA87 | No |  | checked and cleaned all valves; Annual maintenance carried out
ABL825 FLEX | Reboot X2-No change As I was touching the screen found the calibration was way off | Did a touch screen calibration in windows Did the same for the other 2 ABL800's Tested screen and all good | CDBF218A983F79FC | No | Reboot X2-No change As I was touching the screen found the calibration was way off | Did a touch screen calibration in windows Did the same for the other 2 ABL800's Tested screen and all good
AQT90 FLEX |  |  | CDBD943E2215809C | No
ABL80 FLEX Basic version |  |  | CDBEC743981B70F2 | No
ABL90 | siehe SLR | siehe SLR | CDBEA7D59835D2BF | No | see SLR | see SLR | Cleaned | Fluidic pathway | Connector gasket
ABL90 |  |  | CDBEBBDB98129D69 | No |  |  | Defect | SP | FT
ABL835 FLEX | Waste detector csere, kalibrálás, QC mérés. | András Hegedüs | SWO-1110 | No | Waste detector replacement, calibration, QC measurement. | Madhu A
ABL837 FLEX |  |  | CDBEC56E98180AC7 | No
ABL827 FLEX |  |  | CDC28418FA1292B1 | No
ABL837 FLEX | Replaced blue tubing with clear tubing and observed air bubble coming out of inlet probe(Intermitent) | Replaced inlet probe, performed multiple blood sample with Chris, all OK | CDBEF9E3985C0F91 | No | Replaced blue tubing with clear tubing and observed air bubble coming out of inlet probe(Intermitent) | Replaced inlet probe, performed multiple blood sample with Chris, all OK
ABL837 FLEX | Tried lot#R0533 and R0534, same error | Tried Lot#536, sensitivity is now normal. Downloaded data files for TechLine case. Made TechLine case | CDBEF62598490F8F | No | Tried lot#R0533 and R0534, same error | Tried Lot#536, sensitivity is now normal. Downloaded data files for TechLine case. Made TechLine case
AQT90 FLEX |  |  | CDBEEF219852F1A5 | No
ABL80 FLEX Basic version |  |  | CDC2930AFA019585 | No
ABL90 |  |  | CDC28499FA1F9581 | No
ABL90 | The SC's were either completely empty or most of the KCL was crystallized on the outsude of the SC body. all lot # R0568N69. | All 5 will have to be replaced. three were cleaned up and dried and installed for short term untill new SC arrive. SC's shipped on clients memo account. New installation. | CDC63D725C5813C0 | No | The SC's were either completely empty or most of the KCL was crystallized on the outsude of the SC body. all lot # R0568N69. | All 5 will have to be replaced. three were cleaned up and dried and installed for short term untill new SC arrive. SC's shipped on clients memo account. New installation. | Defect | SC
ABL820 FLEX | Error 581 Spectrum mismatch on qc results. | Replaced hemolyzer and performed lamp adjustment | CDC2C945FA4CAF41 | No | Error 581 Spectrum mismatch on qc results. | Replaced hemolyzer and performed lamp adjustment
ABL825 FLEX | Pump tubing on pH BG module had leaked where connected to pre-heater. | Replaced all pump tubings. Replaced Air Valve membrane on pH BG Module. | CDC2807FFA4A9DA0 | No | Pump tubing on pH BG module had leaked where connected to pre-heater. | Replaced all pump tubings. Replaced Air Valve membrane on pH BG Module.
AQT90 FLEX |  |  | CDC2C9B0FA1DF16B | No
AQT90 FLEX |  |  | CDC2D205FA41F16E | No
ABL90 |  |  | CDC3C04670142004 | No |  |  | Replaced | Probe
ABL90 |  |  | CDC68C875C389652 | Yes |  |  | Re classified RMED | LEH Mail
ABL90 | CHECK LOGS FOR MESSAGES. INNHOMOGENOUS CAL 3, CAL 2, & ALL QC, INCONSISTANTY. OXI & PO2 TEST PASSED IN SVC. FLUIDICS TEST:SP FAILED RINSE. RAN ALL TEST 2X+. PASSED. CALS & INTERNAL QC ACCEPTED. | SHUTDOWN & RESTART CLEARED DB ERROR, BUT INHOMOGENIOUS CAL 3. REPEATED CALS | CDC6D0EB5C578674 | No | CHECK LOGS FOR MESSAGES. INNHOMOGENOUS CAL 3, CAL 2, & ALL QC, INCONSISTANTY. OXI & PO2 TEST PASSED IN SVC. FLUIDICS TEST:SP FAILED RINSE. RAN ALL TEST 2X+. PASSED. CALS & INTERNAL QC ACCEPTED. | SHUTDOWN & RESTART CLEARED DB ERROR, BUT INHOMOGENIOUS CAL 3. REPEATED CALS | Adjustment | Run Service programs | power off/on | OBS Previous UpSWO & Email
ABL90 |  | Service Check, diverse Fehler bezgl. Lösungspack, Ersatz des LP, Kal, QC Nadeln und Einlassdichtungen neue Version 1:1 getauscht | CDC3E6B5704C260C | No |  | Service Check, various errors regarding Solution pack, replacement of LP, Kal, QC needles and inlet seals new version 1:1 exchanged | Defect | SP | FT
ABL90 |  |  | CDC3AD41702636A9 | No |  |  | Defect | HW | CPU
ABL820 FLEX |  | Austausch CO Elektrode | CDC40F30700254CD | No |  | Replacement CO electrode
ABL825 FLEX |  | Austausch Ref Elektrode durch Kunden | CDC4103470087681 | No |  | Replacement of Ref Electrode by Customers
ABL815 FLEX |  |  | CDCBC85F3414C994 | No
ABL825 FLEX |  |  | CDC3FA567053859A | No
ABL80 FLEX Basic version |  |  | CDC80672D24B1DF8 | No
ABL80 | I inspected roller and rubings and all lookede fine. | I replaced the pump roller. | CDC50AB5E6634CFD | No | I inspected roller and rubings and all lookede fine. | I replaced the pump roller.
ABL90 |  |  | CDC53FD3E61FA971 | No
ABL90 |  |  | CDC56FCDE624002B | No |  |  | Defect | SP | FT
ABL90 | not able to pump solution | flush the system and replaced inlet gasket with customer stock installed software 3.1 MR1 | CDC53A0BE637823C | No | not able to pump solution | flush the system and replaced inlet gasket with customer stock installed software 3.1 MR1 | Cleaned | Fluidic pathway
ABL90 | ACTUALIZACIÓN DEL SOFTWARE A LA V.3.1MR1, ACTUALIZACION DE MANUALES DE USUARIO Y REFERENCIA, CONFIGURACIÓN E INSTALACIÓN DE ADAPTADOR WIFI PARA CONECTAR EL EQUIPO VÍA WIFI A LA RED DEL HOSPITAL Y A RADIANCE. |  | CDC532E1E622F6DA | Yes | UPDATE OF THE SOFTWARE TO V.3.1MR1, UPDATE OF USER AND REFERENCE MANUALS, CONFIGURATION AND INSTALLATION OF WIFI ADAPTER TO CONNECT WIFI EQUIPMENT TO THE HOSPITAL NETWORK AND RADIANCE. |  | Re classified RMED | FAN / TU
ABL800Basic | Kontrolle der Elektroden, die heute membraniert wurden, bei Lac hatte sich die untere Kappe quergeschoben, wieder in richtige Lage gebracht, Kal, QC |  | CDC548F5E60B4E15 | No | Control of the electrodes, which were membraneed today, at Lac had pushed the lower cap transversely, put it back in the right position, Kal, QC
ABL800Basic |  |  | CDC5210DE6435341 | No
AQT90 FLEX |  |  | CDC506B6E63F9F7B | Yes |  |  | Re classified RMED | OOB | Email : From: Volf, Jiri JIRVO Sent: 23. september 214 8:54 To: Petersen, John M. JMP Cc: Tarp, Mads Agervig MTARP Subject: Re: upSWO | Hi John, Yes BR Jiří Volf Service manager
AQT90 FLEX | Fehlersuche | defekte TnI-Kassette LOT 9619 entfernt | CDC531BCE64293A9 | No | Troubleshooting | defective TnI cassette LOT 9619 removed
TCM-COMBIM | Installation of a new Combi M unit required | Performed installation | CDC513C7E65E58B8 | Yes | Installation of a new Combi M unit required | Performed installation | Re classified RMED | Installation
ABL80 CO-OX |  |  | CDC63D205C1EDF9B | No
ABL80 FLEX Basic version | Entrega del material urgente solicitado por la Dra. Isabel Guitart. | Entrega de material urgente. | CDCCA21CAA511D2B | Yes | Delivery of the urgent material requested by Dr. Isabel Guitart. | Delivery of urgent material. | Re classified RMED | Logistical upSWO
ABL80 OSM | Checked data logs for errors. Ran oximetry tests with pass results on all on-site ABL80 OSM's. | Ran tHB calibrations on this and older analyzers with pass results. Continue to monitor next set of comparison samples. | CDC6FC955C0435EE | No | Checked data logs for errors. Ran oximetry tests with pass results on all on-site ABL80 OSM's. | Ran tHB calibrations on this and older analyzers with pass results. Continue to monitor next set of comparison samples.
ABL80 OSM | Checked data logs for errors. Ran oximetry tests with pass results. Contacted product specialist to discuss issue. | Ran tHB calibration. Will monitor next comparison samples. | CDC6F7235C4435EB | No | Checked data logs for errors. Ran oximetry tests with pass results. Contacted product specialist to discuss issue. | Ran tHB calibration. Will monitor next comparison samples.
ABL80 CO-OX |  |  | CDC686175C2E598D | No
ABL90 |  |  | CDC651B05C18027B | No
ABL90 |  | SC ersetzt, Kal, QC | CDC62FF05C183577 | No |  | SC ersetzt, Kal, QC | Defect | SC | x 2
ABL90 |  |  | CDC675905C180282 | No
ABL835 FLEX | reference membrane issue no Oring on Pco2 electrode tip Glu lot# R1916 | replaced ref membrane new lot replaced tip oring Inspected GLU electrode connector and measuing well will check with REME regarding issue with GLU membranes.or electrode. this unit has had the GLU electrode replaced , and may need to have the EI MET pcb replaced | CDC670215C2808DF | No | reference membrane issue no Oring on Pco2 electrode tip Glu lot# R1916 | replaced ref membrane new lot replaced tip oring Inspected GLU electrode connector and measuing well will check with REME regarding issue with GLU membranes.or electrode. this unit has had the GLU electrode replaced , and may need to have the EI MET pcb replaced
ABL825 FLEX | The chassis for the inlet module is broken at the inlet flaps. | Replaced the chassis for the inlet module. Drained the analyzer as the customer does not have this unit in use and the instrument will be shut down for an extended period of time during the completetion of the building construction. | CDC68F615C4142C3 | No | The chassis for the inlet module is broken at the inlet flaps. | Replaced the chassis for the inlet module. Drained the analyzer as the customer does not have this unit in use and the instrument will be shut down for an extended period of time during the completetion of the building construction.
ABL815 FLEX |  |  | CDC6C8635C58592F | No
ABL825 FLEX | Tried new lot of membranes and cleaned electrode contacts-No Change Cleaned pCO2 electrode-No change | Replaced pCO2 electrode pCO2 electrode lot # LK01 defective and under warranty | CDC6A3E65C2CB98E | No | Tried new lot of membranes and cleaned electrode contacts-No Change Cleaned pCO2 electrode-No change | Replaced pCO2 electrode pCO2 electrode lot # LK01 defective and under warranty
ABL835 FLEX | Pco2 valeurs basse sur tout les niveaux et non-homogene inlet coule | remplacé joint introduction flush V7.V8 remplacé joint PH electrode installé nouvelle electrode de PCo2 et membrane calcium nettoyé chambre de mesure et configuré nettoyage plus fréquent remplacé waste tube | CDC64E685C060E79 | No | Pco2 low values on all levels and non-homogene inlet flows | replaced joint introduction flush V7. V8 replaced joint PH electrode installed new PCo2 electrode and calcium membrane cleaned measuring chamber and configured more frequent cleaning replaced tube waste
ABL825 FLEX | Reboot and did a touch screen calibration-No change | Replaced the touch screen Test and calibrated in windows-OK | CDC6E4F65C47683E | No | Reboot and did a touch screen calibration-No change | Replaced the touch screen Test and calibrated in windows-OK
ABL827 FLEX | Removed the covers and inspected tubing and valves. I do not see any problems. I suspect a membrane valve or a pump tubing. | Removed all three of the measuring module's dual valves and cleaned. Replaced all three of the measuring module's pump tubings. | CDC686885C005AAF | No | Removed the covers and inspected tubing and valves. I do not see any problems. I suspect a membrane valve or a pump tubing. | Removed all three of the measuring module's dual valves and cleaned. Replaced all three of the measuring module's pump tubings.
AQT90 FLEX | Conmprobado tensiones, temperatura y configuración del equipo. El equipo se queda parado pendiente de piezas. |  | CDC6AAAF5C5E44D3 | No | Recommended voltages, temperature and equipment configuration. The equipment stands pending parts.
ABL80 FLEX Basic version |  |  | CDD057A70C320CBA | No
ABL90 | siehe SLR | siehe SLR | CDC7F34AD254F38B | No | see SLR | see SLR | Replaced | Inlet gasket
ABL90 |  |  | CDC7BE66D205F77F | No
ABL800 BASIC | kalibrációk, QC mérés, 1doboz tévedésből kiszállított GLU membrán (942-065 R2019) visszaszállítása Radiometer Mo. Kft. telephelyére | Miklós Czirják | SWO-1175 | No | calibrations, QC Measurement, 1carton of the GLU membrane (942-065 R2019) shipped out of error to Radiometer Mo. Kft. 's site | Nicholas
ABL827 FLEX |  |  | CDC7CFAED22F7C8C | No
ABL835 FLEX | Explicacion del funcionameinto Sustitucion del bote de cal1. Cal , qc y muestras |  | CDC7A42AD25FB440 | No | Explanation of the functionof Substitution of the cal1 boat. Cal , qc and samples
ABL825 FLEX | Verifed the problem. Had already tried a decontamination of the unit, b ut the problem came back a few days later. | Replaced the electrode with a new one. LOT # RB01, and replaced the membrane from customerr stock. | CDC7D81ED20AD1D4 | No | Verifed the problem. Had already tried a decontamination of the unit, b ut the problem came back a few days later. | Replaced the electrode with a new one. LOT # RB01, and replaced the membrane from customerr stock.
ABL827 FLEX |  |  | CDCBD2E934552222 | No
AQT90 FLEX |  |  | CDC7933ED2288D61 | No
TCM400 | Zodoende het volgende heb gecontroleerd, - TCM 400 wordt correct opgestart. - Ref. module wordt correct opgestart. - De beide elektrode (kalibratie meet) worden wel herkend door de TCM. Oorzaak ligt waarschijnlijk aan de verlengkabel tussen de TCM apparaat en de elektrode. De verlengkabel moet worden vervangen. | Reza Gharib | SWO-1177 | No | Have checked the following; -TCM 400 is booted correctly. -Ref. module is booted correctly. -The two electrode (calibration measures) are recognized by the TCM. Cause is probably the extension cable between the TCM device and the electrode. The extension cable must be replaced. | Reza Gharib
ABL90 | Unit measurement of Lactate= 8.5-8.8. Low end of range 9.2.Muliple lots of SCs' have been tried.Fails after 7-10 days after intial install of SC. | Replacement of SCs'.' | CDCBAE22344E19C5 | No | Unit measurement of Lactate= 8.5-8.8. Low end of range 9.2.Muliple lots of SCs' have been tried. Fails after 7-10 days after intial install of SC. | Replacement of SCs'.' | Defect | SC | Lactate
ABL800Basic |  |  | CDCB73D5346325FD | No
ABL815 FLEX |  |  | CDCB8D99344E0D7D | No
ABL820 FLEX | Checked data logs for errors. Found tHB readings drifting during two point calibrations. Ran complete oximetry tests with pass results. Checked halogen lamp intensity, output was slightly low. Found additive (dosi cap) not mixed into cleaning solution. | Adjusted halogen intensity. Ran tHB calibration in service mode. Additive mixed back into cleaning solution. Ran additional cleaning program. Monitor calibrations for tHB drift. After multiple calibrations, tHB drifted again. Replace lamp unit . Run service tHB calibration and restart calibrations. | CDCB77E4342C22DF | No | Checked data logs for errors. Found tHB readings drifting during two point calibrations. Ran complete oximetry tests with pass results. Checked halogen lamp intensity, output was slightly low. Found additive (dosi cap) not mixed into cleaning solution. | Adjusted halogen intensity. Ran tHB calibration in service mode. Additive mixed back into cleaning solution. Ran additional cleaning program. Monitor calibrations for tHB drift. After multiple calibrations, tHB drifted again. Replace lamp unit . Run service tHB calibration and restart calibrations.
ABL825 FLEX | Unit also has 316 and 258 neon lamp errors in activity log. Unit has also been leaking rinse solution at tubing connections between modules. No obvious reason why. | Removed and cleaned all modules. Recalibrated inlet module. Tried to align oxi module but could not. Replaced oxi module PCB. Realigned oxi module. | CDCB59F434417C27 | No | Unit also has 316 and 258 neon lamp errors in activity log. Unit has also been leaking rinse solution at tubing connections between modules. No obvious reason why. | Removed and cleaned all modules. Recalibrated inlet module. Tried to align oxi module but could not. Replaced oxi module PCB. Realigned oxi module.
AQT90 FLEX |  |  | CDCB5728342893DF | No
AQT90 FLEX |  |  | CDCB79F4344A0DDA | No
ABL80 CO-OX |  |  | CDCCA46DAA1149DB | No
ABL80 FLEX Basic version |  |  | CDCCD5A8AA4E0077 | No
ABL80 CO-OX | CO-OX blank failing. tHb cal failures in data logs. Unit failed Cavitation absorb portion of service test. | Replaced hemolyzer unit. Unit leaking at cassette waste port. Repositioned waste tube. Aligned oximeter. | CDCCBE7CAA23C3BB | No | CO-OX blank failing. tHb cal failures in data logs. Unit failed Cavitation absorb portion of service test. | Replaced hemolyzer unit. Unit leaking at cassette waste port. Repositioned waste tube. Aligned oximeter.
ABL90 | instrument is at 3.1 software | i installed software 3.1MR1, reloaded wifi info, changed ip address saved it , put new cassette and solution pack on | CDCCE6E8AA288F0C | No | instrument is at 3.1 software | i installed software 3.1MR1, reloaded wifi info, changed ip address saved it , put new cassette and solution pack on | Defect | HW | Power supply
ABL90 |  |  | CDCCA592AA3CAF81 | No |  |  | Adjustment | New Data base
ABL90 | Battery Option needs replaced. | Replaced Battery Option. Test battery operation. | CDCCF226AA55A5A9 | No | Battery Option needs replaced. | Replaced Battery Option. Test battery operation. | Defect | HW | Battery module
TCM4 | Inicialização. Inicio de calibração. Garrafa de gás vazia. Substituição da garrafa de gás. Calibração ok. |  | CDCCDA7BAA155D48 | No | Startup. Start calibration. Empty gas bottle. Replacing the gas bottle. Calibration OK. |  | User Error | Replacement process | Gas cylinder empty
ABL800 BASIC | pH BG modul felső levegő szelep membránjának cseréje. Kalibrálás, QC mérés. | András Hegedüs | SWO-1198 | No | pH BG Module Upper air valve membrane replacement. Calibration, QC measurement. | Madhu A
ABL835 FLEX | Comprobar módulo de pH gases, reparar el motor de la bomba del módulo, ajustar sensores y bombas. Calibrar y probar con controles. | Reparar el motor de la bomba del módulo pH gases. | CDCCC65AAA09AC6A | No | Check pH gas module, repair module pump motor, adjust sensors and pumps. Calibrate and test with controls. | Repair the pump motor of the pH gas module.
ABL800Basic |  |  | CDCC92C7AA1AD8B6 | No
ABL90 |  |  | CDCE03E8205E1E11 | No |  |  | Defect | HW | SP PCB id
ABL835 FLEX | CAROUSEL GOES UP AND THEN JUST GOES BACK DOWN | REPLACED AUTOCHECK RIBBON CABLE .STILL NOT WORKING .WILL ORDER PARTS AND RETURN TO CONTINUE WORKING | CDCE3B3020020398 | No | CAROUSEL GOES UP AND THEN JUST GOES BACK DOWN | REPLACED AUTOCHECK RIBBON CABLE . STILL NOT WORKING . WILL ORDER PARTS AND RETURN TO CONTINUE WORKING
TCM400 | Volgens de gebruikershandleiding (pagina 109) mag er niet meer dan 10 meter kabel gebruikt worden tussen de sensorkop en het meetinstrument. In de huidige situatie is de lengte ca. 12 meter en dat is dus te veel. Daarnaast is de kabel doorgeknipt en kennelijk weer aan elkaar gesoldeerd. Deze verbinding is niet geschikt om de uiterst kleine signalen van de sensor naar het meetinstrument ongestoord te kunnen laten verlopen. Mogelijk is daarnaast ook een verbinding niet goed gesoldeerd. ADVIES: Nieuwe kabel(s) doorvoeren, eventueel door middel van te leveren kabel(s) met aan een kant een losse connector. Daarnaast de lengte van de kabel beperken tot één lengte van 6 meter + de sensor met aansluitkabel (3 meter). | Reza Gharib | SWO-1218 | No | According to the User manual (page 109), no more than 10 metres of cable can be used between the sensor head and the measuring instrument. In the current situation the length is approx. 12 meters and that is too much. In addition, the cable is cut through and apparently soldered back together. This connection is not suitable for the extremely small signals from the sensor to the measuring instrument to be able to run undisturbed. In addition, a connection may not be properly soldered. ADVICE: Make new cable (s), possibly by means of cable (s) to be supplied with a loose connector on one side. In addition, limit the length of the cable to one length of 6 meters + the sensor with connection cable (3 meters). | Reza Gharib
TCM400 | Bij onderzoek boven het plafond blijkt dat er gebruik is gemaakt van 2 stuks verlengkabels. Volgens de gebruikershandleiding (pagina 109) mag er niet meer dan 10 meter kabel gebruikt worden tussen de sensorkop en het meetinstrument. In de huidige situatie de lengte ca. 12 meter en dat is te veel. ADVIES: Nieuwe kabel doorvoeren, eventueel door middel van te leveren kabel met aan één kant een losse connector. Daarnaast de lengte van de kabel beperken tot één lengte van 6 meter + de sensor met aansluitkabel (3 meter). | Reza Gharib | SWO-1219 | No | Research above the ceiling shows that 2 pieces of extension cables are used. According to the User manual (page 109), no more than 10 metres of cable can be used between the sensor head and the measuring instrument. In the current situation the length is approx. 12 meters and that is too much. ADVICE: Make a new cable, possibly by means of a cable with a single connector on one side. In addition, limit the length of the cable to one length of 6 meters + the sensor with connection cable (3 meters). | Reza Gharib
ABL80 CO-OX | Found customer was out of solution packs and sensor cassettes. Ordered replacement solution pack and cassette for customer and will submit credit form. Performed multiple rinses and flushed out analyzer and sensor cassette. Reinstalled sensor cassette and could not get Level 1 QC to pass for PCO2. Will replace sensor cassette. | Replaced solution pack and cassette and submitted credit form for customer. Verified instrument operation. | CDCF3C68962B3B4B | No | Found customer was out of solution packs and sensor cassettes. Ordered replacement solution pack and cassette for customer and will submit credit form. Performed multiple rinses and flushed out analyzer and sensor cassette. Reinstalled sensor cassette and could not get Level 1 QC to pass for PCO2. Will replace sensor cassette. | Replaced solution pack and cassette and submitted credit form for customer. Verified instrument operation.
ABL90 | On Sept 14 th they had a situation where a patient's Ca was reported as critical, 3.0 mg dL. Approximately 30 minutes later while running one of our routine patient correlations between analyzers we discovered that this analyzer s Ca was not correlating with the other ABL90s. critical patient sample was pulled and rerun on a second analyzer with a result of 4.5 mg dL, consistent with what she had been running. The unit was notified of the correction and stated that they had already treated the patient based on the 3.0 result. | I have collected data and will submit it. | CDCFC8F396334829 | No | On Sept 14 th they had a situation where a patient's Ca was reported as critical, 3.0 mg dL. Approximately 30 minutes later while running one of our routine patient correlations between analyzers we discovered that this analyzer s Ca was not correlating with the other ABL90s. critical patient sample was pulled and rerun on a second analyzer with a result of 4.5 mg dL, consistent with what she had been running. The unit was notified of the correction and stated that they had already treated the patient based on the 3.0 result. | I have collected data and will submit it. | Performance | Ca to low
ABL90 FLEX | Einlassnadel dichtung ausgetauscht, Nacheinweisung von Sr. Raffaela |  | SWO-1506 | No | Inlet pin seal replaced, instruction by Sr. Raffaela
ABL90 FLEX | nach SW Neustart, LP wegen Fälligkeit gew., Startphase abgewartet = Systemtemp.-fehler, div. Einstelluungen im Service durchgef., tHb-Kal, Kal. + QC i.O. |  | SWO-1493 | No | after SW restart; LP due due; Start phase waited = system temp.-error; div. settings in the service; tHb-Kal; Kal. + QC in the first place
ABL835 FLEX |  |  | CDD088880C1E01A7 | No
ABL820 FLEX | The analyzer indicated that several tubes needed to replaced. The items listed the account did not have in stock. Contacted the area FSE. |  | CDCF529296597E44 | No | The analyzer indicated that several tubes needed to replaced. The items listed the account did not have in stock. Contacted the area FSE.
ABL835 FLEX | Intermittant erreur 374 air bubble dans patient et controles,t inlet calibration erreur frequent inspecte le module d'entrée aucune anomalie visuel | Remplacé inlet preheateur et le friction detecteur unit la probe et les tubes de probe, et inlet aussi remplacé air valve | CDCF2CC1965A5019 | No | Intermiting error 374 air bubble in patient and controls,t inlet calibration frequent error inspects input module no visual anomaly | Replaced inlet preheater and friction detector unites probe and probe tubes, and inlet also replaced air valve
ABL825 FLEX | Screen is non responsive when it freezes up. Screen needs to be replaced. | Replace analyzer screen. After a Calibration, get sample problem on Autocheck QC. Cleaned pH Sealing ring and replaced PO2 Sealing ring for customer stock - was missing. Collect log files and a service dump. | CDCF0EDF965479A1 | No | Screen is non responsive when it freezes up. Screen needs to be replaced. | Replace analyzer screen. After a Calibration, get sample problem on Autocheck QC. Cleaned pH Sealing ring and replaced PO2 Sealing ring for customer stock - was missing. Collect log files and a service dump.
ABL810 FLEX |  |  | CDCF3C239659F20C | No
AQT90 FLEX |  |  | CDD43A096E2FC83D | No
ABL80 CO-OX | 2106 ERROR POSSIBLE FLUID BLOCKAGE. | BROUGHT LOANER FROM RCAN TO CUSTOMER S N 302179. SETUP LONER AT CUSTOMER SITE. TRANSFERED ALL DATA TO CUSTOMER USB DRIVE. LEFT USB DRIVE BESIDE ANALYZER. CUSTOMER ACCEPTED LOANER. | CDD099050C1BDD6B | No | 2106 ERROR POSSIBLE FLUID BLOCKAGE. | BROUGHT LOANER FROM RCAN TO CUSTOMER S N 302179. SETUP LONER AT CUSTOMER SITE. TRANSFERED ALL DATA TO CUSTOMER USB DRIVE. LEFT USB DRIVE BESIDE ANALYZER. CUSTOMER ACCEPTED LOANER.
ABL80 FLEX Basic version | Analise dos registos de erros. Sair de standby. Erro: Temperatura da cassete de eléctrodos instável. Substituição da cassete de eléctrodos. Verificação do correcto funcionamento. |  | CDD05BC50C2B3811 | No | Analyze the error logs. Exit standby. Error: Unstable electrode cassette temperature. Replacing the electrode cassette. Verification of proper operation.
ABL825 FLEX | Checked the data logs and verified problem. | Replaced the lactate electrode. | CDD0BF4C0C19B534 | No | Checked the data logs and verified problem. | Replaced the lactate electrode.
ABL820 FLEX | 376 and 377 cal drift errors in cal log. 588 QC lower than statistical range errors occurring for a few days. | Removed slight amount of air from membrane face. Replaced pH seal ring. | CDD06B190C33A0FB | No | 376 and 377 cal drift errors in cal log. 588 QC lower than statistical range errors occurring for a few days. | Removed slight amount of air from membrane face. Replaced pH seal ring.
TCM-COMBIM |  |  | CDD0A8940C296812 | No |  |  | Defect | HW | PCB, Main CPU | PCB, Main CPU
ABL90 FLEX | Nach Telefonat mit Kunden hatte dieser die SC getauscht. Nach Rückruf hatte die 2. SC einen Substrattemparaturfehler. Beim erneuten Tausch der SC war noch ein Lac Problem vorhanden. | Bei Ankunft hatte Lac noch immer ein QC Problem. Nochmals eine Kalibration durchgeführt. LP (aus GS-Bestand erneuert). Danach Gerät wieder ok. Kunden SC für Lagerhaltung überlassen. Keine Berechnung lt. Hr Lager wegen SC Problem. | SWO-1269 | No | After a phone call with customers, he had exchanged the SC. After the recall, the 2nd SC had a substrate temperature defect. When the SC was re-exchanged, there was still a Lac problem. | Upon arrival, Lac still had a QC problem. Once again a calibration was carried out. LP (renewed from GS stock). After that device is ok again. sc for warehousing. No calculation according to Mr. Lager due to SC problem. | Defect | SC | Lac pCO2
ABL90 |  |  | CDE64590E240C3CD | No |  |  | Defect | SP | Clot
ABL90 FLEX | Kal.- u. QC-Ergebnisse geprüft, Flüssigkeitswege überprüft, Datei Betriebsmittel geprüft | Einlassdichtung u. Anschlussdichtung erneuert; LP aus GS19-Bestand erneuert (Anwender hat zuvor LP mit gleicher LOT - RB04 - getauscht), Anschließende QC i.O. | SWO-1262 | No | Kal.- and QC results checked, fluid paths checked, file equipment checked | Inlet seal and connection seal renewed; LP from GS19 stock renewed (user has previously exchanged LP with the same LOT - RB04), subsequent QC i.O. | Defect | SP | pO2 QC-C Low
ABL90 FLEX | ABL steht auf Offline | ABL ausgeschaltet u. vom Netz getrennt, Eingeschaltet, Startphase abgewartet, im Servicetest: Fluid Transport, Temperatur u. Oxi pO2 überprüft - i.O., Kal u. QC i.O. | SWO-1392 | No | ABL is offline | ABL switched off and disconnected from the mains, switched on, start phase awaited, checked in service test: Fluid Transport, temperature and Oxi pO2 - i.O., Kal and QC i.O. | Defect | SW | Offline
ABL90 | checked to see if solution pack is installed correct. | Performed a back flush | CDD428A76E602C25 | No | checked to see if solution pack is installed correct. | Performed a back flush | Cleaned | Fluidic pathway
ABL90 FLEX | Flüssigkeit im Display. | Touchscreen ersetzt. Startphase. | SWO-1484 | No | Liquid in the display. | Touchscreen replaced. Start. | User Error | Cleaning Display
ABL825 FLEX | Status AutoCheck: nur 2 Ampullen von 20 eingelesen | AutoCheck Modul ausgebaut und gereinigt. Fehler bei anschließender Funktionskontrolle noch vorhanden. Laser Adjustment durchgeführt, 14-popliges Flachbandkabel erneuert - Fehler bleibt bestehen. Folgebesuch wg. Ersatzteilbeschaffung notwendig. | SWO-1298 | No | Status AutoCheck: only 2 ampoules of 20 read in | AutoCheck module removed and cleaned. Failed to occur during subsequent function control. Laser adjustment performed, 14-pop ribbon cable renewed - error remains. Follow-up visit wg. spare parts procurement is necessary.
ABL825 FLEX |  |  | CDD6C7E85A231B4B | No
ABL835 FLEX | VISUSAL INSPECTION SHOWS BROKEN FRONT COVER | REPLACED AUTO CHECK MODULE AND FRONT COVER | CDD42B0A6E152588 | No | VISUSAL INSPECTION SHOWS BROKEN FRONT COVER | REPLACED AUTO CHECK MODULE AND FRONT COVER
AQT90 FLEX |  |  | CDD429336E53F035 | No
AQT90 FLEX |  |  | CDD4500A6E331817 | No
TCM-COMBIM | THE INSTRUMENT WAS POWERE ON, THE MONITOR WOULD CALIBRATE THE SENSOR AND THEN IT WOULD POST A GENERAL ERROR | THE INTERFACE BOARD WAS CHANGED AND THE SAME GENERAL ERROR CODE WAS OCCURRING. THE MAIN CPU UNIT WAS CHANGED AND THE ORIGINAL INTERFACE BOARD WAS REINSTALLED. THE ERROR MESSAGE DID NOT OCCUR AGAIN. | CDD411B86E476156 | No | THE INSTRUMENT WAS POWERE ON, THE MONITOR WOULD CALIBRATE THE SENSOR AND THEN IT WOULD POST A GENERAL ERROR | THE INTERFACE BOARD WAS CHANGED AND THE SAME GENERAL ERROR CODE WAS OCCURRING. THE MAIN CPU UNIT WAS CHANGED AND THE ORIGINAL INTERFACE BOARD WAS REINSTALLED. THE ERROR MESSAGE DID NOT OCCUR AGAIN. | Defect | hw | PCB, Main CPU | PCB, Main CPU
Tosca-500 |  |  | CDD594B3E427151F | No |  |  | Defect | Consumables | Electrode
ABL90 FLEX | Kunde meldet das der ABL 90 des öftern nach LP bzw SC Tausch verlangt. | Gerät intensiv gespült und Überprüft. LP ersetzt. Konfiguration überprüft, Software nochmals Installiert, brachte keinen Erfolg. Eine Erneute Komplette Software Install war nicht mehr möglich Gerät wurde benötigt. Gerät ist Funktionsfähig. | SWO-1390 | No | Customer reports that the ABL 90 often asks for LP or SC exchange. | Device intensively rinsed and checked. LP replaced. Configuration checked, software reinstalled, did not succeed. A Re-complete software install was no longer possible device was needed. Device is functional. | Cleaned | Fluidic pathway | SC and SP replacement required?
ABL90 FLEX | Error 2, Größe Datenbank 74MB Fehler trat nach Notstromtest auf. | Alte roma.ini und roma.db ins Verzeichnis DaSi_Datenbank verschoben. Alte Konfiguration eingespielt. | SWO-1320 | No | Error 2, size database 74MB error occurred after emergency power test. | Old roma.ini and roma.db moved to the DaSi_Database directory. Old configuration. | Defect | SW | Roma.ini | Data bank
ABL90 FLEX | Fehlersuche | Kontakte ersetzt, im Service Menü div. Justagen und Funktionen durchgeführt geprüft. Neustart, Kalibration, QC-kontrolle. Auf Kundenwunsch Layout Konfigurationen angepasst. | SWO-1409 | No | Troubleshooting | Contacts replaced, checked in the service menu div. adjustments and functions performed. Restart, calibration, QC control. Adapted layout configurations at customer request. | Defect | HW | Solution Pack PCB id
ABL90 | Comprobar unidad OXI, limpiar la cubeta de medida de la Hb, calibrar y probar con controles. | Limpiar la cubeta de medida de la Hb. | CDD576DDE42E13C9 | No | Check OXI unit, clean Hb measuring bucket, calibrate and test with controls. | Clean the Hb measuring bucket. | Cleaned | Fluidic pathway | Oxi Parameters
ABL90 | Comprobar el circuito neumático, ajustar sensores y bombas, apagar y encender el equipo. | Apagar y encender el equipo. | CDD54EF0E42313C6 | No | Check the pneumatic circuit, adjust sensors and pumps, turn off and on the equipment. | Turn the equipment off and on. | Adjustment | Power off/on | Wetability ?
ABL825 FLEX | Folgebesuch weg. Ersatzteilbeschaffung. | AutoCheck Modul erneuert. | SWO-1424 | No | Follow-up visit away. Spare parts. | AutoCheck module renewed.
ABL825 FLEX | pO2 AQC Level 2 grenzweertig | Messkammer gesäubert, Elektroden kontrolliert, Nullpunktkalibrierung durchgeführt tHb-Kal durchgeführt | SWO-1362 | No | pO2 AQC Level 2 grenzweertig | Measuring chamber cleaned, electrodes controlled, zero-point calibration performed tHb-Kal performed
ABL90 FLEX | Op 29 september, om 08:00 uur, SC (lot: R0260) werd geplaatst. Hierna diverse calibratieproblemen zijn ontstaan met de lactaat en de PCO2. Op ons advies werd 12:40 uur een nieuwe SC (R0263) geplaatst. Deze SC gaf ook calibratieproblemen (te lage gevoeligheid) bij de PCO2. Op 30 september om 14:30 uur werd (op ons advies) opnieuw een nieuwe SC (lot: R0263) geplaatst maar deze SC blijft ook calibratiteproblemen (te lage gevoeligheid) geven bij de PCO2. | Datalogs heb bekeken, zie dat alle calibraties sinds 29 9 zijn afgekeurd wegens vooral te lage gevoeligheid bij de PCO2. Zodoende het volgende heb uitgevoerd: - Flush procedure compleet heb uitgevoerd, geen afwijking heb geconstateerd. - Inletgasket, inletconnector en oxivalve tube heb vervangen, heeft niet geholpen. - Nieuwe RP lot; RB-10 (uit eigen voorraad) heb geplaatst, heeft niet geholpen. - Huidige SC heb doorgespoten en 45 min. heb laten conditioneren, heeft niet geholpen. - Een nieuwe SC lot: R0573 (uit eigen voorraad) heb geplaatst, nu is de PCO2 gevoeligheid wel goed (85 %). Defecte SC's 1X lot: R0260 en 2X lot: R0263 heb meegenomen voor onderzoek. Opmerking 1: Glucose Lactaat gevoeligheden zijn nu te laag, ze hebben meer tijd nodig om te stabiliseren. Opmerking 1: De vervanging logs heb gekopieerd t.b.v. crediteren van defecte SC's. | SWO-1534 | No | On September 29th, at 08:00 pm, SC (Lot: R0260) was placed. Following various calibration problems have arisen with the lactate and the PCO2. On our advice 12:40 hours a new SC (R0263) was placed. This SC also gave calibration problems (too low sensitivity) to the PCO2. On 30 September at 14:30 pm (on our advice) again a new SC (lot: R0263) was placed, but this SC also keeps calibratite problems (too low sensitivity) at the PCO2. | Data logs, see that all calibrations since 29 9 have been frowned on due to especially low sensitivity in the PCO2. So the following have performed:-Flush procedure have performed completely, no deviation have been observed. -Inletgasket, Inletconnector and oxivalve tube have replaced, did not help. -New RP lot; RB-10 (from its own stock) has not helped. -Current SC have been sprayed and 45 min. Have it conditioning, has not helped. -A new SC lot: R0573 (from own stock) have placed, now the PCO2 sensitivity is fine (85%). Defective Sca'S 1X lot: R0260 and 2X lot: R0263 have taken for research. Note 1: Glucose lactate sensitivities are now too low, they need more time to stabilize. Note 1: The replacement logs have been copied for crediting defective Scs.
ABL90 FLEX | Tubuing for Valve und Waste Tubing undicht | Tubuing for Valve und Waste Tubing ausgetauscht | SWO-1553 | No | Tubuing for Valve und Waste Tubing undicht | Tubuing for Valve und Waste Tubing ausgetauscht | Defect | HW | Oxi Module | Pinch Valve tube
ABL90 |  |  | CDD6F08F5A2C6971 | No |  |  | Defect | HW | Sensor cassette PCB
ABL90 FLEX | ABL auf Undichtigkeit überprüft Abfallschlauch an waste - connector undicht | Jahreswartungskit eingebaut. ABL manuell gespült. Inlet probe erneuert (leicht verbogen). Nach Startphase QC Probleme bei pO2 und inhomgene QC 3 Lösung. Lösungspack gewechselt. Kalibration und QC's durcheführt. ABL wieder einsatzbereit. Kunden LP für Lager dagelassen | SWO-1483 | No | ABL for leak ingested waste hose at waste - connector leaks | Annual maintenance kit. ABL manually rinsed. Inlet probe renewed (slightly bent). After start-up qc problems with pO2 and inhomgene QC 3 solution. solution pack. calibration and QC's. ABL ready for use again. Customers LP left for warehouse | Defect | HW | Waste tube
ABL90 | Sustitucion del casete de electrodos. cal y qc, | Sustitucion del casete de electrodos. cal y qc, | CDD74B715A12DD9B | No | Replacement of the electrode cassette. lime and qc, | Replacement of the electrode cassette. lime and qc, | Defect | SC | Reference
ABL80 | se procedio a sustituir manifold completo. tras esto se siguen presentando problemas en el CL en ciclo de calidad (QC· bajo) se realizan diversas pruebas con varios consumibles. se extraen datos y se traslada el caso a soporte. equipo sin el CL, resto de parámetros ok. |  | CDDE1CF81E022E4C | No | proceeded to replace full manifold. after this, problems continue to arise in the quality cycle CL (QC- low) various tests are carried out with various consumables. data is extracted and the case is transferred to support. equipment without the CL, other parameters ok.
ABL815 FLEX | Antrieb für Ampullenbrecher defekt | Motor ersetzt | SWO-1515 | No | Drive for ampoule crusher defective | Engine replaced
ABL827 FLEX |  |  | CDD6DB5D5A1456C5 | No
ABL825 FLEX | Checked data logs. Error code 376 on 1 point calibrations. | Replaced electrode pump tubing on EI Met module. Cleaned top of Lactate electrode with alcohol pad. Remembrane electrode. Also, remembraned 5 electrodes(pO2, NA, CI, K+ and CA) that were due in 2 days. | CDD704565A028D45 | No | Checked data logs. Error code 376 on 1 point calibrations. | Replaced electrode pump tubing on EI Met module. Cleaned top of Lactate electrode with alcohol pad. Remembrane electrode. Also, remembraned 5 electrodes(pO2, NA, CI, K+ and CA) that were due in 2 days.
ABL827 FLEX |  |  | CDE2142B805965D2 | No
Tosca-500 |  |  | CDD815D4D007DBEA | No |  |  | Defect | HW | Calibration
ABL80 | Revisión de la cassette de electrodos (limpieza de pins y contactos de la misma), aleatoriamente entran los valores en los Ciclos de Calidad (5 Ok x 1 mal) Se procede al cambio del Pack de soluciones por estar al punto de agotarse y se realizan diferentes ciclos de calidad y sigue repitiendo el problema del Cl por lo que se sustituye la cassette por una nueva. Se realizan diferentes verificaciones del correcto funcionamiento del equipo | Cambio de la SC y SP (este por estar agotado). | CDD813C1D003EC14 | No | Review of the electrode cassette (cleaning of pins and contacts of the same), randomly enter the values in the Quality Cycles (5 Ok x 1 wrong) We proceed to change the Solution Pack because they are at the point of exhaustion and different quality cycles and follow and repeating the Cl problem so the cassette is replaced with a new one. Different checks are carried out on the proper functioning of the equipment | Change of SC and SP (this one because it is exhausted).
ABL80 FLEX BASIC | QC Fehler in pCO2 und pO2 in Lvl.2 AC4+, pO2 falsch hoch, pCO2 falsch niedrig. Fehler beim Durchführen der QC. | 2.Kal. und QCs durchgeführt, QCs i.O. Konfiguration der QC Temp. auf Eingabe bei jeder Messung geändert. Anwender eingewiesen. | SWO-1640 | No | QC error in pCO2 and pO2 in Lvl.2 AC4+, pO2 incorrectly high, pCO2 incorrectly low. Failed to perform QC. | 2.Kal. QCs, i.O. configuration of the QC Temp. changed to input with each measurement. users.
ABL80 FLEX BASIC | Gerät überprüft und Kanäle kalibriert -> kein Fehler am Gerät feststellbar. SP80 erneuert und Qualitätskontrollen durchgeführt -> ohne Erfolg. Weitere SP80 werden durch GS12-Kollegen (am gleichen Tag, wenn möglich) geliefert, erneuert und getestet. |  | SWO-3736 | No | Device checked and channels calibrated -> no error can be detected on the device. SP80 renewed and quality controls carried out -> to no avail. Additional SP80s are delivered, renewed and tested by GS12 colleagues (on the same day, if possible).
ABL90 | Removed the cover and checked for loose connections. Replacing the computer for diagnostic purposes. | Replaced the CPU unit. | CDD7F6B0D0577F91 | No | Removed the cover and checked for loose connections. Replacing the computer for diagnostic purposes. | Replaced the CPU unit. | Defect | HW | CPU
ABL90 | Opened the bay and observed fluid around the hemolyzer. | Removed the oxi module and cleaned it up. Replaced the valve tubing. | CDD838B3D0397F95 | No | Opened the bay and observed fluid around the hemolyzer. | Removed the oxi module and cleaned it up. Replaced the valve tubing. | Defect | HW | Oxi Module | Pinch Valve tube
ABL90 | performed backflush | replaced the pinch valve tubing, found the waste line leaking so I replaced the waste line. | CDD7E1D9D04C2602 | No | performed backflush | replaced the pinch valve tubing, found the waste line leaking so I replaced the waste line. | Defect | HW | Oxi Module | Pinch Valve tube
ABL90 FLEX | Keine Messung möglich | Lösungspack und Sensorkassette erneuert | SWO-1596 | No | No measurement possible | Solution pack and sensor cartridge renewed | Defect | SW | LS error
ABL90 FLEX | Ableitstrom am Lösungspack | Wastschlauch am Anschluss gereinigt | SWO-1630 | No | Discharge current on the solution pack | Wast hose cleaned at the connection | Defect | HW | Waste tube
ABL800Basic |  |  | CDD80D18D0634E7A | No
ABL800 BASIC | PH CA kal Probleme | PH Elektrode erneuert und membraniert, Kal, s QC,s | SWO-1685 | No | PH CA kal problems | PH electrode renewed and membraneed, Kal, s QC,s
ABL827 FLEX |  |  | CDE0FCB90A1921BA | No
ABL837 FLEX |  |  | CDD80555D0462E4E | No
ABL827 FLEX |  |  | CDD9756C465CC1C2 | No
AQT90 FLEX |  |  | CDD813ABD030C66F | No
ABL80 FLEX Basic version | Realizacion de limpieza con sp de limpieza. Ajuste del cero electronico. Sustitucion de casete y sp cal y qc | Realizacion de limpieza con sp de limpieza. Ajuste del cero electronico. Sustitucion de casete y sp cal y qc | CDD9729F465FDDA8 | No | Cleaning with cleaning sp. Setting the electronic zero. Substitution of cassette and sp cal and qc | Cleaning with cleaning sp. Setting the electronic zero. Substitution of cassette and sp cal and qc
ABL90 | The inlet probe got bent while running samples. | Replaced the inlet probe for the customer. | CDD93AEF464CBEEA | No | The inlet probe got bent while running samples. | Replaced the inlet probe for the customer. | User Error | Probe bend
ABL827 FLEX |  |  | CDD97BAA462DC1C7 | No
ABL825 FLEX | pH Electrode has blood residue on end. | Cleaned Measuring Chamber. Clean end of pH Electrode. Replaced Reference Membrane. | CDD9229D4611D5D5 | No | pH Electrode has blood residue on end. | Cleaned Measuring Chamber. Clean end of pH Electrode. Replaced Reference Membrane.
ABL825 FLEX |  |  | CDD941E3462425EE | Yes |  |  | Re classified RMED | Logistical upSWO
ABL825 FLEX | DID COMPLETE OXIMETRY TEST AND FOUND THE HALOGEN LAMP TO BE AT 2900 .NEW LOT OF QC HAS FC02 LEVEL 2 S7745 LOT 582 RUNNING AT 91.6 | ADJUST THE LAMP TO 3450.PERFORMED THB CAL IN SERVICE AND THE CUSTOMER MODE.QC RUNNING AT 91,7.CUSTOMER ADJUSTED THE QC RANGE | CDD947DC4618BDB9 | No | DID COMPLETE OXIMETRY TEST AND FOUND THE HALOGEN LAMP TO BE AT 2900 . NEW LOT OF QC HAS FC02 LEVEL 2 S7745 LOT 582 RUNNING AT 91.6 | ADJUST THE LAMP TO 3450.PERFORMED THB CAL IN SERVICE AND THE CUSTOMER MODE. QC RUNNING AT 91,7.CUSTOMER ADJUSTED THE QC RANGE
AQT90 FLEX |  |  | CDD991B646527F8A | No
TCM-COMBIM |  |  | CDEDE133A6463565 | No |  |  | Defect | HW | Battery module
TCM-COMBIM |  |  | CDDCF888A85C6C72 | No |  |  | Defect | Consumables | Electrode | Sensor 84
TCM400 | Was not able to duplicate unresponsive screen. pO2 calibration would fail due to temperature error. | Reseated cable to touch screen and returning to replace pO2 electrode | CDDC0D87321DD56D | No | Was not able to duplicate unresponsive screen. pO2 calibration would fail due to temperature error. | Reseated cable to touch screen and returning to replace pO2 electrode | Defect | hw | Touch screen
ABL90 FLEX | Alle OXI Parameter fehlerhaft. OXI Spectrum mismatch, SHb ermittelt und Trübung zu hoch. | System mit Proteinentferner gereinigt, tHb Kal. durchgeführt. | SWO-1742 | No | All OXI parameters are faulty. OXI Spectrum mismatch, SHb determined and turbidity too high. | System cleaned with protein remover, tHb Kal. Performed. | Cleaned | Fluidic pathway | Oxi Parameters
ABL90 FLEX | Häufig Intervention erforderlich. Einweisung, dabei festgestellt, dass Einlassnadel verbogen ist. | Einlassnadel aus Kundenbestand erneuert. | SWO-1851 | No | Frequently intervention required. Instruction, thereby determined that inlet needle is bent. | Inlet needle from customer base renewed. | Replaced | Probe bend
ABL90 FLEX | Lösungspack ist abgelaufen. Undichtigkeit an Abfallschlauch und Tubing for valve festgestelltfestgestellt. | Jahreswartungskit eingebaut. Lösungspack aus GS Bestand erneuert. Kalibration und QC's durchgeführt. | SWO-1835 | No | Solution pack has expired. Leakage detected on waste hose and tubing for valve. | Annual maintenance kit. Solution pack from GS stock renewed. calibration and QC's. | Defect | HW | Waste tube
ABL90 | SWAPPED SOLUTION PACK WITH KNOWN WORKING UNIT (NO HELP)COMPLETED ALL SERVICE FUNCTIONS EXCEPT SOLUTION PACK CHECK DID NOT PASS | REPLACED PROBE AND GASKET, THOROUGHLY FLUSHED SYSTEM. COMPLETED ALL STARUP PROCEDURES AND SERVICE TEST OF SOLUTION PACK PASSED | CDDE6A411E432998 | No | SWAPPED SOLUTION PACK WITH KNOWN WORKING UNIT (NO HELP)COMPLETED ALL SERVICE FUNCTIONS EXCEPT SOLUTION PACK CHECK DID NOT PASS | REPLACED PROBE AND GASKET, THOROUGHLY FLUSHED SYSTEM. COMPLETED ALL STARUP PROCEDURES AND SERVICE TEST OF SOLUTION PACK PASSED | Cleaned | Fluidic pathway
ABL90 | Comprobación del estado de funcionamiento del equipo, se comprueba que el módulo de entrada está roto con la aguja torcida y el gasket en mal estado, se sustituye el módulo de entrada completo con el Inlet probe y el gasket, se sustituye el SC por uno nuevo ya que el que está puesto tiene la fecha sobrepasada, se calibra el equipo y verificamos el funcionamiento final con QC. Se da formación de usuario a las dos supervisoras. |  | CDDD6F01A83957B1 | No | Checking the operating status of the equipment, it is checked that the input module is broken with the twisted needle and the gasket in bad condition, the complete input module is replaced with the Inlet probe and the gasket, the SC is replaced with a new one since the one that is it has the date exceeded, the equipment is calibrated and we verify the final operation with QC. The two supervisors are trained as a user. |  | Replaced | Probe bend
ABL90 | Comprobación del estado de funcionamiento del equipo, sustitución de la power supply, instalación de un SC nuevo, verificación final de funcionamiento con QC. |  | CDDD744DA84F57B5 | No | Checking the operating status of the equipment, replacing the power supply, installing a new SC, final operating check with QC. |  | Defect | HW | Power supply
ABL815 FLEX | Gerät kontrolliert, Gasventilmembranen und HJW wurden bereits durchgeführt | El Met Modul probeweise erneuert, Elektroden kontrolliert, Luftauslass gespült | SWO-1866 | No | device has been inspected, gas valve membranes and HJW have already been carried out | El Met module renewed on a trial basis, controlled electrodes, air outlet flushed
ABL815 FLEX | Gerät überprüft: - 0367: pH BG - Modul nicht innerhalb Zeitrahmen gefüllt - 0878: keine Spüllsöung während Kalibrierung | HJW durchgef., Messkammer gereinigt, Gerät kalibriert, QC´s durchgef. | SWO-1848 | No | Device checked: - 0367: pH BG - module not filled within time frame - 0878: no flushing during calibration | HJW, measuring chamber cleaned, device calibrated, QC's through.
ABL800 BASIC | Häufig Glu Nullstromfehler bei Kal. | Reinigen Meßkammer, Kontakte Elektrode membraniert (Fehler tritt bereits länger auf, Elektrode wurde bereits erneuert, siehe SLR) | SWO-1811 | No | Common Glu zero current error at Kal. | Cleaning measuring chamber, contacts electrode membraneed (error has been around for a long time, electrode has already been renewed, see SLR)
ABL835 FLEX |  |  | CDDD74C2A8053AA9 | No
ABL825 FLEX | pCO2 electrode had lost approx half of the dielectric fluid from membrane.Inspected Chamber Gas Connector Electrode. Lot 0580 membrane was leaking fluid from membrane. | Replaced membrane, Checked Gas operation. | CDDD28FEA8190CC7 | No | pCO2 electrode had lost approx half of the dielectric fluid from membrane. Inspected Chamber Gas Connector Electrode. Lot 0580 membrane was leaking fluid from membrane. | Replaced membrane, Checked Gas operation.
AQT90 FLEX | Gerät gemäß Vorgabe überprüft | keine, alles i.O. Blank test und TnT-QC ergaben sehr gute Werte! | SWO-2089 | No | Device checked according to default | none, all i.O. Blank test and TnT-QC yielded very good values!
ABL90 |  |  | CDDE56E91E452AAE | No |  |  | Defect | HW | CF card
ABL90 | Error 521 - Inhemogeneous sample during analysis. The error occurs on both analyzers and only happens on samples with pO2 values above 120. Sample consists of some from of a cell culture media. | Possible cause maybe aspiration rate of sample or relationship between pO2 and SO2. | CDDEA3CC1E5F942D | Yes | Error 521 - Inhemogeneous sample during analysis. The error occurs on both analyzers and only happens on samples with pO2 values above 120. Sample consists of some from of a cell culture media. | Possible cause maybe aspiration rate of sample or relationship between pO2 and SO2. | Re classified RMED | Not Human blood
ABL90 FLEX | Intervention erforderlich Thermistor Kalibrationsfehler => Sensorkassetten defekt | Sensorkassetten erneuert | SWO-1997 | No | Intervention required Thermistor Calibration error => Sensor cartridges defective | Sensor cassettes renewed | Defect | SC | Termistor
ABL90 | Battery Not charging. Stuck at 0%. Tried to reset the jumper on battery but no change on charging after reboot. Found leakage around waste connector unit. Screen Accidentally gets rotated by user because of keyboard attached to the analyzer. | Replaced Battery, Replaced Waste Connector unit and Replaced Waste tubing. Disabled Hot keys under windows to avoid screen rotation. | CDDE66981E02604F | No | Battery Not charging. Stuck at 0%. Tried to reset the jumper on battery but no change on charging after reboot. Found leakage around waste connector unit. Screen Accidentally gets rotated by user because of keyboard attached to the analyzer. | Replaced Battery, Replaced Waste Connector unit and Replaced Waste tubing. Disabled Hot keys under windows to avoid screen rotation. | Defect | HW | Battery module
ABL90 | verified the pack was properly installed | replaced the solution pack id pcb, also sensor casette | CDDE2C1E1E2AF71E | No | verified the pack was properly installed | replaced the solution pack id pcb, also sensor casette | Defect | SC | Fluid transport
ABL90 | I was able to reproduce the error the user was seeing with the unable to pump solutions. I tried replacing the sp. it did not correct the issue. I also tried a back flush to clear the error but did not work. | replacing of the sp and the sc did correct the issue though. | CDDE7DD71E13E091 | No | I was able to reproduce the error the user was seeing with the unable to pump solutions. I tried replacing the sp. it did not correct the issue. I also tried a back flush to clear the error but did not work. | replacing of the sp and the sc did correct the issue though. | Defect | SC | Fluid transport
ABL90 | was able to reproduce the error the user was seeing with the unable to pump solutions. I tried replacing the sp. it did not correct the issue. I also tried a back flush to clear the error but did not work. | replacing of the sp and the sc did correct the issue though. | CDDE7B0D1E28E08B | No | was able to reproduce the error the user was seeing with the unable to pump solutions. I tried replacing the sp. it did not correct the issue. I also tried a back flush to clear the error but did not work. | replacing of the sp and the sc did correct the issue though. | Defect | SC | Fluid transport
ABL90 | I was able to reproduce the error the user was seeing with the unable to pump solutions. I tried replacing the sp. it did not correct the issue. I also tried a back flush to clear the error but did not work. | replacing of the sp and the sc did correct the issue though. | CDDE7CFA1E37E08E | No | I was able to reproduce the error the user was seeing with the unable to pump solutions. I tried replacing the sp. it did not correct the issue. I also tried a back flush to clear the error but did not work. | replacing of the sp and the sc did correct the issue though. | Defect | SC | Fluid transport
ABL90 | FLUSHED SYSTEM AND RAN SERVICE CHECKS FOR THE WET SECTION (PASSED) | REPLACED GASKET AND RESTARTED | CDDE868C1E2929A1 | No | FLUSHED SYSTEM AND RAN SERVICE CHECKS FOR THE WET SECTION (PASSED) | REPLACED GASKET AND RESTARTED | Cleaned | Fluidic pathway | Replaced Gasket
ABL827 FLEX |  |  | CDF37F5BF41C2164 | No
ABL827 FLEX |  |  | CDFC88572E15F161 | No
ABL827 FLEX |  |  | CDDE300E1E05E6A4 | No
ABL835 FLEX | FOUND INLET LIQUID SENSOR CRACKED | REPLACED LIQUID SENSOR NIPPLE AND ALL CONNECTING TUBINGS INCLUDING INLET PROBE TUBING, NIPPLE CONNECTOR TUBING AND ALL TUBINGS ACROSS THE MEASURING SYSTEM. | CDDE839D1E49299D | No | FOUND INLET LIQUID SENSOR CRACKED | REPLACED LIQUID SENSOR NIPPLE AND ALL CONNECTING TUBINGS INCLUDING INLET PROBE TUBING, NIPPLE CONNECTOR TUBING AND ALL TUBINGS ACROSS THE MEASURING SYSTEM.
ABL837 FLEX | Found leftover QC on membrane and in measuring chamber. | Replaced Na electrode and new lot of Na membranes. All QC is in. User to monitor. | CDDE73121E1824E9 | No | Found leftover QC on membrane and in measuring chamber. | Replaced Na electrode and new lot of Na membranes. All QC is in. User to monitor.
ABL825 FLEX | checked data logs for errors, compared PCO2 results to other analyzers, checked electrode chamber and electrode membranes, | removed electrodes and cleaned electrode chambers, replaced PCO2 and reference membranes, repositioned Ph and Dummy seal rings, removed and cleaned dual housing valves, | CDDECC2B1E00334D | No | checked data logs for errors, compared PCO2 results to other analyzers, checked electrode chamber and electrode membranes, | removed electrodes and cleaned electrode chambers, replaced PCO2 and reference membranes, repositioned Ph and Dummy seal rings, removed and cleaned dual housing valves,
AQT90 FLEX | Barcode scanner is not working | Replaced the Barcode scanner with one that Brian Vasquez sent me. | CDDE9A5D1E44BE91 | No | Barcode scanner is not working | Replaced the Barcode scanner with one that Brian Vasquez sent me.
AQT90 FLEX |  |  | CDDE65961E3C9927 | No
Tosca-500 |  |  | CDDFA27D940B3BDD | No |  |  | Not a upSWO
ABL80 CO-OX | 2106 ERROR POSSIBLE FLUID BLOCKAGE. Flushed sample path. Disassembled unit and inspected all tubing. Reassembled unit and installed previously used SC. Obtained pO2 calibration issues. Zeroed instrument. Re-installed cassette. Installed new cassette | Zeroed instrument and flushed sample path. | CDDFC5F49406BAF9 | No | 2106 ERROR POSSIBLE FLUID BLOCKAGE. Flushed sample path. Disassembled unit and inspected all tubing. Reassembled unit and installed previously used SC. Obtained pO2 calibration issues. Zeroed instrument. Re-installed cassette. Installed new cassette | Zeroed instrument and flushed sample path.
ABL90 FLEX | Ca (1070: Sensor-Ansprachefehler), bei Patientenmessungen tHb-Kal. fällig | tHb-Kal. durchgeführt, Kal. und QC durchgeführt, Kunden geraten, wenn das Problem wieder auftritt die Sensorkassette zu erneuern, Reklamationsformular ausgefüllt | SWO-2060 | No | Ca (1070: Sensor address error); patient measurements tHb-Kal. Due | tHb-Kal. carried out, Kal. and QC performed to advise customers if the problem recurs to renew the sensor cartridge, complaint form filled out | Defect | SC | Ca, Lac
ABL90 | Lavagem do circuit, substituição do tubo da válvula, da entada e da sonda, establização e verificaçãoo do funcionamento. |  | CDDFCCB6942AB305 | No | Washing the circuit, replacing the valve tube, the pipe and the probe, establishing and checking the operation. |  | Defect | HW | Oxi Module | Pinch Valve tube
ABL80 | se procedio a sustituir placa analógica y conector del casete de electrodos. se realizo ajuste del cero. se instalaron nuevos consumibles y el equipo quedo operativo. se extraen datos para su envio a soporte. se realizan qC correctamente. |  | CDF79001562F518B | No | proceeded to replace analog plate and electrode cassette connector. zero adjustment. new consumables were installed and the equipment was operational. data is extracted for shipment to support. are performed qC correctly.
ABL800 BASIC | Glu Nullstromfehler tritt wiederholt auf | Meßkammer gereinigt Nullpunktkalibrierung Austausch Glu -Elektrode | SWO-2134 | No | Glu zero current error occurs repeatedly | Measuring chamber cleaned zero point calibration replacement Glu electrode
ABL825 FLEX | Tried a new lot of rinse solution-No change Checked sealing o-ring on pH and glucose electrode-Appear fine | Cleaned measuring chamber and electrode contacts Replaced pH electrode pH electrode lot #LN03 is defective and under warranty Replaced glucose membrane | CDDF78D794305EF8 | No | Tried a new lot of rinse solution-No change Checked sealing o-ring on pH and glucose electrode-Appear fine | Cleaned measuring chamber and electrode contacts Replaced pH electrode pH electrode lot #LN03 is defective and under warranty Replaced glucose membrane
ABL817 FLEX | Netzwerkanbindung durchgeführt | Netzwerkdatenä am Gerät überprüft und verändert, KIS LIS Schnittstelle konfiguriert, Anbindung mit Herrn Gongoll von der IT kontrolliert. | SWO-2106 | No | Network connection performed | Network data on the device checked and modified, KIS LIS interface configured, connection with Mr. Gongoll controlled by IT.
ABL820 FLEX | Sales plugged in FlexQ module and received errors. Techline says possible blown fuse. Found 7 bent pins on FlexQ Module and a blown fuse on AU PCB. | Fix bent pins on FlexQ Module and replaced 1.5 A fuse on AU PCB. Performed FlexQ Sampler Tray alignment in service program. | CDDF7B979450C1CA | No | Sales plugged in FlexQ module and received errors. Techline says possible blown fuse. Found 7 bent pins on FlexQ Module and a blown fuse on AU PCB. | Fix bent pins on FlexQ Module and replaced 1.5 A fuse on AU PCB. Performed FlexQ Sampler Tray alignment in service program.
AQT90 FLEX | Gerät startet nicht | Daten gesichert, software update erneut durchgeführt, Gerät gestartet, QC durchgeführt, Gerät zum Test erneut gestartet | SWO-2091 | No | Device does not start | Data backed up, software update re-run, device started, QC performed, device restarted for test
AQT90 FLEX | Netzwerkanbindung durchgeführt | Netzwerkdaten überprüft und geändert, KIS LIS Verbindung konfiguriert, Gerät mit Speicherbefehl neu gestartet, mit Herrn Gongoll von der IT die Anbindung überprüft. | SWO-2109 | No | Network connection performed | Network data checked and changed, KIS LIS connection configured, device with memory command restarted, with Mr. Gongoll from IT checking the connection.
TCM-COMBIM | Checked fuses found all ok. | Replaced power supply. | CDDF99CB94576EF1 | No | Checked fuses found all ok. | Replaced power supply. | Defect | HW | Power supply
ABL90 FLEX | Nach Lösungspack Tausch ist der ABL nicht mehr Einsatzbereit. Der Lösungspack wurde nicht richtig aktiviert. | Lösungspack vollständig aktiviert, Fließwege gespült. Startphase. | SWO-2200 | Yes | After solution pack exchange, the ABL is no longer ready for use. The solution pack was not activated correctly. | Solution pack fully activated, flow paths rinsed. Start. | Re classified RMED
ABL90 FLEX | ABL 90 nicht Einsatzbereit, bei der Installation des Lösungspack wurde der Deckel nicht ausreichend nach unten gedrückt. | Lösungspack entsprechend RD aktiviert und Installiert. Kompletter Neustart. Fließwege gespült. | SWO-2205 | No | ABL 90 not ready for use, the lid was not pushed down sufficiently during the installation of the solution pack. | Solution pack according to RD activated and installed. Complete reboot. Flow paths rinsed. | User Error | SP activate | Incorrect activation
ABL90 | Ca and lac also faliing internal QC. Checked QC, activity and cal logs. | Replaced SP and SC. Flushed the analyzer. Cleaned inlet. | CDE78F8858575B08 | No | Ca and lac also faliing internal QC. Checked QC, activity and cal logs. | Replaced SP and SC. Flushed the analyzer. Cleaned inlet. | Cleaned | Fluidic pathway | Gasket SC SP
ABL90 | no investigation needed replacing the wet section board | replaced the wet section board, I also replaced the tubing for valve. | CDE0AFC80A2E6F6D | No | no investigation needed replacing the wet section board | replaced the wet section board, I also replaced the tubing for valve. | Defect | HW | Wet PCB
ABL90 FLEX | Oxi QC Fehler | Fehlversuche Thb Kalibration durchgeführt Fcuv 97,6% Analysator kalibriert und überprüft Interne QC ok | SWO-2298 | No | Oxi QC Fehler | Failed tests Thb calibration performed Fcuv 97.6% analyzer calibrated and checked internal QC ok | Adjustment | Oximoduler | tHb cal
ABL90 FLEX | pCO2-Problem Empf. Zu niedrig (40,6%) | ABL überprüft, Einlassdichtung und Einlasssonde erneuert. SC und LP erneuert. Keine Verbesserung. Vergleichskassette aus ZNA eingesetzt => Kal und QC OK! Da keine Kassetten mehr im Haus vorhanden, Gerät mitgenommen. | SWO-2290 | No | pCO2 problem Empf. Too low (40.6%) | ABL checked, inlet seal and inlet probe renewed. SC and LP. No improvement. Comparison cassette from ZNA used => Kal and QC OK! Since there are no more cassettes in the house, device is taken with you. | Defect | SC | pCO2
ABL825 FLEX |  |  | CDE0EE830A4948D8 | No
ABL800 BASIC | Ca Kalibrier Fehler | Ca und Ref Elektrode erneuert. Glu und Lac membraniert. Nach mehrere Kalibration Lactat QC Level 2 ohne Beanstandung. | SWO-2783 | No | Ca Calibration Error | Ca and Ref electrode renewed. Glu and lac membraned. After several calibration lactate QC level 2 without complaint.
ABL825 FLEX | thb Kalibrationsfehler | Fehlersuche Halogen Lampe zu hoch eingestellt Halogen Lampe eingestellt Oxi Complete Test OK | SWO-2259 | No | thb Kalibrationsfehler | Troubleshooting Halogen Lamp Set Too High Halogen Lamp Oxi Complete Test OK
ABL820 FLEX | QC Daten überprüft, pCO2 im Level 3 zu niedrig, sporadisch Fehler 0581 - Oxi spectrum mismatch, Oxi-Complet-Test | pCO2 Elektrode gereinigt, Messkammer gereinigt, Lampenspannung eingestellt, Service tHb-Kal, Inlet-Kal, Pumpenkal., Pumpenschlauch Lösung erneuert, Hemolyzer erneuert | SWO-2257 | No | QC data checked, pCO2 at level 3 too low, sporadic error 0581 - Oxi spectrum mismatch, Oxi-Complet test | pCO2 electrode cleaned, measuring chamber cleaned, lamp voltage set, service tHb-Kal, inlet-kal, pump scale,pump hose solution renewed, Hemolyzer renewed
AQT90 FLEX |  |  | CDE0ACA60A505D62 | No
ABL80 CO-OX | Upon arrival, verified PO2 not coming in range on QCs botth internal and external. | Replaced all above parts listed. Ran zeroing cassette on unit. Unit now exhibiting V C failure, Solution Pack not accepted. Upon speaking to tech support, was advised to replace both Sensor cassette and solution pack in Tandem. Will order solution pack to complete repair. | CDE200A680295727 | No | Upon arrival, verified PO2 not coming in range on QCs botth internal and external. | Replaced all above parts listed. Ran zeroing cassette on unit. Unit now exhibiting V C failure, Solution Pack not accepted. Upon speaking to tech support, was advised to replace both Sensor cassette and solution pack in Tandem. Will order solution pack to complete repair.
ABL90 FLEX | Tel. Rücksprache durch Frau Rieger und Info Hotline. | pO2 Wert wurde bei einem Patienten nicht ausgegeben, da der CO2 Wert über dem berichtbaren Bereich von 110 mmHg war. Bei Nachmessungen war der CO2 Wert unterhalb. 110 mmHg. Fall ist nicht mehr aufgetreten. Erweiterung des Berichtbaren Bereiches über 110 mmHg gewünscht. Dies wird durch Vertriebskollegen durchgeführt. Meldung 1405 Inkonsistente Datenbank. Applikationssoftware neu auf ABL gespielt. Meldung nicht mehr vorhanden. | SWO-2429 | No | Tel. Consultation by Mrs. Rieger and Info Hotline. | pO2 value was not issued in a patient because the CO2 value was above the reportable range of 110 mmHg. For remeasurements, the CO2 value was below. 110 mmHg. Case no longer occurred. extension of the reportable range over 110 mmHg desired. This is done by sales colleagues. Message 1405 Inconsistent database. application software replayed on ABL. Message no longer exists. | Setup | Reportable range
ABL90 | inspected hemolyzer for leaks and looks like ther is some leaking | Replaced the hemolyze unit. | CDE1FE3B8047DB5E | No | inspected hemolyzer for leaks and looks like ther is some leaking | Replaced the hemolyze unit. | Defect | HW | Hemolyzer leak
ABL90 FLEX | Analysator geöffnet, Kontakte und Verbindungen geprüft. Fehler bleibt bestehen Sensor Interface PCB bestellt. |  | SWO-2471 | No | Analyzer opened, contacts and connections checked. Error remains on sensor interface PCB ordered. |  | Defect | SC | Temperature
ABL90 FLEX | pCO2 Probleme auch mit neuer Kassette | ABL in Niederlassung überprüft. Sensorkassette erneuert. Probelauf. | SWO-2344 | Yes | pCO2 problems even with new cassette | ABL in branch office. Sensor cartridge is renewed. Rehearsal. | Re classified RMED | Double entry
ABL90 FLEX | Weiterhin pCO2-Probleme | Erneut Sensorcasette erneuert. Probelauf über Nacht. Gerät wieder ausgeliefert. | SWO-2372 | Yes | Continued pCO2 problems | Renewed sensor casette. Test run overnight. device is delivered again. | Re classified RMED | Logistical upSWO | See CAS-5316-F5Q8
ABL837 FLEX | Found left over QC on tip of Na electrode. | Replaced measuring chamber. Remembraned Ca and K electrodes.Multiple 2 pt cals. Now reads 159 on level 1 | CDE231D38047FA76 | No | Found left over QC on tip of Na electrode. | Replaced measuring chamber. Remembraned Ca and K electrodes. Multiple 2 pt cals. Now reads 159 on level 1
ABL837 FLEX | Reviewed all data logs. | Cleaned and adjusted electrode connectors. Cleaned measuring chamber.Performed multiple 2pt Cals. Usesr to monitor WDC reports for any improvements. | CDE234618052FA7B | No | Reviewed all data logs. | Cleaned and adjusted electrode connectors. Cleaned measuring chamber. Performed multiple 2pt Cals. Usesr to monitor WDC reports for any improvements.
ABL827 FLEX |  |  | CDE1F276801AB8E4 | No
ABL835 FLEX | QC Level 3 oberhalb Bereich. | Proteinentfernung, Kalibrationen und tHb-Kalibration. | SWO-2457 | Yes | QC Level 3 above range. | Proteinentfernung, Kalibrationen und tHb-Kalibration. | Re classified RMED | Hot line service
ABL90 FLEX | Lösungspack erneut installiert, Service-Test-Programme, Visual Fluid Control, Satz Wartungsteile eingebaut | Lösungspack defekt, kein Lösungspack vor Ort, ABL90 ohne Funktion | SWO-2472 | No | Solution pack reinstalled; Service test programs, Visual Fluid Control; Set of maintenance parts installed | Solution pack defective, no solution pack on site, ABL90 without function | Defect | SP
ABL825 FLEX | Kalventile in Fluidic Modul und Ventile im pH BG- un El Met- Modul überprüft. | Ventilmembran von v2 und Flüssigkeitsventil erneuert (klebten im Ventilgehäuse). Ventilgehäuse vom pH BG- und El Met- Modul gereinigt. Ref-, Lac-, Cl-, Ca-, Ka- und Na-Elektrode neu membraniert. Kontakte gereinigt. Kalibrationen und QC's vorgenommen | SWO-2477 | No | Cal valves in Fluidic module and valves in the pH BG and El Met module checked. | Valve membrane of v2 and liquid valve renewed (glued in valve housing). Valve housing cleaned by pH BG and El Met module. Ref, Lac, Cl, Ca, Ka and Na electrodes. Contacts cleaned. Calibrations and QC's made
ABL80 FLEX CO-OX OSM |  | Sinds enkele weken geeft de analyser de melding dat het lotnummer van QC2 vervallen is. Hierop heeft de klant reeds twee keer een ander RC geïnstalleerd (Lot 205459 + lot 206541 + lot 206805). De foutmelding blijft steeds terugkomen. Andere foutmeldingen worden niet weergegeven. Heb de RC gereset:NOK Heropstarten analyser:NOK Herinladen sw:NOK Voorlopig afwachten op instructies van de fabrikant. | SWO-2582 | No |  | For several weeks the analyser has reported that the lot number of QC2 has expired. The customer has already installed another RC two times (Lot 205459 + lot 206541 + lot 206805). The error message always keeps coming back. Other error messages are not displayed. Have the RC reset: NOK reboot analyser: NOK re-loading SW: NOK tentatively await instructions from the manufacturer.
ABL80 CO-OX | Solution pack gets disconnected due to the latch door. | Replaced the latch brought system to ready. | CDE5C80BE207A6E8 | No | Solution pack gets disconnected due to the latch door. | Replaced the latch brought system to ready.
ABL90 FLEX | Info über -R- Hotline + Telefonische Rücksprache Station. | Nach Tausch der SC durch Anwender regelmäßige Fehlermeldung: Temperatur im Substrat außer Bereich . Temperatur im Service überprüft; Temp. Analysator o.K. Tausch der SC notwendig. keine auf Lager. Vertriebskollege bringt eine vorbei. Alte SC wird reklamiert. Fließwege durchgespült; neuen Quetschschlauch eingebaut. Einlassdichtung erneuert. Mehrere Fehler vorhanden, Pumpenkalibrierfehler und Lösungen können nicht getauscht werden, LP erneuert. pCO2 Fehler nicht mehr vorhanden. LOT Nr. SC: R0252. | SWO-2570 | No | Info about -R- Hotline + Telephone consultation station. | After replacing the SC by users regular error message: Temperature in the substrate out of range . Temperature checked in service; Temp. Analyzer o.K. Exchange of SC necessary. none in stock. Sales colleague brings one over. Old SC is complained about. Flow paths rinsed through; new crushing hose. Inlet seal renewed. Several errors present, pump calibration errors and solutions cannot be replaced, LP renewed. pCO2 error no longer exists. LOT No SC: R0252. | Defect | SC | Termistor
ABL90 FLEX | Der ABL 90 meldet ständig das der LP bzw. die SC getauscht werden soll. Diese Meldungen verschwinden auch nach dem Tausch nicht. | Der ABL Software nochmal neu Installiert und die Datenbank gelöscht. | SWO-2754 | No | The ABL 90 constantly reports that the LP or SC should be replaced. These messages do not disappear even after the exchange. | Reinstalled the ABL software and deleting the database. | Defect | SW
ABL90 FLEX | Halbjahreswartung | Halbjahreswartung durchgeführt. | SWO-2624 | Yes | Halbjahreswartung | half-yearly maintenance. | Re classified RMED | PM
ABL90 FLEX | Software Upgrade | Neue Software 3.1MR1 installiert | SWO-2625 | Yes | Software Upgrade | New software 3.1MR1 installed | Re classified RMED | FAN / TU
ABL90 FLEX | 1066: Sensor Temp. außer Bereich. 1079: Fehler: Sensor-Impedanz Fehler bei Sensorkassetten der Run 229-118 und -126 | Sensor Interface Board eingebaut, Fehler bleibt bestehen. Sensor einer anderen Run eingesezt. Startphase i.O. Rückbau des Sensor Interface Boards. | SWO-2567 | No | 1066: Sensor Temp. out of range. 1079: Error: Sensor impedance error on sensor cartridges of Run 229-118 and -126 | Sensor interface board built-in, error remains. sensor of another run. Start-up phase i.O. dismantling of the sensor interface board. | Defect | SC | Termistor
ABL90 FLEX | Interner Auftrag über Vertrieb und Info -R- Hotline. | Tausch des SC vorab durch Vertriebskollege. SC LOT Nr.: R0235N030. pCO2 QC Grenzverletzung weiterhin vorhanden; Level3. Quetschventilschlauch getauscht, verschmutzt am Hemolyzer, gereinigt. Tausch des LC aus Kundenbestand. QC innerhalb Grenzwerte, nur Temperatur des Substrates außer Bereich. LC war im Kühlschrank! Beobachtung des Analysators. telefonische Rücksprache morgen mit Anwendern. | SWO-2605 | No | Internal order via sales and info -R- hotline. | Exchange of the SC in advance by sales colleague. SC LOT No.: R0235N030. pCO2 QC border violation still present; Level3. Squeeze valve hose, dirty on the hemolyzer, cleaned. Replacement of the LC from customer stock. QC within limits, only temperature of the substrate out range. LC was in the fridge! Observation of the analyzer. consultation with users by telephone tomorrow. | Defect | SC | pCO2
ABL90 FLEX | Intervension erforderlich Senorcassette erneuern | Sensorkassette ausgebaut und gereinigt (Auskristalliersierung) Sensorkassette neu installiert. Kaltstartphase Gerät Messbereit | SWO-2576 | No | Intervension Required To Renew Senorcassette | Sensor cassette removed and cleaned (crystallization) sensor cassette reinstalled. Cold start phase device ready for measurement | Defect | SC | Leak | Ref solution
ABL90 FLEX | Kal Lösung 3 inhomogen Fehlersuche | LP erneuert | SWO-2610 | Yes | Kal Solution 3 inhomogeneous troubleshooting | LP renewed | Re classified RMED | Logistical upSWO
ABL90 FLEX | Messabrüche | Einlassnadel verbogen, getauscht, Startphase durchgeführt Messbereit | SWO-2509 | No | Messabrüche | Inlet needle bent, exchanged, start phase carried out Ready to measure | Replaced | Probe bend
ABL90 |  |  | CDE5DC51E230EEA5 | No |  |  | Defect | SP | Fluid transport
ABL825 FLEX |  |  | CDE838D3CE05F2CD | No
ABL835 FLEX | Erreur 515 neon voltage to high intermittant full oxi test passe mais remarqué Fneon valeur élevé 1.995 valeur maximum 2.0 | remplacé spectrometer full oxi test passe valeur Fneon .405 | CDE5DAC9E2148E1F | No | Error 515 neon voltage to high interlude full oxi test passes but noticed Fneon high value 1.995 maximum value 2.0 | replaced spectrometer full oxi test passes Fneon value .405
ABL835 FLEX | Air valve BG module perc | Remplacé la air valve du module des gas | CDE5CD7BE2058E1B | No | Air valve BG module perc | Replaced the gas module valve air
ABL825 FLEX | Kalibrationen pH und Lyte außer Bereich, sporadisch | Beschichtung der Ref Elektrode fehlt(4105), erneuert. Vermischung von Kalibrationslösungen aufgrund von defektem V2 und undichten Membranen im Fluidic Modul. pH instabil, neuen Dichtring (834-682) aufgesetzt. Lac membraniert, Empf. 70pA Kalibrationen mussten mehrfach durchgeführt werden. QC ohne Beanstandung | SWO-2639 | No | Calibrations pH and lyte out range; Sporadic | Coating of the ref electrode is missing(4105), renewed. Mixing of calibration solutions due to defective V2 and leaky membranes in the Fluidic module. pH unstable, new sealing ring (834-682) set up. Lac membraneed, empf. 70pA calibrations had to be performed several times. QC without complaint
ABL835 FLEX |  |  | CDEE73C91C2133A6 | No
ABL825 FLEX | QC Verletzung im Level 4 bei Cl Kal2 und QC Level 1-3 alle i.O. | Überprüfung QC, Kal und Aktivitäten Datensätze => letzte Membranierung von Referenz am 07.10. (davor am 11.09. und 13.08.); letzte Membranierung von Cl am 24.07. Neumembranierung von Cl => Kal2 i.O. => QC bleibt fehlerhaft Austausch von Cl Elektrode => Kal2 i.O. => QC i.O. Cl bei 93mmol L | SWO-2543 | No | QC Verletzung im Level 4 bei Cl Kal2 und QC Level 1-3 alle i.O. | Review QC, Kal and Activities records => last membraneing of reference on 07.10. (before 11.09. and 13.08.); last membraneing of Cl on 24.07. Remembraneation of Cl => Kal2 i.O. => QC remains faulty replacement of Cl electrode => Kal2 i.O. => QC i.O. Cl at 93mmol L
ABL80 CO-OX | Looking through the data logs I can see that prior to the current problem a sensor cassette was replaced due to high pO2 on solution C8303 on internal QC. After the replacement, the current sensor cassette failed for high pO2 on external QC level 3. Then the customer performed a system cycle, & it failed for N E (No Endpoint). | Performed a zero check, and all results were within manufactures specifications. Ran another system cycle and all levels of external QC. All results were within manufactures specifications. | CDE717A958439176 | No | Looking through the data logs I can see that prior to the current problem a sensor cassette was replaced due to high pO2 on solution C8303 on internal QC. After the replacement, the current sensor cassette failed for high pO2 on external QC level 3. Then the customer performed a system cycle, & it failed for N E (No Endpoint). | Performed a zero check, and all results were within manufactures specifications. Ran another system cycle and all levels of external QC. All results were within manufactures specifications.
ABL90 | The staff was able to intervene by verifying that the patients Ca was normal before resulting it to the floor as critical. | I am opening a CQC and submitting files for RMED to review. | CDE75F0E582DF4E9 | No | The staff was able to intervene by verifying that the patients Ca was normal before resulting it to the floor as critical. | I am opening a CQC and submitting files for RMED to review. | Performance | Ca++ | Techline
ABL90 FLEX | Halbjahreswartung durchgeführt. Gerät überprüft. |  | SWO-2942 | Yes | half-yearly maintenance. device checked. |  | Re classified RMED | PM
ABL90 FLEX | Software Upgrade | 3.1MR1 | SWO-2943 | Yes | Software Upgrade | 3.1MR1 | Re classified RMED | FAN / TU
ABL90 FLEX | Fließwegen im Serviceprogramm überprüft. | Anwender hat Einlaßdichtung auf anraten Fehlersuchroutine erneuert. Dichtung wurde nicht richtig im Halter positioniert, dadurch keine Verbindung zum Flowselector. Einlaßdichtung eingebaut, SP Test im Serviceprogramm durchgeführt. OK. | SWO-2695 | No | Flow paths in the service program. | User has renewed inlet seal on recommended error search routine. Seal was not positioned correctly in the holder, so no connection to the flowselector. Inlet seal installed, SP test carried out in the service program. Ok. | Defect | HW | Inlet Gasket
ABL90 FLEX | Blutmessungen durchgeführt | Gerät in GSWKST abgezogen. | SWO-3018 | No | Blood measurements carried out | device is subtracted into GSWKST. | Unknown | Analyzer to W.shop
ABL90 | Sustituido casete de electrodos. El cliente gasta y tira dos paquetes de soluciones adicionales. |  | CDE764F658552ECA | No | Replaced electrode cassette. The customer spends and throws two additional solution packages. |  | Defect | SC | Wetability
ABL825 FLEX | Checked AutoCheck module: found fan blade assembly blocking carousel movement. | Removed obstacle. | CDE81CBCCE4646E2 | No | Checked AutoCheck module: found fan blade assembly blocking carousel movement. | Removed obstacle.
ABL835 FLEX | Lac membranes are not lasting longer than 10 days, failing calibration | customer replaced the electrode with one sent by FSE; Keith Clark picked up old electrode and will send back to westlake | CDE6F66B581A5360 | No | Lac membranes are not lasting longer than 10 days, failing calibration | customer replaced the electrode with one sent by FSE; Keith Clark picked up old electrode and will send back to westlake
ABL825 FLEX | Herr Egger hat mich angerufen weil der ABL Probleme bei den Kals macht. Bei Eintreffen war der ABL Messbereit aber bei einigen vorherigen Kals waren Driften bei den Elektrolyten festzustellen, welche zu Kalibrationsfehlern Geführt haben. Fehlversuche durchgeführt. | PH BG Modul und ELE MET Modul ausgebaut, Spül- und Luftventil aus beiden Modulen ausgebaut und durch neue Ventile ersetzt. Ventilsitze gereinigt. Gasflaschen erneuert weil bei Gas 1 nur noch geringer Druck vorhanden. Mehrere 2 Punkt und 1 Punkt Kals durchgeführt --> i.O. QCs für alle Level durchgeführt--> i.O. Gerät wieder Messbereit. | SWO-2689 | No | Mr Egger called me because the ABL is causing problems with the kals. Upon arrival, the ABL was ready to measure, but in some previous kals drifts were detected in the electrolytes, which led to calibration errors. failed attempts. | PH BG module and ELE MET module removed, flushing and air valve removed from both modules and replaced by new valves. Valve seats cleaned. Gas cylinders renewed because gas 1 has little pressure. Several 2 point and 1 point kals performed --> i.O. QCs performed for all levels--> i.O. device ready to measure again.
AQT90 FLEX |  |  | CDE82BE8CE47DE05 | No
ABL80 CO-OX | Ran a system cycle & all three levels of external QC as a baseline. Performed a zero check, then ran the system cycle & QC again. I saw no significant change. | Replaced the sensor connector board and spine PCB. Performed a system cycle, then ran two blood samples. First pO2 was 90 & second was 138. Both samples came from the same patient & same syringe. I spoke with my specialist & found that the solution pack provides the reference for the pO2 calculation. I looked at 15 patient samples after the SP change & 8 of the 15 were over 100. I then looked at the 13 prior to the SP change and only 1 was over 100. (102) The customer is out of SP's so I am going to order one over night along with an analog board. | CDE8507ACE05CE04 | No | Ran a system cycle & all three levels of external QC as a baseline. Performed a zero check, then ran the system cycle & QC again. I saw no significant change. | Replaced the sensor connector board and spine PCB. Performed a system cycle, then ran two blood samples. First pO2 was 90 & second was 138. Both samples came from the same patient & same syringe. I spoke with my specialist & found that the solution pack provides the reference for the pO2 calculation. I looked at 15 patient samples after the SP change & 8 of the 15 were over 100. I then looked at the 13 prior to the SP change and only 1 was over 100. (102) The customer is out of SP's so I am going to order one over night along with an analog board.
ABL90 |  |  | CDE8494FCE18AF5E | No |  |  | Defect | SW | Data upgrade failed
ABL90 FLEX | IP Adresse läuft auf DHCP | eine feste IP Adresse eingegeben | SWO-2899 | No | IP address runs on DHCP | a fixed IP address entered | Setup | IP
ABL90 | Upon arrival, found unit not Giving any Oximetry results for All QCs run. | Advised by Technical Support to completely reload all software including operating system. Reloaded. Operating system and Application software. Installed setup. | CDE8589FCE14F0F0 | No | Upon arrival, found unit not Giving any Oximetry results for All QCs run. | Advised by Technical Support to completely reload all software including operating system. Reloaded. Operating system and Application software. Installed setup. | Defect | SW | No OXI QC | Re installed SW
ABL815 FLEX | AQC Modul ausgebaut | Ampullen Scanner einstellen und gereinigt. | SWO-2912 | No | AQC module removed | Ampoule scanner set and cleaned.
ABL827 FLEX |  |  | CDE86361CE63D766 | Yes |  |  | Re classified RMED | Logistical upSWO
ABL815 FLEX | Fehlersuche | Lac Elektrode ersetzt, Kal QC | SWO-2869 | No | Troubleshooting | Lac electrode replaced, Kal QC
ABL805 FLEX | I was here working on another unit and she said she was having issues with this one. Wanda said she had problems with it all weekend. She said this one was doing the same as the other one. Cals were out. The cals and QC are OK now but getting unstable response fault on pH for patient samples. Wanda said the syringe side of the inlet gasket was cracked so I replaced it with customer stock. Also changed the pH seal ring with new style seal ring per FAN 915-318. See SWO B9AN3LUXI for FAN. | Cleaned pH electrode and chamber. Membraned REF with customer stock. Replaced inlet gasket with customer stock. Put on new pH seal ring. | CDE84B9ECE0B8F56 | No | I was here working on another unit and she said she was having issues with this one. Wanda said she had problems with it all weekend. She said this one was doing the same as the other one. Cals were out. The cals and QC are OK now but getting unstable response fault on pH for patient samples. Wanda said the syringe side of the inlet gasket was cracked so I replaced it with customer stock. Also changed the pH seal ring with new style seal ring per FAN 915-318. See SWO B9AN3LUXI for FAN. | Cleaned pH electrode and chamber. Membraned REF with customer stock. Replaced inlet gasket with customer stock. Put on new pH seal ring.
ABL815 FLEX | pO2-Nullwertfehler | Fehlersuche; alle Elektroden entfernt; Elektroden u. MK gereinigt; pO2 u. Ref. neu membraniert; Spüllsg. gew.; div. Kal. = Fehler; neue pO2-Elektrode in gl. Membran getestet; Kal. + QC i.O. | SWO-2866 | No | pO2-Nullwertfehler | Troubleshooting; removed all electrodes; electrodes & MK cleaned; pO2 and Ref. re-membraneed; Rinse sg. w.; div. Kal. = Error; new pO2 electrode in gl. membrane tested; Kal. + QC in the first place
ABL827 FLEX | 657, 529, 672, 657, 535,532, 675, errors in activity log | Replaced inlet gasket and Met II lower liquid sensor | CDE87FFFCE1A42E2 | No | 657, 529, 672, 657, 535,532, 675, errors in activity log | Replaced inlet gasket and Met II lower liquid sensor
ABL815 FLEX | Fließwege überprüft. | Pumpenschläuche ersetzt. Fehlender O-Ring PH Elektrode, die Parameter pO, pCo2, | SWO-2914 | No | Flow paths checked. | pump hoses. Missing O-ring PH electrode, the parameters pO, pCo2,
ABL825 FLEX | QCAmpullen werden nicht aufgebrochen | Brechermotor fährt nicht raus, Spindel am Motor wieder in den Arm eingeführt, Probemessung OK | SWO-2840 | No | QCampullen will not be broken up | Crusher motor does not drive out, spindle on the motor reintroduced into the arm, test measurement OK
AQT90 FLEX | Investigación del problema neumático. Tubo 2 atascado. Sustitución del mismo. Tests de ST Ok. Puesta en marcha, cal tHb, atípicos de blancos. Sustitución unidad óptica. Calibración TnI. Realización blancos. OK. Susitución carcasa trasera (arañazo) | Investigación del problema neumático. Tubo 2 atascado. Sustitución del mismo. Tests de ST Ok. Puestq en marcha, cal tHb, atípicos de blancos. Sustitución unidad óptica. | CDE988964454A852 | No | Investigation of the pneumatic problem. Tube 2 jammed. Substitution of the same. ST Ok tests. Start-up, lime tHb, atypical white. Optical drive replacement. Calibration TnI. White realization. Ok. Rear shell susitution (scratch) | Investigation of the pneumatic problem. Tube 2 jammed. Substitution of the same. ST Ok tests. Well, tq on the go, cal tHb, atypical white. Optical drive replacement.
ABL80 CO-OX | During the prior visit, I discovered that the high pO2 samples all occurred after the most recent solution pack (SP) replacement. | Replaced the SP for diagnostic reasons, compared the system cycle results with the baseline, & the high level pO2 was ten points lower. I ran 4 patient samples & the pO2 ranged from 185-194. Replaced the analog board, & ran 4 samples from the same blood. The pO2 ranged from 147-154. Although this is a significant change, but without any patient information, I can't say if the new values are correct. Customer will run a blood gas on Monday, to verify the fix. | CDE9A4FE44613F7E | No | During the prior visit, I discovered that the high pO2 samples all occurred after the most recent solution pack (SP) replacement. | Replaced the SP for diagnostic reasons, compared the system cycle results with the baseline, & the high level pO2 was ten points lower. I ran 4 patient samples & the pO2 ranged from 185-194. Replaced the analog board, & ran 4 samples from the same blood. The pO2 ranged from 147-154. Although this is a significant change, but without any patient information, I can't say if the new values are correct. Customer will run a blood gas on Monday, to verify the fix.
ABL80 CO-OX | Performed the service zero check with no errors. Reinitialized the sensor cassette sever times but had same the same problem. Replaced the sensor cassette still had the same problem. Both sensor cassettes were lot number 207885. Installed sensor cassette in customers other ABL80 CO-OX. Glucose initialization also failed in that ABL. | Will order new sensor cassette and replace when available.. | CDE9A77D440E7A91 | Yes | Performed the service zero check with no errors. Reinitialized the sensor cassette sever times but had same the same problem. Replaced the sensor cassette still had the same problem. Both sensor cassettes were lot number 207885. Installed sensor cassette in customers other ABL80 CO-OX. Glucose initialization also failed in that ABL. | Will order new sensor cassette and replace when available.. | Re classified RMED | On site installing analyzer is not upSWO
ABL90 FLEX | ABL im Stillstand Durchflußbehinderung Fehler Pumpe, LS | Fehlversuche Überprüfung Fließwege Ventilschlauch erneuert Einlaufverfahren | SWO-3040 | No | ABL in standstill Flow obstruction error pump, LS | Failed tests Inspection flow paths Valve hose renewed inlet procedure | Defect | HW | Oxi Module | Pinch Valve tube
ABL90 |  |  | CDEB1DA1BA034589 | No
ABL90 FLEX | Touchscreen ohne Funktion | Touchscreen getauscht, Startphase, Kal's und QC's durchgeführt, System Messbereit und Touchsreen mit voller Funktion | SWO-3013 | No | Touchscreen without function | Touchscreen swapped, start phase, Kal's and QC's performed, system ready to measure and touchsreen with full function | Defect | HW | Touch screen
ABL90 |  |  | CDE9BA96444914CB | No |  |  | Cleaned | Fluidic pathway | Clot
ABL825 FLEX | Probleme mit Nullpunkt bei Parameter PO2 trotz mehrmaliger nullpunkkalibrierung keine Veränderung. Nach Tausch der PO2 Elektrode durch Herrn Egger (Verwendung der Leihelektrode) war Nullpunkt wieder ok. | Leihelektrode durch neue Elektrode von Techniker ersetzt (Lot LW01) nullpunkkalibrierung durchgeführt. --> i.O. Mehrere Kalibrierungen durchgeführt. --> i.O QCs für alle Level durchgeführt-->i.O. ABL wieder messbereit | SWO-3006 | No | Problems with zero point with parameter PO2 despite several times nullpunk calibration no change. After replacing the PO2 electrode by Mr. Egger (use of the rental electrode), zero point was ok again. | Rental electrode replaced by new electrode replaced by technicians (Lot LW01) zeropunk calibration performed. --> i.O. Multiple calibrations performed. --> i.O QCs performed for all levels-->i.O. ABL ready for measurement again
ABL825 FLEX | Tried cleaning the pCO2 electrode and a new lot of membranes-No change | Replaced the pCO2 electrode | CDE970EB441CA398 | No | Tried cleaning the pCO2 electrode and a new lot of membranes-No change | Replaced the pCO2 electrode
ABL825 FLEX | Going into HOLD and then coming out of HOLD, unit displays Inlet Position error. | Replaced defective Friction Detector. | CDE9A6F9441C1E05 | No | Going into HOLD and then coming out of HOLD, unit displays Inlet Position error. | Replaced defective Friction Detector.
TCM-COMBIM | Customer was suspecting a defective board or module since another sensor was tried. Tried a loaner Sensor 54 on the instrument and left it running for several days without any occurence of the instrument requesting a new membrane. Also tried a Sensor 84 on the instrument, and did not see any membrane change requests over the period of a few hours. Changed membrane on customer's sensor and ran calibration. Calibration failed. Sensor 84 s n 84-32752. | The issue was isolated to the sensor. The sensor will be replaced and shipped directly to the customer as per Customer Service. | CDE9CC764406BB43 | No | Customer was suspecting a defective board or module since another sensor was tried. Tried a loaner Sensor 54 on the instrument and left it running for several days without any occurence of the instrument requesting a new membrane. Also tried a Sensor 84 on the instrument, and did not see any membrane change requests over the period of a few hours. Changed membrane on customer's sensor and ran calibration. Calibration failed. Sensor 84 s n 84-32752. | The issue was isolated to the sensor. The sensor will be replaced and shipped directly to the customer as per Customer Service. | Defect | Consumables | Electrode | Sensor 84
ABL90 FLEX | Überprüfung Analysator | Nach Tausch der SC durch Anwender verlangte der ABL auch den Tausch des LP, obwohl erst seit 4 Tagen in Betrieb. Fehlermeldung nach SC Einsatz Temperatur in Substrat der SC außer Bereich. Beim Service war ABL messbereit, ohne Fehler. Überprüfung des ABL auch im Service o.K. , keine Beanstandung. Tausch des Ventilschlauches, da stark Verschmutzt, durch alten Schlauch, gereinigt. Analysator funktionsbereit. | SWO-3151 | No | Review Analyzer | After the SC was exchanged by users, the ABL also required the replacement of the LP, although only in operation for 4 days. Error message after SC use temperature in substrate of SC out range. In the service, ABL was ready to measure, without errors. Checking the ABL also in the service o.K. , no complaint. Replacement of the valve hose, as heavily polluted, cleaned by old hose. Analyzer ready for use. | No fault found | Temp SC
ABL90 FLEX | Überprüfung Analysator | Bei Eintreffen war Analysator messbereit, ohne Fehler. Überprüfung des ABL90 auch im Service. Keine Beanstandung. Durchspülen der Fließstrecke. Erneuerung der Einlasssichtung, da stark verschmutzt und des Ventilschlauches. Testmessung durchgeführt mit Pico50 Spritze; MNessung in der ZNA und im Kreißsaal. Messergebnisse im Toleranzbereich, keine Beanstandung. Rücksprache mit Oberarzt und Gerätebeauftragten. Klärung der Diskrepanz bei QC Ergebnissen mit Vertriebskollege nächste Woche. | SWO-3148 | Yes | Review Analyzer | Upon arrival, analyzer was ready to measure, without errors. Checking the ABL90 also in service. No complaint. Rinse the flow line. Renewal of inlet sighting, as heavily polluted and the valve hose. Test measurement carried out with Pico50 syringe; MNessung in the ZNA and in the Kreißsaal. Measurement results in the tolerance range, no complaint. Consultation with senior physician and equipment officer. Clarification of the discrepancy in QC results with fellow sales man next week. | Re classified RMED | Courtesy
ABL90 FLEX | pCO2 seit 13:00 Uhr gesperrt. QC wurde vom Anwender wiederholt war OK. Sporadisch Probleme in den Fließwegen und QC Verletzungen im pCO2. Anwender hat SC gegen 17:00 getauscht. | pCO2 nach Startphase vom Labor entsperren lassen. Verbogene Einlaßnadel ersetzt. SC-zurückgetauscht da neue SC pCO2 Empf. 18%. Funktion z. Zt. OK. | SWO-3242 | No | pCO2 has been closed since 1 p.m. QC was repeated by the user was OK. Sporadic problems in the flow paths and QC injuries in pCO2. User has swapped SC for 17:00. | pCO2 after start-up phase from the laboratory. Replaced bent inlet needle. SC-backed since new SC pCO2 empf. 18%. Function currently OK. | Defect | SC | pCO2
ABL90 FLEX | Häufig Probenfehler bei Kap. Messung Probe wir z.t. abgebrochen bzw. Meldung inhomogene Probe Kunde hatte gerade neue SC eingesetzt | Fehlersuche Fließwege gespült Mehrere Testmessungen Spritze u. Kapillare ohne Beanstandung | SWO-3152 | No | Frequently sample errors at Cape. Measurement sample we partly aborted or message inhomogeneous sample customer had just inserted new SC | Troubleshooting Flow paths flushed Several test measurements syringe and capillary without complaint | Cleaned | Fluidic pathway
ABL90 |  |  | CDF1C3D47E546186 | No |  |  | Cleaned | Fluidic pathway
ABL90 FLEX | Flüssigkeitstransportfehler und Probenabbruch | Gerät überprüft, Nadel stark verbogen, getauscht aus Kundenbestand,, Startphase, Kal und QC's durchgeführt, System Messbereit | SWO-3216 | No | Liquid transport error and sample demolition | Device checked, needle strongly bent, exchanged from customer stock,, start phase, Kal and QC carried out, system ready to measure | Replaced | Probe bend
ABL90 | Checked data logs for errors. Found multiple failures on sensor cassettes. | Replaced customer provided sensor cassette. Verify analyzer completes start-up on new sensor. Complete credit form for customer. | CDEDD5F4A6413DA6 | Yes | Checked data logs for errors. Found multiple failures on sensor cassettes. | Replaced customer provided sensor cassette. Verify analyzer completes start-up on new sensor. Complete credit form for customer. | Re classified RMED | Courtesy
ABL90 | Checked data logs for errors. Found multiple errors with different sensor cassettes installed on analyzer. | Replaced customer provided sensor cassette. Verify operation after replacement. Fill out credit form for defective sensor cassette. | CDEDD2F4A6143DA4 | Yes | Checked data logs for errors. Found multiple errors with different sensor cassettes installed on analyzer. | Replaced customer provided sensor cassette. Verify operation after replacement. Fill out credit form for defective sensor cassette. | Re classified RMED | Courtesy
ABL90 | Analyzer reported solution pack as the problem. Solution pack replaced but did not correct issues. | Analyzer flushed and solution pack replaced. Had to replace sensor cassette provided by customer to correct issues. Credit forms will be completed. | CDEDCE90A64A3DA0 | No | Analyzer reported solution pack as the problem. Solution pack replaced but did not correct issues. | Analyzer flushed and solution pack replaced. Had to replace sensor cassette provided by customer to correct issues. Credit forms will be completed. | Defect | SC
ABL90 | Found some inhomogeneous rinse errors in the activity log | Back flush and replaced the inlet gasket with holder Training with customer, ran QC and samples | CDEB266DBA1AE106 | No | Found some inhomogeneous rinse errors in the activity log | Back flush and replaced the inlet gasket with holder Training with customer, ran QC and samples | Replaced | HW | Inlet Gasket
ABL825 FLEX | checked logs | replaced pH oring with new style, CHANGED CAL FREQUENCY for 1pt cal from 0.5hrs to 2.0 hrs to test | CDEACCBFBA413FF3 | No | checked logs | replaced pH oring with new style, CHANGED CAL FREQUENCY for 1pt cal from 0.5hrs to 2.0 hrs to test
ABL825 FLEX | Druckbild schlecht | Gerät aus; Druckkopf gew.; Gerät ein; Kal. = pH- u. Lytefehler; Ref. neu membraniert; Kal. + QC i.O. | SWO-3161 | No | Print image bad | device off; print head. device; Kal. = pH and lyte errors; Ref. re-membraneed; Kal. + QC in the first place
ABL827 FLEX | 657, 792, 529,675 errors in log. | Replaced inlet mdule with R1604N0005 module. Original module # R2048N0002. | CDEAD70EBA099104 | No | 657, 792, 529,675 errors in log. | Replaced inlet mdule with R1604N0005 module. Original module # R2048N0002.
ABL800 BASIC | Mit Frau Korinth nach telefonischer Anleitung QC chargen eingelesen aber ABL hat trotzdem QC nicht durchgeführt. Vor Ort bei Überprüfung gesehen das Ampulle nicht aufgebrochen wird. Autocheckmodul getauscht. Gerät wurde erst gestern (16.10.14) neu aufgestellt. 902-821R2120N0010 |  | SWO-3219 | No | Ms Corinth read QC's instructions but ABL did not perform QC anyway. On site during inspection, the ampoule is not broken up. Autocheck module. Device was repositioned only yesterday (16.10.14). 902-821R2120N0010
TCM-COMBIM | Customer was suspecting a defective board or module since another sensor was tried. Changed membrane on customer's sensor and ran calibration. Calibration passed. Sensor 84 s n 84-32784 | Changed membrane on sensor | CDEAB3C5BA0BBB47 | No | Customer was suspecting a defective board or module since another sensor was tried. Changed membrane on customer's sensor and ran calibration. Calibration passed. Sensor 84 s n 84-32784 | Changed membrane on sensor | Defect | Consumables | Electrode | Sensor 84
ABL805 FLEX | Einlass Positionierungsfehler | Friction Detector ersetz. | SWO-3223 | No | Inlet positioning error | Friction Detector ersetz.
ABL80 | BAS only shows a credit form processed for one leaking SP back in Aug 2014. |  | CDEE87751C48BCC7 | Yes | BAS only shows a credit form processed for one leaking SP back in Aug 2014. |  | Re classified RMED | Not upSWO
ABL90 FLEX | Analysatorkonfiguration überprüft und berichtigt, Datenübertragung überprüft,System- und Funktionskontrolle, System und Funktion i.O. |  | SWO-3388 | No | Analyzer configuration checks and corrects, data transfer checks,system and function control, system and function i.O. |  | Setup | HIS/LIS
ABL90 FLEX | Vakuumfehler, Intervention erforderlich Datenbank und Flüssigkeitswege geprüft | Anwender hatte bereits mehrfach SP und Dichtungen erneuert und den ABL gespült--> Fehler bleibt bestehen--> System gespült, Anschlussdichtung erneuert, Neustart--> ABL geht erneut in Halt mit diversen Fehlermeldungen ( inhomogene Lösungen )--> nochmals die Flüssigkeitswege gespült--> Neustart | SWO-3296 | No | Vacuum error, intervention required database and fluid paths checked | User had already renewed SP and seals several times and flushed the ABL--> error remains--> system flushed, connection seal renewed, restart--> ABL goes back into hold with various error messages (inhomogeneous solutions )--> again the liquid paths rinsed--> Restart | Cleaned | Fluidic pathway
ABL90 | Customer had multiple issues causing SPs being prematurely replaced. Inspected all tubing. Traces of QC and blood found waste connector where the waste tubing connects. Waste tubing covered with QC and blood on both ends. SW at 3.1. Reinserted SP and SC. | Replaced all tubing and gasket. SW upgraded to 3.1 MR1 | CDEEA97B1C0185CB | No | Customer had multiple issues causing SPs being prematurely replaced. Inspected all tubing. Traces of QC and blood found waste connector where the waste tubing connects. Waste tubing covered with QC and blood on both ends. SW at 3.1. Reinserted SP and SC. | Replaced all tubing and gasket. SW upgraded to 3.1 MR1 | Defect | HW | Waste tube
ABL90 FLEX | Einlasssonde verbogen. Aufgrund der Verbogenen Einlasssonde konnte der Flüssigkeitstransport nicht ordnungsgemäß durchgeführt werden. Sensor Kassette aufgrund pCO2 Empf. Kal Fehler defekt. | Defekte Einlasssonde und Einlassdichtung erneuert. Gerät mit einer Spritze mit warmen Wasser durchgespült um die Verstopfungen im Gerät zu lösen. Sensor Kassette (alte Lot: R0236N156, neue Lot: R246N016) durch Kundeneigene Sensorkassette erneuert, Gerät kalibriert und QC durchgeführt. | SWO-3253 | No | inlet probe bent. Due to the bent inlet probe, the fluid transport could not be carried out properly. Sensor cassette due to pCO2 empf. Kal error defective. | Defective inlet probe and inlet seal renewed. Device rinsed with a syringe of warm water to release the blockages in the device. Sensor cassette (old lot: R0236N156, new lot: R246N016) renewed by customer own sensor cassette, device calibrated and QC performed. | Replaced | Probe bend
ABL90 FLEX | Haltbarkeit der Sensorkassette überschritten. | Sensorkassette durch Kundeneigene erneuert. Absprache mit Anwender getroffen das in Zukunft die Radiometer kapillaren. | SWO-3254 | No | the shelf life of the sensor cartridge has been exceeded. | Sensor cartridge renewed by customer's own. Consultation with users made this in the future the radiometer capillaries. | User Error | SC exp
ABL90 FLEX | De HbF QC's zijn te hoog. | Datalogs heb bekeken, zie dat niet alleen de HBf QC's maar ook de sO2 en de FCOHb QC's zijn te aan de hoge kant. Zodoende in het service menu complete oximetrie test heb uitgevoerd. Vervolgens een kalibratie en een tHb calibratie gedaan. Hierna diverse oximetrie QC's heb uitgevoerd, OK. | SWO-3350 | No | De HbF QC's zijn te hoog. | Data logs, see that not only the HBf Qcs but also the sO2 and the FCOHb Qcs are on the high side. Performed complete oximetry test in the service menu. Then a calibration and a tHb calibration done. After running various oximetry Qcs, OK.
ABL90 FLEX | 1326: Kalibrierung des Thermistor - Sensors gescheitert - Thermistor Fehlfunktion, nach erneuern der SC ( Lot r0232N166) Es erfolgt kein Zugriff auf den LP | Sensorkassette (Lot R0233N073) eingebaut und Sensorkassette (Lot R0232N166) ausgebaut, LP (Lot RH-03) eingebaut und LP (RB-02) ausgebaut, Federkontakte für den LP überprüft, System- und Funktionskontrolle, System und Funktion i.O. | SWO-3283 | No | 1326: Calibration of the thermistor sensor failed - Thermistor malfunction, after renewing the SC ( Lot r0232N166) There is no access to the LP | Sensor cassette (Lot R0233N073) installed and sensor cassette (Lot R0232N166) removed, LP (Lot RH-03) installed and LP (RB-02) removed, spring contacts checked for the LP, system and function control, system and function i.O. | Defect | SC | Termistor
ABL90 |  |  | CDEEB7421C475B63 | No
ABL90 FLEX | Bij de hemolyzer unit is een lekkage geconstateerd. | Zie een lekkage voor de hemolyzer unit. Zodoende oximetrie unit heb gedemonteerd, lekkage was waarschijnlijk veroorzaakt door de oxivalve tubing want ik zie geen lekkage in de cuvette. Lekkage heb schoongemaakt, oxivalve tubing en inletconnector tubing heb vervangen. Inlet gasket heb doorgespoten. ABL90 heb voorzien van nieuwe SC RP (beide waren vervallen). Vervolgens complete oximetrie test en tHb kalibratie heb uitgevoerd. Opmerking: Lactaat parameter is nog niet groen, heeft meer tijd nodig om te stabileseren. | SWO-3352 | No | A leakage has been detected in the Hemolyzer unit. | See a leakage for the Hemolyzer unit. So oximetry unit have disassembled, leakage was probably caused by the oxivalve tubing because I see no leakage in the cuvette. Leaks have been cleaned, oxivalve tubing and inletconnector tubing have been replaced. Inlet gasket have been sprayed. ABL90 have provided new SC RP (both were expired). Then complete oximetry test and have performed tHb calibration. Note: Lactate parameter is not yet green, needs more time to stabilate.
ABL90 FLEX | Wiederholt auftretende Flüssigkeitstransportfehler, Sensorkassette wurde wg. Analysatorstatus Benutzerintervention länger nicht konditioniert, dadurch Kal. u. QC-Fehler (K: 1070, pCO2: 0210 u.w.) LP LOT RH-01 SC LOT 255-207 | LP aus Kd. Bestand erneuert, Flüssigkeitsweg gespült, div. Flüssigkeitstransportfehler bestehen weiter, SC aus. Kd.-Bestand erneuert, Startphase abgewartet, keine weiteren Flüssigkeitstransportfehler, jedoch Na-Backlog Fehler, Kalibrierung wiederholt, danach i.O. QC Level B erst nach wiederholter Messung i.O. (Ca-Fehler - zu hoch). | SWO-3374 | No | Repeated liquid transport error, sensor cassette was wg. Analyzer status User intervention longer unconditioned, thereby Kal. and QC errors (K: 1070, pCO2: 0210 etc.) LP LOT RH-01 SC LOT 255-207 | LP from Kd. Stock renewed, liquid path rinsed, various liquid transport defects continue, SC out. Kd.stock renewed, start phase awaited, no further liquid transport errors, but Na-Backlog error, calibration repeated, then i.O. QC Level B only after repeated measurement i.O. (Ca error - too high). | Defect | SC | Fluid transport
ABL805 FLEX | Fehlersuche | Airventil ersetzt da nicht einwandfrei, pCO2 Membranen im Labor geholt und membraniert, Konfiguration angepasst, Kalibrationen | SWO-3287 | No | Troubleshooting | Airvalve does not replace flawlessly, pCO2 membranes are fetched in the laboratory and membraneed, configuration adapted, calibrations
ABL827 FLEX | Customer has had to keep close eye on Waste Bottle.detector was connected to pins 3&4 upon arrival. Replaced Waste Detector with no luck. | Ordered parts. | CDEE9ACA1C1782E6 | No | Customer has had to keep close eye on Waste Bottle.detector was connected to pins 3&4 upon arrival. Replaced Waste Detector with no luck. | Ordered parts.
ABL800 BASIC | Leckage erkannt, QC Fehler vorhanden | Proteinentfernung durchgeführt Pumpenschlauch im pH BG Modul getauscht Verbindungsschlauch von Messkammer zur Pumpe in beiden Modulen getauscht Gerät kalibriert und überprüft | SWO-3336 | No | Leakage detected, QC error present | Protein removal carried out pump hose in pH BG module exchanged connection hose from measuring chamber to pump exchanged device calibrated and checked in both modules
AQT90 FLEX |  |  | CDEEE4F01C0DF727 | No
ABL80 FLEX Basic version | Chequeo completo de equipo y registro de datos, chequeo de reactivos, conexiones y cables, limpieza de waste drain, waste manifol y circuito y tubos, realización de varias calibraciones y pasados varios qcs de distintos lotes, sustitución de SC y sustitución de SP, pasado QCs saliendo OK. Revisión preventiva | Chequeo completo de equipo y registro de datos, chequeo de reactivos, conexiones y cables, limpieza de waste drain, waste manifol y circuito y tubos, realización de varias calibraciones y pasados varios qcs de distintos lotes, sustitución de SC y sustitución de SP, pasado QCs saliendo OK. Revisión preventiva según procedimiento | CDEFDB2D9242A58C | No | Complete equipment check and data logging, reagent check, connections and cables, cleaning of waste drain, waste manifol and circuit and tubes, performing various calibrations and past several qcs of different batches, SC replacement and SP replacement, past QCs coming out OK. Preventive review | Complete equipment check and data logging, reagent check, connections and cables, cleaning of waste drain, waste manifol and circuit and tubes, performing various calibrations and past several qcs of different batches, SC replacement and SP replacement, past QCs coming out OK. Preventive review as procedure
ABL80 CO-OX | ABL80 leaking clear fluid on to counter. Appears to be coming from one of the clear solution ports on SP none of the colored QC material is leaking. SP is suspect but replaced misc parts from PM kit. | Replaced SP from customer stock. I will process credit form for customer. | CDEFC85C925D9331 | No | ABL80 leaking clear fluid on to counter. Appears to be coming from one of the clear solution ports on SP none of the colored QC material is leaking. SP is suspect but replaced misc parts from PM kit. | Replaced SP from customer stock. I will process credit form for customer.
ABL90 | Unit would fail several cal and qc solution checked during solution pack check. Bi passing the sensor cassette with a piece of tubing and performing the solution pack check, the pack woun pass each time |  | CDF0345E924191E7 | No | Unit would fail several cal and qc solution checked during solution pack check. Bi passing the sensor cassette with a piece of tubing and performing the solution pack check, the pack woun pass each time |  | Defect | SP | Fluid transport
ABL90 | Intermittent error 1324 ( Inhomogeneous rinse solution), Failure of multiply solutions during solution pack check. Unit does pass check when cassette is bi passed. |  | CDF04E5C92307994 | No | Intermittent error 1324 ( Inhomogeneous rinse solution), Failure of multiply solutions during solution pack check. Unit does pass check when cassette is bi passed. |  | Defect | SC | Fluid transport
ABL90 FLEX | Fixe Zielwerte | QC Werte werden vom Labor beanstandet. Überprüfung des Eintrags nach fixen Zielwerten im Analysator. Zeile in roma.ini eingetragen, aber fehlerhaft, korrigiert., gespeichert. Neuen LP aktiviert. | SWO-3438 | No | Fixed target values | QC values are objected to by the laboratory. Check the entry for fixed target values in the analyzer. Line entered in roma.ini, but incorrect, corrected., saved. New LP activated. | Setup | Roma.ini | Assigned values
ABL90 | Temp. errors are primarily for substrate temp. low out of range but error's 1221 & 1066 occur intermittently. Customer had replaced both SC & SP multiple times with different lots with no change. Replacement of sensor PCB did not change symptoms. Replaced wet section PCB and interconnecting cable. Reloaded software and calibrated during initialization intermittent substrate temp error would occur causing brief hold and automatic restart. Replaced consumables and restarted. Bottom and top heaters would not turn on, 0% output. Rebooted, power to heaters came back and had to reload consumables. | Replacement of PCB's, consumables, and reboots evenually proved to return the instrument to functionality. Calibration and QC are in control, instrument operates normally. | CDEFFFD79205FC11 | No | Temp. errors are primarily for substrate temp. low out of range but error's 1221 & 1066 occur intermittently. Customer had replaced both SC & SP multiple times with different lots with no change. Replacement of sensor PCB did not change symptoms. Replaced wet section PCB and interconnecting cable. Reloaded software and calibrated during initialization intermittent substrate temp error would occur causing brief hold and automatic restart. Replaced consumables and restarted. Bottom and top heaters would not turn on, 0% output. Rebooted, power to heaters came back and had to reload consumables. | Replacement of PCB's, consumables, and reboots evenually proved to return the instrument to functionality. Calibration and QC are in control, instrument operates normally.
ABL90 | Checked data logs for errors. Analyzer was getting pCO2, Lac and Na failed to calibrate or pass internal QC. Also notice error code 1324: Inhomogeneous rinse solution in system messages. | Restart calibrations and internal QC until all parameters pass. Open TechlineCase for intermittent 1324: Inhomogeneous rinse solution. | CDF05382924ECC53 | No | Checked data logs for errors. Analyzer was getting pCO2, Lac and Na failed to calibrate or pass internal QC. Also notice error code 1324: Inhomogeneous rinse solution in system messages. | Restart calibrations and internal QC until all parameters pass. Open TechlineCase for intermittent 1324: Inhomogeneous rinse solution. | Defect | SC | pCO2
ABL90 | Checked data logs for errors. Analyzer was getting 1254: Failed to aspirate sample and pO2 and Na failed to calibrate or pass QC. Also notice error code 1324: Inhomogeneous rinse solution in system messages. | Flushed analyzer and reseat solution pack. Started calibrations and internal QC until all passed. Open techline case for intermittent inhomogeneous issues. | CDF04E7A925ACC50 | No | Checked data logs for errors. Analyzer was getting 1254: Failed to aspirate sample and pO2 and Na failed to calibrate or pass QC. Also notice error code 1324: Inhomogeneous rinse solution in system messages. | Flushed analyzer and reseat solution pack. Started calibrations and internal QC until all passed. Open techline case for intermittent inhomogeneous issues. | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway | 1324
ABL90 | UIR says to replace SP. Customer replaced and getting same error. Went into service program fluidics and check status - Fails pressure test - all #'s say 0. | Ran Fluidic tests in service. After the pressure test was done, says passed. Ran the rest of the Fluidics tests in service. All green now. | CDF01C4C925FF9CE | No | UIR says to replace SP. Customer replaced and getting same error. Went into service program fluidics and check status - Fails pressure test - all #'s say 0. | Ran Fluidic tests in service. After the pressure test was done, says passed. Ran the rest of the Fluidics tests in service. All green now. | Adjustment | Service FT pgm
ABL90 | Service Fluidics SP test fails QC3 and sometimes changes to Cal2. Replaced Inlet Gasket - still fails. | Performed all Fluidics tests stil UIR. Performed manual backflush - still UIR. Replaced Inlet Gasket - still fails. Replaced SC - no errors now. | CDF020019251F9D1 | No | Service Fluidics SP test fails QC3 and sometimes changes to Cal2. Replaced Inlet Gasket - still fails. | Performed all Fluidics tests stil UIR. Performed manual backflush - still UIR. Replaced Inlet Gasket - still fails. Replaced SC - no errors now. | Defect | SC | Fluid transport
ABL800 BASIC | Druckbild schlecht | Druckkopf getauscht | SWO-3473 | No | Print image bad | Print head exchanged
ABL825 FLEX | Kalibrierfehler PH und Elektrolyte | Beim eintreffen nur Drift bei CL Kunde hat seit mehreren Tagen Probleme trotz mehrmals neu membraniert Messkammer undicht Referenz Elektrode getauscht Analysator kalibriert und überprüft QC Messung durchgeführt | SWO-3544 | No | Calibration error PH and electrolytes | When arriving only drift at CL customer has performed problems for several days despite several times re-membraneed measuring chamber leaking reference electrode exchanged analyzer calibrated and checked QC measurement
ABL820 FLEX | pH seal ring upgrade had been performed. Found broken nipple under Probe Inlet. | Replaced broken Nipple.Also replaced pCO2 electrode, and tried different Lot of Membranes. | CDF00C0D9249A7F0 | No | pH seal ring upgrade had been performed. Found broken nipple under Probe Inlet. | Replaced broken Nipple.Also replaced pCO2 electrode, and tried different Lot of Membranes.
ABL825 FLEX | checked flow of solutions, checked valves and liquid sensors, checked electrodes and membranes, ca membrane missing white boot at tip | removed and replaced Ca membrane, performed pump cal and liquid senssor adjust, performed cleaning and refill, | CDEFD95D9235E963 | No | checked flow of solutions, checked valves and liquid sensors, checked electrodes and membranes, ca membrane missing white boot at tip | removed and replaced Ca membrane, performed pump cal and liquid senssor adjust, performed cleaning and refill,
TCM400 | Temperature error during calibration of electrode | Replaced electrode. | CDF02DBF926391E3 | No | Temperature error during calibration of electrode | Replaced electrode. | Defect | Consumables | Electrode | E525
ABL80 CO-OX | Failed the SO2 & FO2HB on level 1 of the VK kit during install. I tried a hemolyzer but did not solve it. I am installed the spectrometer and COOX control board because either one could be the failure. Replace both so I can this installed. I used my VK kit. | Replaced the spectrometer & COOX Control board. Ran cals, tHb cal, external QC and VK kit. IVW completed. Setup reports and communication. | CDF10EB6085B8C04 | No | Failed the SO2 & FO2HB on level 1 of the VK kit during install. I tried a hemolyzer but did not solve it. I am installed the spectrometer and COOX control board because either one could be the failure. Replace both so I can this installed. I used my VK kit. | Replaced the spectrometer & COOX Control board. Ran cals, tHb cal, external QC and VK kit. IVW completed. Setup reports and communication.
ABL90 FLEX | Undichtigkeit Schlauch bei Quetschventil | Erneuerung Schlauch - Servicetests - Spülung Flüssigkeitstransport - Erneuerung Nadel und Dichtung notwendig - Funktionsüberprüfung | SWO-3709 | No | Leaky hose at crushing valve | Renewal hose - Service tests - Flushing liquid transport - Renewal needle and seal necessary - Functional check | Cleaned | Fluidic pathway | Wetability
ABL90 FLEX | Regional setings nicht korrekt. | Regional setings auf Germany gestellt, Tastatur eingaben überprüft. Änderungen durch Commit Befehl gespeichert. | SWO-3725 | No | Regional setings not correct. | Regional setings set on Germany, keyboard input checked. Changes saved by commit command. | Setup | Regional setting
ABL90 | At 10:15, analyzer went into UIR for above issues. Tried same steps that resolved issue yesterday - Analyzer went green but had a 1324 error during the process. | Replaced SC from customer stock. All ok. | CDF0FAE90827D34E | No | At 10:15, analyzer went into UIR for above issues. Tried same steps that resolved issue yesterday - Analyzer went green but had a 1324 error during the process. | Replaced SC from customer stock. All ok. | Defect | SC | Wetability
ABL90 | describe the error and list the replacement cassette. don't forget to send in a credit form. |  | CDF1780A082DCE22 | No | describe the error and list the replacement cassette. don't forget to send in a credit form.
ABL5 | Fehlversuche, pO2-Elektrode ausgebaut, Ventiltest | Beide Dichtungen der pO2- Elektrode ersetzt, Kal, QC | SWO-3652 | No | failed attempts; pO2 electrode removed; Valve test | Both seals of the pO2 electrode replaced, Kal, QC
ABL800 BASIC | Überprüfung Analysator | pCO2 starker Statusabfall. Status wieder schnell erholt. Wert wieder o.K. Messkammern und Elektrodenkontakte gereinigt und gebürstet. QC bei pCO2 wieder zielwertig. Bei der pH den unteren Elektrodendichtring mit neuem ersetzt. Einlassdichtungsschlauch erneuert, da stark abgerubbelt. QC und Kal. o.K., Analysator funktionsbereit. | SWO-3655 | No | Review Analyzer | pCO2 sharp status drop. Status recovered quickly. Value again o.K. measuring chambers and electrode contacts cleaned and brushed. QC at pCO2 again. In the pH, replace the lower electrode sealing ring with a new one. Inlet sealing hose renewed, as heavily scrubbed. QC and Kal. o.K., analyzer ready for use.
ABL800 BASIC | Einlasskalibrierfehler | Einlasssonde vom Anwender abgebrochen, keine neue auf Lager; neue eingesetzt. Halbjahreswartung fällig durchgeführt. Analysator funktionsbereit. | SWO-3614 | No | Einlasskalibrierfehler | Inlet probe aborted by the user, no new ones in stock; new ones. half-yearly maintenance due. Analyzer ready for use.
ABL810 FLEX |  |  | CDF152110842E819 | No
ABL825 FLEX | QCpCO2 unterhalb RiLiBÄK-Bereich. Kalibrationen ohne Driften. Fussel an Außenseite pCO2-Membran. | PCO2-Membran gereinigt. Schwarzen O-Ring pH-Elektrode durch blauen ersetzt. Kalibrationen und Qualitätskontrollen in Ordnung. Einweisung Herr Dinkel durchgeführt. | SWO-3737 | No | QCpCO2 below RiLiBÄK range. Calibrations without drifts. Lint on the outside pCO2 membrane. | PCO2 membrane. Black O-ring pH electrode replaced by blue. calibrations and quality controls in order. Instruction carried out by Mr. Dinkel.
ABL837 FLEX | Removed pump rotor from El Met II module and found two rollers that were frozen. | Replaced pump rotor. | CDF141D308228392 | No | Removed pump rotor from El Met II module and found two rollers that were frozen. | Replaced pump rotor.
AQT90 FLEX |  |  | CDF298857E29BFCE | No
AQT90 FLEX |  |  | CDF36E2FF44FDEA6 | No
TCM-COMBIM | Attempted to recreate the issue with the supplied equipment. Issue is intermittent. Tried using a different CombiM module and attempted to recreate the error with it (Service Tip 552962824). Unable to recreate the error despite cycling the power several times and letting the analyzer run overnight. |  | CDF0F81208620645 | No | Attempted to recreate the issue with the supplied equipment. Issue is intermittent. Tried using a different CombiM module and attempted to recreate the error with it (Service Tip 552962824). Unable to recreate the error despite cycling the power several times and letting the analyzer run overnight. |  | Defect | HW | PC module | Sound corrupted
ABL90 | All tests and tubing check good. Replacing the hemolyzer for diagnostic purposes. | Replaced the hemolyzer. | CDF2B4127E1FAD07 | No | All tests and tubing check good. Replacing the hemolyzer for diagnostic purposes. | Replaced the hemolyzer. | Defect | HW | Hemolyzer
ABL90 FLEX | Einlaßnadel und -Dichtung erneuert, Probelauf |  | SWO-3797 | No | Inlet needle and seal renewed, test run |  | Cleaned | Fluidic pathway | Wetability
ABL90 FLEX | Analysator gespült - Servicetests etc. - Funktionsüberprüfung Erneuerungsteile geliefert |  | SWO-3968 | No | Analyzer rinsed - service tests etc. - functional check renewal parts supplied |  | Cleaned | Fluidic pathway
ABL90 FLEX | Gerät geliefert, SC und LP eingesetzt, Kaltstartphase |  | SWO-3776 | No | Device supplied, SC and LP used, cold start phase |  | Unknown | Several issues
ABL815 FLEX | FlexQ Modul gereinigt und justiert. |  | SWO-3866 | No | FlexQ module cleaned and adjusted.
ABL825 FLEX | Referenz Leihelektrode durch neue Referenzelektrode ersetzt. (Leihelektrode wieder Medizintechnik übergeben.) Referenzelektrode neu membraniert, mehrere Kalibrierungen durchgeführt --> i.O. QCs für alle Level durchgeführt --> i.O. ABL wieder messbereit. |  | SWO-3883 | No | Reference rental electrode replaced by new reference electrode. (Loan electrode again handed over medical technology.) Reference electrode re-membraneed, several calibrations performed --> i.O. QCs performed for all levels --> i.O. ABL ready for measurement again.
ABL825 FLEX | Cleaned hemolyzer, cal, alignment and tHb service calibration-No change | Replaced hemolyzer Alignment and tHb calibration in service Complete oxi test-OK | CDF2739F7E104628 | No | Cleaned hemolyzer, cal, alignment and tHb service calibration-No change | Replaced Hemolyzer Alignment and tHb calibration in service Complete oxygen test-OK
ABL827 FLEX |  |  | CDFC89A62E46F164 | No
ABL827 FLEX |  |  | CDF2A7D77E290355 | No
TCM-COMBIM | Analyzer received with touch screen behaving abnormally. Discovered fluid between the two layers of the touch screen while attempting to clean it. After replacing the touch screen, noticed that unit wouldn't calibrate, requesting the membrane to be changed, despite just having changed it. | Replaced touch screen, replaced Sensor 84. | CDF299C57E5BDB1B | No | Analyzer received with touch screen behaving abnormally. Discovered fluid between the two layers of the touch screen while attempting to clean it. After replacing the touch screen, noticed that unit wouldn't calibrate, requesting the membrane to be changed, despite just having changed it. | Replaced touch screen, replaced Sensor 84. | Defect | HW | Touch Screen | Touch Screen
ABL90 FLEX | Bildschirm kalibriert Einstellungen gesichert Analysator kalibriert und überprüft Interne QC OK |  | SWO-3899 | No | Screen calibrated settings backed analyzer calibrated and checked internal QC OK |  | Adjustment | Screen alignment
ABL90 FLEX | Gerät mit warmen Wasser mehrfach durchgespült. Einlassdichtung inkl. Halterung sowie die Einlass Anschlussdichtung durch Kundeneigene ersetzt. Gerät kalibriert und QC durchgeführt. Kal ist OK, QC beim Glu im Level 3 unterhalb unterer Grenze alle anderen Parameter OK. Lösungspack Lot : RH06 erneuert, Gerät kalibriert. QC durchgeführt. Glu. QC Level 3 unterhalb unterer Grenze. Sensor Kassette durch Kundeneigene ersetzt. (Neue Lot :0239, alte Lot: 246) |  | SWO-4007 | No | Device rinsed with warm water several times. Inlet seal incl. bracket as well as the inlet connection seal replaced by customer's own. calibrated and QC performed. Kal is OK, QC at Glu at level 3 below the lower limit all other parameters OK. Solution Pack Lot : RH06 renewed, device calibrated. QC. Glu. QC Level 3 below lower limit. Sensor cassette replaced by customer's own. (New lot :0239, old lot: 246) |  | Cleaned | Fluidic pathway | Clot
ABL90 FLEX | gr. Dichtung ersetzt Service Programme durchlaufen lassen |  | SWO-3980 | No | Gr. Seal replaces letting service programs go through |  | Defect | HW | Inlet Gasket
ABL90 FLEX | Fehlersuche SC erneut installiert - zunächst ohne Beanstandung, Nach einigen Minuten Meldung Temp. Außer Bereich Ersatzgerät Inbetriebnahme genommen mit selbiger SC und LP - Fehler wandert mit, daraufhin am Kundengerät SC und LP wiederholt erneuert - anschl. i.O. |  | SWO-3949 | No | Troubleshooting SC reinstalled - initially without complaint, After a few minutes message Temp. Out of area replacement device commissioning taken with the same SC and LP - error migrates with; then repeatedly renewed on the customer device SC and LP - followed by Ok. |  | Defect | SC | Termistor
ABL825 FLEX | Ref.Elektrode getauscht und membraniert Kal,s und QC,s durchgeführt |  | SWO-3965 | No | Ref.Electrode swapped and membraneed Kal,s and QC,s performed
AQT90 FLEX |  |  | CDFF85A81A2981DA | No
AQT90 FLEX | Fehlersuche. Process Unit ohne Funktion verkantet. Process Unit gangbar gemacht. Neustart, Systemreinigung, LQC durchgeführt. |  | SWO-4019 | No | Troubleshooting. Process Unit without function. Process Unit. Restart, system cleaning, LQC performed.
ABL80 FLEX Basic version | Revisión del estado del SC (Sensor Cassette) se revisa el historial y se sustituye por uno nuevo, se realizan varias comprobaciones con QC´s persistiendo los problemas y precisando del cambio del SP (Solution Pack). Calibración y pasar QC´s. | Sustitución del SC Lote:206668 Ref: 945-790 y SC Lote: 206790 Ref: 944-309 | CDF74A1B56008F98 | No | SC Status Review (Sensor Cassette) is reviewed and replaced with a new one, several checks are performed with QC's persisting problems and requiring the change of the SP (Solution Pack). Calibration and pass QC's. | Replacing the SC Lot:206668 Ref: 945-790 and SC Lot: 206790 Ref: 944-309
ABL90 FLEX | Einlassende und Einlassdichtungen erneuert. Analysator gespült und alten SP 944-197 RI-06 wieder installiert, kein Erfolg. SP 944-197 erneuert mit gleicher Lot., alle Werte und Funktionen wieder o.k. |  | SWO-4168 | No | Inlet ends and inlet seals. Analyzer rinsed and old SP 944-197 RI-06 reinstalled, no success. SP 944-197 renewed with the same lot. all values and functions again ok. |  | Defect | SP
ABL90 |  |  | CDF852C2CC1E635C | No |  |  | Unknown | Fluid transport | Wetability ?
ABL90 | Instrument is not pumping QC3. Customer tried multiple SP's without improvement. Flushing did not help. Found evidence of small leak at hemolyzer inlet all other tubings OK. | Replaced all instrument tubings and inlet probe. Instruemnt came up normally. Ran QC's without error. Verified network settings with IT Paul but instrument is not showing up in Meditek. He will continue to work interface. Instrument operates normally. Filed credit form for replacement of SP lot#RH-14 | CDF779B65646BB75 | No | Instrument is not pumping QC3. Customer tried multiple SP's without improvement. Flushing did not help. Found evidence of small leak at hemolyzer inlet all other tubings OK. | Replaced all instrument tubings and inlet probe. Instruemnt came up normally. Ran QC's without error. Verified network settings with IT Paul but instrument is not showing up in Meditek. He will continue to work interface. Instrument operates normally. Filed credit form for replacement of SP lot#RH-14 | Defect | HW | Oxi Module | Pinch Valve tube
ABL90 FLEX | Leihgerät ABL90 abgezogen. Kundengerät installiert und mit div.Proben überprüft, OK. |  | SWO-4099 | Yes | ABL90. Customer device installed and checked with div.samples, OK. |  | Re classified RMED | Re installation after repair
ABL825 FLEX |  |  | CDFC4C852E0B81FD | No
ABL835 FLEX |  |  | CDF75A4456269E25 | No
ABL827 FLEX |  |  | CDF8965CCC1B01D8 | Yes |  |  | Re classified RMED
ABL835 FLEX | Pumpenkalwerte stark unterschiedlich. (12%) Pumpenschläuche Elektrodenmodule erneuert. |  | SWO-4175 | No | pump calorify values vary widely. (12%) Pump hoses electrode modules renewed.
ABL835 FLEX | O-Ring pH-Elektrode ersetzt. Spüllösungsschlauch zu BG-Ventil um 1,8cm zu lang, gekürzt. Kalwerte Elektrodenmodulschläuche bei 94%, Gegenrichtung 83%. Pumpenschläuche ersetzt. |  | SWO-4173 | No | O-ring pH electrode replaced. Rinsing solution hose to BG valve by 1.8cm too long, shortened. Cal values electrode module hoses at 94%, opposite direction 83%. pump hoses.
ABL80 FLEX BASIC | Csere készülék kihelyezése a javítás időtartamára. SN:301106 | Gábor Pap | SWO-4225 | No | Replacing the replacement device for the duration of the repair. SN: 301106 | Parthiban g
ABL90 FLEX | Sensor Interface PCB zodoende vervangen. Afstelling van de kontaktpinnen gecontroleerd. Na herstart gaat het toestel eerst kalibreren en QC's meten: OK. Hierna de electrode simulator test uitgevoerd. Alle testen worden nu goedgekeurd. | John Smeets | SWO-4303 | No | Replace the Sensor Interface PCB. Adjustment of the contact pins is checked. After restart, the device first calibrating and measuring Qcs: OK. Run the electrode simulator test. All tests are now approved. | John Smeets
ABL90 FLEX | Sensor Interface PCB zodoende vervangen. Afstelling van de kontaktpinnen gecontroleerd. Na herstart gaat het toestel eerst kalibreren en QC's meten: OK. Hierna de electrode simulator test uitgevoerd. Alle testen worden nu goedgekeurd. | John Smeets | SWO-4534 | No | Replace the Sensor Interface PCB. Adjustment of the contact pins is checked. After restart, the device first calibrating and measuring Qcs: OK. Run the electrode simulator test. All tests are now approved. | John Smeets
ABL90 | DATA LOGS SHOW 1343 UNABLE TO PUMP SOLUTIONS AND 1324 INHOMGENOUS RINSE ERRORS.TRIED TO REINSTALL THE SOLUTION PACK THE CUSTOMER COULD NOT GET TO ENGAGE AND THE ANALYZER SAID SOLUTION PACK HAD EXCEEDED THE LIFE ON BOARD THE ANALYZER.I FOUND OUT THAT THE SAME SOLUTION PACK WAS INSATLLED 60 DAYS AGO ON ANOTHER ANALYZER.EXPLAINED TO THE CUSTOMER | FLUSHED ANALYZER ,REPLACED INLET GASKET HOLDER AND INLET PROBE.PERFORMED FLUID TRANSPORT ADJUSTMENTS IN SERVICEREPLACED THE SOLUTION PACK .PERFORMED STARTUP AND CALIBRATIONS. | CDF8CF84CC0C6E9B | No | DATA LOGS SHOW 1343 UNABLE TO PUMP SOLUTIONS AND 1324 INHOMGENOUS RINSE ERRORS. TRIED TO REINSTALL THE SOLUTION PACK THE CUSTOMER COULD NOT GET TO ENGAGE AND THE ANALYZER SAID SOLUTION PACK HAD EXCEEDED THE LIFE ON BOARD THE ANALYZER. I FOUND OUT THAT THE SAME SOLUTION PACK WAS INSATLLED 60 DAYS AGO ON ANOTHER ANALYZER. EXPLAINED TO THE CUSTOMER | FLUSHED ANALYZER ,REPLACED INLET GASKET HOLDER AND INLET PROBE. PERFORMED FLUID TRANSPORT ADJUSTMENTS IN SERVICEREPLACED THE SOLUTION PACK . PERFORMED STARTUP AND CALIBRATIONS. | User Error | SP expired
ABL90 |  |  | CDF9CB9E42597C20 | Yes |  |  | Re classified RMED
ABL90 |  |  | CDF9CA8E420B7C1B | No |  |  | Defect | HW | Printer
ABL90 |  |  | CDF9CC49420E7C24 | Yes |  |  | Re classified RMED
ABL90 |  |  | CDF9CCA6424F7C28 | Yes |  |  | Re classified RMED
ABL90 FLEX | Fehlersuche am Analysator, kein Neustart mehr möglich ->Status Code 2!, Roma.db ist beschädigt->ersetzt gegen Roma.db.empty->Neustart ok, Konfig.Radiance und Pat.Ausdruck angepasst, LP nochmal installiert da inkonsistente Datenbank und Neustart System-und Funktionskontrolle i.O. |  | SWO-4276 | No | Troubleshooting on the analyzer, no restart possible ->Status Code 2!, Roma.db is corrupted->replaced with Roma.db.empty->Restart ok, Config.Radiance and Pat.Expression customized, LP reinstalled since inconsistent database and restart system and Functional control in the |  | Defect | SW | Restart analyzer | Status Code 2
ABL90 FLEX | LP90 ausgetauscht, Funktionskontrolle OK |  | SWO-4208 | No | LP90 replaced, function control OK |  | Defect | SP | Glu
ABL815 FLEX | HjW Kit verbaut. Messkammer gereinigt. Gerät kalibriert und QC durchgeführt. |  | SWO-4305 | No | HjW Kit. measuring chamber. calibrated and QC performed.
ABL835 FLEX | Valeur Qc sur ca++ varie beaucoup et maintenance pr | Effecué la maintenance bi-annuel selon les normes remplacé les tubes de pompe nettoyé pompes et flush v7.v8 verifier electrode o'ringfull oxi test et calibration service THB calibration nettoyé les pompes vérifié electronique calibration zero po2 ,lactate.glucose . remplacé electrode de ca++ | CDF8903ECC1F4375 | No | Qc value on ca varies a lot and maintenance pr | Effecué bi-annual maintenance according to standards replaced pump tubes cleaned pumps and flush v7.v8 check electronico'ringfull oxi test and calibration service THB calibration cleaned pumps verified electronic calibration zero po2, lactate.glucose . replaced ca electrode
ABL805 FLEX | Gabel am amp. Breaker modul befestigt |  | SWO-4326 | No | Fork on amp. Breaker module attached
ABL800 FLEX | Konfiguration eingestellt, div Fragen beantwortet |  | SWO-4324 | No | Configuration set, div questions answered
ABL80 FLEX Basic version | Se revisa la cassette de electrodos y se reinstala, se comprueban con diferentes QC´s (level 1) y fallan el pH y la pCO2, por lo que se sustituye el SC. El mismo presenta fallos de pH y pCO2 por lo que se requiere cambio del SP. Se comprueba el equipo con direfentes QC y esta ocasión todo está dentro de los márgenes. | Cambio del SC y SP. | CDFC6BC42E279D72 | No | The electrode cassette is checked and reinstalled, checked with different QC's (level 1) and pH and pCO2 fail, replacing the SC. It has pH and pCO2 failures so SP change is required. The equipment is checked with QC switches and this time everything is within the margins. | SC and SP change.
ABL80 FLEX CO-OX | Nullpunktabgleich durchgeführt, Sc ersetzt, Kal und QC durchgeführt |  | SWO-4404 | No | Zero-point matching performed, Sc replaced, Kal and QC performed
ABL80 CO-OX | Examined the wet section. Clots in tubing and sensor cassette. | Performed yearly maintennce. Replaced sensor cassette with customer spare. Verified operation of ABL80 COOX. | CDFA0474423D8822 | No | Examined the wet section. Clots in tubing and sensor cassette. | Performed yearly maintennce. Replaced sensor cassette with customer spare. Verified operation of ABL80 COOX.
ABL90 FLEX | Bij aankomst zie diverse errors zoals error 0508, 1324 en 1169. In het service menu oximetrie clot detectie waarde zijn allemaal in het rood. Zodoende flush procedure compleet heb uitgevoerd, waste, oxi module, SC en inlet module los heb doorgespoten, hierna zijn alle oxi clot detectie waarden OK. Vervolgens een calibratie en drie QC levels heb bepaald, ze zijn allemaal OK. | Reza Gharib | SWO-4383 | No | On arrival see various errors such as Error 0508, 1324 and 1169. In the Service menu Oximetry clot detection value are all in red. To complete flush procedure, waste, Oxi module, SC and inlet module have been sprayed, then all oxi clot detection values are OK. Then a calibration and three QC levels have determined, they are all OK. | Reza Gharib
ABL90 FLEX | Bij aankomst zie diverse errors zoals error 0508, 1324 en 1169. In het service menu oximetrie clot detectie waarde zijn allemaal in het rood. Zodoende flush procedure compleet heb uitgevoerd, waste, oxi module, SC en inlet module los heb doorgespoten, hierna zijn alle oxi clot detectie waarden OK. Vervolgens een calibratie en drie QC levels heb bepaald, OK. De eerste bloedmeting wordt meteen afgebroken vervolgens verschijnt het Troubelshooting scherm . Zodoende flush procedure opnieuw heb uitgevoerd, nieuwe oxi tubevalve, connector en inletgasket heb gemonteerd en flush procedure meerdere malen heb uitgevoerd, zonder resultaat. Het probleem is uiteindelijk opgelost door een nieuwe SC te installeren. Alle benodigde logs heb verzameld voor verder onderzoek. SC lot. nr. 272-61 en RP lot. nr. RB-02 heb meegenomen voor onderzoek. | Reza Gharib | SWO-4456 | No | On arrival see various errors such as Error 0508, 1324 and 1169. In the Service menu Oximetry clot detection value are all in red. To complete flush procedure, waste, Oxi module, SC and inlet module have been sprayed, then all oxi clot detection values are OK. Then a calibration and three QC levels have determined, OK. The first blood measurement is immediately interrupted then the Troubelshooting screen appears. So flush procedure has been performed again, new Oxi tubevalve, connector and inletgasket have mounted and flush procedure have performed several times, with no result. The problem is finally solved by installing a new SC. All the necessary logs have been collected for further investigation. SC lot. No. 272-61 and RP lot. No. RB-02 included for research. | Reza Gharib
ABL90 FLEX | Roma.ini 07=1 55=1 22=0 23= 0 eingetragen. |  | SWO-4508 | No | Roma.ini 07=1 55=1 22=0 23= 0. |  | Setup | Roma.ini
ABL90 FLEX | Nach Fehlversuche und testweise getauschten Teilen am Kundengerät, Ersatzgerät (0379N005) ABL 90 aufgestellt. Kundengerät (0425N008) abgebaut und zur Überprüfung mit in Geschäftsstelle genommen. Frau Wichert (Herzkatheter) weiß bescheid. Am Ersatzgerät mehrere Kalibrierungen durchgeführt. QCs für alle Level durchgeführt. ABL90 Messbereit. |  | SWO-4413 | Yes | After failed tests and test-replaced parts on the customer device, replacement device (0379N005) ABL 90 installed. Customer device (0425N008) and taken to the office for inspection. Ms. Wichert (heart catheter) knows. Several calibrations are performed on the replacement device. QCs for all levels. ABL90 Ready to measure. |  | Re classified RMED | ABL removed | to repair
ABL90 FLEX | Fließwege gespült, Einlaßnadel und Dichtung ersetzt, Testmessung mit S7770 durchgeführt OK. tHB Kal durchgeführt. |  | SWO-4414 | No | Flow paths rinsed, inlet needle and seal replaced, test measurement performed with S7770 OK. tHB Kal. |  | User Error | Pre analytical | Clot Probe
ABL90 |  |  | CDFA31444243C955 | Yes |  |  | Re classified RMED
ABL90 | Inspection of battery again verifies it does not funtion correctly. | The battery pack was replaced to address the battery pack failure. | CDFB526DB839A49A | No | Inspection of battery again verifies it does not funtion correctly. | The battery pack was replaced to address the battery pack failure. | Defect | HW | Battery module
ABL90 | COMPROBACION DEL ESTADO DE FUNCIONAMIENTO DEL EQUIPO, SE VERIFICA EL LA BAHÍA DEL SC ESTÁ SUELTA O FLOJA EN SU CAVIDAD, SE DESMONTA EL MODULO COMPLETO PARA SU REVISIÓN, SE SUSTITUYE DE NUEVO LA BAHÍA POR UN ERROR EN SL, SE SUSTITUYE TAMBIEN EL MODULO DE INLET, TUBO DE VALVULA Y EL WASTE TUBE, SE REALIZAN TODOS LOS TEST POR SERVICIO, SE CALIBRA LA THB, SE VERIFICA EL FUNCIONAMIENTO FINAL CON QC. |  | CDFA1E64424717E9 | No | CHECKING THE STATE OF OPERATION OF THE EQUIPMENT, THE SC'S BAY IS CHECKED OR LOOSE IN ITS CAVITY, THE COMPLETE MODULE IS DISMOUNTED FOR REVIEW, THE BAY IS REPLACED AGAIN BY An ERROR IN SL, THE MODULE OF INLET IS SUBSTITUTE , VALVE TUBE AND WASTE TUBE, ALL TEST IS MADE BY SERVICE, THB IS CALIBRATED, THE FINAL OPERATION WITH QC IS VERIFIED. |  | Defect | HW | Retainer | Wetability ?
ABL800Basic |  |  | CDFA0C07421A6DE9 | No
ABL817 FLEX | System überprüft kein Fehler fest gestellt |  | SWO-4505 | No | System does not check error found
ABL837 FLEX |  |  | CDFA22AB42109037 | No
ABL825 FLEX | Review data logs and review setup. Found unable to access the calibration schedule, analyzer went into a system error message. | Remove application software, reboot analyzer and install software. Transfer settings from other analyzer and add original database. | CDFA4C1342033F63 | No | Review data logs and review setup. Found unable to access the calibration schedule, analyzer went into a system error message. | Remove application software, reboot analyzer and install software. Transfer settings from other analyzer and add original database.
ABL80 FLEX CO-OX | Analysator geöffnet, blockierte Schläuche ausgetauscht, Fluss behindert. Schlauch-Kit eingebaut. Flüssigkeitstransport ohne Probleme. Kalibrationen durchgeführt, QC wiederholt. Ersatzlösungspack über Zentrale an Kunden senden lassen. Buchung auf GS16. Bei Verwendung des LP wird Material auf Kunden gebucht, ansonsten Rücknahme nach Miete. |  | SWO-4790 | No | Analyzer opened, blocked hoses replaced, flow obstructed. Hose kit installed. Liquid transport without problems. Calibrations performed, QC repeated. Send a replacement solution pack to customers via the head office. Booking on GS16. When using the LP, material is posted to customers, otherwise return after rental.
ABL90 | FSE recieved email from other oncall FSE. Problems with pO2. Analyzer on bench in lab when FSE arrived. Stuck on Change Solution Pack screen. | Checked solution pack and found it was not crunched down all the way. Crunched the SP down. Tried to re-install SP. No luck. Installed a new Sensor Cassette. Completed dayligth saving procedure. Set up time on analyzer. Completed all servie tests. | CDFB2B6CB8564D31 | No | FSE recieved email from other oncall FSE. Problems with pO2. Analyzer on bench in lab when FSE arrived. Stuck on Change Solution Pack screen. | Checked solution pack and found it was not crunched down all the way. Crunched the SP down. Tried to re-install SP. No luck. Installed a new Sensor Cassette. Completed dayligth saving procedure. Set up time on analyzer. Completed all servie tests. | Defect | SC | Wetability
ABL825 FLEX | HJw (siehe SWO) |  | SWO-4463 | No | HJw (see SWO)
ABL825 FLEX | Na failing QC all levels | Replaced Na electrode. | CE0CD1CD2C236C87 | No | Na failing QC all levels | Replaced Na electrode.
ABL800 BASIC | Reinigen Luftansaugstutzen LS Justagefehler |  | SWO-4472 | No | Cleaning air intake nozzle LS adjustment error
ABL825 FLEX | Inhomogenous sample errors. | Replaced Inlet tube. Inlet probe replaced from customer stock | CE0E6F7DA2287836 | No | Inhomogenous sample errors. | Replaced Inlet tube. Inlet probe replaced from customer stock
ABL825 FLEX | Checked data logs. Also check with 5th floor lab (had experienced same issue) and they stated no problems since visit on 10 1 | Replaced electrode pump tubing on EI Met module. Ran zero calibration. Cleaned top of Lactate electrode with alcohol pad. Remembrane electrode (Used mebranes for troubleshooting purposes). | CDFB71F3B85DDD10 | No | Checked data logs. Also check with 5th floor lab (had experienced same issue) and they stated no problems since visit on 10 1 | Replaced electrode pump tubing on EI Met module. Ran zero calibration. Cleaned top of Lactate electrode with alcohol pad. Remembrane electrode (Used mebranes for troubleshooting purposes).
ABL825 FLEX | Checked data logs. | Replaced electrode pump tubing on EI Met module. Ran Zero current calibration. Cleaned top of Lactate electrode with alcohol pad. Remembrane electrode. | CDFB87B6B863DD15 | No | Checked data logs. | Replaced electrode pump tubing on EI Met module. Ran Zero current calibration. Cleaned top of Lactate electrode with alcohol pad. Remembrane electrode.
ABL835 FLEX |  |  | CE110DBE8E638C49 | No
ABL80 FLEX Basic version | Equipo apagado (el cliente se encuentra el equipo apagado). Se le dan instrucciones y arranca pero da fallo al no reconocer la SC. Se comprueba la cassette, se hace un apagado controlado del equipo y se reinstala la cassette siendo esta reconocida. Se le realizan dos calibraciones y se comprueban las corrientes de la bateria, así como los diferentes programas de servicio y configuración. Se le entrega material (SC y SP). Revisión y comprobación del equipo con QC´s. | Cheque del equipo y reinstalación del SC. | CDFC6CED2E079D75 | No | Computer off (client is off). You are given instructions and it starts but fails to recognize the SC. The cassette is checked, the computer is switched off controlled, and the cassette is reinstalled as it is recognized. It is performed with two calibrations and the battery currents are checked, as well as the different service and configuration programs. Material (SC and SP) is delivered to you. Review and check the equipment with QC's. | Check the equipment and reinstall the SC.
ABL90 FLEX | Daten überprüft, SC erneuert (altes Lot: 227-74, neues Lot: 257-14), Startphase, tHb-Kal, Kal i.O., iQC |  | SWO-4625 | No | Data checked, SC renewed (old lot: 227-74, new lot: 257-14), start phase, tHb-Kal, Kal i.O., iQC |  | Defect | SC | Glu
ABL90 FLEX | ABL durchgespült, Jahreswartungskit verbaut, Hemolyzer, Ventil, Waste connector, Einlassdichtung getauscht. mehrmals kalibriert QCs für alle Level durchgeführt ABL wieder messbereit. |  | SWO-4551 | No | ABL rinsed, annual maintenance kit installed, hemolyzer, valve, waste connector, inlet seal exchanged. several times calibrated QCs performed for all levels ABL again ready to measure. |  | Cleaned | Fluidic pathway
ABL90 | KEITH INVESTIGATED AND ORDERED PARTS FOR ERROR CODE AND LEAKAGE ISSUES. I SHOWED UP ON SITE FOUND A SMALL AMOUNT OF BUILT UP QC MATERIAL. I WENT INTO SERVICE MODE AND FAILED SERVICE TEST FOR FLUID TRANSPORT ON SOLUTION PACK. | REPLACED EVERYTHING EXCLUDING THE YELLOW TUBE IN THE PM KIT. ALSO REPLACED CONSUMABLES AND IT RESOLVED THE PROBLEM. | CDFCA5F12E24D89B | No | KEITH INVESTIGATED AND ORDERED PARTS FOR ERROR CODE AND LEAKAGE ISSUES. I SHOWED UP ON SITE FOUND A SMALL AMOUNT OF BUILT UP QC MATERIAL. I WENT INTO SERVICE MODE AND FAILED SERVICE TEST FOR FLUID TRANSPORT ON SOLUTION PACK. | REPLACED EVERYTHING EXCLUDING THE YELLOW TUBE IN THE PM KIT. ALSO REPLACED CONSUMABLES AND IT RESOLVED THE PROBLEM. | Defect | SC | Wetability
ABL90 | Keith investigated and ordered parts for error code and leakage test. I showed up on site and found a small amount of QC material by the pinch tube. I went into service and it failed the fluid transport test for the solution pack. | I replaced all but the yellow tube in the Pm kit and replaced consumables. I ran a backflush and the fluid flow was smooth. This resolved the issue. | CDFCAA022E454C93 | No | Keith investigated and ordered parts for error code and leakage test. I showed up on site and found a small amount of QC material by the pinch tube. I went into service and it failed the fluid transport test for the solution pack. | I replaced all but the yellow tube in the Pm kit and replaced consumables. I ran a backflush and the fluid flow was smooth. This resolved the issue. | Cleaned | Fluidic pathway
ABL90 | KEITH INVESTIGATED AND ORDERED PARTS FOR ERROR CODE AND LEAKAGE ISSUES. I SHOWED UP ON SITE FOUND A SMALL AMOUNT OF BUILT UP QC MATERIAL. I WENT INTO SERVICE MODE AND FAILED SERVICE TEST FOR FLUID TRANSPORT ON SOLUTION PACK. | REPLACED EVERYTHING EXCLUDING THE YELLOW TUBE IN THE PM KIT. ALSO REPLACED CONSUMABLES AND IT RESOLVED THE PROBLEM. | CDFCAB4B2E37D8A1 | No | KEITH INVESTIGATED AND ORDERED PARTS FOR ERROR CODE AND LEAKAGE ISSUES. I SHOWED UP ON SITE FOUND A SMALL AMOUNT OF BUILT UP QC MATERIAL. I WENT INTO SERVICE MODE AND FAILED SERVICE TEST FOR FLUID TRANSPORT ON SOLUTION PACK. | REPLACED EVERYTHING EXCLUDING THE YELLOW TUBE IN THE PM KIT. ALSO REPLACED CONSUMABLES AND IT RESOLVED THE PROBLEM. | Defect | SC | Wetability
ABL90 FLEX | Schulung, Einweisung, div Probleme erörtert, Anbindung an OSM bez anpassung der Lot Nummer QC etc...... Spüladapter, Ersatznadel, Einlasdichtung nun vor Ort |  | SWO-4621 | No | Training, instruction, div problems discussed, connection to OSM matching the lot number QC etc...... Rinsing adapter, replacement needle, lasket now on site |  | Defect | SC | Wetability
ABL815 FLEX | Vérifier power supply OK ABL est fonctionnel mais aucun display video driver défectueux sur CPU. et PM due | Remplacé CPU unit,Effecué la maintenance bi-annuel selon les normes remplacé les tubes de pompe nettoyé pompes et flush v7.v8 verifier electrode o'ringfull oxi test et calibration service THB calibration nettoyé les pompes vérifié electronique calibration zero po2 ,lactate.glucose . nettoyé autocheck laser. | CDFC46C82E44BA27 | No | Check power supply OK ABL is functional but no faulty video driver display on CPU. and PM due | Replaced CPU unit,Effecué bi-annual maintenance according to standards replaced pump tubes cleaned pumps and flush v7.v8 check er electrode o'ringfull oxi test and calibration service THB calibration cleaned pumps verified electronic calibration zero po2, lactate.glucose. cleaned laser autocheck.
ABL827 FLEX | Confirmed the problem. | Replaced the friction detector unit, inlet probe and inlet gasket. | CDFCD9932E3432D1 | No | Confirmed the problem. | Replaced the friction detector unit, inlet probe and inlet gasket.
ABL90 FLEX | System im Serviceprogramm überprüft, Einzelschritte durchgeführt, Funktion im Serviceprogramm OK. Nach Neustart wieder Flowselektorfehler. Flowselector Detector PCB ausgebaut und erneuert. Systemstart durchgeführt. Funktion OK. |  | SWO-4630 | No | System checked in the service program, individual steps performed, function in the service program OK. After rebooting again flow selector error. Flowselector Detector PCB expanded and renewed. System boot. Function OK. |  | Defect | SP | Flow selector
ABL80 FLEX CO-OX |  |  | SWO-5052 | No
ABL90 FLEX | Touchscreenbedienung verschoben --> Touchscreen neu kalibriert, i.O. |  | SWO-4723 | No | Touchscreen operation shifted --> Touchscreen recalibrated, i.O. |  | Adjustment | Touch screen | Alignment
ABL90 |  |  | CE004421902C291C | No |  |  | Cleaned | Fluidic pathway | Clot
ABL90 | We had tried a new CF card last Friday and reloading the software, but it continues to go into a blue screen. This is an OUT OF BOX FAILURE | Don tried another CF card today and it stills go into a blue screen and reboots itself. This seems to be a Computer related problem. | CE00A1719042BC97 | No | We had tried a new CF card last Friday and reloading the software, but it continues to go into a blue screen. This is an OUT OF BOX FAILURE | Don tried another CF card today and it stills go into a blue screen and reboots itself. This seems to be a Computer related problem. | Unknown | Revisit required | CPU
ABL825 FLEX | Skiftet ref. elektrode Renset Lak - Glu forstærker Nul strøms kal. |  | SWO-4794 | No | Changed ref. electrode Purified Varnish-Glu amplifier Zero current KAL.
ABL815 FLEX | Flex Q Module 753R0947N002 abgebaut. |  | SWO-4726 | Yes | Flex Q Module 753R0947N002. |  | Re classified RMED | Analyzer setup | FlexQ Module
ABL90 FLEX | Klant heeft al Flush procedure uitgevoerd maar probleem is daarmee nog niet opgelost. Klant heb aangeraden om de ABL via zwart schakelaar (hard) uitzetten. Na opstart procedure van 45 min. staan alle parameters weer op groen. Bij aankomst zie dat de metingen ook goed verlopen. Het probleem begint meestal na een arteriele meting. Capillaire metingen gaan goed. Klant gebruikt nu PICO50 maar ze gaan binnenkort safePICO spuiten van Radiometer te gebruiken. Tevens de inletgasket blijft hangen in de arteriele stand. Zodoende inletmodule compleet heb vervangen vervolgens vier arteriele metingen heb uitgevoerd, OK. Tevens software upgrade versie 3.1 naar 3.1 MR1 heb uitgevoerd. Hierna diverse calibratie QC's heb uitgevoerd, OK. Datalogs heb gekopieerd voor verder onderzoek. |  | SWO-4815 | No | Customer has already performed Flush procedure but problem has not been solved. Customer recommended to turn off the ABL via Black switch (hard). After the start-up procedure of 45 min. All parameters are returned to green. On arrival see that the measurements also went well. The problem usually starts after an arterial measurement. Capillary measurements go well. Customer is now using PICO50 but they are going to use safePICO spray from Radiometer soon. Also the Inletgasket hangs in the arterial position. So the inletmodule has completely replaced then four arterial measurements have performed, OK. Also implemented software upgrade version 3.1 to 3.1 MR1. Have performed several calibration Qcs, OK. Data logs have been copied for further investigation.
ABL90 | Leak detected errors present 1342. Unable to clear vacuum error in fluids service mode. | Performed several flushes. Ran through fluidic service programs but could not get all tests to pass. Replaced sensor cossette. No help. Replaced solution pack. | CE018078062FA1CD | No | Leak detected errors present 1342. Unable to clear vacuum error in fluids service mode. | Performed several flushes. Ran through fluidic service programs but could not get all tests to pass. Replaced sensor cossette. No help. Replaced solution pack. | Defect | SP | Stuck in "replace Inlet Gasket"
ABL90 | Unit was flushed and reinstalled several times by client. On arrival units was aggressively and SP reinstalled. | Flushed unit and SP reinstalled. SP empty will order replacement. | CE01D7380654AFF3 | No | Unit was flushed and reinstalled several times by client. On arrival units was aggressively and SP reinstalled. | Flushed unit and SP reinstalled. SP empty will order replacement. | Defect | SC | Wetability
ABL827 FLEX |  |  | CE01993D0608398A | No
ABL820 FLEX | Found ampoule breaker not breaking ampoules open. | Replaced ampoule opener stepping motor per technical notice. | CE01AEAA06589CA8 | No | Found ampoule breaker not breaking ampoules open. | Replaced ampoule opener stepping motor per technical notice.
ABL827 FLEX |  |  | CE01851F061901CA | Yes |  |  | Re classified RMED | Preventive maintenance
ABL820 FLEX | Checked data logs for errors. Found sample errors 0719: Insufficient sample. Checked ls adjust and pump calibrations. Found pump calibration factors around the 83 to 86% range for both pH BG module and solution pumps. | Replaced measuring chamber, pH sealing ring and multiple pre-measured tubings from PM kit which included inlet gasket. | CE017E5206511717 | No | Checked data logs for errors. Found sample errors 0719: Insufficient sample. Checked ls adjust and pump calibrations. Found pump calibration factors around the 83 to 86% range for both pH BG module and solution pumps. | Replaced measuring chamber, pH sealing ring and multiple pre-measured tubings from PM kit which included inlet gasket.
ABL825 FLEX | pCO2 Elektrode ersetzt. Kal. Und QC durchgeführt. Werte im RiliBÄK Kontrollbereich. |  | SWO-4900 | No | pCO2 electrode. Kal. And QC performed. Values in the RiliBÄK control area.
ABL837 FLEX |  |  | CE14825BF03920E3 | No
AQT90 FLEX | Reinigung des Barcode Scanners und Neueinstellung damit Laserstrahl mittig in storage wheel Öffnung Barcodes scannen kann. Eine neue Kalibration TnT wurde angestoßen. Kalibration o.K. D-Dimer keine Vorort. Anwender bestellt neue und kalibriert sie. |  | SWO-4825 | No | Cleaning the barcode scanner and resetting so that laser beam can scan in the middle of storage wheel opening barcodes. A new calibration TnT has been initiated. Calibration o.K. D-Dimer no suburb. User orders new ones and calibrates them.
ABL90 | Intervention required when arrived on screen | Back flushed instrument and performed PM. Replaced all parts in PM kit | CE02E4B87C5B87BC | No | Intervention required when arrived on screen | Back flushed instrument and performed PM. Replaced all parts in PM kit | Cleaned | Fluidic pathway
ABL90 | Back flushed instrument but did not resolve | Performed PM. Replaced all parts in PM kit | CE02E5D37C0987C0 | No | Back flushed instrument but did not resolve | Performed PM. Replaced all parts in PM kit | Defect | HW | Waste tube
ABL90 | sensor temp errors , analyzer was shut down for 3 months when start up the analyzer had temp errors, Dan Hanson on site replaced the sensor cassette and solution pack, emailed credit form for the replacement of sensor and solution pack | performed an extensive back flush replaced the sensor cassette and solution pack | CE02C1507C5BD26A | No | sensor temp errors ; analyzer was shut down for 3 months when start up the analyzer had temp errors; Dan Hanson on site replaced the sensor cassette and solution pack; emailed credit form for the replacement of sensor and solution pack | performed an extensive back flush replaced the sensor cassette and solution pack | Defect | SC | Termistor
ABL90 |  |  | CE0299F47C1953BD | No |  |  | Defect | SC | Reference
ABL90 FLEX | Nach Probelauf zurück. SW Upgrade 3.1MR1 durchgef. |  | SWO-4982 | No | Return after trial run. SW Upgrade 3.1MR1. |  | Defect | SC | Termistor
ABL800Basic |  |  | CE02E1EC7C09EFDA | Yes |  |  | Re classified RMED | Planned TU
ABL800 BASIC | System überprüft und gereinigt, Kontakte gereinigt und justiert, Elektrode neu membraniert, Empfindlichkeit zu gering, Elektrode gewaschen, keine Besserung, neue Elektrode wird bestellt und ausgetauscht. |  | SWO-5022 | No | System checked and cleaned, contacts cleaned and adjusted, electrode re-membraneed, sensitivity too low, electrode washed, no improvement, new electrode is ordered and replaced.
ABL825 FLEX | Found pH BG calibration upgraded to a 1 point calibration pCO2 drigting on calibrations Found some upper liquid sensor detected rinse during calibration errors | Installed a measuring chamber-No change-Removed Installed the pH BG air valve based on the upper liquid sensor errors Cleaned the pH and pCO2 electrodes | CE043E9EF23041D3 | No | Found pH BG calibration upgraded to a 1 point calibration pCO2 drigting on calibrations Found some upper liquid sensor detected rinse during calibration errors | Installed a measuring chamber-No change-Removed Installed the pH BG air valve based on the upper liquid sensor errors Cleaned the pH and pCO2 electrodes
ABL825 FLEX | I am writing to see how I go about getting reimbursement of lactate membranes and ordering a lactate electrode. I am not sure if you are aware of how unstable lactate is. I have documented troubleshooting for some time now, as to how many times I had to change lactate more frequently than suggested due to numerous unstable measurements, for calibrations, qc and patient results. The issue is still continuing and myself and service has done everything that we can possibly do to rectify the situation. I do have all the supporting data to show the influx in replacement of lactate membranes as well as all qc, calibration and patient data showing unstable measurements support my findings. We are unaware of why this is happening, but we will like to have this investigated further | ESCALATED THE CALL.ARRANGED TO HAVE 2 BOXES OF MEMBRANES OVERNIGHTED TO CHILDRENS AT NO COST.APPROVED BY JIM GIBBONS.PLEASE KEEP IN MIND THESE 4 ANALYZERS ARE ALL LESS THAN 4 MONTHES OLD AND THE PROBLEM EXISTS ON ALL 4 ANALYZERS.ALSO VERY IMPORTANT IS THAT THEIRRANGES FOR LACTATE ARE VERY TIGHT 1.4-1.8 AND 4.3 TO 4.7 WHICH ONLY MAKES THE ISSUE SEEM WORSE. | CE02FCDA7C396020 | No | I am writing to see how I go about getting reimbursement of lactate membranes and ordering a lactate electrode. I am not sure if you are aware of how unstable lactate is. I have documented troubleshooting for some time now, as to how many times I had to change lactate more frequently than suggested due to numerous unstable measurements, for calibrations, qc and patient results. The issue is still continuing and myself and service has done everything that we can possibly do to rectify the situation. I do have all the supporting data to show the influx in replacement of lactate membranes as well as all qc, calibration and patient data showing unstable measurements support my findings. We are unaware of why this is happening, but we will like to have this investigated further | ESCALATED THE CALL. ARRANGED TO HAVE 2 BOXES OF MEMBRANES OVERNIGHTED TO CHILDRENS AT NO COST. APPROVED BY JIM GIBBONS. PLEASE KEEP IN MIND THESE 4 ANALYZERS ARE ALL LESS THAN 4 MONTHES OLD AND THE PROBLEM EXISTS ON ALL 4 ANALYZERS. ALSO VERY IMPORTANT IS THAT THEIRRANGES FOR LACTATE ARE VERY TIGHT 1.4-1.8 AND 4.3 TO 4.7 WHICH ONLY MAKES THE ISSUE SEEM WORSE.
ABL827 FLEX |  |  | CE05225A6837ED24 | Yes |  |  | Re classified RMED | Not upSWO
ABL827 FLEX |  |  | CE0521C16835ED20 | Yes |  |  | Re classified RMED | Setup / Install
ABL825 FLEX | Checked the calibration logs and noticed that glucose and lactate are presenting unstable measurement and drifting errors. | Cleaned the glucose and lactate electrodes and replaced new membranes. Cleaned the measuring chambers. Cleaned the electrode connectors. Performed the service zero calibration. | CE02EC737C6164DA | No | Checked the calibration logs and noticed that glucose and lactate are presenting unstable measurement and drifting errors. | Cleaned the glucose and lactate electrodes and replaced new membranes. Cleaned the measuring chambers. Cleaned the electrode connectors. Performed the service zero calibration.
ABL80 OSM | went on site and calibrated both units and ran QC as samples and both units compaired well but on blood they did not. | will discuss options with client on evaluating the 800 to make sure the hemolyzer is working to spec. | CE045665F254B000 | No | went on site and calibrated both units and ran QC as samples and both units compaired well but on blood they did not. | will discuss options with client on evaluating the 800 to make sure the hemolyzer is working to spec.
ABL90 FLEX | Fehlersuche SP richtig aktiviert Probenweg gereinigt THb-Kalibrierung SC aufgrund Kalibrierfehler erneuert Gerät steht bei Verlassen noch in Startphase |  | SWO-5109 | No | Troubleshooting SP correctly activated sample path cleaned THb calibration SC due to calibration error renewed device is still in start phase when leaving |  | User Error | SP activate
ABL90 FLEX | Austausch Sensorkassette Kal- und QC Level A,B,C |  | SWO-5146 | No | Replacement sensor cassette Cal- and QC Level A,B,C |  | Defect | SC | pCO2
ABL817 FLEX | Gaseinheit und Ventile zerlegt und gereinigt (Feuchtigkeit vorhanden) Gaseinheit im Sevicemenü getestet - Offset-Spannung mit 0,064 unterhalb des Minimalwetes von 0,07. Ersatzteile bestellt. |  | SWO-5086 | No | Gas unit and valves disassembled and cleaned (moisture available) gas unit tested in the Sevice menu - offset voltage with 0.064 below the minimum wet of 0.07. Spare parts ordered.
ABL825 FLEX | Proteinentfernung durchgeführt - Messkammer gereinigt pH Dichtring sowie Luftventil im El. Met. Modul erneuert Verbindungsschläuche vom Einlass zur Messkammer erneuert Gerät kalibriert und überprüft |  | SWO-5095 | No | Protein removal performed - measuring chamber cleaned pH sealing ring as well as air valve in the el. Met. Module renews connecting hoses from inlet to measuring chamber renewed device calibrated and checked
ABL825 FLEX | Remembraning ref electrode did not make a difference | Replaced Ref electrode | CE0CD9202C526C97 | No | Remembraning ref electrode did not make a difference | Replaced Ref electrode
ABL825 FLEX | inspected oxi module. Ran service checks | Replaced Hemolyzer | CE0CD8102C166C94 | No | inspected oxi module. Ran service checks | Replaced Hemolyzer
TCM-COMBIM |  |  | CE043515F210D7B8 | No |  |  | Defect | HW | Battery module
TCM-COMBIM |  |  | CE0431B2F22B0BD6 | No |  |  | Defect | HW | PCB
TCM-COMBIM |  |  | CE050A0E685BC712 | Yes |  |  | Re classified RMED | No facts
ABL80 CO-OX | Air detected during calibration. Hemolyzer leaking. | Cleaned analyzer. Replaced and aligned hemolyzer unit. Installed maintenance kit. Installed new sensor cassette. | CE055F09680CD94A | No | Air detected during calibration. Hemolyzer leaking. | Cleaned analyzer. Replaced and aligned hemolyzer unit. Installed maintenance kit. Installed new sensor cassette.
ABL90 FLEX | Tausch des Hemolyzers brachte kein Erfolg, ihn zurück gebaut. Fehler vmtl. im Spectrometer oder wet section control pcb. Per Nachtexpress bestellt. |  | SWO-5205 | No | Swap of the Hemolyzer brought no success, he built back. Error vmtl. in spectrometer or wet section control pcb. Ordered by Night Express. |  | Unknown | Revisit required
ABL90 | FOUND ANNUAL KIT PINCH VALVE TUBING HAS A HOLE | REPLACED ALL TUBINGS | CE05942C680EFDD7 | Yes | FOUND ANNUAL KIT PINCH VALVE TUBING HAS A HOLE | REPLACED ALL TUBINGS | Re classified RMED
ABL90 FLEX | Fehlersuche am Analysator, SC bereits 2x durch Anwender erneuert (einmal Routine) und 1xLP, pCO2 Kal. nicht bestanden, da die Empf. zu niedrig (25-49,5%), Einlassdichtung erneuert, Spülung, SC nochmal erneuert->Kal. bestanden, defekte SCs und LP werden reklamiert System-und Funktionskontrolle i.O. | No | SWO-5402 | No | Troubleshooting on the analyzer, SC already renewed 2x by users (once routine) and 1xLP, pCO2 Kal. did not pass, as the mpf. too low (25-49.5%), inlet seal renewed, flushing, SC renewed again->Kal. passed, defective SCs and LP are complained system and function control i.O. | No | Defect | SC | pCO2
ABL90 FLEX | Kundengerät zurückgetauscht Startphase von 55 Min abgewartet Kal und QC ok Reparaturersatzgerät mitgenommen |  | SWO-5273 | Yes | Customer device exchanged back Start phase of 55 min waited Kal and QC ok repair replacement device taken |  | Re classified RMED | Analyzer returned to site
ABL90 FLEX | LP vom Anwender nich korrekt aktiviert, LP aktiviert und eingebaut, Software Downgrade auf 3.1 durcchgeführt. Nach Startphase QC bei pCO2 weiterhin außerhalb. SC90 aus Kundenbestand ersetzt. |  | SWO-5263 | No | LP not correctly activated by the user, LP activated and installed, software downgrade to 3.1 durcch guided. After start-up phase QC at pCO2 still outside. SC90 from customer base. |  | Defect | SC | pCO2
ABL90 FLEX | System Status Datei ausgelesen. Funktionen im Service Menu überprüft. Oxy Schlauch undicht => ausgetauscht! Funktionskontrolle, Kalibrationen Software neu installiert, Konfiguration wieder hergestellt, Funktionskontrolle |  | SWO-5271 | No | System status file read out. functions in the Service Menu. Oxy hose leaking => replaced! Functional control, calibration software reinstalled, configuration restored, function control |  | Defect | HW | Oxi Module | Pinch Valve tube
ABL90 |  |  | CE055743683FF614 | Yes |  |  | Re classified RMED | Logistical upSWO
ABL90 | checked data logs for errors, pH and Cl errors began after 36 hours of operation, | performed back flush, flushed senspr cassette, perfored transport functions manifold test, pressure test liquid sensor and pump calibrations, tested solution pack and clot detection program, no change to analyzer. | CE0835EE5433A055 | No | checked data logs for errors, pH and Cl errors began after 36 hours of operation, | performed back flush, flushed senspr cassette, perfored transport functions manifold test, pressure test liquid sensor and pump calibrations, tested solution pack and clot detection program, no change to analyzer. | Unknown | Revisit required | SC?
ABL90 | approved order. this is for credit to be sent in |  | CE05728F684BF61D | Yes | approved order. this is for credit to be sent in |  | Re classified RMED | Logistical upSWO
ABL815 FLEX | FlexQ Probennehmer werden nicht erkannt. FlexQ Barcode Laser gereinigt. Nach erfolgreichen 40 Versuchen den Probennehmer einzulesen tritt der Effekt auf das der Barcode auf dem Probennehmer nicht erkannt wird. |  | SWO-5257 | No | FlexQ samplers are not recognized. FlexQ barcode laser cleaned. After successful 40 attempts to read the sampler, the effect occurs that the barcode is not recognized on the sampler.
ABL825 FLEX |  |  | CE052A8F682B8C01 | No
ABL825 FLEX | Ran complete oxi test | Lamp unit adjusted | CE0CE8F92C126C9B | No | Ran Complete Oxygen Test | Lamp unit adjusted
ABL80 CO-OX | The fan was not making noise when I arrived. The fan vibrates when compressed air is blown through the fan blades. The fan needs to be replaced. | Replaced the fan. | CE0936B6CA3BB495 | No | The fan was not making noise when I arrived. The fan vibrates when compressed air is blown through the fan blades. The fan needs to be replaced. | Replaced the fan.
ABL90 |  |  | CE08FBA4CA03411B | No |  |  | Defect | HW | CF card | Status Code 2
ABL90 | Checked data logs for errors. Ran test under service mode. Pressure tests were failing, but when rerun the test passed. | Replaced Inlet gasket for fault isolation. Also exchanged inlet connector gasket. Check fluid flows. | CE098A30CA2A93E0 | No | Checked data logs for errors. Ran test under service mode. Pressure tests were failing, but when rerun the test passed. | Replaced Inlet gasket for fault isolation. Also exchanged inlet connector gasket. Check fluid flows. | Defect | HW | Inlet Gasket
ABL90 | Checked analyzer QC results. pCO2 high on all levels. Sensor just replaced. Tried addition calibrations. QC would still not come into range. | Replaced sensor cassette with customer stock. Wait for all parameters to come into range. Fill out credit form for replacement sensor cassette. | CE098EE3CA5A93E4 | No | Checked analyzer QC results. pCO2 high on all levels. Sensor just replaced. Tried addition calibrations. QC would still not come into range. | Replaced sensor cassette with customer stock. Wait for all parameters to come into range. Fill out credit form for replacement sensor cassette. | Defect | SC | pCO2
ABL90 FLEX | Flüssigkeitsweg gespült, Programme Fluid Transport auf Serviceprogrammebene ausgeführt, zunächst keine Veränderung, Systemreset durchgeführt, anschließende Startphase ohne Flüssigkeitstransportfehler durchlaufen, jedoch QC-Fehler pCO2 u. pO2 bei Level A u. B, SC aus GS19 Bestand erneuert | No | SWO-5401 | No | Fluid path flushed, programs Fluid Transport executed at service program level, initially no change; System reset performed, subsequent start phase without liquid transport error, but QC error pCO2 and pO2 at level A &b, SC from GS19 stock renewed | No | Defect | SC | Wetability
ABL90 | checked data logs for errors, | removed and replaced sensor cassette, performed linearity using VK-R5, ran 6-7 waste bloods | CE099095CA19A05D | No | checked data logs for errors, | removed and replaced sensor cassette, performed linearity using VK-R5, ran 6-7 waste bloods | Defect | SC | pH
ABL90 | Checked data logs for errors, since friday calibration errors and or internal QC errors for multiple parameters, customer could not run external QC because of cal and internal QC errors | checked flow of solutions, performed transport programs, replaced solution pack and sensor cassette, ran 6-7 waste bloods. | CE098BBECA1DA058 | No | Checked data logs for errors, since friday calibration errors and or internal QC errors for multiple parameters, customer could not run external QC because of cal and internal QC errors | checked flow of solutions, performed transport programs, replaced solution pack and sensor cassette, ran 6-7 waste bloods. | Defect | SC | Wetability
ABL835 FLEX | Einstellung korrigiert. | No | SWO-5329 | No | Adjustment corrected. | No
AQT90 FLEX |  |  | CE091668CA3D65A6 | No
ABL90 FLEX | Gerät überprüft, manuell gespült, Sensorkassette verstopft, getauscht, Kaltstartphase, mehrere Kal,s durchgeführt (lac ausser Bereich), Funktion des Systems wieder voll vorhanden |  | SWO-5479 | No | Device checked, manually rinsed, sensor cartridge clogged, exchanged, cold start phase, multiple cal,s performed (lac out of area), function of the system fully available again |  | Defect | SC | Wetability/Clot
ABL825 FLEX | Gas 2 Ventil erneuert Ref, Glu und Lac Elektrode neu membraniert Gerät kalibriert und überprüft |  | SWO-5496 | No | Gas 2 valve renews ref, glu and lac electrode re-membraneed device calibrated and checked
ABL825 FLEX | Kunde muß im LIS KIS neuen Gerätenamen anlegen. | Kunde muß im LIS KIS neuen Gerätenamen anlegen. | SWO-5622 | No | Customer must create new device name in LIS KIS. | Customer must create new device name in LIS KIS.
ABL820 FLEX | Finding no apparent problem. Replacing Measure Chamber, Fluid, and Air valve, Upper and lower Liquid sensor.No errors other than Sample error due to liquid in chamber.Calibrates OK. | Replace suspect parts. | CE0A544940378942 | No | Finding no apparent problem. Replacing Measure Chamber, Fluid, and Air valve, Upper and lower Liquid sensor. No errors other than Sample error due to liquid in chamber. Calibrates OK. | Replace suspect parts.
ABL80 FLEX Basic version | Revisión del SC, comprobación del mismo en el equipo junto a otro SP diferente, debiendo ser sustituido por fallo del mismo. Se comprueba el SP con otro equipo siendo de igual manera necesaria su sustitución. Revisión y comprobación del equipo con QC´s. | Cambio del SP y SC. | CE0BABE8B64543CE | No | Review of the SC, checking the same on the computer together with another sp, must be replaced by failure of the same. The SP is checked with another computer, and its replacement is also necessary. Review and check the equipment with QC's. | Switching SP and SC.
ABL90 FLEX | Tausch des Spectrometer Einheit. Tausch auf Gewährleistung, da Analysator 12.2013 beim Anwender aufgestellt. | Tausch des Hemolyzers brachte kein Erfolg, ihn zurück gebaut. Fehler vmtl. im Spectrometer oder wet section control pcb. Per Nachtexpress bestellt. | SWO-5576 | No | Swap the Spectrometer unit. Exchange on warranty, since Analyzer 12.2013 installed at the user. | Swap of the Hemolyzer brought no success, he built back. Error vmtl. in spectrometer or wet section control pcb. Ordered by Night Express. | LEH Email
ABL90 FLEX | Telefonisch klant heb aangeraden om het RP te vervangen, het probleem is nog niet opgelost. Bij aankomst zie veel lekkage bij de waste connector tubing. Zodoende wetsection module heb gedemonteerd, waste connector tubing heb vervangen. ABL heb schoongemaakt hierna zie diverse vloeistoftransport foutmeldingen. Bij nader onderzoek blijkt dat inletgasket was verstopt. Inletgasket naald heb doorgespoten vervolgens kalibratie QC's heb uitgevoerd, OK. | Telefonisch klant heb aangeraden om het RP te vervangen, het probleem is nog niet opgelost. Bij aankomst zie veel lekkage bij de waste connector tubing. Zodoende wetsection module heb gedemonteerd, waste connector tubing heb vervangen. ABL heb schoongemaakt hierna zie diverse vloeistoftransport foutmeldingen. Bij nader onderzoek blijkt dat inletgasket was verstopt. Inletgasket naald heb doorgespoten vervolgens kalibratie QC's heb uitgevoerd, OK. | SWO-5657 | No | Telephone customer recommended to replace RP, the problem has not yet been solved. On arrival see a lot of leakage at the waste connector tubing. So wetsection module got disassembled, waste connector tubing got replaced. ABL have cleaned hereafter see various liquid transport error messages. Further investigation reveals that Inletgasket was blocked. Inletgasket needle have been sprayed then calibration Qcs have performed, OK. | Telephone customer recommended to replace RP, the problem has not yet been solved. On arrival see a lot of leakage at the waste connector tubing. So wetsection module got disassembled, waste connector tubing got replaced. ABL have cleaned hereafter see various liquid transport error messages. Further investigation reveals that Inletgasket was blocked. Inletgasket needle have been sprayed then calibration Qcs have performed, OK.
ABL90 FLEX | De SC heb gecontroleerd, geen afwijkingen heb geconstateerd. ABL heb aan uit gezet maar het probleem is daarmee nog niet opgelost. SC heb 55 min. opnieuw geconditioneerd, nog steeds diverse temp. errors. Een nieuwe SC lot. nr. R0588 N134 heb geplaatst, hierna is alles weer OK. De logs heb gecopieerd voor verder onderzoek. | De SC heb gecontroleerd, geen afwijkingen heb geconstateerd. ABL heb aan uit gezet maar het probleem is daarmee nog niet opgelost. SC heb 55 min. opnieuw geconditioneerd, nog steeds diverse temp. errors. Een nieuwe SC lot. nr. R0588 N134 heb geplaatst, hierna is alles weer OK. De logs heb gecopieerd voor verder onderzoek. | SWO-5665 | No | The SC have checked, no abnormalities have been noted. ABL have turned it off but the problem has not yet been solved. SC have 55 min. Re conditioned, still various temp. Errors. A new SC lot. No. R0588 N134 have placed, after that everything is OK again. The logs have been copied for further investigation. | The SC have checked, no abnormalities have been noted. ABL have turned it off but the problem has not yet been solved. SC have 55 min. Re conditioned, still various temp. Errors. A new SC lot. No. R0588 N134 have placed, after that everything is OK again. The logs have been copied for further investigation.
ABL90 |  |  | CE0E3D7DA2425949 | No |  |  | User Error | Analyzer setup
ABL90 |  |  | CE0BD111B644D94C | No |  |  | Defect | SP
ABL90 |  |  | CE1AE1833E1A9C20 | No
ABL90 FLEX | Netzwerk Konfiguration telefonisch ermittelt. Netzwerkverbindung hergestellt, Verbindung per Ping überprüft, Time Server eingestellt. Daten auf CF Karte übertragen. LIS und Aqure Verbindungen hergestellt. Kommunikation überprüft. | Netzwerk Konfiguration telefonisch ermittelt. Netzwerkverbindung hergestellt, Verbindung per Ping überprüft, Time Server eingestellt. Daten auf CF Karte übertragen. LIS und Aqure Verbindungen hergestellt. Kommunikation überprüft. | SWO-5710 | No | network configuration by telephone. Network connection established, connection checked by ping, Time Server set. Transfer data to CF card. LIS and Aqure connections. communication checked. | network configuration by telephone. Network connection established, connection checked by ping, Time Server set. Transfer data to CF card. LIS and Aqure connections. communication checked. | Setup | HIS/LIS
ABL90 | Calibrada hemoglobina y comprobado buen funcionamiento del equipo. |  | CE0B9236B62FB48A | No | Calibrated hemoglobin and tested proper operation of the equipment. |  | Adjustment | tHb cal
ABL817 FLEX | Geräteüberprüfung->Gerät ohne Fehler | Geräteüberprüfung->Gerät ohne Fehler | SWO-5607 | No | Device verification>device without errors | Device verification>device without errors
ABL825 FLEX | Tried changing the inlet gasket and decontamination programs-No change | Flushed waste lines and replaced waste pump tubings | CE0BDF17B64050A1 | No | Tried changing the inlet gasket and decontamination programs-No change | Flushed waste lines and replaced waste pump tubings
ABL825 FLEX | Skiftet Lak elektrode og Membran Service complete og tHb Kontroller. HB kontrol fra sysmex | Skiftet Lak elektrode og Membran Service complete og tHb Kontroller. HB kontrol fra sysmex | SWO-5580 | No | Changed varnish electrode and membrane Service complete and tHb controls. HB control by at | Changed varnish electrode and membrane Service complete and tHb controls. HB control by at
ABL800 BASIC | Abdeckung muss komplett ausgetauscht werden. Keine Anfahrtsberechnung aufgrund Garantie-Halbjahreswartung |  | SWO-5557 | No | Cover must be completely replaced. No travel calculation due to warranty half-year maintenance
ABL820 FLEX | Unit was displaying Inlet and Lower Inlet sensor failure. Replacement of Inlet heater corrected Inlet error, but Lower Inlet sensor still failing. Cleared path around Lowere Inlet to see if fluid was causin error. Cleared lines and unit still displaying Lower Inlet sensor error. | Ordered pHBG PCB. | CE0B8349B6479D69 | No | Unit was displaying Inlet and Lower Inlet sensor failure. Replacement of Inlet heater corrected Inlet error, but Lower Inlet sensor still failing. Cleared path around Lowere Inlet to see if fluid was causin error. Cleared lines and unit still displaying Lower Inlet sensor error. | Ordered pHBG PCB.
ABL800 FLEX | PC-Einheit ersetzt, Neustart, Funktionskontrolle | PC-Einheit ersetzt, Neustart, Funktionskontrolle | SWO-5565 | No | PC unit replaced, restart, function control | PC unit replaced, restart, function control
ABL80 OSM | QC compairs well but blood does not will change the hemolyzer and rerun samples. | The Thb looks better, Fcohb looks ok, the met remains 2% difference, we will have to look at the 820. | CE0D2EE02C019860 | No | QC compairs well but blood does not will change the hemolyzer and rerun samples. | The Thb looks better, Fcohb looks ok, the met remains 2% difference, we will have to look at the 820.
ABL80 FLEX CO-OX | Anbindung muss von Dorner realisiert werden. Verbindung zu Aqure in Ordnung. | nach dem die Datenabfrage anfänglich einwandfrei funktioniert hatte, funktioniert diese jetzt bereits schon seit über 2 Wochen nicht mehr; dies wurde bereits schon gegenüber Herrn Lünenbürger berichtet, allerdings besteht das Problem nach wie vor; der Anwender wünscht einen Vorort-Service zur Beseitigung der Geräte-Störung; | SWO-5856 | No | Connection must be realized by Dorner. Connection to Aqure all right. | after the data query had initially worked properly, it has not been working for more than 2 weeks; this has already been reported to Mr Lünenbürger, but the problem remains; the user wants a suburban service to eliminate the device malfunction;
ABL90 FLEX | Jeg har kørt blod ind i ABL90 og ladet det stå i apparatet i 2-3 minnutter. Jeg har brugt Visual Flow Check i service, til at køre blod ind i apparatet. Efter dette kan der nu køres ikke humant materiale . Der kommer stadig fejl på pO2, 0521: Uhomogen prøve. Dette skyldes at overfladespændningen på det materiale som man måler på er høj eller lav. Ved at køre blod i mellem de ikke humant materiale prøver burde problemet være løst. |  | SWO-5917 | No | I have driven blood into the ABL90 and left it in the apparatus in 2-3 minnutter. I have used Visual flowcheck in service, to drive blood into the device. After this, no human material can now be run. There are still errors on pO2, 0521: Unhomogeneous Sample. This is because the surface of the material that you measure is high or low. By driving blood in between the non-human material samples, the problem should be solved. |  | Defect | SC | Wetability
ABL90 FLEX | Sichtkontrolle ist ok. Oxy Complete Test ok. Kalibrationen durchgeführt ok. THb Kal ist erneut fehlgeschlagen => Fehler 581 Oxi spectrum mismatch. Ersatzteil muss bestellt werden. | Defekten Hemolyzer ausgetauscht, Startverfahren, Kalibrationen, tHb Kal, Funktionskontrolle | SWO-5838 | No | Visual inspection is ok. Oxy Complete Test ok. Calibrations performed ok. THb Kal failed again => Error 581 Oxi spectrum mismatch. Spare part must be ordered. | Defective Hemolyzer replaced, start-up procedures, calibrations, tHb Kal, function control | Unknown | Revisit required | Hemolyzer
ABL90 FLEX | ABL zerlegt--> Flüssigkeitsaustritt am Abfallanschluss --> ABL gereinigt, Tubing for Valve with Ring erneuert, Station achtet auf erneuten Flüssigkeitsaustritt am ABL Kal. und QC | ABL zerlegt--> Flüssigkeitsaustritt am Abfallanschluss --> ABL gereinigt, Tubing for Valve with Ring erneuert, Station achtet auf erneuten Flüssigkeitsaustritt am ABL Kal. und QC | SWO-5771 | No | ABL disassembles--> Liquid leak at the waste connection --> ABL cleaned, Tubing for Valve with Ring renewed, station pays attention to re-discharge of liquid at The ABL Kal. and QC | ABL disassembles--> Liquid leak at the waste connection --> ABL cleaned, Tubing for Valve with Ring renewed, station pays attention to re-discharge of liquid at The ABL Kal. and QC | Defect | HW | Oxi Module | Pinch Valve tube
ABL90 FLEX | Software stehen geblieben in Spülung notwendig, Gerät neu gestartet, nach Neustart i.O. | Software stehen geblieben in Spülung notwendig, Gerät neu gestartet, nach Neustart i.O. | SWO-5834 | No | Software stopped in flushing necessary, device restarted, after restart i.O. | Software stopped in flushing necessary, device restarted, after restart i.O. | Defect | SC | Wetability
ABL90 FLEX | Nach SC Tausch Kal und QC Fehlermeldungen, bei Eintrefen noch Lac QC fehlerhaft, Kal und QC wiederholt, Anwender vorgehen bei Fehlersuche erläutert | Nach SC Tausch Kal und QC Fehlermeldungen, bei Eintrefen noch Lac QC fehlerhaft, Kal und QC wiederholt, Anwender vorgehen bei Fehlersuche erläutert | SWO-5833 | No | After SC swap Cal and QC error messages, in case of insertion still Lac QC faulty, Kal and QC repeated, user proceed in troubleshooting explained | After SC swap Cal and QC error messages, in case of insertion still Lac QC faulty, Kal and QC repeated, user proceed in troubleshooting explained | Defect | SC | Lac
ABL90 |  |  | CE0CC2662C584CD3 | Yes |  |  | Re classified RMED | Setup / Install
ABL800 BASIC | Fehlersuche Lac-Elektrode membraniert Kalibrierung QC-Messung mit AC5+ Level 2 durchgeführt |  | SWO-5774 | No | Troubleshooting Lac electrode membranes calibration QC measurement performed with AC5+ Level 2
ABL800 BASIC | Fehlersuche Ref.-Elektrode membraniert, 2-Pkt.-Kalibration, QC-Messungen mit AC5+ Level 1 und 3 durchgeführt |  | SWO-5776 | No | Troubleshooting Ref.-electrode membraneed, 2-pt.-calibration, QC measurements performed with AC5+ Level 1 and 3
ABL820 FLEX | Techline indicated this was the sensor in the Inlet module.Replacement of Inlet PCB resolved Sensor issue. Unit still contaminating sample. | Replaced Inlet PCB. | CE0CB6592C37D8A8 | No | Techline indicated this was the sensor in the Inlet module. Replacement of Inlet PCB resolved Sensor issue. Unit still contaminating sample. | Replaced Inlet PCB.
ABL827 FLEX |  |  | CE0E3BF4A20A5942 | Yes |  |  | Re classified RMED | Analyzer setup
ABL80 FLEX CO-OX | Kalibrierung mit anschließender QC in Ordnung. Kein Fehler feststellbar. Einweisung durchgeführt. | Kunde hat keine Sensorkassette mehr vorrätig. Neue SC für 08:00Uhr bestellt. | SWO-5854 | No | Calibration with subsequent QC in order. No error detectable. instruction carried out. | Customer no longer has a sensor cartridge in stock. New SC ordered for 08:00o.
ABL90 FLEX | Abfallschlauch erneuert, Einlassdichtung (war verschmutzt), Einlasszwischendichtung und Schlauch zwischen Messkammer und Oxymetrie erneuert, System- und Funktionskontrolle, System und Funktion i.O. | Ableitstromfehler ist wieder aufgetreten. Hr. Breuer hat gestern das Material wie mit dem Techniker besprochen gewechselt. Hr. Breuer möchte eine Vorortservice und einen Rückruf wann dieser durchgeführt wird. | SWO-5909 | No | Waste hose renewed, inlet seal (was dirty), inlet intermediate seal and hose between measuring chamber and oxymetry renewed, system and function control, system and function i.O. | Leak current error has occurred again. Yesterday, Mr. Breuer changed the material as discussed with the technician. Mr. Breuer wants a suburban service and a recall when it will be performed. | Unknown | Revisit required | Leak current
ABL90 FLEX | Defekten Hemolyzer ausgetauscht, Startverfahren, Kalibrationen, tHb Kal, Funktionskontrolle | Defekten Hemolyzer ausgetauscht, Startverfahren, Kalibrationen, tHb Kal, Funktionskontrolle | SWO-5864 | No | Defective Hemolyzer replaced, start-up procedures, calibrations, tHb Kal, function control | Defective Hemolyzer replaced, start-up procedures, calibrations, tHb Kal, function control | Defect | HW | Hemolyzer
ABL90 FLEX | Gerät überprüft, Gerät te zeigen bei diesem Ringversuch OXI Mismatch Fehler an, es muss im Pat.bericht die Unterdrückung ausgeschltet werden und somit werden die Ergebnisse ausgegeben, nach dem Ringversuch muss das wieder eingeschaltet werden. | mit Herrn Haase telefoniert, haben neues Ringversuch mat. betsellt, er meldet sich wieder bei mir wenn es da ist, am ABL muß dann in Unterdrückung ausgestellt werden und dann nochmal gemessen werden, die Ergebnisse zu Ingo Sprenger schicken..die Messfehler kommen durch das Tierblut als Testblut | SWO-5855 | No | Device checked, device te indicate errors in this ring test OXI Mismatch, the suppression must be eliminated in the Pat.report and thus the results are output, after the ring test this must be switched on again. | with Mr. Haase on the phone, have new ring test mat. Betsellt, he reports back to me when it is there, then the ABL has to be exhibited in suppression and then measured again, send the results to Ingo Sprenger. the measuring errors come through the animal blood as test blood | No fault found | Ring test
ABL90 FLEX | Das gesamte System neu aufgesetzt. Gesamte Einstellungen erneuert. | Der Kunde bittet um einen dringenden Rückruf, wann ein Techniker zur Reparatur kommt. | SWO-6062 | No | The entire system is redesigned. Entire settings are renewed. | The customer asks for an urgent recall when a technician comes to repair. | Defect | SW | Status code 2
ABL90 |  |  | CE107CF88E4DD9EE | No |  |  | Defect | SC | Wetability
ABL90 FLEX | Mit Herrn Beider die LP Wechsel besprochen. Die Wechsel sind verbrauchsgerecht. Klemmschlauch erneuert, tHb Kal. durchgeführt. Herrn Beider eingewiesen. | Immer wieder fordert der ABL90 zum SP Wechsel auf. OA Dr. Beider hat möchte jetzt keinen SP mehr wechseln. Er verlangt einen Vorortservice. Er hat auch kein SP mehr im Vorrat. | SWO-5890 | No | Discussed with Mr. Beider the LP change. The bills of exchange are fuel-efficient. Clamping hose renewed, tHb Kal. Performed. Mr. Beider instructed. | Again and again, the ABL90 calls for a change to the SP. OA Dr. Beider does not want to change any more SP. He requires a suburban service. He also no longer has an SP in stock. | Defect | HW | Oxi Module | Pinch Valve tube | SC (SP replaced several times)
ABL825 FLEX | Random oxi spectrum mismatch on QC. 3-4 per week, almost all on Blue QC. | Replaced hemolyzer and completed lamp adjustment and total Hb cal. Cals and Qc checking OK. Blood leaking around hemolyzer - left part at cust site | CE0E499AA224A135 | No | Random oxi spectrum mismatch on QC. 3-4 per week, almost all on Blue QC. | Replaced hemolyzer and completed lamp adjustment and total Hb cal. Cals and Qc checking OK. Blood leaking around hemolyzer - left part at cust site
ABL835 FLEX | Checked the activity log and found: message 248, Spectrometer missing trigger, 258 Neon lamp intensity too low and, 316 Neon number too low. | Replaced a new Oximetry module PCB board. Checked and adjusted the hemolyzer halogen lamp. Performed a Complete Oximetry test, passing within ranges. Performed a service and, customer mode ctHb calibration, passing within specifications. | CE0E2626A22A469C | No | Checked the activity log and found: message 248, Spectrometer missing trigger, 258 Neon lamp intensity too low and, 316 Neon number too low. | Replaced a new Oximetry module PCB board. Checked and adjusted the hemolyzer halogen lamp. Performed a Complete Oximetry test, passing within ranges. Performed a service and, customer mode ctHb calibration, passing within specifications.
ABL837 FLEX | The analyzer is not printing legible when patient results are printed. | Replaced a new printer unit. | CE0E290EA23B46A0 | No | The analyzer is not printing legible when patient results are printed. | Replaced a new printer unit.
ABL835 FLEX | A L'arrivé l'appareil était fonctionnel sauf pompe calibration error | test gas modules ok reset gas modules reinstallé clip pompe solution calibration pompe ok | CE0DEAA5A233CF13 | No | On arrival the device was functional except pump calibration error | test gas modules ok reset gas modules reinstalled clip pump solution calibration pump ok
ABL800 BASIC | Papier wurde eingelegt ohne Kunststoffrolle zu entfernen, Deckel wurde zu stark heruntergedrückt. CPU ausgebaut und Monitor entfernt, Deckel erneuert. Im Anschluss Garantie-Halbjahreswartung | Papier wurde eingelegt ohne Kunststoffrolle zu entfernen, Deckel wurde zu stark heruntergedrückt. CPU ausgebaut und Monitor entfernt, Deckel erneuert. Im Anschluss Garantie-Halbjahreswartung | SWO-5955 | No | Paper was loaded without removing plastic roll, lid was pushed down too much. CPU removed and monitor removed, lid renewed. Following warranty half-year maintenance | Paper was loaded without removing plastic roll, lid was pushed down too much. CPU removed and monitor removed, lid renewed. Following warranty half-year maintenance
ABL825 FLEX | Gerät-Elektrode und Kontakt ok E788 mit neuer Membran anderes Lot Membraniert.. | Gerät-Elektrode und Kontakt ok E788 mit neuer Membran anderes Lot Membraniert.. | SWO-6135 | No | Device electrode and contact ok E788 with new membrane other solder membranes. | Device electrode and contact ok E788 with new membrane other solder membranes.
ABL825 FLEX | Level 1 ist sehr tief unterer Grenze. Messkammer sauber Kontakte ok..Elektrode mit anderer Lot 465 Membraniert ZW Überprüfung S7735--Lot539 laut Ingo Sprenger ok Zur weiteren Beobachtung. | PCO2-Membran gereinigt. Schwarzen O-Ring pH-Elektrode durch blauen ersetzt. Kalibrationen und Qualitätskontrollen in Ordnung. Einweisung Herr Dinkel durchgeführt. | SWO-5980 | No | Level 1 is very low lower limit. Measuring chamber clean contacts ok.. Electrode with other lot 465 membraneed ZW check S7735--Lot539 according to Ingo Sprenger ok For further observation. | PCO2 membrane. Black O-ring pH electrode replaced by blue. calibrations and quality controls in order. Instruction carried out by Mr. Dinkel.
ABL825 FLEX | Data logs show waste full after replacement of bottle. Detector is failing. | Replaced detector and tested opertaion. Instrument operats normally. | CE0E1AB9A228A4E3 | No | Data logs show waste full after replacement of bottle. Detector is failing. | Replaced detector and tested opertaion. Instrument operats normally.
AQT90 FLEX |  |  | CE0DFDA1A21B3A08 | No
AQT90 FLEX |  |  | CE1215C2040847DC | No
ABL90 |  |  | CE1076A98E2B3D06 | Yes |  |  | Re classified RMED | Installation
ABL825 FLEX |  |  | CE1067A08E2D6875 | No
ABL90 |  |  | CE16D39CDC073234 | Yes |  |  | Re classified RMED | Phone support
ABL90 FLEX | Jeg har skiftet alle slanger i ABL90. Der er udført det der svarer til et årsservice. Jeg har startet ABL90 op igen med ny sensorkassette. Jeg har kørt alle test i service, Fluid transport og de har alle været OK. ABL90 er testet med blod, mælk og kundens buffer der bruges til test. Alle test er kørt uden problemer. Alt er OK igen. | Jeg har skiftet alle slanger i ABL90. Der er udført det der svarer til et årsservice. Jeg har startet ABL90 op igen med ny sensorkassette. Jeg har kørt alle test i service, Fluid transport og de har alle været OK. ABL90 er testet med blod, mælk og kundens buffer der bruges til test. Alle test er kørt uden problemer. Alt er OK igen. | SWO-6070 | No | I changed all the snakes in ABL90. This is the equivalent of an annual service. I started ABL90 up again with new sensor cassette. I have run all tests in service, Fluid transport and they have all been OK. ABL90 is tested with blood, milk and the customer's buffer used for testing. All tests have been run smoothly. Everything is OK again. | I changed all the snakes in ABL90. This is the equivalent of an annual service. I started ABL90 up again with new sensor cassette. I have run all tests in service, Fluid transport and they have all been OK. ABL90 is tested with blood, milk and the customer's buffer used for testing. All tests have been run smoothly. Everything is OK again. | Defect | SC | Wetability
ABL90 FLEX | Tausch des Hemplyzers, hat Erfolg. Keine Fehlermeldung mehr. tHb-Kalibration mit eigenem Kalibrator durchgeführt. | Tausch des Hemplyzers, hat Erfolg. Keine Fehlermeldung mehr. tHb-Kalibration mit eigenem Kalibrator durchgeführt. | SWO-6091 | No | Swap the Hemplyzer, succeeds. No more error messages. tHb calibration with its own calibrator. | Swap the Hemplyzer, succeeds. No more error messages. tHb calibration with its own calibrator. | Defect | HW | Waste tube
ABL90 | checked data logs for errors, checked flow of solutions, lactate drifting for 2-3 days | replaced solution pack, changed sensor cassette, ran 6-7 waste bloods using safe pico, left addidtional sensor to complete validation | CE124BAA0454397C | No | checked data logs for errors, checked flow of solutions, lactate drifting for 2-3 days | replaced solution pack, changed sensor cassette, ran 6-7 waste bloods using safe pico, left addidtional sensor to complete validation | Defect | SC | Lac
ABL810 FLEX | Anschlussnippel Einlassdichtung gespült. Einlassdichtung aus Kundenbestand erneuert. | Mit dem Kunden Optionen besprochen, Nadel lt. Kunden i.O., neue Einlassdichtung nicht vorhanden, Kunde wünscht Kontaktaufnahme durch zuständigen FSE am Montag | SWO-5984 | No | Connection nipple inlet seal rinsed. Inlet seal from customer stock renewed. | Discussed with the customer options, needle according to customer i.O., new inlet seal not available, customer wishes contact by competent FSE on Monday
ABL835 FLEX | Gas erreur intermittant sur les deux gas gas calibration hors limite v gas module OK | remplacé analyser controles PCB. reinstallé logiciel 6.13 installé clé de l'appareil | CE11B61A04198C2C | No | Gas error intermittently on both gas calibration off limit v gas module OK | replaced analyze PCB controls. reinstalled software 6.13 installed key of the device
ABL815 FLEX | Druckerkopf erneuert. pH Dichtungsring erneuert. Gerät kalibriert und QC durchgeführt. |  | SWO-6137 | No | Printer head renewed. pH sealing ring renewed. calibrated and QC performed.
ABL825 FLEX |  |  | CE121897041F47E1 | No
ABL820 FLEX | Waste roller had the pin for roller dropping down out of position from inside roller assy. This caused the Pin to drag on plastic Fluidic assy.Also located leak in Inlet assy. | Successful running of QC sample without fliud contamination. | CE11CF41041FB500 | No | Waste roller had the pin for roller dropping down out of position from inside roller assy. This caused the Pin to drag on plastic Fluidic assy. Also located leak in Inlet assy. | Successful running of QC sample without fliud contamination.
ABL825 FLEX | System überprüft, Einlaßsonde und Einlassdichtung erneuert, System gereinigt, Kalibration und blutige Probe ohne Beanstandungen durchgeführt. | System überprüft, Einlaßsonde und Einlassdichtung erneuert, System gereinigt, Kalibration und blutige Probe ohne Beanstandungen durchgeführt. | SWO-6124 | No | System checked, inserting probe and inlet seal renewed, system cleaned, calibration and bloody sample carried out without complaints. | System checked, inserting probe and inlet seal renewed, system cleaned, calibration and bloody sample carried out without complaints.
ABL825 FLEX | Transportschutz aus dem Flex-Q Modul entfernt - Modul getestet - OK | Transportschutz aus dem Flex-Q Modul entfernt - Modul getestet - OK | SWO-6078 | No | Transport protection removed from Flex-Q module - Module tested - OK | Transport protection removed from Flex-Q module - Module tested - OK
ABL80 FLEX BASIC | CPU Unit csere, Szoftver telelpítés, QC mérés. Csere készülék visszavétele. |  | SWO-6256 | No | CPU Unit replacement, software loading, QC measurement. Return the replacement device.
ABL90 FLEX | Der Kunde ist nicht bereit vier tage auf den ABL zu verzichten. LP in der GS abgeholt und beim Kunden installiert. Dem Kunden mitgeteilt, dass er immer je2 Kassetten und 2Lp im Vorrat halten soll. Startphase QC und Kalibration. | Nach Absprache mit der Bestellabteilung ist eine Nachtexpresslieferung nicht möglich. Die Bestellabteilung versuch die Bestellung per REDLINE zu bestellen, allerdings ist es dafür schon zu spät. Der Kunde muss deswegen min. 4Tagen warten und ist solange nicht meßbereit. Darum bittet der Kunde den Liefermodus zu überprüfen und so zu organisieren, dass ihm ein LP immer zur Verfügung steht. Der Kunde bittet um einen Rückruf, um die Lagerhaltung des LPs zu optimieren. | SWO-6187 | Yes | The customer is not prepared to give up the ABL for four days. LP in the GS and installed at the customer's station. The customer is informed that he should always keep 2 cassettes and 2Lp in stock. Start phase QC and calibration. | After consultation with the ordering department, night express delivery is not possible. The ordering department tries to order the order via REDLINE, but it is already too late for this. The customer therefore has to wait at least 4 days and is not ready for measurement until then. Therefore, the customer asks to check the delivery mode and organize it in such a way that an LP is always available to him. The customer asks for a callback in order to optimize the storage of the LP. | Re classified RMED | Logistical upSWO
ABL5 | Ref. Elektrode 53-04 erneuert durch 2102. |  | SWO-6243 | No | Ref. Electrode 53-04 renewed by 2102.
ABL825 FLEX | De analyser gaf bij aankomst de melding tussenkomst vereist . Bij controle van de kalibratiegegevens bleek dat er geen Cal1-vloeistof werd opgepompt. De oorzaak hiervoor was dat het flesje van de Cal1 leeg was. Dit werd niet aangegeven door het verkeerslicht omdat er een correctie werd ingesteld voor de Cal1 (stond op 20% terwijl deze leeg was). Na het vervangen van de Cal1 werden de kalibraties weer volledig uitgevoerd. |  | SWO-6171 | No | The analyser required the intervention notification upon arrival. Checking the calibration data showed that no Cal1 liquid was pumped up. The cause for this was that the bottle of the Cal1 was empty. This was not indicated by the traffic light because a correction was set for the Cal1 (stood at 20% while it was empty). After replacing the Cal1 the calibrations were completed again.
ABL835 FLEX | The printer paper level is not locking so the pronyer cannot feed paper. | Replaced a new printer unit. | CE13066B7A4CEF24 | No | The printer paper level is not locking so the pronyer cannot feed paper. | Replaced a new printer unit.
AQT90 FLEX | Neuen Pack Kunden vorbeigebracht und gewechselt. Danach war der AQT wieder messbereit. Kunden noch einige Fragen zum Gerät beantwortet. | Nach den Fehlern von gestern (siehe SWO) wird nun die Puffer- Abfallkassette nicht erkannt. Fr Mariano möchte einen Vorortservice. | SWO-6270 | No | New Pack customers brought in and changed. After that, the AQT was ready for measurement again. customers still answered some questions about the device. | After yesterday's errors (see SWO), the buffer waste cartridge is now not recognized. Fr Mariano wants a suburban service.
ABL80 FLEX Basic version | Comprobado tubos y funcionaminento del equipo. El equipo pasa las calibraciones y QCs externos. |  | CE1487D2F0024042 | No | Checked tubes and equipment construction. The equipment passes external calibrations and QCs.
ABL80 FLEX Basic version |  |  | CE14C172F00C2C70 | No
ABL90 FLEX | Hemolyzer, Silikon Disk und Klemmschlauch erneuert. tHb Kal. durchgeführt. Fehler 0581: OXI Spectrum mismatch bei tHb Kal. aufgetreten. original Hemolyzer überprüft, Lötpunkt am Hemolyzer Draht gebrochen. Hemolyzer nachgelötet und zurück gebaut. OXI LS Kabel eingeklemmt, OXI LS erneuert. Neuen Hemolyzer bestellt. tHb Kal. und QCs i.O. Blutmessungen ohne Fehler, Austausch des Hemolyzers geplant. |  | SWO-6388 | No | Hemolyzer, silicone disc and clamping hose renewed. tHb Kal. Performed. Error 0581: OXI Spectrum mismatch at tHb Kal. Occurred. original Hemolyzer checked, soldering point broken on Hemolyzer wire. Hemolyzer soldered and built back. OXI LS cable clamped, OXI LS renewed. New Hemolyzer ordered. tHb Kal. and QC's i.O. blood measurements without errors, replacement of the hemolyzer planned. |  | Defect | HW | Hemolyzer
ABL90 |  |  | CE148598F0443E59 | No
ABL90 |  |  | CE14AE07F01A034B | No
ABL90 | Need to remote into ABL90 and check setup. | Remote to analyzer after installing NetOp on radiance server. After several phone calls and looking at setup of ABL90, it was determined that Cerner was wasn't creating barcodes for production enviorment. Test enviornment barcodes scan properly. Cerner person made changes to Prod barcode and recycled server. Could not correct badge barcode at this time. Need to research some more. | CE147F57F0073B18 | No | Need to remote into ABL90 and check setup. | Remote to analyzer after installing NetOp on radiance server. After several phone calls and looking at setup of ABL90, it was determined that Cerner was wasn't creating barcodes for production enviorment. Test enviornment barcodes scan properly. Cerner person made changes to Prod barcode and recycled server. Could not correct badge barcode at this time. Need to research some more. | Setup | Barcode setup
ABL90 | Descontaminação do circuito, substituição da entrada e da sonda, vários testes, establização e passagem do QC. Tudo OK |  | CE149AFDF04C5861 | No | Circuit decontamination, replacement of inlet and probe, various tests, establishment and passage of QC. Everything OK |  | Cleaned | Fluidic pathway
ABL835 FLEX | REVISIÓN GENERAL DE TODO EL SISTEMA CON LIMPIEZA Y DESCONTAMINACION DE CIRCUITOS. ENGRASE Y LIMPIEZA DE PIEZAS MOVILES Y ENGRANAJES. COMPROBACIONES DE CONECTIVIDAD HIS SIL. BACKUP DE CONFIGURACIONES DEL SISTEMA. LIMPIEZA Y O SUSTITUCION DE FILTROS. TEST DE FUNCIONAMIENTO ELECTRÓNICO Y NEUMÁTICO. SUTITUCION DE PIEZAS DETERIORADAS POR USO. CALIBRAR Y VERIFICAR MEDIANTE CONTROL DE CALIDAD Q.C. | CABEZAL IMPRESORA | CE146C5BF04F9B40 | No | GENERAL REVIEW OF ALL SYSTEM WITH CLEANING AND DISCONTAMINATION OF CIRCUITS. WELCOME AND CLEANING OF MOVIL PARTS AND GRAINS. HIS SIL CONNECTIVITY CHECKS. BACKUP OF SYSTEM SETTINGS. CLEANING AND REPLACEMENT OF FILTERS. ELECTRONIC AND TIRE OPERATION TEST. REPLACEMENT OF DETERIORATED PARTS BY USE. CALIBRATE AND CHECK BY Q.C QUALITY CONTROL. | PRINTER HEAD
ABL835 FLEX | DATA LOGS SHOWS 356 EL MET TEMP ERROR DIAGNOSTICS SHOWS TEMPWERATURE AT 32 DEGREES | REPLACED EL MET MODULE.TEMP AT 37 DEGREES RAN CALIBRATIONS | CE144413F0171AA2 | No | DATA LOGS SHOWS 356 EL MET TEMP ERROR DIAGNOSTICS SHOWS TEMPWERATURE AT 32 DEGREES | REPLACED EL MET MODULE. TEMP AT 37 DEGREES RAN CALIBRATIONS
ABL825 FLEX | Option Key installiert | Option Key installiert | SWO-6409 | No | Option Key installed | Option Key installed
ABL800 BASIC | Neue Elektrode | Neue Elektrode | SWO-6578 | No | New electrode | New electrode
AQT90 FLEX | NT-proBNP-Kassette, Lot 09792 wurde nicht zurückgezogen, da eine Kunststoffnase der Kassette abgebrochen war. Messingschiene überprüft, war bereits V. 3.2. Führung der Kassetten auf Leichtgängigkeit überprüft. Gerät mit QC getestet. |  | SWO-6417 | No | NT-proBNP cassette, lot 09792 was not withdrawn because a plastic nose of the cassette was broken off. brass rail, V. 3.2. Guide of the cassettes was checked for smoothness. Device tested with QC.
ABL80 FLEX Basic version |  |  | CE15C48E661C58AB | No
ABL80 CO-OX | the problem is intermittent so the spectrometer may be unstable, the hemolyzer has been flushed. | replaced spectrometer and ran THB cal and replace SP. | CE15C669663BD357 | No | the problem is intermittent so the spectrometer may be unstable, the hemolyzer has been flushed. | replaced spectrometer and ran THB cal and replace SP.
ABL90 FLEX | Defektes Neuteil Hemolyzer Unit 902-980R0467N001 erneuert durch 902-980R0467N003. Probelauf, tHb Kal., QCs und Blutmessungen i.O. |  | SWO-6455 | No | Defective new part Hemolyzer Unit 902-980R0467N001 renewed by 902-980R0467N003. Test run, tHb Cal., QCs and blood measurements in the |  | Defect | HW | Hemolyzer
ABL90 |  |  | CE15913E6613711D | No |  |  | Defect | HW | Oxi Module | Pinch Valve tube
ABL90 FLEX | Gerät gespült und überprüft,OK | Gerät gespült und überprüft,OK | SWO-6494 | No | Device flushed and checked,OK | Device flushed and checked,OK | Cleaned | Fluidic pathway | Wetability
ABL835 FLEX | Druckerkopf erneuert, Lampe Unit justiert. Gerät kalibriert und QC durchgeführt. |  | SWO-6617 | No | Printer head renewed, lamp unit adjusted. calibrated and QC performed.
ABL827 FLEX | 602 inlet not filled during rinse, 501 rinse not aspirated, 508 rinse error, 923, 340 inlet not filling or emptying within time limit and other errors. | Replaced friction detector in inlet module. Cleaned all electrical contact points for pO2 channel. | CE15995C66166350 | No | 602 inlet not filled during rinse, 501 rinse not aspirated, 508 rinse error, 923, 340 inlet not filling or emptying within time limit and other errors. | Replaced friction detector in inlet module. Cleaned all electrical contact points for pO2 channel.
ABL90 |  |  | CE16CA50DC45936A | No |  |  | Defect | HW | Probe bend
ABL90 FLEX | Abl geht immer wieder in geführte Fehlersuche --> Flüssigkeitsstrecke gespült --> verstopft Abfallschlauch verstopft, erneuert AbL nochmals gespült , Neustart i.O Probennummer 9285 hatte das Problem verursacht | Abl geht immer wieder in geführte Fehlersuche --> Flüssigkeitsstrecke gespült --> verstopft Abfallschlauch verstopft, erneuert AbL nochmals gespült , Neustart i.O Probennummer 9285 hatte das Problem verursacht | SWO-6628 | No | Abl goes again and again in guided troubleshooting --> liquid section flushed --> clogged waste hose clogged, renewed AbL rinsed again, restart i.O sample number 9285 had caused the problem | Abl goes again and again in guided troubleshooting --> liquid section flushed --> clogged waste hose clogged, renewed AbL rinsed again, restart i.O sample number 9285 had caused the problem | Cleaned | Fluidic pathway | Clot
ABL90 FLEX | Lösungspack, Dichtung und SC liefern. Dichtungen, Lösungspack, SC und Einlassnadel ersetzt, scrubbing with water Kal, QC |  | SWO-6663 | No | solution pack, seal and SC. Seals, solution pack, SC and inlet needle replaced, scrubbing with water Kal, QC |  | Cleaned | Fluidic pathway
ABL825 FLEX |  |  | CE16D509DC473237 | No
ABL835 FLEX | Problême varié sur appareil intermittant depuis le debut de l'installation client completement insatisfait de ABL voudrais un remplacement le client utilise pas ABL il ne le fais pas confiance | Replace CPU and Power supply techline will be open | CE16C311DC33BFF7 | No | Varied problem on device intertermitating since the beginning of the customer installation completely dissatisfied with ABL would like a replacement the customer uses not ABL he does not trust it | Replace CPU and Power supply techline will be open
ABL820 FLEX | Unit behaving in same manner with fluid leftover in Chamber contaminating Sample. | Using Modules from old ABL to troubleshoot problem. After replacement of Inlet and pHBG modules unit back to Operational. | CE1712A3DC182A20 | No | Unit behaving in same manner with fluid leftover in Chamber contaminating Sample. | Using Modules from old ABL to troubleshoot problem. After replacement of Inlet and pHBG modules unit back to Operational.
ABL837 FLEX | Several membrane changes does not help | Replaced defective Crea A electrode, sensitivity is now normal | CE16E47FDC3BE81F | No | Several membrane changes does not help | Replaced defective Crea A electrode, sensitivity is now normal
AQT90 FLEX | JW nach R-DEU spezifik. durchgeführt, Reinigungsprogr. durchgeführt, serv.dump gezogen, | JW nach R-DEU spezifik. durchgeführt, Reinigungsprogr. durchgeführt, serv.dump gezogen, | SWO-6655 | No | JW according to R-DEU specific. carried out, cleaning progr. performed, serv.dump pulled, | JW according to R-DEU specific. carried out, cleaning progr. performed, serv.dump pulled,
ABL80 CO-OX | Bad Hemolyzer will replace | Replaced Hemolyzer and ran THB cal, all passed. Ran 5 blood samples on three different bloods, 15 total and compared all three to the GEM. ok. | CE1842605236D366 | No | Bad Hemolyzer will replace | Replaced Hemolyzer and ran THB cal, all passed. Ran 5 blood samples on three different bloods, 15 total and compared all three to the GEM. ok.
ABL80 CO-OX | Performed a zero check and a system cycle. pO2 sensitivity was at 0.025 before & after. Ran one ampoule of all three levels of the calibration verification kit and found that pO2 was out low on level one & high on level three. | Replaced the analog board, then ran a system cycle & pO2 sensitivity is still at 0.025. I also ran one ampoule of each level again, but level three was still high. Customer is low on sensor cassettes, so I need to order one to troubleshoot with. Analyzer is operational. | CE1ADC0C3E623099 | No | Performed a zero check and a system cycle. pO2 sensitivity was at 0.025 before & after. Ran one ampoule of all three levels of the calibration verification kit and found that pO2 was out low on level one & high on level three. | Replaced the analog board, then ran a system cycle & pO2 sensitivity is still at 0.025. I also ran one ampoule of each level again, but level three was still high. Customer is low on sensor cassettes, so I need to order one to troubleshoot with. Analyzer is operational.
ABL90 |  |  | CE1A8FF03E1DEFAE | No |  |  | Cleaned | Fluidic pathway
ABL90 |  |  | CE1AC41D3E4EF769 | No |  |  | Defect | HW | SC ID PCB | Temp
ABL90 FLEX | Einweisung zum Tema Preanalytik die verwendung von Mischdrähtchen gezeigt, Blutmessungen durchgeführt. Zusammen mit Frau Deichert Dichtungstausch praktiziert etc. | Einweisung zum Tema Preanalytik die verwendung von Mischdrähtchen gezeigt, Blutmessungen durchgeführt. Zusammen mit Frau Deichert Dichtungstausch praktiziert etc. | SWO-6996 | No | Instruction to Tema preanalysis shown the use of mixed wires, blood measurements performed. Together with Ms. Deichert, poetry exchange practices etc. | Instruction to Tema preanalysis shown the use of mixed wires, blood measurements performed. Together with Ms. Deichert, poetry exchange practices etc. | User Error | Pre analytical
ABL90 FLEX | Analysator gespült, SP erneuert, Neustart, Kal und QC iO. |  | SWO-6921 | No | Analyzer rinsed, SP renewed, restart, Kal and QC iO. |  | Cleaned | Fluidic pathway
ABL90 |  |  | CE1AE2863E16A5EB | Yes |  |  | Re classified RMED | No info
ABL90 | Performed visual inspection of wet section, inlet probe, solution pack. No issues noted. Activity log indicates several messages 1340 1396 regarding sensor cassette requires replacement. Sensibility values not stable for Lac, internal QC out of range. | Swapped sensor cassette from another ABL 90 that is performing to specifications. Problem moved with cassette, confirming issue. Replacement sensor cassettes have been ordered by Sales to arrive next day. Customer will install replacement sensor cassette. | CE1AA0373E2BE740 | No | Performed visual inspection of wet section, inlet probe, solution pack. No issues noted. Activity log indicates several messages 1340 1396 regarding sensor cassette requires replacement. Sensibility values not stable for Lac, internal QC out of range. | Swapped sensor cassette from another ABL 90 that is performing to specifications. Problem moved with cassette, confirming issue. Replacement sensor cassettes have been ordered by Sales to arrive next day. Customer will install replacement sensor cassette. | Defect | SC | Lac
ABL837 FLEX |  |  | CE1F84E7164FE2E3 | No
ABL825 FLEX | Checked waste system-found waste pump roller sticking Checked QC plots | Replaced waste pump and waste pump tubing's Flushed waste system Replaced pH sealing ring | CE1A83113E0301E0 | No | Checked waste system-found waste pump roller sticking Checked QC plots | Replaced waste pump and waste pump tubing's Flushed waste system Replaced pH sealing ring
ABL825 FLEX |  | Replaced ref electrode | CE23EBE87847DF94 | No |  | Replaced ref electrode
ABL815 FLEX | Customer recently remembraned all electrodes.Believe other issues were at root of cal issues.Blockage,tubing, etc. | Cleaned all contacts and electrodes, Found debris residue on Reference connector arm contact. | CE1AA5FA3E1617C8 | No | Customer recently remembraned all electrodes. Believe other issues were at root of cal issues. Blockage,tubing, etc. | Cleaned all contacts and electrodes, Found debris residue on Reference connector arm contact.
ABL825 FLEX | Pump asymmetry error has been occurring intermittently since 11-18. Verified pH BG & fluidic modules valve operation. Liquid sensors and inlet operation is normal. Pump rollers and tubings are in good condition. | Replaced pump tubings and swapped pump rollers between pH BG & EL MET modules. Pump calibration is normal. Calibrated and ran QC. Instrument operates normally. | CE1AAB633E141985 | No | Pump asymmetry error has been occurring intermittently since 11-18. Verified pH BG & fluidic modules valve operation. Liquid sensors and inlet operation is normal. Pump rollers and tubings are in good condition. | Replaced pump tubings and swapped pump rollers between pH BG & EL MET modules. Pump calibration is normal. Calibrated and ran QC. Instrument operates normally.
AQT90 FLEX |  |  | CE1AB57C3E3AF023 | Yes |  |  | Re classified RMED | Moved analyzer
AQT90 FLEX | Einstellungen sind OK, es funktioniert alles wie es soll, Anwender hat keine Beanstandung mehr, mit EDV, DR. und MicKo telefoniert zwecks Datenübertragung Von seitens Radiometer ist alles richtig eingestellt Systemreinigung durchgeführt Systemüberprüfung OK QC Messung durchgeführt | Einstellungen sind OK, es funktioniert alles wie es soll, Anwender hat keine Beanstandung mehr, mit EDV, DR. und MicKo telefoniert zwecks Datenübertragung Von seitens Radiometer ist alles richtig eingestellt Systemreinigung durchgeführt Systemüberprüfung OK QC Messung durchgeführt | SWO-6765 | No | Settings are OK, everything works as it should, user no longer has complaint, with COMPUTER, DR. and MicKo phoned for data transmission From by radiometer everything is properly set system cleaning system check OK QC measurement Performed | Settings are OK, everything works as it should, user no longer has complaint, with COMPUTER, DR. and MicKo phoned for data transmission From by radiometer everything is properly set system cleaning system check OK QC measurement Performed
ABL80 CO-OX | Inspected fluid pathing, pumps, etc. No apparent problem. | Power failed unit. Upon reboot error was able to clear.Cleaned, purged lines. | CE1BFCB9B40DA9F9 | No | Inspected fluid pathing, pumps, etc. No apparent problem. | Power failed unit. Upon reboot error was able to clear. Cleaned, purged lines.
ABL90 FLEX | Neuen Lösungspack eingesetzt. Einlassnadel erneuert. | Der Kunde möchte keine weitere telefonische Unterstützung. Der Kunde bittet um einen Service. Der Kunde möchte informiert werden, wann ein Techniker kommt. | SWO-7004 | No | new solution pack. Inlet needle renewed. | The customer does not want any further telephone support. The customer asks for a service. The customer wants to be informed when a technician is coming. | Replaced | Probe | Bend
ABL90 |  |  | CE2353DF78421212 | No |  |  | Defect | SP | FT | Wetability ?
ABL90 FLEX | Allgemeine Kontrolle ok. Softwareupdate auf Version 3.1 MR1 ok. System i.o. Kal. + QC i.o. |  | 2851 | No | General control ok. Software update to version 3.1 MR1 ok. System i.o. Kal. + QC i.o.
ABL825 FLEX | Sostituito modulo BGa da testare in sede. Calibrazione e controlli di qualità in ordine |  | 4903 | No | Replaced BGa module to be tested on-site. Calibration and quality controls in order
ABL825 FLEX | checked logs | none calibration frequency at 0.5 and 8 hrs. open techline case | CE1C34BBB460F132 | No | checked logs | none calibration frequency at 0.5 and 8 hrs. open techline case
ABL835 FLEX | gas erreur sur les deux bonbonnes gas flow =0 sur le modules | replaced dual valve housing rinse valve and air valve BG | CE1C0A38B44DC571 | No | gas error on both gas flow canisters on the modules | replaced dual valve housing rinse valve and air valve BG
ABL825 FLEX |  | Lamp unit adjusted | CE23EE4B780FDF99 | No |  | Lamp unit adjusted
ABL820 FLEX | Found that during a rinse the last column of rinse would backed up in to chamber, due to the probe being misaligned. | Performed inlet calibration | CE1D18492A502675 | No | Found that during a rinse the last column of rinse would backed up in to chamber, due to the probe being misaligned. | Performed inlet calibration
ABL825 FLEX | Autocheck barcode reader failing to read barcodes on QC ampoules. Checked all cable connections associated with Autocheck PCB and Laseralignment.Replaced FPC cable, issue persisted. Ordered Autocheck PCB and wiil return on 11-25-2014 to replace. | Replaced Autocheck PCB Laser scanner continued to read intermittently. Replaced Laser Scanner. Scanned Autocheck QC multiplr times with no errors. System operational. | CE1BC841B44E8F5F | No | Autocheck barcode reader failing to read barcodes on QC ampoules. Checked all cable connections associated with Autocheck PCB and Laseralignment.Replaced FPC cable, issue persisted. Ordered Autocheck PCB and wiil return on 11-25-2014 to replace. | Replaced Autocheck PCB Laser scanner continued to read intermittently. Replaced Laser Scanner. Scanned Autocheck QC multiplr times with no errors. System operational.
AQT90 FLEX | Due to an oscillation of the mains, the CPU was damaged on a cable |  | CE24DF8DEE301772 | No | Due to an oscillation of the mains, the CPU was damaged on a cable
ABL80 FLEX Basic version |  |  | CE1E6EF8A033CFD8 | No
ABL80 FLEX Basic version |  |  | CE1D75192A27869B | No
ABL90 |  |  | CE1E7021A057CFDA | No |  |  | No fault found | SC | Wetability
ABL90 FLEX | system mit Blut versorgt. mehrfach neu gestartet. Lösungspack ausgetauscht. Sensorkasette ausgetauscht. | system mit Blut versorgt. mehrfach neu gestartet. Lösungspack ausgetauscht. Sensorkasette ausgetauscht. | SWO-7047 | No | system is supplied with blood. restarted several times. solution pack. Sensor catheter replaced. | system is supplied with blood. restarted several times. solution pack. Sensor catheter replaced. | Defect | SC | Wetability
ABL825 FLEX | Allgemeine Kontrolle ok. Ventile von PH BG, EL MET und Fluidmodul ausgebaut, gereinigt und membrane ersetzt ok. Ventilblock EL MET Modul ausgetauscht ok. Elektrode Connector Lac ersetzt ok. -> unstabil Einlassnadel und Einlassdichtung ersetzt ok. Tubing nipple satzt bei Einlass ersetzt ok. System i.o. Kal. + QC i.o. |  | 2854 | No | General control ok. Valves of PH BG, EL MET and fluid module removed, cleaned and membrane replaced ok. Valve block EL MET module replaced ok. Electrode Connector Lac replaces ok. -> unstable inlet needle and inlet seal replaced ok. Tubing nipple sets replaced at inlet ok. System i.o. Kal. + QC i.o.
ABL827 FLEX |  |  | CE1E6310A034C157 | Yes |  |  | Re classified RMED | Analyzer setup
ABL827 FLEX |  |  | CE1E6381A03EC15B | Yes |  |  | Re classified RMED | Analyzer setup
ABL820 FLEX | Unit was in HOLD. Performed service inlet calibration and failed on Probe position check. | Replaced inlet probe. | CE1D3A762A11927E | No | Unit was in HOLD. Performed service inlet calibration and failed on Probe position check. | Replaced inlet probe.
AQT90 FLEX |  |  | CE1D55562A2281BC | No
ABL90 FLEX | Display neu justiert | Display neu justiert | SWO-7228 | No | Display re-adjusted | Display re-adjusted | Adjustment | Touch screen | Alignment
ABL825 FLEX | Error 374 inhomogenous sample errors. | Replaced Liquid sensor and inlet tube. | CE23E47E7836DF7B | No | Error 374 inhomogenous sample errors. | Replaced Liquid sensor and inlet tube.
ABL825 FLEX | Error 258,316 and 247 | Replaced Spectrometer | CE23E34E782FDF78 | No | Error 258,316 and 247 | Replaced Spectrometer
ABL90 FLEX | Akku ausgetauscht. Funktion in Ordnung. | Kunde berichtet, dass der Akku im Betrieb angefangen hat zu rauchen stinken, der Kunde bittet um einen Einsatz des zuständigen Techniker, bitte um Planung eines Einsatzes. | SWO-7414 | No | Battery replaced. Function in order. | Customer reports that the battery has started smoking in operation, the customer asks for the use of the responsible technician, please to plan an operation. | Defect | HW | Battery module | Burned
ABL90 FLEX | Netzkabel am Computer richtig angesteckt QC Level B,C | Netzkabel am Computer richtig angesteckt QC Level B,C | SWO-7318 | No | Power cord properly plugged into the computer QC Level B,C | Power cord properly plugged into the computer QC Level B,C | Adjustment | Analyzer assembly | Cable seating
ABL835 FLEX | inlet coule intermittant flush waste line remplacé replace inlet gasket OK | Remplacé rotor suspect waste tubing et inlet nipple | CE1FBBD216420FE2 | No | inlet flows intertermittant flush waste line replaced replace inlet gasket OK | Remplacé rotor suspect waste tubing et inlet nipple
ABL825 FLEX | Luftstrecke gespült Ventile V7und V8 ausgebaut und Membrane getauscht Anschließend Halbjahreswartung separater SWO | Luftstrecke gespült Ventile V7und V8 ausgebaut und Membrane getauscht Anschließend Halbjahreswartung separater SWO | SWO-7329 | No | Air section flushed valves V7 and V8 removed and membrane replaced Subsequent half-year maintenance of separate SWO | Air section flushed valves V7 and V8 removed and membrane replaced Subsequent half-year maintenance of separate SWO
ABL825 FLEX | There is no software keys in option.Tried power supply did not help.Drove to meet Steve to get an AU PCB.Downloaded permanet keys and installed keys,Tried to load AU software, cannot. | Steve too come in AM to reload software. | CE1FD30D16186E35 | No | There is no software keys in option. Tried power supply did not help. Drove to meet Steve to get an AU PCB. Downloaded permanet keys and installed keys,Tried to load AU software, cannot. | Steve too come in AM to reload software.
ABL815 FLEX | Fehlersuche = Spüllsg. war leer, wurde v. Kunden gew., Kal. + QC i.O. | Fehlersuche = Spüllsg. war leer; wurde v. Kunden gew.; Kal. + QC i.O. | SWO-7379 | No | Troubleshooting = Rinse Sg. was empty; was awarded by customers; Kal. + QC in the first place | Troubleshooting = Rinse Sg. was empty; was awarded by customers; Kal. + QC in the first place
ABL800 FLEX | pH BG Modul - Ventile gereinigt, Messkammer gereinigt, Gerät kalibriert, QCs durchgef. | pH BG Modul - Ventile gereinigt, Messkammer gereinigt, Gerät kalibriert, QCs durchgef. | SWO-7396 | No | pH BG Module - Valves cleaned, measuring chamber cleaned, device calibrated, QCs | pH BG Module - Valves cleaned, measuring chamber cleaned, device calibrated, QCs
AQT90 FLEX | Gerät geöffnet, def. Kassette D- Dimmer 09084 (Nase abgebrochen) entfernt, QC durchgeführt | Gerät geöffnet, def. Kassette D- Dimmer 09084 (Nase abgebrochen) entfernt, QC durchgeführt | SWO-7472 | No | Device open, def. Cassette D- Dimmer 09084 (nose aborted) removed, QC performed | Device open, def. Cassette D- Dimmer 09084 (nose aborted) removed, QC performed
AQT90 FLEX | Incubator wheel mechanisch blockiert durch Cup Cup entfernt Positions Sensor von Incubator Wheel gereinigt Incubator Wheel kalibriert Troponin Kalibration durchgeführt Systemüberprüfung OK | Incubator wheel mechanisch blockiert durch Cup Cup entfernt Positions Sensor von Incubator Wheel gereinigt Incubator Wheel kalibriert Troponin Kalibration durchgeführt Systemüberprüfung OK | SWO-7412 | No | Incubator wheel mechanically blocked by Cup Cup removed position sensor cleaned by Incubator Wheel Calibrated Troponin Calibration System Review OK | Incubator wheel mechanically blocked by Cup Cup removed position sensor cleaned by Incubator Wheel Calibrated Troponin Calibration System Review OK
ABL825 FLEX | Take over from yesterday. Found AUEngine.exe in another folder. Able to get AU running. Now getting a error code status #2. Back up everything I could to USB drive. Directory structure can't be repaired. | Uninstall Kira and delete 700 folder. Re-install Kira. Stop Roma and drop old Roma.db, Roma.ini and Archive folder into the 700 folder . Correct system clock. Scan in gas tanks. | CE20E1238C07D761 | No | Take over from yesterday. Found AUEngine.exe in another folder. Able to get AU running. Now getting a error code status #2. Back up everything I could to USB drive. Directory structure can't be repaired. | Uninstall Kira and delete 700 folder. Re-install Kira. Stop Roma and drop old Roma.db, Roma.ini and Archive folder into the 700 folder . Correct system clock. Scan in gas tanks.
AQT90 FLEX | Auf Ersatzteil von Transofelx gewartet, vergeblich In GS gefahren um Ersatzgerät zu holen D-Dimer auf Ersatzgerät kalibriert LP installiert Leihgerät S N: 838R0075N004 aufgestellt und in Betrieb genommen Troponin T aus Kundenbestand kalibriert QC Messungen durchgeführt Kundengerät zur Reparatur mit genommen | Auf Ersatzteil von Transofelx gewartet, vergeblich In GS gefahren um Ersatzgerät zu holen D-Dimer auf Ersatzgerät kalibriert LP installiert Leihgerät S N: 838R0075N004 aufgestellt und in Betrieb genommen Troponin T aus Kundenbestand kalibriert QC Messungen durchgeführt Kundengerät zur Reparatur mit genommen | SWO-7851 | No | Waiting for spare part from Transofelx, in vain driven in GS to fetch replacement unit D-Dimer calibrated on replacement device LP installed loaner S N: 838R0075N004 set up and put into operation Troponin T from customer stock calibrated QC measurements performed Customer device taken with him for repair | Waiting for spare part from Transofelx, in vain driven in GS to fetch replacement unit D-Dimer calibrated on replacement device LP installed loaner S N: 838R0075N004 set up and put into operation Troponin T from customer stock calibrated QC measurements performed Customer device taken with him for repair
ABL90 | Confirmed problem. | Replaced the oxi valve tubing. | CE2402447822FABE | No | Confirmed problem. | Replaced the oxi valve tubing. | Defect | HW | Oxi Module | Pinch Valve tube
ABL90 |  |  | CE23B330780A9D31 | No |  |  | Cleaned | Fluidic pathway
TCM/Tosca | Comprobar fusibles y fuente de alimentación, limpiar contactos electricos de la placa base calibrar y comprobar. | Limpiar los contactos de la placa base | CE237B9078159070 | No | Check fuses and power supply, clean electrical contacts from the motherboard calibrate and check. | Cleaning the contacts on the motherboard | Defect | HW | Switch | cleaned
ABL825 FLEX | I inspected the air line for crystals and none was found. | I flushed out the air line with warm water and performed a liquid sensor calibration in the service mode. | CE2359387840FB41 | No | I inspected the air line for crystals and none was found. | I flushed out the air line with warm water and performed a liquid sensor calibration in the service mode.
ABL825 FLEX | The hard drive is bad. | Ordering the latest version of XP to load on the hard drive. Will return on 12-2-2014 to complete repair. | CE236DBD78087A48 | No | The hard drive is bad. | Ordering the latest version of XP to load on the hard drive. Will return on 12-2-2014 to complete repair.
ABL825 FLEX |  |  | CE2FFE2B14610820 | No
AQT90 FLEX | Neuen Lösungspack installiert, Funktionskontrolle | Neuen Lösungspack installiert, Funktionskontrolle | SWO-7559 | No | New solution pack installed, function control | New solution pack installed, function control
ABL90 |  |  | CE24830FEE20A0E8 | Yes |  |  | Re classified RMED | Logistical upSWO
ABL90 | Customer replaced 2 sensor cassettes | Replaced sensor PCB. Performed sensor test in service. Also replaced sensor cassette | CE24A96DEE0D3B2C | No | Customer replaced 2 sensor cassettes | Replaced sensor PCB. Performed sensor test in service. Also replaced sensor cassette | Defect | SC | Termistor
ABL90 FLEX | Neue SC eingesetzt Kal und QC waren OK. | Neue SC eingesetzt Kal und QC waren OK. | SWO-8594 | No | New SC used Kal and QC were OK. | New SC used Kal and QC were OK. | Defect | SC | Termistor
ABL90 FLEX | Tausch des LP ID PCB. Nach erneuter Installation des LP wird er vom Analysator erkannt. Initialisierung und Startphase durchgeführt. Analysator Funktion o.K. Auf Wunsch der Anwender, Frau Assauer eine Einlassdichtung und eine Einlasssonde am Analysator gelassen, da keine Vorort. SC erneuert. Erneut 50 min Initialisierungsphase. Anwender überprüfen Analysator und geben ihn frei wenn alle QC im Grenzbereich. | Tausch des LP ID PCB. Nach erneuter Installation des LP wird er vom Analysator erkannt. Initialisierung und Startphase durchgeführt. Analysator Funktion o.K. Auf Wunsch der Anwender, Frau Assauer eine Einlassdichtung und eine Einlasssonde am Analysator gelassen, da keine Vorort. SC erneuert. Erneut 50 min Initialisierungsphase. Anwender überprüfen Analysator und geben ihn frei wenn alle QC im Grenzbereich. | SWO-7758 | No | Replacing the LP ID PCB. After reinstalling the LP, it is recognized by the analyzer. initialization and start-up phase. Analyzer function o.K. At the request of the users, Ms. Assauer left an inlet seal and an inlet probe at the analyzer, since no suburb. SC renewed. Again 50 min initialization phase. Users check analyzer and release it when all QC in the limit range. | Replacing the LP ID PCB. After reinstalling the LP, it is recognized by the analyzer. initialization and start-up phase. Analyzer function o.K. At the request of the users, Ms. Assauer left an inlet seal and an inlet probe at the analyzer, since no suburb. SC renewed. Again 50 min initialization phase. Users check analyzer and release it when all QC in the limit range. | Defect | HW | SP interface PCB
ABL90 FLEX | ABL zerlegt und alle Schläuche ersetzt. Software Downgrade auf 3.1, SC Kassette ersetzt weil die Kal. in pCO2 und Cl. nicht Bestanden wird. Startphase QC und Kalibration. | ABL zerlegt und alle Schläuche ersetzt. Software Downgrade auf 3.1, SC Kassette ersetzt weil die Kal. in pCO2 und Cl. nicht Bestanden wird. Startphase QC und Kalibration. | SWO-7708 | No | ABL disassembles and replaces all hoses. Software downgrade to 3.1, SC cassette replaced because the Kal. in pCO2 and Cl. is not passed. Start phase QC and calibration. | ABL disassembles and replaces all hoses. Software downgrade to 3.1, SC cassette replaced because the Kal. in pCO2 and Cl. is not passed. Start phase QC and calibration. | Defect | SC | Wetability
ABL90 FLEX | Analysator zerlegt und auf Undichtigkeit geprüft Einlassdichtung und Anschlussdichtung ausgetauscht System gespült Kein offensichtlicher Fehler feststellbar Analysator kalibriert und überprüft Fcuv 98,6% Interne QC Ok Mit Anwender Fehleranalyse am Gerät besprochen, Möglichkeit die Kalibration zu wiederholen und manuelle QC Messung Kapillare Messung gezeigt, Anwender hat zu viel Druck auf den Gerinselfänger ausgeübt und da knickt die Spitze unter Umständen ab | Analysator zerlegt und auf Undichtigkeit geprüft Einlassdichtung und Anschlussdichtung ausgetauscht System gespült Kein offensichtlicher Fehler feststellbar Analysator kalibriert und überprüft Fcuv 98,6% Interne QC Ok Mit Anwender Fehleranalyse am Gerät besprochen, Möglichkeit die Kalibration zu wiederholen und manuelle QC Messung Kapillare Messung gezeigt, Anwender hat zu viel Druck auf den Gerinselfänger ausgeübt und da knickt die Spitze unter Umständen ab | SWO-7762 | No | Analyzer disassembled and checked for leakage inlet seal and connection seal replaced system flushed No obvious error detectable analyzer calibrated and checked Fcuv 98.6% Internal QC Ok With user error analysis discussed on the device, Possibility to repeat the calibration and manual QC measurement capillary measurement shown, user has exerted too much pressure on the scrawler and there the tip bends under certain circumstances | Analyzer disassembled and checked for leakage inlet seal and connection seal replaced system flushed No obvious error detectable analyzer calibrated and checked Fcuv 98.6% Internal QC Ok With user error analysis discussed on the device, Possibility to repeat the calibration and manual QC measurement capillary measurement shown, user has exerted too much pressure on the scrawler and there the tip bends under certain circumstances | User Error | Clot catcher
ABL835 FLEX | TRASLADO DEL EQUIPO AL NUEVO LABORATORIO, SE CALIBRA EL EQUIPO EN VARIAS OCASIONES, VERIFICAMOS EL CORRECTO FUNCIONAMIENTO CON CUATRO NIVELES DE CONTROL DE CALIDAD. COMPROBACIÓN DE CONFIGURACION DEL PROTOCOLO DE COMUNICACIONES DE RED. |  | CE24C7D9EE3B823C | Yes | TRANSFER FROM THE EQUIPMENT TO THE NEW LABORATORY, THE EQUIPMENT IS CALIBRATED IN SEVERAL OCCASIONS, WE VERIFY THE CORRECT OPERATION WITH FOUR LEVELS OF QUALITY CONTROL. NETWORK COMMUNICATIONS PROTOCOL CONFIGURATION CHECK. |  | Re classified RMED | Analyzer transferred to other department
ABL830 FLEX | TRASLADO DEL EQUIPO AL NUEVO LABORATORIO, SE CALIBRA EL EQUIPO EN VARIAS OCASIONES, VERIFICAMOS EL CORRECTO FUNCIONAMIENTO CON CUATRO NIVELES DE CONTROL DE CALIDAD. COMPROBACIÓN DE CONFIGURACION DEL PROTOCOLO DE COMUNICACIONES DE RED. |  | CE24AD83EE3C536A | Yes | TRANSFER FROM THE EQUIPMENT TO THE NEW LABORATORY, THE EQUIPMENT IS CALIBRATED IN SEVERAL OCCASIONS, WE VERIFY THE CORRECT OPERATION WITH FOUR LEVELS OF QUALITY CONTROL. NETWORK COMMUNICATIONS PROTOCOL CONFIGURATION CHECK. |  | Re classified RMED | Analyzer transferred to other department
ABL835 FLEX | Un ou plusieur electrolyte Ca++ et Na++ d une ou deux fois par jour | Remplacè électrode de référence | CE24A830EE2414B8 | No | One or more electrolyte electrolyte ca and Na-a-by-one once or twice a day | Reference electrode replace
ABL815 FLEX | AC Modulplatine ersetzt. | AC Modulplatine ersetzt. | SWO-7812 | No | AC module board. | AC module board.
ABL825 FLEX | Folder needs to b e created in 700 folder for backup | Create backup folder in 700 folder in windows explorer via service program. Point backup to new folder. | CE24BDC7EE001DEF | No | Folder needs to b e created in 700 folder for backup | Create backup folder in 700 folder in windows explorer via service program. Point backup to new folder.
ABL815 FLEX | Endgültigen Key eingespielt. Überprüft, das verbleibende Demo-Keys deaktiviert sind. | der Anwender wünscht einen Rückruf wie in der Angelegenheit weiter zu verfahren ist; | SWO-7775 | Yes | Final key. Checked that remaining demo keys are disabled. | the user wishes a recall as to how to proceed in the matter; | Re classified RMED | Planned activity | Key installation
ABL825 FLEX | pCO2 looks good, low on some QC results based on peer but are well within specs. Na within specs but very borderline low. | Completed maintenance and changed Cal 1 solution. Customer notes the Na QC drifts uop after cal 1 bottle change. This concern was duplicated and will be passed to tech support. | CE253BCBEE4CBF7A | No | pCO2 looks good, low on some QC results based on peer but are well within specs. Na within specs but very borderline low. | Completed maintenance and changed Cal 1 solution. Customer notes the Na QC drifts uop after cal 1 bottle change. This concern was duplicated and will be passed to tech support.
ABL815 FLEX | Fehlersuche = Kal.2-Flasche nicht korrekt angebracht, alle Ventile ausgebaut u. gereinigt, div. Test`s im Service durchgef., div. Kal. i.O., QC i.O. | Fehlersuche = Kal.2-Flasche nicht korrekt angebracht; alle Ventile ausgebaut u. gereinigt; div. Test`s im Service durchgef.; div. Kal. i.O.; QC i.O. | SWO-7723 | No | Troubleshooting = Kal.2 bottle not mounted correctly; all valves removed and cleaned; div. Test it in the service. div. Kal. i.O.; QC in the first | Troubleshooting = Kal.2 bottle not mounted correctly; all valves removed and cleaned; div. Test it in the service. div. Kal. i.O.; QC in the first
ABL825 FLEX | Software needed loading on hard drive. | Loaded software on hard drive. Scanned the QC back into the hard drive and loaded the setup on the hard drive. | CE24AA9DEE0A6F56 | No | Software needed loading on hard drive. | Loaded software on hard drive. Scanned the QC back into the hard drive and loaded the setup on the hard drive.
ABL800 FLEX | pCO2 Elektrode ern., Gerät kalibriert, QC´s durchgef. | pCO2 Elektrode ern., Gerät kalibriert, QC´s durchgef. | SWO-8483 | No | pCO2 electrode ern., device calibrated, QC's through. | pCO2 electrode ern., device calibrated, QC's through.
ABL90 | Customer was failing sensitivity on all electrolytes. Customer has had multiple cassette failures in the past week. | Checked data logs and found calibration errors were cassette related. Customer was using lot 580. Customer installed new sensor cassette and verified instrument operation. | CE2B69F73C38514E | No | Customer was failing sensitivity on all electrolytes. Customer has had multiple cassette failures in the past week. | Checked data logs and found calibration errors were cassette related. Customer was using lot 580. Customer installed new sensor cassette and verified instrument operation. | Defect | SC | Reference?
ABL90 FLEX | Programm durchgeklickt, LP wird sofort und problemlos erkannt und akzeptiert! Beim vorletzten LP Wechsel wurde der LP nicht erkannt => LP-ID-PCB mit verbogenem Kontakt => ersetzt und LP initialisieren lassen. Kalibrationen und QC Messungen => OK | Programm durchgeklickt, LP wird sofort und problemlos erkannt und akzeptiert! Beim vorletzten LP Wechsel wurde der LP nicht erkannt => LP-ID-PCB mit verbogenem Kontakt => ersetzt und LP initialisieren lassen. Kalibrationen und QC Messungen => OK | SWO-7900 | No | Program clicked through, LP is recognized and accepted immediately and easily! On the penultimate LP change, the LP was not recognized => LP-ID-PCB with bent contact => replaced and LP initialized. Calibrations and QC measurements => OK | Program clicked through, LP is recognized and accepted immediately and easily! On the penultimate LP change, the LP was not recognized => LP-ID-PCB with bent contact => replaced and LP initialized. Calibrations and QC measurements => OK | Defect | HW | SP interface PCB
ABL90 |  |  | CE261DB564019F32 | No |  |  | Defect | HW | Oxi Module | Pinch Valve tube
ABL90 FLEX | Beide rechten Kontakte des LP leicht eingedrückt, Contact PCB erneuert. LP neu installiert und Analysator neu gestartet. Analysator Messbereit. Keine Meldung vorhanden. QC ohne Beanstandung | Beide rechten Kontakte des LP leicht eingedrückt, Contact PCB erneuert. LP neu installiert und Analysator neu gestartet. Analysator Messbereit. Keine Meldung vorhanden. QC ohne Beanstandung | SWO-7934 | No | Both right contacts of the LP slightly pressed, Contact PCB renewed. REinstalled LP and restarted analyzer. Analyzer Ready to measure. There is no message. QC without complaint | Both right contacts of the LP slightly pressed, Contact PCB renewed. REinstalled LP and restarted analyzer. Analyzer Ready to measure. There is no message. QC without complaint | Defect | HW | SP interface PCB
ABL820 FLEX | pCO2 is at lower end of peer group on all levels. Less than 2SD but low. (64.4, low peer is 63.8, 38.5, low peer is 38.1, 12.0, low peer is 11.6, 91.8, low peer is 91.1) Values are stable. | Replaced pH o-ring and installed loaner pCO2 electrode, lot RI04. Customer will monitor and advise in 7 days. | CE26163E6418147D | No | pCO2 is at lower end of peer group on all levels. Less than 2SD but low. (64.4, low peer is 63.8; 38.5, low peer is 38.1; 12.0, low peer is 11.6; 91.8, low peer is 91.1) Values are stable. | Replaced pH o-ring and installed loaner pCO2 electrode, lot RI04. Customer will monitor and advise in 7 days.
ABL825 FLEX | Last visit, when I was here on Saturday, I was able to get option keys to reload. Analyzer PCB was suspect. Ordered new Analyzer Control PCB. | Install new AC PCB. Reload AU software and original keys. Configure parameters. Update type in service. Perform PO2 Zero Current calibration in service. | CE25E639643382FF | No | Last visit, when I was here on Saturday, I was able to get option keys to reload. Analyzer PCB was suspect. Ordered new Analyzer Control PCB. | Install new AC PCB. Reload AU software and original keys. Configure parameters. Update type in service. Perform PO2 Zero Current calibration in service.
ABL825 FLEX | All Oxi values are questioned and have no values. | Recalibrated the pumps and liquid sensors. Lubricated the tube valve. | CE265A0064112896 | No | All Oxi values are questioned and have no values. | Recalibrated the pumps and liquid sensors. Lubricated the tube valve.
ABL820 FLEX | Unit has been performing OK since last visit. | Reinstalled pHBG module. | CE2627DC642C9136 | Yes | Unit has been performing OK since last visit. | Reinstalled pHBG module. | Re classified RMED | Courtesy
ABL827 FLEX | Electrodes needing cleaning and remembraning. | Clean and remembrane Lactate and CREA electrodes. | CE263BC56419CAD5 | No | Electrodes needing cleaning and remembraning. | Clean and remembrane Lactate and CREA electrodes.
TCM-COMBIM |  |  | CE25E6676444F20F | No |  |  | Defect | HW | Switch
ABL90 |  |  | CE27586ADA361C5E | No |  |  | Defect | HW | Actuator unit | OOB
ABL90 | SW asking for a replacement solution pack. Authorized one to ship overnight. Sent him to CSS Melissa. | Replaced sensor cassette with bypass tubing, no fluids are flowing from pack in Visual flow check in service.Replaced SP, still no fluid. Took several flow check programs for fluids to start to flow. Multiple cals , ran waste blood. reran VK-R5 kit. All OK now | CE2786F4DA2AC721 | No | SW asking for a replacement solution pack. Authorized one to ship overnight. Sent him to CSS Melissa. | Replaced sensor cassette with bypass tubing, no fluids are flowing from pack in Visual flow check in service. Replaced SP, still no fluid. Took several flow check programs for fluids to start to flow. Multiple cals , ran waste blood. reran VK-R5 kit. All OK now | Defect | SP | FT | Wetability ?
ABL815 FLEX | FlexQ Modul (R0438N002) demontiert und Ersatz Modul (R0955N005) zur Reparatur Überbrückung eingebaut. Gerät kalibriert und QC durchgeführt. | FlexQ Modul (R0438N002) demontiert und Ersatz Modul (R0955N005) zur Reparatur Überbrückung eingebaut. Gerät kalibriert und QC durchgeführt. | SWO-8542 | No | FlexQ module (R0438N002) disassembled and replacement module (R0955N005) installed for repair bridging. calibrated and QC performed. | FlexQ module (R0438N002) disassembled and replacement module (R0955N005) installed for repair bridging. calibrated and QC performed.
ABL825 FLEX |  |  | CE286FB95015F4C5 | No
ABL835 FLEX | Bulle d air avec les controles intermittant | Inspecté les chambres flush waste line remplacé inlet tube | CE271422DA38B00D | No | Air bubble with interterm controls | Inspected flush waste line rooms replaced inlet tube
ABL825 FLEX | 1 point pH BG Cal is upgraded to 1 point cal because of pH drift. | Cleaned pH chamber and tip of pH Electrode. Also inspect data logs for performance. | CE271991DA49152A | No | 1 point pH BG Cal is upgraded to 1 point cal because of pH drift. | Cleaned pH chamber and tip of pH Electrode. Also inspect data logs for performance.
ABL835 FLEX | Referenz Leihelektrode(4101) aus dem ABL ausgebaut und Herrn Höllriegl übergeben. Neue Referenzelektrode (2707) membraniert und eingebaut Mehrmals kalibriert QCs für alle Level durchgeführt. ABL ist wieder Messbereit. | Referenz Leihelektrode(4101) aus dem ABL ausgebaut und Herrn Höllriegl übergeben. Neue Referenzelektrode (2707) membraniert und eingebaut Mehrmals kalibriert QCs für alle Level durchgeführt. ABL ist wieder Messbereit. | SWO-8095 | No | Reference rental electrode(4101) removed from the ABL and handed over to Mr Höllriegl. New reference electrode (2707) membraneed and built-in multiple calibrated QCs performed for all levels. ABL is ready to measure again. | Reference rental electrode(4101) removed from the ABL and handed over to Mr Höllriegl. New reference electrode (2707) membraneed and built-in multiple calibrated QCs performed for all levels. ABL is ready to measure again.
ABL825 FLEX | Analyzer running well other than cross over study | Replaced CL and NA electrodes Service checks done | CE3A2F07C422D483 | No | Analyzer running well other than cross over study | Replaced CL and NA electrodes Service checks done
ABL825 FLEX | Austausch Drucker Testausdruck i.O. | Austausch Drucker Testausdruck i.O. | SWO-8102 | No | Replacement printer test expression i.O. | Replacement printer test expression i.O.
ABL825 FLEX | Austausch pH BG Modul Umsetzen der Kundenelektroden Schlauchnachfüllung und Durchführung Nullpunktkalibrierung Mehrere Kal2 bis alles i.O. | Austausch pH BG Modul Umsetzen der Kundenelektroden Schlauchnachfüllung und Durchführung Nullpunktkalibrierung Mehrere Kal2 bis alles i.O. | SWO-8103 | No | Replacement pH BG module Conversion of customer electrodes hose refill and execution of zero point calibration Several Kal2 to everything in the order. | Replacement pH BG module Conversion of customer electrodes hose refill and execution of zero point calibration Several Kal2 to everything in the order.
ABL837 FLEX | Upon arrival, I found the analyzer leaking around the right. Also, the analyzer is presenting pump errors. | Inspected the analyzer, found the analyzer leaking out of the EI Met module measuring chambers. Found the K membrane missing the bottom ring. Replaced a new K membrane. ran pump calibrations and liquid sensor adjustments. Replaced the Cal 2 and the cleaning solutions, found expired. Replaced the reference and the creatinine membranes, found expired. | CE271E47DA60E5AA | No | Upon arrival, I found the analyzer leaking around the right. Also, the analyzer is presenting pump errors. | Inspected the analyzer, found the analyzer leaking out of the EI Met module measuring chambers. Found the K membrane missing the bottom ring. Replaced a new K membrane. ran pump calibrations and liquid sensor adjustments. Replaced the Cal 2 and the cleaning solutions, found expired. Replaced the reference and the creatinine membranes, found expired.
AQT90 FLEX | Nach Ersatzteilbestellung Lichtschranke PI6 636-550 an Process Unit erneuert, Analysator gereinigt (relativ stark verschmutzt), Neustart des Analysators ok, im Service Safe Positions ok, Systemüberprüfung ok, Reinigungsprogramm durchgeführt und QC ok System-und Funktionskontrolle i.O. | Nach Ersatzteilbestellung Lichtschranke PI6 636-550 an Process Unit erneuert, Analysator gereinigt (relativ stark verschmutzt), Neustart des Analysators ok, im Service Safe Positions ok, Systemüberprüfung ok, Reinigungsprogramm durchgeführt und QC ok System-und Funktionskontrolle i.O. | SWO-8402 | Yes | After spare parts order light barrier PI6 636-550 to process unit renewed, analyzer cleaned (relatively heavily polluted), restart of the analyzer ok, in service safe positions ok, system check ok, cleaning program performed and QC ok system and Functional control in the | After spare parts order light barrier PI6 636-550 to process unit renewed, analyzer cleaned (relatively heavily polluted), restart of the analyzer ok, in service safe positions ok, system check ok, cleaning program performed and QC ok system and Functional control in the | Re classified RMED | Double entry
AQT90 FLEX | Fehlersuche, Safe Positions über Service->Process Unit und Process Wheel Error, Bewegung beider Einheiten überprüft ok, allerdings lässt sich der Process Unit Error nicht beheben, beide Lichtschranken bei Process Unit überprüft->PI 6 636-550 defekt!, PI 6 für nächsten Tag bestellt Anwender informiert, dass der Analysator am Folgetag repariert wird | Nach Ersatzteilbestellung Lichtschranke PI6 636-550 an Process Unit erneuert, Analysator gereinigt (relativ stark verschmutzt), Neustart des Analysators ok, im Service Safe Positions ok, Systemüberprüfung ok, Reinigungsprogramm durchgeführt und QC ok System-und Funktionskontrolle i.O. | SWO-8399 | No | Troubleshooting, Safe Positions via Service->Process Unit and Process Wheel Error, movement of both units checked ok, but the Process Unit Error cannot be fixed, both light barriers checked at Process Unit->PI 6 6 636-550 defective!, PI 6 for next day orders user informs that the analyzer will be repaired the following day | After spare parts order light barrier PI6 636-550 to process unit renewed, analyzer cleaned (relatively heavily polluted), restart of the analyzer ok, in service safe positions ok, system check ok, cleaning program performed and QC ok system and Functional control in the
ABL90 FLEX | Zeitzone falsch->geändert->siehe SWO8252 | Zeitzone falsch->geändert->siehe SWO8252 | SWO-8257 | No | Time zone incorrect->changed->see SWO8252 | Time zone incorrect->changed->see SWO8252 | Setup | Windows | Time Zone | We send out all this information’s via e-mail and additional we upload this to our SharePoint and/or CRM. But again, we work with 54 humans.
ABL90 FLEX | Steckerverbindung zum Bildschirm lose Stecker neu befestigt Gerät kalibriert und überprüft | Steckerverbindung zum Bildschirm lose Stecker neu befestigt Gerät kalibriert und überprüft | SWO-8288 | No | Plug connection to the screen loose plug reattached device calibrated and checked | Plug connection to the screen loose plug reattached device calibrated and checked | Adjustment | HW | Cable seating | We can call for a signature, but we get this for arrival and reading, not for remember.
ABL90 FLEX | Anwender konnte keine Angaben zum Patienten machen. Das Problem soll angeblich bei Dialysepatienten und bei Ambulanten Patienten aufgetreten sein. Im Archiv waren Patienten die im September erstellt worden sind. Anschlusszeit auf 400 Tage gesetzt. Problem wird vom Anwender beobachtet und die Patienten bei denen das Problem auftritt, sollen dokumentiert werden. | Kundin berichtet, dass sie Pat. Namen immer händisch eingeben und bei der nächsten Messung, nach eintippen der Pat. Nummer der Name im Bericht mit allen Daten erscheint, letztens muss es vorgekommen sein, dass nach der Kal. keinerlei Namen mehr verfügbar waren - sie musste alle wieder manuell eingeben, jetzt passiert es, dass eingegebene Namen bei Folgemessungen nicht autom. nach eintippen der Pat. Nummer eingefügt werden, hier müssen wieder manuell alle Daten eingetragen werden, dies passiert in den letzten Tagen immer öfter, Kundin bittet um einen Einsatz des zuständigen Technikers, bitte um Kontaktaufnahme. | SWO-8348 | No | user was unable to provide information about the patient. The problem is said to have occurred in dialysis patients and outpatients. The archive included patients who were created in September. Connection time set to 400 days. Problem is observed by the user and the patients with whom the problem occurs should be documented. | Customer reports that she has Pat. Always enter the name manually and at the next measurement, after typing the pat. Number the name appears in the report with all the data, last it must have happened that after the Cal. no names were available anymore - they had to enter them all manually again, now it happens that entered names are not autom during follow-up measurements. after typing in the Pat. Number must be inserted here again manually all data, this happens more and more often in the last days, customer asks for an assignment of the responsible technician, please contact. | Unknown | Revisit required | SETUP
ABL90 FLEX | Sensor Interface Board und Sensorkassette ersetzt. AU Software installiert. Startphase QC und Kalibration. | Sensor Interface Board und Sensorkassette ersetzt. AU Software installiert. Startphase QC und Kalibration. | SWO-8213 | No | Sensor Interface Board and Sensor Cassette. AU software installed. Start phase QC and calibration. | Sensor Interface Board and Sensor Cassette. AU software installed. Start phase QC and calibration. | Defect | SC | Eckhi mail | How do we ensure that all the bulletins, service tips, technical information’s, instructions, internal information’s are read and followed? That’s really a very good question. We work with humans and I think that it is normal to forget something in some cases. I wish I or we have a solution for this, but I do not have it at the moment.
ABL90 | will reload software and replace CF card | Loaded SW windows and application 3.1 MR1 and replaced CF card, loaded setup and ran commit program. | CE28F45E505584E2 | No | will reload software and replace CF card | Loaded SW windows and application 3.1 MR1 and replaced CF card, loaded setup and ran commit program. | Defect | HW | CF card | Corrupt file
ABL800 BASIC | Einlass Dichtung erneuert. LS Justage, Pumpen Kal | Einlass Dichtung erneuert. LS Justage, Pumpen Kal | SWO-8278 | No | Inlet seal renewed. LS Adjustment, Pumps Kal | Inlet seal renewed. LS Adjustment, Pumps Kal
ABL827 FLEX |  |  | CE28EA09503E94B5 | Yes |  |  | Re classified RMED | Planned activity | Install permanent keys
ABL800 BASIC | ABL überprüft. OXI-Complete-Check durchgeführt. OXI-Modul erneuert. LS-Justage durchegführt, Service-tHb-Kalibration durchgeführt, tHb-Kalibration durchgeführt. THb-Leistungsmessung durchgeführt. Gerät kalibriert. QC durchgeführt. | ABL überprüft. OXI-Complete-Check durchgeführt. OXI-Modul erneuert. LS-Justage durchegführt, Service-tHb-Kalibration durchgeführt, tHb-Kalibration durchgeführt. THb-Leistungsmessung durchgeführt. Gerät kalibriert. QC durchgeführt. | SWO-8301 | No | ABL. OXI Complete Check. OXI module renewed. LS-Adjustment, service tHb calibration performed, tHb calibration performed. THb power measurement. device calibrated. QC. | ABL. OXI Complete Check. OXI module renewed. LS-Adjustment, service tHb calibration performed, tHb calibration performed. THb power measurement. device calibrated. QC.
ABL805 FLEX | I came in because Wanda called in with the electrolytes & Glu are out on cals. She said the status was out. I suggested replacing the REF electrode. When I arrived Wanda had the analyzer up and running. While I was here I went ahead and replaced the pH seal ring with new style. See SWO B9C52R7MF for FAN. Glu zero went out on cals so I replaced the Glu membrane with customer stock. The Glu is showing unstable response fault on QC. Replaced Glu electrode with customer stock. Glu is still giving unstable response fault. | Membraned the Glu electrode and cleaned chamber. Replaced Glu electrode with customer stock. Performed zero calibration. Replaced El Met chamber and Glu connector and solution pump and El Met pump tubes with customer stock. | CE2855C6500A8E07 | No | I came in because Wanda called in with the electrolytes & Glu are out on cals. She said the status was out. I suggested replacing the REF electrode. When I arrived Wanda had the analyzer up and running. While I was here I went ahead and replaced the pH seal ring with new style. See SWO B9C52R7MF for FAN. Glu zero went out on cals so I replaced the Glu membrane with customer stock. The Glu is showing unstable response fault on QC. Replaced Glu electrode with customer stock. Glu is still giving unstable response fault. | Membraned the Glu electrode and cleaned chamber. Replaced Glu electrode with customer stock. Performed zero calibration. Replaced El Met chamber and Glu connector and solution pump and El Met pump tubes with customer stock.
ABL825 FLEX | Logs show PCO2 QC being unstable shifting from 58 to 65 in one day, customer has remembraned PCO2, blue oring has been installed | FSE going to be onsite on monday to check out the analyzer | CE2898EA505A0011 | No | Logs show PCO2 QC being unstable shifting from 58 to 65 in one day, customer has remembraned PCO2, blue oring has been installed | FSE going to be onsite on monday to check out the analyzer
ABL835 FLEX |  |  | CE284EB05014DCE6 | No
ABL825 FLEX | Activity log indicates several occurrences of error 793 - pH BG pump asymmetry errors. Upon visual inspection additional bubbles in fluidic path observed. | Replaced pH BG module pump tubing, inspected associated fluidic and air values. Replaced black tubing from V1 (pH BG module) to fluidic module as a precautionary measure. | CE28717350118115 | No | Activity log indicates several occurrences of error 793 - pH BG pump asymmetry errors. Upon visual inspection additional bubbles in fluidic path observed. | Replaced pH BG module pump tubing, inspected associated fluidic and air values. Replaced black tubing from V1 (pH BG module) to fluidic module as a precautionary measure.
AQT90 FLEX |  |  | CE2872265038B985 | No
ABL90 FLEX | Fehlersuche... Ventilschlauch ausgetauscht. Neustart. |  | SWO-8518 | No | Troubleshooting... valve hose. Restart. |  | Defect | HW | Oxi Module | Pinch Valve tube
ABL90 |  |  | CE3FBEF11208993E | No |  |  | Defect | HW | Display
ABL80 FLEX Basic version |  |  | CE2C3834B2126580 | No
ABL90 FLEX | Serviceprogramm Check, Lösungspack demontiert und überprüft, Kal, QC | Serviceprogramm Check, Lösungspack demontiert und überprüft, Kal, QC | SWO-8467 | No | Service program Check, solution pack dismantled and checked, Kal, QC | Service program Check, solution pack dismantled and checked, Kal, QC | Defect | SC | Wetability
ABL90 | Temp of substrate is low( 36.45) Tried new sensor assette, did not help. Cleaned sensor contacts, did not help. | Replaced with third sensor cassette, temp is now good. Ran both internal and external QC all OK. | CE2C74E7B25A34AA | No | Temp of substrate is low( 36.45) Tried new sensor assette, did not help. Cleaned sensor contacts, did not help. | Replaced with third sensor cassette, temp is now good. Ran both internal and external QC all OK. | Defect | SC | Termistor | Do you have an idea?
ABL90 FLEX | JW durchgeführt. Gelber Ablasschlauch undicht. |  | SWO-8520 | No | JW. Yellow reading hose leaks. |  | Defect | SC | Lac
ABL90 FLEX | Fehlersuche: Diverse Untersuchungen, offenbar keine Spannungsversorgung mehr vorhanden. Power PCB überprüft, defekt -ersetzt. Offene Fragen Herrn Dr. Villa beantwortet und Bestand an Zubehör überprüft. | Fehlersuche: Diverse Untersuchungen, offenbar keine Spannungsversorgung mehr vorhanden. Power PCB überprüft, defekt -ersetzt. Offene Fragen Herrn Dr. Villa beantwortet und Bestand an Zubehör überprüft. | SWO-8581 | No | Troubleshooting: Various investigations, apparently no power supply is available anymore. Power PCB checked, defective -replaced. Open questions Answered By Dr. Villa and inventory of accessories checked. | Troubleshooting: Various investigations, apparently no power supply is available anymore. Power PCB checked, defective -replaced. Open questions Answered By Dr. Villa and inventory of accessories checked. | Defect | HW | Power supply
ABL90 |  |  | CE2C69E7B24D27CC | No |  |  | LEH undersøger
ABL90 | Noted error 1123: The sensor chip data could not be read or written too. Observed area inside sensor cassette with unknown material across several electrode contact points. | Performed a visual inspection, cleaned probe and inlet gasket, replaced sensor cassette due to error code and observation. Replaced solution Pack as a precautionary measure. First Sensor cassette failed due to pCO2 internal QC level C high out of range. Replace with second sensor cassette. Both lot R0595. All parameters now in range | CE2C66F7B263B679 | No | Noted error 1123: The sensor chip data could not be read or written too. Observed area inside sensor cassette with unknown material across several electrode contact points. | Performed a visual inspection, cleaned probe and inlet gasket, replaced sensor cassette due to error code and observation. Replaced solution Pack as a precautionary measure. First Sensor cassette failed due to pCO2 internal QC level C high out of range. Replace with second sensor cassette. Both lot R0595. All parameters now in range | Defect | SC | Leaking Ref solution?
ABL825 FLEX | Computereinheit demontiert, Netzteil und CPU geprüft, Leih CPU installiert, Funktionskontrolle ok | die Netzspannungs-Versorgung ist OK; der Anwender sieht keine Möglichkeit das Problem selber zu beheben; der Anwender wünscht einen Vorort-Service zur Beseitigung der Geräte-Störung; | SWO-8589 | No | Computer unit dismantled, power supply and CPU checked, loan CPU installed, function control ok | the mains voltage supply is OK; the user sees no way to fix the problem himself; the user wants a suburban service to eliminate the device malfunction;
ABL825 FLEX | Complete oxi test gave no errors | Replaced PCO2 electrode to bring QC level up tHb cal done, no errors | CE2EB52A9E49155E | No | Complete oxygen test gave no errors | Replaced PCO2 electrode to bring QC level up tHb cal done, no errors
ABL815 FLEX | Fehlersuche = Nadel verbogen, Nadel gerichtet, QC aufgefüllt, Kal. + QC i.O. | Fehlersuche = Nadel verbogen; Nadel gerichtet; QC aufgefüllt; Kal. + QC i.O. | SWO-8481 | No | Troubleshooting = needle bent; Needle directed; QC replenished; Kal. + QC in the first place | Troubleshooting = needle bent; Needle directed; QC replenished; Kal. + QC in the first place
ABL825 FLEX | Checked network settings, all checked ok. Checked network cable, found corrosion on pins.Checked data logs for errors. Liquid sensor and pump calibrations failing. | Customer supplied replacement network cable. Network connect now ok. Cleaned leakage around inlet probe and reseat inlet tubing. Replaced inlet liquid sensor because of incorrect readings in diagnostics. | CE2CBBEBB24395FC | No | Checked network settings, all checked ok. Checked network cable, found corrosion on pins. Checked data logs for errors. Liquid sensor and pump calibrations failing. | Customer supplied replacement network cable. Network connect now ok. Cleaned leakage around inlet probe and reseat inlet tubing. Replaced inlet liquid sensor because of incorrect readings in diagnostics.
ABL800 BASIC | Auto Check Modul ausgebaut Ampullenbrechermotor getauscht Modul eingebaut - Gerät überprüft | Auto Check Modul ausgebaut Ampullenbrechermotor getauscht Modul eingebaut - Gerät überprüft | SWO-8420 | No | Auto Check module removed ampoule crusher engine replaced module built-in - device checked | Auto Check module removed ampoule crusher engine replaced module built-in - device checked
ABL825 FLEX | Ventile überprüft und gereinigt Kal2- und QC Level 2 | Ventile überprüft und gereinigt Kal2- und QC Level 2 | SWO-8406 | No | Valves checked and cleaned Kal2 and QC Level 2 | Valves checked and cleaned Kal2 and QC Level 2
ABL820 FLEX | Investigated logs and found issue retured a few weeks after my initial correction | Replaced Ampoule acuator Replaced 14 pole cable Stepped through service programs and monitored AQC module. Found replaced parts to not resolve the issue. Ampoule break keeps trying to extend further and hesitates to return some times. Will return with an AQC module. None in stock, all on back order. | CE2C306BB2102770 | No | Investigated logs and found issue retured a few weeks after my initial correction | Replaced Ampoule acuator Replaced 14 pole cable Stepped through service programs and monitored AQC module. Found replaced parts to not resolve the issue. Ampoule break keeps trying to extend further and hesitates to return some times. Will return with an AQC module. None in stock, all on back order.
AQT90 FLEX | Neustart AQT brachte keine Verbesserung, Vergleich der Patientendatenabfrage mit den ABL8xx im Haus-> es wird die Probennehmer ID abgefragt, am AQT nicht verfügbar, Server angepingt-> Rückmeldung in 2ms, Service Dump u. Data Logs erstellt, Gateway umgestellt auf 255.255.0.0 Nach Rücksprache mit POC IT ist keine Patientendatenabfrage am Radiance konfiguriert-> Konfig entsprechend geändert | Neustart AQT brachte keine Verbesserung, Vergleich der Patientendatenabfrage mit den ABL8xx im Haus-> es wird die Probennehmer ID abgefragt, am AQT nicht verfügbar, Server angepingt-> Rückmeldung in 2ms, Service Dump u. Data Logs erstellt, Gateway umgestellt auf 255.255.0.0 Nach Rücksprache mit POC IT ist keine Patientendatenabfrage am Radiance konfiguriert-> Konfig entsprechend geändert | SWO-8380 | No | Restart AQT brought no improvement, comparison of the patient data query with the ABL8xx in the house-> it is queried the sampler ID, not available at the AQT, server pinged-> feedback created in 2ms, service dump and data logs, gateway switched to 255.255.0.0 After consultation with POC IT, no patient data query is configured at the radiance-> Config changed accordingly | Restart AQT brought no improvement, comparison of the patient data query with the ABL8xx in the house-> it is queried the sampler ID, not available at the AQT, server pinged-> feedback created in 2ms, service dump and data logs, gateway switched to 255.255.0.0 After consultation with POC IT, no patient data query is configured at the radiance-> Config changed accordingly
ABL80 FLEX CO-OX | ersetzt | ersetzt | SWO-9091 | No | Replaced | Replaced
ABL80 CO-OX | Set up unit for evaluation with shop consumables. Unit came up with no errors. Dismantled and discovered a crimped waste line. | Ran Service Cleaning Pack; Installed parts from the Annual Maintenance kit (Waste tubing, fluidic diode and pump). | CE2D9A41285B7120 | No | Set up unit for evaluation with shop consumables. Unit came up with no errors. Dismantled and discovered a crimped waste line. | Ran Service Cleaning Pack; Installed parts from the Annual Maintenance kit (Waste tubing, fluidic diode and pump).
ABL90 |  |  | CE2D6B272851883B | No |  |  | Defect | HW | Hemolyzer | Temp
ABL90 FLEX | Serviceprogramme durchgeführt, dabei Probenweg gespült mit Adapter, Kal, QC | Serviceprogramme durchgeführt, dabei Probenweg gespült mit Adapter, Kal, QC | SWO-8889 | No | Service programs carried out, sample path flushed with adapter, Kal, QC | Service programs carried out, sample path flushed with adapter, Kal, QC | Defect | SC | Termistor
ABL90 FLEX | Wartung lt. Vertrag Lösungspack erneuert | Wartung lt. Vertrag Lösungspack erneuert | SWO-8584 | Yes | Maintenance according to contract solution pack renewed | Maintenance according to contract solution pack renewed | Re classified RMED | Preventive maintenance
ABL90 |  |  | CE362562622836D4 | No |  |  | Defect | SC | Wetability
ABL90 FLEX | ASTM.INI neu aufgespielt. Kommunikation mit Herrn Mecking KFH überprüft, der QC Analysatorname immer noch nicht korrekt übertragen. Mit RDEU IT und Conworx nochmal die INI Version geprüft, INI Version ist korrekt. Conworx muss noch nacharbeiten. | ASTM.INI neu aufgespielt. Kommunikation mit Herrn Mecking KFH überprüft, der QC Analysatorname immer noch nicht korrekt übertragen. Mit RDEU IT und Conworx nochmal die INI Version geprüft, INI Version ist korrekt. Conworx muss noch nacharbeiten. | SWO-9051 | No | Astm. INI replayed. Communication with Mr. Mecking KFH verified, the QC analyzer name still not transferred correctly. With RDEU IT and Conworx the INI version is checked again, INI version is correct. Conworx has yet to work. | Astm. INI replayed. Communication with Mr. Mecking KFH verified, the QC analyzer name still not transferred correctly. With RDEU IT and Conworx the INI version is checked again, INI version is correct. Conworx has yet to work. | No fault found | HIS/LIS | Conworx
ABL835 FLEX |  |  | CE34DBD2EC0FFC2B | No
ABL825 FLEX | Verified the problem and foun that the waste rotor was not staying down. | Replaced the waste rotor. | CE2D50ED284D5427 | No | Verified the problem and foun that the waste rotor was not staying down. | Replaced the waste rotor.
ABL827 FLEX |  |  | CE2DAD3E28233EBC | No
AQT90 FLEX | Lieferung von TnT und CRP Test Kit. Kalibrierung und Durchführung QC erfolgte durch das Labor alle QC i.O Lieferung und Rücknahme erfolgte mit Lieferschein | Lieferung von TnT und CRP Test Kit. Kalibrierung und Durchführung QC erfolgte durch das Labor alle QC i.O Lieferung und Rücknahme erfolgte mit Lieferschein | SWO-8731 | Yes | Delivery of TnT and CRP Test Kit. Calibration and execution QC was carried out by the laboratory all QC i.O delivery and return took place with delivery note | Delivery of TnT and CRP Test Kit. Calibration and execution QC was carried out by the laboratory all QC i.O delivery and return took place with delivery note | Re classified RMED | SWU
ABL90 FLEX | oldat csomag cseréje |  | SWO-8792 | No | Solution Package Replacement
ABL90 | Examined Hemolyzer | Replaced Hemolyzer. Performed yearly maintenance. Verified operation of ABL90. | CE2EFA3B9E19C933 | No | Examined Hemolyzer | Replaced Hemolyzer. Performed yearly maintenance. Verified operation of ABL90. | Defect | HW | Hemolyzer
ABL835 FLEX | inlet coule | replaced friction detector pompe tubing et inlet gasket du client | CE2ECE989E3720FA | No | inlet sinks | replaced friction detector pompe tubing et inlet gasket du client
ABL815 FLEX | Ventile ausgebaut und zerlegt, in Ordnung. Messkammer zerlegt und gereinigt. | Ventile ausgebaut und zerlegt, in Ordnung. Messkammer zerlegt und gereinigt. | SWO-8913 | No | Valves removed and disassembled, fine. Measuring chamber disassembled and cleaned. | Valves removed and disassembled, fine. Measuring chamber disassembled and cleaned.
ABL825 FLEX | Service call prior the FSE performed bleaching, tHb Cal and complete oxi test | Replaced hemolyzer | CE3A359EC431D487 | No | Service call prior the FSE performed bleaching, tHb Cal and complete oxi test | Replaced hemolyzer
ABL837 FLEX |  |  | CE34D875EC06FC29 | No
ABL817 FLEX | Gerät überprüft, Einlasssonde erneuert, Gerät neu gestartet, Spüllösung erneuert, Kal durchgeführt. | Gerät überprüft, Einlasssonde erneuert, Gerät neu gestartet, Spüllösung erneuert, Kal durchgeführt. | SWO-8895 | No | Device checked, inlet probe renewed, device restarted, rinsing solution renewed, Kal performed. | Device checked, inlet probe renewed, device restarted, rinsing solution renewed, Kal performed.
ABL837 FLEX |  |  | CE34D55EEC18FC25 | No
ABL820 FLEX | Unit has been performing OK.Reinstalled Inlet module. | Returned unit back to original configuration. | CE2EAAEA9E261141 | Yes | Unit has been performing OK. Reinstalled Inlet module. | Returned unit back to original configuration. | Re classified RMED | Courtesy
ABL825 FLEX |  |  | CE2FE30E141A1D94 | Yes |  |  | Re classified RMED | Planned activity
ABL810 FLEX |  |  | CE2EA7859E3B480D | No
ABL820 FLEX | Came to a conclusion that the AQC PCB needed repalced | Repalced AQC PCB Reprogramed AQC PCB Tried to get analyzer to fail again with no avail. Analyzer AQC module is now performing within specifications. AQC PCB corrected the Breaker arm position issues. | CE2EB7BA9E419CEF | No | Came to a conclusion that the AQC PCB needed repalced | Repalced AQC PCB Reprogramed AQC PCB Tried to get analyzer to fail again with no avail. Analyzer AQC module is now performing within specifications. AQC PCB corrected the Breaker arm position issues.
AQT90 FLEX | Instalación del equipo. |  | CE2EDA6D9E486AE4 | Yes | Installation of the equipment. |  | Re classified RMED | Instruction only
ABL80 CO-OX | Tried re-installing SC. | Replaced SC (customer's stock). | CE301EA3141AA80E | No | Tried re-installing SC. | Replaced SC (customer's stock).
ABL80 FLEX CO-OX | Fehlersuche Sensorkassette gespült Schlauchsystem visuell geprüft QC-Messungen mit interner QC Level 1 und 2 | Fehlersuche Sensorkassette gespült Schlauchsystem visuell geprüft QC-Messungen mit interner QC Level 1 und 2 | SWO-9014 | No | Troubleshooting sensor cartridge flushed hose system visually tested QC measurements with internal QC level 1 and 2 | Troubleshooting sensor cartridge flushed hose system visually tested QC measurements with internal QC level 1 and 2
ABL90 FLEX | mehrfach Lösungstransportprobleme gelöst durch LP Wechsel und inlet connector Dichtung erneuert und richtig eingesetzt SC ersetzt und Startup abgewartet, danach wieder gl. Thermister Fehler, auch nach Sensor Interface Board Tausch Gerät mit in die GS genommen zwecks Reparatur | Serviceprogramme durchgeführt, dabei Probenweg gespült mit Adapter, Kal, QC | SWO-9121 | No | multiple solution transport problems solved by LP change and inlet connector seal renewed and properly inserted SC replaced and startup waited, then gl. Thermister error, even after sensor interface board exchange device taken into the GS for repair | Service programs carried out, sample path flushed with adapter, Kal, QC | Unknown | Revisit required | In GS for repair
ABL90 | Found evidence of a leak at the oxi valve tubing. | Replaced the oxi valve tubing. | CE300371145AA304 | No | Found evidence of a leak at the oxi valve tubing. | Replaced the oxi valve tubing. | Defect | HW | Oxi Module | Pinch Valve tube
ABL815 FLEX | Gerät überprüft, El Met Modul R1627N002 zurückgetauscht, Meßkammer und Luftauslass gesäubert, Elektroden membraniert, E1001 ersetzt | El Met Modul probeweise erneuert, Elektroden kontrolliert, Luftauslass gespült | SWO-9139 | No | Device checked, El Met module R1627N002 replaced, measuring chamber and air outlet cleaned, electrodes membraneed, E1001 replaced | El Met module renewed on a trial basis, controlled electrodes, air outlet flushed
ABL825 FLEX | pCO2 Elektrode getauscht (RK04) Gerät kalibriert und überprüft | pCO2 Elektrode getauscht (RK04) Gerät kalibriert und überprüft | SWO-9076 | No | pCO2 electrode exchanged (RK04) device calibrated and checked | pCO2 electrode exchanged (RK04) device calibrated and checked
ABL825 FLEX | Tausch der Messmodule, pH BG + EL MET Modul. Einkalibrieren der notwendigen Bereiche, Nullstromkalibrierung usw. Tausch der Lac-Elektrode. Tausch der Einlasssonde und der Einlassdichtung. | Tausch der Messmodule, pH BG + EL MET Modul. Einkalibrieren der notwendigen Bereiche, Nullstromkalibrierung usw. Tausch der Lac-Elektrode. Tausch der Einlasssonde und der Einlassdichtung. | SWO-9016 | No | Replacement of the measuring modules, pH BG + EL MET module. Calibration of the necessary areas, zero current calibration, etc. Replacement of the lac electrode. Replacement of the inlet probe and the inlet seal. | Replacement of the measuring modules, pH BG + EL MET module. Calibration of the necessary areas, zero current calibration, etc. Replacement of the lac electrode. Replacement of the inlet probe and the inlet seal.
ABL825 FLEX |  |  | CE2FE26614361D91 | Yes |  |  | Re classified RMED | Planned activity
ABL825 FLEX | QC ampoule barcodes were not being read. All other functions of the Autocheck were operational. Laser scanner would turn on and was aligned but would only read at best several QC Barcodes from a full carousel. Determined cause to be faulty Laser scanner | Replaced Laser scanner. Tested scanner. System operational. | CE2FF55C142F46AA | No | QC ampoule barcodes were not being read. All other functions of the Autocheck were operational. Laser scanner would turn on and was aligned but would only read at best several QC Barcodes from a full carousel. Determined cause to be faulty Laser scanner | Replaced Laser scanner. Tested scanner. System operational.
ABL825 FLEX | checked hemolyzer and lamp unit | performed lamp adjust and tHb cal, performed complete Oxi test, instrument fails NEON POSITION TEST 2.200 high | CE35CB10625E1F41 | No | checked hemolyzer and lamp unit | performed lamp adjust and tHb cal, performed complete Oxi test, instrument fails NEON POSITION TEST 2.200 high
ABL80 OSM | Fluid is not flowing. | Flushed the internal tubings. | CE31A0808A0E4EA7 | No | Fluid is not flowing. | Flushed the internal tubings.
ABL80 CO-OX | Customer had already changed the sensor cassette. Performed zero check and all within spec but zeroed analyzer anyway, ok. Level 1 still out low on external QC. Replaced sensor cassette with customers stock. Replacement failed on pO2 sensitivity too high. Replaced sensor cassette. | Replacing sensor cassette X2 corrected pO2 external QC errors. | CE3155318A118350 | No | Customer had already changed the sensor cassette. Performed zero check and all within spec but zeroed analyzer anyway, ok. Level 1 still out low on external QC. Replaced sensor cassette with customers stock. Replacement failed on pO2 sensitivity too high. Replaced sensor cassette. | Replacing sensor cassette X2 corrected pO2 external QC errors.
ABL90 | Noted messages in Activity log for 1324 - Inhomogeneous rinse solution (LS sensors), Inhomogeneous QC1, QC3 solution. | Solution pack was replaced by customer with one I delivered, no further reported messages. Inspected, sensor interface PCB, no visual issues observed, performed associated services tests, all parameters within specifications. Ran all 3 levels of QC (12 11 & 12 12 2014) for customers as part of QC study as a courtesy, all results within specifications. Due to initial issues with sensor cassette, provided customer with spare solution packs (qty 2 - RK-01) and sensor cassettes (qty 2 - R0599) to allow completion of 20 day QC study without interruption. | CE3132238A09DD90 | No | Noted messages in Activity log for 1324 - Inhomogeneous rinse solution (LS sensors), Inhomogeneous QC1, QC3 solution. | Solution pack was replaced by customer with one I delivered, no further reported messages. Inspected, sensor interface PCB, no visual issues observed, performed associated services tests, all parameters within specifications. Ran all 3 levels of QC (12 11 & 12 12 2014) for customers as part of QC study as a courtesy, all results within specifications. Due to initial issues with sensor cassette, provided customer with spare solution packs (qty 2 - RK-01) and sensor cassettes (qty 2 - R0599) to allow completion of 20 day QC study without interruption.
ABL825 FLEX | Checked hemolyzer, performed complete oxi test, all failed | removed Oxi module and checked all connections, performed re start, instrument rebooted without errors | CE315B148A2C1F28 | No | Checked hemolyzer, performed complete oxi test, all failed | removed Oxi module and checked all connections, performed re start, instrument rebooted without errors
ABL825 FLEX | checked flow of solutions, checked valves and sensors, checked electrodes and electrode camber, | removed modules and cleaned dual housing valves, replaced pH seal ring with new style, cleaned electrode chambers, replaced electrode membrane | CE315E498A1C1F2B | No | checked flow of solutions, checked valves and sensors, checked electrodes and electrode camber, | removed modules and cleaned dual housing valves, replaced pH seal ring with new style, cleaned electrode chambers, replaced electrode membrane
ABL835 FLEX |  |  | CE3159358A392B4F | No
ABL800 BASIC | Gerät überprüft, Meßkammer und Luftauslass gesäubert, pCO2 und Lac-Elektrode membraniert, Dichtungsring pH ersetzt, extra Dezimalstellen programmiert |  | SWO-9274 | No | Device checked, measuring chamber and air outlet cleaned, pCO2 and lac electrode membraned, sealing ring pH replaced, extra decimal places programmed
ABL825 FLEX | Checked hemolyzer and lamp unit, performed complete oxi test | replaced Oxi Control board(903-125), performed tHb cal in service mode, performed complete oxi test for more information, Fneon value at 2.2 exceeding limp of 2.0 | CE32FB00002D1F2F | No | Checked hemolyzer and lamp unit, performed complete oxi test | replaced Oxi Control board(903-125), performed tHb cal in service mode, performed complete oxi test for more information, Fneon value at 2.2 exceeding limp of 2.0
ABL825 FLEX | Checked hemolyzer and lamp unit, checked data logs for errors | removed and replaced spectrometer, performed complete Oxi test, ran tHb cal in service and user, | CE32FD1800301F32 | No | Checked hemolyzer and lamp unit, checked data logs for errors | removed and replaced spectrometer, performed complete Oxi test, ran tHb cal in service and user,
ABL80 CO-OX | he would like to have the analyzers checked out. | preformed zero check on the analyzer the po2 channel was slightly off. I cleaned, and tighten the valves on the manifold. | CE3521C2EC329DEC | No | he would like to have the analyzers checked out. | preformed zero check on the analyzer the po2 channel was slightly off. I cleaned, and tighten the valves on the manifold.
ABL80 CO-OX | pO2 fails linearity but passes system cycle and QC. On a prior visit I replaced the analog board, but this did not resolve the problem. The customer had changed the SC several times since the failure, but the pO2 is still failing linearity. | Replaced the sensor connector board and the spine board. A linearity kit is not available, so I will return to perform another linearity kit. Analyzer is fully operational. | CE354467EC1CFDDE | No | pO2 fails linearity but passes system cycle and QC. On a prior visit I replaced the analog board, but this did not resolve the problem. The customer had changed the SC several times since the failure, but the pO2 is still failing linearity. | Replaced the sensor connector board and the spine board. A linearity kit is not available, so I will return to perform another linearity kit. Analyzer is fully operational.
ABL90 FLEX | nach Rücksprache K.Rudolph ist Problem eindeutig auf die SC zurückzuführen, Anwender erklärt, dass es momentan qualitative Probleme mit den SC gibt, im Fehlerfall gleiche SC erneut installieren oder ggf. neue SC einsetzen, Kal, QC, Probemessung Datensicherung gemacht (wird mit SC) zu K.Rudolph gesandt | nach Rücksprache K.Rudolph ist Problem eindeutig auf die SC zurückzuführen, Anwender erklärt, dass es momentan qualitative Probleme mit den SC gibt, im Fehlerfall gleiche SC erneut installieren oder ggf. neue SC einsetzen, Kal, QC, Probemessung Datensicherung gemacht (wird mit SC) zu K.Rudolph gesandt | SWO-9421 | No | after consultation K.Rudolph problem is clearly due to the SC, user explains that there are currently qualitative problems with the SC, in case of error reinstall the same SC or if necessary use new SC, Kal, QC, sample measurement data backup (is made with SC) sent to K.Rudolph | after consultation K.Rudolph problem is clearly due to the SC, user explains that there are currently qualitative problems with the SC, in case of error reinstall the same SC or if necessary use new SC, Kal, QC, sample measurement data backup (is made with SC) sent to K.Rudolph | Defect | SC | Termistor
ABL90 FLEX | Halbjahreswartung tHb kal. | Halbjahreswartung tHb kal. | SWO-9402 | Yes | Halbjahreswartung tHb kal. | Halbjahreswartung tHb kal. | Re classified RMED | Preventive maintenance | tHb Cal
ABL90 FLEX | De zelfvullende safe-pico spuit is iets groter van diameter en maakt geen tikkend geluid (kogeltje) bij mengen, de safe-pico aspirator wel. Heb met beide spuiten en verschillende volumina uitgetest waarbij steeds GOED gemengd wordt. Als de verlichting goed brandt is de spuit correct ingelegd. Als door niet strak opplakken van etiketten de spuit niet niet vlak ingelegd wordt kan het voorkomen dat de spuit niet herkend wordt, flikkerde verlichting. Dan zal mengen ook niet goed gaan. Er is ook altijd een safe-tip cap nodig als afsluiting van een spuit. Er is geen afstelling voor menger beschikbaar. | De zelfvullende safe-pico spuit is iets groter van diameter en maakt geen tikkend geluid (kogeltje) bij mengen, de safe-pico aspirator wel. Heb met beide spuiten en verschillende volumina uitgetest waarbij steeds GOED gemengd wordt. Als de verlichting goed brandt is de spuit correct ingelegd. Als door niet strak opplakken van etiketten de spuit niet niet vlak ingelegd wordt kan het voorkomen dat de spuit niet herkend wordt, flikkerde verlichting. Dan zal mengen ook niet goed gaan. Er is ook altijd een safe-tip cap nodig als afsluiting van een spuit. Er is geen afstelling voor menger beschikbaar. | SWO-9412 | No | The self-filling safe-pico syringe is slightly larger in diameter and does not make a ticking sound (ball) when mixing, the safe-pico aspirator though. Have tested with both syringes and different volumes, which always blends well. If the lighting is properly lit, the syringe is correctly inlaid. If, by not tightly sticking labels, the syringe does not become flat inlaid, it can prevent the syringe from being detected, flickering lighting. Then mixing will not go well either. There is also always a safe-tip cap needed to close a syringe. No adjustment is available for mixer. | The self-filling safe-pico syringe is slightly larger in diameter and does not make a ticking sound (ball) when mixing, the safe-pico aspirator though. Have tested with both syringes and different volumes, which always blends well. If the lighting is properly lit, the syringe is correctly inlaid. If, by not tightly sticking labels, the syringe does not become flat inlaid, it can prevent the syringe from being detected, flickering lighting. Then mixing will not go well either. There is also always a safe-tip cap needed to close a syringe. No adjustment is available for mixer.
ABL90 FLEX | Einlaßdichtung ausgetauscht, im Serviceprogramm alle Fließwege und Sensoren getestet => alles ok. Startverfahren wieder abgebrochen. LP ausgetauscht => gleicher Fehler. Warmstart durchgeführt, Startverfahren => Fehler bleibt bestehen. SC ausgetauscht. Startverfahren, Kalibrationen, Funktionskontrolle | Einlaßdichtung ausgetauscht, im Serviceprogramm alle Fließwege und Sensoren getestet => alles ok. Startverfahren wieder abgebrochen. LP ausgetauscht => gleicher Fehler. Warmstart durchgeführt, Startverfahren => Fehler bleibt bestehen. SC ausgetauscht. Startverfahren, Kalibrationen, Funktionskontrolle | SWO-9538 | No | Inlet seal replaced, all flow paths and sensors tested in the service program => all ok. Start-up procedure aborted again. LP replaced => same error. Warm start, start procedure => error remains. SC. Start-up procedures, calibrations, function control | Inlet seal replaced, all flow paths and sensors tested in the service program => all ok. Start-up procedure aborted again. LP replaced => same error. Warm start, start procedure => error remains. SC. Start-up procedures, calibrations, function control | Defect | SC | Wetability
ABL90 FLEX | LP und SC nicht richtig initialisiert, Test wiederholen durchgeführt, Kal und QC durchgeführt | LP und SC nicht richtig initialisiert, Test wiederholen durchgeführt, Kal und QC durchgeführt | SWO-9548 | No | LP and SC not properly initialized, test repeat performed, Kal and QC performed | LP and SC not properly initialized, test repeat performed, Kal and QC performed | User Error | SP | Not activated
ABL825 FLEX | New membranes did not resolve parameter issues Protien remover and cleaning programs did not help sensor issues | Repalced inlet liquid sensor Replaced Lac and Na electrodes. | CE4073FC124A350C | No | New membranes did not resolve parameter issues Protien remover and cleaning programs did not help sensor issues | Repalced inlet liquid sensor Replaced Lac and Na electrodes.
ABL825 FLEX | Found no errors with complete oxi test | Replaced hemolyzer adjusted lamp and performed service thb cal | CE40732112233509 | No | Found no errors with complete OXI test | Replaced hemolyzer adjusted lamp and performed service thb cal
ABL825 FLEX | INSPECTED ELECTRODE AND FOUND BLOOD CLOT ON ELECTRODE. | TRIED TO REMOVE BLOOD CLOT AND ELECTRODE SHATTERED DURING MAINTENANCE. | CE354085EC5E2EF4 | No | INSPECTED ELECTRODE AND FOUND BLOOD CLOT ON ELECTRODE. | TRIED TO REMOVE BLOOD CLOT AND ELECTRODE SHATTERED DURING MAINTENANCE.
ABL837 FLEX | Chequeo de la señal de electrodos de CRE-A y CRE-B cambio de membranas, sustitución de las soluciones de CAL1, CAL2 limpieza y solución de lavado, realización del zero de las señales de electrodos por servicio, calibración y test controles de calidad, funcionando correctamente. | Cambio de membranas y soluciones. | CE35244FEC3C182C | No | DNA and CRE-B electrode signal check membrane change, replacement of CAL1 solutions, CAL2 cleaning and washing solution, zero realization of electrode signals by service, calibration and test quality controls, operating Correct. | Change of membranes and solutions.
ABL825 FLEX |  |  | CE350958EC3D8671 | No
AQT90 FLEX |  |  | CE350E32EC0A7902 | No
AQT90 FLEX | Im Storage wheele Halter abgebrochen ( durch vom Anwender auf Kopf eingelegte Kassette ), Abgebrochenes Teil ins Gerät gedrückt dadurch Folgefehler Process wheele und Process unit, E-Teil bestellt - Folge Einsatz notwendig | Kundin berichtet, dass sie eine Testkassette einlegen wollte und diese mit dem Barcode nach oben eingelegt hat, sie konnte die Testkassette wieder herausnehmen, danach stand er AQT aber auf rot, sie hat den AQT neu gestartet, nach dem Neustart zeigte er die angegeben Fehler an und verlangt nach einem erneuten Neustart, Kundin bittet um einen Einsatz des zuständigen Technikers, bitte um Kontaktaufnahme. | SWO-9547 | No | In the storage wheele holder aborted (by the user inserted on head ), Aborted part pressed into the device thereby subsequent error Process wheele and Process unit, E-part ordered - follow-up use necessary | Kundin reports that she wanted to insert a test cartridge and inserted it with the barcode up, she was able to take out the test cartridge again, but after that he stood AQT on red, she restarted the AQT, after the restart he showed the specified errors and requests a restart again, customer asks for the appropriate technician, please contact us.
TCM-COMBIM |  |  | CE34F8D4EC25C107 | No |  |  | Defect | Consumables | Electrode | Sensor 84
ABL80 FLEX BASIC | Gerät in GSWKST abgezogen | Gerät in GSWKST abgezogen | SWO-9664 | No | Device subtracted in GSWKST | Device subtracted in GSWKST
ABL90 |  |  | CE366F28620BAA1F | No |  |  | Defect | HW | Waste tube
ABL90 FLEX | Impedanzmessung im Serviceprog: 63,7mmV, Gerät zerlegt, Wasteschlauch, Ventilschlauch, Einlassdichtung ersetzt, Kontakte an Solution PCB verbogen->ersetzt, Gerät durchgespült nach Startphase erneut Impedanzmessung im Serviceprog: 217mV->LP ersetzt, danach inhomogene Spüllösung->Blutmessung, Impedanzmessung 221mV->Reinigung Waste anschluß->danach 50mV->Waste connector Unit sollte ersetzt werden->Ersatzteil bestellt | Hr Wyzisk bittet um eien Service - Einsatz zur Beseitigung der Störung. Bitte um Kontaktaufnahme zur Terminabsprache. | SWO-9606 | No | Impedance measurement in the service program: 63.7mmV, device disassembled, waste hose, valve hose, inlet seal replaced, contacts bent to Solution PCB,>replaced, device flushed through after start phase again impedance measurement in service mode: 217mV->LP replaced, then inhomogeneous Rinsing solution->blood measurement, impedance measurement 221mV->Cleaning Waste connection->after 50mV->Waste connector Unit should be replaced->Spare part ordered | Mr Wyzisk asks for a service - commitment to remedy the fault. Please contact us to make an appointment. | Defect | SP | Leak current
ABL90 FLEX | Akku defekt neuer Akku eingebaut Analysator lief kurze Zeit dann ist der Akku abgeraucht und der Analysator startet nicht mehr Netzteil und Akku bestellt Erneuter Kundenbesuch notwendig 2x Akku defekt eingeschickt P0035N024 und P0050N004 | Akku defekt neuer Akku eingebaut Analysator lief kurze Zeit dann ist der Akku abgeraucht und der Analysator startet nicht mehr Netzteil und Akku bestellt Erneuter Kundenbesuch notwendig 2x Akku defekt eingeschickt P0035N024 und P0050N004 | SWO-9621 | No | Battery defective new battery built-in analyzer ran short time then the battery is smoked and the analyzer no longer starts power supply and battery ordered Re-customer visit necessary 2x battery defective sent in P0035N024 and P0050N004 | Battery defective new battery built-in analyzer ran short time then the battery is smoked and the analyzer no longer starts power supply and battery ordered Re-customer visit necessary 2x battery defective sent in P0035N024 and P0050N004 | Defect | HW | Battery module
ABL90 FLEX | Fehlersuche Probenweg gespült LS-Kalibrierung und Pumpenkalibrierung im Serviceprogramm durchgeführt QC-Messungen mit QCAutoABL90 Level 1 bis 3 | Fehlersuche Probenweg gespült LS-Kalibrierung und Pumpenkalibrierung im Serviceprogramm durchgeführt QC-Messungen mit QCAutoABL90 Level 1 bis 3 | SWO-9610 | No | Troubleshooting sample path flushed LS calibration and pump calibration performed in service program QC measurements with QCAutoABL90 Level 1 to 3 | Troubleshooting sample path flushed LS calibration and pump calibration performed in service program QC measurements with QCAutoABL90 Level 1 to 3 | Cleaned | Fluidic pathway | Wetability?
ABL90 FLEX | Fehlersuche >>> System durchgespült. In Seviceprogramm alle vorhandene Optionen getestet. Sensorkassette ausgetauscht. neustart. | Fehlersuche >>> System durchgespült. In Seviceprogramm alle vorhandene Optionen getestet. Sensorkassette ausgetauscht. neustart. | SWO-9766 | No | Troubleshooting >>> System flushed through. Tested all existing options in Sevice program. Sensor cartridge. Restart. | Troubleshooting >>> System flushed through. Tested all existing options in Sevice program. Sensor cartridge. Restart. | Defect | SC | Wetability
ABL825 FLEX | checked logs | replaced pump rotor AND Ph TUBING | CE36CB84620EBDBC | No | checked logs | replaced pump rotor AND Ph TUBING
ABL825 FLEX | Hb reading low on WDC | Performed Lamp adjustment. Complete oxi test | CE3654C06245664E | No | Hb reading low on WDC | Performed Lamp adjustment. Complete Oxy Test
ABL825 FLEX | Had replaced Na electrode on previous visit, still an issue | Replaced ref electorde | CE40779D12543518 | No | Had replaced Na electrode on previous visit, still an issue | Replaced ref electorde
AQT90 FLEX | Teilereste entfernt, Halter ersetzt, def. Sensor ermittelt, Besorgungsfahrt, Sensor eingebaut Gerät abgeglichen, Systemtest durchgeführt, QC durchgeführt | Kundin berichtet, dass sie eine Testkassette einlegen wollte und diese mit dem Barcode nach oben eingelegt hat, sie konnte die Testkassette wieder herausnehmen, danach stand er AQT aber auf rot, sie hat den AQT neu gestartet, nach dem Neustart zeigte er die angegeben Fehler an und verlangt nach einem erneuten Neustart, Kundin bittet um einen Einsatz des zuständigen Technikers, bitte um Kontaktaufnahme. | SWO-9749 | No | Parts leftovers removed, holder replaced, def. Sensor detected, errand drive, sensor-installed device matched, system test performed, QC performed | Kundin reports that she wanted to insert a test cartridge and inserted it with the barcode up, she was able to take out the test cartridge again, but after that he stood AQT on red, she restarted the AQT, after the restart he showed the specified errors and requests a restart again, customer asks for the appropriate technician, please contact us.
AQT90 FLEX | Sensor ern., Gerät neu gestartet, QCs durchgef. |  | SWO-9697 | No | Sensor ern., device restarted, QCs through.
TCM-COMBIM |  |  | CE3F02E29C63383F | No |  |  | Defect | Consumables | Electrode
ABL80 FLEX CO-OX | Stark verunreinigtes Gerät gereinigt. Flüssigkeit an Unterseite lässt sich nicht komplett entfernen. Kalibrationen und Qualitätskontrollen in Ordnung. | Nachtrag Fahrzeit und Kilometer zu geschlossenem SWO-9810. | SWO-9810 | No | Highly contaminated device cleaned. Liquid on the bottom cannot be completely removed. calibrations and quality controls in order. | Supplement Travel time and kilometers to closed SWO-9810.
ABL80 FLEX CO-OX | Nachtrag Fahrzeit und Kilometer zu geschlossenem SWO-9810. | Nachtrag Fahrzeit und Kilometer zu geschlossenem SWO-9810. | SWO-9938 | No | Supplement Travel time and kilometers to closed SWO-9810. | Supplement Travel time and kilometers to closed SWO-9810.
ABL90 |  |  | CE37AEDBD80685E1 | No |  |  | Cleaned | Fluidic pathway | Clot
ABL90 FLEX | Lösungspack ist unbenutzt, Haltbarkeit auf Umverpackung 15.10.2015 -> Aufdruck direkt auf LP ist er nur bis 15.01.15 haltbar -> Fotos gemacht, RTP erstellt | Lösungspack ist unbenutzt; Haltbarkeit auf Umverpackung 15.10.2015 -> Aufdruck direkt auf LP ist er nur bis 15.01.15 haltbar -> Fotos gemacht, RTP erstellt | SWO-10003 | No | Solution pack is unused; Shelf life on outer packaging 15.10.2015 -> Imprint directly on LP it can only be stored until 15.01.15 -> Photos taken, RTP created | Solution pack is unused; Shelf life on outer packaging 15.10.2015 -> Imprint directly on LP it can only be stored until 15.01.15 -> Photos taken, RTP created | Defect | SP | Mis Labeled | Exp date
ABL90 FLEX | Main Board DMS 945 eingebaut Software installiert, Gerät konfiguriert ABL kalibriert und überprüft |  | SWO-9885 | No | Main Board DMS 945 built-in software installed, device configured ABL calibrated and checked |  | Defect | HW | CPU
ABL90 FLEX | Austausch der Waste connector Unit->Impedanztest im Serviceprog: 12mV | Austausch der Waste connector Unit->Impedanztest im Serviceprog: 12mV | SWO-9777 | No | Replacement of the Waste connector Unit->Impedance test in service prog: 12mV | Replacement of the Waste connector Unit->Impedance test in service prog: 12mV | Defect | HW | Waste connector Unit
ABL90 FLEX | am 16.12.2014 um 21.00 Uhr Einlasssonde erneuern sowie Gerät nochmals spülen und neu starten lassen | am 16.12.2014 um 21.00 Uhr Einlasssonde erneuern sowie Gerät nochmals spülen und neu starten lassen | SWO-9747 | No | on 16.12.2014 at 9 p.m. renew the inlet probe and have the device rinsed and restarted again | on 16.12.2014 at 9 p.m. renew the inlet probe and have the device rinsed and restarted again | Cleaned | Fluidic pathway | Inlet Probe
ABL90 | Upon arrival, found unit giving spectrum mismatch errors on internal QCs. Found Hemolyzer leaking. | Replaced all above parts listed. Flushed out unit. Ran all service fluid flow tests. Ran service Oximetry test. Ran liquid sensor adjustment program. | CE37A60CD852F517 | No | Upon arrival, found unit giving spectrum mismatch errors on internal QCs. Found Hemolyzer leaking. | Replaced all above parts listed. Flushed out unit. Ran all service fluid flow tests. Ran service Oximetry test. Ran liquid sensor adjustment program. | Defect | HW | Hemolyzer | Leak
ABL90 | Loud noise when connecting battery Barcode not working properly Wireless not connecting | I tried for several hours to resolved the issues, but was unable. The unit has been replaced by another Radiometer analyzer. I am taking the analyzer and shipping to Radiometer. | CE37A6A1D8420D2D | Yes | Loud noise when connecting battery Barcode not working properly Wireless not connecting | I tried for several hours to resolved the issues, but was unable. The unit has been replaced by another Radiometer analyzer. I am taking the analyzer and shipping to Radiometer. | Re classified RMED | OOB | Analyzer removed | HW | Don Masters: This was an out of box failure. The analyzer was installed on 12/4/14 and initially was OK but had multiple issues during the implementation process over the next several weeks with battery, wireless and barcode problems. It was replaced with another analyzer on 12/18/14.
ABL825 FLEX | Jahreswartung durchgeführt. Schläuche ersetzt. neustart. | Es könnte das FluidicModul undicht sein (schlecht verklebt?). Der Kunde bittet um einen Service. Der Kunde hat das Gerät stillgelegt, weil auch Na und Ca außerhalb QC-Bereich ist. Der Kunde wartet auf einen Rückruf, um zu erfahren, wann ein Techniker kommt. | SWO-10035 | No | annual maintenance. replaced hoses. Restart. | It could be the FluidicModul leaking (badly glued?). The customer asks for a service. The customer has shut down the device because Na and Ca are also outside the QC area. The customer waits for a callback to know when a technician is coming.
ABL815 FLEX | AU Control Board als Test eingebaut. Keine Verbesserungen der QC Ergebnisse beim pCO2. Kundeneigenes AU Board wieder eingebaut. |  | SWO-9736 | No | AU Control Board as a test. No improvements to QC results in pCO2. Customer-owned AU board reinstalled.
ABL825 FLEX | NONE NECESSARY | INSTALLED NEW ELECTRODE WITH FRESH MEMBRANE | CE37BE6CD818442D | No | NONE NECESSARY | INSTALLED NEW ELECTRODE WITH FRESH MEMBRANE
AQT90 FLEX | SW 8.6.118 neu installiert. | Kunde berichtet, dass das Gerät wohl heute Nacht die ersten Anzeichen des Fehler gezeigt hat, gemeldet wird Error in Loader, Roma load Errer, Loading data from disc failed, Kunde bittet um einen Einsatz des zuständigen Technikers, bitte um Kontaktaufnahme. | SWO-9846 | No | SW 8.6.118. | Customer reports that the device probably showed the first signs of the error tonight, Error is reported in Loader, Roma load Errer, Loading data from disc failed, customer asks for an operation of the responsible technician, please contact us.
AQT90 FLEX | SW neu installiert, Neustart, LQC durchgeführt. | Kunde berichtet, dass das Gerät wohl heute Nacht die ersten Anzeichen des Fehler gezeigt hat, gemeldet wird Error in Loader, Roma load Errer, Loading data from disc failed, Kunde bittet um einen Einsatz des zuständigen Technikers, bitte um Kontaktaufnahme. | SWO-10076 | No | SW reinstalled, reboot, LQC performed. | Customer reports that the device probably showed the first signs of the error tonight, Error is reported in Loader, Roma load Errer, Loading data from disc failed, customer asks for an operation of the responsible technician, please contact us.
TCM-COMBIM | THE MONITOR IS ASKING FOR THE SENSOR TO BE REMEMBRANED EVERY TWO DAYS | THE SENSOR WAS CONNECTED TO ANOTHER INSTRUMENT AND LAURA WILL MONITOR THE OPERATION TO SEE IF THE SENSOR IS ASKED TO BE REMEMBRANED AFTER TWO DAYS OF OPERATION | CE37CB84D81F4EFD | No | THE MONITOR IS ASKING FOR THE SENSOR TO BE REMEMBRANED EVERY TWO DAYS | THE SENSOR WAS CONNECTED TO ANOTHER INSTRUMENT AND LAURA WILL MONITOR THE OPERATION TO SEE IF THE SENSOR IS ASKED TO BE REMEMBRANED AFTER TWO DAYS OF OPERATION | No fault found
ABL80 FLEX CO-OX | Hemolyzer und Spectrometer gestaucht . | Hemolyzer und Spectrometer gestaucht . | SWO-10016 | No | Hemolyzer and Spectrometer compressed. | Hemolyzer and Spectrometer compressed.
ABL80 FLEX BASIC | Fehlersuche, LP und SC erneuert Gerät zurück geliefert | Fehlersuche, LP und SC erneuert Gerät zurück geliefert | SWO-10173 | No | Troubleshooting, LP and SC renewed device returned supplied | Troubleshooting, LP and SC renewed device returned supplied
ABL80 FLEX CO-OX | Heb de installatie button gereset en daarna geïnstalleerd. Nu komt de analyzer wel door de opstart procedure. De gevraagde tHb calibratie met succes uitgevoerd. Waarschijnlijk is de ontstane kristallisatie in de loop van de nacht opgelost. | Heb de installatie button gereset en daarna geïnstalleerd. Nu komt de analyzer wel door de opstart procedure. De gevraagde tHb calibratie met succes uitgevoerd. Waarschijnlijk is de ontstane kristallisatie in de loop van de nacht opgelost. | SWO-9997 | No | Have reset the installation button and then installed it. Now the analyzer does the startup process. The requested tHb calibration performed successfully. The resulting crystallization was probably dissolved during the night. | Have reset the installation button and then installed it. Now the analyzer does the startup process. The requested tHb calibration performed successfully. The resulting crystallization was probably dissolved during the night.
ABL80 FLEX CO-OX | Heb de installatie button gereset en daarna geïnstalleerd. Nu komt de analyzer wel door de opstart procedure. De gevraagde tHb calibratie met succes uitgevoerd. Waarschijnlijk is de ontstane kristallisatie in de loop van de nacht opgelost. | Heb de installatie button gereset en daarna geïnstalleerd. Nu komt de analyzer wel door de opstart procedure. De gevraagde tHb calibratie met succes uitgevoerd. Waarschijnlijk is de ontstane kristallisatie in de loop van de nacht opgelost. | SWO-9999 | No | Have reset the installation button and then installed it. Now the analyzer does the startup process. The requested tHb calibration performed successfully. The resulting crystallization was probably dissolved during the night. | Have reset the installation button and then installed it. Now the analyzer does the startup process. The requested tHb calibration performed successfully. The resulting crystallization was probably dissolved during the night.
ABL80 FLEX Basic version |  |  | CE3B7E223A0EA376 | No
ABL90 | Error 1343 cant pump solutions replace SP Ran diag with SC out with a tube installed across were SC is mounted and all test pass clott, SP check everything. Put tthe SC back in and SP check fails multiple parameters. flushed SC with syringe while out of analyzer, reinstalled SC ran clot test and SP check and all passed. This unit is not running any samples other than internal QC and rinse so no outside cloting factors are relevant, retested and failed with 1343 and request change SP. | Reinstalled another app software and retested. | CE39282D4E0B1318 | No | Error 1343 cant pump solutions replace SP Ran diag with SC out with a tube installed across were SC is mounted and all test pass clott, SP check everything. Put tthe SC back in and SP check fails multiple parameters. flushed SC with syringe while out of analyzer, reinstalled SC ran clot test and SP check and all passed. This unit is not running any samples other than internal QC and rinse so no outside cloting factors are relevant, retested and failed with 1343 and request change SP. | Reinstalled another app software and retested. | Defect | SC | Wetability
ABL90 |  |  | CE3FE20512503AD0 | Yes |  |  | Re classified RMED | Not upSWO | no info
ABL90 FLEX | Netzteil 910-285P0015N590 ausgebaut - defekt (siehe SWO 9621 GuiFa) Netzteil 910-285P0014N459 eingebaut Akku 910-311P0050N006 eingebaut Widerstand Schutzleiter - Gehäuse 0,9 Ohm Gerät kalibriert und überprüft | Netzteil 910-285P0015N590 ausgebaut - defekt (siehe SWO 9621 GuiFa) Netzteil 910-285P0014N459 eingebaut Akku 910-311P0050N006 eingebaut Widerstand Schutzleiter - Gehäuse 0,9 Ohm Gerät kalibriert und überprüft | SWO-10008 | No | Power supply 910-285P0015N590 removed - defective (see SWO 9621 GuiFa) Power supply 910-285P0014N459 built-in battery 910-311P0050N006 built-in resistance protective conductor - housing 0.9 Ohm device calibrated and checked | Power supply 910-285P0015N590 removed - defective (see SWO 9621 GuiFa) Power supply 910-285P0014N459 built-in battery 910-311P0050N006 built-in resistance protective conductor - housing 0.9 Ohm device calibrated and checked | Defect | HW | Power supply
ABL90 FLEX | System gespült, kassette erneut initialisiert, Impedanzfehler 1079, Kassette ersetz und Neustartphase abgewartet, Kal und QC | System gespült, kassette erneut initialisiert, Impedanzfehler 1079, Kassette ersetz und Neustartphase abgewartet, Kal und QC | SWO-10264 | No | System flushed, cassette reinitialized, impedance error 1079, cassette replace and restart phase awaited, Kal and QC | System flushed, cassette reinitialized, impedance error 1079, cassette replace and restart phase awaited, Kal and QC | Defect | SC | pCO2
ABL800 BASIC | Messkammer gereinigt, PCO2 Elektrode erneut membraniert (Lot 0556), Kal, QC | Messkammer gereinigt, PCO2 Elektrode erneut membraniert (Lot 0556), Kal, QC | SWO-10150 | No | Measuring chamber cleaned, PCO2 electrode re-membraneed (lot 0556), Kal, QC | Measuring chamber cleaned, PCO2 electrode re-membraneed (lot 0556), Kal, QC
ABL825 FLEX | Inspected autocheck ampoule motor for damage and found the motor not connected to the breaker bar. | Reinstalled the motor to the ampoule breaker. Did same thing so I replaced the ampoule breaker motor. | CE38AC674E4A16E6 | No | Inspected autocheck ampoule motor for damage and found the motor not connected to the breaker bar. | Reinstalled the motor to the ampoule breaker. Did same thing so I replaced the ampoule breaker motor.
ABL800 BASIC | pCO2 membraniert Einlaßdichtung erneuert Einlaßnadel gerichtet | pCO2 membraniert Einlaßdichtung erneuert Einlaßnadel gerichtet | SWO-10111 | No | pCO2 membraneed inlet seal renewed inlet needle directed | pCO2 membraneed inlet seal renewed inlet needle directed
ABL827 FLEX |  |  | CE3DB7E426565019 | No
ABL825 FLEX | Oxi Modul Lamp justage durchgef. Lac El. neu Membraniert Mehrfach Kalibriert. | Oxi Modul Lamp justage durchgef. Lac El. neu Membraniert Mehrfach Kalibriert. | SWO-10152 | No | Oxi Module Lamp adjustment. Lac El. new Membraneized Multiple Calibrated. | Oxi Module Lamp adjustment. Lac El. new Membraneized Multiple Calibrated.
ABL80 CO-OX | Machine sent by tech support. This is a depot service repair. Went through data logs. Found multiple PO2 problems. No obvious problems upon inspection other than errors. | Performed overhaul on machine. Performed zero current calibration with zeroing cassette. Replaced all of the hemolyzer tubing and the pump. Performed many manual system cycles. | CE3A1478C447DF02 | No | Machine sent by tech support. This is a depot service repair. Went through data logs. Found multiple PO2 problems. No obvious problems upon inspection other than errors. | Performed overhaul on machine. Performed zero current calibration with zeroing cassette. Replaced all of the hemolyzer tubing and the pump. Performed many manual system cycles.
ABL80 FLEX CO-OX | Fehlersuche Sensorkassette gespült 2-Pkt.-Kalibrierung, QC-Messungen mit interner QC Level 1 und 2 Patientenproben gemessen Sollte Fehler wieder auftreten, wird der Kunde die SC erneuern, da der Fehler erst mit neuer SC auftrat | Fehlersuche Sensorkassette gespült 2-Pkt.-Kalibrierung, QC-Messungen mit interner QC Level 1 und 2 Patientenproben gemessen Sollte Fehler wieder auftreten, wird der Kunde die SC erneuern, da der Fehler erst mit neuer SC auftrat | SWO-10253 | No | Troubleshooting sensor cartridge flushed 2-pcs. calibration, QC measurements with internal QC level 1 and 2 patient samples measured If errors occur again, the customer will renew the SC, as the error only occurred with new SC | Troubleshooting sensor cartridge flushed 2-pcs. calibration, QC measurements with internal QC level 1 and 2 patient samples measured If errors occur again, the customer will renew the SC, as the error only occurred with new SC
ABL90 FLEX | Gerät überprüft: Gerät im Serviceprog. gespült, Einlassdichtung & Anschlussdichtung ersetzt keine Verbesserung->Austausch des LP->Startphase, Kal QC: i.O. | Gerät überprüft: Gerät im Serviceprog. gespült, Einlassdichtung & Anschlussdichtung ersetzt keine Verbesserung->Austausch des LP->Startphase, Kal QC: i.O. | SWO-10177 | No | Device checked: Device in service prog. flushed, inlet seal & connection seal does not replace improvement>replacement of the LP->start phase, Kal QC: i.O. | Device checked: Device in service prog. flushed, inlet seal & connection seal does not replace improvement>replacement of the LP->start phase, Kal QC: i.O. | Defect | SP | FT
ABL90 FLEX | Servicemenue aufgerufen, ID Chipdaten wurden nicht gelesen => alle verfügbaren Tests (fehlerfrei) durchgeführt, Servicemenue verlassen, Test wiederholen aufgerufen, keine Fehler mehr. Kalibrationen, QC Messungen drei Level: OK | Servicemenue aufgerufen, ID Chipdaten wurden nicht gelesen => alle verfügbaren Tests (fehlerfrei) durchgeführt, Servicemenue verlassen, Test wiederholen aufgerufen, keine Fehler mehr. Kalibrationen, QC Messungen drei Level: OK | SWO-10284 | No | Service menu called, ID chip data was not read => all available tests (error-free) performed, service menu exit, test retry called, no more errors. Calibrations, QC measurements three levels: OK | Service menu called, ID chip data was not read => all available tests (error-free) performed, service menu exit, test retry called, no more errors. Calibrations, QC measurements three levels: OK | No fault found
ABL90 FLEX | Akku getauscht, Startphase, System Messbereit | Akku getauscht, Startphase, System Messbereit | SWO-10258 | No | Battery exchanged, start phase, system ready for measurement | Battery exchanged, start phase, system ready for measurement | Defect | Hw | Battery module
ABL90 | Customer replaced the sensor cassette, two solution packs, and inlet gasket. Went into service mode and saw that there were several failures in fluid transport. I could get all to pass except the solution pack (Cal 3 solution). | Performed an aggressive back flush, and replaced the inlet connector gasket and inlet probe, but this did not fully resolve the problem. The analyzer still fails for error code 1343 Unable to Pump Solutions . Further troubleshooting required. | CE3A761CC4222E69 | No | Customer replaced the sensor cassette, two solution packs, and inlet gasket. Went into service mode and saw that there were several failures in fluid transport. I could get all to pass except the solution pack (Cal 3 solution). | Performed an aggressive back flush, and replaced the inlet connector gasket and inlet probe, but this did not fully resolve the problem. The analyzer still fails for error code 1343 Unable to Pump Solutions . Further troubleshooting required. | Unknown | Revisit required | FT
ABL800 BASIC | Bereichsausfall der Lichtschranken bei jeder Funktion. Neustart des Analysators. Initialisierung. Bereich wird wieder erkannt, anwählen der Bereichszeile ohne Beanstandung. Kunde meldet sich wenn Problem wieder auftritt. Ggf. Display austauschen (910-320) | Bereichsausfall der Lichtschranken bei jeder Funktion. Neustart des Analysators. Initialisierung. Bereich wird wieder erkannt, anwählen der Bereichszeile ohne Beanstandung. Kunde meldet sich wenn Problem wieder auftritt. Ggf. Display austauschen (910-320) | SWO-10309 | No | Area failure of the light barriers at each function. Restart the analyzer. Initialization. Range is recognized again, selecting the range line without complaint. Customer reports when problem recurs. Replace display as appropriate (910-320) | Area failure of the light barriers at each function. Restart the analyzer. Initialization. Range is recognized again, selecting the range line without complaint. Customer reports when problem recurs. Replace display as appropriate (910-320)
ABL825 FLEX | Checked carousel and alignment,alignment way off and won't realign in service | Replaced 14 pole flex cable Alignment of auto check in service-OK | CE3A2550C455C7E1 | No | Checked carousel and alignment;alignment way off and won't realign in service | Replaced 14 pole flex cable Alignment of auto check in service-OK
ABL805 FLEX | When the analyzer ran the 2 pt cals all the parameters are showing parameters not installed. I checked and the Blood gas key was never installed. The blood gas key was not included in the option keys. It will need to be ordered. | Mark Ziegler will need to order the blood gas key. | CE39E9ACC45C1B74 | No | When the analyzer ran the 2 pt cals all the parameters are showing parameters not installed. I checked and the Blood gas key was never installed. The blood gas key was not included in the option keys. It will need to be ordered. | Mark Ziegler will need to order the blood gas key.
ABL815 FLEX | QC and Blodd samples giving error 374. Inspected Fluid path, and associated components. | Replaced Inlet nipples. located pinched Inlet to Waste tubing. | CE3A5731C403A02B | No | QC and Blodd samples giving error 374. Inspected Fluid path, and associated components. | Replaced Inlet nipples. located pinched Inlet to Waste tubing.
ABL825 FLEX | CAL VERIFICATION SOLUTION TOO COLD TO PASS. PO2 RESULTS | BROUGHT IN FRESH PACK AND RAN | CE3A0990C4122B05 | No | CAL VERIFICATION SOLUTION TOO COLD TO PASS. PO2 RESULTS | BROUGHT IN FRESH PACK AND RAN
AQT90 FLEX |  |  | CE3A49BDC40AC816 | No
TCM4 |  |  | CE3A18E5C429B626 | No |  |  | Defect | HW | Combi M Module
ABL800 BASIC | Messkammer und Schlauchsystem geprüft und gereinigt, Dichtung an der pH Elektrode ersetzt, Reinigungslösung aktiviert, AC's aufgefüllt, Funktionskontrolle ok | Messkammer und Schlauchsystem geprüft und gereinigt, Dichtung an der pH Elektrode ersetzt, Reinigungslösung aktiviert, AC's aufgefüllt, Funktionskontrolle ok | SWO-10319 | No | Measuring chamber and hose system tested and cleaned, seal replaced at pH electrode, cleaning solution activated, AC's refilled, function control ok | Measuring chamber and hose system tested and cleaned, seal replaced at pH electrode, cleaning solution activated, AC's refilled, function control ok
ABL810 FLEX | Computereinheit geprüft, Harddisk Drive demontiert und ausgetauscht, Win. Software und Betriebsystem 6.13 neu installiert, Konfiguration neu eingegeben, Funktionskontrolle ok | Computereinheit geprüft, Harddisk Drive demontiert und ausgetauscht, Win. Software und Betriebsystem 6.13 neu installiert, Konfiguration neu eingegeben, Funktionskontrolle ok | SWO-10332 | No | Computer unit checked, Harddisk Drive dismantled and replaced, Win. Software and operating system 6.13 reinstalled, configuration re-entered, function control ok | Computer unit checked, Harddisk Drive dismantled and replaced, Win. Software and operating system 6.13 reinstalled, configuration re-entered, function control ok
ABL90 FLEX | neuen Akku, original verpackt, eingebaut, ebenfalls mit gleichem Fehlerbild defekt | neuen Akku, original verpackt, eingebaut, ebenfalls mit gleichem Fehlerbild defekt | SWO-10480 | No | new battery, original packaged, built-in, also defective with same error picture | new battery, original packaged, built-in, also defective with same error picture | Defect | HW | Battery module
ABL90 |  |  | CE3E2A1A2616C573 | No |  |  | Cleaned | Fluidic pathway
ABL90 FLEX | Einlassdichtungen erneuert, mehr Tätigkeiten waren nicht möglich da laut Frau Pohlmann der Raum anschließend belegt sei, Frau Phlmann meldet sich, sollte es weiterhin Probleme geben | Einlassdichtungen erneuert, mehr Tätigkeiten waren nicht möglich da laut Frau Pohlmann der Raum anschließend belegt sei, Frau Phlmann meldet sich, sollte es weiterhin Probleme geben | SWO-10515 | No | Inlet seals were renewed, more activities were not possible because according to Ms. Pohlmann the room was occupied afterwards, Ms. Phlmann reports, should there still be problems | Inlet seals were renewed, more activities were not possible because according to Ms. Pohlmann the room was occupied afterwards, Ms. Phlmann reports, should there still be problems | Cleaned | Fluidic pathway
ABL90 | New analyzers installed one week ago. Problem occurred during training and method comparisons. Data logs show unable to pump solution error1169. All fluidics pass in service program except solution pack check. Performed multiple backflushes and blood samples but problem still there. | Problem only resolved by replacing the sensor cassette. CQC case opened | CE3F12159C305162 | No | New analyzers installed one week ago. Problem occurred during training and method comparisons. Data logs show unable to pump solution error1169. All fluidics pass in service program except solution pack check. Performed multiple backflushes and blood samples but problem still there. | Problem only resolved by replacing the sensor cassette. CQC case opened | Defect | SC | Wetability
ABL90 | Entrega y sustitucion de casete de electrodos. Cal, qc y muestras |  | CE3E65C926058DE4 | No | Delivery and replacement of electrode cassette. Cal, qc and samples |  | Defect | SC | Termistor
ABL835 FLEX | Remembrane did not fix issue | Replaced PO2 electrode | CE4081FB124F3522 | No | Remembrane did not fix issue | Replaced PO2 electrode
ABL825 FLEX | Confirmed problem. | Removed all of the modules. I tightened all of the valves. I installed the blue sealing ring on the pH electrode. I replaced the Na electrode connector. The current results are 160, 141 and 128. The results match their peers on the most recent WDC report. | CE3E06BA261ECF78 | No | Confirmed problem. | Removed all of the modules. I tightened all of the valves. I installed the blue sealing ring on the pH electrode. I replaced the Na electrode connector. The current results are 160, 141 and 128. The results match their peers on the most recent WDC report.
ABL820 FLEX | Checked data logs for errors. Found 0791 during patient samples. Also noted one occurrence of oxi liquid sensor error. Ran complete oximetry tests in diagnostics with pass results. Noticed air segments not occurring during rinse and fill. Pump calibration showed low cal factors on solution pump. | Clear air line through V7. Checked valves. Adjusted halogen lamp intensity up slightly. Replaced solution pump tubing with one from customer stock. Ran LS adjust and pump calibrations with pass results. Ran a couple of patient samples with good results reported from oximetry module. | CE3DF04C263AF117 | No | Checked data logs for errors. Found 0791 during patient samples. Also noted one occurrence of oxi liquid sensor error. Ran complete oximetry tests in diagnostics with pass results. Noticed air segments not occurring during rinse and fill. Pump calibration showed low cal factors on solution pump. | Clear air line through V7. Checked valves. Adjusted halogen lamp intensity up slightly. Replaced solution pump tubing with one from customer stock. Ran LS adjust and pump calibrations with pass results. Ran a couple of patient samples with good results reported from oximetry module.
ABL825 FLEX |  |  | CE3DD3E7262BF831 | No
ABL80 OSM | Found SP was close to expiration. Had customer replace SC. Unit still indicating Not Ready . | Customer later callede back and after a few Cals unit came Ready. | CE5095B010607065 | No | Found SP was close to expiration. Had customer replace SC. Unit still indicating Not Ready . | Customer later callede back and after a few Cals unit came Ready.
ABL90 |  |  | CE3EFE949C4D23D5 | No |  |  | User Error | Inlet gasket misaligned
ABL90 |  |  | CE50EDA7103BCE63 | No
ABL90 FLEX | Überprüfung der QC. Alle QC innerhalb der Grenzen. Keine QC grenzwertig. Zwei Vergleichsmessungen mit Blut durchgeführt ( Vergleichsanalysator ist der ABL 800 BASIC der Intensivstation). Es wurden keinerlei Auffälligkeiten bemerkt. Dr. L. Ist leider nicht vor Ort. Hebamme gibt die Ausdrücke weiter. | Überprüfung der QC. Alle QC innerhalb der Grenzen. Keine QC grenzwertig. Zwei Vergleichsmessungen mit Blut durchgeführt ( Vergleichsanalysator ist der ABL 800 BASIC der Intensivstation). Es wurden keinerlei Auffälligkeiten bemerkt. Dr. L. Ist leider nicht vor Ort. Hebamme gibt die Ausdrücke weiter. | SWO-10580 | No | Review of the QC. All Qc within the limits. No QC borderline. Two comparative measurements were carried out with blood (comparative analyzer is the ABL 800 BASIC of the intensive care unit). No abnormalities were noticed. Dr. L. Sorry, it's not on site. Midwife passes on the expressions. | Review of the QC. All Qc within the limits. No QC borderline. Two comparative measurements were carried out with blood (comparative analyzer is the ABL 800 BASIC of the intensive care unit). No abnormalities were noticed. Dr. L. Sorry, it's not on site. Midwife passes on the expressions. | No fault found | pH ?
ABL800 BASIC | Ref.Elektrode getauscht (aus Technikerbestand 4103) Kal,s und QC durchgeführt | Ref.Elektrode getauscht (aus Technikerbestand 4103) Kal,s und QC durchgeführt | SWO-10493 | No | Ref.Electrode exchanged (from technician stock 4103) Kal,s and QC performed | Ref.Electrode exchanged (from technician stock 4103) Kal,s and QC performed
ABL825 FLEX | Fluidic Module ausgetauscht. Po2 Nullwert fehler. Nullpunkt kalibration. pO2 ausgetauscht. Kal und qc in Ordnung | Es könnte das FluidicModul undicht sein (schlecht verklebt?). Der Kunde bittet um einen Service. Der Kunde hat das Gerät stillgelegt, weil auch Na und Ca außerhalb QC-Bereich ist. Der Kunde wartet auf einen Rückruf, um zu erfahren, wann ein Techniker kommt. | SWO-10563 | No | Fluidic modules. Po2 Null value error. Zero calibration. pO2. Kal and qc okay | It could be the FluidicModul leaking (badly glued?). The customer asks for a service. The customer has shut down the device because Na and Ca are also outside the QC area. The customer waits for a callback to know when a technician is coming.
ABL825 FLEX | Austausch Gas 2 Flasche Kal2- und QC Level 2,3 | Austausch Gas 2 Flasche Kal2- und QC Level 2,3 | SWO-10496 | No | Replacement gas 2 bottle Kal2- and QC Level 2.3 | Replacement gas 2 bottle Kal2- and QC Level 2.3
ABL835 FLEX | HAS ERROR CODES 328, 581, 664, AND 653. FARIBA ALREADY CHANGED INLET GASKET AND BACKFLUSHED, CLEANED INLET NEEDLE, DID LIQUID ADJUSTMENT, PUMP CALIBRATION AND COMPLETE OXI TEST | TOLD HERE TO DO 3 PROTEIN REMOVALS AND 3 CLEANINGS AND CHANGE OUT INLET PROBE. Went to site and changed out two valves and all 3 inlet nipples and inlet tubbing and inlet gasket. | CE3F26A39C23E235 | No | HAS ERROR CODES 328, 581, 664, AND 653. FARIBA ALREADY CHANGED INLET GASKET AND BACKFLUSHED, CLEANED INLET NEEDLE, DID LIQUID ADJUSTMENT, PUMP CALIBRATION AND COMPLETE OXI TEST | TOLD HERE TO DO 3 PROTEIN REMOVALS AND 3 CLEANINGS AND CHANGE OUT INLET PROBE. Went to site and changed out two valves and all 3 inlet nipples and inlet tubbing and inlet gasket.
ABL800 BASIC | Pumpenschläuche, Spüllösungsflasche, Kal1 Lösung ersetzt, allg. Sicht- und Funktionskontrolle | Pumpenschläuche, Spüllösungsflasche, Kal1 Lösung ersetzt, allg. Sicht- und Funktionskontrolle | SWO-10611 | No | Pump hoses, rinsing solution bottle, Kal1 solution replaced, general visual and functional control | Pump hoses, rinsing solution bottle, Kal1 solution replaced, general visual and functional control
ABL820 FLEX | checked data logs for errors, checked flow of solutions and samples, checked valves and liquid sensors, | removed module and cleaned valves, replaced pH seal ring, inlet probe restricted removed and replaced, performed pump cal and liquid sensor adjust, | CE3F9EF09C4C2C85 | No | checked data logs for errors, checked flow of solutions and samples, checked valves and liquid sensors, | removed module and cleaned valves, replaced pH seal ring, inlet probe restricted removed and replaced, performed pump cal and liquid sensor adjust,
TCM-COMBIM | THE INSTRUMENT WAS ASKING TO BE REMEMBRANED AVERY 24 HOURS | THE TC54 SENSOR WAS REPLACED | CE3F202B9C30A700 | No | THE INSTRUMENT WAS ASKING TO BE REMEMBRANED AVERY 24 HOURS | THE TC54 SENSOR WAS REPLACED | Defect | Consumables | Electrode | Sensor 54
ABL90 |  |  | CE4076B81261D8D6 | No |  |  | Defect | HW | Waste tube
ABL835 FLEX | Inlet leaking and PCO2 Qc trop bas | Replaced inlet gasket probe electrode PCO2,new lot of membrane | CE404BC112159F77 | No | Inlet leaking and PCO2 Qc trop bas | Replaced inlet gasket probe electrode PCO2,new lot of membrane
ABL825 FLEX | Electrode was remembraned the week before | Remembraned electrode again with a different lot number | CE407A881253351B | No | Electrode was remembraned the week before | Remembraned electrode again with a different lot number
ABL827 FLEX |  |  | CE469A7C605F4C01 | No
ABL90 | Verificação do estado do equipamento, substituição de SC e SP por um novo, verifica-se que continua com o mesmo problema, teste no equipamento de backup e continua exatamente com o mesmo problema, substituição de otra SC e deixa de dar o erro, establização e passagem do QC. |  | CE42FEF8FE5A7617 | No | Checking the state of the equipment, replacing SC and SP with a new one, it turns out that it continues with the same problem, test in the backup equipment and continues exactly with the same problem, replacing the SC Otra and ceases to give the error, establishment and passage of the Qc. |  | Defect | SC | Termistor
ABL825 FLEX | Decontamination, protein removal, cleaning, changed reference and pCO2 membrane-No change Suspect gel in system | Replace all listed parts and cleaned inlet preheater | CE42F0DBFE5B0400 | No | Decontamination, protein removal, cleaning, changed reference and pCO2 membrane-No change Suspect gel in system | Replace all listed parts and cleaned inlet preheater
TCM-COMBIM |  |  | CE42D37AFE366C78 | No |  |  | Defect | HW | PCB
ABL827 FLEX | Confirmed problem. | Replaced the inlet nipple for the syringe mode. | CE45B098EA249D8E | No | Confirmed problem. | Replaced the inlet nipple for the syringe mode.
ABL90 | Customer constantly getting fluid blockage unable to pump solution and low presssure alarm. Local FSE unable to get to site for medical reasons. RCAN FSE sent to site. | Replaced contents of the ABL90 Maintenance kit as the waste line was plugged up atih red QC or blood. Installed Solution Pack from customer stock (LOT RK-08). Replacved Sensor Cassette from customer stock (LOT R0604) Replaced new inlet gasket from customer stock as the one on the analyzer was split and fallling apart. Power supply failed, replaced it. | CE46A7456062C256 | No | Customer constantly getting fluid blockage unable to pump solution and low presssure alarm. Local FSE unable to get to site for medical reasons. RCAN FSE sent to site. | Replaced contents of the ABL90 Maintenance kit as the waste line was plugged up atih red QC or blood. Installed Solution Pack from customer stock (LOT RK-08). Replacved Sensor Cassette from customer stock (LOT R0604) Replaced new inlet gasket from customer stock as the one on the analyzer was split and fallling apart. Power supply failed, replaced it. | Defect | HW | Power supply
ABL90 FLEX | Hemolyzer ersetzt. THb Kal. Blutprobe gemessen keine keine Meldungen mehr zu dem Ergebnis. | Hemolyzer ersetzt. THb Kal. Blutprobe gemessen keine keine Meldungen mehr zu dem Ergebnis. | SWO-10812 | No | Hemolyzer. THb Kal. Blood sample no longer measured any reports of the result. | Hemolyzer. THb Kal. Blood sample no longer measured any reports of the result. | Defect | HW | Hemolyzer
ABL800 BASIC | ABL übrerprüft. Druckerpapier nicht korrekt eingelegt. Druckerpapier richtig eingelegt. Sr. Heike informiert. QC durchgeführt. | ABL übrerprüft. Druckerpapier nicht korrekt eingelegt. Druckerpapier richtig eingelegt. Sr. Heike informiert. QC durchgeführt. | SWO-10765 | No | ABL. Printer paper not loaded correctly. Printer paper loaded correctly. Sr. Heike. QC. | ABL. Printer paper not loaded correctly. Printer paper loaded correctly. Sr. Heike. QC.
ABL835 FLEX | Gerinnsel in Abfallstrecke unter Einlassdichtung. Verbindungsstutzen entfernt und durchgespült. Restliche Schlauchwege überprüft und durchgespült. Kalibration durchgeführt. Analysator messbereit. | Gerinnsel in Abfallstrecke unter Einlassdichtung. Verbindungsstutzen entfernt und durchgespült. Restliche Schlauchwege überprüft und durchgespült. Kalibration durchgeführt. Analysator messbereit. | SWO-10775 | No | Clots in waste section under inlet seal. connection nozzle is removed and rinsed through. Remaining hoseways checked and rinsed through. calibration. Analyzer ready to measure. | Clots in waste section under inlet seal. connection nozzle is removed and rinsed through. Remaining hoseways checked and rinsed through. calibration. Analyzer ready to measure.
AQT90 FLEX | Optischer Sensor für den Punch out defekt, keinen auf dem Fahrzeug. Neuer bereits bestellt, noch nicht geliefert. Nach tel. Anruf wird er morgen geliefert. Dann erneute Anfahrt. Alten Sensor ausgebaut und an den TS für RDEU geschickt. | Optischer Sensor für den Punch out defekt, keinen auf dem Fahrzeug. Neuer bereits bestellt, noch nicht geliefert. Nach tel. Anruf wird er morgen geliefert. Dann erneute Anfahrt. Alten Sensor ausgebaut und an den TS für RDEU geschickt. | SWO-10769 | No | Optical sensor for punch ing out defective, none on the vehicle. New one already ordered, not yet delivered. After the phone call, it will be delivered tomorrow. Then again. old sensor removed and sent to the TS for RDEU. | Optical sensor for punch ing out defective, none on the vehicle. New one already ordered, not yet delivered. After the phone call, it will be delivered tomorrow. Then again. old sensor removed and sent to the TS for RDEU.
ABL80 CO-OX | Spectrometer failed diag voltage test and and other diag test failed on lamp test, will need new spectrometer | Replaced spectrometer and ran THB cal ( out of box failure) | CE482A14D657665E | No | Spectrometer failed diag voltage test and and other diag test failed on lamp test, will need new spectrometer | Replaced spectrometer and ran THB cal ( out of box failure)
ABL90 FLEX | Einlasssonde nicht in Einlassdichtung arretiert, dadurch lag die Einlassdichtung nicht über Anschluss des LP. Korrigiert, Einlasssonde und Einlassdichtung o.K. Funktion danach o.K. Kal o.K. und QC o.K. | Einlasssonde nicht in Einlassdichtung arretiert, dadurch lag die Einlassdichtung nicht über Anschluss des LP. Korrigiert, Einlasssonde und Einlassdichtung o.K. Funktion danach o.K. Kal o.K. und QC o.K. | SWO-10877 | No | inlet probe is not locked in inlet seal, so the inlet seal was not above the connection of the LP. Corrected, inlet probe and inlet seal o.K. function after o.K. Kal o.K. and QC o.K. | inlet probe is not locked in inlet seal, so the inlet seal was not above the connection of the LP. Corrected, inlet probe and inlet seal o.K. function after o.K. Kal o.K. and QC o.K. | User Error | Inlet gasket misaligned
ABL90 FLEX | Kurz nach Geräteaufstellung am 19.12.14 Probleme mit pCO2. Kalibrierungen immer ok => Empfindlichkeit um die 85% QC Werte von pCO2 immer am driften nach oben, bis oberhalb außer Bereich erreicht wird => deswegen schon LP Tausch am 23.12. (Lot: RK-07 gegen Lot: RK-07) und 27.12. (Lot: RK-07 gegen Lot: RK-12)und SC Tausch am 29.12. (Lot: 278-50 gegen Lot: 278-115) Überprüfung Sensorkassettenfach auf Beschädigung => alle Pins i.O., keine Ablagerungen Reinstallierung SC => QC Level A bleibt fehlerhaft => Reinstallierung LP => QC Level A bleibt fehlerhaft Neuen LP installiert (Lot: RK-07 gegen Lot: RM-04) => pCO2 bleibt fehlerhaft => Gerät zur Reparatur mit in GS genommen Nochmals LP und SC in GS getauscht => dann alles i.O. Gerät zurück zum Kunden | Kurz nach Geräteaufstellung am 19.12.14 Probleme mit pCO2. Kalibrierungen immer ok => Empfindlichkeit um die 85% QC Werte von pCO2 immer am driften nach oben, bis oberhalb außer Bereich erreicht wird => deswegen schon LP Tausch am 23.12. (Lot: RK-07 gegen Lot: RK-07) und 27.12. (Lot: RK-07 gegen Lot: RK-12)und SC Tausch am 29.12. (Lot: 278-50 gegen Lot: 278-115) Überprüfung Sensorkassettenfach auf Beschädigung => alle Pins i.O., keine Ablagerungen Reinstallierung SC => QC Level A bleibt fehlerhaft => Reinstallierung LP => QC Level A bleibt fehlerhaft Neuen LP installiert (Lot: RK-07 gegen Lot: RM-04) => pCO2 bleibt fehlerhaft => Gerät zur Reparatur mit in GS genommen Nochmals LP und SC in GS getauscht => dann alles i.O. Gerät zurück zum Kunden | SWO-10907 | No | Shortly after device installation on 19.12.14 problems with pCO2. Calibrations always ok => Sensitivity around the 85% QC values of pCO2 always drift upwards until above range is reached => therefore already LP exchange on 23.12. (Lot: RK-07 vs. Lot: RK-07) and 27.12. (Lot: RK-07 vs. Lot: RK-12) and SC Exchange on 29.12. (Lot: RK-07) : 278-50 vs. Lot: 278-115) Checking sensor cartridge compartment for damage => all pins i.O., no deposits reinstall SC => QC Level A remains faulty => Reinstall LP => QC Level A remains faulty New LP installed (Lot: RK-07 vs. Lot: RM-04) => pCO2 remains faulty => Device taken for repair in GS Again LP and SC exchanged in GS => then all i.O. device back to the customer | Shortly after device installation on 19.12.14 problems with pCO2. Calibrations always ok => Sensitivity around the 85% QC values of pCO2 always drift upwards until above range is reached => therefore already LP exchange on 23.12. (Lot: RK-07 vs. Lot: RK-07) and 27.12. (Lot: RK-07 vs. Lot: RK-12) and SC Exchange on 29.12. (Lot: RK-07) : 278-50 vs. Lot: 278-115) Checking sensor cartridge compartment for damage => all pins i.O., no deposits reinstall SC => QC Level A remains faulty => Reinstall LP => QC Level A remains faulty New LP installed (Lot: RK-07 vs. Lot: RM-04) => pCO2 remains faulty => Device taken for repair in GS Again LP and SC exchanged in GS => then all i.O. device back to the customer | Defect | SC | pCO2 | QC
ABL815 FLEX | System überprüft und gereinigt, Zerro current Kalibration durchgeführt, Lactat Elektrode gereinigt, Kalibrationen und Qualitätskontrollen ohne Beanstandungen durchgeführt. | System überprüft und gereinigt, Zerro current Kalibration durchgeführt, Lactat Elektrode gereinigt, Kalibrationen und Qualitätskontrollen ohne Beanstandungen durchgeführt. | SWO-10905 | No | System checked and cleaned, cerrocomposition performed, lactate electrode cleaned, calibrations and quality controls carried out without complaints. | System checked and cleaned, cerrocomposition performed, lactate electrode cleaned, calibrations and quality controls carried out without complaints.
ABL825 FLEX |  |  | CE47F411D6540079 | No
ABL805 FLEX | When the option keys were loaded I did not catch that the blood gas key was included. The Demo blood gas option key expired and the parameters showed not installed. It was alsxo due a Ref membrane change, inlet gasket change and decomination. I took care of these. | Installed blood gas option key. | CE47D8D4D64858FF | No | When the option keys were loaded I did not catch that the blood gas key was included. The Demo blood gas option key expired and the parameters showed not installed. It was alsxo due a Ref membrane change, inlet gasket change and decomination. I took care of these. | Installed blood gas option key.
ABL827 FLEX | Confirmed the problem. The cleaning program was moved to a different time and the failure now occurs after the new time. | Removed all of the valves. Cleaned the membranes and reseated. Installed the new style blue sealing ring on the pH electrode. Also performed the annual preventative maintenance, which is documented on another SWO. | CE4926604C5B8B1B | No | Confirmed the problem. The cleaning program was moved to a different time and the failure now occurs after the new time. | Removed all of the valves. Cleaned the membranes and reseated. Installed the new style blue sealing ring on the pH electrode. Also performed the annual preventative maintenance, which is documented on another SWO.
ABL827 FLEX | Confirmed the problem. The cleaning program was moved to a different time and the failure now occurs after the new time. | Removed all of the valves. Cleaned the membranes and reseated. Installed the new style blue sealing ring on the pH electrode. Also performed the annual preventative maintenance, which is documented on another SWO. | CE4926704C008B1C | No | Confirmed the problem. The cleaning program was moved to a different time and the failure now occurs after the new time. | Removed all of the valves. Cleaned the membranes and reseated. Installed the new style blue sealing ring on the pH electrode. Also performed the annual preventative maintenance, which is documented on another SWO.
ABL90 |  |  | CE5A14A24A6034F5 | No |  |  | Defect | SW | SW upgrade failure | & Defect SC Cl
ABL80 CO-OX | checked data logs for srrors, checked sensor for clots, checked pumps and waste lines, checked solution pack, | removed blockage from waste line, flushed sensor, perform flush of system, power off, remove sensor, power on and replace sensor. | CE4C3D9D382556E0 | No | checked data logs for srrors, checked sensor for clots, checked pumps and waste lines, checked solution pack, | removed blockage from waste line, flushed sensor, perform flush of system, power off, remove sensor, power on and replace sensor.
ABL90 | Customer installed a new solution pack and the cycle count was zero. Trouble shot and flushed and reincerrted the solution pack. Then the PCO2 failed cal and internal QC's. Recommended replacing the sensor cassette. After replacing the sensor cassette, the new cassette failed level 3 QC. The customer does not have another cassette. I ran 3 blood samples through the cassette to try to conditionn the cassette. The unit still failed level 3 QC, both internal and external. We are letting the cassette settle over night to see if it will come in . | Flushed unit and reinstalled the solution pack. Replaced the sensor cassette. | CE4B96493804F071 | No | Customer installed a new solution pack and the cycle count was zero. Trouble shot and flushed and reincerrted the solution pack. Then the PCO2 failed cal and internal QC's. Recommended replacing the sensor cassette. After replacing the sensor cassette, the new cassette failed level 3 QC. The customer does not have another cassette. I ran 3 blood samples through the cassette to try to conditionn the cassette. The unit still failed level 3 QC, both internal and external. We are letting the cassette settle over night to see if it will come in . | Flushed unit and reinstalled the solution pack. Replaced the sensor cassette. | User Error | SP install | Chip error
TCM400 | Set up with QA electrodes and new membranes. Calibrated first time with no errors. Let run for two days with auto-calibration. | Completed Performance Test and Service Checkout. No faults found. | CE4B9055382EB360 | No | Set up with QA electrodes and new membranes. Calibrated first time with no errors. Let run for two days with auto-calibration. | Completed Performance Test and Service Checkout. No faults found.
ABL820 FLEX | Upon inspection, found the autocheck carousel housing moving up and the carousel spinning but the qc ampoule is not opened and the housing goes back down cancelling the qc. | Checked the activity log, found error 742 Autocheck carousel failure. Inspected the autocheck unit FPC cable and found it loose and creased. Replaced a new FPC cable. | CE4BBA10383701EA | No | Upon inspection, found the autocheck carousel housing moving up and the carousel spinning but the qc ampoule is not opened and the housing goes back down cancelling the qc. | Checked the activity log, found error 742 Autocheck carousel failure. Inspected the autocheck unit FPC cable and found it loose and creased. Replaced a new FPC cable. | Defect | HW | AutoCheck Module | FPC cable
ABL825 FLEX | Mischmagnet justiert, blutige Messung durchgef. - OK | Mischmagnet justiert, blutige Messung durchgef. - OK | SWO-11045 | No | Mixed magnet adjusted, bloody measurement through. - OK | Mixed magnet adjusted, bloody measurement through. - OK | Defect | HW | FlexQ Module | Magnet loose
AQT90 FLEX | SP war nicht zu entfernen. Verklemmt-Punch up ? Gerät zur rep in GS | Ersatzgerät 838R0170N006 aufgestellt und D-Dimer Messung mit Kunde durchgef. | SWO-11548 | No | SP could not be removed. Stuck punch up ? Device for rep in GS | Replacement device 838R0170N006 installed and D-Dimer measurement with customer.
AQT90 FLEX | Sensor für Process Einheit eingebaut. Analysator nach Start Funktionsbereit. Start Kalibration Parameter. Anwender wird alle Parameter kalibrieren und QC messen. | Sensor für Process Einheit eingebaut. Analysator nach Start Funktionsbereit. Start Kalibration Parameter. Anwender wird alle Parameter kalibrieren und QC messen. | SWO-11012 | No | Sensor for process unit installed. Analyzer ready for use after start. Start calibration parameter. User will calibrate all parameters and measure QC. | Sensor for process unit installed. Analyzer ready for use after start. Start calibration parameter. User will calibrate all parameters and measure QC.
AQT90 FLEX | Barcode PCB ersetzt, | Barcode PCB ersetzt, | SWO-11051 | No | Barcode PCB replaces, | Barcode PCB replaces,
ABL835 FLEX | ABL ohne Fehlermeldung vorgefunden Compl. Oximetrie Test- i.O. Kal seit ca.10 h o.k.,tritt aber immer wieder auf (Archiv) Austausch Oxi Modul PCB SW installiert, Einlaufverfahren | ABL ohne Fehlermeldung vorgefunden Compl. Oximetrie Test- i.O. Kal seit ca.10 h o.k.;tritt aber immer wieder auf (Archiv) Austausch Oxi Modul PCB SW installiert, Einlaufverfahren | SWO-11074 | No | ABL found without error message Compl. Oximetry test- i.O. Kal since about 10 h o.k.; occurs again and again (archive) exchange Oxi module PCB SW installed, inlet procedure | ABL found without error message Compl. Oximetry test- i.O. Kal since about 10 h o.k.; occurs again and again (archive) exchange Oxi module PCB SW installed, inlet procedure | Defect | HW | Oxi module | OXI PCB
AQT90 FLEX | Sensor Punch down reagiert nicht immer auf Wert 1. - Sensor ersetzt, QC durchgeführt | Sensor Punch down reagiert nicht immer auf Wert 1. - Sensor ersetzt, QC durchgeführt | SWO-11238 | No | Sensor punch down does not always respond to value 1. - Sensor replaced, QC performed | Sensor punch down does not always respond to value 1. - Sensor replaced, QC performed
ABL80 CO-OX | INSPECTED DATA LOGS AND ZERO CURRENT. ZERO CURRENT OUT OF RANGE AND FLUXUATING. | REPLACED ANALOG PCB AND COMPLETED REALIGNMENT OF THE TEMP AND THE ZERO. REPLACED LOW SOLUTION PACK AND TESTED. | CE4F87B59A08B794 | No | INSPECTED DATA LOGS AND ZERO CURRENT. ZERO CURRENT OUT OF RANGE AND FLUXUATING. | REPLACED ANALOG PCB AND COMPLETED REALIGNMENT OF THE TEMP AND THE ZERO. REPLACED LOW SOLUTION PACK AND TESTED.
ABL90 FLEX | Lösungspack und Sensorkassette geprüft, den Parameter Ca. deaktiviert, Funktionskontrolle ok | Lösungspack und Sensorkassette geprüft, den Parameter Ca. deaktiviert, Funktionskontrolle ok | SWO-11156 | No | Solution pack and sensor cassette checked, the parameter Approx. deactivated, function control ok | Solution pack and sensor cassette checked, the parameter Approx. deactivated, function control ok | Cleaned | Fluidic pathway
ABL90 FLEX | Fehlersuche, div. Test`s im Service durchgef., LP u. SC gew., Einlaßnadel u. -halter aus Kundenbestand erneuert, Kal. u. QC i.O. | Fehlersuche; div. Test`s im Service durchgef.; LP u. SC gew.; Einlaßnadel u. -halter aus Kundenbestand erneuert; Kal. u. QC i.O. | SWO-11207 | No | Troubleshooting; div. Test it in the service. LP and SC. Inlet needle and holder renewed from customer base; Kal. QC in the world | Troubleshooting; div. Test it in the service. LP and SC. Inlet needle and holder renewed from customer base; Kal. QC in the world | Assisted Replacement | SP | FT | Wetability
ABL825 FLEX | PC Modul ausgebaut Druckereinheit erneuert, Modul eingebaut Gerät kalibriert und überprüft | PC Modul ausgebaut Druckereinheit erneuert, Modul eingebaut Gerät kalibriert und überprüft | SWO-11086 | No | PC module removed printer unit renewed, module built-in device calibrated and checked | PC module removed printer unit renewed, module built-in device calibrated and checked | Defect | HW | Printer module | Printer Head
ABL825 FLEX | PC Modul ausgebaut Druckereinheit erneuert, Modul eingebaut Gerät kalibriert und überprüft | PC Modul ausgebaut Druckereinheit erneuert, Modul eingebaut Gerät kalibriert und überprüft | SWO-11084 | No | PC module removed printer unit renewed, module built-in device calibrated and checked | PC module removed printer unit renewed, module built-in device calibrated and checked | Defect | HW | Printer module | Printer Head
ABL820 FLEX | Random errors on patient samples, pH BG module not filled. | Cleaned all tubing, checked operation of all valves, removed electrodes and cleaned measuring chamber. Pink fluid in reference membrane, installed new membrane. Cleaned electrolyte from reference membrane that was present around the pH electrode. All checking OK at this time. | CE4F92B69A385DFE | No | Random errors on patient samples, pH BG module not filled. | Cleaned all tubing, checked operation of all valves, removed electrodes and cleaned measuring chamber. Pink fluid in reference membrane, installed new membrane. Cleaned electrolyte from reference membrane that was present around the pH electrode. All checking OK at this time. | Assisted Replacement | Electrode | Membrane | Reference
ABL815 FLEX | Fehlersuche, Gerät aus, Membran d. V2 v. pH BG-Modul gew., Gerät ein, Startphase abgewartet, LS-Justage u. Pumpenkal., Kal. + QC i.O. | Fehlersuche; Gerät aus; Membran d. V2 v. pH BG-Modul gew.; Gerät ein; Startphase abgewartet; LS-Justage u. Pumpenkal.; Kal. + QC i.O. | SWO-11134 | No | Troubleshooting; device off; membrane d. V2 v. pH BG module. device; start-up phase awaited; LS-Adjustment and Pump Scale.; Kal. + QC in the first place | Troubleshooting; device off; membrane d. V2 v. pH BG module. device; start-up phase awaited; LS-Adjustment and Pump Scale.; Kal. + QC in the first place | Defect | HW | pH/BG Module | Valve, air
ABL837 FLEX | Cleaning solution is not being pumped. | Replaced cleaning solution. Cleaned measuring chamber, replaced sealing ring, cleaned tip of PCO2 electrode with cleaning solution. | CE4FACB29A2B3524 | No | Cleaning solution is not being pumped. | Replaced cleaning solution. Cleaned measuring chamber, replaced sealing ring, cleaned tip of PCO2 electrode with cleaning solution. | Assisted Replacement | Electrode | Electrode | pCO2
ABL837 FLEX | Reviewed data logs, appears sensitivity decreasing with time. | Replacwed membrane, cleaned contacts, cleaned measuring chamber. | CE4FAF719A463527 | No | Reviewed data logs, appears sensitivity decreasing with time. | Replacwed membrane, cleaned contacts, cleaned measuring chamber. | Assisted Replacement | Electrode | Membrane | Ca
ABL820 FLEX | Analyzer would not run autocheck QC due temperature error. Customer stated that analyzer was leaking out of inlet. Checked autocheck in service and temp was at 55 degrees C. Replaced carousel spin temp. pcb but temp in service mode was at 8 degrees C. Replaced carousel interface board due to evidence of spillage. Scanner would not work. Replaced autocheck module pcb. | Replaced carousel spin temp. pcb, carousel interface pcb and autocheck module pcb | CE4F86CE9A55F26C | No | Analyzer would not run autocheck QC due temperature error. Customer stated that analyzer was leaking out of inlet. Checked autocheck in service and temp was at 55 degrees C. Replaced carousel spin temp. pcb but temp in service mode was at 8 degrees C. Replaced carousel interface board due to evidence of spillage. Scanner would not work. Replaced autocheck module pcb. | Replaced carousel spin temp. pcb, carousel interface pcb and autocheck module pcb | Defect | HW | AutoCheck Module | Autocheck Distribution PCB
AQT90 FLEX | AQT open gemaakt. Storage wheel wordt geblokkerd door een afgebroken lipje van een reagens cassette, deze verwijderd. AQT opgestart, Deze doet het nu weer. Cup transport check uitgevoerd --> OK. | AQT open gemaakt. Storage wheel wordt geblokkerd door een afgebroken lipje van een reagens cassette, deze verwijderd. AQT opgestart, Deze doet het nu weer. Cup transport check uitgevoerd --> OK. | SWO-11178 | No | AQT opened. Storage Wheel is blocked by a broken tab of a reagent cassette, which is removed. AQT started up, it does it again now. Cup transport Check performed--> OK. | AQT opened. Storage Wheel is blocked by a broken tab of a reagent cassette, which is removed. AQT started up, it does it again now. Cup transport Check performed--> OK.
AQT90 FLEX | Ersatzgerät 838R0170N006 aufgestellt und D-Dimer Messung mit Kunde durchgef. | Ersatzgerät 838R0170N006 aufgestellt und D-Dimer Messung mit Kunde durchgef. | SWO-11549 | No | Replacement device 838R0170N006 installed and D-Dimer measurement with customer. | Replacement device 838R0170N006 installed and D-Dimer measurement with customer.
ABL90 |  |  | CE50B2491026A5D8 | No |  |  | Assisted Replacement | SC | Glu/Lac
ABL90 | Confirmed problem. | Removed the oximetry unit and the actuator unit and looked for evidence of a leak. I did not find any leaks in any of the tubing or modules. There was a small buildup around the waste connector. I cleaned that up. Replaced tubing with Yearly Maintenance Service Kit. Cleaned analyzer. | CE511BA410369F3C | No | Confirmed problem. | Removed the oximetry unit and the actuator unit and looked for evidence of a leak. I did not find any leaks in any of the tubing or modules. There was a small buildup around the waste connector. I cleaned that up. Replaced tubing with Yearly Maintenance Service Kit. Cleaned analyzer. | Defect | HW | Waste tube
ABL90 FLEX | Flüssigkeitswege gespült, Dichtung und Einlassnadel gereinigt--> keine Besserung, Vakuumfehler bleibt bestehen Einlassdichtung, -nadel und Anschlussdichtung erneuert, System erneut gespült --> Vakuumfehler bleibt bestehen, inhomogene Kal3-Lösung Neustart brachte keine Besserung SP und SC erneuert--> alle Lösungen inhomogen, Drucküberprüfung fehlerhaft Jahreswartung durchgeführt, Neustart ABL führt Neustart durch --> Flüssigkeitstransport i.O | Flüssigkeitswege gespült, Dichtung und Einlassnadel gereinigt--> keine Besserung, Vakuumfehler bleibt bestehen Einlassdichtung, -nadel und Anschlussdichtung erneuert, System erneut gespült --> Vakuumfehler bleibt bestehen, inhomogene Kal3-Lösung Neustart brachte keine Besserung SP und SC erneuert--> alle Lösungen inhomogen, Drucküberprüfung fehlerhaft Jahreswartung durchgeführt, Neustart ABL führt Neustart durch --> Flüssigkeitstransport i.O | SWO-11275 | No | Liquid paths rinsed, seal and inlet needle cleaned--> no improvement, vacuum fault remains inlet seal, needle and connection seal renewed, system re-rinsed --> Vacuum fault remains, inhomogeneous Kal3 solution restart brought no improvement SP and SC renews--> all solutions inhomogeneous, pressure check faulty annual maintenance performed, restart ABL performs restart --> liquid transport i.O | Liquid paths rinsed, seal and inlet needle cleaned--> no improvement, vacuum fault remains inlet seal, needle and connection seal renewed, system re-rinsed --> Vacuum fault remains, inhomogeneous Kal3 solution restart brought no improvement SP and SC renews--> all solutions inhomogeneous, pressure check faulty annual maintenance performed, restart ABL performs restart --> liquid transport i.O | Defect | HW | Waste tube
ABL90 FLEX | Sensorkassette (aus Kundenbestand) ausgetauscht, Funktionskontrolle ok |  | SWO-11318 | No | Sensor cassette (from customer stock) replaced, function control ok |  | Assisted Replacement | SC | Glu
ABL90 FLEX | Akku abgetrennt damit mit dem Gerät erst gearbeitet werden kann. | Akku abgetrennt damit mit dem Gerät erst gearbeitet werden kann. | SWO-11393 | No | battery disconnected so that the device can only be worked on. | battery disconnected so that the device can only be worked on. | Defect | HW | Battery module
ABL825 FLEX | Replaced AutoCheck Laser Scanner. Scanner turns on but fails to reset between each ampoule. Same issue was exhibited on 11 25 2014.Replacement of Laser Scanner Laser scanner failed to read QC in Diagnostics. Scanner alignment was correct. | Replaced the Laser Scanner SN 910-310R0133 with Laser Scanner SN R0107. System operational. | CE50DD1C103D2996 | No | Replaced AutoCheck Laser Scanner. Scanner turns on but fails to reset between each ampoule. Same issue was exhibited on 11 25 2014.Replacement of Laser Scanner Laser scanner failed to read QC in Diagnostics. Scanner alignment was correct. | Replaced the Laser Scanner SN 910-310R0133 with Laser Scanner SN R0107. System operational. | Defect | HW | AutoCheck Module | Barcode reader
ABL825 FLEX | Removed the solution valve module and inspect the cal solution valves. Found the Clean solution valve wire loose in the connector and cal 2 seal having a groove in the top. | Replace cleaning solution and replace seals in the calibration valve seals. | CE50D2E91006D03F | No | Removed the solution valve module and inspect the cal solution valves. Found the Clean solution valve wire loose in the connector and cal 2 seal having a groove in the top. | Replace cleaning solution and replace seals in the calibration valve seals. | Defect | HW | Fluidic manifold | Valve
ABL90 FLEX | ABL überprüft. ABL gespült (Spülutensilien beim Gerät gefunden und Anwendern gezeigt), Fließwege intern fehlerhaft. Fließwegtest visuell OK! Sensorcassette erneuert. Lösungspack erneuert. Startpahse durchgeführt. Lac kommt in der Empf. nicht. Sensorcassette erneut gewechselt. Startphase durchgeführt. | ABL überprüft. ABL gespült (Spülutensilien beim Gerät gefunden und Anwendern gezeigt), Fließwege intern fehlerhaft. Fließwegtest visuell OK! Sensorcassette erneuert. Lösungspack erneuert. Startpahse durchgeführt. Lac kommt in der Empf. nicht. Sensorcassette erneut gewechselt. Startphase durchgeführt. | SWO-11591 | No | ABL. ABL rinsed (flushing utensils found on the device and shown to users), flow paths internally faulty. Flow path test visually OK! Sensorcassette renewed. Solution pack renewed. Startpahse. Lac comes in the Empf. Not. Sensor cassette changed again. started phase. | ABL. ABL rinsed (flushing utensils found on the device and shown to users), flow paths internally faulty. Flow path test visually OK! Sensorcassette renewed. Solution pack renewed. Startpahse. Lac comes in the Empf. Not. Sensor cassette changed again. started phase. | Cleaned | Fluidic pathway
ABL90 FLEX | System überprüft, Einlass gereinigt, System durch gespült, Kalibration und Interne Qualitätskontrollen ohne Beanstandungen durchgeführt. Gemäß Aussage des Anwender sind schon blutige Messungen ohne Probleme wieder gelaufen. |  | SWO-11811 | No | System checked, inlet cleaned, system flushed, calibration and internal quality controls carried out without complaints. According to the user, bloody measurements have already run without any problems. |  | Assisted Replacement | SC | Wetability
ABL90 | LIMPIEZA DE JUNTA DE ENTRADA FORMACIÓN DEL PERSONAL DE LIMPIEZA DE CIRCUITO NEUMATICO. | LIMPIEZA DE JUNTA DE ENTRADA FORMACIÓN DEL PERSONAL DE LIMPIEZA DE CIRCUITO NEUMATICO. | CE5231D286327E3E | No | CLEANING OF INPUT TRAINING OF THE NEUMATIC CIRCUIT CLEANING STAFF. | CLEANING OF INPUT TRAINING OF THE NEUMATIC CIRCUIT CLEANING STAFF. | Cleaned | Fluidic pathway | sample failure | training
ABL800 BASIC | Proteinentfernungsprogramm durchgeführt, Lac-Elektrode neu membraniert, mehrere Kal, QC | Proteinentfernungsprogramm durchgeführt, Lac-Elektrode neu membraniert, mehrere Kal, QC | SWO-11576 | No | Protein removal program performed, Lac electrode re-membraneed, multiple Kal, QC | Protein removal program performed, Lac electrode re-membraneed, multiple Kal, QC | Assisted Replacement | Electrode | Membrane | Lac
ABL800 BASIC | Gerät demontiert und gereinigt, Verstopfung beseitigt, Funktionskontrolle ok | Die Einlassdichtung soll ersetzt werden, Hinweis auf Tutorium gegeben | SWO-11470 | No | Device dismantled and cleaned, clogging eliminated, function control ok | The inlet seal should be replaced, indication given to tutorial | Cleaned | Fluidic pathway
ABL835 FLEX | ORDER REPLACEMENT ELECTRODE FOR LAC |  | CE51D95F8622630D | No | ORDER REPLACEMENT ELECTRODE FOR LAC |  | Assisted Replacement | Electrode | Electrode | Lac
ABL835 FLEX |  | Das Gerät wurde über Taste in Analysator-Stop gebracht und Spül- und Gas1-Flaschen wurden eingescannt. Gerät geht in Nachfüllung. Der Kunde bittet um einen Service. | SWO-11630 | No |  | The device was brought into analytic stop via button and rinse and Gas1 bottles were scanned. Device goes into refill. The customer asks for a service. | Unknown
AQT90 FLEX |  |  | CE52137286076F8A | No
ABL90 FLEX | Abl90 mehrmals gespült, Abfallschlauch, Ventilschlauch und Lösungspack getauscht, Danach war Fehler behoben. Sensorkassette getauscht ( Fälligkeit vorhanden) welche von Medizintechniker vorbeigebracht wurde. Lösungspack im ABL nur ca. 10 Tage haltbar. Medizintechniker darüber informiert. Abl90 nun wieder Messbereit. Ersatzgerät aus Augustinum welches vor Ort bereit Stand in Medizintechnik abgegeben (Herr riesiger Weißstörche bescheid) | Abl90 mehrmals gespült, Abfallschlauch, Ventilschlauch und Lösungspack getauscht, Danach war Fehler behoben. Sensorkassette getauscht ( Fälligkeit vorhanden) welche von Medizintechniker vorbeigebracht wurde. Lösungspack im ABL nur ca. 10 Tage haltbar. Medizintechniker darüber informiert. Abl90 nun wieder Messbereit. Ersatzgerät aus Augustinum welches vor Ort bereit Stand in Medizintechnik abgegeben (Herr riesiger Weißstörche bescheid) | SWO-11714 | No | Abl90 rinsed several times, exchanged waste hose, valve hose and solution pack, after that error was fixed. Sensor cassette exchanged (due date present) which was brought by medical technicians. Solution pack in the ABL can only be used for about 10 days. medical technicians informed about this. Abl90 now ready to measure again. Replacement device from Augustinum which was delivered on site ready stand in medical technology (Mr. huge white stork sat notice) | Abl90 rinsed several times, exchanged waste hose, valve hose and solution pack, after that error was fixed. Sensor cassette exchanged (due date present) which was brought by medical technicians. Solution pack in the ABL can only be used for about 10 days. medical technicians informed about this. Abl90 now ready to measure again. Replacement device from Augustinum which was delivered on site ready stand in medical technology (Mr. huge white stork sat notice) | Defect | HW | Oxi module | Pinch Valve tube
ABL90 FLEX | THB kalibreret. | THB kalibreret. | SWO-11751 | No | THB kalibreret. | THB kalibreret. | Adjustment | tHb | missing sO2 | Rat blood!
ABL90 FLEX | Sensorkassette aus GS19-Bestand ersetzt, anschließende Stertphase u. Systemkontrolle i.O. | Sensorkassette aus GS19-Bestand ersetzt, anschließende Stertphase u. Systemkontrolle i.O. | SWO-11721 | No | Sensor cassette replaced from GS19 stock, subsequent death phase and system control in the o. | Sensor cassette replaced from GS19 stock, subsequent death phase and system control in the o. | Assisted Replacement | SC | pCO2 | disconnected
ABL90 FLEX | Fehlersuche, div. Reinigungen durchgef., SC u. LP neu installiert, Kal. + QC i.O. | Fehlersuche; div. Reinigungen durchgef.; SC u. LP neu installiert; Kal. + QC i.O. | SWO-11643 | No | Troubleshooting; div. Cleanings through; SC and LP reinstalled; Kal. + QC in the first place | Troubleshooting; div. Cleanings through; SC and LP reinstalled; Kal. + QC in the first place | Cleaned | Fluidic pathway
ABL90 FLEX | Doorgespoeld met alkohol en demiwater. Pinchvalvetubing preventief vervangen. Laten opstarten, kalibreren en QC meten. | Doorgespoeld met alkohol en demiwater. Pinchvalvetubing preventief vervangen. Laten opstarten, kalibreren en QC meten. | SWO-11777 | No | Soaked with alcohol and demi water. Pinchvalvetubing Preventive replacement. Allow booting, calibrating and QC measurement. | Soaked with alcohol and demi water. Pinchvalvetubing Preventive replacement. Allow booting, calibrating and QC measurement. | Cleaned | Fluidic pathway | Inhomogenous sample | & preventive pinch valve tube
ABL90 FLEX | Vloeistofkanaal doorgespoten met demiwater, alkohol en demiwater. Pinchvalvetubing vervangen, beetje materiaal opbouw aan oximetriezijde. Instrument weer opgestart en laten calibreren. Alle drie controles bepaald. Een vijftal bloemonsters gemeten zonder melding. | Vloeistofkanaal doorgespoten met demiwater, alkohol en demiwater. Pinchvalvetubing vervangen, beetje materiaal opbouw aan oximetriezijde. Instrument weer opgestart en laten calibreren. Alle drie controles bepaald. Een vijftal bloemonsters gemeten zonder melding. | SWO-11770 | No | Liquid channel sprayed with demi water, alcohol and demi water. Pinchvalvetubing replaced, bit of material buildup on oximetriezijde. Instrument will be restarted and calibrating. All three controls are determined. Five floweronsters measured without notification. | Liquid channel sprayed with demi water, alcohol and demi water. Pinchvalvetubing replaced, bit of material buildup on oximetriezijde. Instrument will be restarted and calibrating. All three controls are determined. Five floweronsters measured without notification. | Defect | HW | Oxi module | Pinch Valve tube | Inhomo samples & cleaned fluidic pathway
ABL90 FLEX | Vloeistofkanaal doorgespoten met alkohol en demiwater. Pinchvalvetubing vervangen. Opgestart en gecalibreerd. QC gemeten. | Vloeistofkanaal doorgespoten met alkohol en demiwater. Pinchvalvetubing vervangen. Opgestart en gecalibreerd. QC gemeten. | SWO-11787 | No | Liquid duct sprayed with alcohol and demi water. Pinchvalvetubing replaced. Booted and calibrated. QC measured. | Liquid duct sprayed with alcohol and demi water. Pinchvalvetubing replaced. Booted and calibrated. QC measured. | Defect | HW | Oxi module | Pinch Valve tube | inhomo samples & cleaned fluidic
ABL90 | found that the analyzer had a clot in the system | back flushed the analyzer to clear clot, and so flush it out with Hypochlorite solution. After that analyzer having issue coming in on LAC, and CA replaced sc and sp to correct the issue. | CE53C8B7FC5E301D | No | found that the analyzer had a clot in the system | back flushed the analyzer to clear clot, and so flush it out with Hypochlorite solution. After that analyzer having issue coming in on LAC, and CA replaced sc and sp to correct the issue. | Cleaned | Fluidic pathway | Clot
ABL90 FLEX | System ohne SC mit Hypochlorid gespült SC mit Wasser gespült Alle Fluidic Programme ohne Beanstandung Kal und QC Level A i.O. Anwender verwendet Fremdprobennehmer (kabe)=> safeClinitubes und clot catcher dagelassen und um Verwendung gebeten, um auszuschließen das SC durch Fremdprobennehmer verschmutzt wird Durch erhöhten Verbrauch von LP und SC die vergangenen Tage 2mal LP und eine SC beim Anwender gelassen | System ohne SC mit Hypochlorid gespült SC mit Wasser gespült Alle Fluidic Programme ohne Beanstandung Kal und QC Level A i.O. Anwender verwendet Fremdprobennehmer (kabe)=> safeClinitubes und clot catcher dagelassen und um Verwendung gebeten, um auszuschließen das SC durch Fremdprobennehmer verschmutzt wird Durch erhöhten Verbrauch von LP und SC die vergangenen Tage 2mal LP und eine SC beim Anwender gelassen | SWO-11766 | No | System without SC rinsed SC with water rinsed All Fluidic programs without complaint Kal and QC Level A i.O. User uses third-party sampler (kabe)=> safeClinitubes and clot catcher left there and asked for use to exclude the SC by Third-party sampler is polluted By increased consumption of LP and SC the past days 2 times LP and a SC left at the user | System without SC rinsed SC with water rinsed All Fluidic programs without complaint Kal and QC Level A i.O. User uses third-party sampler (kabe)=> safeClinitubes and clot catcher left there and asked for use to exclude the SC by Third-party sampler is polluted By increased consumption of LP and SC the past days 2 times LP and a SC left at the user | Cleaned | Fluidic pathway
ABL90 | SUSTITUCIÓN DE SP Y SC POR CADUCIDAD FORMACIÓN DE PÈRSONAL DE SUSTITUCIÓN DE REACTIVOS | SUSTITUCIÓN DE SP Y SC POR CADUCIDAD FORMACIÓN DE PÈRSONAL DE SUSTITUCIÓN DE REACTIVOS | CE5350CBFC527E45 | No | REPLACING SP AND SC FOR REACTIVE SUBSTITUTE PERSONAL TRAINING | REPLACING SP AND SC FOR REACTIVE SUBSTITUTE PERSONAL TRAINING | User Error | incorrect SP installation | User training | SP out of test?
ABL825 FLEX | checked flow of solutions, checked data logs for errors, checked valves and liquid sensors, | removed inlet system and cleaned inlet probe, replaced tubing as needed, replaced inlet gasket, replaced pH seal ring | CE53B52DFC14C210 | No | checked flow of solutions, checked data logs for errors, checked valves and liquid sensors, | removed inlet system and cleaned inlet probe, replaced tubing as needed, replaced inlet gasket, replaced pH seal ring | Assisted Replacement | HW | Inlet Module | Tubing, blue
ABL825 FLEX | Austausch Einlass- und El Met-Modul aus GS Bestand Kundenelektroden umgesetzt Cl und Ref neu membraniert Neues QC Material Level 4 eingesetzt (Lot: 396) Proteinentfernung durchgeführt Mehrere Kal2 bis alles i.O. Kal1 Lösung und Gas1 getauscht, da kurz vor Abverbauch Aufgrund vom hohen QC Material und Cl Membranverbrauch jeweils 2mal beim Anwender gelassen | Austausch Einlass- und El Met-Modul aus GS Bestand Kundenelektroden umgesetzt Cl und Ref neu membraniert Neues QC Material Level 4 eingesetzt (Lot: 396) Proteinentfernung durchgeführt Mehrere Kal2 bis alles i.O. Kal1 Lösung und Gas1 getauscht, da kurz vor Abverbauch Aufgrund vom hohen QC Material und Cl Membranverbrauch jeweils 2mal beim Anwender gelassen | SWO-11764 | No | Exchange inlet and El Met module from GS stock customer electrodes implemented Cl and Ref newly membraned New QC Material Level 4 used (Lot: 396) Protein removal performed Several Kal2 to all I.O. Kal1 solution and gas1 exchanged, since shortly before shutting due to the high QC material and Cl membrane consumption each left 2 times by the user | Exchange inlet and El Met module from GS stock customer electrodes implemented Cl and Ref newly membraned New QC Material Level 4 used (Lot: 396) Protein removal performed Several Kal2 to all I.O. Kal1 solution and gas1 exchanged, since shortly before shutting due to the high QC material and Cl membrane consumption each left 2 times by the user | Unknown | -
ABL835 FLEX |  |  | CE547FD772171978 | No |  |  | Cleaned | Fluidic pathway
ABL825 FLEX | PH Elektrode ersetzt, Kal und QC durchgeführt | PH Elektrode ersetzt, Kal und QC durchgeführt | SWO-11768 | No | PH electrode replaced, Kal and QC performed | PH electrode replaced, Kal and QC performed | Assisted Replacement | Electrode | Electrode | pH
ABL825 FLEX | Review option keys , the option for FHBF is not available. The parameter was turned on and available in the parameter bar. | Remove the parameter from the active parameters, perform maintenance that was due on analyzer. Re-calibrate the Oximetery system with THB calibration. | CE532631FC564E64 | No | Review option keys , the option for FHBF is not available. The parameter was turned on and available in the parameter bar. | Remove the parameter from the active parameters, perform maintenance that was due on analyzer. Re-calibrate the Oximetery system with THB calibration. | User Error | Sw | SW key exp
ABL835 FLEX | Fehlermeldung tritt erneut auf, trotz Austausch Oxi ModulPCB, Fehlersuche Wechsel komplettes Oximetrie Modul (außer Hemolyser Unit) Oximetrie Test, Oxi justage, tHb Kal Einlaufverfahren | Fehlermeldung tritt erneut auf, trotz Austausch Oxi ModulPCB; Fehlersuche Wechsel komplettes Oximetrie Modul (außer Hemolyser Unit) Oximetrie Test; Oxi justage, tHb Kal Einlaufverfahren | SWO-11752 | No | Error message reappears, despite replacing Oxi modulePCB; Troubleshooting Change Complete Oximetry Module (except Hemolyser Unit) Oximetry Test; Oxi adjustment, tHb Kal inlet method | Error message reappears, despite replacing Oxi modulePCB; Troubleshooting Change Complete Oximetry Module (except Hemolyser Unit) Oximetry Test; Oxi adjustment, tHb Kal inlet method | Defect | HW | Oxi module
ABL825 FLEX |  |  | CE5965324A0D97FA | No |  |  | Assisted Replacement | Electrode | Electrode | Reference
ABL820 FLEX | log shows 3 instances of 643 dms not connected.Checked all voltages and fuses on analyzer control board and all are correct.Tried installing a new Analyzer control board but it was defective from stock.Unable to replicate the problem. | Installed a new Friction Detector in the Inlet module and performed an inlet calibration.Performed two point calibrations. | CE535347FC1C5C50 | No | log shows 3 instances of 643 dms not connected. Checked all voltages and fuses on analyzer control board and all are correct. Tried installing a new Analyzer control board but it was defective from stock. Unable to replicate the problem. | Installed a new Friction Detector in the Inlet module and performed an inlet calibration. Performed two point calibrations. | Defect | HW | Inlet Module | Friction detector unit
ABL80 FLEX CO-OX | LP und SC ersetzt |  | SWO-11884 | No | LP and SC replaced
ABL80 FLEX CO-OX | Fehlersuche am Analysator, Anwender hat bereits den LP SC erneuert da pO2 bei der Kal. IQC Probleme macht(SC Initialisierung ist fehlgeschlagen->vor Ort Service), SC nochmals neu initialisiert-> Initialisierungsphase bestanden System-und Funktionskontrolle i.O. | Fehlersuche am Analysator, Anwender hat bereits den LP SC erneuert da pO2 bei der Kal. IQC Probleme macht(SC Initialisierung ist fehlgeschlagen->vor Ort Service), SC nochmals neu initialisiert-> Initialisierungsphase bestanden System-und Funktionskontrolle i.O. | SWO-11970 | No | Troubleshooting on the analyzer, user has already renewed the LP SC da pO2 at the Kal. IQC problems (SC initialization failed->on-site service), SC re-initialized-> initialization phase passed system and function control i.O. | Troubleshooting on the analyzer, user has already renewed the LP SC da pO2 at the Kal. IQC problems (SC initialization failed->on-site service), SC re-initialized-> initialization phase passed system and function control i.O.
ABL80 FLEX CO-OX | Reinigung - Erneuerung Waste Manifold und Sensorkassette etc. - Erneuerung Lösungspack - Funktionsüberprüfung - Erneuerung Verschlauchung etc. - Medizintechnik ( Hr. Binder ) meldet sich bei weiteren Problemen ! | Reinigung - Erneuerung Waste Manifold und Sensorkassette etc. - Erneuerung Lösungspack - Funktionsüberprüfung - Erneuerung Verschlauchung etc. - Medizintechnik ( Hr. Binder ) meldet sich bei weiteren Problemen ! | SWO-12000 | No | Cleaning - Renewal Waste Manifold and sensor cassette etc. - Renewal solution pack - Functional check - Renewal hose etc. - Medical technology ( Mr. Binder ) reports further problems ! | Cleaning - Renewal Waste Manifold and sensor cassette etc. - Renewal solution pack - Functional check - Renewal hose etc. - Medical technology ( Mr. Binder ) reports further problems !
ABL90 | LOOKED IN DATA LOGS AND HAD CODE 1350, CLOT SUSPECTED IN INLET. ALSO SAW THAT THE CUSTOMER HAS CHANGED THE SOLUTION PACK AND SENSOR CASSETTE TWICE AND DAN HANSON DID MULITIPLE BACK FLUSHES. ALSO DISPLAY HAS SLIGHT DAMAGE. | REPLACED DISPLAY UNIT, BACKFLUSHED UNIT 5 TIMES AND COMPLETED ANNUAL MAINTENANCE KIT AND THEN BACKFLUSHED UNIT 5 MORE TIMES. REPLACED SOLUTION PACK AND SENSOR CASSETTE FROM CUSTOMER SUPPLY. | CE54A838725C698C | No | LOOKED IN DATA LOGS AND HAD CODE 1350, CLOT SUSPECTED IN INLET. ALSO SAW THAT THE CUSTOMER HAS CHANGED THE SOLUTION PACK AND SENSOR CASSETTE TWICE AND DAN HANSON DID MULITIPLE BACK FLUSHES. ALSO DISPLAY HAS SLIGHT DAMAGE. | REPLACED DISPLAY UNIT, BACKFLUSHED UNIT 5 TIMES AND COMPLETED ANNUAL MAINTENANCE KIT AND THEN BACKFLUSHED UNIT 5 MORE TIMES. REPLACED SOLUTION PACK AND SENSOR CASSETTE FROM CUSTOMER SUPPLY. | Cleaned | Fluidic pathway | Clot
ABL815 FLEX | Drucker ersetzt | Drucker ersetzt | SWO-12019 | No | Printer replaced | Printer replaced | Defect | HW | Printer module | Printer Head
ABL825 FLEX | Halbjahreswartung durchgeführt, Elektrode abgewaschen und membraniert |  | SWO-11983 | No | Half-year maintenance carried out, electrode washed off and membraneed |  | Assisted Replacement | Electrode | Membrane | pCO2
ABL815 FLEX | Netzteil erneuert. | Netzteil erneuert. | SWO-11906 | No | power supply is renewed. | power supply is renewed. | Defect | HW | Power supply
ABL825 FLEX | Found blood in fluidic module, samples are moving straight to oxi module. | Removed and cleaned fluidic module and gas valves. Completed LS adjust. 10 samples run - no errors. | CE549C3E720D6C6A | No | Found blood in fluidic module, samples are moving straight to oxi module. | Removed and cleaned fluidic module and gas valves. Completed LS adjust. 10 samples run - no errors. | Cleaned | HW | Fluidic manifold
ABL825 FLEX | When the autocheck unit runs qc, the carousel moves up and it spin but it does not stop at the ampoule breaker position so the carousel goes back to home position and the qc is aborted. | Inspected the autocheck unit and found the FPC 14-pole cable positioned incorrectly and it showed creases. Replaced a new 14-pole FPC cable. | CE54B489720A4F18 | No | When the autocheck unit runs qc, the carousel moves up and it spin but it does not stop at the ampoule breaker position so the carousel goes back to home position and the qc is aborted. | Inspected the autocheck unit and found the FPC 14-pole cable positioned incorrectly and it showed creases. Replaced a new 14-pole FPC cable. | Defect | HW | AutoCheck Module | FPC cable
ABL825 FLEX | Laser Scanner was replaced on !2 11 2014 for same issue. Scanners show evidence of failure. | Replaced Laser Scanner 910-310R0108 with new Laser Scanner R0126. System Operational. | CE547DB77217F836 | No | Laser Scanner was replaced on !2 11 2014 for same issue. Scanners show evidence of failure. | Replaced Laser Scanner 910-310R0108 with new Laser Scanner R0126. System Operational. | Defect | HW | AutoCheck Module | Barcode reader
ABL820 FLEX | Error 357- Barometer temp error. Barometer temp reading 52. Checked regulation, found no fault with regulation. | Re booted anaylzer. Barometer temp at 37.0 | CE548717723EC3CD | No | Error 357- Barometer temp error. Barometer temp reading 52. Checked regulation, found no fault with regulation. | Re booted anaylzer. Barometer temp at 37.0 | Unknown | HW | Analyzer Control PCB | Barometer
ABL800 BASIC | Einlasssonde erneuert. Vmtl . ist sie beim Tausch der Einlassdichtung verbogen worden. Anwenderfehler. Einweisung notwendig auch der Medizintechnik. Berechnung des Serviceeinsatzes. | Einlasssonde erneuert. Vmtl . ist sie beim Tausch der Einlassdichtung verbogen worden. Anwenderfehler. Einweisung notwendig auch der Medizintechnik. Berechnung des Serviceeinsatzes. | SWO-11975 | No | Inlet probe renewed. Vmtl . it has been bent when the inlet seal is exchanged. User error. Instruction also necessary for medical technology. Calculation of the service use. | Inlet probe renewed. Vmtl . it has been bent when the inlet seal is exchanged. User error. Instruction also necessary for medical technology. Calculation of the service use. | Defect | HW | Inlet Module | Probe bend
AQT90 FLEX | Gerät geöffnet, es liegt ein Probenröhrchen im Gerät, Nadel mehrfach verbogen & zerbrochen, Nadel ersetzt, Patientenprobenmessung durchgeführt:i.O. | Gerät geöffnet, es liegt ein Probenröhrchen im Gerät, Nadel mehrfach verbogen & zerbrochen, Nadel ersetzt, Patientenprobenmessung durchgeführt:i.O. | SWO-11947 | No | Device opened, there is a sample tube in the device, needle bent several times & broken, needle replaced, patient sample measurement performed:i.O. | Device opened, there is a sample tube in the device, needle bent several times & broken, needle replaced, patient sample measurement performed:i.O.
TCM400 | The monitor looses memory intermittently when powered down. | Replaced the main battery. If problem persists, then the electronics will need to be replaced. | CE54B4C472379B53 | No | The monitor looses memory intermittently when powered down. | Replaced the main battery. If problem persists, then the electronics will need to be replaced. | Defect | HW | Battery module | Battery module
ABL90 FLEX | Telefonat: Kundin geraten über Benutzerführung LP nochmals herauszunehmen und wieder einzulegen. ABL erkannte LP nicht und blieb hängen. Kundin ABL mit Resetknopf herunterfahren lassen und neu starten lassen. Anschließend konnte LP korrekt installiert werden. Allerdings trat pH und Elektrolyt Kal-Fehler auf (SC wurde von Kundin bereits getauscht). Kundin geraten SC nochmals zu tauschen. Vorort: bei Ankunft nur noch Lac Kaibrierungsfehler und QC Fehler. Mehrere Kalibrationen uns QC's vorgenommen. Anschließend Gerät wieder einsatzbereit. Reklamationsausdruck erstellt und Labor übergeben. |  | SWO-12016 | No | Phone call: Customer advises to remove LP again via user guidance and insert it again. ABL did not recognize LP and got stuck. Let the customer shut down ABL with the reset button and restart it. Lp could then be installed correctly. However, pH and electrolyte Kal errors occurred (SC has already been replaced by customer). Customer advised to swap SC again. Suburb: on arrival only Lac Kaibration error and QC error. Several calibrations made to us QC's. Then the device is ready for use again. Complaint expression created and handed over to laboratory. |  | User Error | SP install
ABL827 FLEX | Techline has ask me to replace the Reference electrode. | Removed Reference electrode 7107 and replaced with 1710. | CE55E6C8E816408E | No | Techline has ask me to replace the Reference electrode. | Removed Reference electrode 7107 and replaced with 1710. | Assisted Replacement | Electrode | Electrode | Reference
ABL80 FLEX CO-OX | Gerät geöffnet, Hemolyser ersetzt, Gerät kalibriert, tHb Kal durchgeführt, QC durgeführt | Gerät geöffnet, Hemolyser ersetzt, Gerät kalibriert, tHb Kal durchgeführt, QC durgeführt | SWO-12221 | No | Device opened, Hemolyser replaced, device calibrated, tHb Kal performed, QC led | Device opened, Hemolyser replaced, device calibrated, tHb Kal performed, QC led
ABL90 FLEX | Gerät durchgespült, Einlass gereinigt, tHb Kal durchgeführt, Temp. Korrigierte Werte erläutert, Konfig. Einstellung unter Diverses verändert | Gerät durchgespült, Einlass gereinigt, tHb Kal durchgeführt, Temp. Korrigierte Werte erläutert, Konfig. Einstellung unter Diverses verändert | SWO-12125 | No | Device rinsed, inlet cleaned, tHb Kal performed, Temp. Corrected values explained, Config. Setting changed under Miscellaneous | Device rinsed, inlet cleaned, tHb Kal performed, Temp. Corrected values explained, Config. Setting changed under Miscellaneous | Cleaned | Fluidic pathway
ABL90 FLEX | SC und LP dem Kunden übergeben | SC und LP dem Kunden übergeben | SWO-12745 | Yes | SC and LP handed over to the customer | SC and LP handed over to the customer | Re classified RMED | Logistical upSWO
ABL835 FLEX |  |  | CE5861FAD40DBE4C | No |  |  | Defect | HW | Oxi module | Valve
ABL835 FLEX | non homogene sur controles | Vérifier et remplacé air valve et rotor bg module replaced waste line tubing | CE582E8ED45A05C6 | No | Non Homogeneous Sur controls | Check and replace air valve and rotor bg module replaced waste line tubing | Defect | HW | pH/BG Module | Pump Rotor
ABL815 FLEX | Sicherungen AU ersetzt. Sicherungen sitzen auch nach Überarbeitung der Sicherungssitze locker. Bei festem Einsetzen der Sicherungen geht Verbindung zur AU verloren. Feinsicherungen provisorisch fixiert. AU-Platine bestellt. | Sicherungen AU ersetzt. Sicherungen sitzen auch nach Überarbeitung der Sicherungssitze locker. Bei festem Einsetzen der Sicherungen geht Verbindung zur AU verloren. Feinsicherungen provisorisch fixiert. AU-Platine bestellt. | SWO-12181 | No | Backups AU replaced. Fuses sit loose even after the safety seats have been overhauled. If the fuses are firmly inserted, connection to the AU is lost. Fine fuses fixed provisionally. AU board. | Backups AU replaced. Fuses sit loose even after the safety seats have been overhauled. If the fuses are firmly inserted, connection to the AU is lost. Fine fuses fixed provisionally. AU board. | Defect | HW | Analyzer Control PCB
ABL825 FLEX | mehrfach Membranen (Cl und Ref) getauscht Tausch von Elektroden (Cl und Ref) Tausch von Ventilen im pH BG-Modul Austausch komplettes pH BG Modul Austausch komplettes Inlet und El Met Modul Austausch pH Dichtring alles ohne Erfolg Erstellung Rückmeldung techn. Probleme | mehrfach Membranen (Cl und Ref) getauscht Tausch von Elektroden (Cl und Ref) Tausch von Ventilen im pH BG-Modul Austausch komplettes pH BG Modul Austausch komplettes Inlet und El Met Modul Austausch pH Dichtring alles ohne Erfolg Erstellung Rückmeldung techn. Probleme | SWO-12153 | No | multiple membranes (Cl and Ref) exchange of electrodes (Cl and Ref) exchange of valves in pH BG module replacement complete pH BG module exchange complete Inlet and El Met module exchange pH sealing ring all without success creation feedback techn. Problems | multiple membranes (Cl and Ref) exchange of electrodes (Cl and Ref) exchange of valves in pH BG module replacement complete pH BG module exchange complete Inlet and El Met module exchange pH sealing ring all without success creation feedback techn. Problems | Unknown | - | -
ABL825 FLEX | pCO2 Elektrode ern., Gerät kalibriert, QCs durchgef. | Der Kunde bittet um einen Service. Der Kunde wird die PCO2-Elektrode noch einmal aus der Membranhülse nehmen und die Elektrodenspitze ins Elektrolyt ein- und austauchen. Der Kunde wartet auf einen Rückruf des zuständigen Technikers. | SWO-12208 | No | pCO2 electrode ern., device calibrated, QCs through. | The customer asks for a service. The customer will take the PCO2 electrode out of the membrane sleeve once again and immerse the electrode tip into and out of the electrolyte. The customer is waiting for a callback from the technician responsible. | Assisted Replacement | Electrode | Electrode | pCO2
ABL835 FLEX | Fehlersuche Oxi- Modul PCB erneuert THb-Kalibrierung durchgeführt QC-Messungen mit AC5+Level 1 und 3 durchgeführt | Technikereinsatz am 12.01.2015 | SWO-12093 | No | Troubleshooting Oxi module PCB renewed THb calibration performed QC measurements performed with AC5+Level 1 and 3 | Technician assignment on 12.01.2015 | Defect | HW | Oxi module | PCB
ABL800 BASIC | Gerät und Einlaß überprüft, Einlaßnadel, Dichtung und Einlaßschlauch ersetzt, Einlaßmodul ausgebaut und kontrolliert, Glukose membraniert, tHb -Kal durchgeführt | Mit Herrn Koch die Einlassdichtung ersetzt. Nach Neustart zeigt der ABL sofort Einlasskalibrier-Fehler. | SWO-12162 | No | Device and inlet checked, inlet needle, seal and inlet hose replaced, inlet module removed and controlled, glucose membraneed, tHb-Kal performed | Replaced with Mr. Koch the inlet seal. After rebooting, the ABL immediately displays inlet calibration errors. | Defect | HW | Inlet Module | Probe
ABL837 FLEX | Checked hemolyzer and lamp unit, performed complete oxi test, | removed and replaced spectrometer, performed lamp unit adjustment and complete oxi test, performed tHb cal | CE58C198D45BC23C | No | Checked hemolyzer and lamp unit, performed complete oxi test, | removed and replaced spectrometer, performed lamp unit adjustment and complete oxi test, performed tHb cal | Defect | HW | Oxi module | Spectrometer
ABL825 FLEX | Halogen Intensity too high Lamp Unit neu justiert | Halogen Intensity too high Lamp Unit neu justiert | SWO-12010 | No | Halogen Intensity too high Lamp Unit re-adjusted | Halogen Intensity too high Lamp Unit re-adjusted | Adjustment | HW | Oxi module | Lamp unit
AQT90 FLEX | Photo Inrerrupter Process Unit ersetzt. alle Funktionen überprüft --ok | Photo Inrerrupter Process Unit ersetzt. alle Funktionen überprüft --ok | SWO-12052 | No | Photo Inrerrupter Process Unit. all functions checked --ok | Photo Inrerrupter Process Unit. all functions checked --ok
AQT90 FLEX | Fehlersuche Optical Sensor an Process Unit defekt Optical Sensor getauscht Systemüberprüfung Ok Sp aus Kundenbestand erneuert da verbraucht Systemreinigung durchgeführt QC Messung durchgeführt Karolin.kraus@kliniken-oal-kf.de | Kundin berichtet, dass der AQT die beschriebenen Symptome zeigt, der Kundin geraten den AQT einmal neu zu starten, sollten danach wieder Probleme angezeigt werden wird sie sich wieder melden. | SWO-12120 | No | Troubleshooting Optical Sensor to Process Unit defective Optical Sensor replaced system check Ok Sp renewed from customer stock since consumed system cleaning performed QC measurement carried out Karolin.kraus@kliniken-oal-kf.de | Kundin reports that the AQT shows the described symptoms, the customer advises to restart the AQT once, should problems be displayed again after that she will report again.
ABL90 |  |  | CE59A15D4A4E22A3 | No |  |  | Defect | HW | Waste tube | & pinch valve tube
ABL827 FLEX | Techline has ask me to replace the Reference electrode. | Removed Reference electrode 7107 and replaced it with 1712. | CE59B1164A32642D | No | Techline has ask me to replace the Reference electrode. | Removed Reference electrode 7107 and replaced it with 1712. | Assisted Replacement | Electrode | Electrode | Reference
AQT90 FLEX |  |  | SWO-12294 | No
ABL90 FLEX | PC Modul und Akku bestellt. | PC Modul und Akku bestellt. | SWO-12530 | No | PC module and battery. | PC module and battery. | Defect | HW | Battery module
ABL90 FLEX | LP und SC über Service entnommen und optische Geräteüberprüfung auf Mängel => am Quetschventilschlauch auf Hemolyzerseite starker Flüssigkeitsaustritt Oximodul ausgebaut und gereinigt Fließwege gespült Alle Serviceprogramme von Fluid Transport ohne Beanstandung LP reinstalliert Servicebox per Lieferschein übergeben tHb-Kal durchgeführt da QC Level B und C beim tHb am unteren Grenzwert | LP und SC über Service entnommen und optische Geräteüberprüfung auf Mängel => am Quetschventilschlauch auf Hemolyzerseite starker Flüssigkeitsaustritt Oximodul ausgebaut und gereinigt Fließwege gespült Alle Serviceprogramme von Fluid Transport ohne Beanstandung LP reinstalliert Servicebox per Lieferschein übergeben tHb-Kal durchgeführt da QC Level B und C beim tHb am unteren Grenzwert | SWO-12740 | No | LP and SC removed via service and optical device check for defects => on the crushing valve hose on hemolyzer side strong liquid outlet Oximodul removed and cleaned flow paths flushed All service programs of Fluid Transport without complaint LP reinstalled service box by delivery note pass tHb-Kal carried out since QC Level B and C at tHb at the lower limit | LP and SC removed via service and optical device check for defects => on the crushing valve hose on hemolyzer side strong liquid outlet Oximodul removed and cleaned flow paths flushed All service programs of Fluid Transport without complaint LP reinstalled service box by delivery note pass tHb-Kal carried out since QC Level B and C at tHb at the lower limit | Defect | HW | Oxi module | Pinch Valve tube
ABL90 FLEX | Im Service Menü Flüssigkeitstransport geprüft. War I.O. Trotzdem Fehler im LP. Danach LP getauscht Analysator wieder OK. | Im Service Menü Flüssigkeitstransport geprüft. War I.O. Trotzdem Fehler im LP. Danach LP getauscht Analysator wieder OK. | SWO-12479 | No | Checked in the Liquid Transport service menu. Was I.O. Nevertheless, errors in the LP. After that LP swaps analyzer again OK. | Checked in the Liquid Transport service menu. Was I.O. Nevertheless, errors in the LP. After that LP swaps analyzer again OK. | Assisted Replacement | SP | FT
ABL825 FLEX | Fehlersuche am Analysator, pCO2 Elektrode wurde bereits mehrfach durch den Anwender neu membraniert, Messkammern gereinigt( Dummy Elektrode über pCO2 verschmutzt) -> keine Besserung nach neuer Kal. (Empf.14%)->pCO2 Elektrode LOT LB-02 defekt->erneuert gegen LOT RK-04 System- und Funktionskontrolle i.O. | Fehlersuche am Analysator, pCO2 Elektrode wurde bereits mehrfach durch den Anwender neu membraniert, Messkammern gereinigt( Dummy Elektrode über pCO2 verschmutzt) -> keine Besserung nach neuer Kal. (Empf.14%)->pCO2 Elektrode LOT LB-02 defekt->erneuert gegen LOT RK-04 System- und Funktionskontrolle i.O. | SWO-12624 | No | Troubleshooting on the analyzer, pCO2 electrode has already been re-membraneed several times by the user, measuring chambers cleaned (dummy electrode polluted via pCO2) -> no improvement after new cal. (Empf.14%)->pCO2 electrode LOT LB-02 defective->renewed against LOT RK-04 system and function control in the entire | Troubleshooting on the analyzer, pCO2 electrode has already been re-membraneed several times by the user, measuring chambers cleaned (dummy electrode polluted via pCO2) -> no improvement after new cal. (Empf.14%)->pCO2 electrode LOT LB-02 defective->renewed against LOT RK-04 system and function control in the entire | Assisted Replacement | Electrode | Membrane | pCO2
ABL800 BASIC | Proteinentfernung und LS Justage durchgeführt kompletten Oximetrietest durchgeführt, Stdev Value = 1,4 mV Oximodul Platine erneuert tHb Kal durchgeführt Gerät kalibriert und überprüft | Proteinentfernung und LS Justage durchgeführt kompletten Oximetrietest durchgeführt, Stdev Value = 1,4 mV Oximodul Platine erneuert tHb Kal durchgeführt Gerät kalibriert und überprüft | SWO-12565 | No | Protein removal and LS adjustment performed complete oximetry test performed, Stdev Value = 1.4 mV Oximodul board renewed tHb Kal performed device calibrated and checked | Protein removal and LS adjustment performed complete oximetry test performed, Stdev Value = 1.4 mV Oximodul board renewed tHb Kal performed device calibrated and checked | Adjustment | HW | Oxi module
ABL825 FLEX | Brass connector from Actuater arm to motor came off. | Replaced defective actuater motor. | CE5AEDDAC016F906 | No | Brass connector from Actuater arm to motor came off. | Replaced defective actuater motor. | Defect | HW | AutoCheck Module | Breaker arm
ABL825 FLEX | Gerät überprüft, HJW durchgeführt, pO2 gereinigt und membraniert | Gerät überprüft, HJW durchgeführt, pO2 gereinigt und membraniert | SWO-12864 | No | Device checked, HJW performed, pO2 cleaned and membraned | Device checked, HJW performed, pO2 cleaned and membraned | Defect | SW | Error 11
ABL835 FLEX |  |  | CE5A9D5CC0511F60 | No |  |  | Defect | HW | Printer module | Printer Head
ABL825 FLEX | Data logs show 509 ph bg module did not recieve rinse from above | Replaced ph bg and el met rinse and air valves.performed five 2 point calibrations successfully with no errors | CE5AF841C00D7FFE | No | Data logs show 509 ph bg module did not recieve rinse from above | Replaced ph bg and el met rinse and air valves.performed five 2 point calibrations successfully with no errors | Defect | HW | pH/BG Module | Valve
AQT90 FLEX | Bjj testen lijkt het goed te gaan. Gisteren door Erna ook een systemclean uitgevoerd zonder meldingen. Heb deur mechanisme gedemonteerd. Scherpe randje weggevijld en geheel licht gesmeerd met vet. Geheel weer geassembleerd. Uitvoerig getest en een vijftal malen een cupje laten transporteren waarbij steeds de cartridge geplaatst is, geen problemen. Datalogs en servicedump verzameld. | Bjj testen lijkt het goed te gaan. Gisteren door Erna ook een systemclean uitgevoerd zonder meldingen. Heb deur mechanisme gedemonteerd. Scherpe randje weggevijld en geheel licht gesmeerd met vet. Geheel weer geassembleerd. Uitvoerig getest en een vijftal malen een cupje laten transporteren waarbij steeds de cartridge geplaatst is, geen problemen. Datalogs en servicedump verzameld. | SWO-12511 | No | BJJ testing seems to be going well. Yesterday by Erna also performed a system clean without notifications. Have disassembled door mechanism. Sharp edge and completely lightly lubricated with grease. Assembled entirely again. Extensively tested and have five times a chamber that always has the cartridge installed, no problems. Data logs and service dump collected. | BJJ testing seems to be going well. Yesterday by Erna also performed a system clean without notifications. Have disassembled door mechanism. Sharp edge and completely lightly lubricated with grease. Assembled entirely again. Extensively tested and have five times a chamber that always has the cartridge installed, no problems. Data logs and service dump collected.
ABL80 FLEX CO-OX | Gerät gereinigt und überprüft, Schlauch und Fließwege überprüft, Elektrodenkanäle genullt, Sensorkassette neu installiert, Lösungspack installiert, neue Sensorkassette installiert, Kal und QCs durchgeführt. |  | SWO-12877 | No | Device cleaned and checked, hose and flow paths checked, electrode channels nullified, sensor cartridge reinstalled, solution pack installed, new sensor cartridge installed, Kal and QCs performed.
ABL80 FLEX Basic version | Comprobaciónd de las acciones realizadas por parte del cliente. Borrado y reinicio del SP, calibración del mismo y pasar el QC correspondiente. | Reinicio del SP y comprobación del equipo con QC | CE5BDE9836372FE5 | No | Check the actions performed by the customer. Deleting and restarting the SP, calibrating it and passing the corresponding QC. | Rebooting the SP and checking the equipment with QC
ABL80 FLEX Basic version |  |  | CE613A320E16C0C6 | No
ABL90 |  |  | CE64E0C07057E428 | No |  |  | Assisted Replacement | SC
ABL90 FLEX | System manuell gespült, Einlassnadel verbogen und Einlassdichtung verstopft, beides aus Kundenbestand getauscht, Kaltstartphase, Kal und QC,s IO | System manuell gespült, Einlassnadel verbogen und Einlassdichtung verstopft, beides aus Kundenbestand getauscht, Kaltstartphase, Kal und QC,s IO | SWO-12708 | No | System manually rinsed, inlet needle bent and inlet seal clogged, both exchanged from customer stock, cold start phase, Kal and QC,s IO | System manually rinsed, inlet needle bent and inlet seal clogged, both exchanged from customer stock, cold start phase, Kal and QC,s IO | Replaced | probe | Bend
ABL90 | inhomgeous qc1 and solution 3 | i repalced the sensor cassette | CE5C30D7362709BF | No | inhomgeous qc1 and solution 3 | i repalced the sensor cassette | Assisted Replacement | SC | Wetability
ABL835 FLEX | Limpieza de salida de desperdicios. Limpieza de rodillos de la bomba de desechos y cambio de los tubos de bomba desechos. Sustitucion de la puerta de entrada. Cal, qc y muestras. |  | CE5C54D13613A293 | No | Cleaning of waste outlet. Cleaning of waste pump rollers and change of debris pump tubes. Replacement of the front door. Cal, qc and samples. |  | Assisted Replacement | HW | Inlet Module | Inlet Gasket
ABL800 BASIC | Neuen Demo-Key eingelesen QC-Messung mit AC5+ Level 1 durchgeführt | Neuen Demo-Key eingelesen QC-Messung mit AC5+ Level 1 durchgeführt | SWO-12773 | No | New demo key read in QC measurement performed with AC5+ Level 1 | New demo key read in QC measurement performed with AC5+ Level 1 | Setup | SW | Demo Key Exp
ABL805 FLEX | 1-14-15 Phone Earl said there was pink fluid in the REF membrane. I asked him to rinse the REF electrode off with water throughly and wipe it off. 1-15-15 On site Chris said the Cl is still running a little low. I cleaned the pH chamber with protein remover and water. I membraned the Cl electrode. PCO2 sen. would not come up so I membraned it. Also did pH seal ring upgrade. See SWO B9GV2JING for upgrade. | Cleaned pH BG chamber and membrane Cl & pCO2 with customer stock. | CE5BF2FB3607CC36 | No | 1-14-15 Phone Earl said there was pink fluid in the REF membrane. I asked him to rinse the REF electrode off with water throughly and wipe it off. 1-15-15 On site Chris said the Cl is still running a little low. I cleaned the pH chamber with protein remover and water. I membraned the Cl electrode. PCO2 sen. would not come up so I membraned it. Also did pH seal ring upgrade. See SWO B9GV2JING for upgrade. | Cleaned pH BG chamber and membrane Cl & pCO2 with customer stock. | Assisted Replacement | Electrode | Membrane | Reference
ABL815 FLEX | AU-Platine erneuert. AU-Software und Keys eingespielt. Probenzähler überprüft. | AU-Platine erneuert. AU-Software und Keys eingespielt. Probenzähler überprüft. | SWO-12720 | No | AU board renewed. AU software and keys. sample counter checked. | AU board renewed. AU software and keys. sample counter checked. | Defect | HW | Analyzer Control PCB
ABL825 FLEX | PCO2 on their peer 2 report was low. | The operator had replaced the membrane. Cleaned all electrical connections. | CE5BFFCE36151A2C | No | PCO2 on their peer 2 report was low. | The operator had replaced the membrane. Cleaned all electrical connections. | Assisted Replacement | Electrode | Membrane | pCO2
ABL820 FLEX | Customer thinks Cal. Verification mode is choosing wrong Format. Customer indicating outage is due to Radiometer personnel not creating correct Cal Ver setup. | Convinced Customer that a previous Cal Ver had been ran which started clock cycle. Unit now in Blackout as indicated by Techline to customer. | CE5BFC91361FFDBD | No | Customer thinks Cal. Verification mode is choosing wrong Format. Customer indicating outage is due to Radiometer personnel not creating correct Cal Ver setup. | Convinced Customer that a previous Cal Ver had been ran which started clock cycle. Unit now in Blackout as indicated by Techline to customer. | Defect | SW | VRK
ABL837 FLEX | All three sensitivities are out of range. | Remembraned electrodes, cleaned contacts, cleaned measuring chamber. Took multiple 2pt cals for sensitivities to come in. | CE5C0C0A3646CC46 | No | All three sensitivities are out of range. | Remembraned electrodes, cleaned contacts, cleaned measuring chamber. Took multiple 2pt cals for sensitivities to come in. | Assisted Replacement | Electrode | Membrane | Crea
ABL827 FLEX | Inspection of data logs confirmed. | Restore analyzer to Radiometer default settings and reenter customer setup manually. | CE69CC0348471406 | No | Inspection of data logs confirmed. | Restore analyzer to Radiometer default settings and reenter customer setup manually. | Setup | QC Schedule
ABL827 FLEX | Inspection of QC data logs confirmed analyzer running QCs outside fo schedule. | Restore analyzer to Radiometer default settings and reenter customer setup manually. | CE69DC28485C1409 | No | Inspection of QC data logs confirmed analyzer running QCs outside fo schedule. | Restore analyzer to Radiometer default settings and reenter customer setup manually. | Setup | QC Schedule
ABL817 FLEX | verified neon lamp position as too low | replaced oxi module pcb | CE5C3AE43604B442 | No | verified neon lamp position as too low | Replaced Oxygen module PCB | Defect | HW | Oxi module | PCB
ABL825 FLEX | RAN INLET CALIBRATION IN SERVICE AND FAILED DUE TO UNKNOWN POSITION OF PROBE BY SENSOR.DATA LOGS SHOW 676 INLET SENSITIVITY ERROR. | REPLACED FRICTION DETECTOR SENSOR ASSEMBLY.RAN MULTIPLE INLET CALIBRATIONS SUCCESSFULLY.REPLACED INLET GASKET AND INLET PROBE TUBING.PERFORMED LIQUID SENSOR ADJUSTMENT.PERFORMED 2 POINT CALIBRATIONS. | CE5C42A5364CB55D | No | RAN INLET CALIBRATION IN SERVICE AND FAILED DUE TO UNKNOWN POSITION OF PROBE BY SENSOR. DATA LOGS SHOW 676 INLET SENSITIVITY ERROR. | REPLACED FRICTION DETECTOR SENSOR ASSEMBLY. RAN MULTIPLE INLET CALIBRATIONS SUCCESSFULLY. REPLACED INLET GASKET AND INLET PROBE TUBING. PERFORMED LIQUID SENSOR ADJUSTMENT. PERFORMED 2 POINT CALIBRATIONS. | Defect | HW | Inlet Module | Friction detector unit
AQT90 FLEX | Nach öffnen des Analysators gesehen das Storage doof das Storage Wheel blockiert weil Storage door aus Halterung gesprungen. Storagedoor wieder eingehängt und Safe Positions angefahren --> i.O. Systemüberprüfung, mehrere Systemreinigungen und QC Multicheck Level 2 durchgeführt --> i.O. AQT wieder messbereit | Nach öffnen des Analysators gesehen das Storage doof das Storage Wheel blockiert weil Storage door aus Halterung gesprungen. Storagedoor wieder eingehängt und Safe Positions angefahren --> i.O. Systemüberprüfung, mehrere Systemreinigungen und QC Multicheck Level 2 durchgeführt --> i.O. AQT wieder messbereit | SWO-12706 | No | After opening the analyzer seen the storage doof blocked the storage wheel because storage door jumped out of bracket. Storagedoor remounted and Safe Positions started --> i.O. System check, multiple system cleanings and QC Multicheck Level 2 performed --> i.O. AQT ready for measurement again | After opening the analyzer seen the storage doof blocked the storage wheel because storage door jumped out of bracket. Storagedoor remounted and Safe Positions started --> i.O. System check, multiple system cleanings and QC Multicheck Level 2 performed --> i.O. AQT ready for measurement again
AQT90 FLEX | Einstellungen am Gerät überprüft - Software erneut aufgespielt - Gerät macht den Fehler weiterhin, CF Karte erneuert, Geräte Software ausfgespielt, Kundendaten eingespielt, QC durchgef. |  | SWO-12753 | No | Settings checked on the device - software replayed - device continues to make the error, CF card renewed, device software played out, customer data played in, QC through.
ABL80 FLEX Basic version |  |  | CE5D349EAC0F55B4 | No
ABL90 |  |  | CE5D25F9AC352BDD | No |  |  | Defect | HW | touch screen
ABL90 | found that analyzer states there is a leak in the system | check out system nothing is leaking. Replaced the sp did not correct the issue. Replaced the SC did correct the issue. | CE5DB4F1AC071C97 | No | found that analyzer states there is a leak in the system | check out system nothing is leaking. Replaced the sp did not correct the issue. Replaced the SC did correct the issue. | Assisted Replacement | SC | wetability
ABL90 |  |  | CE60502698094400 | No |  |  | Unknown
ABL825 FLEX | I verified that the printer was not working well and it is not printing all characters. | I replaced the printer and it resolved the problem | CE5D805AAC388359 | Yes | I verified that the printer was not working well and it is not printing all characters. | I replaced the printer and it resolved the problem | Re classified RMED | SenDx
ABL825 FLEX | Fehlersuche Mechanisch defekt, Sicherungsring hatte sich gelöst und dadurch ist der Steckbolzen rausgefallen Sicherungsring montiert Funktionstest Analysator kalibriert und überprüft QC Messung durchgeführt | Fehlersuche Mechanisch defekt, Sicherungsring hatte sich gelöst und dadurch ist der Steckbolzen rausgefallen Sicherungsring montiert Funktionstest Analysator kalibriert und überprüft QC Messung durchgeführt | SWO-12970 | No | Troubleshooting Mechanically defective, fuse ring had loosened and as a result the plug-in bolt failed fuse ring mounted function test analyzer calibrated and checked QC measurement performed | Troubleshooting Mechanically defective, fuse ring had loosened and as a result the plug-in bolt failed fuse ring mounted function test analyzer calibrated and checked QC measurement performed | Defect | HW | AutoCheck Module | Breaker arm
ABL805 FLEX | Found the analyzer leaking from the inlet module syringe port. | Flushed the inlet waste lines but, found back pressure on one of the lines. Found a piece of cotton inside one of the waste connectors. Replaced the wastye connectors. | CE5D313AAC1A8729 | No | Found the analyzer leaking from the inlet module syringe port. | Flushed the inlet waste lines but, found back pressure on one of the lines. Found a piece of cotton inside one of the waste connectors. Replaced the wastye connectors. | Defect | HW | Inlet Module | Leaking | Nipple Kit
ABL825 FLEX | Tried protein removal, inlet gasket, liquid sensor calibration and waste pump tubes-No change | Flushed entire system for clots,flushed air vent lines, replaced listed parts and inlet gasket Installed new pH 0-ring Training with customer | CE5D2EF2AC2C146E | No | Tried protein removal, inlet gasket, liquid sensor calibration and waste pump tubes-No change | Flushed entire system for clots,flushed air vent lines, replaced listed parts and inlet gasket Installed new pH 0-ring Training with customer | Defect | HW | Inlet Module | Leaking | Nipple Kit
ABL835 FLEX | Bottle Illumination PCB ausgebaut, Sichtfenster und Lichtschranken Übertragung am ABL Gehäuse lose, dadurch keine Erkennung. Loses Fenster befestigt und fixiert. PCB nicht erneuert. Testerkennung der Lösungsflaschen durchgeführt, ohne Beanstandung. Kalibration und QC in Lac nicht i.O. Es konnte keine Membranierung durchgeführt werden, kein Material vor Ort. Kunden informiert. Anbindung mit Herrn Germer besprochen. Dosen im KRS noch nicht gepatcht. EDV gibt Dosen frei und meldet sich. | Bottle Illumination PCB ausgebaut; Sichtfenster und Lichtschranken Übertragung am ABL Gehäuse lose, dadurch keine Erkennung. Loses Fenster befestigt und fixiert. PCB nicht erneuert. Testerkennung der Lösungsflaschen durchgeführt, ohne Beanstandung. Kalibration und QC in Lac nicht i.O. Es konnte keine Membranierung durchgeführt werden, kein Material vor Ort. Kunden informiert. Anbindung mit Herrn Germer besprochen. Dosen im KRS noch nicht gepatcht. EDV gibt Dosen frei und meldet sich. | SWO-13049 | No | Bottle Illumination PCB expanded; Visual windows and light barriers transmission on the ABL housing loose, thus no detection. Loose window fixed and fixed. PCB not renewed. Test detection of the solution bottles, without complaint. Calibration and QC in Lac not i.O. No membraneing could be carried out, no material on site. informed customers. with Mr. Germer. doses in KRS not yet patched. EDP releases doses and reports. | Bottle Illumination PCB expanded; Visual windows and light barriers transmission on the ABL housing loose, thus no detection. Loose window fixed and fixed. PCB not renewed. Test detection of the solution bottles, without complaint. Calibration and QC in Lac not i.O. No membraneing could be carried out, no material on site. informed customers. with Mr. Germer. doses in KRS not yet patched. EDP releases doses and reports. | Defect | HW | Basic Unit | Solution Window
ABL815 FLEX | pCO2 Elektrode gereinigt, membraniert, Gerät kalibriert und QC durchgeführt. pCO2 Elektrode zur Überprüfung mitgenommen. Kundeneigene pCO2 Elektrode zur Überbrückung eingesetzt. |  | SWO-13056 | No | pCO2 electrode cleaned, membraneed, device calibrated and QC performed. pCO2 electrode for inspection. Customer-owned pCO2 electrode used for bridging. |  | Assisted Replacement | Electrode | Electrode | pCO2
AQT90 FLEX | Service Dump erstellt, Fließwege überprüft - Auswertung der Fließwege grenzwertig, alle Daten gesichert, P2, P5 und P4 ern., Gerät gereinigt, diverse Schlauchwege ern., Messungen durchgef., Service Dump erstellt und ausgewertet - Fließwege OK | Service Dump erstellt, Fließwege überprüft - Auswertung der Fließwege grenzwertig, alle Daten gesichert, P2, P5 und P4 ern., Gerät gereinigt, diverse Schlauchwege ern., Messungen durchgef., Service Dump erstellt und ausgewertet - Fließwege OK | SWO-12923 | No | Service Dump created, flow paths checked - evaluation of flow paths borderline, all data secured, P2, P5 and P4 ern., device cleaned, various hose paths ern., measurements carried out, service dump created and evaluated - flow paths OK | Service Dump created, flow paths checked - evaluation of flow paths borderline, all data secured, P2, P5 and P4 ern., device cleaned, various hose paths ern., measurements carried out, service dump created and evaluated - flow paths OK
TCM4 | Monitor goes through a tank of cal gas once a day. | Replaced Module | CE5D78CAAC0BF360 | No | Monitor goes through a tank of cal gas once a day. | Replaced Module | Defect | HW | Gas module
ABL827 FLEX | checked data logs for errors, checked flow of solutions, checked for clots, checked inlet and inlet nipples | removed and replaced cracked inlet nipple, replaced clear inlet tubing, | CE5ED3F5224480C5 | No | checked data logs for errors, checked flow of solutions, checked for clots, checked inlet and inlet nipples | removed and replaced cracked inlet nipple, replaced clear inlet tubing, | Defect | HW | Inlet Module | Leaking | Nipple Kit
ABL90 FLEX | Beim Eintreffen Gerät in Startphase Gerät wurde provisorisch ohne Spül-Kit über Abfallanschluß gespült Überprüfung Dateiarchiv Aktivitäten, Kal, QC und Patientenergebnisse => heut früh 05:27 Uhr abgebrochene Patientenmessung, danach Fehler beim Flüssigkeitstransport => wahrscheinlich Gerinnsel ins Gerät gezogen => durch provisorische Spülung vom Anwender wieder beseitigt Startphase, Kal und QC i.O. | Beim Eintreffen Gerät in Startphase Gerät wurde provisorisch ohne Spül-Kit über Abfallanschluß gespült Überprüfung Dateiarchiv Aktivitäten, Kal, QC und Patientenergebnisse => heut früh 05:27 Uhr abgebrochene Patientenmessung; danach Fehler beim Flüssigkeitstransport => wahrscheinlich Gerinnsel ins Gerät gezogen => durch provisorische Spülung vom Anwender wieder beseitigt Startphase, Kal und QC i.O. | SWO-13048 | No | Upon arrival device in start-up device was provisionally flushed without flushing kit via waste connection check file archive activities, Kal, QC and patient results => today early 05:27 oorted patient measurement; thereafter error during liquid transport => probably clots pulled into the device => by provisional flushing by the user again eliminated start phase, Kal and QC i.O. | Upon arrival device in start-up device was provisionally flushed without flushing kit via waste connection check file archive activities, Kal, QC and patient results => today early 05:27 oorted patient measurement; thereafter error during liquid transport => probably clots pulled into the device => by provisional flushing by the user again eliminated start phase, Kal and QC i.O. | No fault found | User solved by backflush | No flush device
ABL825 FLEX | checked fuses and power supply loaded au software | unable to correct- parts needed - au pcb | CE6100630E1A7890 | No | checked fuses and power supply loaded au software | unable to correct- parts needed - au pcb | Unknown | HW | Analyzer Control PCB
ABL90 FLEX | Es konnte kein Geräteprobleme ausgemacht werden, dass die niedrigen PH Messergebnisse erklären könnte. Vielleicht ist der Fehler durch die Abnahme entstanden. Bei dem Säugling wurde das Blut an der Verse entnommen . | Es konnte kein Geräteprobleme ausgemacht werden, dass die niedrigen PH Messergebnisse erklären könnte. Vielleicht ist der Fehler durch die Abnahme entstanden. Bei dem Säugling wurde das Blut an der Verse entnommen . | SWO-13434 | No | No device problems could be identified that could explain the low PH measurement results. Perhaps the error was caused by the decrease. In the infant, the blood was taken from the verse. | No device problems could be identified that could explain the low PH measurement results. Perhaps the error was caused by the decrease. In the infant, the blood was taken from the verse. | No fault found | Pre analytical
ABL90 FLEX | Waste connector gereinigt, Kaltstartphase, LP getauscht, Systemüberprüfung I.O | Waste connector gereinigt, Kaltstartphase, LP getauscht, Systemüberprüfung I.O | SWO-13138 | No | Waste connector cleaned, cold start phase, LP exchanged, system check I.O | Waste connector cleaned, cold start phase, LP exchanged, system check I.O | Assisted Replacement | SP | Leak current
ABL90 FLEX | Fließwege gespült, Einl.nadel getauscht, Schlauch zw. Hemolysator und SC getauscht, Abf.wege gespült, Einl.dichtung getauscht, Inl.gasket getauscht, alle Programme im Service durchgeführt, SC aus Kundenbestand getauscht, Gerät kalibriert, | Fließwege gespült, Einl.nadel getauscht, Schlauch zw. Hemolysator und SC getauscht, Abf.wege gespült, Einl.dichtung getauscht, Inl.gasket getauscht, alle Programme im Service durchgeführt, SC aus Kundenbestand getauscht, Gerät kalibriert, | SWO-13262 | No | Flow paths rinsed, single needle exchanged, hose between Hemolysator and SC exchanged, drains rinsed, inlet gasket exchanged, Inl.gasket exchanged, all programs carried out in service, SC exchanged from customer stock, device calibrated, | Flow paths rinsed, single needle exchanged, hose between Hemolysator and SC exchanged, drains rinsed, inlet gasket exchanged, Inl.gasket exchanged, all programs carried out in service, SC exchanged from customer stock, device calibrated, | Replaced | Probe | Bend
ABL90 | inhomegeneous solutions unable to pump solutions | flushed and changed solutoin pack | CE613A450E5834A7 | No | inhomegeneous solutions unable to pump solutions | flushed and changed solutoin pack | Cleaned | Fluidic pathway | - | Wetability
ABL90 FLEX | LP ist am Wochenende leer gelaufen und konnte mangels Material nicht ausgetauscht werden. Neuen LP installiert, Startverfahren => SC ist ausgetrocknet und läßt sich nicht mehr richtig reaktivieren. SC ausgetauscht, Startverfahren, Kalibrationen Je eine SC und LP als Vorrat vor Ort gelassen Einweisungstermin wird mit dem Vertrieb abgeklärt. | LP ist am Wochenende leer gelaufen und konnte mangels Material nicht ausgetauscht werden. Neuen LP installiert, Startverfahren => SC ist ausgetrocknet und läßt sich nicht mehr richtig reaktivieren. SC ausgetauscht, Startverfahren, Kalibrationen Je eine SC und LP als Vorrat vor Ort gelassen Einweisungstermin wird mit dem Vertrieb abgeklärt. | SWO-13228 | Yes | LP ran empty at the weekend and could not be replaced due to lack of material. New LP installed, startup procedure => SC has dried out and can no longer be reactivated properly. SC exchanged, start procedures, calibrations Each one SC and LP left as stock on site Instruction date is clarified with the sales. | LP ran empty at the weekend and could not be replaced due to lack of material. New LP installed, startup procedure => SC has dried out and can no longer be reactivated properly. SC exchanged, start procedures, calibrations Each one SC and LP left as stock on site Instruction date is clarified with the sales. | Re classified RMED | Logistical upSWO
ABL815 FLEX | Gerät abgeschaltet und in die GS transportiert. | Gerät abgeschaltet und in die GS transportiert. | SWO-13400 | No | device and transported to the GS. | device and transported to the GS. | Unknown
ABL825 FLEX | Lactat Elektrode ausgetauscht. | Lactat Elektrode ausgetauscht. | SWO-13502 | No | Lactate electrode. | Lactate electrode. | Assisted Replacement | Electrode | Electrode | Lac
ABL825 FLEX | Analyzer not progressing through AU software install. Fellow FSE has tried sanding the fuses and changing the power supply. | Replaced the Analyzer Control board and re-installed AU software and analyzer keys. | CE6129C50E14A0D0 | No | Analyzer not progressing through AU software install. Fellow FSE has tried sanding the fuses and changing the power supply. | Replaced the Analyzer Control board and re-installed AU software and analyzer keys. | Defect | HW | Analyzer Control PCB
ABL800 BASIC | Proteinentfernung durchgeführt Oximetriemodul sowie Pumpenschläuche erneuert Anwender- sowie Service-tHb-Kalibrierung durchgeführt Gerät kalibriert und überprüft | Proteinentfernung durchgeführt Oximetriemodul sowie Pumpenschläuche erneuert Anwender- sowie Service-tHb-Kalibrierung durchgeführt Gerät kalibriert und überprüft | SWO-13093 | No | Protein removal performed oximetry module as well as pump hoses renewed user and service tHb calibration performed device calibrated and checked | Protein removal performed oximetry module as well as pump hoses renewed user and service tHb calibration performed device calibrated and checked | Defect | HW | Oxi module | PCB
ABL815 FLEX | Proteinentfernung durchgeführt, Messkammer gereinigt Glucose Elektrode getauscht Gerät kalibriert und überprüft | Proteinentfernung durchgeführt, Messkammer gereinigt Glucose Elektrode getauscht Gerät kalibriert und überprüft | SWO-13142 | No | Protein removal performed, measuring chamber cleaned glucose electrode exchanged device calibrated and checked | Protein removal performed, measuring chamber cleaned glucose electrode exchanged device calibrated and checked | Assisted Replacement | Electrode | Electrode | Glu
ABL837 FLEX | Found electrode is missing the sealing ring | Replaced sealing ring, left two spares.QC is now in. | CE6117190E32EA88 | No | Found electrode is missing the sealing ring | Replaced sealing ring, left two spares. QC is now in. | Assisted Replacement | Electrode | Electrode | Sealing ring(pCO2)
AQT90 FLEX | error code 1067 and 1218. Insufficient sample. | Cleaned storage, process wheels and cleaned sample pipette | CE61181D0E08624D | No | error code 1067 and 1218. Insufficient sample. | Cleaned storage, process wheels and cleaned sample pipette
AQT90 FLEX | Sensor up down für Process Einheit defekt. Keinen auf dem Fahrzeug. Neuen per Nachtexpress bestellt. Falls Lieferschwierigkeiten Abklärung mit Kollegen, wie weitere Vorgehensweise. | Gebrauchten Sensor eingebaut, da neuer Sensor zu späterem Zeitpunkt geliefert wird. Daher auch keine Aufführung im SWO. Systemüberprüfung o.K., QC o.K. Analysator ohne Beanstandung. | SWO-13261 | No | Sensor up down for process unit defective. None on the vehicle. New ordered by Night Express. If delivery difficulties clarify with colleagues, such as further procedure. | Used sensor installed as new sensor will be delivered at a later time. Therefore no performance in the SWO. System check o.K., QC o.K. analyzer without complaint.
ABL80 CO-OX | Per the data logs the po2 on patient sample on analyzer 1 was 76 and on analyzer 2 was 85 | I preformed zero check the analyzer checked out fine. The only common piece to the puzzle was the sp was the same on both sc that had issues with po2 replaced the sp for diagnostic purposes. | CE6240A6840B3388 | No | Per the data logs the po2 on patient sample on analyzer 1 was 76 and on analyzer 2 was 85 | I preformed zero check the analyzer checked out fine. The only common piece to the puzzle was the sp was the same on both sc that had issues with po2 replaced the sp for diagnostic purposes.
ABL90 | none needed | Installed using Matt Ackerman's stock. | CE6286208427AA79 | No | none needed | Installed using Matt Ackerman's stock.
ABL90 | came in found that error message leak current in sp found. | Reinstalled sp which corrected the issue on the analyzer. | CE626E3A842EA77F | No | came in found that error message leak current in sp found. | Reinstalled sp which corrected the issue on the analyzer. | Assisted Replacement | SP | Leak current
ABL90 | I came on site after working with the customer over the phone. They had already replaced multiple solution packs and Sensors. The problem remained with the PCO2 and I came on site. | I backflushed the system and replaced the consumables with ones they had on site. I ran a half dozen waste bloods and ran many calibrations and external QC's and it still would not pass external level 1 or 3. It was high on level 3 and low on level 1. I contacted tech support and they are sending replacment consumables overnight. | CE627358840508EE | No | I came on site after working with the customer over the phone. They had already replaced multiple solution packs and Sensors. The problem remained with the PCO2 and I came on site. | I backflushed the system and replaced the consumables with ones they had on site. I ran a half dozen waste bloods and ran many calibrations and external QC's and it still would not pass external level 1 or 3. It was high on level 3 and low on level 1. I contacted tech support and they are sending replacment consumables overnight. | Assisted Replacement | SC | pCO2
ABL90 | instrument says to change solution pack | i changed the sensor cassette | CE62519D8414D58A | No | instrument says to change solution pack | i changed the sensor cassette | Assisted Replacement | SC | Wetabiliy
ABL800 BASIC | Nullstromkalibrierung durchgeführt, Einlassnadel aus Kundenbestand eingebaut, System- und Funktionskontrolle, System und Funktion i.O. | Nullstromkalibrierung durchgeführt, Einlassnadel aus Kundenbestand eingebaut, System- und Funktionskontrolle, System und Funktion i.O. | SWO-13453 | No | Zero current calibration performed, inlet needle installed from customer stock, system and function control, system and function in the end | Zero current calibration performed, inlet needle installed from customer stock, system and function control, system and function in the end | Adjustment | HW | pH/BG Module | pO2 Zero
ABL835 FLEX |  |  | CE6273CA843CDB64 | No |  |  | Defect | HW | AutoCheck Module | FPC cable
ABL827 FLEX | checked data logs for errors, checked flow of solutions, checked valves and liquid sensors, checked pumps and pump tubing | removed modules and cleaned valves, replaced ph seal ring, replaced pump tubing, | CE62CC36842C80DB | No | checked data logs for errors, checked flow of solutions, checked valves and liquid sensors, checked pumps and pump tubing | removed modules and cleaned valves, replaced ph seal ring, replaced pump tubing, | Cleaned | Fluidic pathway | -
ABL825 FLEX | checked data logs for errors, checked lamp and hemolyzer, performed complete Oxi test, lamp intensity above 3900 | performed Lamp unit adjustment, performed complete oxi test, replaced pH seal ring | CE62C9A5845480D7 | No | checked data logs for errors, checked lamp and hemolyzer, performed complete Oxi test, lamp intensity above 3900 | performed Lamp unit adjustment, performed complete oxi test, replaced pH seal ring | Adjustment | HW | Oxi module | Lamp unit
ABL825 FLEX | Several inlet calibration errors occurred after maintenance was completed on analyzer and initialization was started. Inspected inlet assembly and all cable connections. Found Cable from friction detector unit to Inlet module PCB was not correctly fitted into cable connection. | Checked all cable connections on Inlet assembly. Cleaned inlet probe. Completed multiple inlet calibrations with out any errors. | CE6256F68443369B | No | Several inlet calibration errors occurred after maintenance was completed on analyzer and initialization was started. Inspected inlet assembly and all cable connections. Found Cable from friction detector unit to Inlet module PCB was not correctly fitted into cable connection. | Checked all cable connections on Inlet assembly. Cleaned inlet probe. Completed multiple inlet calibrations with out any errors. | Defect | HW | Inlet Module | Cable
AQT90 FLEX | Gebrauchten Sensor eingebaut, da neuer Sensor zu späterem Zeitpunkt geliefert wird. Daher auch keine Aufführung im SWO. Systemüberprüfung o.K., QC o.K. Analysator ohne Beanstandung. | Gebrauchten Sensor eingebaut, da neuer Sensor zu späterem Zeitpunkt geliefert wird. Daher auch keine Aufführung im SWO. Systemüberprüfung o.K., QC o.K. Analysator ohne Beanstandung. | SWO-13457 | No | Used sensor installed as new sensor will be delivered at a later time. Therefore no performance in the SWO. System check o.K., QC o.K. analyzer without complaint. | Used sensor installed as new sensor will be delivered at a later time. Therefore no performance in the SWO. System check o.K., QC o.K. analyzer without complaint.
ABL80 FLEX CO-OX | Fehlersuche->Kal. wurde wieder bestanden, allerdings trat der pO2 Kal. Fehler erneut auf(kein Endpunkt), Lösungspack wurde bereits durch den Anwender erneuert, Sensorkassette aus GS Bestand erneuert Reklamationsformular wird ausgefüllt, Kontaktaufnahme mit SIEBA bezüglich weiterer Vorgehensweise System-und Funktionskontrolle i.O. | Fehlersuche->Kal. wurde wieder bestanden, allerdings trat der pO2 Kal. Fehler erneut auf(kein Endpunkt), Lösungspack wurde bereits durch den Anwender erneuert, Sensorkassette aus GS Bestand erneuert Reklamationsformular wird ausgefüllt, Kontaktaufnahme mit SIEBA bezüglich weiterer Vorgehensweise System-und Funktionskontrolle i.O. | SWO-13640 | No | Troubleshooting->Kal. was passed again, but the pO2 Kal. Error again on (no endpoint), solution pack has already been renewed by the user, sensor cassette from GS stock renewed complaint form is completed, contact with SIEBA regarding further procedure system and function control i.O. | Troubleshooting->Kal. was passed again, but the pO2 Kal. Error again on (no endpoint), solution pack has already been renewed by the user, sensor cassette from GS stock renewed complaint form is completed, contact with SIEBA regarding further procedure system and function control i.O.
ABL90 |  |  | CE63B273FA3D4BC7 | No |  |  | Defect | HW | Wet sec control
ABL90 |  |  | CE6C4A37345C2C8F | Yes |  |  | Re classified RMED | Training
ABL90 FLEX | Konfiguration geändert, automatische Datensendung der Pat Daten nur für die HLM. Senden der Pat Daten ins Medico erfolgt jetzt manuell. Einlaßdichtung ausgetauscht. Probennadel für den Vorrat vor Ort gelassen. | Herr Plesch hat bereits das Spülverfahren durchgeführt und installiert das selbe SP noch einmal. | SWO-13600 | No | Configuration changed, automatic data transmission of Pat data only for the HLM. Sending the Pat data to The Medico is now done manually. inlet seal replaced. Sample needle for the stock left on site. | Mr Plesch has already carried out the rinsing procedure and is reinstalling the same SP. | Cleaned | Fluidic pathway
ABL90 | Printer is not working |  | CE63F5DCFA0DE1D1 | No | Printer is not working |  | Defect | HW | Printer
ABL90 | Customer replaced SP & SC, flushed system and still gets UIR screen prompting for SP change. Customer spent many hours troubleshooting and is frustrated & doesn't have any more time to spend on this issue and wants service on site. I ran fluidic transport test and nearly all fluidic transport tests are failing. | I replaced SP & SC again and flushed system, then reran all fluidic transport tests which now passed. | CE63F660FA31E30E | No | Customer replaced SP & SC, flushed system and still gets UIR screen prompting for SP change. Customer spent many hours troubleshooting and is frustrated & doesn't have any more time to spend on this issue and wants service on site. I ran fluidic transport test and nearly all fluidic transport tests are failing. | I replaced SP & SC again and flushed system, then reran all fluidic transport tests which now passed. | Assisted Replacement | SC | Wetability
ABL825 FLEX | This error came and went after the last service visit. Had customer try rebooting instrument, but this did not help. Changed travel arrangements and ordered parts. | Replaced Analyzer control board again. | CE63A48BFA43A0E5 | No | This error came and went after the last service visit. Had customer try rebooting instrument, but this did not help. Changed travel arrangements and ordered parts. | Replaced Analyzer control board again. | Defect | HW | Analyzer Control PCB
ABL825 FLEX |  |  | CE63DD8CFA2964F7 | No |  |  | Assisted Replacement | Electrode | Electrode | Reference
ABL835 FLEX | PO2 QC OUT OF PEER GROUP RANGE ON LEVEL 2 S7745 LOT 582 LOW AT 106 AND LEVEL 3 S7755 LOT 514 AT 68.5. PERFORMED GAS CHECK BOTH REGULATORS STARTING AND STOPPING PROPERLY.INSPECT PH BG CHAMBER NO CARRYOVER.ALSO CHECKED PH ELECTRODE FOR BLUE SEALING RING ALREADY INSTALLED. | REPLACED BOTH GAS TANKS,PERFORMED GAS CHECK.REPLACED NEW PO2 ELECTRODE AND MEMBRANES.REPLACED PH BG MEASURING CHAMBER,SEALING RINGS AND HEAT DISTRIBUTION PLATE.REPLACED WASTE AND AIR VALVE AIN FLUIDIC MODULE,REPLACED PH BG AND MET EL RINSE AND AIR VALVES,REPLACED DUAL HOUSING IN PH BG MODULE.REPLACED INLET PROBE AND INLET PREHEATER BLUE TUBINGS.REPLACED ALL BLACK TUBINGS ON TOP OF PH AND EL MET MODULES AND GAS TUBINGS.PERFORMED MULTIPLE CALIBRATIONS | CE63BE5AFA1FB6CF | No | PO2 QC OUT OF PEER GROUP RANGE ON LEVEL 2 S7745 LOT 582 LOW AT 106 AND LEVEL 3 S7755 LOT 514 AT 68.5. PERFORMED GAS CHECK BOTH REGULATORS STARTING AND STOPPING PROPERLY. INSPECT PH BG CHAMBER NO CARRYOVER. ALSO CHECKED PH ELECTRODE FOR BLUE SEALING RING ALREADY INSTALLED. | REPLACED BOTH GAS TANKS,PERFORMED GAS CHECK. REPLACED NEW PO2 ELECTRODE AND MEMBRANES. REPLACED PH BG MEASURING CHAMBER,SEALING RINGS AND HEAT DISTRIBUTION PLATE. REPLACED WASTE AND AIR VALVE AIN FLUIDIC MODULE,REPLACED PH BG AND MET EL RINSE AND AIR VALVES,REPLACED DUAL HOUSING IN PH BG MODULE. REPLACED INLET PROBE AND INLET PREHEATER BLUE TUBINGS. REPLACED ALL BLACK TUBINGS ON TOP OF PH AND EL MET MODULES AND GAS TUBINGS. PERFORMED MULTIPLE CALIBRATIONS | Unknown | - | - | pO2
ABL80 | Screen would not turn on and TEMP E errors on printout Confirmed Customer's findings. Found Power Supply to be faulty | Replaced Power Supply Confirmed proper seating of all cables Confirmed proper use of EMI sheilding on unit Confirmed Battery Pack Tested Diagnaustics and 0 point configuration Installed new Solution Pack Installed ne Sensor Cassette All tests passed | CE64C26A7042868E | No | Screen would not turn on and TEMP E errors on printout Confirmed Customer's findings. Found Power Supply to be faulty | Replaced Power Supply Confirmed proper seating of all cables Confirmed proper use of EMI sheilding on unit Confirmed Battery Pack Tested Diagnaustics and 0 point configuration Installed new Solution Pack Installed ne Sensor Cassette All tests passed
ABL90 | Recommendation to turn on electrolytes as if affects algorithms | Enabled electrolytes to change algorithms in Lactate calculations | CE6552B87032C68F | Yes | Recommendation to turn on electrolytes as if affects algorithms | Enabled electrolytes to change algorithms in Lactate calculations | Re classified RMED | Techline | Enable Electrolyte
ABL90 | Recommendation to enable all electrolytes to change Lactate calculation | Enabled electrolytes to change algorithms in Lactate calculation. | CE655699700D647E | Yes | Recommendation to enable all electrolytes to change Lactate calculation | Enabled electrolytes to change algorithms in Lactate calculation. | Re classified RMED | Techline | Enable Electrolyte
ABL90 FLEX | ABL überprüft. Lösungspackaktivitäten kontrolliert. Immer wieder Fehler 1134: Chipdaten können nicht gelesen werden. Neuen Lösungspack installiert. Keine Fehlermeldung während der Installation. Gerätestartphase durchgeführt. QC durchgeführt. | ABL überprüft. Lösungspackaktivitäten kontrolliert. Immer wieder Fehler 1134: Chipdaten können nicht gelesen werden. Neuen Lösungspack installiert. Keine Fehlermeldung während der Installation. Gerätestartphase durchgeführt. QC durchgeführt. | SWO-13848 | No | ABL. solution packing activities. Again and again Error 1134: Chip data cannot be read. New solution pack installed. No error message during installation. device start phase. QC. | ABL. solution packing activities. Again and again Error 1134: Chip data cannot be read. New solution pack installed. No error message during installation. device start phase. QC. | Assisted Replacement | SP | Chip error
ABL800 BASIC | Elektrodenkappe Lac ersetzt, da häufiges Membranieren keine Besserung mit unterschiedlichen Lot's brachte, Kal, QC | Kundin berichtet, dass sie wieder Probleme mit dem Lac, haben, die Empf. liegt hier bei nur 130,8 pA, Kundin geraten die Lac neu zu membranieren, danach mind 2-3 mal kalibrieren, die fehlerhafte QC manuell zu starten, sollte es weitere Probleme geben wird sich die Kundin wieder melden. | SWO-13957 | No | Electrode cap Lac replaced as frequent membraneing did not bring any improvement with different lot, Kal, QC | Customer reports that they have problems with the Lac, the Empf again. here is only 130.8 pA, customer advises to re-diaphragm the lac, then calibrate at least 2-3 times, to start the faulty QC manually, should there be further problems the customer will report again. | Defect | HW | El/Met Module | electrode connector | Lac
ABL825 FLEX | Fluidic Modul. V2+V3 werden sehr heiß Valves ersetzt | Fluidic Modul. V2+V3 werden sehr heiß Valves ersetzt | SWO-14374 | No | Fluidic module. V2+V3 will be replaced very hot Valves | Fluidic module. V2+V3 will be replaced very hot Valves | Defect | HW | Fluidic manifold | Valves
ABL835 FLEX | intermittant Erreur,735,354,259.270,355.356,vérifié voltage OK, intallé nouveau O'ring PH | Remplacé power supply | CE64B3E4704F5FE8 | No | intertermittant Error,735,354,259.270,355.356,verified voltage OK, new O'ring PH | Replaced power supply | Defect | HW | Power supply
ABL825 FLEX | Inspection of inlet module and AutoCheck module found no reason for inlet leak or AutoCheck issue.. | Replace waste tubings from inlet to tubing organizer. Replace waste pump tubings. Replace inlet nipples for gasket. Replace inlet gasket. | CE64CCB4702B0E39 | No | Inspection of inlet module and AutoCheck module found no reason for inlet leak or AutoCheck issue.. | Replace waste tubings from inlet to tubing organizer. Replace waste pump tubings. Replace inlet nipples for gasket. Replace inlet gasket. | Defect | HW | Inlet Module | Inlet Gasket
ABL835 FLEX | PCO2 WAS 13.3 ON LEVEL 3 S7755 AND PO2 WAS 70.5 | REPLACED GAS 2 REGULATOR.INSTALLED CUSTOMERS ORIGINAL PO2 ELECTRODE.PERFORMED MULTIPLE CALIBRATIONS PO2 LEVEL 3 NOW AT 12.6 ON PEER MEAN AND PCO2 LEVEL 3 NOW A T73.1 ALSO NEAR MEAN. | CE64E7C87028E2CB | No | PCO2 WAS 13.3 ON LEVEL 3 S7755 AND PO2 WAS 70.5 | REPLACED GAS 2 REGULATOR. INSTALLED CUSTOMERS ORIGINAL PO2 ELECTRODE. PERFORMED MULTIPLE CALIBRATIONS PO2 LEVEL 3 NOW AT 12.6 ON PEER MEAN AND PCO2 LEVEL 3 NOW A T73.1 ALSO NEAR MEAN. | Defect | HW | Gas module | Gas 2
ABL835 FLEX | LOOKED FOR CODE IN MANUAL, CHECKED FOR LEAKS ALONG FLUID TRANSPORT. TURNED MACHINE OFF AND LOOKED OVER BOTH MEASURING CHAMBERS, AND FLUIDIC MODULE. CHAMBERS LOOKED GOOD BUT FLUIDIC MODULE HAD #7 VALVE THAT LOOKED SUSPECIOUS. ALSO FOUND A BENT INLET PROBE. | REPLACED SOLUTION PUMP TUBING (CUSTOMERS), REPLACED #7 VALVE, REPLACED INLET GASKET AND INLET PROBE (CUSTOMERS). COULDNT GET GLUCOSE IN ON QC LEVEL 3. SO REMEMBRANED AND DID 2 PT CAL 3 TIMES. | CE64E649705CC04F | No | LOOKED FOR CODE IN MANUAL, CHECKED FOR LEAKS ALONG FLUID TRANSPORT. TURNED MACHINE OFF AND LOOKED OVER BOTH MEASURING CHAMBERS, AND FLUIDIC MODULE. CHAMBERS LOOKED GOOD BUT FLUIDIC MODULE HAD #7 VALVE THAT LOOKED SUSPECIOUS. ALSO FOUND A BENT INLET PROBE. | REPLACED SOLUTION PUMP TUBING (CUSTOMERS), REPLACED #7 VALVE, REPLACED INLET GASKET AND INLET PROBE (CUSTOMERS). COULDNT GET GLUCOSE IN ON QC LEVEL 3. SO REMEMBRANED AND DID 2 PT CAL 3 TIMES. | Defect | HW | Inlet Module | Probe bend
ABL835 FLEX | DID COMPLETED OXIMETRY CALIBRATION. ALL PASSED. DID 1 PT CAL AND THEN DID THB IN SERVICE, AND IN USER, AND DID IT AS A SAMPLE. LOOKED GOOD AND THEN QC WAS FAILING ALL PERIMETERS AND GOT CODE 0739, SOLUTIONS COMPOSITION NOT IDENTIFICABLE. | LOOKED TO SEE IF ANYTHING WAS LEAKING, TOOK PH MODULE OUT AND LOOKED AT VALVES. THEY ALL LOOKED GOOD, THEN I LOOKED AT INLET MODULE AND FOUND THE SMALL BLUE TUBING FULL OF SALT DEPOSITS. REPLACED SMALL BLUE TUBING. | CE651DE17041C054 | No | DID COMPLETED OXIMETRY CALIBRATION. ALL PASSED. DID 1 PT CAL AND THEN DID THB IN SERVICE, AND IN USER, AND DID IT AS A SAMPLE. LOOKED GOOD AND THEN QC WAS FAILING ALL PERIMETERS AND GOT CODE 0739, SOLUTIONS COMPOSITION NOT IDENTIFICABLE. | LOOKED TO SEE IF ANYTHING WAS LEAKING, TOOK PH MODULE OUT AND LOOKED AT VALVES. THEY ALL LOOKED GOOD, THEN I LOOKED AT INLET MODULE AND FOUND THE SMALL BLUE TUBING FULL OF SALT DEPOSITS. REPLACED SMALL BLUE TUBING. | Cleaned | Fluidic pathway
ABL837 FLEX | The analyzer is leaking from the inlet syringe port. | Found the left inlet waste connector broken. Replaced the connector using the inlet nipple kit. Flushed the system lines. ran a liquid sensor adjsut and a pump calibration. The calibration 1 solution and the gas 2 tank were low, replaced both items for the customer. | CE64DD687025676C | No | The analyzer is leaking from the inlet syringe port. | Found the left inlet waste connector broken. Replaced the connector using the inlet nipple kit. Flushed the system lines. ran a liquid sensor adjsut and a pump calibration. The calibration 1 solution and the gas 2 tank were low, replaced both items for the customer. | Defect | HW | Inlet Module | Leaking | Nipple Kit
TCM/Tosca |  |  | CE64C0B9702004BF | No |  |  | Defect | HW | Power Supply | Power Supply
TCM-COMBIM |  |  | CE650ECB705A04CF | No |  |  | Defect | HW | Touch Screen | Touch Screen
ABL80 FLEX CO-OX | Fehlersuche Probenweg auf Undichtheit kontrolliert Probemessungen durchgeführt ==> wieder o.k. QC-Messungen mit interner QC durchgeführt |  | SWO-13975 | No | Troubleshooting Sample path carried out on leakage controlled sample measurements ==> again o.k. QC measurements carried out with internal QC
ABL90 FLEX | 2. SP geht auch nicht. Pressuretest failt.Tubings und Kontakte ersetzt. | 2. SP geht auch nicht. Pressuretest failt.Tubings und Kontakte ersetzt. | SWO-14375 | No | 2. SP is not possible either. Pressuretest fails. Tubings and contacts replaced. | 2. SP is not possible either. Pressuretest fails. Tubings and contacts replaced. | Defect | HW | SP-PCB | LEH
ABL90 FLEX | Thermistor ersetzt. Temperatur I.O. | Thermistor ersetzt. Temperatur I.O. | SWO-13912 | No | Thermistor replaced. Temperature I.O. | Thermistor replaced. Temperature I.O. | Defect | HW | Termistor cable | LEH
ABL90 FLEX | PC Modul Eingebaut SW installiert. | PC Modul Eingebaut SW installiert. | SWO-13916 | No | PC Module Built-in SW installed. | PC Module Built-in SW installed. | Defect | HW | Battery module | LEH
ABL90 FLEX | LP und Sensorkassette getauscht, Kaltstartphase, mehrere zusätzliche Kal,s durchgeführt, QC,s vermessen | LP und Sensorkassette getauscht, Kaltstartphase, mehrere zusätzliche Kal,s durchgeführt, QC,s vermessen | SWO-13915 | No | LP and sensor cartridge exchanged, cold start phase, several additional cal,s performed, QC,s measured | LP and sensor cartridge exchanged, cold start phase, several additional cal,s performed, QC,s measured | Assisted Replacement | SC | pCO2 | LEH
ABL90 | Checked Cal and Internal QC logs for any failures - none. Status for pCO2 dropped from 54.1 (1-21-15) to 18.2 today. Sensitivity went up from 82.1% (1-21-15) to 87.1%. | Reran VK kit. Filled out all info in BAS. | CE6647A7E659F042 | No | Checked Cal and Internal QC logs for any failures - none. Status for pCO2 dropped from 54.1 (1-21-15) to 18.2 today. Sensitivity went up from 82.1% (1-21-15) to 87.1%. | Reran VK kit. Filled out all info in BAS. | Assisted Replacement | SC | pCO2
ABL815 FLEX | FlexQ Modul in der GS zerlegt, gereinigt und getestet. Kundeneigenes FlexQ Modul (R0438N002) wieder zurück getauscht. Gerät kalibriert und QC durchgeführt. | FlexQ Modul in der GS zerlegt, gereinigt und getestet. Kundeneigenes FlexQ Modul (R0438N002) wieder zurück getauscht. Gerät kalibriert und QC durchgeführt. | SWO-13979 | No | FlexQ module in GS disassembled, cleaned and tested. Customer-owned FlexQ module (R0438N002) replaced again. calibrated and QC performed. | FlexQ module in GS disassembled, cleaned and tested. Customer-owned FlexQ module (R0438N002) replaced again. calibrated and QC performed. | Unknown | HW | FlexQ Module
AQT90 FLEX |  |  | CE662E3AE626A35A | No
AQT90 FLEX | Fehlersuche Lichtschranke von optische Einheit defekt Lichtschranke getauscht Systemüberprüfung OK | Fehlersuche Lichtschranke von optische Einheit defekt Lichtschranke getauscht Systemüberprüfung OK | SWO-13855 | No | Troubleshooting light barrier of optical unit defective light barrier replaced system check OK | Troubleshooting light barrier of optical unit defective light barrier replaced system check OK
AQT90 FLEX | Überprüfung Process Unit mittels Serviceprogramm Safe Positions => scheitert bei Process Unit und Process Wheel Ansteuerung Process Unit nicht möglich => Position der Process Unit wird nicht richtig erkannt => vermutlich Gabellichtschranke defekt Austausch der Gabellichtschranke aus GS11 Geräteneustart und Startphase i.O. Safe Positions ohne Fehler durchführbar 2x Systemreinigung durchgeführt QC CRP Level 2 i.O. | Überprüfung Process Unit mittels Serviceprogramm Safe Positions => scheitert bei Process Unit und Process Wheel Ansteuerung Process Unit nicht möglich => Position der Process Unit wird nicht richtig erkannt => vermutlich Gabellichtschranke defekt Austausch der Gabellichtschranke aus GS11 Geräteneustart und Startphase i.O. Safe Positions ohne Fehler durchführbar 2x Systemreinigung durchgeführt QC CRP Level 2 i.O. | SWO-14059 | No | Review Process Unit by service program Safe Positions => fails in Process Unit and Process Wheel control Process Unit not possible => Position of the Process Unit is not recognized correctly => presumably fork light barrier defective Replacement of the Fork light barrier made of GS11 device restart and start phase in the safe positions without error feasible 2x system cleaning carried out QC CRP Level 2 i.O. | Review Process Unit by service program Safe Positions => fails in Process Unit and Process Wheel control Process Unit not possible => Position of the Process Unit is not recognized correctly => presumably fork light barrier defective Replacement of the Fork light barrier made of GS11 device restart and start phase in the safe positions without error feasible 2x system cleaning carried out QC CRP Level 2 i.O.
ABL80 CO-OX | COMPLETE SYSTEM INSPECTION FOUND FAULTY PUMP MOTOR | NONE | CE6DB6EDAA03F7D6 | No | COMPLETE SYSTEM INSPECTION FOUND FAULTY PUMP MOTOR | NONE
ABL90 | Customer had flow selector error and replaced solution pack with no resolve. Suspect is malfunctioning flow selector detector pcb. | Replaced flow selector detector pcb. Verified instrument operation. | CE6AE30CBE50CE2D | No | Customer had flow selector error and replaced solution pack with no resolve. Suspect is malfunctioning flow selector detector pcb. | Replaced flow selector detector pcb. Verified instrument operation. | Defect | HW | Flow Selector PCB
ABL90 |  |  | CE6E7E12204C296D | No |  |  | Cleaned | Fluidic pathway | Clot
ABL825 FLEX | ABL genstartet derefter ok | ABL genstartet derefter ok | SWO-14274 | No | ABL then restarted OK | ABL then restarted OK | Defect | HW | Basic Unit | Barcode reader
ABL825 FLEX | Mixertray erneuert. | Mixertray erneuert. | SWO-14237 | No | Mixertray renewed. | Mixertray renewed. | Defect | HW | FlexQ Module | Tray
ABL820 FLEX | Performed visual inspection and examined the data logs. Confirmed erratic 1 and 2 pt calibrations beginning on 1 15, drift and sens out of spec for pCO2. pCO2 QC erratic. Noted white debris on gas air valve for pH BG module, air bubbles in bottom of pCO2 electrode membrane, possible issue with pH blue sealing ring. | Replaced gas air valve for pH BG module, replaced rinse valve as precautionary measure, replaced pH blue sealing ring, replaced bubbles from pCO2 electrode membrane, cleaned electrodes and measuring chamber. | CE6C6D8534110A44 | No | Performed visual inspection and examined the data logs. Confirmed erratic 1 and 2 pt calibrations beginning on 1 15, drift and sens out of spec for pCO2. pCO2 QC erratic. Noted white debris on gas air valve for pH BG module, air bubbles in bottom of pCO2 electrode membrane, possible issue with pH blue sealing ring. | Replaced gas air valve for pH BG module, replaced rinse valve as precautionary measure, replaced pH blue sealing ring, replaced bubbles from pCO2 electrode membrane, cleaned electrodes and measuring chamber. | Defect | HW | pH/BG Module | Valve, air
ABL825 FLEX | checked logs,ran completee oxi x3, adjusted lamp unit,rebooted | unable to correct | CE6A07E1483D7389 | No | checked logs,ran completee oxi x3, adjusted lamp unit,rebooted | unable to correct | Adjustment | HW | Oxi module | Lamp unit
AQT90 FLEX | Comprobar unidad de secado, reparar sensor de temperatura de la unidad de secado. | Reparar sensor de temperatura del módulo de secado. | CE69D1F048179C9E | No | Check drying unit, repair temperature sensor of the drying unit. | Repair temperature sensor of the drying module.
ABL90 | Confirmed problem. Removed oxi module and found a leak at the valve tubing. | Cleaned up leak. Replaced the valve tubing and silicone disc. | CE6B5404BE241F2D | No | Confirmed problem. Removed oxi module and found a leak at the valve tubing. | Cleaned up leak. Replaced the valve tubing and silicone disc. | Defect | HW | Oxi module | Pinch Valve tube
ABL90 | Reviewed data logs. Problem started with SP change | Replaced solution pack, tHb calibration. Ran both internal and external QCs all pass. | CE6B3A45BE582C88 | No | Reviewed data logs. Problem started with SP change | Replaced solution pack, tHb calibration. Ran both internal and external QCs all pass. | Assisted Replacement | SP | Ca, Na
ABL90 | CALLED KIM LEE AND DAN HANSON AND WENT INTO CUSTOMER TO SEE WHAT ISSUES THE CUSTOMER HAD. ONCE I GOT THERE THE ANALYZER IN THE LIVER TRANSPLANT OPERATING ROOM HAD ISSUE WITH SOLUTION PACK. THEY HAD ALREADY CHANGED IT OUT. DID TRAINING WITH CUSTOMER ON INLET CLEANING, AND CALIBRATIONS AND QC AND ALSO HOW TO INSTALL THE INLET GASKET ADAPTER, SO PROBE WILL NOT BOTTOM OUT ON SHRINGE. ALL 3 ANALYZERS WERE ALL GREEN AND GOOD. | N A | CE6B5A30BE519965 | Yes | CALLED KIM LEE AND DAN HANSON AND WENT INTO CUSTOMER TO SEE WHAT ISSUES THE CUSTOMER HAD. ONCE I GOT THERE THE ANALYZER IN THE LIVER TRANSPLANT OPERATING ROOM HAD ISSUE WITH SOLUTION PACK. THEY HAD ALREADY CHANGED IT OUT. DID TRAINING WITH CUSTOMER ON INLET CLEANING, AND CALIBRATIONS AND QC AND ALSO HOW TO INSTALL THE INLET GASKET ADAPTER, SO PROBE WILL NOT BOTTOM OUT ON SHRINGE. ALL 3 ANALYZERS WERE ALL GREEN AND GOOD. | N A | Re classified RMED | Courtesy | Training
ABL90 |  |  | CE6E7841201A2960 | Yes |  |  | Re classified RMED | Training
ABL835 FLEX | Air dans patient et controles and inlet calibration erreur | remplac air valve inlet probe et inlet gasket du client | CE6B37BCBE476100 | No | Air in patient and controls and inlet calibration error | remplac air valve inlet probe et inlet gasket du client | Defect | HW | pH/BG Module | Valve, air
ABL835 FLEX |  |  | CE6AF265BE1D22E4 | No |  |  | Defect | HW | Inlet Module | Probe bend
ABL825 FLEX | Gerät überprüft, Nullpunktkalibration durchgeführt, pO-Elektrode (LT-03) ausgetauscht, Glukose- und Lactatelektrode abgewaschen, Meßkammer und Luftauslass gesäubert, Gasventile kontrolliert |  | SWO-14442 | No | Device checked, zero-point calibration performed, pO electrode (LT-03) replaced, glucose and lactate electrodes washed off, measuring chamber and air outlet cleaned, gas valves checked |  | Assisted Replacement | Electrode | Electrode | pO2
ABL827 FLEX | Unable to determine cause of leakage. determined unit had Decontamination ran without removing electrodes. | Replaced tubings, remembraned electrodes, ran calibrations. | CE6B5891BE18568F | No | Unable to determine cause of leakage. determined unit had Decontamination ran without removing electrodes. | Replaced tubings, remembraned electrodes, ran calibrations. | Unknown | HW | Tubings
ABL90 FLEX | SP90 verriegelt richtig. | Der Kunde bittet um einen Service. | SWO-14495 | No | SP90 locks correctly. | The customer asks for a service. | User Error | SP install
ABL90 | Battery dead | replaced battery from a loaner | CE6C49063438A69C | No | Battery dead | replaced battery from a loaner | Defect | HW | Battery module
ABL90 | Error 1254 find damage inlet gasket movement of inlet and spring not free | Replaced inlet connector and inlet assembly | CE6C61A53451A69F | No | Error 1254 find damage inlet gasket movement of inlet and spring not free | Replaced inlet connector and inlet assembly | Replaced | Inlet gasket
ABL90 |  |  | CE6D88B6AA44A5A9 | No |  |  | Cleaned | Fluidic pathway | Clot
ABL90 | Customer replaced sensor cassette and failed Q.C. | Shipped customer new sensor cassette | CE6C30FA3418856A | No | Customer replaced sensor cassette and failed Q.C. | Shipped customer new sensor cassette
ABL90 | Lactate response fault. | Replaced the sensor cassette, but the external QC failed to aspirate. Cleaned the inlet gasket and recalibrated the system. | CE6C778F343F14F2 | No | Lactate response fault. | Replaced the sensor cassette, but the external QC failed to aspirate. Cleaned the inlet gasket and recalibrated the system. | Assisted Replacement | SC | Lactate
ABL90 | Clot suspected, unable to pump solutions, Inhomogeneous cal 2-3-QC1, etc. | Replaced the sensor cassette and ran some waste blood, found the inlet spring to be faulty-replaced the inlet module and gasket. I also found a clot in the system-cleared the clot Brought system to ready after an hour start up. | CE6C7EBF344714F5 | No | Clot suspected, unable to pump solutions, Inhomogeneous cal 2-3-QC1, etc. | Replaced the sensor cassette and ran some waste blood, found the inlet spring to be faulty-replaced the inlet module and gasket. I also found a clot in the system-cleared the clot Brought system to ready after an hour start up. | Cleaned | Fluidic pathway | Clot
ABL90 |  |  | CE6E7ABA205E2962 | Yes |  |  | Re classified RMED | Installation
ABL827 FLEX | Error 351- Neon number to high and Neon position failing during oxi test. | Replaced oxi module PCB. Performed software install. Lamp adjustment. | CE6C716A3426601B | No | Error 351- Neon number to high and Neon position failing during oxi test. | Replaced oxi module PCB. Performed software install. Lamp adjustment. | Defect | HW | Oxi module | PCB
AQT90 FLEX | CHEQUEO DE TODO EL EQUIPO, LIMPIEZA DE TODAS US PARTES, CHEQUEO Y AJUSTE DE TODAS LAS PIEZAS POR SERVICIO, ENGRASE DE TODOS LOS EJES DE LOS MOTORES, LIMPIEZA DE FILTRO, CALIBRACION DE PIEZAS, CALIBRACIÓN DE DE LOTE DE MYO SALIENDO ABORTADO CON ERROR 1061 | CHEQUEO DE TODO EL EQUIPO, LIMPIEZA DE TODAS US PARTES, CHEQUEO Y AJUSTE DE TODAS LAS PIEZAS POR SERVICIO, ENGRASE DE TODOS LOS EJES DE LOS MOTORES, LIMPIEZA DE FILTRO, CALIBRACION DE PIEZAS, CALIBRACIÓN DE DE LOTE DE MYO SALIENDO ABORTADO CON ERROR 1061 | CE6E185BAA0D4BF6 | No | CHECK ALL EQUIPMENT, CLEANING ALL US PARTIES, CHECKING AND ADJUSTING ALL PARTS BY SERVICE, GRASE ALL MOTOR EXAMPLES, FILTER CLEANING, PART CALIBRATION, MYO SAL OS | CHECK ALL EQUIPMENT, CLEANING ALL US PARTIES, CHECKING AND ADJUSTING ALL PARTS BY SERVICE, GRASE ALL MOTOR EXAMPLES, FILTER CLEANING, PART CALIBRATION, MYO SAL OS
AQT90 FLEX |  |  | CE6C77253401BDB9 | No
AQT90 FLEX | Instalation capteur (vieux type). Le capteur original a le numéro du batch L09. Controle du système : OK | Instalation capteur (vieux type). Le capteur original a le numéro du batch L09. Controle du système : OK | SWO-14517 | No | Sensor instalation (old type). The original sensor has the number of batch L09. System control: OK | Sensor instalation (old type). The original sensor has the number of batch L09. System control: OK
TCM/Tosca | Found unit will not power on | checked all connections and powersupply analyzer will still not power on | CE6C657434215344 | No | Found unit will not power on | checked all connections and powersupply analyzer will still not power on | Revisit required | HW | Power Supply | Power Supply
ABL80 FLEX CO-OX | Fehlersuche Einstellungen mit Speicherbefehl abgespeichert | Fehlersuche Einstellungen mit Speicherbefehl abgespeichert | SWO-14781 | No | Troubleshooting settings saved with memory command | Troubleshooting settings saved with memory command
ABL80 CO-OX | STEPPER MOTOR FOUND FROZEN AT LAST VISIT | REPLACED MOTOR AND SOLUTION PACK | CE6DB87BAA4EF7D9 | No | STEPPER MOTOR FOUND FROZEN AT LAST VISIT | REPLACED MOTOR AND SOLUTION PACK
ABL90 | Analyzer at intervention screen indicating SP needs to be changed. Worked through diagnostics and noted inhomogenious error codes. Changed SP, inlet gasket, connecting tubing. Error codes at random inclusive of 1324,1083,1085,1087 all related to inhomogenous. SP changed a total of 6 times with scenario of intervention screen and some form of inhomogenous erroro code. | Sensor cassette bay unit is common point for LLS related codes. Advised AWM for further disposition. | CE6E00F1AA33C83E | No | Analyzer at intervention screen indicating SP needs to be changed. Worked through diagnostics and noted inhomogenious error codes. Changed SP, inlet gasket, connecting tubing. Error codes at random inclusive of 1324,1083,1085,1087 all related to inhomogenous. SP changed a total of 6 times with scenario of intervention screen and some form of inhomogenous erroro code. | Sensor cassette bay unit is common point for LLS related codes. Advised AWM for further disposition. | Assisted Replacement | SC | Wetability
ABL825 FLEX | Drucker getauscht, Mehrere Kalibrierungen durchgeführt, QCs für alle Level durchgeführt -->i.O. ABL wieder Messbereit. | Drucker getauscht, Mehrere Kalibrierungen durchgeführt, QCs für alle Level durchgeführt -->i.O. ABL wieder Messbereit. | SWO-14797 | No | Printer exchanged, several calibrations performed, QCs performed for all levels -->i.O. ABL ready to measure again. | Printer exchanged, several calibrations performed, QCs performed for all levels -->i.O. ABL ready to measure again. | Defect | HW | Printer module | Printer Head
ABL827 FLEX | Spectrum mismatch & phase regulation errors on samples. OXI module diagnostic fails tuning voltage limits. | Replaced hemolyzer per Radiometer specifications. Calibrated and ran QC. Instrument operates normally. | CE6D8F73AA4B1764 | No | Spectrum mismatch & phase regulation errors on samples. OXI module diagnostic fails tuning voltage limits. | Replaced hemolyzer per Radiometer specifications. Calibrated and ran QC. Instrument operates normally. | Defect | HW | Oxi module | Hemolyzer
ABL800 BASIC | Proteinentfernung durchgeführt Lampenabgleich im Serviceprogramm vorgenommen Gerät kalibriert und überprüft | Proteinentfernung durchgeführt Lampenabgleich im Serviceprogramm vorgenommen Gerät kalibriert und überprüft | SWO-14688 | No | Protein removal performed lamp matching in the service program made device calibrated and checked | Protein removal performed lamp matching in the service program made device calibrated and checked | Cleaned | Fluidic pathway | - | Protein Removal
ABL827 FLEX | Error 315- Neon number high. Fneon reading high 2.09 during oxi test. | Replaced spectrometer | CE6D9900AA3E7F44 | No | Error 315-Neon number high. Fneon reading high 2.09 during oxygen test. | Replaced spectrometer | Defect | HW | Oxi module | Spectrometer
ABL837 FLEX | checked data logs for errors, checked hemolyzer and lamp unit, performed complete oxi test, | removed and replaced lamp unit, no change to lamp adjustment, removed and replaced oxi module board, replaced hemolyzer and optical fiber, performed tHb cal and complete oxi test, | CE6E216EAA3C8804 | No | checked data logs for errors, checked hemolyzer and lamp unit, performed complete oxi test, | removed and replaced lamp unit, no change to lamp adjustment, removed and replaced oxi module board, replaced hemolyzer and optical fiber, performed tHb cal and complete oxi test, | Defect | HW | Oxi module | PCB
ABL825 FLEX | Tried loading AU software multiple times. Tried other module to elimintate that as problem. | Replaced the AU board. Replaced the air valve in the PH BG module as it was fused closed. Performed zero current cal and reset all data according to manual after replacing Control board. | CE6DF67EAA0282B7 | No | Tried loading AU software multiple times. Tried other module to elimintate that as problem. | Replaced the AU board. Replaced the air valve in the PH BG module as it was fused closed. Performed zero current cal and reset all data according to manual after replacing Control board. | Defect | HW | Analyzer Control PCB
ABL825 FLEX | Installation. | Visit to install 5 analyzer at Bellevue. Returning 01-30-2015 to complete calibration verification. | CE6EA360201117EE | No | Installation. | Visit to install 5 analyzer at Bellevue. Returning 01-30-2015 to complete calibration verification. | Unknown | Should have been reclassified
AQT90 FLEX | Fließwege mit Hypochlorid gespült, alle Fließstrecken im Servicemode überprüft, kein Fehler festgestellt, mehrere Messungen mit ungenügender Probenmenge im Speicher ( Proben wurden dann evt aus selben Röhrchen noch einmal gemessen und abgebrochen), LP und LP Anschlüsse überprüft, QC durchgeführt | Fließwege mit Hypochlorid gespült, alle Fließstrecken im Servicemode überprüft, kein Fehler festgestellt, mehrere Messungen mit ungenügender Probenmenge im Speicher ( Proben wurden dann evt aus selben Röhrchen noch einmal gemessen und abgebrochen), LP und LP Anschlüsse überprüft, QC durchgeführt | SWO-14847 | No | Flow paths rinsed with hypochloride, all flow lines in service mode checked, no error detected, several measurements with insufficient sample quantity in the memory (samples were then measured and aborted again from the same tube), LP and LP connections checked, QC performed | Flow paths rinsed with hypochloride, all flow lines in service mode checked, no error detected, several measurements with insufficient sample quantity in the memory (samples were then measured and aborted again from the same tube), LP and LP connections checked, QC performed
ABL90 | Confirmed problem | Replaced the valve tubing. | CE6F4FF8204AA270 | No | Confirmed problem | Replaced the valve tubing. | Defect | HW | Oxi module | Pinch Valve tube
ABL90 | Confirmed problem | Replaced the valve tubing. | CE6F6341202AA273 | No | Confirmed problem | Replaced the valve tubing. | Defect | HW | Oxi module | Pinch Valve tube
ABL90 | Reviewed logs and confirmed problem. | Replaced the spectrometer and valve tubing for diagnostic purposes. | CE6F2AD8200CA26D | No | Reviewed logs and confirmed problem. | Replaced the spectrometer and valve tubing for diagnostic purposes. | Defect | HW | Spectrometer
ABL90 FLEX | Nadel und LP getauscht | Nadel und LP getauscht | SWO-14929 | No | Needle and LP exchanged | Needle and LP exchanged | Replaced | Probe
ABL90 | pH QC has been running high in and out of 2 SD range since 1-22. Customer tried new SC and had some improvement but results were still high. Changed SP and pH QC improved but pCO2 calibration sensitivity dropped low out of limits. | Replaced SC and all parameters perform in range for calibration & QC. Filed customer credit forms for customer. Instrument is operating normally. | CE6EFFA920027413 | No | pH QC has been running high in and out of 2 SD range since 1-22. Customer tried new SC and had some improvement but results were still high. Changed SP and pH QC improved but pCO2 calibration sensitivity dropped low out of limits. | Replaced SC and all parameters perform in range for calibration & QC. Filed customer credit forms for customer. Instrument is operating normally. | Assisted Replacement | SC | several parameter
ABL820 FLEX | pCO2 for level 3 and 4 trading towards the low end of package insert. Found that a droplet of rinse solution being left in front of reference electrode | Replaced o ring on pH electrode. Replaced reference membrane. | CE6EEE472063BBA4 | No | pCO2 for level 3 and 4 trading towards the low end of package insert. Found that a droplet of rinse solution being left in front of reference electrode | Replaced o ring on pH electrode. Replaced reference membrane. | Assisted Replacement | Electrode | Membrane | pH
ABL820 FLEX | pCO2 on level 3 amd 4 reading low, barely within package insert | Replaced o ring on pH electrode. Replaced pCO2 electrode. reading on levels were both at mean of package insert. Having relacement elecrode sent. | CE6EEB202006BBA0 | No | pCO2 on level 3 amd 4 reading low, barely within package insert | Replaced o ring on pH electrode. Replaced pCO2 electrode. reading on levels were both at mean of package insert. Having relacement elecrode sent. | Assisted Replacement | Electrode | Membrane | pH
ABL835 FLEX | Glu Elektrode & Messkammer gereinigt und Membraniert. Zusätzlich Lactat Fehler. Nach mehreren Kalibierungen sind die Parameter wieder O.K. | Glu Elektrode & Messkammer gereinigt und Membraniert. Zusätzlich Lactat Fehler. Nach mehreren Kalibierungen sind die Parameter wieder O.K. | SWO-15320 | No | Glu electrode & measuring chamber cleaned and membraneed. In addition, lactate errors. After several calibrations, the parameters are O.K. again. | Glu electrode & measuring chamber cleaned and membraneed. In addition, lactate errors. After several calibrations, the parameters are O.K. again. | Assisted Replacement | Electrode | Membrane | Glu
ABL90 |  |  | CE700D2B962E7C8E | No |  |  | Cleaned | Fluidic pathway | Clot
ABL825 FLEX | Tried protien remover, cleaning , LS adjust and pump cal-No change Replaced inlet gasket-No change Replaced pH o-ring- No change | Replaced listed parts Ran 4 QC samples and 2 patient samples-OK | CE709C89964354DB | No | Tried protien remover, cleaning , LS adjust and pump cal-No change Replaced inlet gasket-No change Replaced pH o-ring- No change | Replaced listed parts Ran 4 QC samples and 2 patient samples-OK | Defect | HW | pH/BG Module | Measuring chamber
ABL90 | Original error began 1-10 with SC change. Instrument instructed to change successive SP's but did not change error. Also found leak at hemolyzer pinch valve tubing. | Changed SC and replaced instrument tubings. Calibrated and ran QC. Instrument operates normally. | CE72D32A826146B8 | No | Original error began 1-10 with SC change. Instrument instructed to change successive SP's but did not change error. Also found leak at hemolyzer pinch valve tubing. | Changed SC and replaced instrument tubings. Calibrated and ran QC. Instrument operates normally. | Assisted Replacement | SC | Wetability
ABL90 |  |  | CE72C65E823BC517 | No |  |  | Cleaned | Fluidic pathway
ABL725 | pH gumigyűrű, El Met - pH BG oldatcső csere. Kalibrálás, QC mérés. |  | SWO-15044 | No | pH rubber Ring, El Met-pH BG solution pipe replacement. Calibration, QC measurement.
ABL825 FLEX | Key installiert und Remote eingerichtet. | Key installiert und Remote eingerichtet. | SWO-15130 | Yes | Key installed and set up remotely. | Key installed and set up remotely. | Re classified RMED | Setup for remote support
ABL825 FLEX | Remote key installiert und eingerichtet. | Remote key installiert und eingerichtet. | SWO-15132 | Yes | Remote key installed and set up. | Remote key installed and set up. | Re classified RMED | Setup for remote support
ABL827 FLEX |  |  | CE7315FD82456A71 | Yes |  |  | Re classified RMED | Training
ABL825 FLEX | completee oxi test | replaced spectrometer | CE72A59E82439828 | No | completee oxi test | replaced spectrometer | Defect | HW | Oxi module | Spectrometer
AQT90 FLEX | Se sustituyeron los componentes de Fluidica y el equipo sigue sin poder calibrar TNI. Se necesita sustituir óptica y process Wheel. |  | CE7DD817A861F572 | No | Fluidica components were replaced and the equipment still cannot calibrate TNI. Optical and process wheel need to be replaced.
AQT90 FLEX | CRP-QC von alter Lot war nicht i.O. Weil außerhalb Bereich. Deshalb mehrere Systemreinigungen durchgeführt. Mehrere CRP QCs, mit neuer Lot durchgeführt. Ampelstatus aber noch auf Gelb weil diverse QCs verfallen sind. Diese müssen von Kunden erst gelöscht werden, dann Ampel wieder auf grün. Reagenzienpack erneuert weil Verfallsdatum erreicht. AQT wieder Messbereit. | CRP-QC von alter Lot war nicht i.O. Weil außerhalb Bereich. Deshalb mehrere Systemreinigungen durchgeführt. Mehrere CRP QCs, mit neuer Lot durchgeführt. Ampelstatus aber noch auf Gelb weil diverse QCs verfallen sind. Diese müssen von Kunden erst gelöscht werden, dann Ampel wieder auf grün. Reagenzienpack erneuert weil Verfallsdatum erreicht. AQT wieder Messbereit. | SWO-15018 | No | CRP-QC of old lot was not i.O. Because outside area. Therefore, several system cleanings are carried out. Multiple CRP QCs, performed with new lot. Traffic light status but still on yellow because various QCs have expired. These must first be deleted by customers, then traffic lights back to green. Reagent pack renewed because expiry date is reached. AQT ready to measure again. | CRP-QC of old lot was not i.O. Because outside area. Therefore, several system cleanings are carried out. Multiple CRP QCs, performed with new lot. Traffic light status but still on yellow because various QCs have expired. These must first be deleted by customers, then traffic lights back to green. Reagent pack renewed because expiry date is reached. AQT ready to measure again.
AQT90 FLEX | Fehlersuche. Sensor von Optical Unit erneuert. | Der Kunde hat im Servicemode Safe Positions die Taste Run All betätigt. Fehler bei Processunit und Proxesswheel. Der Kunde bittet um einen Service. Der Kunde wartet auf Rückruf, um zu erfahren, wann ein Techniker kommt. | SWO-15899 | No | Troubleshooting. sensor from Optical Unit. | The customer has pressed the Button Run All in the Safe Positions service mode. Processunit and Proxesswheel failed. The customer asks for a service. The customer waits for a callback to know when a technician is coming.
ABL90 |  |  | CE75FB78E40DB9AA | No |  |  | Defect | HW | Probe bend
ABL90 FLEX | Hemolyzer, Pinch Valve sowie Solution Pack PCB erneuert Abfallschlauch und Ventilschlauch erneuert Gerät kalibriert, tHb kalibriert und überprüft | Hemolyzer, Pinch Valve sowie Solution Pack PCB erneuert Abfallschlauch und Ventilschlauch erneuert Gerät kalibriert, tHb kalibriert und überprüft | SWO-15217 | No | Hemolyzer, Pinch Valve as well as Solution Pack PCB renews waste hose and valve hose renewed device calibrated, tHb calibrated and checked | Hemolyzer, Pinch Valve as well as Solution Pack PCB renews waste hose and valve hose renewed device calibrated, tHb calibrated and checked | Defect | HW | Hemolyzer
ABL90 FLEX | Schlauch getauscht zwischen Hemolyzer und SC, Kaltstartphase, Kal und QC,s IO | Schlauch getauscht zwischen Hemolyzer und SC, Kaltstartphase, Kal und QC,s IO | SWO-15260 | No | Hose exchanged between Hemolyzer and SC, cold start phase, Kal and QC,s IO | Hose exchanged between Hemolyzer and SC, cold start phase, Kal and QC,s IO | Defect | HW | Oxi module | Pinch Valve tube
ABL90 FLEX | Gerät manuell gespült und Nadel getauscht, Kal und QC IO | Gerät manuell gespült und Nadel getauscht, Kal und QC IO | SWO-15261 | No | Device manually rinsed and needle exchanged, Kal and QC IO | Device manually rinsed and needle exchanged, Kal and QC IO | Defect | HW | Probe bend
ABL90 FLEX | LP im Servicemenue erneut entfernt und montiert, Serviceprogramme gestartet und alle erforderlichen Programme problemlos mit dem eingesetzten LP absolviert, keine Fehlermeldungen. Komplettes Initialisierungsverfahren laufen lassen, OK. Zweites, angebrochenes, LP installiert, Startphase, einmal Spüllösung inhomogen , danach Startphase OK. Bei beiden LPs Meldung 1405 Datenbank Fehler , Neustart ausgelöst, Startphase 50 Minuten, danach Ampel auf grün, OK. | LP im Servicemenue erneut entfernt und montiert, Serviceprogramme gestartet und alle erforderlichen Programme problemlos mit dem eingesetzten LP absolviert, keine Fehlermeldungen. Komplettes Initialisierungsverfahren laufen lassen, OK. Zweites, angebrochenes, LP installiert, Startphase, einmal Spüllösung inhomogen , danach Startphase OK. Bei beiden LPs Meldung 1405 Datenbank Fehler , Neustart ausgelöst, Startphase 50 Minuten, danach Ampel auf grün, OK. | SWO-15276 | No | Removed and mounted LP in the service menu, started service programs and completed all required programs without any problems with the inserted LP, no error messages. Complete initialization procedure, OK. Second, started, LP installed, start phase, once rinsing solution inhomogeneous, then start phase OK. For both LPs message 1405 database error, restart triggered, start phase 50 minutes, then traffic light to green, OK. | Removed and mounted LP in the service menu, started service programs and completed all required programs without any problems with the inserted LP, no error messages. Complete initialization procedure, OK. Second, started, LP installed, start phase, once rinsing solution inhomogeneous, then start phase OK. For both LPs message 1405 database error, restart triggered, start phase 50 minutes, then traffic light to green, OK. | Assisted Replacement | SP | Chip error
ABL90 | Checked solution pack status in the service program and found numerous flow errors. Found contamination all around the inlet gasket. | Performed solution pack test numerous times and the test passed. Performed startup and each time the startup failed with error 1343. Flushed with water and replaced the inlet gasket, but the analyzer still fails the startup. | CE73EC4BF81EB727 | No | Checked solution pack status in the service program and found numerous flow errors. Found contamination all around the inlet gasket. | Performed solution pack test numerous times and the test passed. Performed startup and each time the startup failed with error 1343. Flushed with water and replaced the inlet gasket, but the analyzer still fails the startup. | Assisted Replacement | SC | Wetability
ABL90 | Found contamination around the inlet gasket and fluid blockage messages. | Flushed the analyzer with water and replaced the inlet gasket. Numerous inhomogeneous Cal 2 solution errors. No replacement solution pack or sensor cassettes available. | CE73EE14F80AB72B | No | Found contamination around the inlet gasket and fluid blockage messages. | Flushed the analyzer with water and replaced the inlet gasket. Numerous inhomogeneous Cal 2 solution errors. No replacement solution pack or sensor cassettes available. | Assisted Replacement | SC | Wetability
ABL90 | Performed the service checkout. Flushed the analyzer with water and resolved the fluid flow errors. Analyzer passes calibrations but fails internal QC levels A and C, especially the pCO2, but other parameter fail intermittently. | Replaced the solution pack and sensor cassette. | CE73EFD1F84BB72E | No | Performed the service checkout. Flushed the analyzer with water and resolved the fluid flow errors. Analyzer passes calibrations but fails internal QC levels A and C, especially the pCO2, but other parameter fail intermittently. | Replaced the solution pack and sensor cassette. | Assisted Replacement | SC | Wetability
ABL5 | Fehlersuche V3 reagiert nicht, Ventil ausgetauscht Einlaufverfahren pO2 membraniert | Fehlersuche V3 reagiert nicht; Ventil ausgetauscht Einlaufverfahren pO2 membraniert | SWO-15213 | No | Troubleshooting V3 does not respond; Valve replaced inlet method pO2 membraneed | Troubleshooting V3 does not respond; Valve replaced inlet method pO2 membraneed
ABL817 FLEX | Blauen Nupsi freigespült, Proteinentfernung, Justagen, Schlauchnachfüllung, Kalibrationen, Qualitätskontrollen und Testmessungen durchgeführt. | Blauen Nupsi freigespült, Proteinentfernung, Justagen, Schlauchnachfüllung, Kalibrationen, Qualitätskontrollen und Testmessungen durchgeführt. | SWO-15179 | No | Blue Nupsi rinsed, protein removal, adjustments, hose refill, calibrations, quality controls and test measurements performed. | Blue Nupsi rinsed, protein removal, adjustments, hose refill, calibrations, quality controls and test measurements performed. | Cleaned | Fluidic pathway
ABL825 FLEX | Druckerkopf ausgetauscht | Druckerkopf ausgetauscht | SWO-15344 | No | Printer head replaced | Printer head replaced | Defect | HW | Printer module | Printer Head
ABL825 FLEX |  |  | CE73BF6BF8580A2A | No |  |  | Defect | HW | Oxi module | Hemolyzer
ABL837 FLEX |  |  | CE740D96F8218284 | No |  |  | Defect | HW | Oxi module | Spectrometer
ABL827 FLEX |  |  | CE76512CE43DB79D | No |  |  | Cleaned | Fluidic pathway | -
AQT90 FLEX | Se sustituyo óptica y process Wheel según protocolo del fabricante. Se calibro TNI y Dimer satisfactoriamente. Se comprobó correcto funcoinamiento del equipo mediante QC. |  | CE7DDAF3A842F57B | No | Optical and process wheel are replaced according to manufacturer's protocol. TNI and Dimer are successfully calibro. The equipment was proven to be correct by QC.
AQT90 FLEX |  |  | CE7404B8F80A1C30 | No
ABL80 FLEX BASIC | Altes Material (Lösungspack und Sensorkassette) übergeben, eingesetzt und geprüft | Altes Material (Lösungspack und Sensorkassette) übergeben, eingesetzt und geprüft | SWO-15435 | No | Old material (solution pack and sensor cassette) handed over, inserted and tested | Old material (solution pack and sensor cassette) handed over, inserted and tested
ABL80 FLEX Basic version |  |  | CE755E9D6E0A52B0 | No
ABL90 | AFTER REMOVAL OF BATTERY, ABL 90 WORKS FINE. ITS A BAD BATTERY. | REPLACED BATTERY | CE7566F76E4E0414 | No | AFTER REMOVAL OF BATTERY, ABL 90 WORKS FINE. ITS A BAD BATTERY. | REPLACED BATTERY | Defect | HW | Battery module
ABL90 FLEX | Die Geräteüberprüfung hat ergeben das die Fehlerursache (Bildschirm Flimmern) durch eine lockere Steckverbindung im. Display lag. ABL 90 wieder beim Kunden aufgebaut. | Herr Quick meldet Schneegestöber auf dem ABL Bildschirm. Er hat den Analysator neu gestartet. der Fehler bleibt bestehen. Herr Quick wünscht einen Vorortservice. | SWO-15594 | No | The device check has shown that the cause of the error (screen flicker) was due to a loose plug-in connection in the display. ABL 90 at the customer's place. | Mr Quick reports snowdrops on the ABL screen. He has restarted the analyzer. the error remains. Mr Quick wants a suburban service. | Defect | Hw | Cable seating
ABL90 | Examined Display and barcode reader | Replaced Barcode reader. Verified operation of ABL90. | CE752D456E0CAA0B | No | Examined Display and barcode reader | Replaced Barcode reader. Verified operation of ABL90. | Defect | HW | Barcode Reader
ABL90 |  |  | CE75466E6E49BE66 | No |  |  | Unknown
ABL90 |  |  | CE754F696E3DBE67 | No |  |  | Cleaned | Fluidic pathway
ABL90 FLEX | waste Schlauch getauscht, Kaltstartphase, Kal und QC IO | Waste connector gereinigt, Kaltstartphase, LP getauscht, Systemüberprüfung I.O | SWO-15419 | No | waste hose exchanged, cold start phase, Kal and QC IO | Waste connector cleaned, cold start phase, LP exchanged, system check I.O | Defect | HW | Waste tube
ABL90 | Found error code 978 flow selector calibration error. Replaced flow selector pcb on 1 26 2015 and will replace flow selector pcb again in addition to the wet section control pcb which controls the flow selector pcb since it failed a second time. | Replaced flow selector pcb and wet section control board. Verified instrument operation. | CE765F68E42E96A0 | No | Found error code 978 flow selector calibration error. Replaced flow selector pcb on 1 26 2015 and will replace flow selector pcb again in addition to the wet section control pcb which controls the flow selector pcb since it failed a second time. | Replaced flow selector pcb and wet section control board. Verified instrument operation. | Defect | HW | Flow Selector PCB
ABL800 BASIC | Lactate muss sehr häufig aller 6-10 Tage membraniert werden, Lac-Elektrode ersetzt, mehrere Kal, QC | Lactate muss sehr häufig aller 6-10 Tage membraniert werden, Lac-Elektrode ersetzt, mehrere Kal, QC | SWO-15458 | No | Lactates must be membraneed very frequently every 6-10 days, lac-electrode replaced, several Kal, QC | Lactates must be membraneed very frequently every 6-10 days, lac-electrode replaced, several Kal, QC | Assisted Replacement | Electrode | Electrode | Lac
ABL835 FLEX | COMPROBACIÓN DEL ESTADO DE FUNCIONAMIENTO DEL EQUIPO, COMPROBACIÓN DE MEMBRANAS DE PCO2, SE SUSTITUYE LA MEMBRANA DE PCO2 POR DEFECTO, COMPROBACION DEL ELECTRODO DE PH, SUSTITUCION DE LA MEMBRANA DE REF, COMPROBACION DE REGISTROS DE DATOS, COMPROBACION DE CONECTORES, |  | CE7569A76E06BC68 | No | CHECKING THE STATE OF OPERATION OF THE EQUIPMENT, CHECKING PCO2 MEMBRANES, REPLACEMENT THE PCO2 MEMBRANE FOR DEFECT, CHECK OF THE PH ELECTROGRAPH, REPLACEMENT OF THE REF MEMBRANE, CHECKING DATA REGISTRATIONS, CHECKING CONNECTORS, |  | Assisted Replacement | Electrode | Membrane | pCO2
ABL800 BASIC | Membrane Gasventil PH BG defekt durchstossen, getauscht, Kal und QC durchgeführt |  | SWO-15533 | No | Membrane gas valve PH BG defective lytered, exchanged, Kal and QC performed |  | Defect | HW | pH/BG Module | Valve, air | Membrane
ABL835 FLEX |  | Replaced Glu electrode,measuring chamber and connector | CE7F98FA1E4021FB | No |  | Replaced Glu electrode,measuring chamber and connector | Assisted Replacement | Electrode | Electrode | Glu
ABL815 FLEX | Kal-Lösung 1 korrekt montiert, Nachfüllung, pCO2-Elektrode membraniert, Printer kontrolliert, mehrere Kal abgewartet, CO2–Elektrode ersetzt, korrekte Kal abgewartet | Kal-Lösung 1 korrekt montiert, Nachfüllung, pCO2-Elektrode membraniert, Printer kontrolliert, mehrere Kal abgewartet, CO2–Elektrode ersetzt, korrekte Kal abgewartet | SWO-15736 | No | Kal solution 1 correctly mounted, refill, pCO2 electrode membraned, printer controlled, several cal waits, CO2 electrode replaced, correct cal awaiting | Kal solution 1 correctly mounted, refill, pCO2 electrode membraned, printer controlled, several cal waits, CO2 electrode replaced, correct cal awaiting | User Error | Replacement support
ABL827 FLEX | Checked waste lines: found syringe nipple partially blocked. | Replaced syringe nipple. | CE7540686E001ADA | No | Checked waste lines: found syringe nipple partially blocked. | Replaced syringe nipple. | Defect | HW | Inlet Module | Leaking | Nipple Kit
ABL825 FLEX | E788 erneuert defekt: RI02 | E788 erneuert defekt: RI02 | SWO-15416 | No | E788 renewed defect: RI02 | E788 renewed defect: RI02 | Assisted Replacement | Electrode | Electrode | pCO2
ABL820 FLEX | Time zone is set to Central time. Removed and reentered the QC schedule. | Changed the ABL time to the time of each scheduled QC. All levels of QC were running at the scheduled times. | CE752E506E03397E | No | Time zone is set to Central time. Removed and reentered the QC schedule. | Changed the ABL time to the time of each scheduled QC. All levels of QC were running at the scheduled times. | Setup | Time Setting
ABL820 FLEX | Time zone is set to Central time. Removed and reentered the QC schedule. | Changed the ABL time to the time of each scheduled QC. All levels of QC were running at the scheduled times. | CE7529126E5B397B | No | Time zone is set to Central time. Removed and reentered the QC schedule. | Changed the ABL time to the time of each scheduled QC. All levels of QC were running at the scheduled times. | Setup | Time Setting
ABL820 FLEX | Performed visual inspection, ran applicable service diagnostics. Tested gas system for leaks, none observed, removed and inspected electrodes membranes orings seals, all impacted had been recently replaced at least once, no issues noted. Checked temperature system, no issues noted. | Replaced black tubing from gas to valve manifold on pH BG module. Ran all 3 levels of QC, all parameters are now reading close to the mean for each level. | CE7537F66E22D3E5 | No | Performed visual inspection, ran applicable service diagnostics. Tested gas system for leaks, none observed, removed and inspected electrodes membranes orings seals, all impacted had been recently replaced at least once, no issues noted. Checked temperature system, no issues noted. | Replaced black tubing from gas to valve manifold on pH BG module. Ran all 3 levels of QC, all parameters are now reading close to the mean for each level. | Assisted Replacement | Electrode | Electrode | pCO2
ABL810 FLEX | pCo2 Elektrode getauscht mehrere Kals durchgeführt und QC vermessen | pCo2 Elektrode getauscht mehrere Kals durchgeführt und QC vermessen | SWO-15399 | No | pCo2 electrode swapped several kals performed and QC measured | pCo2 electrode swapped several kals performed and QC measured | Assisted Replacement | Electrode | Electrode | pCO2
ABL90 |  |  | CE767DD9E452456A | No |  |  | Cleaned | Fluidic pathway | Clot
ABL837 FLEX | I spoke with Chris Miller and he said the pCO2 is dropping low after a couple weeks of replacing the membrane. They are having to replace the membranes too often. This is already under investigation by RMED. They have things that we are supposed to check and I have done these. I will gather the data logs and open a techline case. I was not able to remove the pCO2 electrode to check the membrane for leaks. It is due for a membrane change and they do not have any membranes and I did not want to take a chance of removing the electrode from the chamber and it not coming back in. Customer will check it when they get membranes and can change it. | Checked the pCO2 well for bubbles and checked the valves to make sure they have the e-clips on the bottom. I saw no bubbles and the e-clips are in place. Gathered the data for techline case. | CE765554E41C3B0D | No | I spoke with Chris Miller and he said the pCO2 is dropping low after a couple weeks of replacing the membrane. They are having to replace the membranes too often. This is already under investigation by RMED. They have things that we are supposed to check and I have done these. I will gather the data logs and open a techline case. I was not able to remove the pCO2 electrode to check the membrane for leaks. It is due for a membrane change and they do not have any membranes and I did not want to take a chance of removing the electrode from the chamber and it not coming back in. Customer will check it when they get membranes and can change it. | Checked the pCO2 well for bubbles and checked the valves to make sure they have the e-clips on the bottom. I saw no bubbles and the e-clips are in place. Gathered the data for techline case. | Assisted Replacement | Electrode | Electrode | pCO2
ABL805 FLEX | Wanda said the pCO2 is dropping low after a couple weeks of replacing the membrane. This is already under investigation by RMED. They have things that we are supposed to check and I have done these. I will gather the data logs and open a techline case. Found a little contamination on the Cl membrane so I chnged the membrane with customer stock. | Checked the pCO2 for bubbles, the upper valves for the e-clip and pCO2 membrane for leaking. All checked good. | CE7635B0E45E3B07 | No | Wanda said the pCO2 is dropping low after a couple weeks of replacing the membrane. This is already under investigation by RMED. They have things that we are supposed to check and I have done these. I will gather the data logs and open a techline case. Found a little contamination on the Cl membrane so I chnged the membrane with customer stock. | Checked the pCO2 for bubbles, the upper valves for the e-clip and pCO2 membrane for leaking. All checked good. | Assisted Replacement | Electrode | Electrode | pCO2
ABL825 FLEX |  | Replaced ref Electrode | CE7F9AEC1E0B2203 | No |  | Replaced ref Electrode | Assisted Replacement | Electrode | Electrode | Reference
ABL825 FLEX |  | Replaced ref Electrode | CE7F9A301E472200 | No |  | Replaced ref Electrode | Assisted Replacement | Electrode | Electrode | Reference
ABL825 FLEX |  | Replaced Ref electrode | CE7F9CAB1E062209 | No |  | Replaced Ref electrode | Assisted Replacement | Electrode | Electrode | Reference
ABL825 FLEX |  | Replaced ref electrode | CE7F9BE11E212206 | No |  | Replaced ref electrode | Assisted Replacement | Electrode | Electrode | Reference
ABL837 FLEX | Inspection of the lower pH sealing ring (P#834-660 previous style)revealed the lower side of the pH lower seal in worn or damaged causing fluid and b ubbles to accumulate in the same lower portion of the well. | The lower pH seal was replaced with the new style lower pH seal (P3834-682). | CE766C51E435EC4B | No | Inspection of the lower pH sealing ring (P#834-660 previous style)revealed the lower side of the pH lower seal in worn or damaged causing fluid and b ubbles to accumulate in the same lower portion of the well. | The lower pH seal was replaced with the new style lower pH seal (P3834-682). | Assisted Replacement | Electrode | Electrode | pH | Sealing ring
ABL835 FLEX |  |  | CE7660CDE40017E1 | No |  |  | Defect | HW | Inlet Module | Friction detector unit
ABL805 FLEX | Kein Gerätefehler feststellbar, Gerinnsel entfernt. Da es diese Woche bereits Probleme mit pCO2 gab zusätzlich Ventile überprüft und pCO2 membraniert um das Einschwingverhalten zu bewerten. | Kein Gerätefehler feststellbar, Gerinnsel entfernt. Da es diese Woche bereits Probleme mit pCO2 gab zusätzlich Ventile überprüft und pCO2 membraniert um das Einschwingverhalten zu bewerten. | SWO-15614 | No | No device error detectable, clots removed. Since there were already problems with pCO2 this week additionally checked valves and pCO2 membraneed to evaluate the settling behavior. | No device error detectable, clots removed. Since there were already problems with pCO2 this week additionally checked valves and pCO2 membraneed to evaluate the settling behavior. | Cleaned | Fluidic pathway | - |  | pCO2
ABL825 FLEX | pCO2 unstable during qc measurements. Membrane has already been replaced. | Replaced electrode | CE767B4EE45C0776 | No | pCO2 unstable during qc measurements. Membrane has already been replaced. | Replaced electrode | Assisted Replacement | Electrode | Electrode | pCO2
ABL825 FLEX | LOOKED THROUGH DATA LOGS AND FOUND ABL DMS PC DISCONNECTED FROM AU 0643. IT HAPPENED 11 TIMES IN THE LAST 3 DAYS | REPLACED POWER SUPPLY AND LOOKED AT ALL CONNECTIONS. | CE76581DE44E74E1 | No | LOOKED THROUGH DATA LOGS AND FOUND ABL DMS PC DISCONNECTED FROM AU 0643. IT HAPPENED 11 TIMES IN THE LAST 3 DAYS | REPLACED POWER SUPPLY AND LOOKED AT ALL CONNECTIONS. | Defect | HW | Power Supply
AQT90 FLEX | Prepared unit for removal. | Performed long term shutdown and crated. | CE7806075A34C392 | No | Prepared unit for removal. | Performed long term shutdown and crated.
ABL80 FLEX Basic version |  |  | CE77CC865A0DE9EF | No
ABL80 CO-OX | Analyzer had leaked clear fluid from SP area and on to a book and the floor. Customer had replaced the SP and leaking stopped. Power down analyzer and inspect. | No correction needed at this time as the rep laced SP appears to have resolved the leaking issue. | CE7791005A281E4A | No | Analyzer had leaked clear fluid from SP area and on to a book and the floor. Customer had replaced the SP and leaking stopped. Power down analyzer and inspect. | No correction needed at this time as the rep laced SP appears to have resolved the leaking issue.
ABL90 | Having different substrate errors. Can see leakage by hemolyzer pinch valve tubing. Disassemble analyzer. Leak was older and dried up. Leak is from pinch valve tubing. SC is still suspect for 1221 error. | PM is due beginning of next month but since analyzer is apart, replaced parts from PM kit. Ran all Fluidic tests in service. Getting 1221 temp error even after recalibrating thermistor in service. Reinstall original SC. SC has resolved temp issue. Now have OXI issues. Correct loose cable on Wet Section PCB. Level2 internal QC out on all CO-OX parameters but external QC has no CO-Ox issues - replace SP. Internal QC ok now. | CE77D85B5A0D7B6D | No | Having different substrate errors. Can see leakage by hemolyzer pinch valve tubing. Disassemble analyzer. Leak was older and dried up. Leak is from pinch valve tubing. SC is still suspect for 1221 error. | PM is due beginning of next month but since analyzer is apart, replaced parts from PM kit. Ran all Fluidic tests in service. Getting 1221 temp error even after recalibrating thermistor in service. Reinstall original SC. SC has resolved temp issue. Now have OXI issues. Correct loose cable on Wet Section PCB. Level2 internal QC out on all CO-OX parameters but external QC has no CO-Ox issues - replace SP. Internal QC ok now. | Assisted Replacement | SC | Thermistor
ABL90 | Customer replaced sensor cassette and sloution pack | Performed sensor simulator test and passed. Performed PM. Ran calibrations and Q.C Level 3 still low for internal Q.C. Swapped solution pack with instrument next to it. Ran Q.C. still low for glucose level 3. Replaced sensor cassette to resolve. Will submit credit form for sensor cassette | CE77E5265A085C49 | No | Customer replaced sensor cassette and sloution pack | Performed sensor simulator test and passed. Performed PM. Ran calibrations and Q.C Level 3 still low for internal Q.C. Swapped solution pack with instrument next to it. Ran Q.C. still low for glucose level 3. Replaced sensor cassette to resolve. Will submit credit form for sensor cassette | Assisted Replacement | SC | Glu
ABL90 |  |  | CE77CC855A3AECB6 | No |  |  | Defect | HW | Probe bend
ABL90 FLEX | Serviceprogramm Test, Pressure Fehler, Wechsel des LP, Kal, Wechsel der SC, da Aufruf unter Betriebsmittel, Kal, QC Ausgebauter fehlerhafte Lösungspack Lot RN07 | Serviceprogramm Test, Pressure Fehler, Wechsel des LP, Kal, Wechsel der SC, da Aufruf unter Betriebsmittel, Kal, QC Ausgebauter fehlerhafte Lösungspack Lot RN07 | SWO-15720 | No | Service program test, pressure error, change of LP, Cal, change of SC, since call under equipment, Kal, QC Removed faulty solution pack Lot RN07 | Service program test, pressure error, change of LP, Cal, change of SC, since call under equipment, Kal, QC Removed faulty solution pack Lot RN07 | Assisted Replacement | SC | Wetability
ABL825 FLEX | Na dropping lower on all QC. Exchanges reference and Na electrodes with little improvement. Swapped membrane. | Very little improvement, but still within insert, and just on the edge of peer. Customer will monitor in next 14 days to see if improvements ore ok. | CE77D3595A111BBC | No | Na dropping lower on all QC. Exchanges reference and Na electrodes with little improvement. Swapped membrane. | Very little improvement, but still within insert, and just on the edge of peer. Customer will monitor in next 14 days to see if improvements ore ok. | Assisted Replacement | Electrode | Electrode | Reference
ABL827 FLEX |  |  | CE77CCF45A597FE9 | No |  |  | Cleaned | Fluidic pathway | - |  | pH
AQT90 FLEX | Fluidic path gereinigt. Filter und P4 erneuert. CRP neu kalibriert. LQC's durchgeführt. CRP noch immer zu hoch. Fehler Vorort nicht lokalisierbar. Nach Rücksprache mit Kunden CRP gesperrt. Kunde erhält Montag einen Anruf über die weitere vorgehensweise. | Kunde wollte Service erst am 6.2. gegen 9:00 Uhr | SWO-15719 | No | Fluidic path cleaned. filter and P4. CRP recalibrated. LQC's. CRP still too high. Error suburb not localizable. After consultation with customers CRP blocked. Customer will receive a call Monday about what to do next. | Customer wanted service only on 6.2. at 9:00 a.m.
AQT90 FLEX |  |  | CE7816E15A53A20C | No |  |  | Defect | Consumables | SP | Chip error
tc sensor TOSCA 92 | Vervangen van sensor door nieuwe. Deze werd naar de klant opgestuurd onder waarborg. Techline case RBNL 708333096 gemaakt. |  | SWO-15728 | No | Replacing sensor with new. This was sent to the customer under guarantee. Techline case RBNL 708333096 made.
ABL80 FLEX Basic version |  |  | CE7CE20C3206077D | No
ABL90 FLEX | LP ersetzt, Kalibration, QC Kontrolle, tHb Kal. | LP ersetzt, Kalibration, QC Kontrolle, tHb Kal. | SWO-15921 | No | LP replaced, calibration, QC control, tHb Kal. | LP replaced, calibration, QC control, tHb Kal. | Assisted Replacement | SP
ABL90 |  |  | CE7B8DACBC16BE65 | No |  |  | Defect | HW | Oxi module
ABL90 | Comprobación del estado del analizador. Se comprueba el origen del problema, siendo la inlet gasket rota en el interior del m ulo de entrada de muestras. Desmontaje y extracci de la pieza rota y posterior montaje de entrada de muestras. Revisi y comprobación del equipo con Cal y QC's. | Se comprueba el origen del problema, siendo la inlet gasket rota en el interior del m ulo de entrada de muestras. Desmontaje y extracci de la pieza rota y posterior montaje de entrada de muestra | CE7BA906BC199E8D | No | Checking the status of the scanner. The source of the problem is checked, with the inlet gasket being broken inside the sample input m ulo. Disassembly and removal of the broken part and subsequent assembly of sample input. Check and check the equipment with Cal and QC's. | The source of the problem is checked, with the inlet gasket being broken inside the sample input m ulo. Disassembly and removal of the broken part and subsequent sample input assembly | Defect | Inlet gasket
ABL90 | Chequeo de circuito neumatico , limpieza a fondo,, ejecuación de programas de servicio, realización de varias calibraciones, pasados qc, cumplimentación de parte de trabajo. | Chequeo de circuito neumatico , limpieza a fondo,, ejecuación de programas de servicio, realización de varias calibraciones, pasados qc. | CE7B4B8CBC3D31B8 | No | Pneumatic circuit check, thorough cleaning, operation of service programs, performing various calibrations, past qc, filling part of work. | Pneumatic circuit check, thorough cleaning, operation of service programs, performing various calibrations, passed qc. | Cleaned | Fluidic pathway
ABL827 FLEX |  |  | CE7B4BBBBC63A437 | No |  |  | Defect | HW | Printer module | Printer Head
ABL800 BASIC | Reinigung, Proteinentfernung Meßkammer gereinigt | Reinigung, Proteinentfernung Meßkammer gereinigt | SWO-15879 | No | Cleaning, protein removal measuring chamber cleaned | Cleaning, protein removal measuring chamber cleaned | Cleaned | Fluidic pathway
ABL825 FLEX |  | Replaced ref electrode | CE7F9F3D1E252218 | No |  | Replaced ref electrode | Assisted Replacement | Electrode | Electrode | Reference
ABL800 BASIC | Fehlersuche Verbogene Einlaßnadel ersetzt Anpassung Rilibäk Grenzen | Fehlersuche Verbogene Einlaßnadel ersetzt Anpassung Rilibäk Grenzen | SWO-15875 | No | Troubleshooting Bent Inlet Needle Replaces Adjustment Rilibäk Limits | Troubleshooting Bent Inlet Needle Replaces Adjustment Rilibäk Limits | Defect | HW | Inlet module | Probe bend
ABL820 FLEX | checked QC log, 6 S7355 QC aborted in a row, checked AutoCheck module, carousel and ampoul breaker | removed and cleaned stuck carousel, straightened inlet probe, removed and replaced inlet probe tube | CE7BE27CBC4C4501 | No | checked QC log, 6 S7355 QC aborted in a row, checked AutoCheck module, carousel and ampoul breaker | removed and cleaned stuck carousel, straightened inlet probe, removed and replaced inlet probe tube | Defect | HW | Inlet Module | Probe bend
ABL827 FLEX | I was headed to Greensboro this morning for a courtesy cal,l and turned to come here. I tried a hemolyzer but could not get the lamp to adjust above 2500. i replaced the OXI control PCB, fibre optic cable and lamp unit with hemolyzer. Was able to adjust lamp properly. 0791 error still there. Replaced the OXI liquid sensor and error went away. I started getting 0581-OXI spectrum mismatch on blood. Put their hemolyzer and OXI PCB back in. I was able to adjust lamp. | Replaced the OXI fibre optic cable, lamp unit and liquid sensor. Replaced solution pump tube and electrode pump tubes with custoemr stock. Adjusted lamp. Ran Complete OXI test, pump calibration and liquid sensor adjustment. | CE7B75AEBC513B0F | No | I was headed to Greensboro this morning for a courtesy cal,l and turned to come here. I tried a hemolyzer but could not get the lamp to adjust above 2500. i replaced the OXI control PCB, fibre optic cable and lamp unit with hemolyzer. Was able to adjust lamp properly. 0791 error still there. Replaced the OXI liquid sensor and error went away. I started getting 0581-OXI spectrum mismatch on blood. Put their hemolyzer and OXI PCB back in. I was able to adjust lamp. | Replaced the OXI fibre optic cable, lamp unit and liquid sensor. Replaced solution pump tube and electrode pump tubes with custoemr stock. Adjusted lamp. Ran Complete OXI test, pump calibration and liquid sensor adjustment. | Unknown | HW | Oxi module | Hemolyzer
ABL825 FLEX | Laser scanner failed to read QC ampoules. Power to scanner but the scanner does not reset. Possible issue may stem from power supplied to analyzers. Analyzers were not connected to a UPS. | Replaced AutoCheck PCB and Laser scanner. Installed a UPS received from IT. | CE7B81D7BC436AD4 | No | Laser scanner failed to read QC ampoules. Power to scanner but the scanner does not reset. Possible issue may stem from power supplied to analyzers. Analyzers were not connected to a UPS. | Replaced AutoCheck PCB and Laser scanner. Installed a UPS received from IT. | Defect | HW | AutoCheck Module | Barcode reader
ABL825 FLEX | LOOKED IN DATA LOGS AND FOUND MULTIPLE ERRORS FOR 643 | REPLACED ANALYZER CONTROL PCB, LOADED SOFTWARE TO BOARD, LOADED KEYS, RAN 2PT CALS, DID ZERO CALIBRATION. | CE7B7988BC27A494 | No | LOOKED IN DATA LOGS AND FOUND MULTIPLE ERRORS FOR 643 | REPLACED ANALYZER CONTROL PCB, LOADED SOFTWARE TO BOARD, LOADED KEYS, RAN 2PT CALS, DID ZERO CALIBRATION. | Defect | HW | Analyzer Control PCB
ABL815 FLEX |  |  | CE7B6948BC15A6CD | Yes |  |  | Re classified RMED | FlexQ installation
AQT90 FLEX | Beste Erry, Ik heb zojuist de AQT van Schagen even nagelopen. Allereerst kan ik de fout van gisteren niet terug vinden in de logfiles van de AQT. Ik vermoed daarom dat de fout in de windows omgeving (besturingssysteem van de AQT) heeft plaatsgevonden en niet in de AQT software. Ik zou je willen adviseren om het de komende tijd even in de gaten te houden, mocht het vaker voorkomen dan zal één van onze technici er even naar moeten kijken. Ik heb de nieuwe LQC level 1 (lotnr 09445) ingelezen en één keer gedraaid, deze was goed. Ik vraag mij alleen af hoe en of je in het verleden nieuwe lotnummers LQC's hebt ingelezen? Want, zoals gisteren telefonisch al aangegeven, kan je, als je niet ingelogd bent, niet bij deze functies in de software. Als je nog geen eigen inlogcode hebt, is het wellicht voor jou als verantwoordelijke voor de AQT in Schagen, handig om deze aan te maken zodat je wel dit soort aanpassingen kunt doen in de software. Ik help je daar graag bij. Het gebruikte LQC buisje zet ik in een bekertje in de koelkast, deze is vanaf vandaag nog 5 dagen houdbaar in de koelkast. Het doosje met de overige LQC buisjes staat weer in de vriezer. Ik hoop je voor nu voldoende geïnformeerd te hebben. Met vriendelijke groet, Kind regards, Erwin Oeben Application Specialist Specialty Testing Radiometer Benelux BV | Beste Erry, Ik heb zojuist de AQT van Schagen even nagelopen. Allereerst kan ik de fout van gisteren niet terug vinden in de logfiles van de AQT. Ik vermoed daarom dat de fout in de windows omgeving (besturingssysteem van de AQT) heeft plaatsgevonden en niet in de AQT software. Ik zou je willen adviseren om het de komende tijd even in de gaten te houden, mocht het vaker voorkomen dan zal één van onze technici er even naar moeten kijken. Ik heb de nieuwe LQC level 1 (lotnr 09445) ingelezen en één keer gedraaid, deze was goed. Ik vraag mij alleen af hoe en of je in het verleden nieuwe lotnummers LQC's hebt ingelezen? Want, zoals gisteren telefonisch al aangegeven, kan je, als je niet ingelogd bent, niet bij deze functies in de software. Als je nog geen eigen inlogcode hebt, is het wellicht voor jou als verantwoordelijke voor de AQT in Schagen, handig om deze aan te maken zodat je wel dit soort aanpassingen kunt doen in de software. Ik help je daar graag bij. Het gebruikte LQC buisje zet ik in een bekertje in de koelkast, deze is vanaf vandaag nog 5 dagen houdbaar in de koelkast. Het doosje met de overige LQC buisjes staat weer in de vriezer. Ik hoop je voor nu voldoende geïnformeerd te hebben. Met vriendelijke groet, Kind regards, Erwin Oeben Application Specialist Specialty Testing Radiometer Benelux BV | SWO-15760 | No | Dear Erry, I have just nagelopen the AQT van Schagen. First of all, I cannot find the error of yesterday in the log files of the AQT. I therefore suspect that the error in the Windows environment (operating system of the AQT) has occurred and not in the AQT software. I would like to advise you to keep an eye on it in the near future, if it is more common then one of our technicians will have to look at it. I have read the new LQC Level 1 (lotnr 09445) and rotated once, this one was good. I just wonder how and if you've read new lottery numbers in the past? Because, as indicated by phone yesterday, if you are not logged in, you will not be able to do these features in the software. If you don't already have your own login code, it might be for you to be responsible for the AQT in Schagen, handy to create these so you can do this kind of customization in the software. I would like to help you with that. I put the used LQC tube in a cup in the refrigerator, which is still available for 5 days in the refrigerator from today. The box with the other LQC tubes is in the freezer again. I hope you have been adequately informed for now. Sincerely, child regards, Erwin Oeben Application Specialist Specialty Testing radiometer Benelux BV | Dear Erry, I have just nagelopen the AQT van Schagen. First of all, I cannot find the error of yesterday in the log files of the AQT. I therefore suspect that the error in the Windows environment (operating system of the AQT) has occurred and not in the AQT software. I would like to advise you to keep an eye on it in the near future, if it is more common then one of our technicians will have to look at it. I have read the new LQC Level 1 (lotnr 09445) and rotated once, this one was good. I just wonder how and if you've read new lottery numbers in the past? Because, as indicated by phone yesterday, if you are not logged in, you will not be able to do these features in the software. If you don't already have your own login code, it might be for you to be responsible for the AQT in Schagen, handy to create these so you can do this kind of customization in the software. I would like to help you with that. I put the used LQC tube in a cup in the refrigerator, which is still available for 5 days in the refrigerator from today. The box with the other LQC tubes is in the freezer again. I hope you have been adequately informed for now. Sincerely, child regards, Erwin Oeben Application Specialist Specialty Testing radiometer Benelux BV
AQT90 FLEX |  |  | CE7B7D3BBC2C9367 | No
ABL80 CO-OX | FOUND ANALYZER OPERATIONAL. RAN 3 LEVELS OF QC, BROUGHT ANALYZER TO A GREEN LIGHT. NO MECHANICAL OR ELCTRONIC ISSUE FOUND | DISCUSSED DETAILS OF FAILURES WITH CUSTOMER. CUSTOMER AGREED TO WORK THROUGH TROUBLESHOOTING WITH KEITH CLARK DIRECTLY AND FURTHER TRAINING WILL BE CONSIDERED. | CE7CCEDA322E8BF1 | No | FOUND ANALYZER OPERATIONAL. RAN 3 LEVELS OF QC, BROUGHT ANALYZER TO A GREEN LIGHT. NO MECHANICAL OR ELCTRONIC ISSUE FOUND | DISCUSSED DETAILS OF FAILURES WITH CUSTOMER. CUSTOMER AGREED TO WORK THROUGH TROUBLESHOOTING WITH KEITH CLARK DIRECTLY AND FURTHER TRAINING WILL BE CONSIDERED.
ABL80 CO-OX | remove solution pack and cassette and still same. | Installed smart chip | CE7C99213202CEF5 | No | remove solution pack and cassette and still same. | Installed smart chip
ABL90 |  |  | CE7CA9B03239C4CF | No |  |  | Setup | Time Setting
ABL90 FLEX | Keine IP Adresse am ABL hinterlegt. Adresse eingegeben und abgespeichert, Neustart, Datenübertragung im LIS überprüft --> iO. | Keine IP Adresse am ABL hinterlegt. Adresse eingegeben und abgespeichert, Neustart, Datenübertragung im LIS überprüft --> iO. | SWO-16155 | No | No IP address stored on the ABL. Address entered and saved, restart, data transfer checked in LIS --> iO. | No IP address stored on the ABL. Address entered and saved, restart, data transfer checked in LIS --> iO. | Setup | HIS/LIS
ABL90 | Analyzer requires excessive pressure to flush the waste system. Analyzer has had numerous fluid type errors. | Replaced the Waste Connector, Waste Line, Hemolyzer Unit, Pinch Tubing and Sealing Gasket and analyzer now flushes much easier. | CE7CB57432145C05 | No | Analyzer requires excessive pressure to flush the waste system. Analyzer has had numerous fluid type errors. | Replaced the Waste Connector, Waste Line, Hemolyzer Unit, Pinch Tubing and Sealing Gasket and analyzer now flushes much easier. | Assisted Replacement | SP | Chip error
ABL820 FLEX | készülék újraindítása, teljes OXI teszt, OXI PCB csere, kalibráció, tHb kalibráció, QC mérés |  | SWO-15994 | No | Device restart, full OXY test, OXY PCB replacement, calibration, tHb calibration, QC measurement |  | Defect | HW | Oxi module | PCB
ABL820 FLEX | will replaced CPU | Replaced CPU, during QC an Inhomogeneous sample error came up so replaced inlet gasket with client stock. | CE7CC6683262FB4D | No | will replaced CPU | Replaced CPU, during QC an Inhomogeneous sample error came up so replaced inlet gasket with client stock. | Defect | HW | CPU
ABL825 FLEX | Ventilschlauch erneuert. Gerät von Blutaustritt gereinigt. | Ventilschlauch erneuert. Gerät von Blutaustritt gereinigt. | SWO-16115 | No | Valve hose renewed. device cleaned of blood leakage. | Valve hose renewed. device cleaned of blood leakage. | Defect | HW | Oxi module | Valve tubing
ABL825 FLEX | Überprüfung Einlass und pH BG Modul Einlassnadel stark verbogen, verschmutzte Einlassdichtung => Austausch Nadel und Dichtung Flüssigkeitsrückstände an pH und Ref Elektrode => Reinigung Fällige Elektroden membraniert => pCO2, pO2, Na, K, Cl, Ca | Kundin berichtet, dass beide Geräte der Station ausgefallen sind und die identischen Fehler zeigen, Kundin ist nicht am Gerät und wünscht dringenden Einsatz des zuständigen Technikers, sie meinte, sie habe schon die Abdeckung runter genommen und gesehen, dass Flüssigkeit ausläuft, Kundin bittet um Rückruf. | SWO-16002 | No | Check inlet and pH BG module inlet needle strongly bent, dirty inlet seal => Exchange needle and seal liquid residues to pH and ref electrode => cleaning Due electrodes membraneed => pCO2, pO2, Na, K, Cl, Ca | Customer reports that both devices of the station have failed and the identical faults show, customer is not on the device and wishes urgent use of the responsible technician, she said she has already taken down the cover and seen that liquid expirs, customer asks for recall. | Defect | HW | Inlet module | Probe bend
ABL825 FLEX | Überprüfung El Met Modul Unterer Dichtring an Glu Elektrode lose => Dichtring ordentlich angebracht und Messkammer gereinigt Fällige Betriebsmittel getauscht => pCO2, pO2, Na, Kund Cl membraniert Einlassdichtung getauscht | Überprüfung El Met Modul Unterer Dichtring an Glu Elektrode lose => Dichtring ordentlich angebracht und Messkammer gereinigt Fällige Betriebsmittel getauscht => pCO2, pO2, Na, Kund Cl membraniert Einlassdichtung getauscht | SWO-16001 | No | Check El Met Module Lower sealing ring to Glu electrode loose => Sealing ring neatly attached and measuring chamber cleaned Due equipment exchanged => pCO2, pO2, Na, Kund Cl membraneed inlet seal exchanged | Check El Met Module Lower sealing ring to Glu electrode loose => Sealing ring neatly attached and measuring chamber cleaned Due equipment exchanged => pCO2, pO2, Na, Kund Cl membraneed inlet seal exchanged | Assisted Replacement | Electrode | Membrane | Glu
AQT90 FLEX | Jahreswartung durchgeführt | Jahreswartung durchgeführt | SWO-16448 | Yes | Annual maintenance carried out | Annual maintenance carried out | Re classified RMED | JMP mail: There is one upSWO to be removed (has been agreed with Michael Brand)
AQT90 FLEX |  |  | CE7F24781E193E23 | No
ABL80 FLEX Basic version |  |  | CE7E1A4CA818C926 | No
ABL90 FLEX | Fehlersuche Probenweg gespült Einlassdichtungshalter und Einlassverbindungsdichtung erneuert QC-Messungen mit QCAutoABL90 Level 1 bis 3 durchgeführt | der Anwender wünscht einen Vorort-Service zur Beseitigung der Geräte-Störung; | SWO-16228 | No | Troubleshooting sample path flushed inlet seal holder and inlet connection seal renewed QC measurements performed with QCAutoABL90 Level 1 to 3 | the user wants a suburban service to eliminate the device malfunction; | Cleaned | Fluidic pathway
ABL80 CO-OX | Substituição do modulo frontal completo (retirado do equipamento 302029). Inicialização. Calibração de ctHb. Realização de ciclo de Qualidade. Verificação com realização de amostras. |  | CE7DE9D0A842B547 | No | Replacement of the full frontal module (removed from the 302029 equipment). Startup. Calibration of ctHb. Realization of quality cycle. Verification with the realization of samples.
ABL825 FLEX |  |  | CE7E088CA822898E | No |  |  | Defect | HW | Oxi module | Hemolyzer
ABL837 FLEX | Statistics show a N:0 | Reloaded 6.13 software. Ran three levels of QC statistics arenow being updated. N:1 | CE7E0990A813CF5D | No | Statistics show a N:0 | Reloaded 6.13 software. Ran three levels of QC statistics arenow being updated. N:1 | Defect | SW | Windows
TCM/Tosca | found that unit will not power on. | It was a combination issue with both the Power supply and the front button. | CE7DCD0CA85BE8D9 | No | found that unit will not power on. | It was a combination issue with both the Power supply and the front button. | Defect | HW | Power Supply | Power Supply
ABL90 |  |  | CE7F3C481E223A8B | No |  |  | Cleaned | Fluidic pathway
ABL90 | Confirmed corrupt file. Reseated CF card and re-attempted to start analyzer, failed during program load, the file or directory D: is corrupt and Windows cannot find d: | Reloaded Windows and ABL 90 application, restored setup parameters. | CE7FA6CB1E49EBEE | No | Confirmed corrupt file. Reseated CF card and re-attempted to start analyzer, failed during program load; the file or directory D: is corrupt and Windows cannot find d: | Reloaded Windows and ABL 90 application, restored setup parameters. | Defect | SW | Roma.ini
ABL90 | as stated above | i altered astm.ini file by deelting patientDepartment per instruction | CE7F39961E147BFF | No | as stated above | i altered astm.ini file by deelting patientDepartment per instruction | Defect | HIS/LIS
ABL90 FLEX | Sostituiti tutti i tubi e apparecchio testato. | Sostituiti tutti i tubi e apparecchio testato. | SWO-16444 | No | Replaced all tubes and tested appliance. | Replaced all tubes and tested appliance. | Cleaned | Fluidic pathway | replace Tubings
ABL90 FLEX | THb-Kalibrierung durchgeführt QC-Messung Level 2 | THb-Kalibrierung durchgeführt QC-Messung Level 2 | SWO-16468 | No | THb calibration performed QC measurement level 2 | THb calibration performed QC measurement level 2 | User Error | tHb calibration | Expired
ABL810 FLEX |  |  | CE7F65691E5E327A | No |  |  | Defect | HW | Fluidic manifold | Pump motor, right
ABL815 FLEX | Kundeneigene Elektrode in der GS gereinigt, einkalibriert und wieder zurück getauscht. Absprache mit Kunden getroffen das die QC Ergebbnisse weiterhin überwacht werden und bei Problemen wird sich der Kunde wieder melden. | pCO2 Elektrode gereinigt, membraniert, Gerät kalibriert und QC durchgeführt. pCO2 Elektrode zur Überprüfung mitgenommen. Kundeneigene pCO2 Elektrode zur Überbrückung eingesetzt. | SWO-16480 | No | Customer-owned electrode cleaned in the GS, calibrated and replaced again. Consultation with customers has been made that the QC's recovering orders will continue to be monitored and in case of problems the customer will come back. | pCO2 electrode cleaned, membraneed, device calibrated and QC performed. pCO2 electrode for inspection. Customer-owned pCO2 electrode used for bridging. | Assisted Replacement | Electrode | Electrode | pCO2
ABL835 FLEX | Jeg har forsøgt at skifte Inlet PCB, men dette har ikke løst problemet. Jeg har skiftet Friction detector og nu virker inlet igen. Inlet er testet og fundet ok. Der er kørt kal. og QC. Alt er OK. | Jeg har forsøgt at skifte Inlet PCB, men dette har ikke løst problemet. Jeg har skiftet Friction detector og nu virker inlet igen. Inlet er testet og fundet ok. Der er kørt kal. og QC. Alt er OK. | SWO-16701 | No | I have tried to change the Inlet PCB, but this has not solved the problem. I changed the Friction detector and now the inlet works again. Inlet is tested and found OK. There's a run of the Kal. and QC. Everything is OK. | I have tried to change the Inlet PCB, but this has not solved the problem. I changed the Friction detector and now the inlet works again. Inlet is tested and found OK. There's a run of the Kal. and QC. Everything is OK. | Defect | HW | Inlet Module | Friction detector unit
ABL820 FLEX | VISUAL INSPECTION SHOWS PROBE IS SNAPPED OFF | REPLACED INLET PROBE.PERFORMED PUMP CALIBRATION AND LIQUID SENSOR ADJUST.PERFORM INLET CALIBRATION. | CE7F66351E2A0A4E | No | VISUAL INSPECTION SHOWS PROBE IS SNAPPED OFF | REPLACED INLET PROBE. PERFORMED PUMP CALIBRATION AND LIQUID SENSOR ADJUST. PERFORM INLET CALIBRATION. | Defect | HW | Inlet Module | Probe
AQT90 FLEX | Gerät überprüft, Gerät ausgeschaltet, Barcodescanner und Kabel überprüft, gerät neu gestartet. Scanner i.O. | Gerät überprüft, Gerät ausgeschaltet, Barcodescanner und Kabel überprüft, gerät neu gestartet. Scanner i.O. | SWO-16486 | No | Device checked, device turned off, barcode scanner and cable checked, restarts. Scanner i.O. | Device checked, device turned off, barcode scanner and cable checked, restarts. Scanner i.O.
TCM/Tosca |  |  | CE7F45551E351E12 | No
ABL80 FLEX | Bildschirm ersetzt und geprüft. |  | SWO-16551 | No | screen replaced and checked.
ABL80 FLEX CO-OX | CO-OX Control Board augetauscht tHb Kal Funktionskontrolle i.o Datei Aktivitäten und QA Zyklen zum auswerten aus dem Analysator gezogen 1 Schachtel 944-021 tHb Kal kostenlos dem Kunden übergeben | CO-OX Control Board augetauscht tHb Kal Funktionskontrolle i.o Datei Aktivitäten und QA Zyklen zum auswerten aus dem Analysator gezogen 1 Schachtel 944-021 tHb Kal kostenlos dem Kunden übergeben | SWO-16571 | No | CO-OX Control Board exchanged tHb Kal function control i.o file activities and QA cycles for evaluation pulled from the analyzer 1 box 944-021 tHb Kal free of charge passed to the customer | CO-OX Control Board exchanged tHb Kal function control i.o file activities and QA cycles for evaluation pulled from the analyzer 1 box 944-021 tHb Kal free of charge passed to the customer
ABL90 FLEX | Vloeistofkanaal doorgespoten met demiwater. Pinchvalvetubing vervangen. Wastetubing gecontroleerd, geen lekkage. Instrument weer opgestart en laten calibreren. Alle drie controles gemeten, OK. | Vloeistofkanaal doorgespoten met demiwater. Pinchvalvetubing vervangen. Wastetubing gecontroleerd, geen lekkage. Instrument weer opgestart en laten calibreren. Alle drie controles gemeten, OK. | SWO-16567 | No | Liquid channel sprayed with demi water. Pinchvalvetubing replaced. Wastetubing checked, no leakage. Instrument will be restarted and calibrating. All three checks measured, OK. | Liquid channel sprayed with demi water. Pinchvalvetubing replaced. Wastetubing checked, no leakage. Instrument will be restarted and calibrating. All three checks measured, OK. | Cleaned | Fluidic pathway
ABL90 |  |  | CE807AF794611FD1 | No |  |  | Cleaned | Fluidic pathway
ABL90 FLEX | Das Labor meldet sehr häufig Abbrüche und inhomogene Probe bei der Patientenmessung.--> Mit dem Anwender folgende Probleme besprochen: Es wird überwiegend ohne Gerinselfänger gemessen Es werden 2 versch. Kapillaren benutzt ( Firma Hämacont und Radiometer ) Abbrüche sind prozentual nicht im Vergleich zum letzten Zählerstand gestiegen. Es gibt aber bei ca. 100 Messungen 10 Messungen mit inhomogener Probe. Nach Absprache mit dem Labor wurde entschieden, dass bis zum 22.02.2015 nur mit Kap. der Firma Radiometer gemessen wird. Weiterhin wird ein Gerinselfänger benutzt. Rückruf Besuch erfolgt durch Radiometer am 23.02.2015. | Das Labor meldet sehr häufig Abbrüche und inhomogene Probe bei der Patientenmessung.--> Mit dem Anwender folgende Probleme besprochen: Es wird überwiegend ohne Gerinselfänger gemessen Es werden 2 versch. Kapillaren benutzt ( Firma Hämacont und Radiometer ) Abbrüche sind prozentual nicht im Vergleich zum letzten Zählerstand gestiegen. Es gibt aber bei ca. 100 Messungen 10 Messungen mit inhomogener Probe. Nach Absprache mit dem Labor wurde entschieden, dass bis zum 22.02.2015 nur mit Kap. der Firma Radiometer gemessen wird. Weiterhin wird ein Gerinselfänger benutzt. Rückruf Besuch erfolgt durch Radiometer am 23.02.2015. | SWO-16602 | No | The laboratory very often reports cancellations and inhomogeneous sample in patient measurement.--> The following problems are discussed with the user: It is mainly measured without a device catcher. Capillaries used (company Hemacont and radiometer ) cancellations have not increased in percentage compared to the last meter reading. However, there are 10 measurements with inhomogeneous sample in about 100 measurements. After consultation with the laboratory, it was decided that until 22.02.2015 only with chap. measured by the company Radiometer. Furthermore, a Gerinsel catcher is used. Recall visit is made by radiometer on 23.02.2015. | The laboratory very often reports cancellations and inhomogeneous sample in patient measurement.--> The following problems are discussed with the user: It is mainly measured without a device catcher. Capillaries used (company Hemacont and radiometer ) cancellations have not increased in percentage compared to the last meter reading. However, there are 10 measurements with inhomogeneous sample in about 100 measurements. After consultation with the laboratory, it was decided that until 22.02.2015 only with chap. measured by the company Radiometer. Furthermore, a Gerinsel catcher is used. Recall visit is made by radiometer on 23.02.2015. | User Error | clot catcher
ABL90 | THIS FUNCTION WAS NOT ENABLED ON STANDARD CONFIGURATION. | CHANGED CONFIGURATION IN ROMA.INI FILE TO ENABLE results below Reportable Ranges will be shown as < x.xxx and result above Reportable Ranges will be shown as > x.xxx where x.xxx is the actually exceeded limit. | CE80A0D9940DCCAA | Yes | THIS FUNCTION WAS NOT ENABLED ON STANDARD CONFIGURATION. | CHANGED CONFIGURATION IN ROMA. INI FILE TO ENABLE results below Reportable Ranges will be shown as < x.xxx and result above Reportable Ranges will be shown as > x.xxx where x.xxx is the actually exceeded limit. | Re classified RMED | Setup by user request
ABL90 | THIS FUNCTION WAS NOT ENABLED ON STANDARD CONFIGURATION. | CHANGED CONFIGURATION IN ROMA.INI FILE TO ENABLE results below Reportable Ranges will be shown as < x.xxx and result above Reportable Ranges will be shown as > x.xxx where x.xxx is the actually exceeded limit. | CE80A2889442CCAD | Yes | THIS FUNCTION WAS NOT ENABLED ON STANDARD CONFIGURATION. | CHANGED CONFIGURATION IN ROMA. INI FILE TO ENABLE results below Reportable Ranges will be shown as < x.xxx and result above Reportable Ranges will be shown as > x.xxx where x.xxx is the actually exceeded limit. | Re classified RMED | Setup by user request
ABL800 BASIC | Flachbandkabelkabel von der Auto QC PCB zur Platine Lid Stellung def, Flachbandkabel bestellt und ersetzt, Kal und QC durchgeführt |  | SWO-16522 | No | Ribbon cable from the Auto QC PCB to the board eye position def, ribbon cable ordered and replaced, Kal and QC performed |  | Defect | HW | AutoCheck Module | FPC Cable
ABL800 BASIC | Proteinentfernung durchgeführt Lampeneinheit erneuert, Lampenabgleich durchgeführt Halogen Intensity im Complete Oxi Test 3330 tHb kalibriert (Fcuv 0,9849), Gerät überprüft | Proteinentfernung durchgeführt Lampeneinheit erneuert, Lampenabgleich durchgeführt Halogen Intensity im Complete Oxi Test 3330 tHb kalibriert (Fcuv 0,9849), Gerät überprüft | SWO-16546 | No | Protein removal performed lamp unit renewed, lamp matching performed halogen intensity in Complete Oxi Test 3330 tHb calibrated (Fcuv 0,9849), device checked | Protein removal performed lamp unit renewed, lamp matching performed halogen intensity in Complete Oxi Test 3330 tHb calibrated (Fcuv 0,9849), device checked | Defect | HW | Oxi module | Lamp unit
ABL835 FLEX |  |  | CE84680FF6077BC2 | No |  |  | Unknown | HW | Oxi module | Spectrometer | Mismatch
ABL827 FLEX |  |  | CE8070AF94199F42 | No |  |  | Assisted Replacement | Electrode | Electrode | pCO2
ABL825 FLEX | Multiple parameters out. Sample error on all parameters. Could not get the rinse to pass. Flushed everything and finnally found that the el met pump tube had a small leak in it. | Replaced the inlet probe. Replaced pump tube from customer stock. | CE80943C9424A1C5 | No | Multiple parameters out. Sample error on all parameters. Could not get the rinse to pass. Flushed everything and finnally found that the el met pump tube had a small leak in it. | Replaced the inlet probe. Replaced pump tube from customer stock. | Defect | HW | El/Met Module | Pump tubing | Leakage
ABL800 BASIC | Defekte Referenzelektrode ausgetauscht. Elektrodendeckel gereinigt. Na und Ca überprüft. Kalibrationen durchgeführt. Im Anschluss Halbjahreswartung Scannereinstellungendes FlexQ überprüft und nachjustiert (Position 2 nicht zuverlässig) | Defekte Referenzelektrode ausgetauscht. Elektrodendeckel gereinigt. Na und Ca überprüft. Kalibrationen durchgeführt. Im Anschluss Halbjahreswartung Scannereinstellungendes FlexQ überprüft und nachjustiert (Position 2 nicht zuverlässig) | SWO-16541 | No | Replaced defective reference electrode. Electrode lid cleaned. Na and Ca checked. calibrations. Following half-year maintenance, the FlexQ is checked and readjusted (position 2 not reliable) | Replaced defective reference electrode. Electrode lid cleaned. Na and Ca checked. calibrations. Following half-year maintenance, the FlexQ is checked and readjusted (position 2 not reliable) | Assisted Replacement | Electrode | Electrode | Reference
ABL815 FLEX | Schlauchweg am EL Met Modul zwischen der Messkammer und dem Rollenkranz korrekt angebracht. Prophylaktisch die Membranventile in den Messmodulen ausgebaut und gereinigt. Gerät kalibriert und QC durchgeführt. | Schlauchweg am EL Met Modul zwischen der Messkammer und dem Rollenkranz korrekt angebracht. Prophylaktisch die Membranventile in den Messmodulen ausgebaut und gereinigt. Gerät kalibriert und QC durchgeführt. | SWO-16620 | No | Hose path on the EL Met module between the measuring chamber and the roller ring correctly attached. Prophylactic, the diaphragm valves in the measuring modules are removed and cleaned. calibrated and QC performed. | Hose path on the EL Met module between the measuring chamber and the roller ring correctly attached. Prophylactic, the diaphragm valves in the measuring modules are removed and cleaned. calibrated and QC performed. | Unknown | Electrode | Electrode | Glu
ABL820 FLEX | pCO2 reading on the low end of qc range. Found a small droplet of rinse solution being left in reference chamber during analysis. | Replaced reference and pCO2 membrane. | CE808D969456E599 | No | pCO2 reading on the low end of qc range. Found a small droplet of rinse solution being left in reference chamber during analysis. | Replaced reference and pCO2 membrane. | Assisted Replacement | Electrode | Membrane | pCO2
ABL820 FLEX | pCO2 reading on the low end of package insert. | Replaced electrode | CE80DC0B940A74F1 | No | pCO2 reading on the low end of package insert. | Replaced electrode | Assisted Replacement | Electrode | Electrode | pCO2
ABL825 FLEX | Customer rebooted and got error code 0734 | Replacced power supply but did not resolve | CE80825F94360191 | No | Customer rebooted and got error code 0734 | Replacced power supply but did not resolve | Unknown | HW | Power Supply
ABL820 FLEX | Touch screen out of calibration | Recalibrated touch screen, reloaded 6.13 software. Touch screen now works. | CE808930940333C7 | No | Touch screen out of calibration | Recalibrated touch screen, reloaded 6.13 software. Touch screen now works. | Adjustment | HW | DMS_PC | Touch screen
ABL820 FLEX | Reviewed QC data logs and verifed that incorrect QC levels were run at the auto scheduled times. | Reload of version 6.13 software. Uploaded QC schedule from working ABL 820 on site. Correct QC ampule was run at auto scheduled time of 2pm. | CE808AC7943D33CA | No | Reviewed QC data logs and verifed that incorrect QC levels were run at the auto scheduled times. | Reload of version 6.13 software. Uploaded QC schedule from working ABL 820 on site. Correct QC ampule was run at auto scheduled time of 2pm. | Setup | QC Schedule
ABL800 BASIC | Pumpen Kal und LS Justage durchgeführt. Gerät kalibriertund QC durchgeführt. | Pumpen Kal und LS Justage durchgeführt. Gerät kalibriertund QC durchgeführt. | SWO-16623 | No | Pumps Kal and LS adjustment performed. device calibrated and QC performed. | Pumps Kal and LS adjustment performed. device calibrated and QC performed. | User Error | Sample volume
AQT90 FLEX | Mechanisch def CRP Kassette Lot 10091 beim Punch Out hängengeblieben, Kassette entfernt, QC durchgeführt | Mechanisch def CRP Kassette Lot 10091 beim Punch Out hängengeblieben, Kassette entfernt, QC durchgeführt | SWO-16761 | No | Mechanical def CRP Cassette Lot 10091 Stuck at Punch Out, Cassette Removed, QC Performed | Mechanical def CRP Cassette Lot 10091 Stuck at Punch Out, Cassette Removed, QC Performed
ABL90 | Customerchanged SP and back flushed-No change Would like service on Monday I found the problem to be either the parts I changed or the SC | Changed listed parts In service the problem seemed to follow the SC but left on because 267 tests left and all OK now Customer will change SC if problem returns | CE843545F654C2F4 | No | Customerchanged SP and back flushed-No change Would like service on Monday I found the problem to be either the parts I changed or the SC | Changed listed parts In service the problem seemed to follow the SC but left on because 267 tests left and all OK now Customer will change SC if problem returns | Assisted Replacement | SC | Wetability
ABL90 FLEX | Gerät manuell gespült, Einlassanschlussdichtung erneuert Visual Flow Check im Service durchgeführt, kein Lösungstransport neuen Lösungspack installiert, Sensorkassette erneut installiert Gerät kalibriert und überprüft | Gerät manuell gespült, Einlassanschlussdichtung erneuert Visual Flow Check im Service durchgeführt, kein Lösungstransport neuen Lösungspack installiert, Sensorkassette erneut installiert Gerät kalibriert und überprüft | SWO-16711 | No | Device manually rinsed, inlet connection seal renewed Visual Flow Check carried out in service, no solution transport new solution pack installed, sensor cartridge reinstalled device calibrated and checked | Device manually rinsed, inlet connection seal renewed Visual Flow Check carried out in service, no solution transport new solution pack installed, sensor cartridge reinstalled device calibrated and checked | Cleaned | Fluidic pathway
ABL90 | Found cause to be defective battery pack. When pack was removed unit operated normally. | Replaced battery pack | CE846B8BF61C536D | No | Found cause to be defective battery pack. When pack was removed unit operated normally. | Replaced battery pack | Defect | HW | Battery module
ABL825 FLEX | Elektroden Kontakte gereinigt. Messkammer gereinigt. Kalinration mehrfach durchgeführt. Messkammer der PH- BG Modul ausgetauscht. Alle Elektroden neu bezogen | Frau Buch geleitet und die Messkammern gereinigt. | SWO-16706 | No | Electrode contacts cleaned. measuring chamber. calinration performed several times. measuring chamber of the PH- BG module. All electrodes re-sourced | Mrs. Buch and cleaned the measuring chambers. | Cleaned | HW | pH/BG Module | electrode connector | Reference
ABL825 FLEX |  |  | CE845495F60B8D55 | No |  |  | Unknown | HW | Oxi module | HBF
ABL835 FLEX |  |  | CE846A12F6507BC3 | No |  |  | Unknown | HW | Oxi module | Unknown
ABL825 FLEX | Inspected cabling, Module PCB, reload software,Module appears to be offline, Service mode EI Met will not respond. | Ordered AU Bd. and EI Met Mod PCB. | CE849888F622E658 | No | Inspected cabling, Module PCB, reload software,Module appears to be offline, Service mode EI Met will not respond. | Ordered AU Bd. and EI Met Mod PCB. | Defect | HW | - | Unknown
ABL825 FLEX | Returned to site with part | Replaced AU PCB. tried to load AU software but would not download. | CE8436FCF650C1A2 | No | Returned to site with part | Replaced AU PCB. tried to load AU software but would not download. | Unknown | HW
AQT90 FLEX |  |  | CE8479E0F64E60D7 | No
AQT90 FLEX |  |  | CE86BC4EE20DAFA0 | No
TCM/Tosca |  |  | CE841D2FF65428B8 | No
ABL90 | C3 has inhomogenous failure. | Service test mode, did 2 clot removal cycles, vacuum test cycle. Did solution pack check. Retested SP. Rassed. Unit went through internal QC. All levels passed. Verified operation of ABL90. | CE858DAB6C00BEE5 | No | C3 has inhomogenous failure. | Service test mode, did 2 clot removal cycles, vacuum test cycle. Did solution pack check. Retested SP. Rassed. Unit went through internal QC. All levels passed. Verified operation of ABL90. | Cleaned | Fluidic pathway
ABL90 |  |  | CE85B6126C36310C | Yes |  |  | Re classified RMED | Removed by RSUB
ABL90 FLEX | Oximetrie unit heb opnieuwe afgesteld en daarbij het volgende heb uitgevoerd: - Complete oximetrie test - 2-pnut Calibratie - tHb calibratie - 2-punt Calibratie Hierna heb vier QC metingen S9040 en S9050 heb uitgevoerd, OK. | Oximetrie unit heb opnieuwe afgesteld en daarbij het volgende heb uitgevoerd: - Complete oximetrie test - 2-pnut Calibratie - tHb calibratie - 2-punt Calibratie Hierna heb vier QC metingen S9040 en S9050 heb uitgevoerd, OK. | SWO-16999 | No | Oximetry unit have retuned and performed the following:-Complete oximetry Test-2-pnut calibration-tHb Calibration-2-point calibration after that, four QC measurements S9040 and S9050 have performed, OK. | Oximetry unit have retuned and performed the following:-Complete oximetry Test-2-pnut calibration-tHb Calibration-2-point calibration after that, four QC measurements S9040 and S9050 have performed, OK. | Adjustment | Oximetry module | tHb
ABL90 FLEX | THb Kal durchgeführt. | THb Kal durchgeführt. | SWO-16795 | No | THb Kal. | THb Kal. | Adjustment | tHb calibration
ABL90 FLEX | System überprüft, tHb Kalibration durchgeführt, Vergleichsmessungen durchgeführt, kein Fehler feststellbar. |  | SWO-16770 | No | System checked, tHb calibration carried out, comparative measurements carried out, no error detectable. |  | Unknown | No errors found
ABL827 FLEX | REVISIÓN GENERAL DE TODO EL SISTEMA CON LIMPIEZA Y DESCONTAMINACION DE CIRCUITOS. ENGRASE Y LIMPIEZA DE PIEZAS MOVILES Y ENGRANAJES. COMPROBACIONES DE CONECTIVIDAD HIS SIL. BACKUP DE CONFIGURACIONES DEL SISTEMA. LIMPIEZA Y O SUSTITUCION DE FILTROS. TEST DE FUNCIONAMIENTO ELECTRÓNICO Y NEUMÁTICO. SUTITUCION DE PIEZAS DETERIORADAS POR USO. CALIBRAR Y VERIFICAR MEDIANTE CONTROL DE CALIDAD Q.C. | PCB INLET | CE8570696C1D5D38 | No | GENERAL REVIEW OF ALL SYSTEM WITH CLEANING AND DISCONTAMINATION OF CIRCUITS. WELCOME AND CLEANING OF MOVIL PARTS AND GRAINS. HIS SIL CONNECTIVITY CHECKS. BACKUP OF SYSTEM SETTINGS. CLEANING AND REPLACEMENT OF FILTERS. ELECTRONIC AND TIRE OPERATION TEST. REPLACEMENT OF DETERIORATED PARTS BY USE. CALIBRATE AND CHECK BY Q.C QUALITY CONTROL. | PCB INLET | Defect | HW | Inlet Module | PCB
ABL815 FLEX | Fuite d inlet tubing percé et waste line bouché | remplacé joint introduction inlet tubing et nettoyé waste line a partir inlet et nettoyé V7,V8, installé nouveau o'ring PH et nouvelle membrane de PC02 | CE8551816C4C2161 | No | Breakthrough tubing inlet leak and clogged waste line | replaced joint introduction inlet tubing and cleaned waste line from inlet and cleaned V7,V8, installed new o'ring PH and new PC02 membrane | Defect | HW | Inlet Module | Tubing | Leakage
ABL835 FLEX |  |  | CE8E433CA61BC637 | No |  |  | Defect | HW | Oxi module
ABL825 FLEX | Checked electrode chamber, checked zero calibration and cleaned original electrode. | Replaced Lactate electrode. | CE8604D56C164A42 | No | Checked electrode chamber, checked zero calibration and cleaned original electrode. | Replaced Lactate electrode. | Assisted Replacement | Electrode | Electrode | Lac
ABL835 FLEX | N A | REPLACED INLET CLIP | CE85AB016C01A61E | No | N A | REPLACED INLET CLIP | Defect | HW | Inlet module | Probe clip
ABL800 BASIC | Restliche Finale Keys installiert. Update Type. Abschaltung Demokeys. Initialisierung. Kal i.O. | Restliche Finale Keys installiert. Update Type. Abschaltung Demokeys. Initialisierung. Kal i.O. | SWO-16771 | Yes | Remaining Final Keys installed. Update Type. Shutdown Demokeys. Initialization. Kal i.O. | Remaining Final Keys installed. Update Type. Shutdown Demokeys. Initialization. Kal i.O. | Re classified RMED | Demo key exp
ABL837 FLEX | checked data logs for errors, checked hemolyzer and lamp unit, checked IP address, checked connections | replaced lamp unit and was not able to adjust lamp voltage above 2900, replaced lamp unit with original lamp, lamp voltage would not adjust above 3000, lamp voltage within specifications, connected analyzer to another port, instrument esatblished connection | CE86163F6C0A58D3 | No | checked data logs for errors, checked hemolyzer and lamp unit, checked IP address, checked connections | replaced lamp unit and was not able to adjust lamp voltage above 2900, replaced lamp unit with original lamp, lamp voltage would not adjust above 3000, lamp voltage within specifications, connected analyzer to another port, instrument esatblished connection | Adjustment | HW | Oxi module | Lamp unit
ABL825 FLEX | Checked power supply for problems. It passed. | Replaced the AU control board. Performed all necessary calibrations. | CE8598CF6C0A3246 | No | Checked power supply for problems. It passed. | Replaced the AU control board. Performed all necessary calibrations. | Defect | HW | Analyzer Control PCB
ABL80 CO-OX | Customer said clear fluid was leaking. Also found some internal qc leaked in the rear of analyzer. | Replaced Manifold. Replaced all parts from PM kit. | CE86D908E256E41F | No | Customer said clear fluid was leaking. Also found some internal qc leaked in the rear of analyzer. | Replaced Manifold. Replaced all parts from PM kit.
ABL90 FLEX | Beide Schläuche getauscht, wasteconnector gereinigt, manuell gespült, Kaltstartphase, | Beide Schläuche getauscht, wasteconnector gereinigt, manuell gespült, Kaltstartphase, | SWO-17092 | No | Both hoses are replaced, wasteconnector cleaned, manually rinsed, cold start phase, | Both hoses are replaced, wasteconnector cleaned, manually rinsed, cold start phase, | Cleaned | Fluidic pathway | replaced tubes
ABL90 |  |  | CE893FABCE4F5CFC | No |  |  | Unknown | SP
ABL90 FLEX | Gerät mangels Ersatzteile zur Rep. mitgenommen. | Gerät mangels Ersatzteile zur Rep. mitgenommen. | SWO-17053 | No | Device due to lack of spare parts for Rep. Taken. | Device due to lack of spare parts for Rep. Taken. | Assisted Replacement | SP
ABL90 FLEX | Allgemeine Kontrolle ok. Softwareupdate auf Version 3.1 MR3 ok. System i.o. Kal. + QC i.o. | Allgemeine Kontrolle ok. Softwareupdate auf Version 3.1 MR3 ok. System i.o. Kal. + QC i.o. | SWO-17040 | Yes | General control ok. Software update to version 3.1 MR3 ok. System i.o. Kal. + QC i.o. | General control ok. Software update to version 3.1 MR3 ok. System i.o. Kal. + QC i.o. | Re classified RMED | SWU
ABL825 FLEX | Gerät kontrolliert, Elektroden kontrolliert, pCO2 Elektrode abgewaschen, mehrfach membraniert, Gasventilmembran pH BG Modul ersetzt, Oximetrieabgleich durchgeführt | Gerät kontrolliert, Elektroden kontrolliert, pCO2 Elektrode abgewaschen, mehrfach membraniert, Gasventilmembran pH BG Modul ersetzt, Oximetrieabgleich durchgeführt | SWO-17104 | No | Device controlled, electrodes controlled, pCO2 electrode washed off, multiple membranes, gas valve membrane pH BG module replaced, oximetry matching performed | Device controlled, electrodes controlled, pCO2 electrode washed off, multiple membranes, gas valve membrane pH BG module replaced, oximetry matching performed | Defect | HW | pH/BG Module | Valve, air
ABL825 FLEX | error 210 calibration errors or measurement unstable on intermittent blood gases | Replaced both Lactate and Glucose membranes with new lot # for troubleshooting purposes and cleaned electrode | CE86D57EE258DF57 | No | error 210 calibration errors or measurement unstable on intermittent blood gases | Replaced both Lactate and Glucose membranes with new lot # for troubleshooting purposes and cleaned electrode | Assisted Replacement | Electrode | Membrane | Lac
ABL825 FLEX | Module has no activity. Part arrival delayed due to Bad weather in Memphis(FedEx). | Replaced EI Met module PCB. | CE86D616E22DF9A7 | No | Module has no activity. Part arrival delayed due to Bad weather in Memphis(FedEx). | Replaced EI Met module PCB. | Defect | HW | El/Met Module | PCB
ABL825 FLEX | Returned to site with parts | Replaced AU board and power supply. Downloaded AU software. Downloaded option keys | CE86D8FCE242EE1A | No | Returned to site with parts | Replaced AU board and power supply. Downloaded AU software. Downloaded option keys | Defect | HW | Analyzer Control PCB
ABL835 FLEX | Oximetriemodul repariert, Justagen, Kalibrationen und Qualitätskontrollen durchgeführt. | Oximetriemodul repariert, Justagen, Kalibrationen und Qualitätskontrollen durchgeführt. | SWO-17070 | No | Oximetry module repaired, adjustments, calibrations and quality controls carried out. | Oximetry module repaired, adjustments, calibrations and quality controls carried out. | Defect | HW | Oxi module
ABL820 FLEX | Isolation test in service shows hiher the 500 mV | Found waste valve stuck open. Remembraned all valves. Put fluidic module from old ABL, this solved the problem. Will order new PCB board for the fluidic module. | CE8710B1E2420813 | No | Isolation test in service shows hiher the 500 mV | Found waste valve stuck open. Remembraned all valves. Put fluidic module from old ABL, this solved the problem. Will order new PCB board for the fluidic module. | Defect | HW | Fluidic manifold | Valve V8
ABL820 FLEX | Review QC log and confirmed reported complaint. performed visual inspection and ran service diagnostics. No failures noted in diagnostics. Inspected all electrodes including pCO2 and pH. electrodes | Replaced rinse and air gas valves in pH BG module, replace black tubing from air gas valve to gas. Replaced blue sealing ring for pH electrode, noted and cleaned a small streak of blood on the end of the electrode inside the area the blue seal covers at the bottom of the electrode. Cleaned mesuring chamber as precautionary measure. Remembraned Ref electrode. | CE86C581E226851A | No | Review QC log and confirmed reported complaint. performed visual inspection and ran service diagnostics. No failures noted in diagnostics. Inspected all electrodes including pCO2 and pH. electrodes | Replaced rinse and air gas valves in pH BG module, replace black tubing from air gas valve to gas. Replaced blue sealing ring for pH electrode, noted and cleaned a small streak of blood on the end of the electrode inside the area the blue seal covers at the bottom of the electrode. Cleaned mesuring chamber as precautionary measure. Remembraned Ref electrode. | Assisted Replacement | Electrode | Membrane | Electrode
AQT90 FLEX |  |  | CE8712B6E2277D04 | No
ABL80 FLEX CO-OX | Alle Schläche ersetzt, Gerät zerlegt und gereinigt, Kalibration, QC Kontrolle Leider pO2 zu hoch in der QC, daher per Redline Lösungspack und Sensor bestellt | Alle Schläche ersetzt, Gerät zerlegt und gereinigt, Kalibration, QC Kontrolle Leider pO2 zu hoch in der QC, daher per Redline Lösungspack und Sensor bestellt | SWO-17381 | No | Replaced all sleeps, device disassembled and cleaned, calibration, QC control Unfortunately pO2 too high in the QC, therefore ordered by Redline solution pack and sensor | Replaced all sleeps, device disassembled and cleaned, calibration, QC control Unfortunately pO2 too high in the QC, therefore ordered by Redline solution pack and sensor
ABL80 CO-OX |  |  | CE87FCAE5858D375 | No
ABL90 FLEX | 2lösungspack geliefert, | 2lösungspack geliefert, | SWO-17296 | Yes | 2 solution pack delivered, | 2 solution pack delivered, | Re classified RMED | Logistical upSWO
ABL90 FLEX | Mit Fr. Bertram-Bruns die Pat. Berichte durchsucht, es werden auffällig viele Kap.BGAs durch Probenprobleme abgebrochen. Das Gerät führt häufig die Reinigung durch, es werden Gerinnsel erkannt. Mit der Kundin das manuelle Spülverfahren durchgeführt. 1VE 942-892 zur Probe da gelassen. Einlassdichtungen mit geliefert. | Verschiedene Ursachen abgefragt. Es werden keine -R- Probennehmer verwendet. Dazu wurden alle Einlassdichtungen aufgebraucht. Der Kundin einen Service am nächsten Tag und die Belieferung mit ZVBM zugesagt. | SWO-17365 | No | With Ms. Bertram-Bruns the Pat. Reports are searched, a striking number of Kap.BGAs are aborted by rehearsal problems. The device often performs the cleaning, clots are detected. With the customer the manual rinsing procedure is carried out. 1VE 942-892 for rehearsal. Inlet seals supplied. | Various causes queried. No -R samplers are used. For this purpose, all inlet seals were used up. The customer promised a service the next day and the delivery with ZVBM. | Unknown
ABL90 |  |  | CE8809B1581026E6 | No |  |  | Unknown | -
ABL90 |  |  | CE8D1EEC3062A04A | Yes |  |  | Re classified RMED | Logistical upSWO
ABL90 |  |  | CE8D1FE4304AA04F | Yes |  |  | Re classified RMED | Logistical upSWO
ABL90 |  |  | CE880B52580367C4 | No |  |  | Assisted Replacement | SC
ABL825 FLEX | Fehlersuche>>>>>Ref-Elektrode ausgetauscht. Kontakt überprüft. Alle Elektroden in der Messkammer neu bezogen. PH-BG pcb. ausgetauscht. | Fehlersuche>>>>>Ref-Elektrode ausgetauscht. Kontakt überprüft. Alle Elektroden in der Messkammer neu bezogen. PH-BG pcb. ausgetauscht. | SWO-17301 | No | Troubleshooting>>>>>Ref electrode replaced. Contact checked. All electrodes in the measuring chamber are re-sourced. PH-BG pcb. Exchanged. | Troubleshooting>>>>>Ref electrode replaced. Contact checked. All electrodes in the measuring chamber are re-sourced. PH-BG pcb. Exchanged. | Defect | HW | pH/BG Module | PCB
ABL800 BASIC | Austausch der E788 danach Kal wieder stabil | Austausch der E788 danach Kal wieder stabil | SWO-17280 | No | Replacement of the E788 then Kal stable again | Replacement of the E788 then Kal stable again | Assisted Replacement | Electrode | Electrode | pCO2
ABL820 FLEX | Confirmed reported issued. Determined cause to be clogged Blue vent port for V7 on fluidic module. | Flushed cleaned Blue port and V7. Ran service diagnostics, no other issues reported. | CE8857355800C604 | No | Confirmed reported issued. Determined cause to be clogged Blue vent port for V7 on fluidic module. | Flushed cleaned Blue port and V7. Ran service diagnostics, no other issues reported. | Cleaned | HW | Fluidic manifold | Vent channel
ABL800 BASIC | Reinigen Einlaß, Fließwege, Verbogene Einlaßnadel gerichtet | Reinigen Einlaß, Fließwege, Verbogene Einlaßnadel gerichtet | SWO-17294 | No | Cleaning inlet, flow paths, bent inlet needle directed | Cleaning inlet, flow paths, bent inlet needle directed | Defect | HW | Inlet Module | Probe bend
ABL800 BASIC | Fehlersuche Fließwege gereinigt Verbogene Einlaßnadel gerichtet | Fehlersuche Fließwege gereinigt Verbogene Einlaßnadel gerichtet | SWO-17266 | No | Troubleshooting flow paths cleaned bent inlet needle directed | Troubleshooting flow paths cleaned bent inlet needle directed | Defect | HW | Inlet Module | Probe bend
ABL835 FLEX |  |  | CE8E43A1A60CC639 | No |  |  | Defect | HW | Oxi module
ABL800 BASIC | Fehlersuche Spüllösung erneuert, obwohl noch ca 15 % gefüllt Pumpenkalibrierung., LS justage Demontage und Wartung Ventile am BG und EL Modul Einlaß incl. Nadel gereinigt Cl membraniert | Fehlersuche Spüllösung erneuert, obwohl noch ca 15 % gefüllt Pumpenkalibrierung., LS justage Demontage und Wartung Ventile am BG und EL Modul Einlaß incl. Nadel gereinigt Cl membraniert | SWO-17265 | No | Troubleshooting rinsing solution renewed, although still about 15% filled pump calibration,, LS adjustment disassembly and maintenance valves on the BG and EL module inlet incl. needle cleaned clmembrane | Troubleshooting rinsing solution renewed, although still about 15% filled pump calibration,, LS adjustment disassembly and maintenance valves on the BG and EL module inlet incl. needle cleaned clmembrane | Cleaned | Fluidic pathway
ABL825 FLEX | Fließwege gespült, Ventile überprüft | Herr Dürr stellte nach Meßkammerreinigung fest, daß der Schlauch der Einlassnadel nicht korrekt befestigt war. Bei Messung keine Fehler mehr. Tropft noch vereinzelt aus Kapillareinlass. Herr Dürr wünscht Besuch durch zuständigen Techniker. | SWO-17293 | No | Flow paths rinsed, valves checked | After cleaning the measuring chamber, Mr Dürr found that the hose of the inlet needle was not properly attached. No more errors when measured. Drips sporadically from capillary inlet. Mr Dürr wishes to see the technicians responsible. | Defect | HW | Inlet Module | Tubing, blue
ABL820 FLEX | Verified problem via WDC reports. WDC report shows that December was on the edge and January was out. Checked electrode for air bubbles, none present. Checked for o-ring, present. Remembraned pCO2 electrode, out even higher on level 3. | Installed service electrode and pCO2 now reads well within range @ 13.2 with the mean @ 12.7. | CE8823345858A6C4 | No | Verified problem via WDC reports. WDC report shows that December was on the edge and January was out. Checked electrode for air bubbles, none present. Checked for o-ring, present. Remembraned pCO2 electrode, out even higher on level 3. | Installed service electrode and pCO2 now reads well within range @ 13.2 with the mean @ 12.7. | Assisted Replacement | Electrode | Membrane | pCO2
ABL835 FLEX |  |  | CE8E476DA61BC64E | No |  |  | Defect | HW | Oxi module | Oxi module
ABL825 FLEX | Replace inlet probe and gasket did not correct problem. Fails inlet calibration in service mode. | Replace friction detector unit. Install software upgrade 6.14. | CE882B54585A6426 | No | Replace inlet probe and gasket did not correct problem. Fails inlet calibration in service mode. | Replace friction detector unit. Install software upgrade 6.14. | Defect | HW | Inlet Module | Friction detector unit
ABL825 FLEX | Jeg har skiftet printer og nu kan der printes igen. Alt er OK. | Jeg har skiftet printer og nu kan der printes igen. Alt er OK. | SWO-17373 | No | I changed the printer and now you can print again. Everything is OK. | I changed the printer and now you can print again. Everything is OK. | Defect | HW | Printer module | Printer Head
AQT90 FLEX | Öffnen des Analysators. Nachjustieren des Barcodescanners, Fläche gereinigt. Abscannen aller Positionen erfolgreich. Initialisierung. Betriebsmittelstände werden komplett angezeigt. Analysator Messbereit | Öffnen des Analysators. Nachjustieren des Barcodescanners, Fläche gereinigt. Abscannen aller Positionen erfolgreich. Initialisierung. Betriebsmittelstände werden komplett angezeigt. Analysator Messbereit | SWO-17308 | No | Open the analyzer. Readjustment of the barcode scanner, surface cleaned. Scan all positions successfully. Initialization. Equipment levels are displayed in full. Analyzer Ready to Measure | Open the analyzer. Readjustment of the barcode scanner, surface cleaned. Scan all positions successfully. Initialization. Equipment levels are displayed in full. Analyzer Ready to Measure
ABL80 OSM | The analyzer keeps presenting errors 1100, 1101 and, 1035 related to solution pack errors. | Replaced the opto sensor and the SP door latch. Also, replaced the solution pack due to the errors described. The O2HB, COHB and, METHB manual qc ranges were incorrect, showed the customer how to adjust them properly . Note: The customer has replaced two additional solution packs due to the above listed errors. I wil submit 3 credit forms for the customer. | CE8917F6CE46EE39 | No | The analyzer keeps presenting errors 1100, 1101 and, 1035 related to solution pack errors. | Replaced the opto sensor and the SP door latch. Also, replaced the solution pack due to the errors described. The O2HB, COHB and, METHB manual qc ranges were incorrect, showed the customer how to adjust them properly . Note: The customer has replaced two additional solution packs due to the above listed errors. I wil submit 3 credit forms for the customer.
ABL90 FLEX | Stark verschmutzte Einlassdichtung durch Kundeneigene ersetzt. Gerät kalibriert und QC durchgeführt. Kal OK L3 Glu unterhalb unterer QC Grenze. Lösungspack erneuert, QC erneut durchgeführt. Sensor Kassette erneuert, Gerät mehrfach kalibriert und QC durchgeführt. | Stark verschmutzte Einlassdichtung durch Kundeneigene ersetzt. Gerät kalibriert und QC durchgeführt. Kal OK L3 Glu unterhalb unterer QC Grenze. Lösungspack erneuert, QC erneut durchgeführt. Sensor Kassette erneuert, Gerät mehrfach kalibriert und QC durchgeführt. | SWO-17478 | No | Highly soiled inlet seal replaced by customer's own. calibrated and QC performed. Kal OK L3 Glu below lower QC limit. Solution pack renewed, QC performed again. Sensor cassette renewed, device repeatedly calibrated and QC performed. | Highly soiled inlet seal replaced by customer's own. calibrated and QC performed. Kal OK L3 Glu below lower QC limit. Solution pack renewed, QC performed again. Sensor cassette renewed, device repeatedly calibrated and QC performed. | Assisted Replacement | SC | Cleaned FT
ABL90 |  |  | CE891AF1CE507C24 | No |  |  | Defect | HW | Waste tube
ABL90 FLEX | Flüssigkeitsstrecke gespült SP und Einlassdichtung getauscht Neustart | Flüssigkeitsstrecke gespült SP und Einlassdichtung getauscht Neustart | SWO-17438 | No | Liquid line flushed SP and inlet seal exchanged restart | Liquid line flushed SP and inlet seal exchanged restart | Cleaned | Fludic pathway
ABL90 | Inspected sensor casette for leakage and none was found. | Performed a back flush on system and conditioned the cassette. | CE8930E8CE04213C | No | Inspected sensor casette for leakage and none was found. | Performed a back flush on system and conditioned the cassette. | Cleaned | Fluidic pathway
ABL827 FLEX | SUSTITUIR EL MATERIAL ABAJO DESCRITO Y PRUEBAS DE FUNCIONAMIENTO Y CALIBRACIONES. CALIBRADO Y VERIFICADO CON Q.C. | SUSTITUCION THERMISTORES | CE8992E5CE3A5623 | No | SUBSTITUTE THE MATERIAL DESCRIBED BELOW AND TESTS OF OPERATION AND CALIBRATIONS. CALIBRATED AND VERIFIED WITH Q.C. | THERMISTORES SUBSTITUTE | Defect | HW | El/Met Module | Thermistor
ABL820 FLEX | New pCO2 eletrode needs installed. | Replaced pCO2 electrode under warranty. | CE894C6DCE0FA6C7 | No | New pCO2 eletrode needs installed. | Replaced pCO2 electrode under warranty. | Assisted Replacement | Electrode | Electrode | pCO2
AQT90 FLEX | Sustituido inlet por rotura de la misma. Puesta protección en la pipette. Comprobado buen funcionamiento del equipo. |  | CE896108CE3CB2AB | No | Replaced inlet by breakage. Putting protection on the pipette. Proven good operation of the equipment.
AQT90 FLEX |  |  | CE8D3F20301FF90C | No
AQT90 FLEX | defekten Sensor ersetzt, Gerät zusammengesetzt, Neustart abgewartet, Puffermessungen, QC |  | SWO-17459 | No | defective sensor replaced, device assembled, restart awaited, buffer measurements, QC
ABL90 |  |  | CE8A83034427408C | Yes |  |  | Re classified RMED | No info
ABL90 FLEX | Akku ersetzt, Notstromgerät abgebaut | Akku ersetzt, Notstromgerät abgebaut | SWO-17495 | No | Battery replaced, emergency power unit dismantled | Battery replaced, emergency power unit dismantled | Defect | HW | Battery module
ABL80 FLEX CO-OX | Mit neuer Sensorkassette, pO2 Fehler in QC daher gebrauchte SC wider verwendet. Nachdem das Gerät im Büro einwandfrei gelaufen ist wurde nach Transport ca 30min Sensor Schaltkreisfehler angezeigt. Sensor reinitialisiert nun kal. Und QC im Bereich | Mit neuer Sensorkassette, pO2 Fehler in QC daher gebrauchte SC wider verwendet. Nachdem das Gerät im Büro einwandfrei gelaufen ist wurde nach Transport ca 30min Sensor Schaltkreisfehler angezeigt. Sensor reinitialisiert nun kal. Und QC im Bereich | SWO-17570 | No | With new sensor cartridge, pO2 error in QC therefore used SC against. After the device has run flawlessly in the office, after transport about 30min sensor circuit error was displayed. Sensor now reinitializes kal. And QC in the field | With new sensor cartridge, pO2 error in QC therefore used SC against. After the device has run flawlessly in the office, after transport about 30min sensor circuit error was displayed. Sensor now reinitializes kal. And QC in the field
ABL80 CO-OX | Comprobación aspirado correcto con varias muestras. Se realiza trabajo pendiente sin Na K Ca. Muestras OK. Sustitución SP (acabado). Ciclo de calidad sigue dando fallo en citados parámetros. Sustitución SC. Ciclo de calidad. OK |  | CE8D13EF300E9AAE | No | Correct vacuum check with multiple samples. Pending work is done without Na K Ca. Samples OK. SP substitution (finishing). Quality cycle continues to fail in these parameters. SC substitution. Quality cycle. Ok
ABL90 FLEX | Nach einer Woche bei 90 Proben treten folgende Fehlmessungen auf: 2x inhomogen 7 x ungenügend 6 x Abbruch Lt. Datenbank wurden 53 Probenmessungen am Stück ohne Fehler gemessen Es erfolgte eine Rückmeldung an Hr. Fordran Die Jahreswartung wird zeitnah durchgeführt | Nach einer Woche bei 90 Proben treten folgende Fehlmessungen auf: 2x inhomogen 7 x ungenügend 6 x Abbruch Lt. Datenbank wurden 53 Probenmessungen am Stück ohne Fehler gemessen Es erfolgte eine Rückmeldung an Hr. Fordran Die Jahreswartung wird zeitnah durchgeführt | SWO-17566 | No | After one week with 90 samples the following incorrect measurements occur: 2x inhomogeneous 7 x insufficient 6 x demolition Lt. Database 53 sample measurements were measured at a time without errors. | After one week with 90 samples the following incorrect measurements occur: 2x inhomogeneous 7 x insufficient 6 x demolition Lt. Database 53 sample measurements were measured at a time without errors. | User Error
ABL90 FLEX | Akku ausgetauscht. Zeiteinstellung. Nach dem SW Update auf 3.1 MR3 wurde das Gerät automatisch herunter gefahren und nicht mehr gebootet. Nur ein lang anhaltender Peepton. Akku wieder ausgebaut. | Akku ausgetauscht. Zeiteinstellung. Nach dem SW Update auf 3.1 MR3 wurde das Gerät automatisch herunter gefahren und nicht mehr gebootet. Nur ein lang anhaltender Peepton. Akku wieder ausgebaut. | SWO-17561 | No | Battery replaced. Time setting. After the SW update to 3.1 MR3, the device was automatically shut down and no longer booted. Just a long-lasting peepton. Battery is removed again. | Battery replaced. Time setting. After the SW update to 3.1 MR3, the device was automatically shut down and no longer booted. Just a long-lasting peepton. Battery is removed again. | Defect | HW | Battery module
ABL90 FLEX | Verbogene Nadel ausgetauscht. Testprogramme im Servicemenü durchgeführt. Drucktest bestanden. Kalibration und QC ohne Beanstandung | Verbogene Nadel ausgetauscht. Testprogramme im Servicemenü durchgeführt. Drucktest bestanden. Kalibration und QC ohne Beanstandung | SWO-17558 | No | Bent needle replaced. test programs in the service menu. Print test passed. Calibration and QC without complaint | Bent needle replaced. test programs in the service menu. Print test passed. Calibration and QC without complaint | Replaced | Probe | Bend
ABL90 FLEX | Konfiguration manuell vom Leihgerät übertragen. Scanner auf 2 von 5 eingestellt. Fixe Zielwerte eingestellt. ABL ans Radiance angebunden. | Konfiguration manuell vom Leihgerät übertragen. Scanner auf 2 von 5 eingestellt. Fixe Zielwerte eingestellt. ABL ans Radiance angebunden. | SWO-17615 | No | manually transfer configuration from the loaner. Scanner set to 2 of 5. Fixed target values set. ABL connected to the radiance. | manually transfer configuration from the loaner. Scanner set to 2 of 5. Fixed target values set. ABL connected to the radiance. | Defect | HW | Battery module
ABL90 |  |  | CE8CF69B302A1500 | No |  |  | Defect | SW | Reboot
ABL90 FLEX | Flüssigkeitswege wurden durch Medizintechniker bereits gespült. Oxiventil öffnet verzögert, geölt. Software auf Version 3.1 MR3 aktualisiert. | Flüssigkeitswege wurden durch Medizintechniker bereits gespült. Oxiventil öffnet verzögert, geölt. Software auf Version 3.1 MR3 aktualisiert. | SWO-17563 | No | Fluid pathways have already been flushed by medical technicians. Oxivalve opens delayed, oiled. Software updated to version 3.1 MR3. | Fluid pathways have already been flushed by medical technicians. Oxivalve opens delayed, oiled. Software updated to version 3.1 MR3. | Defect | HW | Pinch valve
ABL90 FLEX | Kunde hat sehr viel Messabbrüche bei Kapilärer Messung Einlassdichtungshalter getauscht da verschmutzt Analysator kalibriert und überprüft Fcuv 99,0% Interne QC OK Da der Fehler nicht eindeutig festgestellt werden kann wird zum Ausschluss einen Geräte Fehlers ein Leihgerät für zwei Wochen parallel aufgestellt. Erneuter Kundenbesuch wenn Leihgerät verfügbar. | Kunde hat sehr viel Messabbrüche bei Kapilärer Messung Einlassdichtungshalter getauscht da verschmutzt Analysator kalibriert und überprüft Fcuv 99,0% Interne QC OK Da der Fehler nicht eindeutig festgestellt werden kann wird zum Ausschluss einen Geräte Fehlers ein Leihgerät für zwei Wochen parallel aufgestellt. Erneuter Kundenbesuch wenn Leihgerät verfügbar. | SWO-17514 | No | Customer has replaced a lot of measurement interruptions in capilary measurement inlet seal holder as dirty analyzer calibrates and checks Fcuv 99.0% Internal QC OK Since the error cannot be clearly determined, a device error is used to exclude a loaner for two weeks in parallel. Renewed customer visit when rental device is available. | Customer has replaced a lot of measurement interruptions in capilary measurement inlet seal holder as dirty analyzer calibrates and checks Fcuv 99.0% Internal QC OK Since the error cannot be clearly determined, a device error is used to exclude a loaner for two weeks in parallel. Renewed customer visit when rental device is available. | User Error | Pre analytical | Analyzer loaner | Capillary abort
ABL90 FLEX | Fetttest, Justagen, Kalibrationen und Qualitätskontrollen durchgeführt | Fetttest, Justagen, Kalibrationen und Qualitätskontrollen durchgeführt | SWO-17659 | No | Fat test, adjustments, calibrations and quality controls carried out | Fat test, adjustments, calibrations and quality controls carried out | Unknown | -
ABL90 | Back flushed system | Replaced sensor cassette. | CE8CE4C330071CF4 | No | Back flushed system | Replaced sensor cassette. | Assisted Replacement | SC | Wetability
ABL835 FLEX |  |  | SWO-17588 | Yes |  |  | Re classified RMED | HW | Inlet Module | Inlet Gasket
ABL835 FLEX | OXI modul ombyttet | OXI modul ombyttet | SWO-17590 | No | OXY ombyttet Mode | OXY ombyttet Mode | Defect | HW | Oxi module | Hemolyzer
AQT90 FLEX | Gerät überprüft, Fehlermeldungen geprüft, kein Fehler feststellbar. Gerätefunktion OK |  | SWO-18091 | No | Device checked, error messages checked, no error detected. Device function OK
ABL80 OSM | Checked sensor cassette and solution pack, checked pumps and valves | Performed Pouch cycles, valves cycles, waste pump clear, ran liquid sensor adjust(63), removed hemolyzer and replaced hemolyzer tubing, | CE8EC56FA63BC003 | No | Checked sensor cassette and solution pack, checked pumps and valves | Performed Pouch cycles, valves cycles, waste pump clear, ran liquid sensor adjust(63), removed hemolyzer and replaced hemolyzer tubing,
ABL800 BASIC | Messkammer gereinigt, Ref.-Elektrode neu membraniert, Ca-Elektrode überprüft, da Status zu niedrig, Kal, QC | Messkammer gereinigt, Ref.-Elektrode neu membraniert, Ca-Elektrode überprüft, da Status zu niedrig, Kal, QC | SWO-17839 | No | Measuring chamber cleaned, ref electrode re-membraneed, Ca electrode checked as status too low, Kal, QC | Measuring chamber cleaned, ref electrode re-membraneed, Ca electrode checked as status too low, Kal, QC | Assisted Replacement | Electrode | Membrane | Reference
ABL825 FLEX | Gerät kontrolliert, Meßmodulventile ausgebaut und geprüft, pCO2 Membran (0595) ausgetauscht gegen (0490), Service tHb-Kal durchgeführt, Schlauchsystem kontrolliert, Schläuche im Einlaß ersetzt, E788 ausgetauscht. | Gerät kontrolliert, Elektroden kontrolliert, pCO2 Elektrode abgewaschen, mehrfach membraniert, Gasventilmembran pH BG Modul ersetzt, Oximetrieabgleich durchgeführt | SWO-17859 | No | Device inspected, measuring module valves removed and tested, pCO2 membrane (0595) replaced with (0490), service tHb-Kal carried out, hose system inspected, hoses replaced in inlet, E788 replaced. | Device controlled, electrodes controlled, pCO2 electrode washed off, multiple membranes, gas valve membrane pH BG module replaced, oximetry matching performed | Assisted Replacement | Electrode | Membrane | pCO2
ABL835 FLEX |  |  | CE8E9D26A6406367 | No |  |  | Unknown | Setup
ABL800 BASIC | Gerät gereinigt und überprüft, Proteinentfernung durchgeführt, CO2 neu membraniert, Messkammern gereinigt, Ventile im Serviceprogramm überprüft, das Gasventil im pH BG Modul erneuert, Kals und QCs durchgeführt. | Gerät überprüft, Meßkammer und Luftauslass gesäubert, pCO2 und Lac-Elektrode membraniert, Dichtungsring pH ersetzt, extra Dezimalstellen programmiert | SWO-18095 | No | Device cleaned and checked, protein removal performed, CO2 re-logged, measuring chambers cleaned, valves checked in the service program, the gas valve renewed in the pH BG module, Kals and QCs performed. | Device checked, measuring chamber and air outlet cleaned, pCO2 and lac electrode membraned, sealing ring pH replaced, extra decimal places programmed | Assisted Replacement | Electrode | Electrode | pCO2
ABL825 FLEX | Komplettes OXI Modul ausgetauscht, justiert und tHb Kal ausgeführt QC i.o Funktionskontrolle i.o | Komplettes OXI Modul ausgetauscht, justiert und tHb Kal ausgeführt QC i.o Funktionskontrolle i.o | SWO-17779 | No | Complete OXI module replaced, adjusted and tHb Kal executed QC i.o function control i.o | Complete OXI module replaced, adjusted and tHb Kal executed QC i.o function control i.o | Defect | HW | Oxi module | Oxi module
ABL820 FLEX | Syringe port blocked on Inlet. | Replaced Inlet Nipples. | CE8E479EA62CFDD1 | No | Syringe port blocked on Inlet. | Replaced Inlet Nipples. | Defect | HW | Inlet Module | Leaking | Nipple Kit
ABL80 FLEX CO-OX | Automatische Lieferung nicht erfolgt, keine SK verfügbar | Frau Feldmann vermutet eine Verstopfung der SC. SC demontiert und mit Luft, dann mit Wasser durchspült. SC wieder installiert. Frau F. meldet sich wenn sie weiter Unterstützung benötigt. | SWO-17972 | No | Automatic delivery not made, no SK available | Ms. Feldmann suspects a blockage of the SC. SC dismantled and rinsed with air, then with water. SC reinstalled. Ms. F. reports if she needs further support.
ABL80 FLEX CO-OX | Analysator gereinigt | Es kommt immer wieder Flüssigkeit dazu. Der ABL zeigt keinen Fehler. Messungen funktionieren. Frau Feldmann bittet um einen Vorortservice. Bitte zurückrufen. | SWO-17975 | No | Analyzer cleaned | There is always fluid. The ABL shows no error. Measurements work. Ms Feldmann is asking for a suburban service. Please call back.
ABL90 | Error 1069 unable to pump solutions . I also noticed calibration errors that started after a NICU sample was run(a clot might have been introduced). When flushing the system the customer should use tap water only, otherwise damage to the sensor cassette might occur. | Tried flushing the system, but it did not work. Replaced the sensor cassette from customer stock. I recommend customer switch to a smaller size tube. | CE8FCB151C3A8E04 | No | Error 1069 unable to pump solutions . I also noticed calibration errors that started after a NICU sample was run(a clot might have been introduced). When flushing the system the customer should use tap water only, otherwise damage to the sensor cassette might occur. | Tried flushing the system, but it did not work. Replaced the sensor cassette from customer stock. I recommend customer switch to a smaller size tube. | Assisted Replacement | SC | ?
ABL90 | Bay door opens right after it tries closing Found debris in the way of optical sensor | Removed Debris from Optical sensor switch Manullay closed door by adjusting set screw back to home position. Stepped through opening and closing door with success. | CE8FBF581C4A236E | No | Bay door opens right after it tries closing Found debris in the way of optical sensor | Removed Debris from Optical sensor switch Manullay closed door by adjusting set screw back to home position. Stepped through opening and closing door with success. | Adjustment | HW | SC Door
ABL5 | Gerät geprüft. Gas1 Flasche erneuert. Gas neu konfiguriert. Defekten Gasschlauch gekürzt. Elektroden und Membranen geprüft. Gerät kalibriert und QC durchgeführt. - Gasschläuche müssen beim nächsten Service erneuert werden! |  | SWO-18007 | No | device checked. Gas1 bottle renewed. Gas reconfigured. Defective gas hose shortened. electrodes and membranes. calibrated and QC performed. - Gas hoses must be renewed at the next service!
ABL825 FLEX | Druckereinheit ausgetauscht. Neustart. | Druckereinheit ausgetauscht. Neustart. | SWO-18204 | No | replaced by the printer unit. Restart. | replaced by the printer unit. Restart. | Defect | HW | Printer module | Printer Head
ABL825 FLEX | Gerät überprüft, E788 zurückgetauscht, Oximetriemodul probeweise gewechselt, Fluidichannel ausgebaut, Ventile kontrolliert, Ventilmembran V8 ersetzt, Alle Elektroden membraniert, Dichtungsring pH und Glukose getauscht, alles zurückgebaut | Gerät kontrolliert, Elektroden kontrolliert, pCO2 Elektrode abgewaschen, mehrfach membraniert, Gasventilmembran pH BG Modul ersetzt, Oximetrieabgleich durchgeführt | SWO-18086 | No | Device checked, E788 replaced, oximetry module changed on a trial basis, fluidichannel removed, valves checked, valve membrane V8 replaced, all electrodes membraneed, sealing ring pH and glucose exchanged, everything dismantled | Device controlled, electrodes controlled, pCO2 electrode washed off, multiple membranes, gas valve membrane pH BG module replaced, oximetry matching performed | Defect | HW | Fluidic manifold | Valve, V8
ABL825 FLEX |  |  | CE8F9AA41C28D32B | No |  |  | Assisted Replacement | Electrode | Electrode | pCO2
ABL825 FLEX | Alle Parametern membraniert, Protein Entfernung, tHb Kal durchgeführt. PH Elektrode und Messkammer gereinigt. Mehrere Kalibration durchgeführt. | Alle Parametern membraniert, Protein Entfernung, tHb Kal durchgeführt. PH Elektrode und Messkammer gereinigt. Mehrere Kalibration durchgeführt. | SWO-18052 | No | All parameters membraned, protein removal, tHb Kal performed. PH electrode and measuring chamber. Multiple calibrationperformed. | All parameters membraned, protein removal, tHb Kal performed. PH electrode and measuring chamber. Multiple calibrationperformed. | Assisted Replacement | Electrode | Membrane | All
TCM/Tosca | Setup the unit with customer's sensor and it powered up normally and calibrated without error. Monitored self for 30 minutes. |  | CE8FDBD41C41E306 | No | Setup the unit with customer's sensor and it powered up normally and calibrated without error. Monitored self for 30 minutes. |  | No fault found | No fault found | No fault found | No fault found
ABL80 FLEX BASIC | neue SC installiert (auf Lieferschein), Kal, QC Level 2 und 3 i.O. | Die SC wurde reinstalliert, das Gerät geht in die Startphase und verlangt ständig, dass am P-Schlauch gezupft wird. Nach 15% bricht das Gerät die Kal ab. Der Kunde wurde an die Bestellannahme weitergeleitet um eine Nachtexpresslieferung für eine SC evtl. zu starten. Der Kunde bittet um einen Service und möchte dass noch heute eine SC vor Ort geliefert wird. | SWO-18299 | No | new SC installed (on delivery note), Kal, QC Level 2 and 3 I.O. | The SC has been reinstalled, the device goes into the start-up phase and constantly requires plucking on the P-tube. After 15%, the device breaks off the cal. The customer has been forwarded to the order acceptance in order to start a night express delivery for a SC possibly. The customer asks for a service and would like to have a SC delivered on site today.
ABL80 FLEX Basic version |  |  | CE90B5329201B297 | No
ABL90 FLEX | Software upgrade 3.1MR3 | Software upgrade 3.1MR3 | SWO-18205 | Yes | Software upgrade 3.1MR3 | Software upgrade 3.1MR3 | Re classified RMED | SWU
ABL90 FLEX | Einlassdichtung erneuert. Siehe Jahreswartung. | Einlassdichtung erneuert. Siehe Jahreswartung. | SWO-18319 | No | Inlet seal renewed. See Annual Maintenance. | Inlet seal renewed. See Annual Maintenance. | Defect | HW | Inlet Gasket
ABL90 FLEX | Fehlersuche Fließwege gespült Einlaßdichtung erneuert, da verstopft Einlaufverfahren | Fehlersuche Fließwege gespült Einlaßdichtung erneuert, da verstopft Einlaufverfahren | SWO-18234 | No | Troubleshooting flow paths flushed inlet seal renewed, as clogged inlet method | Troubleshooting flow paths flushed inlet seal renewed, as clogged inlet method | Cleaned | Fluidic pathway | Clot
ABL90 FLEX | Software upgrading uitgevoerd die een robustere monster afhandeling verzorgt. Vloeistofkanaal doorgespoten met demiwater. Wastetubing vervangen daar en enige paarse verkleuring zichtbaar is aan de connector zijde. Instrument weer opgestart en laten kalibreren. Problemen bij dit opstarten. Spoelen etc. helpt niet. Reagenspack omgewisseld met die van analyzer bestemd voor Harlingen. Nu functioneren beide analyzers zonder problemen. | Software upgrading uitgevoerd die een robustere monster afhandeling verzorgt. Vloeistofkanaal doorgespoten met demiwater. Wastetubing vervangen daar en enige paarse verkleuring zichtbaar is aan de connector zijde. Instrument weer opgestart en laten kalibreren. Problemen bij dit opstarten. Spoelen etc. helpt niet. Reagenspack omgewisseld met die van analyzer bestemd voor Harlingen. Nu functioneren beide analyzers zonder problemen. | SWO-18203 | No | Software upgrading that provides Robustere sample handling. Liquid channel sprayed with demi water. Wastetubing replaced there and some purple discoloration is visible on the connector side. The Instrument and calibrate it. Problems with this startup. Rinse etc. does not help. Reagent pack exchanged with that of analyzer intended for Harlingen. Now both analyzers function without problems. | Software upgrading that provides Robustere sample handling. Liquid channel sprayed with demi water. Wastetubing replaced there and some purple discoloration is visible on the connector side. The Instrument and calibrate it. Problems with this startup. Rinse etc. does not help. Reagent pack exchanged with that of analyzer intended for Harlingen. Now both analyzers function without problems. | Defect | HW | Waste tube
ABL90 FLEX | Einlasssondenhalter ersetzt, Test wiederholen -AABL i.O., Kal u. QC i.O. |  | SWO-18302 | No | Inlet probe holder replaced, test repeat -AABL i.O., Kal & QC i.O. |  | Defect | HW | Inlet Gasket
ABL90 FLEX | Bei Eintreffen LP nicht im Gerät verriegelt Lösungen Können nicht gepumpt werden, LP neu installiert, Nadel verbogen -aus Kundenbestand ersetzt, Einlassdichtung und Einlassanschlussdichtung ersetzt, Fließwege gespült, SC abgelaufen-SC neu installiert | Bei Eintreffen LP nicht im Gerät verriegelt Lösungen Können nicht gepumpt werden, LP neu installiert, Nadel verbogen -aus Kundenbestand ersetzt, Einlassdichtung und Einlassanschlussdichtung ersetzt, Fließwege gespült, SC abgelaufen-SC neu installiert | SWO-18305 | No | On arrival LP not locked in the device solutions Can not be pumped, LP reinstalled, needle bent -replaced from customer stock, inlet seal and inlet connection seal replaced, flow paths flushed, SC expired-SC reinstalled | On arrival LP not locked in the device solutions Can not be pumped, LP reinstalled, needle bent -replaced from customer stock, inlet seal and inlet connection seal replaced, flow paths flushed, SC expired-SC reinstalled | User Error | SP install
ABL90 | Sustituido gasket, inlet y tubos de fluidos. Revisado y comprobado buen funcionamiento del equipo. |  | CE90BBFB923B18D3 | No | Replaced gasket, inlet and fluid tubes. Checked and tested proper operation of the equipment. |  | Defect | HW | Oxi module | Pinch Valve tube
ABL820 FLEX | Unit working upon arrival. Dicussed with Biomed about use of UPS backup power. Also Winter Dry StaticElectric. | Replaced suspect Display. | CE90E2079206EE81 | No | Unit working upon arrival. Dicussed with Biomed about use of UPS backup power. Also Winter Dry StaticElectric. | Replaced suspect Display. | Defect | HW | DMS_PC | Touch screen | Display
ABL825 FLEX | Printer will not print. After replacing the printer the unit would not print. Checked all the connections gouing to the printer and the unit did not work. Called tech support and was told to check the fuse on the back of the power supply. Found that the fuse holder wasd loose. | Replaced the printer assembly. Tightend the fuse holder. | CE90D5ED9233CCBC | No | Printer will not print. After replacing the printer the unit would not print. Checked all the connections gouing to the printer and the unit did not work. Called tech support and was told to check the fuse on the back of the power supply. Found that the fuse holder wasd loose. | Replaced the printer assembly. Tightend the fuse holder. | Defect | HW | Power Supply | Fuse for Printer
ABL90 FLEX | Lösungspack im Servicemode überprüft Pumpenkal zeitweise fehlerhaft, LP ersetzt (LP alt Lot RN -07), Temperaturfehler Substrat - SC erneuert ( SC alt Lot R0290), Kal und QC durchgeführt, Pat Ergebnisse im Speicher mit Fr. Dr. Polievka durchgesprochen - Probenfehler , inhommogen Probe ungenügende Probe vorhanden - Einweisung und Probennehmer Demo angeboten | Bei Eintreffen LP nicht im Gerät verriegelt Lösungen Können nicht gepumpt werden, LP neu installiert, Nadel verbogen -aus Kundenbestand ersetzt, Einlassdichtung und Einlassanschlussdichtung ersetzt, Fließwege gespült, SC abgelaufen-SC neu installiert | SWO-18468 | No | Solution pack in service mode checks pump scale temporarily faulty, LP replaced (LP alt Lot RN -07), temperature error substrate - SC renewed ( SC old lot R0290), Kal and QC performed, Pat results performed in memory with Fr. Dr. Polievka - sample error, inhommogen Sample insufficient sample available - instruction and sampler demo offered | On arrival LP not locked in the device solutions Can not be pumped, LP reinstalled, needle bent -replaced from customer stock, inlet seal and inlet connection seal replaced, flow paths flushed, SC expired-SC reinstalled | User Error | SP install | Probe bend
ABL90 FLEX | Leihgerät aufgestellt 090R0538N002 Analysator kalibriert und überprüft Interne QC OK Gerät ans Netzwerk angebunden und an Aqure 1x SC und 1x LP für Leihgerät dagelassen Leihgerät für zwei Wochen zum testen |  | SWO-18349 | Yes | Loanr set up 090R0538N002 Analyzer calibrated and checked Internal QC OK device connected to the network and left to Aqure 1x SC and 1x LP for loaner loaner for two weeks |  | Re classified RMED | Install loaner
ABL90 | The analyzer continues to ask for a flush or the replacement of the Solution pack. The SP SC has been replaced two times, so I it is not the source of the problem. When I flushed the system the water was not flowing properly from the probe. | I replaced the probe and checked the flow it was much better Brougth the system up. | CE920A9F08368E08 | No | The analyzer continues to ask for a flush or the replacement of the Solution pack. The SP SC has been replaced two times, so I it is not the source of the problem. When I flushed the system the water was not flowing properly from the probe. | I replaced the probe and checked the flow it was much better Brougth the system up. | Replaced | Probe | Clot
AQT90 FLEX | Mechanik geprüft, Motoren gefettet, Processwheel Sensor gereinigt, Reinigung durchgef., QCs durchgef. Bisher hat der Kunde noch keine QCs am Gerät gemessen. Die meisten QC´s beim Kunden sind verfallen. Herrn Schweitzer darauf hingewiesen, der Kunde hat neue QCs bestellt Zudem wurden keine Systemreinigungen am Gerät durchgeführt. | Mechanik geprüft, Motoren gefettet, Processwheel Sensor gereinigt, Reinigung durchgef., QCs durchgef. Bisher hat der Kunde noch keine QCs am Gerät gemessen. Die meisten QC´s beim Kunden sind verfallen. Herrn Schweitzer darauf hingewiesen, der Kunde hat neue QCs bestellt Zudem wurden keine Systemreinigungen am Gerät durchgeführt. | SWO-18417 | No | Mechanics tested, motors greased, Processwheel sensor cleaned, cleaning through, QCs carried through. So far, the customer has not measured any QCs on the device. Most of the QQes at the customer have fallen into disrepair. Mr. Schweitzer pointed out that the customer has ordered new QCs and that no system cleaning was carried out on the device. | Mechanics tested, motors greased, Processwheel sensor cleaned, cleaning through, QCs carried through. So far, the customer has not measured any QCs on the device. Most of the QQes at the customer have fallen into disrepair. Mr. Schweitzer pointed out that the customer has ordered new QCs and that no system cleaning was carried out on the device.
ABL90 FLEX | Fehlersuche LP verbraucht, keine Aktivitäten mehr vorhanden LP geliefert und eingelegt Analysator kalibriert und überprüft Interne QC OK | Fehlersuche LP verbraucht, keine Aktivitäten mehr vorhanden LP geliefert und eingelegt Analysator kalibriert und überprüft Interne QC OK | SWO-18530 | Yes | Troubleshooting LP consumed, no more activities available LP delivered and inserted analyzer calibrated and checked Internal QC OK | Troubleshooting LP consumed, no more activities available LP delivered and inserted analyzer calibrated and checked Internal QC OK | Re classified RMED | Logistical upSWO
ABL90 |  |  | CE9796B3E0545428 | Yes |  |  | Re classified RMED | Preventive maintenance
ABL90 FLEX | Gerät aus- u. eingeschaltet, Einlass gespült, im Service Fluid Transport überprüft u. Fehler behoben Kal u. QC i.O. Probenmessung - i.O. | Gerät aus- u. eingeschaltet, Einlass gespült, im Service Fluid Transport überprüft u. Fehler behoben Kal u. QC i.O. Probenmessung - i.O. | SWO-18702 | No | Device off and on, inlet flushed, checked in service fluid transport and bug fixed Kal & QC i.O. sample measurement - i.O. | Device off and on, inlet flushed, checked in service fluid transport and bug fixed Kal & QC i.O. sample measurement - i.O. | Cleaned | Fluidic pathway
ABL825 FLEX | Gerät überprüft, Hemolysator ausgetauscht und justiert, pH BG Modul erneuert | Gerät überprüft, Hemolysator ausgetauscht und justiert, pH BG Modul erneuert | SWO-18683 | No | Device checked, hemolysator replaced and adjusted, pH BG module renewed | Device checked, hemolysator replaced and adjusted, pH BG module renewed | Assisted Replacement | Electrode | Membrane | pCO2 | pH/BG replaced
ABL815 FLEX | Keine weiteren Maßnahmen durchgeführt. Fällige Halbjahreswartung durchgeführt. | Keine weiteren Maßnahmen durchgeführt. Fällige Halbjahreswartung durchgeführt. | SWO-18627 | Yes | No further action has been taken. Due half-year maintenance. | No further action has been taken. Due half-year maintenance. | Re classified RMED
ABL815 FLEX |  |  | CE95E2706A39AB95 | Yes |  |  | Re classified RMED | Installation
ABL80 FLEX CO-OX | Gerät demontiert und Gereinigt. Schlauchset erneuert. Nach Neustart sind QC Fehler in sO2 und pCO2 aufgetreten, Lösungspack erneuert. |  | SWO-18969 | No | Device dismantled and cleaned. Hose set renewed. After reboot, QC errors occurred in sO2 and pCO2, solution pack renewed.
ABL90 FLEX | Gerät überprüft, Pumpenschlauch defekt, getauscht, waste connector verschmutz, gereinigt, mit Proteinentfernung manuell gespült, Kaltstartphase, LP und SC getauscht, Kal und QC io | Gerät überprüft, Pumpenschlauch defekt, getauscht, waste connector verschmutz, gereinigt, mit Proteinentfernung manuell gespült, Kaltstartphase, LP und SC getauscht, Kal und QC io | SWO-18796 | No | Device checked, pump hose defective, exchanged, waste connector dirty, cleaned, with protein removal manually rinsed, cold start phase, LP and SC exchanged, Kal and QC io | Device checked, pump hose defective, exchanged, waste connector dirty, cleaned, with protein removal manually rinsed, cold start phase, LP and SC exchanged, Kal and QC io | Assisted Replacement | SP | Pump tube
ABL90 FLEX | Software upgrade V3.1MR3 selon l'AVIS DE SÉCURITÉ ABL90 du 13 janvier 2015 | Software upgrade V3.1MR3 selon l'AVIS DE SÉCURITÉ ABL90 du 13 janvier 2015 | SWO-18900 | Yes | Software upgrade V3.1MR3 according to THE ABL90 SAFETY AVIS of January 13, 2015 | Software upgrade V3.1MR3 according to THE ABL90 SAFETY AVIS of January 13, 2015 | Re classified RMED | FAN / TU
ABL835 FLEX |  | Ref electrode replaced | CEA3EF8D7C03470D | No |  | Ref electrode replaced | Assisted Replacement | Electrode | Electrode | Reference
ABL835 FLEX |  | Remembraned pco2 | CEA3EEC87C2F470B | No |  | Remembraned pco2 | Assisted Replacement | Electrode | Membrane | pCO2
ABL835 FLEX |  |  | CE97901AE05A5422 | No |  |  | Other | Adjustment | Oxi module | Lamp unit
ABL800 BASIC | Dummy durch Elektrode ersetzt, Dichtungsring pH getauscht, HJW-Kit eingebaut, Meßkammer und Luftauslass gesäubert, Elektroden membraniert, Ventile kontrolliert, | Dummy durch Elektrode ersetzt, Dichtungsring pH getauscht, HJW-Kit eingebaut, Meßkammer und Luftauslass gesäubert, Elektroden membraniert, Ventile kontrolliert, | SWO-18948 | No | Dummy replaced by electrode, sealing ring pH exchanged, HJW kit installed, measuring chamber and air outlet cleaned, electrodes membraneed, valves controlled, | Dummy replaced by electrode, sealing ring pH exchanged, HJW kit installed, measuring chamber and air outlet cleaned, electrodes membraneed, valves controlled, | Assisted Replacement | Electrode | Membrane | pCO2
ABL820 FLEX | Determined has RALS network and that the astm.ini file must be modified. | Copied necessary file files into ABL700 folder, also entered Static IP requested by customer. | CEA386117C5E3B02 | No | Determined has RALS network and that the astm.ini file must be modified. | Copied necessary file files into ABL700 folder, also entered Static IP requested by customer. | Setup | HIS/LIS | -
ABL815 FLEX |  |  | CE9771F9E01F57AF | No |  |  | Other | Adjustment | Oxi module | Lamp unit
AQT90 FLEX | Gerät eingeschalten - Röhrchen aus Gerät entfernt und Neustartphase inkl Temperierung abgewartet, Puffermessungen und PatMessungen i.O., kurze Einweisung durchgeführt | Gerät eingeschalten - Röhrchen aus Gerät entfernt und Neustartphase inkl Temperierung abgewartet, Puffermessungen und PatMessungen i.O., kurze Einweisung durchgeführt | SWO-18828 | No | Device switched on - tubes removed from device and restart phase incl. temperature control awaited, buffer measurements and pat measurements i.O., short instruction carried out | Device switched on - tubes removed from device and restart phase incl. temperature control awaited, buffer measurements and pat measurements i.O., short instruction carried out
TCM-COMBIM | THE INSTRUMENT WAS POSTING A HEATER FAILURE. IT WAS OVERHEATING THE SENSOR. | THE ELECTRODE MODULE WAS REPLACED | CE970F7DE02C9083 | No | THE INSTRUMENT WAS POSTING A HEATER FAILURE. IT WAS OVERHEATING THE SENSOR. | THE ELECTRODE MODULE WAS REPLACED | Defect | HW | Electrode module
ABL80 FLEX BASIC | LP u. SC erneut installiert -> weiter Fehler, neue SC (alt: 208838) inst., div. Kal., QC = Fehler, LP (alt: 207847) gew., div. Kal., Kal. + QC i.O. | LP u. SC erneut installiert -> weiter Fehler; neue SC (alt: 208838) inst.; div. Kal.; QC = Fehler; LP (alt: 207847) gew.; div. Kal.; Kal. + QC i.O. | SWO-18968 | No | LP and SC reinstalled -> further error; new SC (old: 208838) inst.; div. Kal.; QC = Error; LP (old: 207847) (old: 207847) div. Kal.; Kal. + QC in the first place | LP and SC reinstalled -> further error; new SC (old: 208838) inst.; div. Kal.; QC = Error; LP (old: 207847) (old: 207847) div. Kal.; Kal. + QC in the first place
ABL90 FLEX | Szennyes elfolyó cső szivárgott a szennyes csatlakozónál. Csatlakozó tisztítása, szennyes rendszer tisztítása. Szivárgás teszt, pumpa, folyadékérzékelő beállítás. Kalibrálás, QC mérés. |  | SWO-19012 | No | Effluent drain pipe leaked at the dirty coupling. Cleaning the connector, cleaning the dirty system. Leakage test, pump, fluid detection setting. Calibration, QC measurement. |  | Cleaned | Fluidic pathway
ABL90 FLEX | Spülung und neustart. | Spülung und neustart. | SWO-19209 | No | Rinse and restart. | Rinse and restart. | Cleaned | Fluidic pathway | Clot
ABL90 | Checked the data logs and found the problem has happened intermittently over the last few days. To date the solution pack and sensor cassette have been changed twice and the inlet probe once. While in service, I found that the solution pack had failed. | Upgraded analyzer software to version 3.1 MR3, and replaced the inlet gasket and Inlet connector tube. I ran the clot detection removal then tested the solution pack again the solution pack passed. | CE98A3DA56206A6A | No | Checked the data logs and found the problem has happened intermittently over the last few days. To date the solution pack and sensor cassette have been changed twice and the inlet probe once. While in service, I found that the solution pack had failed. | Upgraded analyzer software to version 3.1 MR3, and replaced the inlet gasket and Inlet connector tube. I ran the clot detection removal then tested the solution pack again the solution pack passed. | Assisted Replacement | SC | Wetability | Bulletin not read
ABL825 FLEX | Verificação do funcionamento da válvula de oximetria, verificação de funcionamento deficiente, desmontagem e lubrificação da mesma, montagem e verificação do funcioanamento. Passagem do QC. tudo Ok. |  | CE9893B2563989BA | No | Check the operation of the oximetry valve, check for poor operation, disassembly and lubrication thereof, assembly and verification of the Funcioanamento. QC Pass. All right. |  | Defect | HW | Oxi module | Tube valve
ABL825 FLEX | Pumpen Schläuche erneuert, Pre -heater schlauch erneuert. LS - Justage, Mehrere Kal. | Pumpen Schläuche erneuert, Pre -heater schlauch erneuert. LS - Justage, Mehrere Kal. | SWO-19221 | No | Pumps hoses renewed, pre-heater hose renewed. LS - Adjustment, Multiple Cal. | Pumps hoses renewed, pre-heater hose renewed. LS - Adjustment, Multiple Cal. | Other | Adjustment | Tube reconnected
ABL827 FLEX |  |  | CE9AE81742315D26 | No |  |  | Defect | HW | Inlet module | Probe
ABL815 FLEX | Server Ping durchgeführt (172.16.10.60) durchgeführt - ok Blutgasgerät Ping durchgeführt (172.16.60.60) - ok Kommunikatioseinstellungen überprüft - ok Gerät verliert sporadisch die Verbindung zu Orbis (ca. 02:50 Uhr) und baut diese verzögert wieder auf Verbindung mit ASTM.NET überprüft - ok | Analysator hat nach Neustart einen Oximetriefehler, Analysator neu starten lassen. Lac Kal Fehler. Datenübertragung immer noch nicht möglich. Laborsoftware wurde neu gestartet, immer noch keine Übertragung der Befunde. | SWO-18977 | No | Server Ping performed (172.16.10.60) - ok blood gas device ping performed (172.16.60.60) - ok communicative settings checked - ok device sporadically loses the connection to Orbis (approx. 02:50) and reconnects to ASTM. NET checked - ok | Analyzer has had an oximetry error after reboot, analyser has restarted. Lac Kal error. Data transfer still not possible. Laboratory software has been restarted, still no transmission of the findings. | Defect | SW | His/lis
ABL835 FLEX | Fehlersuche Sicherungen überprüft, Aufnähme gewartet Software neu installiert Überprüfung Fließwege, Keys eingelesen Einlaufverfahren | Fehlersuche Sicherungen überprüft, Aufnähme gewartet Software neu installiert Überprüfung Fließwege, Keys eingelesen Einlaufverfahren | SWO-18991 | No | Troubleshooting backups checked, sewing serviced software reinstalled verification flow paths, keys read in run-in method | Troubleshooting backups checked, sewing serviced software reinstalled verification flow paths, keys read in run-in method | User Error | SW | SW Key | 15 min to initialize
ABL820 FLEX | Performed service diagnostics, lamp status check, Neon position failed @ 2.035, range = 0.100 to 2.00. Complete Oxi test indicates Lamp unit failure. | Replaced Lamp unit, performed required service diagnostics and adjustments. | CE98CDEF564B3034 | No | Performed service diagnostics, lamp status check, Neon position failed @ 2.035, range = 0.100 to 2.00. Complete Oxi test indicates Lamp unit failure. | Replaced Lamp unit, performed required service diagnostics and adjustments. | Defect | HW | Oxi module | Neon lamp
ABL820 FLEX | Performed service diagnostics, lamp status check, Neon position failed @ 2.035, range = 0.100 to 2.00. Complete Oxi test indicates Lamp unit failure. checked main power supply, found to be at spec. | Replaced Oxi Module PCB and Lamp unit, did not resolve problem. additional troubleshooting suggest potential malfunction of another component. Will order additional items and return next day. | CEA14F279040B59F | No | Performed service diagnostics, lamp status check, Neon position failed @ 2.035, range = 0.100 to 2.00. Complete Oxi test indicates Lamp unit failure. checked main power supply, found to be at spec. | Replaced Oxi Module PCB and Lamp unit, did not resolve problem. additional troubleshooting suggest potential malfunction of another component. Will order additional items and return next day. | Unknown | HW | Oxi module
ABL90 | CUSTOMER TRIED A NEW 2 SENSOR CASSETTES AND 1 SOLUTION PACK, CUSTOMER FLUSHED UNIT NUMEROUS TIMES COULD NOT REMOVE ERROR. | CHANGED ALL TUBING WITH ANNUAL MAINTENANCE KIT AND BROUGHT SOLUTION PACK AND SENSOR CASSETTE AND CHANGED BOTH. ALSO UPGRADED SOFTWARE TO 3.1 MR3. | CE99B6E5CC549BAB | No | CUSTOMER TRIED A NEW 2 SENSOR CASSETTES AND 1 SOLUTION PACK, CUSTOMER FLUSHED UNIT NUMEROUS TIMES COULD NOT REMOVE ERROR. | CHANGED ALL TUBING WITH ANNUAL MAINTENANCE KIT AND BROUGHT SOLUTION PACK AND SENSOR CASSETTE AND CHANGED BOTH. ALSO UPGRADED SOFTWARE TO 3.1 MR3. | Assisted Replacement | SC | Wetability | Bulletin not read
ABL90 FLEX | Fehlersuche Solutionpack erneuert, Kalibrierung und QC-Messungen durchgeführt | Fehlersuche Solutionpack erneuert; Kalibrierung und QC-Messungen durchgeführt | SWO-19550 | No | Troubleshooting Solutionpack renewed; Calibration and QC measurements performed | Troubleshooting Solutionpack renewed; Calibration and QC measurements performed | Assisted Replacement | SP | unknown
ABL825 FLEX | Found small hole in in let tube, customer does not have one | Replaced inlet tube Gave customer part # and they will order | CE997C96CC0FA883 | No | Found small hole in in let tube; customer does not have one | Replaced inlet tube Gave customer part # and they will order | Defect | HW | Inlet module | Inlet tube
ABL825 FLEX | Pump error in BG module. | Replaced all pump tubing, cleaned measuring chamber and replaced o ring on ph electrode, removed, cleaned and re-installed valves in ph module. | CE99E2E5CC00BB5D | No | Pump error in BG module. | Replaced all pump tubing, cleaned measuring chamber and replaced o ring on ph electrode, removed, cleaned and re-installed valves in ph module. | Defect | HW | Pump tube
ABL835 FLEX |  |  | CE99CF8BCC54E81B | No |  |  | Assisted Replacement | Electrode | Electrode | Reference
ABL820 FLEX | Autocheck will intermittently run the incorrect level of QC. Verified error by calling for a Level 3 and a level 1 was broken and ran. | Replaced 14 pole cable, checked and reseated all cables in autocheck system. | CE998B5ACC48A5EF | No | Autocheck will intermittently run the incorrect level of QC. Verified error by calling for a Level 3 and a level 1 was broken and ran. | Replaced 14 pole cable, checked and reseated all cables in autocheck system. | Defect | HW | AutoCheck Module | FPC Cable
ABL820 FLEX | Fails leak current in service. | Tried multiple module replacements, isolated to fluidic module.Replaced PCB now passes leak current test. | CE99C99DCC568AB5 | No | Fails leak current in service. | Tried multiple module replacements, isolated to fluidic module. Replaced PCB now passes leak current test. | Defect | HW | Fluidic manifold PCB | Is it valve??
ABL825 FLEX | System überprüft und gereinigt, pH Elektrode gereinigt, Gas Elektroden überprüft, Kalibrationen und Qualitätskontrollen ohne Beanstandungen durchgeführt, blutige Messungen ohne Beanstandungen. | System überprüft und gereinigt, pH Elektrode gereinigt, Gas Elektroden überprüft, Kalibrationen und Qualitätskontrollen ohne Beanstandungen durchgeführt, blutige Messungen ohne Beanstandungen. | SWO-19991 | No | System checked and cleaned, pH electrode cleaned, gas electrodes checked, calibrations and quality controls carried out without complaints, bloody measurements without complaints. | System checked and cleaned, pH electrode cleaned, gas electrodes checked, calibrations and quality controls carried out without complaints, bloody measurements without complaints. | Cleaned | Fluidic pathway | -
ABL825 FLEX | Looked through history and found numbers have improved slighlty over last week. Checked electrode for o-ring. Inspected machine for other signs of problems. | Re-membraned PCO2 electrode. I also re-membraned ref electrode as it was due on schedule. Performed necessary calibrations. | CE99D02CCC3541C7 | No | Looked through history and found numbers have improved slighlty over last week. Checked electrode for o-ring. Inspected machine for other signs of problems. | Re-membraned PCO2 electrode. I also re-membraned ref electrode as it was due on schedule. Performed necessary calibrations. | Assisted Replacement | Electrode | Membrane | pCO2
ABL825 FLEX | Der Fehler ist nicht mehr vorhanden und lässt sich nicht reproduzieren. | Der Fehler ist nicht mehr vorhanden und lässt sich nicht reproduzieren. | SWO-19670 | No | The error no longer exists and cannot be reproduced. | The error no longer exists and cannot be reproduced. | Unknown | HW | Display
ABL80 FLEX CO-OX | Service Test COOX i.O, tHb Kal durchgeführt, QC i.O, Manifold unit ersetzt, Kal und QC durchgeführt, Kunde beobachtet weiter und meldet sich bei Problemen |  | SWO-19960 | No | Service test COOX i.O, tHb Kal performed, QC i.O, Manifold unit replaced, Kal and QC performed, customer continues to monitor and reports in case of problems
ABL90 |  | Replaced Inlet pos detector PCB | CEA4CEABF2364713 | No |  | Replaced Inlet pos detector PCB | Defect | HW | Inlet PCB
ABL90 FLEX | Gerät überprüft, Nadel verbogen und Einlassdichtung verstopft, beides getauscht, Kal und QC Vermessen | Gerät überprüft, Nadel verbogen und Einlassdichtung verstopft, beides getauscht, Kal und QC Vermessen | SWO-19669 | No | Device checked, needle bent and inlet seal clogged, both exchanged, Kal and QC measured | Device checked, needle bent and inlet seal clogged, both exchanged, Kal and QC measured | Replaced | Probe | Probe bend
ABL90 |  |  | CE9FC51D1A57909E | No |  |  | Defect | HW | Inlet Gasket | LEH Email
ABL800 BASIC | Hibás pCO2 elektróda cseréje. Mérőkamra tisztítása. Kalibrálás, QC mérés. |  | SWO-19729 | No | Faulty pCO2 electrode replacement. Cleaning the measuring chamber. Calibration, QC measurement. |  | Assisted Replacement | Electrode | Electrode | pCO2
ABL827 FLEX |  |  | CEA3E1DB7C3B1AFD | No |  |  | Unknown | - | -
ABL825 FLEX |  | Replaced Hemloyzer | CEA4D032F2164716 | No |  | Replaced Hemloyzer | Defect | HW | Oxi module | Hemolyzer
ABL825 FLEX | System überprüft, Proteinentfernung durchgeführt, QC Eingescannt, LS-Justage durchgeführt, Reinigungslösung erneuert, Elektroden membraniert, Kalibrationen und Qualitätskontrollen ohne Beanstandungen durchgeführt. | System überprüft, Proteinentfernung durchgeführt, QC Eingescannt, LS-Justage durchgeführt, Reinigungslösung erneuert, Elektroden membraniert, Kalibrationen und Qualitätskontrollen ohne Beanstandungen durchgeführt. | SWO-19986 | No | System checked, protein removal performed, QC scanned, LS-adjustment performed, cleaning solution renewed, electrodes membraneed, calibrations and quality controls performed without complaints. | System checked, protein removal performed, QC scanned, LS-adjustment performed, cleaning solution renewed, electrodes membraneed, calibrations and quality controls performed without complaints. | Cleaned | Fluidic pathway | -
ABL835 FLEX | Comprobar el módulo oxi, probar cambiando la placa del módulo sin exito. Cambiar el espectrometro, calibrar las hemoglobinas y probar con muestras y controles. | Cambiar el espectrometro. | CE9AEABA4249FED9 | No | Check the oxy module, test by changing the module board without success. Change the spectrometer, calibrate hemoglobins and test with samples and controls. | Change the spectrometer. | Defect | HW | Oxi module | Spectrometer
ABL820 FLEX | Verified errors in cal log. Found lamp reading borderline high. | Adjusted lamp unit to ~3400. | CE9B3B4B423AA5F3 | No | Verified errors in cal log. Found lamp reading borderline high. | Adjusted lamp unit to ~3400. | Other | Adjustment | Lamp unit
ABL815 FLEX | System überprüft, Flex.-Q-Modul kalibriert, keine Besserung, Modul erneuert, Justagen durchgeführt, System OK. | System überprüft, Flex.-Q-Modul kalibriert, keine Besserung, Modul erneuert, Justagen durchgeführt, System OK. | SWO-19987 | No | System checked, Flex.-Q module calibrated, no improvement, module renewed, adjustment performed, system OK. | System checked, Flex.-Q module calibrated, no improvement, module renewed, adjustment performed, system OK. | Defect | HW | FlexQ Module
ABL90 |  |  | CEA9303C544A66AB | No |  |  | Cleaned | Fluidic pathway |  | LEH Email
ABL90 | REMOVED BATTERY AND ANALYZER STOPPED CYCLING ON AND OFF. BATTERY UNDER WARRANTY. | REPLACED BATTERY UNIT. | CE9EE862A42BD60E | No | REMOVED BATTERY AND ANALYZER STOPPED CYCLING ON AND OFF. BATTERY UNDER WARRANTY. | REPLACED BATTERY UNIT. | Defect | HW | Battery module
ABL90 FLEX | Fehler überprüft, Datensicherung erstellt. und an TS versendet. |  | SWO-19837 | No | Error checked, data backup created. and sent to TS. |  | Defect | SW | Repeated Cal 15 min
ABL90 | Clock needs to be set for 24hr. | Change clock settings in Control Panel Regional settings for 24hr format. Commit changes and perform reboot from command prompt. | CE9E6D7EA463881F | No | Clock needs to be set for 24hr. | Change clock settings in Control Panel Regional settings for 24hr format. Commit changes and perform reboot from command prompt. | Setup | Windows setting | 12H
ABL90 | Went into service fluidic tests. Could see where failed 5 SP solutions and also Clot detect failed Inlet and SC - # high. Ran all tests in service. All okay until running a QC - got 1250 error. | Replaced Inlet Probe from customer stock. Also replaced Inlet gasket from my stock. | CE9EE307A4432A44 | No | Went into service fluidic tests. Could see where failed 5 SP solutions and also Clot detect failed Inlet and SC - # high. Ran all tests in service. All okay until running a QC - got 1250 error. | Replaced Inlet Probe from customer stock. Also replaced Inlet gasket from my stock. | Cleaned | Fluidic pathway | Clot
ABL90 |  |  | CE9EB652A459F43B | No |  |  | Cleaned | Fluidic pathway
ABL825 FLEX | De FPC kabel van de autocheck werd vervangen. Wanneer deze defect is kunnen er problemen ontstaan zoals beschreven. Afregelen van de autocheck : OK Kalibraties en QC's : OK | De FPC kabel van de autocheck werd vervangen. Wanneer deze defect is kunnen er problemen ontstaan zoals beschreven. Afregelen van de autocheck : OK Kalibraties en QC's : OK | SWO-19819 | No | The car check's FPC cable was replaced. If this is defective, problems may arise as described. How to adjust the Auto check: OK calibrations and Qcs: OK | The car check's FPC cable was replaced. If this is defective, problems may arise as described. How to adjust the Auto check: OK calibrations and Qcs: OK | Defect | HW | AutoCheck Module | FPC Cable | Old version
ABL825 FLEX | Checked Auto check in service, alignment in service off Tried to upgrade software and screen froze again | Returned on 3-10-15 with USB bridge to retrieve data Instaleed new hard drive and complete software install Performed annual overhaul Replaced all parts in PM kit Alignment of auto check in service | CE9EAFCAA4291ACA | No | Checked Auto check in service, alignment in service off Tried to upgrade software and screen froze again | Returned on 3-10-15 with USB bridge to retrieve data Instaleed new hard drive and complete software install Performed annual overhaul Replaced all parts in PM kit Alignment of auto check in service | Defect | HW | DMS_PC | SATA HD
ABL825 FLEX |  | Replaced OXI PCB | CEA4D10DF2554719 | No |  | Replaced OXYGEN PCB | Defect | HW | Oxi module | PCB
ABL825 FLEX |  | Ref Electrode Replaced | CEA4D283F251471F | No |  | Ref Electrode Replaced | Assisted Replacement | Electrode | Electrode | Reference
ABL825 FLEX |  | Replaced Ref Electrode | CEA4D1E6F22E471C | No |  | Replaced Ref Electrode | Assisted Replacement | Electrode | Electrode | Reference
ABL810 FLEX |  |  | CE9FB81B1A5EF4E2 | No |  |  | Defect | HW | Fluidic manifold | Stepper motor
ABL827 FLEX | Fließwege und Messkammer gereinigt, Elektroden noch einmal neu bezogen, | Fließwege und Messkammer gereinigt, Elektroden noch einmal neu bezogen, | SWO-19874 | No | flow paths and measuring chamber cleaned, electrodes re-sourced, | flow paths and measuring chamber cleaned, electrodes re-sourced, | Assisted Replacement | Electrode | Membrane | Crea | Courtesy??
ABL815 FLEX | Fließwege und Ref Elektrodenconnector gereinigt, Luftwege gespült. Ventile ausgebaut und Membranen überprüft, Funktion OK. Problem trat während Service nicht auf. | Fließwege und Ref Elektrodenconnector gereinigt, Luftwege gespült. Ventile ausgebaut und Membranen überprüft, Funktion OK. Problem trat während Service nicht auf. | SWO-20058 | No | Flow ways and ref electrode connector cleaned, airways flushed. Valves removed and membranes checked, function OK. Problem did not occur during service. | Flow ways and ref electrode connector cleaned, airways flushed. Valves removed and membranes checked, function OK. Problem did not occur during service. | No fault found | - | - | pCO2
ABL817 FLEX | Found customer keyboard pushed under solution bottles. Possibly causing the bottles to loosen. Ran 2 point calibrations with no fluid transport errors. Ph out of range on calibrations | Moved keyboard out of the way. Remembraned reference for customer and will let analyzer calibrate over night. | CE9EE0EBA426548E | No | Found customer keyboard pushed under solution bottles. Possibly causing the bottles to loosen. Ran 2 point calibrations with no fluid transport errors. Ph out of range on calibrations | Moved keyboard out of the way. Remembraned reference for customer and will let analyzer calibrate over night. | User Error | Solution bottle | -
ABL825 FLEX | Rinse left in inlet after rinse. Waste tubing pinched. pH seal ring not sealing properly. | Replace seal ring for pH electrode. Properly route waste tubings. | CEA77084DE17622E | No | Rinse left in inlet after rinse. Waste tubing pinched. pH seal ring not sealing properly. | Replace seal ring for pH electrode. Properly route waste tubings. | Defect | HW | Fluidic manifold | Waste tube
AQT90 FLEX | AQT überprüft. Abgebrochene CRP-Kalibrationskassette entfernt. Startphase durchgeführt. QC's durchgeführt. | AQT überprüft. Abgebrochene CRP-Kalibrationskassette entfernt. Startphase durchgeführt. QC's durchgeführt. | SWO-20182 | No | AQT. Removed aborted CRP calibration cartridge. started phase. QC's. | AQT. Removed aborted CRP calibration cartridge. started phase. QC's.
AQT90 FLEX |  |  | CE9ECA12A41BB57F | No
ABL80 CO-OX | Reviewed data logs for any issues. Performed a service zero check with no problems. Could find no problem with the analyzer. Started up customers old ABL80 CO-OX and installed the sensor cassette on that ABL. Glucose also failed to calibrate on that analyzer. Replaced the sensor cassette but glucose failed to calibrate. Tried that sensor cassette on the other ABL had same problem. Tried a second new sensor cassette glucose failed to calibrations on both analyzers. | Problem not corrected. | CEA003161A376422 | No | Reviewed data logs for any issues. Performed a service zero check with no problems. Could find no problem with the analyzer. Started up customers old ABL80 CO-OX and installed the sensor cassette on that ABL. Glucose also failed to calibrate on that analyzer. Replaced the sensor cassette but glucose failed to calibrate. Tried that sensor cassette on the other ABL had same problem. Tried a second new sensor cassette glucose failed to calibrations on both analyzers. | Problem not corrected.
ABL90 |  |  | CEA00A241A1299C5 | No |  |  | Assisted Replacement | SC | Wetability
ABL90 FLEX | Gerät geprüft. Sämtliche Diagnoseprogramme im Service aufgerufen. Fliesswege mehrmals gespült. Schlauch vom Quetschventil erneuert. Einlassdichtung erneuert. Lösungspack und Sensorcassette installiert. Gerät demontiert und Abfallschlauch erneuert. Ganze Prozedur wiederholt. | Gerät geprüft. Sämtliche Diagnoseprogramme im Service aufgerufen. Fliesswege mehrmals gespült. Schlauch vom Quetschventil erneuert. Einlassdichtung erneuert. Lösungspack und Sensorcassette installiert. Gerät demontiert und Abfallschlauch erneuert. Ganze Prozedur wiederholt. | SWO-20720 | No | device checked. All diagnostic programs in the service are called. Flow paths rinsed several times. hose from the crushing valve. Inlet seal renewed. solution pack and sensor cassette. device and redesigned waste hose. Entire procedure repeated. | device checked. All diagnostic programs in the service are called. Flow paths rinsed several times. hose from the crushing valve. Inlet seal renewed. solution pack and sensor cassette. device and redesigned waste hose. Entire procedure repeated. | Defect | HW | Waste tube
ABL90 FLEX | Cassette erneuert, LP erneuert, Fließwege gespült Einl.dichtung erneuert, | Cassette erneuert, LP erneuert, Fließwege gespült Einl.dichtung erneuert, | SWO-20019 | No | Cassette renewed, LP renewed, flow paths flushed einl. seal renewed, | Cassette renewed, LP renewed, flow paths flushed einl. seal renewed, | Assisted Replacement | SC | Thermistor
ABL90 |  |  | CEA92C83543866A8 | No |  |  | Defect | Probe | Probe bend | LEH Email
ABL825 FLEX | se detecto códigos de lotes de QC mal cargados. se le explico al cliente como cargarlos correctamente. se solventaron dudas de funcionamiento. se comprobó correcto funcionamiento del equipo mediante qc. |  | CEB2B8AE0421F0FC | No | poorly loaded QC batch codes were detected. we explained to the customer how to load them correctly. doubts of operation were resolved. proper operation of the equipment using qc. |  | User Error | QC Setup | -
ABL800 BASIC | System überprüft, Verbindungskabel Monitor PC überprüft, Netzteil versuchsweise getauscht ohne Besserung, Fehler liegt am Mainboard, Ersatz PC Einheit eingebaut und Konfiguration größten teils hergestellt, Radiance Verbindung hergestellt. Kalibrationen und Qualitätskontrollen ohne Beanstandungen durchgeführt. PC Einheit wird nach Reparatur zurück getauscht. Radiance Verbindung nicht sichtbar, mit Herrn Azim Verbindung überprüft, noch einmal zurück in den OP, Verbindung OK! Server lässt sich an Pingen, zurück zum Labor, Gerät IST wieder in der Ansicht sichtbar. | System überprüft, Verbindungskabel Monitor PC überprüft, Netzteil versuchsweise getauscht ohne Besserung, Fehler liegt am Mainboard, Ersatz PC Einheit eingebaut und Konfiguration größten teils hergestellt, Radiance Verbindung hergestellt. Kalibrationen und Qualitätskontrollen ohne Beanstandungen durchgeführt. PC Einheit wird nach Reparatur zurück getauscht. Radiance Verbindung nicht sichtbar, mit Herrn Azim Verbindung überprüft, noch einmal zurück in den OP, Verbindung OK! Server lässt sich an Pingen, zurück zum Labor, Gerät IST wieder in der Ansicht sichtbar. | SWO-20422 | No | System checked, connection cable monitor PC checked, power supply exchanged experimentally without improvement, fault lies on the motherboard, replacement PC unit built-in and configuration mostly partially manufactured, radiance connection established. calibrations and quality controls without complaints. PC unit is replaced after repair. Radiance connection not visible, checked with Mr. Azim connection, back in the operating room, connection OK! Server can be pinged, back to lab, device IS back visible in view. | System checked, connection cable monitor PC checked, power supply exchanged experimentally without improvement, fault lies on the motherboard, replacement PC unit built-in and configuration mostly partially manufactured, radiance connection established. calibrations and quality controls without complaints. PC unit is replaced after repair. Radiance connection not visible, checked with Mr. Azim connection, back in the operating room, connection OK! Server can be pinged, back to lab, device IS back visible in view. | Defect | HW | DMS_PC
ABL827 FLEX |  |  | CEA250DF0663A013 | No |  |  | Other | User error | Solution replacement | Crea cleaning not mixed
ABL825 FLEX | Zur Überbrückung einen anderen Barcode Reader eingesetzt, Ersatzteil bestellt. | Zur Überbrückung einen anderen Barcode Reader eingesetzt, Ersatzteil bestellt. | SWO-20039 | No | Used to bridge a different barcode reader, spare part ordered. | Used to bridge a different barcode reader, spare part ordered. | Defect | HW | Barcode Reader | Front
ABL815 FLEX | Clucose and lactate intermittently failing calibrations due to high drift. pCO2 sharp downward shift on level 3 QC, pO2 level 3 QC riding consistently below insert mean, Na level 2 and 3 consistently running higher than insert mean. Checked zero current cal, inspected sample path. Removed and inspected measuring module valves, found both air valves stuck in open position. Also noticed daylight savings time had not taken effect on the instrument. | Replaced both measuring module air valves Properly configured Control Panel time settings | CEA010391A063405 | No | Clucose and lactate intermittently failing calibrations due to high drift. pCO2 sharp downward shift on level 3 QC, pO2 level 3 QC riding consistently below insert mean, Na level 2 and 3 consistently running higher than insert mean. Checked zero current cal, inspected sample path. Removed and inspected measuring module valves, found both air valves stuck in open position. Also noticed daylight savings time had not taken effect on the instrument. | Replaced both measuring module air valves Properly configured Control Panel time settings | Assisted Replacement | Electrode | Membrane | Glu
AQT90 FLEX |  |  | CEA014FF1A4DDAA6 | No
TCM-COMBIM | Upon seeing unit, found unit power supply and CF card corrupted. | Removed unit from Institution to perform Home repair. Ordered needed parts to complete repair. | CEA008011A22929F | No | Upon seeing unit, found unit power supply and CF card corrupted. | Removed unit from Institution to perform Home repair. Ordered needed parts to complete repair. | Defect | HW | Power Supply | Power Supply
ABL80 FLEX CO-OX | LP Lot 211436 läuft aus, aus Kundenbestand ersetzt, Initialisierung abgewartet, kein weiterer Austritt von Flüssigkeiten festgestellt | LP Lot 211436 läuft aus, aus Kundenbestand ersetzt, Initialisierung abgewartet, kein weiterer Austritt von Flüssigkeiten festgestellt | SWO-20522 | No | LP Lot 211436 expires, replaced from customer stock, initialization awaited, no further leakage of liquids detected | LP Lot 211436 expires, replaced from customer stock, initialization awaited, no further leakage of liquids detected
ABL80 FLEX CO-OX | Neues manifold, Funktionstest |  | SWO-20214 | No | New manifold, function test
ABL80 CO-OX | Checked all tubing and electrical connections everything was plugged in and connected ok, but the hemolyzer pinch tube looked to long , bearly in the pinch valve, may have not being pinched all the way. | shortened the tube and flushed the hemolyzer, all green | CEA160A2904A992E | No | Checked all tubing and electrical connections everything was plugged in and connected ok, but the hemolyzer pinch tube looked to long , bearly in the pinch valve, may have not being pinched all the way. | shortened the tube and flushed the hemolyzer, all green
ABL80 FLEX Basic version |  |  | CEA2575306629781 | No
ABL90 | FLUSHED UNIT NUMEROUS TIMES. ALREADY REPLACED SOLUTION PACK AND SENSOR. CASSETTE. REPLACED SENSOR CASSETTE FROM MY CONSIGNMENT AND PROBLEM WENT AWAY. ALSO UPGRADED SOFTWARE TO 3.1 MR3. | FLUSHED ANALYZER 10 TIMES AND REPLACED SENSOR CASSETTE. INSTALLED SOFTWARE 3.1MR3. ALSO REMOVED SHOW VALUES OUT OF REPORTABLE RANGES IN ROMA.INI. | CEA1432B90400EB6 | No | FLUSHED UNIT NUMEROUS TIMES. ALREADY REPLACED SOLUTION PACK AND SENSOR. CASSETTE. REPLACED SENSOR CASSETTE FROM MY CONSIGNMENT AND PROBLEM WENT AWAY. ALSO UPGRADED SOFTWARE TO 3.1 MR3. | FLUSHED ANALYZER 10 TIMES AND REPLACED SENSOR CASSETTE. INSTALLED SOFTWARE 3.1MR3. ALSO REMOVED SHOW VALUES OUT OF REPORTABLE RANGES IN ROMA. INI. | Assisted Replacement | SC | Wetability
ABL825 FLEX |  | Replaced Spectrometer unit and lamp unit. | CEA4DF36F2634732 | No |  | Replaced Spectrometer unit and lamp unit. | Defect | HW | Oxi module | Spectrometer
ABL820 FLEX | Performed visual inspection and ran service diagnostics. No errors were reported by assy test. | Performed cleaning procedure on pCO2 electrode. Reinstalled, allowed multiple 2 pt calibration cycles until stable calibration results were observed. Ran QC, observed pCO2 values now recovering at or close to mean (WDC) and within package range. | CEA15E48902CB5A2 | No | Performed visual inspection and ran service diagnostics. No errors were reported by assy test. | Performed cleaning procedure on pCO2 electrode. Reinstalled, allowed multiple 2 pt calibration cycles until stable calibration results were observed. Ran QC, observed pCO2 values now recovering at or close to mean (WDC) and within package range. | Other | No fault found | - | pCO2
ABL800 BASIC | Lampe neu justiert plus alle nötigen Kalibrierungen Cl im Lvl 4 fehlerhaft Ref. und Cl membraniert | Lampe neu justiert plus alle nötigen Kalibrierungen Cl im Lvl 4 fehlerhaft Ref. und Cl membraniert | SWO-20385 | No | Lamp re-adjusted plus all necessary calibrations Cl in lvl 4 faulty ref. and Cl membraned | Lamp re-adjusted plus all necessary calibrations Cl in lvl 4 faulty ref. and Cl membraned | Other | Adjustment | Oxi module | Lamp unit
AQT90 FLEX | Erneuerung und Umbau Frontklappe - Funktionsüberprüfung etc. | Erneuerung und Umbau Frontklappe - Funktionsüberprüfung etc. | SWO-20451 | No | Renewal and conversion front flap - function check etc. | Renewal and conversion front flap - function check etc.
AQT90 FLEX | Funktionsüberprüfung etc. - Software Upgrade siehe TU127 | Funktionsüberprüfung etc. - Software Upgrade siehe TU127 | SWO-20452 | No | Functional verification etc. - Software upgrade see TU127 | Functional verification etc. - Software upgrade see TU127
AQT90 FLEX |  |  | CEA0F3E890494823 | No
ABL90 |  |  | CEA2762A0651A950 | No |  |  | Defect | HW | Battery module
ABL90 | Checked oxi complete and found noise limit a little low on the spectrometer | Adjusted with oxi test complete Diag. all passed, recall and ran QC, all good | CEA290C506509932 | No | Checked oxi complete and found noise limit a little low on the spectrometer | Adjusted with oxi test complete Diag. all passed, recall and ran QC, all good | Adjustment | Oxi module | tHb cal
ABL90 |  |  | CEA3003906122383 | No |  |  | Assisted Replacement | SC | Wetability
ABL90 FLEX | Software neu installiert - Funktionsüberprüfung etc. ohne Beanstandung ! | Software neu installiert - Funktionsüberprüfung etc. ohne Beanstandung ! | SWO-20645 | No | Software reinstalled - function check etc. without complaint ! | Software reinstalled - function check etc. without complaint ! | Defect | SW | Touch did not function
ABL800 BASIC | Level 1Rot war in der Konfiguration nicht eingescannt, durchgeführt, AC Karussell eingesetzt, Kal, QC | Level 1Rot war in der Konfiguration nicht eingescannt, durchgeführt, AC Karussell eingesetzt, Kal, QC | SWO-20650 | No | Level 1Red was not scanned in the configuration, performed, AC carousel inserted, Kal, QC | Level 1Red was not scanned in the configuration, performed, AC carousel inserted, Kal, QC | User Error | QC Setup | -
ABL800 BASIC | Gerät dekontaminiert, LS Justage durchgeführt Complete Oximetrie Test durchgeführt - ok Halogen Lamp Intensity 3554, Stdev Dark Value 1,4 mV Gerät kalibriert und überprüft | Gerät dekontaminiert, LS Justage durchgeführt Complete Oximetrie Test durchgeführt - ok Halogen Lamp Intensity 3554, Stdev Dark Value 1,4 mV Gerät kalibriert und überprüft | SWO-20613 | No | Device Decontaminated, LS Adjustment Performed Complete Oximetry Test Performed - ok Halogen Lamp Intensity 3554, Stdev Dark Value 1.4 mV Device Calibrated and Checked | Device Decontaminated, LS Adjustment Performed Complete Oximetry Test Performed - ok Halogen Lamp Intensity 3554, Stdev Dark Value 1.4 mV Device Calibrated and Checked | Cleaned | Fluidic pathway | Oxi module
ABL835 FLEX |  | Checked for clots. Reset autocheck setup | CEA4D47FF219472A | No |  | Checked for clots. Reset autocheck setup | Cleaned | Fluidic pathway | -
ABL825 FLEX |  | Replaced AUTOCHECK FAN | CEA4D539F245472C | No |  | Replaced AUTOCHECK FAN | Defect | HW | AutoCheck Module | FAN
ABL815 FLEX | pCO2 Elektrode neu membraniert. Kammer gereinigt. Elektroden im PH BG- Modul überprüftund die Kammern gereinigt. Systemneustart. Ventile ausgebaut und getestet. | pCO2 Elektrode neu membraniert. Kammer gereinigt. Elektroden im PH BG- Modul überprüftund die Kammern gereinigt. Systemneustart. Ventile ausgebaut und getestet. | SWO-20517 | No | pCO2 electrode. Chamber cleaned. Electrodes in the PH BG module checked and the chambers cleaned. System restart. Valves removed and tested. | pCO2 electrode. Chamber cleaned. Electrodes in the PH BG module checked and the chambers cleaned. System restart. Valves removed and tested. | Assisted Replacement | Electrode | Membrane | pCO2
ABL825 FLEX | Laser scanner was replaced in January due to failure of scanner to read ampoules. | Replaced AutoCheck module and Laser Scanner. Verified Autocheck module operation. System operational. | CEA27FEC061C432C | No | Laser scanner was replaced in January due to failure of scanner to read ampoules. | Replaced AutoCheck module and Laser Scanner. Verified Autocheck module operation. System operational. | Defect | HW | AutoCheck Module | Barcode reader
ABL825 FLEX | Review system process, date time settings and the Windows time. | Determined the analyzer time was different from the hospital time by 7 minutes. When the information was input with time advanced from the analyzer it would populate the draw time field with the previous day. Advanced the time to coordinate with the hospital time. | CEA26ACA06177131 | No | Review system process, date time settings and the Windows time. | Determined the analyzer time was different from the hospital time by 7 minutes. When the information was input with time advanced from the analyzer it would populate the draw time field with the previous day. Advanced the time to coordinate with the hospital time. | Defect | SW | Time
ABL825 FLEX |  |  | CEA2943606182506 | No |  |  | Cleaned | Fluidic pathway | Oxi module | tHb cal
ABL825 FLEX |  |  | CEA297F3064F250A | No |  |  | Cleaned | Fluidic pathway | Oxi module | tHb cal
ABL825 FLEX | Messkammer und Schlauchsystem geprüft, Dichtung an der pH Elektrode ersetzt, Modulventile V1-V4 demontiert und geprüft, Gasversorgung geprüft, Funktionskontrolle ok | Messkammer und Schlauchsystem geprüft, Dichtung an der pH Elektrode ersetzt, Modulventile V1-V4 demontiert und geprüft, Gasversorgung geprüft, Funktionskontrolle ok | SWO-21719 | No | Measuring chamber and hose system tested, seal replaced on pH electrode, module valves V1-V4 dismantled and tested, gas supply tested, function control ok | Measuring chamber and hose system tested, seal replaced on pH electrode, module valves V1-V4 dismantled and tested, gas supply tested, function control ok | Defect | Electrode | Electrode | pCO2
AQT90 FLEX | Mehrere Reinigungen durchgeführt, Fließwege mit Hypochlorid gespült, Filter gereinigt, | Mehrere Reinigungen durchgeführt, Fließwege mit Hypochlorid gespült, Filter gereinigt, | SWO-20553 | No | Several cleanings performed, flow paths rinsed with hypochloride, filter cleaned, | Several cleanings performed, flow paths rinsed with hypochloride, filter cleaned,
ABL80 CO-OX | Analyzer down for a couple days. Has had 2 different SP's installed. Leakage behind SP all down manifold. | Clean leakage behind SP. Perform PM on analyzer. Replaced parts from PM kit. Performed alignment. | CEA380B57C1611E9 | No | Analyzer down for a couple days. Has had 2 different SP's installed. Leakage behind SP all down manifold. | Clean leakage behind SP. Perform PM on analyzer. Replaced parts from PM kit. Performed alignment.
ABL80 CO-OX | Checked the network connections, and found that the Radiance connection was not enabled. All other settings were the same as their 4 other analyzers, but I still could not connect. I took the analyzer to a different location with a known good analyzer connected. At first the analyzer would not connect, but I found the IP address was set to static in windows. I made the change to automatic and the analyzer connected. I took it back to its correct location and it still would not connect. I tried another analyzer in the bad location and it also would not connect. | The problem is either with the cable or the port in the wall. The customer is going to get with the hospital IT department to correct the problem. | CEA3F8887C345CB1 | No | Checked the network connections, and found that the Radiance connection was not enabled. All other settings were the same as their 4 other analyzers, but I still could not connect. I took the analyzer to a different location with a known good analyzer connected. At first the analyzer would not connect, but I found the IP address was set to static in windows. I made the change to automatic and the analyzer connected. I took it back to its correct location and it still would not connect. I tried another analyzer in the bad location and it also would not connect. | The problem is either with the cable or the port in the wall. The customer is going to get with the hospital IT department to correct the problem.
ABL90 FLEX | SC geliefert | SC geliefert | SWO-21077 | Yes | SC delivered | SC delivered | Re classified RMED | Logistical upSWO
ABL90 FLEX | Konfiguration angepasst | Konfiguration angepasst | SWO-20783 | No | Configuration adapted | Configuration adapted | Setup | User config
ABL90 FLEX | Installata e testata versione software 3.1 MR3 | Installata e testata versione software 3.1 MR3 | SWO-20813 | Yes | Installed and tested software version 3.1 MR3 | Installed and tested software version 3.1 MR3 | Re classified RMED | SWU
ABL90 | Unit stuck in the trouble shooting loop. Could not get the sensor cassette to pass the internal QC's. | Flushed the unit. Ran 5 test blood samples through the sensor. Performed a software upgrade. Replaced the sensor cassette from customers stock. | CEA3CD7F7C1612BB | No | Unit stuck in the trouble shooting loop. Could not get the sensor cassette to pass the internal QC's. | Flushed the unit. Ran 5 test blood samples through the sensor. Performed a software upgrade. Replaced the sensor cassette from customers stock. | Assisted Replacement | SC | unknown
ABL90 | A) Need to replace the wet section control board. B) Need to install the new style inlet probe and gasket. | A) Installed new board, but the analyzer keeps saying to close the inlet even when it is closed and the lights indicate the flap is down and closed. Original board reinstalled. B) Need to install the new style inlet probe and gasket. | CEA3A1F27C5A77EB | No | A) Need to replace the wet section control board. B) Need to install the new style inlet probe and gasket. | A) Installed new board, but the analyzer keeps saying to close the inlet even when it is closed and the lights indicate the flap is down and closed. Original board reinstalled. B) Need to install the new style inlet probe and gasket. | Defect | HW | Wet section PCB
ABL90 FLEX | Überprüfung Sensorfach, versucht zu schließen, bleibt aber hängen und geht in die offen Stellung zurück Demontage , manuelle Schließung Mgl. Schließsensor nachgebogen, Test anschl. i.O. Einlaufverfahren Lac Empf.fehler LP erneuert | Überprüfung Sensorfach, versucht zu schließen, bleibt aber hängen und geht in die offen Stellung zurück Demontage , manuelle Schließung Mgl. Schließsensor nachgebogen, Test anschl. i.O. Einlaufverfahren Lac Empf.fehler LP erneuert | SWO-20858 | No | Check sensor compartment, tries to close, but gets stuck and goes into the open position back disassembly, manual closing Mgl. locking sensor bent, test followed. i.O. inlet method Lac Empf.fehler LP renewed | Check sensor compartment, tries to close, but gets stuck and goes into the open position back disassembly, manual closing Mgl. locking sensor bent, test followed. i.O. inlet method Lac Empf.fehler LP renewed | Adjustment | Sensor Gate
ABL90 FLEX | Kalibrieren Messung QC Lösungspack wurde durch Kunden bereitserneuert | Kalibrieren Messung QC Lösungspack wurde durch Kunden bereitserneuert | SWO-20766 | No | Calibration measurement QC solution pack has already been renewed by customers | Calibration measurement QC solution pack has already been renewed by customers | Assisted Replacement | SP | Glu
ABL825 FLEX | Im Servicemenü überprüft, in Ordnung. Barcodescanner staubfrei. | Im Servicemenü überprüft, in Ordnung. Barcodescanner staubfrei. | SWO-20828 | No | Checked in the service menu, all right. Barcode scanner dust-free. | Checked in the service menu, all right. Barcode scanner dust-free. | Cleaned | HW | FlexQ Module | Dust
ABL825 FLEX |  | Replaced Hemolyzer | CEA4E04EF25F4736 | No |  | Replaced Hemolyzer | Defect | HW | Oxi module | Hemolyzer
ABL800 BASIC | Gerät überprüft, pH BG Modul erneuert und abgeglichen, Lactat membraniert, E788 ausgetauscht, Kal und QC o.k. | Dummy durch Elektrode ersetzt, Dichtungsring pH getauscht, HJW-Kit eingebaut, Meßkammer und Luftauslass gesäubert, Elektroden membraniert, Ventile kontrolliert, | SWO-20893 | No | Device checked, pH BG module renewed and matched, lactate membraneed, E788 replaced, Kal and QC ok. | Dummy replaced by electrode, sealing ring pH exchanged, HJW kit installed, measuring chamber and air outlet cleaned, electrodes membraneed, valves controlled, | Assisted Replacement | Electrode | Electrode | pCO2
ABL835 FLEX | Garantieaustausch des OXI-Moduls, Proteinentfernung durchgeführt. 902-833R1812N010 ersetzt durch 902-8341824N007. OXI Complete Test und Service tHb Kal, durchgeführt. tHb Test i.O. | Garantieaustausch des OXI-Moduls, Proteinentfernung durchgeführt. 902-833R1812N010 ersetzt durch 902-8341824N007. OXI Complete Test und Service tHb Kal, durchgeführt. tHb Test i.O. | SWO-20845 | Yes | Warranty replacement of the OXI module, protein removal carried out. 902-833R1812N010 replaced by 902-8341824N007. OXI Complete Test and service tHb Kal. tHb Test i.O. | Warranty replacement of the OXI module, protein removal carried out. 902-833R1812N010 replaced by 902-8341824N007. OXI Complete Test and service tHb Kal. tHb Test i.O. | Re classified RMED | out of box | Oxi module
AQT90 FLEX | Mehrfache Systemreinigung - Fehlersuche und Reinigung etc. laut Wartungsliste - Funktionsüberprüfung Praxis schloß um 14 Uhr - LQC wird von Kunde am Montag durchgeführt - telef. Rücksprache erfolgt ! | Mehrfache Systemreinigung - Fehlersuche und Reinigung etc. laut Wartungsliste - Funktionsüberprüfung Praxis schloß um 14 Uhr - LQC wird von Kunde am Montag durchgeführt - telef. Rücksprache erfolgt ! | SWO-20819 | No | Multiple system cleaning - troubleshooting and cleaning etc. according to maintenance list - function verification practice closed at 2 pm - LQC is performed by customer on Monday - telef. Consultation is carried out ! | Multiple system cleaning - troubleshooting and cleaning etc. according to maintenance list - function verification practice closed at 2 pm - LQC is performed by customer on Monday - telef. Consultation is carried out !
ABL90 FLEX | Gerät geprüft. Sämtliche Diagnoseprogramme im Service aufgerufen. Einlassnadel und Einlassdichtung erneuert. Lösungspack installiert. Neustart durchgeführt. QC und blutige Messung durchgeführt. |  | SWO-21165 | No | device checked. All diagnostic programs in the service are called. Inlet needle and inlet seal renewed. solution pack installed. Restart. QC and bloody measurement performed. |  | Cleaned | Fluidic pathway
ABL90 | Screen flashing, Mixer light flashing, burnt smell coming from battery | Replaced battery | CEA77504DE531424 | No | Screen flashing, Mixer light flashing, burnt smell coming from battery | Replaced battery | Defect | HW | Battery module
ABL90 FLEX | Fehlersuche Überprüfung im Service Programm Fließwege manuell durchgespült Einlaßdichtung und-Nadel gereinigt Einlaufverfahren Anschl. Backlog Fehler Na, k,ca, pH Austausch SC , da die im ABL befindliche noch keine ordentlich. Kal abgeliefert hat(wurde vermutlich im Tiefkühlschrank gelagert) | Fehlersuche Überprüfung im Service Programm Fließwege manuell durchgespült Einlaßdichtung und-Nadel gereinigt Einlaufverfahren Anschl. Backlog Fehler Na, k,ca, pH Austausch SC , da die im ABL befindliche noch keine ordentlich. Kal abgeliefert hat(wurde vermutlich im Tiefkühlschrank gelagert) | SWO-21058 | No | Troubleshooting check in the service program flow paths manually rinsed through the inlet seal and needle cleaned inlet method. Backlog error Na, k,ca, pH exchange SC , since the one in the ABL is not yet neat. Kal (probably stored in the freezer) | Troubleshooting check in the service program flow paths manually rinsed through the inlet seal and needle cleaned inlet method. Backlog error Na, k,ca, pH exchange SC , since the one in the ABL is not yet neat. Kal (probably stored in the freezer) | Assisted Replacement | SC | Wetability
ABL90 FLEX | LP richtig aktiviert, Fluid Transport ( Service) - i.O., Test wiederholen ausgeführt, Kal u. QC i.O. Zeit korrigiert | LP richtig aktiviert, Fluid Transport ( Service) - i.O., Test wiederholen ausgeführt, Kal u. QC i.O. Zeit korrigiert | SWO-21207 | No | LP correctly activated, Fluid Transport (Service) - i.O., Test Repeat executed, Kal & QC i.O. time corrected | LP correctly activated, Fluid Transport (Service) - i.O., Test Repeat executed, Kal & QC i.O. time corrected | No fault found | unknown
ABL90 FLEX | Sensor erneuert, Tests im Service Menü, allg. Sicht- und Funktionskontrolle, Kalibration, QC Kontrolle | Sensor erneuert, Tests im Service Menü, allg. Sicht- und Funktionskontrolle, Kalibration, QC Kontrolle | SWO-20975 | No | Sensor renewed, tests in the service menu, general visual and function control, calibration, QC control | Sensor renewed, tests in the service menu, general visual and function control, calibration, QC control | Assisted Replacement | SC | Wetability
ABL90 |  |  | CEA76AC0DE4E557E | No |  |  | Assisted Replacement | SC | Glu
ABL800 BASIC | Sprengring ab, Gasventil getauscht | Sprengring ab, Gasventil getauscht | SWO-21016 | No | Blasting ring off, gas valve replaced | Blasting ring off, gas valve replaced | Defect | HW | Gasmodule
ABL825 FLEX | Getting error 579 oxi hardware error. | Ran complete oxi test and did a tHB cal. | CEA75801DE1BC7F5 | No | Getting error 579 oxy hardware error. | Ran Complete oxygen test and did a tHB cal. | Adjustment | HW | Oxi module | tHb cal
ABL825 FLEX | Überprüfung Druckerrolle auf Verschmutzung oder Papierreste =>alles i.O. => schlechter Ausdruck durch fehlerhaften Druckkopf Austausch Druckkopf Testausdruck i.O. | Überprüfung Druckerrolle auf Verschmutzung oder Papierreste =>alles i.O. => schlechter Ausdruck durch fehlerhaften Druckkopf Austausch Druckkopf Testausdruck i.O. | SWO-21032 | No | Check printer roll for contamination or paper residue =>everything i.O. => poor printout due to incorrect print head exchange print head test printout i.O. | Check printer roll for contamination or paper residue =>everything i.O. => poor printout due to incorrect print head exchange print head test printout i.O. | Defect | HW | Printer module
ABL820 FLEX | Reviewed logs, and determined values were change to default value after a reboot had taken place on 1 22 15. No explanation as to why the unit rebooted or why it reset values to default.It was suggested the use a UPS for Protection and Filtering. | Unit to be monitored. | CEA7984CDE06E6A6 | No | Reviewed logs, and determined values were change to default value after a reboot had taken place on 1 22 15. No explanation as to why the unit rebooted or why it reset values to default. It was suggested the use a UPS for Protection and Filtering. | Unit to be monitored. | Defect | SW | Analyzer reset after power down
ABL837 FLEX | Found inlet probe is bent upward causing probe to hit on the top of the plastic collar | Replaced inlet probe, adjusted carasoul. Crea was out on QC.Cledaned measuring chamber. Multiple 2Pt cals required to bet Crea in .Changed the pH sealing ring with new blue sealing ring. | CEA76B2CDE342771 | No | Found inlet probe is bent upward causing probe to hit on the top of the plastic collar | Replaced inlet probe, adjusted carasoul. Crea was out on QC. Cledaned measuring chamber. Multiple 2Pt cals required to bet Crea in . Changed the pH sealing ring with new blue sealing ring. | Defect | HW | Inlet module | Probe bend
ABL825 FLEX | Bei Eintreffen Gerät in Ruhestellung da Abfallbehälter voll, Abfallbehälter erneuert, Gerät aus der Ruhestellung genommen, Kalibration anschließend OK, blutige Messungen abgearbeitet, Tray erneuert. | Bei Eintreffen Gerät in Ruhestellung da Abfallbehälter voll, Abfallbehälter erneuert, Gerät aus der Ruhestellung genommen, Kalibration anschließend OK, blutige Messungen abgearbeitet, Tray erneuert. | SWO-21522 | No | Upon arrival device in resting position as waste container full, waste container renewed, device taken from resting position, calibration then OK, bloody measurements processed, tray renewed. | Upon arrival device in resting position as waste container full, waste container renewed, device taken from resting position, calibration then OK, bloody measurements processed, tray renewed. | User Error | Replacement support | -
ABL820 FLEX | Inspected Probe and Inlet Gasket. Customer mentioned still having issues with RALS system. ABL will connect to system then disconnect.Called placed to Radance. | Replaced Inlet Gasket. | CEB07EC0185B4B85 | No | Inspected Probe and Inlet Gasket. Customer mentioned still having issues with RALS system. ABL will connect to system then disconnect. Called placed to Radance. | Replaced Inlet Gasket. | Assisted Replacement | HW | Inlet module | Inlet Gasket
ABL80 FLEX Basic version |  |  | CEA8FAF8544911D7 | No
ABL80 CO-OX | Investigation documented on SWO B9II4HMBB. Replaced the Sensor Connector PCB and the Analog PCB still had the same problem. | Unable to correct problem during this service call. | CEA8CBEA541F6732 | No | Investigation documented on SWO B9II4HMBB. Replaced the Sensor Connector PCB and the Analog PCB still had the same problem. | Unable to correct problem during this service call.
ABL90 | AFTER DUPLICATING THE PROBLEM I SUCCESSFULLY RAN A MANUAL QC SAMPLE.ANALYZER WAS GREEN WITH ALL CALIBRATIONS AND INTERNAL QC PASSING.PERFORMED ALL FLUID FLOW ADJUSTMENTS IN SERVICE.ERROR CODE 1254 FAILED TO ASPIRATE SAMPLES.TRIED REBOOTING ANALYZER.STILL ABORTING SAMPLES BUT RUNNING EXTERNAL QC FINE. | REPLACED SENSOR CASETTE AND PROBLEM WENT AWAY | CEA8DEC85409F1F2 | No | AFTER DUPLICATING THE PROBLEM I SUCCESSFULLY RAN A MANUAL QC SAMPLE. ANALYZER WAS GREEN WITH ALL CALIBRATIONS AND INTERNAL QC PASSING. PERFORMED ALL FLUID FLOW ADJUSTMENTS IN SERVICE. ERROR CODE 1254 FAILED TO ASPIRATE SAMPLES. TRIED REBOOTING ANALYZER. STILL ABORTING SAMPLES BUT RUNNING EXTERNAL QC FINE. | REPLACED SENSOR CASETTE AND PROBLEM WENT AWAY | Assisted Replacement | SC | Wetability
ABL90 | out of box failure | Installed new battery in ABL 90 at start up no battery icon reset battery then still no battery icon and abl bip continiusly reinstall old battery | CEA8E93D54601523 | Yes | out of box failure | Installed new battery in ABL 90 at start up no battery icon reset battery then still no battery icon and abl bip continiusly reinstall old battery | Re classified RMED | Install new battery
ABL90 | ABL90 would not asirate cap tube. | Reset the cap tube key and unit started to work. | CEA89F55545F1D8F | No | ABL90 would not asirate cap tube. | Reset the cap tube key and unit started to work. | No fault found | unknown | CAP testing?
ABL90 FLEX | Analysator geprüft. SC90 erneuert. Probelauf und Initialisierung i.O. SC Reklamiert. QC Fehler in pCO2 beseitigt. | Analysator geprüft. SC90 erneuert. Probelauf und Initialisierung i.O. SC Reklamiert. QC Fehler in pCO2 beseitigt. | SWO-21112 | No | Analyzer. SC90. Test run and initialization in sc complaint. QC error in pCO2 fixed. | Analyzer. SC90. Test run and initialization in sc complaint. QC error in pCO2 fixed. | Assisted Replacement | SC | Wetability
ABL825 FLEX | Found white residue on Reference Electrode Connector finger. | Cleaned Reference Electrode Connector finger with Alcohol. Also cleaned CA contacts as well. Cleaned contacts on Reference and CA Electrodes. Installed new blue Sealing Ring on pH Electrode as long as I was here. | CEA884F5540E9655 | No | Found white residue on Reference Electrode Connector finger. | Cleaned Reference Electrode Connector finger with Alcohol. Also cleaned CA contacts as well. Cleaned contacts on Reference and CA Electrodes. Installed new blue Sealing Ring on pH Electrode as long as I was here. | Cleaned | HW | pH/BG Module | Ref connector
ABL825 FLEX | AQC-Modul ausgeblasen, Luftauslass gespült, Kal und QC o.k. |  | SWO-21299 | No | AQC module blown out, air outlet rinsed, Kal and QC ok |  | Cleaned | HW | AutoCheck Module | Barcode reader
ABL800 BASIC | Leihgerät 754R0093N008 meßbereit aufgestellt, Konfiguration und Onlineanbindung eingestellt | Dummy durch Elektrode ersetzt, Dichtungsring pH getauscht, HJW-Kit eingebaut, Meßkammer und Luftauslass gesäubert, Elektroden membraniert, Ventile kontrolliert, | SWO-21298 | No | Rental device 754R0093N008 set up ready for measurement, configuration and online connection set | Dummy replaced by electrode, sealing ring pH exchanged, HJW kit installed, measuring chamber and air outlet cleaned, electrodes membraneed, valves controlled, | Assisted Replacement | Electrode | Membrane | pCO2
ABL835 FLEX | The printer is not printing. | Replaced the printer unit. Analyzer now prints as it should. | CEA8B802543D47A4 | No | The printer is not printing. | Replaced the printer unit. Analyzer now prints as it should. | Defect | HW | Printer module
TCM-COMBIM |  |  | CEA8EC8E5430E3DE | No |  |  | Defect | HW | Touch Screen | Touch Screen
ABL90 FLEX | Lösunggspack PCB ersetzt. Startphase QC und Kalibration. | Lösunggspack PCB ersetzt. Startphase QC und Kalibration. | SWO-21422 | No | Solution pack PCB. Start phase QC and calibration. | Solution pack PCB. Start phase QC and calibration. | Defect | HW | SP PCB
ABL90 | Analyzer turning off and beeping from CPU whenever it is powered up. Screen is blank. | Suspect battery has a problem with the Smart Battery Charger PCB. Checked battery unit as per ABL90 Flex battery unit procedure. Battery s n is P0057N023. Oredered Smart Battery PCB for unit. Will put battery back in unit when parts are recieved. | CEA9E361CA071BEE | No | Analyzer turning off and beeping from CPU whenever it is powered up. Screen is blank. | Suspect battery has a problem with the Smart Battery Charger PCB. Checked battery unit as per ABL90 Flex battery unit procedure. Battery s n is P0057N023. Oredered Smart Battery PCB for unit. Will put battery back in unit when parts are recieved. | Defect | HW | Battery module
ABL90 FLEX | Allgemeine Kontrolle ok. Alle Schläuche inkl. Einlassdichtung ersetzt ok. Einlassnadel (Mat. von Kunde) ersetzt ok. Komplettest i.o. System i.o. Kal. + QC | Allgemeine Kontrolle ok. Alle Schläuche inkl. Einlassdichtung ersetzt ok. Einlassnadel (Mat. von Kunde) ersetzt ok. Komplettest i.o. System i.o. Kal. + QC | SWO-21397 | No | General control ok. All hoses incl. inlet seal replaced ok. Inlet pin (mat. from customer) replaces ok. Complete i.o. system i.o. Kal. + QC | General control ok. All hoses incl. inlet seal replaced ok. Inlet pin (mat. from customer) replaces ok. Complete i.o. system i.o. Kal. + QC | Defect | HW | Oxi module | Pinch Valve tube
ABL90 FLEX | Einlassnadel verbogen - Einlassnadel aus Kundenbestand ersetzt, Fehlerspeicher ausgewertet , Einweisungstermin für Fehlersuche und Vermeidung wird vereinbart, Kal und QC durchgeführt | Einlassnadel verbogen - Einlassnadel aus Kundenbestand ersetzt, Fehlerspeicher ausgewertet , Einweisungstermin für Fehlersuche und Vermeidung wird vereinbart, Kal und QC durchgeführt | SWO-21493 | No | Inlet needle bent - inlet needle replaced from customer stock, error memory evaluated, instruction date for troubleshooting and avoidance is agreed, Kal and QC performed | Inlet needle bent - inlet needle replaced from customer stock, error memory evaluated, instruction date for troubleshooting and avoidance is agreed, Kal and QC performed | Replaced | Probe | Bend | SW ver ?
ABL90 | All communication is via the Wet Section Control Board. Need to replace the board. | Replaced the Wet Section Control Board and loaded software. | CEA9E899CA541147 | No | All communication is via the Wet Section Control Board. Need to replace the board. | Replaced the Wet Section Control Board and loaded software. | Defect | HW | Wet section PCB
ABL90 | Above error message appeared during install, brought unit back to repair depot. | Installed new SC and let unit run. | CEA9F5CDCA5F118B | No | Above error message appeared during install, brought unit back to repair depot. | Installed new SC and let unit run. | Assisted Replacement | SC | Wetability
ABL90 | after calibrations and waste blood sensitivity climbed to 43 and then dropped to 42 |  | CEAA27FFCA3EDA22 | No | after calibrations and waste blood sensitivity climbed to 43 and then dropped to 42 |  | Assisted Replacement | SC | pCO2
ABL835 FLEX | Kijelző csere. | Kijelző csere. | SWO-21496 | No | Display replacement. | Display replacement. | Defect | HW | DMS_PC | Touch screen
ABL825 FLEX | Trotz Kalibration keine Besserung Austausch und Kalibration FlexQ Modul | Trotz Kalibration keine Besserung Austausch und Kalibration FlexQ Modul | SWO-21353 | No | Despite calibration no improvement exchange and calibration FlexQ module | Despite calibration no improvement exchange and calibration FlexQ module | Defect | HW | FlexQ Module | Barcode reader
AQT90 FLEX | Laser am Storage wheele gereinigt, QC durchgeführt, | Laser am Storage wheele gereinigt, QC durchgeführt, | SWO-21512 | No | Laser cleaned on the storage wheele, QC performed, | Laser cleaned on the storage wheele, QC performed,
TCM-COMBIM | Powered monitor up ran cals and tested on multiple people and was unable to recreate errors. | No Fault Found | CEA9E154CA2750C5 | No | Powered monitor up ran cals and tested on multiple people and was unable to recreate errors. | No Fault Found | No fault found | No fault found | No fault found | No fault found
TCM400 | When the electrodes are connected to the patient skin the electrode does not read the patient oxigenation. it would only show the room air reading. The customer has replaced the membrane on all leads. This is occurring with all 3 modules. | Replaced all 3 O2 modules. | CEA9E5D0CA3F04B1 | No | When the electrodes are connected to the patient skin the electrode does not read the patient oxigenation. it would only show the room air reading. The customer has replaced the membrane on all leads. This is occurring with all 3 modules. | Replaced all 3 O2 modules. | Defect | HW | O2 Module
ABL90 FLEX | Gerät überprüft: am gerät ist unter den Druckoptionen die Auswahlliste für installierte Drucker angewählt, dieses Auswahlfenster wird beim Status an Druckers senden unter die Meldung Ausdruck wird erstellt gelegt dadurch ensteht der Eindruck das Gerät ist nicht mehr bedienbar ->Einstellung abgwählt, Ausdruck erfolgt fehlerfrei, Rep. erfolgt im Rahmen der FAN TU | Gerät überprüft: am gerät ist unter den Druckoptionen die Auswahlliste für installierte Drucker angewählt, dieses Auswahlfenster wird beim Status an Druckers senden unter die Meldung Ausdruck wird erstellt gelegt dadurch ensteht der Eindruck das Gerät ist nicht mehr bedienbar ->Einstellung abgwählt, Ausdruck erfolgt fehlerfrei, Rep. erfolgt im Rahmen der FAN TU | SWO-21864 | No | Device checked: the selection list for installed printers is selected under the print options, this selection window is created when the status to send to printer is created under the message expression is created, creating the impression the device is no longer usable -> Setting is selected, expression is error-free, Rep. within the framework of the FAN TU | Device checked: the selection list for installed printers is selected under the print options, this selection window is created when the status to send to printer is created under the message expression is created, creating the impression the device is no longer usable -> Setting is selected, expression is error-free, Rep. within the framework of the FAN TU | Defect
ABL90 | SUSPECTED BATTERY ISSUE AS ANALZYER BEEPS AND CREEN IS BLANK. | INSTALLED SMART BATTERY CHARGER. | CEAB79D940266663 | No | SUSPECTED BATTERY ISSUE AS ANALZYER BEEPS AND CREEN IS BLANK. | INSTALLED SMART BATTERY CHARGER. | Defect | HW | Battery module
ABL90 FLEX | Datensicherung - Software installiert ( Clean Install ) - Initialisierung und Dateneingabe - Funktionsüberprüfung etc. | Datensicherung - Software installiert ( Clean Install ) - Initialisierung und Dateneingabe - Funktionsüberprüfung etc. | SWO-21757 | No | Data backup - Software installed ( Clean Install ) - Initialization and data entry - Function verification etc. | Data backup - Software installed ( Clean Install ) - Initialization and data entry - Function verification etc. | Defect | SW | Analyzer
ABL825 FLEX |  |  | CEAB27B7403FDC1C | Yes |  |  | Re classified RMED | Pre install OOB
ABL820 FLEX | pCO2 reading on the low end of package insert. | Replaced measuring chamber. Checked air valve on pH module. Checked gas flow. | CEAB2E5C403092FF | No | pCO2 reading on the low end of package insert. | Replaced measuring chamber. Checked air valve on pH module. Checked gas flow. | Assisted Replacement | Electrode | Electrode | pCO2
ABL80 CO-OX | Unable to determine the cause. | Performed system data dump and submit a TechLine case. | CEB00BB0182571B9 | No | Unable to determine the cause. | Performed system data dump and submit a TechLine case.
ABL90 |  |  | CEBAA6F2C82D6497 | No |  |  | Replaced | Inlet gasket
ABL90 FLEX | IP Einstellung vorgenommen BCR Kommunikation vorbereitet. Datensicherung erstellt. | IP Einstellung vorgenommen BCR Kommunikation vorbereitet. Datensicherung erstellt. | SWO-21998 | No | IP setting made BCR communication prepared. data backup. | IP setting made BCR communication prepared. data backup. | Setup | Barcode setup
ABL90 FLEX | Probenweg gespült, Einlassdichtung erneuert | Probenweg gespült; Einlassdichtung erneuert | SWO-21894 | No | Sample path rinsed; Inlet seal renewed | Sample path rinsed; Inlet seal renewed | Cleaned | Fluidic pathway
ABL90 FLEX | IP Einstellung vorgenommen. Serielle Kommunikation für BCR am Analysator eingestellt. Die Barcode Reader sind nicht auf Serielle Kommunikation voreingestellt. Einstellung in der Zentrale veranlasst. 2x Honeywell MK9540-37C41 SN: 2C14430309 und 2C14420166 abgezogen. Diese werden an die Zentrale versendet und dort Konfiguriert und an den Kunden zurück gesendet. Den Anschluss an die ABLs übernimmt Hr. Kühn Dräger TGM. Datensicherung erstellt. |  | SWO-21907 | Yes | IP setting. Serial communication for BCR set on the analyzer. The barcode readers are not preset to serial communication. setting at the head office. 2x Honeywell MK9540-37C41 SN: 2C14430309 and 2C14420166 subtracted. These are sent to the head office, where they are configured and sent back to the customer. Mr. Kühn Dräger TGM takes over the connection to the ABLs. data backup. |  | Re classified RMED | Barcode setup
ABL90 FLEX | Allgemeine Kontrolle ok. JW Kit ersetzt ok. Test im Servicemenü -> Manifold und Vakuum error. Mit neuer Nadel, Solutionpack -> gleiches Problem. ABL 90 Ersatzgerät am Kunde hingestellt ok. Kundengerät wird in Thalwil repariert. (Ersatz Kassette abgegeben und Solutionpack neu). |  | SWO-22241 | No | General control ok. JW Kit replaces ok. Test in the service menu -> Manifold and Vacuum error. With new needle, Solutionpack -> same problem. ABL 90 replacement device placed on the customer ok. Customer device is repaired in Thalwil. (Replacement cassette delivered and Solutionpack new). |  | Unknown | Analyzer replaced
TCM400 | Could not duplicate problem. opened techline case based on pictures. performed alignment check and performance test |  | CEAC6AB0B61FDB3D | No | Could not duplicate problem. opened techline case based on pictures. performed alignment check and performance test
ABL820 FLEX | Checked the autocheck carousel for proper function and found a faulty 14 pin cable | Replaced 14 pin FPC cable and verified proper autocheck function by successfully running all QC ampoules. Instrument is within acceptable operating specifications. | CEAC9CA4B613B979 | No | Checked the autocheck carousel for proper function and found a faulty 14 pin cable | Replaced 14 pin FPC cable and verified proper autocheck function by successfully running all QC ampoules. Instrument is within acceptable operating specifications. | Defect | HW | AutoCheck Module | FPC Cable
AQT90 FLEX |  |  | CEAE10A12C6122D4 | No
TCM/Tosca |  |  | CEACAD9BB62ECD73 | No |  |  | Defect | HW | Power Supply | Power Supply
ABL90 | Found waste line leaking at waste connector. | I replaced the waste line. | CEADA1142C11F483 | No | Found waste line leaking at waste connector. | I replaced the waste line. | Defect | HW | Waste tube
TCM-COMBIM | Display showing different lines of color. Found flat cable faryed. | Replaced cable | CEAF5308A2065576 | No | Display showing different lines of color. Found flat cable faryed. | Replaced cable | Defect | HW | Cable | Cable
ABL80 CO-OX | Investigation performed on previous service work orders. | Replaced the sensor cassette with a known good sensor cassette. | CEB03D431825E9D1 | No | Investigation performed on previous service work orders. | Replaced the sensor cassette with a known good sensor cassette.
ABL80 CO-OX | analyzer install due | preformed analyzer install per the IVW | CEB0614C1832B228 | No | analyzer install due | preformed analyzer install per the IVW
ABL90 | Inspected casette and bay for damage and none was found. | Installed customers casette. | CEB0673818167EC5 | No | Inspected casette and bay for damage and none was found. | Installed customers Casette. | Assisted Replacement | SC | Wetability
ABL90 FLEX | Ich habe mit Pfleger Meier vereinbart, dass sie die Pico 50 benutzen sollten und ich werde mich die Woche jeden Tag melden und Nachfaragen ob der ABL funktioniert. | Sensor interface PCB, Hemolysator, alle Schläuche, Dichtung, Einlasssonde und Einlasshalter ersetzt. Startphase QC und Kalibration. Am 27.03 mit Hr. Meier telefoniert der ABL Funktioniert wenn Bedarf besteht würde sich Hr. Meier melden. | SWO-22910 | No | I have agreed with Pfleger Meier that they should use the Pico 50 and I will report every day every week and see if the ABL works. | Sensor interface PCB, hemolyser, all hoses, seal, inlet probe and inlet holder replaced. Start phase QC and calibration. On 27.03 with Mr. Meier the ABL Works if there is need, Mr. Meier would contact us. | No fault found | Very unsatisfied customer
TCM-COMBIM |  |  | CEB2E57B042D3B0E | No
ABL90 |  |  | CEB088C4182DEF61 | Yes |  |  | Re classified RMED | SWU
ABL800 BASIC | Fehlersuche Demontage und Wartung Ventile Meßmodule Reinigen Einlaßnadel und - Dichtung Lac membraniert Meßkammer gereinigt LS justage, Pumpenkalibrierung Fließwege überprüft Pumpenschlauch BG Modul erneuert | Fehlersuche Demontage und Wartung Ventile Meßmodule Reinigen Einlaßnadel und - Dichtung Lac membraniert Meßkammer gereinigt LS justage, Pumpenkalibrierung Fließwege überprüft Pumpenschlauch BG Modul erneuert | SWO-22134 | No | Troubleshooting Disassembly and maintenance valves Measuring modules Cleaning inlet needle and - Seal Lac membraneed measuring chamber cleaned LS adjustment, pump calibration flow paths checked pump hose BG module renewed | Troubleshooting Disassembly and maintenance valves Measuring modules Cleaning inlet needle and - Seal Lac membraneed measuring chamber cleaned LS adjustment, pump calibration flow paths checked pump hose BG module renewed | Cleaned | Fluidic pathway | - | Cl
ABL820 FLEX | checked flow of solutions, checked waste system | removed and replaced broken inlet nipple, | CEB0BD1E1856BC6C | No | checked flow of solutions, checked waste system | removed and replaced broken inlet nipple, | Defect | HW | Inlet module | Leaking | Nipple Kit
ABL835 FLEX | Neuen Dichtungsring bei PH eingesetzt. Messkammer gereinigt. | Neuen Dichtungsring bei PH eingesetzt. Messkammer gereinigt. | SWO-21994 | No | New sealing ring used at PH. measuring chamber. | New sealing ring used at PH. measuring chamber. | Defect | Electrode | Electrode | pCO2
AQT90 FLEX | Gemäß Anweisung, Gerät aus der GS geholt und vor Ort Gerät ausgetauscht. Ersatzgerät 838R0108N00100. | Gemäß Anweisung, Gerät aus der GS geholt und vor Ort Gerät ausgetauscht. Ersatzgerät 838R0108N00100. | SWO-22028 | No | According to instructions, device was taken out of the GS and replaced on site device. Replacement device 838R0108N00100. | According to instructions, device was taken out of the GS and replaced on site device. Replacement device 838R0108N00100.
AQT90 FLEX | Datensicherung und Service Dump erstellt Data Collection Guide ausgefüllt Rückmeldung techn. Probleme nach Willich Systemüberprüfung und Systemreinigung QC Messung D-Dimer Level 2 |  | SWO-22106 | No | Data backup and service dump created Data Collection Guide filled out feedback technical. Problems after Willich system verification and system cleaning QC measurement D-Dimer Level 2
ABL80 FLEX Basic version |  |  | CEB177C78E204281 | No
ABL90 |  |  | CEB1764C8E58304A | No |  |  | Assisted Replacement | SC | Wetability
ABL825 FLEX | pH BG és El Met modul szelepeinek tisztítása, kalibráció, QC mérés |  | SWO-22211 | No | pH BG and El Met module cleaning, calibration, QC measurement |  | Cleaned | Fluidic pathway | pH/BG Module
ABL825 FLEX | Laser failed to read Ampoule Barcode. Inspection of Laser showed damage to reflective mirror. Mirror was discolored and had a film that wiped off with Q tip. | Replaced AutoCheck Laser Scanner. | CEB196358E5B0B04 | No | Laser failed to read Ampoule Barcode. Inspection of Laser showed damage to reflective mirror. Mirror was discolored and had a film that wiped off with Q tip. | Replaced AutoCheck Laser Scanner. | Defect | HW | AutoCheck Module | Barcode reader
ABL827 FLEX | Crea Messkammer gereinigt, Glaskopf der Crea A+B Elektroden mit warmen Wasser gespült und mit einem Watteträger poliert, anschließend die Crea A+B Elektroden nochmals mit warmen Wasser gespült und membraniert, Gerät kalibriert und QC durchgeführt. Crea Kal 11,3 und OK. QC auch OK. | Crea Messkammer gereinigt, Glaskopf der Crea A+B Elektroden mit warmen Wasser gespült und mit einem Watteträger poliert, anschließend die Crea A+B Elektroden nochmals mit warmen Wasser gespült und membraniert, Gerät kalibriert und QC durchgeführt. Crea Kal 11,3 und OK. QC auch OK. | SWO-23061 | No | Crea measuring chamber cleaned, glass head of the Crea A+B electrodes rinsed with warm water and polished with a cotton carrier, then the Crea A+B electrodes rinsed again with warm water and membraned, device calibrated and QC performed. Crea Kal 11,3 and OK. QC also OK. | Crea measuring chamber cleaned, glass head of the Crea A+B electrodes rinsed with warm water and polished with a cotton carrier, then the Crea A+B electrodes rinsed again with warm water and membraned, device calibrated and QC performed. Crea Kal 11,3 and OK. QC also OK. | Assisted Replacement | Electrode | Membrane | Crea
ABL825 FLEX | FlexQ Modul demontiert, Schraube montiert, Schlittentransport geprüft und justiert, Funktionskontrolle ok | FlexQ Modul demontiert, Schraube montiert, Schlittentransport geprüft und justiert, Funktionskontrolle ok | SWO-22322 | No | FlexQ module disassembled, screw mounted, slide transport tested and adjusted, function control ok | FlexQ module disassembled, screw mounted, slide transport tested and adjusted, function control ok | Adjustment | HW | FlexQ Module
AQT90 FLEX | Fehlersuche, Nicht mehr benötigte Kalibrierungen LQC Daten gelöscht->Verbindung ok Es wird keine Fehlermeldung (Verbindungsproblem) bei den Aktivitätendateien angezeigt System-und Funktionskontrolle i.O. | Fehlersuche, Nicht mehr benötigte Kalibrierungen LQC Daten gelöscht->Verbindung ok Es wird keine Fehlermeldung (Verbindungsproblem) bei den Aktivitätendateien angezeigt System-und Funktionskontrolle i.O. | SWO-22368 | No | Troubleshooting, Unnecessary calibrations LQC data deleted->Connection ok There is no error message (connection problem) displayed in the activity files system and function control i.O. | Troubleshooting, Unnecessary calibrations LQC data deleted->Connection ok There is no error message (connection problem) displayed in the activity files system and function control i.O.
AQT90 FLEX | Fehlersuche, Radiance Connect Disconnect Error(0133 0134)->nicht mehr benötigte Kal. LQCs gelöscht annulliert, Verbindung ok System- und Funktionskontrolle i.O. | Fehlersuche, Radiance Connect Disconnect Error(0133 0134)->nicht mehr benötigte Kal. LQCs gelöscht annulliert, Verbindung ok System- und Funktionskontrolle i.O. | SWO-22366 | No | Troubleshooting, Radiance Connect Disconnect Error(0133 0134)->Kal no longer needed. LQCs deleted cancelled, connection ok system and function control i.O. | Troubleshooting, Radiance Connect Disconnect Error(0133 0134)->Kal no longer needed. LQCs deleted cancelled, connection ok system and function control i.O.
ABL80 CO-OX | Check active parameters, and solution pack (SP). | Turned on pH, pCO2 and p02 parameters, in active parameters. Changed defective SP lot: 211498 and cleaned up bottom of unit. Taught John Vestal about software on ABL80. Verified operation of ABL80 COOX. | CEB51FA0F010AAEB | No | Check active parameters, and solution pack (SP). | Turned on pH, pCO2 and p02 parameters, in active parameters. Changed defective SP lot: 211498 and cleaned up bottom of unit. Taught John Vestal about software on ABL80. Verified operation of ABL80 COOX.
ABL90 | found that when battery is connected unit does not power all the way up and beeps | tried two new battery options, and they did not correct the issue. I have parts on order when i receive them i will come back. | CEB2EB91045A5B4B | No | found that when battery is connected unit does not power all the way up and beeps | tried two new battery options, and they did not correct the issue. I have parts on order when i receive them i will come back. | Defect | HW | Battery module
ABL90 FLEX | Datenbank überprüft Abl weisst öfters den Fehler Sensor hat inhomogene Lösung erkannt und Kein Gerinsel feststellbar Sensor interface PCB bestellt. | Sensor interface PCB, Hemolysator, alle Schläuche, Dichtung, Einlasssonde und Einlasshalter ersetzt. Startphase QC und Kalibration. Am 27.03 mit Hr. Meier telefoniert der ABL Funktioniert wenn Bedarf besteht würde sich Hr. Meier melden. | SWO-22912 | No | Database checks Abl whites often the error sensor has detected inhomogeneous solution and no Gerinsel detectable sensor interface PCB ordered. | Sensor interface PCB, hemolyser, all hoses, seal, inlet probe and inlet holder replaced. Start phase QC and calibration. On 27.03 with Mr. Meier the ABL Works if there is need, Mr. Meier would contact us. | Assisted Replacement | SC | Wetability
ABL800 BASIC | System überprüft und gereinigt, Referenzelektrode erneuert, Kalibrationen und Qualitätskontrollen ohne Beanstandungen durchgeführt. | System überprüft und gereinigt, Referenzelektrode erneuert, Kalibrationen und Qualitätskontrollen ohne Beanstandungen durchgeführt. | SWO-22665 | No | System checked and cleaned, reference electrode renewed, calibrations and quality controls carried out without complaints. | System checked and cleaned, reference electrode renewed, calibrations and quality controls carried out without complaints. | Assisted Replacement | Electrode | Electrode | Reference
ABL827 FLEX |  |  | CEB2BA000409A33C | No |  |  | Defect | HW | Inlet module | Friction detector unit
ABL835 FLEX | Eine krumme Einlaßnadel und eine def. Einlassdichtung waren die Ursache. Dichtung und Nadel sind aus dem Kundenbestand. Mehrere Kallibrationen waren nötig bis alles wieder OK war. | Eine krumme Einlaßnadel und eine def. Einlassdichtung waren die Ursache. Dichtung und Nadel sind aus dem Kundenbestand. Mehrere Kallibrationen waren nötig bis alles wieder OK war. | SWO-22679 | No | A crooked inlet needle and a def. Inlet seal was the cause. Seal and needle are from the customer base. Several calibrations were necessary until everything was OK again. | A crooked inlet needle and a def. Inlet seal was the cause. Seal and needle are from the customer base. Several calibrations were necessary until everything was OK again. | Defect | HW | Inlet module | Probe bend
ABL825 FLEX | Halogenlampe eingestellt StHbKal-, tHb Kal-, Kal2- und QC | Halogenlampe eingestellt StHbKal-, tHb Kal-, Kal2- und QC | SWO-22464 | No | Halogen lamp discontinued StHbKal,tHb Kal-, Kal2- and QC | Halogen lamp discontinued StHbKal,tHb Kal-, Kal2- and QC | Adjustment | HW | Oxi module | Lamp unit
ABL825 FLEX | Fehlersuche StHb Kal ohne Erfolg--> Austausch OXI Modul PCB--> ohne Erfolg Austausch OXI Modul, OXI Complete Test-, StHb Kal-, tHb Kal-, Kal2- und QC durchgeführt und i.O. | Fehlersuche StHb Kal ohne Erfolg--> Austausch OXI Modul PCB--> ohne Erfolg Austausch OXI Modul, OXI Complete Test-, StHb Kal-, tHb Kal-, Kal2- und QC durchgeführt und i.O. | SWO-22510 | Yes | Troubleshooting StHb Kal without success--> Replacement OXI Module PCB--> without success Replacement OXI Module, OXI Complete Test-, StHb Kal-, tHb Kal-, Kal2- and QC performed and i.O. | Troubleshooting StHb Kal without success--> Replacement OXI Module PCB--> without success Replacement OXI Module, OXI Complete Test-, StHb Kal-, tHb Kal-, Kal2- and QC performed and i.O. | Re classified RMED | out of box | Oxi module
AQT90 FLEX | Tests im Servicemenü durchgeführt. LP entfernt. Save Position durchgeführt, Fehler bei Prozess Unit. Weitere Tests durchgeführt und Fehler bei Gabellichtschranke für Position der Optischen Einheit festgestellt. Durch TS erfahren das es eine Charge L09 gibt die von Kopenhagen als Fehlerhaft deklariert wurde. Diese Charge ist im AQT verbaut und im Lager Willich gibt es auch nur die L09. Ersatzgerät angefordert für 26.3.15. Gerät muß zur weiteren Reparatur in GS17. | Gabellichtschranke und Ersatzgerät konnte nicht in die GS17 geliefert werden. Sensor 636-550 aus Demogerät ausgebaut und in defektes Kundengerät montiert. Leider handelt es sich wieder um die L009 Serie, hat aber bisher noch funktioniert. Kunde ( Chefarzt Dr. Koch ) drängt auf eine schnellstmögliche Lösung bez. Verfügbarkeit des AQT. | SWO-23023 | No | tests in the service menu. LP removed. Save position performed, error at process unit. Further tests were carried out and errors in fork light barrier for position of the optical unit detected. Through TS learn that there is a batch L09 that has been declared faulty by Copenhagen. This batch is installed in the AQT and in the Willich warehouse there is only the L09. Replacement device requested for 26.3.15. Device must be in GS17 for further repair. | Fork light barrier and replacement device could not be supplied to the GS17. Sensor 636-550 removed from demo device and mounted in defective customer device. Unfortunately, it is again the L009 series, but has still worked. Customer ( Chief Physician Dr. Koch ) urges a solution as soon as possible, e. availability of the AQT.
ABL90 | LOOKED THROUGH DATA LOGS AND SEE THAT LACTATE HAS BEEN OUT SINCE MARCH 23, 2015. ALSO HAD INLET ERRORS. HAS MESSAGE 1396: SENSOR CASSETTE REPLACEMENT IS RECOMMENDED. | REPLACED BATTERY, REPLACED INLET GASKET TO VERSION 3 AS WELL AS REPLACED INLET PROBE WITH VERSION 3. REPLACED SENSOR CASSETTE WITH CUSTOMERS SUPPLY. | CEB40A3C7A3410F8 | No | LOOKED THROUGH DATA LOGS AND SEE THAT LACTATE HAS BEEN OUT SINCE MARCH 23, 2015. ALSO HAD INLET ERRORS. HAS MESSAGE 1396: SENSOR CASSETTE REPLACEMENT IS RECOMMENDED. | REPLACED BATTERY, REPLACED INLET GASKET TO VERSION 3 AS WELL AS REPLACED INLET PROBE WITH VERSION 3. REPLACED SENSOR CASSETTE WITH CUSTOMERS SUPPLY. | Defect | HW | Battery module
ABL90 FLEX | Sensor interface PCB, Hemolysator, alle Schläuche, Dichtung, Einlasssonde und Einlasshalter ersetzt. Startphase QC und Kalibration. Am 27.03 mit Hr. Meier telefoniert der ABL Funktioniert wenn Bedarf besteht würde sich Hr. Meier melden. | Sensor interface PCB, Hemolysator, alle Schläuche, Dichtung, Einlasssonde und Einlasshalter ersetzt. Startphase QC und Kalibration. Am 27.03 mit Hr. Meier telefoniert der ABL Funktioniert wenn Bedarf besteht würde sich Hr. Meier melden. | SWO-22920 | No | Sensor interface PCB, hemolyser, all hoses, seal, inlet probe and inlet holder replaced. Start phase QC and calibration. On 27.03 with Mr. Meier the ABL Works if there is need, Mr. Meier would contact us. | Sensor interface PCB, hemolyser, all hoses, seal, inlet probe and inlet holder replaced. Start phase QC and calibration. On 27.03 with Mr. Meier the ABL Works if there is need, Mr. Meier would contact us. | Assisted Replacement | SC | Wetability
ABL90 |  |  | CEB413127A16A385 | No |  |  | Assisted Replacement | SC | Wetability
ABL90 FLEX | Fließdynamik mit Testschlauch und SC in Wechselposition OK. Im Meßmodus traten Probleme bei der LP Überprüfung auf. Funktion mit neuer SC im Serviceprogramm OK. SC hätte heute gewechselt werden müssen. SC erneuert, Funktion OK. | Fließdynamik mit Testschlauch und SC in Wechselposition OK. Im Meßmodus traten Probleme bei der LP Überprüfung auf. Funktion mit neuer SC im Serviceprogramm OK. SC hätte heute gewechselt werden müssen. SC erneuert, Funktion OK. | SWO-22808 | No | Flow dynamics with test hose and SC in alternating position OK. In measurement mode, problems occurred during the LP check. Function with new SC in the OK service program. SC should have been changed today. SC renewed, function OK. | Flow dynamics with test hose and SC in alternating position OK. In measurement mode, problems occurred during the LP check. Function with new SC in the OK service program. SC should have been changed today. SC renewed, function OK. | Assisted Replacement | SC | Wetability
ABL90 | Rebooted the system twice. The analyzer screen goes black after couple hours of reboot. Power supply seems faulty. Battery not charging as well. Battery ordered. | Replaced Power supply. | CEB446157A1FF31D | No | Rebooted the system twice. The analyzer screen goes black after couple hours of reboot. Power supply seems faulty. Battery not charging as well. Battery ordered. | Replaced Power supply. | Defect | HW | PS
ABL835 FLEX | LOOKED THROUGH DATA LOGS AND LACTATE LEVEL 3 IS READING LOW, READS 7.4, RANGE IS 10.1 TO 11.8. LOOKING AT THE QC IT LOOKS LIKE THEY GET ABOUT 6 DAYS OUT OF A MEMBRANE. NICK YANG ALSO E-MAILED ME ABOUT 0581 OXI MISMATCH. DID COMPLETE OXIMETRY TEST AND DIDN'T FIND ANYTHING WRONG. ALSO LOOKED AT QC AND CALS FOR THE LAST FEW DAYS AND HAD NO ERRORS. | DID A COMPLETE DECONTAMINATION. REPLACED ALL ELECTRODES MEMBRANES AND ALL SOLUTIONS. | CEB4388E7A201100 | No | LOOKED THROUGH DATA LOGS AND LACTATE LEVEL 3 IS READING LOW, READS 7.4, RANGE IS 10.1 TO 11.8. LOOKING AT THE QC IT LOOKS LIKE THEY GET ABOUT 6 DAYS OUT OF A MEMBRANE. NICK YANG ALSO E-MAILED ME ABOUT 0581 OXI MISMATCH. DID COMPLETE OXIMETRY TEST AND DIDN'T FIND ANYTHING WRONG. ALSO LOOKED AT QC AND CALS FOR THE LAST FEW DAYS AND HAD NO ERRORS. | DID A COMPLETE DECONTAMINATION. REPLACED ALL ELECTRODES MEMBRANES AND ALL SOLUTIONS. | Assisted Replacement | Electrode | Electrode | Lac
ABL825 FLEX |  |  | CEB458577A5D94B2 | No |  |  | Cleaned | Fluidic pathway | Fluidic Module | Valve, air
ABL825 FLEX |  |  | CEB459D47A2994B4 | No |  |  | Cleaned | Fluidic pathway | Fluidic Module | Valve, air
AQT90 FLEX |  |  | CEB3C4E77A1868EB | No
AQT90 FLEX |  |  | CEBA7BA9C85502CC | No
AQT90 FLEX | Aus Ersatzgerät Photointerrupter Proccess Unit ausgebaut. Testlauf und QC erfolgreich. Gerät am 27.03. wieder zurück. | Gabellichtschranke und Ersatzgerät konnte nicht in die GS17 geliefert werden. Sensor 636-550 aus Demogerät ausgebaut und in defektes Kundengerät montiert. Leider handelt es sich wieder um die L009 Serie, hat aber bisher noch funktioniert. Kunde ( Chefarzt Dr. Koch ) drängt auf eine schnellstmögliche Lösung bez. Verfügbarkeit des AQT. | SWO-22799 | No | Removed from replacement device Photointerrupter Proccess Unit. Test run and QC successful. device on 27.03. | Fork light barrier and replacement device could not be supplied to the GS17. Sensor 636-550 removed from demo device and mounted in defective customer device. Unfortunately, it is again the L009 series, but has still worked. Customer ( Chief Physician Dr. Koch ) urges a solution as soon as possible, e. availability of the AQT.
AQT90 FLEX | Gabellichtschranke und Ersatzgerät konnte nicht in die GS17 geliefert werden. Sensor 636-550 aus Demogerät ausgebaut und in defektes Kundengerät montiert. Leider handelt es sich wieder um die L009 Serie, hat aber bisher noch funktioniert. Kunde ( Chefarzt Dr. Koch ) drängt auf eine schnellstmögliche Lösung bez. Verfügbarkeit des AQT. | Gabellichtschranke und Ersatzgerät konnte nicht in die GS17 geliefert werden. Sensor 636-550 aus Demogerät ausgebaut und in defektes Kundengerät montiert. Leider handelt es sich wieder um die L009 Serie, hat aber bisher noch funktioniert. Kunde ( Chefarzt Dr. Koch ) drängt auf eine schnellstmögliche Lösung bez. Verfügbarkeit des AQT. | SWO-23820 | Yes | Fork light barrier and replacement device could not be supplied to the GS17. Sensor 636-550 removed from demo device and mounted in defective customer device. Unfortunately, it is again the L009 series, but has still worked. Customer ( Chief Physician Dr. Koch ) urges a solution as soon as possible, e. availability of the AQT. | Fork light barrier and replacement device could not be supplied to the GS17. Sensor 636-550 removed from demo device and mounted in defective customer device. Unfortunately, it is again the L009 series, but has still worked. Customer ( Chief Physician Dr. Koch ) urges a solution as soon as possible, e. availability of the AQT. | Re classified RMED | JMP(215-4-7) : Re instal of Analyzer after repair in GS
ABL90 FLEX | Analysator überprüft, QC Fehler seit Wechsel der SC. Fällige Einlassdichtung erneuert. Intervall der tHb Kal ausgeschaltet. Nach Austausch der Sensorkassette traten Fehler bei Glu auf. Sc sind bestellt und werden nach Erhalt ausgetauscht. | Analysator überprüft, QC Fehler seit Wechsel der SC. Fällige Einlassdichtung erneuert. Intervall der tHb Kal ausgeschaltet. Nach Austausch der Sensorkassette traten Fehler bei Glu auf. Sc sind bestellt und werden nach Erhalt ausgetauscht. | SWO-22997 | No | Analyzer checked, QC error since changing the SC. Due inlet seal renewed. interval of tHb Kal switched off. After replacing the sensor cartridge, Errors occurred with Glu. Sc are ordered and will be replaced upon receipt. | Analyzer checked, QC error since changing the SC. Due inlet seal renewed. interval of tHb Kal switched off. After replacing the sensor cartridge, Errors occurred with Glu. Sc are ordered and will be replaced upon receipt. | Assisted Replacement | SC | Glu
ABL90 FLEX | SC ausgetauscht, mehrere Kal und QC Messungen durchgeführt. Die Empf. der Lac Elektrode ist noch zu gering, steigt aber langsam an. Personal informiert und gebeten sich wieder zu melden, falls sich die Empf. der Lac Elektrode nicht einstellen sollte. | SC ausgetauscht, mehrere Kal und QC Messungen durchgeführt. Die Empf. der Lac Elektrode ist noch zu gering, steigt aber langsam an. Personal informiert und gebeten sich wieder zu melden, falls sich die Empf. der Lac Elektrode nicht einstellen sollte. | SWO-23036 | No | SC replaced, several Kal and QC measurements performed. The Mpf. the lac electrode is still too low, but is slowly rising. staff informed and asked to come back if the mpf. should not adjust the lac electrode. | SC replaced, several Kal and QC measurements performed. The Mpf. the lac electrode is still too low, but is slowly rising. staff informed and asked to come back if the mpf. should not adjust the lac electrode. | Assisted Replacement | SC | Lac
ABL815 FLEX | Überprüfung und Reinigung Elektrodenkontakte etc. - Erneuerung pH Elektrode - Funktionsüberprüfung etc. | Überprüfung und Reinigung Elektrodenkontakte etc. - Erneuerung pH Elektrode - Funktionsüberprüfung etc. | SWO-22890 | No | Checking and cleaning electrode contacts etc. - Renewal pH electrode - function check etc. | Checking and cleaning electrode contacts etc. - Renewal pH electrode - function check etc. | Assisted Replacement | Electrode | Electrode | pH
ABL800 BASIC | Demontage und Reinigung AC Modul Test ampullenerkennung i.O. | Demontage und Reinigung AC Modul Test ampullenerkennung i.O. | SWO-22835 | No | Disassembly and cleaning AC module test ampoule detection in the o. | Disassembly and cleaning AC module test ampoule detection in the o. | Cleaned | HW | AutoCheck Module | Barcode reader
ABL827 FLEX |  |  | CEB936A052319719 | No |  |  | Unknown | HW | fluidic pathway
ABL815 FLEX | Fehlersuche, div. Programme im Service durchgef., Spüllsg. + div. Schlauchmaterial gew., QC Level 2 neu eingelesen, QC-Plan angepaßt, Kal. + QC i.O. | Fehlersuche; div. Programme im Service durchgef.; Spüllsg. + div. Schlauchmaterial gew.; QC Level 2 neu eingelesen; QC-Plan angepaßt; Kal. + QC i.O. | SWO-22924 | No | Troubleshooting; various programs in the service; Rinse sg. + various hose material. QC Level 2 re-read; QC plan; Kal. + QC in the first place | Troubleshooting; various programs in the service; Rinse sg. + various hose material. QC Level 2 re-read; QC plan; Kal. + QC in the first place | Cleaned | Fluidic pathway | -
ABL815 FLEX | Es war eine alte Elektrode verbaut, Membran leer gelaufen, Membran ersetzt, Gas neu eingelesen, weiterhin Kal Fehler, Elektrode ersetzt und einlaufen lassen, Ventile BG Modul überprüft, Kontakt Konnektor PCO2 justiert, Luftkanal gespült, Kal und QC durchgeführt | Es war eine alte Elektrode verbaut, Membran leer gelaufen, Membran ersetzt, Gas neu eingelesen, weiterhin Kal Fehler, Elektrode ersetzt und einlaufen lassen, Ventile BG Modul überprüft, Kontakt Konnektor PCO2 justiert, Luftkanal gespült, Kal und QC durchgeführt | SWO-23008 | No | It was an old electrode installed, membrane run empty, membrane replaced, gas re-read, continue kal error, electrode replaced and let run in, valves BG module checked, contact connector PCO2 adjusted, air duct flushed, Kal and QC performed | It was an old electrode installed, membrane run empty, membrane replaced, gas re-read, continue kal error, electrode replaced and let run in, valves BG module checked, contact connector PCO2 adjusted, air duct flushed, Kal and QC performed | Assisted Replacement | Electrode | Electrode | pCO2
ABL815 FLEX | E1001 Lot 1701 und E788 Lot RI02 ausgebaut und erneuert, Kal und QC z.Zt. OK | E1001 Lot 1701 und E788 Lot RI02 ausgebaut und erneuert, Kal und QC z.Zt. OK | SWO-23056 | No | E1001 Lot 1701 and E788 Lot RI02 expanded and renewed, Kal and QC currently OK | E1001 Lot 1701 and E788 Lot RI02 expanded and renewed, Kal and QC currently OK | Assisted Replacement | Electrode | Electrode | pCO2
ABL825 FLEX | QC Karussell war nicht in der richtigen Lage ( etwas zu hoch), Lac-Elektrode neu membraniert (Lot 2548), Kal, QC | QC Karussell war nicht in der richtigen Lage ( etwas zu hoch), Lac-Elektrode neu membraniert (Lot 2548), Kal, QC | SWO-22987 | Yes | QC carousel was not in the right position (slightly too high), Lac electrode re-logged (lot 2548), Kal, QC | QC carousel was not in the right position (slightly too high), Lac electrode re-logged (lot 2548), Kal, QC | Re classified RMED | out of box | AutoCheck Module | carousel
ABL90 FLEX | Gerät überprüft und gespült. Div. Solutions werden nicht erkannt.Inlet ok. SP entnommen und nochmal aktiviert.Geräusch deutlich hörbar. SP erneut eingesetzt. | Gerät überprüft und gespült. Div. Solutions werden nicht erkannt.Inlet ok. SP entnommen und nochmal aktiviert.Geräusch deutlich hörbar. SP erneut eingesetzt. | SWO-23024 | No | device checked and rinsed. Div. Solutions are not recognized. Inlet ok. SP and activated again. Noise clearly audible. SP re-inserted. | device checked and rinsed. Div. Solutions are not recognized. Inlet ok. SP and activated again. Noise clearly audible. SP re-inserted. | User Error | SP activation
ABL835 FLEX | Checked machines. Both in good working order. Service THB calibrations will tighten parameters to customer specs. | Performed Protein remover. Performed 1 pt calibration. Performed service THB calibration. | CEB8040EDC42334E | No | Checked machines. Both in good working order. Service THB calibrations will tighten parameters to customer specs. | Performed Protein remover. Performed 1 pt calibration. Performed service THB calibration. | Adjustment | HW | Oxi module | tHb cal
ABL80 OSM | Analyzer failed level one external Qc. | I replaced the hemolyzer Performed THB. Brought system to ready. This completes the installation. | CEB957DC524D6265 | No | Analyzer failed level one external Qc. | I replaced the hemolyzer Performed THB. Brought system to ready. This completes the installation.
ABL90 |  |  | CEB949B2521BAA5D | Yes |  |  | Re classified RMED | Software upgrade
ABL800 BASIC | Einlaßnadel ausgetauscht Einlaßdichtung erneuert | Einlaßnadel ausgetauscht Einlaßdichtung erneuert | SWO-23286 | No | Inlet needle replaced inlet seal renewed | Inlet needle replaced inlet seal renewed | Defect | HW | Inlet module | Probe bend
ABL835 FLEX | supsect issue with power supply | replaced power supply, same error inspected and tested fuses, passed supsect issue with AU board, replaced Au board reloaded software and options | CEB9701F520E8F2A | No | supsect issue with power supply | replaced power supply, same error inspected and tested fuses, passed supsect issue with AU board, replaced Au board reloaded software and options | Defect | HW | Analyzer Control PCB
ABL835 FLEX | Kijelző kábel cseréje, kijelző kalibrálás, belső kábelek ellenőrzése, QC mérés, kalibrálások |  | SWO-23401 | No | Display cable replacement, display calibration, internal cable control, QC measurement, calibration |  | Defect | HW | DMS_PC | Touch screen | Cable connection
AQT90 FLEX |  |  | CEB92BEE52303ADB | No
ABL80 FLEX CO-OX | Hemolyser tubing kit ersetzt und geprüft. | Hemolyser tubing kit ersetzt und geprüft. | SWO-23340 | No | Hemolyser tubing kit replaced and tested. | Hemolyser tubing kit replaced and tested.
ABL80 CO-OX | Found leakage around manifold ports. Current cassette had no tests remaining. | Cleaned manifold. Installed cassette. | CEBA4E9AC83791DA | No | Found leakage around manifold ports. Current cassette had no tests remaining. | Cleaned manifold. Installed cassette.
ABL90 FLEX | Fehlersuche, Dichtung + Nadel ersetzt, Neustart, Kal. + QC i.O. | Fehlersuche; Dichtung + Nadel ersetzt; Neustart; Kal. + QC i.O. | SWO-23391 | No | Troubleshooting; Seal + needle replaced; Restart; Kal. + QC in the first place | Troubleshooting; Seal + needle replaced; Restart; Kal. + QC in the first place | Defect | HW | Inlet Gasket
ABL90 FLEX | Drucker erneuert | Drucker erneuert | SWO-23465 | No | Printer renewed | Printer renewed | Defect | HW | Printer
ABL825 FLEX | Lamp is intensity is high. | Adjusted lamp output and calibrated OXI module per Radiometer specifications. All QC is on or near mean. Instrument operates normally. | CEBA1E84C85FE77A | No | Lamp is intensity is high. | Adjusted lamp output and calibrated OXI module per Radiometer specifications. All QC is on or near mean. Instrument operates normally. | Other | Adjustment | Lamp unit
ABL837 FLEX | Reviewed data logs, cals went out after CAL1 bottle was changed. It still had the foil on the dosi cap. Also Cleaning solution bottle still had its foil still on | Replaced CAL1 and cleaning with new bottles | CEBA5AE5C84DAF7F | No | Reviewed data logs, cals went out after CAL1 bottle was changed. It still had the foil on the dosi cap. Also Cleaning solution bottle still had its foil still on | Replaced CAL1 and cleaning with new bottles | Assisted Replacement | Solution bottle | - | Crea
ABL837 FLEX | Reviewed data logs found the error occurred after CAL 1 solution was changed. Dosi cap was not inverted, tablet nevr went into solution | Replaced CAL 1 solution, also Cleaning bottle still had the foil on thr dosi cap | CEBA57FFC85CAF7C | No | Reviewed data logs found the error occurred after CAL 1 solution was changed. Dosi cap was not inverted, tablet nevr went into solution | Replaced CAL 1 solution, also Cleaning bottle still had the foil on thr dosi cap | Assisted Replacement | Solution bottle | - | Crea
AQT90 FLEX | Bei Ankunft AQT Messbereit. Freu Breilling hat mir das Problem geschildert das am Vortag aufgetreten ist. Frontcover abgebaut und Türe im Service mehrmals auf und zu gefahren. Alles i.O. Systemüberprüfung im Zuge des hardwareupdates mit durchgeführt. Keine Fehler feststellbar. AQT Messbereit. | Bei Ankunft AQT Messbereit. Freu Breilling hat mir das Problem geschildert das am Vortag aufgetreten ist. Frontcover abgebaut und Türe im Service mehrmals auf und zu gefahren. Alles i.O. Systemüberprüfung im Zuge des hardwareupdates mit durchgeführt. Keine Fehler feststellbar. AQT Messbereit. | SWO-23359 | No | Ready for measurement on arrival AQT. Freu Breilling told me the problem that occurred the day before. Front cover dismantled and door in service several times up and to drive. All i.O. system check performed in the course of the hardware update. No errors detectable. AQT Ready to measure. | Ready for measurement on arrival AQT. Freu Breilling told me the problem that occurred the day before. Front cover dismantled and door in service several times up and to drive. All i.O. system check performed in the course of the hardware update. No errors detectable. AQT Ready to measure.
AQT90 FLEX |  |  | CEBCCC3FB4349739 | Yes |  |  | Re classified RMED | FAN / TU
ABL90 |  |  | CEBBAE713E044957 | Yes |  |  | Re classified RMED | Software upgrade
ABL90 FLEX | Einlass und Schlauchsystem geprüft, Einlassdichtung ausgetauscht, Funktionskontrolle ok | Einlass und Schlauchsystem geprüft, Einlassdichtung ausgetauscht, Funktionskontrolle ok | SWO-23592 | No | Inlet and hose system tested, inlet seal replaced, function control ok | Inlet and hose system tested, inlet seal replaced, function control ok | Assisted Replacement | SC | Wetability
ABL90 FLEX | Haken DST muss aber gesetzt werden an allen Geräten Einstellung vorgenommen und gestestet | Haken DST muss aber gesetzt werden an allen Geräten Einstellung vorgenommen und gestestet | SWO-23725 | No | Hook DST must be set however on all devices setting and tested | Hook DST must be set however on all devices setting and tested
ABL825 FLEX |  |  | CEBB8F103E048335 | No |  |  | Cleaned | Fluidic pathway | pH/BG Module
ABL825 FLEX |  |  | CEBB91093E5C8338 | No |  |  | Unknown
ABL827 FLEX |  |  | CEC39869024D8D27 | No |  |  | Assisted Replacement | Electrode | Electrode | pCO2
ABL825 FLEX | Left side of display un responsive to touch. | Replaced display. | CEBC05103E05007E | No | Left side of display un responsive to touch. | Replaced display. | Defect | HW | Display
ABL820 FLEX | Glean states that the external barcode reader intermittently fails when a reagent has to be replaced. it will not scan the barcodes. The customer has to reboot the analyzer to make it work. | Replaced a new external barcode reader. Replaced the pump tubing from customer's stock. Flushed system lines. Ran a service liquid sensor adjust . | CEBB72703E4C955E | Yes | Glean states that the external barcode reader intermittently fails when a reagent has to be replaced. it will not scan the barcodes. The customer has to reboot the analyzer to make it work. | Replaced a new external barcode reader. Replaced the pump tubing from customer's stock. Flushed system lines. Ran a service liquid sensor adjust . | Re classified RMED | HW | External Barcode
ABL825 FLEX | PO2 Elektrode ersetzt..neue Membran anders Lot....Gas1--2 ersetzt. Empf.10 pA Nullwert 9,4 Elektrode zurückgetauscht mit altem Lot. 3xx Membraniert Nullwert 0,4 | PO2 Elektrode ersetzt..neue Membran anders Lot....Gas1--2 ersetzt. Empf.10 pA Nullwert 9,4 Elektrode zurückgetauscht mit altem Lot. 3xx Membraniert Nullwert 0,4 | SWO-23730 | No | PO2 electrode replaced. new membrane different lot.... Gas1--2. Empf.10 pA zero value 9.4 electrode replaced with old lot. 3xx Membraneed zero value 0.4 | PO2 electrode replaced. new membrane different lot.... Gas1--2. Empf.10 pA zero value 9.4 electrode replaced with old lot. 3xx Membraneed zero value 0.4 | Assisted Replacement | Electrode | Membrane | pO2
ABL815 FLEX | Inlet Kal mehrfach durchgeführt. Aus Gerät Sportmedizin. Inlet Modul ausgebaut in ABL eingesetzt. Initialisierung Kal und QC ok Spare Part bestellt. | Inlet Kal mehrfach durchgeführt. Aus Gerät Sportmedizin. Inlet Modul ausgebaut in ABL eingesetzt. Initialisierung Kal und QC ok Spare Part bestellt. | SWO-23729 | No | Inlet Kal performed several times. From device sports medicine. Inlet module extended used in ABL. Initialization Kal and QC ok Spare Part ordered. | Inlet Kal performed several times. From device sports medicine. Inlet module extended used in ABL. Initialization Kal and QC ok Spare Part ordered. | Defect | HW | Inlet module | Spinlet Control PCB MKII
ABL825 FLEX |  |  | CEBC7CDBB4613863 | No |  |  | Defect | HW | Inlet module | Leaking | Nipple Kit
AQT90 FLEX |  |  | CEBB97523E026BE2 | Yes |  |  | Re classified RMED | FAN / TU
AQT90 FLEX | QC Messungen im Bereich. Via Analysetools (Kundenanschreiben) und SD Analysetool Messungen und Fließgeschwindigkeit geprüft. Keine Fehler oder Abweichungen ersichtlich. Analysator messbereit. Kein Fehler feststellbar | QC Messungen im Bereich. Via Analysetools (Kundenanschreiben) und SD Analysetool Messungen und Fließgeschwindigkeit geprüft. Keine Fehler oder Abweichungen ersichtlich. Analysator messbereit. Kein Fehler feststellbar | SWO-23697 | No | QC measurements in the range. Measurements and flow rate are checked via analysis tools (customer letter) and SD analysis tool. No errors or deviations can be seen. Analyzer ready to measure. No error detectable | QC measurements in the range. Measurements and flow rate are checked via analysis tools (customer letter) and SD analysis tool. No errors or deviations can be seen. Analyzer ready to measure. No error detectable
AQT90 FLEX |  |  | CEBBA2523E443384 | No
ABL90 FLEX | Fehlersuche SW aktualisiert Neue SC90 installiert | Fehlersuche SW aktualisiert Neue SC90 installiert | SWO-23813 | No | Troubleshooting SW updated New SC90 installed | Troubleshooting SW updated New SC90 installed | Assisted Replacement | SC | Glu
ABL90 |  |  | CEC2F03F023322B9 | No
ABL90 | Screen is blank | Replaced power supply installed new sensor cassette and solution Pack | CEBCC336B44F40E4 | No | Screen is blank | Replaced power supply installed new sensor cassette and solution Pack | Defect | HW | PS
ABL837 FLEX | Confirmed by visual inspection V8 leaking. | Removed and cleaned valve position, installed new V8. Tested using service diagnostics, no errors reported in diagnostics. | CEBCFC74B45F10A3 | No | Confirmed by visual inspection V8 leaking. | Removed and cleaned valve position, installed new V8. Tested using service diagnostics, no errors reported in diagnostics. | Defect | HW | Fluidic manifold | V8
ABL835 FLEX |  |  | CEBD0B17B45DA418 | Yes |  |  | Re classified RMED | HW | 6 month service
ABL820 FLEX | ABL820 would not run QC. The carousel would spin and stop thin go back down, it did not attempt to open a ampoule. | Replaced the 14 pole FPC cable. | CEBD2D0FB44806EA | No | ABL820 would not run QC. The carousel would spin and stop thin go back down, it did not attempt to open a ampoule. | Replaced the 14 pole FPC cable. | Defect | HW | AutoCheck Module | FPC Cable
ABL825 FLEX | AC Modul geprüft, Barcode Reader geprüft, Funktionskontrolle ok | AC Modul geprüft, Barcode Reader geprüft, Funktionskontrolle ok | SWO-23859 | No | AC module checked, barcode reader checked, function control ok | AC module checked, barcode reader checked, function control ok | Adjustment | HW | AutoCheck Module | carousel
ABL825 FLEX | Messkammer und pO2 Elektrode geprüft, pO2 Elektrode (E799 RI03) ausgetauscht, Funktionskontrolle ok | Messkammer und pO2 Elektrode geprüft, pO2 Elektrode (E799 RI03) ausgetauscht, Funktionskontrolle ok | SWO-23833 | No | Measuring chamber and pO2 electrode tested, pO2 electrode (E799 RI03) replaced, function control ok | Measuring chamber and pO2 electrode tested, pO2 electrode (E799 RI03) replaced, function control ok | Assisted Replacement | Electrode | Electrode | pO2
AQT90 FLEX | Calibratie blancometingen zijn te hoog (hoger dan 1000). Reiniging van optical unit, process wheel en Refelectorschijf lost het probleem ook niet op. Zodoende de optical unit heb vervangen en afgesteld. Vervolgens een TNT-calibratie heb uitgevoerd, OK. | Calibratie blancometingen zijn te hoog (hoger dan 1000). Reiniging van optical unit, process wheel en Refelectorschijf lost het probleem ook niet op. Zodoende de optical unit heb vervangen en afgesteld. Vervolgens een TNT-calibratie heb uitgevoerd, OK. | SWO-23853 | No | Calibration blancommeasurements are too high (higher than 1000). Cleaning of optical unit, process wheel and Refelector disk also does not solve the problem, so the optical unit has been replaced and adjusted. Then performed a TNT calibration, OK. | Calibration blancommeasurements are too high (higher than 1000). Cleaning of optical unit, process wheel and Refelector disk also does not solve the problem, so the optical unit has been replaced and adjusted. Then performed a TNT calibration, OK.
ABL80 FLEX Basic version |  |  | CEBE3C402A1EF883 | No
ABL800 BASIC | Fehlersuche = Netzteil def., Gerät aus, Netzteil gew., Gerät ein, lamp unit neu eingestellt, Service- u. Anwender-tHb-Kal. durchgef., Spüllsg. gew., AC aufgefüllt, Kal. = pCO2-Fehler, pCO2 neu membr., Kal. weiter Fehler, neue Elektrode in gl. Membran getestet, Kal. i.O., QC-Fehler tHb + pCO2, wieder Oxitest -> i.O., Hemolyzer gew., pH BGF-modul ausgebaut u. Ventile gereinigt, wieder div. Kal., Kal. = Fehler Ca + Lac, Ca neu membraniert, Kal außer Lac i.O., Lac abgeschaltet u. Anwender informiert, QC i.O. | Fehlersuche = Netzteil def.; Gerät aus; Netzteil gew.; Gerät ein; lamp unit neu eingestellt; Service- u. Anwender-tHb-Kal. durchgef.; Spüllsg. gew.; AC aufgefüllt; Kal. = pCO2-Fehler; pCO2 neu membr.; Kal. weiter Fehler; neue Elektrode in gl. Membran getestet; Kal. i.O.; QC-Fehler tHb + pCO2; wieder Oxitest -> i.O.; Hemolyzer gew.; pH BGF-modul ausgebaut u. Ventile gereinigt; wieder div. Kal.; Kal. = Fehler Ca + Lac; Ca neu membraniert; Kal außer Lac i.O.; Lac abgeschaltet u. Anwender informiert; QC i.O. | SWO-23966 | No | Troubleshooting = power supply def.; device off; power supply. device; lamp unit re-hired; Service and user tHb-Kal. through it; Rinse sg. w.; AC replenished; Kal. = pCO2 error; pCO2 new membr.; Kal. further errors; new electrode in gl. membrane tested; Kal. i.O.; QC error tHb + pCO2; again Oxitest -> i.O.; Hemolyzer. pH BGF module removed and valves cleaned; again div. Kal.; Kal. = Error Ca + Lac; Ca re-membraneed; Kal except Lac i.O.; Lac switched off and informs users; QC in the first | Troubleshooting = power supply def.; device off; power supply. device; lamp unit re-hired; Service and user tHb-Kal. through it; Rinse sg. w.; AC replenished; Kal. = pCO2 error; pCO2 new membr.; Kal. further errors; new electrode in gl. membrane tested; Kal. i.O.; QC error tHb + pCO2; again Oxitest -> i.O.; Hemolyzer. pH BGF module removed and valves cleaned; again div. Kal.; Kal. = Error Ca + Lac; Ca re-membraneed; Kal except Lac i.O.; Lac switched off and informs users; QC in the first | Defect | HW | Power Supply
ABL90 |  |  | CECD7D00B258CC6B | No |  |  | Defect | HW | Inlet Gasket
ABL90 | LOOKED AT DATA LOGS AND FOUND ON APR 4, 2015 THE ANALYZER HAD A OXI SPECTRUM MISMATCH 0581. ALSO AFTER LOOKED LIKE LIQUID SENSOR HAD SOME LEAKAGE AROUND IT. | REPLACED HEMOLYZER, LIQUID SENSOR AND INLET PROBE V3. | CEC1DF458C1E3662 | No | LOOKED AT DATA LOGS AND FOUND ON APR 4, 2015 THE ANALYZER HAD A OXI SPECTRUM MISMATCH 0581. ALSO AFTER LOOKED LIKE LIQUID SENSOR HAD SOME LEAKAGE AROUND IT. | REPLACED HEMOLYZER, LIQUID SENSOR AND INLET PROBE V3. | Defect | HW | Hemolyzer
ABL837 FLEX | Reviewed data logs sensitivity error occured after membrane change | Remembraned with same lot # still high sensitivity. Cleaned contacts,replaced the membranes and still have sentitivity error, too high.Tried another lot and after multiple 2pt Cals sensitivity is within range. Complete OXI alignment, Hb cal in service, User to monitor WDC reports. | CEC205B88C16B533 | No | Reviewed data logs sensitivity error occured after membrane change | Remembraned with same lot # still high sensitivity. Cleaned contacts,replaced the membranes and still have sentitivity error, too high. Tried another lot and after multiple 2pt Cals sensitivity is within range. Complete OXI alignment, Hb cal in service, User to monitor WDC reports. | Defect | Membrane | Crea
ABL827 FLEX | Upon arrival, found Autocheck module housing not staying in place to run QCs. | Replaced FPC cable. | CEC1B26A8C1AF68B | No | Upon arrival, found Autocheck module housing not staying in place to run QCs. | Replaced FPC cable. | Defect | HW | AutoCheck Module | FPC Cable
ABL820 FLEX | Verified non power on by power switch. Changed receptacles, no change. | Replaced power supply. After power supply replacement, analyzer had 0537 Oxi LS error. Flushed air intake & performed LS adjust, all ok. | CEC1C7E98C292BE4 | No | Verified non power on by power switch. Changed receptacles, no change. | Replaced power supply. After power supply replacement, analyzer had 0537 Oxi LS error. Flushed air intake & performed LS adjust, all ok. | Defect | HW | Power Supply
ABL820 FLEX | ANALYZER WORKING UPON ARRIVAL.DATA LOGS SHOW MULTIPLE ERRORS 751 AND 508 ALSO 602 INLET NOT FILLED AND 765 INLET FLOW IMPEDED.VISUALLY OBSERVED LOOSE INLET PROBE TUBING AND BENT PROBE | REPLACED INLET PROBE AND TUBING.REPLACED SOLUTION PUMP TUBING.PERFORMED INLET CALIBRATION,PUMP CALIBRATION AND LIQUID SENSOR ADJUSTMENT.PERFORMED 1 AND 2 POINT CALIBRATIONS | CEC209688C0D5844 | No | ANALYZER WORKING UPON ARRIVAL. DATA LOGS SHOW MULTIPLE ERRORS 751 AND 508 ALSO 602 INLET NOT FILLED AND 765 INLET FLOW IMPEDED. VISUALLY OBSERVED LOOSE INLET PROBE TUBING AND BENT PROBE | REPLACED INLET PROBE AND TUBING. REPLACED SOLUTION PUMP TUBING. PERFORMED INLET CALIBRATION,PUMP CALIBRATION AND LIQUID SENSOR ADJUSTMENT. PERFORMED 1 AND 2 POINT CALIBRATIONS | Defect | Inlet | Inlet Module | Probe bend
ABL825 FLEX | Rinse error was caused by the rinse valve in the PH BG module failing to open. | Replaced valve and normal operation was restored. Calibrated and ran QC. Instrument operates normally. | CEC1EEAD8C30B051 | No | Rinse error was caused by the rinse valve in the PH BG module failing to open. | Replaced valve and normal operation was restored. Calibrated and ran QC. Instrument operates normally. | Defect | HW | pH/BG Module | Valve rinse
ABL80 OSM | found that they keep getting it on system cycle | open analyzer up and moved the tube that supplies fuild from sp to not get pinched | CEC3407A025E93EC | No | found that they keep getting it on system cycle | open analyzer up and moved the tube that supplies fuild from sp to not get pinched
ABL90 |  |  | CEC32CA402552B6C | No |  |  | Assisted Replacement | SC | Wetability
ABL90 |  |  | CEC320250223D907 | No |  |  | Replaced | Probe | Bend
ABL90 FLEX | Teile im Umfang einer Jahreswartung erneuert. Lösungspack und Sensorkassette entnommen und wieder installiert. Fehlerfreies und schnelles Einlaufverhalten. Alle Funktionen ( Servicetests ) Kal. und QC o.k. | Teile im Umfang einer Jahreswartung erneuert. Lösungspack und Sensorkassette entnommen und wieder installiert. Fehlerfreies und schnelles Einlaufverhalten. Alle Funktionen ( Servicetests ) Kal. und QC o.k. | SWO-24809 | No | Parts to the extent of an annual maintenance. Solution pack and sensor cartridge removed and reinstalled. Error-free and fast inlet behavior. All functions ( service tests ) Kal. and QC o.k. | Parts to the extent of an annual maintenance. Solution pack and sensor cartridge removed and reinstalled. Error-free and fast inlet behavior. All functions ( service tests ) Kal. and QC o.k. | Unknown | HW | Yearly maintenance kit
ABL5 | Schläuche der internen Gasversorgung erneuert, Systemkontrolle iO. | Schläuche der internen Gasversorgung erneuert, Systemkontrolle iO. | SWO-24091 | No | Hoses of internal gas supply renewed, system control iO. | Hoses of internal gas supply renewed, system control iO.
ABL800 BASIC | Lac-Elektrode neu membraniert, Kal, QC | Lac-Elektrode neu membraniert, Kal, QC | SWO-24126 | No | Lac electrode re-membraneed, Kal, QC | Lac electrode re-membraneed, Kal, QC | Assisted Replacement | Electrode | Membrane | Lac
ABL827 FLEX |  |  | CEC3596902512242 | No |  |  | Defect | Inlet | Inlet Module | Probe bend
ABL825 FLEX | The ampoule opener was intermittantly not opening the ampoules. | Replaced the motor to the ampoule opener. | CEC3073502156E30 | No | The ampoule opener was intermittantly not opening the ampoules. | Replaced the motor to the ampoule opener. | Defect | HW | AutoCheck Module | Motor
ABL835 FLEX | ABL neu gebootet ,Kalibration und QC durchgeführt |  | SWO-24037 | No | ABL rebooted,calibration and QC performed |  | Defect | HW | Power Supply
ABL820 FLEX | Performed service diagnostics, Oximetry complete test, indicates Lamp unit failure . Neon position Fneon: 2.093, limit = 0.100 - 2.000. | Replaced Spectrometer to correct error 315 problem. Performed lamp adjustment, Complete Oximetry Test - Passed, tHB calibration. No errors observed. Proactively performed Inspection of fluid transport system, verified rinse and waste tubing connections and routing were acceptable, confirmed valve performance, cleaned measuring chambers and electrodes as needed. Returned 4 9 for a courtesy visit to confirm performance of analyzer and answer any questions. No further problems reported. | CEC3551702631B16 | No | Performed service diagnostics, Oximetry complete test, indicates Lamp unit failure . Neon position Fneon: 2.093, limit = 0.100 - 2.000. | Replaced Spectrometer to correct error 315 problem. Performed lamp adjustment, Complete Oximetry Test - Passed, tHB calibration. No errors observed. Proactively performed Inspection of fluid transport system, verified rinse and waste tubing connections and routing were acceptable, confirmed valve performance, cleaned measuring chambers and electrodes as needed. Returned 4 9 for a courtesy visit to confirm performance of analyzer and answer any questions. No further problems reported. | Defect | HW | Spectrometer | Neon high
ABL825 FLEX | Oxi Complete Test zunächst ohne Beanstandung - bei 2 Punkt KAL wieder Fehlermeldung 315 - Erneuerung Oxi Modul,Einlassnadel und Einlassdichtung ( Einlass Kalibrierfehler ) - Einstellungen und Justagen etc. - Funktionsüberprüfung etc. - Lac und Glu Kalibrierfehler - Membranierungen etc. | Oxi Complete Test zunächst ohne Beanstandung - bei 2 Punkt KAL wieder Fehlermeldung 315 - Erneuerung Oxi Modul,Einlassnadel und Einlassdichtung ( Einlass Kalibrierfehler ) - Einstellungen und Justagen etc. - Funktionsüberprüfung etc. - Lac und Glu Kalibrierfehler - Membranierungen etc. | SWO-24057 | No | Oxi Complete Test first without complaint - at 2 point KAL again error message 315 - Renewal Oxi module, inlet needle and inlet seal ( inlet calibration error ) - Settings and adjustments etc. - Function check etc. - Lac and Glu calibration error - membranes, etc. | Oxi Complete Test first without complaint - at 2 point KAL again error message 315 - Renewal Oxi module, inlet needle and inlet seal ( inlet calibration error ) - Settings and adjustments etc. - Function check etc. - Lac and Glu calibration error - membranes, etc. | Defect | HW | Oxi module | Spectrometer
ABL835 FLEX | Checked fluid in front of pco2 electrode (none), Checked pH electrode (blue) ran calibrations and qc. | unable to correct problem- open techline case | CEC3A0F802007A19 | No | Checked fluid in front of pco2 electrode (none), Checked pH electrode (blue) ran calibrations and qc. | unable to correct problem- open techline case | Assisted Replacement | Electrode | Electrode | PCO2
ABL825 FLEX | Checked pH oring,checked for fliud in front on pco2 electrode,ran calibrations and external qc | unable to correct - open techline case | CEC3A2E902157A1D | No | Checked pH oring,checked for fliud in front on pco2 electrode,ran calibrations and external qc | unable to correct - open techline case | Assisted Replacement | Electrode | Electrode | PCO2
AQT90 FLEX | Neue Pipettennadel eingebaut, Reinigung durchgeführt, |  | SWO-24293 | No | New pipette needle installed, cleaning carried out,
AQT90 FLEX | Arbeiten: AQT90Flex lokal per Crosskabel mit Test LIS-Simulator Laptop verbunden. Test Query vom AQT90Flex an den LIS Simulator erfolgreich => OK Original ASTM Response vom LIS-Molis über den LIS Simulator an AQT90 gesendet =>OK AQT90Flex wieder an das Netzwerk angeschlossen Trace-Level Debugging am AQT90Flex aktiviert. Mitschnitt der Datenkommunikation zur Fehlersuche. Kommunikationsanalyse des Datenverkehrs. ASTM HighLevel =>OK ASTM LowLevel => fehlerhaft Ursache: In der Response vom LIS-Treiber zum AQT90Flex fehlte beim Abschluss- L Segment eine <13> Steuerzeichen. Trace: molis: Apr 07 11:08:18 - NETWORK - [L] - <1>H¦ ^&¦¦¦SYN-MOLIS¦¦¦¦¦¦¦¦1¦20150407110858<13>P¦1¦¦209000193901¦209000193901¦TEST¦¦¦¦¦¦¦¦¦¦¦¦¦¦¦¦¦¦¦¦<13>L¦1¦N<4> Aktuell: L¦1¦N<4> eine <13> fehlt !! Richtig wäre: L¦1¦N<13><4> Rücksprache mit Frau Leoanhard IT Synlab: Gesendetes Trace.log besprochen und erklärt. Entsprechende fehlerhafte Stelle wurde seitens des LIS Treiber korrigiert Jetzt mit <13> Testlauf am AQT90: -Neuen Auftrag im LIS angelegt -Probe am AQT90 eingelegt -Scan und Query der Auftrags-ID -Response (Stammdaten) wurden jetzt vom AQT90 übernommen und während der Analyse im Display angezeigt. -Befund wurde nach der Messung autom. an den LIS Treiber gesendet. Prüfung der Datenübernahme im LIS System durch Laborpersonal => OK KH Trostberg: Änderung wurde auch am LIS Treiber für den AQT90Flex im KH Trostberg durch Frau Leonhard durchgeführt. Kontakt mit Labor in Trostberg hergestellt: -Testmessung und Query mit Rückmeldung der Stammdaten am AQT90Flex erfolgreich |  | SWO-24097 | No | Work: AQT90Flex locally connected via cross cable to test LIS simulator laptop. Test query from AQT90Flex to the LIS Simulator successful => OK Original ASTM Response from the LIS-Molis sent via the LIS Simulator to AQT90 =>OK AQT90Flex reconnected to the network Trace-Level Debugging on the AQT90Flex activated. Recording of data communication for troubleshooting. Communication analysis of traffic. ASTM HighLevel =>OK ASTM LowLevel => Faulty Cause: In the response from the LIS driver to the AQT90Flex, the closing L segment lacked a .13> control character. Trace: molis: Apr 07 11:08:18 - NETWORK - [L] - .H¦ SYN-MOLIS¦¦¦¦¦¦¦ 1¦20150407110858 en 13>P¦1¦¦ 209000193901¦209000193901¦ TEST¦¦¦¦¦¦¦¦¦¦¦¦¦¦¦¦¦¦¦¦¦¦¦ "L¦1¦ N'4> Current: L¦1¦ N.4> a "13> is missing !! Correct would be: L¦1¦ N-13>-4> Consultation with Ms. Leoanhard IT Synlab: Sent Trace.log discussed and explained. Corresponding erroneous location has been corrected by the LIS driver Now with "13> Test run on AQT90: -New order created in the LIS -Sample inserted on The AQT90 -Scan and Query of the order ID response (master data) were now taken over by the AQT90 and during the analysis in the display. -Findings were made after the measurement autom. sent to the LIS driver. Testing of the data transfer in the LIS system by laboratory personnel => OK KH Trostberg: Change was also carried out on the LIS driver for the AQT90Flex in the KH Trostberg by Ms. Leonhard. Contact with laboratory in Trostberg: -Test measurement and query with feedback of master data on the AQT90Flex successfully
AQT90 FLEX |  |  | CEC33B9302235C5A | No
AQT90 FLEX | Nadel erneuert, Reinigung durchgeführt, |  | SWO-24282 | No | Needle renewed, cleaning carried out,
ABL80 FLEX CO-OX | Lösungspack und Sensorkassette geliefert, wird durch Frau Kersten bei Gelegenheit ersetzt. Weitere Massnahmen dann ggf. telefonisch. | Lösungspack und Sensorkassette geliefert, wird durch Frau Kersten bei Gelegenheit ersetzt. Weitere Massnahmen dann ggf. telefonisch. | SWO-24286 | No | solution pack and sensor cartridge will be replaced by Ms. Kersten on occasion. Further measures can then be taken by telephone if necessary. | solution pack and sensor cartridge will be replaced by Ms. Kersten on occasion. Further measures can then be taken by telephone if necessary.
ABL80 CO-OX | all diag passing but may have a moving clot | Flushed everything and replaced the spectrometer but still failed PCO2 PH Replaced Hemolyzerand ran THB cal - resolved | CEC48FF978339C7D | No | all diag passing but may have a moving clot | Flushed everything and replaced the spectrometer but still failed PCO2 PH Replaced Hemolyzerand ran THB cal - resolved
ABL80 CO-OX | CO's don't compare well between old analyzer Rapidpoint 405. First time running level 1 QC each day, PO2 reads low. Ok next time. | Replaced parts from PM kit. Performed Alignment. Run zeroing SC. Perform tHb calibration. | CEC439697811BC8D | No | CO's don't compare well between old analyzer Rapidpoint 405. First time running level 1 QC each day, PO2 reads low. Ok next time. | Replaced parts from PM kit. Performed Alignment. Run zeroing SC. Perform tHb calibration.
ABL90 FLEX | Roma.ini, Log und db gesichert. Pc Modul und Sensorkassette erneuert. Akku ausgetauscht. | Der Kunde bittet um einen Service. | SWO-24205 | No | Roma.ini; Log and db backed up. Pc module and sensor cartridge renewed. Battery replaced. | The customer asks for a service. | Defect | HW | CPU / Battery
ABL90 FLEX | System manuell gespült Einlassdichtung und Nadel aus Kundenbestand getauscht, SC getauscht, Kaltstartphase, mehrere man.Kal,s durchgeführt,tHb Kal durchgeführt | System manuell gespült Einlassdichtung und Nadel aus Kundenbestand getauscht, SC getauscht, Kaltstartphase, mehrere man.Kal,s durchgeführt,tHb Kal durchgeführt | SWO-24204 | No | System manually rinsed inlet seal and needle swapped from customer stock, SC exchanged, cold start phase, several man. Kal,s performed,tHb Kal performed | System manually rinsed inlet seal and needle swapped from customer stock, SC exchanged, cold start phase, several man. Kal,s performed,tHb Kal performed | Assisted Replacement | SC | Wetability
ABL90 FLEX | Nettoyage du système fluidique e.o Calibration e.o contrôle de qualité e.o | Nettoyage du système fluidique e.o Calibration e.o contrôle de qualité e.o | SWO-24168 | No | Cleaning the fluid system e.o Calibration e.o quality control e.o | Cleaning the fluid system e.o Calibration e.o quality control e.o | Cleaned | Fluidic pathway
ABL90 FLEX | Fehlersuche Kalibrierung QC Messung | Fehlersuche Kalibrierung QC Messung | SWO-24297 | No | Troubleshooting Calibration QC Measurement | Troubleshooting Calibration QC Measurement | Assisted Replacement | SP | FT
ABL90 |  |  | CEC4854B7826A0C3 | No |  |  | Assisted Replacement | SC | Glu
ABL825 FLEX |  |  | CEC45F43780DC2E8 | No |  |  | Defect | HW | OXI | Lamp unit
ABL835 FLEX | multiple errors on calibration #238 248 spectrometer bias and trigger issues 316 258 neon lamp errors suspect issue with power supply and oxi pcb | replaced power supply resolved spectrometer issues replaced oxi pcb and lamp unit to resolve neon lamp errors | CEC471277809FF7D | No | multiple errors on calibration #238 248 spectrometer bias and trigger issues 316 258 neon lamp errors suspect issue with power supply and oxi pcb | replaced power supply resolved spectrometer issues replaced oxi pcb and lamp unit to resolve neon lamp errors | Defect | HW | OXI | PCB
AQT90 FLEX | Comprobación de HTC mediante tool de radiometer. No se detecta incidencia alguna. Sustituida pipeta y sensor por rotura de los mismos. Entregada carta de la FAN a la abajo firmante. |  | CEC45CEE7800D6BF | Yes | Checking HTC using radiometer tool. No incident is detected. Replaced pipette and sensor by breakage. Delivered fan letter to the undersigned. |  | Re classified RMED | FAN / TU
AQT90 FLEX |  |  | CEC6FDCA643F4EA0 | Yes |  |  | Re classified RMED | FAN / TU
AQT90 FLEX |  |  | CEC472377853C970 | No
ABL90 FLEX | Nadel richtig in der Einlassdichtung eingesetzt. Gerät ging in der Start Phase, mehrere Kal wiederholt, QC wiederholt, pO2 complete Test durchgeführt. QC 1 und 2 ok. QC 3 Fehler bei pO2. Kundin wiederholt sich die QC 3 heute Abend noch mal . Kundin meldet sich wieder wenn weitere Unterstützung notwendig sein wird. | Nadel richtig in der Einlassdichtung eingesetzt. Gerät ging in der Start Phase, mehrere Kal wiederholt, QC wiederholt, pO2 complete Test durchgeführt. QC 1 und 2 ok. QC 3 Fehler bei pO2. Kundin wiederholt sich die QC 3 heute Abend noch mal . Kundin meldet sich wieder wenn weitere Unterstützung notwendig sein wird. | SWO-24611 | No | needle correctly inserted in the inlet seal. Device went into the start phase, several kal repeated, QC repeated, pO2 complete test performed. QC 1 and 2 ok. QC 3 error with pO2. Customer repeats the QC 3 tonight. Customer will come back if further support will be needed. | needle correctly inserted in the inlet seal. Device went into the start phase, several kal repeated, QC repeated, pO2 complete test performed. QC 1 and 2 ok. QC 3 error with pO2. Customer repeats the QC 3 tonight. Customer will come back if further support will be needed. | User Error | Inlet gasket
ABL90 | It had 1343 unable to pump solution with user intervetion. | Cleaned inlet gasket and probe. Flushed unit 15 times, replaced solution pack and still having problem. Called Tech support and told me to changed out all tubing from PM kit and to replace the solution pack and sensor cassette. So replaced all tubing, inlet gasket. I also replaced sensor cassette and solution pack both from customers supplies. | CEC5AFC2EE39013A | No | It had 1343 unable to pump solution with user intervetion. | Cleaned inlet gasket and probe. Flushed unit 15 times, replaced solution pack and still having problem. Called Tech support and told me to changed out all tubing from PM kit and to replace the solution pack and sensor cassette. So replaced all tubing, inlet gasket. I also replaced sensor cassette and solution pack both from customers supplies. | Assisted Replacement | SC | Wetability
ABL90 | ANALYZER WAS ON AND HAD A 1063 ERROR MESSAGE. FLUSHED UNIT, ERROR CLEAR. ANALYZER YELLOW BECAUSE SENSOR CASSETTE NEEDS TO BE REPLACED TODAY, ITS 30 DAYS ARE UP. | CLEANED INLET GASKET, BACK FLUSHED 10 TIMES, THEN REPLACED SENSOR CASSETTE FROM CUSTOMERS STOCK. | CEC58E0EEE610137 | No | ANALYZER WAS ON AND HAD A 1063 ERROR MESSAGE. FLUSHED UNIT, ERROR CLEAR. ANALYZER YELLOW BECAUSE SENSOR CASSETTE NEEDS TO BE REPLACED TODAY, ITS 30 DAYS ARE UP. | CLEANED INLET GASKET, BACK FLUSHED 10 TIMES, THEN REPLACED SENSOR CASSETTE FROM CUSTOMERS STOCK. | Cleaned | Fluidic pathway
ABL90 FLEX | Fehlersuche Fließwege durchgespült Einlaßdichtung erneuert Kontrolle Pumpenrotor Test im Service Programm. wird bei einem LP keine Lösung erkannt nach einsetzen eines anderen LP dieser o.k. Einlaufverfahren | Fehlersuche Fließwege durchgespült Einlaßdichtung erneuert Kontrolle Pumpenrotor Test im Service Programm. wird bei einem LP keine Lösung erkannt nach einsetzen eines anderen LP dieser o.k. Einlaufverfahren | SWO-24425 | No | Troubleshooting flow paths flushed through inlet seal renewed control pump rotor test in the service program. no solution is detected in an LP after inserting another LP of this o.k. encrocity method | Troubleshooting flow paths flushed through inlet seal renewed control pump rotor test in the service program. no solution is detected in an LP after inserting another LP of this o.k. encrocity method | Assisted Replacement | SP | FT
ABL90 |  |  | CEC5A066EE04B6B0 | No |  |  | Assisted Replacement | SC | Reference
ABL90 | Entrenar a Mariano Gómez para que en el programa de servicio pueda pasar una muestra con sangre para solventar la incidencia. Calibrar y probar con controles. | Pasar muestras repetidas con sangre. | CEC5AC45EE202147 | No | Train Mariano Gómez so that in the service program she can pass a blood sample to solve the problem. Calibrate and test with controls. | Pass repeated samples with blood. | Assisted Replacement | SC | Wetability
ABL825 FLEX | SE CAMBIA UNIDAD DE LAMPARA, PLACA OXI Y TUBO VALVULA, SE HACE CALIBRACION 2 PUNTOS CAL CTHB Y SE COMPRUEBA FUNCIONAMIENTO CON QC | SE CAMBIA UNIDAD DE LAMPARA, PLACA OXI Y TUBO VALVULA, SE HACE CALIBRACION 2 PUNTOS CAL CTHB Y SE COMPRUEBA FUNCIONAMIENTO CON QC | CEC58078EE60E46C | No | CHANGE LAMPARA UNIT, OXI PLACE AND VALVE TUBE, CALIBRATION 2 CAL CTHB POINTS IS CHECKED AND OPERATION IS CHECKED WITH QC | CHANGE LAMPARA UNIT, OXI PLACE AND VALVE TUBE, CALIBRATION 2 CAL CTHB POINTS IS CHECKED AND OPERATION IS CHECKED WITH QC | Defect | HW | Tube Valve
ABL837 FLEX | Inspection of the fluidic channel revealed stress cracks near V8 opening. | Replaced the fluidic channel. | CEC5D619EE06D1BD | No | Inspection of the fluidic channel revealed stress cracks near V8 opening. | Replaced the fluidic channel. | Defect | HW | Fluidic manifold
ABL820 FLEX | Powered unit off. Vibration still present. Suspect change in Air venting. | Customer is calling in Hospital Maintenance to eval vibration. | CEC5C02AEE33253A | Yes | Powered unit off. Vibration still present. Suspect change in Air venting. | Customer is calling in Hospital Maintenance to eval vibration. | Re classified RMED | Hospital vibrating
ABL827 FLEX | found in data logs that lot number was not transmitting | rebooted analyzer and is now transmitting. Opening a customer quality case. Performed a service dump and loaded analyzer information. | CEC59913EE08FA18 | No | found in data logs that lot number was not transmitting | rebooted analyzer and is now transmitting. Opening a customer quality case. Performed a service dump and loaded analyzer information. | Defect | SW | His/lis
ABL820 FLEX | Checked the ampoule actuator in the autocheck service program and found in would mot move. Removed the autocheck module and found the ampoule actuator motor shaft had disconnected from the ampoule actuator. | Reconnected the ampoule actuator motor shaft to the ampoule actuator. | CEC58C52EE35F197 | No | Checked the ampoule actuator in the autocheck service program and found in would mot move. Removed the autocheck module and found the ampoule actuator motor shaft had disconnected from the ampoule actuator. | Reconnected the ampoule actuator motor shaft to the ampoule actuator. | Defect | HW | AutoCheck Module | Actuator shaft
ABL825 FLEX |  |  | CEC5DC14EE54CB84 | No |  |  | Defect | HW | AutoCheck Module | Actuator shaft
ABL835 FLEX | Jeg har logget på ABL835 og fjernet markering for at ABL skal kommunikere med Radiance. Ventet 30 sek. og sat markering igen. Nu virker alt igen. | Jeg har logget på ABL835 og fjernet markering for at ABL skal kommunikere med Radiance. Ventet 30 sek. og sat markering igen. Nu virker alt igen. | SWO-24737 | No | I have logged on ABL835 and removed selection for ABL to communicate with Radiance. Waited 30 sec. and set selection again. Now everything works again. | I have logged on ABL835 and removed selection for ABL to communicate with Radiance. Waited 30 sec. and set selection again. Now everything works again. | Setup | HIS/LIS | His/lis
ABL90 FLEX | Gerät überprüft, manuell gespült, waste connector gereinigt, Kaltstartphase und manuell kal,s und QC durchgeführt da pO2 nach oben raus ohne erfolg LP getauscht Kal und QC I.O System messbereit | Gerät überprüft, manuell gespült, waste connector gereinigt, Kaltstartphase und manuell kal,s und QC durchgeführt da pO2 nach oben raus ohne erfolg LP getauscht Kal und QC I.O System messbereit | SWO-24794 | No | Device checked, manually rinsed, waste connector cleaned, cold start phase and manually kal,s and QC performed as pO2 up out without success LP swapped Kal and QC I.O system ready to measure | Device checked, manually rinsed, waste connector cleaned, cold start phase and manually kal,s and QC performed as pO2 up out without success LP swapped Kal and QC I.O system ready to measure | Cleaned | Fluidic pathway
ABL90 FLEX | Sämtliche Flüssigkeitswege überprüft. Verbindungsschlauch zwischen SC und Hemolyzer ersetzt (undicht). Neue SC installiert und SW upgrade auf 3.1MR3 durchgeführt. System, Kal. und QC i.o. Funktionskontrolle i.o. | Sämtliche Flüssigkeitswege überprüft. Verbindungsschlauch zwischen SC und Hemolyzer ersetzt (undicht). Neue SC installiert und SW upgrade auf 3.1MR3 durchgeführt. System, Kal. und QC i.o. Funktionskontrolle i.o. | SWO-24770 | No | Check all fluid pathways. Replaced (leaked) connection hose between SC and Hemolyzer. New SC installed and SW upgrade performed to 3.1MR3. System, Kal. and QC i.o. functional control i.o. | Check all fluid pathways. Replaced (leaked) connection hose between SC and Hemolyzer. New SC installed and SW upgrade performed to 3.1MR3. System, Kal. and QC i.o. functional control i.o. | Defect | HW | Oxi module | Pinch Valve tube
ABL90 FLEX | Zeiteinstellung nach dem SW Update überprüft. Automatische Zeiteinstellung heraus genommen. Neustart. | Den Kunden bei der Konfiguration der Uhrzeit telefonisch geführt. | SWO-24803 | Yes | Time setting checked after the SW update. Automatic time setting taken out. Restart. | Guided the customer by phone when configuring the time. | Re classified RMED | Phone support
ABL90 FLEX |  |  | SWO-24677 | No |  |  | Defect | HW | Battery module
ABL90 FLEX | Analysator teilweise dem. und Steckkontakte überprüft etc.- Ersatzteilebestellung erfolgt | Analysator teilweise dem. und Steckkontakte überprüft etc.- Ersatzteilebestellung erfolgt | SWO-24760 | No | Analyzer partially the. and plug-in contacts checked etc.- spare parts order is made | Analyzer partially the. and plug-in contacts checked etc.- spare parts order is made | Unknown | Keyboard??
ABL90 |  |  | CEC6E06A640C0F69 | No |  |  | Defect | HW | CPU
ABL90 FLEX | Datensicherung, Verdacht auf fehlerhafte LP. Sind aber derzeit nur Lot RT07 eingesetzt und vorhanden. Es sollte eine andere Lot.-Nummer getestet werden. | Datensicherung, Verdacht auf fehlerhafte LP. Sind aber derzeit nur Lot RT07 eingesetzt und vorhanden. Es sollte eine andere Lot.-Nummer getestet werden. | SWO-24787 | Yes | Data backup, suspicion of faulty LP. However, only Lot RT07 is currently used and available. A different lot number should be tested. | Data backup, suspicion of faulty LP. However, only Lot RT07 is currently used and available. A different lot number should be tested. | Re classified RMED | Hotline call w actually no visit
ABL827 FLEX |  |  | CECA6A1BC6071CBE | No |  |  | Defect | HW | unknown
ABL835 FLEX | THIS HAS BEEN AN ONGOING PROBLEM WITH THIS ANALYZER. THIS IS TO DOCUMENT THE WARRANTY ELECTRODE AND THE BOX OF LACTATE MEMBRANES WERE AUTHORIZED BY CHUCK RALL. | REPLACED ELECTRODE AND MEMBRANE AND SHIPPED OLD ELECTRODE BACK. | CEC9A1CC504E8ABF | No | THIS HAS BEEN AN ONGOING PROBLEM WITH THIS ANALYZER. THIS IS TO DOCUMENT THE WARRANTY ELECTRODE AND THE BOX OF LACTATE MEMBRANES WERE AUTHORIZED BY CHUCK RALL. | REPLACED ELECTRODE AND MEMBRANE AND SHIPPED OLD ELECTRODE BACK. | Assisted Replacement | Electrode | Electrode | Lac
ABL825 FLEX | Modul kalibriert etc. - Erkennung Sampler Tray 1 defekt - Modul sollte ausgetauscht werden ! | Modul kalibriert etc. - Erkennung Sampler Tray 1 defekt - Modul sollte ausgetauscht werden ! | SWO-24759 | No | Module calibrated etc. - detection Sampler Tray 1 defective - module should be replaced ! | Module calibrated etc. - detection Sampler Tray 1 defective - module should be replaced ! | Defect | HW | FlexQ Module | Tray1
ABL825 FLEX | Gerät überprüft, Ref.Elektrode getauscht, und defekte eingeklemmte PH Dichtung getauscht, mehrere Kal,s durchgeführt und QC vermessen | Gerät überprüft, Ref.Elektrode getauscht, und defekte eingeklemmte PH Dichtung getauscht, mehrere Kal,s durchgeführt und QC vermessen | SWO-24710 | No | Device checked, ref.electrode exchanged, and defective clamped PH seal exchanged, multiple kal,s performed and QC measured | Device checked, ref.electrode exchanged, and defective clamped PH seal exchanged, multiple kal,s performed and QC measured | Assisted Replacement | Electrode | Membrane | Reference
AQT90 FLEX | SE HACE COMPROBACION DEL EQUIPO Y SE LE NOTIFICA AL CLIENTE QUE LA PRUEBA A SIDO SATISFACTORIA. SE DETECTA SP DEFECTUOSO LOTE RP-16 3 UNIDADES, SE CAMBIA SP Y SE LES VERIFICA EL RESTO DE SP | SE HACE COMPROBACION DEL EQUIPO Y SE LE NOTIFICA AL CLIENTE QUE LA PRUEBA A SIDO SATISFACTORIA. SE DETECTA SP DEFECTUOSO LOTE RP-16 3 UNIDADES, SE CAMBIA SP Y SE LES VERIFICA EL RESTO DE SP | CEC6B76E6423B87B | Yes | EQUIPMENT IS CHECKED AND THE CUSTOMER IS NOTIFIED THAT THE TEST HAS BEEN SATISFACTORY. SP DEFECT LOT RP-16 3 UNITS IS DETECTED, SP IS CHANGED AND THE REST OF SP IS CHECKED | EQUIPMENT IS CHECKED AND THE CUSTOMER IS NOTIFIED THAT THE TEST HAS BEEN SATISFACTORY. SP DEFECT LOT RP-16 3 UNITS IS DETECTED, SP IS CHANGED AND THE REST OF SP IS CHECKED | Re classified RMED | FAN / TU
AQT90 FLEX |  |  | CEC9CEA5505AC51A | No
ABL90 |  |  | CED292F58A359EEC | Yes |  |  | Re classified RMED | SWU
ABL90 FLEX | Lösungspack ersetzt, Kal, QC, gleicher Fehler, Wechsel der SC, Kal, QC | Lösungspack ersetzt, Kal, QC, gleicher Fehler, Wechsel der SC, Kal, QC | SWO-24990 | No | Solution pack replaced, Kal, QC, same error, change of SC, Kal, QC | Solution pack replaced, Kal, QC, same error, change of SC, Kal, QC | Assisted Replacement | SC | pCO2
ABL90 | LOOKED THROUGH DATA LOGS AND FOUND THAT STEPHEN WESTON LOGGED ON TO FIX THE INHOMOGENEOUS RINSE SOLUTION ERROR AND THEN JUST GOT A DIFFERENT ANALYZER. | DID 10 BACKFLUSHES, DID ALL TESTS IN FLUID TRANSPORT IN SERVICE. REPLACED SOLUTION PACK, RAN CALIBRATIONS 5 TIMES BEFORE IT CAME IN. | CECAB905C62D98BE | No | LOOKED THROUGH DATA LOGS AND FOUND THAT STEPHEN WESTON LOGGED ON TO FIX THE INHOMOGENEOUS RINSE SOLUTION ERROR AND THEN JUST GOT A DIFFERENT ANALYZER. | DID 10 BACKFLUSHES, DID ALL TESTS IN FLUID TRANSPORT IN SERVICE. REPLACED SOLUTION PACK, RAN CALIBRATIONS 5 TIMES BEFORE IT CAME IN. | Cleaned | Fluidic pathway
ABL90 | DID EXTERNAL QC AND IT WAS HIGH ON LV 1 AND LOW ON LV 3. | RECALIBRATED IT AND THEN DID EXTERNAL QC AGAIN. | CECAAA1AC62F98BB | No | DID EXTERNAL QC AND IT WAS HIGH ON LV 1 AND LOW ON LV 3. | RECALIBRATED IT AND THEN DID EXTERNAL QC AGAIN. | Adjustment | Calibration | Lac
ABL90 |  |  | CED2999F8A2F5D68 | Yes |  |  | Re classified RMED | SWU
ABL90 FLEX | Hemolyser erneuert | Hemolyser erneuert | SWO-24887 | No | Hemolyser renewed | Hemolyser renewed | Defect | HW | Hemolyzer
ABL90 |  |  | CECD12B4B232B6C6 | Yes |  |  | Re classified RMED | Telephone support
ABL90 | Comprobar SP y SC, cambiar SC, calibrar y probar con controles. | Cambiar SC. | CECB13B8C62A0689 | No | Check SP and SC, change SC, calibrate and test with controls. | Change SC. | Assisted Replacement | SC | Wetability
ABL837 FLEX |  |  | CECE8524280BA4B4 | No |  |  | Defect | HW | Inlet module | Probe cal.
ABL837 FLEX |  |  | CECEBA21285CB059 | No |  |  | Defect | HW | Inlet module | LS
ABL837 FLEX |  |  | CECDA992B240CC8C | Yes |  |  | Re classified RMED | OOB
AQT90 FLEX |  |  | CECAC15BC61C2DAB | Yes |  |  | Re classified RMED | FAN / TU
TCM/Tosca | TECHLINE SUGGESTS REPLACEMENT OF MODULE | REPLACED MODULE AND RECALIBRATED SYSTEM. | CECA90B7C61B646E | No | TECHLINE SUGGESTS REPLACEMENT OF MODULE | REPLACED MODULE AND RECALIBRATED SYSTEM.
ABL90 FLEX | Lösungspack (RT-08) geliefert und eingesetzt | Lösungspack (RT-08) geliefert und eingesetzt | SWO-25116 | Yes | Solution pack (RT-08) supplied and used | Solution pack (RT-08) supplied and used | Re classified RMED | Logistical upSWO
ABL90 FLEX | Fehlersuche Demontage Einlaß, verkantete Feder gelöst, Test i.O. | Fehlersuche Demontage Einlaß, verkantete Feder gelöst, Test i.O. | SWO-25160 | No | Troubleshooting Disassembly inlet, edged spring solved, test i.O. | Troubleshooting Disassembly inlet, edged spring solved, test i.O. | Defect | HW | Inlet module
ABL90 FLEX | Accu vervangen voor een nieuwe accu. Getest of accu correct werkt via servicemenu --> OK. Getest of accu correct werkt na verwijderen stekker, en weer aansluiten stekker --> OK. | Accu vervangen voor een nieuwe accu. Getest of accu correct werkt via servicemenu --> OK. Getest of accu correct werkt na verwijderen stekker, en weer aansluiten stekker --> OK. | SWO-25195 | No | Replace the battery for a new battery. Tested if battery works correctly via service menu--> OK. Tested if battery works correctly after removing plug, and reconnect plug--> OK. | Replace the battery for a new battery. Tested if battery works correctly via service menu--> OK. Tested if battery works correctly after removing plug, and reconnect plug--> OK. | Defect | HW | Battery module
ABL90 FLEX | Accu vervangen voor een nieuwe accu. Getest of accu correct werkt via servicemenu --> OK. Getest of accu correct werkt na verwijderen stekker, en weer aansluiten stekker --> OK. | Accu vervangen voor een nieuwe accu. Getest of accu correct werkt via servicemenu --> OK. Getest of accu correct werkt na verwijderen stekker, en weer aansluiten stekker --> OK. | SWO-25192 | No | Replace the battery for a new battery. Tested if battery works correctly via service menu--> OK. Tested if battery works correctly after removing plug, and reconnect plug--> OK. | Replace the battery for a new battery. Tested if battery works correctly via service menu--> OK. Tested if battery works correctly after removing plug, and reconnect plug--> OK. | Defect | HW | Battery module
ABL90 FLEX | Accu vervangen voor een nieuwe accu. Getest of accu correct werkt via servicemenu --> OK. Getest of accu correct werkt na verwijderen stekker, en weer aansluiten stekker --> OK. | Accu vervangen voor een nieuwe accu. Getest of accu correct werkt via servicemenu --> OK. Getest of accu correct werkt na verwijderen stekker, en weer aansluiten stekker --> OK. | SWO-25193 | No | Replace the battery for a new battery. Tested if battery works correctly via service menu--> OK. Tested if battery works correctly after removing plug, and reconnect plug--> OK. | Replace the battery for a new battery. Tested if battery works correctly via service menu--> OK. Tested if battery works correctly after removing plug, and reconnect plug--> OK. | Defect | HW | Battery module
ABL90 FLEX | Accu vervangen voor een nieuwe accu. Getest of accu correct werkt via servicemenu --> OK. Getest of accu correct werkt na verwijderen stekker, en weer aansluiten stekker --> OK. | Accu vervangen voor een nieuwe accu. Getest of accu correct werkt via servicemenu --> OK. Getest of accu correct werkt na verwijderen stekker, en weer aansluiten stekker --> OK. | SWO-25194 | No | Replace the battery for a new battery. Tested if battery works correctly via service menu--> OK. Tested if battery works correctly after removing plug, and reconnect plug--> OK. | Replace the battery for a new battery. Tested if battery works correctly via service menu--> OK. Tested if battery works correctly after removing plug, and reconnect plug--> OK. | Defect | HW | Battery module
ABL90 FLEX | div. Test`s im Service durchgef., Einlaßdichtung gew., div. Kal. ohne Erfolg, SC reinstalliert, LP gew., Kal. + QC i.O. | div. Test`s im Service durchgef.; Einlaßdichtung gew.; div. Kal. ohne Erfolg; SC reinstalliert; LP gew.; Kal. + QC i.O. | SWO-25084 | No | div. Test it in the service. inlet seal. div. Kal. to no avail; SC reinstalled; LP. Kal. + QC in the first place | div. Test it in the service. inlet seal. div. Kal. to no avail; SC reinstalled; LP. Kal. + QC in the first place | Assisted Replacement | SP | pO2
ABL800 BASIC | Kira startet nicht, Kira deinstalliert, Kira neu aufgespielt, Gerät startet wieder, Konfiguration wieder eingestallt, Einlassnadel verbogen - Einlassnadel ersetzt, QC Temperatur bei 55Grad keine Regelung, Auto QC PCB besorgt und eingebaut - Fehler bleibt, QC Modul ausgebaut QC händisch durchgeführt i.O Station informiert und eingewiesen ETeil bestellt | Auto QC Modul als Ersatz eingesetzt bis Ersatzteil da, Kal und QC durchgeführt | SWO-25206 | No | Kira does not start, Kira uninstalled, Kira replayed, device restarts, configuration re-installed, inlet needle bent - inlet needle replaced, QC temperature at 55 degrees no control, Auto QC PCB concerned and built-in - error remains, QC module removed QC manually carried out i.O station informed and instructed EPart ordered | Auto QC module used as replacement until spare part there, Kal and QC performed | Defect | HW | AutoCheck Module
ABL837 FLEX | Cleaned fuses, replaced power supply,did not help. Ordered AU PCB overnight, will return with part in AM |  | CECC1C0F3C56B597 | No | Cleaned fuses, replaced power supply,did not help. Ordered AU PCB overnight, will return with part in AM |  | Defect | HW | Analyzer Control PCB
ABL825 FLEX |  |  | CECC0DB13C20076F | No |  |  | Defect | HW | Oxi module | Spectrometer
ABL825 FLEX | After powering off customer could not get instrument back on. | Reloaded software. Did not resolve. Replaced AU board and power supply. Did not resolve. Replaced CPU and could not get past replacement screen. Tried original AU board and still DMS not connected. | CECBE3893C257E35 | No | After powering off customer could not get instrument back on. | Reloaded software. Did not resolve. Replaced AU board and power supply. Did not resolve. Replaced CPU and could not get past replacement screen. Tried original AU board and still DMS not connected. | Unknown | HW | Analyzer Control PCB
ABL825 FLEX | Messkammer und Schlauchsystem geprüft, Lac. Elektrode (E7077 RN05) ausgetauscht, Funktionskontrolle ok | Messkammer und Schlauchsystem geprüft, Lac. Elektrode (E7077 RN05) ausgetauscht, Funktionskontrolle ok | SWO-25331 | No | Measuring chamber and hose system tested, Lac. Electrode (E7077 RN05) replaced, function control ok | Measuring chamber and hose system tested, Lac. Electrode (E7077 RN05) replaced, function control ok | Assisted Replacement | Electrode | Electrode | Lac
AQT90 FLEX | Comprobación de la célula de hematocrito correcta. Se entrega carta al abajo firmante. |  | CECBD85B3C01F3B7 | Yes | Checking the correct hematocrit cell. Letter is delivered to the undersigned. |  | Re classified RMED | FAN / TU
TCM/Tosca | TRIED MODULE WITH NO HELP. | REPLACED ELECTRODE UNDER WARRANTY. | CECF91D59E583A33 | No | TRIED MODULE WITH NO HELP. | REPLACED ELECTRODE UNDER WARRANTY.
ABL80 | Troubleshoot the Power supply, the Lower LED was flahing RED and when the system was unplugged to operate on battery it turned off. Perform system checkout, checking the electronics in service using system test cassette. | Perform Set 0 routine in service with test cassette, and verify operation of circuit. Replace the Power supply for charging issue and test analyzer after replacement. | CECD4DC9B222778B | No | Troubleshoot the Power supply, the Lower LED was flahing RED and when the system was unplugged to operate on battery it turned off. Perform system checkout, checking the electronics in service using system test cassette. | Perform Set 0 routine in service with test cassette, and verify operation of circuit. Replace the Power supply for charging issue and test analyzer after replacement.
ABL80 OSM | Possible fluid blockage error while running Quality Control. I had replaced the hemolyzer before for an unrelated problem. | I calibrated the LS and the value is 61 (OK). I replaced the Spectrometer Oxi Pcb. | CECD1EA3B209A419 | No | Possible fluid blockage error while running Quality Control. I had replaced the hemolyzer before for an unrelated problem. | I calibrated the LS and the value is 61 (OK). I replaced the Spectrometer Oxi Pcb.
ABL80 CO-OX |  |  | CECDAA1BB256CC8F | No
ABL90 | tried running cap mode no video and not able to, mixer not working, worked with IT for and hour trying to load the Static IP address and they were not sure of the DNS server number , Loaded 10.0.8.101 ---- 255.255.255.0 ----- 10.0.8.254, this will require a follow up visit after April 27 when the IT returns from overseas trip | replaced the inlet position dectector PCB; replaced the sample mixer Unit; Added the IP address with Dan from IT; FSE Jerry Kelly assisted; | CECD123EB2422509 | No | tried running cap mode no video and not able to; mixer not working; worked with IT for and hour trying to load the Static IP address and they were not sure of the DNS server number ; Loaded 10.0.8.101 ---- 255.255.255.0 ----- 10.0.8.254; this will require a follow up visit after April 27 when the IT returns from overseas trip | replaced the inlet position dectector PCB; replaced the sample mixer Unit; Added the IP address with Dan from IT; FSE Jerry Kelly assisted; | Defect | HW | Sample Mixer
ABL820 FLEX | szelep ellenőrzése és tisztítása, kalibráció, QC mérés |  | SWO-25278 | No | Valve inspection and cleaning, calibration, QC measurement |  | Defect | HW | pH/BG Module | Valve
ABL800 BASIC | Schläuche, pumpenschläuche und Airventil ersetzt, Messkammern gereinigt, Kalibration, QC Kontrolle | Schläuche, pumpenschläuche und Airventil ersetzt, Messkammern gereinigt, Kalibration, QC Kontrolle | SWO-25360 | No | Hoses, pump hoses and air valve replaced, measuring chambers cleaned, calibration, QC control | Hoses, pump hoses and air valve replaced, measuring chambers cleaned, calibration, QC control | Defect | HW | El/Met Module | Air Valve
ABL827 FLEX |  |  | CECD1CB2B2219AE2 | No |  |  | Unknown | Replaced analyzer | unknown
ABL800 BASIC | Auto QC Modul als Ersatz eingesetzt bis Ersatzteil da, Kal und QC durchgeführt | Auto QC Modul als Ersatz eingesetzt bis Ersatzteil da, Kal und QC durchgeführt | SWO-25465 | No | Auto QC module used as replacement until spare part there, Kal and QC performed | Auto QC module used as replacement until spare part there, Kal and QC performed | Revisit required | HW | AutoCheck Module
ABL837 FLEX | Tried new AU PCB and now connects | Reloaded AU softwware and option keys. Glucose has on QC unstable .Remembraned Glucose electrode. Multiple Cals required. | CECD341FB227B877 | No | Tried new AU PCB and now connects | Reloaded AU softwware and option keys. Glucose has on QC unstable . Remembraned Glucose electrode. Multiple Cals required. | Revisit required | Replace PCB | Replace PCB
ABL837 FLEX |  |  | CECD73A8B2279AEE | Yes |  |  | Re classified RMED | Yearly maintenance
ABL835 FLEX | Massive Verstopfung im Ablaßbereich. | Massive Verstopfung im Ablaßbereich. | SWO-25461 | No | Massive constipation in the drain area. | Massive constipation in the drain area. | Cleaned | Fluidic pathway | Waste tube
ABL825 FLEX | CUSTOMER HAD JUST CHANGED CA+ AND REFERENCE MEMBRANES 3 TIMES.DATA LOGS SHOW CA+ SENSITIVITY TO BE RUNNING BETWEEN 105.1 AND 111% | I REPLACED CALIBRATION 1 AND 2 SOLUTIONS. IORDERED NEW LOT NUMBERS OF CA+ AND REFERENCE MEMBRANES AND ELECTRODES.REPLACED CA+AND REFERENCE ELECTRODES AND MEMBRANES.RAN MULTIPLE 1 AND 2 PONT CALIBRATIONS SENSITIVITY WAS AT 104.5% | CECD1667B23E1061 | No | CUSTOMER HAD JUST CHANGED CA+ AND REFERENCE MEMBRANES 3 TIMES. DATA LOGS SHOW CA+ SENSITIVITY TO BE RUNNING BETWEEN 105.1 AND 111% | I REPLACED CALIBRATION 1 AND 2 SOLUTIONS. IORDERED NEW LOT NUMBERS OF CA+ AND REFERENCE MEMBRANES AND ELECTRODES. REPLACED CA+AND REFERENCE ELECTRODES AND MEMBRANES. RAN MULTIPLE 1 AND 2 PONT CALIBRATIONS SENSITIVITY WAS AT 104.5% | Assisted Replacement | Electrode | Membrane | Ca | Maybe replace El/met PCB
ABL815 FLEX | Waste- und Entlüftungswege gespült. | Waste- und Entlüftungswege gespült. | SWO-25425 | No | Waste and vent paths. | Waste and vent paths. | Cleaned | Fluidic pathway | Fluid manifold | Air intake
ABL835 FLEX | Mehrfach versucht den Fehler durch Modulkalibration zu beseitigen, ohne Erfolg Ersatzmodul eingebaut. 902-753 R1072N002 | Mehrfach versucht den Fehler durch Modulkalibration zu beseitigen, ohne Erfolg Ersatzmodul eingebaut. 902-753 R1072N002 | SWO-25411 | No | Multiple attempts to eliminate the error by module calibration, without success replacement module built in. 902-753 R1072N002 | Multiple attempts to eliminate the error by module calibration, without success replacement module built in. 902-753 R1072N002 | Defect | HW | FlexQ Module | Sampler not recognized
ABL825 FLEX | FOUND PROBE DRIVE STEPPER MOTOR CABLE PARTIALLY CONNECTED. NO OTHER DAMAGE FOUND. | RECONNECTED CABLE TO PCB AND TESTED. | CECEBB6B283B92B0 | No | FOUND PROBE DRIVE STEPPER MOTOR CABLE PARTIALLY CONNECTED. NO OTHER DAMAGE FOUND. | RECONNECTED CABLE TO PCB AND TESTED. | Other | Adjustment | Inlet module | Cable connection
ABL820 FLEX | Unit failing all three areas when performing complete oxi test | Reseated cabling on oxi module, performed lamp adjustment | CECD8264B24CE865 | Yes | Unit failing all three areas when performing complete oxi test | Reseated cabling on oxi module, performed lamp adjustment | Re classified RMED | OOB
ABL820 FLEX | Unit failing a three checks during oxi test. | Checked an reseated cabling, performed lamp adjustment | CECD85C5B215E86C | No | Unit failing a three checks during oxi test. | Checked an reseated cabling, performed lamp adjustment | Adjustment | HW | Oxi module | Lamp unit
AQT90 FLEX | In het service menu de inletdoors getest, deze werken naar behoren. Toch de inletdoors met de aansturings motoren vervangen. Tevens het Cabinet PCB ( artnr 903-117) vervangen. Ook nu weer de werking getest in het servicemenu --> OK. Testen met meten, en AQT loopt weer vast bij verwijderen buisje. Derhalve de originele onderdelen weer terug geplaatst. Na even gewacht te hebben logt de AQT de gebruiker uit, inloggen is weer mogelijk. Echter de balk bovenin is niet te bedienen, en enkele knoppen in het menu ( plaats buis, en verwijder buis) zijn niet te bedienen. Back-up gemaakt, en hierna een clean install uitgevoerd van de software versie 8.7.6. Database en overige instellingen klant teruggezet. Testen met meten, en AQT loopt weer vast bij verwijderen van buisje. Overlegd met Kopenhagen, zij vragen voor de servicenummer, en zullen dan advies geven. Wordt verder opgevolgd | In het service menu de inletdoors getest, deze werken naar behoren. Toch de inletdoors met de aansturings motoren vervangen. Tevens het Cabinet PCB ( artnr 903-117) vervangen. Ook nu weer de werking getest in het servicemenu --> OK. Testen met meten, en AQT loopt weer vast bij verwijderen buisje. Derhalve de originele onderdelen weer terug geplaatst. Na even gewacht te hebben logt de AQT de gebruiker uit, inloggen is weer mogelijk. Echter de balk bovenin is niet te bedienen, en enkele knoppen in het menu ( plaats buis, en verwijder buis) zijn niet te bedienen. Back-up gemaakt, en hierna een clean install uitgevoerd van de software versie 8.7.6. Database en overige instellingen klant teruggezet. Testen met meten, en AQT loopt weer vast bij verwijderen van buisje. Overlegd met Kopenhagen, zij vragen voor de servicenummer, en zullen dan advies geven. Wordt verder opgevolgd | SWO-25383 | No | In the Service menu The inletdoors tested, these work properly. Replace the inletdoors with the steering motors. Also replace the Cabinet PCB (Artnr 903-117). Also now again the operation tested in the Service menu--> OK. Test with measurement, and AQT crashes again when removing tube. Therefore the original parts are repositioned. After waiting for a while, the AQT logs the user out, login is possible again. However the bar at the top is not operable, and some buttons in the menu (place tube, and remove tube) are not operable. Backup, followed by a clean install of the software version 8.7.6. Database and other customer settings. Testing with measurement, and AQT is fixed again when removing the tube. Consulted with Copenhagen, they ask for the service number, and will then give advice. Is further monitored | In the Service menu The inletdoors tested, these work properly. Replace the inletdoors with the steering motors. Also replace the Cabinet PCB (Artnr 903-117). Also now again the operation tested in the Service menu--> OK. Test with measurement, and AQT crashes again when removing tube. Therefore the original parts are repositioned. After waiting for a while, the AQT logs the user out, login is possible again. However the bar at the top is not operable, and some buttons in the menu (place tube, and remove tube) are not operable. Backup, followed by a clean install of the software version 8.7.6. Database and other customer settings. Testing with measurement, and AQT is fixed again when removing the tube. Consulted with Copenhagen, they ask for the service number, and will then give advice. Is further monitored
ABL90 FLEX | QC Level A und B bei Parameter pO2 fehlerhaft --> SP Lot RT-07 inst. --> alle QC Level gemessen --> i.O | QC Level A und B bei Parameter pO2 fehlerhaft --> SP Lot RT-07 inst. --> alle QC Level gemessen --> i.O | SWO-25565 | No | QC Level A and B faulty for parameter pO2 --> SP Lot RT-07 inst. --> all QC levels measured --> i.O | QC Level A and B faulty for parameter pO2 --> SP Lot RT-07 inst. --> all QC levels measured --> i.O | Assisted Replacement | SP | pO2
ABL90 | In the syringe mode setup the analyzer defaults to the upper right selection box. I this case it happened to by Ampoule QC , but changing that box to a different mode, and putting Ampoule QC in a different box, this changed the default selection. Printed analyzer setup, and saved the setup and database to a thumb drive. Then I restored the analyzer defaults. This resolved the problem, so I reloaded the customer settings, and the problem returned. | Returned the analyzer to the default settings, and loaded the customer settings taken from one of their other analyzers, then changed the differences by hand. | CECE8F492846D6DF | No | In the syringe mode setup the analyzer defaults to the upper right selection box. I this case it happened to by Ampoule QC , but changing that box to a different mode, and putting Ampoule QC in a different box, this changed the default selection. Printed analyzer setup, and saved the setup and database to a thumb drive. Then I restored the analyzer defaults. This resolved the problem, so I reloaded the customer settings, and the problem returned. | Returned the analyzer to the default settings, and loaded the customer settings taken from one of their other analyzers, then changed the differences by hand. | Defect | SW | Setup file??
ABL90 FLEX | Zichtplaatsing in UZ LEUVEN (MIG) ter vervanging van ABL90 (R0305N0010) die ter herstelling meegenomen werd gisteren. Ip adres ingegeven en gekoppeld met radiance. Configuratie ingelezen en heropgestart. Calibratie en controle : OK | Zichtplaatsing in UZ LEUVEN (MIG) ter vervanging van ABL90 (R0305N0010) die ter herstelling meegenomen werd gisteren. Ip adres ingegeven en gekoppeld met radiance. Configuratie ingelezen en heropgestart. Calibratie en controle : OK | SWO-25962 | Yes | View placement in UZ LEUVEN (MIG) to replace ABL90 (R0305N0010) which was taken for repair yesterday. Ip address entered and linked with radiance. Configuration is read and restarted. Calibration and Control: OK | View placement in UZ LEUVEN (MIG) to replace ABL90 (R0305N0010) which was taken for repair yesterday. Ip address entered and linked with radiance. Configuration is read and restarted. Calibration and Control: OK | Re classified RMED | Re install analyzer
ABL835 FLEX | Datensicherung etc. - Installation komplette Software - Einstellungen und Funktionsüberprüfung etc. | Datensicherung etc. - Installation komplette Software - Einstellungen und Funktionsüberprüfung etc. | SWO-25607 | No | Data backup etc. - Installation of complete software - settings and function verification etc. | Data backup etc. - Installation of complete software - settings and function verification etc. | Defect | HW | Display
ABL835 FLEX | pO2 Elektrode getauscht | pO2 Elektrode getauscht | SWO-25618 | No | pO2 electrode exchanged | pO2 electrode exchanged | Assisted Replacement | Electrode | Electrode | pO2
AQT90 FLEX |  |  | CECE87F72820D1AA | Yes |  |  | Re classified RMED | FAN / TU
AQT90 FLEX |  |  | CECEA106282E8D54 | No
ABL80 CO-OX | Checked over machine for obvious problems. Found demo keys expired. Printer not functioning at all. | Replaced printer and control board. Installed permanent radiance keys. | CECFE0729E54B300 | No | Checked over machine for obvious problems. Found demo keys expired. Printer not functioning at all. | Replaced printer and control board. Installed permanent radiance keys.
ABL90 FLEX | Sensor interface PCB ersetzt, Neustart, Kalibration, QC Kontrolle | Sensor interface PCB ersetzt, Neustart, Kalibration, QC Kontrolle | SWO-25851 | No | Sensor interface PCB replaced, restart, calibration, QC control | Sensor interface PCB replaced, restart, calibration, QC control | Defect | HW | SC ID PCB
ABL90 FLEX | ABL neu gestartet und Einstellungen überprüft. | ABL neu gestartet und Einstellungen überprüft. | SWO-25726 | No | ABL restarted and settings checked. | ABL restarted and settings checked. | Setup | Time Setting
ABL90 FLEX | div. Reinigungen durchgef., Einstellungen im Serviceprogr. überprüft, Startphase abgewartet, Kal. + QC i.O. | div. Reinigungen durchgef.; Einstellungen im Serviceprogr. überprüft; Startphase abgewartet; Kal. + QC i.O. | SWO-25813 | No | div. Cleanings through; Settings in the ServiceProgr. checked; start-up phase awaited; Kal. + QC in the first place | div. Cleanings through; Settings in the ServiceProgr. checked; start-up phase awaited; Kal. + QC in the first place | Assisted Replacement | SC | Wetability
ABL90 FLEX | Einlaß hakt nach wie vor bei einigen Proben Einlaßmodul erneuert | Einlaß hakt nach wie vor bei einigen Proben Einlaßmodul erneuert | SWO-25716 | Yes | Inlet still hooks up to some samples inlet module renewed | Inlet still hooks up to some samples inlet module renewed | Re classified RMED | Double entry, diff date
ABL90 FLEX | Remplacé le kit de maintenance e.o Remplacé l'aiguille d'aspiration e.o Remplacé le pack de solutions + la cassette de mesure e.o Calibration e.o contrôle de qualité e.o | Setup configuration Time Serveur + fichiers patients (numéro de travail) Cal+CQ e.o | SWO-25826 | No | Replaced the maintenance kit e.o Replaced the suction needle e.o Replaced the solution pack - the e.o calibration e.o calibration tape e.o quality control e.o | Setup configuration Time Server - patient files (work number) Cal-CQ e.o | Assisted Replacement | SC | Wetability
ABL815 FLEX | Oxi Test im Serviceprogramm durchgeführt, Darkvoltage außerhalb Bereich, Neonvoltage zu hoch. Vermutlich Netzteil defekt. Da zur Zeit 2Hyperthermieanwendungen laufen Termin auf 24.4 gelegt. | Telefonische Hilfe abgelehnt, Uweba soll bei einem Fehler informiert werden. Kundin bittet um Rückruf. | SWO-25867 | Yes | Oxi test performed in the service program, darkvoltage outside area, neonvoltage too high. Probably power supply defective. Since currently 2hyperthermy applications run date set to 24.4. | Telephone help rejected, Uweba should be informed in case of an error. Customer asks for a call back. | Re classified RMED | Telephone support
ABL825 FLEX |  |  | CECFC0439E4B7E47 | No |  |  | Defect | HW | pH/BG Module | electrode connector
ABL825 FLEX | Ausserbereichunterdrückung FHbF deaktiviert. HbF-Korrektion für alle Anteile aktiviert. tHb- und Service-tHb-Kal durchgeführt. Verifikationen in Ordnung. | Ausserbereichunterdrückung FHbF deaktiviert. HbF-Korrektion für alle Anteile aktiviert. tHb- und Service-tHb-Kal durchgeführt. Verifikationen in Ordnung. | SWO-25841 | Yes | Out-of-range suppression FHbF disabled. HbF correction for all shares. tHb and service tHb-Kal. Verifications in order. | Out-of-range suppression FHbF disabled. HbF correction for all shares. tHb and service tHb-Kal. Verifications in order. | Re classified RMED | Analyzer setup
ABL825 FLEX | FOUND BAD CABLE 14 POLE | REPLACED 14 POLE CABLE | CECFAC1B9E1F13D7 | No | FOUND BAD CABLE 14 POLE | REPLACED 14 POLE CABLE | Defect | HW | AutoCheck Module | Cable
ABL827 FLEX |  |  | CED7386A6236FEF8 | No |  |  | Assisted Replacement | Electrode | Electrode | Reference
ABL825 FLEX | Bei Eintreffen, Gerät OK, System überprüft, Service Aufkleber angebracht. | Der Kunde lehnt ab, weitere Wartungsmaßnahmen ab und möchte keine telefonische Unterstützung. Der Kunde bittet um einen Service. | SWO-25924 | No | On arrival, device OK, system checked, service sticker attached. | The customer refuses further maintenance and does not want telephone support. The customer asks for a service. | Unknown | HW | Oxi module | LS error
ABL90 FLEX | Manuell gespült,LP, waste tubing, SC und Anschlussdichtung getauscht, Hemolysator mit Hypoclorid manuell gespült, Kaltstartphase, Kal und QC,s vermessen |  | SWO-25877 | No | Manually rinsed,LP, waste tubing, SC and connection seal exchanged, hemolysator with hypoclorid manually rinsed, cold start phase, kal and QC,measured |  | Cleaned | Fluidic pathway
ABL80 FLEX CO-OX | Analysator überprüft, kompletter Ventilblock ausgetauscht (undicht), Wastemanifold ausgetauscht Fluidtransport siliconschlauch ausgetauscht (knick) tHb Kal ausgeführt QA zyklen i.o Funktionskontrolle i.o | Analysator überprüft, kompletter Ventilblock ausgetauscht (undicht), Wastemanifold ausgetauscht Fluidtransport siliconschlauch ausgetauscht (knick) tHb Kal ausgeführt QA zyklen i.o Funktionskontrolle i.o | SWO-25975 | No | Analyzer checked, complete valve block replaced (leaked), Wastemanifold replaced fluid transport silicon hose replaced (kink) tHb Kal executed QA cycles i.o function control i.o | Analyzer checked, complete valve block replaced (leaked), Wastemanifold replaced fluid transport silicon hose replaced (kink) tHb Kal executed QA cycles i.o function control i.o
ABL80 CO-OX | Found Computer cooling fan to be defective. | Cleaned Replaced Fan Assy. | CED36314003A0573 | No | Found Computer cooling fan to be defective. | Cleaned Replaced Fan Assy.
ABL80 FLEX BASIC | Anwender hat bereits Betriebsmittel bestellt. Lieferung, durch Zentrale am kommenden Werktag. Bei Lieferung Einsatz der neuen SC. Falls noch Fehler vorhanden sein sollten, meldet sich der Anwender erneut. Keine Rückmeldung bis heute erhalten. Lieferung erfolgt und Funktion Analysator o.K. | Anwender hat bereits Betriebsmittel bestellt. Lieferung, durch Zentrale am kommenden Werktag. Bei Lieferung Einsatz der neuen SC. Falls noch Fehler vorhanden sein sollten, meldet sich der Anwender erneut. Keine Rückmeldung bis heute erhalten. Lieferung erfolgt und Funktion Analysator o.K. | SWO-26025 | No | User has already ordered equipment. Delivery, by headquarters on the coming business day. When using the new SC. If there are still errors, the user logs back in. No feedback received to date. delivery and function analyzer o.K. | User has already ordered equipment. Delivery, by headquarters on the coming business day. When using the new SC. If there are still errors, the user logs back in. No feedback received to date. delivery and function analyzer o.K.
ABL90 |  |  | CED89658D837BF1E | Yes |  |  | Re classified RMED | Phone support
ABL90 FLEX | Fließwege gespült, Serviceprogramme durchgeführt , Einlass gespült Einlassdichtungshalter erneuert, Gerät kalibriert, SC aus Kundenbestand erneuert, Danach O2 QC Fehler, LP ersetzt, | Fließwege gespült, Serviceprogramme durchgeführt , Einlass gespült Einlassdichtungshalter erneuert, Gerät kalibriert, SC aus Kundenbestand erneuert, Danach O2 QC Fehler, LP ersetzt, | SWO-26289 | No | Flow paths flushed, service programs performed, inlet flushed inlet seal holder renewed, device calibrated, SC renewed from customer stock, After that O2 QC error, LP replaced, | Flow paths flushed, service programs performed, inlet flushed inlet seal holder renewed, device calibrated, SC renewed from customer stock, After that O2 QC error, LP replaced, | Cleaned | Fluidic pathway
ABL90 | Reseteo de la cpu y del equipo. Desobtrucion pasando agua en ambos sentidos. Chequeo por servicio. Cal, qc, muestras |  | CED35739002C0557 | No | Reset the cpu and equipment. Unobstructing by passing water in both directions. Check for service. Cal, qc, samples |  | Cleaned | Fluidic pathway
ABL835 FLEX | Desmontaje y limpieza de las valvulas de los dos modulos. Cal, qc y muestras |  | CED3B712004D0563 | No | Disassembly and cleaning of the valves of the two modules. Cal, qc and samples |  | Cleaned | HW | pH/BG, El/Met | valve
ABL825 FLEX | Barcode reader gedemonteerd en schoongemaakt. Veel stof verwijdert. Na terugplaatsing werkt barcode reader weer zoals het hoort. Instellingen FLEXQ en Autocheck gecontroleerd en bijgesteld. | Barcode reader gedemonteerd en schoongemaakt. Veel stof verwijdert. Na terugplaatsing werkt barcode reader weer zoals het hoort. Instellingen FLEXQ en Autocheck gecontroleerd en bijgesteld. | SWO-26096 | No | Barcode reader disassembled and cleaned. Removes a lot of dust. After reinstallation, barcode reader works again as it should. Settings FLEXQ and AutoCheck checked and adjusted. | Barcode reader disassembled and cleaned. Removes a lot of dust. After reinstallation, barcode reader works again as it should. Settings FLEXQ and AutoCheck checked and adjusted. | Cleaned | HW | Barcode Reader | front
ABL825 FLEX |  |  | CED3A77F00264C70 | No |  |  | Adjustment | HW | Oxi module | Lamp unit
ABL827 FLEX |  |  | CED34BE0000A7250 | Yes |  |  | Re classified RMED | Logistical upSWO
ABL800 BASIC | Fehlersuche Luftansaugstutzen gereinigt Reinigung Fließwege überprüft Ventile der Module demontiert und gewartet Wartung Pumpenrotor Pumpen Kalibrierung , LS Justage | Fehlersuche Luftansaugstutzen gereinigt Reinigung Fließwege überprüft Ventile der Module demontiert und gewartet Wartung Pumpenrotor Pumpen Kalibrierung , LS Justage | SWO-26010 | No | Troubleshooting air intake nozzle cleaned cleaning flow paths checked valves of modules dismantled and maintained maintenance pump rotor pumps calibration, LS adjustment | Troubleshooting air intake nozzle cleaned cleaning flow paths checked valves of modules dismantled and maintained maintenance pump rotor pumps calibration, LS adjustment | Cleaned | HW | Fluidic manifold | Air intake
ABL820 FLEX | Neue Keys geladen. | Neue Keys geladen. | SWO-26100 | No | New keys loaded. | New keys loaded. | Setup | Software key exp. | -
ABL835 FLEX | Es gibt keine Eintragungen im QC Plan. QC Plan neuangelegt. | Es gibt keine Eintragungen im QC Plan. QC Plan neuangelegt. | SWO-26113 | No | Es gibt keine Eintragungen im QC Plan. QC Plan neuangelegt. | Es gibt keine Eintragungen im QC Plan. QC Plan neuangelegt. | Other | User error | QC schedule lost
ABL825 FLEX | discrepancy between Po2 and So2, cause of air bubble detected in system | removed both ph bg and ei met moduals inspected valves, both air valves seized, removed cleaned and tested passed replaced inlet pre heater speaker was making a vibrating sound when in operation replaced speaker | CED376EC0024D80B | No | discrepancy between Po2 and So2, cause of air bubble detected in system | removed both ph bg and ei met moduals inspected valves, both air valves seized, removed cleaned and tested passed replaced inlet pre heater speaker was making a vibrating sound when in operation replaced speaker | Defect | HW | Inlet module | Heater
ABL825 FLEX | Inlet Modul zerlegt Tubing Nipple ersetz -> ohne Erfolg. Friction Detector unit ersetzt ok. Getestet ok Inlet Modul retour an Kunde | Inlet Modul zerlegt Tubing Nipple ersetz -> ohne Erfolg. Friction Detector unit ersetzt ok. Getestet ok Inlet Modul retour an Kunde | SWO-27748 | No | Inlet module decomposes Tubing Nipple -> to no avail. Detector Friction unit replaces ok. Tested ok Inlet module back to customer | Inlet module decomposes Tubing Nipple -> to no avail. Detector Friction unit replaces ok. Tested ok Inlet module back to customer | Defect | HW | Inlet module | Friction detector unit | LEH
ABL825 FLEX | Lampeneinstellung auf ca. 3300 durchgeführt. | Lampeneinstellung auf ca. 3300 durchgeführt. | SWO-26591 | Yes | lamp setting to approx. 3300. | lamp setting to approx. 3300. | Re classified RMED | Normal Installation
AQT90 FLEX | Demontaje y limpieza del waste module. Chequeo por servicio. Comprobacion con blancos. | Demontaje y limpieza del waste module. Chequeo por servicio. Comprobacion con blancos. | CED3B18B0037055C | No | Disassembly and cleaning of the waste module. Check for service. Target check. | Disassembly and cleaning of the waste module. Check for service. Target check.
AQT90 FLEX | Fehler Rotate Storage Wheel, Set Safe 22, Gerät geöffnet, steckengebliebene Cassette entfernt ( Nase abgebrochen), Safe Position i.O. QC durchgeführt | Fehler Rotate Storage Wheel, Set Safe 22, Gerät geöffnet, steckengebliebene Cassette entfernt ( Nase abgebrochen), Safe Position i.O. QC durchgeführt | SWO-26164 | No | Error Rotate Storage Wheel, Set Safe 22, Device opened, Stuck Cassette Removed (Nose Aborted), Safe Position performed in i.O. QC | Error Rotate Storage Wheel, Set Safe 22, Device opened, Stuck Cassette Removed (Nose Aborted), Safe Position performed in i.O. QC
ABL80 | The analyzer failed overnight calibrations. Afetr 2 system cycyles the system passed and then after several samples there was air detected errors. | Instructed to install a new SP, current pack was low on solution. needed to run several system cycles before the Calcium stabilized and passed external QC. | CED539CD76355DCF | No | The analyzer failed overnight calibrations. Afetr 2 system cycyles the system passed and then after several samples there was air detected errors. | Instructed to install a new SP, current pack was low on solution. needed to run several system cycles before the Calcium stabilized and passed external QC.
ABL90 FLEX | Fehlersuche Überprüfung im Service Programm Fließwege gespült LP reinitialisiert Austausch SC aufgrund Spülfehler, Fehlerhafter Kal pCO2 SW V3.1MR3 installiert | Fehlersuche Überprüfung im Service Programm Fließwege gespült LP reinitialisiert Austausch SC aufgrund Spülfehler, Fehlerhafter Kal pCO2 SW V3.1MR3 installiert | SWO-26260 | No | Troubleshooting check in service program flow paths flushed LP reinitialized replacement SC due to flushing error, faulty Kal pCO2 SW V3.1MR3 installed | Troubleshooting check in service program flow paths flushed LP reinitialized replacement SC due to flushing error, faulty Kal pCO2 SW V3.1MR3 installed | Cleaned | Fluidic pathway
ABL90 FLEX | Fehlerquelle nicht ersichtlich ! Gelegentliches Einfrieren des Bildschirms bei Patientenmessung (siehe SWO) ! | Fehlerquelle nicht ersichtlich ! Gelegentliches Einfrieren des Bildschirms bei Patientenmessung (siehe SWO) ! | SWO-26324 | No | Source of error not visible ! Occasional screen freezing for patient measurement (see SWO) ! | Source of error not visible ! Occasional screen freezing for patient measurement (see SWO) ! | Unknown | Occasional freeze
ABL90 | SOFTWARE WAS NOT LOADING | SHANE HAD REPLACED THE CPU AND THE CF CARD BUT WAS UNABLE TO GET THE THE SOFTWARE TO LOAD. THE APPLICATION SOFTWARE WAS RELOADED AND THE SP HAD TO BE REPLACED BECAUSE OF AN INTERVENTION ERROR. ALSO THE SETUP WAS LOADED FROM THE OTHER ANALYZER TO MAKE THEM THE SAME. | CED4BDA2761BB05B | No | SOFTWARE WAS NOT LOADING | SHANE HAD REPLACED THE CPU AND THE CF CARD BUT WAS UNABLE TO GET THE THE SOFTWARE TO LOAD. THE APPLICATION SOFTWARE WAS RELOADED AND THE SP HAD TO BE REPLACED BECAUSE OF AN INTERVENTION ERROR. ALSO THE SETUP WAS LOADED FROM THE OTHER ANALYZER TO MAKE THEM THE SAME. | Defect | SW
ABL90 with Auto Inlet | Einlassdichtungseinheit erneuert (903-585) - Funktionsüberprüfung etc. - KAL in mehreren Parametern nicht bestanden - Sensorkassette erneuert - späterer Rückruf mit Fr. Schindler erfolgt ! | Einlassdichtungseinheit erneuert (903-585) - Funktionsüberprüfung etc. - KAL in mehreren Parametern nicht bestanden - Sensorkassette erneuert - späterer Rückruf mit Fr. Schindler erfolgt ! | SWO-26326 | No | Inlet seal unit renewed (903-585) - function check etc. - KAL failed in several parameters - sensor cassette renewed - later recall with Ms. Schindler is carried out ! | Inlet seal unit renewed (903-585) - function check etc. - KAL failed in several parameters - sensor cassette renewed - later recall with Ms. Schindler is carried out ! | Replaced | Inlet Gasket
ABL90 FLEX PLUS | INSTALACION Y PRUEBAS DE FUNCIONAMIENTO ASI COMO CONFIGURACION Y FORMACION AL CLIENTE EN MANEJO Y FUNCIONAMIENTO. |  | CED50285760168A6 | Yes | INSTALLATION AND TESTS OF OPERATION AS SETTINGS AND TRAINING TO THE CUSTOMER IN MANAGEMENT AND OPERATION. |  | Re classified RMED | Installation and training
ABL825 FLEX |  |  | CED4CF417617EFB8 | No |  |  | unknown | Radiance
ABL835 FLEX | Belüftungskanal gespült, alle Elektroden neu membraniert, da lt. Plan fällig, Kal, QC | Belüftungskanal gespült, alle Elektroden neu membraniert, da lt. Plan fällig, Kal, QC | SWO-26331 | No | Ventilation duct flushed, all electrodes re-membraneed, as according to plan due, Kal, QC | Ventilation duct flushed, all electrodes re-membraneed, as according to plan due, Kal, QC | User Error | Fluidic pathway | Cleaned
ABL825 FLEX |  |  | CED499C276086910 | No |  |  | Adjustment | HW | Oxi module | Lamp unit
ABL825 FLEX | Ventile überprüft und gesäubert, Pumpenschläuche der Elektrodenmodule erneuert, Justagen, Kalibrationen und Qualitätskontrollen durchgeführt. | Ventile überprüft und gesäubert, Pumpenschläuche der Elektrodenmodule erneuert, Justagen, Kalibrationen und Qualitätskontrollen durchgeführt. | SWO-26215 | No | Valves checked and cleaned, pump hoses of electrode modules renewed, adjustments, calibrations and quality controls carried out. | Valves checked and cleaned, pump hoses of electrode modules renewed, adjustments, calibrations and quality controls carried out. | Defect | HW | pH/BG Module | Pump tubing
ABL825 FLEX | Customer has been experiencing above issue. Inspected sample path and measuring module valves. Noticed that both measuring module air valves were jammed in open position. | Replaced pH BG and El Met air valves. | CED4D1DF760A9DEC | No | Customer has been experiencing above issue. Inspected sample path and measuring module valves. Noticed that both measuring module air valves were jammed in open position. | Replaced pH BG and El Met air valves. | Defect | HW | pH/BG, El/Met | Valve
ABL825 FLEX | Gerätestatus mit Kunden besprochen. Systemabsturz ist In der Nacht vom 20.04 zum 21.04 drei mal aufgetreten. Fehler ist zum ersten Mal in der letzten Woche aufgetreten. Kunde behilft sich immer wieder mit einem Neustart. Kunde wünscht einen Vor Service. | Gerätestatus mit Kunden besprochen. Systemabsturz ist In der Nacht vom 20.04 zum 21.04 drei mal aufgetreten. Fehler ist zum ersten Mal in der letzten Woche aufgetreten. Kunde behilft sich immer wieder mit einem Neustart. Kunde wünscht einen Vor Service. | SWO-26200 | Yes | Device status discussed with customers. System crash occurred three times during the night from 20.04 to 21.04. Error occurred for the first time last week. Customer always helps himself with a reboot. Customer wants a pre-service. | Device status discussed with customers. System crash occurred three times during the night from 20.04 to 21.04. Error occurred for the first time last week. Customer always helps himself with a reboot. Customer wants a pre-service. | Re classified RMED | HW | unknown
ABL90 FLEX | System gespült, Fließwege überprüft - i.O., Einlasssondenhalter ersetzt, Test wiederholen , akz. Kal abgewartet, Kal u. QC i.O. | System gespült, Fließwege überprüft - i.O., Einlasssondenhalter ersetzt, Test wiederholen , akz. Kal abgewartet, Kal u. QC i.O. | SWO-26367 | No | System flushed, flow paths checked - i.O., inlet probe holder replaced, test repeat, akz. Kal waited, Kal & QC i.O. | System flushed, flow paths checked - i.O., inlet probe holder replaced, test repeat, akz. Kal waited, Kal & QC i.O. | Assisted Replacement | SP | Leak
ABL90 |  |  | CED64465EC532CB6 | No |  |  | Defect | HW | Flow Selector PCB
ABL90 | VERIFIED THAT WHEN YOU TRY TO DO A CAPILLARY SAMPLE THERE IS NO VIDEO. TRYIED TO RELOAD WINDOWS AND THIS ANALYZER REQUIRES NEW WINDOWS FOR UIM 945. ALSO THE PRINTER REQUIRES A USB CABLE TO LOAD ON ANALYZER. | N A | CED5FEB4EC267A27 | No | VERIFIED THAT WHEN YOU TRY TO DO A CAPILLARY SAMPLE THERE IS NO VIDEO. TRYIED TO RELOAD WINDOWS AND THIS ANALYZER REQUIRES NEW WINDOWS FOR UIM 945. ALSO THE PRINTER REQUIRES A USB CABLE TO LOAD ON ANALYZER. | N A | Defect | SW
ABL815 FLEX | PCO 2 Elektrode neu membraniert, Kal ok, Lac instabil, mit Haustechnik alle Elektroden membraniert, Kal und QC durchgeführt | das pCO2 Problem besteht weiterhin; Anwender wünscht den Einbau einer neuen pCO2 Elektrode E788; der Anwender wünscht einen Vorort-Service zur Beseitigung der Geräte-Störung; | SWO-26541 | No | PCO 2 electrode re-membraneed, Kal ok, Lac unstable, with building technology all electrodes membraneed, Kal and QC performed | the pCO2 problem persists; User wants the installation of a new pCO2 electrode E788; the user wants a suburban service to eliminate the device malfunction; | Assisted Replacement | Electrode | Electrode | pCO2
ABL825 FLEX | Barcode Reader ausgetauscht | Barcode Reader ausgetauscht | SWO-26600 | No | Barcode Reader replaced | Barcode Reader replaced | Defect | HW | Barcode Reader | Internal
ABL835 FLEX | Målekammer skiftet | Målekammer skiftet | SWO-26594 | No | Measuring chamber Change | Measuring chamber Change | Defect | HW | El/Met Module | Measuring chamber
ABL825 FLEX | Visited customer yesterday for the same issue (lactate drift flagging frequently). Inspected measuring module valves and found measuring module air valves jammed in open position again. Replaced measuring module air valves, using 902-700 instead of 902-701. Performed zero current calibration, cleaned dummy electrodes and El Met measuring chamber. | Replaced Lactate electrode connector. Lactate eventually started acting up again, checked measuring module valves again, and the air valves were stuck. | CED5FCB8EC0B9DF3 | No | Visited customer yesterday for the same issue (lactate drift flagging frequently). Inspected measuring module valves and found measuring module air valves jammed in open position again. Replaced measuring module air valves, using 902-700 instead of 902-701. Performed zero current calibration, cleaned dummy electrodes and El Met measuring chamber. | Replaced Lactate electrode connector. Lactate eventually started acting up again, checked measuring module valves again, and the air valves were stuck. | Defect | HW | pH/BG, El/Met | Valve
ABL90 FLEX | Manuell gespült im Service Flüssigkeitstransport getestet, Kaltstartphase wird nicht bestanden, Strecke Hemolyzer nicht voll durchgängig, Hemolyzer bestellt | Manuell gespült im Service Flüssigkeitstransport getestet, Kaltstartphase wird nicht bestanden, Strecke Hemolyzer nicht voll durchgängig, Hemolyzer bestellt | SWO-26744 | No | Manually rinsed in service liquid transport tested, cold start phase is not passed, route Hemolyzer not fully continuous, Hemolyzer ordered | Manually rinsed in service liquid transport tested, cold start phase is not passed, route Hemolyzer not fully continuous, Hemolyzer ordered | Unknown | Revisit required
ABL90 FLEX | Änderung der Einheit für Parameter tBil von µmol dl auf mg dl nicht durchgeführt, da diese Einheit im Gerät schon hinterlegt ist. | Änderung der Einheit für Parameter tBil von µmol dl auf mg dl nicht durchgeführt, da diese Einheit im Gerät schon hinterlegt ist. | SWO-26806 | Yes | Modification of the unit for parameter tBil from smol dl to mg dl not carried out, because this unit is already stored in the device. | Modification of the unit for parameter tBil from smol dl to mg dl not carried out, because this unit is already stored in the device. | Re classified RMED | No complaint, SETUP
ABL90 | Se retira el equipo del servicio y se lleva a revisar al taller. Se comprueba los tubos, bahía, sensores bomba y entrada de muestra sin detectar fallo. Se cambia el SC y el equipo funciona, realiza 18 muestras y vuelve a aparecer el mismo problema, se coloca otro SC y después de realizar 12 muestras vuelve a aparecer el mismo problema, se pasa en el circuito neumatico agua a presión sin éxito, se pasa después solución limpiadora de electrodos y el equipo queda operativo. | Lavar el SC con solución detergente. | CED7219C62344024 | No | The equipment is removed from service and taken to the workshop. Tubes, bay, pump sensors and sample input are checked without detecting fault. The SC is changed and the equipment works, it performs 18 samples and the same problem reappears, another SC is placed and after performing 12 samples the same problem reappears, it is passed in the pressure water system without success, then cleaning solution is passed electrodes and the equipment is operational. | Wash the SC with detergent solution. | Cleaned | Fluidic pathway | Hypo cl.
ABL835 FLEX |  |  | CED887A8D825B50E | No |  |  | Assisted Replacement | Electrode | Electrode | Reference | LEH
ABL825 FLEX | Clot in the ABL825. | Ran rinse until clot came out. | CED7AD5F62382A68 | No | Clot in the ABL825. | Ran rinse until clot came out. | Cleaned | Fluidic pathway | - |  | LEH
ABL815 FLEX | Fließwege gereinigt, Elektrodenkontakt. mit Cl getauscht, Pumpenschlauch ersetzt, alle Ventile in den Messmodulen ausgebaut, gereinigt, Spüllös.ventil im BG Modul getauscht, Gerät mehrmals. kalibriert und QC durchgeführt, | Fließwege gereinigt, Elektrodenkontakt. mit Cl getauscht, Pumpenschlauch ersetzt, alle Ventile in den Messmodulen ausgebaut, gereinigt, Spüllös.ventil im BG Modul getauscht, Gerät mehrmals. kalibriert und QC durchgeführt, | SWO-26957 | No | Flow paths cleaned, electrode contact. with Cl exchanged, pump hose replaced, all valves in the measuring modules removed, cleaned, rinsing valve in the BG module exchanged, device several times. calibrated and QC performed, | Flow paths cleaned, electrode contact. with Cl exchanged, pump hose replaced, all valves in the measuring modules removed, cleaned, rinsing valve in the BG module exchanged, device several times. calibrated and QC performed, | Defect | HW | Valve, rinse |  | LEH
ABL825 FLEX | Put in a new hard drive and loaded Windows and 6.13 | Now showing disc boot error Now it goes to the blue screen of death.Ordered new CPU and will return. | CED745D3622FDCB7 | No | Put in a new hard drive and loaded Windows and 6.13 | Now showing disc boot error Now it goes to the blue screen of death. Ordered new CPU and will return. | Defect | HW | HDD |  | LEH
ABL835 FLEX |  |  | CED77DEC622D2301 | No |  |  | Adjustment | HW | AutoCheck Module | Laser
ABL825 FLEX | Issue was isolated to be caused by jammed measuring module air valves, but replacing the valves did not resolve the issue Replaced Analyzer Control PCB, re- installed SW and software keys. Issue eventually came back after successfully calibrating instrument. Inspected measuring module valves again, both air valves were jammed. | Replaced both El Met and pH BG modules with used ones for investigative purposes. Noticed the air valve on only one of them was jamming, so replaced it. | CED735D2625E4C9B | No | Issue was isolated to be caused by jammed measuring module air valves, but replacing the valves did not resolve the issue Replaced Analyzer Control PCB, re- installed SW and software keys. Issue eventually came back after successfully calibrating instrument. Inspected measuring module valves again, both air valves were jammed. | Replaced both El Met and pH BG modules with used ones for investigative purposes. Noticed the air valve on only one of them was jamming, so replaced it. | Unknown | HW | Unknown |  | Team
ABL815 FLEX | Modul untersucht Friction Unit ersetzt. ABL wieder ok. | Modul untersucht Friction Unit ersetzt. ABL wieder ok. | SWO-26956 | No | Module examines Friction Unit replaced. ABL ok again. | Module examines Friction Unit replaced. ABL ok again. | Defect | HW | Inlet friction detection unit |  | LEH
ABL800 BASIC | FlexQ Modul muss getauscht werden, da sporadisch Mixer Tray nicht verfährt zusätzlich bei der Kal pO2 Nullstromfehler -> Zero Kal durchgeführt Lac membraniert, da QC Fehler | FlexQ Modul muss getauscht werden, da sporadisch Mixer Tray nicht verfährt zusätzlich bei der Kal pO2 Nullstromfehler -> Zero Kal durchgeführt Lac membraniert, da QC Fehler | SWO-26765 | No | FlexQ module must be replaced, as sporadic mixer tray does not proceed additionally with the Kal pO2 zero current error -> Zero Kal performed Lac membranes, since QC error | FlexQ module must be replaced, as sporadic mixer tray does not proceed additionally with the Kal pO2 zero current error -> Zero Kal performed Lac membranes, since QC error | Defect | HW | FlexQ Module
ABL90 | LOOKED THROUGH DATA LOGS AND FROM APRIL 10 TO APRIL 17, THE UNIT HAD MULTIPLE ERRORS FOR 1297 PRINTER OFFLINE AND 0216, GENERAL PRINTER ERROR. | REPLACED PRINTER, CABLE AND USB CABLE | CED9DD2C4E5CD366 | No | LOOKED THROUGH DATA LOGS AND FROM APRIL 10 TO APRIL 17, THE UNIT HAD MULTIPLE ERRORS FOR 1297 PRINTER OFFLINE AND 0216, GENERAL PRINTER ERROR. | REPLACED PRINTER, CABLE AND USB CABLE | Defect | HW | Printer
ABL90 FLEX | IP Konfiguration geprüft an beiden Geräten. Ping an Mirth Server gesendet, Ping i.O. Geräteseitige Einstellung i.O. POCT-IT kontaktiert. Die Konfiguration des Servers wird durch Herrn Azim überprüft. | IP Konfiguration geprüft an beiden Geräten. Ping an Mirth Server gesendet, Ping i.O. Geräteseitige Einstellung i.O. POCT-IT kontaktiert. Die Konfiguration des Servers wird durch Herrn Azim überprüft. | SWO-26882 | No | IP configuration checked on both devices. Ping sent to Mirth Server, Ping i.O. device-side setting i.O. POCT-IT contacted. The configuration of the server is checked by Mr. Azim. | IP configuration checked on both devices. Ping sent to Mirth Server, Ping i.O. device-side setting i.O. POCT-IT contacted. The configuration of the server is checked by Mr. Azim. | Setup | HIS/LIS
ABL90 FLEX | Anweisung nach Erhalt eines neuen LP diesen zu tauschen, und RT-06 (219 Aktivitäten) zu reklamieren. Service bei Software Update. | Anweisung nach Erhalt eines neuen LP diesen zu tauschen, und RT-06 (219 Aktivitäten) zu reklamieren. Service bei Software Update. | SWO-26911 | No | Instruction after receiving a new LP to exchange it, and to complain to RT-06 (219 activities). Service at Software Update. | Instruction after receiving a new LP to exchange it, and to complain to RT-06 (219 activities). Service at Software Update. | Assisted Replacement | SP | pCO2
ABL90 FLEX | Hemolyzer getauscht, Sensorbay gereinigt, Startphase, tHb Kal durchgeführt, Kal und QC vermessen | Hemolyzer getauscht, Sensorbay gereinigt, Startphase, tHb Kal durchgeführt, Kal und QC vermessen | SWO-26842 | No | Hemolyzer exchanged, Sensorbay cleaned, start phase, tHb Kal performed, Kal and QC measured | Hemolyzer exchanged, Sensorbay cleaned, start phase, tHb Kal performed, Kal and QC measured | Defect | HW | Hemolyzer
ABL90 FLEX | Systeem doorgespoel met demiwater. pO2 lensje schoongemaakt. Aan dhr. O Terpstra van de MID technische uitleg gegeven en daarbij alle servicefunctie uitgevoerd. Analyzer geeft bij opstarten lekkage problemen, inlet gasket vervangen. Nu goed maar uitgevoerd controle geeft bij Hb aan dat er geen nulcalibratie is gedaan. Analyzer uitgezet en daarna weer aan. Na volledige startup (45 min) zijn er geen foutmeldingen meer. Tijdsinstelling aangepast in windows, stond GMT+1 zonder daylight sayving maar een uur vooruit. Met commando bevestigd in windows. Na herstarten analyzer is de tijd goed. | Systeem doorgespoel met demiwater. pO2 lensje schoongemaakt. Aan dhr. O Terpstra van de MID technische uitleg gegeven en daarbij alle servicefunctie uitgevoerd. Analyzer geeft bij opstarten lekkage problemen, inlet gasket vervangen. Nu goed maar uitgevoerd controle geeft bij Hb aan dat er geen nulcalibratie is gedaan. Analyzer uitgezet en daarna weer aan. Na volledige startup (45 min) zijn er geen foutmeldingen meer. Tijdsinstelling aangepast in windows, stond GMT+1 zonder daylight sayving maar een uur vooruit. Met commando bevestigd in windows. Na herstarten analyzer is de tijd goed. | SWO-26942 | No | System is flush with demi water. pO2 Lens cleaned. To Mr. O Terpstra of the MID technical explanation given and all the service function performed. Analyzer indicates leakage problems at startup, inlet gasket replacement. Now good but carried out control at HB indicates that no zero calibration is done. Analyzer turned off and then on again. After full startup (45 min) There are no more error messages. Time setting adjusted in Windows, GMT + 1 without daylight sayving was only one hour ahead. With command confirmed in Windows. After restarting analyzer the time is good. | System is flush with demi water. pO2 Lens cleaned. To Mr. O Terpstra of the MID technical explanation given and all the service function performed. Analyzer indicates leakage problems at startup, inlet gasket replacement. Now good but carried out control at HB indicates that no zero calibration is done. Analyzer turned off and then on again. After full startup (45 min) There are no more error messages. Time setting adjusted in Windows, GMT + 1 without daylight sayving was only one hour ahead. With command confirmed in Windows. After restarting analyzer the time is good. | Cleaned | Fluidic pathway | Clot
ABL815 FLEX | Netzteil ausgebaut und erneuert, stdevdark Value mit 1,5 mV OK. Bei Oxi complete Test trat Neon Voltage to High auf. Im Betrieb z.Zt. OK. | Netzteil ausgebaut und erneuert, stdevdark Value mit 1,5 mV OK. Bei Oxi complete Test trat Neon Voltage to High auf. Im Betrieb z.Zt. OK. | SWO-26972 | No | Power supply removed and renewed, stdevdark value with 1.5 mV OK. At Oxi complete Test, Neon Voltage to High occurred. In operation at the right time OK. | Power supply removed and renewed, stdevdark value with 1.5 mV OK. At Oxi complete Test, Neon Voltage to High occurred. In operation at the right time OK. | Defect | HW | Power Supply |  | LEH
ABL825 FLEX | checked fluid flow and electrode | Noted micro bubbles traveling from V6 rinse valve in fluidic module.Removed module and replaced v6. Removed all fluidic module valves and inspected cleaned. Reinstalled original ph and electrolyte modules. | CED88A79D816E7A9 | No | checked fluid flow and electrode | Noted micro bubbles traveling from V6 rinse valve in fluidic module. Removed module and replaced v6. Removed all fluidic module valves and inspected cleaned. Reinstalled original ph and electrolyte modules. | Cleaned | Fluidic pathway | Valve, rinse | V6 | LEH
ABL835 FLEX | Erneuerung Display Unit - Funktionsüberprüfung etc. | Erneuerung CPU Unit - Rücktausch Display Unit - Funktionsüberprüfung etc. | SWO-26869 | No | Renewal Display Unit - Function check etc. | Renewal CPU Unit - Exchange Display Unit - Function check etc. | Defect | HW | DMS_PC | Touch screen | LEH
TCM TOSCA | Controle des cables et contacts : OK Remplacement du batterie. Si on débrancher le cable d'alimentation le batterie fonctionne bien. Après il charge normallement. Test : OK | Controle des cables et contacts : OK Remplacement du batterie. Si on débrancher le cable d'alimentation le batterie fonctionne bien. Après il charge normallement. Test : OK | SWO-26943 | No | Cable and contact control: OK Battery replacement. If one unplugthe power cable the battery works well. After it normally charges. Test: OK | Cable and contact control: OK Battery replacement. If one unplugthe power cable the battery works well. After it normally charges. Test: OK | Defect | HW | Power Supply | Power Supply
ABL80 FLEX Basic version |  |  | CEDC77853A405EF2 | No
ABL90 | Analyzer is rebooting due to the internal battery. | Disconnected the battery Ordered a new battery. Installed a new cassette, brought system to ready. | CEDC40973A278292 | No | Analyzer is rebooting due to the internal battery. | Disconnected the battery Ordered a new battery. Installed a new cassette, brought system to ready. | Defect | HW | Battery module
ABL90 FLEX | Scanner demontiert und gereinigt Plastischere ist leicht beschlagen anschl. Test mehrere Barcodes -i.O. | Scanner demontiert und gereinigt Plastischere ist leicht beschlagen anschl. Test mehrere Barcodes -i.O. | SWO-27151 | No | Scanner disassembled and cleaned Plasticiser is slightly fogged attached. Test multiple barcodes -i.O. | Scanner disassembled and cleaned Plasticiser is slightly fogged attached. Test multiple barcodes -i.O. | Adjustment | HW | Cleaned barcode reader
ABL90 FLEX | Fehlersuche Kalibrierung und QC durchgeführt, Glu Wert im Lev C liegt konstant zu tief Erneuter Austausch LP Fließwege gespült, Einlaßdichtung erneuert Geräteneustart QC Glu im Lev C bleibt trotz aller Maßnahmen zu tief anschl. SC erneuert , danach Kal und QC i.O. | Fehlersuche Kalibrierung und QC durchgeführt, Glu Wert im Lev C liegt konstant zu tief Erneuter Austausch LP Fließwege gespült, Einlaßdichtung erneuert Geräteneustart QC Glu im Lev C bleibt trotz aller Maßnahmen zu tief anschl. SC erneuert , danach Kal und QC i.O. | SWO-27152 | No | Troubleshooting calibration and QC performed, Glu value in Lev C is constantly flushed to deep re-exchange LP flow paths, inlet seal renewed device restart QC Glu in the Lev C remains too deep in spite of all measures. SC renewed, then Kal and QC i.O. | Troubleshooting calibration and QC performed, Glu value in Lev C is constantly flushed to deep re-exchange LP flow paths, inlet seal renewed device restart QC Glu in the Lev C remains too deep in spite of all measures. SC renewed, then Kal and QC i.O. | Assisted Replacement | SC | Glu
ABL90 FLEX | Setup configuration Time Serveur + fichiers patients (numéro de travail) Cal+CQ e.o | Setup configuration Time Serveur + fichiers patients (numéro de travail) Cal+CQ e.o | SWO-27214 | No | Setup configuration Time Server - patient files (work number) Cal-CQ e.o | Setup configuration Time Server - patient files (work number) Cal-CQ e.o | Setup | HIS/LIS | Time setting
ABL800 BASIC | Fehlersuche Überprüfung Fließwege, Wartung Ventile BG Modul Überprüfung und WartungTransducer -Wert Transducer Offset Voltage liegt außer Bereich zu tief (0.064) , könnte trotz Wartungsmaßnahmen und Geräteneustart nicht angehoben werden, Ersatzteil nicht vorhanden , wird geordert verbogene Einlaßnadel gerichtet, Einlaß gereinigt LS justage | ABL ist o.k., offsetwerte wieder imBereich, Austausch der Gasunit deshalb zurück gestellt, abwarten ob Fehler wieder auftritt | SWO-27150 | No | Troubleshooting check flow paths, maintenance valves BG module inspection and maintenanceTransducer value Transducer Offset Voltage lies out of range to deep (0.064), could not be raised despite maintenance measures and device restart; Spare part not present, is ordered bent inlet needle directed, inlet cleaned LS adjustment | ABL is OK, offset values back in the range, replacement of the gas unit therefore postponed, waiting if errors occur again | Defect | HW | Inlet Module | Probe bend | LEH
ABL825 FLEX | Installed new CPU, did not help. Tried new power supply, did not help. | Put back their power supply, CPU and HD, same problem. Escalated to Jamie to get a loaner DMS. | CEDC61B73A50CD5C | No | Installed new CPU, did not help. Tried new power supply, did not help. | Put back their power supply, CPU and HD, same problem. Escalated to Jamie to get a loaner DMS. | Unknown | Unknown | Unknown |  | LEH
ABL800 BASIC | PC Modul ausgebaut, Drucker erneuert Modul eingebaut Gerät kalibriert und überprüft | PC Modul ausgebaut, Drucker erneuert Modul eingebaut Gerät kalibriert und überprüft | SWO-27047 | No | PC module removed, printer renewed module built-in device calibrated and checked | PC module removed, printer renewed module built-in device calibrated and checked | Defect | HW | Printer |  | LEH
ABL837 FLEX | Found that the unit is not print all the time, and not print correctly. I tried just the printer head to see if that would correct the issue it did not. Had to come back today with the whole printer unit. | replaced printer unit everything is working fine now. | CEDC34493A0601E8 | No | Found that the unit is not print all the time, and not print correctly. I tried just the printer head to see if that would correct the issue it did not. Had to come back today with the whole printer unit. | replaced printer unit everything is working fine now. | Defect | HW | Printer module |  | LEH
ABL800 BASIC | Fehlenden Stützring der Glu Elektrode ersetzt, Lac,pCO2 und pO2 überprüft. Referenzelektrode überprüft und gereinigt. Leicht verbogene Nadel gerichtet. Positionierung des FlexQ Schlittens überprüft und justiert. Kalibrationen mehrfach durchgeführt. QC ohne Beanstandung (Lv3 Lot 549) | Fehlenden Stützring der Glu Elektrode ersetzt, Lac,pCO2 und pO2 überprüft. Referenzelektrode überprüft und gereinigt. Leicht verbogene Nadel gerichtet. Positionierung des FlexQ Schlittens überprüft und justiert. Kalibrationen mehrfach durchgeführt. QC ohne Beanstandung (Lv3 Lot 549) | SWO-27091 | No | Missing support ring of the Glu electrode replaced, Lac,pCO2 and pO2 checked. Reference electrode checked and cleaned. Slightly bent needle directed. Positioning of the FlexQ slide checked and adjusted. calibrations performed several times. QC without complaint (Lv3 Lot 549) | Missing support ring of the Glu electrode replaced, Lac,pCO2 and pO2 checked. Reference electrode checked and cleaned. Slightly bent needle directed. Positioning of the FlexQ slide checked and adjusted. calibrations performed several times. QC without complaint (Lv3 Lot 549) | Assisted Replacement | Electrode | Membrane | Glu | LEH
ABL827 FLEX | Remplacé la carte d'alimentation (power supply) e.o Calibration e.o contrôle de qualité e.o | Remplacé la carte d'alimentation (power supply) e.o Calibration e.o contrôle de qualité e.o | SWO-27143 | No | Replaced the power supply e.o calibration e.o quality control e.o | Replaced the power supply e.o calibration e.o quality control e.o | Defect | HW | Power Supply | Sporadic temp error | LEH
AQT90 FLEX | Reagenz-Kassette war nicht richtig in Storagewheel eingelegt und somit konnte die Türe nur halb geschlossen werden. Die Kassette, welche blockiert hat, entfernt. Systemüberprüfung durchgeführt. --> i.O. Systemreinigung durchgeführt --> i.O. QC Multicheck Level,2 durchgeführt--> i.O. AQT wieder Messbereit. | Reagenz-Kassette war nicht richtig in Storagewheel eingelegt und somit konnte die Türe nur halb geschlossen werden. Die Kassette, welche blockiert hat, entfernt. Systemüberprüfung durchgeführt. --> i.O. Systemreinigung durchgeführt --> i.O. QC Multicheck Level,2 durchgeführt--> i.O. AQT wieder Messbereit. | SWO-27162 | No | Reagent cartridge was not inserted correctly into Storagewheel and thus the door could only be half closed. Removes the cartridge that has blocked. system verification. --> i.O. system cleaning performed --> i.O. QC Multicheck Level,2 performed--> i.O. AQT again ready to measure. | Reagent cartridge was not inserted correctly into Storagewheel and thus the door could only be half closed. Removes the cartridge that has blocked. system verification. --> i.O. system cleaning performed --> i.O. QC Multicheck Level,2 performed--> i.O. AQT again ready to measure.
TCM/Tosca | Electrode module is causing sensor not to be detected. Found spO2 sensor had failed while pCO2 was operating normally. | Replaced electrode module and sensor. Calibrated and verified operation. Monitor operates normally | CEDC3D993A1CA195 | No | Electrode module is causing sensor not to be detected. Found spO2 sensor had failed while pCO2 was operating normally. | Replaced electrode module and sensor. Calibrated and verified operation. Monitor operates normally | Defect | HW | Electrode Module
ABL90 FLEX | Beim Eintreffen war das Fach offen Im Service Menü das SC Fach mehrmals auf und zu gemacht SC getauscht da fällig Med Techniker wartet die Konditionierung der SC ab und meldet sich bei Bedarf | Beim Eintreffen war das Fach offen Im Service Menü das SC Fach mehrmals auf und zu gemacht SC getauscht da fällig Med Techniker wartet die Konditionierung der SC ab und meldet sich bei Bedarf | SWO-27466 | No | Upon arrival, the compartment was open In the service menu the SC compartment was exchanged several times on and made SC as due Med technician waits for the conditioning of the SC and reports if necessary | Upon arrival, the compartment was open In the service menu the SC compartment was exchanged several times on and made SC as due Med technician waits for the conditioning of the SC and reports if necessary | Unknown | HW | SC gate
ABL90 FLEX | Fehlersuche LP reinitialisiert, richtig aktiviert( war nicht vollständig runter gedrückt) Einlaßdichtung gereinigt LP erneuert , da Verfallsdatum erreicht Einlaufverfahren | Fehlersuche LP reinitialisiert, richtig aktiviert( war nicht vollständig runter gedrückt) Einlaßdichtung gereinigt LP erneuert , da Verfallsdatum erreicht Einlaufverfahren | SWO-27395 | No | Troubleshooting LP reinitialized, properly activated(was not fully pressed down) Inlet seal cleaned LP renewed, since expiry date reaches inlet procedure | Troubleshooting LP reinitialized, properly activated(was not fully pressed down) Inlet seal cleaned LP renewed, since expiry date reaches inlet procedure | User Error | SP activation
ABL90 FLEX | Lösungspack im Serviceprogramm ausgebaut, Lösungspack war nicht eingerastet. Lösungspack aktiviert. Während der Startphase diverse Fehler im Lösungstransport, SC90 aus Kundenbestand ersetzt. Funktion OK. | Lösungspack im Serviceprogramm ausgebaut, Lösungspack war nicht eingerastet. Lösungspack aktiviert. Während der Startphase diverse Fehler im Lösungstransport, SC90 aus Kundenbestand ersetzt. Funktion OK. | SWO-27429 | No | Solution pack in the service program, solution pack was not locked. Solution pack activated. During the start-up phase, various errors in solution transport, SC90 replaced from customer inventory. Function OK. | Solution pack in the service program, solution pack was not locked. Solution pack activated. During the start-up phase, various errors in solution transport, SC90 replaced from customer inventory. Function OK. | User Error | SP activation
ABL827 FLEX | SE VERIFICA ESTADO GENERAL DEL EQUIPO, SE HACE DESPROTEINIZACION DEL SISTEMA FLUIDICO, SE MIRAN REGISTROS, SE HACE TEST DE SERVICIO, SE COMPRUEBA FUNCINAMIENTO CON CAL Y QC | SE VERIFICA ESTADO GENERAL DEL EQUIPO, SE HACE DESPROTEINIZACION DEL SISTEMA FLUIDICO, SE MIRAN REGISTROS, SE HACE TEST DE SERVICIO, SE COMPRUEBA FUNCINAMIENTO CON CAL Y QC | CEDDAF36B00618CC | Yes | THE GENERAL STATE OF THE EQUIPMENT IS VERIFIED, THE FLUIDIC SYSTEM IS DEPROTEINIZATION, RECORDS ARE WATCHED, SERVICE TEST IS CHECKED, CAL AND QC FUNCINATION CHECKS | THE GENERAL STATE OF THE EQUIPMENT IS VERIFIED, THE FLUIDIC SYSTEM IS DEPROTEINIZATION, RECORDS ARE WATCHED, SERVICE TEST IS CHECKED, CAL AND QC FUNCINATION CHECKS | Re classified RMED | Courtesy
ABL815 FLEX | Module demontiert, gereinigt und getrocknet - Isolationsüberprüfung im Service - Funktionsüberprüfung etc. - QC abgebrochen ( kein vorderes Luftsegment im LS erkannt ) - Proteinentfernung - Spülung Airventil im Fluidic Modul - LS Justage - Pumpenkalibrierung - Funktionsüberprüfung ohne Beanstandung ! | Module demontiert, gereinigt und getrocknet - Isolationsüberprüfung im Service - Funktionsüberprüfung etc. - QC abgebrochen ( kein vorderes Luftsegment im LS erkannt ) - Proteinentfernung - Spülung Airventil im Fluidic Modul - LS Justage - Pumpenkalibrierung - Funktionsüberprüfung ohne Beanstandung ! | SWO-27444 | No | Modules dismantled, cleaned and dried - insulation check in service - function check etc. - QC aborted (no front air segment detected in LS) - protein removal - flushing air valve in fluidic module - LS adjustment - pump calibration - Functional check without complaint ! | Modules dismantled, cleaned and dried - insulation check in service - function check etc. - QC aborted (no front air segment detected in LS) - protein removal - flushing air valve in fluidic module - LS adjustment - pump calibration - Functional check without complaint ! | Cleaned | Fluidic pathway | Leak current |  | LEH
ABL800 BASIC | ABL ist o.k., offsetwerte wieder imBereich, Austausch der Gasunit deshalb zurück gestellt, abwarten ob Fehler wieder auftritt | ABL ist o.k., offsetwerte wieder imBereich, Austausch der Gasunit deshalb zurück gestellt, abwarten ob Fehler wieder auftritt | SWO-27357 | No | ABL is OK, offset values back in the range, replacement of the gas unit therefore postponed, waiting if errors occur again | ABL is OK, offset values back in the range, replacement of the gas unit therefore postponed, waiting if errors occur again | Defect | HW | Gas Module |  | LEH
ABL825 FLEX | System überprüft und gereinigt, Lufteinlass gereinigt, Proteinentfernung durchgeführt, diverse Elektroden membraniert, Tray justiert, LS Justage und Pumpen Kalibration durchgeführt, tHb Kalibration durchgeführt, Qualitätskontrollen Eingescannt, Kalibrationen und Qualitätskontrollen ohne Beanstandungen durchgeführt. | System überprüft und gereinigt, Lufteinlass gereinigt, Proteinentfernung durchgeführt, diverse Elektroden membraniert, Tray justiert, LS Justage und Pumpen Kalibration durchgeführt, tHb Kalibration durchgeführt, Qualitätskontrollen Eingescannt, Kalibrationen und Qualitätskontrollen ohne Beanstandungen durchgeführt. | SWO-27396 | No | System checked and cleaned, air inlet cleaned, protein removal performed, various electrodes membraneed, tray adjusted, LS adjustment and pump calibration performed, tHb calibration performed, quality controls scanned, calibrations and quality controls without complaints. | System checked and cleaned, air inlet cleaned, protein removal performed, various electrodes membraneed, tray adjusted, LS adjustment and pump calibration performed, tHb calibration performed, quality controls scanned, calibrations and quality controls without complaints. | Cleaned | Fluidic pathway | Air vent |  | LEH
ABL825 FLEX | Try new HD, loaded complete software | Replaced all pump tubings, complete calibration. remembraned all electrodes. Ran VK-R7 linearity kit, results to be sent to Gary | CEDDAE92B031538C | No | Try new HD, loaded complete software | Replaced all pump tubings, complete calibration. remembraned all electrodes. Ran VK-R7 linearity kit, results to be sent to Gary | Defect | HW | HDD |  | LEH
ABL825 FLEX | System überprüft und gereinigt, pH Elektrode gereinigt, pCO2 Elektrode neu membraniert, Kalibrationen und Qualitätskontrollen ohne Beanstandungen durchgeführt. | System überprüft und gereinigt, pH Elektrode gereinigt, pCO2 Elektrode neu membraniert, Kalibrationen und Qualitätskontrollen ohne Beanstandungen durchgeführt. | SWO-27956 | No | System checked and cleaned, pH electrode cleaned, pCO2 electrode re-membraneed, calibrations and quality controls carried out without complaints. | System checked and cleaned, pH electrode cleaned, pCO2 electrode re-membraneed, calibrations and quality controls carried out without complaints. | Assisted Replacement | Electrode | Membrane | pCO2 | LEH
ABL825 FLEX | DATA LOGS SHOW 794 EL MET MODULE ASYMMETTRY EROOR | REPLACED EL MET,PH BG .WASTE AND SOLUTION PUMP TUBINGS.PERFORMED SEVERAL SUCCESSFUL PUMP CALIBRATIONS.PERFORMED 2 POINT CALIBRATION | CEDD84EDB02837AF | No | DATA LOGS SHOW 794 EL MET MODULE ASYMMETTRY EROOR | REPLACED EL MET,PH BG . WASTE AND SOLUTION PUMP TUBINGS. PERFORMED SEVERAL SUCCESSFUL PUMP CALIBRATIONS. PERFORMED 2 POINT CALIBRATION | Defect | HW | Pump tubing |  | LEH
ABL800 BASIC | FlexQ Modul ausgetauscht (S N alt: R0699N004) , da Mixer Tray sporadisch nicht verfährt bzw. festfährt und erst im Service resetet werden muss notwendige Kalibrierungen durchgeführt | FlexQ Modul ausgetauscht (S N alt: R0699N004) , da Mixer Tray sporadisch nicht verfährt bzw. festfährt und erst im Service resetet werden muss notwendige Kalibrierungen durchgeführt | SWO-27435 | No | FlexQ module replaced (S N old: R0699N004) , because Mixer Tray does not drive or boots sporadically and must only be reset in the service necessary calibrations | FlexQ module replaced (S N old: R0699N004) , because Mixer Tray does not drive or boots sporadically and must only be reset in the service necessary calibrations | Defect | HW | FlexQ Module | tray does not move | LEH
ABL80 CO-OX | Could not an fluid blockage error, found a LS error and ran a LS cal and all was OK unit green and ready to run samples, ran level 3 QC, all passed. looked like a small leak in the LS tube on hemolyzer maybe. had leaked down on the Hemolyzer. | Replaced Hemolyzer and ran THB cal, all good. | CEDEA61126148FF8 | No | Could not an fluid blockage error, found a LS error and ran a LS cal and all was OK unit green and ready to run samples, ran level 3 QC, all passed. looked like a small leak in the LS tube on hemolyzer maybe. had leaked down on the Hemolyzer. | Replaced Hemolyzer and ran THB cal, all good.
ABL90 FLEX | Remplacé l'hémolyseur de l'oxymétrie e.o garantie Remplacé le kit de maintenance annuel e.o Calibration e.o contrôle de qualité e.o | Remplacé l'hémolyseur de l'oxymétrie e.o garantie Remplacé le kit de maintenance annuel e.o Calibration e.o contrôle de qualité e.o | SWO-27704 | No | Replaced the oxymetry hemolyizer e.o warranty Replaced the annual maintenance kit e.o Calibration e.o quality control e.o | Replaced the oxymetry hemolyizer e.o warranty Replaced the annual maintenance kit e.o Calibration e.o quality control e.o | Defect | HW | Hemolyzer
ABL90 FLEX | Gerät zerlegt Wastetubing def. JW Kit Teile eingebaut. System wieder dicht. SP ID PCB ersetzt..Kontakte def. | Gerät zerlegt Wastetubing def. JW Kit Teile eingebaut. System wieder dicht. SP ID PCB ersetzt..Kontakte def. | SWO-27595 | No | Device disassembles wastetubing def. JW Kit parts built in. system tight again. SP ID PCB replaced. Contacts def. | Device disassembles wastetubing def. JW Kit parts built in. system tight again. SP ID PCB replaced. Contacts def. | Defect | HW | Waste tube
ABL90 FLEX | Gerät unter zu Hilfenahme der Serivceprog. in den einzelnen Baugruppen durchgespült, Startphase abgewartet, danach Kal & QC i.O. | Gerät unter zu Hilfenahme der Serivceprog. in den einzelnen Baugruppen durchgespült, Startphase abgewartet, danach Kal & QC i.O. | SWO-27749 | No | device at the serivceprog. rinsed in the individual assemblies, waited for the start phase, then Kal & QC i.O. | device at the serivceprog. rinsed in the individual assemblies, waited for the start phase, then Kal & QC i.O. | Cleaned | Fluidic pathway
ABL90 FLEX | Roma.ini Eintrag vorgenommen. Schnittstelle auf extern gesetzt. Qc's vorgenommen | Roma.ini Eintrag vorgenommen. Schnittstelle auf extern gesetzt. Qc's vorgenommen | SWO-27654 | No | Roma.ini entry. interface set to external. Qc's done | Roma.ini entry. interface set to external. Qc's done | Setup | Roma.ini | Ext Barcode scanner
ABL80 CO-OX |  |  | CEDECD892652C64A | Yes |  |  | Re classified RMED | Preparing for Demo
ABL90 | SUSTITUCIÓN DE PANEL TACTIL PANTALLA | SUSTITUCIÓN DE PANEL TACTIL PANTALLA | CEDF54EF263FE73A | No | REPLACING SCREEN TACTIL PANEL | REPLACING SCREEN TACTIL PANEL | Defect | HW | Touch screen
ABL815 FLEX | OXI PCB ausgebaut und erneuert. System justiert. Funktion OK. | Netzteil ausgebaut und erneuert, stdevdark Value mit 1,5 mV OK. Bei Oxi complete Test trat Neon Voltage to High auf. Im Betrieb z.Zt. OK. | SWO-27755 | No | OXI PCB expanded and renewed. system is adjusted. Function OK. | Power supply removed and renewed, stdevdark value with 1.5 mV OK. At Oxi complete Test, Neon Voltage to High occurred. In operation at the right time OK. | Defect | HW | Power Supply |  | LEH
ABL800 BASIC | CPU - Unit getauscht, Startphase, konfig. erneuert, Kal und QC durchgeführt | CPU - Unit getauscht, Startphase, konfig. erneuert, Kal und QC durchgeführt | SWO-27564 | No | CPU unit swapped, start phase, config. renewed, Kal and QC carried out | CPU unit swapped, start phase, config. renewed, Kal and QC carried out | Defect | HW | DMS_PC | CPU | LEH
ABL800 BASIC | Gerät im technischen Service überprüft, Fehler kommt immer wieder, komplette PC Einheit leihweise eingebaut, Startphase Kal und QC vermessen | Gerät im technischen Service überprüft, Fehler kommt immer wieder, komplette PC Einheit leihweise eingebaut, Startphase Kal und QC vermessen | SWO-27634 | No | Device checked in technical service, error comes again and again, complete PC unit built-in on loan, start phase Kal and QC measured | Device checked in technical service, error comes again and again, complete PC unit built-in on loan, start phase Kal and QC measured | Defect | HW | HDD |  | LEH
ABL800 BASIC | Ventile überprüft und gereinigt Defektes Spülventil und defekte Airventilmembran im pH BG Modul ersetzt LS Justage und Pumpenkal Flüssigkeit sackt im pH BG Modul ab--> Austausch diverser Schläuche Kal2- und QC Patientenprobe weiterhin problematisch mit Fehler Cl, pCO2, pO2, Ref neu membraniert und Messkammer gereinigt System gespült Austausch Teile siehe Liste Abfallflasche getauscht LS Justage Pumpenkal Kal2- und QC Zehn Patientenmessungen ohne Probleme | Ventile überprüft und gereinigt Defektes Spülventil und defekte Airventilmembran im pH BG Modul ersetzt LS Justage und Pumpenkal Flüssigkeit sackt im pH BG Modul ab--> Austausch diverser Schläuche Kal2- und QC Patientenprobe weiterhin problematisch mit Fehler Cl, pCO2, pO2, Ref neu membraniert und Messkammer gereinigt System gespült Austausch Teile siehe Liste Abfallflasche getauscht LS Justage Pumpenkal Kal2- und QC Zehn Patientenmessungen ohne Probleme | SWO-27664 | No | Valves checked and cleaned Defective flushing valve and defective air valve membrane in the pH BG module replaces LS adjustment and pump scale liquid sags in the pH BG module --> Replacement of various hoses Kal2 and QC patient sample continues problematic with error Cl, pCO2, pO2, Ref new membraneand measuring chamber cleaned system rinsed replacement parts see list Waste bottle exchanged LS Adjustment Pump scale Kal2- and QC Ten patient measurements without problems | Valves checked and cleaned Defective flushing valve and defective air valve membrane in the pH BG module replaces LS adjustment and pump scale liquid sags in the pH BG module --> Replacement of various hoses Kal2 and QC patient sample continues problematic with error Cl, pCO2, pO2, Ref new membraneand measuring chamber cleaned system rinsed replacement parts see list Waste bottle exchanged LS Adjustment Pump scale Kal2- and QC Ten patient measurements without problems | Defect | HW | Valve, air |  | LEH
ABL825 FLEX | Beide Pumpenschläuche der Elektrodenmodule ersetzt, Pumpenkalibration, Werte des El Met-Moduls nicht gut, Wechsel des Rollenkranzes des El Met-Moduls, Pumpenkalibration, Kal, QC | Beide Pumpenschläuche der Elektrodenmodule ersetzt, Pumpenkalibration, Werte des El Met-Moduls nicht gut, Wechsel des Rollenkranzes des El Met-Moduls, Pumpenkalibration, Kal, QC | SWO-27677 | No | Replaced both pump hoses of the electrode modules, pump calibration, values of the El Met module not good, change of the roller ring of the El Met module, pump calibration, Kal, QC | Replaced both pump hoses of the electrode modules, pump calibration, values of the El Met module not good, change of the roller ring of the El Met module, pump calibration, Kal, QC | Defect | HW | Pump tubing |  | LEH
TCM/Tosca | Remplacé l' hémolyseur de l'oxymétrie e.o garantie Remplacé le kit de maintenance annuel e.o Calibration e.o contrôle de qualité e.o | Remplacé l' hémolyseur de l'oxymétrie e.o garantie Remplacé le kit de maintenance annuel e.o Calibration e.o contrôle de qualité e.o | SWO-27699 | No | Replaced the oxymetry hemolyizer e.o warranty Replaced the annual maintenance kit e.o Calibration e.o quality control e.o | Replaced the oxymetry hemolyizer e.o warranty Replaced the annual maintenance kit e.o Calibration e.o quality control e.o
ABL80 CO-OX | will replace the OXI LS | replace the OXI LS and checked all diag. | CEE02B519C399003 | No | will replace the OXI LS | replace the OXI LS and checked all diag.
ABL90 FLEX | Archiv ausgelesen, Problem trat nach SC Wechsel am 24.4 erstmalig auf. SC90 945-009 R0320N 122 durch SC90 946-009 R0327N146 aus Kundenbestand ersetzt. Dokumentation an RDEUService geschickt. | Fr. Kirschbaum bittet um einen Service | SWO-27926 | No | Archive read out, problem occurred for the first time after SC change on 24.4. SC90 945-009 R0320N 122 replaced by SC90 946-009 R0327N146 from customer base. documentation sent to RDEUService. | Ms. Kirschbaum asks for a service | Assisted Replacement | SC | Leak current
ABL90 FLEX | ABL90 start niet meer op. Alleen bij de mixer is activiteit zichtbaar. Zodoende het volgende heb uitgevoerd: - Voeding vervangen, heeft niet geholpen. - CF card heb gecontroleerd, OK. Bij nader onderzoek blijkt dat CPU UIM910 defect te zijn. Defecte UIM910 heb vervangen door UIM945 (910-342P1506N0742) vervolgens nieuwe batterij option heb geinstalleerd, OK. Defecte onderdelen: - CPU UIM910, 910-287P1440N5290 - Defecte Batterij Module (brandlucht), 910-311P0049N029 - Defecte Batterij Module (verschillende pieptonen), 910-311P0063N017 |  | SWO-27987 | No | ABL90 does not start anymore. Only the mixer is showing activity. So have done the following:-Power supply replacement, did not help. -CF card checked, OK. Further investigation reveals that CPU UIM910 are defective. Defective UIM910 have replaced by UIM945 (910-342P1506N0742) then new battery option installed, OK. Defective parts:-CPU UIM910; 910-287P1440N5290-Defective battery Module (fire air); 910-311P0049N029-Faulty battery Module (different beeps); 910-311P0063N017 |  | Defect | HW | Battery module
ABL90 | Inspection of fluidic path revealed there is a small leak near the pinch tubing connection to the hemolyzer. Inlet gasket rubber appears to have a lot of debris on it that isn't rinsing off. | Replaced pinch valve tubing & inlet gasket. | CEE010369C63753E | No | Inspection of fluidic path revealed there is a small leak near the pinch tubing connection to the hemolyzer. Inlet gasket rubber appears to have a lot of debris on it that isn't rinsing off. | Replaced pinch valve tubing & inlet gasket. | Defect | HW | Oxi module | Pinch Valve tube
ABL825 FLEX | Neuen Drucker eingebaut. | Neuen Drucker eingebaut. | SWO-27842 | No | New printer installed. | New printer installed. | Defect | HW | Printer head |  | LEH
ABL820 FLEX | Level 1 and 2 running low pCO2.Performed Gas Reset,Checked Blue O-ring,Cleaned Chamber Contacts. Inspected Complete Unit Logs. No apparent issues. | Tried different Lot of Membranes, and pCO2 Electrode.Same results. | CEE008949C3E0FC4 | No | Level 1 and 2 running low pCO2.Performed Gas Reset,Checked Blue O-ring,Cleaned Chamber Contacts. Inspected Complete Unit Logs. No apparent issues. | Tried different Lot of Membranes, and pCO2 Electrode.Same results. | Assisted Replacement | Electrode | Electrode | pCO2 | LEH
ABL835 FLEX | Found inlet positioning error after changing consumables | Performed inlet calibration by extending the probe and hitting restart. Re-initialized analyzer, performed all calibrations and QC's | CEDFEFBB9C2D9B19 | No | Found inlet positioning error after changing consumables | Performed inlet calibration by extending the probe and hitting restart. Re-initialized analyzer, performed all calibrations and QC's | Adjustment | HW | Inlet module |  | LEH
ABL825 FLEX | Bildschirm demontiert und auf Lösungseintritt überprüft => alles i.O. Touchscreen kalibriert Gerätnestart Touchscreen i.O. | Bildschirm demontiert und auf Lösungseintritt überprüft => alles i.O. Touchscreen kalibriert Gerätnestart Touchscreen i.O. | SWO-27922 | No | Screen disassembled and checked for solution entry => all i.O. touchscreen calibrated device launch touchscreen i.O. | Screen disassembled and checked for solution entry => all i.O. touchscreen calibrated device launch touchscreen i.O. | Defect | HW | DMS_PC | Touch screen | LEH
ABL835 FLEX | Erneuerung CPU Unit - Rücktausch Display Unit - Funktionsüberprüfung etc. | Erneuerung CPU Unit - Rücktausch Display Unit - Funktionsüberprüfung etc. | SWO-27931 | No | Renewal CPU Unit - Exchange Display Unit - Function check etc. | Renewal CPU Unit - Exchange Display Unit - Function check etc. | Defect | HW | DMS_PC | Touch screen | LEH
ABL80 OSM | Liquid sensor cal error. | Replaced the liquid sensor. | CEE14F22122F2D91 | No | Liquid sensor cal error. | Replaced the liquid sensor.
ABL90 FLEX | ABL90 heb gerepareerd en voorzien van een nieuwe batterij, complte oximetrie test, tHb calibratie en QC's heb uitgevoerd vervolgens heb meegenomen naar de lokatie. De Instellingen heb teruggezet. Opmerking: na vervangen van CF card, datalogs zijn verloren gegaan maar de ABL90 is nog niet in gebruik en niet gekoppeld. |  | SWO-28082 | No | ABL90 have repaired and equipped a new battery, complte oximetry test, tHb calibration and Qcs have performed then have taken to the location. Reset the settings. Note: After replacing CF card, data logs have been lost but the ABL90 is not yet in use and not linked.
ABL90 FLEX | Bei Eintreffen steht Gerät mit geöffnetem Sensorfach, Gerät überprüft: Sensorfach lässt sich mechanisch schließen, aber mit öffnen über Serviceprog. schließt Fach nicht vollständig ( geht wieder auf), Austausch der Sensor. Interface PCB ( aus Ersatzteilmangel gegen eine gebrauchte aus GS11 Leihgerät), danach Sensorfachschließung i.O., Meldung Ref.-Elektrode in SC defekt-> Einbau einer Vor Ort am Gerät liegenden alten SC ( keine weitere SC mehr vorrätig lt. Kundenaussage), Startphase abgewartet, pCO2 Kal Empf. Fehler -> keine weitere SC Vor Ort->Kunde besorgt neu SC & meldet sich ggf. wieder | Bei Eintreffen steht Gerät mit geöffnetem Sensorfach, Gerät überprüft: Sensorfach lässt sich mechanisch schließen, aber mit öffnen über Serviceprog. schließt Fach nicht vollständig ( geht wieder auf), Austausch der Sensor. Interface PCB ( aus Ersatzteilmangel gegen eine gebrauchte aus GS11 Leihgerät), danach Sensorfachschließung i.O., Meldung Ref.-Elektrode in SC defekt-> Einbau einer Vor Ort am Gerät liegenden alten SC ( keine weitere SC mehr vorrätig lt. Kundenaussage), Startphase abgewartet, pCO2 Kal Empf. Fehler -> keine weitere SC Vor Ort->Kunde besorgt neu SC & meldet sich ggf. wieder | SWO-27986 | No | Upon arrival, device with the sensor compartment is open, device checked: Sensor compartment can be closed mechanically, but with open via service test. does not complete compartment (resumes), replacing the sensor. Interface PCB (due to lack of spare parts against a used GS11 rental device), thereafter sensor compartment closure i.O., message Ref.-Electrode in SC defective-> Installation of an on-site on-site old SC (no more SC in stock according to customer message), start phase awaited, pCO2 Kal Empf. Error -> no more SC On site->Customer re-cares SC & reports again if necessary | Upon arrival, device with the sensor compartment is open, device checked: Sensor compartment can be closed mechanically, but with open via service test. does not complete compartment (resumes), replacing the sensor. Interface PCB (due to lack of spare parts against a used GS11 rental device), thereafter sensor compartment closure i.O., message Ref.-Electrode in SC defective-> Installation of an on-site on-site old SC (no more SC in stock according to customer message), start phase awaited, pCO2 Kal Empf. Error -> no more SC On site->Customer re-cares SC & reports again if necessary | Defect | HW | SC ID PCB
ABL825 FLEX |  |  | CEE5329B74494D39 | No |  |  | Defect | HW | Inlet module | Leaking | Nipple kit
ABL80 FLEX CO-OX | Analysator ausgeschaltet, Schlauchwege überprüft --> nichts festgestellt, Analysator neu gestartet --> Startphase, SC (Lot 211246) eingebaut und SC (Lot 211836) ausgebaut, System- und Funktionskontrolle, System und Funktion i.O. |  | SWO-28179 | No | Analyzer switched off, hose pathways checked --> nothing detected, analyzer restarted --> start phase, SC (lot 211246) installed and SC (lot 211836) removed, system and function control, system and function i.O.
ABL90 FLEX | Gerät geprüft sensor cassette ersetzt i.o. | Gerät geprüft sensor cassette ersetzt i.o. | SWO-28176 | No | Device tested sensor cassette replaces i.o. | Device tested sensor cassette replaces i.o. | Assisted Replacement | SC | Wetability |  | mail to RSCH
ABL90 FLEX | Sensorkassette und Lösungspack neu initialisiert Kal- und QC i.O. | Sensorkassette und Lösungspack neu initialisiert Kal- und QC i.O. | SWO-28074 | No | Sensor cassette and solution pack reinitialized Kal- and QC i.O. | Sensor cassette and solution pack reinitialized Kal- and QC i.O. | Assisted Replacement | SP | Chip error
TCM400 |  |  | CEE63AE4EA01F052 | No
ABL825 FLEX | Valve 3-6-7-8 ersetzt. Gerät Kalibriert. | Valve 3-6-7-8 ersetzt. Gerät Kalibriert. | SWO-28078 | No | Valve 3-6-7-8 replaced. Device Calibrated. | Valve 3-6-7-8 replaced. Device Calibrated. | Defect | HW | Fluidic manifold | Valve | Leak current
ABL827 FLEX |  |  | CEE50F7D7459CF32 | No |  |  | Unknown | HW | pH/BG Module
ABL835 FLEX | Analysator wird getauscht gegen 754-R2015N004 - Installation etc. | Analysator wird getauscht gegen 754-R2015N004 - Installation etc. | SWO-28139 | No | Analyzer is exchanged for 754-R2015N004 - installation etc. | Analyzer is exchanged for 754-R2015N004 - installation etc. | Defect | HW | DMS_PC | Touch screen
AQT90 FLEX | Reajustado lector código de barras de la storage wheel por desajuste del mismo no leyendo quesitos de NTBMP |  | CEE4F2D0745D9AD2 | No | Reset storage wheel barcode reader by mismatching it not reading NTBMP cheeses
AQT90 FLEX |  |  | CEE57AA07419211A | No
AQT90 FLEX |  |  | CEE57B8B745E211D | No
ABL90 |  |  | CEE64754EA13944A | No |  |  | Defect | HW | Barcode Reader
ABL90 FLEX | Neue Sensorkassette inst. - Erneuerung Einlassdichtung notwendig - Analysator gespült - Serviceprogramme durchlaufen - Überprüfung Verschlauchung - Erneuerung Lösungspack - Funktionsüberprüfung etc. | Neue Sensorkassette inst. - Erneuerung Einlassdichtung notwendig - Analysator gespült - Serviceprogramme durchlaufen - Überprüfung Verschlauchung - Erneuerung Lösungspack - Funktionsüberprüfung etc. | SWO-28346 | No | New sensor cartridge inst. - Renewal inlet seal necessary - analyzer rinsed - service programs go through - inspection hose - renewal solution pack - function check etc. | New sensor cartridge inst. - Renewal inlet seal necessary - analyzer rinsed - service programs go through - inspection hose - renewal solution pack - function check etc. | Assisted Replacement | SP | Fluid transport
ABL90 | No battery Icon and PM due | replaced Battery same issue performed reset battery same problem replaced power supply now CPU bip not able repair on site ABL will be send to London Office | CEE65512EA3EF625 | No | No battery Icon and PM due | replaced Battery same issue performed reset battery same problem replaced power supply now CPU bip not able repair on site ABL will be send to London Office | Unknown | HW | Battery module
ABL90 |  |  | CEE63F46EA34A706 | No |  |  | Defect | HW | CPU
ABL90 | Confirmed report message on display upon arrival. Reviewed logs and observed 1184 (leak detected), 1063 (pressure test failure), 1369 (inlet gasket holder replacement failed), 1342 (leak detected). Ran service diagnostics, confirmed consistant pressure test failure @ 0.95?, intemittent failure of other areas, clot detection, pump calibration. | Installed Complete annual maintenance kit including all associated tubing, inlet gasket holder, black pinch valve tube (see associated SWO). Inspected all components, realigned wet section to ensure SP interface was not compromised. Replacing the sensor cassette resolved the reported problem and all related error codes. | CEE67507EA206F46 | No | Confirmed report message on display upon arrival. Reviewed logs and observed 1184 (leak detected), 1063 (pressure test failure), 1369 (inlet gasket holder replacement failed), 1342 (leak detected). Ran service diagnostics; confirmed consistant pressure test failure @ 0.95?, intemittent failure of other areas; clot detection, pump calibration. | Installed Complete annual maintenance kit including all associated tubing, inlet gasket holder, black pinch valve tube (see associated SWO). Inspected all components, realigned wet section to ensure SP interface was not compromised. Replacing the sensor cassette resolved the reported problem and all related error codes. | Assisted Replacement | SC | Pressure error
ABL90 FLEX | Gerät gespült, Lösungspack erneut eingesetzt, Einlassdichtung erneuert, LS Justage im Service durchgeführt, Anschließender Neustart ohne Probleme seitens des Lösungspack, pCO2 von der Empfindlichkeit zu niedrig (3,3%), Sensorkassette nach mehreren Kalibrationen von anderer Abteilung besorgt und eingebaut, | Frau F. wechselt das SP. Der Arretierungsmechanismuss hörte sich am Telefon sehr laut an. Frau F. wünscht einen Vorortservice. | SWO-28232 | Yes | Device rinsed, solution pack reinserted, inlet seal renewed, LS adjustment performed in service, subsequent restart without problems on the part of the solution pack, pCO2 from sensitivity to low (3.3%), sensor cartridge after several calibrations of other department concerned and installed, | Ms. F. changes the SP. The locking mechanism sounded very loud on the phone. Ms. F. wants a suburban service. | Re classified RMED | Phone support
ABL90 | Advised that probe was lifted all night. SC and SP were changed cleaned inlet gasket and backflushed instrument. performed fluid flow check. | Replaced sensor cassette again. | CEE64094EA42855B | No | Advised that probe was lifted all night. SC and SP were changed cleaned inlet gasket and backflushed instrument. performed fluid flow check. | Replaced sensor cassette again. | User Error | Inlet open
ABL827 FLEX | SE FORMATEA HD Y SE REINSTALA WINDOWS Y SW ABL8XX, PERO ES IMPOSIBLE LA CONEXION CON AU, SE REVISAN TENSIONES DE F.A. SIENDO ESTAS CORRECTAS SE CAMBIA CPU Y SE EFECTUAN PRUEBAS DE FUNCIONAMIENTO, SIENDO ESTAS CORRECTAS. CONFIGURACION DE GASOMETRO DESDE INICIO Y PARAMETROS DE RED, CONEXIONES CON RADIANCE NETOP ETC. CALIBRADO Y VERIFICADO CON Q.C. |  | CEE6951AEA1F3105 | No | HD IS FORMED AND WINDOWS AND SW ABL8XX RE-INSTALLED, BUT THE CONNECTION WITH AU IS IMPOSSIBLE, F.A. TENSIONS ARE REVIEWED BEING THESE CORRECT CPU IS CHANGED AND OPERATING TESTS ARE EFFECTED, BEING THESE CORRECT. GASOMETRO CONFIGURATION FROM HOME AND NETWORK PARAMETERS, CONNECTIONS WITH RADIANCE NETOP ETC. CALIBRATED AND VERIFIED WITH Q.C. |  | Defect | HW | CPU
ABL800 BASIC | ABL überprüft. Fällige Elektroden neu membraniert. PH Bg-Modul-Ventile und El Met-Modul-Ventile ausgebaut, gereinigt und geölt. El Met-Air-Ventil hing fest. Ls-Justage und Pumpen-Kal durchgeführt. QC durchgeführt. -> keine Verbesserung. Kal1 und Kal2-Lösung erneuert. Drift in Kal1 = alte Lösung war nicht OK. Gerät kalibriert. QC's durchgeführt. Alle 3 Level OK! Auto-Check-Karussell aufgefüllt. | ABL überprüft. Fällige Elektroden neu membraniert. PH Bg-Modul-Ventile und El Met-Modul-Ventile ausgebaut, gereinigt und geölt. El Met-Air-Ventil hing fest. Ls-Justage und Pumpen-Kal durchgeführt. QC durchgeführt. -> keine Verbesserung. Kal1 und Kal2-Lösung erneuert. Drift in Kal1 = alte Lösung war nicht OK. Gerät kalibriert. QC's durchgeführt. Alle 3 Level OK! Auto-Check-Karussell aufgefüllt. | SWO-28352 | No | ABL. Due electrodes re-membrane. PH Bg module valves and El Met module valves removed, cleaned and oiled. El Met Air valve hung tightly. Ls-adjustment and pump-scale. QC. -> no improvement. Kal1 and Kal2 solution. Drift in Kal1 = old solution was not OK. device calibrated. QC's. All 3 levels OK! Auto-check carousel replenished. | ABL. Due electrodes re-membrane. PH Bg module valves and El Met module valves removed, cleaned and oiled. El Met Air valve hung tightly. Ls-adjustment and pump-scale. QC. -> no improvement. Kal1 and Kal2 solution. Drift in Kal1 = old solution was not OK. device calibrated. QC's. All 3 levels OK! Auto-check carousel replenished. | Assisted Replacement | Solution bottle | Solution bottle | Kal1 and Kal2
ABL835 FLEX | THIS HAS BEEN AN ONGOING PROBLEM WITH THIS ANALYZER. ALSO VERIFIED DOING SAME SAMPLES AS RG12 AND THAT LACTATE MEMBRANE LASTS 20 DAYS. RG 11 RAN 1,108 SAMPLES IN LAST 30 DAYS. RG 12 RAN 1,071 SAMPLES IN LAST 30 DAYS. | REPLACED MEASURING CHAMBER, ELECTRODE CONNECTOR, AND EL MET PCB, DID ZERO CURRENT CALIBRATION AS WELL. | CEE645AEEA50EB88 | No | THIS HAS BEEN AN ONGOING PROBLEM WITH THIS ANALYZER. ALSO VERIFIED DOING SAME SAMPLES AS RG12 AND THAT LACTATE MEMBRANE LASTS 20 DAYS. RG 11 RAN 1,108 SAMPLES IN LAST 30 DAYS. RG 12 RAN 1,071 SAMPLES IN LAST 30 DAYS. | REPLACED MEASURING CHAMBER, ELECTRODE CONNECTOR, AND EL MET PCB, DID ZERO CURRENT CALIBRATION AS WELL. | Assisted Replacement | Electrode | Membrane | Lac
ABL820 FLEX | QC still running just above 2SD. Replaced Electrode, and new lot of Membranes No help change.Checked Gases. | Replaced Tbuings,Measure Chamber. | CEE64681EA1FC539 | No | QC still running just above 2SD. Replaced Electrode, and new lot of Membranes No help change. Checked Gases. | Replaced Tbuings,Measure Chamber. | Assisted Replacement | Electrode | Membrane | pCO2
ABL800 BASIC | Auto QC Modul wieder eingesetzt, 4. Elelktrolyt Key eingelesen, Kal und QC durchgeführt | Auto QC Modul wieder eingesetzt, 4. Elelktrolyt Key eingelesen, Kal und QC durchgeführt | SWO-28353 | No | Auto QC module reinstated, 4. Elelktrolyt Key read in, Kal and QC performed | Auto QC module reinstated, 4. Elelktrolyt Key read in, Kal and QC performed | Defect | HW | AutoCheck Module | Distribution PCB
ABL800 BASIC | Membranventile demontiert und gereinigt, pCO2 Elektrode erneuert, Gerät kalibriert. | Membranventile demontiert und gereinigt, pCO2 Elektrode erneuert, Gerät kalibriert. | SWO-28662 | No | Diaphragm valves disassembled and cleaned, pCO2 electrode renewed, device calibrated. | Diaphragm valves disassembled and cleaned, pCO2 electrode renewed, device calibrated. | Assisted Replacement | Electrode | Electrode | pCO2
AQT90 FLEX |  |  | CEE66437EA280BD6 | No
AQT90 FLEX |  |  | CEE67AB2EA3F4D7E | No
ABL80 | Customer had already replaced the sensor cassette and solution pack without resolving the problem. Troubleshot to the manifold unit. | Replaced the manifold unit, but this did not resolve the problem. Reseated the SC and SP. Removed all the waste tubing, flushed and reinstalled. None of this resolved the problem. | CEE7BECB603E1FD6 | No | Customer had already replaced the sensor cassette and solution pack without resolving the problem. Troubleshot to the manifold unit. | Replaced the manifold unit, but this did not resolve the problem. Reseated the SC and SP. Removed all the waste tubing, flushed and reinstalled. None of this resolved the problem.
ABL80 FLEX Basic version | Comprobar SC, limpiar restos de líquidos y circuito del desperdicio. | Limpiar el circuito del desperdicio. | CEE7AC446046B2C8 | No | Check SC, clean liquid residues and waste circuit. | Clean the waste circuit.
ABL90 | Analyzer needs to be setup from scratch. | Got a Patient sample from Nicole to set up report layout. Create 9 custom fields and 7 lists. Configure Radiance connection. | CEE7964C601FBEDB | No | Analyzer needs to be setup from scratch. | Got a Patient sample from Nicole to set up report layout. Create 9 custom fields and 7 lists. Configure Radiance connection. | Defect | SW | Data lost
ABL90 |  |  | CEE777F560271FBA | No |  |  | Cleaned | Fluidic pathway
ABL90 |  |  | CEEE4006AE0D321D | No |  |  | Defect | HW | Inlet Gasket
ABL90 FLEX | Neue Sensor Interface PCB eingebaut QC-Messungen mit QCAutoABL90 Level 1 bis 3 | Neue Sensor Interface PCB eingebaut QC-Messungen mit QCAutoABL90 Level 1 bis 3 | SWO-28420 | No | New Sensor Interface PCB built-in QC measurements with QCAutoABL90 Level 1 to 3 | New Sensor Interface PCB built-in QC measurements with QCAutoABL90 Level 1 to 3 | Defect | HW | SC PCB
ABL90 | VERIFIED THAT WHEN YOU TRY TO DO A CAPILLARY SAMPLE THERE IS NO VIDEO. | RE-INSTALLED MICROSOFT WINDOWS, RE-INSTALLED RADIOMETER 1.3.1 SOFTWARE. RESTORED ALL SETUP AND DATA. SET TIME AND DID SCREEN ALIGNMENT. WORKED WITH BOB ON COMPLETING THE STATIC IP ADDRESS. ALSO DID COMMIT COMMAND. ALSO ADDED ANOTHER CAPILLARY BUTTON. | CEE7931160092012 | No | VERIFIED THAT WHEN YOU TRY TO DO A CAPILLARY SAMPLE THERE IS NO VIDEO. | RE-INSTALLED MICROSOFT WINDOWS, RE-INSTALLED RADIOMETER 1.3.1 SOFTWARE. RESTORED ALL SETUP AND DATA. SET TIME AND DID SCREEN ALIGNMENT. WORKED WITH BOB ON COMPLETING THE STATIC IP ADDRESS. ALSO DID COMMIT COMMAND. ALSO ADDED ANOTHER CAPILLARY BUTTON. | Defect | SW | No video
ABL90 | The analyzer are not being used and they are waiting for training. They walked by and saw the error unable to pump solution. We checked and found 1324-inhomogeneous rinse solution LS sensors. It said to replace SP. We swapped the SP with the other analyzer and got the same message. We swapped the SC and the issue followed the SC. | I filled out the credit form and sent it in for SC replacement. They have no supplies on hand. | CEE7BB716030455C | No | The analyzer are not being used and they are waiting for training. They walked by and saw the error unable to pump solution. We checked and found 1324-inhomogeneous rinse solution LS sensors. It said to replace SP. We swapped the SP with the other analyzer and got the same message. We swapped the SC and the issue followed the SC. | I filled out the credit form and sent it in for SC replacement. They have no supplies on hand. | Assisted Replacement | SC | Wetability
ABL825 FLEX |  |  | CEE7A9076008E2F2 | No |  |  | Defect | HW | CPU
ABL820 FLEX | WDC reports inidcated pCO2 Low Level 1. Almost 2SD. Inspected Unit for possible pCO2 contamination, malfunction, etc. pCO2 electrode from stock proved to show same results. | Cleaned Calibrations. | CEE7E04960401188 | No | WDC reports inidcated pCO2 Low Level 1. Almost 2SD. Inspected Unit for possible pCO2 contamination, malfunction, etc. pCO2 electrode from stock proved to show same results. | Cleaned Calibrations. | Cleaned | Fluidic pathway | - | pCO2
ABL815 FLEX | QC Werte überprüft und OK. Labor beauftrag die Firma UKM zu häufig wenn die QC Ergebnis nicht wie gewünscht erzielt wurden aber innerhalb der Rilibäk Richtlinien liegen. Entsprechende Handhabung mit dem Labor ( Frau Schmitt), Herrn Gödecke und Herrn Gellert besprochen. Kal und QC sind OK. | QC Werte überprüft und OK. Labor beauftrag die Firma UKM zu häufig wenn die QC Ergebnis nicht wie gewünscht erzielt wurden aber innerhalb der Rilibäk Richtlinien liegen. Entsprechende Handhabung mit dem Labor ( Frau Schmitt), Herrn Gödecke und Herrn Gellert besprochen. Kal und QC sind OK. | SWO-28667 | No | QC values checked and OK. Labor orders the company UKM too often if the QC result was not achieved as desired but lie within the Rilibäk guidelines. Appropriate handling discussed with the laboratory (Mrs Schmitt), Mr Gödecke and Mr Gellert. Kal and QC are OK. | QC values checked and OK. Labor orders the company UKM too often if the QC result was not achieved as desired but lie within the Rilibäk guidelines. Appropriate handling discussed with the laboratory (Mrs Schmitt), Mr Gödecke and Mr Gellert. Kal and QC are OK. | Assisted Replacement | Electrode | Electrode | pCO2
ABL800 FLEX | Optionkey installiert | Optionkey installiert | SWO-28685 | Yes | Optionkey installed | Optionkey installed | Re classified RMED
ABL835 FLEX | Messkammer und Schlauchsystem geprüft, Fluid-und Modulventile geprüft, tHb Kalibration, Funktionskontrolle ok | Frau Domberg 05115322529 möchte dass der neue Analysator nochmals durch einen KDB geprüft wird! Level 2 QC Verletzungen in pCO2, Lyte und Oximetrie. Bitte mit Kundin Termin absprechen | SWO-28707 | No | Measuring chamber and hose system tested, fluid and module valves tested, tHb calibration, function control ok | Ms. Domberg 05115322529 wants the new analyzer to be checked again by a KDB! Level 2 QC injuries in pCO2, lyte and oximetry. Please arrange an appointment with customer | Assisted Replacement | Electrode | Electrode | pCO2
AQT90 FLEX | Storage Wheel wird durch eine Defekte Testkassetten blockiert. Kunststoffteil ist aus Testkassettten so ausgebrochen, dass es das Storage Wheel beim drehen hindert. Ausgebrochenes Teil ist nicht durch Produktion entstanden, sondern vmtl. Durch Handhabubg durch Anwender, z.B. Sturz der Testkassetten. Systemreinigung fällig, durchgeführt, erfolgreich. Systemüberprüfung durchgeführt, ohne Beanstandung. Analysator funktionsbereit. | Fr. Beilstein bittet Rückruf und Service | SWO-28683 | No | Storage Wheel is blocked by a defective test cartridge. Plastic part has broken out of test coffers in such a way that it prevents the storage wheel from turning. The broken part was not created by production, but vmtl. By hand-habubg by users, e.g. fall of the test cassettes. System cleaning due, performed, successful. system verification, without complaint. Analyzer ready for use. | Ms. Beilstein asks for recall and service
ABL80 FLEX BASIC | Einsatz des neuen LP, ohne Beanstandung. Analysator fordert auch neue SC an, erneuert. Kalibration o.K. QC konnte nicht durchgeführt werden, da kein QC Material Vorort. Analysator funktionsbereit. | Anwender hat bereits Betriebsmittel bestellt. Lieferung, durch Zentrale am kommenden Werktag. Bei Lieferung Einsatz der neuen SC. Falls noch Fehler vorhanden sein sollten, meldet sich der Anwender erneut. Keine Rückmeldung bis heute erhalten. Lieferung erfolgt und Funktion Analysator o.K. | SWO-28680 | No | Use of the new LP, without complaint. Analyzer also requests new SC, renewed. Calibration o.K. QC could not be carried out as no QC material on the ground. Analyzer ready for use. | User has already ordered equipment. Delivery, by headquarters on the coming business day. When using the new SC. If there are still errors, the user logs back in. No feedback received to date. delivery and function analyzer o.K.
ABL90 FLEX | Austausch des Touchscreens | Austausch des Touchscreens | SWO-28833 | No | Replacement of the touchscreen | Replacement of the touchscreen | Defect | HW | Touch screen
ABL90 FLEX | Uhrzeit eingestellt und gespeichert Analysator kalibriert und überprüft Interne QC OK | Uhrzeit eingestellt und gespeichert Analysator kalibriert und überprüft Interne QC OK | SWO-28704 | No | Time set and saved analyzer calibrated and checked Internal QC OK | Time set and saved analyzer calibrated and checked Internal QC OK | Setup | Time Setting
ABL90 FLEX | Einlass blockiert da Nadel sehr verbogen Einlassnadel getauscht Uhrzeit eingestellt und gespeichert Analysator kalibriert und überprüft Interne QC OK | Einlass blockiert da Nadel sehr verbogen Einlassnadel getauscht Uhrzeit eingestellt und gespeichert Analysator kalibriert und überprüft Interne QC OK | SWO-28703 | No | Inlet blocked since needle very bent inlet needle exchanged time set and saved analyzer calibrated and checked internal QC OK | Inlet blocked since needle very bent inlet needle exchanged time set and saved analyzer calibrated and checked internal QC OK | Setup | Time Setting
ABL90 | Checked units. Consumables needed replaced. | Replaced customer consumables, both solution pack and sensor cassette from customer stock. Provide additional troubleshoot to key operators. | CEE8EB34D65676EF | No | Checked units. Consumables needed replaced. | Replaced customer consumables, both solution pack and sensor cassette from customer stock. Provide additional troubleshoot to key operators. | User Error | Consumables exp
ABL90 FLEX | System mit Kondensmilch gespült, Einlassdichtung erneuert. | System mit Kondensmilch gespült, Einlassdichtung erneuert. | SWO-28632 | No | System rinsed with condensed milk, inlet seal renewed. | System rinsed with condensed milk, inlet seal renewed. | Assisted Replacement | SC | Wetability
ABL837 FLEX | Screen is blank, current errors, THB will not cal, sounds like DMS is not connecting. | replaced screen with the one from the old demo unit, cleaned chambers and checked wiring, still getting current errors, found loose cable on DMS | CEE8B6F0D65BAAB1 | Yes | Screen is blank, current errors, THB will not cal, sounds like DMS is not connecting. | replaced screen with the one from the old demo unit, cleaned chambers and checked wiring, still getting current errors, found loose cable on DMS | Re classified RMED | Demo setup
ABL825 FLEX | Gerät überprüft, Monitor und Kabelverbindung getestet, Computerunit zur Reparatur mitgenommen | Computerboard erneuert, Kal und QC o.k., tHb-Kal o.k. | SWO-28977 | No | Device checked, monitor and cable connection tested, computer unit taken for repair | Computerboard erneuert, Kal und QC o.k., tHb-Kal o.k. | Defect | HW | Revisit
ABL820 FLEX | QC was on Mean at time of last membrane replacement, lasted a couple of days then dropped off low and stayed low. QC returned to Mean after Cleaning. Customer to monitor unit. | Cleaned Measuring Chamber, Connector contacts, Etc. | CEE8C8AAD61F244B | No | QC was on Mean at time of last membrane replacement, lasted a couple of days then dropped off low and stayed low. QC returned to Mean after Cleaning. Customer to monitor unit. | Cleaned Measuring Chamber, Connector contacts, Etc. | Cleaned | Fluidic pathway | - | pCO2
ABL805 FLEX | OptionKey installiert | OptionKey installiert | SWO-28687 | Yes | OptionKey installed | OptionKey installed | Re classified RMED
ABL835 FLEX | Found Inlet probe to have a slight bend upward after extening it out (showed Technician on site) Found inlet gasket life to be failing in service programs. Found slight agitation while extening probe arm out during service cal test. Found Inlet Nupples to be full of debris | Repalced Inet Probe Replaced Inlet Gasket Replaced Inlet Nipples Replaced Friction Detection Unit to help mitigate further inlet calibration issues. Upon starting up, analyzer passed initialization, inlet calibration, and leakage test with no issues. Stepped analyzer through multiple ampoule breaks without carousel in to test with passing results. | CEE8DE6DD65BBB16 | No | Found Inlet probe to have a slight bend upward after extening it out (showed Technician on site) Found inlet gasket life to be failing in service programs. Found slight agitation while extening probe arm out during service cal test. Found Inlet Nupples to be full of debris | Repalced Inet Probe Replaced Inlet Gasket Replaced Inlet Nipples Replaced Friction Detection Unit to help mitigate further inlet calibration issues. Upon starting up, analyzer passed initialization, inlet calibration, and leakage test with no issues. Stepped analyzer through multiple ampoule breaks without carousel in to test with passing results. | Defect | HW | Inlet module | Friction detector unit
ABL835 FLEX | FlexQ Modul demontiert und geprüft, Ersatzmodul installiert, Funktionskontrolle ok |  | SWO-28693 | No | FlexQ module disassembled and tested, replacement module installed, function control ok |  | Defect | HW | FlexQ Module
ABL80 | Troubleshot to a faulty signal data board. Replaced sensor board for diagnostic reasons. | Replaced the signal data board, and this resolved the problem. | CEEA0F004C438736 | No | Troubleshot to a faulty signal data board. Replaced sensor board for diagnostic reasons. | Replaced the signal data board, and this resolved the problem.
ABL90 FLEX | A meglévő oldatpakk töbszöri instalálása, Öblítés, alvadék eltávolítása, Kalibrálás, QC |  | SWO-29026 | No | Töbszöri instalation of existing solution pack, rinse, clot removal, calibration, QC |  | Assisted Replacement | SP
ABL90 FLEX | Jahreswartung vorgezogen und alle Teile erneuert. Gerät daraufhin wieder funktionsbereit. | Jahreswartung vorgezogen und alle Teile erneuert. Gerät daraufhin wieder funktionsbereit. | SWO-28935 | No | annual maintenance and renewed all parts. device will then be ready for use again. | annual maintenance and renewed all parts. device will then be ready for use again. | Defect | HW | Inlet Gasket
ABL90 FLEX | Softwareupgrade 3.1MR3 ausgeführt Sensor erneuert da nur noch 2 Messungen vorhanden Funktionskontrolle i.o | Softwareupgrade 3.1MR3 ausgeführt Sensor erneuert da nur noch 2 Messungen vorhanden Funktionskontrolle i.o | SWO-28867 | Yes | Software upgrade 3.1MR3 executed sensor renewed since only 2 measurements still exist function control i.o | Software upgrade 3.1MR3 executed sensor renewed since only 2 measurements still exist function control i.o | Re classified RMED
ABL90 FLEX | Bewegliche Zunge nachjustiert. Öffnen und Schließen der Abdeckklappe mehrfach durchgeführt. In Ordnung. Lid Closed Indicator bestellt. | Bewegliche Zunge nachjustiert. Öffnen und Schließen der Abdeckklappe mehrfach durchgeführt. In Ordnung. Lid Closed Indicator bestellt. | SWO-28945 | No | Movable tongue readjusted. open and close the cover flap several times. All right. Lid Closed Indicator. | Movable tongue readjusted. open and close the cover flap several times. All right. Lid Closed Indicator. | Adjustment | HW | Sensor Cassette Door
ABL815 FLEX | Einlaßdichtung erneuert, Wege zum Abfall durchgespült, Einlaßnadel gespült, Dichtung pH erneuert, Glu Membran ausgetrocknet - div andere Parameter + Glu membraniert, mehrere Kal, QC |  | SWO-29243 | No | Inlet seal renewed, paths to waste rinsed, inlet needle rinsed, seal pH renewed, Glu membrane dried out - div other parameters + glu membrane, several Kal, QC |  | Assisted Replacement | HW | Inlet module | Inlet Gasket
ABL825 FLEX | Computerboard erneuert, Kal und QC o.k., tHb-Kal o.k. | Computerboard erneuert, Kal und QC o.k., tHb-Kal o.k. | SWO-28976 | No | Computerboard erneuert, Kal und QC o.k., tHb-Kal o.k. | Computerboard erneuert, Kal und QC o.k., tHb-Kal o.k. | Defect | HW | CPU
ABL835 FLEX | Halogenlampe justiert (3312). Oxi Komplett Test i.O. Kalibrationen durchgeführt. Weiterhin Oximetriefehler. Vermutlich Hemolyzer defekt. Ersatzteil wird beschafft und eingebaut | Laufende Sektio, Wartezeit 30 min. Austausch des Hemolysators (R3676N0005). Entfernen des Fremdkörpers in Ventilmembrane des EL Moduls. Initialisierung. Lamp Unit neu justiert (3250). Oxi Test durchgeführt. Während Kalibration Fehlermeldung '0255 Photodiode current setting too low', Oxi PcB ausgetauscht AU Software musste neu installiert werden (Oxi Modul inaktiv). Erneute Einstellung der Lamp Unit. Weiterhin '0255 Photodiode current setting too low' TonKr wird im Laufe des Tages ein neues Oxi Modul installieren. | SWO-29000 | No | Halogen lamp adjusted (3312). Oxi Complete Test i.O. Calibrations performed. Further oximetry errors. Probably Hemolyzer defective. Spare part is procured and installed | Ongoing sektio, waiting time 30 min. replacement of the hemolyser (R3676N0005). Removal of the foreign body in the valve membrane of the EL module. Initialization. Lamp Unit re-adjusted (3250). Oxi test. While calibration error message '0255 photodiode current setting too low', Oxi PcB exchanged AU software had to be reinstalled (Oxi module inactive). Re-adjustment of the Lamp Unit. Furthermore ,0255 Photodiode current setting too low' TonKr will install a new Oxi module during the day. | Defect | HW | Oxi module | Unknown
ABL835 FLEX | Laufende Sektio, Wartezeit 30 min. Austausch des Hemolysators (R3676N0005). Entfernen des Fremdkörpers in Ventilmembrane des EL Moduls. Initialisierung. Lamp Unit neu justiert (3250). Oxi Test durchgeführt. Während Kalibration Fehlermeldung '0255 Photodiode current setting too low', Oxi PcB ausgetauscht AU Software musste neu installiert werden (Oxi Modul inaktiv). Erneute Einstellung der Lamp Unit. Weiterhin '0255 Photodiode current setting too low' TonKr wird im Laufe des Tages ein neues Oxi Modul installieren. | Laufende Sektio, Wartezeit 30 min. Austausch des Hemolysators (R3676N0005). Entfernen des Fremdkörpers in Ventilmembrane des EL Moduls. Initialisierung. Lamp Unit neu justiert (3250). Oxi Test durchgeführt. Während Kalibration Fehlermeldung '0255 Photodiode current setting too low', Oxi PcB ausgetauscht AU Software musste neu installiert werden (Oxi Modul inaktiv). Erneute Einstellung der Lamp Unit. Weiterhin '0255 Photodiode current setting too low' TonKr wird im Laufe des Tages ein neues Oxi Modul installieren. | SWO-29001 | Yes | Ongoing sektio, waiting time 30 min. replacement of the hemolyser (R3676N0005). Removal of the foreign body in the valve membrane of the EL module. Initialization. Lamp Unit re-adjusted (3250). Oxi test. While calibration error message '0255 photodiode current setting too low', Oxi PcB exchanged AU software had to be reinstalled (Oxi module inactive). Re-adjustment of the Lamp Unit. Furthermore ,0255 Photodiode current setting too low' TonKr will install a new Oxi module during the day. | Ongoing sektio, waiting time 30 min. replacement of the hemolyser (R3676N0005). Removal of the foreign body in the valve membrane of the EL module. Initialization. Lamp Unit re-adjusted (3250). Oxi test. While calibration error message '0255 photodiode current setting too low', Oxi PcB exchanged AU software had to be reinstalled (Oxi module inactive). Re-adjustment of the Lamp Unit. Furthermore ,0255 Photodiode current setting too low' TonKr will install a new Oxi module during the day. | Re classified RMED | HW | Oxi module | Unknown
ABL835 FLEX | Erneuerung Oxi Modul - Oxi Complete Test - Funktionsüberprüfung etc. | Laufende Sektio, Wartezeit 30 min. Austausch des Hemolysators (R3676N0005). Entfernen des Fremdkörpers in Ventilmembrane des EL Moduls. Initialisierung. Lamp Unit neu justiert (3250). Oxi Test durchgeführt. Während Kalibration Fehlermeldung '0255 Photodiode current setting too low', Oxi PcB ausgetauscht AU Software musste neu installiert werden (Oxi Modul inaktiv). Erneute Einstellung der Lamp Unit. Weiterhin '0255 Photodiode current setting too low' TonKr wird im Laufe des Tages ein neues Oxi Modul installieren. | SWO-28948 | Yes | Erneuerung Oxi Module-Oxy Complete Test-Funktionsüberprüfung etc. | Ongoing sektio, waiting time 30 min. replacement of the hemolyser (R3676N0005). Removal of the foreign body in the valve membrane of the EL module. Initialization. Lamp Unit re-adjusted (3250). Oxi test. While calibration error message '0255 photodiode current setting too low', Oxi PcB exchanged AU software had to be reinstalled (Oxi module inactive). Re-adjustment of the Lamp Unit. Furthermore ,0255 Photodiode current setting too low' TonKr will install a new Oxi module during the day. | Re classified RMED | HW | Oxi module | Unknown
ABL825 FLEX |  |  | CEF31E0986099211 | No |  |  | Assisted Replacement | Electrode | Electrode | Reference
ABL825 FLEX | Inlet Diags failing on Gasket Life Check , Optical Point Check , Probe Position Check . | Replaced defective Friction Detector. | CEEA12964C3B3B15 | No | Inlet Diags failing on Gasket Life Check , Optical Point Check , Probe Position Check . | Replaced defective Friction Detector. | Defect | HW | Inlet module | Friction detector unit
AQT90 FLEX |  |  | CEEA2FE44C0A8A5C | No
ABL90 FLEX | Gerät manuell gespült Waste Connection Unit, Waste Tubing und Tubing for Valve erneuert Sensorkassette aus Kundenbestand erneuert Lösungspack (RV-08) geliefert Gerät kalibriert und überprüft |  | SWO-29125 | No | Device manually rinsed Waste Connection Unit, Waste Tubing and Tubing for Valve renewed sensor cartridge from customer stock renewed solution pack (RV-08) delivered device calibrated and checked |  | Defect | HW | Waste Connector Unit
ABL90 FLEX | Rettet kommunikationssetup i roma.ini | Rettet kommunikationssetup i roma.ini | SWO-29110 | No | Directed communication setup in Roma. ini | Directed communication setup in Roma. ini | Setup | Roma.ini | Ext Barcode scanner
ABL90 | LOOKED IN DATA LOGS. CALIBRATION FAILS FOR SODIUM. READS 225MV, 75.9 %, ALSO ALL 3 QC WHERE FAILING FOR SODIUM AND LACTATE. | DID BACKFLUSH 10 TIMES, DID SOME WASTE BLOOD ON ANALYZER AND ALL PASSED | CEEDED2FAE3299DB | No | LOOKED IN DATA LOGS. CALIBRATION FAILS FOR SODIUM. READS 225MV, 75.9 %, ALSO ALL 3 QC WHERE FAILING FOR SODIUM AND LACTATE. | DID BACKFLUSH 10 TIMES, DID SOME WASTE BLOOD ON ANALYZER AND ALL PASSED | Assisted Replacement | SC | Wetability
ABL90 | DID 10 BACKFLUSHES, CLEANED INLET GASKET, CLEANED INLET PROBE. RAN WASTE BLOOD, STILL HAVING INHOMOGENOUS RINSE. | REPLACED CUSTOMERS SOLUTION PACK WITH CUSTOMERS SUPPLY. ALSO COMPLETED WARRANTY FORM FOR SOLUTION PACK. | CEEDEA08AE1799D7 | No | DID 10 BACKFLUSHES, CLEANED INLET GASKET, CLEANED INLET PROBE. RAN WASTE BLOOD, STILL HAVING INHOMOGENOUS RINSE. | REPLACED CUSTOMERS SOLUTION PACK WITH CUSTOMERS SUPPLY. ALSO COMPLETED WARRANTY FORM FOR SOLUTION PACK. | Assisted Replacement | SP | Fluid transport
ABL90 FLEX | Installation und Funktionsüberprüfung etc. | Installation und Funktionsüberprüfung etc. | SWO-29223 | Yes | installation and function verification, etc. | installation and function verification, etc. | Re classified RMED | Returned from Willich
ABL90 FLEX | Lösungspack im Serviceprogramm ausgeworfen und neu installiert, Installationsschritte im Serviceprogramm durchgeführt, OK. Nach verlassen des Serviceprogramms war Installation möglich. Testmessung durchgeführt, NA Werte OK. | Lösungspack im Serviceprogramm ausgeworfen und neu installiert, Installationsschritte im Serviceprogramm durchgeführt, OK. Nach verlassen des Serviceprogramms war Installation möglich. Testmessung durchgeführt, NA Werte OK. | SWO-29218 | No | Solution pack ejected in the service program and reinstalled, installation steps performed in the service program, OK. After leaving the service program, installation was possible. Test measurement performed, NA values OK. | Solution pack ejected in the service program and reinstalled, installation steps performed in the service program, OK. After leaving the service program, installation was possible. Test measurement performed, NA values OK. | User Error | SP activation
ABL90 | Ran a calibration and pCO2 would not register after the calibration. | Replaced sensor cassette, and ran 5 waste blood samples. | CEEE1313AE20D0D1 | No | Ran a calibration and pCO2 would not register after the calibration. | Replaced sensor cassette, and ran 5 waste blood samples. | Assisted Replacement | SC | pCO2
ABL815 FLEX | Power Supply ersetzt, keine Änderung. Versucht PC von USB neu zu installieren, nicht möglich. Ersatzmodul aus Katharinenhospital installiert. Festplattenschaden an Ersatzgerät nach Booten. Ersatz-PC-Modul vom Krankenhaus vom Roten Kreuz eingebaut. Konfiguration von OP-Gerät übernommen und manuell ergänzt. Kalibrationen und Qualitätskontrollen in Ordnung. Datenübertragung zu Radiance und Copra in Ordnung. | Power Supply ersetzt, keine Änderung. Versucht PC von USB neu zu installieren, nicht möglich. Ersatzmodul aus Katharinenhospital installiert. Festplattenschaden an Ersatzgerät nach Booten. Ersatz-PC-Modul vom Krankenhaus vom Roten Kreuz eingebaut. Konfiguration von OP-Gerät übernommen und manuell ergänzt. Kalibrationen und Qualitätskontrollen in Ordnung. Datenübertragung zu Radiance und Copra in Ordnung. | SWO-29216 | No | Power Supply replaced, no change. Attempts to reinstall PC from USB, not possible. Replacement module from Katharinenhospital installed. Hard drive damage to replacement device after booting. Replacement PC module installed by the Red Cross Hospital. configuration of surgical device and manually supplemented. calibrations and quality controls in order. Data transfer to Radiance and Copra fine. | Power Supply replaced, no change. Attempts to reinstall PC from USB, not possible. Replacement module from Katharinenhospital installed. Hard drive damage to replacement device after booting. Replacement PC module installed by the Red Cross Hospital. configuration of surgical device and manually supplemented. calibrations and quality controls in order. Data transfer to Radiance and Copra fine. | Defect | HW | DMS_PC | Unknown
ABL825 FLEX |  |  | CEEEEC0D244FE369 | No |  |  | Defect | HW | Inlet module | Unknown
ABL825 FLEX | Autocheckmodul ausgebaut, an Unterseite des Autocheckmodul sind anstatt Senkkopfschrauben, erhöhte Schrauben von RMED eingebaut. Autocheckmodul steht am oberen Rand ca 5mm ab. Schrauben sind per Handbuch nicht verfügbar, werden von Altmodul ausgebaut. | Autocheckmodul ausgebaut, an Unterseite des Autocheckmodul sind anstatt Senkkopfschrauben, erhöhte Schrauben von RMED eingebaut. Autocheckmodul steht am oberen Rand ca 5mm ab. Schrauben sind per Handbuch nicht verfügbar, werden von Altmodul ausgebaut. | SWO-29011 | No | Autocheck module removed, at the bottom of the car check module instead of lowering head screws, raised screws are installed by RMED. Autocheck module is about 5mm at the top. Screws are not available via manual, are removed by old module. | Autocheck module removed, at the bottom of the car check module instead of lowering head screws, raised screws are installed by RMED. Autocheck module is about 5mm at the top. Screws are not available via manual, are removed by old module. | Defect | HW | AutoCheck Module | Screw type
AQT90 FLEX | Fehlersuche, div. Komponenten überprüft = div. cups nicht ausgeworfen, cups entfernt, Neustart, Service-dump gezogen, LQC i.O. | Fehlersuche; div. Komponenten überprüft = div. cups nicht ausgeworfen; cups entfernt; Neustart; Service-dump gezogen; LQC i.O. | SWO-29072 | No | Troubleshooting; div. Components checked = div. cups not ejected; cups removed; Restart; Service dump dragged; LQC i.O. | Troubleshooting; div. Components checked = div. cups not ejected; cups removed; Restart; Service dump dragged; LQC i.O.
AQT90 FLEX | De sensor voor de process unit blijkt niet te functioneren. Deze vervangen. Nu start de AQT wel zonder foutmelding door. | De sensor voor de process unit blijkt niet te functioneren. Deze vervangen. Nu start de AQT wel zonder foutmelding door. | SWO-29189 | No | The sensor for the process unit does not appear to function. Replace it. Now the AQT will start without an error message. | The sensor for the process unit does not appear to function. Replace it. Now the AQT will start without an error message.
AQT90 FLEX |  |  | CEEF225E245BFDB5 | No
ABL80 FLEX CO-OX | Lüfter an- und abgesteckt. Keine Änderung. CPU-Unit ausgebaut und Steckkontakt überprüft, in Ordnung. Nach Neustart erkennt ABL CF-Karte nicht mehr. Gerät zur weiteren Überarbeitung mitgenommen. Lüfter bestellt. Case CF-Karte erstellt. | Lüfter erneuert. | SWO-29436 | No | fan plugged in and off. No change. CPU unit removed and plug-in contact checked, all right. After rebooting, ABL no longer recognizes CF card. device for further revision. Fan ordered. Case CF map created. | Fan renewed.
ABL90 FLEX | System überprüft, Spülfehler während der Kalibration, Kondensmilch mehrfach gemessen, anschließende Kalibration ohne Beanstandungen durchgeführt, Interne Qualitätskontrollen ohne Beanstandungen durchgeführt, blutige Messungen ebenfalls ohne Beanstandungen. | System überprüft, Spülfehler während der Kalibration, Kondensmilch mehrfach gemessen, anschließende Kalibration ohne Beanstandungen durchgeführt, Interne Qualitätskontrollen ohne Beanstandungen durchgeführt, blutige Messungen ebenfalls ohne Beanstandungen. | SWO-29334 | No | System checked, flushing error during calibration, condensed milk measured several times, subsequent calibration carried out without complaints, internal quality checks carried out without complaints, bloody measurements also without complaints. | System checked, flushing error during calibration, condensed milk measured several times, subsequent calibration carried out without complaints, internal quality checks carried out without complaints, bloody measurements also without complaints. | Assisted Replacement | SC | Wetability
ABL90 FLEX | Service Menü, pO2 recalibration. Pumpentest erneut durchgeführt. LS Justage und Clot detection remove. Test wiederholen und Initialisierung. SC verfällt am 12.05. um 13:00, wird durch Labor erneuert. Kal und QC ohne Beanstandung. Analysator wieder funktionsbereit. | Service Menü, pO2 recalibration. Pumpentest erneut durchgeführt. LS Justage und Clot detection remove. Test wiederholen und Initialisierung. SC verfällt am 12.05. um 13:00, wird durch Labor erneuert. Kal und QC ohne Beanstandung. Analysator wieder funktionsbereit. | SWO-29300 | No | Service menu, pO2 recalibration. pump test again. LS Adjustment and Clot detection remove. Repeat test and initialize. SC expires on May 12 at 13:00, is renewed by laboratory. Kal and QC without complaint. Analyzer ready for use again. | Service menu, pO2 recalibration. pump test again. LS Adjustment and Clot detection remove. Repeat test and initialize. SC expires on May 12 at 13:00, is renewed by laboratory. Kal and QC without complaint. Analyzer ready for use again. | Cleaned | Fluidic pathway | Clot
ABL90 | Inspection of the unit revealed inlet value Clot detection of 0.01, SPfailing on Rinse, & Pressure test failing in service menu. Inlet probe appears to have a small dent in it. Customer recently replaced SC & I replaced the SP from the customer's stock. Time shifted on the analyzer intermittently today before the system S W (software) froze. | The system was rebooted, time adjusted to correct date & time, analyzer flushed & Inlet probe was replaced. | CEEF5F22244A7595 | No | Inspection of the unit revealed inlet value Clot detection of 0.01, SPfailing on Rinse, & Pressure test failing in service menu. Inlet probe appears to have a small dent in it. Customer recently replaced SC & I replaced the SP from the customer's stock. Time shifted on the analyzer intermittently today before the system S W (software) froze. | The system was rebooted, time adjusted to correct date & time, analyzer flushed & Inlet probe was replaced. | Replaced | Probe | Bend
ABL827 FLEX | SE CAMBIA ANALYZER CONTROL PCB, SE REINSTALA SW A MODULOS SE INSTALAN LLAVES, Y SE COMPRUEBA LA CONEXION ENTRE PC Y ANALYZER CONTROL PCB | SE CAMBIA ANALYZER CONTROL PCB, SE REINSTALA SW A MODULOS SE INSTALAN LLAVES, Y SE COMPRUEBA LA CONEXION ENTRE PC Y ANALYZER CONTROL PCB | CEEEEABD240D7D59 | No | ANALYZER CONTROL PCB IS CHANGED, SW IS RE-INSTALLED TO MODULES KEYS ARE INSTALLED, AND THE CONNECTION BETWEEN PC AND ANALYZER CONTROL PCB IS CHECKED | ANALYZER CONTROL PCB IS CHANGED, SW IS RE-INSTALLED TO MODULES KEYS ARE INSTALLED, AND THE CONNECTION BETWEEN PC AND ANALYZER CONTROL PCB IS CHECKED | Defect | HW | Analyzer Control PCB
ABL825 FLEX | Gerät überprüft, Oximetrietest durchgeführt, Meßkammer und Luftauslass gesäubert, Elektroden kontrolliert, pO2 Elektrode membraniert, Service tHb-Kal durchgeführt |  | SWO-29405 | No | Device checked, oximetry test performed, measuring chamber and air outlet cleaned, electrodes controlled, pO2 electrode membraneed, service tHb-Kal performed |  | Cleaned | Fluidic pathway | - | Oximodule
ABL835 FLEX | PCO2 LEVEL 1 S7735 LOT 587 FAILED NUMEROUS TIMES. 62.2, 62.7, 54.1: RANGE IS 63.1 TOT 69.1 | REPLACED MEASURING CHAMBER, ELECTRODE CONNECTOR, MEMBRANE VALVE, FLAT MEMBRANE VALVE AND PH BG PCB | CEEF29B3245823E7 | No | PCO2 LEVEL 1 S7735 LOT 587 FAILED NUMEROUS TIMES. 62.2, 62.7, 54.1: RANGE IS 63.1 TOT 69.1 | REPLACED MEASURING CHAMBER, ELECTRODE CONNECTOR, MEMBRANE VALVE, FLAT MEMBRANE VALVE AND PH BG PCB | Assisted Replacement | Electrode | Electrode | pCO2
ABL800 BASIC | Fehlersuche pO2-, Ref.-, Lac- und Glu-Elektroden membraniert Einlasssonde und Einlassdichtung erneuert 2-Pkt.-Kalibrierung mehrmals durchgeführt QC-Messungen mit AC5+ Level 1 und 3 | Fehlersuche pO2-, Ref.-, Lac- und Glu-Elektroden membraniert Einlasssonde und Einlassdichtung erneuert 2-Pkt.-Kalibrierung mehrmals durchgeführt QC-Messungen mit AC5+ Level 1 und 3 | SWO-29343 | No | Troubleshooting pO2, Ref.- Lac- and Glu-electrodes membrane-inlet probe and inlet seal renews 2-pcs. calibration performed several times QC measurements with AC5+ Level 1 and 3 | Troubleshooting pO2, Ref.- Lac- and Glu-electrodes membrane-inlet probe and inlet seal renews 2-pcs. calibration performed several times QC measurements with AC5+ Level 1 and 3 | Defect | HW | Inlet module | Probe bend
ABL827 FLEX |  |  | CEF2CFD38613054F | No |  |  | Assisted Replacement | Electrode | Electrode | Reference
ABL820 FLEX | CHECKED AUTOCHECK MODULE FOR ANY ABNORMAL FUNCTION. | CLEANED AUTOCHECK MIRROR AND REPLACED 14 POLE RIBBON CABLE CONTROLLING CAROUSEL FUNCTION. VERIFIED THAT AMPOULES ARE PROPERLY RECOGNIZED. INSTRUMENT IS WITHIN ACCEPTABLE OPERATING SPECIFICATIONS. | CEEF16EA245D0663 | No | CHECKED AUTOCHECK MODULE FOR ANY ABNORMAL FUNCTION. | CLEANED AUTOCHECK MIRROR AND REPLACED 14 POLE RIBBON CABLE CONTROLLING CAROUSEL FUNCTION. VERIFIED THAT AMPOULES ARE PROPERLY RECOGNIZED. INSTRUMENT IS WITHIN ACCEPTABLE OPERATING SPECIFICATIONS. | Cleaned | HW | AutoCheck Module | Barcode reader
ABL815 FLEX | Reinigung, Fehlersuche, Datensicherung, Modify KIRA im WebInstaller gewählt, Neustart, Datenrücksicherung, Inbetriebnahme, Kal & QK ok | Reinigung, Fehlersuche, Datensicherung, Modify KIRA im WebInstaller gewählt, Neustart, Datenrücksicherung, Inbetriebnahme, Kal & QK ok | SWO-29359 | No | Cleaning, troubleshooting, data backup, Modify KIRA selected in the WebInstaller, restart, data backup, commissioning, Kal & QK ok | Cleaning, troubleshooting, data backup, Modify KIRA selected in the WebInstaller, restart, data backup, commissioning, Kal & QK ok | Defect | SW | - | Unknown
ABL80 FLEX CO-OX | Neue CF-Karte eingesetzt und bespielt. Gerät bootet. | Neue CF-Karte eingesetzt und bespielt. Gerät bootet. | SWO-29435 | No | New CF card inserted and played. device boots. | New CF card inserted and played. device boots.
ABL80 FLEX CO-OX | Lüfter erneuert. | Lüfter erneuert. | SWO-29546 | No | Fan renewed. | Fan renewed.
ABL80 FLEX CO-OX | Lösungspack mit der Lot 212798 durch Lot: 0211940 | Bitte um Kontaktaufnahme. | SWO-29537 | No | Solution pack with the lot 212798 by lot: 0211940 | Please contact us.
ABL90 | Checked the sensor open close motor.Sensor door adjusted.Cleaned motor shaft. Error1081,1089 and 1176 . Ran all service checks. Liquid sensor checks failing. | Replaced Inlet probe,all tubings and gasket. | CEFADD3F4A4C8B51 | No | Checked the sensor open close motor. Sensor door adjusted. Cleaned motor shaft. Error1081,1089 and 1176 . Ran all service checks. Liquid sensor checks failing. | Replaced Inlet probe,all tubings and gasket. | Defect | HW | SC door
ABL90 | NETOP SETUP INCOMPLETE. | ORDERING ADDITIONAL SOFTWARE FOR NETOP REMOTE SETUP COMPLETION WILL RETURN TO COMPETE. | CEF09D5B9A029B25 | No | NETOP SETUP INCOMPLETE. | ORDERING ADDITIONAL SOFTWARE FOR NETOP REMOTE SETUP COMPLETION WILL RETURN TO COMPETE. | Setup | HIS/LIS | NetOp
ABL90 FLEX | Klemmschlauch und Abfallschlauch erneuert. | Klemmschlauch und Abfallschlauch erneuert. | SWO-29717 | No | clamping hose and waste hose. | clamping hose and waste hose. | Defect | HW | waste tube
ABL90 | Due to the recent installation, the customer has not been trained on the flush procedure. | Performed the flush procedure, and this resolved the problem. | CEF07C9E9A5BC463 | No | Due to the recent installation, the customer has not been trained on the flush procedure. | Performed the flush procedure, and this resolved the problem. | Cleaned | Fluidic pathway
ABL725 | Nettoyage des tuyaux du drainage : OK Le pompe du drainage se blocque parfois (si il tourne vite). Il était deux roue qui ne tourne pas bien. Graissage des roue : OK Remplacement preventive des tuyaux de drainage. Test : OK Calibration + CQ : OK | Nettoyage des tuyaux du drainage : OK Le pompe du drainage se blocque parfois (si il tourne vite). Il était deux roue qui ne tourne pas bien. Graissage des roue : OK Remplacement preventive des tuyaux de drainage. Test : OK Calibration + CQ : OK | SWO-29543 | No | Drain pipe cleaning: OK The drainage pump sometimes becomes blocked (if it rotates quickly). It was two wheels that didn't turn well. Wheel grease: OK Preventive replacement of drainage pipes. Test: OK Calibration - CQ: OK | Drain pipe cleaning: OK The drainage pump sometimes becomes blocked (if it rotates quickly). It was two wheels that didn't turn well. Wheel grease: OK Preventive replacement of drainage pipes. Test: OK Calibration - CQ: OK
ABL800 BASIC | Messkammer und Schlauchsystem geprüft und gereinigt, Einlassnadel geprüft und ausgetauscht, Reinigungslösung aktiviert, AC Modul geprüft, Funktionskontrolle ok | Messkammer und Schlauchsystem geprüft und gereinigt, Einlassnadel geprüft und ausgetauscht, Reinigungslösung aktiviert, AC Modul geprüft, Funktionskontrolle ok | SWO-29692 | No | Measuring chamber and hose system tested and cleaned, inlet needle tested and replaced, cleaning solution activated, AC module tested, function control ok | Measuring chamber and hose system tested and cleaned, inlet needle tested and replaced, cleaning solution activated, AC module tested, function control ok | Defect | HW | Inlet module | Probe
ABL827 FLEX | Seal rings missing from Lactate and Glucose electrodes. | Replace seal rings for electrodes. Clean measuring chambers. | CEF073419A214EF7 | No | Seal rings missing from Lactate and Glucose electrodes. | Replace seal rings for electrodes. Clean measuring chambers. | Assisted Replacement | Electrode | Electrode | Lac
ABL827 FLEX | Seal ring missing from electrode. | Replace seal ring. Clean measuring chamber. | CEF07D4C9A254EFB | No | Seal ring missing from electrode. | Replace seal ring. Clean measuring chamber. | Assisted Replacement | Electrode | Electrode | Lac
ABL825 FLEX | During Complete OXI Test phase regulation failed. | Replace hemolyzer unit. Perfor lamp alignment and Complete OXI Test. Perform service tHb calibration. | CEF0A4B29A099F27 | No | During Complete OXI Test phase regulation failed. | Replace hemolyzer unit. Perfor lamp alignment and Complete OXI Test. Perform service tHb calibration. | Defect | HW | Oxi module | Hemolyzer
ABL825 FLEX | Leaking from inlet due to hole in waste line. | Replace waste tubings and waste pump tubings. Replace inlet nipples and inlet gasket. Perform Complete OXI Test. Perform service tHb calibration. | CEF09F5C9A169F20 | No | Leaking from inlet due to hole in waste line. | Replace waste tubings and waste pump tubings. Replace inlet nipples and inlet gasket. Perform Complete OXI Test. Perform service tHb calibration. | Defect | HW | Fluidic manifold | Waste pump tube
ABL825 FLEX | In-house IT and Host IT are indicating that some Services associated with Data transmissions does not have several Services that need to be included with data was not present. | Complete Software install and Setup. | CEF0A7479A22019E | No | In-house IT and Host IT are indicating that some Services associated with Data transmissions does not have several Services that need to be included with data was not present. | Complete Software install and Setup. | Setup | HIS/LIS
ABL825 FLEX | 5 5 Examined Oxi module. 5 13 Examined Oxi module. | 5 5 Replaced Hemolyzer and lamp unit. Adjusted Oxi module and run service tHb. Run level 5 clear vial same problem. Ordered part. 5 13 Replaced spectrometer. Adjusted Oxi Module and run service tHb. Reloaded widows and application software version 6.13. Setup unit to run VKR-7 kit lot: 425. All levels in the kit ran fine. Verified operation of ABL825. | CEF08BD59A26E806 | No | 5 5 Examined Oxy module. 5 13 Examined Oxy module. | 5 5 Replaced Hemolyzer and lamp unit. Adjusted Oxi module and run service tHb. Run level 5 clear vial same problem. Ordered part. 5 13 Replaced spectrometer. Adjusted Oxi Module and run service tHb. Reloaded widows and application software version 6.13. Setup unit to run VKR-7 kit lot: 425. All levels in the kit ran fine. Verified operation of ABL825. | Defect | HW | Oxi module | Spectrometer
ABL827 FLEX | Proteinentfernung durchgeführt - Messkammer gereinigt Vorverstärkerkalibration durchgeführt, neue Lösungen angebracht Gerät kalibriert und überprüft | Proteinentfernung durchgeführt - Messkammer gereinigt Vorverstärkerkalibration durchgeführt, neue Lösungen angebracht Gerät kalibriert und überprüft | SWO-29492 | No | Protein removal performed - measuring chamber cleaned preamplifier calibration performed, new solutions attached device calibrated and checked | Protein removal performed - measuring chamber cleaned preamplifier calibration performed, new solutions attached device calibrated and checked | Assisted Replacement | Solution bottle | - | Crea
AQT90 FLEX | Anwender gewünschte Kalibrations-Kassette geliefert, da Lieferung an Technikbüro des Servictechnikers. QC Grenzverletzung bei TNT, Level blau. Werte wurden zu hoch gemessen. Bei der dritten Messung war Wert im Grenzbereich. Service Dunp gezogen. Analysator funktionsbereit. | Anwender gewünschte Kalibrations-Kassette geliefert, da Lieferung an Technikbüro des Servictechnikers. QC Grenzverletzung bei TNT, Level blau. Werte wurden zu hoch gemessen. Bei der dritten Messung war Wert im Grenzbereich. Service Dunp gezogen. Analysator funktionsbereit. | SWO-29521 | No | User desired calibration cassette, since delivery to technical office of the service technician. QC boundary violation at TNT, level blue. Values were measured too high. In the third measurement, value was in the boundary range. Service Dunp pulled. Analyzer ready for use. | User desired calibration cassette, since delivery to technical office of the service technician. QC boundary violation at TNT, level blue. Values were measured too high. In the third measurement, value was in the boundary range. Service Dunp pulled. Analyzer ready for use.
ABL80 CO-OX | They had a fluid blockage error and then a LS error | I opened the analyzer and found red fluid on top the Hemolyzer. I replaced the hemolyzer and ran setup | CEF1BD3D105B8494 | No | They had a fluid blockage error and then a LS error | I opened the analyzer and found red fluid on top the Hemolyzer. I replaced the hemolyzer and ran setup
ABL825 FLEX | Previous Service call performed Complete Software install=OK. Unit is able to Ping Host Server Reply.All settings in HIS LIS were to Host specs, all Windows TcpIp settings were to Host specs. Local IT is going back and forth from Static IP settings to DHCP for testing purposes. | Region Radiance rep suggested to Local IT to replace NIC board. | CEF215D6102E71D3 | Yes | Previous Service call performed Complete Software install=OK. Unit is able to Ping Host Server Reply.All settings in HIS LIS were to Host specs, all Windows TcpIp settings were to Host specs. Local IT is going back and forth from Static IP settings to DHCP for testing purposes. | Region Radiance rep suggested to Local IT to replace NIC board. | Re classified RMED | Radiance setup
ABL825 FLEX | pO2 membrán csere, zero beállítása. Kijelzős csatlakozó kábel csere. Bemenet csere, levegőző szelep tisztítás. |  | SWO-29584 | No | pO2 membrane replacement, zero adjustment. Display connection cable replacement. Inlet replacement, ventilation valve cleaning. |  | Unknown |  |  |  | 429
ABL90 | Sensor cassette bay filled with crystal build up. Leak suspected in sensor bay unit. Contact probes found bent. Ran service checks multiple times. Checked all tubings. LS failing during service checks.Pump calibration failing. system temperature out of range. rebooted the analyzer multiple times. Ran sensor simulator. Hemolyzer plugged. Flushed the analyzer multiple times. Inspected oxi valve and waste connector. | Replaced Sensor bay unit. Replaced Hemolyzer. Replaced all tubings and gaskets. Installed contact probes from old PCB for testing purpose. | CEF2F51B861A8445 | No | Sensor cassette bay filled with crystal build up. Leak suspected in sensor bay unit. Contact probes found bent. Ran service checks multiple times. Checked all tubings. LS failing during service checks. Pump calibration failing. system temperature out of range. rebooted the analyzer multiple times. Ran sensor simulator. Hemolyzer plugged. Flushed the analyzer multiple times. Inspected oxi valve and waste connector. | Replaced Sensor bay unit. Replaced Hemolyzer. Replaced all tubings and gaskets. Installed contact probes from old PCB for testing purpose. | Defect | HW | SC retainer
ABL90 |  |  | CEF2CFF98638308D | Yes |  |  | Re classified RMED | FAN / TU | SW install
ABL835 FLEX |  |  | CEF2B54C86232580 | Yes |  |  | Re classified RMED | Downgrade
ABL90 FLEX | LP und Sensorkassette erneuert System gespült und gereinigt QC durchgeführt | LP und Sensorkassette erneuert System gespült und gereinigt QC durchgeführt | SWO-29703 | No | LP and sensor cassette renewed system flushed and cleaned QC performed | LP and sensor cassette renewed system flushed and cleaned QC performed | Assisted Replacement | SP | Leak current
ABL90 | LOOKED IN DATA LOGS FOUND INHOMOGENEOUS QC1 SOLUTION, CODE 1083, INHOMOGENEOUS CAL 2 SOLUTION, CODE 1085. FLUSHED UNIT 10 TIMES, CLEANED INLET GASKET, CLEANED INLET PROBE. RAN WASTE BLOOD ON ANALYZER AND STILL WOULD NOT CLEAR. | REPLACED SOLUTION PACK FROM CUSTOMER'S SUPPLY. ALSO COMPLETED CREDIT FORM FOR CUSTOMER. | CEF6C543E826A5F4 | No | LOOKED IN DATA LOGS FOUND INHOMOGENEOUS QC1 SOLUTION, CODE 1083, INHOMOGENEOUS CAL 2 SOLUTION, CODE 1085. FLUSHED UNIT 10 TIMES, CLEANED INLET GASKET, CLEANED INLET PROBE. RAN WASTE BLOOD ON ANALYZER AND STILL WOULD NOT CLEAR. | REPLACED SOLUTION PACK FROM CUSTOMER'S SUPPLY. ALSO COMPLETED CREDIT FORM FOR CUSTOMER. | Assisted Replacement | SP | FT |  | 183
ABL90 FLEX | Sensorkassette erneuert. Nach Kundenwunsch gewartet bis die Kassette installiert ist und kein Fehler auftritt. | Sensorkassette erneuert. Nach Kundenwunsch gewartet bis die Kassette installiert ist und kein Fehler auftritt. | SWO-29882 | No | Sensor cartridge is renewed. Waiting for the cartridge to be installed and no error. | Sensor cartridge is renewed. Waiting for the cartridge to be installed and no error. | Assisted Replacement | SC | pCO2
ABL800 BASIC | PC Einheit zurück getauscht, Startphase Konfig eingespielt, mehrere Kal,s durchgeführt, QC vermessen | PC Einheit zurück getauscht, Startphase Konfig eingespielt, mehrere Kal,s durchgeführt, QC vermessen | SWO-29796 | No | PC unit swapped back, start phase config recorded, several kal,s performed, QC measured | PC unit swapped back, start phase config recorded, several kal,s performed, QC measured | Defect | HW | DMS_PC | HDD
ABL800 BASIC | Fehlersuche Überprüfung Ventile, Gasfluß, LS Demontage Ventile am BG Modul, Airvalve gereinigt LS Justage, Pumpenkalibrierung Kal 2 Lösung erneuert , da Abbruch Kal nach einziehen Kal 2 Lsg. Pumpenschläuche ausgetauscht Glu membraniert , da bei QC instabil | Fehlersuche Überprüfung Ventile, Gasfluß, LS Demontage Ventile am BG Modul, Airvalve gereinigt LS Justage, Pumpenkalibrierung Kal 2 Lösung erneuert , da Abbruch Kal nach einziehen Kal 2 Lsg. Pumpenschläuche ausgetauscht Glu membraniert , da bei QC instabil | SWO-29885 | No | Troubleshooting check valves, gas flow, LS disassembly valves on the BG module, airvalve cleaned LS adjustment, pump calibration Kal 2 solution renewed, since demolition Kal after retracting Kal 2 Lsg. Pump hoses replaced Glu membranes, since at QC unstable | Troubleshooting check valves, gas flow, LS disassembly valves on the BG module, airvalve cleaned LS adjustment, pump calibration Kal 2 solution renewed, since demolition Kal after retracting Kal 2 Lsg. Pump hoses replaced Glu membranes, since at QC unstable | Defect | HW | pH/BG Module | Pump tubing
ABL825 FLEX | Lactate would not calibrate. Called tech support and was told how to clean the tip of the electrode. | Cleaned the connector, electrode measuring well, and connections on the EL MET board. Cleaned the tip of the Electrode. | CEF6B41DE83E45D6 | No | Lactate would not calibrate. Called tech support and was told how to clean the tip of the electrode. | Cleaned the connector, electrode measuring well, and connections on the EL MET board. Cleaned the tip of the Electrode. | Cleaned | Electrode | Electrode | Lac
ABL825 FLEX | Lactate would not calibrate. Called tech support and was told how to clean the tip of the electrode. | Cleaned the connector, electrode measuring well, and connections on the EL MET board. Cleaned the tip of the Electrode. | CEF6B6B5E81145D9 | No | Lactate would not calibrate. Called tech support and was told how to clean the tip of the electrode. | Cleaned the connector, electrode measuring well, and connections on the EL MET board. Cleaned the tip of the Electrode. | Cleaned | Electrode | Electrode | Lac
ABL825 FLEX | Lactate would not calibrate. Cleaned the electrode and connctor. Cleaned the measuring chamber. | Cleaned the electrode and connctor. Cleaned the measuring chamber. Replaced the membrane from customers stock. | CEF6EDF4E85D6294 | No | Lactate would not calibrate. Cleaned the electrode and connctor. Cleaned the measuring chamber. | Cleaned the electrode and connctor. Cleaned the measuring chamber. Replaced the membrane from customers stock. | Cleaned | Electrode | Electrode | Lac
ABL825 FLEX | Referenz-Elektrode erneuert. | Referenz-Elektrode erneuert. | SWO-29795 | No | Reference electrode renewed. | Reference electrode renewed. | Assisted Replacement | Electrode | Electrode | Reference
ABL825 FLEX | Friction Unit ersetzt. | Friction Unit ersetzt. | SWO-29874 | No | Friction Unit replaced. | Friction Unit replaced. | Defect | HW | Inlet module | Friction detector unit
ABL830 FLEX | Fließwegegereinigt, div. Schläuche getauscht, Messkammer gereinigt, Pumpenkal. Und LS justage durchgeführt, Einl.dichtung ersetzt, BG-Modul ausgebaut, die Ventile gereinigt und Funktion überprüft, Ventil im Oximodul ausgebaut und gängig gemacht, Gerät geprüft und Probemessungen gemacht, | Fließwegegereinigt, div. Schläuche getauscht, Messkammer gereinigt, Pumpenkal. Und LS justage durchgeführt, Einl.dichtung ersetzt, BG-Modul ausgebaut, die Ventile gereinigt und Funktion überprüft, Ventil im Oximodul ausgebaut und gängig gemacht, Gerät geprüft und Probemessungen gemacht, | SWO-29895 | No | Flow paths cleaned, various hoses exchanged, measuring chamber cleaned, pump scale. And LS adjustment performed, single seal replaced, BG module removed, valves cleaned and function checked, valve removed in oximodule and made common, device tested and sample measurements made, | Flow paths cleaned, various hoses exchanged, measuring chamber cleaned, pump scale. And LS adjustment performed, single seal replaced, BG module removed, valves cleaned and function checked, valve removed in oximodule and made common, device tested and sample measurements made, | Cleaned | Fluidic pathway
ABL825 FLEX | Gerät neu gestartet. Eingaben wieder möglich. | Gerät neu gestartet. Eingaben wieder möglich. | SWO-29923 | No | Device restarted. Inputs possible again. | Device restarted. Inputs possible again. | Defect | HW | DMS_PC | Touch screen
ABL825 FLEX | Falsche Schrauben durch Senkkopfschrauben ersetzt. | Falsche Schrauben durch Senkkopfschrauben ersetzt. | SWO-29921 | No | Incorrect screws replaced by vertical head screws. | Incorrect screws replaced by vertical head screws. | Defect | HW | AutoCheck Module | Screw type
AQT90 FLEX | In de activity log is te zien dat beide foutmeldingen voorkomen. error 4587 heeft te mkaen met een time out van de inletdoors. Geen afwijkingen kunnen constateren. In het service menu diverse checkouts uitgevoerd, deze zijn allemaal goed. Incubator Wheel temp reg -->OK. CPU temp check -->OK. Inletdoors gecontroleerd -->OK. Na het verlaten van het servicemenu staat de AQT meetklaar, derhalve 2 QC's van RM gemeten --> OK. 2 fictieve patientmetingen uitgevoerd -->OK. De klant zelf 2 QC's van de klant laten meten --> OK. Met klant overlegd, en aangezien alle metingen nu goed gaan, zal de klant morgenvroeg weer opstarten en meten, en bij afwijkingen contact met RM opnemen. | In de activity log is te zien dat beide foutmeldingen voorkomen. error 4587 heeft te mkaen met een time out van de inletdoors. Geen afwijkingen kunnen constateren. In het service menu diverse checkouts uitgevoerd, deze zijn allemaal goed. Incubator Wheel temp reg -->OK. CPU temp check -->OK. Inletdoors gecontroleerd -->OK. Na het verlaten van het servicemenu staat de AQT meetklaar, derhalve 2 QC's van RM gemeten --> OK. 2 fictieve patientmetingen uitgevoerd -->OK. De klant zelf 2 QC's van de klant laten meten --> OK. Met klant overlegd, en aangezien alle metingen nu goed gaan, zal de klant morgenvroeg weer opstarten en meten, en bij afwijkingen contact met RM opnemen. | SWO-30229 | No | The activity log shows that both error messages occur. Error 4587 has to employ with a time out of the inletdoors. Cannot detect deviations. In the service menu various checkouts performed, these are all good. Incubator Wheel Temp Reg--> OK. CPU Temp Check--> OK. Inletdoors checked--> OK. After leaving the service menu, the AQT measurement is ready, therefore 2 Qcs of RM are measured--> OK. 2 fictitious Patientmeasurements performed--> OK. Let the customer measure 2 Qcs of the customer--> OK. With customer, and since all measurements are now going well, the customer will start and measure early tomorrow, and contact RM in case of deviations. | The activity log shows that both error messages occur. Error 4587 has to employ with a time out of the inletdoors. Cannot detect deviations. In the service menu various checkouts performed, these are all good. Incubator Wheel Temp Reg--> OK. CPU Temp Check--> OK. Inletdoors checked--> OK. After leaving the service menu, the AQT measurement is ready, therefore 2 Qcs of RM are measured--> OK. 2 fictitious Patientmeasurements performed--> OK. Let the customer measure 2 Qcs of the customer--> OK. With customer, and since all measurements are now going well, the customer will start and measure early tomorrow, and contact RM in case of deviations.
ABL80 FLEX CO-OX | Fehlersuche am Analysator, Kein akuter Flüssigkeitsaustritt festgestellt( mehrere Kal. QCs angestoßen)->vermutlich noch Rückstände eines in der Vergangenheit undichten LPs, Schlauchwege überprüft ok, Analysator gereinigt, Herrn Bär über zukünftige Vorgehensweise informiert System-und Funktionskontrolle i.O. | Fehlersuche am Analysator, Kein akuter Flüssigkeitsaustritt festgestellt( mehrere Kal. QCs angestoßen)->vermutlich noch Rückstände eines in der Vergangenheit undichten LPs, Schlauchwege überprüft ok, Analysator gereinigt, Herrn Bär über zukünftige Vorgehensweise informiert System-und Funktionskontrolle i.O. | SWO-30298 | No | Troubleshooting on the analyzer, No acute fluid leak age detected(several cal. QCs initiated)->probably still residues of a leaking LPs, hose ways checked ok, analyzer cleaned, Mr. Bär informed about future procedure system and function control i.O. | Troubleshooting on the analyzer, No acute fluid leak age detected(several cal. QCs initiated)->probably still residues of a leaking LPs, hose ways checked ok, analyzer cleaned, Mr. Bär informed about future procedure system and function control i.O.
ABL90 FLEX | CF Karte getauscht, Windows und AU neu installiert, in Roma ini div Eintragungen vorgenommen Barcode fixe Zielwerte, Kaltstartphase, Kundenkonfig aufgespielt, Kal und QC durchgeführt | CF Karte getauscht, Windows und AU neu installiert, in Roma ini div Eintragungen vorgenommen Barcode fixe Zielwerte, Kaltstartphase, Kundenkonfig aufgespielt, Kal und QC durchgeführt | SWO-30554 | No | CF card swapped, Windows and AU reinstalled, made in Roma ini div entries barcode fixed target values, cold start phase, customer confectionery played, Kal and QC performed | CF card swapped, Windows and AU reinstalled, made in Roma ini div entries barcode fixed target values, cold start phase, customer confectionery played, Kal and QC performed | Defect | HW | CF card
ABL90 FLEX | Lösungspack kontrolliert und erneut eingesetzt, Chipdaten können nicht gelesen werden, Lösungspack aus Technikerbestand eingebaut, Gerät Macht Kaltstart, pCO2 Probleme in der Kalibration, Empfindlichkeit außer Bereich, mehrere Kalibrationen durchgeführt, minimale Besserung aber außer Bereich, Sensorkassette aus Technikerbestand eingebaut, Kalibrationen und Interne Qualitätskontrollen ohne Beanstandungen durchgeführt. | Lösungspack kontrolliert und erneut eingesetzt, Chipdaten können nicht gelesen werden, Lösungspack aus Technikerbestand eingebaut, Gerät Macht Kaltstart, pCO2 Probleme in der Kalibration, Empfindlichkeit außer Bereich, mehrere Kalibrationen durchgeführt, minimale Besserung aber außer Bereich, Sensorkassette aus Technikerbestand eingebaut, Kalibrationen und Interne Qualitätskontrollen ohne Beanstandungen durchgeführt. | SWO-30307 | No | Solution pack controlled and re-inserted, chip data cannot be read, solution pack built-in from technician stock, device power cold start, pCO2 problems in calibration, sensitivity out range, multiple calibrations performed, minimal calibrations performed, minimal Improvement but out of range, sensor cassette installed from technician stock, calibrations and internal quality controls carried out without complaints. | Solution pack controlled and re-inserted, chip data cannot be read, solution pack built-in from technician stock, device power cold start, pCO2 problems in calibration, sensitivity out range, multiple calibrations performed, minimal calibrations performed, minimal Improvement but out of range, sensor cassette installed from technician stock, calibrations and internal quality controls carried out without complaints. | Assisted Replacement | SP | Chip error
ABL90 FLEX | Einlassdichtung und -Nadel überprüft. In Ordnung. Einzelne Fluidictransportprogramme im Servicemenü durchgeführt. Gerätestart in Ordnung. 3 Testmessungen durchgeführt. | Einlassdichtung und -Nadel überprüft. In Ordnung. Einzelne Fluidictransportprogramme im Servicemenü durchgeführt. Gerätestart in Ordnung. 3 Testmessungen durchgeführt. | SWO-29996 | No | inlet seal and needle. All right. Individual fluidic transport programs are carried out in the service menu. Device start in order. 3 test measurements carried out. | inlet seal and needle. All right. Individual fluidic transport programs are carried out in the service menu. Device start in order. 3 test measurements carried out. | Assisted Replacement | SP | Vacum
ABL815 FLEX | Defekten Druckerkopf ausgetauscht. Kal und QC ohne. Beanstandung | Defekten Druckerkopf ausgetauscht. Kal und QC ohne. Beanstandung | SWO-30037 | No | Replaced defective printer head. Kal and QC without. Complaint | Replaced defective printer head. Kal and QC without. Complaint | Defect | HW | Printer module | Printer Head
ABL825 FLEX | Lactate is failing calibrations. | Replaced the measuring chamber, the electrode connector, and both heat plates. | CEF7EB9D5E177B56 | No | Lactate is failing calibrations. | Replaced the measuring chamber, the electrode connector, and both heat plates. | Defect | HW | El/Met Module | Measuring chamber
ABL820 FLEX | Ran a control and confirmed problem. | Replaced the above parts. Also replaced the inlet gasket. | CEF826AA5E056CED | No | Ran a control and confirmed problem. | Replaced the above parts. Also replaced the inlet gasket. | Assisted Replacement | HW | Inlet module | Inlet Gasket
AQT90 FLEX |  |  | CEF82FF75E1830F6 | No
AQT90 FLEX | Analyzer opnieuw opgestart en getracht de fout te herhalen. niet gelukt. SD gemaakt voor verder onderzoek en SW opnieuw geladen (clean install). AQT werkt naar behoren en volgende dag extra controle. | geen correctie meer nodig. Dinsdag 26 mei nogmaals met klant eerste start in de ochtend gezamenlijk doen en dan case afsluiten. | SWO-30576 | No | Analyzer restarts and attempted to repeat the error. did not succeed. SD made for further investigation and SW Reloaded (clean install). AQT works properly and next day extra control. | No more correction needed. Tuesday May 26th again with customer first start in the morning do joint and then close case.
ABL80 CO-OX | NO BLOCKAGE FOUND | FLUSHED SENSOR PACK AND REPPLACED FLUIDIC PACK | CEF9A1FAD411E665 | No | NO BLOCKAGE FOUND | FLUSHED SENSOR PACK AND REPPLACED FLUIDIC PACK
ABL90 FLEX | LP Verriegelung im Service geprüft, kein Fehler feststellbar. LP installiert. Startphase 55 min. unterbrochen durch inhomogene Spüllösung. Analysator gespült, Startphase läuft ohne Probleme. pCO2 Empf. Fehler. | LP Verriegelung im Service geprüft, kein Fehler feststellbar. LP installiert. Startphase 55 min. unterbrochen durch inhomogene Spüllösung. Analysator gespült, Startphase läuft ohne Probleme. pCO2 Empf. Fehler. | SWO-30353 | No | LP locking checked in service, no error detectable. LP installed. Start-up phase 55 min. interrupted by inhomogeneous rinsing solution. Analyzer flushed, start-up phase runs without problems. pCO2 Mpf. Error. | LP locking checked in service, no error detectable. LP installed. Start-up phase 55 min. interrupted by inhomogeneous rinsing solution. Analyzer flushed, start-up phase runs without problems. pCO2 Mpf. Error. | No fault found | SP install
ABL90 FLEX | In service bij fluidic transport staat de waarde bij 'cloth detect' voor de OXI module te hoog, ook bij herhaling. Na vervanging van oxi-valve tubing en waste tubing opnieuw opgestart met veel beter waarden in het service programma (oxi-module waarde ligt op target). Na calibratie en 3 levels QC staat analyzer opnieuw meetklaar. | In service bij fluidic transport staat de waarde bij 'cloth detect' voor de OXI module te hoog, ook bij herhaling. Na vervanging van oxi-valve tubing en waste tubing opnieuw opgestart met veel beter waarden in het service programma (oxi-module waarde ligt op target). Na calibratie en 3 levels QC staat analyzer opnieuw meetklaar. | SWO-30236 | No | In service at Fluidic Transport, the value of ' cloth detect ' for the OXI module is too high, even when repeated. After replacing oxi-valve tubing and waste tubing rebooted with much better values in the service program (OXI-module value is on target). After calibration and 3 levels QC, Analyzer is again measuring ready. | In service at Fluidic Transport, the value of ' cloth detect ' for the OXI module is too high, even when repeated. After replacing oxi-valve tubing and waste tubing rebooted with much better values in the service program (OXI-module value is on target). After calibration and 3 levels QC, Analyzer is again measuring ready. | Defect | HW | Oxi module | Pinch Valve tube
ABL835 FLEX | Sustitucion de la fuente de alimentación. Limpieza de salida de aire y desperdicios.Cal sl. Cambio de membrana y oring de po2. Sustitucion de la puerta de entrada. Cal, qc y muestras | Sustitucion de la fuente de alimentación. Limpieza de salida de aire y desperdicios.Cal sl. Cambio de membrana y oring de po2. | CEF926EDD4446C53 | No | Power supply replacement. Cleaning of air outlet and waste. Cal sl. Membrane change and po2 oring. Replacement of the front door. Cal, qc and samples | Power supply replacement. Cleaning of air outlet and waste. Cal sl. Membrane change and po2 oring. | Defect | HW | Power Supply
ABL825 FLEX |  |  | CEF98206D448B15B | No |  |  | Adjustment | HW | Oxi module | Lamp unit
ABL827 FLEX | Reinigungsarbeiten, Fehlersuche, stark verschlissenen oder falschen Pumpenschlauch des Creamoduls durch neuen Modulpumpenschlauch ersetzt, mehrere Kal ok, QK ok | Reinigungsarbeiten, Fehlersuche, stark verschlissenen oder falschen Pumpenschlauch des Creamoduls durch neuen Modulpumpenschlauch ersetzt, mehrere Kal ok, QK ok | SWO-30252 | No | Cleaning work, troubleshooting, heavily worn or incorrect pump hose of the Creamodule replaced by new module pump hose, several cal ok, QK ok | Cleaning work, troubleshooting, heavily worn or incorrect pump hose of the Creamodule replaced by new module pump hose, several cal ok, QK ok | Defect | HW | Met II Module | Pump tubing
ABL825 FLEX | Probleme mit Bildschirm ( keine touchfunktion) Frau Kürzinger hat mich angerufen das der Bildschirm nicht funktioniert. Bildschirm und Verbindungskabel getauscht. Bei Neustart festgestellt dass CPU Unit nicht richtig funktioniert ( Bluescreen) Leihweise CPU Unit aus ABL 800 (754R1380N003) von GS eingebaut und Konfiguration von Kundengerät angepasst. CPU Unit aus Kundengerät 754R2005N004 zur Reparatur mitgenommen. Alten Bildschirm mit neuer CPU probiert. Leider immer noch keine Touch Funktion. Neuen Bildschirm wider angeschlossen. Touch Funktion wieder hergestellt. Mehrere Kalibrierungen durchgeführt. tHb Kalibration durchgeführt i.O. Kalverifizierung durchgeführt. --> i.O. Messung bei Glu QC instabil deswegen Glu neu überzogen. Nochmals mehrmals kalibriert. QCs für alle Level durchgeführt. | Probleme mit Bildschirm ( keine touchfunktion) Frau Kürzinger hat mich angerufen das der Bildschirm nicht funktioniert. Bildschirm und Verbindungskabel getauscht. Bei Neustart festgestellt dass CPU Unit nicht richtig funktioniert ( Bluescreen) Leihweise CPU Unit aus ABL 800 (754R1380N003) von GS eingebaut und Konfiguration von Kundengerät angepasst. CPU Unit aus Kundengerät 754R2005N004 zur Reparatur mitgenommen. Alten Bildschirm mit neuer CPU probiert. Leider immer noch keine Touch Funktion. Neuen Bildschirm wider angeschlossen. Touch Funktion wieder hergestellt. Mehrere Kalibrierungen durchgeführt. tHb Kalibration durchgeführt i.O. Kalverifizierung durchgeführt. --> i.O. Messung bei Glu QC instabil deswegen Glu neu überzogen. Nochmals mehrmals kalibriert. QCs für alle Level durchgeführt. | SWO-30304 | No | Problems with screen (no touch function) Mrs. Kürzinger called me that the screen does not work. screen and connection cable. When restarted, it was determined that CPU Unit was not working properly (Bluescreen) Loaned CPU Unit from ABL 800 (754R1380N003) built by GS and customized configuration of customer device. CPU unit from customer device 754R2005N004 for repair. Old screen with new CPU tried. Unfortunately still no touch function. New screen connected. Touch function restored. Multiple calibrations performed. tHb calibration carried out in the i.O. kalverification. --> i.O. measurement in Glu QC unstable therefore Glu re-coated. Calibrated again several times. QCs for all levels. | Problems with screen (no touch function) Mrs. Kürzinger called me that the screen does not work. screen and connection cable. When restarted, it was determined that CPU Unit was not working properly (Bluescreen) Loaned CPU Unit from ABL 800 (754R1380N003) built by GS and customized configuration of customer device. CPU unit from customer device 754R2005N004 for repair. Old screen with new CPU tried. Unfortunately still no touch function. New screen connected. Touch function restored. Multiple calibrations performed. tHb calibration carried out in the i.O. kalverification. --> i.O. measurement in Glu QC unstable therefore Glu re-coated. Calibrated again several times. QCs for all levels. | Defect | HW | DMS_PC | Touch screen
AQT90 FLEX | In de instellingen een paar wijzigingen aangebracht. Alle printeropdrachten uitgeschakeld. Volgende dag wederom gezamenlijk opstarten met klant voor controle. | geen correctie meer nodig. Dinsdag 26 mei nogmaals met klant eerste start in de ochtend gezamenlijk doen en dan case afsluiten. | SWO-30579 | No | Made a few changes In the settings. All printer commands are disabled. Next day again jointly start up with customer for control. | No more correction needed. Tuesday May 26th again with customer first start in the morning do joint and then close case.
AQT90 FLEX |  |  | CEF9125DD40DFDCA | No
ABL80 CO-OX | Confirmed findings | Replaced Hemolyzer Assembly Performed Zero Cal Performed tHb calibration | CEFA5EC14A63D2AE | No | Confirmed findings | Replaced Hemolyzer Assembly Performed Zero Cal Performed tHb calibration
ABL80 OSM | found that solution pump rotor was not on motor all the way | reseated pump rotor head to make a better connection to motor. | CEFA73094A5CB5C6 | No | found that solution pump rotor was not on motor all the way | reseated pump rotor head to make a better connection to motor.
ABL80 OSM | found that the solution tube has come off its connection to the fluidic block | reconnect tube to fluidic block which has corrected the issue. | CEFA712C4A49B5C2 | No | found that the solution tube has come off its connection to the fluidic block | reconnect tube to fluidic block which has corrected the issue.
ABL90 FLEX | Scanner gereinigt | Scanner gereinigt | SWO-30341 | No | Scanner cleaned | Scanner cleaned | Adjustment | Cleaned | Barcode Reader
ABL90 |  |  | CEFA5A364A3AF8EF | No |  |  | Defect | HW | Hemolyzer | & install SW
ABL90 FLEX | Nettoyage de la section fluidique e.o Test à l'aide du programme de service e.o Remplacé la k7 de mesure e.o Service dump e.o Calibration e.o contrôle de qualité e.o | Nettoyage de la section fluidique e.o Test à l'aide du programme de service e.o Remplacé la k7 de mesure e.o Service dump e.o Calibration e.o contrôle de qualité e.o | SWO-30503 | No | Cleaning the fluidsection e.o Test using the service program e.o Replaced the measurement k7 e.o Service dump e.o Calibration e.o quality control e.o | Cleaning the fluidsection e.o Test using the service program e.o Replaced the measurement k7 e.o Service dump e.o Calibration e.o quality control e.o | Cleaned | Fluidic pathway
ABL90 FLEX | roma.ini überprüft u. angepaßt | roma.ini überprüft u. angepaßt | SWO-30385 | No | roma.ini checked and adapted | roma.ini checked and adapted | Setup | roma.ini
ABL90 | Unable to pump solution. Analyzer frozen. | Rebooted analyzer. Replaced Solution Pack and Sensor cassette. Completed all service tests. | CEFAD5AF4A106941 | No | Unable to pump solution. Analyzer frozen. | Rebooted analyzer. Replaced Solution Pack and Sensor cassette. Completed all service tests. | Assisted Replacement | SP | FT
ABL90 with Auto Inlet | Erneuerung Einlassnadel und Dichtung - Analysator gespült - Serviceprogramm Justagen und Kalibration etc. - Funktionsüberprüfung etc. Analysator zeigt pO2 KAL Fehler und springt in Intervention erforderlich (Spülung !) - Neustart ohne Erfolg - Serviceproramm bei Clot Detection Status Fehler vorhanden - pO2 Überprüfung (Service-KAL) ohne Beanstandung - Fluid Transport Schritte wiederholt ohne Beanstandung - Erneute Installation Sensorkassette (LOT R0321N021) notwendig - Installation neuer Sensorkassette notwendig (LOT R0328N192) - Funktionsüberprüfung ohne Beanstandung ! | Erneuerung Einlassnadel und Dichtung - Analysator gespült - Serviceprogramm Justagen und Kalibration etc. - Funktionsüberprüfung etc. Analysator zeigt pO2 KAL Fehler und springt in Intervention erforderlich (Spülung !) - Neustart ohne Erfolg - Serviceproramm bei Clot Detection Status Fehler vorhanden - pO2 Überprüfung (Service-KAL) ohne Beanstandung - Fluid Transport Schritte wiederholt ohne Beanstandung - Erneute Installation Sensorkassette (LOT R0321N021) notwendig - Installation neuer Sensorkassette notwendig (LOT R0328N192) - Funktionsüberprüfung ohne Beanstandung ! | SWO-30468 | No | Renewal inlet needle and seal - analyzer rinsed - service program adjustment and calibration etc. - function check etc. Analyzer shows pO2 KAL error and jumps in intervention required (flushing !) - Restart without success - Serviceproramm at Clot Detection Status error present - pO2 review (service-KAL) without complaint - Fluid transport steps repeated without complaint - Reinstallation sensor cassette (LOT R0321N021) necessary - installation of new sensor cassette necessary (LOT R0328N192) - Functional check without complaint ! | Renewal inlet needle and seal - analyzer rinsed - service program adjustment and calibration etc. - function check etc. Analyzer shows pO2 KAL error and jumps in intervention required (flushing !) - Restart without success - Serviceproramm at Clot Detection Status error present - pO2 review (service-KAL) without complaint - Fluid transport steps repeated without complaint - Reinstallation sensor cassette (LOT R0321N021) necessary - installation of new sensor cassette necessary (LOT R0328N192) - Functional check without complaint ! | Replaced | Probe | Bend
ABL90 | Comprobación del estado de funcionamiento del equipo, realización de kit anual de tubos, inspección de todos los mecanos internos del equipo, se desmonta totalmente el módulo de sujección del SP, la bahía del SC, se lubrica el motor actuator del SP, ajuste de precisión de todos estos mecanos, limpieza del módulo Oxi, comprobación de registros, cal de la tHb, verificación final con QC y muestra de paciente. |  | CEFA71CE4A0F910E | No | Checking the operating status of the equipment, performing annual pipe kit, inspecting all internal equipment mechanists, fully disassembled the SP clamping module, SC bay, lubricated SP actuator motor, precision adjustment of all these mecans, Cleaning the Oxi module, checking records, tHb lime, final verification with QC and patient sample. |  | Defect | HW | waste tube
ABL835 FLEX | Proteinentfernung durchgeführt - Messkammer gereinigt pH Elektrode erneuert Gerät kalibriert und überprüft | Proteinentfernung durchgeführt - Messkammer gereinigt pH Elektrode erneuert Gerät kalibriert und überprüft | SWO-30444 | No | Protein removal performed - measuring chamber cleaned pH electrode renewed device calibrated and checked | Protein removal performed - measuring chamber cleaned pH electrode renewed device calibrated and checked | Assisted Replacement | Electrode | Electrode | pH
ABL805 FLEX | PCO2 Elektrode ersetzt. | PCO2 Elektrode ersetzt. | SWO-30736 | No | PCO2 electrode replaced. | PCO2 electrode replaced. | Assisted Replacement | Electrode | Electrode | pCO2
ABL825 FLEX | Unable to run controls Ampoule breaker is not been activated. Annalyzer running without any errors. | Found the breaker activator disconnected from the motor. Installed the breaker back in place Secured the connector in place. | CEFA8AF64A156F1D | No | Unable to run controls Ampoule breaker is not been activated. Annalyzer running without any errors. | Found the breaker activator disconnected from the motor. Installed the breaker back in place Secured the connector in place. | Defect | HW | AutoCheck Module | Breaker arm
ABL825 FLEX | The analyzer is failing calibrations due to message 251 error code 'Halogen lamp intensity too high. | Checked the hemolyzer lamp intensity and it was at 3905, high. Adjusted the hemolyzer lamp, performed a complete oximetry test, passing within ranges. Performed the service and customer mode ctHb calibration. Flushed system lines. | CEFA82F64A5181CD | No | The analyzer is failing calibrations due to message 251 error code 'Halogen lamp intensity too high. | Checked the hemolyzer lamp intensity and it was at 3905, high. Adjusted the hemolyzer lamp, performed a complete oximetry test, passing within ranges. Performed the service and customer mode ctHb calibration. Flushed system lines. | Adjustment | HW | Oxi module | Lamp unit
ABL830 FLEX | QC-Modul im serv.Programm getestet, PCB zum testen getauscht, Funktionen überprüft, FPC cable ersetzt, | QC-Modul im serv.Programm getestet, PCB zum testen getauscht, Funktionen überprüft, FPC cable ersetzt, | SWO-30397 | No | QC module in serv. program tested, PCB exchanged for testing, functions checked, FPC cable replaced, | QC module in serv. program tested, PCB exchanged for testing, functions checked, FPC cable replaced, | Defect | HW | AutoCheck Module | FPC cable
ABL835 FLEX | Bei Versuch Altgerät anzubinden (754R0142N010) friert Bildschirm bei Anschluss Netzwerkstecker ein - Analysator automatischer Neustart - Bildschirm noch eingefroren, bis Tastatur eingesteckt, dann Touchscreen wieder Funktionsbereit ! EDV (Hr. Truch) lässt Netzwerkdose überprüfen und neuen Switch Port einrichten ! Kunde meldet sich und berichtet über weitere Entwicklung ! Rücksprache Medizintechnik (Hr. Stein) | Bei Versuch Altgerät anzubinden (754R0142N010) friert Bildschirm bei Anschluss Netzwerkstecker ein - Analysator automatischer Neustart - Bildschirm noch eingefroren, bis Tastatur eingesteckt, dann Touchscreen wieder Funktionsbereit ! EDV (Hr. Truch) lässt Netzwerkdose überprüfen und neuen Switch Port einrichten ! Kunde meldet sich und berichtet über weitere Entwicklung ! Rücksprache Medizintechnik (Hr. Stein) | SWO-30369 | No | When trying to connect old device (754R0142N010) screen freezes when connected network plug - analyzer automatic restart - screen still frozen until keyboard plugged in, then touchscreen ready for use again ! EDP (Mr. Truch) allows to check network socket and set up new switch port ! Customer reports and reports on further development ! Consultation of medical technology (Mr. Stein) | When trying to connect old device (754R0142N010) screen freezes when connected network plug - analyzer automatic restart - screen still frozen until keyboard plugged in, then touchscreen ready for use again ! EDP (Mr. Truch) allows to check network socket and set up new switch port ! Customer reports and reports on further development ! Consultation of medical technology (Mr. Stein) | Unknown | HW | DMS_PC | Touch screen
AQT90 FLEX |  |  | CEFA9BF04A096632 | No
AQT90 FLEX | Ophalen van instellingen van locatie Blakeburg, om dan in AQT Nieuwe Laan te plaatsen. Dit gedaan einde ochtend en AQT herkent barcode veel gemakkelijker dan voorheen. Ook geen foutmelding waarneembaar. Resultaat komt door in Cobas IT. Morgen weer bezoek om samen met klant op te starten van eerste testen. | geen correctie meer nodig. Dinsdag 26 mei nogmaals met klant eerste start in de ochtend gezamenlijk doen en dan case afsluiten. | SWO-30582 | No | Retrieving location settings from Bladenburg, to place in AQT New Avenue. This done end of morning and AQT recognizes barcode much easier than before. Also no error message perceptible. Result is in Cobas IT. Visit tomorrow again to start the first tests together with the customer. | No more correction needed. Tuesday May 26th again with customer first start in the morning do joint and then close case.
TCM-COMBIM | Installed an electrode module, but this did not resolve the problem. Removed the electrode and replaced it with one from another unit, and the problem followed the electrode. | Placed and electrode on order for overnight delivery. Customer will return the faulty electrode to Radiometer after the new electrode is received. | CEFAA7B84A0B1C24 | No | Installed an electrode module, but this did not resolve the problem. Removed the electrode and replaced it with one from another unit, and the problem followed the electrode. | Placed and electrode on order for overnight delivery. Customer will return the faulty electrode to Radiometer after the new electrode is received.
TCM-COMBIM | Troubleshot to a faulty electrode | Placed an electrode on order for overnight delivery | CEFAAF604A211C27 | No | Troubleshot to a faulty electrode | Placed an electrode on order for overnight delivery
ABL90 FLEX |  |  | SWO-30544 | No |  |  | Defect | SW | -
ABL90 FLEX | Gerät gespült, Lösungspack geliefert und ersetzt, test im Service nun erfolgreich, Neustart, Kalibration, QC- Kontrolle | Gerät gespült, Lösungspack geliefert und ersetzt, test im Service nun erfolgreich, Neustart, Kalibration, QC- Kontrolle | SWO-30676 | No | Device rinsed, solution pack delivered and replaced, test in service now successful, restart, calibration, QC control | Device rinsed, solution pack delivered and replaced, test in service now successful, restart, calibration, QC control | Assisted Replacement | SP | QC2
ABL90 FLEX | Maßnahmen durchgeführt, Gerät mehrfach gespült, Hämolysator überprüft --> i.O., Verunreinigung beseitigt, Gerät i.O., Sensorkassete vom Kunden ausgetauscht | Die Chipdaten des SP werden nicht gelesen. Frau Siegel hat bereits ein weiteres SP mit der gleichen Fehlermeldung installiert. Die Kontakte zum SP wurden gereinigt. Der Fehler bleibt. Frau Siegel wünscht einen Service vor Ort. | SWO-30784 | No | Measures carried out, device rinsed several times, hemolyser checked --> i.O., contamination eliminated, device i.O., sensor cash exchanged by the customer | The chip data of the SP is not read. Ms. Siegel has already installed another SP with the same error message. The contacts to the SP have been cleaned. The error remains. Ms. Siegel wants a service on site. | Defect | HW | SP PCB
ABL90 FLEX | CF Card ersetzt, Software installiert, allg. Sicht- und Funktionskontrolle, Konfiguration rekonstruiert, kalibration, QC Kontrolle durch Anwender und oder Remote | CF Card ersetzt, Software installiert, allg. Sicht- und Funktionskontrolle, Konfiguration rekonstruiert, kalibration, QC Kontrolle durch Anwender und oder Remote | SWO-30675 | No | CF Card replaces, software installed, general visual and function control, configuration reconstructed, calibration, QC control by user and or remote | CF Card replaces, software installed, general visual and function control, configuration reconstructed, calibration, QC control by user and or remote | Defect | HW | CF card
ABL90 FLEX | Jahreswartungs Satz verbaut. Mehrere Messungen durchgeführt. Test i.O. | Jahreswartungs Satz verbaut. Mehrere Messungen durchgeführt. Test i.O. | SWO-30717 | No | Annual maintenance set installed. Multiple measurements performed. Test i.O. | Annual maintenance set installed. Multiple measurements performed. Test i.O. | Defect | HW | Inlet Gasket
ABL90 FLEX | Analysator gespült, div. Maßnahmen und Test im Service durchgeführt. Fehler kommt immer wieder. Einlassdichtung erneuert, Fehler behobenen. | Analysator gespült, div. Maßnahmen und Test im Service durchgeführt. Fehler kommt immer wieder. Einlassdichtung erneuert, Fehler behobenen. | SWO-30716 | No | Analyzer rinsed, div. measures and test carried out in the service. Mistakes keep coming. Inlet seal renewed, errors fixed. | Analyzer rinsed, div. measures and test carried out in the service. Mistakes keep coming. Inlet seal renewed, errors fixed. | Defect | HW | Inlet Gasket
ABL90 FLEX | 2. Gerät ABL 90 288 003 aufgestellt zur direkten Vergleichsmessung von Ca und Lac . -> 3 Vergleichsmessungen mit Patientenblut durchgeführt, Keine Abweichungen festgestellt. Weitere Vergleichsmessungen werden vom Stationspersonal durchgeführt und später verglichen und analysiert. | 2. Gerät ABL 90 288 003 aufgestellt zur direkten Vergleichsmessung von Ca und Lac . -> 3 Vergleichsmessungen mit Patientenblut durchgeführt, Keine Abweichungen festgestellt. Weitere Vergleichsmessungen werden vom Stationspersonal durchgeführt und später verglichen und analysiert. | SWO-30619 | No | 2. Device ABL 90 288 003 set up for direct comparison measurement of Ca and Lac . -> 3 comparative measurements performed with patient blood, no deviations detected. Further comparative measurements are carried out by the station staff and later compared and analysed. | 2. Device ABL 90 288 003 set up for direct comparison measurement of Ca and Lac . -> 3 comparative measurements performed with patient blood, no deviations detected. Further comparative measurements are carried out by the station staff and later compared and analysed. | No fault found | Ca & Lac results OK?
ABL90 FLEX | SC. Ersetzt. | SC. Ersetzt. | SWO-30735 | No | SC. Replaced. | SC. Replaced. | Assisted Replacement | SC | Sealing | Gasket
ABL90 | Reviewed logs and confirmed problem. | A level 1 of S7730 QC was due to be run so the analyzer was locked. I ran the level 1 of external QC to unlock the analyzer then performed the tHb calibration. I then ran the level B and the other two levels of external QC. | CEFC0A86C027297F | No | Reviewed logs and confirmed problem. | A level 1 of S7730 QC was due to be run so the analyzer was locked. I ran the level 1 of external QC to unlock the analyzer then performed the tHb calibration. I then ran the level B and the other two levels of external QC. | Adjustment | tHb cal | sO2 fail QC
ABL825 FLEX | Inspected fuses and all looked good. | I replaced the power supply. | CEFE1D1BAC20A2D0 | No | Inspected fuses and all looked good. | I replaced the power supply. | Unknown | HW | Power Supply
ABL825 FLEX | The Cal 1 bottle port was loose causing bubbles in the cal solution. The pH lower sealing ring was pushed up into the electrode jacket causing fluid to fill in the pH well. The LAC electrode lower seal was folded under the electrode glass bottom causing a protruding section of the LAC membrane. Fluid is leaking out of the CA well drain spout due to the white Teflon is pushed up into the electrode jacket. The PO2 membrane looks wrinkled. | The pH lower seal ring, LAC lower seal were replaced. The PO2, LAC & CA membranes were replaced. | CEFBEEB4C0336351 | No | The Cal 1 bottle port was loose causing bubbles in the cal solution. The pH lower sealing ring was pushed up into the electrode jacket causing fluid to fill in the pH well. The LAC electrode lower seal was folded under the electrode glass bottom causing a protruding section of the LAC membrane. Fluid is leaking out of the CA well drain spout due to the white Teflon is pushed up into the electrode jacket. The PO2 membrane looks wrinkled. | The pH lower seal ring, LAC lower seal were replaced. The PO2, LAC & CA membranes were replaced. | Assisted Replacement | Electrode | Membrane | Lac
AQT90 FLEX | geen correctie meer nodig. Dinsdag 26 mei nogmaals met klant eerste start in de ochtend gezamenlijk doen en dan case afsluiten. | geen correctie meer nodig. Dinsdag 26 mei nogmaals met klant eerste start in de ochtend gezamenlijk doen en dan case afsluiten. | SWO-30583 | No | No more correction needed. Tuesday May 26th again with customer first start in the morning do joint and then close case. | No more correction needed. Tuesday May 26th again with customer first start in the morning do joint and then close case.
TCM-COMBIM |  |  | CEFBCC39C0034458 | No |  |  | Defect | HW | Touch Screen | Touch Screen
ABL90 | Verify issue. Capture screen shot of error message. Also saved a service dump for a Techline case. | Back up all data. Reload Kira. Save DB, Roma.ini and Communication settings. Re-install software. After rebooting, Calibrations were failing because of Na issues. Tried many things. Also couldn't aspirate samples. possible air in sample.Perfomed backflush 3X times. Didn't help. Replaced parts from PM kit and installed a new probe. Replaced SC. PCO2 won't calibrate. Sensitivity Now getting inhomo solutions. Replaced SP. Still bad. Replaced another SC. Will return on Saturday to finish up. Return on Saturday. Calibrations were fine when arrived. Ran all 3 levels of external QC without errors. | CEFCC215364BFF1C | No | Verify issue. Capture screen shot of error message. Also saved a service dump for a Techline case. | Back up all data. Reload Kira. Save DB, Roma.ini and Communication settings. Re-install software. After rebooting, Calibrations were failing because of Na issues. Tried many things. Also couldn't aspirate samples. possible air in sample. Perfomed backflush 3X times. Didn't help. Replaced parts from PM kit and installed a new probe. Replaced SC. PCO2 won't calibrate. Sensitivity Now getting inhomo solutions. Replaced SP. Still bad. Replaced another SC. Will return on Saturday to finish up. Return on Saturday. Calibrations were fine when arrived. Ran all 3 levels of external QC without errors. | Defect | SW | shuts down
ABL825 FLEX | AU board needs to be replaced. | I replaced the AU board. | CEFE1DF2AC2EA2D3 | No | AU board needs to be replaced. | I replaced the AU board. | Defect | HW | Analyzer Control PCB
ABL825 FLEX | AU board needs replaced. | I replaced the AU board. | CEFF42222210EB41 | No | AU board needs replaced. | I replaced the AU board. | Defect | HW | Analyzer Control PCB
ABL825 FLEX |  |  | CEFFC46D2221C983 | No |  |  | Assisted Replacement | Electrode | Membrane | Ca++
ABL90 |  |  | CF008B2898270302 | No |  |  | Replaced | probe | Bend
ABL90 | LOOKED IN DATA LOGS AND AND FOUND THAT INLET CLEANING WAS DUE ON MAY 22, 2015 WHICH WAS NEVER DONE UNTIL I DID IT TODAY. ALSO FOUND THAT SENSOR CASSETTE WOULD EXPIRE TONIGHT AT 7:00PM. | CLEANED INLET GASKET AND REPLACED SENSOR CASSETTE WITH CUSTOMERS STOCK. DID 5 WASTE BLOODS. | CF00D03F9843513B | Yes | LOOKED IN DATA LOGS AND AND FOUND THAT INLET CLEANING WAS DUE ON MAY 22, 2015 WHICH WAS NEVER DONE UNTIL I DID IT TODAY. ALSO FOUND THAT SENSOR CASSETTE WOULD EXPIRE TONIGHT AT 7:00PM. | CLEANED INLET GASKET AND REPLACED SENSOR CASSETTE WITH CUSTOMERS STOCK. DID 5 WASTE BLOODS. | Re classified RMED | Courtesy
ABL90 | FOUND ANALYZER IN PUMP ROOM NOT PLUGGED IN AND ANALYZER OFF. PLUGGED ANALYZER IN AND FOUND THAT SOLUTION PACK HAD ZERO SOLUTION LEFT. | PLUGGED ANALYZER IN AND CHANGED SOLUTION PACK FROM CUSTOMERS STOCK. | CF00C402981F0971 | Yes | FOUND ANALYZER IN PUMP ROOM NOT PLUGGED IN AND ANALYZER OFF. PLUGGED ANALYZER IN AND FOUND THAT SOLUTION PACK HAD ZERO SOLUTION LEFT. | PLUGGED ANALYZER IN AND CHANGED SOLUTION PACK FROM CUSTOMERS STOCK. | Re classified RMED | Courtesy
ABL90 FLEX | div. Möglichkeiten der Präanalytik mit Schwester Susanne besprochen, 1 Packung Pico50 zu Vergleichsmessungen gebracht | div. Möglichkeiten der Präanalytik mit Schwester Susanne besprochen; 1 Packung Pico50 zu Vergleichsmessungen gebracht | SWO-30871 | No | div. Possibilities of preanalysis discussed with Sister Susanne; 1 pack of Pico50 brought to comparative measurements | div. Possibilities of preanalysis discussed with Sister Susanne; 1 pack of Pico50 brought to comparative measurements | Unknown | Pre analytical | Pico5
ABL825 FLEX | Reparierte CPU Unit wieder eingebaut. Mehrere Kalibrierungen durchgeführt. Complete Oximetrietest durchgeführt --> i.O. tHb kal durchgeführt -->i.O Fcuv.:0,9729 Kalverifizierung durchgeführt --> i.O. QCs für alle Level durchgeführt. -> i.O.w ABL wieder Messbereit. | Reparierte CPU Unit wieder eingebaut. Mehrere Kalibrierungen durchgeführt. Complete Oximetrietest durchgeführt --> i.O. tHb kal durchgeführt -->i.O Fcuv.:0,9729 Kalverifizierung durchgeführt --> i.O. QCs für alle Level durchgeführt. -> i.O.w ABL wieder Messbereit. | SWO-30888 | No | Repaired CPU unit reinstalled. Multiple calibrations performed. Complete oximetry test performed --> i.O. tHb kal performed -->i.O Fcuv.:0,9729 Calverification performed --> i.O. QCs performed for all levels. -> i.O.w ABL ready to measure again. | Repaired CPU unit reinstalled. Multiple calibrations performed. Complete oximetry test performed --> i.O. tHb kal performed -->i.O Fcuv.:0,9729 Calverification performed --> i.O. QCs performed for all levels. -> i.O.w ABL ready to measure again. | Defect | HW | DMS_PC | Touch screen
AQT90 FLEX |  |  | CF00E0A99820583D | No
TCM-COMBIM |  |  | CF034F2C840E9EEA | No
ABL80 FLEX CO-OX | Lösungspack und Sensorkassette geliefert und ersetzt | Lösungspack und Sensorkassette geliefert und ersetzt | SWO-31058 | No | Solution pack and sensor cartridge supplied and replaced | Solution pack and sensor cartridge supplied and replaced
ABL90 FLEX | Touchscreen erneuert. | Touchscreen erneuert. | SWO-30993 | No | Touchscreen renewed. | Touchscreen renewed. | Defect | HW | Touch screen
ABL90 FLEX | Schläuche manuell mehrfach gespült, Gerinnsel entfernt, Kaltstartphase, Kal und QC durchgeführt | Schläuche manuell mehrfach gespült, Gerinnsel entfernt, Kaltstartphase, Kal und QC durchgeführt | SWO-31144 | No | Hoses manually rinsed several times, clots removed, cold start phase, Kal and QC performed | Hoses manually rinsed several times, clots removed, cold start phase, Kal and QC performed | Cleaned | Fluidic pathway | Clot
ABL90 FLEX | Fehlersuche Probenweg gespült, Einlassdichtung gereinigt Startphase | Fehlersuche Probenweg gespült; Einlassdichtung gereinigt Startphase | SWO-30931 | No | Troubleshooting sample path flushed; Inlet seal cleaned start phase | Troubleshooting sample path flushed; Inlet seal cleaned start phase | Cleaned | Fluidic pathway
ABL90 FLEX | Alle Schläuche und Dichtungen ersetzt, Reinigung Spülung, Lösungspack und Sensorkassette ersetzt, Kalibration, QC-Kontrolle | Alle Schläuche und Dichtungen ersetzt, Reinigung Spülung, Lösungspack und Sensorkassette ersetzt, Kalibration, QC-Kontrolle | SWO-31057 | No | Replaced all hoses and seals, cleaning rinse, solution pack and sensor cartridge, calibration, QC control | Replaced all hoses and seals, cleaning rinse, solution pack and sensor cartridge, calibration, QC control | Assisted Replacement | SC | Wetability
ABL815 FLEX | Gaseinheit und Messsystem geprüft. pCO2 Membran erneuert. QCs unwesentlich besser aber im Bereich. | Gaseinheit und Messsystem geprüft. pCO2 Membran erneuert. QCs unwesentlich besser aber im Bereich. | SWO-31349 | No | gas unit and measuring system. pCO2 membrane. QCs are marginally better but in the range. | gas unit and measuring system. pCO2 membrane. QCs are marginally better but in the range. | Assisted Replacement | Electrode | Membrane | pCO2
ABL835 FLEX |  |  | CF0321AC84331F53 | No |  |  | Assisted Replacement | Electrode | Membrane | pO2
ABL835 FLEX | Sampler Detektor PCB ausgetauscht, keine Veränderung, neues FlexQ Modul eingesetzt, Pos. 1-2 werden nicht erkannt, Modul zurück getauscht, Sampler Sensor Kalibrierung durchgeführt, Gerät i.O. | Sampler Detektor PCB ausgetauscht, keine Veränderung, neues FlexQ Modul eingesetzt, Pos. 1-2 werden nicht erkannt, Modul zurück getauscht, Sampler Sensor Kalibrierung durchgeführt, Gerät i.O. | SWO-31129 | No | Sampler detector PCB exchanged, no change, new FlexQ module inserted, Pos. 1-2 are not detected, module is replaced, sampler sensor calibration performed, device i.O. | Sampler detector PCB exchanged, no change, new FlexQ module inserted, Pos. 1-2 are not detected, module is replaced, sampler sensor calibration performed, device i.O. | Defect | HW | FlexQ Module | Position detector PCB
ABL837 FLEX |  |  | CF03CCDD840CC665 | No |  |  | Cleaned | Fluidic pathway
ABL90 | I inspected the casette connections, connections for the solution pak, checked for leaks. | I performed a flush, calibrations in the service mode replaced the inlet gasket. | CF035F8B8420FF86 | No | I inspected the casette connections, connections for the solution pak, checked for leaks. | I performed a flush, calibrations in the service mode replaced the inlet gasket. | Cleaned | Fluidic pathway
ABL90 | Customer complained the time changed when they moved the analyzer I also found the battery not active. | I verified that Windows setting were correct and made proper modification Ordered a new battery. Allowed analyzer to go through the 50 minutes startup time. | CF032EE584197B3C | No | Customer complained the time changed when they moved the analyzer I also found the battery not active. | I verified that Windows setting were correct and made proper modification Ordered a new battery. Allowed analyzer to go through the 50 minutes startup time. | Defect | SW | Time setting
ABL90 | I checked the log but could not find anything | I did pull all the files and ran QC | CF033DB384378C20 | No | I checked the log but could not find anything | I did pull all the files and ran QC | No fault found | -
ABL835 FLEX | Sustitución de la membrana de lactato y del electrodo, realización de varios cambios del analizador explicación de uso del mismo, calibración y test controles de calidad, funcionando correctamente. | Cambio electrodo de LACT. | CF03465C840AAB73 | No | Replacement of the lactate membrane and electrode, making several changes of the analyzer explanation of use of it, calibration and test quality controls, operating properly. | LACT electrode change. | Assisted Replacement | Electrode | Electrode | Lac
ABL825 FLEX |  |  | CF032EB584010399 | No |  |  | Defect | HW | Oxi module | Hemolyzer
ABL820 FLEX | Customer has been having to remembrane often to get QC to come in.They had cleaned chambers contacts.I checked Gases Chamber Electrode Valves contacts. | Re-membraned pCO2 electrode with different lot of membranes. | CF0357FD84066425 | No | Customer has been having to remembrane often to get QC to come in. They had cleaned chambers contacts. I checked Gases Chamber Electrode Valves contacts. | Re-membraned pCO2 electrode with different lot of membranes. | Assisted Replacement | Electrode | Membrane | pCO2
ABL825 FLEX | Einlassnadel manuell vorgeschoben und Einlass im Service kalibriert, Fehler behoben. Net Op. neu gestartet. Fehler bleib bestehen, no response from 172.17.68.85. Alle NetOp Fenster schließen lassen. Alle Einstellungen überprüft und neu eingestellt, Fehler bleibt bestehenden. | Einlassnadel manuell vorgeschoben und Einlass im Service kalibriert, Fehler behoben. Net Op. neu gestartet. Fehler bleib bestehen, no response from 172.17.68.85. Alle NetOp Fenster schließen lassen. Alle Einstellungen überprüft und neu eingestellt, Fehler bleibt bestehenden. | SWO-31345 | No | Inlet pin manually pushed forward and inlet calibrated in the service, error fixed. Net Op. restarted. Errors remain, no response from 172.17.68.85. Let all NetOp windows close. All settings checked and re-set, error remains existing. | Inlet pin manually pushed forward and inlet calibrated in the service, error fixed. Net Op. restarted. Errors remain, no response from 172.17.68.85. Let all NetOp windows close. All settings checked and re-set, error remains existing. | Unknown | HW | Inlet module
ABL835 FLEX | OXI Modul 902-833 R1812N001 ersetzt durch 902-833 R1164N003. OXI complete test, Service tHb Kal. und tHB Test i.O. | OXI Modul 902-833 R1812N001 ersetzt durch 902-833 R1164N003. OXI complete test, Service tHb Kal. und tHB Test i.O. | SWO-31348 | No | OXI Module 902-833 R1812N001 replaced by 902-833 R1164N003. OXI complete test, service tHb Kal. and tHB Test in the o. | OXI Module 902-833 R1812N001 replaced by 902-833 R1164N003. OXI complete test, service tHb Kal. and tHB Test in the o. | Defect | HW | Oxi module | Hemolyzer
ABL835 FLEX | Analysator überprüft, Messungen beobachtet und mit verschiedenen Benutzern gesprochen (derzeit kein Fehler feststellbar) - Konfiguration neu aufgesetzt (A12 angepasst) - Analysator mit EDV (Hr. Truch) angebunden etc. - Funktionsüberprüfung etc. | Evt. statische Aufladung möglich - Standortwechsel Analysator - Radiance Anbindung und Funktionsüberprüfung etc. | SWO-31206 | No | Analyzer checked, measurements observed and spoken to different users (currently no error detectable) - configuration re-set (A12 adjusted) - Analyzer connected with computer (Mr. Truch) etc. - Function check etc. | Evt. Static charging possible - location change analyzer - Radiance connection and function verification etc. | Defect | HW | DMS_PC | Touch screen
AQT90 FLEX | Ich wurde von Herrn Radloff darüber informiert das der AQT im Fehler steht. Bei Eintreffen beim Kunden festgestellt das mehrere QCs nicht i.O. waren und eine TnI Kalibrierung gescheitert war. Es wurde der Fehler ( 1035 Prüfung maximale dunkelstrommessung gescheitert) angezeigt. Dieser Fehler war laut Archiv schon länger vorhanden. Optische Einheit nach Manual getauscht. Systemüberprüfung durchgeführt. Mehrere Systemreinigungen durchgeführt.--> i.O. Es konnten keine Parameter kalibriert werden da keine Kalibrationskassette der Parameter vorhanden waren und leider auch keine kal Kassette nachbestellt werden konnte. vorhandene Kalibrierungen entfernt da diese auf die alte optische Einheit programmiert waren, und nun nicht mehr gültig sind. Vorhandenes Material vom Kunden mitgenommen und dem Kunden gutgeschrieben. Kunden von jedem genutzten Parameter ( TnI, NTproBNP, D-Dimmer) einen neue Schachtel via e-Order bestellt. Diese pastellnen Testkits werden voraussichtlich am nächsten Tag (29.05.15) vor 10 Uhr geliefert. Kunde kalibriert die Parameter selbstständig und führt die QCs nach der Kalibrierung auch selbstständig aus. Das Praxispersonal darauf hingewiesen das in Zukunft wenn ein Fehler auftritt oder Unklarheiten sind und diese nicht selbst behoben werden können, in der HOTLINE angerufen werden sollte. Das Praxispersonal nochmals die Menüstruktur, Kalibrierung der Parameter und wie eine QC sowie eine Systemreinigungen zu machen ist, erklärt. AQT IST IM MOMENT NICHT MESSBEREIT WEIL KEINE PARAMETER KALIBRIERT SIND. ER STEHT IN ERNEUERUNG. ES WURDEN DESHALB AUCH KEINE QCS DURCHGEFÜHRT. | Ich wurde von Herrn Radloff darüber informiert das der AQT im Fehler steht. Bei Eintreffen beim Kunden festgestellt das mehrere QCs nicht i.O. waren und eine TnI Kalibrierung gescheitert war. Es wurde der Fehler ( 1035 Prüfung maximale dunkelstrommessung gescheitert) angezeigt. Dieser Fehler war laut Archiv schon länger vorhanden. Optische Einheit nach Manual getauscht. Systemüberprüfung durchgeführt. Mehrere Systemreinigungen durchgeführt.--> i.O. Es konnten keine Parameter kalibriert werden da keine Kalibrationskassette der Parameter vorhanden waren und leider auch keine kal Kassette nachbestellt werden konnte. vorhandene Kalibrierungen entfernt da diese auf die alte optische Einheit programmiert waren, und nun nicht mehr gültig sind. Vorhandenes Material vom Kunden mitgenommen und dem Kunden gutgeschrieben. Kunden von jedem genutzten Parameter ( TnI, NTproBNP, D-Dimmer) einen neue Schachtel via e-Order bestellt. Diese pastellnen Testkits werden voraussichtlich am nächsten Tag (29.05.15) vor 10 Uhr geliefert. Kunde kalibriert die Parameter selbstständig und führt die QCs nach der Kalibrierung auch selbstständig aus. Das Praxispersonal darauf hingewiesen das in Zukunft wenn ein Fehler auftritt oder Unklarheiten sind und diese nicht selbst behoben werden können, in der HOTLINE angerufen werden sollte. Das Praxispersonal nochmals die Menüstruktur, Kalibrierung der Parameter und wie eine QC sowie eine Systemreinigungen zu machen ist, erklärt. AQT IST IM MOMENT NICHT MESSBEREIT WEIL KEINE PARAMETER KALIBRIERT SIND. ER STEHT IN ERNEUERUNG. ES WURDEN DESHALB AUCH KEINE QCS DURCHGEFÜHRT. | SWO-31258 | No | I have been informed by Mr Radloff that the AQT is in error. Upon arrival at the customer, it was determined that several QCs were not in the o.o. and a TnI calibration had failed. The error (1035 check maximum dark current measurement failed) was displayed. According to the archive, this error had been around for some time. Optical unit swapped to manual. system verification. Several system cleanings performed.--> i.O. No parameters could be calibrated as there was no calibration cassette of the parameters and unfortunately no kal cassette could be reordered. removed existing calibrations because they were programmed to the old optical unit and are no longer valid. Existing material taken from the customer and credited to the customer. Customers of each parameter used ( TnI, NTproBNP, D-Dimmer) ordered a new box via e-order. These pastel test kits are expected to be delivered the next day (29.05.15) before 10 a.m. Customer calibrates the parameters independently and executes the QCs independently after calibration. The practice staff pointed out that in the future, if an error occurs or is unclear and these cannot be corrected themselves, HOTLINE should be called in the future. The practice staff once again explained the menu structure, calibration of the parameters and how to make a QC as well as a system cleaning. AQT IS NOT MESSBEREIT WEIL NO PARAMETERS ARE CALIBRATED IN THE MOMENT. IT IS IN RENEWAL. IT IS ALSO NOT PERFORMED QCS. | I have been informed by Mr Radloff that the AQT is in error. Upon arrival at the customer, it was determined that several QCs were not in the o.o. and a TnI calibration had failed. The error (1035 check maximum dark current measurement failed) was displayed. According to the archive, this error had been around for some time. Optical unit swapped to manual. system verification. Several system cleanings performed.--> i.O. No parameters could be calibrated as there was no calibration cassette of the parameters and unfortunately no kal cassette could be reordered. removed existing calibrations because they were programmed to the old optical unit and are no longer valid. Existing material taken from the customer and credited to the customer. Customers of each parameter used ( TnI, NTproBNP, D-Dimmer) ordered a new box via e-order. These pastel test kits are expected to be delivered the next day (29.05.15) before 10 a.m. Customer calibrates the parameters independently and executes the QCs independently after calibration. The practice staff pointed out that in the future, if an error occurs or is unclear and these cannot be corrected themselves, HOTLINE should be called in the future. The practice staff once again explained the menu structure, calibration of the parameters and how to make a QC as well as a system cleaning. AQT IS NOT MESSBEREIT WEIL NO PARAMETERS ARE CALIBRATED IN THE MOMENT. IT IS IN RENEWAL. IT IS ALSO NOT PERFORMED QCS.
AQT90 FLEX |  |  | CF0350D5843F3256 | No
AQT90 FLEX | Pipettennadel ersetzt, Systemreinigung durchgeführt, Fließwege gereinigt, Funktionen im Serviceprogramme getestet, Probemessungen durchgeführt, | Pipettennadel ersetzt, Systemreinigung durchgeführt, Fließwege gereinigt, Funktionen im Serviceprogramme getestet, Probemessungen durchgeführt, | SWO-31181 | No | Pipette needle replaced, system cleaning performed, flow paths cleaned, functions tested in service programs, sample measurements performed, | Pipette needle replaced, system cleaning performed, flow paths cleaned, functions tested in service programs, sample measurements performed,
AQT90 FLEX | Systemreinigung und -Überprüfung durchgeführt TropI (Lot10556) und NTproBNP (Lot10433) kalibriert Gerät überprüft, LQC durchgeführt offene Fragen mit Frau Ketzer geklärt | Systemreinigung und -Überprüfung durchgeführt TropI (Lot10556) und NTproBNP (Lot10433) kalibriert Gerät überprüft, LQC durchgeführt offene Fragen mit Frau Ketzer geklärt | SWO-31107 | No | System cleaning and verification performed TropI (Lot10556) and NTproBNP (Lot10433) calibrated device checked, LQC performed open questions with Mrs. Ketzer clarified | System cleaning and verification performed TropI (Lot10556) and NTproBNP (Lot10433) calibrated device checked, LQC performed open questions with Mrs. Ketzer clarified
ABL80 CO-OX | Customer had replaced several SC's and SP's. Found that internal PO2 C8303 out high. Check Zero with Zeroing SC. Zero # is not stable and goes high and low. Sensor Connector board is suspect. | Replaced Sensor Connector Board. Perform Zero check. Set Zero. Also perform Annual PM as it is due next week. Replaced parts from PM kit. Perform additional alignment. Internal QC fine but L 1 external QC out low. Replaced SC. L1 QC fine now. | CF04775FFA03F746 | No | Customer had replaced several SC's and SP's. Found that internal PO2 C8303 out high. Check Zero with Zeroing SC. Zero # is not stable and goes high and low. Sensor Connector board is suspect. | Replaced Sensor Connector Board. Perform Zero check. Set Zero. Also perform Annual PM as it is due next week. Replaced parts from PM kit. Perform additional alignment. Internal QC fine but L 1 external QC out low. Replaced SC. L1 QC fine now.
ABL90 FLEX | Versuch das Gerät manuell zu spülen scheitert neue, mitgebrachte Kassette (Run0316) eingebaut Gerät manuell gespült, kalibriert und überprüft fällige tHb Kal durchgeführt | Versuch das Gerät manuell zu spülen scheitert neue, mitgebrachte Kassette (Run0316) eingebaut Gerät manuell gespült, kalibriert und überprüft fällige tHb Kal durchgeführt | SWO-31368 | No | Attempting to flush the device manually fails new, brought cassette (Run0316) built-in device manually rinsed, calibrated and checked due tHb Kal performed | Attempting to flush the device manually fails new, brought cassette (Run0316) built-in device manually rinsed, calibrated and checked due tHb Kal performed | Assisted Replacement | SC | Leak
ABL820 FLEX | pH BG folyadék szelep nem üzemel, szelep tisztítása, esedékes Ref membrán cseréje, kalibrációk, QC mérés |  | SWO-31308 | No | pH BG Liquid valve Not in operation, Valve cleaning, due Ref membrane replacement, calibrations, QC measurement |  | Defect | HW | pH/BG Module | Valve, rinse
ABL835 FLEX | Evt. statische Aufladung möglich - Standortwechsel Analysator - Radiance Anbindung und Funktionsüberprüfung etc. | Evt. statische Aufladung möglich - Standortwechsel Analysator - Radiance Anbindung und Funktionsüberprüfung etc. | SWO-31338 | No | Evt. Static charging possible - location change analyzer - Radiance connection and function verification etc. | Evt. Static charging possible - location change analyzer - Radiance connection and function verification etc. | Unknown | HW | DMS_PC | Touch screen
ABL90 FLEX | Gebrauchten Touchscreen (P1110N0574) eingebaut und (P1333N3273) ausgebaut, System- und Funktionskontrolle, System und Funktion i.O. --> warten auf Neuteil | Gebrauchten Touchscreen (P1110N0574) eingebaut und (P1333N3273) ausgebaut, System- und Funktionskontrolle, System und Funktion i.O. --> warten auf Neuteil | SWO-31626 | No | Used touchscreen (P1110N0574) installed and removed (P1333N3273), system and function control, system and function in the me --> waiting for new part | Used touchscreen (P1110N0574) installed and removed (P1333N3273), system and function control, system and function in the me --> waiting for new part | Defect | HW | Touch screen
ABL90 FLEX | Bereits eingesetzten Lösungspack installiert. pco2 und Lac. QC Fehler Kal. O.k. SC 945-009 und LP 944-197 eingesetzt. LP wird spontan nicht erkannt, Solution Pack PCB erneuert. | Bereits eingesetzten Lösungspack installiert. pco2 und Lac. QC Fehler Kal. O.k. SC 945-009 und LP 944-197 eingesetzt. LP wird spontan nicht erkannt, Solution Pack PCB erneuert. | SWO-31619 | No | already used solution pack. pco2 and Lac. QC Error Kal. O.K. SC 945-009 and LP 944-197. LP is not detected spontaneously, Solution Pack PCB is renewed. | already used solution pack. pco2 and Lac. QC Error Kal. O.K. SC 945-009 and LP 944-197. LP is not detected spontaneously, Solution Pack PCB is renewed. | Defect | HW | SP PCB
ABL90 | Troubleshot with customer. Attempted fluid blockage removal through service programs. | Replacement of solution pack and sensor cassette. Calibrations come in. | CF0872E45C5EBEAD | No | Troubleshot with customer. Attempted fluid blockage removal through service programs. | Replacement of solution pack and sensor cassette. Calibrations come in. | Cleaned | Fluidic pathway
ABL835 FLEX | Proteinentfernung durchgeführt - Messkammer gereinigt pO2 Elektrode erneuert Gerät kalibriert und überprüft | Proteinentfernung durchgeführt - Messkammer gereinigt pO2 Elektrode erneuert Gerät kalibriert und überprüft | SWO-31518 | No | Protein removal performed - measuring chamber cleaned pO2 electrode renewed device calibrated and checked | Protein removal performed - measuring chamber cleaned pO2 electrode renewed device calibrated and checked | Assisted Replacement | Electrode | Electrode | pO2
ABL837 FLEX |  |  | CF10DE53965798C3 | No |  |  | User Error | HW | Inlet module | Inlet Gasket | Broken Cap
ABL815 FLEX |  |  | CF084A355C33AE04 | No |  |  | Assisted Replacement | Electrode | Electrode | Lac
ABL825 FLEX | PCO2 running right on lower end of radiometer range QC for level 1 and 2. Not acceptable to the customer. | Replaced PCO2 membrane, checked pH Oring, Checked fluid flow in pH BG module in regards to bulletin. Still on low end | CF0821505C633B2E | Yes | PCO2 running right on lower end of radiometer range QC for level 1 and 2. Not acceptable to the customer. | Replaced PCO2 membrane, checked pH Oring, Checked fluid flow in pH BG module in regards to bulletin. Still on low end | Re classified RMED | Within Specs
AQT90 FLEX | Systemreinigung, LQC wiederholt. | Systemreinigung, LQC wiederholt. | SWO-31742 | No | System cleaning, LQC repeated. | System cleaning, LQC repeated.
ABL80 CO-OX |  |  | CF23F790802B243A | No
ABL80 FLEX Basic version |  |  | CF09E4EFD20D843E | No
ABL80 FLEX CO-OX | Keine weiteren corrective Aktionen möglich, da SC nicht vorhanden ist. PO2 in Kal. Fehlerhaft, für Dialyse nicht vorrangig wichtig. Alle weiteren Parameter o.k. PO2 wurde ausgeblendet. Ansonsten keine weiteren Funktionsstörungen. Bei Lieferung und Einbau der neuen Sensorkassetten sollte der Fehler behoben sein. | Keine weiteren corrective Aktionen möglich, da SC nicht vorhanden ist. PO2 in Kal. Fehlerhaft, für Dialyse nicht vorrangig wichtig. Alle weiteren Parameter o.k. PO2 wurde ausgeblendet. Ansonsten keine weiteren Funktionsstörungen. Bei Lieferung und Einbau der neuen Sensorkassetten sollte der Fehler behoben sein. | SWO-31990 | No | No further corrective actions possible because SC does not exist. PO2 in Kal. Incorrect, not primarily important for dialysis. All other parameters ok PO2 has been hidden. Otherwise no further malfunctions. When the new sensor cartridges are delivered and installed, the error should be corrected. | No further corrective actions possible because SC does not exist. PO2 in Kal. Incorrect, not primarily important for dialysis. All other parameters ok PO2 has been hidden. Otherwise no further malfunctions. When the new sensor cartridges are delivered and installed, the error should be corrected. |  | To fix the calibration error he replaced the SC.
ABL90 FLEX | Apparecchio testato in programmi servizio riattivato. | Apparecchio testato in programmi servizio riattivato. | SWO-31680 | No | Device tested in reactivated service programs. | Device tested in reactivated service programs. | Assisted Replacement | SC | Wetability
ABL90 | Upon arrival, viewed data records and verified above issue. | Upon arrival, performed service tip to commit Clock Zone changes. Rebooted unit. | CF099247D22746E5 | No | Upon arrival, viewed data records and verified above issue. | Upon arrival, performed service tip to commit Clock Zone changes. Rebooted unit. | Defect | SW | Time setting
ABL820 FLEX | Hemolyser vervangen. Lamp afgeregeld, 3327. Complete oxitest uitgevoerd. Service tHb calibratie uitgevoerd, fcuv 0,9873. Terugmeting voor alle parameters OK. Alle vier controles bepaald. | Hemolyser vervangen. Lamp afgeregeld, 3327. Complete oxitest uitgevoerd. Service tHb calibratie uitgevoerd, fcuv 0,9873. Terugmeting voor alle parameters OK. Alle vier controles bepaald. | SWO-31721 | No | Hemolyser replacement. Lamp adjusted, 3327. Complete Oxitest run. Service tHb Calibration performed, FCUV 0.9873. Return measurement for all parameters OK. All four controls determined. | Hemolyser replacement. Lamp adjusted, 3327. Complete Oxitest run. Service tHb Calibration performed, FCUV 0.9873. Return measurement for all parameters OK. All four controls determined. | Defect | HW | Oxi module | Hemolyzer
ABL825 FLEX | checked data logs for errors, checked qc log, checked vales and liquid sensors, | removed dual housing and cleaned valves, replaced connecting tubing, cleaned pH electrode and chamber | CF0A0777D24D79F9 | No | checked data logs for errors, checked qc log, checked vales and liquid sensors, | removed dual housing and cleaned valves, replaced connecting tubing, cleaned pH electrode and chamber | Defect | HW | Inlet module | Tubings
ABL825 FLEX | Lac Elektrodenkontakte nachjustiert. nach mehreren Kal ok QC ok | Lac Elektrodenkontakte nachjustiert. nach mehreren Kal ok QC ok | SWO-32138 | No | Lac electrode contacts adjusted. after several Kal ok QC ok | Lac electrode contacts adjusted. after several Kal ok QC ok | Adjustment | Eckhard | El/Met Module | electrode connector
ABL827 FLEX | Reinigungsarbeiten, Fehlersuche, Ventile der Messmodule demontiert und kontrolliert, Reinigung des Lufteinlassventils am MetII-Modul, Probelauf, mehrere korrekte Kal abgewartet, QK ok | Reinigungsarbeiten, Fehlersuche, Ventile der Messmodule demontiert und kontrolliert, Reinigung des Lufteinlassventils am MetII-Modul, Probelauf, mehrere korrekte Kal abgewartet, QK ok | SWO-31709 | No | Cleaning work, troubleshooting, valves of the measuring modules disassembled and controlled, cleaning of the air inlet valve on the MetII module, test run, several correct cal awaiting, QK ok | Cleaning work, troubleshooting, valves of the measuring modules disassembled and controlled, cleaning of the air inlet valve on the MetII module, test run, several correct cal awaiting, QK ok | Defect | HW | Met II Module | Valve, air
ABL835 FLEX |  |  | CF21EC4B9462590B | No |  |  | Defect | HW | Oxi module | Oxi module
ABL835 FLEX |  |  | CF09CB06D23BCE0B | No |  |  | Adjustment | HW | Oxi module | Lamp unit
ABL80 FLEX Basic version |  |  | CF0AC86B48113421 | No
ABL80 FLEX CO-OX | mit Test Cass. Zero Kal durchgeführt anschließend Kal und beide Lvl QC i.O. 2x SC 2x LP beim Kunden gelassen | mit Test Cass. Zero Kal durchgeführt anschließend Kal und beide Lvl QC i.O. 2x SC 2x LP beim Kunden gelassen | SWO-31932 | No | with Test Cass. Zero Kal then performed Kal and both Lvl QC in the 2x SC 2x LP left at the customer's | with Test Cass. Zero Kal then performed Kal and both Lvl QC in the 2x SC 2x LP left at the customer's
ABL90 FLEX | LP wurde schon ersetzt, war aber vor einer Woche schon einmal installiert, Einlassdichtung fällig, Einlass- und Einlassanschluss Dichtung ersetzt, Kal und QC durchgeführt | Kundin geraten erst einmal den LP90 zu tauschen da dieser nur noch 45 Aktivitäten hat, sollte dies nicht zur Problemlösung führen wird die Kundin noch die SC90 tauschen - dies ist dann die letzte die die Kundin noch vor Ort hat, sollte es hier weitere Probleme geben wird sich die Kundin wieder melden. | SWO-31982 | Yes | LP has already been replaced, but was already installed a week ago, inlet seal due, inlet and inlet connection seal replaced, Kal and QC performed | Customer first advised to exchange the LP90 because it has only 45 activities, if this does not lead to problem solving the customer will still exchange the SC90 - this is then the last one that the customer still has on site, if there are further problems here the customer wi eder. | Re classified RMED
ABL90 FLEX | Diverse Fehler, u.A. Substrat Temp. Fehler. Servicemenue aufgerufen, Sensorbay fährt bis auf etwa 2cm herunter und sofort wieder auf - immer wieder nachvollziehbar. Sensormodul ausgebaut und Öffnungsmechanik manuell überprüft => Lid Closed Indicator mit Spannungsrissen im Gehäuse => neu justiert und Mechanik mehrfach manuell gefahren, Modul eingebaut, Öffnen Schließen nun auch per Servicemenue ohne Probleme möglich. Neue SC eingesetzt. SW Update auf V3.1MR3. Startverfahren, Kal und QC Messungen: Na und Lac noch nicht stabil, bitte später noch einmal messen! | Diverse Fehler, u.A. Substrat Temp. Fehler. Servicemenue aufgerufen, Sensorbay fährt bis auf etwa 2cm herunter und sofort wieder auf - immer wieder nachvollziehbar. Sensormodul ausgebaut und Öffnungsmechanik manuell überprüft => Lid Closed Indicator mit Spannungsrissen im Gehäuse => neu justiert und Mechanik mehrfach manuell gefahren; Modul eingebaut, Öffnen Schließen nun auch per Servicemenue ohne Probleme möglich. Neue SC eingesetzt. SW Update auf V3.1MR3. Startverfahren, Kal und QC Messungen: Na und Lac noch nicht stabil, bitte später noch einmal messen! | SWO-32002 | No | Various errors, including Substrate Temp. Error. Service menu called, Sensorbay shuts down to about 2cm and immediately reopens - again and again. Sensor module removed and opening mechanism manually checked => Lid Closed Indicator with voltage cracks in the housing => re-adjusted and mechanics driven manually several times; Module built-in, open closing now also possible via service menu without any problems. New SC. SW Update to V3.1MR3. Start method, Kal and QC measurements: Well and Lac not yet stable, please measure again later! | Various errors, including Substrate Temp. Error. Service menu called, Sensorbay shuts down to about 2cm and immediately reopens - again and again. Sensor module removed and opening mechanism manually checked => Lid Closed Indicator with voltage cracks in the housing => re-adjusted and mechanics driven manually several times; Module built-in, open closing now also possible via service menu without any problems. New SC. SW Update to V3.1MR3. Start method, Kal and QC measurements: Well and Lac not yet stable, please measure again later! | Adjustment |  | SC Door
ABL90 |  |  | CF0B08D5485C8566 | No |  |  | Defect | HW | Oxi module | Pinch Valve tube
ABL90 FLEX | Allgemeine Kontrolle ok. Alle Tests im Servicemenü ok. Thb gescheitert 2x mit Fehler 1011. Hemolyzer ersetzt ok. System i.o Kal. + QC + Thb i.o. | Allgemeine Kontrolle ok. Alle Tests im Servicemenü ok. Thb gescheitert 2x mit Fehler 1011. Hemolyzer ersetzt ok. System i.o Kal. + QC + Thb i.o. | SWO-31848 | No | General control ok. All tests in the service menu ok. Thb failed 2x with error 1011. Hemolyzer replaces ok. System i.o Kal. + QC + Thb i.o. | General control ok. All tests in the service menu ok. Thb failed 2x with error 1011. Hemolyzer replaces ok. System i.o Kal. + QC + Thb i.o. | Defect | HW | Hemolyzer
ABL90 | O cliente já tinha trocado o cartucho de soluções e os erros mantinham-se. Limpeza do circuito da amostra. Reinstalação do cartucho de soluções. Inicialização. Verificação do correcto funcionamento. |  | CF0AE62C4840BF5E | No | The client had already switched the solution cartridge and the errors kept. cleaning the sample circuit. Reinstallation of the solution cartridge. Startup. Verification of proper operation. |  | Cleaned | Fluidic pathway
ABL810 FLEX |  |  | CF0A9DC4483E49E3 | No |  |  | Defect | HW | pH/BG Module | Valve, air
ABL825 FLEX |  |  | CF0AB3EF4844C9B4 | No |  |  | Defect | HW | Barcode Reader
ABL800 BASIC | Key eingespielt | Key eingespielt | SWO-31951 | Yes | Key recorded | Key recorded | Re classified RMED
ABL825 FLEX | Mangels Spare Part Ersatz Oxi Modul eingebaut Oxi PCB bestellt Kal und QC ok | Mangels Spare Part Ersatz Oxi Modul eingebaut Oxi PCB bestellt Kal und QC ok | SWO-32077 | No | Lack of Spare Part Replacement Oxi Module Built-in Oxi PCB Ordered Kal and QC ok | Lack of Spare Part Replacement Oxi Module Built-in Oxi PCB Ordered Kal and QC ok | Defect | yes sure, I try to help. | Oxi module | PCB
ABL800 BASIC | Luftwege freigespült, Gerät kalibriert, Kappiläre Messung durchgeführt, QC Chargen konfiguriert. | Luftwege freigespült, Gerät kalibriert, Kappiläre Messung durchgeführt, QC Chargen konfiguriert. | SWO-32085 | No | Airways cleared, device calibrated, kappile measurement performed, QC batches configured. | Airways cleared, device calibrated, kappile measurement performed, QC batches configured. | Cleaned | The analyzer has a calibration error at pCO2 (not pO2 as describe in the case title) and additional to this the error message 145 was present. To solve this issue the FSE replace the SP, but the message was still there. After that he shut down the analyzer and restarts again. The message was not present. | Air vent
ABL825 FLEX | Ersatzmodul aus Altgerät eingebaut Sparepart bestellt | Ersatzmodul aus Altgerät eingebaut Sparepart bestellt | SWO-32149 | No | Replacement module from old device built-in spare part ordered | Replacement module from old device built-in spare part ordered | Defect |  | AutoCheck Module | PCB
ABL825 FLEX | Ersatzmodul installiert, Ersatzteile bestellt | Einlassnadel manuell vorgeschoben und Einlass im Service kalibriert, Fehler behoben. Net Op. neu gestartet. Fehler bleib bestehen, no response from 172.17.68.85. Alle NetOp Fenster schließen lassen. Alle Einstellungen überprüft und neu eingestellt, Fehler bleibt bestehenden. | SWO-32048 | No | Replacement module installed, spare parts ordered | Inlet pin manually pushed forward and inlet calibrated in the service, error fixed. Net Op. restarted. Errors remain, no response from 172.17.68.85. Let all NetOp windows close. All settings checked and re-set, error remains existing. | Defect | HW | Inlet module
ABL825 FLEX | System gereinigt und dekontamination durchgeführt | System gereinigt und dekontamination durchgeführt | SWO-31896 | No | System cleaned and decontamination carried out | System cleaned and decontamination carried out | Cleaned | Fluidic pathway
ABL835 FLEX | Der Scanner konnte nicht funktionieren da in der Roma.ini Einträge fehlten z.b. BCR.Parameter usw. Das Verzögerte Ansprechverhalten des Bildschirms wird vermutlich durch ein IP Adressenkonflikt verursacht. Wenn der Haken Radiance Verbindung herstellen gewählt wird kommt es zu dem beschriebenen Bildschirm einfrieren. Wenn nun versucht wird den Haken zu entfernen reagiert der Bildschirm mit starker Verzögerung ca. 30-40 sec. Ich habe nun folgendes gemacht. Bei dem Alt Bestandsgerät die fehlenden Roma. Einträge eingepflegt so das der Scanner wieder funktioniert. Am Montag wird bei dem neuen ABL S N 754R2015N004 die Windows und ABL Software neu aufgesetzt. Außerdem werden die Radiance Anbindungs Daten aller an Radiance hängenden Geräten überprüft. Alles ist mit der Med. Technik so abgestimmt. Zusätzlich werden wir Eventl. Ein Leihgerät aufgestellten. | Der Scanner konnte nicht funktionieren da in der Roma.ini Einträge fehlten z.b. BCR.Parameter usw. Das Verzögerte Ansprechverhalten des Bildschirms wird vermutlich durch ein IP Adressenkonflikt verursacht. Wenn der Haken Radiance Verbindung herstellen gewählt wird kommt es zu dem beschriebenen Bildschirm einfrieren. Wenn nun versucht wird den Haken zu entfernen reagiert der Bildschirm mit starker Verzögerung ca. 30-40 sec. Ich habe nun folgendes gemacht. Bei dem Alt Bestandsgerät die fehlenden Roma. Einträge eingepflegt so das der Scanner wieder funktioniert. Am Montag wird bei dem neuen ABL S N 754R2015N004 die Windows und ABL Software neu aufgesetzt. Außerdem werden die Radiance Anbindungs Daten aller an Radiance hängenden Geräten überprüft. Alles ist mit der Med. Technik so abgestimmt. Zusätzlich werden wir Eventl. Ein Leihgerät aufgestellten. | SWO-32110 | No | The scanner could not work because in the Roma.ini entries were missing e.g. BCR. parameters, etc. The delayed response of the screen is probably caused by an IP address conflict. When the check mark To establish Radiance connection is selected, it occurs to freeze the screen described. When trying to remove the hook, the screen reacts with strong delay about 30-40 sec. I have now done the following. In the old stock device the missing Roma. entries so that the scanner works again. On Monday, the new ABL S N 754R2015N004 will re-install the Windows and ABL software. In addition, the Radiance connection data of all devices attached to Radiance is checked. Everything is coordinated with med. technology. In addition, we will be Eventl. A rental device set up. | The scanner could not work because in the Roma.ini entries were missing e.g. BCR. parameters, etc. The delayed response of the screen is probably caused by an IP address conflict. When the check mark To establish Radiance connection is selected, it occurs to freeze the screen described. When trying to remove the hook, the screen reacts with strong delay about 30-40 sec. I have now done the following. In the old stock device the missing Roma. entries so that the scanner works again. On Monday, the new ABL S N 754R2015N004 will re-install the Windows and ABL software. In addition, the Radiance connection data of all devices attached to Radiance is checked. Everything is coordinated with med. technology. In addition, we will be Eventl. A rental device set up. | Unknown | Best regards
ABL90 FLEX | Fehlersuche Altes LP erneut installiert ( Fehlermeldung bleibt), neues LP installiert Einlaß gereinigt, Fließwege gespült Datenbankfehler könnt mit längerer Geräte Aus Phase behoben werden Neue SC installiert -Einlaufverfahren | Fehlersuche Altes LP erneut installiert ( Fehlermeldung bleibt), neues LP installiert Einlaß gereinigt, Fließwege gespült Datenbankfehler könnt mit längerer Geräte Aus Phase behoben werden Neue SC installiert -Einlaufverfahren | SWO-32018 | No | Troubleshooting Old LP reinstalled (error message remains), new LP installed inlet cleaned, flow paths flushed database error can be fixed with longer devices Off phase New SC installs inlet method | Troubleshooting Old LP reinstalled (error message remains), new LP installed inlet cleaned, flow paths flushed database error can be fixed with longer devices Off phase New SC installs inlet method | Assisted Replacement | SC | Wetability
ABL90 | Verificação do mau contacto, substituição do cabo da alimentação, reinicialização, estabilização e passagem dos QCs. |  | CF0C4AF0BE32F0C2 | No | Verification of bad contact, replacement of power cord, reboot, stabilization and passage of QCs. |  | Defect | HW | Power cable
ABL825 FLEX | 841-774 ausgetauscht da von neu aus einLoch drinn pH Elektrode ausgetauscht ( ohne Erfolg und neue EL undicht) 834-660 montiert, pH Elektrode Dicht. Messkammer entleerung einwandfrei pCO2 mit Lot R0601, R0666 und R0630 neu membraniert. (KeinErfolg) Funktionskontrolle i.o QC i.o aber immer noch Grenzwertig Techline! | 841-774 ausgetauscht da von neu aus einLoch drinn pH Elektrode ausgetauscht ( ohne Erfolg und neue EL undicht) 834-660 montiert, pH Elektrode Dicht. Messkammer entleerung einwandfrei pCO2 mit Lot R0601, R0666 und R0630 neu membraniert. (KeinErfolg) Funktionskontrolle i.o QC i.o aber immer noch Grenzwertig Techline! | SWO-32058 | No | 841-774 replaced since replaced from new onehole drinn pH electrode (without success and new EL leak) 834-660 mounted, pH electrode seal. Measuring chamber emptying flawlessly re-membraneed pCO2 with solder R0601, R0666 and R0630. (Nosuccess) Functional control i.o QC i.o but still limitig Techline! | 841-774 replaced since replaced from new onehole drinn pH electrode (without success and new EL leak) 834-660 mounted, pH electrode seal. Measuring chamber emptying flawlessly re-membraneed pCO2 with solder R0601, R0666 and R0630. (Nosuccess) Functional control i.o QC i.o but still limitig Techline! | Defect |  | pH/BG Module | Tubings
AQT90 FLEX |  |  | CF0C1FD1BE0BF219 | No
ABL80 FLEX CO-OX | Gerät überprüft. Anschlusspin der neuen Sensorkassette beim instalieren verbogen, geradebiegen nicht möglich. Sensorkassette aus Kundenbestand erneuert. Sensor Initalisierung durchgeführt. Erste Initalisierung nicht erfolgreich. Gerät neu gebootet und selbe Sensorkassette initalisiert. OK. Kalibrierung durchgeführt. QC durchgeführt. | neues LP und SC zum Kunden versendet - der Kunde hat die SC heute erneuert - alles wieder OK | SWO-32218 | No | device checked. Connection pin of the new sensor cassette bent when instaling, straight bending not possible. Sensor cassette from customer base renewed. sensor initalization. Initial initiation not successful. Device rebooted and inituated the same sensor cartridge. Ok. calibration performed. QC. | new LP and SC shipped to the customer - the customer renewed the SC today - everything is OK again
ABL90 FLEX | Essai avec un autre pack de solution : NOK Remplacement des tuyaux a l'interieur. Mis en route et vérification dans le programme de service : OK Calibration et contrle : OK Retiré le câble de la batterie. | Essai avec un autre pack de solution : NOK Remplacement des tuyaux a l'interieur. Mis en route et vérification dans le programme de service : OK Calibration et contrle : OK Retiré le câble de la batterie. | SWO-32180 | No | Try with another solution pack: NOK Replacement of pipes inside. Started and checked in the service program: OK Calibration and Check: OK Remove the battery cable. | Try with another solution pack: NOK Replacement of pipes inside. Started and checked in the service program: OK Calibration and Check: OK Remove the battery cable. | Defect | HW | Waste Tube
ABL90 | Order new supplies. Found the errors follows the SC | Replaced the SC; pCO2 sensetivity very low Tried a different SC, inhomogeneous solution errors Back to 1st SC and sensitivityy still low | CF0D5B0234258257 | No | Order new supplies. Found the errors follows the SC | Replaced the SC; pCO2 sensetivity very low Tried a different SC, inhomogeneous solution errors Back to 1st SC and sensitivityy still low | Assisted Replacement | SC | Wetability
ABL90 FLEX | Uhrzeit im Control Panel korrigiert, auf CF-Card gespeichert, Neustart, Uhrzeit überprüft - i.O., SC erneuert, Kal u. QC i.O. | Uhrzeit im Control Panel korrigiert, auf CF-Card gespeichert, Neustart, Uhrzeit überprüft - i.O., SC erneuert, Kal u. QC i.O. | SWO-32233 | No | Time corrected in control panel, saved to CF card, restart, time checked - i.O., SC renewed, Kal & QC i.O. | Time corrected in control panel, saved to CF card, restart, time checked - i.O., SC renewed, Kal & QC i.O. | Defect | SW | Time setting
ABL90 FLEX | Lösungspack richtig aktiviert, Schlauch am Quetschventil, Abfallschlauch, Einlassanschlussdichtung und Einlassdichtung erneuert, System- und Funktionskontrolle, System und Funktion i.O. -->Lösungspack (RV04) beim Kunden vor Ort gelassen, da dieser keinen mehr hat | Lösungspack richtig aktiviert, Schlauch am Quetschventil, Abfallschlauch, Einlassanschlussdichtung und Einlassdichtung erneuert, System- und Funktionskontrolle, System und Funktion i.O. -->Lösungspack (RV04) beim Kunden vor Ort gelassen, da dieser keinen mehr hat | SWO-32153 | No | Solution pack correctly activated, hose on the squeegee valve, waste hose, inlet connection seal and inlet seal renewed, system and function control, system and function as a.o. -->Solution pack (RV04) left on site at the customer's premises, as this one no longer has one | Solution pack correctly activated, hose on the squeegee valve, waste hose, inlet connection seal and inlet seal renewed, system and function control, system and function as a.o. -->Solution pack (RV04) left on site at the customer's premises, as this one no longer has one | User Error | SP activation
ABL90 | 24 hour Clock not enabled.Checked Time Zone Internet Time Roma Time Daylight Savings Unchecked. | Performed 24 hour clock set. | CF0D83DE3459E12C | No | 24 hour Clock not enabled. Checked Time Zone Internet Time Roma Time Daylight Savings Unchecked. | Performed 24 hour clock set. | Setup | 24H setting
ABL90 | Attempted backflush. Disassembled unit and flushed each part individually, in both directions. Something was stuck between connector gasket and sensor cassette bay. | Managed to free whatever was stuck (but wasn't able to see it) | CF0D370F3444C80B | No | Attempted backflush. Disassembled unit and flushed each part individually, in both directions. Something was stuck between connector gasket and sensor cassette bay. | Managed to free whatever was stuck (but wasn't able to see it) | Cleaned | Fluidic pathway
ABL817 FLEX | Gerät überprüft, Meßkammer und Luftauslass gesäubert, Ventile kontrolliert, Halogenlampe eingestellt, Ca membraniert LOT (566), Einlaßdichtung ausgetauscht, tHb-Kal durchgeführt, Oxi Test durchgeführt, Hemolysator erneuert und justiert, Einlaßnadel, Schlauch und Modulpumpenschläuche ersetzt, Dichtungsring pH getauscht, Ref.-Elektrode membraniert, Kontakte kontrolliert, probeweise Ca und Ref.- Elektrode ersetzt, tHb-Kal durchgeführt |  | SWO-32184 | No | Device checked, measuring chamber and air outlet cleaned, valves checked, halogen lamp adjusted, Ca membraniert LOT (566), inlet seal replaced, tHb-Kal performed, Oxi test performed, Hemolysator renewed and adjusted, inlet needle, hose and Module pump hoses replaced, sealing ring pH replaced, ref.-electrode membraned, contacts controlled, test Ca and Ref.- electrode replaced, tHb-Kal performed |  | Defect | HW | Oxi module | Hemolyzer
ABL825 FLEX | 237 and 248 errors present spectrometer voltage high and missing trigger. | Problem would recur after unit was rebooted. Replaced and programmed oxi module PCB. Oxi module not active error. Checked connections. Reprogrammed X2.Oxi module still not active. Reinstalled original Oxi PCB. Oxi module is active but original problem remains. | CF0D937A343DF180 | No | 237 and 248 errors present spectrometer voltage high and missing trigger. | Problem would recur after unit was rebooted. Replaced and programmed oxi module PCB. Oxi module not active error. Checked connections. Reprogrammed X2. Oxi module still not active. Reinstalled original Oxi PCB. Oxi module is active but original problem remains. | Unknown | HW | Oxi module | Spectrometer
ABL835 FLEX | Allgemeine Kontrolle ok. Einlassdichtung ersetzt ok. PH Modul Messkammer gereinigt ok. PCO2 Membrane ersetzt ok. System i.o. Kal. + QC i.o. |  | SWO-32176 | No | General control ok. Inlet seal replaced ok. PH module measuring chamber cleaned ok. PCO2 membrane replaced ok. System i.o. Kal. + QC i.o. |  | Assisted Replacement | HW | Inlet module | Inlet Gasket
ABL90 | Bad consumables and ordered more pCO2 sensitivity low and inhomogeneous errors | Replaced SP and left customer a back up SC Ran blood samples and test flow in service-OK | CF101AB9201D098B | No | Bad consumables and ordered more pCO2 sensitivity low and inhomogeneous errors | Replaced SP and left customer a back up SC Ran blood samples and test flow in service-OK | Assisted Replacement | SP | FT
ABL835 FLEX | Inlet modul csere a szervízes készülékből, Na membrán csere, Kalibrációk (Inlet, LS, Pumpa, tHb,Elektródák) QC mérés hitelesítés Későbbi kiszálás során új modul beépítése szükséges |  | SWO-32210 | No | Inlet module replacement from the service unit, Na membrane replacement, calibrations (Inlet, LS, pump, tHb, electrodes) QC measurement Verification The installation of a new module is required during subsequent |  | Defect | HW | Inlet module
ABL90 FLEX | PC Unit erneuert. tHb Kal, Kalibration und QC's vorgenommen. Neue MAC 00-04-5F-97-16-48 | PC Unit erneuert. tHb Kal, Kalibration und QC's vorgenommen. Neue MAC 00-04-5F-97-16-48 | SWO-32213 | No | PC Unit. tHb Kal, Calibration and QC's. New MAC 00-04-5F-97-16-48 | PC Unit. tHb Kal, Calibration and QC's. New MAC 00-04-5F-97-16-48 | Defect | HW | CPU
ABL90 | User has flushed the analyzer and replaced the sensor cassette and solution pack and continues to have multiple errors. Found signs of leakage around the hemolyzer. | Replaced Liquid Sensor 3 and the Solution Pack I O PCB. Flushed system with hot water and noticed excessive back pressure. Replaced the waste tubing, pinch valve tubing, and inlet connector gasket and inlet gasket. Unable to resolve the errors. | CF114EA4965833D3 | No | User has flushed the analyzer and replaced the sensor cassette and solution pack and continues to have multiple errors. Found signs of leakage around the hemolyzer. | Replaced Liquid Sensor 3 and the Solution Pack I O PCB. Flushed system with hot water and noticed excessive back pressure. Replaced the waste tubing, pinch valve tubing, and inlet connector gasket and inlet gasket. Unable to resolve the errors. | Unknown | Revisit required
ABL90 | LOOKED THROUGH DATA LOGS AND SAW WHERE THE CALIBRATION FAILED FOR PO2 IN THE EARLY MORNING 7 TIMES. CUSTOMER PUT NEW SOLUTION PACK AND SONIA SAID THE ANALYZER DIDN'T RECOGNIZE NEW SOLUTION PACK. SO SHE CHANGED IT AGAIN. THEN IT STARTED WORKING BUT WAS GETTING 1324: INHOMOGENEOUS RINSE SOLUTION. SO I FLUSHED UNIT 15 TIMES AND RE-INTALLED SOLUTION PACK. DID IT TWICE BEFORE FINALLY GOT A CHANCE TO DO 5 WASTE BLOODS. | ONCE COMPLETED ALL 5 WASTE BLOODS. INHOMOGENOUS SAMPLE WENT AWAY. | CF1112BB9628E2B0 | No | LOOKED THROUGH DATA LOGS AND SAW WHERE THE CALIBRATION FAILED FOR PO2 IN THE EARLY MORNING 7 TIMES. CUSTOMER PUT NEW SOLUTION PACK AND SONIA SAID THE ANALYZER DIDN'T RECOGNIZE NEW SOLUTION PACK. SO SHE CHANGED IT AGAIN. THEN IT STARTED WORKING BUT WAS GETTING 1324: INHOMOGENEOUS RINSE SOLUTION. SO I FLUSHED UNIT 15 TIMES AND RE-INTALLED SOLUTION PACK. DID IT TWICE BEFORE FINALLY GOT A CHANCE TO DO 5 WASTE BLOODS. | ONCE COMPLETED ALL 5 WASTE BLOODS. INHOMOGENOUS SAMPLE WENT AWAY. | Assisted Replacement | SC | Wetability
ABL90 | Se sustituye el casette de electrodos, se realiza calibración y se comprueban diferentes consumibles, se accede a servicio y se realiza chequeo de la señal electronica observandose que no cambia. | Fallo de calibración del Na+ | CF1122E0963C2820 | No | The electrode cassette is replaced, calibration is performed and different consumables are checked, service is accessed and the electronic signal is checked not to change. | Na+ calibration failure | Assisted Replacement | SC | Na+
ABL817 FLEX | Gerät überprüft, El Met Modul erneuert, Gasventilmembran pH BG Modul ausgetauscht, Elyte und Ref. Elektrode membraniert, Rilibäkgrenzen eingegeben | Gerät überprüft, El Met Modul zurückgetauscht, AQC 5+ vermessen (keine Änderung), Fließwege mit Proteinentferner gespült, Dekontamination durchgeführt, Ref. Elektrode (7106) ersetzt, , Abfall-und Lösungspumpenschlauch erneuert, Crea membraniert, Lösungen ausgetauscht | SWO-32430 | No | Device checked, El Met module renewed, gas valve membrane pH BG module replaced, Elyte and Ref. Electrode membraneed, Rilibäk boundaries entered | Device checked, El Met module exchanged back, AQC 5+ measured (no change), flow paths rinsed with protein remover, decontamination performed, ref. Electrode (7106) replaced, waste and solution pump hose renewed, Crea membraneed, solutions replaced | Defect | HW | El/Met Module | Rilibäk
ABL835 FLEX | LOOKED IN DATA LOGS AND ANALYZER HAS PREVIOUS LACTATE ISSUES ONLY LASTING 5 DAYS TO 8 DAYS. LAST WEEK JERRY KELLY SUGGESTED CHANGING V1,V2,V3, V7 AND LOOKING AT ALL THE VALVES ON PH BG MODULE AND EL MET MODULE. FOUND WHITE DEPOSIT ON V1,V2, AND V3. ALSO FOUND OTHER VALVES MEMBRANES WORE V6, V7. ALSO BOTH EL MET VALVES. AND BLACK VALVE ON PH BG MODULE. | REPLACED 4 MEMBRANE VALVES, 1 FLAT MEMBRANE VALVE AND 3 RUBBER MEMBRANES. ALSO USED CUSTOMERS STOCK TO CHANGE ALL PUMP TUBING. | CF1151699633E2B3 | No | LOOKED IN DATA LOGS AND ANALYZER HAS PREVIOUS LACTATE ISSUES ONLY LASTING 5 DAYS TO 8 DAYS. LAST WEEK JERRY KELLY SUGGESTED CHANGING V1,V2,V3, V7 AND LOOKING AT ALL THE VALVES ON PH BG MODULE AND EL MET MODULE. FOUND WHITE DEPOSIT ON V1,V2, AND V3. ALSO FOUND OTHER VALVES MEMBRANES WORE V6, V7. ALSO BOTH EL MET VALVES. AND BLACK VALVE ON PH BG MODULE. | REPLACED 4 MEMBRANE VALVES, 1 FLAT MEMBRANE VALVE AND 3 RUBBER MEMBRANES. ALSO USED CUSTOMERS STOCK TO CHANGE ALL PUMP TUBING. | Defect | HW | Fluidic manifold | Valve, rinse
ABL827 FLEX | Tried new 14 pole cable did not help. Powewr distribution cable badly bent in several places | Ordered new cable overnight, will be back in AM. Ran QC manually. | CF111093962418DC | No | Tried new 14 pole cable did not help. Powewr distribution cable badly bent in several places | Ordered new cable overnight, will be back in AM. Ran QC manually. | Unknown | HW | AutoCheck Module | FPC cable
ABL800 BASIC | Demo key installiert. Noch fehlen Glu und one Elektrolyte Key, wurden bestellt ! | Demo key one elektrolyte und Glu installiert. Na und K membraniert. Kal mehrfach durchgeführt | SWO-32387 | Yes | Demo key installed. Still missing Glu and one electrolyte key, were ordered ! | Demo key one electrolyte and Glu installed. Na and K membraned. Kal performed multiple times | Re classified RMED | Demo key exp
ABL827 FLEX | Remplacement de la carte d'alimentation ( power supply ) e.o Remplacement de la platine ( Measuring light heat PCB du module PH ) e.o Calibration e.o contrôle de qualité e.o | Remplacement de la carte d'alimentation ( power supply ) e.o Remplacement de la platine ( Measuring light heat PCB du module PH ) e.o Calibration e.o contrôle de qualité e.o | SWO-32373 | No | Replacement of power supply e.o Platinum Replacement (Measuring light heat PCB of the PH module) e.o Calibration e.o quality control e.o | Replacement of power supply e.o Platinum Replacement (Measuring light heat PCB of the PH module) e.o Calibration e.o quality control e.o | Defect | HW | Power Supply
ABL825 FLEX | Examined barcode reader. | Replaced barcode reader. Scanned barcode. Barcode sanned correctly. Verified operation of ABL 825. | CF11199B9652A4E5 | No | Examined barcode reader. | Replaced barcode reader. Scanned barcode. Barcode sanned correctly. Verified operation of ABL 825. | Defect | HW | Basic Unit | Barcode reader
ABL825 FLEX | Does not continue to show the carousel packing after a few minutes. | Replaced the auto check module. | CF1114649626F7E5 | No | Does not continue to show the carousel packing after a few minutes. | Replaced the auto check module. | Defect | HW | AutoCheck Module
ABL835 FLEX | Neuinstallation Windows und Applikation - Software - Konfigurationseinstellungen etc. - Funktionsüberprüfung etc. und Radiance-Anbindung | Neuinstallation Windows und Applikation - Software - Konfigurationseinstellungen etc. - Funktionsüberprüfung etc. und Radiance-Anbindung | SWO-32353 | No | New installation Windows and application - Software - configuration settings etc. - Function check etc. and radiance connection | New installation Windows and application - Software - configuration settings etc. - Function check etc. and radiance connection | Defect | SW | Windows
AQT90 FLEX | Systemreinigung mit dem Kunden zusammen durchgeführt - Kunden zehn Minuten später angerufen - Gerät ist wieder OK | Systemreinigung mit dem Kunden zusammen durchgeführt - Kunden zehn Minuten später angerufen - Gerät ist wieder OK | SWO-32284 | No | System cleaning carried out together with the customer - customer called ten minutes later - device is OK again | System cleaning carried out together with the customer - customer called ten minutes later - device is OK again
ABL80 CO-OX |  |  | CF2416D880592474 | No
ABL80 FLEX CO-OX | Hemolyser schwingt nicht, Hemolyser ersetzt, Kal na Q durchgeführt | Hemolyser schwingt nicht, Hemolyser ersetzt, Kal na Q durchgeführt | SWO-32510 | No | Hemolyser does not vibrate, Hemolyser replaces, Kal na Q performed | Hemolyser does not vibrate, Hemolyser replaces, Kal na Q performed
ABL90 | Confirmed results. | Performed a tHb calibration. | CF12D9080C0B62F8 | No | Confirmed results. | Performed a tHb calibration. | Adjustment | tHb cal
ABL90 | Confirmed results. | Performed a tHb calibration. | CF12D9510C4B62F9 | No | Confirmed results. | Performed a tHb calibration. | Adjustment | tHb cal
ABL5 | 401-008 Halogenlampe ausgetauscht. Nach Neustart war die Konfiguration ungültig. Speicher Batterie erneuert | 401-008 Halogenlampe ausgetauscht. Nach Neustart war die Konfiguration ungültig. Speicher Batterie erneuert | SWO-32618 | No | 401-008 Halogen lamp replaced. After rebooting, the configuration was invalid. Memory battery renewed | 401-008 Halogen lamp replaced. After rebooting, the configuration was invalid. Memory battery renewed
ABL817 FLEX | Gerät überprüft, El Met Modul zurückgetauscht, AQC 5+ vermessen (keine Änderung), Fließwege mit Proteinentferner gespült, Dekontamination durchgeführt, Ref. Elektrode (7106) ersetzt, , Abfall-und Lösungspumpenschlauch erneuert, Crea membraniert, Lösungen ausgetauscht | Gerät überprüft, El Met Modul zurückgetauscht, AQC 5+ vermessen (keine Änderung), Fließwege mit Proteinentferner gespült, Dekontamination durchgeführt, Ref. Elektrode (7106) ersetzt, , Abfall-und Lösungspumpenschlauch erneuert, Crea membraniert, Lösungen ausgetauscht | SWO-32578 | No | Device checked, El Met module exchanged back, AQC 5+ measured (no change), flow paths rinsed with protein remover, decontamination performed, ref. Electrode (7106) replaced, waste and solution pump hose renewed, Crea membraneed, solutions replaced | Device checked, El Met module exchanged back, AQC 5+ measured (no change), flow paths rinsed with protein remover, decontamination performed, ref. Electrode (7106) replaced, waste and solution pump hose renewed, Crea membraneed, solutions replaced | Unknown | - | - | Rilibäk
ABL825 FLEX | Spectrometer errors had cleared on arrival. Unit is now aborting QC but no messages as to why. | Rebooted analyzer and tried to rerun QC. Alignment seemed off. Realigned auto check module. 237 error recurred. Replaced oxi controller with a differnt PCB and the AU engine loaded normally. QC intermittently aborting. Adjusted laser from 195 to 193. 237 error recurred again and autocheck module is continuously conditioning.Replaced analyzer control PCB. | CF1276ED0C33D57F | No | Spectrometer errors had cleared on arrival. Unit is now aborting QC but no messages as to why. | Rebooted analyzer and tried to rerun QC. Alignment seemed off. Realigned auto check module. 237 error recurred. Replaced oxi controller with a differnt PCB and the AU engine loaded normally. QC intermittently aborting. Adjusted laser from 195 to 193. 237 error recurred again and autocheck module is continuously conditioning. Replaced analyzer control PCB. | Defect | HW | Analyzer Control PCB
ABL827 FLEX | Will not pass autocheck module tests in service | Replaced power cable as old had multiple kinks . Did not help, changed AutoCheck PCB, now passes all tests. Cable probably shorted the PCB. | CF1249500C398FDC | No | Will not pass autocheck module tests in service | Replaced power cable as old had multiple kinks . Did not help, changed AutoCheck PCB, now passes all tests. Cable probably shorted the PCB. | Defect | HW | AutoCheck Module | FPC cable
ABL80 CO-OX | LOOKED AT BOTTOM OF ANALYZER, HAS EXCESSIVE LEAK at WASTE. IT S A CLEAR TAN COLOR THAT IS BUBBLY. | CLEANED LEAK, AND COMPLETED ANNUAL MAINTENANCE WITH ALL NEW TUBING AND PUMPS. | CF13C2058202125F | No | LOOKED AT BOTTOM OF ANALYZER, HAS EXCESSIVE LEAK at WASTE. IT S A CLEAR TAN COLOR THAT IS BUBBLY. | CLEANED LEAK, AND COMPLETED ANNUAL MAINTENANCE WITH ALL NEW TUBING AND PUMPS.
ABL80 CO-OX | LOOKED AT BOTTOM OF ANALYZER, HAS EXCESSIVE LEAK WASTE. ITS A CLEAR TAN COLOR THAT IS BUBBLY. | CLEANED LEAK, REPLACED MANIFOLD UNIT, AND COMPLETED ANNUAL MAINTENANCE WITH ALL NEW TUBING. | CF139B14822F125A | No | LOOKED AT BOTTOM OF ANALYZER, HAS EXCESSIVE LEAK WASTE. ITS A CLEAR TAN COLOR THAT IS BUBBLY. | CLEANED LEAK, REPLACED MANIFOLD UNIT, AND COMPLETED ANNUAL MAINTENANCE WITH ALL NEW TUBING.
ABL90 | Found hemolyzer leakage. LS3 and tubings and gaskets have already been replaced. | Replaced the hemolyzer unit. The sensor cassette did not work well because of the amount of time it was inactive. Replaced the sensor cassette. | CF13BF908216B475 | No | Found hemolyzer leakage. LS3 and tubings and gaskets have already been replaced. | Replaced the hemolyzer unit. The sensor cassette did not work well because of the amount of time it was inactive. Replaced the sensor cassette. | Defect | HW | Hemolyzer | Clot
ABL90 | Checked entire unit for fluid leakage as reported by night shift operator. Replacement sensor cassette failing pCO2 sensitivity. | No leakage found. Started annual maintenance. Sensor cassette will need replaced. Will replace from customer stock. Credit form will need completed. | CF13CCAF8256656E | No | Checked entire unit for fluid leakage as reported by night shift operator. Replacement sensor cassette failing pCO2 sensitivity. | No leakage found. Started annual maintenance. Sensor cassette will need replaced. Will replace from customer stock. Credit form will need completed. | Assisted Replacement | SC | pCO2
ABL90 FLEX | Reste des Dichtungshalters aus dem Einlassmodul entfernt, Einlassnadel stark verbogen - Einlassnadel und Einlassdichtung mit Halter aus Kundenbestand ersetzt | Reste des Dichtungshalters aus dem Einlassmodul entfernt, Einlassnadel stark verbogen - Einlassnadel und Einlassdichtung mit Halter aus Kundenbestand ersetzt | SWO-32751 | No | Remains of the seal holder removed from the inlet module, inlet needle strongly bent - inlet needle and inlet seal with holder from customer stock replaced | Remains of the seal holder removed from the inlet module, inlet needle strongly bent - inlet needle and inlet seal with holder from customer stock replaced | Defect | HW | Inlet Gasket | Broken
ABL90 FLEX | Gerät steht in Intervention erforderlich . Servicemenue aufgerufen und alle erforderlichen Tests und Programme aufgerufen, alle einwandfrei bestanden und nach Startphase Ampel grün inkl. QC Messungen | Gerät steht in Intervention erforderlich . Servicemenue aufgerufen und alle erforderlichen Tests und Programme aufgerufen, alle einwandfrei bestanden und nach Startphase Ampel grün inkl. QC Messungen | SWO-32680 | No | Device is required in intervention . Service menu called and all necessary tests and programs called, all passed flawlessly and after start phase traffic light green incl. QC measurements | Device is required in intervention . Service menu called and all necessary tests and programs called, all passed flawlessly and after start phase traffic light green incl. QC measurements | No fault found | - | Wetability
ABL90 | 6 9 Examined the Sensor cassette | 6 9 Replaced inlet probe, inlet gasket w holder and sensor cassette. 6 10Checked on ABL90 and talked to B. Medina. Verified operation of ABL90. | CF13C07482447634 | No | 6 9 Examined the Sensor cassette | 6 9 Replaced inlet probe, inlet gasket w holder and sensor cassette. 6 10Checked on ABL90 and talked to B. Medina. Verified operation of ABL90. | Assisted Replacement | SC | - | pO2
TCM-COMBIM | Comprobar batería y conector, anular el conector y soldar directamente los cables a la batería, calibrar y comprobar. | Soldar los cables directamente a la batería. | CF136DC88245C412 | No | Check battery and connector, override the connector and directly weld the cables to the battery, calibrate and check. | Weld the cables directly to the battery.
ABL815 FLEX | Defekter Wastesensor ersetzt. | Defekter Wastesensor ersetzt. | SWO-32720 | No | Replaced the defective waste sensor. | Replaced the defective waste sensor. | Defect | HW | Basic Unit | Waste detector
ABL835 FLEX | Fehlersuche Messkammer und Kontakte gereinigt PH Elektrode erneuert alte Lot VU-02 Kal Werte deutlich besser aber immer noch außer Bereich Ref Elektrode erneuert alte Lot 71-12 Analysator kalibriert und überprüft QC Messung durchgeführt | Fehlersuche Messkammer und Kontakte gereinigt PH Elektrode erneuert alte Lot VU-02 Kal Werte deutlich besser aber immer noch außer Bereich Ref Elektrode erneuert alte Lot 71-12 Analysator kalibriert und überprüft QC Messung durchgeführt | SWO-32696 | No | Troubleshooting measuring chamber and contacts cleaned PH electrode renewed old lot VU-02 cal values significantly better but still out of range ref electrode renewed old solder 71-12 analyzer calibrated and checked QC measurement performed | Troubleshooting measuring chamber and contacts cleaned PH electrode renewed old lot VU-02 cal values significantly better but still out of range ref electrode renewed old solder 71-12 analyzer calibrated and checked QC measurement performed | Assisted Replacement | Electrode | Electrode | pH
ABL825 FLEX | Aborted QC's have no error code and No message present in the message area. | Replaced AutoCheck module PCB. | CF13AD12824C966B | No | Aborted QC's have no error code and No message present in the message area. | Replaced AutoCheck module PCB. | Defect | HW | AutoCheck Module | PCB
ABL800 BASIC | Demo key one elektrolyte und Glu installiert. Na und K membraniert. Kal mehrfach durchgeführt | Demo key one elektrolyte und Glu installiert. Na und K membraniert. Kal mehrfach durchgeführt | SWO-32744 | Yes | Demo key one electrolyte and Glu installed. Na and K membraned. Kal performed multiple times | Demo key one electrolyte and Glu installed. Na and K membraned. Kal performed multiple times | Re classified RMED | Demo key exp
ABL80 CO-OX |  |  | CF241AB880632476 | No
ABL80 FLEX CO-OX | Fehlersuche vor Ort. Ersatzteile für den 11.06. bis 10 Uhr bestellt. Gerät mit in die GS genommen. Hemolyzer erneuert. Neue Sensorkassette eingebaut. | Fehlersuche vor Ort. Ersatzteile für den 11.06. bis 10 Uhr bestellt. Gerät mit in die GS genommen. Hemolyzer erneuert. Neue Sensorkassette eingebaut. | SWO-32772 | No | Troubleshooting on site. Spare parts for the 11th of May to 10 a.m. device into the GS. Hemolyzer. New sensor cartridge installed. | Troubleshooting on site. Spare parts for the 11th of May to 10 a.m. device into the GS. Hemolyzer. New sensor cartridge installed.
ABL80 CO-OX | Analyzers not comparing. Returning to continue repair | Comparrisons are off using green top with gel seperator. Ran 3-4 samples per machine with values way off. Replaced Spectrometer thinking a wavelength was off in the spectrum with slight change in results but still out. Changed Hemolyzer unit to confirm first repalcement was not bad with no change in results. Changed Specimen to green top lithium heprin without gel and from non smoker. Analyzer dropped 3% but still on the high end of 1.8. Most results from gel tube being around 5-5.2%. Replaced CO-OX Control Board due to skewed results. Obtained Data logs and DB to open Case | CF14F116F83C220F | No | Analyzers not comparing. Returning to continue repair | Comparrisons are off using green top with gel seperator. Ran 3-4 samples per machine with values way off. Replaced Spectrometer thinking a wavelength was off in the spectrum with slight change in results but still out. Changed Hemolyzer unit to confirm first repalcement was not bad with no change in results. Changed Specimen to green top lithium heprin without gel and from non smoker. Analyzer dropped 3% but still on the high end of 1.8. Most results from gel tube being around 5-5.2%. Replaced CO-OX Control Board due to skewed results. Obtained Data logs and DB to open Case
ABL90 FLEX | Lösungspack und Sensorcassette angeliefert. Gerät geprüft. Fliesswege gespült. Einlasshalter und Nadel erneuert. Lösungspack und Sensorcassette installiert. Pco2 Empf. nach Kal viel zu niedrig.4% Sensorcassette erneuert und installiert. 1x Lösungspack als Reserve beim Kunden gelassen. | Kundin hat kein Material vor Ort! Fr. Klante bittet um einen Service | SWO-33068 | No | solution pack and sensor cassette. device checked. Flow paths rinsed. Inlet holder and needle renewed. solution pack and sensor cassette. Pco2 Empf. to Kal far too low.4% sensor cassette renewed and installed. 1x solution pack as a reserve left at the customer' site. | Customer has no material on site! Ms. Klante asks for a service | User Error | Inlet module | installation
ABL90 | Verify error messages by looking at logs. SP test fails in service fluidics program. Put clear tubing in place of SC and SP passes test every time. Put SC back in and SP test fails. | Replaced waste port and parts from PM kit. Still have errors. Replaced SC and all is okay. Collect log files and a service dump. | CF14C56DF8526C24 | No | Verify error messages by looking at logs. SP test fails in service fluidics program. Put clear tubing in place of SC and SP passes test every time. Put SC back in and SP test fails. | Replaced waste port and parts from PM kit. Still have errors. Replaced SC and all is okay. Collect log files and a service dump. | Assisted Replacement | SC | Wetability | why no Waste Blood
ABL90 | Upon arrival found unit asking for solution pack change. | Reinstalled solution pack. Unit then went to troubleshooting mode. | CF14BE59F82B7C19 | No | Upon arrival found unit asking for solution pack change. | Reinstalled solution pack. Unit then went to troubleshooting mode. | Adjustment | Service Cal
ABL90 FLEX | Fehlersuche Probenweg gespült, Dichtungen und Schläuche erneuert | Fehlersuche Probenweg gespült; Dichtungen und Schläuche erneuert | SWO-32868 | No | Troubleshooting sample path flushed; Seals and hoses renewed | Troubleshooting sample path flushed; Seals and hoses renewed | Cleaned | Fluidic pathway
ABL90 FLEX | LP eingesetzt Analysator kalibriert SC getauscht da CO2 Empfindlichkeit bei 38% Analysator kalibriert und überprüft Interne QC OK | LP eingesetzt Analysator kalibriert SC getauscht da CO2 Empfindlichkeit bei 38% Analysator kalibriert und überprüft Interne QC OK | SWO-32785 | No | LP used analyzer calibrated SC exchanged as CO2 sensitivity calibrated at 38% analyzer and checked internal QC OK | LP used analyzer calibrated SC exchanged as CO2 sensitivity calibrated at 38% analyzer and checked internal QC OK | User Error | SP activation
ABL90 FLEX | Remplacé le kit de maintenance e.o Remplacé l'aiguille d'aspiration e.o Nettoyage de la section fluidique e.o Remplacé cassette (nouveau lot) et pack de solutions (nouveau lot) .o Calibration e.o CQ e.o GARANTIE | Remplacé le kit de maintenance e.o Remplacé l'aiguille d'aspiration e.o Nettoyage de la section fluidique e.o Remplacé cassette (nouveau lot) et pack de solutions (nouveau lot) .o Calibration e.o CQ e.o GARANTIE | SWO-32822 | No | Replaced the maintenance kit e.o Replaced the suction needle e.o Cleaning the fluidsection e.o Replaced cassette (new batch) and solution pack (new batch) .o Calibration e.o CQ e.o GARANTIE | Replaced the maintenance kit e.o Replaced the suction needle e.o Cleaning the fluidsection e.o Replaced cassette (new batch) and solution pack (new batch) .o Calibration e.o CQ e.o GARANTIE | Cleaned | Fluidic pathway
ABL825 FLEX | Data logs show COHb QC results to be erratic. Diagnostic shows normal OXI module operation. | Replaced hemolyzer per Radiometer specifications. Calibrated and ran multiple QC's of all levels. Instrument results were on or near mean with expected repeatability. Instrument operates normally. | CF14C2E6F8600045 | No | Data logs show COHb QC results to be erratic. Diagnostic shows normal OXI module operation. | Replaced hemolyzer per Radiometer specifications. Calibrated and ran multiple QC's of all levels. Instrument results were on or near mean with expected repeatability. Instrument operates normally. | Defect | HW | Oxi module | Hemolyzer
AQT90 FLEX | Tausch der Optischen Einheit und des Process Wheels. Factory Reference neu kalibriert, Wert o.K. 1919168,41. Alle drei Parameter neu kalibriert. Kalibrationskassetten auf Lager beim Anwender. Bestell-Nummer Proccess Wheel nicht im CRM, Nr. 903-236. Durchführung von QC, Werte o.K. Bei CRP gab es eine QC Grenzverletzung. Systemreinigung durchgeführt. Erneut QC durchgeführt Auch mit neuer AC Charge. Analysator funktionsbereit. | Tausch der Optischen Einheit und des Process Wheels. Factory Reference neu kalibriert, Wert o.K. 1919168,41. Alle drei Parameter neu kalibriert. Kalibrationskassetten auf Lager beim Anwender. Bestell-Nummer Proccess Wheel nicht im CRM, Nr. 903-236. Durchführung von QC, Werte o.K. Bei CRP gab es eine QC Grenzverletzung. Systemreinigung durchgeführt. Erneut QC durchgeführt Auch mit neuer AC Charge. Analysator funktionsbereit. | SWO-32953 | No | Replacement of the Optical Unit and the Process Wheel. Factory Reference recalibrated, value o.K. 1919168.41. Recalibrates all three parameters. Calibration cassettes in stock at the user's disposal. Order number Proccess Wheel not in CRM, No. 903-236. Execution of QC, values o.K. CrP had a QC. system cleaning. Again QC performed Also with new AC Charge. Analyzer ready for use. | Replacement of the Optical Unit and the Process Wheel. Factory Reference recalibrated, value o.K. 1919168.41. Recalibrates all three parameters. Calibration cassettes in stock at the user's disposal. Order number Proccess Wheel not in CRM, No. 903-236. Execution of QC, values o.K. CrP had a QC. system cleaning. Again QC performed Also with new AC Charge. Analyzer ready for use.
AQT90 FLEX | Fünf Messungen mit Spüllösung durchgeführt und ausgewertet. Alle Messungen liegen bei 1250-1355. Max. bei 1600. Cup-Wash-Unit und Nadel ausgebaut und gereinigt. Nochmals div. Schläuche und Pumpe1 getauscht. Reinigungsprogramm durchgeführt. Danach nochmals fünf Blankmessungen durchgeführt und ausgewertet. Alle Messungen lagen unter dem Grenzbereich. Gerät mit QC überprüft. | Fünf Messungen mit Spüllösung durchgeführt und ausgewertet. Alle Messungen liegen bei 1250-1355. Max. bei 1600. Cup-Wash-Unit und Nadel ausgebaut und gereinigt. Nochmals div. Schläuche und Pumpe1 getauscht. Reinigungsprogramm durchgeführt. Danach nochmals fünf Blankmessungen durchgeführt und ausgewertet. Alle Messungen lagen unter dem Grenzbereich. Gerät mit QC überprüft. | SWO-32835 | No | Five measurements with rinsing solution carried out and evaluated. All measurements are 1250-1355. Max. at 1600. Cup-Wash-Unit and needle removed and cleaned. Replaced again in a various hoses and pump1. cleaning program. Then another five blank measurements were carried out and evaluated. All measurements were below the limit range. Device checked with QC. | Five measurements with rinsing solution carried out and evaluated. All measurements are 1250-1355. Max. at 1600. Cup-Wash-Unit and needle removed and cleaned. Replaced again in a various hoses and pump1. cleaning program. Then another five blank measurements were carried out and evaluated. All measurements were below the limit range. Device checked with QC.
ABL90 | Redid annual overhaul Replaced SP | Replaced SC | CF1609C86E09656F | No | Redid annual overhaul Replaced SP | Replaced SC | Assisted Replacement | SC | Wetability
ABL90 FLEX | In Control Panel Uhrzeit korrigiert und Einstellungen kontrolliert - Reconnect Intervall auf 120 Sekunden erhöht - Funktionsüberprüfung und Erläuterungen | In Control Panel Uhrzeit korrigiert und Einstellungen kontrolliert - Reconnect Intervall auf 120 Sekunden erhöht - Funktionsüberprüfung und Erläuterungen | SWO-33051 | No | In Control Panel time corrected and settings controlled - Reconnect interval increased to 120 seconds - Function check and explanations | In Control Panel time corrected and settings controlled - Reconnect interval increased to 120 seconds - Function check and explanations | Defect | SW | Freeze
ABL90 FLEX | Kalibrierung und QC Werte i.O. Vergleichsmessungen durchgeführt(ABL Kreißsaal ) EL Ergebnisse identisch Techn. Probleme am Gerät unwahrscheinlich Problematik Kap.Messung bei Elektrolyten angesprochen | Kalibrierung und QC Werte i.O. Vergleichsmessungen durchgeführt(ABL Kreißsaal ) EL Ergebnisse identisch Techn. Probleme am Gerät unwahrscheinlich Problematik Kap.Messung bei Elektrolyten angesprochen | SWO-33022 | No | Calibration and QC values in the case of comparative measurements performed (ABL Kreißsaal ) EL Results identical Techn. Problems on the device unlikely Problem Cap.Measurement addressed in electrolytes | Calibration and QC values in the case of comparative measurements performed (ABL Kreißsaal ) EL Results identical Techn. Problems on the device unlikely Problem Cap.Measurement addressed in electrolytes | User Error | Capillary tube w/o Balanced Hep.
ABL90 FLEX | Einlassdichtung und Ventilschlauch erneuert - Fehler weiter vorhanden. Hemolyzer Unit erneuert. Oxi Complete Test ok. tHb Kalibration ausgeführt. Kalibration und QC's vorgenommen | Einlassdichtung und Ventilschlauch erneuert - Fehler weiter vorhanden. Hemolyzer Unit erneuert. Oxi Complete Test ok. tHb Kalibration ausgeführt. Kalibration und QC's vorgenommen | SWO-33066 | No | Inlet seal and valve hose renewed - fault spresent. Hemolyzer Unit. Oxi Complete Test ok. tHb calibration. Calibration and QC's made | Inlet seal and valve hose renewed - fault spresent. Hemolyzer Unit. Oxi Complete Test ok. tHb calibration. Calibration and QC's made | Cleaned | Fluidic pathway | Replaced Hemo
ABL90 FLEX | Vom Kunden durchgeführte Vergleichsmessungen ( ABL90 zu ABL90 ) analysiert. Zusätzlich eigene Vergleichsmessungen durchgeführt. Dazu Daten und Informationen vom Labor und HKL eingeholt. Unsere ( Klaus und Zvonko ) Ergebnisse Stat. -Leitung Frau Goll erläutert. Ebenso wurde Herr Hübner aus der Med. -Technik über den aktuellen Stand informiert. Bezüglich des IT-Problems wird nun überprüft ob die Lizens für das Dorner IT-System für denABL noch gültig ist. Es wird von uns analysiert ob die engen Ca Berichte wie vom Kunden gewünscht eingehalten und mit anderen Systemen verglichen werden können. | Vom Kunden durchgeführte Vergleichsmessungen ( ABL90 zu ABL90 ) analysiert. Zusätzlich eigene Vergleichsmessungen durchgeführt. Dazu Daten und Informationen vom Labor und HKL eingeholt. Unsere ( Klaus und Zvonko ) Ergebnisse Stat. -Leitung Frau Goll erläutert. Ebenso wurde Herr Hübner aus der Med. -Technik über den aktuellen Stand informiert. Bezüglich des IT-Problems wird nun überprüft ob die Lizens für das Dorner IT-System für denABL noch gültig ist. Es wird von uns analysiert ob die engen Ca Berichte wie vom Kunden gewünscht eingehalten und mit anderen Systemen verglichen werden können. | SWO-33098 | No | Customer-based comparative measurements ( ABL90 to ABL90 ) analyzed. In addition, own comparative measurements were carried out. For this purpose, data and information were obtained from the laboratory and HKL. Our ( Klaus and Zvonko ) results Stat. -Director Ms. Goll explained. Mr Hübner was also informed about the current state of play by Med. Technology. With regard to the IT problem, it is now checked whether the license for the Dorner IT system is still valid for the ABL. We analyze whether the narrow Ca reports can be adhered to as desired by the customer and compared with other systems. | Customer-based comparative measurements ( ABL90 to ABL90 ) analyzed. In addition, own comparative measurements were carried out. For this purpose, data and information were obtained from the laboratory and HKL. Our ( Klaus and Zvonko ) results Stat. -Director Ms. Goll explained. Mr Hübner was also informed about the current state of play by Med. Technology. With regard to the IT problem, it is now checked whether the license for the Dorner IT system is still valid for the ABL. We analyze whether the narrow Ca reports can be adhered to as desired by the customer and compared with other systems. | Unknown | Performance | Ca, Lac
ABL90 FLEX | Gerät gespült, Einlass separat gespült, Kal & QC i.O., techn. Einweisung der div. Anwender | Gerät gespült, Einlass separat gespült, Kal & QC i.O., techn. Einweisung der div. Anwender | SWO-33103 | No | Device rinsed, inlet rinsed separately, Kal & QC i.O., technical instruction of the div. user | Device rinsed, inlet rinsed separately, Kal & QC i.O., technical instruction of the div. user | Cleaned | Fluidic pathway
TCM-COMBIM | Substituição da membrana. Calibração, o equipamento pede para substituir a membrana sucessivamente sem calibrar. Substituição do electrodo. Calibração e ectivação da função SmartCal. |  | CF161DF66E4F4743 | No | Membrane replacement. Calibration, the equipment asks to replace the membrane successively without calibrating. Replacing the Electrodo. Calibration and ectification of the SmartCal function.
ABL800 BASIC | Fehlersuche Luftweg gespült Funktion der Ventile überprüft Kal 1 Lösung nicht richtig angesteckt Gas 2 erneuert Ablassbehälter erneuert Analysator kalibriert und überprüft QC Messung durchgeführt Darocha@ukmp.de | Fehlersuche Luftweg gespült Funktion der Ventile überprüft Kal 1 Lösung nicht richtig angesteckt Gas 2 erneuert Ablassbehälter erneuert Analysator kalibriert und überprüft QC Messung durchgeführt Darocha@ukmp.de | SWO-32979 | No | Troubleshooting Airway flushed function of the valves checked Kal 1 solution not properly plugged gas 2 renewed drain tank renewed analyzer calibrated and checked QC measurement performed Darocha@ukmp.de | Troubleshooting Airway flushed function of the valves checked Kal 1 solution not properly plugged gas 2 renewed drain tank renewed analyzer calibrated and checked QC measurement performed Darocha@ukmp.de | Cleaned | Fluidic pathway
ABL825 FLEX | System überprüft, Flex-Q-Modul ausgetauscht und justiert, System OK. | System überprüft, Positionierungsfehler des Moduls, nach Rücksprache Modul zurück getauscht und Kundenmodul gangbar gemacht. | SWO-33237 | No | System checked, Flex-Q module replaced and adjusted, system OK. | System checked, positioning error of the module, exchanged after consultation module and customer module made usable. | Adjustment | HW | FlexQ Module | Sample tray(same analyzer)
ABL825 FLEX | System überprüft, Positionierungsfehler des Moduls, nach Rücksprache Modul zurück getauscht und Kundenmodul gangbar gemacht. | System überprüft, Positionierungsfehler des Moduls, nach Rücksprache Modul zurück getauscht und Kundenmodul gangbar gemacht. | SWO-33238 | No | System checked, positioning error of the module, exchanged after consultation module and customer module made usable. | System checked, positioning error of the module, exchanged after consultation module and customer module made usable. | Adjustment | HW | FlexQ Module | Sample tray(same analyzer)
ABL835 FLEX | Den grå slange er skiftet. Den blå o-ring er skiftet. Der er kørt dekontaminering 2 gange. Der er kørt cleaning 1 gang. Der er kørt væske sensor justering. Der er kørt pumpe kalibrering. Der er kørt 2 punkts kal. Der er kørt QCer. Der er kørt patient prøver. Alt er OK. | Den grå slange er skiftet. Den blå o-ring er skiftet. Der er kørt dekontaminering 2 gange. Der er kørt cleaning 1 gang. Der er kørt væske sensor justering. Der er kørt pumpe kalibrering. Der er kørt 2 punkts kal. Der er kørt QCer. Der er kørt patient prøver. Alt er OK. | SWO-33319 | No | The grey snake has changed. The Blue O-ring is switched. Decontamination has been made 2 times. Cleaning 1 time has been run. Fluid sensor adjustment has been made. Pump calibration has been run. Two-point Kal has been run. QCer has been run. Patient tests have been carried out. Everything is OK. | The grey snake has changed. The Blue O-ring is switched. Decontamination has been made 2 times. Cleaning 1 time has been run. Fluid sensor adjustment has been made. Pump calibration has been run. Two-point Kal has been run. QCer has been run. Patient tests have been carried out. Everything is OK. | Defect | HW | El/Met Module | Tubing connection
ABL825 FLEX | Photo Interrupter ausgebaut und überprüft. Housing down ok Housing up getrocknete Flüssigkeit nach Reinigung Funktionen ok. Lamp Unit Justage durchgeführt. | Photo Interrupter ausgebaut und überprüft. Housing down ok Housing up getrocknete Flüssigkeit nach Reinigung Funktionen ok. Lamp Unit Justage durchgeführt. | SWO-33031 | No | Photo Interrupter removed and checked. Housing down ok Housing up dried liquid after cleaning functions ok. Lamp Unit Adjustment. | Photo Interrupter removed and checked. Housing down ok Housing up dried liquid after cleaning functions ok. Lamp Unit Adjustment. | Defect | HW | AutoCheck Module | PCB
ABL830 FLEX | Display Unit kompl. ausgetauscht, | Display Unit kompl. ausgetauscht, | SWO-32969 | No | Display Unit kompl. Exchanged | Display Unit kompl. Exchanged | Defect | HW | Basic Unit | Display
ABL830 FLEX | QC- Modul im Serviceprogramm getestet, Verbindungskabel ersetzt, alle Funktionen getestet, | QC- Modul im Serviceprogramm getestet, Verbindungskabel ersetzt, alle Funktionen getestet, | SWO-32970 | No | QC module tested in the service program, connection cable replaced, all functions tested, | QC module tested in the service program, connection cable replaced, all functions tested, | Defect | HW | AutoCheck Module | FPC cable
ABL825 FLEX |  |  | CF23A00D8025F352 | No |  |  | Defect | Electrode | Electrode | pH
ABL90 |  |  | CF19E2FFD0366CBC | No |  |  | Defect | HW | Inlet Gasket
ABL90 | HAD PREVIOUS PROBLEM AND WAS WAITING FOR SOFTWARE INSTALL. FOUND BAD CF CARD. | REPLACED CF CARD, RELOADED WINDOWS, RELOADED RADIOMETER SOFTWARE, RELOADED WIRELESS, RE-INSTALLED ALL DATA AND SETUP. | CF1A15C5D03F0FF3 | No | HAD PREVIOUS PROBLEM AND WAS WAITING FOR SOFTWARE INSTALL. FOUND BAD CF CARD. | REPLACED CF CARD, RELOADED WINDOWS, RELOADED RADIOMETER SOFTWARE, RELOADED WIRELESS, RE-INSTALLED ALL DATA AND SETUP. | Defect | HW | CF card
ABL90 FLEX | Service Programme durchlaufen lassen neue graue Dichtung eingesetzt neue SC, da Gerät seit 3 Tagen ohne Spülung Startup abgewartet | Service Programme durchlaufen lassen neue graue Dichtung eingesetzt neue SC, da Gerät seit 3 Tagen ohne Spülung Startup abgewartet | SWO-33156 | No | Service programs go through let new gray seal inserted new SC as device waited for 3 days without flushing startup | Service programs go through let new gray seal inserted new SC as device waited for 3 days without flushing startup | Assisted Replacement | SC | Wetability
ABL90 | Verificação dos registros, limpeza do circuito, inspecção da sonda, substituição da entrada, testes de serviço OK, reinicialização, estabilização e passagem dos QCs. |  | CF1A1EAAD01B0934 | No | Checking records, cleaning the circuit, inspecting the probe, replacing the input, testing of OK service, reboot, stabilization and passage of the QCs. |  | Defect | HW | Inlet Gasket
ABL825 FLEX | 1) Lactaat probleem: Hiervoor werd de volledige EL MET module omgewisseld. Reiniging van de contacten en de dichtingen. Bij de Glu was de dichting niet meer aanwezig : vervangen. 2) pH lek : De blauwe dichting werd weer vervangen door de vorige versie (zwarte dichting). Extra reiniging van de meetkamers : OK 3) pO2 drift : Hier werd de meetkamer ook extra gereinigd. Contacten zijn OK. De mogelijkheid bestaat dat de driften ontstaan door de vervuiling veroorzaakt door het lek in de pH meetkamer. Voorlopig even afwachten of de bovenstaande acties positief resultaat geven. | 1) Lactaat probleem: Hiervoor werd de volledige EL MET module omgewisseld. Reiniging van de contacten en de dichtingen. Bij de Glu was de dichting niet meer aanwezig : vervangen. 2) pH lek : De blauwe dichting werd weer vervangen door de vorige versie (zwarte dichting). Extra reiniging van de meetkamers : OK 3) pO2 drift : Hier werd de meetkamer ook extra gereinigd. Contacten zijn OK. De mogelijkheid bestaat dat de driften ontstaan door de vervuiling veroorzaakt door het lek in de pH meetkamer. Voorlopig even afwachten of de bovenstaande acties positief resultaat geven. | SWO-33174 | No | 1) Lactate problem: For this the complete EL WITH module was exchanged. Cleaning of the contacts and seals. At the Glu the seal was no longer present: replaced. 2) pH leak: The blue seal was replaced again by the previous version (Black seal). Additional cleaning of the measuring chambers: OK 3) pO2 drift: Here the measuring chamber was also cleaned extra. Contacts are OK. The possibility exists that the drifts arise due to the pollution caused by the leak in the pH measuring chamber. For the time being, wait and see if the above actions give positive results. | 1) Lactate problem: For this the complete EL WITH module was exchanged. Cleaning of the contacts and seals. At the Glu the seal was no longer present: replaced. 2) pH leak: The blue seal was replaced again by the previous version (Black seal). Additional cleaning of the measuring chambers: OK 3) pO2 drift: Here the measuring chamber was also cleaned extra. Contacts are OK. The possibility exists that the drifts arise due to the pollution caused by the leak in the pH measuring chamber. For the time being, wait and see if the above actions give positive results. | Assisted Replacement | Electrode | Electrode | pH | Sealing ring
ABL827 FLEX |  |  | CF1B318D463AB6D1 | No |  |  | Assisted Replacement | Electrode | Electrode | Lac
ABL825 FLEX | Error328. No leading air segment in inlet liquid sensor. Found nipple on inlet module broken. | Replaced nipple on inlet module. | CF19F666D05A88D4 | No | Error328. No leading air segment in inlet liquid sensor. Found nipple on inlet module broken. | Replaced nipple on inlet module. | Defect | HW | Inlet module | Nipple kit | Nipple kit
ABL827 FLEX | checked inlet probe and inlet gasket, | removed and replaced inlet PCB, installed AU software,no change, replaced inlet friction detector and original PCB | CF1A6322D02FD874 | No | checked inlet probe and inlet gasket, | removed and replaced inlet PCB, installed AU software,no change, replaced inlet friction detector and original PCB | Defect | HW | Inlet module | Friction detector unit
ABL80 CO-OX |  |  | CF241C7680092477 | No
ABL80 FLEX CO-OX | Hämolysator erneuert. Kalibrationen und Qualitätskontrollen in Ordnung. | Hämolysator erneuert. Kalibrationen und Qualitätskontrollen in Ordnung. | SWO-33459 | No | Hemolyser renewed. calibrations and quality controls in order. | Hemolyser renewed. calibrations and quality controls in order.
ABL90 | Symptoms indicate sensor cassette contamination. | Flushed the analyzer with hot water. Resolved the pump solution issue, but now the pO2 and pCO2 will not work. Replaced the sensor cassette. Analyzer worked for about two hours, then failed with a Replace Solution Pack, Unable to Pump Solutions error. | CF1B55FD465BE1B5 | No | Symptoms indicate sensor cassette contamination. | Flushed the analyzer with hot water. Resolved the pump solution issue, but now the pO2 and pCO2 will not work. Replaced the sensor cassette. Analyzer worked for about two hours, then failed with a Replace Solution Pack, Unable to Pump Solutions error. | Assisted Replacement | SC | Wetability
ABL90 | ABL90 keeps going into the service inntervention screen. Cannot pump solution error message is presant. After rplacing parts started getting inhomoginous sample error. Then the sensor cassette would not calibrate. | Replaced the Sensor Cassette bay assembly, waste connector, and put in a new annual PM kit. Replaced the solution pack from customerstock. Replaced the sensor cassette from customer stock. | CF1C1E3BBC2F6E7A | No | ABL90 keeps going into the service inntervention screen. Cannot pump solution error message is presant. After rplacing parts started getting inhomoginous sample error. Then the sensor cassette would not calibrate. | Replaced the Sensor Cassette bay assembly, waste connector, and put in a new annual PM kit. Replaced the solution pack from customerstock. Replaced the sensor cassette from customer stock. | Defect | HW | Cassette retainer
ABL90 |  |  | CF1B1D7B462A4E26 | No |  |  | Defect | HW | CPU
ABL90 |  |  | CF1D83873242CA0A | No |  |  | Assisted Replacement | SC | Wetability
ABL90 FLEX | Wechsel SC | Wechsel SC | SWO-33461 | Yes | Change SC | Change SC | Re classified RMED | Customer replaced SC, OK
ABL90 | Trouble shot and found a recessed pin on the Solution Pack ID PCB. | Replaced the Solution Pack ID PCB, and this resolved the problem. I completed the leftover checks required during the PM, that could not be done due to the malfunction. | CF1B2A7346176E99 | No | Trouble shot and found a recessed pin on the Solution Pack ID PCB. | Replaced the Solution Pack ID PCB, and this resolved the problem. I completed the leftover checks required during the PM, that could not be done due to the malfunction. | Defect | HW | SP PCB
ABL825 FLEX | WDC report is showing pO2 to be below Peer Group.Inspected logs. pO2 has been within 1SD on unit. | Cleaned Chambers Contacts Ran decontamination. | CF1B05CF464B086A | No | WDC report is showing pO2 to be below Peer Group.Inspected logs. pO2 has been within 1SD on unit. | Cleaned Chambers Contacts Ran decontamination. | Defect | Electrode | Membrane | pO2
ABL827 FLEX |  |  | CF1C6315BC28A481 | No |  |  | Cleaned | Fluidic pathway
ABL825 FLEX | Kabelverbindungen vom Display überprüft, Analysator runter und wieder hochgefahren, QC Level 1 neu eingespannt, da nur 2 stellig sichtbar, QC Karussell aus und wieder eingebaut, Glukose Elektrode neu membraniert, da Empfindlichkeitsfehler, Kal, QC | Kabelverbindungen vom Display überprüft, Analysator runter und wieder hochgefahren, QC Level 1 neu eingespannt, da nur 2 stellig sichtbar, QC Karussell aus und wieder eingebaut, Glukose Elektrode neu membraniert, da Empfindlichkeitsfehler, Kal, QC | SWO-33465 | No | Cable connections checked from the display, analyzer down and restarted, QC Level 1 re-engaged as only 2 digits visible, QC carousel off and reinstalled, glucose electrode re-membranes as sensitivity error, Kal, QC | Cable connections checked from the display, analyzer down and restarted, QC Level 1 re-engaged as only 2 digits visible, QC carousel off and reinstalled, glucose electrode re-membranes as sensitivity error, Kal, QC | Defect | SW | Windows | Restart
AQT90 FLEX |  |  | CF1C406BBC5E53D2 | No
ABL80 CO-OX | Found Data reveiling possible bias of 4-5% on CO's asparating from Vacutainer of any green top with Gel. | Installed all original parts Performed tHb calibration Performed all system calibrations Obtained 25 EDTA tubes for testing across AVOX, 90, and 80 with customer. Will verify and confirm analyzer was in working condition from initial visit, and the method of collection will need to be suggested for CO's. Customer stated that Syringes will be mainly used, but will obtain a tube once in a while. He will discuss with group and make a final decision. | CF1C5AA9BC239270 | No | Found Data reveiling possible bias of 4-5% on CO's asparating from Vacutainer of any green top with Gel. | Installed all original parts Performed tHb calibration Performed all system calibrations Obtained 25 EDTA tubes for testing across AVOX, 90, and 80 with customer. Will verify and confirm analyzer was in working condition from initial visit, and the method of collection will need to be suggested for CO's. Customer stated that Syringes will be mainly used, but will obtain a tube once in a while. He will discuss with group and make a final decision.
ABL80 CO-OX | unit will not pass Sc initiilation, air detected errors | replaced tubing and pump assy using 80 pm kit, same error replaced conector board, and anolog pcb, same error inspected valve assy, hemo unit and spectrometer for possible leaks or cracks none found. new Sc and Sp were tested same issue | CF1C7F04BC0D99EB | No | unit will not pass Sc initiilation, air detected errors | replaced tubing and pump assy using 80 pm kit, same error replaced conector board, and anolog pcb, same error inspected valve assy, hemo unit and spectrometer for possible leaks or cracks none found. new Sc and Sp were tested same issue
ABL90 FLEX | Maintenant tout les paramètres sont OK (calibrations et CQ). J'ai pris un service dump pour envoyer a Radiometer. Test en service : OK Calibrations et CQ : OK | Maintenant tout les paramètres sont OK (calibrations et CQ). J'ai pris un service dump pour envoyer a Radiometer. Test en service : OK Calibrations et CQ : OK | SWO-33568 | No | Maintenant tout les paramètres sont OK (calibrations et CQ). J'ai pris un service dump pour envoyer a Radiometer. Test en service : OK Calibrations et CQ : OK | Maintenant tout les paramètres sont OK (calibrations et CQ). J'ai pris un service dump pour envoyer a Radiometer. Test en service : OK Calibrations et CQ : OK | Assisted Replacement | SC | Chip error
ABL90 FLEX | Neues Lösungspack angeschlossen und Gerät kalibriert, | Neues Lösungspack angeschlossen und Gerät kalibriert, | SWO-33561 | No | New solution pack connected and device calibrated, | New solution pack connected and device calibrated, | Assisted Replacement | SP | FT
ABL90 |  |  | CF1C57A1BC086209 | No |  |  | Assisted Replacement | SC | Wetability
ABL90 FLEX | System überprüft, Anbindung hergestellt, mit Herrn Falabu vom OSM dieVerbindung überprüft, mehrfach Patientendaten und Qualitätskontrollen gesendet, Daten kommen an, System OK. Gespräch mit Frau Jaganow und HerrnJesse. | System überprüft, Anbindung hergestellt, mit Herrn Falabu vom OSM dieVerbindung überprüft, mehrfach Patientendaten und Qualitätskontrollen gesendet, Daten kommen an, System OK. Gespräch mit Frau Jaganow und HerrnJesse. | SWO-33553 | Yes | System checked, connection established, with Mr. Falabu from the OSM the connection checked, multiple patient data and quality controls sent, data arrive, system OK. Interview with Ms. Yaganov and Mr. Jesse. | System checked, connection established, with Mr. Falabu from the OSM the connection checked, multiple patient data and quality controls sent, data arrive, system OK. Interview with Ms. Yaganov and Mr. Jesse. | Re classified RMED | Setup
ABL90 | All diagnostics, always Pass. Visual inspection of fluid path shows nothing improper. Unit will indicate a Failure of the SP that certain Fluids are not being detected during initialization. | Ordered Hemolyzer and LS3 Sensor for replacement. | CF1C8880BC328052 | No | All diagnostics, always Pass. Visual inspection of fluid path shows nothing improper. Unit will indicate a Failure of the SP that certain Fluids are not being detected during initialization. | Ordered Hemolyzer and LS3 Sensor for replacement. | Unknown | Revisit required
ABL820 FLEX | Customer remembraned the reference membrane, completed an unsuccessful 2 point calibration and performed a tHb calibration. None of these trouble shooting methods resolved the issue | Replaced cal solution 1, performed successful complete oximetery test and 2 point calibrations. Instrument is within acceptable operating specifications. | CF1C72FABC4DBB72 | No | Customer remembraned the reference membrane, completed an unsuccessful 2 point calibration and performed a tHb calibration. None of these trouble shooting methods resolved the issue | Replaced cal solution 1, performed successful complete oximetery test and 2 point calibrations. Instrument is within acceptable operating specifications. | Unknown | HW | Oxi module | tHb cal
ABL825 FLEX | EI Met Pump Asymetry.Ran several Pump Cals, LS cals, unable to replicate problem.Inspected Roller Assy. Problem has occured only once since Install. | Customer to Monitor unit for Error. | CF1C8B46BC288055 | No | EI Met Pump Asymetry.Ran several Pump Cals, LS cals, unable to replicate problem. Inspected Roller Assy. Problem has occured only once since Install. | Customer to Monitor unit for Error. | Unknown | HW | El/Met Module | Pump tubing
ABL835 FLEX |  |  | CF23A6EE803DF366 | No |  |  | Adjustment | HW | Oxi module | Lamp unit
AQT90 FLEX | ACTUALIZACION DE SW A VERSION 8.7.6.ES Y COMPROBACION FUNCIONAMIENTO. CALIBRADO Y VERIFICADO CON Q.C. |  | CF1C53FEBC5DCEE1 | Yes | SW TO VERSION UPDATE 8.7.6.ES AND OPERATION CHECK. CALIBRATED AND VERIFIED WITH Q.C. |  | Re classified RMED | FAN / TU
AQT90 FLEX | Sustitución de diferentes piezas del analizador, la tarjeta control del laser , el cable de conexión del lector y el lector de codios de barras de los quesitos, chequeo y ajuste, funcionando correctamente. |  | CF1C8452BC2ABE16 | No | Replacing different analyzer parts, laser control card, reader connection cable and bar codios reader of the hold, checking and adjusting, working properly.
ABL80 CO-OX | Unit will not intilize SC air detected errors | As per ASM techline case opened | CF1D690A320D7ADB | No | Unit will not intilize SC air detected errors | As per ASM techline case opened
ABL90 FLEX | Nach Aktuellem Service Manual wird Tausch des Sensor boards empfohlen. Tausch des Sensor Boards. Einlassdichtung verschmutzt, erneuert. Startup und Kalibration mit alter SC bei pCO2 weiterhin keine Werte. Einsatz einer neuen SC. | Zusammen mit Kundin die SC90 erneut eingelegt und Meldung kam wieder, Kundin ist nicht gewillt eine weiter (bereits die Dritte) SC90 einzusetzen, Kundin bittet um einen dringenden Einsatz des zuständigen Technikers, bitte um Kontaktaufnahme. | SWO-33763 | No | According to the current Service Manual, replacement of the sensor board is recommended. Replacing the sensor board. Inlet seal dirty, renewed. Startup and calibration with old SC at pCO2 still no values. Use of a new SC. | Together with customer the SC90 re-inserted and notification came back, customer is not willing to use a further (already the third) SC90, customer asks for an urgent use of the responsible technician, please contact. | Assisted Replacement | SC | pCO2
ABL90 FLEX | Gerät bei Eintreffen Messbereit, Kunde hatte gestern Gerät gespült, Einlass geprüft: Einlassnadel teilweise nicht durchgängig , Nadel ersetzt, Kal & QC i.O., Patientenmessung: i.O. | Gerät bei Eintreffen Messbereit, Kunde hatte gestern Gerät gespült, Einlass geprüft: Einlassnadel teilweise nicht durchgängig , Nadel ersetzt, Kal & QC i.O., Patientenmessung: i.O. | SWO-33721 | No | Device ready to be measured on arrival, customer had flushed device yesterday, inlet checked: inlet needle partially not continuous, needle replaced, Kal & QC i.O., patient measurement: i.O. | Device ready to be measured on arrival, customer had flushed device yesterday, inlet checked: inlet needle partially not continuous, needle replaced, Kal & QC i.O., patient measurement: i.O. | No fault found | Probe bend
ABL825 FLEX | Sensor getauscht. | Sensor getauscht. | SWO-33680 | No | Sensor is exchanged. | Sensor is exchanged. | Defect | HW | Basic Unit | Waste detector
ABL825 FLEX | Messkammern und Kontakte gereinigt, Ref. Membran ist seitlich am Deckel undicht, Ref. Lösung tritt in die Messkammer aus. Ref. Memran nochmal erneuert, Membran bleibt dicht, Kal. wieder i.O. | Messkammern und Kontakte gereinigt, Ref. Membran ist seitlich am Deckel undicht, Ref. Lösung tritt in die Messkammer aus. Ref. Memran nochmal erneuert, Membran bleibt dicht, Kal. wieder i.O. | SWO-33728 | No | Measuring chambers and contacts cleaned, ref. Membrane is leaking on the side of the lid, ref. solution enters the measuring chamber. Ref. Memran renewed again, membrane remains tight, Kal. again i.O. | Measuring chambers and contacts cleaned, ref. Membrane is leaking on the side of the lid, ref. solution enters the measuring chamber. Ref. Memran renewed again, membrane remains tight, Kal. again i.O. | Defect | Electrode | Membrane | Reference
ABL820 FLEX | Checked calibration logs, sensitivity and status values not reading. Another electrode tried by customer. Values still not correct. | Replaced pH electrode and electrode connector. Repaired inside connector grounding tab. New electrode not reading proper signal values in diagnostics. Had to utilize older electrode from site. Electrode will be replaced on next visit with new. | CF1DB7013263428E | No | Checked calibration logs, sensitivity and status values not reading. Another electrode tried by customer. Values still not correct. | Replaced pH electrode and electrode connector. Repaired inside connector grounding tab. New electrode not reading proper signal values in diagnostics. Had to utilize older electrode from site. Electrode will be replaced on next visit with new. | Unknown | Electrode | Electrode | pH
ABL827 FLEX | Drift bei Ca und Na, Crea Status Fehler, Fr. Dr. Brocci hat die Kal und Reinigungsflaschen gestern getauscht und die Tabletten nicht aktiviert, Flaschen aktiviert, Kal durchgeführt i.O. QC bei Crea noch Fehlerhaft, Anwenderfehler erklärt, bisher angefallene Fragen beantwortet, Fr. Dr. Brocci beobachtet und tauscht die Flaschen bei anhaltendem QC Fehler noch einmal. KDB meldet sich Montag noch einmal wegen Einweisung und Anbindung bei Ihr, Anbindung noch nicht möglich da Netzwerkdosen fehlen | Drift bei Ca und Na, Crea Status Fehler, Fr. Dr. Brocci hat die Kal und Reinigungsflaschen gestern getauscht und die Tabletten nicht aktiviert, Flaschen aktiviert, Kal durchgeführt i.O. QC bei Crea noch Fehlerhaft, Anwenderfehler erklärt, bisher angefallene Fragen beantwortet, Fr. Dr. Brocci beobachtet und tauscht die Flaschen bei anhaltendem QC Fehler noch einmal. KDB meldet sich Montag noch einmal wegen Einweisung und Anbindung bei Ihr, Anbindung noch nicht möglich da Netzwerkdosen fehlen | SWO-33801 | No | Drift at Ca and Na, Crea Status Error, Fr. Dr. Brocci exchanged the Kal and Cleaning Bottles yesterday and did not activate the tablets, activated bottles, Kal performed i.O. QC at Crea still Faulty, user error explained, previously answered questions Dr. Brocci observes and exchanges the bottles again in case of a persistent QC error. KDB reports again on Monday because of instruction and connection with her, connection not yet possible because network boxes are missing | Drift at Ca and Na, Crea Status Error, Fr. Dr. Brocci exchanged the Kal and Cleaning Bottles yesterday and did not activate the tablets, activated bottles, Kal performed i.O. QC at Crea still Faulty, user error explained, previously answered questions Dr. Brocci observes and exchanges the bottles again in case of a persistent QC error. KDB reports again on Monday because of instruction and connection with her, connection not yet possible because network boxes are missing | User Error | HW | Met II Module | Solution activation
AQT90 FLEX | SE COLOCA LOS NUEVOS TUBE HOLDER, SE COLOCA LA NUEVA GUIA, SE HACEN PRUEBAS CON MUESTRAS | SE COLOCA LOS NUEVOS TUBE HOLDER, SE COLOCA LA NUEVA GUIA, SE HACEN PRUEBAS CON MUESTRAS | CF2F2163A65A9C19 | No | THE NEW TUBE HOLDER IS PLACED, THE NEW GUIDE IS PLACED, TESTS ARE MADE WITH SAMPLES | THE NEW TUBE HOLDER IS PLACED, THE NEW GUIDE IS PLACED, TESTS ARE MADE WITH SAMPLES
AQT90 FLEX | Kein Fehler feststellbar - Leihgerät aufgestellt, zur Klärung der Messerergebnisse durch doppelbestimmungen | Kein Fehler feststellbar - Leihgerät aufgestellt, zur Klärung der Messerergebnisse durch doppelbestimmungen | SWO-33778 | No | No error detectable - loaner set up to clarify knife results by double determinations | No error detectable - loaner set up to clarify knife results by double determinations
AQT90 FLEX | Data Collection Guide für D-dimer und TnT ausgefüllt Gerät beim Kunden abgeholt Nach Anleitung zur Aufarbeitung von falsch positiven Messungen Gerät aufgearbeitet Gerät beim Kunden wieder aufgebaut | Data Collection Guide für D-dimer und TnT ausgefüllt Gerät beim Kunden abgeholt Nach Anleitung zur Aufarbeitung von falsch positiven Messungen Gerät aufgearbeitet Gerät beim Kunden wieder aufgebaut | SWO-33766 | No | Data Collection Guide for D-dimer and TnT filled out device picked up at the customer After manual for processing false positive measurements device reassembled device at the customer | Data Collection Guide for D-dimer and TnT filled out device picked up at the customer After manual for processing false positive measurements device reassembled device at the customer
ABL90 FLEX | ABL zeigt ständig SP -Wechsel an. Beim Eintreffen ist der Bildschirmschoner aktiv. Im Hintergrund wird die Aufforderung zum SP-Wechsel angezeigt--> ABL neu gestartet --> Fehler nicht mehr vorhanden Uhrzeit fehlerhaft--> neu gestellt, ABL neu gestartet Medizintechnik geschult in Gerätebedienung und Fehlerbehebung | ABL zeigt ständig SP -Wechsel an. Beim Eintreffen ist der Bildschirmschoner aktiv. Im Hintergrund wird die Aufforderung zum SP-Wechsel angezeigt--> ABL neu gestartet --> Fehler nicht mehr vorhanden Uhrzeit fehlerhaft--> neu gestellt, ABL neu gestartet Medizintechnik geschult in Gerätebedienung und Fehlerbehebung | SWO-33894 | No | ABL constantly displays SP changes. When you arrive, the screen saver is active. In the background the prompt for SP change is displayed--> ABL is restarted --> Error no longer present Time faulty--> restarted, ABL restarted medical technology trained in device operation and troubleshooting | ABL constantly displays SP changes. When you arrive, the screen saver is active. In the background the prompt for SP change is displayed--> ABL is restarted --> Error no longer present Time faulty--> restarted, ABL restarted medical technology trained in device operation and troubleshooting | Defect | SW | Time setting
ABL825 FLEX | Getriebe mit Motor für Autochekmodul ersetzt. Startphase Qc und Kalibration. | Getriebe mit Motor für Autochekmodul ersetzt. Startphase Qc und Kalibration. | SWO-33946 | No | Gearbox with motor for autochek module replaced. Start phase Qc and calibration. | Gearbox with motor for autochek module replaced. Start phase Qc and calibration. | Defect | HW | AutoCheck Module | Stepper motor
ABL825 FLEX | Level1 running just below mean and above 1Sd on unit. | Tried different pO2 electrode and membrane without success. | CF1F2E2EA85EF6F1 | No | Level1 running just below mean and above 1Sd on unit. | Tried different pO2 electrode and membrane without success. | Unknown | Electrode | Electrode | pO2
ABL825 FLEX | Some ampoules are not read Pm is due on this analyzer | I tried performing the Laser calibration in the Service Mode, but it did not work. I had to perform the rough adjustment Performed Pm Package the old units to ship to Radiometer. | CF1EF109A863864A | No | Some ampoules are not read Pm is due on this analyzer | I tried performing the Laser calibration in the Service Mode, but it did not work. I had to perform the rough adjustment Performed Pm Package the old units to ship to Radiometer. | Unknown | HW | AutoCheck Module | Barcode reader
ABL825 FLEX | Leaking out of the Ph Electrode Drift errors. Pm is due on this analyzer. | Replaced the Ph gasket Cleaned analyzer Performed Pm. | CF1EEFDEA84E8645 | No | Leaking out of the Ph Electrode Drift errors. Pm is due on this analyzer. | Replaced the Ph gasket Cleaned analyzer Performed Pm. | Assisted Replacement | Electrode | Electrode | pH
ABL825 FLEX | Gerät demontiert, Einlass Modul PCB (902-781r1004n0038) ausgetauscht, Software 6.13 neu installiert, Friction Detector Unit ersetzt, Funktionskontrolle ok | Ersatzmodul installiert, Ersatzteile bestellt | SWO-33881 | No | Device dismantled, inlet module PCB (902-781r1004n0038) replaced, software 6.13 reinstalled, Friction Detector Unit replaced, function control ok | Replacement module installed, spare parts ordered | Defect | HW | Inlet module | PCB
ABL827 FLEX | Erneut alle Ventile der Messmodule demontiert, gereinigt, montiert, Pumpenschlauch PHBG- & ElMet-Modul ersetzt, RinseOutStopper am Crea-Modul ersetzt, Einlaßdichtung ersetzt, mehrere korrekte Kal abgewartet, QK ok | Reinigungsarbeiten, Fehlersuche, Ventile der Messmodule demontiert und kontrolliert, Reinigung des Lufteinlassventils am MetII-Modul, Probelauf, mehrere korrekte Kal abgewartet, QK ok | SWO-33908 | No | Again all valves of the measuring modules disassembled, cleaned, mounted, pump hose PHBG & ElMet module replaced, RinseOutStopper replaced on Crea module, inlet seal replaced, several correct cal waited, QK ok | Cleaning work, troubleshooting, valves of the measuring modules disassembled and controlled, cleaning of the air inlet valve on the MetII module, test run, several correct cal awaiting, QK ok | Defect | HW | Met II Module | Valve
ABL825 FLEX | Ran pump calibration and LS adjust, did not take away the error. Checked the service programs and found the pump calibration error on the ph bg module pump. I switched pumps with the electrolyte module pump and the error stayed with the ph bg module. | Ordering a new pH BG module PCB | CF1EC662A8223376 | No | Ran pump calibration and LS adjust, did not take away the error. Checked the service programs and found the pump calibration error on the ph bg module pump. I switched pumps with the electrolyte module pump and the error stayed with the ph bg module. | Ordering a new pH BG module PCB | Unknown | HW | pH/BG Module | Pump…
ABL825 FLEX | Examined Oxi module and checked complete Oxi Test. | Performed lamp Adj. and complete Oxi test. Performed a Service tHb cal. Broke pH electrode. Replaced pH electrode. Membraned Ref. pCO2, pO2 and Ca++ electrodes. Verified operation of ABL825. | CF1EEFCEA813D393 | No | Examined by oxygen module and checked complete oxygen Test. | Performed lamp Adj. and complete Oxi test. Performed a Service tHb cal. Broke pH electrode. Replaced pH electrode. Membraned Ref. pCO2, pO2 and Ca++ electrodes. Verified operation of ABL825. | Adjustment | HW | Oxi module | Lamp unit
ABL827 FLEX | Halogenlampe justiert Oxi complete Test durchgeführt Service Thb durchgeführt Analysator kalibriert und überprüft Fcuv QC Messung durchgeführt QC Level 2 Barcode eingelesen Key eingespielt | Halogenlampe justiert Oxi complete Test durchgeführt Service Thb durchgeführt Analysator kalibriert und überprüft Fcuv QC Messung durchgeführt QC Level 2 Barcode eingelesen Key eingespielt | SWO-33854 | No | Halogen lamp adjusted Oxi complete test performed service Thb performed analyzer calibrated and verified Fcuv QC measurement performed QC Level 2 barcode read key imported | Halogen lamp adjusted Oxi complete test performed service Thb performed analyzer calibrated and verified Fcuv QC measurement performed QC Level 2 barcode read key imported | Adjustment | HW | Oxi module | Lamp unit
AQT90 FLEX | Lubricacion del muelle de la puerta. Chequeo con programa de servicio. Comprobacion con qc. |  | CF1F0E6EA8261549 | No | Lubrication of the door spring. Check with service program. Check with qc.
ABL80 FLEX CO-OX | Spülung Fließwege, Reinigungsverfahren, Oxi-Test, tHb-Kal | Spülung Fließwege, Reinigungsverfahren, Oxi-Test, tHb-Kal | SWO-33957 | No | Flushing flow paths, cleaning procedures, Oxi test, tHb-Kal | Flushing flow paths, cleaning procedures, Oxi test, tHb-Kal
ABL90 FLEX | Beim Eintreffen steht der ABL messbereit da mit Kal.-Fehler bei Ca und QC-Fehler bei Lac und Ca. Kal. durchgeführt --> Ca-Fehler bleibt bestehen. Neue SC inst. Station führt selbstständig notwendige Kal. und QC durch | Beim Eintreffen steht der ABL messbereit da mit Kal.-Fehler bei Ca und QC-Fehler bei Lac und Ca. Kal. durchgeführt --> Ca-Fehler bleibt bestehen. Neue SC inst. Station führt selbstständig notwendige Kal. und QC durch | SWO-33962 | No | Upon arrival, the ABL is ready for measurement with Cal. errors in Ca and QC errors at Lac and Ca. Kal. --> Ca error remains. New SC inst. Station independently performs necessary Cal. and QC by | Upon arrival, the ABL is ready for measurement with Cal. errors in Ca and QC errors at Lac and Ca. Kal. --> Ca error remains. New SC inst. Station independently performs necessary Cal. and QC by | No fault found | SC | Wetability
ABL80 FLEX CO-OX | ABL auf technische Mängel überprüft und im Menü sämtliche Test zur Fehlerdiagnose durchgeführt. Alles i.o. Neue SC Installiert und Vergleichsmessungen mit Herr Urs Stolz und Herr Mauros Curti durchgeführt. Alle i.o. System und QA Zyklen i.o. Funktionskontrolle i.o. | Sämtliche Ventile und Pumpen überprüft. Sämtliche Schläuche überprüft und durchgespült. Abfallpumpe und Abfallschläuche ersetzt. System und QA Zyklus i.o. Funktionskontrolle i.o. | SWO-34180 | No | ABL checked for technical defects and carried out all tests for error diagnosis in the menu. Everything i.o. New SC Installed and comparative measurements carried out with Mr. Urs Stolz and Mr. Mauros Curti. All i.o. system and QA cycles i.o. function control i.o. | All valves and pumps checked. All hoses are checked and rinsed through. waste pump and waste hoses. System and QA cycle i.o. function control i.o.
ABL80 FLEX CO-OX | Nocheinmal pO2 Nullpunkt kalibriert. OC Lösung Level 3 i.o Neue Sensorkassette vom 22.06.15 ohne Erfolg reinstalliert. pH und pCO2 Fehler. Alte Sensorkassette zum testen wieder initialisiert. I.o jedoch keine Messungen mehr verfügbar. 3. Kundenkassette initialisiert. Gleicher Fehler wie mit 2. Sensorkassette aufgetreten. Komplette Front ausgetauscht. 3. Sensor mit Wasser gespült und nocheinmal initialisiert. I.o | Nocheinmal pO2 Nullpunkt kalibriert. OC Lösung Level 3 i.o Neue Sensorkassette vom 22.06.15 ohne Erfolg reinstalliert. pH und pCO2 Fehler. Alte Sensorkassette zum testen wieder initialisiert. I.o jedoch keine Messungen mehr verfügbar. 3. Kundenkassette initialisiert. Gleicher Fehler wie mit 2. Sensorkassette aufgetreten. Komplette Front ausgetauscht. 3. Sensor mit Wasser gespült und nocheinmal initialisiert. I.o | SWO-34096 | No | Again pO2 zero point calibrated. OC solution Level 3 i.o New sensor cartridge from 22.06.15 reinstalled without success. pH and pCO2 errors. Old sensor cartridge for testing re-initialized. I.o, however, no more measurements available. 3. Customer cartridge initialized. Same error as with 2nd sensor cartridge occurred. Complete front replaced. 3. Sensor rinsed with water and initialized again. Ok | Again pO2 zero point calibrated. OC solution Level 3 i.o New sensor cartridge from 22.06.15 reinstalled without success. pH and pCO2 errors. Old sensor cartridge for testing re-initialized. I.o, however, no more measurements available. 3. Customer cartridge initialized. Same error as with 2nd sensor cartridge occurred. Complete front replaced. 3. Sensor rinsed with water and initialized again. Ok
ABL80 OSM | All oxi values were questioned on plood samples. | replaced the Hemolyzer. | CF22BC4C0A32405E | No | All oxi values were questioned on plood samples. | replaced the Hemolyzer.
ABL90 FLEX | Effectué la maintenance de l'analyseur (remplacé toutes les conduites souples) e.o Test de la section fluidique à l'aide du programme de service e.o Calibration e.o contrôle de qualité e.o | Effectué la maintenance de l'analyseur (remplacé toutes les conduites souples) e.o Test de la section fluidique à l'aide du programme de service e.o Calibration e.o contrôle de qualité e.o | SWO-34192 | No | Performed the maintenance of the analyzer (replaced all flexible pipes) e.o Testing the fluid section using the service program e.o Calibration e.o quality control e.o | Performed the maintenance of the analyzer (replaced all flexible pipes) e.o Testing the fluid section using the service program e.o Calibration e.o quality control e.o | Assisted Replacement | SC | Wetability
ABL90 FLEX | Fehlersuche: LP ausgebaut und untersucht => LP wurde nicht richtig aktiviert, Deckel nicht komplett aufgedrückt! => Deckel vollständig heruntergedrückt und LP erneut installiert, Startverfahren, Kalibrationen und drei QC Messungen => alle OK Restbestand unbekannter Kapillaren entsorgt und auf die ausschließliche Verwendung von =R= safeClinitubes hingewiesen | Der Kunde hat das Gerät neu gebootet und den LP noch einmal installiert. Das Gerät zeigt Startphase 55Min. an. Der Kunde meldet sich, wenn der Fehler wieder angezeigt wird. | SWO-34170 | No | Troubleshooting: LP removed and examined => LP was not activated correctly, lid not completely pressed! => Lid completely pressed down and LP reinstalled, start method, calibrations and three QC measurements => all OK remaining stock of unknown capillaries disposed of and indicated the exclusive use of =R= safeClinitubes | The customer has rebooted the device and reinstalled the LP. The device displays start-up phase 55min. The customer reports when the error is displayed again. | User Error | SP activation
ABL90 FLEX | Fehlersuche LP entfernt und System durchgespült Systemüberprüfung von Flows durchgeführt OK Einlassdichtung getauscht Analysator kalibriert und überprüft Fcuv 99,2% Interne QC OK | Fehlersuche LP entfernt und System durchgespült Systemüberprüfung von Flows durchgeführt OK Einlassdichtung getauscht Analysator kalibriert und überprüft Fcuv 99,2% Interne QC OK | SWO-34087 | No | Troubleshooting LP removed and system flushed system check performed by flows OK inlet seal exchanged analyzer calibrated and checked Fcuv 99.2% Internal QC OK | Troubleshooting LP removed and system flushed system check performed by flows OK inlet seal exchanged analyzer calibrated and checked Fcuv 99.2% Internal QC OK | Cleaned | Fluidic pathway
ABL90 | Cals fine, internal QC are in. Passes level 1 and 2 external QC but not level 3(S7750) | performed manual back flush with warm water and performed scrubbing cleaning. Still will abort only level 3 external QC or reads it but questions pH, PCO2 and PO2. Steve ordered new sensor and SP overnight. | CF22EE2A0A1D7B9B | No | Cals fine, internal QC are in. Passes level 1 and 2 external QC but not level 3(S7750) | performed manual back flush with warm water and performed scrubbing cleaning. Still will abort only level 3 external QC or reads it but questions pH, PCO2 and PO2. Steve ordered new sensor and SP overnight. | Assisted Replacement | SC | Ext. QC
ABL90 FLEX | SC SP unauffällig tHb Kal durchgeführt. QC wiederholt ok | SC SP unauffällig tHb Kal durchgeführt. QC wiederholt ok | SWO-34183 | No | SC SP inconspicuously tHb Kal performed. QC repeats ok | SC SP inconspicuously tHb Kal performed. QC repeats ok | Adjustment | tHb cal
ABL90 FLEX | SC lid closed Indikator ersetzt. SC-Fach auf korrekte Funktion überprüft. | SC lid closed Indikator ersetzt. SC-Fach auf korrekte Funktion überprüft. | SWO-34150 | No | SC lid closed indicator replaced. SC tray checked for correct function. | SC lid closed indicator replaced. SC tray checked for correct function. | Defect | HW | SC Door
ABL825 FLEX | 237 Spectrometer voltage bias high. | Replaced spectrometer and aligned Oximeter system. | CF22B5940A18CD5A | No | 237 Spectrometer voltage bias high. | Replaced spectrometer and aligned Oximeter system. | Defect | HW | Oxi module | Spectrometer
ABL825 FLEX | System überprüft, Einlaßsonde wurde bereits durch den Anwender gewechselt, Flex-Q-Modul überprüft und komplett justiert, Laser justiert, Qualitätskontrollen Eingescannt, System OK. Mit Anwender zu bestellende Materialien besprochen. | System überprüft, Einlaßsonde wurde bereits durch den Anwender gewechselt, Flex-Q-Modul überprüft und komplett justiert, Laser justiert, Qualitätskontrollen Eingescannt, System OK. Mit Anwender zu bestellende Materialien besprochen. | SWO-34153 | No | System checked, inlet probe has already been changed by the user, Flex-Q module has been checked and completely adjusted, laser adjusted, quality controls scanned, system OK. Materials to be ordered with users. | System checked, inlet probe has already been changed by the user, Flex-Q module has been checked and completely adjusted, laser adjusted, quality controls scanned, system OK. Materials to be ordered with users. | Defect | HW | Inlet module | Probe bend
ABL825 FLEX | Ersatz Oxi Modul eingebaut Spare Parts bestellt. | Ersatz Oxi Modul eingebaut Spare Parts bestellt. | SWO-34182 | No | Replacement Oxi module built-in spare parts ordered. | Replacement Oxi module built-in spare parts ordered. | Defect | HW | Oxi module | Lamp unit
ABL825 FLEX | Hemolysator ersetzt. Justage durchgeführt. Oxi Modul wieder zurückgebaut. 22.06.15 | Hemolysator ersetzt. Justage durchgeführt. Oxi Modul wieder zurückgebaut. 22.06.15 | SWO-34184 | No | Hemolysator. adjustment performed. Oxi module dismantled again. 22.06.15 | Hemolysator. adjustment performed. Oxi module dismantled again. 22.06.15 | Defect | HW | Oxi module | Hemolyzer
ABL825 FLEX | Analysator neu gestartet, Oxi Modul demontiert, Lampeneinheit justiert, tHb Kalibration, Funktionskontrolle ok |  | SWO-34373 | No | Analyzer restarted, Oxi module dismantled, lamp unit adjusted, tHb calibration, function control ok |  | Adjustment | HW | Oxi module | Lamp unit
ABL827 FLEX | ABL netzwerktechnisch angebunden. Verbindung mit Haus-IT (Hr. Gorissen 02151 991533) und IMP-Computersysteme (Hr. Henning 030 75472-200) getestet. Geräte-IP: 192.168.97.51 Sub-Netz: 255.255.255.0 Gateway: 192.168.97.1 Poctopus: High Level: HL7 2,5, Low Level: Netzwerk (TCP IP), IP: 192.168.98.20, Port: 9219 Crea A+B, GLU und LAC neu membraniert. Gerät einkalibriert QC's durchgeführt. | ABL netzwerktechnisch angebunden. Verbindung mit Haus-IT (Hr. Gorissen 02151 991533) und IMP-Computersysteme (Hr. Henning 030 75472-200) getestet. Geräte-IP: 192.168.97.51 Sub-Netz: 255.255.255.0 Gateway: 192.168.97.1 Poctopus: High Level: HL7 2,5, Low Level: Netzwerk (TCP IP), IP: 192.168.98.20, Port: 9219 Crea A+B, GLU und LAC neu membraniert. Gerät einkalibriert QC's durchgeführt. | SWO-34187 | No | ABL network-related. tested with House IT (Mr. Gorissen 02151 991533) and IMP computer systems (Mr. Henning 030 75472-200). Device IP: 192.168.97.51 Sub-network: 255.255.255.0 Gateway: 192.168.97.1 Poctopus: High Level: HL7 2.5, Low Level: Network (TCP IP), IP: 192.168.98.20, Port: 9219 Crea A+B, GLU and re-membrane. device calibrated QC's performed. | ABL network-related. tested with House IT (Mr. Gorissen 02151 991533) and IMP computer systems (Mr. Henning 030 75472-200). Device IP: 192.168.97.51 Sub-network: 255.255.255.0 Gateway: 192.168.97.1 Poctopus: High Level: HL7 2.5, Low Level: Network (TCP IP), IP: 192.168.98.20, Port: 9219 Crea A+B, GLU and re-membrane. device calibrated QC's performed. | Setup | SW | His/lis
ABL835 FLEX | Roma.ini - Eintrag (49=1) in BCR Parameter eingetragen - Funktionsüberprüfung etc. | Roma.ini - Eintrag (49=1) in BCR Parameter eingetragen - Funktionsüberprüfung etc. | SWO-34055 | No | Roma.ini - entry (49=1) entered in BCR parameters - function check etc. | Roma.ini - entry (49=1) entered in BCR parameters - function check etc. | Setup | roma.ini | - | Scanner funtion
ABL835 FLEX |  |  | CF23A8748051F36A | No |  |  | Defect | HW | Oxi module | Spectrometer
ABL825 FLEX | Geräte Konfig. vorgenommen | Geräte Konfig. vorgenommen | SWO-34121 | No | Devices Config. Made | Devices Config. Made | Setup | SW | - | QA Portal
AQT90 FLEX |  |  | CF2283AB0A2369F4 | No
AQT90 FLEX | AQT netzwerktechnisch angebunden. Verbindung mit Haus-IT (Hr. Gorissen 02151 991533) und IMP-Computersysteme (Hr. Henning 030 75472-200) getestet. Gerätename: 838R0360N006 IP: 192.168.97.50 Subnet-Mask: 255.255.255.0 Gateway: 192.168.97.1 Poctopus: High Level: HL7 2,5, IP: 192.168.98.20 Port: 9219 Commit-to-C-Drive Befehl ausgeführt. | AQT netzwerktechnisch angebunden. Verbindung mit Haus-IT (Hr. Gorissen 02151 991533) und IMP-Computersysteme (Hr. Henning 030 75472-200) getestet. Gerätename: 838R0360N006 IP: 192.168.97.50 Subnet-Mask: 255.255.255.0 Gateway: 192.168.97.1 Poctopus: High Level: HL7 2,5, IP: 192.168.98.20 Port: 9219 Commit-to-C-Drive Befehl ausgeführt. | SWO-34186 | Yes | AQT network-connected. tested with House IT (Mr. Gorissen 02151 991533) and IMP computer systems (Mr. Henning 030 75472-200). Device Name: 838R0360N006 IP: 192.168.97.50 Subnet Mask: 255.255.255.0 Gateway: 192.168.97.1 Poctopus: High Level: HL7 2.5, IP: 192.168.98.20 Port: 9219 Commit-to-C-Drive Executed. | AQT network-connected. tested with House IT (Mr. Gorissen 02151 991533) and IMP computer systems (Mr. Henning 030 75472-200). Device Name: 838R0360N006 IP: 192.168.97.50 Subnet Mask: 255.255.255.0 Gateway: 192.168.97.1 Poctopus: High Level: HL7 2.5, IP: 192.168.98.20 Port: 9219 Commit-to-C-Drive Executed. | Re classified RMED | According to JMP-email, not upSWO | Setup
ABL80 CO-OX |  |  | CF241E1480362478 | No
ABL80 FLEX CO-OX | Sämtliche Ventile und Pumpen überprüft. Sämtliche Schläuche überprüft und durchgespült. Abfallpumpe und Abfallschläuche ersetzt. System und QA Zyklus i.o. Funktionskontrolle i.o. | Sämtliche Ventile und Pumpen überprüft. Sämtliche Schläuche überprüft und durchgespült. Abfallpumpe und Abfallschläuche ersetzt. System und QA Zyklus i.o. Funktionskontrolle i.o. | SWO-34326 | No | All valves and pumps checked. All hoses are checked and rinsed through. waste pump and waste hoses. System and QA cycle i.o. function control i.o. | All valves and pumps checked. All hoses are checked and rinsed through. waste pump and waste hoses. System and QA cycle i.o. function control i.o.
ABL80 CO-OX | unit working on arrivel, will run Zero cal test cassette. | ran Zero cal test cassette.all test passed was in spec allready. ran System cycle after and all parameters came in fine, the intermittent pc02 failure is most likely a weak Solution Pack, QC is right on the edge, suggest replacing it and getting credit. | CF24044C80021331 | No | unit working on arrivel, will run Zero cal test cassette. | ran Zero cal test cassette.all test passed was in spec allready. ran System cycle after and all parameters came in fine, the intermittent pc02 failure is most likely a weak Solution Pack, QC is right on the edge, suggest replacing it and getting credit.
ABL90 FLEX | Systemspülungen durchgeführt, Gerät getestet | Systemspülungen durchgeführt, Gerät getestet | SWO-34421 | No | System flushing performed, device tested | System flushing performed, device tested | Assisted Replacement | SC | pCO2
ABL835 FLEX | Comprobar módulo OXI, enfocar la lámpara, calibrar la tHb en el programa de servicio, | Enfocar la lámpara y calibrar la tHb con el programa de servicio. | CF23CE398022C0D8 | No | Check OXI module, focus the lamp, calibrate the tHb in the service program, | Focus the lamp and calibrate the tHb with the service program. | Adjustment | HW | Oxi module | Lamp unit
ABL825 FLEX | 902-833 OXI Modul ausgetauscht. Probelauf, QCs und Blutmessung i.O. | der Oxymetrie Fehler besteht weiterhin; der Anwender wünscht einen Vorort-Service zur Beseitigung der Geräte-Störung; DRINGEND! | SWO-34511 | No | 902-833 OXI module replaced. Test run, QCs and blood measurement in the order of | the oxymetry error persists; the user wants a suburban service to eliminate the device malfunction; urgent! | Defect | HW | Oxi module | Hemolyzer phase regulation
ABL800 BASIC | Messkammer und Schlauchsystem geprüft, Referenz Elektrode (E1001 67-08) ausgetauscht, Funktionskontrolle ok |  | SWO-34374 | No | Measuring chamber and hose system tested, reference electrode (E1001 67-08) replaced, function control ok |  | Assisted Replacement | Electrode | Electrode | Reference
ABL825 FLEX |  |  | CF2452E1801C0DDA | No |  |  | Adjustment | HW | Oxi module | Service Cal(Oxi)
ABL827 FLEX | checked gas regulators and pressure transducer | removed gas bottles and reset regulators, removed and cleaned pressure transducer, | CF24752A8019D898 | No | checked gas regulators and pressure transducer | removed gas bottles and reset regulators, removed and cleaned pressure transducer, | Cleaned | HW | Basic Unit | Gas module
ABL835 FLEX | Comparisons are off more than what is acceptable. | Freed stuck air valves in pH BG module and El Met module. | CF23DFB7801F30E3 | No | Comparisons are off more than what is acceptable. | Freed stuck air valves in pH BG module and El Met module. | Defect | HW | pH/BG Module | Valve
AQT90 FLEX |  |  | CF24021A805B0DD6 | No
AQT90 FLEX | Spannungen, Temperatur und Uhrzeit laut Service und Windows i.O. Test mit Cleaning Solution Röhrchen => Gerät bricht nach Öffnen der Tür und Einlegen des Röhrchens ab => Barcode wird nicht ausgelesen und Tür bleibt geöffnet Überprüfung Inlet Door => lässt sich über Service ohne Probleme Öffnen und Schließen Überprüfung Barcodescanner => über Vision Control lässt sich Barcode auslesen Nochmaliger Test => nach Einlegen des Röhrchens kurzes Aufleuchten vom Scanner, anschließend Fehler 5766 SETINSTRUMENTFANFAILED => Gerät muss neugestartet werden Gerät zur Reparatur mitgenommen | Kundin bittet um einen Einsatz des zuständigen Technikers, bitte um Kontaktaufnahme. | SWO-34290 | No | Voltages, temperature and time according to service and Windows i.O. Test with Cleaning Solution tubes => Device breaks off after opening the door and inserting the tube => Barcode is not read and door remains open Check Inlet Door => can be checked via service without Problems Opening and closing checking Barcode scanner => via Vision Control can be read barcode Again test => after inserting the tube briefly light up from the scanner, then error 5766 SETINSTRUMENTFANFAILED => device must be restarted Device taken with him for repair | Customer asks for the use of the responsible technician, please contact us.
ABL80 FLEX CO-OX | ABL steht mit Fehler --> Flüssigkeitsstrecke gespült und gereinigt--> Neustart Kal.-Fehler --> JW durchgeführt--> Fehler bleibt bestehen ABL zur Rep. mitgenommen | ABL steht mit Fehler --> Flüssigkeitsstrecke gespült und gereinigt--> Neustart Kal.-Fehler --> JW durchgeführt--> Fehler bleibt bestehen ABL zur Rep. mitgenommen | SWO-34553 | No | ABL is rupeated with error --> Liquid track flushed and cleaned--> Restart Cal.-Error --> JW performed--> Error remains ABL to Rep. Taken | ABL is rupeated with error --> Liquid track flushed and cleaned--> Restart Cal.-Error --> JW performed--> Error remains ABL to Rep. Taken
ABL90 | Entered the service program and performed all the fluid transport test, and they all passed. Removed the solution pack and flushed the fluid path. Customer had already replaced the sensor cassette, so I removed and reinstalled the sensor cassette, but this did not help. | Replaced the inlet probe, but this did not resolve the problem. Replaced the sensor cassette (Supplied by customer), and this resolved the problem. | CF2529FEF61DDCD9 | No | Entered the service program and performed all the fluid transport test, and they all passed. Removed the solution pack and flushed the fluid path. Customer had already replaced the sensor cassette, so I removed and reinstalled the sensor cassette, but this did not help. | Replaced the inlet probe, but this did not resolve the problem. Replaced the sensor cassette (Supplied by customer), and this resolved the problem. | Assisted Replacement | SC | Wetability
ABL90 | Troubleshot and found that the time in Windows was an hour off of the time set in the analyzer software. | Corrected the time in Windows, and this resolved the problem. | CF2571A2F651F7CE | No | Troubleshot and found that the time in Windows was an hour off of the time set in the analyzer software. | Corrected the time in Windows, and this resolved the problem. | Defect | SW | Time setting
ABL90 | In the control panel, daylight savings time was checked. | Followed the Daylight Savings Service Tip 663572096 to correct the problem. | CF2564D3F637DCDD | No | In the control panel, daylight savings time was checked. | Followed the Daylight Savings Service Tip 663572096 to correct the problem. | Defect | SW | Time setting
ABL90 | Battery makes a loud noise when connected | Replaced battery Brought system to ready | CF251ABCF6363E62 | No | Battery makes a loud noise when connected | Replaced battery Brought system to ready | Defect | HW | Battery module
ABL90 | Confirmed errors. | Replaced the above parts. | CF256DC5F62CB10F | No | Confirmed errors. | Replaced the above parts. | Defect | HW | Cassette retainer
ABL825 FLEX | Inspected calibration and found halogen lamp intensity set to high. | Performed lamp unit adjustment. Re-membraned PO2 electrode as it was due. Performed necessary calibrations and qc to get analyzer back in to working order. | CF2533F0F62708AB | No | Inspected calibration and found halogen lamp intensity set to high. | Performed lamp unit adjustment. Re-membraned PO2 electrode as it was due. Performed necessary calibrations and qc to get analyzer back in to working order. | Adjustment | HW | Oxi module | Lamp unit
ABL827 FLEX | checked gas regulators and pressure transducer | actual Transducer voltage high, removed and replaced pressure transducer with chassis and analyzer control board, installed option keys and reagents, performed an update type, recalibrated hemolyzer, complete oxi test and tHb cal | CF25B97BF657D8A4 | No | checked gas regulators and pressure transducer | actual Transducer voltage high, removed and replaced pressure transducer with chassis and analyzer control board, installed option keys and reagents, performed an update type, recalibrated hemolyzer, complete oxi test and tHb cal | Defect | HW | Analyzer Control PCB | PCB
ABL810 FLEX | V2-Ventil ausgebaut, zerlegt und erneut eingesetzt. Befüllung und -Lüftung in Ordnung. | V2-Ventil ausgebaut, zerlegt und erneut eingesetzt. Befüllung und -Lüftung in Ordnung. | SWO-34558 | No | V2 valve removed, disassembled and re-inserted. Filling and ventilation in order. | V2 valve removed, disassembled and re-inserted. Filling and ventilation in order. | Defect | HW | Fluidic manifold | Valve
AQT90 FLEX | Einlassnadel inkl. Nadel-Spritzschutz über die Serviceebene erneuert, Systemüberprüfung und Systemreinigung durchgeführt, System- und Funktionskontrolle, System und Funktion i.O. | Einlassnadel inkl. Nadel-Spritzschutz über die Serviceebene erneuert, Systemüberprüfung und Systemreinigung durchgeführt, System- und Funktionskontrolle, System und Funktion i.O. | SWO-34451 | No | Inlet needle incl. needle spray protection renewed via the service level, system inspection and system cleaning carried out, system and function control, system and function in the following way. | Inlet needle incl. needle spray protection renewed via the service level, system inspection and system cleaning carried out, system and function control, system and function in the following way.
AQT90 FLEX | Fehlersuche Gerät zur Reparatur in Werkstatt mitgenommen |  | SWO-34599 | No | Troubleshooting device taken to workshop for repair
ABL80 FLEX CO-OX | Austausch des Gerätes gegen anderen ABL80 COOx 311647, Konfig. eingespielt, Gerät an Haus EDV angebunden, Testmessung: i.O. Zählerstand bei Lieferung 0, einen Satz ZVBM da gelassen | ABL steht mit Fehler --> Flüssigkeitsstrecke gespült und gereinigt--> Neustart Kal.-Fehler --> JW durchgeführt--> Fehler bleibt bestehen ABL zur Rep. mitgenommen | SWO-34574 | No | Replacement of the device with other ABL80 COOx 311647, config. recorded, device connected to house COMPUTER, test measurement: i.O. meter reading at delivery 0, a set of ZVBM left there | ABL is rupeated with error --> Liquid track flushed and cleaned--> Restart Cal.-Error --> JW performed--> Error remains ABL to Rep. Taken
ABL80 OSM | Cleaned entire fluid path.Reinitized unit No fault. | Cleaning. Left SP for customer, they did not have Spare. | CF2DBF46303DF053 | No | Cleaned entire fluid path. Reinitized unit No fault. | Cleaning. Left SP for customer, they did not have Spare.
ABL90 FLEX | Fehlersuche, div. Komponenten ausgetauscht, Startphase abgewartet = pCO2-, Cl- u. Glu-Fehler, SC erneut initialisiert -> weiter Fehler, neue SC eingesetzt, wieder Startphase abgewartet, Kal. + QC i.O. | Fehlersuche; div. Komponenten ausgetauscht; Startphase abgewartet = pCO2-, Cl- u. Glu-Fehler; SC erneut initialisiert -> weiter Fehler; neue SC eingesetzt; wieder Startphase abgewartet; Kal. + QC i.O. | SWO-34627 | No | Troubleshooting; various components replaced; Start-up phase waited for = pCO2, Cl and Glu errors; SC reinitialized -> further error; new SC; again waited for the start phase; Kal. + QC in the first place | Troubleshooting; various components replaced; Start-up phase waited for = pCO2, Cl and Glu errors; SC reinitialized -> further error; new SC; again waited for the start phase; Kal. + QC in the first place | Defect | HW | Tubings
ABL90 |  |  | CF265AF16C3D7D33 | No |  |  | Assisted Replacement | SC | pH | LEH
ABL90 | Instrument was running when FSE showed up. | Customer performed several flushes and it fixed the error. | CF26561A6C4BE3AC | No | Instrument was running when FSE showed up. | Customer performed several flushes and it fixed the error. | No fault found | Customer cleaned FT
ABL837 FLEX | Upon arrival, found unit failing sensitivity 3 on low side. | Cleaned out all Dummy chambers in Crea module. Cleaned and remembraned Crea electrodes. | CF26384A6C0501D9 | No | Upon arrival, found unit failing sensitivity 3 on low side. | Cleaned out all Dummy chambers in Crea module. Cleaned and remembraned Crea electrodes. | Cleaned | Fluidic pathway | Met II Module | Crea
ABL820 FLEX | Checked QC setup. QC's set to run manual instead of auto. | Changed QC to run auto. | CF26AD446C07287E | No | Checked QC setup. QC's set to run manual instead of auto. | Changed QC to run auto. | Setup | QC Schedule
ABL800 BASIC | alle Elektroden kontrolliert, Na- Elektrode war nicht richtig membraniert, pO2Elektrode nicht richtig membraniert, | alle Elektroden kontrolliert, Na- Elektrode war nicht richtig membraniert, pO2Elektrode nicht richtig membraniert, | SWO-34733 | No | controlled all electrodes, Na-electrode was not properly membraneed, pO2electrode not properly membraneed, | controlled all electrodes, Na-electrode was not properly membraneed, pO2electrode not properly membraneed, | Assisted Replacement | Electrode | Electrode | User error
ABL835 FLEX | Laktat Elektrode (E7077 RP01) geprüft und ausgetauscht, Funktionskontrolle ok |  | SWO-34813 | No | Lactate electrode (E7077 RP01) tested and replaced, function control ok |  | Assisted Replacement | Electrode | Electrode | Lac
ABL825 FLEX | checked pumps tubings | replaced pump tubing on ph bg module, performed pump calibration. | CF264BA06C21B15C | No | checked pumps tubings | replaced pump tubing on ph bg module, performed pump calibration. | Defect | HW | pH/BG Module | Pump tubing
ABL825 FLEX | Lampenwert hat sich auf über 4000 verstellt, neu justiert auf knapp 3200, Complete Oxymetrie Test i.o., Kal, Service tHb Kal, Kal, Anwender tHb-Kal, QC | Lampenwert hat sich auf über 4000 verstellt, neu justiert auf knapp 3200, Complete Oxymetrie Test i.o., Kal, Service tHb Kal, Kal, Anwender tHb-Kal, QC | SWO-34699 | No | Lamp value has been adjusted to over 4000, re-adjusted to just under 3200, Complete Oxymetry Test i.o., Kal, Service tHb Kal, Kal, User tHb-Kal, QC | Lamp value has been adjusted to over 4000, re-adjusted to just under 3200, Complete Oxymetry Test i.o., Kal, Service tHb Kal, Kal, User tHb-Kal, QC | Adjustment | HW | Oxi module | Lamp unit
AQT90 FLEX | Fehlersuche am Analysator, TnI Testkassetten lassen sich nicht einlesen->Barcodereader und Analysator gereinigt, 3Testkassetten eingelegt ok(Staubschutz wird bei der JW angebracht->heute nicht möglich) System-und Funktionskontrolle i.O. | Fehlersuche am Analysator, TnI Testkassetten lassen sich nicht einlesen->Barcodereader und Analysator gereinigt, 3Testkassetten eingelegt ok(Staubschutz wird bei der JW angebracht->heute nicht möglich) System-und Funktionskontrolle i.O. | SWO-34642 | No | Troubleshooting on the analyzer, TnI test cassettes cannot be read->barcode reader and analyzer cleaned, 3 test cassettes inserted ok(dust protection is not possible with the JW->not possible today) System and function control i.O. | Troubleshooting on the analyzer, TnI test cassettes cannot be read->barcode reader and analyzer cleaned, 3 test cassettes inserted ok(dust protection is not possible with the JW->not possible today) System and function control i.O.
AQT90 FLEX | Fehlersuche PCB für Incubator Wheel erneuert Blankmessungen LQC-Messungen nicht möglich, da Kunde kein Material vor Ort hat (ist bestellt, aber noch nicht geliefert) Gerät wieder beim Kunden aufgestellt |  | SWO-34600 | No | Troubleshooting PCB for Incubator Wheel renews blank measurements LQC measurements not possible, as customer has no material on site (is ordered but not yet delivered) device re-installed at the customer
ABL80 FLEX CO-OX | Sensor Kassette durch Kundeneigene ersetzt. Frau Geck demonstriert wie das Radiometer Retour Formular ausgefüllt wird. | Sensor Kassette durch Kundeneigene ersetzt. Frau Geck demonstriert wie das Radiometer Retour Formular ausgefüllt wird. | SWO-35588 | No | Sensor cassette replaced by customer's own. Ms. Geck demonstrates how to fill out the Radiometer Return Form. | Sensor cassette replaced by customer's own. Ms. Geck demonstrates how to fill out the Radiometer Return Form.
ABL90 |  |  | CF27D14CE24C0EDD | No |  |  | Assisted Replacement | SC | Wetability | LEH
ABL90 FLEX | Angaben unter Betriebsmittel: installationsdatum 10.07. Haltbar bis 08.08. heute ist 26.06 :-) Eine neue Sensorkassette installiert. status hat sich wieder korrigiert. Der eigentliche Fehler kann erst in 8 Std. abgefragt werden. | Angaben unter Betriebsmittel: installationsdatum 10.07. Haltbar bis 08.08. heute ist 26.06 :-) Eine neue Sensorkassette installiert. status hat sich wieder korrigiert. Der eigentliche Fehler kann erst in 8 Std. abgefragt werden. | SWO-34877 | No | Details under equipment: installationdate 10.07. Durable until 08.08. today is 26.06 :-) A new sensor cartridge installed. status has corrected itself again. The actual error can only be queried in 8 hrs. | Details under equipment: installationdate 10.07. Durable until 08.08. today is 26.06 :-) A new sensor cartridge installed. status has corrected itself again. The actual error can only be queried in 8 hrs. | Assisted Replacement | SC | Date setting
ABL90 FLEX | Konfiguration angepsst | Konfiguration angepsst | SWO-34850 | No | Configuration plugped | Configuration plugped | Setup | HIS/LIS
ABL90 FLEX | Erneuter Tausch des Sensor Boards, da eingebautes Sensor Board erst vor ein paar Tagen eingebaut. Kalibrationen und QC o.K. QC Wert bei pCO2 grenzwertig und manchmal außer Bereich. Anwender empfohlen die SC zu erneuern. Analysator messbereit. | Der Kunde wünscht keine telefonische Unterstützung. Der Kunde bittet um eine Reparatur. | SWO-34940 | No | Replacement of the sensor board again, since built-in sensor board only installed a few days ago. Calibrations and QC o.K. QC value at pCO2 borderline and sometimes out range. recommended for the SC. Analyzer ready to measure. | The customer does not want telephone support. The customer asks for a repair. | Defect | HW | SC PCB
ABL825 FLEX | pH is running low outside customer's qc level 3 range. pH: 7.566 Customer Ranges: 7.567-7.579 | Replaced the pH electrode blue seal ring. Checked the QA portal 2 SD ranges. The pH mean is 7.570 and the ranges 7.565-7.574 Recommended the customer to revise and readjust her ranges. | CF27B77CE20165BF | No | pH is running low outside customer's qc level 3 range. pH: 7.566 Customer Ranges: 7.567-7.579 | Replaced the pH electrode blue seal ring. Checked the QA portal 2 SD ranges. The pH mean is 7.570 and the ranges 7.565-7.574 Recommended the customer to revise and readjust her ranges. | Assisted Replacement | Electrode | Electrode | Sealing Ring
ABL835 FLEX | Messkammer und Schlauchsystem geprüft, diverse Schläuche ausgetauscht, Modulventile V1-V4 demontiert und geprüft, pCO2 und Cl Membrane erneuert, Funktionskontrolle ok | Messkammer und Schlauchsystem geprüft, diverse Schläuche ausgetauscht, Modulventile V1-V4 demontiert und geprüft, pCO2 und Cl Membrane erneuert, Funktionskontrolle ok | SWO-34915 | No | Measuring chamber and hose system tested, various hoses replaced, module valves V1-V4 disassembled and tested, pCO2 and Cl membrane renewed, function control ok | Measuring chamber and hose system tested, various hoses replaced, module valves V1-V4 disassembled and tested, pCO2 and Cl membrane renewed, function control ok | Assisted Replacement | Electrode | Membrane | pCO2
ABL90 | No errors in daiganostics. Techline suggested replacement of Hemolyzer. | Replaced Hemolyzer Assy. | CF2B902D444D5452 | No | No errors in daiganostics. Techline suggested replacement of Hemolyzer. | Replaced Hemolyzer Assy. | Assisted Replacement | Electrode | Electrode | Lac
ABL90 FLEX | De hemolyser werd preventief vervangen. tHb-kalibratie en gewone kalibratie:OK In het serviceprogramma was de oxi-test ook geslaagd. Kalties + QC's : OK | De hemolyser werd preventief vervangen. tHb-kalibratie en gewone kalibratie:OK In het serviceprogramma was de oxi-test ook geslaagd. Kalties + QC's : OK | SWO-34992 | No | The hemolyser was replaced preventively. tHb calibration and regular calibration: OK In The service program, the Oxi test was also successful. Kalties + Qcs: OK | The hemolyser was replaced preventively. tHb calibration and regular calibration: OK In The service program, the Oxi test was also successful. Kalties + Qcs: OK | Defect | HW | Hemolyzer | LEH
ABL90 FLEX | Analysator mehrfach mit Kondensmilch gespült. System ok. | Analysator mehrfach mit Kondensmilch gespült. System ok. | SWO-35015 | No | Analyzer rinsed several times with condensed milk. System ok. | Analyzer rinsed several times with condensed milk. System ok. | Assisted Replacement | SC | Wetability | LEH
ABL90 FLEX PLUS |  |  | CF2CD917BA111ADC | No |  |  | Defect | HW | Display
ABL820 FLEX | Confirmed leak, customer was able to resolve immediate cause by replacing waste pump tubing. Fluidic panel and surrounding area required more extensive cleaning. | Removed fluidic panel, tubing manifold, clean all areas impacted, redressed tubing, installed new waste pump tubing. No additional issues observed. | CF2B88054417B772 | No | Confirmed leak, customer was able to resolve immediate cause by replacing waste pump tubing. Fluidic panel and surrounding area required more extensive cleaning. | Removed fluidic panel, tubing manifold, clean all areas impacted, redressed tubing, installed new waste pump tubing. No additional issues observed. | Assisted Replacement | Electrode | Electrode | Lac
ABL825 FLEX | Laktat Elektrode (E7007 RN05) geprüft und ausgetauscht, Funktionskontrolle ok | Laktat Elektrode (E7007 RN05) geprüft und ausgetauscht, Funktionskontrolle ok | SWO-35029 | No | Lactate electrode (E7007 RN05) tested and replaced, function control ok | Lactate electrode (E7007 RN05) tested and replaced, function control ok | Assisted Replacement | Electrode | Electrode | Lac
ABL825 FLEX | Analysator teilweise demontiert - Sicherungen auf AU Board überprüft und Kontakte gereinigt etc. - Funktionsüberprüfung etc. | Analysator teilweise demontiert - Sicherungen auf AU Board überprüft und Kontakte gereinigt etc. - Funktionsüberprüfung etc. | SWO-35019 | No | Analyzer partially dismantled - fuses checked on AU board and contacts cleaned etc. - function check etc. | Analyzer partially dismantled - fuses checked on AU board and contacts cleaned etc. - function check etc. | Defect | HW | Analyzer Control PCB | PCB
ABL825 FLEX | Checked PCO2 electrode and membrane, no noticable air bubbles or pinhole leaks. Checked lot numbers for PCO2 electrode and membrane and compared with other two analyzers. | Swapped electrodes with another analyzer, unstable errors followed the electrode. | CF2B615044381AEE | No | Checked PCO2 electrode and membrane, no noticable air bubbles or pinhole leaks. Checked lot numbers for PCO2 electrode and membrane and compared with other two analyzers. | Swapped electrodes with another analyzer, unstable errors followed the electrode. | Assisted Replacement | Electrode | Electrode | Lac
ABL825 FLEX | checked electrode and membrane for air bubles or pinhole leaks. Swapped electrode with another analyzer. Error followed the electrode. | Ordering a new PCO2 electrode for customer | CF2B734F443B1AF2 | No | checked electrode and membrane for air bubles or pinhole leaks. Swapped electrode with another analyzer. Error followed the electrode. | Ordering a new PCO2 electrode for customer | Assisted Replacement | Electrode | Electrode | Lac
ABL835 FLEX | Pumpenschlauch ersetzt, nach Pumpenkal: 93%:90%. Fehlerhaften Pumpenschlauch in BG-Modul eingesetzt. Jetzt hier: 78%:92%. Beide 842-328 ersetzt. | Pumpenschlauch ersetzt, nach Pumpenkal: 93%:90%. Fehlerhaften Pumpenschlauch in BG-Modul eingesetzt. Jetzt hier: 78%:92%. Beide 842-328 ersetzt. | SWO-35105 | No | Pump hose replaced, according to pump scale: 93%:90%. Faulty pump hose in BG module. Now playing: 78%:92%. Replaced both 842-328. | Pump hose replaced, according to pump scale: 93%:90%. Faulty pump hose in BG module. Now playing: 78%:92%. Replaced both 842-328. | Defect | HW | pH/BG Module | Pump tubing
ABL90 FLEX | Akku getauscht | Akku getauscht | SWO-35280 | No | Battery replaced | Battery replaced | Defect | HW | Battery module
ABL90 FLEX | Chip ID PCB ersetzt, Dateien überprüft und Dr. Oberhollenzer Tipps zur richtigen Anwendung der safeClinitubes gegeben sowie einen Mischmagnet überreicht. | Chip ID PCB ersetzt, Dateien überprüft und Dr. Oberhollenzer Tipps zur richtigen Anwendung der safeClinitubes gegeben sowie einen Mischmagnet überreicht. | SWO-35329 | No | Chip ID PCB replaced, files checked and Dr. Oberhollenzer gave tips on the correct application of safeClinitubes as well as a mixed magnet handed over. | Chip ID PCB replaced, files checked and Dr. Oberhollenzer gave tips on the correct application of safeClinitubes as well as a mixed magnet handed over. | Defect | HW | SP PCB
ABL90 FLEX | Remplacé le kit de maintenance annuel + l'aiguille e.o garantie Remplacé l'hémolyseur + la valve de l'oxymétrie e.o garantie Calibration e.o contrôle de qualité e.o garantie | Remplacé le kit de maintenance annuel + l'aiguille e.o garantie Remplacé l'hémolyseur + la valve de l'oxymétrie e.o garantie Calibration e.o contrôle de qualité e.o garantie | SWO-35249 | No | Replaced the annual maintenance kit - the e.o guaranteed needle Replaced the hemolyser - the oxymetry valve e.o guarantee Calibration e.o quality control e.o warranty | Replaced the annual maintenance kit - the e.o guaranteed needle Replaced the hemolyser - the oxymetry valve e.o guarantee Calibration e.o quality control e.o warranty | Defect | HW | Hemolyzer
ABL90 | During Initialize process unit displayed Recommend replacing SC. | Replaced SC. | CF2DBD383052F051 | No | During Initialize process unit displayed Recommend replacing SC. | Replaced SC. | Assisted Replacement | SC | Lac
ABL825 FLEX | Instalacion equipo Reparacion Cooximetro Controles y calibraciones Configuracion envio a SIL |  | CF312BEA9262C0CF | Yes | Installation equipment Repair Cooximeter Controls and calibrations Configuration shipping to SIL |  | Re classified RMED | Installation
ABL800 FLEX | looked to see if the keys were missing | Reinstalled the option keys | CF2CDAF9BA4707A7 | No | looked to see if the keys were missing | Reinstalled the option keys | Defect | SW | SW Key missing
ABL837 FLEX | Creatinine failing calibration the sensitivity status are off. New solutions membranes did not solve the problem. | Found Electrode A missing the O-ring. Replaced the O-ring at the tip of the Electrode New Membranes...After several calibrations OK | CF2CFE3DBA58A9F3 | No | Creatinine failing calibration the sensitivity status are off. New solutions membranes did not solve the problem. | Found Electrode A missing the O-ring. Replaced the O-ring at the tip of the Electrode New Membranes... After several calibrations OK | Assisted Replacement | Electrode | Membrane | Sealing Ring | Crea
ABL825 FLEX | Located bent Probe. | Replaced with customers stock Probe. | CF2DBB0C305AF04E | No | Located bent Probe. | Replaced with customers stock Probe. | Defect | HW | Inlet module | Probe bend
ABL825 FLEX | PCO2 electrode found to be the problem from yesterdays troubleshooting | Replaced PCO2 electrode lot # RN01 with lot# RW04 | CF2CA826BA4F17D6 | No | PCO2 electrode found to be the problem from yesterdays troubleshooting | Replaced PCO2 electrode lot # RN01 with lot# RW04 | Assisted Replacement | Electrode | Electrode | pCO2
ABL825 FLEX | PCO2 electrode found to be the problem from yesterdays troubleshooting | Replaced PCO2 electrode lot # RM02 with lot# RW04 | CF2CA932BA3217DA | No | PCO2 electrode found to be the problem from yesterdays troubleshooting | Replaced PCO2 electrode lot # RM02 with lot# RW04 | Assisted Replacement | Electrode | Electrode | pCO2
AQT90 FLEX | Optische Einheit erneuert. Alle Parameter neu kalibriert. Mehrere QC- und Patientenvergleichsmessungen durchgeführt. Alle Trop-Messungen waren unter 0,01ug L. Gerät überprüft. | Optische Einheit erneuert. Alle Parameter neu kalibriert. Mehrere QC- und Patientenvergleichsmessungen durchgeführt. Alle Trop-Messungen waren unter 0,01ug L. Gerät überprüft. | SWO-35208 | No | Optical unit renewed. All parameters are recalibrated. Several QC and patient comparison measurements. All Trop measurements were checked under 0.01ug L. device. | Optical unit renewed. All parameters are recalibrated. Several QC and patient comparison measurements. All Trop measurements were checked under 0.01ug L. device.
ABL80 FLEX BASIC | SC needs to be replaced, informed the customer to install new SC after receiving the same. Informed the dealer to supply new SC. | SC needs to be replaced, informed the customer to install new SC after receiving the same. Informed the dealer to supply new SC. | SWO-35475 | No | SC needs to be replaced, informed the customer to install new SC after receiving the same. Informed the dealer to supply new SC. | SC needs to be replaced, informed the customer to install new SC after receiving the same. Informed the dealer to supply new SC.
ABL90 FLEX | Konfiguration und Einstellungen überprüft, Daten per Laptop ausgelesen, Analysator an der Herz- Lungenmaschine angeschlossen und erfolgter Datentransfer durch Kardiotechniker bestätigt 3 weitere Geräte identisch konfiguriert und Dataflow per Laptop bestätigt | Konfiguration und Einstellungen überprüft, Daten per Laptop ausgelesen, Analysator an der Herz- Lungenmaschine angeschlossen und erfolgter Datentransfer durch Kardiotechniker bestätigt 3 weitere Geräte identisch konfiguriert und Dataflow per Laptop bestätigt | SWO-35476 | No | Configuration and settings checked, data read out via laptop, analyzer connected to the heart-lung machine and data transfer by cardio technician confirmed 3 other devices identically configured and dataflow confirmed by laptop | Configuration and settings checked, data read out via laptop, analyzer connected to the heart-lung machine and data transfer by cardio technician confirmed 3 other devices identically configured and dataflow confirmed by laptop | Setup | HIS/LIS
ABL90 FLEX | Allgemeine Kontrolle ok. Lösungspack ersetzt ok. (wird verrechnet) Sie erhalten eine Gutschrift auf restli. Tests Im Servicemenü alle Tests durchgeführt ok. System neu gestartet ok. Bestehende Sensorkassette installiert ok. kal. + QC i.o. *** kunde behaltet eine Sensorkassette als ersatz *** Wird verrechnet. | Allgemeine Kontrolle ok. Lösungspack ersetzt ok. (wird verrechnet) Sie erhalten eine Gutschrift auf restli. Tests Im Servicemenü alle Tests durchgeführt ok. System neu gestartet ok. Bestehende Sensorkassette installiert ok. kal. + QC i.o. *** kunde behaltet eine Sensorkassette als ersatz *** Wird verrechnet. | SWO-35511 | No | General control ok. Solution pack replaces ok. (will be charged) You will receive a credit on restli. Tests In the service menu all tests performed ok. System restarted ok. Existing sensor cartridge installed ok. Kal. + QC i.o. *** customer retains a sensor cartridge as replacement *** Will be charged. | General control ok. Solution pack replaces ok. (will be charged) You will receive a credit on restli. Tests In the service menu all tests performed ok. System restarted ok. Existing sensor cartridge installed ok. Kal. + QC i.o. *** customer retains a sensor cartridge as replacement *** Will be charged. | Unknown | SC | Wetability
ABL825 FLEX | Printer head ausgebaut und ersetzt, Luftkanal gespült, Kal und QC durchgeführt | Printer head ausgebaut und ersetzt, Luftkanal gespült, Kal und QC durchgeführt | SWO-35421 | No | Printer head removed and replaced, duct flushed, Kal and QC performed | Printer head removed and replaced, duct flushed, Kal and QC performed | Defect | HW | Printer module | Printer Head
ABL837 FLEX | Performed troubleshooting methods with customer to determine why instrument remains in hold. Instructed customer to select replaced for the actions that were completed and then to select restart . The instrument began initializing but returned to hold status after a few seconds. Instructed to the customer to remove the fluidics module covers to further investigate but she was unable to do so. | Phone troubleshooting methods did not correct the issue. A service visit is needed and has been scheduled. | CF2DE5E43001778C | Yes | Performed troubleshooting methods with customer to determine why instrument remains in hold. Instructed customer to select replaced for the actions that were completed and then to select restart . The instrument began initializing but returned to hold status after a few seconds. Instructed to the customer to remove the fluidics module covers to further investigate but she was unable to do so. | Phone troubleshooting methods did not correct the issue. A service visit is needed and has been scheduled. | Re classified RMED | Phone support
ABL837 FLEX | Inlet probe is observed to be bent and can not pass through the inlet gasket | Replaced inlet gasket and insured that instrument performs properly. Instrument is within acceptable operating specifications. | CF2DEB833009778E | No | Inlet probe is observed to be bent and can not pass through the inlet gasket | Replaced inlet gasket and insured that instrument performs properly. Instrument is within acceptable operating specifications. | Defect | HW | Inlet module | Probe bend
ABL825 FLEX | Found actuator arm was missing cotter pin. | Replaced cotter pin with non inventory part. Performed necessary adjustments. | CF2E3C1E301988FE | No | Found actuator arm was missing cotter pin. | Replaced cotter pin with non inventory part. Performed necessary adjustments. | Defect | HW | AutoCheck Module | Breaker arm
ABL835 FLEX | A megjavított cseremodul visszaépítése a készülékbe. Szofver frissítése (933-330 ABL800) v6.14 verzióra. Kalibrálás, QC mérés hitelesítés |  | SWO-35347 | No | Reconstruct the repaired replacement module in the device. Update Software (933-330 ABL800) to v 6.14 version. Calibration, QC measurement authentication |  | Defect | HW | Inlet module | Friction detector unit
ABL825 FLEX | Überprüfung Measurement Ventile - Erneuerung Airmembran von EL MET Modul - Funktionsüberprüfung etc. | Überprüfung Measurement Ventile - Erneuerung Airmembran von EL MET Modul - Funktionsüberprüfung etc. | SWO-35477 | No | Verification Measurement Valves - Renewal Airmembrane of EL MET Module - Function Verification etc. | Verification Measurement Valves - Renewal Airmembrane of EL MET Module - Function Verification etc. | Defect | HW | El/Met Module | Valve, air
ABL825 FLEX | Erneuerung Power Supply - zunächst kein Fehler vorhanden - nach Test Restart Fehler wieder vorhanden - Software erneuert ( keine Fehlermeldung ) - nach Restart, Konfiguration neu geladen ( Fehler wieder vorhanden ) - nach Rücktausch Power Supply wieder kein Fehler vorhanden - Analysator momentan Funktionsbereit - Bestellung Analyzer Control Board ist erfolgt Fazit : Nach jedem Neustart Fehlermeldung wieder vorhanden ( bleibt auch bei erneutem Restart bestehen ) - außer Gerät war längere Zeit ausgeschaltet ( Reparatur ) !? | Erneuerung CPU Unit brachte keine Fehlerbeseitigung ( Rücktausch ) - Erneuerung Connector PCB brachte keine Fehlerbeseitigung ( kein Rücktausch erfolgt, da bei Ausbau Altteil kleine Beschädigung an PCB entstanden ) - Analysator voll Funktionsbereit, lediglich muss der Analysator kurze Zeit ruhen, falls er ausgeschaltet wurde ( einige Minuten ) - Kalibration und QC ohne Beanstandung ! Bem.: Nach Rücksprache erfolgt erneute Ersatzteilbestellung ( evt. Festplatte ) | SWO-35478 | No | Renewal Power Supply - initially no error present - after test restart error again present - software renewed ( no error message ) - after restart, configuration reloaded ( error again present ) - after reconversion Power Supply again no error present - Analyzer currently functional - Order Analyzer Control Board is done Conclusion: After each restart error message is present again ( remains even with restart ) - except device was switched off for a long time ( Repair ) !? | Renewal CPU Unit did not bring any troubleshooting ( exchange ) - Renewal Connector PCB brought no bug removal (no replacement takes place, since during removal old part caused small damage to PCB ) - Analyzer fully functional, only the Analyzer rest for a short time in case it has been switched off (a few minutes ) - calibration and QC without complaint ! Bem.: After consultation, a new order for spare parts will be made (evt. Hard drive ) | Defect | HW | Analyzer Control PCB
ABL835 FLEX |  |  | CF2F349EA63593B6 | Yes |  |  | Re classified RMED | Installation
ABL835 FLEX | checked lamp settings found it was maxed checked logs to varify thb was out | adjusted lamp into range | CF2E0D1C30593B6D | No | checked lamp settings found it was maxed checked logs to varify thb was out | adjusted lamp into range | Adjustment | HW | Oxi module | Lamp unit
ABL90 |  |  | CF3098A41C28BF5A | No |  |  | Defect | HW | Hemolyzer
ABL835 FLEX | Conexion por remoto no permite solucionar averia Desmontaje de todos los modulos por avistamiento de liquido derramado Secado del equipo Montaje de equipo Controles y calibracion |  | CF2F47FBA60BC0CA | No | Remote connection does not allow to solve damage Disassembly of all modules by sighting of spilled liquid Equipment Equipment Assembly Controls and calibration |  | Cleaned | Fluidic pathway
ABL825 FLEX | 537 oxi and 532 El Met liquid sensor errors in activity log. | Freed and cleaned stuck flat membrane valves in electrode modules. Also flushed air vent in fluidic module. Remembraned pO2 and pCO2 electrodes and replaced cleaning solution bottle. | CF2F65F1A65F7714 | No | 537 oxi and 532 El Met liquid sensor errors in activity log. | Freed and cleaned stuck flat membrane valves in electrode modules. Also flushed air vent in fluidic module. Remembraned pO2 and pCO2 electrodes and replaced cleaning solution bottle. | Defect | HW | pH/BG Module | Valve, air
ABL800 BASIC | Parameter Glu membraniert. Kal mehrfach durchgeführt. QCs wiederholt. | Parameter Glu membraniert. Kal mehrfach durchgeführt. QCs wiederholt. | SWO-35723 | No | Parameter Glu membraned. Kal performed several times. QCs. | Parameter Glu membraned. Kal performed several times. QCs. | Assisted Replacement | Electrode | Membrane | Glu
ABL815 FLEX | Oxi PCB erneuert | Oxi PCB erneuert | SWO-35703 | No | Oxygen PCB Erneuert | Oxygen PCB Erneuert | Defect | HW | Oxi module | PCB
ABL825 FLEX | Erneuerung Analysator Control Board - Softwarepflege, Optionkeys, Konfiguration und Roma.ini Datei geändert ( Barcodes nicht lesbar ) - Analysator zunächst Funktionsbereit - Nach Neustart Fehlermeldung 643 - Rücktausch Analysator Control Board - Analysator voll Funktionsbereit, muss aber gegebenenfalls bei Abschaltung einige Zeit ruhen vor Neustart ( 5 Min. ). Bestellung CPU und Connector PCB ist erfolgt ! | Erneuerung CPU Unit brachte keine Fehlerbeseitigung ( Rücktausch ) - Erneuerung Connector PCB brachte keine Fehlerbeseitigung ( kein Rücktausch erfolgt, da bei Ausbau Altteil kleine Beschädigung an PCB entstanden ) - Analysator voll Funktionsbereit, lediglich muss der Analysator kurze Zeit ruhen, falls er ausgeschaltet wurde ( einige Minuten ) - Kalibration und QC ohne Beanstandung ! Bem.: Nach Rücksprache erfolgt erneute Ersatzteilbestellung ( evt. Festplatte ) | SWO-35696 | No | Renewal Analyzer Control Board - Software Maintenance, Optionkeys, Configuration and Roma.ini File Changed ( Barcodes Not Readable ) - Analyzer Initially Functional - After Restart Error Message 643 - Replacement Analyzer Control Board - Analyzer Full Ready for use, but may have to rest some time before restart ( 5 min. if switched off). Ordering CPU and Connector PCB has been made ! | Renewal CPU Unit did not bring any troubleshooting ( exchange ) - Renewal Connector PCB brought no bug removal (no replacement takes place, since during removal old part caused small damage to PCB ) - Analyzer fully functional, only the Analyzer rest for a short time in case it has been switched off (a few minutes ) - calibration and QC without complaint ! Bem.: After consultation, a new order for spare parts will be made (evt. Hard drive ) | Defect | HW | Analyzer Control PCB
ABL90 FLEX | Gerät aus ein, Einlassdichtung ersetzt, s-box geliefert, tHb Kal durchgeführt Intervall für Einlassdichtungerneuerung korrigiert da nächster Termin auf 28.08.15 bleibt | Gerät aus ein, Einlassdichtung ersetzt, s-box geliefert, tHb Kal durchgeführt Intervall für Einlassdichtungerneuerung korrigiert da nächster Termin auf 28.08.15 bleibt | SWO-35813 | No | Device from on, inlet seal replaced, s-box supplied, tHb Kal carried out interval for inlet seal renewal corrected as next appointment remains on 28.08.15 | Device from on, inlet seal replaced, s-box supplied, tHb Kal carried out interval for inlet seal renewal corrected as next appointment remains on 28.08.15 | Defect | SW | Replacement schedule
ABL90 FLEX | Touchscreen lässt sich nicht kalibrieren--> erneuert--> Touchscreen kalibriert--> i.O |  | SWO-35782 | No | Touchscreen cannot be calibrated--> renewed--> Touchscreen calibrated--> i.O |  | Defect | HW | Touch screen
ABL825 FLEX |  |  | CF3056911C1C47A8 | No |  |  | Defect | HW | pH/BG Module | LS
ABL827 FLEX | Ventilblock des CreaModuls (Leihgabe) ersetzt, mehrere korrekte Kal abgewartet | Erneut alle Ventile der Messmodule demontiert, gereinigt, montiert, Pumpenschlauch PHBG- & ElMet-Modul ersetzt, RinseOutStopper am Crea-Modul ersetzt, Einlaßdichtung ersetzt, mehrere korrekte Kal abgewartet, QK ok | SWO-35855 | No | Valve block of the CreaModule (loan) replaced, several correct cal waits | Again all valves of the measuring modules disassembled, cleaned, mounted, pump hose PHBG & ElMet module replaced, RinseOutStopper replaced on Crea module, inlet seal replaced, several correct cal waited, QK ok | Defect | HW | Met II Module | Valve
ABL825 FLEX | Erneuerung CPU Unit brachte keine Fehlerbeseitigung ( Rücktausch ) - Erneuerung Connector PCB brachte keine Fehlerbeseitigung ( kein Rücktausch erfolgt, da bei Ausbau Altteil kleine Beschädigung an PCB entstanden ) - Analysator voll Funktionsbereit, lediglich muss der Analysator kurze Zeit ruhen, falls er ausgeschaltet wurde ( einige Minuten ) - Kalibration und QC ohne Beanstandung ! Bem.: Nach Rücksprache erfolgt erneute Ersatzteilbestellung ( evt. Festplatte ) | Erneuerung CPU Unit brachte keine Fehlerbeseitigung ( Rücktausch ) - Erneuerung Connector PCB brachte keine Fehlerbeseitigung ( kein Rücktausch erfolgt, da bei Ausbau Altteil kleine Beschädigung an PCB entstanden ) - Analysator voll Funktionsbereit, lediglich muss der Analysator kurze Zeit ruhen, falls er ausgeschaltet wurde ( einige Minuten ) - Kalibration und QC ohne Beanstandung ! Bem.: Nach Rücksprache erfolgt erneute Ersatzteilbestellung ( evt. Festplatte ) | SWO-35866 | No | Renewal CPU Unit did not bring any troubleshooting ( exchange ) - Renewal Connector PCB brought no bug removal (no replacement takes place, since during removal old part caused small damage to PCB ) - Analyzer fully functional, only the Analyzer rest for a short time in case it has been switched off (a few minutes ) - calibration and QC without complaint ! Bem.: After consultation, a new order for spare parts will be made (evt. Hard drive ) | Renewal CPU Unit did not bring any troubleshooting ( exchange ) - Renewal Connector PCB brought no bug removal (no replacement takes place, since during removal old part caused small damage to PCB ) - Analyzer fully functional, only the Analyzer rest for a short time in case it has been switched off (a few minutes ) - calibration and QC without complaint ! Bem.: After consultation, a new order for spare parts will be made (evt. Hard drive ) | Unknown | HW | -
ABL825 FLEX | pCO2 Elektrode erneuert. Messkammer und Elektrodenkontakte gereinigt. Proteinentfernung und Service tHB Kal. durchgeführt. Servicetest durchgeführt . kein Erfolg! Oximodul erneuert. Alles o.k. | pCO2 Elektrode erneuert. Messkammer und Elektrodenkontakte gereinigt. Proteinentfernung und Service tHB Kal. durchgeführt. Servicetest durchgeführt . kein Erfolg! Oximodul erneuert. Alles o.k. | SWO-35868 | No | pCO2 electrode. measuring chamber and electrode contacts. Protein removal and service tHB Kal. Performed. Service test performed . no success! Oximodul. Everything OK. | pCO2 electrode. measuring chamber and electrode contacts. Protein removal and service tHB Kal. Performed. Service test performed . no success! Oximodul. Everything OK. | Assisted Replacement | Electrode | Electrode | pCO2
ABL825 FLEX | mit Kunden Sinn der Modustasten und der Layouts besprochen, Ausdrucke der durchgeführten Ringversuche mit anderem Layout incl. Oxymetrie erzeugt | mit Kunden Sinn der Modustasten und der Layouts besprochen, Ausdrucke der durchgeführten Ringversuche mit anderem Layout incl. Oxymetrie erzeugt | SWO-35922 | No | discussed with customers sense of mode keys and layouts, prints of the performed ring tests with different layout incl. oxymetry generated | discussed with customers sense of mode keys and layouts, prints of the performed ring tests with different layout incl. oxymetry generated | Setup | Patient report | -
ABL80 FLEX BASIC | Need to replace solution pack informed the customer to install new SP after receiving the same. Informed the dealer to supply new SP. |  | SWO-36786 | No | Need to replace solution pack informed the customer to install new SP after receiving the same. Informed the dealer to supply new SP.
ABL80 FLEX BASIC | Replaced sensor cassette | Need to replace solution pack informed the customer to install new SP after receiving the same. Informed the dealer to supply new SP. | SWO-37061 | No | Replaced sensor cassette | Need to replace solution pack informed the customer to install new SP after receiving the same. Informed the dealer to supply new SP.
ABL80 FLEX Basic version |  |  | CF347D437E49CA8E | No
ABL90 | Bruce said the analyzer is showing 1184-leak detected. He said there is also QC material leaking where the SP sits. I found QC around the waste port, spring and tube. I put on a maintenance kit. I will claim the kit on another SWO that I will use for PM. See SWO B9LY4SU9J for Annual. | Put on yearly maintenance kit. Cleaned the QC off. | CF3462A67E037C19 | No | Bruce said the analyzer is showing 1184-leak detected. He said there is also QC material leaking where the SP sits. I found QC around the waste port, spring and tube. I put on a maintenance kit. I will claim the kit on another SWO that I will use for PM. See SWO B9LY4SU9J for Annual. | Put on yearly maintenance kit. Cleaned the QC off. | Defect | HW | Waste Tube
ABL90 | LOOKED THROUGH DATA LOGS AND SENSOR CASSETTE EXPIRED ON JULY 4, 2015, ALSO THB CALIBRATION IS REQUIRED ON JULY 6, 2015. | REPLACED SENSOR CASSETTE WITH CUSTOMERS STOCK, DID 5 WASTE BLOODS, LACTATE TOOK SOME TIME TO COME IN. DID CALIBRATION ABOUT 5 TIMES. ALSO DID THB CALIBRATION. | CF345E657E40DF55 | No | LOOKED THROUGH DATA LOGS AND SENSOR CASSETTE EXPIRED ON JULY 4, 2015, ALSO THB CALIBRATION IS REQUIRED ON JULY 6, 2015. | REPLACED SENSOR CASSETTE WITH CUSTOMERS STOCK, DID 5 WASTE BLOODS, LACTATE TOOK SOME TIME TO COME IN. DID CALIBRATION ABOUT 5 TIMES. ALSO DID THB CALIBRATION. | User Error | SC expired
ABL90 FLEX | In service menu alle checkouts uitgevoerd, deze zijn goed. Backflush uitgevoerd, gaat niet helemaal soepel, derhalve naald, en gasket vervangen uit voorraad klant. Tevens de inlet connector gasket vervangen. Backflush gaat nu soepel. tHb-kalibratie uitgevoerd --> OK. Meten van patientmonster wordt wederom afgebroken met melding 1189. Visual flow check uitgevoerd, deze is goed. Meten van monsters via de capillaire modus gaat wel goed. Inlet Position Detector PCB vervangen, hierna gaan de metingen via spuitmodus ook goed. Enkele testmonsters gemeten, dit gaat allemaal goed. ABL Meetklaar achtergelaten. | In service menu alle checkouts uitgevoerd, deze zijn goed. Backflush uitgevoerd, gaat niet helemaal soepel, derhalve naald, en gasket vervangen uit voorraad klant. Tevens de inlet connector gasket vervangen. Backflush gaat nu soepel. tHb-kalibratie uitgevoerd --> OK. Meten van patientmonster wordt wederom afgebroken met melding 1189. Visual flow check uitgevoerd, deze is goed. Meten van monsters via de capillaire modus gaat wel goed. Inlet Position Detector PCB vervangen, hierna gaan de metingen via spuitmodus ook goed. Enkele testmonsters gemeten, dit gaat allemaal goed. ABL Meetklaar achtergelaten. | SWO-36365 | No | In Service menu All checkouts performed, these are good. Backflush implemented, does not go completely smoothly, therefore needle, and gasket replaced from stock customer. Also replace the inlet connector gasket. Backflush now goes smoothly. tHb calibration performed--> OK. Measuring the patient sample is again aborted with notification 1189. Visual flow Check performed, this is good. Measuring samples via capillary mode is fine. Replace the Inlet Position Detector PCB, the measurements will also be carried out via spraying mode. Some test samples measured, this all goes well. ABL measuring ready left behind. | In Service menu All checkouts performed, these are good. Backflush implemented, does not go completely smoothly, therefore needle, and gasket replaced from stock customer. Also replace the inlet connector gasket. Backflush now goes smoothly. tHb calibration performed--> OK. Measuring the patient sample is again aborted with notification 1189. Visual flow Check performed, this is good. Measuring samples via capillary mode is fine. Replace the Inlet Position Detector PCB, the measurements will also be carried out via spraying mode. Some test samples measured, this all goes well. ABL measuring ready left behind. | Defect | HW | Inlet PCB
ABL90 FLEX | Gerät manuell im Serviceprogramm gespült, Einlasssonde erneuert, Einlassdichtung gereinigt, Kal und QCs durchgeführt. | Gerät manuell im Serviceprogramm gespült, Einlasssonde erneuert, Einlassdichtung gereinigt, Kal und QCs durchgeführt. | SWO-36050 | No | Device manually rinsed in the service program, inlet probe renewed, inlet seal cleaned, Kal and QCs performed. | Device manually rinsed in the service program, inlet probe renewed, inlet seal cleaned, Kal and QCs performed. | Cleaned | Fluidic pathway
ABL90 FLEX PLUS |  |  | CF34657A7E0708A8 | No |  |  | Assisted Replacement | SP | LS error
ABL825 FLEX | Sostituita valvola soluzione di calibrazione 2. | Sostituita valvola soluzione di calibrazione 2. | SWO-36082 | No | Replaced calibration solution valve 2. | Replaced calibration solution valve 2. | Defect | HW | Fluidic manifold | Valve
ABL815 FLEX | Fließwege überprüft, Druckverlust von der Lösungspumpe zum Einlaß. Hemolysator ausgebaut, Lichtleiter ließ sich nicht lösen. OXI-Modul in der Zentrale bestellt. | Notdiensttechniker bitte am 04.07., früh gleich Kunden anrufen, wie es weitergehen soll | SWO-36198 | No | Flow paths checked, pressure loss from the solution pump to inlet. Hemolysator removed, light guide could not be detached. OXI module at the head office. | Emergency service technicians please call customers early on 04.07., as to how to proceed | Defect | HW | Oxi module | ?
ABL800 BASIC | Ref Elektrode getauscht, mehrere Kal,s durchgeführt und QC vermessen | Ref Elektrode getauscht, mehrere Kal,s durchgeführt und QC vermessen | SWO-36090 | No | Ref electrode exchanged, several kal,s performed and QC measured | Ref electrode exchanged, several kal,s performed and QC measured | Assisted Replacement | Electrode | Electrode | Reference
ABL825 FLEX |  |  | CF3493597E205806 | No |  |  | Defect | HW | Oxi module | Hemolyzer
ABL800 BASIC | pCO 2 Elektrode mit Membran anderer Lot membraniert, es wurde vormittags eine Dekontamination durchgeführt, ohne beiLac und Gluc Dummy Elektroden einzusetzen, Lac und Gluc Membran ersetzt, Kal und QC durchgeführt, Anwender auf Durchführung der Proteinentfernung und Dekontaminatin hingewiesen | pCO 2 Elektrode mit Membran anderer Lot membraniert, es wurde vormittags eine Dekontamination durchgeführt, ohne beiLac und Gluc Dummy Elektroden einzusetzen, Lac und Gluc Membran ersetzt, Kal und QC durchgeführt, Anwender auf Durchführung der Proteinentfernung und Dekontaminatin hingewiesen | SWO-36199 | No | pCO 2 electrode with membrane membrane of other solder membranes, decontamination was carried out in the morning without using lac and gluc dummy electrodes, Lac and Gluc membrane replaced, Kal and QC performed, users performed on performing protein removal and Decontaminatine noted | pCO 2 electrode with membrane membrane of other solder membranes, decontamination was carried out in the morning without using lac and gluc dummy electrodes, Lac and Gluc membrane replaced, Kal and QC performed, users performed on performing protein removal and Decontaminatine noted | Assisted Replacement | Electrode | Membrane | pCO2
ABL820 FLEX | ACTIVITY LOG SHOWS 248 SPECTROMETER MISSING TRIGGER. 247 A TO D FAILURE AND 316 NEON LAMP ERROR.VISUAL INSPECTION SHOWS MISSING PO2 ELECTRODE O RING.PC02 SENSITIVITY AT 80% | REBOOTED THE ANALYZER.ALL OXI ERRORS WENT AWAY.IF THE PROBLEM REOCCURS THEN THE OXI MODULE BOARD WILL NEED REPLACING .REPLACED MISSING PO2 ELECTRODE O RING.REMEMBRANED THE PC02 ELECTRODE AND SENSITIVITY NOW IN THE 90'S.RAN MULTIPLE CALIBRATIONS | CF3473787E158E90 | No | ACTIVITY LOG SHOWS 248 SPECTROMETER MISSING TRIGGER. 247 A TO D FAILURE AND 316 NEON LAMP ERROR. VISUAL INSPECTION SHOWS MISSING PO2 ELECTRODE O RING. PC02 SENSITIVITY AT 80% | REBOOTED THE ANALYZER. ALL OXI ERRORS WENT AWAY. IF THE PROBLEM REOCCURS THEN THE OXI MODULE BOARD WILL NEED REPLACING . REPLACED MISSING PO2 ELECTRODE O RING. REMEMBRANED THE PC02 ELECTRODE AND SENSITIVITY NOW IN THE 90'S. RAN MULTIPLE CALIBRATIONS | Assisted Replacement | Electrode | Electrode | Sealing Ring
ABL825 FLEX | Elektroden und Kontakte überprüft, Messkammer war sauber. Dekontamination durchgeführt. Ventile überprüft und Einlassdichtung erneuert. Eingeklemmten Schlauch neu verlegt. 3 mal kalibriert, alle einwandfrei. | Alle Schläuche und die Ref Elektrode erneuert. Danach 4 kallibrationen durchgeführt, alle einwandfrei. | SWO-36173 | No | Electrodes and contacts checked, measuring chamber was clean. decontamination. valves checked and inlet seal renewed. Clamped hose re-laid. 3 times calibrated, all flawless. | All hoses and the ref electrode renewed. After that 4 callibrations performed, all flawless. | Cleaned | Fluidic pathway | Electrode | Reference | LEH Mail to RDEU
ABL825 FLEX | Beide Pumpenschläuche erneuert. Pumpenkalibrierung in Ordnung. | Beide Pumpenschläuche erneuert. Pumpenkalibrierung in Ordnung. | SWO-36064 | No | Both pump hoses renewed. Pump calibration in order. | Both pump hoses renewed. Pump calibration in order. | Defect | HW | El/Met Module | Pump tubing
ABL835 FLEX | Checked connections. Everything connected. AU board failure. Neon lamp intensity too low, position error on Neon Lamp. | Replaced the AU board. Performed necessary calibrations required after AU board replacement. I also picked up trade in analyzers from Utah valley and shipped back. Managers are the same for both of these facilities. | CF34A8497E39E73F | No | Checked connections. Everything connected. AU board failure. Neon lamp intensity too low, position error on Neon Lamp. | Replaced the AU board. Performed necessary calibrations required after AU board replacement. I also picked up trade in analyzers from Utah valley and shipped back. Managers are the same for both of these facilities. | Defect | HW | Analyzer Control PCB
AQT90 FLEX | Nadel ausgetauscht, Nadel kalibriert und QC durchgeführt | Nadel ausgetauscht, Nadel kalibriert und QC durchgeführt | SWO-36189 | No | Needle replaced, needle calibrated and QC performed | Needle replaced, needle calibrated and QC performed
ABL90 | Error 1169. Fluid pathe blocked.Flushed Unit.Customer performed several Flushes with no result. | Performed Flush replaced SC and SP. | CF409EAB1A1227F8 | No | Error 1169. Fluid pathe blocked. Flushed Unit.Customer performed several Flushes with no result. | Performed Flush replaced SC and SP. | Assisted Replacement | SC | Wetability
ABL815 FLEX | OXI-Module eingebaut und justiert, Funktion OK. | Notdiensttechniker bitte am 04.07., früh gleich Kunden anrufen, wie es weitergehen soll | SWO-36381 | No | OXI modules built-in and adjusted, function OK. | Emergency service technicians please call customers early on 04.07., as to how to proceed | Unknown | HW | Oxi module | ?
ABL835 FLEX | sustitucion de la placa ph-bg y fusibles de la analyzer board. Chequeo por servicio. Cal de liquidos y bombas. Sustitucion de la membrana de pco2 y glu. Varias cal, qc. | sustitucion de la placa ph-bg y fusibles de la analyzer board. Chequeo por servicio. Cal de liquidos y bombas. Sustitucion de la membrana de pco2 y glu. Varias cal, qc. | CF357E04F43EC0B1 | No | replacement of the ph-bg board and analyzer board fuses. Check for service. Liquid and bomb lime. Replacement of the pco2 and glu membrane. Various lime, qc. | replacement of the ph-bg board and analyzer board fuses. Check for service. Liquid and bomb lime. Replacement of the pco2 and glu membrane. Various lime, qc. | Defect | HW | pH/BG Module | PCB
ABL825 FLEX |  |  | CF3562C8F4497478 | No |  |  | Defect | HW | pH/BG Module | Pump tubing
ABL825 FLEX | Probleme bestehen weiterhin. Kunde bittet um einen Vor Ort Service. | Alle Schläuche und die Ref Elektrode erneuert. Danach 4 kallibrationen durchgeführt, alle einwandfrei. | SWO-36288 | No | Problems remain. Customer asks for an on-site service. | All hoses and the ref electrode renewed. After that 4 callibrations performed, all flawless. | Assisted Replacement | Electrode | Electrode | Reference | LEH Mail to RDEU
ABL825 FLEX | Neue PO2 Elektrode (alt RN-02) neu eingesetzt, da trotz mehrfachen Membranierens verschiedener D799 der Nullwert > 10mmHg ist, Kal, QC | Neue PO2 Elektrode (alt RN-02) neu eingesetzt, da trotz mehrfachen Membranierens verschiedener D799 der Nullwert > 10mmHg ist, Kal, QC | SWO-36376 | No | New PO2 electrode (old RN-02) newly inserted, since despite multiple membraneing of different D799 the zero value > 10mmHg, Kal, QC | New PO2 electrode (old RN-02) newly inserted, since despite multiple membraneing of different D799 the zero value > 10mmHg, Kal, QC | Assisted Replacement | Electrode | Electrode | pO2
ABL825 FLEX | Messkammer gereinigt. Gerinsel in pH Messkammer. Oxitest durchgeführt. Lamp Voltage ( 3900 )zu hoch. Werte justiert und tHB Kal. durchgeführt. Neustart durchgeführt. Alle Funktionen und QC o.k. | Messkammer gereinigt. Gerinsel in pH Messkammer. Oxitest durchgeführt. Lamp Voltage ( 3900 )zu hoch. Werte justiert und tHB Kal. durchgeführt. Neustart durchgeführt. Alle Funktionen und QC o.k. | SWO-36379 | No | measuring chamber. Gerinsel in pH measuring chamber. Oxitest. Lamp Voltage ( 3900 ) too high. values adjusted and tHB Kal. Performed. Restart. All functions and QC ok. | measuring chamber. Gerinsel in pH measuring chamber. Oxitest. Lamp Voltage ( 3900 ) too high. values adjusted and tHB Kal. Performed. Restart. All functions and QC ok. | Adjustment | HW | Oxi module | Lamp unit
ABL825 FLEX | No investigation necessary. Install work order. |  | CF35F3C7F45919A9 | No | No investigation necessary. Install work order.
ABL80 FLEX CO-OX | Lüfter ersetzt. | Lüfter ersetzt. | SWO-36518 | No | replaced fan. | replaced fan.
ABL80 FLEX Basic version |  |  | CF3692F46A35E60E | No
ABL80 FLEX CO-OX | Fließwege im Serviceprogramm überprüft, Testmessung durchgeführt, Ansaugverhalten OK. Präanalytik und Probevolumen mit Sr. Lucia besprochen. Funktion z.Zt. OK. | Fließwege im Serviceprogramm überprüft, Testmessung durchgeführt, Ansaugverhalten OK. Präanalytik und Probevolumen mit Sr. Lucia besprochen. Funktion z.Zt. OK. | SWO-36556 | No | Flow paths in the service program checked, test measurement performed, suction behavior OK. Preanalysis and sample volume discussed with Sr. Lucia. Function currently OK. | Flow paths in the service program checked, test measurement performed, suction behavior OK. Preanalysis and sample volume discussed with Sr. Lucia. Function currently OK.
ABL80 FLEX BASIC | Replaced the third SC and calibrated. Tests were resumed |  | SWO-37049 | No | Replaced the third SC and calibrated. Tests were resumed
ABL90 |  |  | CF36ABA36A0BC6FC | No |  |  | Assisted Replacement | SP | pO2
ABL90 FLEX | Austausch des Gerätes gegen Leihgerät ( 090R0523N007) aus GS11 Bestand, Mitnahme des Kundengerätes zur Rep. in GS11 | Gerät aus ein, Einlassdichtung ersetzt, s-box geliefert, tHb Kal durchgeführt Intervall für Einlassdichtungerneuerung korrigiert da nächster Termin auf 28.08.15 bleibt | SWO-36425 | No | Replacement of the device with rental device ( 090R0523N007) from GS11 stock, carrying the customer device to The Rep. in GS11 | Device from on, inlet seal replaced, s-box supplied, tHb Kal carried out interval for inlet seal renewal corrected as next appointment remains on 28.08.15 | Unknown | Analyzer replaced
ABL90 FLEX | Defekten Touchscreen ausgetauscht. Kalibrierung des Displays durchgeführt. Initialisierungsphase. Kal und QC ohne Beanstandung | Defekten Touchscreen ausgetauscht. Kalibrierung des Displays durchgeführt. Initialisierungsphase. Kal und QC ohne Beanstandung | SWO-36486 | No | Defective touchscreen replaced. calibration of the display. Initialization phase. Kal and QC without complaint | Defective touchscreen replaced. calibration of the display. Initialization phase. Kal and QC without complaint | Defect | HW | Touch screen
ABL90 FLEX | Dateien gemäß Handbuch gesichert und gelöscht, leere Datenbank eingepflegt. SC und LP neu installiert. ABL nach Angaben konfiguriert. Der ABL baut keine Verbindung zum Radianceserver . ABL Software komplett neu installiert und konfiguriert. | Dateien gemäß Handbuch gesichert und gelöscht, leere Datenbank eingepflegt. SC und LP neu installiert. ABL nach Angaben konfiguriert. Der ABL baut keine Verbindung zum Radianceserver . ABL Software komplett neu installiert und konfiguriert. | SWO-36509 | No | Files according to the manual backed up and deleted, empty database entered. SC and LP. ABL according to the specifications. The ABL does not connect to the radiance server . ABL software completely reinstalled and configured. | Files according to the manual backed up and deleted, empty database entered. SC and LP. ABL according to the specifications. The ABL does not connect to the radiance server . ABL software completely reinstalled and configured. | Defect | SW | Status Code 2
ABL90 FLEX | Kunde hat Sensor Kassette gewechselt und Kal Fehler bei pCO2 aufgetreten. Alte Sensor Kassette wieder eingesetzt (Kein Neue SC vorhandene). SP ID PCB ausgetauscht, Einlassdichtung ausgetauscht. | Kunde hat Sensor Kassette gewechselt und Kal Fehler bei pCO2 aufgetreten. Alte Sensor Kassette wieder eingesetzt (Kein Neue SC vorhandene). SP ID PCB ausgetauscht, Einlassdichtung ausgetauscht. | SWO-36646 | No | Customer has changed sensor cassette and Kal error occurred with pCO2. Old sensor cassette reinserted (No New SC available). SP ID PCB replaced, inlet seal replaced. | Customer has changed sensor cassette and Kal error occurred with pCO2. Old sensor cassette reinserted (No New SC available). SP ID PCB replaced, inlet seal replaced. | Assisted Replacement | SC | pCO2
ABL90 FLEX | Gerät manuell gespült. Visual flow test durchgeführt. Jahreswartungskit eingebaut. LP und Hemolyzer (Flüssigkeitsaustritt an Hemolyzer) erneuert. Oxi Complete Test durchgeführt - ok. tHb Kalibration vorgenommen. Kalibrationen und QC's vorgenommen. | Gerät manuell gespült. Visual flow test durchgeführt. Jahreswartungskit eingebaut. LP und Hemolyzer (Flüssigkeitsaustritt an Hemolyzer) erneuert. Oxi Complete Test durchgeführt - ok. tHb Kalibration vorgenommen. Kalibrationen und QC's vorgenommen. | SWO-36763 | No | device manually rinsed. Visual flow test performed. Annual maintenance kit. LP and Hemolyzer (liquid leakage to hemolyzer). Oxi Complete Test performed - ok. tHb calibration. calibrations and QC's. | device manually rinsed. Visual flow test performed. Annual maintenance kit. LP and Hemolyzer (liquid leakage to hemolyzer). Oxi Complete Test performed - ok. tHb calibration. calibrations and QC's. | Defect | HW | Hemolyzer
ABL90 | No investigation needed. | I replaced the waste line, pinch valve tubing and performed a flush. | CF36C8ED6A31828B | No | No investigation needed. | I replaced the waste line, pinch valve tubing and performed a flush. | Cleaned | Fluidic pathway
ABL90 | Spectrum mismatch and turbidity errors on blood samples. Errors occured while unit was being validated by product spec. | Replaced hemolyzer | CF36CB8D6A59D90B | No | Spectrum mismatch and turbidity errors on blood samples. Errors occured while unit was being validated by product spec. | Replaced hemolyzer | Defect | HW | Hemolyzer
ABL80 FLEX Basic version | Confirmed problem. | Tested all of the valves and pump. Inspected the tubing for leaks. I did not find any defects. The sensor cassette needs to be replaced. I ordered a new one to be delivered non priority. | CF3702B46A2BD6C1 | No | Confirmed problem. | Tested all of the valves and pump. Inspected the tubing for leaks. I did not find any defects. The sensor cassette needs to be replaced. I ordered a new one to be delivered non priority.
ABL800 BASIC | Bildschirm abgezogen und Neustart durchgeführt ohne Erfolg, Mainbord bestellt | Bildschirm abgezogen und Neustart durchgeführt ohne Erfolg, Mainbord bestellt | SWO-36553 | No | Screen subtracted and restart performed without success, Mainbord ordered | Screen subtracted and restart performed without success, Mainbord ordered | Defect | HW | DMS_PC | Main Cpu
ABL835 FLEX | Vom Kunden aus (Ersatz-)Gerät Einlassmodul ausgebaut und zur Rep.-Überbrückung eingesetzt, Kalibrierungen und QC ohne Beanstandung | Vom Kunden aus (Ersatz-)Gerät Einlassmodul ausgebaut und zur Rep.-Überbrückung eingesetzt, Kalibrierungen und QC ohne Beanstandung | SWO-36568 | No | Removed from the customer (replacement) device inlet module and used for rep. bridging, calibrations and QC without complaint | Removed from the customer (replacement) device inlet module and used for rep. bridging, calibrations and QC without complaint | Defect | HW | Inlet module | ?
ABL825 FLEX |  |  | CF36DF106A618661 | No |  |  | Cleaned | HW | AutoCheck Module | Barcode reader
ABL825 FLEX | Alle Schläuche und die Ref Elektrode erneuert. Danach 4 kallibrationen durchgeführt, alle einwandfrei. | Alle Schläuche und die Ref Elektrode erneuert. Danach 4 kallibrationen durchgeführt, alle einwandfrei. | SWO-36458 | No | All hoses and the ref electrode renewed. After that 4 callibrations performed, all flawless. | All hoses and the ref electrode renewed. After that 4 callibrations performed, all flawless. | Cleaned | Fluidic pathway | Electrode | Reference | LEH Mail to RDEU
ABL825 FLEX | Verschlauchung EL MET Modul erneuert - Messkammer überprüft - mehrfache Kalibrationen - Funktionsüberprüfung etc. ohne Beanstandung ! | Verschlauchung EL MET Modul erneuert - Messkammer überprüft - mehrfache Kalibrationen - Funktionsüberprüfung etc. ohne Beanstandung ! | SWO-36564 | No | Hose EL MET module renewed - measuring chamber checked - multiple calibrations - function check etc. without complaint ! | Hose EL MET module renewed - measuring chamber checked - multiple calibrations - function check etc. without complaint ! | Cleaned | Fluidic pathway | El/Met Module | Tubings
ABL835 FLEX | Blutreste beseitigt - Erneuerung pH Dichtring - Funktionsüberprüfung etc. | Blutreste beseitigt - Erneuerung pH Dichtring - Funktionsüberprüfung etc. | SWO-36563 | No | Blood residues eliminated - renewal pH sealing ring - function check etc. | Blood residues eliminated - renewal pH sealing ring - function check etc. | Assisted Replacement | Electrode | Electrode | Sealing Ring
AQT90 FLEX | Se detecto que el equipo marcaba en stock 5 determinaciones de TNi y físicamente tenia 2. Se retiro cartucho de TNI y se limpio internamente el equipo. Se comprobó correcto funcionamiento del equipo mediante QC. |  | CF485FA8DE2D1358 | No | It was detected that the equipment marked in stock 5 determinations of TNi and physically had 2. TNI cartridge is removed and the equipment is cleaned internally. The equipment was checked to be properly functioning using QC.
AQT90 FLEX | Incubator PCB bestellt | Der Kunde bittet um einen Service. | SWO-36452 | No | Incubator PCB ordered | The customer asks for a service.
ABL80 FLEX CO-OX | Kalibrationen und QC führen nicht zum Erfolg, daher Lösungspack ersetzt, pO2nun im Bereich, leider tritt läuft rote QC Lösung aus dem Gerät, Lösungspack entfernt, Anschlüsse geprüft Gerät gereinigt, Lösungspack installiert, Kal und QC im Bereich aber klare Flüssigkeit unter dem Gerät. Lösungspack entfernt Gerät und Möbel gereinigt, Lösungspack installiert, Kal. QC im Bereich und zeitgleich kein Lösungsaustritt. Sollte der Fehler wieder auftreten wird das Gerät abgeholt. | Kalibrationen und QC führen nicht zum Erfolg, daher Lösungspack ersetzt, pO2nun im Bereich, leider tritt läuft rote QC Lösung aus dem Gerät, Lösungspack entfernt, Anschlüsse geprüft Gerät gereinigt, Lösungspack installiert, Kal und QC im Bereich aber klare Flüssigkeit unter dem Gerät. Lösungspack entfernt Gerät und Möbel gereinigt, Lösungspack installiert, Kal. QC im Bereich und zeitgleich kein Lösungsaustritt. Sollte der Fehler wieder auftreten wird das Gerät abgeholt. | SWO-36673 | No | Calibrations and QC do not lead to success, therefore solution pack replaced, pO2nun in the range, unfortunately runs red QC solution from the device, solution pack removed, connectors tested device cleaned, solution pack installed, Kal and QC in the range but clear liquid under the Device. Solution pack removes device and furniture cleaned, solution pack installed, Kal. QC in the area and at the same time no solution exit. If the error occurs again, the device will be picked up. | Calibrations and QC do not lead to success, therefore solution pack replaced, pO2nun in the range, unfortunately runs red QC solution from the device, solution pack removed, connectors tested device cleaned, solution pack installed, Kal and QC in the range but clear liquid under the Device. Solution pack removes device and furniture cleaned, solution pack installed, Kal. QC in the area and at the same time no solution exit. If the error occurs again, the device will be picked up.
ABL80 FLEX | suggested them the proper QC of Biorad as they want to use 3rd party QC. |  | SWO-37062 | No | suggested them the proper QC of Biorad as they want to use 3rd party QC.
ABL90 FLEX | Display Unit (Rep. Überbrückung) getauscht gegen neue Display Unit, Nacharbeit siehe SWO 31626 System-und Funktionskontrolle i.O. | Display Unit (Rep. Überbrückung) getauscht gegen neue Display Unit, Nacharbeit siehe SWO 31626 System-und Funktionskontrolle i.O. | SWO-36977 | No | Display Unit (Rep. bridging) is exchanged for new display unit, rework see SWO 31626 system and function control in the following. | Display Unit (Rep. bridging) is exchanged for new display unit, rework see SWO 31626 system and function control in the following. | Defect | HW | Touch screen
ABL90 | LOOKED THROUGH DATA LOGS AND SEEN IT HAD A FEW INHOMOGENOUS CAL 3 SOLUTION ERROR 1084. CUSTOMER HAS ALREADY CHANGED SOLUTION PACK I WOULD CHANGE SENSOR CASSETTE BUT CUSTOMER ONLY HAS 1 SENSOR CASSETTE LEFT FOR 4 ANALYZERS. | CLEANED INLET GASKET, INLET PROBE, AND DID 15 BACKFLUSHES. USED SAME SOLUTION PACK. ALSO RESET RADIANCE WITH HELP FROM JERRY KELLY (SR FSE) | CF385803E04F703D | No | LOOKED THROUGH DATA LOGS AND SEEN IT HAD A FEW INHOMOGENOUS CAL 3 SOLUTION ERROR 1084. CUSTOMER HAS ALREADY CHANGED SOLUTION PACK I WOULD CHANGE SENSOR CASSETTE BUT CUSTOMER ONLY HAS 1 SENSOR CASSETTE LEFT FOR 4 ANALYZERS. | CLEANED INLET GASKET, INLET PROBE, AND DID 15 BACKFLUSHES. USED SAME SOLUTION PACK. ALSO RESET RADIANCE WITH HELP FROM JERRY KELLY (SR FSE) | Cleaned | Fluidic pathway | 15 times
ABL90 | Inaspected casette | Replaced the SC | CF37ED20E04A7056 | No | Inaspected cottages | Replaced the SC | Assisted Replacement | SC | pCO2
ABL90 | Confirmed problem. | Replaced the above parts. Analyzer restarted with the same problems. Flushed analyzer with 20CC syringe filled with hot tap water. Analyzer restarted and calibrated but pCO2 failed on internal QC. Replaced the sensor cassette. | CF3899E7E04849D7 | No | Confirmed problem. | Replaced the above parts. Analyzer restarted with the same problems. Flushed analyzer with 20CC syringe filled with hot tap water. Analyzer restarted and calibrated but pCO2 failed on internal QC. Replaced the sensor cassette. | Assisted Replacement | SC | Wetability
ABL90 FLEX | Oximetriefehler bei Pat.-Messung--> Fehler tritt sporadisch auf Oximetrietest durchgeführt Hemolyzer bestellt | Hemolyzer erneuert--> Oxi-Test durchgeführt--> Kal., tHb-Kal. und QC | SWO-36909 | No | Oximetry error in pat.-measurement--> error occurs sporadically on oximetry test performed hemolyzer ordered | Hemolyzer renews--> Oxi-Test performed--> Kal., tHb-Kal. and QC | Defect | HW | Hemolyzer
ABL90 | Unit would not accept SC. | Replaced SC. | CF382E75E01F4469 | No | Unit would not accept SC. | Replaced SC. | Assisted Replacement | SC | Chip error
ABL90 | Replaced SP. | Replace Sensor Cassette. | CF382BF8E02F4465 | No | Replaced SP. | Replace Sensor Cassette. | Assisted Replacement | SC | Wetability
ABL825 FLEX | checked data logs for errors, checked electrodes and electrode membranes, checked valves and liquid sensors | removed pH BG elctrodes and cleaned electrode chambers, replaced PCO2 membrane, removed modules and cleaned dual housing valves, replaced connecting tubing | CF389B37E03BD8D2 | No | checked data logs for errors, checked electrodes and electrode membranes, checked valves and liquid sensors | removed pH BG elctrodes and cleaned electrode chambers, replaced PCO2 membrane, removed modules and cleaned dual housing valves, replaced connecting tubing | Cleaned | Fluidic pathway | -
ABL827 FLEX | MET-II-Modul (alt:R200N004 neu: R248N002) leihweise ersetzt, Inbetriebnahme, Lösungspumpenschlauch ersetzt, korrekte Kal abgewartet, QK | MET-II-Modul (alt:R200N004 neu: R248N002) leihweise ersetzt, Inbetriebnahme, Lösungspumpenschlauch ersetzt, korrekte Kal abgewartet, QK | SWO-36647 | No | MET-II module (old:R200N004 new: R248N002) replaces on loan, commissioning, solution pump hose replaced, correct calf waited, QK | MET-II module (old:R200N004 new: R248N002) replaces on loan, commissioning, solution pump hose replaced, correct calf waited, QK | Defect | HW | Met II Module | Pump tubing
ABL825 FLEX |  |  | CF383953E00E253F | No |  |  | Defect | HW | Inlet module | Friction detector unit
ABL835 FLEX | Bildschirmrahmen sauber. Neu Verbinden der Kabel kein Erfolg. Nach Neustart lässt sich Touchscreen bedienen. Neuen Bildschirm bestellt, da Fehler bereits mehrfach aufgetreten ist. | Bildschirmrahmen sauber. Neu Verbinden der Kabel kein Erfolg. Nach Neustart lässt sich Touchscreen bedienen. Neuen Bildschirm bestellt, da Fehler bereits mehrfach aufgetreten ist. | SWO-36866 | No | Screen frame clean. Reconnecting the cables no success. After reboot, touchscreen can be operated. New screen ordered because errors have already occurred multiple times. | Screen frame clean. Reconnecting the cables no success. After reboot, touchscreen can be operated. New screen ordered because errors have already occurred multiple times. | Defect | HW | Touch screen
ABL835 FLEX | Umgebungsdruck angepasst. Gemessen wird 742mmHg. Druck laut Wetteramt 765mmHg. Angepasst auf 762mmHg. | Umgebungsdruck angepasst. Gemessen wird 742mmHg. Druck laut Wetteramt 765mmHg. Angepasst auf 762mmHg. | SWO-36864 | No | ambient pressure. Measured is 742mmHg. Pressure according to weather office 765mmHg. Adapted to 762mmHg. | ambient pressure. Measured is 742mmHg. Pressure according to weather office 765mmHg. Adapted to 762mmHg. | Setup | SW | Barometric pressure | Barometer
ABL815 FLEX | Fehlenden O-Ring Laktatelektrode eingesetzt. Elektrode neu membraniert. Kalibrationen und Qualitätskontrollen in Ordnung. | Fehlenden O-Ring Laktatelektrode eingesetzt. Elektrode neu membraniert. Kalibrationen und Qualitätskontrollen in Ordnung. | SWO-36863 | No | Missing O-ring lactate electrode. electrode re-membraneed. calibrations and quality controls in order. | Missing O-ring lactate electrode. electrode re-membraneed. calibrations and quality controls in order. | Assisted Replacement | Electrode | Electrode | Lac
ABL835 FLEX | Troubleshot problem through service mode. Tried replacing other parts. Spectrometer is bad. | Replaced spectrometer. Performed service thb and 2 pt calibrations. | CF3823BAE019093F | No | Troubleshot problem through service mode. Tried replacing other parts. Spectrometer is bad. | Replaced spectrometer. Performed service thb and 2 pt calibrations. | Defect | HW | El/Met Module | Tubings
ABL810 FLEX | Umgebungsdruck angepasst. Gemessen wird 742mmHg. Druck laut Wetteramt 765mmHg. Angepasst auf 762mmHg. | Umgebungsdruck angepasst. Gemessen wird 742mmHg. Druck laut Wetteramt 765mmHg. Angepasst auf 762mmHg. | SWO-36854 | No | ambient pressure. Measured is 742mmHg. Pressure according to weather office 765mmHg. Adapted to 762mmHg. | ambient pressure. Measured is 742mmHg. Pressure according to weather office 765mmHg. Adapted to 762mmHg. | Setup | SW | Barometric pressure | Barometer
AQT90 FLEX | Gerät zerlegt und Platine gewechselt Incubator Calibration durchgeführt, inkl. mechanischem Testlauf Gerät temperiert | Gerät zerlegt und Platine gewechselt Incubator Calibration durchgeführt, inkl. mechanischem Testlauf Gerät temperiert | SWO-36769 | No | Device disassembled and board changed Incubator Calibration performed, incl. mechanical test run device tempered | Device disassembled and board changed Incubator Calibration performed, incl. mechanical test run device tempered
ABL80 CO-OX | Performed oxi diagnostics and neon position failed. Rebooted and performed oxi diagnostics, still failing for neon position. | Replaced spectrometer which corrected the 1021 error. | CF3D1B31B8336FBD | No | Performed oxi diagnostics and neon position failed. Rebooted and performed oxi diagnostics, still failing for neon position. | Replaced spectrometer which corrected the 1021 error.
ABL90 FLEX | Scanner erneuert | Scanner erneuert | SWO-36962 | No | Scanner renewed | Scanner renewed | Defect | HW | Barcode Reader
ABL90 |  |  | CF574836663B4521 | No
ABL90 FLEX | Hemolyzer erneuert--> Oxi-Test durchgeführt--> Kal., tHb-Kal. und QC | Hemolyzer erneuert--> Oxi-Test durchgeführt--> Kal., tHb-Kal. und QC | SWO-36905 | Yes | Hemolyzer renews--> Oxi-Test performed--> Kal., tHb-Kal. and QC | Hemolyzer renews--> Oxi-Test performed--> Kal., tHb-Kal. and QC | Re classified RMED | Double entry
ABL90 | SE REVISA TODO EL SISTEMA NEUMATICO DESMONTAJE E INSPECCION VISUAL, NO ENCONTRANDO NINGUN PROBLEMA SE SUSTITUYE SP AUNQUE YA FUESE CAMBIADO POR EL CLIENTE Y SIGUE EL PROBLEMA CON DIFERENTES MENSAJES SE SUSTITUYE SC Y NO SE SOLUCIONA, SE MOINTA SP ANTIGUO Y DESAPARECEN LOS PROBLEMAS. |  | CF392E7D563E4560 | No | ALL DISASSEMBLY AND VISUAL INSPECTION NEUMATIC SYSTEM IS REVIEWED, NOT FINDING ANY PROBLEM IS SUBSTITUTE SP EVEN IF IT WAS CHANGED BY THE CUSTOMER AND FOLLOWS THE PROBLEM WITH DIFFERENT MESSAGES IS SUBSTITUTES AND IS NOT SOLVED, MOINTA SP ANTIGUO AND Problems. |  | Unknown | Fluid Transport | Leak
ABL825 FLEX | checked cal log and valve | replaced upper flat valve in el ml modual | CF39469256243747 | No | checked cal log and valve | replaced upper flat valve in el ml modual | Defect | HW | El/Met Module | Tubings
ABL825 FLEX | Probehalber Tausch des Displays und Kabel | Probehalber Tausch des Displays und Kabel | SWO-36938 | No | Sample replacement of the display and cable | Sample replacement of the display and cable | Defect | HW | Touch screen
ABL835 FLEX |  |  | CF399B8C560C3897 | No |  |  | Defect | HW | El/Met Module | Tubings
ABL815 FLEX | Umgebungsdruck angepasst. Gemessen wird 742mmHg. Druck laut Wetteramt 765mmHg. Angepasst auf 762mmHg. | Umgebungsdruck angepasst. Gemessen wird 742mmHg. Druck laut Wetteramt 765mmHg. Angepasst auf 762mmHg. | SWO-36865 | No | ambient pressure. Measured is 742mmHg. Pressure according to weather office 765mmHg. Adapted to 762mmHg. | ambient pressure. Measured is 742mmHg. Pressure according to weather office 765mmHg. Adapted to 762mmHg. | Setup | SW | Barometric pressure | Barometer
ABL810 FLEX | Umgebungsdruck angepasst. Gemessen wird 742mmHg. Druck laut Wetteramt 765mmHg. Angepasst auf 762mmHg. Qualitätskontrollen in Ordnung. | Umgebungsdruck angepasst. Gemessen wird 742mmHg. Druck laut Wetteramt 765mmHg. Angepasst auf 762mmHg. Qualitätskontrollen in Ordnung. | SWO-36853 | No | ambient pressure. Measured is 742mmHg. Pressure according to weather office 765mmHg. Adapted to 762mmHg. Quality controls in order. | ambient pressure. Measured is 742mmHg. Pressure according to weather office 765mmHg. Adapted to 762mmHg. Quality controls in order. | Setup | SW | Barometric pressure | Barometer
ABL810 FLEX | Umgebungsdruck angepasst. Gemessen wird 742mmHg. Druck laut Wetteramt 765mmHg. Angepasst auf 762mmHg. | Umgebungsdruck angepasst. Gemessen wird 742mmHg. Druck laut Wetteramt 765mmHg. Angepasst auf 762mmHg. | SWO-36852 | No | ambient pressure. Measured is 742mmHg. Pressure according to weather office 765mmHg. Adapted to 762mmHg. | ambient pressure. Measured is 742mmHg. Pressure according to weather office 765mmHg. Adapted to 762mmHg. | Setup | SW | Barometric pressure | Barometer
TCM-COMBIM |  |  | CF392BF2565AA021 | No
TCM/Tosca |  |  | CF396DA3561F2F06 | No
ABL90 FLEX | ABL überprüft. Gerät mehrfach gespült. Lösungspack erneut erneuert. Gerätestartphase durchgeführt (Kal und QC's). Auf Kundenwunsch SC90 vor Ort gelassen, falls weitere Probleme auftreten. Wochenendeinsatz genehmigt durch Oberarzt Dr. Schooldermann. | ABL überprüft. Gerät mehrfach gespült. Lösungspack erneut erneuert. Gerätestartphase durchgeführt (Kal und QC's). Auf Kundenwunsch SC90 vor Ort gelassen, falls weitere Probleme auftreten. Wochenendeinsatz genehmigt durch Oberarzt Dr. Schooldermann. | SWO-37085 | No | ABL. device rinsed several times. Solution pack renewed. device start phase (Kal and QC's). At the customer's request SC90 left on site in case of further problems. Weekend assignment approved by Senior Physician Dr. Schooldermann. | ABL. device rinsed several times. Solution pack renewed. device start phase (Kal and QC's). At the customer's request SC90 left on site in case of further problems. Weekend assignment approved by Senior Physician Dr. Schooldermann. | Assisted Replacement | SC | Wetability
ABL90 |  |  | CF3D856EB80E9A8F | No |  |  | Cleaned | Fluidic pathway
ABL90 FLEX | Remplacé la FC ( flash card ) garantie Configuration du système e.o Chargé le software 3.1 MR3 e.o Calibration e.o contrôle de qualité e.o | Remplacé la FC ( flash card ) garantie Configuration du système e.o Chargé le software 3.1 MR3 e.o Calibration e.o contrôle de qualité e.o | SWO-37239 | No | Replaced the FC (flash card) guarantee System configuration e.o Loaded software 3.1 MR3 e.o Calibration e.o quality control e.o | Replaced the FC (flash card) guarantee System configuration e.o Loaded software 3.1 MR3 e.o Calibration e.o quality control e.o | Defect | HW | CF card
ABL90 | LOOKED THROUGH DATALOGS, 2 DAYS AFTER SENSOR CASSETTE WAS CHANGED INTERNAL QC LEVEL 3 FAILING READING 103, 100, 99, RANGE IS 106 TO 139. | REPLACED SENSOR CASSETTE WITH CUSTOMERS SUPPLY. DID A FEW WASTE BLOODS. ALSO PRINTER OUT AND SENT CREDIT CLAIM FORM FOR SENSOR CASSETTE. | CF3D6EA2B817A3F1 | No | LOOKED THROUGH DATALOGS, 2 DAYS AFTER SENSOR CASSETTE WAS CHANGED INTERNAL QC LEVEL 3 FAILING READING 103, 100, 99; RANGE IS 106 TO 139. | REPLACED SENSOR CASSETTE WITH CUSTOMERS SUPPLY. DID A FEW WASTE BLOODS. ALSO PRINTER OUT AND SENT CREDIT CLAIM FORM FOR SENSOR CASSETTE. | Assisted Replacement | SC | Glu
ABL90 FLEX | Analysator zerlegt und alle Schläuche ausgetauscht (Verstopfung) Initialisierung neuer Sensor. Medizintechniker beobachtet |  | SWO-37178 | No | Analyzer disassembled and all hoses replaced (constipation) initialization of new sensor. Medical technicians observed |  | Cleaned | Fluidic pathway
ABL90 | HCT on the ABL has a slightly higher bias when compared with the Hemotology Lab Analyzer. | Don and I came in and ran 5 sets of controls for each level on each analyzer Ran ten blood samples. The results were submitted to James Halbleib for further study. | CF3E16712E5CBA86 | No | HCT on the ABL has a slightly higher bias when compared with the Hemotology Lab Analyzer. | Don and I came in and ran 5 sets of controls for each level on each analyzer Ran ten blood samples. The results were submitted to James Halbleib for further study. | Performance | Other | HCT
ABL90 FLEX | Flüssigkeitsstrecke gespült, SP im Serviceprogramm entriegelt --> SP neu eingesetzt --> Neustart | Flüssigkeitsstrecke gespült, SP im Serviceprogramm entriegelt --> SP neu eingesetzt --> Neustart | SWO-37262 | No | Liquid track rinsed, SP unlocked in service program --> SP re-inserted --> Restart | Liquid track rinsed, SP unlocked in service program --> SP re-inserted --> Restart | Cleaned | Fluidic pathway
ABL90 FLEX | 2x LP ABL90 geliefert Während SW-Update getauscht | 2x LP ABL90 geliefert Während SW-Update getauscht | SWO-37254 | Yes | 2x LP ABL90 delivered During SW update exchanged | 2x LP ABL90 delivered During SW update exchanged | Re classified RMED | Logistical upSWO
ABL90 FLEX | Fehlersuche, Hemolyzer mit div. Lösungen gereinigt, div. Tets`s im Serviceprogr. durchgef., Kal. + QC i.O. | Fehlersuche; Hemolyzer mit div. Lösungen gereinigt; div. Tets`s im Serviceprogr. durchgef.; Kal. + QC i.O. | SWO-37298 | No | Troubleshooting; Hemolyzer cleaned with various solutions; div. Tets's in the service through it; Kal. + QC in the first place | Troubleshooting; Hemolyzer cleaned with various solutions; div. Tets's in the service through it; Kal. + QC in the first place | Cleaned | Fluidic pathway | Oxi module
ABL90 | Unable to pump solutions Error 1343. Ran all service checks. Vaccum errors on pressure test. LS errors.SP failed. Tried replacing the Inlet module but no difference. | Replaced probe and gasket. | CF3FA064A401DAC7 | No | Unable to pump solutions Error 1343. Ran all service checks. Vaccum errors on pressure test. LS errors. SP failed. Tried replacing the Inlet module but no difference. | Replaced probe and gasket. | Defect | HW | Connector gasket
ABL827 FLEX |  |  | CF3D322DB85DAD93 | No |  |  | Defect | HW | El/Met Module | Tubings
ABL825 FLEX | Pumprotor ok LS werte ok Einlass Dichtung ersetzt .Einlassbereich gereinigt Pump Kal ok | Pumprotor ok LS werte ok Einlass Dichtung ersetzt .Einlassbereich gereinigt Pump Kal ok | SWO-37439 | No | Pump rotor ok LS values ok inlet seal replaced. Inlet area cleaned Pump Kal ok | Pump rotor ok LS values ok inlet seal replaced. Inlet area cleaned Pump Kal ok | Defect | HW | Oxi module | Pump tubing | Asymetri
ABL820 FLEX | DATA LOGS SHOW 247 ATO D FAILURES AND 316 NEON LAMP LOW AND 248 SPECTROMETER MISSING TRIGGER ERROR | REPLACED OXI MODULE PCB.REPLACED POWER SUPPLY BUT THE NEW POWER SUPPLY WAS DEFECTIVE FROM STOCK AS THE ANALYZER WOULD NOT POWER UP.REINSTALLED ORIGINAL POWER SUPPLY.PERFORMED THE COMPLETE OXIMETRY PROGRAM.PERFORMED CALIBRATIONS.IF ERRORS RETURN THEN THE POWER SUPPLY SHOULD BE CHANGED. | CF3D1BD8B8106EF4 | No | DATA LOGS SHOW 247 ATO D FAILURES AND 316 NEON LAMP LOW AND 248 SPECTROMETER MISSING TRIGGER ERROR | REPLACED OXI MODULE PCB. REPLACED POWER SUPPLY BUT THE NEW POWER SUPPLY WAS DEFECTIVE FROM STOCK AS THE ANALYZER WOULD NOT POWER UP. REINSTALLED ORIGINAL POWER SUPPLY. PERFORMED THE COMPLETE OXIMETRY PROGRAM. PERFORMED CALIBRATIONS. IF ERRORS RETURN THEN THE POWER SUPPLY SHOULD BE CHANGED. | Defect | HW | El/Met Module | Tubings
ABL800 BASIC | Gerät neu gestartet - Display bleibt dunkeln - Bootet nicht - AU arbeitet - Display getauscht - ohne Erfolg DMS def. - E-Teil bestellt | Der Kunde bittet um einen Service. | SWO-37463 | No | Device restarted - display remains dark - Bootet not - AU works - display swapped - without success DMS def. - E-part ordered | The customer asks for a service. | Unknown | HW | DMS_PC | Touch screen
ABL815 FLEX | Abfallsensor gew., Kal. + QC i.O. | Abfallsensor gew.; Kal. + QC i.O. | SWO-37192 | No | Abfallsensor gew.; Kal. + QC i.O. | Abfallsensor gew.; Kal. + QC i.O. | Defect | HW | Basic Unit | Waste detector
AQT90 FLEX | TnT Testkassette ist auf der Barcodeseite beschädigt. Aufgeklebte Barcodeetikette beschädigt und etwas Das Gehäuse der Testkassette. Testkassette entfernt, sie hat das storage wheel beim Drehen behindert. Eine Schwarze Plastik Führung herausgehebelt, wieder eingeklickt. Funktion des storage wheels wieder o.K. Analysator ohne Beanstandung und messbereit. | Frau Tox meldet nach einlegen einer Tropkassette dass der Analysator keine Reagenzien mehr anzeigt und in Fehler geht. Telefonische Hilfe nicht erfolgreich. Frau Tox wünscht keinen kostenpflichtigen Einsatz am Wochenende. Akutservice soll am Montagmorgen, 13.07.15 erfolgen! Bitte Kundin bescheid geben wann sie mit einem Techniker rechnen kann. Analysatro sollte wenn möglich abends wieder laufen! | SWO-37307 | No | TnT test cartridge is damaged on the barcode side. Glued barcode label damaged and somewhat the housing of the test cartridge. removed from the test cartridge, it has obstructed the storage wheel during turning. A black plastic guide is lifted out, clicked back in. Function of the storage wheel again o.K. analyzer without complaint and ready to measure. | After inserting a trophy, Ms. Tox reports that the analyzer no longer displays reagents and makes mistakes. Telephone help not successful. Ms. Tox does not want a paid bet on weekends. Acute service is scheduled to take place on Monday morning, 13.07.15! Please let the customer know when she can count on a technician. Analysatro should run again in the evening if possible!
AQT90 FLEX | Opened the machine . Cleaned inside of the machine . cleaned the barcode reader, pcb and connector wire . Adjusted the position of the barcode reader Completed system cleaning procedure |  | SWO-37932 | No | Opened the machine . Cleaned inside of the machine . cleaned the barcode reader, pcb and connector wire . Adjusted the position of the barcode reader Completed system cleaning procedure
ABL80 FLEX | need to change the data signal board | need to change the data signal board | SWO-38110 | No | need to change the data signal board | need to change the data signal board
ABL90 | Customer had already replaced the solution pack, and performed the back flush, but this did not resolve the problem. Entered the fluid transport program in service and found that the solution pack had failed. I was able to get the solution pack to pass, but this still did not resolve the problem. | Replaced inlet gasket with holder and the inlet connector gasket, but this did not resolve the problem. Replaced the sensor cassette (provided by customer), and this resolved the problem. | CF3E67E12E5F4196 | No | Customer had already replaced the solution pack, and performed the back flush, but this did not resolve the problem. Entered the fluid transport program in service and found that the solution pack had failed. I was able to get the solution pack to pass, but this still did not resolve the problem. | Replaced inlet gasket with holder and the inlet connector gasket, but this did not resolve the problem. Replaced the sensor cassette (provided by customer), and this resolved the problem. | Assisted Replacement | SC | Wetability
ABL827 FLEX | Inlet gasket replaced by customer. Ran inlet calibration diagnostics wtih failed results. Six membranes due for replacement. | Replaced friction detection unit. Ran inlet calibration with pass results. Replaced membranes that needed changed. Crea not calibrating, replaced those membranes also. Restart calibrations. | CF3E6E172E20B9B2 | No | Inlet gasket replaced by customer. Ran inlet calibration diagnostics wtih failed results. Six membranes due for replacement. | Replaced friction detection unit. Ran inlet calibration with pass results. Replaced membranes that needed changed. Crea not calibrating, replaced those membranes also. Restart calibrations. | Defect | HW | Inlet module | Friction detector unit
ABL820 FLEX | Unit failing leakage test. Found inlet tube crimped | Replaced inlet tube. | CF3FD12BA4401D86 | No | Unit failing leakage test. Found inlet tube crimped | Replaced inlet tube. | Defect | HW | Inlet module | Tubings
ABL825 FLEX | Eintrag in Run.Roma.cmd nicht möglich ( Analysator macht Neustart beim Versuch Datei zu öffnen ) - Funktionsüberprüfung etc. - Analysator voll funktionsfähig, nach Abschaltung kurze Ruhepause vor Neustart einhalten ! | Eintrag in Run.Roma.cmd nicht möglich ( Analysator macht Neustart beim Versuch Datei zu öffnen ) - Funktionsüberprüfung etc. - Analysator voll funktionsfähig, nach Abschaltung kurze Ruhepause vor Neustart einhalten ! | SWO-37445 | No | Entry in Run.Roma.cmd not possible ( Analyzer makes restart when trying to open file ) - Function check etc. - Analyzer fully functional, after shutdown observe short rest before restart ! | Entry in Run.Roma.cmd not possible ( Analyzer makes restart when trying to open file ) - Function check etc. - Analyzer fully functional, after shutdown observe short rest before restart ! | Unknown | HW | Analyzer Control PCB | Awaiting Email respons
ABL825 FLEX | ACModul demontiert, Scanner Kabel getrennt und wieder gesteckt, Scanner leuchtet wieder, Messung mit Scannertest i.O. | ACModul demontiert, Scanner Kabel getrennt und wieder gesteckt, Scanner leuchtet wieder, Messung mit Scannertest i.O. | SWO-37502 | No | ACModule dismantled, scanner cable disconnected and plugged back in, scanner lights up again, measurement with scanner test i.O. | ACModule dismantled, scanner cable disconnected and plugged back in, scanner lights up again, measurement with scanner test i.O. | Defect | HW | Basic Unit | Barcode reader | Cable
ABL80 |  |  | CF437940063A08AA | No
ABL90 | bend in probe, softwarte is at 3.1MR1 | i replaced the inlet probe and upgraded the software to 3.1 MR3 | CF3FBFC1A4010A0F | No | bend in probe, softwarte is at 3.1MR1 | i replaced the inlet probe and upgraded the software to 3.1 MR3 | Replaced | HW | Inlet Probe
ABL90 FLEX | SC Kassette und Einlassdichtung ersetzt. Startphase QC und Kalibration. | SC Kassette und Einlassdichtung ersetzt. Startphase QC und Kalibration. | SWO-37678 | No | SC cassette and inlet seal replaced. Start phase QC and calibration. | SC cassette and inlet seal replaced. Start phase QC and calibration. | Assisted Replacement | SC | Wetability
ABL90 FLEX | allgemeine Sicht- und Funktionskontrolle, Reinigung Desinfektion, Einlassdichtung und Anschlussdichtung ersetzt, Lösungspack ersetzt da Lsg Kal3 weiterhin inhomogen, Test im Service Menü nun erfolgreich, Neustart, Kalibration, QC Kontrolle | allgemeine Sicht- und Funktionskontrolle, Reinigung Desinfektion, Einlassdichtung und Anschlussdichtung ersetzt, Lösungspack ersetzt da Lsg Kal3 weiterhin inhomogen, Test im Service Menü nun erfolgreich, Neustart, Kalibration, QC Kontrolle | SWO-37724 | No | general visual and functional control, cleaning disinfection, inlet seal and connection seal replaced, solution pack replaces as Lsg Kal3 continues to be inhomogeneous, test in the service menu now successful, restart, calibration, QC control | general visual and functional control, cleaning disinfection, inlet seal and connection seal replaced, solution pack replaces as Lsg Kal3 continues to be inhomogeneous, test in the service menu now successful, restart, calibration, QC control | Assisted Replacement | SP
ABL90 FLEX | LP aus Kundenbestand ersetzt, Leak Voltage im Serviceprogramm überprüft, Werte OK. Funktion Z.Zt OK. | LP aus Kundenbestand ersetzt, Leak Voltage im Serviceprogramm überprüft, Werte OK. Funktion Z.Zt OK. | SWO-37938 | No | LP replaced from customer inventory, Leak Voltage checked in service program, values OK. Function Zt.Zt OK. | LP replaced from customer inventory, Leak Voltage checked in service program, values OK. Function Zt.Zt OK. | Assisted Replacement | SP | Leak current
ABL90 | Solution pack connecting Disconnecting. | Replaced the Solution Pack Id Reload software(This solved the connection problem). I had to replace the Solution pack also. The analyzer continued to drift on Pc02, because it has been down for over a day. I also found a mis-alightment in the Inlet that was causing air detection errors. | CF3FEA41A41B878A | No | Solution pack connecting Disconnecting. | Replaced the Solution Pack Id Reload software(This solved the connection problem). I had to replace the Solution pack also. The analyzer continued to drift on Pc02, because it has been down for over a day. I also found a mis-alightment in the Inlet that was causing air detection errors. | Defect | HW | SP PCB
ABL90 FLEX | Fehlersuche, Spülung ausgelöst, Kal. + QC i.O. | Fehlersuche; Spülung ausgelöst; Kal. + QC i.O. | SWO-37693 | No | Troubleshooting; rinse triggered; Kal. + QC in the first place | Troubleshooting; rinse triggered; Kal. + QC in the first place | Assisted Replacement | SC | Wetability
ABL90 | Review data logs, observed problem first began on 7 16 2015 @ approximately 4:46PM. Sean Burgess and Frank Castenado were completing installation and training when problem developed. Problem developed running Lac, began to fail calibration, QC and linearity testing due to errors 1310, 377, 378. | On 7 16, several SC and SP were tried with no success. 7 27 - attempted troubleshooting, cleaned SC pins, replaced SC, no resolution. Most likely cause is sensor interface PCB. | CF3FE33EA4169F6D | No | Review data logs, observed problem first began on 7 16 2015 @ approximately 4:46PM. Sean Burgess and Frank Castenado were completing installation and training when problem developed. Problem developed running Lac, began to fail calibration, QC and linearity testing due to errors 1310, 377, 378. | On 7 16, several SC and SP were tried with no success. 7 27 - attempted troubleshooting, cleaned SC pins, replaced SC, no resolution. Most likely cause is sensor interface PCB. | Defect | HW | SC PCB
ABL90 | Inspection revealed display LCD is cracked from top to bottom. No visible damage to bezel or touch screen. | Per customer she is refusing this analyzer. Contacted ASM and sales both indicate another unit will need to be ordered, do not install unit. | CF3FB054A4490DAE | Yes | Inspection revealed display LCD is cracked from top to bottom. No visible damage to bezel or touch screen. | Per customer she is refusing this analyzer. Contacted ASM and sales both indicate another unit will need to be ordered, do not install unit. | Re classified RMED | OOB
ABL800 BASIC | Mainbord getauscht, Startphase, Ref, Glu und Lac neu membraniert, Kal und QC durchgeführt | Mainbord getauscht, Startphase, Ref, Glu und Lac neu membraniert, Kal und QC durchgeführt | SWO-37595 | No | Mainbord swapped, start phase, ref, glu and lac re-membraneed, Kal and QC performed | Mainbord swapped, start phase, ref, glu and lac re-membraneed, Kal and QC performed | Defect | HW | DMS_PC | Main Cpu
ABL827 FLEX |  |  | CF4256AC902578BD | No |  |  | Assisted Replacement | HW | Fluidic manifold | Pump tubing
TCM400-3 | Upon talking to the wound care Tech, Temp was being confused with TcpO2. Explained the operation of the monitor. | Explained the operation of the monitor, when electrodes are initialy mounted on the patient. | CF3FD46CA4541D8A | No | Upon talking to the wound care Tech, Temp was being confused with TcpO2. Explained the operation of the monitor. | Explained the operation of the monitor, when electrodes are initialy mounted on the patient.
ABL90 FLEX | Sensorkassette erneuert Kalibrationsfehler und QC durchgeführt | Sensorkassette erneuert Kalibrationsfehler und QC durchgeführt | SWO-37983 | No | Sensor cassette renews calibration error and QC performed | Sensor cassette renews calibration error and QC performed | Assisted Replacement | SC
ABL90 FLEX | Analysator gespült - Service Tests und Justagen durchgeführt - danach 1245 Flüssigkeitstransport gescheitert und 1086 Inhomogene QC 2 Lösung - Erneuerung Inlet Connector Gasket - Lösungspack nochmals inst. - Funktionsüberprüfung ohne Beanstandung ! | Analysator gespült - Service Tests und Justagen durchgeführt - danach 1245 Flüssigkeitstransport gescheitert und 1086 Inhomogene QC 2 Lösung - Erneuerung Inlet Connector Gasket - Lösungspack nochmals inst. - Funktionsüberprüfung ohne Beanstandung ! | SWO-37917 | No | Analyzer rinsed - service tests and adjustments carried out - then 1245 liquid transport failed and 1086 Inhomogeneous QC 2 solution - renewal Inlet Connector Gasket - solution pack again inst. - Functional check without complaint ! | Analyzer rinsed - service tests and adjustments carried out - then 1245 liquid transport failed and 1086 Inhomogeneous QC 2 solution - renewal Inlet Connector Gasket - solution pack again inst. - Functional check without complaint ! | Defect | HW | Inlet connector gasket
ABL90 FLEX PLUS | Reinigung, Kontrolle Einlaß und SP90, Serviceprogramme kontrolliert, Startphase, Einlaß erneut geprüft und gereinigt, Startphase abgewartet | Reinigung, Kontrolle Einlaß und SP90, Serviceprogramme kontrolliert, Startphase, Einlaß erneut geprüft und gereinigt, Startphase abgewartet | SWO-37790 | No | Cleaning, control inlet and SP90, service programs checked, start-up phase, inlet checked and cleaned again, start phase awaited | Cleaning, control inlet and SP90, service programs checked, start-up phase, inlet checked and cleaned again, start phase awaited | Cleaned | Fluidic pathway
ABL827 FLEX |  |  | CF5139B6181D4255 | Yes |  |  | Re classified RMED | No info
ABL825 FLEX |  |  | CF40F3131A496674 | No |  |  | Assisted Replacement | Electrode | Electrode | pCO2
ABL827 FLEX | Inlet probe is in need of replacement | Replaced inlet probe and ensured that it was properly seated and functioning. Disabled the following parameters after discovering dummy electrodes in the measuring chamber: Cl-, Lac, Na+, K+. Site confirmed that these parameters are not in use but they were not yet disabled. Instrument is within acceptable operating specifications. | CF4112331A48B674 | No | Inlet probe is in need of replacement | Replaced inlet probe and ensured that it was properly seated and functioning. Disabled the following parameters after discovering dummy electrodes in the measuring chamber: Cl-, Lac, Na+, K+. Site confirmed that these parameters are not in use but they were not yet disabled. Instrument is within acceptable operating specifications. | Defect | HW | Inlet module | Probe bend
ABL825 FLEX | Nullstaromkalibration im Service durchgeführt, Luftblase entfernt, Kalibrationen und Qualitätskontrollen ohne Beanstandungen durchgeführt. | Nullstaromkalibration im Service durchgeführt, Luftblase entfernt, Kalibrationen und Qualitätskontrollen ohne Beanstandungen durchgeführt. | SWO-37828 | No | Zero starom calibration performed in service, air bubble removed, calibrations and quality controls performed without complaints. | Zero starom calibration performed in service, air bubble removed, calibrations and quality controls performed without complaints. | Cleaned | Electrode | Membrane | pO2
ABL825 FLEX | Einstellungen angepasst. | Einstellungen angepasst. | SWO-37824 | Yes | Adjusted settings. | Adjusted settings. | Re classified RMED | Option key exp
ABL825 FLEX | Auf Verdacht Pump Tübingen BG Modul ersetzt.pH Dichtring erneuert QC blau jetzt okay | Auf Verdacht Pump Tübingen BG Modul ersetzt.pH Dichtring erneuert QC blau jetzt okay | SWO-37889 | No | Suspected Pump Tübingen BG module ersetzt.pH sealing ring renewed QC blue now okay | Suspected Pump Tübingen BG module ersetzt.pH sealing ring renewed QC blue now okay | Assisted Replacement | Electrode | Electrode | pH Sealing ring
ABL800 BASIC | Gerät weiterhin keine Anzeige, CPU Unit ersetzt, Neustart durchgeführt, keine Verbindung zum Accure, Mac Adresse ausgelesen und Hr. Manfred EDV Abteilung mitgeteilt, Mac Filter wurde frei geschaltet, Verbing i.O. Kal und QC durchgeführt | Gerät neu gestartet - Display bleibt dunkeln - Bootet nicht - AU arbeitet - Display getauscht - ohne Erfolg DMS def. - E-Teil bestellt | SWO-37942 | No | Device still no display, CPU unit replaced, restart performed, no connection to Accure, Mac address read out and Mr. Manfred EDV department informed, Mac filter was activated, verbing i.O. Kal and QC performed | Device restarted - display remains dark - Bootet not - AU works - display swapped - without success DMS def. - E-part ordered | Unknown | HW | DMS_PC | ?
ABL825 FLEX | Unable to locate any problems with Pump Cals. | Replace Roller Assy. | CF40F9C31A417DF2 | No | Unable to locate any problems with Pump Cals. | Replace Roller Assy. | Defect | HW | El/Met Module | Pump rotor
ABL800 BASIC | Po2 membrane need replacement.Replace the po2 membrane.After replace the po2 membrane unit is working properly. | Po2 membrane need replacement.Replace the po2 membrane.After replace the po2 membrane unit is working properly. | SWO-38134 | No | Po2 membrane need replacement. Replace the po2 membrane. After replace the po2 membrane unit is working properly. | Po2 membrane need replacement. Replace the po2 membrane. After replace the po2 membrane unit is working properly. | Assisted Replacement | Electrode | Membrane | pO2
AQT90 FLEX |  |  | CF41FB539002C864 | No
ABL90 FLEX | Fehlersuche Überprüfung fester Sitz Kabel Neustart | Fehlersuche Überprüfung fester Sitz Kabel Neustart | SWO-38048 | No | Troubleshooting Check Fixed Seat Cable Restart | Troubleshooting Check Fixed Seat Cable Restart | Adjustment | Display | Cable
ABL90 FLEX | Gerät wurde zwei Tage vom Kunde getestet, keine Beanstandungen, Leihgerät wieder abgeholt | Gerät wurde zwei Tage vom Kunde getestet, keine Beanstandungen, Leihgerät wieder abgeholt | SWO-37991 | Yes | Device was tested by the customer for two days, no complaints, rental device picked up again | Device was tested by the customer for two days, no complaints, rental device picked up again | Re classified RMED | Analyzer re installed
ABL90 FLEX | Überprüfung im Service, System gespült, Einlasssonde ersetzt, Test wiederholen , LP ersetzt (Kundenbestand), Kal u. QC i.O. | Überprüfung im Service, System gespült, Einlasssonde ersetzt, Test wiederholen , LP ersetzt (Kundenbestand), Kal u. QC i.O. | SWO-38028 | No | Inspection in service, system flushed, inlet probe replaced, test repeat, LP replaced (customer stock), Kal & QC i.O. | Inspection in service, system flushed, inlet probe replaced, test repeat, LP replaced (customer stock), Kal & QC i.O. | Replaced | HW | Inlet Probe
ABL90 FLEX | alle Fließwege mehrfach gespült - keine Verstopfung, alle Sensoren kalibriert, Komponenten im Service geprüft - keine Fehler in allen Servicetest, Neustart abgewartet, korrekte Kal abgewartet | Gestern wurde das LP gewechselt, danach hat der ABL eine neue SC verlangt. Für 1 Messung hat es funktioniert. Jetzt soll erneut das LP getauscht werden. Habe Sr. Grit einen Neustart, bzw. den Tausch der Einlaßdichtung empfohlen. Sie möchte einen Service haben, da erst vor kurzem eine Wartung durchgeführt worden ist. Gibt es dort ein Präanalytikproblem? | SWO-38104 | No | all flow paths rinsed several times - no blockage, all sensors calibrated, components checked in service - no errors in all service tests, restart waited, correct calving waited | Yesterday the LP was changed, after which the ABL demanded a new SC. For 1 measurement it worked. Now the LP is to be replaced again. Have sr. Grit recommended a restart or the exchange of the inlet seal. She wants to have a service because maintenance has only recently been carried out. Is there a pre-analysis problem there? | Cleaned | Fluidic pathway
ABL90 FLEX | Vergleichsmessungen ABL8XX und ABL90 durchgeführt, keine nennenswerte Unterschiede feststellbar Konfiguration angepasst |  | SWO-38117 | No | Comparative measurements ABL8XX and ABL90 performed, no significant differences detectable configuration adapted |  | No fault found | tHb cal
ABL90 | Troubleshot and found that the customer was running the wrong lot of level 3 external QC. I also found that the defualt syringe mode would not change to from QC Ampoule. | Ran the correct lot of level 3 external QC, and this resolved the problem. I also set the setting back to Radiometer default, then copied and loaded the settings from their other analyzer. All problems were resolved. | CF421B61904585C8 | No | Troubleshot and found that the customer was running the wrong lot of level 3 external QC. I also found that the defualt syringe mode would not change to from QC Ampoule. | Ran the correct lot of level 3 external QC, and this resolved the problem. I also set the setting back to Radiometer default, then copied and loaded the settings from their other analyzer. All problems were resolved. | User Error | Wrong QC lot
ABL90 |  | Replaced Inlet position detector PCB | CF4F46712C610AE8 | No |  | Replaced Inlet position detector PCB | Defect | HW | Inlet PCB
ABL90 FLEX PLUS |  |  | CF42449D90208159 | Yes |  |  | Re classified RMED | PER
ABL820 FLEX | Performed visual and service diagnostics inspection. No failures reported though diagnostics menu. Suspect small air leak in sample inlet path. Visually inspected electrodes and membranes, noted pCO2 electrode appears to be loose and not secured completely, not a user issue, problem is evident using Membrane 942-063 lot#R0671, exp 5 2017. Tried several with same observation. | Replace sample inlet tubing nipple, preheater blue tube, sample probe blue tube and lower pH BG blue inlet tube. Cleaned all measuring chambers and dummy electrode and lower seal. Ran QC, level 3 pCO2 now at group mean | CF4202F7905A3866 | No | Performed visual and service diagnostics inspection. No failures reported though diagnostics menu. Suspect small air leak in sample inlet path. Visually inspected electrodes and membranes, noted pCO2 electrode appears to be loose and not secured completely, not a user issue, problem is evident using Membrane 942-063 lot#R0671, exp 5 2017. Tried several with same observation. | Replace sample inlet tubing nipple, preheater blue tube, sample probe blue tube and lower pH BG blue inlet tube. Cleaned all measuring chambers and dummy electrode and lower seal. Ran QC, level 3 pCO2 now at group mean | Assisted Replacement | Electrode | Electrode | pCO2
ABL827 FLEX | ABL überprüft. Kommunikationseinstellungen alle wie bei Anbindung vorhanden. Windows Firewall deaktiviert. Gerät kann wieder angepingt werden. Patientenabfrage funktioniert auch wieder. Mit aktivierter Firewall klappt dies Abfrage nicht mehr. | ABL überprüft. Kommunikationseinstellungen alle wie bei Anbindung vorhanden. Windows Firewall deaktiviert. Gerät kann wieder angepingt werden. Patientenabfrage funktioniert auch wieder. Mit aktivierter Firewall klappt dies Abfrage nicht mehr. | SWO-38015 | No | ABL. communication settings all as available with connection. Windows Firewall disabled. Device can be pinged again. Patient consultation also works again. With the firewall enabled, this query no longer works. | ABL. communication settings all as available with connection. Windows Firewall disabled. Device can be pinged again. Patient consultation also works again. With the firewall enabled, this query no longer works. | Setup | HIS/LIS
ABL825 FLEX | Auf dem Analysator ist die Software 6.14 vorhanden. Neu ab dieser Software Version ist, dass der Anwender ein Verbrauchsmaterial scannen muss, sobald er auf ersetzen gegangen ist. Kurzeinweisung durchgef. | Auf dem Analysator ist die Software 6.14 vorhanden. Neu ab dieser Software Version ist, dass der Anwender ein Verbrauchsmaterial scannen muss, sobald er auf ersetzen gegangen ist. Kurzeinweisung durchgef. | SWO-37987 | Yes | The software 6.14 is available on the analyzer. New from this software version is that the user must scan a consumable as soon as he has gone on replacement. Brief briefing. | The software 6.14 is available on the analyzer. New from this software version is that the user must scan a consumable as soon as he has gone on replacement. Brief briefing. | Re classified RMED | Customer training
ABL825 FLEX | Pumpenschläuche ern., Pumpen Kal durchgef., Gerät kalibriert, QC´s durchgef. | Pumpenschläuche ern., Pumpen Kal durchgef., Gerät kalibriert, QC´s durchgef. | SWO-37994 | No | Pump hoses ern., pumps Kal gef., device calibrated, QC's through. | Pump hoses ern., pumps Kal gef., device calibrated, QC's through. | Defect | HW | El/Met Module | Pump tubing
AQT90 FLEX | AQT überprüft. Kommunikationseinstellungen alle wie bei Anbindung vorhanden. Windows Firewall deaktiviert. Gerät kann wieder angepingt werden. Patientenabfrage funktioniert auch wieder. Mit aktivierter Firewall klappt dies Abfrage nicht mehr. | AQT überprüft. Kommunikationseinstellungen alle wie bei Anbindung vorhanden. Windows Firewall deaktiviert. Gerät kann wieder angepingt werden. Patientenabfrage funktioniert auch wieder. Mit aktivierter Firewall klappt dies Abfrage nicht mehr. | SWO-38012 | No | AQT. communication settings all as available with connection. Windows Firewall disabled. Device can be pinged again. Patient consultation also works again. With the firewall enabled, this query no longer works. | AQT. communication settings all as available with connection. Windows Firewall disabled. Device can be pinged again. Patient consultation also works again. With the firewall enabled, this query no longer works.
AQT90 FLEX | Servicemenü lässt sich öffnen => Safe Postions oder anderen Programme nicht ausführbar, keine Kommunikation zur AU Speicherkarte lässt sich über Windowsoberfläche öffnen => vermutlich Applikationssoftware defekt Software 8.8.15 installiert => nach Geräteneustart wird inkonsistente Software angezeigt Nochmalige Softwareunstallation |  | SWO-38122 | No | Service menu can be opened => Safe Postions or other programs not executable; no communication to the AU memory card can be opened via Windows interface => probably application software defective software 8.8.15 installed => after device restart inconsistent software is displayed Again software unstallation
ABL800 BASIC | Ventile von Kal1, Kal2, Reinigungs- und Spüllösung komplett mit Flüssigkeit gefüllt Air Kanal verstopft => Spülung Kanal über V7 im Fluidic Modul Nachfüllung und LS Justage i.O. | Ventile von Kal1, Kal2, Reinigungs- und Spüllösung komplett mit Flüssigkeit gefüllt Air Kanal verstopft => Spülung Kanal über V7 im Fluidic Modul Nachfüllung und LS Justage i.O. | SWO-38157 | No | Valves of Kal1, Kal2, cleaning and rinsing solution completely filled with liquid Air channel clogged => Flushing channel via V7 in Fluidic module refill and LS Adjustment i.O. | Valves of Kal1, Kal2, cleaning and rinsing solution completely filled with liquid Air channel clogged => Flushing channel via V7 in Fluidic module refill and LS Adjustment i.O. | Cleaned | Fluidic pathway
ABL90 FLEX | No Gerät lief vor Ankunft durch telefonische Unterstützung von ALESEP | No Gerät lief vor Ankunft durch telefonische Unterstützung von ALESEP | SWO-38286 | Yes | No device ran before arrival by telephone support from ALESEP | No device ran before arrival by telephone support from ALESEP | Re classified RMED | Telephone support
ABL90 | inspected measuring chamber, sensors, hemolyzer. | replaced measuring chamber, overhaul kit | CF45CACFF211EF2D | No | inspected measuring chamber, sensors, hemolyzer. | replaced measuring chamber, overhaul kit | Defect | HW | Cassette retainer
ABL90 FLEX | Durch Kunden wurde bereits mehrfach LP erneuert Gerätestandort überprüft, Umgebungstemp. relativ hoch Kunden Standortwechsel empfohlen Anwender hat Lüfter zur Kühlung aufgestellt, da Standortwechsel aufgrund Anbindung schwierig | Durch Kunden wurde bereits mehrfach LP erneuert Gerätestandort überprüft, Umgebungstemp. relativ hoch Kunden Standortwechsel empfohlen Anwender hat Lüfter zur Kühlung aufgestellt, da Standortwechsel aufgrund Anbindung schwierig | SWO-38274 | No | Customers have already checked LP renewed device location several times, ambient temp. relatively high customer location change recommended User has installed fans for cooling, because location change due to connection difficult | Customers have already checked LP renewed device location several times, ambient temp. relatively high customer location change recommended User has installed fans for cooling, because location change due to connection difficult | User Error | Temperature above…
ABL90 FLEX | ABL 90 Konfiguration überprüft und Werte mit Labor EDV verglichen. Es stellte sich heraus, dass die geänderten RiliBÄK Grenzen für pCO2 nicht im Laborsystem von M+S eingepflegt waren. Fixe Zielwerte mit Datenblatt, Labordaten und ABL Ausdruck verglichen. Eingabefehler waren im Lev. A und C bei pCO2. Laborleiter Herr Maier wird veranlassen dass die RiliBäk Grenzen im Laborsystem korrekt eingegeben werden. | ABL 90 Konfiguration überprüft und Werte mit Labor EDV verglichen. Es stellte sich heraus, dass die geänderten RiliBÄK Grenzen für pCO2 nicht im Laborsystem von M+S eingepflegt waren. Fixe Zielwerte mit Datenblatt, Labordaten und ABL Ausdruck verglichen. Eingabefehler waren im Lev. A und C bei pCO2. Laborleiter Herr Maier wird veranlassen dass die RiliBäk Grenzen im Laborsystem korrekt eingegeben werden. | SWO-38323 | No | ABL 90 configuration checked and values compared with laboratory EDP. It turned out that the modified RiliBÄK limits for pCO2 were not included in the M+S laboratory system. Fixed target values compared with data sheet, laboratory data and ABL expression. Input errors were in the Lev. A and C for pCO2. Laboratory manager Mr. Maier will arrange for the RiliBäk boundaries to be entered correctly in the laboratory system. | ABL 90 configuration checked and values compared with laboratory EDP. It turned out that the modified RiliBÄK limits for pCO2 were not included in the M+S laboratory system. Fixed target values compared with data sheet, laboratory data and ABL expression. Input errors were in the Lev. A and C for pCO2. Laboratory manager Mr. Maier will arrange for the RiliBäk boundaries to be entered correctly in the laboratory system. | Setup | HIS/LIS | Rilibäk tolerances
ABL90 |  |  | CF45C895F25F32C3 | No |  |  | Assisted Replacement | SC | Multi
ABL90 FLEX PLUS | Cambio gasket y sigue igual. Se instala nuevo SC. Arranque OK, cal y QC Ok | Cambio gasket y sigue igual. Se instala nuevo SC. Arranque OK, | CF45E08FF22EED4D | No | Gasket change and stays the same. New SC is installed. Start OK, lime and QC Ok | Gasket change and stays the same. New SC is installed. Boot OK, | Assisted Replacement | SC | Wetability
ABL835 FLEX | Friction Detector Unit eingesetzt, Einlass überprüft und einkalibriert, i.O. | Vom Kunden aus (Ersatz-)Gerät Einlassmodul ausgebaut und zur Rep.-Überbrückung eingesetzt, Kalibrierungen und QC ohne Beanstandung | SWO-38214 | No | Detector Unit used, inlet checked and calibrated, i.O. | Removed from the customer (replacement) device inlet module and used for rep. bridging, calibrations and QC without complaint | Defect | HW | Inlet module | Friction detector unit
ABL835 FLEX | Målekammer skiftet | Målekammer skiftet | SWO-38240 | No | Measuring chamber Change | Measuring chamber Change | Defect | HW | El/Met Module | Measuring chamber
ABL825 FLEX | Ersatzmodul eingebaut Gerät wieder in Betrieb genommen. Kal und QC okay | Ersatzmodul eingebaut Gerät wieder in Betrieb genommen. Kal und QC okay | SWO-38490 | No | Replacement module installed device put back into operation. Kal and QC okay | Replacement module installed device put back into operation. Kal and QC okay | Defect | HW | AutoCheck Module | Stepper motor
ABL825 FLEX | Analysator heruntergefahren und teilweise demontiert- Erneuerung EL MET Modul PCB - Rückmontage und Funktionsüberprüfung etc. | Analysator heruntergefahren und teilweise demontiert - Erneuerung EL MET Modul PCB - Rückmontage und Funktionsüberprüfung etc. | SWO-38232 | No | Analyzer shut down and partially dismantled- Renewal EL MET Module PCB - Backassembly and function verification etc. | Analyzer shut down and partially dismantled - Renewal EL MET Module PCB - Backassembly and function verification etc. | Defect | HW | El/Met Module | PCB
ABL825 FLEX | Analysator heruntergefahren und teilweise demontiert - Erneuerung EL MET Modul PCB - Rückmontage und Funktionsüberprüfung etc. | Analysator heruntergefahren und teilweise demontiert - Erneuerung EL MET Modul PCB - Rückmontage und Funktionsüberprüfung etc. | SWO-38233 | Yes | Analyzer shut down and partially dismantled - Renewal EL MET Module PCB - Backassembly and function verification etc. | Analyzer shut down and partially dismantled - Renewal EL MET Module PCB - Backassembly and function verification etc. | Re classified RMED | Doube entry
ABL825 FLEX | Mehrmaliges herunterfahren des Analysators mit anschließendem sofortigem Neustart ohne Beanstandung - Funktionsüberprüfung etc. | Mehrmaliges herunterfahren des Analysators mit anschließendem sofortigem Neustart ohne Beanstandung - Funktionsüberprüfung etc. | SWO-38231 | Yes | Shutdown of the analyzer several times with subsequent immediate restart without complaint - function check, etc. | Shutdown of the analyzer several times with subsequent immediate restart without complaint - function check, etc. | Re classified RMED | See next upSWO | El/Met Module | PCB
AQT90 FLEX | GS Gerät zum Reparaturaustausch vorbereitet Reagenzienpack eingelegt und D-Dimer kalibriert GS Gerät 838R0052N009 beim Kunden aufgestellt Kundengerät zu Reparatur in GS Konfiguration eingespielt Zur Zeit kein TnI vorrätig => Kalibrationskassette in Willich bestellt => Lieferung zu morgen 8 Uhr an Kunden LQC CRP Lot 10493 i.O. | GS Gerät zum Reparaturaustausch vorbereitet Reagenzienpack eingelegt und D-Dimer kalibriert GS Gerät 838R0052N009 beim Kunden aufgestellt Kundengerät zu Reparatur in GS Konfiguration eingespielt Zur Zeit kein TnI vorrätig => Kalibrationskassette in Willich bestellt => Lieferung zu morgen 8 Uhr an Kunden LQC CRP Lot 10493 i.O. | SWO-38170 | No | GS device for repair replacement prepared reagent pack inserted and D-Dimer calibrated GS device 838R0052N009 installed at the customer customer customer device for repair in GS configuration Currently no TnI in stock => Calibration cassette ordered in Willich => Delivery tomorrow 8 o'clock to customers LQC CRP Lot 10493 i.O. | GS device for repair replacement prepared reagent pack inserted and D-Dimer calibrated GS device 838R0052N009 installed at the customer customer customer device for repair in GS configuration Currently no TnI in stock => Calibration cassette ordered in Willich => Delivery tomorrow 8 o'clock to customers LQC CRP Lot 10493 i.O.
ABL80 |  |  | CF47D7E8DE57C586 | No
ABL90 FLEX | Schlauchsystem gespült, Funktionskontrolle ok | Der Kunde ist der Meinung, dass das Gerät überprüft werden sollte. Der Kunde bittet um einen Service und Rückruf. | SWO-38837 | No | Hose system flushed, function control ok | The customer believes that the device should be checked. The customer asks for a service and callback. | Cleaned | Fluidic pathway
ABL90 |  |  | CF473A06684EFE62 | No |  |  | Defect | HW | Cassette retainer
ABL90 FLEX | SC90 & SP90 entnommen & Schwingenden Hemolyzer mit Proteinentferner gereinigt, OxiVentil kontrolliert, erneut Serviceprogramme insb. Clotdetecting durchgeführt, kompletter OxiTest ok, korrekte Kal abgewartet, thb-Kal & QK ok | SC90 & SP90 entnommen & Schwingenden Hemolyzer mit Proteinentferner gereinigt, OxiVentil kontrolliert, erneut Serviceprogramme insb. Clotdetecting durchgeführt, kompletter OxiTest ok, korrekte Kal abgewartet, thb-Kal & QK ok | SWO-38551 | No | SC90 & SP90 taken & vibrating hemolyzer cleaned with protein remover, OxiVentil controlled, again service programs in addition. Clotdetecting performed, complete OxiTest ok, correct kal wait, thb-Kal & QK ok | SC90 & SP90 taken & vibrating hemolyzer cleaned with protein remover, OxiVentil controlled, again service programs in addition. Clotdetecting performed, complete OxiTest ok, correct kal wait, thb-Kal & QK ok | Cleaned | Fluidic pathway | Protein remover
ABL820 FLEX | Proteïne removal en reiniging uitgevoerd. Meetkamers schoongemaakt, lekkage komt vanaf de referentie electrode. Pomprotoren schoongemaakt en voorzien van nieuwe tubings uit voorraad klant. Inlet schoongemaakt. Instellingen gecontroleerd, indien nodig bijgesteld. Instrument weer opgestart en gecalibreerd. Controle S7745 gemeten, OK. | Proteïne removal en reiniging uitgevoerd. Meetkamers schoongemaakt, lekkage komt vanaf de referentie electrode. Pomprotoren schoongemaakt en voorzien van nieuwe tubings uit voorraad klant. Inlet schoongemaakt. Instellingen gecontroleerd, indien nodig bijgesteld. Instrument weer opgestart en gecalibreerd. Controle S7745 gemeten, OK. | SWO-38452 | No | Protein removal and cleaning performed. Measuring chambers cleaned, leakage comes from the reference electrode. Pump rotors cleaned and equipped with new tubings from stock customer. Inlet cleaned. Institutions, if necessary, adjusted. Instrument will be restarted and calibrated. Check S7745 measured, OK. | Protein removal and cleaning performed. Measuring chambers cleaned, leakage comes from the reference electrode. Pump rotors cleaned and equipped with new tubings from stock customer. Inlet cleaned. Institutions, if necessary, adjusted. Instrument will be restarted and calibrated. Check S7745 measured, OK. | Assisted Replacement | Electrode | Electrode | pCO2
ABL825 FLEX | Indskannet gas 1 Kørt kalibrering - ok Ingen fejl fundet | Indskannet gas 1 Kørt kalibrering - ok Ingen fejl fundet | SWO-38429 | No | Scanned gas 1 driven calibration-OK no errors found | Scanned gas 1 driven calibration-OK no errors found | No fault found
ABL835 FLEX | pCO2 Elektrode gereinigt und neu bezogen, Gerät kalibriert, QCs durchgef. | Kunde bittet um einen Einsatz des zuständigen Technikers, da das Gerät erst einige Tage im Betrieb ist. | SWO-38598 | No | pCO2 electrode cleaned and re-covered, device calibrated, QCs through. | The customer asks for the technician to be deployed, as the device is only in operation for a few days. | Assisted Replacement | Electrode | Electrode | pCO2
ABL827 FLEX | Met-II-Modul zurückgetauscht, nach Problem mit Ableitströmen nochmals die Membranventile der Messmodule kontrolliert, sporadischer Driftfehler bei pH Elektrolyten -> probeweise Referenzelektrode (9702) ersetzt, Inbetriebnahme, korrekte Kal abgewartet, QK ok | Met-II-Modul zurückgetauscht, nach Problem mit Ableitströmen nochmals die Membranventile der Messmodule kontrolliert, sporadischer Driftfehler bei pH Elektrolyten -> probeweise Referenzelektrode (9702) ersetzt, Inbetriebnahme, korrekte Kal abgewartet, QK ok | SWO-38553 | No | Met-II module replaced, after problem with dissipation currents again the diaphragm valves of the measuring modules checked, sporadic drift error in pH electrolytes -> test reference electrode (9702) replaced, commissioning, correct cal awaiting, QK ok | Met-II module replaced, after problem with dissipation currents again the diaphragm valves of the measuring modules checked, sporadic drift error in pH electrolytes -> test reference electrode (9702) replaced, commissioning, correct cal awaiting, QK ok | Defect | HW | Met II Module | Valve
AQT90 FLEX | REMOVED CARTRIDGE AND RUN ON SAFE POSITION NOW WORKING SATISFACTORILY | REMOVED CARTRIDGE AND RUN ON SAFE POSITION NOW WORKING SATISFACTORILY | SWO-39187 | No | REMOVED CARTRIDGE AND RUN ON SAFE POSITION NOW WORKING SATISFACTORILY | REMOVED CARTRIDGE AND RUN ON SAFE POSITION NOW WORKING SATISFACTORILY
ABL80 FLEX | ABL kompl. überprüft. (Fliesswege und Programm Fehlerdiagnose), alles i.o. Externe QC4+ Level 1-3 als Messungen durchgeführt. Alle im Bereich gemäss Beipackzettel. |  | SWO-38660 | No | ABL compl. Checked. (Flow paths and fault diagnosis program), all i.o. External QC4+ Level 1-3 performed as measurements. All in the area according to the package leaflet.
ABL80 | Multiple SC have exhibited the issue, testing and verify the system and electronic channels. | Calcium channel drifted, replace the connector and signal data PCB and align to manufactuerer specifications. | CF489FF4DE50FB34 | No | Multiple SC have exhibited the issue, testing and verify the system and electronic channels. | Calcium channel drifted, replace the connector and signal data PCB and align to manufactuerer specifications.
ABL90 |  |  | CF488EF7DE0B9987 | No |  |  | Cleaned | Fluidic pathway
ABL90 | The hemolyer gives a 1069 error message. Verrified the error. Verified the hemolyzer cable connections per the service manual. After the startup the unit kept thermostating and the hemolyzer temp was out of cal. | Ordered a new hemolyzer and will return tomorrow 7-23-15 to finish the repair. | CF485FFEDE167E81 | No | The hemolyer gives a 1069 error message. Verrified the error. Verified the hemolyzer cable connections per the service manual. After the startup the unit kept thermostating and the hemolyzer temp was out of cal. | Ordered a new hemolyzer and will return tomorrow 7-23-15 to finish the repair. | Unknown | Revisit required | Hemolyzer
ABL90 | Reviewed logs and confirmed problems. | Replaced the above parts and continued to have problems. Put a jumper tubing in the SC bay and noticed lots of bubbles. Swapped soltion pack with ABLI and bubbles went away. However, ABLI came up just fine as well. | CF48B9D1DE60A2EF | No | Reviewed logs and confirmed problems. | Replaced the above parts and continued to have problems. Put a jumper tubing in the SC bay and noticed lots of bubbles. Swapped soltion pack with ABLI and bubbles went away. However, ABLI came up just fine as well. | Assisted Replacement | SC | Wetability
ABL90 FLEX | System mehrfach gespült, LS kalibriert, Pumpen kalibriert, Lösungspackzyklen niedrig - LP ersetzt, Einlaßdichtung erneuert, Startphase Kal und QC abgewartet | Der Kunde ist unzufrieden, weil alle Blutmessungen ohne erkennbaren Grund vom Gerät nicht gemessen worden sind. Der Kunde bittet um eine zeitnahe Überprüfung. Bitte den Kunden über weiteres Vorgehen informieren. | SWO-38722 | No | System repeatedly rinsed, LS calibrated, pumps calibrated, solution packing cycles low - LP replaced, inlet seal renewed, start phase Kal and QC awaited | The customer is dissatisfied because all blood measurements have not been measured by the device for no apparent reason. The customer asks for a timely check. Please inform the customer about further action. | Unknown | Aborted samples
ABL825 FLEX | Instrument has not been in use and has not had regular user maintenance performed. Reference and lactat membranes are in need of replacement. | Replaced expired reference membrane. Lactate membrane is in need of replacement but customer currently has none on site. Temporarily disabled lactate parameter until lactate membrane arrives for replacement. Completed tHb calibration and installed instrument's 4 levels of quality control. Instrument is within acceptable operating specifications. | CF486434DE0EF527 | No | Instrument has not been in use and has not had regular user maintenance performed. Reference and lactat membranes are in need of replacement. | Replaced expired reference membrane. Lactate membrane is in need of replacement but customer currently has none on site. Temporarily disabled lactate parameter until lactate membrane arrives for replacement. Completed tHb calibration and installed instrument's 4 levels of quality control. Instrument is within acceptable operating specifications. | User Error | Missed maintenance
ABL825 FLEX | Instrument has not been in use and has not had regular user maintenance performed. Several membranes are in need of replacement. | Replaced Ca++, Ref, PCO2 and O2 membranes. Lactate membrane is in need of replacement but customer currently has none on site.Temporarily disabled Lactate parameter until the lactate membrane arrives for replacement. Replaced expired cleaning solution and performed tHb calibration. Instrument is within acceptable operating specifications. | CF486039DE3DF524 | No | Instrument has not been in use and has not had regular user maintenance performed. Several membranes are in need of replacement. | Replaced Ca++, Ref, PCO2 and O2 membranes. Lactate membrane is in need of replacement but customer currently has none on site. Temporarily disabled Lactate parameter until the lactate membrane arrives for replacement. Replaced expired cleaning solution and performed tHb calibration. Instrument is within acceptable operating specifications. | User Error | Missed maintenance
ABL825 FLEX | Einlassnadel leicht verbogen => Nadel gerichtet Austausch Cl Elektrode Überprüfung Lac Elektrode => Elektrode neu in Membran eingesetzt Mehrfach kalibriert bis i.O. Fehlerhafte QC's Level 2,3 und 4 nachgeholt | Einlassnadel leicht verbogen => Nadel gerichtet Austausch Cl Elektrode Überprüfung Lac Elektrode => Elektrode neu in Membran eingesetzt Mehrfach kalibriert bis i.O. Fehlerhafte QC's Level 2,3 und 4 nachgeholt | SWO-38606 | No | Inlet needle slightly bent => Needle directed replacement Cl electrode checking lac electrode => Electrode newly inserted in membrane Multiple calibrated up to i.O. Faulty QC's Levels 2,3 and 4 | Inlet needle slightly bent => Needle directed replacement Cl electrode checking lac electrode => Electrode newly inserted in membrane Multiple calibrated up to i.O. Faulty QC's Levels 2,3 and 4 | Assisted Replacement | Electrode | Electrode | Cl
ABL825 FLEX | - | - | SWO-38539 | No | - | - | Assisted Replacement | Electrode | Electrode | pCO2
ABL80 CO-OX |  |  | CF497A32541D317E | No
ABL90 | REBOOTED ANALYZER AND PROBLEM CAME UP AGAIN. CHANGED OUT CF CARD AND FOUND THAT THE SETUP AND DATA ARE CORRUPT. SO WILL E-MAIL DATA LOGS FOR LAST 3 MONTHS TO ALICE SEID. ALSO NOTICED THAT THE SOLUTION PACK EXPIRES TOMORROW. | REPLACED CF CARD, RELOADED WINDOWS, RELOADED RADIOMETER SOFTWARE, RELOADED WIRELESS ADAPTER, RELOADED SETUP FROM ANOTHER ANALYZER. LOADED ALL QC. REPLACED SOLUTION PACK FROM CUSTOMER SUPPLY. | CF4A3118544CEA3E | No | REBOOTED ANALYZER AND PROBLEM CAME UP AGAIN. CHANGED OUT CF CARD AND FOUND THAT THE SETUP AND DATA ARE CORRUPT. SO WILL E-MAIL DATA LOGS FOR LAST 3 MONTHS TO ALICE SEID. ALSO NOTICED THAT THE SOLUTION PACK EXPIRES TOMORROW. | REPLACED CF CARD, RELOADED WINDOWS, RELOADED RADIOMETER SOFTWARE, RELOADED WIRELESS ADAPTER, RELOADED SETUP FROM ANOTHER ANALYZER. LOADED ALL QC. REPLACED SOLUTION PACK FROM CUSTOMER SUPPLY. | Defect | HW | CF card
ABL90 | LOOKED THROUGH DATA LOGS AND FOUND INHOMOGENOUS ERRORS, DID BACK FLUSH 15 TIMES, RE-INSTALLED SOLUTION PACK. NO HELP. | REPLACE SOLUTION PACK FROM CUSTOMERS SUPPLY. ALSO DID CREDIT FORM FOR SOLUTION PACK FOR CUSTOMER. | CF4A367F5463EA42 | No | LOOKED THROUGH DATA LOGS AND FOUND INHOMOGENOUS ERRORS, DID BACK FLUSH 15 TIMES, RE-INSTALLED SOLUTION PACK. NO HELP. | REPLACE SOLUTION PACK FROM CUSTOMERS SUPPLY. ALSO DID CREDIT FORM FOR SOLUTION PACK FOR CUSTOMER. | Assisted Replacement | SP | FT
ABL90 | The hemolyzer kept having temp errors (1069). The solution pack was empty. | Replaced the hemolyzer. Replaced solution pack from customers stock. | CF49C7E3541A022C | No | The hemolyzer kept having temp errors (1069). The solution pack was empty. | Replaced the hemolyzer. Replaced solution pack from customers stock. | Defect | HW | Hemolyzer
ABL90 | Analyzer started giving 1343 error 2 days after SP replacement. Ran service checks over the phone first with the customer issue persists.Arrived on site and performed manual back flush multiple times but issue not resolved.Customer Replaced SP as well but no luck. I tried reinserting the same SP but star tup failed. ran service checks repeatedly. SP check failing repeatedly. Noticed that both SP replaced were same LOT SB-02. Borrowed another SP(Lot SB-09) from different department. Installed SP and the analyzer showed similar errors. Performed Manual back flush, replaced connector gasket, cleaned inlet gasket. Flushed the probe.Repeated all service checks. SP check failing and the parameters keep changing, sometimes the QC fail and sometimes the cal solution fail. Visual Flow check Pass. | Replaced SP with new lot. Replaced connector gasket and inlet gasket with holder(customer stock).Ran all service checks. Performed multiple back flushes. Performed extended flush. Inspected sensor bay. Opened analyzer and inspected all tubing. Replaced SC. Installed Inlet Clip. | CF498CA55408A094 | No | Analyzer started giving 1343 error 2 days after SP replacement. Ran service checks over the phone first with the customer issue persists. Arrived on site and performed manual back flush multiple times but issue not resolved. Customer Replaced SP as well but no luck. I tried reinserting the same SP but star tup failed. ran service checks repeatedly. SP check failing repeatedly. Noticed that both SP replaced were same LOT SB-02. Borrowed another SP(Lot SB-09) from different department. Installed SP and the analyzer showed similar errors. Performed Manual back flush, replaced connector gasket, cleaned inlet gasket. Flushed the probe. Repeated all service checks. SP check failing and the parameters keep changing, sometimes the QC fail and sometimes the cal solution fail. Visual Flow check Pass. | Replaced SP with new lot. Replaced connector gasket and inlet gasket with holder(customer stock). Ran all service checks. Performed multiple back flushes. Performed extended flush. Inspected sensor bay. Opened analyzer and inspected all tubing. Replaced SC. Installed Inlet Clip. | User Error | Pre analytical | Inlet Clip
ABL90 FLEX | Sensorkassette erneuert, Kalibrationen und Qualitätskontrollen ohne Beanstandungen durchgeführt. Anwender angeleitet Ausdrucke zu erstellen und einzuschicken. | Sensorkassette erneuert, Kalibrationen und Qualitätskontrollen ohne Beanstandungen durchgeführt. Anwender angeleitet Ausdrucke zu erstellen und einzuschicken. | SWO-38845 | No | Sensor cartridge renewed, calibrations and quality controls carried out without complaints. User-instructed to create and submit printouts. | Sensor cartridge renewed, calibrations and quality controls carried out without complaints. User-instructed to create and submit printouts. | Assisted Replacement | SC | pH
ABL825 FLEX | Flex-Q-Modul ausgetauscht. | Flex-Q-Modul ausgetauscht. | SWO-38785 | No | Flex-Q module. | Flex-Q module. | Defect | HW | FlexQ Module | Postion
ABL825 FLEX |  |  | CF49C3FE540C6A7D | No |  |  | Assisted Replacement | Electrode | Electrode | Reference
ABL835 FLEX | Luft ventil og pumpeslange i EL MET modul skiftet | Luft ventil og pumpeslange i EL MET modul skiftet | SWO-38844 | No | Air valve and pump hose in EL MET module switch | Air valve and pump hose in EL MET module switch | Defect | HW | El/Met Module | Valve, air
ABL825 FLEX | Neu Kalibration des Moduls | Neu Kalibration des Moduls | SWO-38911 | No | New calibration of the module | New calibration of the module | Adjustment | HW | FlexQ Module | Postion
ABL825 FLEX | pH Dichtring ersetzt. Manuelle Dichtigkeitskontrolle okay |  | SWO-38898 | No | pH sealing ring. Manual tightness control okay |  | Defect | Electrode | Electrode | Sealing Ring
ABL825 FLEX | Pulled autocheck apart and found it missing the cotter pin that holds the motor to the gear assembly for the housing unit up position. | Reconnected the gear assembly to motor unit and affixed it with adhesive. Performed necessary calibrations to bring analyzer back in to service. | CF49EB0354137D7E | No | Pulled autocheck apart and found it missing the cotter pin that holds the motor to the gear assembly for the housing unit up position. | Reconnected the gear assembly to motor unit and affixed it with adhesive. Performed necessary calibrations to bring analyzer back in to service. | Defect | HW | AutoCheck Module | Breaker arm
ABL827 FLEX | Lösungen ersetzt und Crea membraniert, mehrere Kal, QC |  | SWO-38970 | No | Solutions replaced and Crea membraneed, several Kal, QC |  | Assisted Replacement | Electrode | Membrane | Crea
ABL90 FLEX | Gerät Netz aus ein, LP mechanisch entfernt, war nicht korrekt aktiviert, Gerät gespült, Startphase abgewartet, Fehler tritt erneut auf, alten Pack wieder eingesetzt, Startphase, Kal & QC i.O. | Gerät Netz aus ein, LP mechanisch entfernt, war nicht korrekt aktiviert, Gerät gespült, Startphase abgewartet, Fehler tritt erneut auf, alten Pack wieder eingesetzt, Startphase, Kal & QC i.O. | SWO-39228 | No | Device mains from on, LP mechanically removed, was not properly activated, device flushed, start phase waited, error reoccurs, old pack reinserted, start phase, Kal & QC i.O. | Device mains from on, LP mechanically removed, was not properly activated, device flushed, start phase waited, error reoccurs, old pack reinserted, start phase, Kal & QC i.O. | User Error | SP activation
ABL90 |  |  | CF4ABD40CA18D286 | No |  |  | Defect | HW | Oxi module | Pinch Valve tube
ABL90 | Issue reoccurred overnight. OIR with errors leak detected and unable to pump solutions. Ran all service checks. Clot Detect Status failing. Solution pack check failing all parameters. Error 1084 1085,1086 and 1087. Pump calibrations failing as well. | Performed multiple back flushes.Flushed the SC.Checked and cleaned all gaskets and probe.Flushed the probe. Reinserted SP and SC. Ran 10 patient samples for testing purposes. | CF4AEE50CA1CAB11 | No | Issue reoccurred overnight. OIR with errors leak detected and unable to pump solutions. Ran all service checks. Clot Detect Status failing. Solution pack check failing all parameters. Error 1084 1085,1086 and 1087. Pump calibrations failing as well. | Performed multiple back flushes. Flushed the SC. Checked and cleaned all gaskets and probe. Flushed the probe. Reinserted SP and SC. Ran 10 patient samples for testing purposes. | Cleaned | Fluidic pathway
ABL90 | Found Temp error to be related to the SC, which was due for replacement. | Customer to replace SC. | CF4EF0C92C27A40F | No | Found Temp error to be related to the SC, which was due for replacement. | Customer to replace SC. | Assisted Replacement | SC | Temp
ABL90 | Unable to pump solutions Error 1169 | Ran service checks and flushed analyzer. | CF51B6AC1811B337 | No | Unable to pump solutions Error 1169 | Ran service checks and flushed analyzer. | Cleaned | Fluidic pathway
ABL90 |  | Replaced Inlet pos PCB | CF51B784180DB339 | No |  | Replaced Inlet pos PCB | Defect | HW | Inlet PCB
ABL835 FLEX | Viewed QC logs and inspected inlet section. | Replaced the probe and inlet gasket. | CF4AC777CA286B5F | No | Viewed QC logs and inspected inlet section. | Replaced the probe and inlet gasket. | Assisted Replacement | HW | Inlet module | Inlet Gasket
ABL825 FLEX | pCO2 Elektrode ausgetauscht und neu membraniert (neue Lot von RSCH) pH Elektrode ausgetauscht (weisse Fleck im Glas) Andere Lot QC getestet(ohne Erfolg) Problematik immer noch vorhanden. | pCO2 Elektrode ausgetauscht und neu membraniert (neue Lot von RSCH) pH Elektrode ausgetauscht (weisse Fleck im Glas) Andere Lot QC getestet(ohne Erfolg) Problematik immer noch vorhanden. | SWO-39063 | No | pCO2 electrode replaced and re-membraneed (new solder from RSCH) pH electrode replaced (white stain in the glass) Other lot QC tested (without success) problems still present. | pCO2 electrode replaced and re-membraneed (new solder from RSCH) pH electrode replaced (white stain in the glass) Other lot QC tested (without success) problems still present. | Unknown | Electrode | Membrane | pCO2
ABL820 FLEX | Continuing to investigate possible causes related to rinse and or air gas issue at pH BG module. | Replaced rinse and air gas valves and dual valve housing, black tubing from housing to gas ports. Examined pCO2 electrode and chamber, no issues observed. Ran multiple calibration cycles. QC is recovering within manufacturers range, low end of peer 2 group occasionally low of peer 2. Suspect pH BG PCB and will need to schedule a time to replace the PCB so customer can run Lin verification test. | CF4B08FCCA4C38BA | No | Continuing to investigate possible causes related to rinse and or air gas issue at pH BG module. | Replaced rinse and air gas valves and dual valve housing, black tubing from housing to gas ports. Examined pCO2 electrode and chamber, no issues observed. Ran multiple calibration cycles. QC is recovering within manufacturers range, low end of peer 2 group occasionally low of peer 2. Suspect pH BG PCB and will need to schedule a time to replace the PCB so customer can run Lin verification test. | Assisted Replacement | Electrode | Electrode | pCO2
ABL825 FLEX | air Kanal gespült LS Justage erfolgreich Neue QC Lot eingelesen | air Kanal gespült LS Justage erfolgreich Neue QC Lot eingelesen | SWO-39052 | No | air channel flushed LS adjustment successfully Reread New QC Lot | air channel flushed LS adjustment successfully Reread New QC Lot | Cleaned | Fluidic pathway
ABL825 FLEX | Fällige Elektroden auf Wunsch Membraniert Ref..Lac. QC Lot nochmal eingelesen. Verdacht auf Modulfehler ?.QC Modul wird nach Erhalt ersetzt, Solange kann QC von Hand gemacht werden. | Konfig. Menu überprüft, Chargen sind richtig eingescannt, AC steht auch auf Auto. THb Kal ist durchgeführt. Kalibrationen sind ok. In den Aktivitäten sind keine Fehlermeldungen bzgl. QC Kontrollen, im Datei Archiv gibt es keine Abbrüche. Trotzdem werden die Kontrollen zu den eingestellten Zeiten nicht durchgeführt. Fr. Krämer hat heute Morgen den ABL neu gestartet und beobachtet das Problem weiter. | SWO-39051 | No | Due electrodes on request Membraneed Ref.. Lac. QC Lot re-read. Suspicion of module error ?. QC module is replaced upon receipt, Solange can be made QC by hand. | Config. Menu checked, batches are scanned correctly, AC is also on car. THb Kal is performed. Calibrations are ok. In the activities there are no error messages regarding QC controls, in the file archive there are no cancellations. Nevertheless, the checks are not carried out at the set times. Ms. Krämer has restarted the ABL this morning and continues to monitor the problem. | Unknown | HW | AutoCheck Module
ABL800 BASIC | Fehlersuche, Anwender hat bereits die Referenzelektrode neu membraniert, die Empf. von pH ist schon seit 2-3 Wochen grenzenwertig niedrig und seit 22.07. unter 92% gefallen, die Referenzelektrode ist ok, pH Elektrode LOT RV-02 defekt->erneuert gegen LOT RM-07->Empf. angestiegen auf 97%, mehrere Kal. ok System-und Funktionskontrolle i.O. | Fehlersuche, Anwender hat bereits die Referenzelektrode neu membraniert, die Empf. von pH ist schon seit 2-3 Wochen grenzenwertig niedrig und seit 22.07. unter 92% gefallen, die Referenzelektrode ist ok, pH Elektrode LOT RV-02 defekt->erneuert gegen LOT RM-07->Empf. angestiegen auf 97%, mehrere Kal. ok System-und Funktionskontrolle i.O. | SWO-39934 | No | Troubleshooting, user has already re-membraneed the reference electrode, the Empf. of pH has been marginally low for 2-3 weeks and has fallen below 92% since 22.07., the reference electrode is ok, pH electrode LOT RV-02 defective->renewed against LOT RM-07->Empf. increased to 97%, several ok System and function control i.O. | Troubleshooting, user has already re-membraneed the reference electrode, the Empf. of pH has been marginally low for 2-3 weeks and has fallen below 92% since 22.07., the reference electrode is ok, pH electrode LOT RV-02 defective->renewed against LOT RM-07->Empf. increased to 97%, several ok System and function control i.O. | Defect | Electrode | Electrode | pH
AQT90 FLEX | Barcodescanner elektrisch und mechanisch justiert, Kal.cassetten eingelegt und gestartet, | Barcodescanner elektrisch und mechanisch justiert, Kal.cassetten eingelegt und gestartet, | SWO-39129 | No | Barcode scanner electrically and mechanically adjusted, Kal.cassetten inserted and started, | Barcode scanner electrically and mechanically adjusted, Kal.cassetten inserted and started,
ABL90 |  |  | CF4EDFB32C07EC79 | No |  |  | Unknown | Removed by RSUB | SC Bay
ABL90 FLEX | Funktion überprüft, abgeknickte Einlassnadel ersetzt, Fehlersuchverfahren durchgeführt, Funktion OK | Funktion überprüft, abgeknickte Einlassnadel ersetzt, Fehlersuchverfahren durchgeführt, Funktion OK | SWO-39899 | No | Function checked, bent inlet needle replaced, error search procedure performed, function OK | Function checked, bent inlet needle replaced, error search procedure performed, function OK | Replaced | Probe | Bend
ABL837 FLEX | Switched off the Crea key in options and restarted the analyzer. performed update type and found analyzer working. Switched ON the Crea key and performed update type, analyzer went to power on mode. Erased the option key and application software in the analyzer and re-installed the software. RE-installed the option keys and found analyzer working, but, conditioning error in Met II Module. Checked all the valves and tubing s and found Ok. Cleaned the Dual valve housing, Valves and valve membranes in the Met II module and found analyzer working satisfactorily without any error. | Switched off the Crea key in options and restarted the analyzer. performed update type and found analyzer working. Switched ON the Crea key and performed update type, analyzer went to power on mode. Erased the option key and application software in the analyzer and re-installed the software. RE-installed the option keys and found analyzer working, but, conditioning error in Met II Module. Checked all the valves and tubing s and found Ok. Cleaned the Dual valve housing, Valves and valve membranes in the Met II module and found analyzer working satisfactorily without any error. | SWO-39839 | No | Switched off the Crea key in options and restarted the analyzer. performed update type and found analyzer working. Switched ON the Crea key and performed update type, analyzer went to power on mode. Erased the option key and application software in the analyzer and re-installed the software. RE-installed the option keys and found analyzer working, but, conditioning error in Met II Module. Checked all the valves and tubing s and found Ok. Cleaned the Dual valve housing, Valves and valve membranes in the Met II module and found analyzer working satisfactorily without any error. | Switched off the Crea key in options and restarted the analyzer. performed update type and found analyzer working. Switched ON the Crea key and performed update type, analyzer went to power on mode. Erased the option key and application software in the analyzer and re-installed the software. RE-installed the option keys and found analyzer working, but, conditioning error in Met II Module. Checked all the valves and tubing s and found Ok. Cleaned the Dual valve housing, Valves and valve membranes in the Met II module and found analyzer working satisfactorily without any error. | Defect | HW | Met II Module | Valve
ABL90 | Checked activity logs. Ran service checks, Clot detect remove failing.Ran clot detect remove multiple times. | SC and SP reinitialized. Installed SW 3.1 MR5. Ran 6 samples for testing. | CF4DA1A7B62BAB1E | No | Checked activity logs. Ran service checks, Clot detect remove failing. Ran clot detect remove multiple times. | SC and SP reinitialized. Installed SW 3.1 MR5. Ran 6 samples for testing. | Assisted Replacement | SC | Wetability
ABL90 |  |  | CF4EDA672C07CCF4 | No |  |  | Adjustment | SC Door | Cleaned pivot
ABL90 FLEX | Verstopfung zwischen Meßkammer und Einlaß durch spülen beseitigt, Dichtung erneuert Einlaufverfahren | Verstopfung zwischen Meßkammer und Einlaß durch spülen beseitigt, Dichtung erneuert Einlaufverfahren | SWO-39936 | No | Blockage between measuring chamber and inlet eliminated by flushing, seal renews inlet method | Blockage between measuring chamber and inlet eliminated by flushing, seal renews inlet method | Cleaned | Fluidic pathway
ABL90 FLEX | Nach Test wiederholen mehrfach die automatische Kal. gescheitert. Analysator geprüft, erneuert manuell gespült und Einlassdichtung gereinigt. LP Check i.O. tHb Kal. Proben Aspiration gescheitert, defekte Anschlussdichtung festgestellt, Anschlussdichtung erneuert. Vakuum Fehler aufgetreten, Einlassdichtung erneuert. | Nach Test wiederholen mehrfach die automatische Kal. gescheitert. Analysator geprüft, erneuert manuell gespült und Einlassdichtung gereinigt. LP Check i.O. tHb Kal. Proben Aspiration gescheitert, defekte Anschlussdichtung festgestellt, Anschlussdichtung erneuert. Vakuum Fehler aufgetreten, Einlassdichtung erneuert. | SWO-40081 | No | After testing, the automatic cal repeat several times. Failed. Analyzer tested, renewed manually rinsed and inlet seal cleaned. LP Check i.O. tHb Kal. Samples aspiration failed, defective connection seal detected, connection seal renewed. Vacuum error occurred, inlet seal renewed. | After testing, the automatic cal repeat several times. Failed. Analyzer tested, renewed manually rinsed and inlet seal cleaned. LP Check i.O. tHb Kal. Samples aspiration failed, defective connection seal detected, connection seal renewed. Vacuum error occurred, inlet seal renewed. | Defect | HW | Connector gasket
ABL90 FLEX | ABL 90 Geräte Kalibration durchgeführt, THb Kalibierung durchgeführt. die gescheiteren Qc nochmals wiederholt. ABL ist wieder einsatzbereit . | ABL 90 Geräte Kalibration durchgeführt, THb Kalibierung durchgeführt. die gescheiteren Qc nochmals wiederholt. ABL ist wieder einsatzbereit . | SWO-40061 | No | ABL 90 devices calibration performed, THb calibration performed. the clever qc again. ABL is ready for use again. | ABL 90 devices calibration performed, THb calibration performed. the clever qc again. ABL is ready for use again. | Adjustment | tHb cal
ABL90 FLEX | Sensor cassette Bay交換 センサーピン調整他 | Sensor cassette Bay交換 センサーピン調整他 | SWO-45090 | No | Sensor cassette Bay replacement Sensor pin adjustment, etc. | Sensor cassette Bay replacement Sensor pin adjustment, etc.
ABL90 | SUSTITUCION DE CONJUNTO DE MONITOR POR NO RESPONDER PANTALLA TACTIL. AJUSTES Y PRUEBAS DE FUNCIONAMIENTO, SIENDO ESTAS CORRECTAS. CALIBRADO Y VERIFICADO CON Q.C. |  | CF4EE8A52C47D9CE | No | REPLACEMENT OF MONITOR SET BY NOT RESPONDING TO TACTILE DISPLAY. OPERATION SETTINGS AND TESTS, BEING THESE CORRECT. CALIBRATED AND VERIFIED WITH Q.C. |  | Defect | HW | Touch screen
ABL835 FLEX | Se reitira y limpia el modulo de entrada se sustituye el nippel roto y se comprueba fucnionamiento, se realiza calibración de bombas y de sensores de liquidos, funcionando correctamente. |  | CF4E75872C4DDAD0 | No | The input module is re-iated and cleaned, the broken nippel is replaced and fucnionation is checked, pump and liquid sensor calibration is performed, working properly. |  | Defect | HW | Inlet module
ABL815 FLEX | pH BG mérőkamra hypo-s tisztítása, pO2 membrán csere, kalibrációk, QC mérés |  | SWO-39912 | No | pH BG Measuring Chamber hypo-S cleaning, pO2 membrane replacement, calibrations, QC measurement |  | Cleaned | Fluidic pathway | pH/BG Module
ABL825 FLEX | Unit in hold on arrival. Several errors in activity log including 1020, 1014 and 1015 volume measurment errors. 541 pump error, 343 LS error, 978 pH error etc. | Cleaned measuring modules and chamber. Checked Valves. Replaced flat valve membrane in El Met module (Sticky). Remembraned chloride electrode to stop chamber from filling. Remembraned Ca electrode. Flushed unit with hot water to treat for possible contamination from separator gel from a clinitube sample. | CF4EE2E32C0A3A05 | No | Unit in hold on arrival. Several errors in activity log including 1020, 1014 and 1015 volume measurment errors. 541 pump error, 343 LS error, 978 pH error etc. | Cleaned measuring modules and chamber. Checked Valves. Replaced flat valve membrane in El Met module (Sticky). Remembraned chloride electrode to stop chamber from filling. Remembraned Ca electrode. Flushed unit with hot water to treat for possible contamination from separator gel from a clinitube sample. | Cleaned | Fluidic pathway | - | Gel Clot
ABL825 FLEX | Ersatz Oxi Modul eingebaut. pH Elektrode undicht..Dichtring ersetzt. Alle Dummydichtringe ersetzt | Ersatz Oxi Modul eingebaut. pH Elektrode undicht..Dichtring ersetzt. Alle Dummydichtringe ersetzt | SWO-40149 | No | Replacement Oxi module built-in. pH electrode leaking.. Sealing ring replaced. Replaced all dummy sealing rings | Replacement Oxi module built-in. pH electrode leaking.. Sealing ring replaced. Replaced all dummy sealing rings | Defect | HW | Oxi module | Neon number high
ABL825 FLEX | cleaned rotor and replaced pump tube, no help | Replaced rotorA | CF5017F0A235CE9B | No | cleaned rotor and replaced pump tube, no help | Replaced rotorA | Defect | HW | pH/BG Module | Pump rotor
ABL825 FLEX | Will need to change the Hemolyzer and retest | Replaced Hemolyzer and adjusted lamp and ran THB service cal. | CF4EFA782C30CE93 | No | Will need to change the Hemolyzer and retest | Replaced Hemolyzer and adjusted lamp and ran THB service cal. | Defect | HW | Oxi module | Hemolyzer
ABL815 FLEX | Überprüfung. Nach Erhöhung des Luftdrucks um 19 mmHg alle Werte stabil innerhalb Kontrollbereichs. | Überprüfung. Nach Erhöhung des Luftdrucks um 19 mmHg alle Werte stabil innerhalb Kontrollbereichs. | SWO-40228 | No | Review. After increasing the air pressure by 19 mmHg all values stable within the control range. | Review. After increasing the air pressure by 19 mmHg all values stable within the control range. | Setup | SW | Barometric pressure | Barometer
ABL827 FLEX | Gerät und Elektroden kontrolliert, Lösungen erneuert, tHb Kal und QC o.k. | Gerät und Elektroden kontrolliert, Lösungen erneuert, tHb Kal und QC o.k. | SWO-40069 | No | Device and electrodes controlled, solutions renewed, tHb Kal and QC o.k. | Device and electrodes controlled, solutions renewed, tHb Kal and QC o.k.
ABL835 FLEX | Unit in standby with empty waste container in place on arrival. | Replaced waste detector. | CF4EBCD02C383A01 | No | Unit in standby with empty waste container in place on arrival. | Replaced waste detector. | Defect | HW | Basic Unit | Waste detector
ABL810 FLEX | Qualitätskontrollen nach Korrektur Luftdruck um 2 mmHg näher an Zielwert, aber noch grenzwertig zu RiLiBÄK-Bereich. Elektrode mit anderer Charge neu membraniert. | Qualitätskontrollen nach Korrektur Luftdruck um 2 mmHg näher an Zielwert, aber noch grenzwertig zu RiLiBÄK-Bereich. Elektrode mit anderer Charge neu membraniert. | SWO-40227 | No | Quality controls after correction air pressure by 2 mmHg closer to target value, but still borderline to RiLiBÄK range. electrode with other batch. | Quality controls after correction air pressure by 2 mmHg closer to target value, but still borderline to RiLiBÄK range. electrode with other batch. | Assisted Replacement | Electrode | Membrane | pCO2
ABL80 CO-OX | Upon arrival the analyzer needed a Solution Pack replacement due to insufficient fluids. Ran Level 1 QC and pO2 was out low. Performed zero check, all ok. To test for leaking sensor cassette at sensor connector luer, I ran 2 system cycles, 3 external QC's, 4 blood samples and manually pumped solutions 1, 2,3 & 4 through sensor cassette without any leaks. | Unable to determine source of leak that customer stated at this time. Only change was the Solution pack due to insufficient samples. | CF4FF0E5A23B6FDD | No | Upon arrival the analyzer needed a Solution Pack replacement due to insufficient fluids. Ran Level 1 QC and pO2 was out low. Performed zero check, all ok. To test for leaking sensor cassette at sensor connector luer, I ran 2 system cycles, 3 external QC's, 4 blood samples and manually pumped solutions 1, 2,3 & 4 through sensor cassette without any leaks. | Unable to determine source of leak that customer stated at this time. Only change was the Solution pack due to insufficient samples.
ABL80 CO-OX | Upon arrival the analyzer needed a Solution Pack replacement due to insufficient fluids. Performed zero check, all ok. To test for leaking sensor cassette at sensor connector luer, I ran 2 system cycles, 3 external QC's, 4 blood samples and manually pumped solutions 1, 2,3 & 4 through sensor cassette without any leaks. | Unable to determine source of leak that customer stated at this time. Only change was the Solution pack due to insufficient samples. | CF4FF450A2496FE0 | No | Upon arrival the analyzer needed a Solution Pack replacement due to insufficient fluids. Performed zero check, all ok. To test for leaking sensor cassette at sensor connector luer, I ran 2 system cycles, 3 external QC's, 4 blood samples and manually pumped solutions 1, 2,3 & 4 through sensor cassette without any leaks. | Unable to determine source of leak that customer stated at this time. Only change was the Solution pack due to insufficient samples.
ABL90 FLEX | Tel. mit Herrn Hein: Beim Patienten Wolfgang Krämer werden keine Werte bei pCO2 und pO2 angezeigt, Probe mehrfach gemessen, immer ohne Werte. Wird DIESELBE Spritze unter einem anderen Namen gemessen, erscheinen die Werte! Serviceeinsatz erfolgt heute oder morgen. | Tel. mit Herrn Hein: Beim Patienten Wolfgang Krämer werden keine Werte bei pCO2 und pO2 angezeigt; Probe mehrfach gemessen, immer ohne Werte. Wird DIESELBE Spritze unter einem anderen Namen gemessen, erscheinen die Werte! Serviceeinsatz erfolgt heute oder morgen. | SWO-40212 | Yes | Tel. with Mr. Hein: In the patient Wolfgang Krämer no values are displayed for pCO2 and pO2; Sample measured several times, always without values. If THE syringe is measured under a different name, the values appear! Service is done today or tomorrow. | Tel. with Mr. Hein: In the patient Wolfgang Krämer no values are displayed for pCO2 and pO2; Sample measured several times, always without values. If THE syringe is measured under a different name, the values appear! Service is done today or tomorrow. | Re classified RMED | Telephone support
ABL90 FLEX | Kein Flüssigkeitsaustritt. Gerät manuell gespült. Einlaßnadel überprüft - ok. LP Wechsel durchgeführt (eingelegten Pack wieder eingesetzt) - ok. Kalibration und QC's durchgeführt. Mit Kunden abgelehnte Reklamation für SC durchgesprochen. Bei 2 SC maximale Testanzahl erreicht. | Kein Flüssigkeitsaustritt. Gerät manuell gespült. Einlaßnadel überprüft - ok. LP Wechsel durchgeführt (eingelegten Pack wieder eingesetzt) - ok. Kalibration und QC's durchgeführt. Mit Kunden abgelehnte Reklamation für SC durchgesprochen. Bei 2 SC maximale Testanzahl erreicht. | SWO-40271 | No | No liquid leakage. device manually rinsed. Inlet needle checked - ok. LP Change performed (inserted pack reinstated) - ok. calibration and QC's. Complaint for SC rejected with customers. Maximum test count reached at 2 SC. | No liquid leakage. device manually rinsed. Inlet needle checked - ok. LP Change performed (inserted pack reinstated) - ok. calibration and QC's. Complaint for SC rejected with customers. Maximum test count reached at 2 SC. | No fault found | Frequent SP replace
ABL800 BASIC | スペクトロメーター交換実施。 | スペクトロメーター交換実施。 | SWO-42948 | No | Spectrometer replacement was implemented. | Spectrometer replacement was implemented.
ABL820 FLEX | ABL überprüft. pCO2-Elektrode gereinigt und mit anderer Charge membraniert -> kein Erfolg PH Bg-Ventile ausgebaut und gereinigt. Mehrer Kalibrierungen OK! Gerät kalibreirt. QC durchgeführt. Anwender beobachten das Gerät die nächsten Tage! | ABL überprüft. pCO2-Elektrode gereinigt und mit anderer Charge membraniert -> kein Erfolg PH Bg-Ventile ausgebaut und gereinigt. Mehrer Kalibrierungen OK! Gerät kalibreirt. QC durchgeführt. Anwender beobachten das Gerät die nächsten Tage! | SWO-40280 | No | ABL. pCO2 electrode cleaned and membraned with other batch -> no success PH Bg valves removed and cleaned. More calibrations OK! Device calibrated. QC. Users will be watching the device for the next few days! | ABL. pCO2 electrode cleaned and membraned with other batch -> no success PH Bg valves removed and cleaned. More calibrations OK! Device calibrated. QC. Users will be watching the device for the next few days! | Defect | HW | pH/BG Module | Valve, air
ABL825 FLEX | Undichtigkeiten an Elektroden festgestellt. Alle Elyt..Met. Elektroden neu Membraniert. Manuelle Dichtigkeitskontrolle okay Kal und QC okay | Undichtigkeiten an Elektroden festgestellt. Alle Elyt..Met. Elektroden neu Membraniert. Manuelle Dichtigkeitskontrolle okay Kal und QC okay | SWO-40132 | No | leaks on electrodes. All Elyt.. Met. Electrodes re-membraneed. Manual tightness control okay Kal and QC okay | leaks on electrodes. All Elyt.. Met. Electrodes re-membraneed. Manual tightness control okay Kal and QC okay | Assisted Replacement | Electrode | Membrane | Leaking membrane
ABL825 FLEX | Nach Reinigung Bildschirmrahmen (leichter, minimaler Schleier) und ein- und Ausstecken des Monitorkabels lässt sich Bildschirm wieder betätigen. Nach Neustart mehrere Kals notwendig, da Oximodul nicht aktiv. | Nach Reinigung Bildschirmrahmen (leichter, minimaler Schleier) und ein- und Ausstecken des Monitorkabels lässt sich Bildschirm wieder betätigen. Nach Neustart mehrere Kals notwendig, da Oximodul nicht aktiv. | SWO-40282 | No | After cleaning the screen frame (light, minimal veil) and plugging in and unplugging the monitor cable, the screen can be re-operated. After rebooting several kals necessary because Oximodul is not active. | After cleaning the screen frame (light, minimal veil) and plugging in and unplugging the monitor cable, the screen can be re-operated. After rebooting several kals necessary because Oximodul is not active. | Cleaned | HW | DMS_PC | Touch screen cable
ABL825 FLEX | Keys, Remotesupport, alternierenden QC Plan |  | SWO-40352 | Yes | Keys, Remotesupport, alternierenden QC Plan |  | Re classified RMED | Setup
ABL835 FLEX | Fehlersuche am Analysator, Kal. Dateien überprüft, Filter gesetzt->immer wieder starke Schwankungen der Empf. bei pCO2, Elektrode wurde vom Anwender mehrmals neu Membraniert auch mit anderer Membrancharge, Messkammern sauber (alle im pH BG Modul überprüft und nochmals gereinigt), Kontake gereinigt mit Glasfaserpinsel gereinigt, Elektrode gereinigt->keine Besserung (Empf. Bei 78,3%), neue pCO2 Elektrode eingesetzt->Empf. bei 93,3% Tendenz steigend , QC ok (vorher auch zum Teil Grenzwertüberschreitungen und stark schwankend), Anwender wird dies weiter beobachten, mehrere Kal. ok System-und Funktionskontrolle i.O. | Fehlersuche am Analysator, Kal. Dateien überprüft, Filter gesetzt->immer wieder starke Schwankungen der Empf. bei pCO2, Elektrode wurde vom Anwender mehrmals neu Membraniert auch mit anderer Membrancharge, Messkammern sauber (alle im pH BG Modul überprüft und nochmals gereinigt), Kontake gereinigt mit Glasfaserpinsel gereinigt, Elektrode gereinigt->keine Besserung (Empf. Bei 78,3%), neue pCO2 Elektrode eingesetzt->Empf. bei 93,3% Tendenz steigend , QC ok (vorher auch zum Teil Grenzwertüberschreitungen und stark schwankend), Anwender wird dies weiter beobachten, mehrere Kal. ok System-und Funktionskontrolle i.O. | SWO-40291 | No | Troubleshooting the Analyzer, Kal. Files checked, filters set->again and again strong fluctuations of the empf. at pCO2, electrode was repeatedly re-logged by the user also with other membrane batch, measuring chambers clean (all checked in the pH BG module and cleaned again), kontake cleaned with glass fiber brush, electrode cleaned->no improvement (empf. At 78.3%), new pCO2 electrode used->Empf. at 93.3% trend increasing, QC ok (previously also partly exceeded limit values and strongly fluctuating), users will continue to observe this, several Cal. ok System and function control i.O. | Troubleshooting the Analyzer, Kal. Files checked, filters set->again and again strong fluctuations of the empf. at pCO2, electrode was repeatedly re-logged by the user also with other membrane batch, measuring chambers clean (all checked in the pH BG module and cleaned again), kontake cleaned with glass fiber brush, electrode cleaned->no improvement (empf. At 78.3%), new pCO2 electrode used->Empf. at 93.3% trend increasing, QC ok (previously also partly exceeded limit values and strongly fluctuating), users will continue to observe this, several Cal. ok System and function control i.O. | Assisted Replacement | Electrode | Electrode | pCO2
ABL825 FLEX | N A | INSTALLED KEYS | CF4FEDF2A21F522B | Yes | N A | INSTALLED KEYS | Re classified RMED | Setup
ABL825 FLEX | N A | INSTALLED KEYS | CF4FED0DA2465228 | Yes | N A | INSTALLED KEYS | Re classified RMED | Setup
AQT90 FLEX | Run safe positions once. Restarted the analyzer |  | SWO-40718 | No | Run safe positions once. Restarted the analyzer
AQT90 FLEX |  |  | CF503BE8A216C2C7 | No
ABL80 FLEX Basic version | Verificar el material sustituido por el cliente en el equipo. Se realizan instalaciones de SP y SC sin éxito por lo que requiere localizar material en otro laboratorio para solucionar el problema. Se instala un nuevo SP y SC solucionandose el problema. | Cambio de SC y SP. | CF513B0518270F50 | No | Verify the material replaced by the customer in the equipment. SP and SC installations are unsuccessful so it requires locating material in another lab to solve the problem. A new SP and SC is installed to resolve the issue. | SC and SP change.
ABL80 FLEX BASIC | バーコードの件はCPU交換実施。初期化エラーについては内部チューブ交換、センサーボード交換、シグナルデータボード交換を実施しました。 | バーコードの件はCPU交換実施。初期化エラーについては内部チューブ交換、センサーボード交換、シグナルデータボード交換を実施しました。 | SWO-42121 | No | The case of the bar code is cpu exchange. For initialization errors, internal tube replacement, sensor board replacement, and signal data board exchange were performed. | The case of the bar code is cpu exchange. For initialization errors, internal tube replacement, sensor board replacement, and signal data board exchange were performed.
ABL90 FLEX | Fehlersuche: Beim Patienten Krämer fehlen am 28.07. zwischen 11:33h und 12:47h sowohl die Ergebnisse für pO2 sowie für pCO2, Referenzbereiche sind bei den entsprechenden Proben NICHT eingetragen. KEINE zugehörigen Fehlermeldungen. Probenmessungen davor und danach sind unauffällig! Datenextrakt an MicBr zur Prüfung gesendet. | Fehlersuche: Beim Patienten Krämer fehlen am 28.07. zwischen 11:33h und 12:47h sowohl die Ergebnisse für pO2 sowie für pCO2; Referenzbereiche sind bei den entsprechenden Proben NICHT eingetragen. KEINE zugehörigen Fehlermeldungen. Probenmessungen davor und danach sind unauffällig! Datenextrakt an MicBr zur Prüfung gesendet. | SWO-40418 | No | Troubleshooting: In the patient Krämer on 28.07. between 11:33h and 12:47h both the results for pO2 and for pCO2 are missing; Reference ranges are NOT entered in the corresponding samples. NO related error messages. Sample measurements before and after are inconspicuous! Data extract sent to MicBr for review. | Troubleshooting: In the patient Krämer on 28.07. between 11:33h and 12:47h both the results for pO2 and for pCO2 are missing; Reference ranges are NOT entered in the corresponding samples. NO related error messages. Sample measurements before and after are inconspicuous! Data extract sent to MicBr for review. | Defect | SW | Result not reported
ABL90 FLEX | Gerät überprüft: kein Fehler feststellbar, QC Werte i.O., Daten gezogen, Ringversuchsergebnisse kopiert, mit Fr. Neumann vereinbart, das bei nächsten ringversuch (Ende Aug15) ein -R- Mitarbeiter anwesend ist | Gerät überprüft: kein Fehler feststellbar, QC Werte i.O., Daten gezogen, Ringversuchsergebnisse kopiert, mit Fr. Neumann vereinbart, das bei nächsten ringversuch (Ende Aug15) ein -R- Mitarbeiter anwesend ist | SWO-40629 | No | Device checked: no error detectable, QC values i.O., data drawn, ring test results copied, agreed with Ms. Neumann, which is present at next ring test (end of Aug15) a -R employee | Device checked: no error detectable, QC values i.O., data drawn, ring test results copied, agreed with Ms. Neumann, which is present at next ring test (end of Aug15) a -R employee | Unknown | External test
ABL825 FLEX | Glukose Elektrode geprüft, Gerät demontiert und dekontaminiert, alle Elektroden neu membraniert, alle Lösungen ausgetauscht, Dichtung an der PH Elektrode ersetzt, Funktionskontrolle ok |  | SWO-40462 | No | Glucose electrode tested, device dismantled and decontaminated, all electrodes re-membraneed, all solutions replaced, seal replaced at PH electrode, function control ok |  | Cleaned | Fluidic pathway
ABL835 FLEX |  |  | SWO-40361 | No |  |  | Defect | HW | El/Met Module | Glu/Lac
ABL837 FLEX | Creatine sensitivities are out Lactate status is out. | Found the Lactate without the Oring Installed an Oring. Membraned Creatine and debubled. | CF517700183359AA | No | Creatine sensitivities are out Lactate status is out. | Found the Lactate without the Oring Installed an Oring. Membraned Creatine and debubled. | Assisted Replacement | Electrode | Electrode
ABL835 FLEX | Keys installiert, Einstellungen vorgenommen | Keys installiert, Einstellungen vorgenommen | SWO-40485 | No | Keys installed, settings made | Keys installed, settings made
ABL825 FLEX | Keys installiert, Remote eingerichtet Lactat membraniert, | Keys installiert, Remote eingerichtet Lactat membraniert, | SWO-40484 | No | Keys installed, remotely set up lactate membraneed, | Keys installed, remotely set up lactate membraneed,
AQT90 FLEX | Se realiza limpieza general del analizador se encuentra el tubo de la bomba P2 suelto se reinstala se realizan todos los test por sistema servicio, se calibra y se comprueba funcionamiento calibrando el sistema. |  | CF5118261834DAD8 | No | General cleaning of the analyzer is performed the p2 pump tube loose is reinstalled all tests are performed by service system, is calibrated and operation is checked by calibrating the system.
AQT90 FLEX |  |  | CF516DAA18349C41 | No
TCM400-6 | found that monitor is pink. checked all the connections work tight. | Replaced the Flat cable and Display to correct the issue. | CF5116FE1815150B | No | found that monitor is pink. checked all the connections work tight. | Replaced the Flat cable and Display to correct the issue.
ABL80 FLEX CO-OX | ABL überprüft. Verstopfung im Hemolyzer vermutet. Verstopfung innerhalb der Jahreswartung (seperates Case) behoben. Gerät kalibriert. Intene QC's durchgeführt. | ABL überprüft. Verstopfung im Hemolyzer vermutet. Verstopfung innerhalb der Jahreswartung (seperates Case) behoben. Gerät kalibriert. Intene QC's durchgeführt. | SWO-40660 | No | ABL. constipation in the Hemolyzer. Fixed constipation within the annual maintenance (seperates case). device calibrated. Intene QC's. | ABL. constipation in the Hemolyzer. Fixed constipation within the annual maintenance (seperates case). device calibrated. Intene QC's.
ABL80 FLEX BASIC | Performed the Qualicheck4 Level 1, 2, 3 found results are with in the range, explained the problem with pre analytical Issue. | Performed the Qualicheck4 Level 1, 2, 3 found results are with in the range, explained the problem with pre analytical Issue. | SWO-40964 | No | Performed the Qualicheck4 Level 1, 2, 3 found results are with in the range, explained the problem with pre analytical Issue. | Performed the Qualicheck4 Level 1, 2, 3 found results are with in the range, explained the problem with pre analytical Issue.
ABL90 FLEX | Roma.ini Einträge überprüft und fehlende Codes 55=1, 07=1 , 23=0, 22=0 nachgetragen | Roma.ini Einträge überprüft und fehlende Codes 55=1, 07=1 , 23=0, 22=0 nachgetragen | SWO-40712 | No | Roma.ini Entries checked and missing codes 55=1, 07=1 , 23=0, 22=0 | Roma.ini Entries checked and missing codes 55=1, 07=1 , 23=0, 22=0 | Setup | Roma.ini | Scanner
ABL90 FLEX | Inletnadel und inletdichtung ersetzt. Solution Pack ersetzt. | Inletnadel und inletdichtung ersetzt. Solution Pack ersetzt. | SWO-40697 | No | Inlet needle and inlet seal replaced. Solution Pack replaced. | Inlet needle and inlet seal replaced. Solution Pack replaced. | Assisted Replacement | SP | FT
ABL90 FLEX | Gerät überprüft: An Solution Pack PCB sind Kontakte verbogen-> ersetzt, Gerät gespült, Anschlussdichtung ersetzt, Einlasnadel gespült, Netz aus ein Startphase, keine Verbesserung, die Serviceprog. für Fluid Transport durchgeführt-> Fehlermedlung, alten LP wieder eingesetzt | Gerät überprüft: An Solution Pack PCB sind Kontakte verbogen-> ersetzt, Gerät gespült, Anschlussdichtung ersetzt, Einlasnadel gespült, Netz aus ein Startphase, keine Verbesserung, die Serviceprog. für Fluid Transport durchgeführt-> Fehlermedlung, alten LP wieder eingesetzt | SWO-40633 | No | Device checked: On Solution Pack PCB contacts are bent-> replaced, device flushed, connection seal replaced, needle flushed, net from a start phase, no improvement, the service prog. for Fluid Transport performed-> Fault medlung, old LP reinstated | Device checked: On Solution Pack PCB contacts are bent-> replaced, device flushed, connection seal replaced, needle flushed, net from a start phase, no improvement, the service prog. for Fluid Transport performed-> Fault medlung, old LP reinstated | Cleaned | Fluidic pathway
ABL90 FLEX | Im Service Pumpenkal. LS justage durchgeführt und schon war die Kal. Lösung nicht mehr inhomogen. Kein ersatz Lösungspack mehr vor Ort, Kalibration, QC Kontrolle | Im Service Pumpenkal. LS justage durchgeführt und schon war die Kal. Lösung nicht mehr inhomogen. Kein ersatz Lösungspack mehr vor Ort, Kalibration, QC Kontrolle | SWO-40698 | No | In the service pump scale. LS adjustment and already the Kal. solution is no longer inhomogeneous. No more replacement solution pack on site, calibration, QC control | In the service pump scale. LS adjustment and already the Kal. solution is no longer inhomogeneous. No more replacement solution pack on site, calibration, QC control | Assisted Replacement | SC | Wetability
ABL90 FLEX | 更换新试剂包 | 更换新试剂包 | SWO-46044 | Yes | Replace a new kit | Replace a new kit | Re classified RMED
ABL820 FLEX |  | Replaced pH BG module PCB. Performed Zero current calibration, 1 and 2 PT cal, ran QC, all values close to peer 2 mean. | CF52AB818E381F8C | No |  | Replaced pH BG module PCB. Performed Zero current calibration, 1 and 2 PT cal, ran QC, all values close to peer 2 mean. | Defect | HW | pH/BG Module
AQT90 FLEX | Gerät geöffnet und das Inletdoor händisch geschlossen, Reinigung durchgeführt, | Gerät geöffnet und das Inletdoor händisch geschlossen, Reinigung durchgeführt, | SWO-40792 | No | device is opened and the Inletdoor is closed manually, cleaning is carried out, | device is opened and the Inletdoor is closed manually, cleaning is carried out,
ABL80 FLEX BASIC | 更换测试卡正常。 | 更换测试卡正常。 | SWO-45779 | Yes | The replacement test card is OK. | The replacement test card is OK. | Re classified RMED
ABL80 FLEX BASIC | 更换测试卡正常。 | 更换测试卡正常。 | SWO-45507 | Yes | The replacement test card is OK. | The replacement test card is OK. | Re classified RMED
ABL90 |  |  | CF53AF0A04482F04 | No
ABL90 | Checked data logs for errors. Pack will not seat properly. Replace inlet gasket with holder because of messages occurring. | Reset Pack insertion procedure in diagnostics. Replaced Inlet Gasket with holder. Start annual maintenance. | CF53B0A7041A26BC | No | Checked data logs for errors. Pack will not seat properly. Replace inlet gasket with holder because of messages occurring. | Reset Pack insertion procedure in diagnostics. Replaced Inlet Gasket with holder. Start annual maintenance. | Defect | HW | Inlet Gasket | leak
ABL90 FLEX | Gerät 090R578N004 installiert, Konfig. eingespielt Netzwerkverbindung hergestellt | Gerät 090R578N004 installiert, Konfig. eingespielt Netzwerkverbindung hergestellt | SWO-40937 | No | Device 090R578N004 installed, Config. recorded network connection established | Device 090R578N004 installed, Config. recorded network connection established | Unknown | Removed by RSUB | intervention required
ABL90 FLEX | Überprüfung LP => richtig aktiviert Überprüfung Datei Archiv Aktivitäten => Fehlermeldung seit 30.07. 13:23 Uhr => Lösungspack wurde 13:23 Uhr manuell entfernt Überprüfung LP über Serviceprogramme Fluid Transport => alles i.O. Reinigung Einlassmodul und Einlassdichtung LP reinstalliert Startphase abgewartet => alles i.O. tHb-Kal seit 26.05. fällig => tHb-Kal i.O. => Status 280,29 pA Empf. 100,3% | Überprüfung LP => richtig aktiviert Überprüfung Datei Archiv Aktivitäten => Fehlermeldung seit 30.07. 13:23 Uhr => Lösungspack wurde 13:23 Uhr manuell entfernt Überprüfung LP über Serviceprogramme Fluid Transport => alles i.O. Reinigung Einlassmodul und Einlassdichtung LP reinstalliert Startphase abgewartet => alles i.O. tHb-Kal seit 26.05. fällig => tHb-Kal i.O. => Status 280,29 pA Empf. 100,3% | SWO-40982 | No | Review LP => Correctly activated verification file archive activities => Error message since 30.07. 13:23 => Solution pack was manually removed 13:23 check LP via service programs Fluid Transport => all i.O. cleaning inlet module and inlet seal LP reinstalled start phase waited => all i.O. tHb-Kal since 26.05. due => tHb-Kal i.O. => Status 280,29 pA Empf. 100,3% | Review LP => Correctly activated verification file archive activities => Error message since 30.07. 13:23 => Solution pack was manually removed 13:23 check LP via service programs Fluid Transport => all i.O. cleaning inlet module and inlet seal LP reinstalled start phase waited => all i.O. tHb-Kal since 26.05. due => tHb-Kal i.O. => Status 280,29 pA Empf. 100,3% | User Error | incorrect sp installation
ABL90 FLEX | 更换测试卡正常。 | 更换测试卡正常。 | SWO-46719 | Yes | The replacement test card is OK. | The replacement test card is OK. | Re classified RMED |  | Also ribbon cable connection Pipette PCB and AU Board chafed and had to be replaced.
ABL800 BASIC | Oxi Modul PCB ausgetauscht, Complete Oximetry Test , Proteinentfernung und tHb Kal durchgeführt, pCO2 membraniert, Kal mehrfach durchgeführt | Oxi Modul PCB ausgetauscht, Complete Oximetry Test , Proteinentfernung und tHb Kal durchgeführt, pCO2 membraniert, Kal mehrfach durchgeführt | SWO-40950 | No | Oxi Module PCB Exchanged, Complete Oximetry Test, Protein Removal and tHb Kal Performed, pCO2 membraneed, Kal performed multiple times | Oxi Module PCB Exchanged, Complete Oximetry Test, Protein Removal and tHb Kal Performed, pCO2 membraneed, Kal performed multiple times | Defect | HW | Oxi module
ABL815 FLEX | Messkammer und Kappe der Referenzelektrode gereinigt, Lösungsverlauf danach OK. QC ohne Kal zur Probe durchgeführt, pCO2 Wert stieg bei S7735 von 63,1 mmHg auf 66,1 mmHg an. Referenzelektrode war mit Lot R1761 membraniert. E1001 mit Lot R1827 neu membraniert. |  | SWO-40862 | No | Measuring chamber and cap of the reference electrode cleaned, solution course then OK. QC performed without kal to the sample, pCO2 value increased at S7735 from 63.1 mmHg to 66.1 mmHg. Reference electrode was membraneed with solder R1761. E1001 with solder R1827. |  | Assisted Replacement | Electrode | Electrode
ABL825 FLEX | Neue Oxi PCB eingebaut.Oxi Kompl Test. tHb Kal durchgeführt. | Ersatz Oxi Modul eingebaut. pH Elektrode undicht..Dichtring ersetzt. Alle Dummydichtringe ersetzt | SWO-40851 | No | New Oxi PCB built-in. Oxi Kompl Test. tHb Kal. | Replacement Oxi module built-in. pH electrode leaking.. Sealing ring replaced. Replaced all dummy sealing rings | Defect | HW | Oxi module
ABL835 FLEX | Ventile überprüft ... i.O., lt Med.-Technik lag gebrochene Glaskapillare vor --> komplettes Schlauchsystem erneuert |  | SWO-40959 | No | Valves checked ... i.O., lt Med.-Technology had broken glass capillaries --> complete hose system renewed |  | User Error | Inlet module | Inlet Gasket
ABL825 FLEX | QC Modul PCB ersetzt.Probelauf erfolgreich Kal und QC abgewartet. | QC Modul PCB ersetzt.Probelauf erfolgreich Kal und QC abgewartet. | SWO-40763 | No | QC module PCB replaced. Test run successfully waited Kal and QC. | QC module PCB replaced. Test run successfully waited Kal and QC. | Defect | HW | AutoCheck Module | PCB
ABL825 FLEX | De instabiliteit van calcium wordt veroorzaakt door een gekreukeld membraan. Bij nader onderzoek blijkt ook dat er bloedresten achtergebleven zijn onder het membraan. Misschien is dit ook een oorzaak van de driften op de lactaat. Na vervangen van het Ca-membraan enkele keren gekalibreerd en instabiliteit niet meer waargenomen. QC gemeten: OK en ligt op target waarde. | De instabiliteit van calcium wordt veroorzaakt door een gekreukeld membraan. Bij nader onderzoek blijkt ook dat er bloedresten achtergebleven zijn onder het membraan. Misschien is dit ook een oorzaak van de driften op de lactaat. Na vervangen van het Ca-membraan enkele keren gekalibreerd en instabiliteit niet meer waargenomen. QC gemeten: OK en ligt op target waarde. | SWO-40979 | No | The instability of calcium is caused by a creased membrane. Further investigation also shows that blood residues have been left under the membrane. Perhaps this is also a cause of the drifts on the lactate. After replacing the Ca membrane several times calibrated and instability no longer observed. QC measured: OK and is at target value. | The instability of calcium is caused by a creased membrane. Further investigation also shows that blood residues have been left under the membrane. Perhaps this is also a cause of the drifts on the lactate. After replacing the Ca membrane several times calibrated and instability no longer observed. QC measured: OK and is at target value. | Assisted Replacement | Electrode | Membrane
ABL815 FLEX | OXI PCB ersetzt, Fehler tritt weiterhin sporadisch auf. Spektrometer erneuert, Funktion OK. | Kundin bittet um Rückruf und Service | SWO-45559 | No | REPLACED BY OXI PCB, error continues to occur sporadically. Spectrometer renewed, function OK. | Customer asks for recall and service
ABL825 FLEX | Flex-Q Modul wird auf Kunden Wunsch wieder ausgebaut. Wirdnicht benötigt. |  | SWO-40897 | Yes | Flex-Q module will be expanded again at the customer's request. Not needed. |  | Re classified RMED | Installation
ABL825 FLEX | Flex-Q Modul wird auf Kunden ausgebaut. |  | SWO-40896 | Yes | Flex-Q module is extended to customers. |  | Re classified RMED | Installation
ABL80 FLEX BASIC | 更换新测试卡 | 更换新测试卡 | SWO-45602 | No | Replace a new test card | Replace a new test card
ABL80 FLEX BASIC | 更换新测试卡 | 更换新测试卡 | SWO-45607 | No | Replace a new test card | Replace a new test card
ABL90 |  |  | CF56B74FF051B6DD | No |  |  | Assisted Replacement | SC | pCO2
ABL825 FLEX | Fehlersuche, Fixierung der Einlasskapillare, Funktionskontrolle, Ersatzteil konnte nicht gestellt werden, da nicht on board | Fehlersuche, Fixierung der Einlasskapillare, Funktionskontrolle, Ersatzteil konnte nicht gestellt werden, da nicht on board | SWO-45490 | No | Troubleshooting, fixing of the inlet capillaries, function control, spare part could not be provided, as not on board | Troubleshooting, fixing of the inlet capillaries, function control, spare part could not be provided, as not on board | Defect | HW | Inlet module | Probe lock
ABL80 FLEX BASIC | 更换测试卡正常。 | 更换测试卡正常。 | SWO-45785 | No | The replacement test card is OK. | The replacement test card is OK.
ABL80 FLEX BASIC | Removed the Sensor Cassette and Solution Pack, restarted the machine, kept the sensor cassette in refrigerator for 30 mins, installed both sensor cassette and solution pack, after sensor initialization machine is in ready mode. | Removed the Sensor Cassette and Solution Pack, restarted the machine, kept the sensor cassette in refrigerator for 30 mins, installed both sensor cassette and solution pack, after sensor initialization machine is in ready mode. | SWO-46643 | No | Removed the Sensor Cassette and Solution Pack, restarted the machine, kept the sensor cassette in refrigerator for 30 mins, installed both sensor cassette and solution pack, after sensor initialization machine is in ready mode. | Removed the Sensor Cassette and Solution Pack, restarted the machine, kept the sensor cassette in refrigerator for 30 mins, installed both sensor cassette and solution pack, after sensor initialization machine is in ready mode.
ABL80 FLEX BASIC | カセット交換 |  | SWO-46196 | No | Cassette replacement
ABL80 FLEX BASIC | 更换903-334一个正常。 | 更换903-334一个正常。 | SWO-45585 | No | Replace 903-334 one normal. | Replace 903-334 one normal.
ABL90 |  |  | CF57DB9F6628BC47 | No |  |  | Assisted Replacement | SC
ABL90 FLEX | ABL durchgespült, Sensorkassette und Ls-Pack erneuert. Startphase QC und Kalibration. Dem Kunden mitgeteilt, dass er Pico 50 verwenden soll>> Laut Kunden wurde sofort eine VE bestellt. | Rückruf des Technikers erwünscht | SWO-45594 | No | ABL flushed through, sensor cassette and Ls-Pack renewed. Start phase QC and calibration. The customer was informed that he should use Pico 50>> According to the customer, a VE was ordered immediately. | Technician recall desired | User Error | SC Expired
ABL90 FLEX | Lösungspack (Lot SB-07) eingebaut und (Lot SB-05) ausgebaut, System und Funktion i.O. --> zusätzlich 1 LP (Lot SK-03) vor Ort gelassen, da der Kunde momentan keinen mehr hat | Schaluchwege komplett erneuert, System- und Funktionskontrolle, System und Funktion i.O. | SWO-45900 | Yes | Solution pack (Lot SB-07) installed and (Lot SB-05) removed, system and function usually --> additionally 1 LP (Lot SK-03) left on site, since the customer currently has no more | Soluchwege completely renewed, system and function control, system and function i.O. | Re classified RMED | Double entry
ABL90 FLEX | 更换测试卡正常。 | 更换测试卡正常。 | SWO-45555 | No | The replacement test card is OK. | The replacement test card is OK. | Assisted Replacement | SC | pCO2
AQT90 FLEX | 更换903-175一个正常。 | 更换903-175一个正常。 | SWO-45518 | No | Replace 903-175 one normal. | Replace 903-175 one normal. | Defect | HW | Power Supply | Power Supply | 3264
ABL827 FLEX |  |  | CF57C0DC664BE4D4 | No |  |  | Adjustment | HW | FlexQ Module | Tray
ABL800 BASIC | Messkammer und Elektrode mit Proteinentferner gereinigt, LS Kal. Pumpen und Ventile geprüft. Probelauf i.O. Mehrere Kapillaren Blut fehlerfrei gemessen. Präanalytik besprochen. | der Anwender hat heute mehrere MBU Blutmessungen durchgeführt, jedes Mal meldet das Gerät die Probenfehler, z.B. Probe inhomogen; an einem anderen Gerät gemessen funktioniert es einwandfrei; | SWO-46290 | No | Measuring chamber and electrode cleaned with protein remover, LS Kal. pumps and valves. Test run in the first place Several capillaries of blood measured error-free. pre-analysis. | the user has today carried out several MBU blood measurements, each time the device reports the sample errors, e.g. sample inhomogeneous; measured on another device, it works perfectly; | Unknown
ABL825 FLEX | Baro Korrektur + 19mmHg Kal und QC ok QC PCO2 Werte nahe ZW Zur weiteren Beobachtung nicht abgeschlossen | Baro Korrektur + 19mmHg Kal und QC ok QC PCO2 Werte nahe ZW Zur weiteren Beobachtung nicht abgeschlossen | SWO-45861 | No | Baro Correction + 19mmHg Kal and QC ok QC PCO2 Values near ZW For further observation not completed | Baro Correction + 19mmHg Kal and QC ok QC PCO2 Values near ZW For further observation not completed | Adjustment | HW | Analyzer Control PCB | Barometer
ABL815 FLEX |  |  | CF589216DC312B21 | No |  |  | Defect | HW | Oxi module | Hemolyzer
ABL815 FLEX | Reinigung der Messmammern und bürsten der Elektrodenkontakte. Einige Elektrodenkontakte verschmutzt, gereinigt. pCO2 und Ref-Elektrode membraniert. Baro-Druck überprüft, korrigiert. QC Werte bei pCO2 in allen Leveln grenzwertig am unteren Grenzwert. Proteinentferung durchgeführt. Einlassdichtung verschmutzt, erneuert. Kalibration und QC o.K. Analysator messbereit. | Reinigung der Messmammern und bürsten der Elektrodenkontakte. Einige Elektrodenkontakte verschmutzt, gereinigt. pCO2 und Ref-Elektrode membraniert. Baro-Druck überprüft, korrigiert. QC Werte bei pCO2 in allen Leveln grenzwertig am unteren Grenzwert. Proteinentferung durchgeführt. Einlassdichtung verschmutzt, erneuert. Kalibration und QC o.K. Analysator messbereit. | SWO-46001 | No | Cleaning of the measuring buntings and brushing the electrode contacts. Some electrode contacts dirty, cleaned. pCO2 and ref-electrode. Baro print checked, corrected. QC values at pCO2 in all levels borderline at the lower limit. protein deferification. Inlet seal dirty, renewed. Calibration and QC o.K. analyzer ready for measurement. | Cleaning of the measuring buntings and brushing the electrode contacts. Some electrode contacts dirty, cleaned. pCO2 and ref-electrode. Baro print checked, corrected. QC values at pCO2 in all levels borderline at the lower limit. protein deferification. Inlet seal dirty, renewed. Calibration and QC o.K. analyzer ready for measurement. | Cleaned | HW | pH/BG Module | Contacts
ABL835 FLEX | Found machine not recognizing operator log in barcode or bottle barcodes when trying to replace solutions. | Copied INI settings from another analyzer that I had just reprogrammed for the same issue. Loading INI setting was a direct initiative from techline. | CF57BA59661C4C55 | No | Found machine not recognizing operator log in barcode or bottle barcodes when trying to replace solutions. | Copied INI settings from another analyzer that I had just reprogrammed for the same issue. Loading INI setting was a direct initiative from techline. | Setup | Ini files | Barcode Reader
AQT90 FLEX |  |  | CF57E60C66379292 | No |  |  | Adjustment | HW | Process module | Process wheel | Reference cup
AQT90 FLEX |  |  | CF58A4F7DC236463 | Yes |  |  | Re classified RMED | Setup | Training
AQT90 FLEX | Messcassette aus dem storagewheel entfernt, Probemessungen durchgeführt, | Kundin bittet um einen Einsatz des zuständigen Technikers, bitte um Kontaktaufnahme. | SWO-46046 | No | Measuring cassette removed from the storagewheel, sample measurements carried out, | Customer asks for the use of the responsible technician, please contact us. | Defect | Consumables | Test Kit | Cartridge | Broken
ABL80 FLEX BASIC | 更换两张测试卡正常。 | 更换两张测试卡正常。 | SWO-45959 | No | Replacing the two test cards is ok. | Replacing the two test cards is ok.
ABL80 FLEX BASIC | 更换新测试卡 | 更换新测试卡 | SWO-45795 | No | Replace a new test card | Replace a new test card
ABL80 CO-OX | Pumped all four solutions with no problems in diagnostics. Unable to perform system cycle, ls adjust or oxi diagnostics due to error. Powered off on, swapped sensor cassettes, no change. | Disassembled analyzer and flushed hemolyzer with water. Powered on, performed oxi diagnostics & ls adjust, all passed. | CF58FC57DC306FF5 | No | Pumped all four solutions with no problems in diagnostics. Unable to perform system cycle, ls adjust or oxi diagnostics due to error. Powered off on, swapped sensor cassettes, no change. | Disassembled analyzer and flushed hemolyzer with water. Powered on, performed oxi diagnostics & ls adjust, all passed.
ABL80 FLEX BASIC | 更换新测试卡 | 更换新测试卡 | SWO-45783 | No | Replace a new test card | Replace a new test card
ABL80 FLEX BASIC |  |  | SWO-45827 | No
ABL80 FLEX BASIC | 更换新测试卡 | 更换新测试卡 | SWO-45824 | No | Replace a new test card | Replace a new test card
ABL80 FLEX BASIC | 更换测试卡 | 更换测试卡 | SWO-45967 | No | Replace the test card | Replace the test card
ABL80 FLEX CO-OX | 更换测试卡正常。 | 更换测试卡正常。 | SWO-45816 | No | The replacement test card is OK. | The replacement test card is OK.
ABL90 FLEX | Dem Kunden eine VE Pico 50 vorbeigebracht und da gelassen, wenn die Lieferung beim Kunden ankommt werde ich die eine Packung Pico 50 wieder abholen. | Rückruf des Technikers erwünscht | SWO-45876 | No | The customer will bring a VE Pico 50 and leave there when the delivery arrives at the customer I will pick up the one pack of Pico 50 again. | Technician recall desired | Assisted Replacement | SC | Wetability
ABL90 | Analyzer failed to open due to a corrupt roma.db or roma.ini file. | Deleted roma.db and roma.ini. Replaced with backup versions of roma.db and roma.ini. This resolved the problem. After restart pCO2, Cl, Glu and Lac all failed calibration. It took several calibrations to bring them all into range. | CF58F7AEDC2A8AD1 | No | Analyzer failed to open due to a corrupt roma.db or roma.ini file. | Deleted roma.db and roma.ini. Replaced with backup versions of roma.db and roma.ini. This resolved the problem. After restart pCO2, Cl, Glu and Lac all failed calibration. It took several calibrations to bring them all into range. | Defect | SW | Roma.ini
ABL90 FLEX | Neue Einlassdichtung eingebaut , Neustart durchgeführt |  | SWO-45904 | No | New inlet seal built-in, restart performed |  | Defect | HW | Inlet Gasket
ABL90 | Reviewed data logs happens every 6-7 samples | Flushed hemolyzer in service with protein removal. Passes all OXI tests in service. | CF58C01BDC41D254 | No | Reviewed data logs happens every 6-7 samples | Flushed hemolyzer in service with protein removal. Passes all OXI tests in service. | Cleaned | Fluidic pathway | Protein remover
ABL90 | Po2 failing QC error 521 inhomogenous sample. Checked all tubing and gaskets. Old SC Run # 668-82 | Replaced waste tubing. Ran all service checks. Replaced SC.Ran the Unit overnight. Ran comparison samples with Respiratory department analyzer. Comparison OK. | CF58EB40DC2CE776 | No | Po2 failing QC error 521 inhomogenous sample. Checked all tubing and gaskets. Old SC Run # 668-82 | Replaced waste tubing. Ran all service checks. Replaced SC. Ran the Unit overnight. Ran comparison samples with Respiratory department analyzer. Comparison OK. | Assisted Replacement | SC | Wetability
TCM CombiM | 更换945-737一个正常。 | 更换945-737一个正常。 | SWO-45596 | No | Replace the 945-737 one normal. | Replace the 945-737 one normal.
ABL90 FLEX | 更换测试卡正常 | 更换测试卡正常 | SWO-45821 | No | Replace the test card ok | Replace the test card ok | Assisted Replacement | SC | K+
ABL90 FLEX | 更换测试卡正常 | 更换测试卡正常 | SWO-45741 | No | Replace the test card ok | Replace the test card ok | Assisted Replacement | SC | pCO2
ABL825 FLEX | OXI szelepcső cseréje, hemolizátor és levegőző csőszakasz hypo-s tisztítása, pH BG mérőkamra tisztítása, LS és pumpa kalibrációk, teljes OXI teszt, protein eltávolítás, esedékesség miatti tHb kalibrálás elvégzése, kalibrációk, QC mérés |  | SWO-45938 | No | OXY Valve Tube Replacement, hypo cleaning of hemolytic and ventilation pipe, pH BG cleaning, LS and pump calibrations, full OXY test, protein removal, due date of tHb calibration, calibrations, QC measurement |  | Defect | HW | Oxi module | Valve tubing
ABL825 FLEX | Review logs, check the valves in fluid module. The V2 valve was slow to close. | Install new V2 and install new seal on Cal 1 solution. replace bent probe and inlet tube. | CF58E9D1DC63CEBC | No | Review logs, check the valves in fluid module. The V2 valve was slow to close. | Install new V2 and install new seal on Cal 1 solution. replace bent probe and inlet tube. | Defect | HW | Fluidic manifold | Valve 2
ABL820 FLEX | pO2 values are too low for level 2 and high for level 1. | Tried new membranes with existing electrode. Replaced p02 electrde(under warranty) and used new membranes. | CF58FD43DC0FBF10 | No | pO2 values are too low for level 2 and high for level 1. | Tried new membranes with existing electrode. Replaced p02 electrde(under warranty) and used new membranes. | Assisted Replacement | Electrode | Electrode | pO2
ABL825 FLEX | 跟换902-885 | 跟换902-885 | SWO-45755 | No | Follow 902-885 | Follow 902-885 | Defect | HW | Inlet module | Friction detector unit
ABL837 FLEX | Creatine continues to fail during Calibrations in Sensivity. | Replaced both Electrodes, since we have tried many other things. Brought system to ready. Travel time from Hotel-Fedex-Hospital-Naples. | CF58F4E3DC089F39 | No | Creatine continues to fail during Calibrations in Sensivity. | Replaced both Electrodes, since we have tried many other things. Brought system to ready. Travel time from Hotel-Fedex-Hospital-Naples. | Assisted Replacement | Electrode | Electrode | Crea
ABL835 FLEX | Messkammer geprüft und gereinigt, Elektroden neu membraniert, Funktionskontrolle ok | Messkammer geprüft und gereinigt, Elektroden neu membraniert, Funktionskontrolle ok | SWO-46035 | No | Measuring chamber tested and cleaned, electrodes re-membraneed, function control ok | Measuring chamber tested and cleaned, electrodes re-membraneed, function control ok | Cleaned | Fluidic pathway
ABL825 FLEX | Inlet Modul geprüft, Einlassdichtung und Nadel ersetzt, Lac. Und pCO2 Elektrode neu membraniert, Service tHb Kalibration durchgeführt, Funktionskontrolle ok | Inlet Modul geprüft, Einlassdichtung und Nadel ersetzt, Lac. Und pCO2 Elektrode neu membraniert, Service tHb Kalibration durchgeführt, Funktionskontrolle ok | SWO-46036 | No | Inlet module tested, inlet seal and needle replaced, Lac. And pCO2 electrode re-membraneed, service tHb calibration performed, function control ok | Inlet module tested, inlet seal and needle replaced, Lac. And pCO2 electrode re-membraneed, service tHb calibration performed, function control ok | Defect | HW | Inlet module | Probe
ABL835 FLEX | RECEIVED KEYS | INPUTTED KEYS, BUT MISSING HIS LIS COMM KEY (WILL BE GOOD FOR 117 DAYS), EMAILED RAME AND WILL UPDATE KEY ONCE RECEIVED. | CF58FDCCDC3CBCA9 | No | RECEIVED KEYS | INPUTTED KEYS, BUT MISSING HIS LIS COMM KEY (WILL BE GOOD FOR 117 DAYS), EMAILED RAME AND WILL UPDATE KEY ONCE RECEIVED.
ABL825 FLEX | Oxi Modul PCB ausgetauscht, SW angepasst, einkalibriert, Meldung erscheint nach ein paar Minuten wieder, Oxi Modul komplett ausgetauscht | Oxi Modul PCB ausgetauscht, SW angepasst, einkalibriert, Meldung erscheint nach ein paar Minuten wieder, Oxi Modul komplett ausgetauscht | SWO-45953 | No | Oxi module PCB replaced, SW adjusted, calibrated, message appears again after a few minutes, Oxi module completely replaced | Oxi module PCB replaced, SW adjusted, calibrated, message appears again after a few minutes, Oxi module completely replaced | Defect | HW | Oxi module
ABL825 FLEX | Fliesswege gespült, Ventile überprüft --> Air Ventil Blutreste entfernt und Airkanal gespült, i.O., pCO2 neu membraniert, pH Dichtungsring kontrolliert, Messkammer gereinigt | Nach Rücksprache hat sich ein Schlauch gelöst, den der Kunde wieder befestigt hat, Gerät wieder i.O. | SWO-45954 | No | Flow paths rinsed, valves checked --> Air valve blood residue removed and air duct rinsed, i.O., pCO2 re-membraneed, pH sealing ring checked, measuring chamber cleaned | After consultation, a hose has loosened, which the customer has reattached, device again i.O. | Cleaned | Fluidic pathway
AQT90 FLEX | Fließwege mit Proteinentferner durchgespült, LQCs mehrmals wiederholt. SD Analyse durchgeführt. | Fließwege mit Proteinentferner durchgespült, LQCs mehrmals wiederholt. SD Analyse durchgeführt. | SWO-45950 | No | Flow paths rinsed with protein remover, LQCs repeated several times. SD analysis. | Flow paths rinsed with protein remover, LQCs repeated several times. SD analysis. | Cleaned | HW | CRP | Fluidic pathway | LQC
AQT90 FLEX | Opened the analyzer, Found that there is dust in the process wheel and reference .Cleaned the process wheel and reference and removed the dust. |  | SWO-46619 | No | Opened the analyzer, Found that there is dust in the process wheel and reference . Cleaned the process wheel and reference and removed the dust. |  | Cleaned | HW | Process module | Process wheel | Reference cup
TCM/Tosca | UNIT WOULD CALIBRATE FIND AND UPON COMPLETION WOULD SAY READY AND SECONDS LATER SCREEN WOULD SAY COMMUNICATION ERROR | IN THIS EXACT ORDER.REPLACED TOSCA MODULE,REPLACED COMMUNICATION BOARD,REPLACED INTERFACE BOARD AND LASTLY REPLACED CPU AND COMMUNICATION ERROR WENT AWAY. | CF58CA2CDC564FC3 | No | UNIT WOULD CALIBRATE FIND AND UPON COMPLETION WOULD SAY READY AND SECONDS LATER SCREEN WOULD SAY COMMUNICATION ERROR | IN THIS EXACT ORDER. REPLACED TOSCA MODULE,REPLACED COMMUNICATION BOARD,REPLACED INTERFACE BOARD AND LASTLY REPLACED CPU AND COMMUNICATION ERROR WENT AWAY. | Defect | HW | PCB, Main CPU | PCB, Main CPU
CombiM | Confirmed the reported problem. | Located and replaced the electrode module which was defective. | CF58D9B6DC4E43BB | No | Confirmed the reported problem. | Located and replaced the electrode module which was defective.
ABL90 FLEX | Einlassdichtungshaler erneuert da Gummi verschoben und verdreht, einmal als Ersatz vor Ort gelassen Einlasssonde einmal als Ersatz vor Ort gelassen System durchgespült Analysator kalibriert und überprüft | Einlassdichtungshaler erneuert da Gummi verschoben und verdreht, einmal als Ersatz vor Ort gelassen Einlasssonde einmal als Ersatz vor Ort gelassen System durchgespült Analysator kalibriert und überprüft | SWO-46113 | No | Inlet seal errater renewed since rubber shifted and twisted, once left as a replacement on site inlet probe once as a replacement on site left system flushed through analyzer calibrated and checked | Inlet seal errater renewed since rubber shifted and twisted, once left as a replacement on site inlet probe once as a replacement on site left system flushed through analyzer calibrated and checked | Defect | HW | Inlet Gasket
ABL90 | client replaced Inlet, inlet gasket and connector gasket, no change. Ran disg all passed | Replaced SC with client stock, resolved | CF5A09D3524E6A7F | No | client replaced Inlet, inlet gasket and connector gasket, no change. Ran disg all passed | Replaced SC with client stock, resolved | Assisted Replacement | SC | Wetability
ABL90 | User keeps getting unable to pump solutions | Preformed annual pm on analyzer to change out tubes along the fluid path to correct the issue. | CF59F885525121D0 | No | User keeps getting unable to pump solutions | Preformed annual pm on analyzer to change out tubes along the fluid path to correct the issue. | Defect | HW | Tubes
ABL90 | This continued from yesterday, and the only change is that Glu will now not pass calibration as well due to high sensitivity. | Replaced the sensor cassette (provided by customer), and this resolved the problem. | CF59EB35520813C5 | No | This continued from yesterday, and the only change is that Glu will now not pass calibration as well due to high sensitivity. | Replaced the sensor cassette (provided by customer), and this resolved the problem. | Assisted Replacement | SC | Glu
ABL90 FLEX | JW Teile eingebaut. Alle Wege durchgespült. Alle Solutions werden angesaugt. Immer noch Fehler --SC erneuert. Gerät ok. | JW Teile eingebaut. Alle Wege durchgespült. Alle Solutions werden angesaugt. Immer noch Fehler --SC erneuert. Gerät ok. | SWO-46233 | No | JW parts built-in. All paths washed through. All solutions are sucked in. Still error --SC renewed. Device ok. | JW parts built-in. All paths washed through. All solutions are sucked in. Still error --SC renewed. Device ok. | Assisted Replacement | SC | Wetability
ABL90 | Unit Out of Service upon arrival. | Replaced SC, Ran Bloods. | CF6DF44FB240E517 | No | Unit Out of Service upon arrival. | Replaced SC, Ran Bloods. | Assisted Replacement | SC | Wetability
ABL90 FLEX | Schaluchwege komplett erneuert, System- und Funktionskontrolle, System und Funktion i.O. | Schaluchwege komplett erneuert, System- und Funktionskontrolle, System und Funktion i.O. | SWO-46149 | No | Soluchwege completely renewed, system and function control, system and function i.O. | Soluchwege completely renewed, system and function control, system and function i.O. | Defect | HW | Tubes | YM
ABL90 FLEX | Gerät überprüft: Timeserver Einstellungen nach RDEU Vorgaben, Update now-> Uhrzeit 1h zurück, Autom. Sommer Winterzeitumstellung Haken raus rein, Update now Uhrzeit i.O., Netz aus ein Uhrzeit i.O., am nachbargerät sind Einstellungen identisch0900639N001 dieses Gerät macht den Fehler nicht Gerät hat SW 3.1 MR5 installiert | Gerät überprüft: Timeserver Einstellungen nach RDEU Vorgaben, Update now-> Uhrzeit 1h zurück, Autom. Sommer Winterzeitumstellung Haken raus rein, Update now Uhrzeit i.O., Netz aus ein Uhrzeit i.O., am nachbargerät sind Einstellungen identisch0900639N001 dieses Gerät macht den Fehler nicht Gerät hat SW 3.1 MR5 installiert | SWO-46086 | No | Device checked: Timeserver settings according to RDEU specifications, update now-> time 1h back, autom. Summer winter time change hook out, update now time i.O., mains from a time i.O., on the neighbordevice are settings identical0900639N001 this device makes the error does not device has SW 3.1 MR5 installed | Device checked: Timeserver settings according to RDEU specifications, update now-> time 1h back, autom. Summer winter time change hook out, update now time i.O., mains from a time i.O., on the neighbordevice are settings identical0900639N001 this device makes the error does not device has SW 3.1 MR5 installed | Defect | SW | Time setting | Time Zone
ABL90 FLEX | Gerät überprüft: Kontaktfahne nachjustiert, Test im Serviceprog. i.O., SC ersetzt, Startphase abgewartet, Hr. kühn beobachtet Verhalten des ABL und meldet sich ggf. wieder | Gerät überprüft: Kontaktfahne nachjustiert, Test im Serviceprog. i.O., SC ersetzt, Startphase abgewartet, Hr. kühn beobachtet Verhalten des ABL und meldet sich ggf. wieder | SWO-46085 | No | Device checked: Contact flag adjusted, test in service test. i.O., SC replaced, start phase awaited, Mr. boldly observes behavior of the ABL and reports again if necessary | Device checked: Contact flag adjusted, test in service test. i.O., SC replaced, start phase awaited, Mr. boldly observes behavior of the ABL and reports again if necessary | Adjustment | HW | Sensor door
ABL90 FLEX |  |  | SWO-45970 | No |  |  | Defect | HW | Touch screen
ABL817 FLEX | ASTM.ini - Datei Version überspielt, Neustart des Analysators durchgeführt, Kontakt zu Convorx Herr Korte (8:30 Uhr) --> Übermittlung funktioniert, System und Funktion i.O. | ASTM.ini - Datei Version überspielt, Neustart des Analysators durchgeführt, Kontakt zu Convorx Herr Korte (8:30 Uhr) --> Übermittlung funktioniert, System und Funktion i.O. | SWO-46042 | No | ASTM.ini - File version overplayed, restart of the analyzer performed, contact to Convorx Mr. Korte (8:30 a.m.) --> Transmission works, system and function i.O. | ASTM.ini - File version overplayed, restart of the analyzer performed, contact to Convorx Mr. Korte (8:30 a.m.) --> Transmission works, system and function i.O. | Setup | HW | DMS_PC | His/Lis
ABL820 FLEX | Flushed Fluid path, valves, etc. | Ran Liquid sensor cal. | CF63D1050258B1E7 | No | Flushed Fluid path, valves, etc. | Ran Liquid sensor cal. | Cleaned | Fluidic pathway | -
ABL825 FLEX | checked unit and ran Diag and found pump cal's barely in spec | replace both tubes with client stock and ran service cal on LS and Pump cal's. ran 6 syringe samples all ok. reviewed ICU log and found no unstable measurements in the last 24 hours. | CF5A394252546A82 | No | checked unit and ran Diag and found pump cal's barely in spec | replace both tubes with client stock and ran service cal on LS and Pump cal's. ran 6 syringe samples all ok. reviewed ICU log and found no unstable measurements in the last 24 hours. | Defect | HW | pH/BG Module | Pump tubing
ABL820 FLEX | Checked connections and monitor. Monitor does light up but does not display anything. | Replaced the CPU and tested. | CF5A0E6F52523EFE | No | Checked connections and monitor. Monitor does light up but does not display anything. | Replaced the CPU and tested. | Defect | HW | DMS_PC
ABL837 FLEX | Pumpenschlauch pH- BG-Modul erneuert. | Pumpenschlauch pH- BG-Modul erneuert. | SWO-46159 | No | Pump hose pH- BG module renewed. | Pump hose pH- BG module renewed. | Assisted Replacement | HW | pH/BG Module | Pump tubing
AQT90 FLEX | CLEANED THE PROCESS WHEEL AND THE OPTICAL UNIT. |  | SWO-46644 | No | CLEANED THE PROCESS WHEEL AND THE OPTICAL UNIT. |  | Cleaned | HW | Process module | Optical Unit
ABL90 FLEX | Fehlersuche: installiertes LP sowie ein weiteres, gebrauchtes, eingesetzt => Pressuretest nicht bestanden. Einlassdichtung ersetzt, Fließwege überprüft, weiteres neues LP eingesetzt, Werte nun OK. Installation des LP, Startphase, Kal und QC Messungen => alle Werte OK. | Fehlersuche: installiertes LP sowie ein weiteres, gebrauchtes, eingesetzt => Pressuretest nicht bestanden. Einlassdichtung ersetzt, Fließwege überprüft, weiteres neues LP eingesetzt, Werte nun OK. Installation des LP, Startphase, Kal und QC Messungen => alle Werte OK. | SWO-46453 | No | Troubleshooting: installed LP as well as another used, used => Pressuretest failed. Inlet seal replaced, flow paths checked, further new LP inserted, values now OK. Installation of the LP, start phase, Kal and QC measurements => all values OK. | Troubleshooting: installed LP as well as another used, used => Pressuretest failed. Inlet seal replaced, flow paths checked, further new LP inserted, values now OK. Installation of the LP, start phase, Kal and QC measurements => all values OK. | Defect | HW | Inlet Gasket
ABL90 FLEX | Gerät durchgespült, Kpl. Oxy & pO2 Test: i.O., Kal durchgeführt, fehlerhaft QC wiederholt:i.O. | Gerät durchgespült, Kpl. Oxy & pO2 Test: i.O., Kal durchgeführt, fehlerhaft QC wiederholt:i.O. | SWO-46361 | No | Device flushed through, Kpl. Oxy & pO2 Test: i.O., Kal performed, faulty QC repeated:i.O. | Device flushed through, Kpl. Oxy & pO2 Test: i.O., Kal performed, faulty QC repeated:i.O. | Cleaned | Fluidic pathway
ABL90 FLEX | Fehlersuche, Fließwege gespült, div. Kleinteile gew., Starphase abgewartet, Servicekoffer überreicht, Kal. + QC i.O. | Fehlersuche; Fließwege gespült; div. Kleinteile gew.; Starphase abgewartet; Servicekoffer überreicht; Kal. + QC i.O. | SWO-46391 | No | Troubleshooting; Flow paths rinsed; various small parts; Star phase awaited; Service case handed over; Kal. + QC in the first place | Troubleshooting; Flow paths rinsed; various small parts; Star phase awaited; Service case handed over; Kal. + QC in the first place | User Error | Inlet opened while measure…
TCM4 |  |  | SWO-46295 | No
AQT90 FLEX | 更换910-304一个正常。 | 更换910-304一个正常。 | SWO-45969 | No | Replace 910-304 one normal. | Replace 910-304 one normal. | Defect | HW | Storage module | Barcode reader, Storage wheel | BR1
AQT90 FLEX | 更换903-171一个正常。 | 更换903-171一个正常。 | SWO-45966 | No | Replace 903-171 one normal. | Replace 903-171 one normal. | Defect | HW | Aspiration Module | Bend needle | 4794
ABL805 FLEX | checked the tubing in inlet module checked the probe and inlet gasket for leaks -found liqid poring out of the probe and inlet gasket | changed tubing on inlet module and replaced nipples also checked waste tubing for clogs | CF5B5DDFC85F08C2 | Yes | checked the tubing in inlet module checked the probe and inlet gasket for leaks -found liqid poring out of the probe and inlet gasket | changed tubing on inlet module and replaced nipples also checked waste tubing for clogs | Re classified RMED | SenDx
ABL800 BASIC | PH Elektrode und Messkammer gereinigt kal,s durchgeführt Empf. zu niedrig und Messung instabil, PH Elektrode getauscht, Kal,s und QC durchgeführt | PH Elektrode und Messkammer gereinigt kal,s durchgeführt Empf. zu niedrig und Messung instabil, PH Elektrode getauscht, Kal,s und QC durchgeführt | SWO-46387 | No | PH electrode and measuring chamber cleaned kal,s performed empf. too low and measurement unstable, PH electrode exchanged, Kal,s and QC performed | PH electrode and measuring chamber cleaned kal,s performed empf. too low and measurement unstable, PH electrode exchanged, Kal,s and QC performed | Assisted Replacement | Electrode | Electrode | pH
ABL825 FLEX | Inlet Modul geprüft, Einlassdichtung und Nadel ersetzt, Funktionskontrolle ok | Inlet Modul geprüft, Einlassdichtung und Nadel ersetzt, Funktionskontrolle ok | SWO-46465 | No | Inlet module tested, inlet seal and needle replaced, function control ok | Inlet module tested, inlet seal and needle replaced, function control ok | Assisted Replacement | HW | Inlet Module | Inlet Gasket
ABL835 FLEX |  |  | CF5B6A28C8401FF2 | No |  |  | Defect | Electrode | Electrode | Sealing Ring
AQT90 FLEX |  |  | CF5C824C3E2D4A09 | Yes |  |  | Re classified RMED
ABL80 FLEX CO-OX | Fliesswege gereinigt. System ok | Fliesswege gereinigt. System ok | SWO-46581 | No | Flow paths cleaned. System ok | Flow paths cleaned. System ok
ABL80 FLEX BASIC | Performed BIORAD Level 2 QC and found all values within range and close to mean value. Met the Dr. Aswini clarified the same | Performed BIORAD Level 2 QC and found all values within range and close to mean value. Met the Dr. Aswini clarified the same | SWO-46626 | No | Performed BIORAD Level 2 QC and found all values within range and close to mean value. Met the Dr. Aswini clarified the same | Performed BIORAD Level 2 QC and found all values within range and close to mean value. Met the Dr. Aswini clarified the same
ABL90 FLEX | ABL manuell gespült aus GS mitgebrachten LP eingesetzt Gerät kalibriert und überprüft | ABL manuell gespült aus GS mitgebrachten LP eingesetzt Gerät kalibriert und überprüft | SWO-46549 | No | ABL manually rinsed from GS brought LP inserted device calibrated and checked | ABL manually rinsed from GS brought LP inserted device calibrated and checked | Assisted Replacement | SP | After replacement Blank-Tests were performed (everything was fine) and analyzer was configured like before. LQC measurement was performed.
ABL90 FLEX | Fehlersuche, alle Fließwege gespült, LP neu inst., Startphase abgewartet = Fehler SC, gl. SC erneut initialisiert = pCO2-Fehler, neue SC eingelesen, Initialisierung abgewartet = Lac-Fehler, Anwender ist mit Lac-Fehler vorerst einverstanden, meldet sich bei Bedarf | Fehlersuche; alle Fließwege gespült; LP neu inst.; Startphase abgewartet = Fehler SC; gl. SC erneut initialisiert = pCO2-Fehler; neue SC eingelesen; Initialisierung abgewartet = Lac-Fehler; Anwender ist mit Lac-Fehler vorerst einverstanden; meldet sich bei Bedarf | SWO-46464 | No | Troubleshooting; all flow paths rinsed; LP new inst.; Start phase waited = error SC; Gl. SC reinitialized = pCO2 error; new SC; Initialization awaited = Lac error; User agrees with Lac errors for the time being; reports if necessary | Troubleshooting; all flow paths rinsed; LP new inst.; Start phase waited = error SC; Gl. SC reinitialized = pCO2 error; new SC; Initialization awaited = Lac error; User agrees with Lac errors for the time being; reports if necessary | Cleaned | Fluidic pathway | For a few days analyzer stayed in our workshop and was restarted a couple of times. No error messages during that time.
ABL90 FLEX | Abzug des ABL90 (090R0593N007) zwecks Reparatur und aufstellen eines Rep.-Überbrückungsgerätes (090R0636N006) --> Rep.-Überbrückungsgerät i.O. | Schaluchwege komplett erneuert, System- und Funktionskontrolle, System und Funktion i.O. | SWO-46670 | No | Deduction of the ABL90 (090R0593N007) for repair and setting up of a rep. bridging device (090R0636N006) --> Rep. bridging device i.O. | Soluchwege completely renewed, system and function control, system and function i.O. | Defect | HW | Unknown
ABL90 | Analyzer in the UIR screen, unable to pump solutions. | Replaced the inlet gasket, inlet probe, solution pack and sensor cassette. | CF5CC2023E297F7C | No | Analyzer in the UIR screen, unable to pump solutions. | Replaced the inlet gasket, inlet probe, solution pack and sensor cassette. | Assisted Replacement | SC | Wetability
ABL90 FLEX | 更换新测试卡 | 更换新测试卡 | SWO-47577 | No | Replace a new test card | Replace a new test card | Assisted Replacement | SC | pH
ABL800 BASIC | 更换902-833一个正常。 | 更换902-833一个正常。 | SWO-45981 | No | Replace 902-833 one normal. | Replace 902-833 one normal. | Defect | HW | Oxi module
ABL825 FLEX | 更换K电极膜正常。 | 更换K电极膜正常。 | SWO-46436 | No | Replace the K electrode film is normal. | Replace the K electrode film is normal. | Assisted Replacement | Electrode | Membrane | K+
AQT90 FLEX |  |  | CF5C7C593E1A1FF5 | No |  |  | Cleaned | HW | Process module | Process Unit Sensor
AQT90 FLEX | Nach Reparatur in Zentral Gerät zurück in GS NTproBNP neu kalibriert und Test-Kit mit zum Kunden genommen Geräteaufstellung beim Kunden | Nach Reparatur in Zentral Gerät zurück in GS NTproBNP neu kalibriert und Test-Kit mit zum Kunden genommen Geräteaufstellung beim Kunden | SWO-46571 | No | After repair in central device back in GS NTproBNP recalibrated and test kit taken to the customer device assembly at the customer | After repair in central device back in GS NTproBNP recalibrated and test kit taken to the customer device assembly at the customer | Defect | HW | Analyzer control Module | Cable | 3264
AQT90 FLEX | Fehlersuche Nadeldefekt Nadel getauscht und kalibriert Systemreinigung durchgeführt Systemüberprüfung OK QC Messung nicht durchgeführt da keine QC vor Ort Diese ist bereits bestellt, Kunden darauf hingewiesen diese vor der Benutzung des Gerätes selbstständig durchzuführen info@kardiologie-innenstadt.com | Fehlersuche Nadeldefekt Nadel getauscht und kalibriert Systemreinigung durchgeführt Systemüberprüfung OK QC Messung nicht durchgeführt da keine QC vor Ort Diese ist bereits bestellt, Kunden darauf hingewiesen diese vor der Benutzung des Gerätes selbstständig durchzuführen info@kardiologie-innenstadt.com | SWO-46448 | No | Troubleshooting needle defect needle exchanged and calibrated system cleaning performed system review OK QC measurement not performed since no QC on site This is already ordered, customers advised to perform this independently before using the device info@kardiologie-innenstadt.com | Troubleshooting needle defect needle exchanged and calibrated system cleaning performed system review OK QC measurement not performed since no QC on site This is already ordered, customers advised to perform this independently before using the device info@kardiologie-innenstadt.com | Defect | HW | Aspiration Module | Bend needle
ABL80 FLEX BASIC | 更换测试卡正常 | 更换测试卡正常 | SWO-46652 | No | Replace the test card ok | Replace the test card ok
ABL80 FLEX BASIC | 更换测试卡正常。 | 更换测试卡正常。 | SWO-46660 | No | The replacement test card is OK. | The replacement test card is OK.
ABL80 FLEX BASIC | Temporarily placed a standby analyzer. |  | SWO-46630 | No | Temporarily placed a standby analyzer.
ABL90 FLEX | Fehlersuche Test im Service Programm Einlaßdichtung und Halterung erneuert Wechsel SP erfolgte bereits durch Kunden Einlaufverfahren | Fehlersuche Test im Service Programm Einlaßdichtung und Halterung erneuert Wechsel SP erfolgte bereits durch Kunden Einlaufverfahren | SWO-46663 | No | Troubleshooting test in the service program inlet seal and bracket renewed Replacement SP already carried out by customer inlet procedure | Troubleshooting test in the service program inlet seal and bracket renewed Replacement SP already carried out by customer inlet procedure | Defect | HW | Inlet Gasket
ABL90 FLEX | 更换910-285修复机器 | 更换910-285修复机器 | SWO-46503 | No | Replace the 910-285 repair machine | Replace the 910-285 repair machine | Defect | HW | Power Supply
AQT90 FLEX | 更换910-339，903-186，636-666各一个正常。 | 更换910-339，903-186，636-666各一个正常。 | SWO-46513 | No | Replace 910-339, 903-186, 636-666 each normal. | Replace 910-339, 903-186, 636-666 each normal. | Defect | HW | Aspiration Module | Barcode reader, Inlet
ABL80 FLEX BASIC | 更换测试卡正常 | 更换测试卡正常 | SWO-46817 | No | Replace the test card ok | Replace the test card ok
ABL80 FLEX BASIC | 更换试剂包正常。 | 更换试剂包正常。 | SWO-46911 | No | The replacement kit is OK. | The replacement kit is OK.
ABL80 FLEX BASIC | 更换测试卡正常。 | 更换测试卡正常。 | SWO-47100 | No | The replacement test card is OK. | The replacement test card is OK.
ABL80 CO-OX | Checked all fluid paths for blockage. Also getting intermittent fluid blockage errors. | Flushed hemolyzer. Started annual maintenance. | CF606950A05A9783 | No | Checked all fluid paths for blockage. Also getting intermittent fluid blockage errors. | Flushed hemolyzer. Started annual maintenance.
ABL80 FLEX BASIC |  |  | SWO-46697 | No
ABL80 FLEX CO-OX | Gerät kompl. gereinigt, Reinigungsprogr. durchgeführt, Abfallwege und Anschlüsse gereinigt, | Gerät kompl. gereinigt, Reinigungsprogr. durchgeführt, Abfallwege und Anschlüsse gereinigt, | SWO-46785 | No | Device compl. cleaned, cleaning progr. carried out, waste paths and connections cleaned, | Device compl. cleaned, cleaning progr. carried out, waste paths and connections cleaned,
ABL80 |  |  | CF60504DA02AAA35 | No
ABL80 CO-OX | found that they are failing calibration side of system cycle | replaced hemolyzer which corrected the issue. | CF606BD2A02D7C81 | No | found that they are failing calibration side of system cycle | replaced hemolyzer which corrected the issue.
ABL90 FLEX | Gerät überprüft: Wasteanschluß am Gerät stark verfärbt, Ventilschlauch am Oxy Modul undicht, Gerät zerlegt, gereinigt, Schlauchsystem erneuert, LP installiert Startphase:i.O., Kal & QC: i.O. | Gerät überprüft: Wasteanschluß am Gerät stark verfärbt, Ventilschlauch am Oxy Modul undicht, Gerät zerlegt, gereinigt, Schlauchsystem erneuert, LP installiert Startphase:i.O., Kal & QC: i.O. | SWO-46909 | No | Device checked: Waste connection on the device strongly discolored, valve hose leaking to Oxy module, device disassembled, cleaned, hose system renewed, LP installs start phase:i.O., Kal & QC: i.O. | Device checked: Waste connection on the device strongly discolored, valve hose leaking to Oxy module, device disassembled, cleaned, hose system renewed, LP installs start phase:i.O., Kal & QC: i.O. | Defect | HW | Oxi module | Pinch Valve tube
ABL90 FLEX | Rückspulung durchgeführt, LS Justage , Pumpen Kal, Pressuretest, Clot detect Remove durchgeführt. Gerät in der Startphase. | Nadel, Einlass Dichtung und Inlet connector gasket erneuert, Rückspülung mehrfach durchgeführt, Service Programme durchgeführt | SWO-46886 | No | Rewind performed, LS adjustment, pumps Kal, Pressuretest, Clot detect Remove performed. device in the start-up phase. | Needle, inlet seal and inlet connector gasket renewed, backwash performed multiple times, service programs performed | Cleaned | Fluidic pathway | inlet
ABL90 FLEX | Fehlersuche Versuch Pointer zu calibrieren scheitert, da Bildschirm (Touch screen ) nicht reagiert Kabelverbindungen überprüft Austausch komplette Bildschirmeinheit incl. Kabel - anschl.i.O. | Fehlersuche Versuch Pointer zu calibrieren scheitert, da Bildschirm (Touch screen ) nicht reagiert Kabelverbindungen überprüft Austausch komplette Bildschirmeinheit incl. Kabel - anschl.i.O. | SWO-46857 | No | Troubleshooting Attempt to caliblate pointer fails because screen (touch screen) unresponsive cable connections checked replacement complete screen unit incl. cable - t.i.O. | Troubleshooting Attempt to caliblate pointer fails because screen (touch screen) unresponsive cable connections checked replacement complete screen unit incl. cable - t.i.O. | Adjustment | Touch screen cable | Screen reinstalled
ABL90 FLEX | Erneuerung nach Urlaub, Spülen der Fließwege, Kalibrationen und Qualitätskontrollen ohne Beanstandungen. | Erneuerung nach Urlaub, Spülen der Fließwege, Kalibrationen und Qualitätskontrollen ohne Beanstandungen. | SWO-46840 | No | Renewal after holidays, rinsing of flow paths, calibrations and quality controls without complaints. | Renewal after holidays, rinsing of flow paths, calibrations and quality controls without complaints. | Cleaned | Fluidic pathway
ABL90 FLEX | Nadel erneuert, Service Programme durchgeführt, LP R reinstalliert, Kal wiederholt nicht bestanden, Sensor Kassette erneuert, der Kunde bekommt einen Rückruf morgen früh | Nadel erneuert, Service Programme durchgeführt, LP R reinstalliert, Kal wiederholt nicht bestanden, Sensor Kassette erneuert, der Kunde bekommt einen Rückruf morgen früh | SWO-46885 | No | Needle renewed, service programs performed, LP R reinstalled, Kal repeatedly failed, sensor cassette renewed, customer gets a recall tomorrow morning | Needle renewed, service programs performed, LP R reinstalled, Kal repeatedly failed, sensor cassette renewed, customer gets a recall tomorrow morning | Replaced | probe
ABL90 FLEX | 更换测试卡正常。 | 更换测试卡正常。 | SWO-47138 | No | The replacement test card is OK. | The replacement test card is OK. | Assisted Replacement | SC | Reference
ABL90 FLEX | 更换902-918一个正常 | 更换902-918一个正常 | SWO-46640 | No | Replace 902-918 one normal | Replace 902-918 one normal | Defect | HW | LS error
AQT90 FLEX |  |  | SWO-46699 | No |  |  | Defect | HW | Analyzer control Module | Barcode reader, display | -
AQT90 FLEX | 更换903-032一个正常。 | 更换903-032一个正常。 | SWO-45617 | No | Replace 903-032 one normal. | Replace 903-032 one normal. | Defect | HW | Incubator Module | Cup lifter | PCB
ABL837 FLEX | Replaced Gas 2 cylinder and performed calibration. | Replaced Gas 2 cylinder and performed calibration. | SWO-47770 | No | Replaced Gas 2 cylinder and performed calibration. | Replaced Gas 2 cylinder and performed calibration. | Assisted Replacement | Gas Cylinder | Gas 2
ABL827 FLEX | El Met modulog pumpeslanger udskiftet, slutafprøvet. | El Met modulog pumpeslanger udskiftet, slutafprøvet. | SWO-46864 | No | El Met Moduland Pump hoses replaced, final tested. | El Met Moduland Pump hoses replaced, final tested. | Defect | HW | El/Met Module | Glu/Lac
ABL827 FLEX | Upon arrival, found unit having Oximetry Spectrum mismatch errors on QCS on intermitent Basis. | Replaced Hemolyzer and Lamp unit. Realigned Oximetry system running all service alignment programs. | CF602BEBA04B6AB1 | No | Upon arrival, found unit having Oximetry Spectrum mismatch errors on QCS on intermitent Basis. | Replaced Hemolyzer and Lamp unit. Realigned Oximetry system running all service alignment programs. | Defect | HW | Oxi module | Hemolyzer
ABL827 FLEX | Upon speaking to customer, found unit running within 2SD range for CAP Survey signifying unit running correctly. Checked unit at customers request. | Checked all Oximetry values in service programs. All values within Radiometer acceptable ranges. | CF62C12D8C30AAE6 | Yes | Upon speaking to customer, found unit running within 2SD range for CAP Survey signifying unit running correctly. Checked unit at customers request. | Checked all Oximetry values in service programs. All values within Radiometer acceptable ranges. | Re classified RMED | Courtesy
ABL825 FLEX | LOOKED IN DATA LOGS. THE ANALYZER HAS BEEN HAVING INTERMITTENT LACTATE DRIFT ERRORS ON BOTH 2PT AND 1 PT CALS. IT HAPPENING 3 TO 5 TIMES A DAY. THE CUSTOMER HAS ALREADY REPLACED DIFFERENT LOT OF LACTATE MEMBRANE. ALSO DID WARRANTY REPLACEMENT OF LACTATE ELECTRODE. WAS ADVISED BY KEITH CLARK TO CHANGE OUT ALL TUGING, CHANGE EL MET PCB, AND MEASURING CHAMBER AND CONNECTOR. | REPLACED ALL TUBING, REPLACED EL MET PCB, REPLACED EL MET MEASURING CHAMBER, REPLACED LACTATE CONNECTOR, REPLACED FLAT MEMBRANE VALVE ON EL MET MODULE. | CF6052B5A01BA866 | No | LOOKED IN DATA LOGS. THE ANALYZER HAS BEEN HAVING INTERMITTENT LACTATE DRIFT ERRORS ON BOTH 2PT AND 1 PT CALS. IT HAPPENING 3 TO 5 TIMES A DAY. THE CUSTOMER HAS ALREADY REPLACED DIFFERENT LOT OF LACTATE MEMBRANE. ALSO DID WARRANTY REPLACEMENT OF LACTATE ELECTRODE. WAS ADVISED BY KEITH CLARK TO CHANGE OUT ALL TUGING, CHANGE EL MET PCB, AND MEASURING CHAMBER AND CONNECTOR. | REPLACED ALL TUBING, REPLACED EL MET PCB, REPLACED EL MET MEASURING CHAMBER, REPLACED LACTATE CONNECTOR, REPLACED FLAT MEMBRANE VALVE ON EL MET MODULE. | Defect | HW | El/Met Module | PCB | Lac
ABL825 FLEX | Siehe SWO 46837 | Kunde hat bereits Pumpenkalibration durchgeführt. Fehler ist weiterhin vorhanden. Der Kunde wünscht einen Serviceeinsatz zur Beseitigung der Gerätestörung. | SWO-46979 | Yes | See SWO 46837 | Customer has already performed pump calibration. Error still exists. The customer wants a service to eliminate the equipment malfunction. | Re classified RMED | Telephone support
ABL825 FLEX | Überprüfung Messkammer und pH und Ref Elektrode Leichter Film auf Glasfläche der pH-Elektrode und Verunreinigungen an unterer Dichtung => Reinigung der Elektrode FlexQ Positionen verschoben => Positionskalibierung FlexQ Modul Mehrere Testmessungen mit Blut ohne Instabilitäten | Überprüfung Messkammer und pH und Ref Elektrode Leichter Film auf Glasfläche der pH-Elektrode und Verunreinigungen an unterer Dichtung => Reinigung der Elektrode FlexQ Positionen verschoben => Positionskalibierung FlexQ Modul Mehrere Testmessungen mit Blut ohne Instabilitäten | SWO-46832 | No | Checking measuring chamber and pH and ref electrode Light film on glass surface of the pH electrode and impurities on lower seal => Cleaning of the electrode FlexQ positions moved => Position calibration FlexQ module Multiple test measurements with blood without Instabilities | Checking measuring chamber and pH and ref electrode Light film on glass surface of the pH electrode and impurities on lower seal => Cleaning of the electrode FlexQ positions moved => Position calibration FlexQ module Multiple test measurements with blood without Instabilities | Cleaned | Fluidic pathway | pH/BG Module
ABL825 FLEX | Reinigung Einlassmodul, Einlassdichtung und Nadel Nadel gerichtet und fehlenden probe lock angebracht (Ersatzteil war zum WE-Einsatz 02.08. nicht verfügbar) Reinigung der Messkammern bei pH, Ref und El Met Modul Durchführung Proteinentfernung Membranierung von Na => Kal2 => Lac und Glu mit zu tiefer Empfindlichkeit, Na i.O. Membranierung von Lac und Glu Mehrmals Kal bis i.O. QC Level 1-3 i.O. => Level 4 mit zu niedrigem pCO2 => pCO2 neu membraniert und Messkammer gesäubert Mehrmals Kal und QC Level 4 | Reinigung Einlassmodul, Einlassdichtung und Nadel Nadel gerichtet und fehlenden probe lock angebracht (Ersatzteil war zum WE-Einsatz 02.08. nicht verfügbar) Reinigung der Messkammern bei pH, Ref und El Met Modul Durchführung Proteinentfernung Membranierung von Na => Kal2 => Lac und Glu mit zu tiefer Empfindlichkeit, Na i.O. Membranierung von Lac und Glu Mehrmals Kal bis i.O. QC Level 1-3 i.O. => Level 4 mit zu niedrigem pCO2 => pCO2 neu membraniert und Messkammer gesäubert Mehrmals Kal und QC Level 4 | SWO-46833 | No | Cleaning inlet module, inlet seal and needle needle directed and missing test lock attached (spare part was not available for WE insert 02.08. not available) Cleaning of the measuring chambers at pH, Ref and El Met module carrying out protein removal membraneing of Na => Kal2 => Lac and Glu with too low sensitivity, Na i.O. membraneing of Lac and Glu Multiple Kal to I.O. QC Level 1-3 i.O. => Level 4 with too low pCO2 => pCO2 re-membraneed and measuring chamber cleaned Several times Kal and QC Level 4 | Cleaning inlet module, inlet seal and needle needle directed and missing test lock attached (spare part was not available for WE insert 02.08. not available) Cleaning of the measuring chambers at pH, Ref and El Met module carrying out protein removal membraneing of Na => Kal2 => Lac and Glu with too low sensitivity, Na i.O. membraneing of Lac and Glu Multiple Kal to I.O. QC Level 1-3 i.O. => Level 4 with too low pCO2 => pCO2 re-membraneed and measuring chamber cleaned Several times Kal and QC Level 4 | Defect | HW | Inlet module | Probe lock
ABL835 FLEX | Recurring 0618 errors. TAmbient voltage high. | Reseated all connections between waste detector and AU PCB. | CF604BA1A019FA30 | No | Recurring 0618 errors. TAmbient voltage high. | Reseated all connections between waste detector and AU PCB. | Adjustment | HW | Cable | Reseated
ABL820 FLEX | Troubleshot and found clots in the inlet gasket, inlet probe, and throughout the waste tubing. | Replaced the inlet gasket and the inlet probe tube (Provided by customer), I cleaned the entire waste path to remove all clots. | CF605370A02DE894 | No | Troubleshot and found clots in the inlet gasket, inlet probe, and throughout the waste tubing. | Replaced the inlet gasket and the inlet probe tube (Provided by customer), I cleaned the entire waste path to remove all clots. | Cleaned | Fluidic pathway | Inlet module | Gasket
ABL80 FLEX BASIC | 更换新测试卡 | 更换新测试卡 | SWO-46933 | No | Replace a new test card | Replace a new test card
ABL80 FLEX BASIC | 更换测试卡正常。 | 更换测试卡正常。 | SWO-46927 | No | The replacement test card is OK. | The replacement test card is OK.
ABL80 FLEX Basic version |  |  | CF61A1C016218967 | No
ABL80 CO-OX | Verified error. Powered on and screen stays black and the computer beeps for ~3 seconds and no beep for 1 second then continually repeats. Removed CF card with the same results. Pressed the Reset button but same error. Ordered parts for overnight delivery. |  | CF61F38D16281F67 | No | Verified error. Powered on and screen stays black and the computer beeps for ~3 seconds and no beep for 1 second then continually repeats. Removed CF card with the same results. Pressed the Reset button but same error. Ordered parts for overnight delivery.
ABL80 FLEX BASIC | 更换测试卡正常。 | 更换测试卡正常。 | SWO-46970 | No | The replacement test card is OK. | The replacement test card is OK.
ABL80 CO-OX | Customer unable to download onto USB drive due to write protection. | Gave customer USB without write protection. Downloaded event data logs. | CF618672162A98CC | No | Customer unable to download onto USB drive due to write protection. | Gave customer USB without write protection. Downloaded event data logs.
ABL90 | Upon arrival, found unit time set correctly by customer. | Checked all service programs related to time. Checked unit not connected to internet time. Checked and commited time software upgrade on unit. Reviewed unit logs with customer and found unit being rebooted a few times over past month. Also found unit not connected to dedicated UPS. Recommended to customer a dedicated UPS and investigate reason for unit being powered down unneccesarrily. | CF61D3331629A23B | No | Upon arrival, found unit time set correctly by customer. | Checked all service programs related to time. Checked unit not connected to internet time. Checked and commited time software upgrade on unit. Reviewed unit logs with customer and found unit being rebooted a few times over past month. Also found unit not connected to dedicated UPS. Recommended to customer a dedicated UPS and investigate reason for unit being powered down unneccesarrily. | Setup | Time Setting
ABL90 FLEX | Nadel, Einlass Dichtung und Inlet connector gasket erneuert, Rückspülung mehrfach durchgeführt, Service Programme durchgeführt | Nadel, Einlass Dichtung und Inlet connector gasket erneuert, Rückspülung mehrfach durchgeführt, Service Programme durchgeführt | SWO-47092 | No | Needle, inlet seal and inlet connector gasket renewed, backwash performed multiple times, service programs performed | Needle, inlet seal and inlet connector gasket renewed, backwash performed multiple times, service programs performed | Cleaned | Fluidic pathway | inlet parts
ABL90 | SC change gave unable to pump solutions error after replacement Performed all fluid transport tests in service, performed extended flush | Found no errors in fluid transport system when tubing replaced SC, SP test was failing QC 2 and 3 with SC installed | CF61BB58161D7063 | No | SC change gave unable to pump solutions error after replacement Performed all fluid transport tests in service, performed extended flush | Found no errors in fluid transport system when tubing replaced SC, SP test was failing QC 2 and 3 with SC installed
ABL90 | Examined sensor cassette | Replaced sensor cassette. Verified operation of ABL90. | CF62BA468C611DBF | No | Examined sensor cassette | Replaced sensor cassette. Verified operation of ABL90. | Assisted Replacement | SC | Wetability
ABL90 | Comprobación lector. O bien no lee o bien no lee el número correcto. Sustitución lector código de barras. Comprobación del mismo. Cal , QC. Prueba transmisión |  | CF61D75A1640F262 | No | Reader check. Either you don't read or you don't read the right number. Replacement reader barcode. Checking the same. Cal , QC. Transmission test |  | Defect | HW | Barcode Reader
AQT90 FLEX |  |  | SWO-46921 | No |  |  | Defect | HW | Storage module | Barcode reader, Storage wheel
AQT90 FLEX |  |  | SWO-46908 | No |  |  | Defect | HW | Process module | Cup dryer, PCB
ABL827 FLEX |  |  | CF6402DA0252426A | No |  |  | Defect | HW | DMS_PC | Speaker
ABL825 FLEX | Electrode is less than a year old. Doesn't appear to be an electrode issue as no damage seen on electrode tip. However I am seeing excessive bubbles in the EL Met module. Inspected El Met valves and found V7 is torn and fluid got into the valve causing damage to the piston. | Installed parts & tubing from a PM kit for diagnostic purposes, same issue. Replaced the El Met V7 flat membrane valve to resolve the issue, cals passing. | CF617B6216637B03 | No | Electrode is less than a year old. Doesn't appear to be an electrode issue as no damage seen on electrode tip. However I am seeing excessive bubbles in the EL Met module. Inspected El Met valves and found V7 is torn and fluid got into the valve causing damage to the piston. | Installed parts & tubing from a PM kit for diagnostic purposes, same issue. Replaced the El Met V7 flat membrane valve to resolve the issue, cals passing. | Defect | HW | Fluidic manifold | Valve, air
ABL825 FLEX | ran qc and checked logs | Replaced hemolyzer assy. | CF61F979160A9606 | No | ran qc and checked logs | Replaced hemolyzer assy. | Defect | HW | Oxi module | Hemolyzer
ABL825 FLEX |  |  | CF6559CA781792FD | No |  |  | Assisted Replacement | Electrode | Electrode | pCO2
AQT90 FLEX | Nadel Kalibriert, Reinigung nach neuen Service Angaben durchgeführt, QC 3 mal wiederholt ohne Beanstandung, QC Lot CRP2 10756 :- 119 mg L , QC Lot CRP 1 10754 :- 44,5 mg L , CRP Test Kassette Lot 10779 | Nadel Kalibriert, Reinigung nach neuen Service Angaben durchgeführt, QC 3 mal wiederholt ohne Beanstandung, QC Lot CRP2 10756 :- 119 mg L , QC Lot CRP 1 10754 :- 44,5 mg L , CRP Test Kassette Lot 10779 | SWO-47093 | No | Needle Calibrated, cleaning carried out according to new service specifications, QC 3 times repeated without complaint, QC Lot CRP2 10756 :- 119 mg L , QC Lot CRP 1 10754 :- 44.5 mg L , CRP Test Cassette Lot 10779 | Needle Calibrated, cleaning carried out according to new service specifications, QC 3 times repeated without complaint, QC Lot CRP2 10756 :- 119 mg L , QC Lot CRP 1 10754 :- 44.5 mg L , CRP Test Cassette Lot 10779 | Cleaned | Fluidic pathway | CRP | LQC
ABL80 FLEX BASIC | 更换新测试卡 | 更换新测试卡 | SWO-47182 | No | Replace a new test card | Replace a new test card
ABL80 FLEX BASIC |  |  | SWO-47107 | No
ABL80 FLEX BASIC |  | 更换903-334和903-924各一个没恢复。再借902-972一个。 | SWO-47120 | No |  | Replacements 903-334 and 903-924 each did not recover. Borrow another 902-972.
ABL80 FLEX BASIC | replaced the sensor cassette with new | replaced the sensor cassette with new | SWO-47817 | No | replaced the sensor cassette with new | replaced the sensor cassette with new
ABL90 FLEX | Wie mit Frau Zehntner vereinbart Reparatur-Leihgerät geliefert. Konfiguration übertragen. Patientenlayout überprüft - ok. Kalibration und QC's durchgeführt. | Wie mit Frau Zehntner vereinbart Reparatur-Leihgerät geliefert. Konfiguration übertragen. Patientenlayout überprüft - ok. Kalibration und QC's durchgeführt. | SWO-47328 | No | As agreed with Mrs. Zehntner, repair rental equipment was delivered. configuration. Patient layout checked - ok. calibration and QC's. | As agreed with Mrs. Zehntner, repair rental equipment was delivered. configuration. Patient layout checked - ok. calibration and QC's. | Unknown | Replaced analyzer
ABL90 | Paul H. determined unit was having Cal errors. Tried several SCs' and SPs'. Issue's remained.Suspect Hemolyzer Cntrl Bd. Pinch Valve with possible sticking issue. | Replaced Hemolyzer Assy.Realingned Pinch Valve. | CF631E5D8C2363C1 | No | Paul H. determined unit was having Cal errors. Tried several SCs' and SPs'. Issue's remained. Suspect Hemolyzer Cntrl Bd. Pinch Valve with possible sticking issue. | Replaced Hemolyzer Assy.Realingned Pinch Valve. | Defect | Hemolyzer | Clot
ABL90 | Examined sensor cassette | Replaced sensor cassette. Verified operation of ABL90. | CF62ACFB8C0E09B1 | No | Examined sensor cassette | Replaced sensor cassette. Verified operation of ABL90. | Assisted Replacement | SC | Wetability
ABL90 FLEX |  | 更换910-324一个后，一半灵敏一半不灵敏。怀疑是液晶屏有问题。 | SWO-47106 | No |  | After replacing 910-324, half sensitive and half insensitive. The suspicion is that there is a problem with the LCD screen. | Defect | HW | Touch screen
ABL90 | Comprobación de la configuración interna (Roma) , pruebas. Tiene la misma configuración que los equipos que si leen sin dígito. Se añaden dos lineas al archivo roma, se graba y reinicia. Lectura de etiquetas OK. Cal QC Ok | Comprobación de la configuración interna (Roma) , pruebas. Tiene la misma configuración que los equipos que si leen sin dígito. Se añaden dos lineas al archivo roma, se graba y reinicia. | CF62D5B98C622E8E | No | Checking the internal configuration (Rome), tests. It has the same settings as computers as if they read without digits. Two lines are added to the roma file, recorded and restarted. Reading OK tags. Cal QC Ok | Checking the internal configuration (Rome), tests. It has the same settings as computers as if they read without digits. Two lines are added to the roma file, recorded and restarted. | Setup | Roma.ini | Barcode Reader
AQT90 FLEX | カートリッジ除去で改善。 | カートリッジ除去で改善。 | SWO-47338 | No | Improved with cartridge removal. | Improved with cartridge removal. | User Error | Consumables | Test kit | Cartridge | CRP
AQT90 FLEX | インレットドア調整実施 |  | SWO-47105 | No | Inlet door adjustment |  | Defect | SW | Aspiration Module | Inlet Door
ABL820 FLEX |  |  | SWO-47153 | No |  |  | Assisted Replacement | Electrode | Electrode | pH
ABL825 FLEX | Air vent in fluidic module blocked. Calibrations and QC OK. | Cleaned air vent nipple. | CF63031A8C477E62 | No | Air vent in fluidic module blocked. Calibrations and QC OK. | Cleaned air vent nipple. | Cleaned | Fluidic pathway | Fluidic manifold | Air vent
ABL827 FLEX | Checked data logs for errors. Review information. While checking El Met module, found passive pre-heater hold down screws loose. | Tighten pre-heater screws. Performed zero current adjustment in service mode. Replaced Lactate electrode under warranty. Restart calibrations. Waiting for both parameters to stabilize. Glucose and Lactate still drifting continue to troubleshoot. Exchanged El Met module for fault isolation. Power up restart calibrations. | CF62FA618C2DA4BD | No | Checked data logs for errors. Review information. While checking El Met module, found passive pre-heater hold down screws loose. | Tighten pre-heater screws. Performed zero current adjustment in service mode. Replaced Lactate electrode under warranty. Restart calibrations. Waiting for both parameters to stabilize. Glucose and Lactate still drifting continue to troubleshoot. Exchanged El Met module for fault isolation. Power up restart calibrations. | Defect | HW | El/Met Module | Glu/Lac
AQT90 FLEX | NEED TO CALIBRATE AGAIN WITH SAME LOT NO I.E 9906 | NEED TO CALIBRATE AGAIN WITH SAME LOT NO I.E 9906 | SWO-47782 | No | NEED TO CALIBRATE AGAIN WITH SAME LOT NO I.E 9906 | NEED TO CALIBRATE AGAIN WITH SAME LOT NO I.E 9906 | User Error | Re calibration | NT-proBNP | 0
ABL80 FLEX BASIC | 更换新测试卡 | 更换新测试卡 | SWO-47395 | No | Replace a new test card | Replace a new test card
ABL90 | TRIED ADDING, TRIED EDITING, TRIED REBOOTING, DID NOT WORK. | SAVED KNOWN GOOD SETUP AND INSTALLED THAT SETUP. | CF642A4E022A3D9B | No | TRIED ADDING, TRIED EDITING, TRIED REBOOTING, DID NOT WORK. | SAVED KNOWN GOOD SETUP AND INSTALLED THAT SETUP. | Defect | SW | Edit Layout
ABL90 | Customer replaced SP twice per unit troubleshooting on his own. Customer called for service, I had him flush system then replace SC, same issue persists. Getting 1084 inhomogeneous cal 3 soluton erros. | Installed PM kit for diagnostic purposes. | CF641DB3022F7B1D | No | Customer replaced SP twice per unit troubleshooting on his own. Customer called for service, I had him flush system then replace SC, same issue persists. Getting 1084 inhomogeneous cal 3 soluton erros. | Installed PM kit for diagnostic purposes. | Assisted Replacement | SC | Wetability
ABL90 | Found the the system was in a loop requesting to have the sensor cassett replaced. | Replaced the sensor cassett, cleaned and readjusted the inlet module as well. | CF6417AE02440314 | No | Found the the system was in a loop requesting to have the sensor cassett replaced. | Replaced the sensor cassett, cleaned and readjusted the inlet module as well. | Assisted Replacement | SC | Wetability
ABL90 FLEX | Einlassdichtung lt. Fehlervorschlag ersetzt, Programme zum Test der Wet Section durchgeführt, vorher Probenweg mit Spülspritze gespült, weiterhin diverse Pumpenkalibrier und Leckagefehler, Wechsel des Lösungspacks, da nur noch 125 Zyklen enthalten, Kal, QC i.O. | Einlassdichtung lt. Fehlervorschlag ersetzt, Programme zum Test der Wet Section durchgeführt, vorher Probenweg mit Spülspritze gespült, weiterhin diverse Pumpenkalibrier und Leckagefehler, Wechsel des Lösungspacks, da nur noch 125 Zyklen enthalten, Kal, QC i.O. | SWO-47532 | No | Replacement of inlet seal according to fault proposal, programs for testing the wet section performed, previously sample path flushed with flushing syringe, further various pump calibration and leakage errors, change of solution pack, since only 125 cycles are still included, Kal, QC i.O. | Replacement of inlet seal according to fault proposal, programs for testing the wet section performed, previously sample path flushed with flushing syringe, further various pump calibration and leakage errors, change of solution pack, since only 125 cycles are still included, Kal, QC i.O. | Assisted Replacement | SP | leak detected
ABL90 FLEX | CF Karte erneuert, Win, SW 3.1MR3 installiert, KfH Konfiguration geladen, IP Adressen eingegeben. | Kundin bittet um Rückruf und Service | SWO-47747 | No | CF card renewed, Win, SW 3.1MR3 installed, KfH configuration loaded, IP addresses entered. | Customer asks for recall and service | Defect | HW | CF card
ABL90 FLEX | QC -Ergebnisse geprüft, im Serviceprogramm pO2-Kanalprüfung durchgeführt --> pO2Dark > 200 --> Ersatzteile bestellt Station und Labor wurde informiert | QC -Ergebnisse geprüft, im Serviceprogramm pO2-Kanalprüfung durchgeführt --> pO2Dark > 200 --> Ersatzteile bestellt Station und Labor wurde informiert | SWO-47490 | No | QC results checked, performed in the service program pO2 channel testing --> pO2Dark > 200 --> Spare parts ordered station and laboratory was informed | QC results checked, performed in the service program pO2 channel testing --> pO2Dark > 200 --> Spare parts ordered station and laboratory was informed | Unknown | Revisit required | parts ordered
ABL80 CO-OX | after replacing door on sp found that the analyzer would not come in for sc or sp keeps stating that air detection failed | replaced the sp door and analog board to correct all the issue with analyzer. | CF640E8402156BEA | No | after replacing door on sp found that the analyzer would not come in for sc or sp keeps stating that air detection failed | replaced the sp door and analog board to correct all the issue with analyzer.
TCM400 | 更换945-605一个正常。 | 更换945-605一个正常。 | SWO-47203 | No | Replace the 945-605 one normal. | Replace the 945-605 one normal.
AQT90 FLEX |  |  | SWO-47401 | No |  |  | Unknown | LEH Mail |  | =AU4442
ABL825 FLEX | Tápegység csere után a tesztek megfelelőek. Kalibrálás, QC mérés. |  | SWO-47746 | No | After replacing the power supply, the tests are appropriate. Calibration, QC measurement. |  | Defect | HW | Power Supply
ABL825 FLEX | checked all data logs, inpected fluid transport system, checked measuring chambers. Checked service programs for liquid sensors, pump cals, oximetry system, measuring and fluidic module valves. | Adjusted lamp unit in the service program, replaced ICA electrode old lot #6707 with new lot # 9704 and Reference electrode old lot #1712 with new lot# 9703. Will send both old electrodes in for warranty replacement. Tracking number for electrodes:806457122074. | CF63E31B020A70C2 | No | checked all data logs, inpected fluid transport system, checked measuring chambers. Checked service programs for liquid sensors, pump cals, oximetry system, measuring and fluidic module valves. | Adjusted lamp unit in the service program, replaced ICA electrode old lot #6707 with new lot # 9704 and Reference electrode old lot #1712 with new lot# 9703. Will send both old electrodes in for warranty replacement. Tracking number for electrodes:806457122074. | Assisted Replacement | Electrode | Electrode | Ca
ABL800 FLEX | Gerät gereinigt und überprüft, Proteinentfernung durchgeführt, Lufteinlass gespült, Dichtungsring der O2 Elektrode erneuert, Ref neu membraniert, Kals und QCs | Gerät gereinigt und überprüft, Proteinentfernung durchgeführt, Lufteinlass gespült, Dichtungsring der O2 Elektrode erneuert, Ref neu membraniert, Kals und QCs | SWO-47554 | No | Device cleaned and checked, protein removal performed, air inlet flushed, sealing ring of The O2 electrode renewed, ref re-membraneed, kals and QCs | Device cleaned and checked, protein removal performed, air inlet flushed, sealing ring of The O2 electrode renewed, ref re-membraneed, kals and QCs | Cleaned | Fluidic pathway
ABL835 FLEX | N A | INSTALLED LIS HIS KEYS | CF6428C7025A3D98 | Yes | N A | INSTALLED LIS HIS KEYS | Re classified RMED | Setup
ABL825 FLEX |  |  | CF6575DF7815BCC9 | No |  |  | Assisted Replacement | Electrode | Electrode | Reference
ABL80 FLEX BASIC | Luftblasen Erkennung abgeschaltet. Kein Lösungspack mehr vorrätig. Reklamationsformular ausgefüllt. | Luftblasen Erkennung abgeschaltet. Kein Lösungspack mehr vorrätig. Reklamationsformular ausgefüllt. | SWO-47641 | No | Air bubble detection switched off. No more solution packs in stock. Complaint form completed. | Air bubble detection switched off. No more solution packs in stock. Complaint form completed.
ABL80 FLEX BASIC | 更换测试卡正常。 | 更换测试卡正常。 | SWO-47609 | No | The replacement test card is OK. | The replacement test card is OK.
ABL80 FLEX BASIC | 更换测试卡正常。 | 更换测试卡正常。 | SWO-47599 | No | The replacement test card is OK. | The replacement test card is OK.
ABL80 FLEX BASIC |  |  | SWO-47565 | No
ABL80 FLEX BASIC | 更换903-334一个正常。 | 更换903-334一个正常。 | SWO-47650 | No | Replace 903-334 one normal. | Replace 903-334 one normal.
ABL90 FLEX | Jahreswartung durchgeführt. | Jahreswartung durchgeführt. | SWO-47662 | No | annual maintenance. | annual maintenance. | Unknown | frequent SC replacement
ABL90 | Entered service mode and performed all the tests in the fluid transport section. All results were within manufactures specifications. | Performed a back flush, and this resolved the problem. | CF65A96478416530 | No | Entered service mode and performed all the tests in the fluid transport section. All results were within manufactures specifications. | Performed a back flush, and this resolved the problem. | Cleaned | Fluidic pathway
ABL90 FLEX | Lieferung 2x LP LP installiert Startphase abgewartet | Lieferung 2x LP LP installiert Startphase abgewartet | SWO-47927 | Yes | Delivery 2x LP LP installed start phase awaited | Delivery 2x LP LP installed start phase awaited | Re classified RMED | Logistical upSWO
ABL90 FLEX | Sensor interface PCB erneuert, pO2-Kanal im Serviceprogramm getestet--> i.O Kal und QC durchgeführt | Sensor interface PCB erneuert, pO2-Kanal im Serviceprogramm getestet--> i.O Kal und QC durchgeführt | SWO-47674 | No | Sensor interface PCB renewed, pO2 channel tested in service program--> i.O Kal and QC performed | Sensor interface PCB renewed, pO2 channel tested in service program--> i.O Kal and QC performed | Defect | HW | SC PCB | dark current
ABL90 | Checked data logs. | Replaced sensor cassette and solution pack. Verified operation of ABL90. | CF6E5CDCB2004D7C | No | Checked data logs. | Replaced sensor cassette and solution pack. Verified operation of ABL90. | Defect | SW | Error 1343 /Wetability
ABL90 | Checked data logs. | Replaced sensor cassette and solution pack. Verified operation of ABL90. | CF6E561FB2454D77 | No | Checked data logs. | Replaced sensor cassette and solution pack. Verified operation of ABL90. | Defect | SW | Error 1343 /Wetability
ABL825 FLEX |  |  | CF6540087809A9C8 | No |  |  | Assisted Replacement | Electrode | Electrode | pCO2
ABL800 BASIC | オキシモジュール交換 持ち帰りチェック |  | SWO-47878 | No | Oxy Module Replacement Take-away Check |  | Defect | HW | Oxi module | Spectrum mismatch
ABL825 FLEX | Halogenlamp Intensity too high Kamp Justage.oxi compl. Test Kal und QC okay | Halogenlamp Intensity too high Kamp Justage.oxi compl. Test Kal und QC okay | SWO-48087 | No | Halogenlamp Intensity too high Kamp Justage.oxi compl. Test Kal und QC okay | Halogenlamp Intensity too high Kamp Justage.oxi compl. Test Kal und QC okay | Adjustment | HW | Oxi module | Lamp unit
ABL825 FLEX | Gerät gereinigt und überprüft, Proteinentfernung durchgeführt, pH-Elektrode überprüft, pH-Elektrode erneuert, Kals und QCs durchgeführt. | Ref. Membran wurde erneuert, Messkammern und Kontakte sind sauber. Wahrscheinlich die pH Elektrode defekt. Service Einsatz notwendig. Der Kunde besteht auf einen Service am selben Tag. Bitte Kontakt aufnehmen und abklären. | SWO-47851 | No | Device cleaned and checked, protein removal performed, pH electrode checked, pH electrode renewed, kals and QCs performed. | Ref. Membrane has been renewed, measuring chambers and contacts are clean. Probably the pH electrode defective. Service use necessary. The customer insists on a service on the same day. Please contact us and clarify. | Assisted Replacement | Electrode | Electrode | pH
ABL835 FLEX | No errors present but some liquid has been dripping from inlet. | Flushed waste lines. Inspected inlet gasket and recalibrated inlet in service program. Replaced waste pump tubes after getting a small bit of leakage after several rinses. No help. Replaced inlet gasket. | CF65BBF9781A322D | No | No errors present but some liquid has been dripping from inlet. | Flushed waste lines. Inspected inlet gasket and recalibrated inlet in service program. Replaced waste pump tubes after getting a small bit of leakage after several rinses. No help. Replaced inlet gasket. | Assisted Replacement | HW | Inlet module | Inlet Gasket
ABL80 FLEX BASIC | 更换两张新测试卡 | 更换两张新测试卡 | SWO-47852 | No | Replace two new test cards | Replace two new test cards
ABL80 FLEX BASIC | 更换新测试卡 | 更换新测试卡 | SWO-47898 | No | Replace a new test card | Replace a new test card
ABL90 FLEX | 更换新测试卡 | 更换新测试卡 | SWO-47856 | No | Replace a new test card | Replace a new test card | Assisted Replacement | SC | pO2
ABL825 FLEX |  |  | SWO-47837 | No |  |  | Defect | HW | Inlet module | Friction detector unit
ABL825 FLEX | Einlassnadel erneuert, Einlass Kal. i.O. |  | SWO-47972 | No | Inlet needle renewed, inlet Kal. Ok. |  | Defect | HW | Inlet module | Probe
ABL80 FLEX BASIC | 更换新测试卡 | 更换新测试卡 | SWO-47853 | No | Replace a new test card | Replace a new test card
ABL80 FLEX BASIC | 更换新测试卡后正常 | 更换新测试卡后正常 | SWO-49338 | No | Normal after replacing a new test card | Normal after replacing a new test card
ABL825 FLEX | AQC Modul wieder eingebaut Ersatzmodul wieder zurück | Kundin bittet um einen Einsatz des zuständigen Technikers, bitte um Kontaktaufnahme. | SWO-47870 | No | AQC module reinstalled replacement module back | Customer asks for the use of the responsible technician, please contact us. | Defect | HW | AutoCheck Module | PCB
ABL80 FLEX CO-OX | Fehlersuche, Schlauchwege überprüft, Anschlussstutzen zum LP gereinigt, SC und LP wurden bereits vom Anwender erneuert, LP erneuert->LQC und Kal. mehrmals wiederholt ok System-und Funktionskontrolle i.O. | Fehlersuche, Schlauchwege überprüft, Anschlussstutzen zum LP gereinigt, SC und LP wurden bereits vom Anwender erneuert, LP erneuert->LQC und Kal. mehrmals wiederholt ok System-und Funktionskontrolle i.O. | SWO-48021 | No | Troubleshooting, hose paths checked, connection plugs to the LP cleaned, SC and LP have already been renewed by the user, LP renewed->LQC and Kal. repeated several times ok system and function control i.O. | Troubleshooting, hose paths checked, connection plugs to the LP cleaned, SC and LP have already been renewed by the user, LP renewed->LQC and Kal. repeated several times ok system and function control i.O.
ABL80 FLEX BASIC | 更换新测试卡 | 更换新测试卡 | SWO-47985 | No | Replace a new test card | Replace a new test card
ABL80 FLEX BASIC | 更换测试卡正常。 | 更换测试卡正常。 | SWO-48115 | No | The replacement test card is OK. | The replacement test card is OK.
ABL80 FLEX BASIC | 更换测试卡正常。 | 更换测试卡正常。 | SWO-48114 | No | The replacement test card is OK. | The replacement test card is OK.
ABL80 FLEX CO-OX | Themp.fehler im Oximetriemodul beseitigt, neuen Hemolysatoreinheit eingebaut, Gerät gereinigt und neues Lösungspack eingebaut, Gerät kalibriert, div. Schläuche erneuert, | Themp.fehler im Oximetriemodul beseitigt, neuen Hemolysatoreinheit eingebaut, Gerät gereinigt und neues Lösungspack eingebaut, Gerät kalibriert, div. Schläuche erneuert, | SWO-48018 | No | Themp.error in the oximetry module eliminated, new hemolyser unit installed, device cleaned and new solution pack installed, device calibrated, various hoses renewed, | Themp.error in the oximetry module eliminated, new hemolyser unit installed, device cleaned and new solution pack installed, device calibrated, various hoses renewed,
ABL80 FLEX BASIC | GIVEN TRAINING ON THE ANALYZER | GIVEN TRAINING ON THE ANALYZER | SWO-49158 | Yes | GIVEN TRAINING ON THE ANALYZER | GIVEN TRAINING ON THE ANALYZER | Re classified RMED | Training only
ABL90 FLEX | SC erneuert im Anschluss die Jahreswartung vorgezogen. | Herr Ch. hat seit zwei Wochen Schwierigkeiten mit seinem ABL. SC & SP wurden mehrfach gewechselt. das Gerät wurde manuell gespült. Da es so nicht weiter gehen kann wünscht Herr Ch. eine Überprüfung durch einen Radiometer Techniker. | SWO-48238 | No | SC then renews the annual maintenance preferred. | Mr. Ch. has been in trouble with his ABL for two weeks. SC & SP have been changed several times. the device was rinsed manually. Since it cannot go on like this, Mr. Ch. wants a check by a radiometer technician. | Assisted Replacement | SC | pCO2
ABL90 FLEX | Warmstart durchgeführt, tHb-Kal o.k. | Warmstart durchgeführt, tHb-Kal o.k. | SWO-48095 | No | warm start, tHb-Kal ok. | warm start, tHb-Kal ok. | No fault found | Internvention required
TCM CombiM |  |  | SWO-47913 | No
AQT90 FLEX |  |  | SWO-47903 | No |  |  | Defect | HW | Process module | Optical Unit | 135
AQT90 FLEX |  |  | SWO-47880 | No |  |  | Defect | HW | Process module | Optical Unit | 131
AQT90 FLEX |  |  | SWO-47951 | No |  |  | Defect | HW | Analyzer control Module | Main Board, PCB
ABL800 BASIC | Referenzelektrode am Elektrodenstift blank, Garantieteparatur. Elektrode ausgetauscht. Kalibrationen durchgeführt. QC ohne Beanstandung | Referenzelektrode am Elektrodenstift blank; Garantieteparatur. Elektrode ausgetauscht. Kalibrationen durchgeführt. QC ohne Beanstandung | SWO-48052 | No | Reference electrode on the electrode pin bare; Guaranteed parsing. electrode. calibrations. QC without complaint | Reference electrode on the electrode pin bare; Guaranteed parsing. electrode. calibrations. QC without complaint | Assisted Replacement | Electrode | Electrode | Reference
ABL800 BASIC | 更换902-833一个 | 更换902-833一个 | SWO-47995 | No | Replace 902-833 one | Replace 902-833 one | Defect | HW | Oxi module
ABL820 FLEX | Verified errors in QC logs. Intermittent for the last month. Performed complete oxi test, all passed. | Replaced hemolyzer & performed all associated adjustments. | CF692418DA281F71 | No | Verified errors in QC logs. Intermittent for the last month. Performed complete oxi test, all passed. | Replaced hemolyzer & performed all associated adjustments. | Defect | HW | Oxi module | Hemolyzer
ABL825 FLEX | Lamp Unit Justage durchgeführt. Oxi compl.Test Kal und QC okay | Lamp Unit Justage durchgeführt. Oxi compl.Test Kal und QC okay | SWO-48086 | No | Lamp Unit Adjustment. Oxi compl. Test Kal and QC okay | Lamp Unit Adjustment. Oxi compl. Test Kal and QC okay | Adjustment | HW | Oxi module | Lamp unit
ABL825 FLEX | QC Modul nach Rep.zurück. Ersatzmodul wieder ausgebaut. | QC Modul nach Rep.zurück. Ersatzmodul wieder ausgebaut. | SWO-48083 | No | QC module to Rep.back. Replacement module. | QC module to Rep.back. Replacement module. | Defect | HW | AutoCheck Module | PCB
ABL827 FLEX | Lampenspannung abgeglichen, diverse Lösungen und Membranen ern., Messkammer gereinigt, Proteinentfernung durchgef., Gerät kalibriert, QC´s durchgef. | Kundin bittet um einen Service. | SWO-48038 | No | Lamp tension matched, various solutions and membranes ern., measuring chamber cleaned, protein removal carried out, device calibrated, QC's through. | Customer asks for a service. | Adjustment | HW | Oxi module | Lamp unit
ABL825 FLEX | cals and Qc's within range | ran precision studies with Qualicheck 5+ 10 samples each on level 1,2,3. Precision studies within range. Ran VK-R7 kit, values within range. | CF6922EBDA55B947 | No | cals and Qc's within range | ran precision studies with Qualicheck 5+ 10 samples each on level 1,2,3. Precision studies within range. Ran VK-R7 kit, values within range. | Unknown | Performing within R spec
ABL825 FLEX | cals and QC's within range | ran precision studies with Qualicheck 5+ 10 samples each on level 1,2,3. Precision studies within range. Ran VK-R7 kit, values within range. | CF6925E0DA02B94B | No | cals and QC's within range | ran precision studies with Qualicheck 5+ 10 samples each on level 1,2,3. Precision studies within range. Ran VK-R7 kit, values within range. | Unknown | Performing within R spec
ABL825 FLEX | checked cals and QC's | ran precision studies with Qualicheck 5+ 10 samples each on level 1,2,3. Precision studies within range. Ran VK-R7 kit, values within range. | CF692862DA35B94E | No | checked cals and QC's | ran precision studies with Qualicheck 5+ 10 samples each on level 1,2,3. Precision studies within range. Ran VK-R7 kit, values within range. | Unknown | Performing within R spec
ABL820 FLEX |  |  | CF694B0ADA1E0FCD | No |  |  | Defect | HW | Gas module | Pressure Transducer
AQT90 FLEX | Actualización del software a la v. 8.8.15, mantenimiento preventivo, aspirado interior del equipo, ajuste del scanner del reactivo, limpieza de la unidad de lavado, limpieza de la inncubadora, engrase de ejes de motores, comprobación de registros de datos, comprobación de la unidad óptica, comprobación de la process wheel, comprobación del funcionamiento con los programas de servicio, reinicio del equipo. |  | CF690A92DA1F7FA5 | No | Software update to v. 8.8.15, preventive maintenance, internal suction of the equipment, adjustment of the reagent scanner, cleaning of the washing unit, cleaning of the incubator, greasing motor shafts, checking data records, Optical drive check, process wheel check, operation check with service programs, computer restart. |  | Cleaned | HW | Storage module | Barcode reader, Storage wheel
ABL80 FLEX BASIC | 收到的备件是903-381。需要再借一个CPU | 更换903-381一个正常。 | SWO-47966 | No | The spare parts received are 903-381. Need to borrow another CPU | Replace 903-381 one normal.
ABL80 FLEX BASIC |  |  | SWO-48109 | No
TCM CombiM | 更换945-737修复机器 | 更换945-737修复机器 | SWO-47968 | No | Replace the 945-737 repair machine | Replace the 945-737 repair machine
TCM CombiM | 模块编号报错，重新借903-111 | 模块编号报错，重新借903-111 | SWO-48185 | No | Module number error, re-borrow edf 903-111 | Module number error, re-borrow edf 903-111
ABL90 FLEX | 更换新试剂包 | 更换新试剂包 | SWO-48144 | No | Replace a new kit | Replace a new kit | Assisted Replacement | SP | FT | Vaccum
ABL80 CO-OX |  |  | CF6D07503C5C26F2 | Yes |  |  | Re classified RMED | Training only
ABL825 FLEX | Found 3 oxi parameters outside of qc range. | Performed multiple protein remover programs and calibrations. Performed service and regular THB calibration. | CF6A67D9501BECDB | No | Found 3 oxi parameters outside of qc range. | Performed multiple protein remover programs and calibrations. Performed service and regular THB calibration. | Cleaned | Fluidic pathway | Oxi module
ABL827 FLEX |  |  | CF6AAF25501A545F | No |  |  | Assisted Replacement | Electrode | Electrode | Glu
ABL815 FLEX | Pumpen Schläuche und Dichtungen erneuert, LS Justage und Pumpen Kal im Service Programm durchgeführt, Kal mehrfach durchgeführt QC i.O. | Pumpen Schläuche und Dichtungen erneuert, LS Justage und Pumpen Kal im Service Programm durchgeführt, Kal mehrfach durchgeführt QC i.O. | SWO-48235 | No | Pumps hoses and seals renewed, LS adjustment and pumps Kal performed in the service program, Kal performed several times QC i.O. | Pumps hoses and seals renewed, LS adjustment and pumps Kal performed in the service program, Kal performed several times QC i.O. | Defect | HW | El/Met Module | Pump tubing
AQT90 FLEX | CALIBRATE THE NTPROBNP | CALIBRATE THE NTPROBNP | SWO-49157 | No | CALIBRATE THE NTPROBNP | CALIBRATE THE NTPROBNP | User Error | Re calibration | NT-proBNP | 0
ABL80 FLEX BASIC | Alles wieder in Ordnung | Alles wieder in Ordnung | SWO-49051 | No | All right again | All right again
ABL80 FLEX BASIC | 更换新测试卡后正常 | 更换新测试卡后正常 | SWO-49348 | No | Normal after replacing a new test card | Normal after replacing a new test card
ABL80 FLEX BASIC | 更换903-381一个正常。 | 更换903-381一个正常。 | SWO-48110 | No | Replace 903-381 one normal. | Replace 903-381 one normal.
ABL80 FLEX BASIC | 更换测试卡正常。 | 更换测试卡正常。 | SWO-49003 | No | The replacement test card is OK. | The replacement test card is OK.
ABL80 FLEX BASIC | 更换新测试卡 | 更换新测试卡 | SWO-49169 | No | Replace a new test card | Replace a new test card
ABL80 FLEX CO-OX | Anschl.schlauch zwischen manifold und Lueranschluss hat sich gelöst, | Anschl.schlauch zwischen manifold und Lueranschluss hat sich gelöst, | SWO-49033 | No | the hose between the manifold and the Luer connection has dissolved, | the hose between the manifold and the Luer connection has dissolved,
ABL80 CO-OX | Confirmed liquid sensor failed. Leakage on the liquid sensor. | Replaced the liquid sensor. | CF6BEE29C63F3B70 | No | Confirmed liquid sensor failed. Leakage on the liquid sensor. | Replaced the liquid sensor.
ABL80 FLEX BASIC | given training for users and explained about Pre-Analytical Errors. and suggested the samplers. | given training for users and explained about Pre-Analytical Errors. and suggested the samplers. | SWO-49576 | No | given training for users and explained about Pre-Analytical Errors. and suggested the samplers. | given training for users and explained about Pre-Analytical Errors. and suggested the samplers.
ABL90 FLEX | OXI compartment, bottom part und die fehlenden Schrauben ersetzt. Keine Ersatznadel vor Ort, Nadel erneuert und eine Ersatznadel vor Ort gelassen. Testmessungen mit Blut durchgeführt. |  | SWO-49275 | No | OXI compartment, bottom part and the missing screws replaced. No replacement needle on site, needle renewed and a replacement needle left on site. Test measurements carried out with blood. |  | Defect | HW | Oxi retainor
ABL90 FLEX | Hemolyzer 902-980R487N009 ersetzt durch 902-980R0481N004. Defektes OXI Gehäuse festgestellt, Austausch notwendig, Ersatzteile bestellt. |  | SWO-49272 | Yes | Hemolyzer 902-980R487N009 replaced by 902-980R0481N004. Defective OXI housing detected, replacement necessary, spare parts ordered. |  | Re classified RMED | Double entry | SWO-49275 | CAS-6123-B8P3
ABL90 FLEX | Rücktausch Display Unit | Austausch Display Unit Funktionstest i.O. | SWO-49048 | No | Rücktausch Display Unit | Replacement Display Unit Functional Test i.O. | Adjustment | touch screen
ABL90 FLEX | Sensorkassette erneuert, Kalibrationen und Qualitätskontrollen ohne Beanstandungen durchgeführt. | Sensorkassette erneuert, Kalibrationen und Qualitätskontrollen ohne Beanstandungen durchgeführt. | SWO-48989 | No | Sensor cartridge renewed, calibrations and quality controls carried out without complaints. | Sensor cartridge renewed, calibrations and quality controls carried out without complaints. | Cleaned | Fluidic pathway | replaced SC
ABL90 FLEX | Gerät signalisiert weit der Auslieferung immer wieder QC Level 2 Probleme bei der Oximetrie. Gerät mehrfach mit warmen Wasser durchgespült, Schlauchwege und Dichtungen aufgrund inhomogener Flüssigkeit erneuert. Anschließend QC erfolgreich durchgeführt. | Gerät signalisiert weit der Auslieferung immer wieder QC Level 2 Probleme bei der Oximetrie. Gerät mehrfach mit warmen Wasser durchgespült, Schlauchwege und Dichtungen aufgrund inhomogener Flüssigkeit erneuert. Anschließend QC erfolgreich durchgeführt. | SWO-49303 | No | Device signals far the delivery repeatedly QC Level 2 problems with oximetry. Device repeatedly rinsed with warm water, hoseways and seals renewed due to inhomogeneous liquid. Subsequently QC successfully performed. | Device signals far the delivery repeatedly QC Level 2 problems with oximetry. Device repeatedly rinsed with warm water, hoseways and seals renewed due to inhomogeneous liquid. Subsequently QC successfully performed. | Cleaned | Fluidic pathway
ABL90 FLEX | Remplacé le kit de maintenance + aiguille e.o Calibration e.o contrôle de qualité e.o | Remplacé le kit de maintenance + aiguille e.o Calibration e.o contrôle de qualité e.o | SWO-49008 | No | Replaced the maintenance kit - needle e.o Calibration e.o quality control e.o | Replaced the maintenance kit - needle e.o Calibration e.o quality control e.o | Defect | HW | waste tube
TCM-COMBIM | Desmontar equipo, desconectar conexiones bateria, comprobar modulo TC. Referenciar piezas para su peticion Anular pO2 en otro monitor a peticion de cliente. Limpiar equipo |  | CF6BA9E5C61D25D0 | No | Disassemble equipment, disconnect battery connections, check TC module. Reference parts for your Override pO2 request on another monitor at customer request. Clean equipment
TCM CombiM | 借903-111一个。 | 模块编号报错，重新借903-111 | SWO-48287 | No | Borrow 903-111 one. | Module number error, re-borrow edf 903-111
ABL90 | Commprobaciónn de los registros, se sustituye Inlet e Inlet gasket, se procede a realizar limpieza del equipo, cliente informa que lo realizo y solo salían gotas, se procede a realizar la limpieza, se busca donnde esta la obstrucción, se determina que esta entra módulo de entrada y primer nipple casete de electrodos, se intenta desobstruir esa zona con agua y hipoclorito, no se consigue nada por lo que se pide material. | no se puede desobstruir el equipo. equipo en intervención requerida. | CF6CC0973C3BD433 | No | Commbation of the records, Inlet and Inlet gasket is replaced, cleaning the equipment is carried out, customer informs that I performed it and only drops were coming out, the cleaning is carried out, the blockage is sought, it is determined that this entermodule of entrance and first nipple electrode cassette, trying to unblock that area with water and hypochlorite, you get nothing so you ask for material. | you can't unblock the equipment. intervention team required. | Unknown | Revisit required | analyzer blocked | clot?
AQT90 FLEX | 退换新机 | 更换AU板，重装XP和软件后可以做样本，但仍不稳定 | SWO-48299 | No | Return ing-up to a new machine | Replace the AU board, reinstall XP and software can do samples, but still unstable | Defect | SW | Analyzer control Module | Reinstall XP
ABL825 FLEX | Waste-Detector erneuert. | Herr Zimmermann bemängelt das bei seinem Überbrückungsgerät (bis zum Neukauf) Glu und Lac nur maximal zwei Wochen halten und oft hohe Driften zu beobachten sind. Die Kalibrationen brechen immer wieder ab. Herr Z. bittet um eine Überprüfung des Analysators. | SWO-48963 | No | Waste Detector renewed. | Mr. Zimmermann criticizes this with his bridging device (until the new purchase) Glu and Lac only hold for a maximum of two weeks and often high drifts can be observed. The calibrations stop again and again. Mr Z. asks for a review of the analyzer. | Defect | HW | Basic Unit | Waste detector
ABL825 FLEX | QC Plan geändert | QC Plan geändert | SWO-49129 | No | QC plan changed | QC plan changed | Setup | QC Schedule
ABL825 FLEX | QC Plan eingegeben und konfiguriert | QC Plan eingegeben und konfiguriert | SWO-49131 | No | QC Plan entered and configured | QC Plan entered and configured | Setup | QC Schedule
ABL825 FLEX | Aqureanbindung durchgeführt, | Aqureanbindung durchgeführt, | SWO-49128 | Yes | Aqure connection carried out, | Aqure connection carried out, | Re classified RMED | Setup
ABL825 FLEX | QC Plan installiert | QC Plan installiert | SWO-49130 | No | QC Plan installed | QC Plan installed | Setup | QC Schedule
ABL825 FLEX | alternierenden QC Plan konfigugiert | alternierenden QC Plan konfigugiert | SWO-49132 | No | alternierenden QC Plan konfigugiert | alternierenden QC Plan konfigugiert | Setup | QC Schedule
ABL825 FLEX | Lactatelektrode erneuert. | Lactatelektrode erneuert. | SWO-49114 | No | Lactate electrode renewed. | Lactate electrode renewed. | Assisted Replacement | Electrode | Electrode | Lac
ABL80 FLEX BASIC |  |  | SWO-49138 | No
ABL80 FLEX BASIC | 更换测试卡正常。 | 更换测试卡正常。 | SWO-49150 | No | The replacement test card is OK. | The replacement test card is OK.
ABL80 FLEX BASIC | 排液ドレインマウント洗浄、クリーニングSP,SC使用し、全流路洗浄。 | 排液ドレインマウント洗浄、クリーニングSP,SC使用し、全流路洗浄。 | SWO-50556 | Yes | Drain drain drain mount cleaning, cleaning SP, SC, full flow path cleaning. | Drain drain drain mount cleaning, cleaning SP, SC, full flow path cleaning. | Re classified RMED | Dublicate in CRM | same as SWO-5558
ABL80 FLEX BASIC | 排液ドレインマウント洗浄、クリーニングSP,SCを使用し全流路洗浄 | 排液ドレインマウント洗浄、クリーニングSP,SC使用し、全流路洗浄。 | SWO-50558 | No | Drain drain mount cleaning, cleaning SP, sc using full flow path cleaning | Drain drain drain mount cleaning, cleaning SP, SC, full flow path cleaning.
ABL80 FLEX BASIC |  |  | SWO-49168 | No
ABL90 | Internal Battery is not charging. | Replaced battery Brought system to ready. | CF6D0D693C42308C | No | Internal Battery is not charging. | Replaced battery Brought system to ready. | Defect | HW | Battery module
ABL90 FLEX | Austausch Display Unit Funktionstest i.O. | Austausch Display Unit Funktionstest i.O. | SWO-49267 | No | Replacement Display Unit Functional Test i.O. | Replacement Display Unit Functional Test i.O. | Defect | HW | Monitor
ABL5 | pH Messungen der Radiometer und der Fremdlösungen durchgeführt. Soll Werte beider Lösungen: Kal 1: 7,38 und Kal 2: 6,84. -R- Kal 1: 7,37, -R- Kal 2: 6,83. Phoenix Kal 1: 7,33, Phoenix Kal 2: 6,79! Fehler ist auf die Reagenzien zurückzuführen! | Kundin bittet um einen Einsatz des zuständigen Technikers, bitte um Kontaktaufnahme. | SWO-49378 | No | pH measurements of radiometers and external solutions. Should be values of both solutions: Kal 1: 7.38 and Kal 2: 6.84. -R- Kal 1: 7.37, -R- Kal 2: 6.83. Phoenix Kal 1: 7.33, Phoenix Kal 2: 6.79! Errors are due to the reagents! | Customer asks for the use of the responsible technician, please contact us.
TCM CombiM |  |  | SWO-49201 | No
TCM CombiM |  |  | SWO-49236 | No
ABL90 | CHEQUEO DE REACTIVOS, LIMPIEZA SEGÚN PROCEDIMIENTO, SUSTITUCIÓN DE INLET GASKET Y SONDA, REINSTALACIÓN SP Y REINICIO COMPLETO,CALIBRACIÓN Y PASADOS CICLOS DE CALIDAD | SUSTITUCIÓN DE SONDA, GASKET Y SP | CF6CEC813C47DA79 | No | REACTIVE CHECK, CLEANING AS PROCEDURE, SUBSTITUTE OF INLET GASKET AND SONDA, REINSTALLATION SP AND COMPLETE RESET, CALIBRATION AND PAST QUALITY CYCLES | SUBSTITUTE OF SOUND, GASKET AND SP | Cleaned | Fluidic pathway
ABL90 FLEX | 更换测试卡正常。 | 更换测试卡正常。 | SWO-49196 | No | The replacement test card is OK. | The replacement test card is OK. | Assisted Replacement | SC | Reference
AQT90 FLEX |  |  | SWO-49152 | No |  |  | Defect | HW | Storage module | Casette extracting mechanism | 5979
AQT90 FLEX | 重新取出测试卡，恢复安全位置 | 重新取出测试卡，恢复安全位置 | SWO-49327 | No | Re-remove the test card and restore the safe position | Re-remove the test card and restore the safe position | Defect | Consumables | Test Kit | Cartridge | 5784 (Broken)
ABL800 BASIC | Gerät dekontaminiert - Messkammer gereinigt Membran von Gasventil in pH BG Messkammer sowie pH Dichtring erneuert pO2 Nullpunktkalibrierung durchgeführt Gerät kalibriert und überprüft | Gerät dekontaminiert - Messkammer gereinigt Membran von Gasventil in pH BG Messkammer sowie pH Dichtring erneuert pO2 Nullpunktkalibrierung durchgeführt Gerät kalibriert und überprüft | SWO-49183 | No | Device decontaminated - measuring chamber cleaned membrane of gas valve in pH BG measuring chamber as well as pH sealing ring renewed pO2 zero point calibration device calibrated and checked | Device decontaminated - measuring chamber cleaned membrane of gas valve in pH BG measuring chamber as well as pH sealing ring renewed pO2 zero point calibration device calibrated and checked | Defect | Electrode | Electrode | Sealing Ring
ABL825 FLEX | アンプルアクチュエーター交換 |  | SWO-49135 | No | Ampoule actuator replacement |  | Defect | HW | AutoCheck Module | Breaker arm
ABL825 FLEX | Alle Ventile überprüft Neu membraniert Pumpenschläuche Module ausgetauscht Barometer überprüft Funktionskontrolle i.o QC Grafik i.o Funktionskontrolle i.o Analysator wieder zurückgebracht zum Kunden | Alle Ventile überprüft Neu membraniert Pumpenschläuche Module ausgetauscht Barometer überprüft Funktionskontrolle i.o QC Grafik i.o Funktionskontrolle i.o Analysator wieder zurückgebracht zum Kunden | SWO-49883 | No | All valves checked New membranated pump hoses modules replaced barometer checked function control i.o QC Graphic i.o Function control i.o Analyzer returned to the customer | All valves checked New membranated pump hoses modules replaced barometer checked function control i.o QC Graphic i.o Function control i.o Analyzer returned to the customer | Assisted Replacement | Electrode | Membrane | pCO2
ABL825 FLEX | PCO2 is out on QC level 2 low. Repeated multiple times with same results. PCO2 has been remembraned multiple times over the last month, the most recent was on 8-17-15. | Replaced PCO2 electrode with lot SK03. Old lot # RM02 will be reutrned under warranty fedex tracking number 806457122063. Left box of PCO2 membranes with 3 remaining for customer. | CF6CCE823C0592F5 | No | PCO2 is out on QC level 2 low. Repeated multiple times with same results. PCO2 has been remembraned multiple times over the last month, the most recent was on 8-17-15. | Replaced PCO2 electrode with lot SK03. Old lot # RM02 will be reutrned under warranty fedex tracking number 806457122063. Left box of PCO2 membranes with 3 remaining for customer. | Assisted Replacement | Electrode | Electrode | pCO2
ABL825 FLEX | Checked, cleaned and remembraned reference electrode. Replaced sealing ring on dummy electrode for Glucose and replaced sealing ring for pH electrode. | Replaced sealing rings for dummy electrode glucose and pH electrode. Verified instrument operation. | CF6D26013C2119EE | No | Checked, cleaned and remembraned reference electrode. Replaced sealing ring on dummy electrode for Glucose and replaced sealing ring for pH electrode. | Replaced sealing rings for dummy electrode glucose and pH electrode. Verified instrument operation. | Defect | Electrode | Electrode | Sealing Ring
ABL825 FLEX | Found 14 pin cable on autocheck module was worn and needed to be replaced. | Replaced 14 pin cable on autocheck module. Verified instrument operation. | CF6D22FA3C5919EA | No | Found 14 pin cable on autocheck module was worn and needed to be replaced. | Replaced 14 pin cable on autocheck module. Verified instrument operation. | Defect | HW | AutoCheck Module | FPC Cable
ABL825 FLEX | Fehlersuche, Kal.- QC- und Aktivitätendateien überprüft, QC Verlauf zeigt, dass pCO2 tendenziell zu niedrig gemessen wird und beim Level 3 auch zum Teil außerhalb der Herstellergrenzen, Messkammer pCO2 auffällig nass->Dichtungsring pH und Dummyelektrode neu, pCO2 Elektrode mit neuer Membran( andere LOT), mehrere Kal.-> keine Besserung Level1->pCO2 Elektrode RV-02 erneuert gegen SK-02 anschließende QCs zeigen deutliche Besserung System-und Funktionskontrolle i.O. | Fehlersuche, Kal.- QC- und Aktivitätendateien überprüft, QC Verlauf zeigt, dass pCO2 tendenziell zu niedrig gemessen wird und beim Level 3 auch zum Teil außerhalb der Herstellergrenzen, Messkammer pCO2 auffällig nass->Dichtungsring pH und Dummyelektrode neu, pCO2 Elektrode mit neuer Membran( andere LOT), mehrere Kal.-> keine Besserung Level1->pCO2 Elektrode RV-02 erneuert gegen SK-02 anschließende QCs zeigen deutliche Besserung System-und Funktionskontrolle i.O. | SWO-49472 | No | Troubleshooting, Kal.- QC and activity files checked, QC history shows that pCO2 tends to be measured too low and at level 3 also partly outside the manufacturer's limits, measuring chamber pCO2 conspicuously wet->sealing ring pH and dummy electrode new, pCO2 electrode with new membrane (other LOT), several cal.-> no improvement Level1->pCO2 electrode RV-02 renewed against SK-02 subsequent QCs show significant improvement system and function control i.O. | Troubleshooting, Kal.- QC and activity files checked, QC history shows that pCO2 tends to be measured too low and at level 3 also partly outside the manufacturer's limits, measuring chamber pCO2 conspicuously wet->sealing ring pH and dummy electrode new, pCO2 electrode with new membrane (other LOT), several cal.-> no improvement Level1->pCO2 electrode RV-02 renewed against SK-02 subsequent QCs show significant improvement system and function control i.O. | Assisted Replacement | Electrode | Electrode | pCO2
ABL80 FLEX BASIC | 更换测试卡正常。 | 更换测试卡正常。 | SWO-49392 | No | The replacement test card is OK. | The replacement test card is OK.
ABL80 FLEX BASIC |  |  | SWO-49344 | No
ABL90 FLEX | Remplacement de la batterie, test: OK S N batterie originale : P0065N020 Livraison d'un adaptateur pour calibration Hb (manquait avec livraison) | Remplacement de la batterie, test: OK S N batterie originale : P0065N020 Livraison d'un adaptateur pour calibration Hb (manquait avec livraison) | SWO-49493 | No | Battery replacement, test: OK S N original battery: P0065N020 Delivery of an adapter for Hb calibration (missed with delivery) | Battery replacement, test: OK S N original battery: P0065N020 Delivery of an adapter for Hb calibration (missed with delivery) | Defect | HW | Battery module
ABL90 |  |  | CF6E3768B24F764C | No |  |  | Assisted Replacement | SP | FT
TCM CombiM |  | 910-250借错，重借910-246一个。 | SWO-49394 | No |  | 910-250 borrows error, reborrows 910-246 one. | Defect | HW | Power Supply | Power Supply
TCM CombiM |  |  | SWO-49388 | No
ABL90 FLEX | 清洁管路，清洗进样口垫片。 | 清洁管路，清洗进样口垫片。 | SWO-49427 | No | Clean the lines and clean the inlet gaskets. | Clean the lines and clean the inlet gaskets. | Defect | HW | Inlet Gasket
AQT90 FLEX |  |  | SWO-49351 | No |  |  | Defect | HW | Process module | Process Unit sensor | 5784
ABL825 FLEX | Neuen Drucker eingebaut. | Neuen Drucker eingebaut. | SWO-49423 | No | New printer installed. | New printer installed. | Defect | HW | Printer module | Printer Head
ABL825 FLEX | Performed visual inspection, observed 14 pole cable had slight kink, no issues observed with cracker motor. Ran service diagnostics, no failures reported. | Removed autocheck module, cleaned dust, replaced 14 pole cable as precautionary measure. Retested and found no errors reported. | CF6E2824B25D4DE4 | No | Performed visual inspection, observed 14 pole cable had slight kink, no issues observed with cracker motor. Ran service diagnostics, no failures reported. | Removed autocheck module, cleaned dust, replaced 14 pole cable as precautionary measure. Retested and found no errors reported. | Defect | HW | AutoCheck Module | FPC Cable
ABL820 FLEX | METHB QC HIGH ON LEVELS S7735 AND S7745 HIGH AT 2.4 AND 5.6 PERFORMED COMPLETE OXIMETRY TEST AND HEMOLYZER TEST FAILS WITH PHASE ERROR | REPLACED HEMOLYZER.ADJUSTED HALOGEN LAMP TO 3500.PERFORMED THB CALIBRATIONS AND SERVICE MODES.PERFORMED THB CAL CHECK. | CF6E14E5B215C470 | No | METHB QC HIGH ON LEVELS S7735 AND S7745 HIGH AT 2.4 AND 5.6 PERFORMED COMPLETE OXIMETRY TEST AND HEMOLYZER TEST FAILS WITH PHASE ERROR | REPLACED HEMOLYZER. ADJUSTED HALOGEN LAMP TO 3500.PERFORMED THB CALIBRATIONS AND SERVICE MODES. PERFORMED THB CAL CHECK. | Defect | HW | Oxi module | Hemolyzer
ABL815 FLEX | Dummy Dichtring ersetzt. zur weiteren Beobachtung | Dummy Dichtring ersetzt. zur weiteren Beobachtung | SWO-49470 | No | Dummy sealing ring replaced. for further observation | Dummy sealing ring replaced. for further observation | Assisted Replacement | Electrode | Electrode | Sealing Ring
ABL825 FLEX |  |  | CF6E0E61B2537BE4 | No |  |  | Adjustment | HW | Oxi module | tHb cal
ABL825 FLEX | Silliconschlauch erneuert, Fließwege gespült. pCO2 Elektrodenkontaktübergänge mit reinem Alkohol gereinigt.. Lactate Elektrode membraniert, Status zu tief. Konfiguration gesichert-gespeichert, Alles wieder o.k. | Silliconschlauch erneuert, Fließwege gespült. pCO2 Elektrodenkontaktübergänge mit reinem Alkohol gereinigt.. Lactate Elektrode membraniert, Status zu tief. Konfiguration gesichert-gespeichert, Alles wieder o.k. | SWO-49593 | No | Sillicon hose renewed, flow paths rinsed. pCO2 electrode contact transitions cleaned with pure alcohol. Lactate electrode membranes, status too deep. Configuration backed-up-saved, all again ok. | Sillicon hose renewed, flow paths rinsed. pCO2 electrode contact transitions cleaned with pure alcohol. Lactate electrode membranes, status too deep. Configuration backed-up-saved, all again ok. | Defect | HW | Basic Unit | Waste tube
TCM TOSCA | Patient data downloaded in PD and handed over. TCM TOSCA calibrated and serviced. Working fine... | Patient data downloaded in PD and handed over. TCM TOSCA calibrated and serviced. Working fine... | SWO-49519 | No | Patient data downloaded in PD and handed over. TCM TOSCA calibrated and serviced. Working fine... | Patient data downloaded in PD and handed over. TCM TOSCA calibrated and serviced. Working fine...
ABL80 FLEX BASIC | 更换新测试卡 | 更换新测试卡 | SWO-49583 | No | Replace a new test card | Replace a new test card
ABL90 | SC causing error. | Replaced Sensor Cassette. Provided by customer. | CF72E9A78A303032 | No | SC causing error. | Replaced Sensor Cassette. Provided by customer. | Assisted Replacement | SC | thermistor
ABL80 FLEX BASIC | 更换新测试卡 | 更换新测试卡 | SWO-49604 | No | Replace a new test card | Replace a new test card
ABL80 FLEX BASIC | 更换新测试卡修复机器 | 更换新测试卡修复机器 | SWO-49639 | No | Replace the new test card repair machine | Replace the new test card repair machine
AQT90 FLEX |  |  | SWO-49605 | No |  |  | Defect | HW | Analyzer control Module | CPU | 3264
AQT90 FLEX | Performed Incubator cup sensor calibration and incubator wheel calibration, Cartridge lock calibration. | Performed Incubator cup sensor calibration and incubator wheel calibration, Cartridge lock calibration. | SWO-50679 | No | Performed Incubator cup sensor calibration and incubator wheel calibration, Cartridge lock calibration. | Performed Incubator cup sensor calibration and incubator wheel calibration, Cartridge lock calibration. | Other | HW | Incubator Module | Cup sensor | 3553
ABL80 FLEX BASIC | 更换新测试卡后正常 | 更换新测试卡后正常 | SWO-49737 | No | Normal after replacing a new test card | Normal after replacing a new test card
ABL80 FLEX BASIC |  |  | SWO-49695 | No
ABL80 FLEX BASIC | 连续更换两张测试卡后正常 | 连续更换两张测试卡后正常 | SWO-49742 | No | Normal after replacing two test cards in a row | Normal after replacing two test cards in a row
ABL80 CO-OX | Power the unit on everything calibrated. I let the unit sit overnight and when i came in the next day the unit was locked. | Replaced the CPU and CF Card | CF71FB8C145FA190 | No | Power the unit on everything calibrated. I let the unit sit overnight and when i came in the next day the unit was locked. | Replaced the CPU and CF Card
ABL80 FLEX CO-OX | QC noch einmal gemessen, LP erneuert da Lösung 1 nur noch 17 Zyklen Lot 214686, alle Parameter wieder grün, Kunde meldet sich wieder wenn nicht o.k. Gerät kalibriert, | QC noch einmal gemessen, LP erneuert da Lösung 1 nur noch 17 Zyklen Lot 214686, alle Parameter wieder grün, Kunde meldet sich wieder wenn nicht o.k. Gerät kalibriert, | SWO-49797 | No | QC measured again, LP renews since solution 1 only 17 cycles lot 214686, all parameters again green, customer reports again if not o.k. device calibrated, | QC measured again, LP renews since solution 1 only 17 cycles lot 214686, all parameters again green, customer reports again if not o.k. device calibrated,
ABL90 | After replacement of the solution pack, tHb parameters are still out of range. | Performed back flush procedure using a bleach water mixture. Ran Oxi complete test and performed a tHb calibration. Instrument is within acceptable operating specifications. | CF71C0C21438C711 | No | After replacement of the solution pack, tHb parameters are still out of range. | Performed back flush procedure using a bleach water mixture. Ran Oxi complete test and performed a tHb calibration. Instrument is within acceptable operating specifications. | Adjustment | thb cal | backflush
ABL90 FLEX | Datensätze gesendet und das Ankommen der Ergebnisse im Labor überprüft. QC Ergebnisse werden gesendet und erscheinen im System. |  | SWO-49891 | No | records and checked the arrival of the results in the lab. QC results are sent and appear in the system. |  | No fault found
ABL90 FLEX | Inhomogene Spüllösung--> Flüssigkeitsstrecke gespült, Einlassdichtung gereinigt--> Fehler bleibt bestehen--> Test im Serviceprogramm durchgeführt--> i.O--> Neustart | Inhomogene Spüllösung--> Flüssigkeitsstrecke gespült, Einlassdichtung gereinigt--> Fehler bleibt bestehen--> Test im Serviceprogramm durchgeführt--> i.O--> Neustart | SWO-49801 | No | Inhomogeneous rinsing solution--> liquid section rinsed, inlet seal cleaned--> error remains--> Test performed in the service program--> i.O--> Restart | Inhomogeneous rinsing solution--> liquid section rinsed, inlet seal cleaned--> error remains--> Test performed in the service program--> i.O--> Restart | Cleaned | Fluidic pathway
ABL90 FLEX | ABL Ein aus--> OXI complete-Test durchgeführt ( Spectrometer noice zu hoch ) --> erneut ausgeschaltet--> Kal. i.O, QC Level C fehlerhaft mit Oximetrie-Hardwarefehler (1034 ) --> Oxicompletetest nochmals durchgeführt--> QC Level C fehlerhaft --> ABL nach Absprache zur Rep mitgenommen | Spectrometer defekt ( Fehler 1022 und 1034 Spectrometer noice zu hoch ) --> Spectrometer getauscht, Initialisierung Spectrometer und Neustart durchgeführt | SWO-49767 | No | ABL An aus--> OXI complete test performed ( Spectrometer noice too high ) --> switched off again--> Kal. i.O, QC Level C faulty with oximetry hardware error (1034 ) --> Oxicompletetest again performed--> QC Level C faulty --> ABL after consultation with rep | Spectrometer defective ( Error 1022 and 1034 Spectrometer noice too high ) --> Spectrometer exchanged, initialization spectrometer and restart performed | Defect | HW | Spectrometer
ABL90 | Unit indicating Replacement of SP, or Flush.SP failing in Diagnostics, System Temp errors 1065 , 1221, Thermostatting error. | Customer Replaced SP, further troubleshooting indicated SC at fault.SC provided by customer. | CF723836141D37F0 | No | Unit indicating Replacement of SP, or Flush.SP failing in Diagnostics, System Temp errors 1065 , 1221, Thermostatting error. | Customer Replaced SP, further troubleshooting indicated SC at fault. SC provided by customer. | Assisted Replacement | sc | Wetability
ABL90 | Examined data logs. Examined the sensor cassette. | Replaced sensor cassette. Verified operaton of ABL 90. | CF720BD0143C69A9 | No | Examined data logs. Examined the sensor cassette. | Replaced sensor cassette. Verified operaton of ABL 90. | Assisted Replacement | sc | Wetability
AQT90 FLEX | 更换903-171一个后正常 |  | SWO-49849 | No | Replace 903-171 one after normal |  | Defect | HW | Aspiration Module | Bend needle | 4794
AQT90 FLEX | 不良試薬品取り除き改善。 | 不良試薬品取り除き改善。 | SWO-51080 | No | Improved removal of defective test chemicals. | Improved removal of defective test chemicals. | Defect | Consumables | Test Kit | Cartridge | 5784 (Broken)
ABL825 FLEX | Performed visual inspection. Noted the blue probe tube not completely installed on the end of the probe and appeared angled. Observed a crack in the inlet gasket holder port for the syringe with a slight leak evident. | Replaced the blue tubing for the probe, replaced the inlet gasket holder. Replaced the inlet gasket holder nipples fitting as precautionary measure. Ran service diagnostics, no observed indications of failure. | CF71E1A9142FBBFD | No | Performed visual inspection. Noted the blue probe tube not completely installed on the end of the probe and appeared angled. Observed a crack in the inlet gasket holder port for the syringe with a slight leak evident. | Replaced the blue tubing for the probe, replaced the inlet gasket holder. Replaced the inlet gasket holder nipples fitting as precautionary measure. Ran service diagnostics, no observed indications of failure. | Defect | HW | Inlet module | Inlet Gasket
ABL800 BASIC | ヘモライザー交換 | ヘモライザー交換 | SWO-49612 | No | Hemolyzer replacement | Hemolyzer replacement | Defect | HW | Oxi module | Hemolyzer
ABL825 FLEX | Erneuerung Referenz Elektrode und Messkammern Überprüfung etc. - Mehrfache Kalibrierung - Funktionsüberprüfung etc. | Erneuerung Referenz Elektrode und Messkammern Überprüfung etc. - Mehrfache Kalibrierung - Funktionsüberprüfung etc. | SWO-49778 | No | Renewal reference electrode and measuring chambers verification etc. - Multiple calibration - function check etc. | Renewal reference electrode and measuring chambers verification etc. - Multiple calibration - function check etc. | Assisted Replacement | Electrode | Electrode | Reference
ABL825 FLEX | Gerät überprüft, Barometer kontrolliert, pCO2 membraniert | Gerät überprüft, Barometer kontrolliert, pCO2 membraniert | SWO-49850 | No | Device checked, barometer controlled, pCO2 membraneed | Device checked, barometer controlled, pCO2 membraneed | Assisted Replacement | Electrode | Membrane | pCO2
ABL825 FLEX | Netwerkkabel besorgt und ersetzt, Funktionskontrolle | Netwerkkabel besorgt und ersetzt, Funktionskontrolle | SWO-49938 | Yes | Netwerk cable provided and replaced, function control | Netwerk cable provided and replaced, function control | Re classified RMED
ABL800 BASIC | Na+ membrane need replacement. Replace the Na+ membrane with new one.After replace the membrane, unit is working properly. | Na+ membrane need replacement. Replace the Na+ membrane with new one.After replace the membrane, unit is working properly. | SWO-49834 | No | Na+ membrane need replacement. Replace the Na+ membrane with new one. After replace the membrane, unit is working properly. | Na+ membrane need replacement. Replace the Na+ membrane with new one. After replace the membrane, unit is working properly. | Assisted Replacement | Electrode | Membrane | Na+
ABL800 BASIC | Gerät überprüft, Kal überprüft, Lampeneinheit justiert, Oxi Complet Check durchgeführt, Service tHb Kal, Kals und QCs durchgeführt. | Gerät überprüft, Kal überprüft, Lampeneinheit justiert, Oxi Complet Check durchgeführt, Service tHb Kal, Kals und QCs durchgeführt. | SWO-50158 | No | Device checked, Kal checked, lamp unit adjusted, Oxi Complet check performed, service tHb Kal, Kals and QCs performed. | Device checked, Kal checked, lamp unit adjusted, Oxi Complet check performed, service tHb Kal, Kals and QCs performed. | Adjustment | HW | Oxi module | Lamp unit
ABL80 FLEX BASIC | 更换测试卡 | 更换测试卡 | SWO-49972 | No | Replace the test card | Replace the test card
ABL80 FLEX BASIC |  |  | SWO-49964 | No
ABL90 |  |  | CF7361438A63BD1B | Yes |  |  | Re classified RMED | SWU | Planned visit
ABL90 FLEX | Fehlersuche: LP ID-PCB defekt, ersetzt. Sevicemenue aufgerufen, Chip Daten jetzt lesbar. Diverse Flüssigkeitstransport - Programme laufen lassen, alle erfolgreich absolviert. Neustart, Startphase, Kalibrationen, Funktionstest und QC Messungen, OK | Fehlersuche: LP ID-PCB defekt, ersetzt. Sevicemenue aufgerufen, Chip Daten jetzt lesbar. Diverse Flüssigkeitstransport - Programme laufen lassen, alle erfolgreich absolviert. Neustart, Startphase, Kalibrationen, Funktionstest und QC Messungen, OK | SWO-50052 | No | Troubleshooting: LP ID-PCB defective, replaced. Sevicemenue called, chip data now readable. Various fluid transport - running programs, all successfully completed. Restart, start-up phase, calibrations, function test and QC measurements, OK | Troubleshooting: LP ID-PCB defective, replaced. Sevicemenue called, chip data now readable. Various fluid transport - running programs, all successfully completed. Restart, start-up phase, calibrations, function test and QC measurements, OK | Defect | HW | SP PCB
ABL90 FLEX | Fehlersuche: Nasssektion sichtkontrolliert, minimale Spuren unterhalb des Ventilschlauches, System gespült, Flüssigkeitsaustritt am Ventilschlauch => Mikroperforation festgestellt, Schlauch ersetzt. LP-ID PCB mit ungleich hohen Kontakten => ersetzt. Alle flüssigkeitsbezogenen Serviceprogramme absolviert, alle ok. Startphase, Kalibrationen, QC Messungen. | Fehlersuche: Nasssektion sichtkontrolliert, minimale Spuren unterhalb des Ventilschlauches, System gespült, Flüssigkeitsaustritt am Ventilschlauch => Mikroperforation festgestellt, Schlauch ersetzt. LP-ID PCB mit ungleich hohen Kontakten => ersetzt. Alle flüssigkeitsbezogenen Serviceprogramme absolviert, alle ok. Startphase, Kalibrationen, QC Messungen. | SWO-50107 | No | Troubleshooting: Wet section visually controlled, minimal traces below the valve hose, system flushed, liquid outlet on the valve hose => Microperforation detected, hose replaced. LP-ID PCB with unequal contacts => replaced. Completed all liquid-related service programs, all ok. Start phase, calibrations, QC measurements. | Troubleshooting: Wet section visually controlled, minimal traces below the valve hose, system flushed, liquid outlet on the valve hose => Microperforation detected, hose replaced. LP-ID PCB with unequal contacts => replaced. Completed all liquid-related service programs, all ok. Start phase, calibrations, QC measurements. | Defect | HW | Oxi module | Pinch Valve tube
ABL90 FLEX | De accu vervangen voor een nieuwe. ABL start nu normaal op, en accu werkt naar behoren. Accu getest, alles werkt naar behoren. |  | SWO-49924 | No | Replace the battery for a new one. ABL now starts up normally, and battery works properly. Battery tested, everything works properly. |  | Defect | HW | Battery module
ABL90 | Troubleshot and found that the customer had replaced the solution pack and sensor cassette, but the ampoule based QC had not been ran. | Ran all levels of ampoule based QC, and this resolved the problem. | CF735D698A2C9D1A | No | Troubleshot and found that the customer had replaced the solution pack and sensor cassette, but the ampoule based QC had not been ran. | Ran all levels of ampoule based QC, and this resolved the problem. | User Error | Manual QC lockout
ABL90 FLEX | Qualitätskontrolle 2* wiederholt. QCs jeweils in Ordnung. Am Vortag vereinzelt instabile Kal3-Lösung. Weiter kein Fehler feststellbar. | Qualitätskontrolle 2* wiederholt. QCs jeweils in Ordnung. Am Vortag vereinzelt instabile Kal3-Lösung. Weiter kein Fehler feststellbar. | SWO-50062 | No | Quality control 2* repeated. QCs all right. The day before, unstable Kal3 solution. No further error can be detected. | Quality control 2* repeated. QCs all right. The day before, unstable Kal3 solution. No further error can be detected. | No fault found
ABL90 FLEX | Verbindung eingerichtet HL7 2.2, TCP IP, 10.38.40.23, Port 20901 |  | SWO-50085 | No | Connection established HL7 2.2; TCP IP; 10.38.40.23; Port 20901 |  | Setup | network | HL7
ABL90 FLEX | Spectrometer defekt ( Fehler 1022 und 1034 Spectrometer noice zu hoch ) --> Spectrometer getauscht, Initialisierung Spectrometer und Neustart durchgeführt | Spectrometer defekt ( Fehler 1022 und 1034 Spectrometer noice zu hoch ) --> Spectrometer getauscht, Initialisierung Spectrometer und Neustart durchgeführt | SWO-49918 | No | Spectrometer defective ( Error 1022 and 1034 Spectrometer noice too high ) --> Spectrometer exchanged, initialization spectrometer and restart performed | Spectrometer defective ( Error 1022 and 1034 Spectrometer noice too high ) --> Spectrometer exchanged, initialization spectrometer and restart performed | Defect | HW | Spectrometer
ABL90 | This is a NICU location, suspect blockage in Cassette. Customer had tried replacing SP, unit then went into Replace SP and Flush. Entering into Service I was able to have unit not going into Replace routine, continued troubleshooting. | Flushed unit with warm water, ran diagnostics utilities. | CF735CAD8A35C4E7 | No | This is a NICU location, suspect blockage in Cassette. Customer had tried replacing SP, unit then went into Replace SP and Flush. Entering into Service I was able to have unit not going into Replace routine, continued troubleshooting. | Flushed unit with warm water, ran diagnostics utilities. | Cleaned | Fluidic pathway
ABL5 |  |  | SWO-50798 | No
TCM CombiM |  |  | SWO-49852 | No
ABL90 | Reinicio del analizador, mantenimiento preventivo, comprobación de registros, comprobación de la impresora, cheking por los programas de servicio, verificación final con QC. | Reinicio del analizador | CF73054E8A4EE4AF | No | Scanner restart, preventive maintenance, log checking, printer check, check by service programs, final verification with QC. | Rebooting the analyzer | Adjustment | restart analyzer | FT
AQT90 FLEX | 更换903-329一个后正常 | 更换903-329一个后正常 | SWO-49711 | No | Replace 903-329 one after normal | Replace 903-329 one after normal | Defect | HW | Aspiration Module | Needle Wash Unit | 1264.1265
ABL825 FLEX | checked tubes probe nipples waste tubes found cracked and blocked nipple | replaced nipples on inlet | CF7328B58A0C7C4A | No | checked tubes probe nipples waste tubes found cracked and blocked nipple | replaced nipples on inlet | Defect | HW | Inlet module | Leaking | Nipple kit
ABL800 BASIC | 更换902-830一个后正常 | 更换902-830一个后正常 | SWO-51075 | No | Replace 902-830 one after normal | Replace 902-830 one after normal
ABL825 FLEX | Lampe korrekt justiert, tHb Kal durchgeführt |  | SWO-50086 | No | Lamp correctly adjusted, tHb Kal performed |  | Adjustment | HW | Oxi module | Lamp unit
ABL825 FLEX | Gerät überprüft, Meßkammer gesäubert, Elektroden kontrolliert, Dichtungsring pO2 ersetzt, Barometer kontrolliert, pO2 und Ref-Elektrode membraniert | Mit Frau Eva Wortmann. Die Einlassdichtung ausgetauscht. Das Gerät bricht die Leckage Test ab bei ca .12 min. | SWO-50087 | No | Device checked, measuring chamber cleaned, electrodes checked, sealing ring pO2 replaced, barometer controlled, pO2 and ref electrode membraneed | With Mrs. Eva Wortmann. Replace the inlet seal. The device cancels the leakage test at about .12 min. | Assisted Replacement | Electrode | Electrode | pO2 | Sealing ring
ABL820 FLEX | Valeur controles Basse < 2sd | Remplac la chambre de mesure ,Pco2 contact, pompe rotor, pompe BG tubing | CF73323C8A1092AD | No | Low Controls Value - 2sd | Replace the measuring chamber ,Pco2 contact, rotor pump, BG tubing pump | Assisted Replacement | Electrode | Electrode | pCO2
ABL825 FLEX | 更换902-832一个后正常 | 更换902-832一个后正常 | SWO-50154 | No | Replace 902-832 one after normal | Replace 902-832 one after normal
ABL80 FLEX BASIC | 更换新测试卡和新试剂包修复机器 | 更换新测试卡和新试剂包修复机器 | SWO-50140 | No | Replace the new test card and the new reagent kit repair machine | Replace the new test card and the new reagent kit repair machine
ABL80 FLEX BASIC | 更换新测试卡 | 更换新测试卡 | SWO-50171 | No | Replace a new test card | Replace a new test card
ABL90 |  |  | CF74A3580058EA3E | Yes |  |  | Re classified RMED | yearly maintenance | Planned visit
ABL90 | Solution pack was not fully seated. Partial blockage in waste connector. | Replace waste connector and waste tubing. Replace inlet components. Replace pinch valve tubing. Test for leaks or blockages - none found. | CF747789003FA1DF | No | Solution pack was not fully seated. Partial blockage in waste connector. | Replace waste connector and waste tubing. Replace inlet components. Replace pinch valve tubing. Test for leaks or blockages - none found.
ABL90 FLEX | Gerät geprüft. Einlass geprüft und gereinigt. LP entfernt und erneut installiert. Diverse Kal und QC durchgeführt. THB Kal. |  | SWO-50739 | No | device checked. Inlet checked and cleaned. REMOVED and reinstalled LP. Various Kal and QC performed. THB Kal. |  | Assisted Replacement | SP | pH backlog?
ABL90 |  |  | CF7595B8765C5E82 | Yes |  |  | Re classified RMED | Installation | Planned visit
ABL90 FLEX | Gerät durchgespült, Lösungspack durch Kundeneigenes Betriebsmittel aufgrund von Inhomoginität der QC2 Lösung erneuert, Gerät kalibriert und QC durchgeführt. Kunden zum Radiometer Retour Verfahren geschult. | Gerät durchgespült, Lösungspack durch Kundeneigenes Betriebsmittel aufgrund von Inhomoginität der QC2 Lösung erneuert, Gerät kalibriert und QC durchgeführt. Kunden zum Radiometer Retour Verfahren geschult. | SWO-50319 | No | Device flushed through, solution pack replaced by customer's own equipment due to inhomoginity of the QC2 solution, device calibrated and QC performed. customers trained to use the radiometer return procedure. | Device flushed through, solution pack replaced by customer's own equipment due to inhomoginity of the QC2 solution, device calibrated and QC performed. customers trained to use the radiometer return procedure. | Assisted Replacement | SP | Inhomo QC2
ABL90 | Customer had replaced SP in an effort to make unit Ready.Upon arrival found unit in Replace SP Flush.Enter Service bypassed Replace SP. Diagnostics=OK. | Replaced SC. | CF7494BE003A5366 | No | Customer had replaced SP in an effort to make unit Ready.Upon arrival found unit in Replace SP Flush.Enter Service bypassed Replace SP. Diagnostics=OK. | Replaced SC. | Assisted Replacement | SC | pCO2
ABL90 FLEX | Fehlersuche, roma.ini angepaßt, alles i.O. | Fehlersuche; roma.ini angepaßt; alles i.O. | SWO-50229 | No | Troubleshooting; roma.ini adapted; everything in the | Troubleshooting; roma.ini adapted; everything in the | Setup | Roma.ini | Barcode Reader
ABL90 FLEX | 更换910-342一个后正常 | 更换910-342一个后正常 | SWO-49715 | No | Replace 910-342 one after normal | Replace 910-342 one after normal | Defect | HW | CPU
ABL90 FLEX | 更换测试卡正常。 | 更换测试卡正常。 | SWO-50284 | No | The replacement test card is OK. | The replacement test card is OK. | Assisted Replacement | SC | Reference
ABL90 FLEX | 更换新测试卡后正常 | 更换新测试卡后正常 | SWO-50543 | No | Normal after replacing a new test card | Normal after replacing a new test card | Assisted Replacement | SC | thermistor
ABL80 CO-OX | Verificação dos registos de dados. Verificação do circuito de fluidos. Foi detectada solução de QC no SL. Limpeza do sensor de liquidos. Lubrificação da valvula de oximetria. Limpeza do circuito de fluidos. Inicialização. Calibração da ctHb. Calibração e verificação do correcto funcionamento com ciclo de qualidade. |  | CF7491A9005FE99D | No | Verification of data records. Fluid circuit check. QC solution was detected in SL. Cleaning the liquid sensor. Oximetry valve lubrication. Cleaning the fluid circuit. Startup. CtHb Calibration. Calibration and verification of the correct operation with quality cycle.
ABL805 FLEX | spoke with Justin Holis, stated that during a virus check the analyzer was flagged for a virus | changed out the hard drive installed windows and 6.13 ABL software. Sent the other back to Ronni Cahoon in Copenhagen to check the drive. DMS not connected error happened after install of 6.13 changed out power supply to fix the error | CF744E2C0021717D | Yes | spoke with Justin Holis, stated that during a virus check the analyzer was flagged for a virus | changed out the hard drive installed windows and 6.13 ABL software. Sent the other back to Ronni Cahoon in Copenhagen to check the drive. DMS not connected error happened after install of 6.13 changed out power supply to fix the error | Re classified RMED | SenDx analyzer
ABL815 FLEX | 更换902-496一个后正常 | 更换902-496一个后正常 | SWO-49848 | No | Replace 902-496 one after normal | Replace 902-496 one after normal
ABL800 BASIC | 更换910-274一个和910-277一个后未修复，需要借910-321一个和945-614一个继续维修 | 更换910-321一个，945-614一个正常。 | SWO-49177 | No | Replace 910-274 one and 910-277 one after not repaired, need to borrow 910-321 one and 945-614 one to continue repair | Replace 910-321 one, 945-614 one normal.
ABL825 FLEX | Customer had inlet positioning error multiple times. Instrument is new, but issue occurs every time customer goes into hold mode and remembrane an electrode. | Replaced friction detector pcb. Verified instrument operation. | CF74B98300341A0A | No | Customer had inlet positioning error multiple times. Instrument is new, but issue occurs every time customer goes into hold mode and remembrane an electrode. | Replaced friction detector pcb. Verified instrument operation.
ABL825 FLEX | pH Elektrode ersetzt | pH Elektrode ersetzt | SWO-50355 | No | pH electrode replaced | pH electrode replaced | Assisted Replacement | Electrode | Electrode | pH
ABL825 FLEX |  |  | CF7449940025982C | Yes |  |  | Re classified RMED | Telephone support
ABL815 FLEX | alle Gaswege geprüft und mit Druckluft durchgespült, Ventile im BG modul ausgebaut und gereinigt, Gerät kalibriert, Gasdruck im Serviceprogramm kontrolliert, Datum steht auf der Flasche, | alle Gaswege geprüft und mit Druckluft durchgespült, Ventile im BG modul ausgebaut und gereinigt, Gerät kalibriert, Gasdruck im Serviceprogramm kontrolliert, Datum steht auf der Flasche, | SWO-50258 | No | all gas paths tested and rinsed with compressed air, valves removed and cleaned in the BG module, device calibrated, gas pressure checked in the service program, date is on the bottle, | all gas paths tested and rinsed with compressed air, valves removed and cleaned in the BG module, device calibrated, gas pressure checked in the service program, date is on the bottle, | Cleaned | Fluidic pathway | - | Gas Valve
ABL80 FLEX BASIC | 更换新测试卡 | 更换新测试卡 | SWO-50852 | No | Replace a new test card | Replace a new test card
ABL80 FLEX BASIC | 更换新测试卡后正常 | 更换新测试卡后正常 | SWO-50519 | No | Normal after replacing a new test card | Normal after replacing a new test card
ABL80 FLEX BASIC | 连续更换三张测试卡后正常 | 连续更换三张测试卡后正常 | SWO-51140 | No | Normal after replacing three test cards in a row | Normal after replacing three test cards in a row |  |  |  | Optical Unit
ABL90 | CUSTOMER FLUSHED 15 TIMES AND NEW SOLUTION PACK. NO HELP. | REPLACED SENSOR CASSETTE. DID 1 WASTE BLOOD. HAD TO DO A FEW CALIBRATIONS TO GET THE SODIUM IN FOR EXTERNAL QC LEVEL 3. IT ALL CAME IN. | CF75D97376057E11 | No | CUSTOMER FLUSHED 15 TIMES AND NEW SOLUTION PACK. NO HELP. | REPLACED SENSOR CASSETTE. DID 1 WASTE BLOOD. HAD TO DO A FEW CALIBRATIONS TO GET THE SODIUM IN FOR EXTERNAL QC LEVEL 3. IT ALL CAME IN.
ABL90 FLEX | Einlassdichtung aus Kundenbestand erneuert, Service Pressure Test i.O. Kundin meldet Fehler bei der tHb Bestimmung, weitere Untersuchungen notwendig. | Einlassdichtung aus Kundenbestand erneuert, Service Pressure Test i.O. Kundin meldet Fehler bei der tHb Bestimmung, weitere Untersuchungen notwendig. | SWO-50515 | No | Inlet seal renewed from customer stock, Service Pressure Test i.O. Customer reports errors in the tHb determination, further examinations necessary. | Inlet seal renewed from customer stock, Service Pressure Test i.O. Customer reports errors in the tHb determination, further examinations necessary. | Defect | HW | Inlet Gasket
ABL90 FLEX | Hemolyzer erneuert. OXI Complete Test i.O. Kal. und tHb Kal. i.O. | Hemolyzer erneuert. OXI Complete Test i.O. Kal. und tHb Kal. i.O. | SWO-50517 | No | Hemolyzer erneuert. OXI Complete Test i.O. Kal. und tHb Kal. i.O. | Hemolyzer erneuert. OXI Complete Test i.O. Kal. und tHb Kal. i.O. | Defect | HW | Hemolyzer
TCM-COMBIM | Comprobar módulos y fuente de alimentación, cambiar el electrodo, calibrar y probar. | Cambiar el eletrodo. | CF75B1E77634B366 | No | Check modules and power supply, change the electrode, calibrate and test. | Change the eletrode.
AQT90 FLEX | 更换903-364一个正常。 | 更换903-364一个正常。 | SWO-50162 | No | Replace 903-364 one normal. | Replace 903-364 one normal. | Defect | HW | Process Module | Optical Unit | 131
AQT90 FLEX |  |  | SWO-50352 | No |  |  | Defect | HW | Incubator Module | Incubator, PCB | 2167
ABL827 FLEX |  |  | CF772503EC15D543 | No
ABL800 BASIC | Oxi PCB ausgetauscht ,Fehler weiterhin vorhanden, Complete Oxi Modul ausgetauscht, Oxi Complete Test, Service tHb Kal durchgeführt. Kal und QC i.O | Oxi PCB ausgetauscht ,Fehler weiterhin vorhanden, Complete Oxi Modul ausgetauscht, Oxi Complete Test, Service tHb Kal durchgeführt. Kal und QC i.O | SWO-50400 | No | Oxi PCB exchanged ,Error still present, Complete Oxi module replaced, Oxi Complete Test, Service tHb Kal performed. Kal and QC i.O | Oxi PCB exchanged ,Error still present, Complete Oxi module replaced, Oxi Complete Test, Service tHb Kal performed. Kal and QC i.O | Defect | HW | Oxi module | PCB
ABL837 FLEX | Intermittent screen problem. Screen black and sometimes CPU boots up. | Screen was tested by exchanging it for another analyzer's screen. Replaced power supply. Replaced CPU. | CF75CC8B76386C89 | No | Intermittent screen problem. Screen black and sometimes CPU boots up. | Screen was tested by exchanging it for another analyzer's screen. Replaced power supply. Replaced CPU.
ABL825 FLEX | Module ausgebaut, Gasventilmembranen erneuert, Ventile, Elektroden und Fließwege kontrolliert, Pumpenschläuche Module ersetzt, Meßkammer und Luftauslass gesäubert, | Module ausgebaut, Gasventilmembranen erneuert, Ventile, Elektroden und Fließwege kontrolliert, Pumpenschläuche Module ersetzt, Meßkammer und Luftauslass gesäubert, | SWO-50531 | No | Modules removed, gas valve membranes renewed, valves, electrodes and flow paths inspected, pump hose modules replaced, measuring chamber and air outlet cleaned, | Modules removed, gas valve membranes renewed, valves, electrodes and flow paths inspected, pump hose modules replaced, measuring chamber and air outlet cleaned, | Defect | HW | El/Met Module | Valve, air
ABL825 FLEX | Gerät überprüft, Meßkammer und Luftauslass gesäubert, pH Elektrode gewaschen, Elektroden kontrolliert, | Gerät überprüft, Meßkammer und Luftauslass gesäubert, pH Elektrode gewaschen, Elektroden kontrolliert, | SWO-50530 | No | device checked, measuring chamber and air outlet cleaned, pH electrode washed, electrodes checked, | device checked, measuring chamber and air outlet cleaned, pH electrode washed, electrodes checked, | Cleaned | Fluidic pathway | -
ABL815 FLEX | Pumpenschläuche erneuert. |  | SWO-50512 | No | Pump hoses renewed. |  | Assisted Replacement | HW | pH/BG Module | Pump tubing
AQT90 FLEX | Aufstellung Austauschgerät 838R0021N005 Einspielung Konfiguration Kundengerät zur Reparatur mit in GS ggf. Zentrale genommen | Aufstellung Austauschgerät 838R0021N005 Einspielung Konfiguration Kundengerät zur Reparatur mit in GS ggf. Zentrale genommen | SWO-50475 | No | Installation replacement device 838R0021N005 recording configuration customer device for repair with taken in GS possibly central | Installation replacement device 838R0021N005 recording configuration customer device for repair with taken in GS possibly central | Defect | SW | AQT9 | Breakdown | 42173
ABL80 FLEX BASIC | 更换测试卡正常 | 更换测试卡正常 | SWO-51343 | No | Replace the test card ok | Replace the test card ok
ABL80 FLEX BASIC | 更换两张测试卡修复机器 | 更换两张测试卡修复机器 | SWO-50629 | No | Replace two test cards to repair the machine | Replace two test cards to repair the machine
ABL80 FLEX BASIC | 更换新测试卡 | 更换新测试卡 | SWO-50612 | No | Replace a new test card | Replace a new test card
ABL80 FLEX BASIC | 更换新测试卡修复机器 | 更换新测试卡修复机器 | SWO-50630 | No | Replace the new test card repair machine | Replace the new test card repair machine
ABL80 FLEX BASIC | 更换新测试卡 | 更换新测试卡 | SWO-50582 | No | Replace a new test card | Replace a new test card
ABL80 FLEX CO-OX | Meldung kam immer wieder, keine interne Spülung möglich, SC oder LP konnte nicht installiert werden Gerät aus- und wieder eingeschaltet alle Notwendigen Kalibrierungen durchgeführt LP ist am Vortag auch ausgelaufen - wird reklamiert, neuer LP wurde vom Kunden vorher schon eingesetzt | Meldung kam immer wieder, keine interne Spülung möglich, SC oder LP konnte nicht installiert werden Gerät aus- und wieder eingeschaltet alle Notwendigen Kalibrierungen durchgeführt LP ist am Vortag auch ausgelaufen - wird reklamiert; neuer LP wurde vom Kunden vorher schon eingesetzt | SWO-50695 | No | Message came again and again, no internal flushing possible, SC or LP could not be installed device off and again switched on all necessary calibrations performed LP is also expired the day before - is complained; new LP has been used by the customer before | Message came again and again, no internal flushing possible, SC or LP could not be installed device off and again switched on all necessary calibrations performed LP is also expired the day before - is complained; new LP has been used by the customer before
ABL80 FLEX Basic version |  |  | CF76E748EC102F89 | No
ABL80 CO-OX | It is not pulling any solutions. | Replaced the tubing. | CF77453BEC077997 | No | It is not pulling any solutions. | Replaced the tubing.
ABL90 |  |  | CF7729DCEC61D550 | No
ABL90 FLEX | Hemolysator erneuert. | Hemolysator erneuert. | SWO-50660 | No | Hemolysator. | Hemolysator. | Defect | HW | Hemolyzer
ABL90 FLEX | Gerät überprüft, mehrfach Vergleichsmessungen zum Labor durchgeführt, Vergleichsmessungen OK |  | SWO-50681 | No | Device checked, multiple comparative measurements carried out to the laboratory, comparative measurements OK |  | No fault found | pH
ABL90 FLEX | Anwender hatte bereits gespült, Einlassdichtung erneuert und SP getestet--> keine Besserung Flüssigkeitswege im Serviceprogramm geprüft--> Einlassdichtung erneut getauscht, keine Besserung --> SP getauscht --> Fehlermeldung: Chipdaten werden nicht gelesen --> auch der alte SP wird nicht erkannt--> aus ein bringt keine Besserung --> Solutionpack ID PCB erneuert--> Neustart mit neuem SP Inkonsistente Datenbank --> nach Abschluss des Neustartes den SP erneut eingelegt--> Fehler bleibt bestehen--> MedTech. ist informiert und beobachtet. Kal und QC fehlerfrei | Anwender hatte bereits gespült, Einlassdichtung erneuert und SP getestet--> keine Besserung Flüssigkeitswege im Serviceprogramm geprüft--> Einlassdichtung erneut getauscht, keine Besserung --> SP getauscht --> Fehlermeldung: Chipdaten werden nicht gelesen --> auch der alte SP wird nicht erkannt--> aus ein bringt keine Besserung --> Solutionpack ID PCB erneuert--> Neustart mit neuem SP Inkonsistente Datenbank --> nach Abschluss des Neustartes den SP erneut eingelegt--> Fehler bleibt bestehen--> MedTech. ist informiert und beobachtet. Kal und QC fehlerfrei | SWO-50667 | No | User had already rinsed, renewed inlet seal and SP tested--> no improvement liquid paths in the service program checked--> inlet seal replaced again, no improvement --> SP exchanged --> Error message: Chip data is not read --> also the old SP is not detected--> from a brings no improvement --> Solutionpack ID PCB renews--> Restart with new SP Inconsistent database --> after the restart the SP is inserted again--> Error remains--> MedTech. is informed and observed. Kal and QC error-free | User had already rinsed, renewed inlet seal and SP tested--> no improvement liquid paths in the service program checked--> inlet seal replaced again, no improvement --> SP exchanged --> Error message: Chip data is not read --> also the old SP is not detected--> from a brings no improvement --> Solutionpack ID PCB renews--> Restart with new SP Inconsistent database --> after the restart the SP is inserted again--> Error remains--> MedTech. is informed and observed. Kal and QC error-free | Defect | HW | SP PCB
ABL90 FLEX | SC geliefert, Lac wieder aktiviert u.v. Labor entsperren lassen, div. Reinigungen, installierte SC erneut eingelesen, Startphase abgewartet, Kal. + QC i.O. | SC geliefert; Lac wieder aktiviert u.v. Labor entsperren lassen; div. Reinigungen; installierte SC erneut eingelesen; Startphase abgewartet; Kal. + QC i.O. | SWO-50713 | No | SC supplied; Lac reactivated and let laboratory unlock; div. cleanings; re-read the installed SC; start-up phase awaited; Kal. + QC in the first place | SC supplied; Lac reactivated and let laboratory unlock; div. cleanings; re-read the installed SC; start-up phase awaited; Kal. + QC in the first place | Assisted Replacement | SC | Lac
ABL90 | the softwares incompatable | i orderd a new cf card and attemped to load windows. my laptop needs administrative access to get software | CF79EC25D8522C57 | No | the softwares incompatable | i orderd a new cf card and attemped to load windows. my laptop needs administrative access to get software
ABL90 FLEX | 更换新测试卡后正常 | 更换新测试卡后正常 | SWO-51095 | No | Normal after replacing a new test card | Normal after replacing a new test card
ABL90 FLEX | 更换新测试卡后正常 | 更换新测试卡后正常 | SWO-51346 | No | Normal after replacing a new test card | Normal after replacing a new test card
AQT90 FLEX |  |  | SWO-50524 | No |  |  | Defect | HW | Aspiration Module | Needle Wash Unit | 1258
AQT90 FLEX | 更换903-364修复机器 | 更换903-364修复机器 | SWO-50144 | No | Replace the 903-364 repair machine | Replace the 903-364 repair machine | Defect | HW | Process Module | Optical Unit | 135
ABL800 BASIC | 更换910-321一个，945-614一个正常。 | 更换910-321一个，945-614一个正常。 | SWO-50142 | No | Replace 910-321 one, 945-614 one normal. | Replace 910-321 one, 945-614 one normal.
ABL820 FLEX | The analyzers are not connecting to radiance. | I contacted our Sernior IT Specialist, Steve McClain. We rebooted the radiance server and tried some changes to the settings. However, I was not able to change any settings becuase of administrative restrictions. Advised the customer to have their IT person with administrative rights contact our IT specialist in order to reconfigure the radiance server. | CF7700C2EC3C9776 | No | The analyzers are not connecting to radiance. | I contacted our Sernior IT Specialist, Steve McClain. We rebooted the radiance server and tried some changes to the settings. However, I was not able to change any settings becuase of administrative restrictions. Advised the customer to have their IT person with administrative rights contact our IT specialist in order to reconfigure the radiance server.
ABL800 BASIC | Gerät überprüft und gereinigt, Ref neu membraniert, Kals und QCs durchgeführt | Fr. Dreyer empfohlen, die Na Elektrode zu membranieren, was nur widerwillig umsetzen möchte, da sie der Meinung, das andauernd die Elektroden membraniert werden müssen! Habe Sie gefragt, ob sie das weiter beschreiben könnte, da muss sie dann passen. Fr. Dreyer möchte gerne die Alarmgrenzen für den Abfallbehälter verändern - was leider nicht vorgesehen ist, das findet sie auch nicht gut. Bei weiteren Fragen wird sie sich wieder melden. | SWO-50680 | No | Device checked and cleaned, ref re-membraneed, kals and QCs performed | Ms. Dreyer recommends membraneing the Na electrode, which is reluctant to implement, as she believes that the electrodes need to be membraneed all the time! If she asked if she could describe this further, she would have to fit in. Ms. Dreyer would like to change the alarm limits for the waste container - which unfortunately is not provided for, which she does not like either. She will contact me again if she has any further questions.
ABL825 FLEX | System überprüft, OK, Kalibrationen und Qualitätskontrollen ohne Beanstandungen durchgeführt. | System überprüft, OK, Kalibrationen und Qualitätskontrollen ohne Beanstandungen durchgeführt. | SWO-51172 | No | System checked, OK, calibrations and quality controls carried out without complaints. | System checked, OK, calibrations and quality controls carried out without complaints.
ABL825 FLEX | 更换945-619一个正常。 | 更换945-619一个正常。 | SWO-50303 | No | Replace the 945-619 one normal. | Replace the 945-619 one normal. | Assisted Replacement | Electrode | Electrode | Lac
ABL825 FLEX | Da in der Vergangenheit schon diverse Maßnahmen getroffen wurden, wird nun das BG Modul testweise ersetzt. Herr Vorwalter wird die nächsten Tage das Kalibrierverhalten verstärkt beobachten. BG Modul erneuert. Elektroden aus Kundenbestand übernommen. Kal .und QC o.k. | Kunde bittet um einen erneuten Besuch. Hr. Vorwalter bittet um einen Rückruf zwecks Terminabsprache. | SWO-50822 | No | Since various measures have already been taken in the past, the BG module will now be replaced on a trial basis. Mr Vorwalter will be watching the calibration behaviour more closely over the next few days. BG module renewed. electrodes from customer base. Kal .and QC ok. | Customer asks for a new visit. Mr. Vorwalter asks for a call back for appointments.
ABL825 FLEX | Fließwege und Messkammer gereinigt, pH-Elektrode ersetzt, Gerät kalibriert, | Fließwege und Messkammer gereinigt, pH-Elektrode ersetzt, Gerät kalibriert, | SWO-50671 | No | Flow paths and measuring chamber cleaned, pH electrode replaced, device calibrated, | Flow paths and measuring chamber cleaned, pH electrode replaced, device calibrated,
ABL825 FLEX | Valve 1 und 2 erneuert. anschließend Halbjahreswartung durchgeführt, siehe SWO Garantie !!!!! | Valve 1 und 2 erneuert. anschließend Halbjahreswartung durchgeführt, siehe SWO Garantie !!!!! | SWO-50818 | No | Valve 1 and 2 renewed. then carried out half-year maintenance, see SWO warranty !!!!! | Valve 1 and 2 renewed. then carried out half-year maintenance, see SWO warranty !!!!!
ABL815 FLEX | Applikation SW deinstalliert und wieder neue installiert, der Fehler ist weiterhin vorhanden. Danach Windows neue installiert, Applikation SW installiert , Das Gerät konfiguriert.tHb Kal durchgeführt. Kal und QC i.O. | Applikation SW deinstalliert und wieder neue installiert, der Fehler ist weiterhin vorhanden. Danach Windows neue installiert, Applikation SW installiert , Das Gerät konfiguriert.tHb Kal durchgeführt. Kal und QC i.O. | SWO-50756 | No | Application SW uninstalled and new ones installed again, the error still exists. After that Windows new installed, application SW installed , the device configured.tHb Kal performed. Kal and QC i.O. | Application SW uninstalled and new ones installed again, the error still exists. After that Windows new installed, application SW installed , the device configured.tHb Kal performed. Kal and QC i.O.
ABL825 FLEX |  |  | CF770A4DEC613D95 | No
ABL825 FLEX | Gerät überprüft, pCO2 membraniert (Charge 639 und 672) | Gerät überprüft, pCO2 membraniert (Charge 639 und 672) | SWO-50824 | No | Device checked, pCO2 membraneed (charge 639 and 672) | Device checked, pCO2 membraneed (charge 639 and 672)
ABL825 FLEX | fehlenden Dichtungsring pO2 ersetzt, Einlaßdichtung ausgetauscht, Elektroden kontrolliert, Luftauslass gesäubert, tHb Kal o.k. | fehlenden Dichtungsring pO2 ersetzt, Einlaßdichtung ausgetauscht, Elektroden kontrolliert, Luftauslass gesäubert, tHb Kal o.k. | SWO-50825 | No | replaced missing sealing ring pO2, inlet seal replaced, electrodes checked, air outlet cleaned, tHb Kal o.k. | replaced missing sealing ring pO2, inlet seal replaced, electrodes checked, air outlet cleaned, tHb Kal o.k.
ABL825 FLEX | Gerät überprüft, Dichtungsring pH, Einlaßdichtung | Module ausgebaut, Gasventilmembranen erneuert, Ventile, Elektroden und Fließwege kontrolliert, Pumpenschläuche Module ersetzt, Meßkammer und Luftauslass gesäubert, | SWO-50639 | No | Device checked, sealing ring pH, inlet seal | Modules removed, gas valve membranes renewed, valves, electrodes and flow paths inspected, pump hose modules replaced, measuring chamber and air outlet cleaned,
ABL825 FLEX | Gerät überprüft, Einlaßdichtung ersetzt | Gerät überprüft, Einlaßdichtung ersetzt | SWO-50823 | No | Device checked, inlet seal replaced | Device checked, inlet seal replaced
ABL810 FLEX | Oxi compleet Test, Kalibration und QC Kontrollen sind einwandfrei. Das Gerät ist wieder voll funktionsbereit. | Fr. Wirth bittet um einen Rückruf und Service. | SWO-50583 | No | Oxi compleet test, calibration and QC controls are flawless. The device is fully functional again. | Ms. Wirth asks for a callback and service. | Unknown | HW | Oxi module | Neon number low
ABL827 FLEX | Einlass gereinigt, Einlasssonde u. - Dichtung ersetzt ( aus Kundenbestand), Reinigungs-, Kal1- und Kal2-Lösung aktiviert und ersetzt, Neustart, akz. Kal abgewartet, Kal u. QC i.O. | Frau D. ist nicht am Gerät und meldet uns eine verbogene Nadel. Der Haustechniker hat leider keine Nadel vorrätig. Frau D bittet um einen Vorortservice | SWO-50729 | No | Inlet cleaned, inlet probe and seal replaced (from customer stock), cleaning, Kal1 and Kal2 solution activated and replaced, restart, etc. Kal waited, Kal & QC i.O. | Ms. D. is not on the device and reports us a bent needle. Unfortunately, the house technician does not have a needle in stock. Ms D asks for a suburban service
AQT90 FLEX | Kalibration durchgeführt, System OK. | Kalibration durchgeführt, System OK. | SWO-50585 | No | Calibration performed, system OK. | Calibration performed, system OK. | Test Kit | Consumables | D-dimer | No Barcode -Test kit
TCM4 |  |  | SWO-50774 | No
AQT90 FLEX |  | 收到的903-138 换上去后2168，怀疑质量有问题。需要再借903-138，636-574各一个。 | SWO-50789 | No |  | After receiving 903-138 on the back of 2168, there was a suspected quality problem. Another 903-138, 636-574 each needs to be borrowed. | Defect | HW | Incubator Module | Incubator, PCB | Cup Sensor
TCM4 | 重新插拔模块，做清洁正常。 | 重新插拔模块，做清洁正常。 | SWO-50803 | No | Replug the module and do clean it properly. | Replug the module and do clean it properly.
ABL80 FLEX BASIC | 更换新测试卡 | 更换新测试卡 | SWO-50865 | No | Replace a new test card | Replace a new test card
ABL80 FLEX CO-OX | Sensorkassette gewechselt | Sensorkassette gewechselt | SWO-50958 | No | Sensor cassette changed | Sensor cassette changed
ABL90 FLEX | Über Service Fließwege gespült => von SC zum Einlass und von SC zum Abfall => alles i.O. Spülung der Fließwege mit SC i.O. Alle Serviceprogramme im Fluid Transport i.O. Kurzeinweisung QC RiliBÄK-Auswertungstabelle | Über Service Fließwege gespült => von SC zum Einlass und von SC zum Abfall => alles i.O. Spülung der Fließwege mit SC i.O. Alle Serviceprogramme im Fluid Transport i.O. Kurzeinweisung QC RiliBÄK-Auswertungstabelle | SWO-51065 | No | Washed via service flow paths => from SC to inlet and from SC to waste => all i.O. flushing of the flow paths with SC i.O. All service programs in Fluid Transport in general briefing QC RiliBÄK evaluation table | Washed via service flow paths => from SC to inlet and from SC to waste => all i.O. flushing of the flow paths with SC i.O. All service programs in Fluid Transport in general briefing QC RiliBÄK evaluation table | Cleaned | Fluidic pathway | ft
ABL90 FLEX | Fehlersuche, Leckage erkannt->unter Service Clot detect Remove fehlgeschlagen, LP wurde bereits erneuert, Einlassdichtung überprüft->Dichtung im Einlassdichtungshalter nach unten verschoben, Einlassdichtung wurde aber bereits vom Kunden getauscht->Einlassdichtung getauscht (vorher eine manuelle Spülung durchgeführt), Clot detect Remove ok, LS Justage etc. bestanden, Startphase durchlaufen, Präanalytik und Tausch ELD mit Anwendern Station und Labor besprochen System-und Funktionskontrolle i.O. | Fehlersuche, Leckage erkannt->unter Service Clot detect Remove fehlgeschlagen, LP wurde bereits erneuert, Einlassdichtung überprüft->Dichtung im Einlassdichtungshalter nach unten verschoben, Einlassdichtung wurde aber bereits vom Kunden getauscht->Einlassdichtung getauscht (vorher eine manuelle Spülung durchgeführt), Clot detect Remove ok, LS Justage etc. bestanden, Startphase durchlaufen, Präanalytik und Tausch ELD mit Anwendern Station und Labor besprochen System-und Funktionskontrolle i.O. | SWO-50993 | No | Troubleshooting, leakage detected->under Service Clot detect Remove failed, LP has already been renewed, inlet seal checked->seal in the inlet seal holder moved down, but inlet seal has already been replaced by the customer->inlet seal exchanged (previously a manual flushing performed), Clot detect Remove ok, LS Adjustment etc. passed, start phase go through, pre-analysis and exchange ELD with users station and laboratory discussed system and function control i.O. | Troubleshooting, leakage detected->under Service Clot detect Remove failed, LP has already been renewed, inlet seal checked->seal in the inlet seal holder moved down, but inlet seal has already been replaced by the customer->inlet seal exchanged (previously a manual flushing performed), Clot detect Remove ok, LS Adjustment etc. passed, start phase go through, pre-analysis and exchange ELD with users station and laboratory discussed system and function control i.O. | Cleaned | Fluidic pathway
ABL90 | Checked system log. Multiple inhomogenous solution and rinse errros. | Scrub fluidics path with lukewarm water, reinstalled inlet connector gasket and assembly. | CF7ACA044E30B3C2 | No | Checked system log. Multiple inhomogenous solution and rinse errros. | Scrub fluidics path with lukewarm water, reinstalled inlet connector gasket and assembly. | Cleaned | Fluidic pathway | FT
ABL90 | status 2 code, swapped cf card | unable to install windows worked with sonny and rick | CF7AD2744E0524CD | No | status 2 code, swapped cf card | unable to install windows worked with sonny and rick | Defect | SW | Status Code 2
ABL90 FLEX PLUS | Software Version 3.1 MR5 neu installiert | Software Version 3.1 MR5 neu installiert | SWO-51055 | No | Software version 3.1 MR5 reinstalled | Software version 3.1 MR5 reinstalled | Defect | sw | database
ABL90 FLEX | 代替機設置し引き上げ |  | SWO-50831 | No | Installation of alternative machines and raising |  | Assisted Replacement | SP | FT
ABL825 FLEX |  |  | CF7ABD234E0A4A60 | No |  |  | Assisted Replacement | Electrode | Membrane | pCO2
ABL835 FLEX | 各リンス エアバルブ動作確認し以上ありませんでした。別のリンスボトルを取り付けた所正常にリンスが掛かりましたので、取り付けていたリンスボトルが原因と思われます。 | 各リンス エアバルブ動作確認し以上ありませんでした。別のリンスボトルを取り付けた所正常にリンスが掛かりましたので、取り付けていたリンスボトルが原因と思われます。 | SWO-51073 | No | Each rinse air valve operation was not more than confirmed. Because the rinse hung normally where another rinse bottle was installed, it seems to be the cause of the rinse bottle which had been installed. | Each rinse air valve operation was not more than confirmed. Because the rinse hung normally where another rinse bottle was installed, it seems to be the cause of the rinse bottle which had been installed. | Assisted Replacement | Solution bottle | Solution bottle | Rinse
ABL820 FLEX | Oxi liquid sensor and lower liquid sensor failing during ls adjust | Flushed air line in fluidic module. Replaced sealing ring on pH electrode and replaced blocked nipple on inlet module. | CF7AB7BE4E3288A8 | No | Oxi liquid sensor and lower liquid sensor failing during ls adjust | Flushed air line in fluidic module. Replaced sealing ring on pH electrode and replaced blocked nipple on inlet module. | Cleaned | Fluidic pathway | Fluidic manifold | Air vent
ABL835 FLEX |  |  | CF7AB6364E01A64F | Yes |  |  | Re classified RMED | Logistical upSWO
ABL825 FLEX | Gerät überprüft, E1001 (PKW) übergangsweise erneuert | Gerät überprüft, E1001 (PKW) übergangsweise erneuert | SWO-51067 | No | Device checked, E1001 (car) temporarily renewed | Device checked, E1001 (car) temporarily renewed | Assisted Replacement | Electrode | Electrode | Reference
ABL825 FLEX | Elektrode erneuert | Elektrode erneuert | SWO-51068 | No | Electrode renewed | Electrode renewed | Assisted Replacement | Electrode | Electrode | pH
ABL825 FLEX | Viewed logs for errors. Ran service and could not adjust hemolyzer back in range. | Replaced the hemolyzer. | CF7A9E094E3155EA | No | Viewed logs for errors. Ran service and could not adjust hemolyzer back in range. | Replaced the hemolyzer. | Defect | HW | Oxi module | Hemolyzer
ABL835 FLEX | Recurring pH sensitivity errors. Sensitivity varies between 91 and 93%. | Replaced pH electrode. Performed maintenance due. | CF7A9CF44E0BDEC5 | No | Recurring pH sensitivity errors. Sensitivity varies between 91 and 93%. | Replaced pH electrode. Performed maintenance due. | Assisted Replacement | Electrode | Electrode | pH
ABL800 BASIC | Lactate membrane replace & run 2-point calibration .calibration pass. | Lactate membrane replace & run 2-point calibration .calibration pass. | SWO-51921 | No | Lactate membrane replace & run 2-point calibration .calibration pass. | Lactate membrane replace & run 2-point calibration .calibration pass. | Assisted Replacement | Electrode | Membrane | Lac
AQT90 FLEX | Fremdkörper entfernt, Probenhalter ersetzt, Gerät startet nicht vollständig, Software neu aufgespielt, Systemüberprüfung - Fehlermeldung bei Flow Temperatur, im Servicemenü Cup Temperatur Abgleich durchgeführt - Temperatur zu niedrig bei 86,540, Fehlersuche durchgeführt, E- Teil aus der Zentrale besorgt, Cup Dryer PCB ersetzt, Fehler bleibt, Gerät mit zur Werkstatt, Kunden informiert | Herr Moser meldet das ein Rädchen an der Probenhalterung abgebrochen sei. Herr Moser wünscht eine Instandsetzung. | SWO-51281 | No | Foreign body removed, sample holder replaced, device does not start completely, software replayed, system check - error message at flow temperature, performed in the service menu Cup temperature matching - temperature too low at 86,540, troubleshooting performed , E- part from the headquarters concerned, Cup Dryer PCB replaced, error remains, device with to the workshop, customer informed | Mr Moser reports that a cog at the sample holder has been broken off. Mr Moser wishes to be repaired. | Revisit required | Revisit required | Revisit required | Revisit required | Finally Main replaced
ABL80 FLEX BASIC | 更换新测试卡后正常 | 更换新测试卡后正常 | SWO-51309 | No | Normal after replacing a new test card | Normal after replacing a new test card
ABL80 FLEX BASIC | 更换新测试卡 | 更换新测试卡 | SWO-51203 | No | Replace a new test card | Replace a new test card
ABL80 FLEX BASIC | 更换测试卡正常 | 更换测试卡正常 | SWO-51293 | No | Replace the test card ok | Replace the test card ok
ABL80 FLEX BASIC | 更换测试卡正常。 | 更换测试卡正常。 | SWO-51991 | No | The replacement test card is OK. | The replacement test card is OK.
ABL80 FLEX BASIC | 更换测试卡正常。 | 更换测试卡正常。 | SWO-51992 | Yes | The replacement test card is OK. | The replacement test card is OK. | Re classified RMED | Double upSWO
ABL80 FLEX BASIC | 更换新测试卡后正常 | 更换新测试卡后正常 | SWO-51307 | No | Normal after replacing a new test card | Normal after replacing a new test card
ABL80 FLEX BASIC | 更换新测试卡 | 更换新测试卡 | SWO-51165 | No | Replace a new test card | Replace a new test card
ABL80 FLEX CO-OX | telef. Rücksprache- Kunde hat LP erneut gewechselt bei Besuch vor Ort war mit erneuern LP Fehler behoben, ABL voll einsatzbereit Überprüfung Daten PCO2 Fehler bei letzten LP kontinuierlich Rücksprache mit Anwender bei erneutem Auftreten Fehlerausdruck wir RDeu zugeschickt | telef. Rücksprache- Kunde hat LP erneut gewechselt bei Besuch vor Ort war mit erneuern LP Fehler behoben, ABL voll einsatzbereit Überprüfung Daten PCO2 Fehler bei letzten LP kontinuierlich Rücksprache mit Anwender bei erneutem Auftreten Fehlerausdruck wir RDeu zugeschickt | SWO-51194 | No | telef. Consultation- Customer has changed LP again during visit on site was fixed with renewing LP error, ABL fully ready to check data PCO2 error at last LP continuous consultation with user in case of recurrence error expression we sent to RDeu | telef. Consultation- Customer has changed LP again during visit on site was fixed with renewing LP error, ABL fully ready to check data PCO2 error at last LP continuous consultation with user in case of recurrence error expression we sent to RDeu
ABL90 FLEX | Einlass überprüft, Nadelhalter arretiert nicht mehr im Einlassbügel => neuen Dichtungshalter montiert. Alle Fließweg bezogenen Serviceprogramme absolviert, alle OK. Neustart, (1h-) Startphase, Kal Fehler, pCO2 Sens. weit unterhalb der Grenze => SC erneuert, Startphase, Kal und QC Messung | Einlass überprüft, Nadelhalter arretiert nicht mehr im Einlassbügel => neuen Dichtungshalter montiert. Alle Fließweg bezogenen Serviceprogramme absolviert, alle OK. Neustart, (1h-) Startphase, Kal Fehler, pCO2 Sens. weit unterhalb der Grenze => SC erneuert, Startphase, Kal und QC Messung | SWO-51228 | No | Inlet checked, needle holder no longer locks in the inlet bracket => new seal holder mounted. Completes all flow path-related service programs, all OK. Restart, (1h)Start Phase, Kal Error, pCO2 Sens. far below the limit => SC renewed, start phase, Kal and QC measurement | Inlet checked, needle holder no longer locks in the inlet bracket => new seal holder mounted. Completes all flow path-related service programs, all OK. Restart, (1h)Start Phase, Kal Error, pCO2 Sens. far below the limit => SC renewed, start phase, Kal and QC measurement | User Error | incorrect inlet gasket replacement
ABL90 FLEX | Kal. und QC durchgeführt Level A: 23.8 kPa, Level B: 40,1 kPa, Level C:1,11 kPa Probe 161: 12,1 kPa Probe 162: 14,1 kPa Als Pat.-Messung durchgeführt --> Temperaturkorrektion funktioniert nicht ( klare Flüssigkeit ) | Kal. und QC durchgeführt Level A: 23.8 kPa; Level B: 40,1 kPa; Level C:1,11 kPa Probe 161: 12,1 kPa Probe 162: 14,1 kPa Als Pat.-Messung durchgeführt --> Temperaturkorrektion funktioniert nicht ( klare Flüssigkeit ) | SWO-51143 | No | Kal. and QC performed level A: 23.8 kPa; Level B: 40.1 kPa; Level C:1,11 kPa Sample 161: 12.1 kPa Sample 162: 14.1 kPa Performed as Pat. Measurement --> Temperature correction does not work ( clear liquid ) | Kal. and QC performed level A: 23.8 kPa; Level B: 40.1 kPa; Level C:1,11 kPa Sample 161: 12.1 kPa Sample 162: 14.1 kPa Performed as Pat. Measurement --> Temperature correction does not work ( clear liquid ) | Performance | pO2 | ringversuch
ABL90 | Upon arrival, found the warning the customer the Sensor Cassette has expired its life time. Also, found fluid transport errors 1351, 1063 and, 1324. | Before replacing a new sensor cassette, I performed a complete back-flush of the fluid transport. Replaced a new inlet gasket, inlet probe and, the inlet connector gasket. Found the vacuum value failing in service, performed the fluid transport check, passing within specifications. Note: As per customer's request, I changed time setting to military time on both ABL90 analyzers at this facilities. | CF7C1F3FC45E1A9E | No | Upon arrival, found the warning the customer the Sensor Cassette has expired its life time. Also, found fluid transport errors 1351, 1063 and, 1324. | Before replacing a new sensor cassette, I performed a complete back-flush of the fluid transport. Replaced a new inlet gasket, inlet probe and, the inlet connector gasket. Found the vacuum value failing in service, performed the fluid transport check, passing within specifications. Note: As per customer's request, I changed time setting to military time on both ABL90 analyzers at this facilities. | Cleaned | Fluidic pathway | FT
ABL90 FLEX | Fließwege gespült, Lagerhaltung und Verfallsdaten besprochen |  | SWO-51316 | No | Flow paths rinsed, storage and expiry dates discussed |  | Cleaned | Fluidic pathway | ft
ABL90 | a cf card was shipped from westlake with 3.1MR5 software | unable to correct | CF7C09CBC46048F4 | No | a cf card was shipped from westlake with 3.1MR5 software | unable to correct | Unknown | CF-Card
ABL90 FLEX | 更换新测试卡后正常 | 更换新测试卡后正常 | SWO-51314 | No | Normal after replacing a new test card | Normal after replacing a new test card | Assisted Replacement | sc | pO2
ABL90 FLEX | QAキット交換 | QAキット交換 | SWO-51101 | No | QA Kit Replacement | QA Kit Replacement | Defect | HW | tubes
ABL90 FLEX | 更换新测试卡 | 更换新测试卡 | SWO-51169 | No | Replace a new test card | Replace a new test card | Assisted Replacement | SC | Unknown
ABL90 FLEX | 更换924-816一个后正常 | 更换924-816一个后正常 | SWO-51394 | No | Replace 924-816 one after normal | Replace 924-816 one after normal | Defect | HW | inlet
AQT90 FLEX | 更换903-364一个正常。 | 更换903-364一个正常。 | SWO-51193 | No | Replace 903-364 one normal. | Replace 903-364 one normal. | Defect | HW | Process Module | Optical Unit | 124
AQT90 FLEX | 更换定标卡944-258 | 更换定标卡944-258 | SWO-50917 | No | Replace the calibration card 944-258 | Replace the calibration card 944-258 | Assisted Replacement | Cartridge | NT-proBNP | Calibration
AQT90 FLEX |  |  | SWO-51133 | No |  |  | Defect | HW | Process Module | Optical Unit | 135
ABL827 FLEX |  |  | CF7E9058B00B3F8C | No |  |  | Assisted Replacement | Electrode | Membrane | Crea
ABL825 FLEX |  |  | CF7BF3EAC44B603F | No |  |  | Defect | HW | Oxi module | Neon number low
ABL815 FLEX | Monitor 910-320 1129 3288 testweise erneuert. PCO2 Elektrode neu membraniert, Elektrodenkontakte gereinigt und nachjustiert. Konfiguration gesichert. | Monitor 910-320 1129 3288 testweise erneuert. PCO2 Elektrode neu membraniert, Elektrodenkontakte gereinigt und nachjustiert. Konfiguration gesichert. | SWO-51287 | No | Monitor 910-320 1129 3288 renewed on a test basis. PCO2 electrode re-membraneed, electrode contacts cleaned and readjusted. Configuration backed up. | Monitor 910-320 1129 3288 renewed on a test basis. PCO2 electrode re-membraneed, electrode contacts cleaned and readjusted. Configuration backed up. | Defect | HW | DMS_PC | Monitor
ABL835 FLEX | ABL war bei Eintreffen Messbereit und alle Parameter sowie Ampel auf grünem Status. Man konnte aber in einigen Kals eine Schwankung der Driften bei Lac erkennen. Nullpunktkalibrierung durchgeführt und Lactat neu membraniert. Mehrere 1 und 2 Punkt Kalibrierungen durchgeführt. QCs für alle vorhandenen Level durchgeführt. ABL wieder Messbereit. | ABL war bei Eintreffen Messbereit und alle Parameter sowie Ampel auf grünem Status. Man konnte aber in einigen Kals eine Schwankung der Driften bei Lac erkennen. Nullpunktkalibrierung durchgeführt und Lactat neu membraniert. Mehrere 1 und 2 Punkt Kalibrierungen durchgeführt. QCs für alle vorhandenen Level durchgeführt. ABL wieder Messbereit. | SWO-51246 | No | ABL was ready to measure when it arrived and all parameters as well as traffic lights were on green status. However, one could see in some kals a fluctuation of the drifts at Lac. zero-point calibration and lactate re-membrane. Several 1 and 2 point calibrations performed. QCs for all existing levels. ABL ready to measure again. | ABL was ready to measure when it arrived and all parameters as well as traffic lights were on green status. However, one could see in some kals a fluctuation of the drifts at Lac. zero-point calibration and lactate re-membrane. Several 1 and 2 point calibrations performed. QCs for all existing levels. ABL ready to measure again. | Assisted Replacement | Electrode | Membrane | Lac
AQT90 FLEX |  |  | CF7C12DBC4425844 | No |  |  | Setup | SW | Analyzer control Module | Auto Backup
AQT90 FLEX | Cup Dryer PCB ern., ohne Erfolg - Inlet PCB ern., ohne Erfolg, E-Teil bestellt | Cup Dryer PCB ern., ohne Erfolg - Inlet PCB ern., ohne Erfolg, E-Teil bestellt | SWO-53084 | No | Cup Dryer PCB ern., without success - Inlet PCB ern., without success, E-part ordered | Cup Dryer PCB ern., without success - Inlet PCB ern., without success, E-part ordered | Revisit required | Revisit required | Revisit required | Revisit required | Finally Main replaced
ABL80 FLEX BASIC | 更换新的试剂包和新的测试卡后，定标未通过 | 更换新测试卡后正常 | SWO-52279 | No | The calibration failed after replacing the new reagent kit and the new test card | Normal after replacing a new test card |  | Double
ABL80 FLEX BASIC | 更换新测试卡后正常 | 更换新测试卡后正常 | SWO-52287 | No | Normal after replacing a new test card | Normal after replacing a new test card |  | Double
ABL80 FLEX BASIC | 更换测试卡正常。 | 更换测试卡 | SWO-51509 | No | The replacement test card is OK. | Replace the test card
ABL80 FLEX BASIC | 更换测试卡正常。 | 更换测试卡 | SWO-51510 | Yes | The replacement test card is OK. | Replace the test card | Re classified RMED | Triple reg
ABL80 FLEX BASIC | 更换测试卡 | 更换测试卡 | SWO-51511 | Yes | Replace the test card | Replace the test card | Re classified RMED | Triple reg
ABL80 FLEX BASIC | 更换新测试卡后未修复，显示温度-A | 更换新测试卡后正常 | SWO-51310 | No | Unrepaired after replacing new test card, display temperature-A | Normal after replacing a new test card
ABL80 FLEX BASIC | 更换新测试卡后正常 | 更换新测试卡后正常 | SWO-51312 | No | Normal after replacing a new test card | Normal after replacing a new test card
ABL80 OSM | Flushed entire system completely.Passed Service Test, but still fails on Blood results. | Replaced Hemolyzer Assy. | CF980A02E839295F | No | Flushed entire system completely. Passed Service Test, but still fails on Blood results. | Replaced Hemolyzer Assy.
ABL80 |  |  | CF8D443D381CB2CA | No
ABL90 | Examined data logs. Examined solution pack and sensor cassette. | Reinstalled solution pack and sensor cassette. Flushed instrument, as per guided troubleshooting procedure. Performed fluid transport calibrations, in service menu. Verified operation of ABL90. | CF7D438C3A315E61 | No | Examined data logs. Examined solution pack and sensor cassette. | Reinstalled solution pack and sensor cassette. Flushed instrument, as per guided troubleshooting procedure. Performed fluid transport calibrations, in service menu. Verified operation of ABL90. | Cleaned | Fluidic pathway | FT
ABL90 | Examined data logs. Examined sensor cassette and solution pack. | Performed guided troubleshooting. Reinstalled solution pack and sensor cassette. Flushed instrument, as per guided troubleshooting. Verified operation of ABL90. | CF7D3FC33A085E5E | No | Examined data logs. Examined sensor cassette and solution pack. | Performed guided troubleshooting. Reinstalled solution pack and sensor cassette. Flushed instrument, as per guided troubleshooting. Verified operation of ABL90. | Cleaned | Fluidic pathway | FT
ABL90 FLEX PLUS | Analysator-SW neu inst. --> Fehler bleibt bestehen --> Speicherkarte formatiert, XP und Analystor-SW inst. SP und SC neu eingesetzt --> Neustart, Konfig. neu eingespielt Kal. und QC bei Glu und Lac. fehlerhaft --> Station ist informiert | Analysator-SW neu inst. --> Fehler bleibt bestehen --> Speicherkarte formatiert, XP und Analystor-SW inst. SP und SC neu eingesetzt --> Neustart, Konfig. neu eingespielt Kal. und QC bei Glu und Lac. fehlerhaft --> Station ist informiert | SWO-51436 | No | Analyzer-SW new inst. --> Error persists --> Memory card formatted, XP and Analystor-SW inst. SP and SC re-installed --> Restart, Config. re-recorded Kal. and QC at Glu and Lac. faulty --> station is informed | Analyzer-SW new inst. --> Error persists --> Memory card formatted, XP and Analystor-SW inst. SP and SC re-installed --> Restart, Config. re-recorded Kal. and QC at Glu and Lac. faulty --> station is informed | Defect | sw | Status Code 2
ABL90 FLEX PLUS | Netzwerkadressen eingegeben, LIS KIS eingerichtet, Datenkommunikation i.O. | Netzwerkadressen eingegeben, LIS KIS eingerichtet, Datenkommunikation i.O. | SWO-51464 | No | Network addresses entered, LIS KIS set up, data communication in the entireion | Network addresses entered, LIS KIS set up, data communication in the entireion | Setup | HIS/LIS
TCM4 |  |  | SWO-51296 | No
ABL90 FLEX | 更换测试卡正常。 | 更换测试卡正常。 | SWO-51324 | No | The replacement test card is OK. | The replacement test card is OK. | Assisted Replacement | sc | Lac
ABL90 FLEX | 更换测试卡，试剂包正常。 | 更换测试卡，试剂包正常。 | SWO-51358 | No | Replace the test card and the reagent package is OK. | Replace the test card and the reagent package is OK. | Assisted Replacement | SC | Wetability
ABL90 | SE HACE PRUEBAS COMPARATIVAS, SE RECOGEN DATOS PARA TECHLINE, SE CAMBIA HEMOLIZADOR, SE VERIFICA QUE LOS CONTROLES DE OXI YA NO DAN BAJOS. SE COMPRUEBA CON CAL Y QC Y MUESTRAS CONPARATIVAS | SE HACE PRUEBAS COMPARATIVAS, SE RECOGEN DATOS PARA TECHLINE, SE CAMBIA HEMOLIZADOR, SE VERIFICA QUE LOS CONTROLES DE OXI YA NO DAN BAJOS. SE COMPRUEBA CON CAL Y QC Y MUESTRAS CONPARATIVAS | CF7D53573A3577E1 | No | COMPARATIC TESTS ARE MADE, DATA IS COLLECTED FOR TECHLINE, HEMOLIZADOR IS CHANGED, OXI CONTROLS ARE VERIFIED NOW LOW. CHECK WITH CAL AND QC AND CONPARATIVE SAMPLES | COMPARATIC TESTS ARE MADE, DATA IS COLLECTED FOR TECHLINE, HEMOLIZADOR IS CHANGED, OXI CONTROLS ARE VERIFIED NOW LOW. CHECK WITH CAL AND QC AND CONPARATIVE SAMPLES | Defect | HW | Hemolyzer
ABL800 BASIC | Gerät aufgestellt, Startphase durchgeführt, Gerät konfiguriert, Leihgerät SN: 754R0093N008 mitgenommen, Nullpunktkalibration im Serviceprogramm durchgeführt, Kals und QCs durchgeführt. | Gerät aufgestellt, Startphase durchgeführt, Gerät konfiguriert, Leihgerät SN: 754R0093N008 mitgenommen, Nullpunktkalibration im Serviceprogramm durchgeführt, Kals und QCs durchgeführt. | SWO-51496 | No | Device set up, start-up phase performed, device configured, loaner SN: 754R0093N008 taken, zero-point calibration performed in the service program, Kals and QCs performed. | Device set up, start-up phase performed, device configured, loaner SN: 754R0093N008 taken, zero-point calibration performed in the service program, Kals and QCs performed. | Assisted Replacement | Electrode | Electrode | pCO2
ABL825 FLEX |  |  | CF7D5B453A2CF0C2 | Yes |  |  | Re classified RMED | Logistical upSWO
ABL825 FLEX | Fluidicmodul testweise ersetzt aus GS17 Rücknahmegerät SLK Klinik Plattenwald Kreißsaal. Alle Schlauchverbindungen auf Beschädigung überprüft. Alle Ventile getestet und optisch kontrolliert. Diverse Verbindungsschläuche erneuert. Kal. 1+2 und Spüllösung erneuert. | Fluidicmodul testweise ersetzt aus GS17 Rücknahmegerät SLK Klinik Plattenwald Kreißsaal. Alle Schlauchverbindungen auf Beschädigung überprüft. Alle Ventile getestet und optisch kontrolliert. Diverse Verbindungsschläuche erneuert. Kal. 1+2 und Spüllösung erneuert. | SWO-51494 | No | Fluidicmodul tested replaced from GS17 readmission device SLK Klinik Plattenwald Kreißsaal. All hose connections checked for damage. All valves tested and optically controlled. Various connecting hoses renewed. Kal. 1+2 and rinsing solution renewed. | Fluidicmodul tested replaced from GS17 readmission device SLK Klinik Plattenwald Kreißsaal. All hose connections checked for damage. All valves tested and optically controlled. Various connecting hoses renewed. Kal. 1+2 and rinsing solution renewed. | Defect | HW | Fluidic manifold | Fluidic manifold
ABL825 FLEX | Gerät überprüft, pH BG Modul R2242N003 erneuert, Ref. Elektrode zurückgetauscht | Gerät überprüft, pH BG Modul R2242N003 erneuert, Ref. Elektrode zurückgetauscht | SWO-51528 | No | Device checked, pH BG module R2242N003 renewed, ref. Electrode replaced | Device checked, pH BG module R2242N003 renewed, ref. Electrode replaced | Defect | HW | pH/BG Module | ph/BG Module
ABL835 FLEX | Customer has tried different lots of membranes but electrode does not remain stable. | Had replacement electrode shipped to customer. | CF7D7B653A077F72 | Yes | Customer has tried different lots of membranes but electrode does not remain stable. | Had replacement electrode shipped to customer. | Re classified RMED | Logistical upSWO
ABL835 FLEX | PAUL GARDNER CAN SEE THE OTHER 7 ABL 800'S JUST NOT THIS ONE. | WENT INTO CONTROL PANEL, AND CHANGED THE FIREWALL EXCEPTIONS. I ADDED THE HOST AND CLICKED THE NETOP HOST ENCRYPTION. | CF7D38783A5A1012 | No | PAUL GARDNER CAN SEE THE OTHER 7 ABL 800'S JUST NOT THIS ONE. | WENT INTO CONTROL PANEL, AND CHANGED THE FIREWALL EXCEPTIONS. I ADDED THE HOST AND CLICKED THE NETOP HOST ENCRYPTION. | Setup | HIS/LIS
ABL825 FLEX | SWV 6.15 wiederholt installiert. PH- Elektrode wegen zuvor aufgetretene Instabille Messungen mit Proteinentferner gereinigt. | SWV 6.15 wiederholt installiert. PH- Elektrode wegen zuvor aufgetretene Instabille Messungen mit Proteinentferner gereinigt. | SWO-51468 | Yes | SWV 6.15 is installed repeatedly. PH electrode cleaned with protein remover due to previously occurring unstable measurements. | SWV 6.15 is installed repeatedly. PH electrode cleaned with protein remover due to previously occurring unstable measurements. | Re classified RMED | FAN / TU
AQT90 FLEX | AQT 90 838R0068N010 aufgestellt konfiguriert, kalibriert, QC durchgeführt | Fremdkörper entfernt, Probenhalter ersetzt, Gerät startet nicht vollständig, Software neu aufgespielt, Systemüberprüfung - Fehlermeldung bei Flow Temperatur, im Servicemenü Cup Temperatur Abgleich durchgeführt - Temperatur zu niedrig bei 86,540, Fehlersuche durchgeführt, E- Teil aus der Zentrale besorgt, Cup Dryer PCB ersetzt, Fehler bleibt, Gerät mit zur Werkstatt, Kunden informiert | SWO-51533 | No | AQT 90 838R0068N010 configured, calibrated, QC performed | Foreign body removed, sample holder replaced, device does not start completely, software replayed, system check - error message at flow temperature, performed in the service menu Cup temperature matching - temperature too low at 86,540, troubleshooting performed , E- part from the headquarters concerned, Cup Dryer PCB replaced, error remains, device with to the workshop, customer informed | Revisit required | Revisit required | Revisit required | Revisit required | Finally Main replaced
ABL80 FLEX BASIC | 连续更换2张测试卡后正常 | 连续更换2张测试卡后正常 | SWO-52484 | No | Normal after 2 consecutive test cards | Normal after 2 consecutive test cards
ABL80 FLEX BASIC | 更换903-334，902-972各一个正常。 | 更换903-334，902-972各一个正常。 | SWO-51512 | No | Replace 903-334, 902-972 each normal. | Replace 903-334, 902-972 each normal.
ABL80 FLEX BASIC | 更换两张测试卡正常。 | 更换两张测试卡正常。 | SWO-51566 | No | Replacing the two test cards is ok. | Replacing the two test cards is ok.
ABL80 CO-OX | Possible fluid blockage error. | Replaced the hemolyzer Spectrometer Performed THB calibration. | CF7EC934B01F80A5 | No | Possible fluid blockage error. | Replaced the hemolyzer Spectrometer Performed THB calibration.
ABL90 FLEX | Remplacé l'hémolyseur + l'électrovanne de oxymétrie e.o Calibration e.o contrôle de qualité e.o | Remplacé l'hémolyseur + l'électrovanne de oxymétrie e.o Calibration e.o contrôle de qualité e.o | SWO-51700 | No | Replaced the hemolyser - the electrovane of oxymetry e.o calibration e.o quality control e.o | Replaced the hemolyser - the electrovane of oxymetry e.o calibration e.o quality control e.o | Defect | HW | Hemolyzer
ABL90 | While in process of enabling Comm log to capture LIS trace, was getting error message An Unexpected problem has occurred. application cannot continue executing . When error closed, Kira shuts down. Must reboot. This happened several times, even when running samples. | Could not resolve issue. Will return in AM. Return to site. Analyzer has been stable overnight and CO is going out as COHB and not DOHB. Will leave comm log enabled and turn off one day next week. Also will create techline case. Worked on Radiance server and have set up NetOp connectivity fro client on server. | CF7ED02AB046746F | No | While in process of enabling Comm log to capture LIS trace, was getting error message An Unexpected problem has occurred. application cannot continue executing . When error closed, Kira shuts down. Must reboot. This happened several times, even when running samples. | Could not resolve issue. Will return in AM. Return to site. Analyzer has been stable overnight and CO is going out as COHB and not DOHB. Will leave comm log enabled and turn off one day next week. Also will create techline case. Worked on Radiance server and have set up NetOp connectivity fro client on server. | Setup | HIS/LIS | parameter transmit
ABL90 | Visited with a new solution pack from the SIUH-North division and, to replace it on this analyzer. | Replaced the new solution pack. Showed the customer how to run the ctHb calibration. | CF7E9CF0B03D1F32 | No | Visited with a new solution pack from the SIUH-North division and, to replace it on this analyzer. | Replaced the new solution pack. Showed the customer how to run the ctHb calibration. | Assisted Replacement | SP | pO2
ABL90 | USB Extender Cable to be at fault. | Replaced Cable.Unit still not working.I will update Richard Thurman. | CF956FAEFC3BEECB | No | USB Extender Cable to be at fault. | Replaced Cable.Unit still not working. I will update Richard Thurman. | Unknown | HIS/LIS | Wireless
ABL90 FLEX | 更换试剂包正常。 | 更换试剂包正常。 | SWO-51561 | No | The replacement kit is OK. | The replacement kit is OK. | Assisted Replacement | SP | FT
ABL825 FLEX |  |  | CF7E7271B051BC36 | No |  |  | Defect | HW | Oxi module | Hemolyzer
ABL835 FLEX |  | プローブ曲がりの為交換。 | SWO-52723 | No |  | Exchange for probe bending. | Assisted Replacement | HW | Inlet module | Probe bend
ABL820 FLEX | Check analyzer control board. Fuse F3 is blown. | Ordering fuse for replacement. | CF7EB674B04D2470 | No | Check analyzer control board. Fuse F3 is blown. | Ordering fuse for replacement. | Defect | HW | Analyzer Control PCB | F3
ABL825 FLEX | Proteinentfernung durchgeführt, Messkammern gereinigt, O2 Elektrode gereinigt und gebürstet, Dichtungsring der O2 Elektrode ersetzt, Nullpunktkalibration im Serviceprogramm durchgeführt, Kals und QCs durchgeführt. | Proteinentfernung durchgeführt, Messkammern gereinigt, O2 Elektrode gereinigt und gebürstet, Dichtungsring der O2 Elektrode ersetzt, Nullpunktkalibration im Serviceprogramm durchgeführt, Kals und QCs durchgeführt. | SWO-51705 | No | Protein removal performed, measuring chambers cleaned, O2 electrode cleaned and brushed, sealing ring replaced by The O2 electrode, zero-point calibration performed in the service program, Kals and QCs performed. | Protein removal performed, measuring chambers cleaned, O2 electrode cleaned and brushed, sealing ring replaced by The O2 electrode, zero-point calibration performed in the service program, Kals and QCs performed. | Adjustment | Electrode | Electrode | pO2
ABL825 FLEX | Einlassmudul, Elytemodul testweise getauscht und Waste Connektor erneuert. Schläuche insgesamt ( siehe vorherige SWO's) im Umfang einer Jahreswartung erneuert. Vermutlich wurde die Flüssigkeitsäule zwischen BG und Elytemodul zerrissen da Cl von dem Problem nicht betroffen ist. | Einlassmudul, Elytemodul testweise getauscht und Waste Connektor erneuert. Schläuche insgesamt ( siehe vorherige SWO's) im Umfang einer Jahreswartung erneuert. Vermutlich wurde die Flüssigkeitsäule zwischen BG und Elytemodul zerrissen da Cl von dem Problem nicht betroffen ist. | SWO-51946 | No | Inlet mudul, elyte module tested and Waste Connektor renewed. Total hoses (see previous SWO's) renewed in the amount of an annual maintenance. Presumably, the fluid column between BG and elyte modulus was torn apart as Cl is not affected by the problem. | Inlet mudul, elyte module tested and Waste Connektor renewed. Total hoses (see previous SWO's) renewed in the amount of an annual maintenance. Presumably, the fluid column between BG and elyte modulus was torn apart as Cl is not affected by the problem. | Defect | HW | Inlet module | Electrolytes
ABL825 FLEX | Erneuerung Oximetrie Modul (mit Halbjahreswartung durchgeführt) - tHb KAL und Funktionsüberprüfung etc. | Erneuerung Oximetrie Modul (mit Halbjahreswartung durchgeführt) - tHb KAL und Funktionsüberprüfung etc. | SWO-51693 | No | Renewal oximetry module (carried out with half-year maintenance) - tHb KAL and function check etc. | Renewal oximetry module (carried out with half-year maintenance) - tHb KAL and function check etc. | Defect | HW | Oxi module | Phase regulation
ABL815 FLEX | pCO 2 Elektrode mit neuer Charge membraniert und Elektrodenkontakte gereinigt. | pCO 2 Elektrode mit neuer Charge membraniert und Elektrodenkontakte gereinigt. | SWO-51947 | No | pCO 2 electrode with new batch membraneed and electrode contacts cleaned. | pCO 2 electrode with new batch membraneed and electrode contacts cleaned. | Cleaned | HW | pH/BG Module | Electrode connector
AQT90 FLEX |  |  | CF7EA639B05210CD | Yes |  |  | Re classified RMED
AQT90 FLEX |  |  | CF7E96CDB00110BE | Yes |  |  | Re classified RMED
ABL80 FLEX BASIC | 换测试卡正常。 | 换测试卡正常。 | SWO-53838 | No | Swap ping test card OK. | Swap ping test card OK.
ABL80 OSM | Examined solution pack and sensor cassette. | Replaced solution pack, with customer stock. Verified operation of ABL80 OSM. | CF7FD6A42633D26C | No | Examined solution pack and sensor cassette. | Replaced solution pack, with customer stock. Verified operation of ABL80 OSM.
ABL90 | I WAS ABLE TO SUCCESSFULLY RUN 3 LEVELS OF QC THE ONLY TARGETS ON THE DISPLAY THAT WOULD WORK WERE THE DATA LOG FOLDER AND THE ANALYZER STATUS TARGET | CALIBRATED THE TOUCH SCREEN AND UPGRADED THE SOFTWARE TO VERSION 3.1MR5.IF THE PROBLEM PERSISTS WE WILL REPLACE THE DISPLAY. | CF7FDAB52601AE89 | No | I WAS ABLE TO SUCCESSFULLY RUN 3 LEVELS OF QC THE ONLY TARGETS ON THE DISPLAY THAT WOULD WORK WERE THE DATA LOG FOLDER AND THE ANALYZER STATUS TARGET | CALIBRATED THE TOUCH SCREEN AND UPGRADED THE SOFTWARE TO VERSION 3.1MR5. IF THE PROBLEM PERSISTS WE WILL REPLACE THE DISPLAY. | Defect | SW | Freeze
ABL90 FLEX | 更换新测试卡后正常 | 更换新测试卡后正常 | SWO-52492 | No | Normal after replacing a new test card | Normal after replacing a new test card | Assisted Replacement | SC | Reference
ABL825 FLEX | Examined data logs. Examined Oxi module. Performed complete Oxi test. | Performed Halogen lamp adjustment. Performed service tHB calibration. Verified operation of ABL825. | CF7F9ED126013A46 | No | Examined data logs. Examined Oxi module. Performed Complete oxy test. | Performed Halogen lamp adjustment. Performed service tHB calibration. Verified operation of ABL825. | Adjustment | HW | Oxi module | Lamp unit
ABL825 FLEX |  |  | CF7FFD76265CF71A | No |  |  | Assisted Replacement | Electrode | Electrode | Reference
AQT90 FLEX |  |  | CF7FB369261F9FC7 | No |  |  | Defect | HW | Analyzer control Module | Cable | Screen
AQT90 FLEX | Tausch der Pipette und Neukalibrierung der Pipette auf die Cups und den Washblock. Keine weitere Beanstandung am Analysator. Service Dump gezogen. Aktuelle Nadelstiche mit der getauschten Nadel 655, Zähler auf Null gesetzt. QC o. K., Analysator funktionsbereit. | Der Kunde bittet um einen Service. | SWO-51873 | No | Swap the pipette and recalibrate the pipette on the cups and the washblock. No further complaint to the analyzer. Service Dump pulled. Current needle stitches with the exchanged needle 655, counter set to zero. QC o. K., analyzer ready for use. | The customer asks for a service. | Defect | HW | Aspiration Module | Bend needle | 4794
AQT90 FLEX | Netzteil ern., ohne Erfolg - neue CF Karte getestet - ohne Erfolg, Main Board bestellt | Cup Dryer PCB ern., ohne Erfolg - Inlet PCB ern., ohne Erfolg, E-Teil bestellt | SWO-53086 | No | Power supply ern., without success - new CF card tested - without success, main board ordered | Cup Dryer PCB ern., without success - Inlet PCB ern., without success, E-part ordered | Revisit required | Revisit required | Revisit required | Revisit required | Finally Main replaced
ABL90 FLEX | 更换新测试卡后正常 | 更换新测试卡后正常 | SWO-51977 | No | Normal after replacing a new test card | Normal after replacing a new test card | Assisted Replacement | SC | Wetability
ABL80 FLEX BASIC | 更换新测试卡后正常 | 更换新测试卡后正常 | SWO-51969 | No | Normal after replacing a new test card | Normal after replacing a new test card
ABL90 FLEX | 更换新测试卡后正常 | 更换新测试卡后正常 | SWO-51965 | No | Normal after replacing a new test card | Normal after replacing a new test card | Assisted Replacement | SC | Wetability
AQT90 FLEX | 更换903-171一个正常。 | 更换903-171一个正常。 | SWO-51958 | No | Replace 903-171 one normal. | Replace 903-171 one normal. | Defect | HW | Aspiration Module | Bend needle | 4794
ABL80 FLEX BASIC | 更换测试卡正常。 | 更换测试卡正常。 | SWO-52326 | No | The replacement test card is OK. | The replacement test card is OK.
ABL80 FLEX BASIC | 更换新测试卡后正常 | 更换新测试卡后正常 | SWO-52252 | No | Normal after replacing a new test card | Normal after replacing a new test card
ABL80 FLEX BASIC | 更换新测试卡后正常 | 更换新测试卡后正常 | SWO-52243 | No | Normal after replacing a new test card | Normal after replacing a new test card
ABL80 FLEX BASIC | 更换924-738一个后修复 | 更换924-738一个后修复 | SWO-51959 | No | Replace 924-738 one post-fix | Replace 924-738 one post-fix
ABL80 FLEX CO-OX | Reinigung wet section i.o Sensor SC80 Neu Installiert i.o Kalibration i.o QC i.o | Reinigung wet section i.o Sensor SC80 Neu Installiert i.o Kalibration i.o QC i.o | SWO-52175 | No | Cleaning wet section i.o Sensor SC80 New Installed i.o Calibration i.o QC i.o | Cleaning wet section i.o Sensor SC80 New Installed i.o Calibration i.o QC i.o
ABL80 FLEX BASIC | 更换新测试卡后正常 | 更换新测试卡后正常 | SWO-52085 | No | Normal after replacing a new test card | Normal after replacing a new test card
ABL80 FLEX CO-OX | kalibriert - ohne Erfolg, Sensor ersetzt, Initialisierung abgewartet, Proben gemessen, QC | kalibriert - ohne Erfolg, Sensor ersetzt, Initialisierung abgewartet, Proben gemessen, QC | SWO-52227 | No | calibrated - without success, sensor replaced, initialization awaited, samples measured, QC | calibrated - without success, sensor replaced, initialization awaited, samples measured, QC
ABL90 FLEX | Neue Einlassdichtung eingebaut. | Neue Einlassdichtung eingebaut. | SWO-52277 | No | New inlet seal installed. | New inlet seal installed. | Defect | HW | Inlet Gasket
ABL90 FLEX | Oxi Systemtest und Valvetest durchgeführt. Funktionen o.k. System gespült, Oxitubing und Einlassdichtungen erneuert. Material aus Jahreswartungskit entwendet. | Oxi Systemtest und Valvetest durchgeführt. Funktionen o.k. System gespült, Oxitubing und Einlassdichtungen erneuert. Material aus Jahreswartungskit entwendet. | SWO-52224 | No | Oxi system test and valve test. Functions o.k. system flushed, oxitubing and inlet seals renewed. material stolen from annual maintenance kit. | Oxi system test and valve test. Functions o.k. system flushed, oxitubing and inlet seals renewed. material stolen from annual maintenance kit. | Cleaned | Fluidic pathway
ABL90 FLEX | Einlassdichtung erneuert. | Einlassdichtung erneuert. | SWO-52274 | No | Inlet seal renewed. | Inlet seal renewed. | Defect | HW | Inlet Gasket
ABL90 FLEX | Flüssigkeitsstrecke gespült, Einlassdichtung gereinigt, --> Neustart Kal.-Fehler bei K --> SC erneuert Station überwacht die Initialisierung der SC | Flüssigkeitsstrecke gespült, Einlassdichtung gereinigt, --> Neustart Kal.-Fehler bei K --> SC erneuert Station überwacht die Initialisierung der SC | SWO-52140 | No | Liquid track flushed, inlet seal cleaned, --> Restart Cal.-Error at K --> SC renews station monitors the initialization of sc | Liquid track flushed, inlet seal cleaned, --> Restart Cal.-Error at K --> SC renews station monitors the initialization of sc | Assisted Replacement | SC | Wetability
ABL90 FLEX | Netzwerkdaten eingegeben, Serverdaten eingegeben, Neustart, Kal u. QC i.O. Verbindung überprüfen lassen - i.O. |  | SWO-52510 | No | Network data entered, server data entered, restart, Kal & QC i.O. connection checked - i.O. |  | Setup | HIS/LIS
ABL90 FLEX | Netzwerkdaten eingegeben, Serverdaten eingegeben, Neustart, Verbindung überprüfen lassen - i.O. |  | SWO-52526 | No | Network data entered, server data entered, restart, connection verified - i.O. |  | Setup | HIS/LIS
ABL90 FLEX | 更换891-051一个后修复 | 还没有修复，发现进样口针角弯，需要再借891-051一个。 | SWO-52050 | No | Replace 891-051 one post-fix | Has not been repaired, found the needle angle bend, need to borrow another 891-051 one. | Replaced | Probe | Bend
ABL90 FLEX | 还没有修复，发现进样口针角弯，需要再借891-051一个。 | 还没有修复，发现进样口针角弯，需要再借891-051一个。 | SWO-51966 | Yes | Has not been repaired, found the needle angle bend, need to borrow another 891-051 one. | Has not been repaired, found the needle angle bend, need to borrow another 891-051 one. | Re classified RMED | Double upSWO
AQT90 FLEX | 更换903-171 | 更换903-171 | SWO-52099 | No | Replace 903-171 | Replace 903-171 | Defect | HW | Aspiration Module | Needle Wash Unit | 1259
ABL837 FLEX | ｱﾝﾌﾟﾙｵｰﾌﾟﾅｰﾓｰﾀｰ交換。 | ｱﾝﾌﾟﾙｵｰﾌﾟﾅｰﾓｰﾀｰ交換。 | SWO-52237 | No | Ampoule opener motor replacement. | Ampoule opener motor replacement. | Defect | HW | AutoCheck Module | Stepper motor
ABL825 FLEX | Referenz Membran neu membraniert. Mehrere Kalibrierungen durchgeführt. QCs für alle Level durchgeführt. ABL wieder Messbereit. | Referenz Membran neu membraniert. Mehrere Kalibrierungen durchgeführt. QCs für alle Level durchgeführt. ABL wieder Messbereit. | SWO-52165 | No | Reference membrane re-membraneed. Multiple calibrations performed. QCs for all levels. ABL ready to measure again. | Reference membrane re-membraneed. Multiple calibrations performed. QCs for all levels. ABL ready to measure again. | Assisted Replacement | Electrode | Membrane | Reference
AQT90 FLEX | removed the tube | removed the tube and train them how to fix the tube properly so that this will not happen in future. | SWO-53200 | No | removed the tube | removed the tube and train them how to fix the tube properly so that this will not happen in future. | User Error | Sample measurement | Aspiration Module | Tube mounting
AQT90 FLEX | Main Board ern., Gerät abgeglichen und kalibriert, Blind Messungen durchgeführt OK, QC´s durchgef. | Cup Dryer PCB ern., ohne Erfolg - Inlet PCB ern., ohne Erfolg, E-Teil bestellt | SWO-53089 | No | Main board ern., device matched and calibrated, blind measurements performed OK, QC's through. | Cup Dryer PCB ern., without success - Inlet PCB ern., without success, E-part ordered | Defect | HW | Analyzer control Module | Main Board, PCB | Finally Main replaced
ABL80 FLEX BASIC | 连续更换两张测试卡后正常 | 连续更换两张测试卡后正常 | SWO-52297 | No | Normal after replacing two test cards in a row | Normal after replacing two test cards in a row
ABL80 FLEX BASIC | 更换测试卡正常。 | 更换测试卡正常。 | SWO-52472 | No | The replacement test card is OK. | The replacement test card is OK.
ABL80 FLEX BASIC | Software reinstalled. Data Cleaned up. Analyser serviced and Calibrated. Working fine. | Software reinstalled. Data Cleaned up. Analyser serviced and Calibrated. Working fine. | SWO-53167 | No | Software reinstalled. Data Cleaned up. Analyser serviced and Calibrated. Working fine. | Software reinstalled. Data Cleaned up. Analyser serviced and Calibrated. Working fine.
ABL90 FLEX | System überprüft Sensorkassette erneuert Kalibrationsfehler und QC durchgeführt | System überprüft Sensorkassette erneuert Kalibrationsfehler und QC durchgeführt | SWO-52305 | No | System checks sensor cartridge renewed calibration error and QC performed | System checks sensor cartridge renewed calibration error and QC performed | Assisted Replacement | SC | Wetability
ABL90 FLEX | Touchscreen versucht im Serviceprogramm zu kalibrieren, keine Verbesserung, Touchscreen reagiert immer schlechter, ABL zur weiteren Überprüfung in die GS mitgenommen. | Der Kunde bittet um einen Service und Rückruf. | SWO-52493 | No | Touchscreen tries to calibrate in the service program, no improvement, touchscreen reacts ever worse, ABL is taken to the GS for further verification. | The customer asks for a service and callback. | Unknown | Touch screen | Analyzer removed
ABL90 FLEX | Einlassdichtung und Einlasssonde erneuert. | Einlassdichtung und Einlasssonde erneuert. | SWO-52397 | No | Inlet seal and inlet probe renewed. | Inlet seal and inlet probe renewed. | Cleaned | Fluidic pathway
ABL90 | Went into Diagnostics ran several diag routines. I was able to clear all values in Red from diagnostics. | Diagnostics, replace SP. SP provided by customer. | CF84D2CCFE60158A | No | Went into Diagnostics ran several diag routines. I was able to clear all values in Red from diagnostics. | Diagnostics, replace SP. SP provided by customer. | Assisted Replacement | SP | FT
ABL90 | instrument in good condition | powered up , i set units and added qc data, ran thb cal | CF84AE74FE26A4A3 | Yes | instrument in good condition | powered up , i set units and added qc data, ran thb cal | Re classified RMED | Re install analyzer
TCM CombiM |  |  | SWO-52333 | No
ABL90 | Comprobar SP y SC, cambiar el SP y el SC, calibrar y probar con controles. | Cambiar el SC y el SP. | CF84B117FE0EC499 | No | Check SP and SC, change SP and SC, calibrate and test with controls. | Change the SC and SP. | Assisted Replacement | SC | Wetability
ABL90 | SE DESMONTA MODULO NO ENCONTRANDOSE NINGUN PROBLEMA EVIDENTE Y SE VUELVEN HACER PRUEBAS DE FUNCIONAMIENTO REPETITIVAS MEDIANTE PROGRAMAS DE SERVICIO TECNICO, SIENDO ESTAS CORRECTAS. SE MONTAN REACTIVOS HASTA SU ESTABILIZACION Y LISTO PARA FUNCIONAR. CALIBRADO Y VERIFICADO MEDIANTE Q.C. |  | CF850F98FE1A7D93 | No | DISMONTATES NOT FIND ANY OBVIOUS PROBLEM AND RETURN TO MAKE REPETITIVE OPERATION TESTS THROUGH TECHNICAL SERVICE PROGRAMS, BEING THESE CORRECT. REACTIVE TO YOUR STABILIZATION AND READY TO WORK. CALIBRATED AND VERIFIED BY Q.C. |  | Adjustment | Sensor Bay
AQT90 FLEX | 更换903-364一个和903-119一个正常。 | 更换903-364一个和903-119一个正常。 | SWO-51997 | No | Replace 903-364 one and 903-119 one normal. | Replace 903-364 one and 903-119 one normal. | Defect | HW | Process Module | Optical Unit | 341
AQT90 FLEX |  |  | SWO-52241 | No |  |  | Assisted Replacement | Cartridge | NT-proBNP | Calibration | 5257
AQT90 FLEX | 更换903-119一个正常。 | 更换903-119一个正常。 | SWO-51993 | No | Replace the 903-119 one normal. | Replace the 903-119 one normal. | Defect | HW | Aspiration Module | Bend needle
AQT90 FLEX | 更换903-031，903-032，580-098各一个正常。 | 更换903-031，903-032，580-098各一个正常。 | SWO-52292 | No | Replace 903-031, 903-032, 580-098 each normal. | Replace 903-031, 903-032, 580-098 each normal. | Defect | HW | Incubator Module | Cup lifter, stepper motor | 5297
AQT90 FLEX | 更换903-113一个后修复 | 更换903-113一个后修复 | SWO-51974 | No | Replace 903-113 one post-fix | Replace 903-113 one post-fix | Defect | HW | Aspiration Module | PCB | 48
ABL825 FLEX | AQC Modul justiert, Messkammer gereinigt, QC´s durchgef. | Technikereinsatz erforderlich | SWO-52368 | No | AQC module adjusted, measuring chamber cleaned, QC's through. | Technician stake required | Defect | HW | AutoCheck Module | unknown
ABL827 FLEX | checked data logs for errors, checked valves, liquid sensors and pumps, | removed modules and cleaned all valves, | CF853CF5FE3E94AF | No | checked data logs for errors, checked valves, liquid sensors and pumps, | removed modules and cleaned all valves, | Cleaned | Fluidic pathway | - | Valve
ABL80 FLEX BASIC | 更换测试卡正常。 | 更换测试卡正常。 | SWO-52596 | No | The replacement test card is OK. | The replacement test card is OK.
ABL80 FLEX BASIC |  |  | SWO-52569 | No
ABL80 FLEX BASIC | 更换新测试卡后正常 | 更换新测试卡后正常 | SWO-52585 | No | Normal after replacing a new test card | Normal after replacing a new test card
ABL80 FLEX BASIC | 更换新测试卡修复机器 | 更换新测试卡修复机器 | SWO-52634 | No | Replace the new test card repair machine | Replace the new test card repair machine
ABL80 FLEX CO-OX | Halbjahreswartung durchgeführt. Einlass der Sensor Kassette (945-722 Lot: 214740) defekt. Sensor Kassette erneuert, Gerät kalibriert und QC durchgeführt. | Halbjahreswartung durchgeführt. Einlass der Sensor Kassette (945-722 Lot: 214740) defekt. Sensor Kassette erneuert, Gerät kalibriert und QC durchgeführt. | SWO-52623 | Yes | half-yearly maintenance. Inlet of the sensor cassette (945-722 Lot: 214740) defective. Sensor cassette renewed, device calibrated and QC performed. | half-yearly maintenance. Inlet of the sensor cassette (945-722 Lot: 214740) defective. Sensor cassette renewed, device calibrated and QC performed. | Re classified RMED | PM
ABL80 CO-OX | Found extra line of code in ASTM.ini file and Device ID not set.Stephen McKeown sent wrong Astm file from RALS. | Enter corrected files at necessary locations. | CF957166FC13EECF | No | Found extra line of code in ASTM.ini file and Device ID not set. Stephen McKeown sent wrong Astm file from RALS. | Enter corrected files at necessary locations.
ABL80 FLEX BASIC | 换测试卡 | 换测试卡 | SWO-52744 | No | Swap test card | Swap test card
ABL90 FLEX | Fehlersuche SW aktualisiert V3.1MR5 Kalibrierung durchgeführt QC Werte pO2 sind mit den letzten Messung deutlich gesunken , nach mehrfachen Kal. im Bereich | Fehlersuche Überprüfung Datein (Kal, QC) Austausch SC ( da mit einsetzen der bisherigen SC der Fehler auftrat) anschl. Werte i.O. | SWO-52659 | No | Troubleshooting SW updated V3.1MR5 calibration performed QC values pO2 are significantly decreased with the last measurement, after multiple cal. in the area of | Troubleshooting Check File (Kal, QC) Replacement SC (since with insert the previous SC the error occurred) followed. Values in the | Assisted Replacement | SP | pO2
ABL90 | I rebooted the analyzer several times in an attempt to pass the start up screen. I tried to recover the data from the CF card and was unable to retrieve any backups or settings. After several failed attempts to reload the windows software, I determined that the CPU is faulty and will need replaced. | I replaced the CPU and booted the analyzer up with the original CF card. The analyzer booted up and all data and settings returned. | CF860B2E74256FBA | No | I rebooted the analyzer several times in an attempt to pass the start up screen. I tried to recover the data from the CF card and was unable to retrieve any backups or settings. After several failed attempts to reload the windows software, I determined that the CPU is faulty and will need replaced. | I replaced the CPU and booted the analyzer up with the original CF card. The analyzer booted up and all data and settings returned. | Defect | HW | CPU
ABL90 FLEX | Touchscreen am ABL ersetz, tHB Kal durchgeführt, ABL messbereit beim Kunden aufgestellt. | Display ausgebaut und erneuert, Scanner umgebaut, Touchscreen kalibriert und Werte gespeichert. Funktion OK. | SWO-52606 | No | Touchscreen on the ABL replaced, tHB Kal performed, ABL ready for measurement at the customer. | Display removed and renewed, scanner converted, touchscreen calibrated and values stored. Function OK. | Unknown | Touch screen | Display replaced next day
ABL90 | Chequeo de equipo, sustitución de sonda doblada, revisión general, limpieza de circuito neumatico, sustitución de SC y calibración, pasado qc. | Sustitución sonda y SC Revisión preventiva | CF860EC4746389C2 | Yes | Equipment check, bent probe replacement, general review, tire circuit cleaning, SC replacement and calibration, past qc. | Probe replacement and SC Preventive review | Re classified RMED | PM
ABL90 FLEX | 更换测试卡正常。 | 更换测试卡正常。 | SWO-52725 | No | The replacement test card is OK. | The replacement test card is OK. | Assisted Replacement | SC | Reference
AQT90 FLEX |  |  | SWO-52485 | No |  |  | Defect | HW | Process Module | Optical Unit | 131
ABL820 FLEX | temp errors, gas voltage errors, neon lamp errors and spectrometer errors | Replaced Oxi module PCB and fuse in AU unit. Reloaded AU software. Adjusted flex Q tray to specifications. | CF862FA47403DF16 | No | temp errors, gas voltage errors, neon lamp errors and spectrometer errors | Replaced Oxi module PCB and fuse in AU unit. Reloaded AU software. Adjusted flex Q tray to specifications. | Defect | HW | Oxi module | PCB
ABL815 FLEX | Investigate analzyer and take data logs. | Investigated analyzer and took data logs. Replaced K, and pO2 membrane from customer stock. Found no oring on bottom of pO2 membrane and replaced from extra stock in FSE kit. Found a bent connector from Met module under Lac electrode connector and fixed by starightening it. | CF85D65B7424DE6F | No | Investigate analzyer and take data logs. | Investigated analyzer and took data logs. Replaced K, and pO2 membrane from customer stock. Found no oring on bottom of pO2 membrane and replaced from extra stock in FSE kit. Found a bent connector from Met module under Lac electrode connector and fixed by starightening it. | Defect | HW | El/Met Module | Electrode connector
ABL825 FLEX | Dichtungsring pH Elektrode ersetzt, pH BG Modul zurückgetauscht 902-823R2242N003 eingebaut, 823R2278N002 zurück im Bestand, Software aktualisiert | Dichtungsring pH Elektrode ersetzt, pH BG Modul zurückgetauscht 902-823R2242N003 eingebaut, 823R2278N002 zurück im Bestand, Software aktualisiert | SWO-52834 | No | Sealing ring pH electrode replaced, pH BG module replaced 902-823R2242N003 built-in, 823R2278N002 back in stock, software updated | Sealing ring pH electrode replaced, pH BG module replaced 902-823R2242N003 built-in, 823R2278N002 back in stock, software updated | Assisted Replacement | Electrode | Electrode | pH | Sealing ring
AQT90 FLEX | Se realizo una limpieza interna del equipo y se sustituyo un kit de tubos anual. Se calibro LOTE nuevo de DIM, Se usaron QC´s nuevos y el equipo no pasa los Dimeros. queda pendiente de instrucciones del fabricante. |  | CF8C6D14C2332ADB | No | Internal cleaning of the equipment was performed and an annual tube kit is replaced. New DIM LOTE cadell is available, new QC's were used and the equipment does not pass the Dimeros. manufacturer's instructions are pending. |  | Cleaned | HW | Aspiration Module | Unknown
AQT90 FLEX | Kalibrierung eines neuen Testkits, keine neue Losnummer vorhanden. Testkit mit neuer Losnummer bestellt für Kunden. Anwender kalibrieren neue Losnummer und führen QC durch. Ist QC nicht mehr grenzwertig, ist kein Service mehr notwendig. Ist QC weiterhin grenzwertig, ist ein Tausch der Optischen Einheit notwendig. Service wird von Kunde angefordert. | Kalibrierung eines neuen Testkits, keine neue Losnummer vorhanden. Testkit mit neuer Losnummer bestellt für Kunden. Anwender kalibrieren neue Losnummer und führen QC durch. Ist QC nicht mehr grenzwertig, ist kein Service mehr notwendig. Ist QC weiterhin grenzwertig, ist ein Tausch der Optischen Einheit notwendig. Service wird von Kunde angefordert. | SWO-52825 | No | Calibration of a new test kit, no new lot number available. Test kit with new lot number ordered for customers. Users calibrate new lot number and perform QC. If QC is no longer borderline, no service is necessary. If QC continues to be borderline, an exchange of the optical unit is necessary. Service is requested by customer. | Calibration of a new test kit, no new lot number available. Test kit with new lot number ordered for customers. Users calibrate new lot number and perform QC. If QC is no longer borderline, no service is necessary. If QC continues to be borderline, an exchange of the optical unit is necessary. Service is requested by customer. | Adjustment | LQC | TnT-LQC | New LQC lot
ABL80 FLEX BASIC | 更换新试剂包修复机器 | 更换新试剂包修复机器 | SWO-52823 | No | Replace the new reagent kit repair machine | Replace the new reagent kit repair machine
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-54718 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL80 FLEX BASIC | refrigerated the sc for some time and reinstalled. got ready |  | SWO-53214 | No | refrigerated the sc for some time and reinstalled. got ready
ABL80 FLEX Basic version |  |  | CF8C18C4C24B487F | No
ABL90 FLEX | Eintrag für fixe Werte vorhanden, jedoch fehlerhaft, Buchstabendreher. Eintrag korrigiert und gespeichert. Neuen Lösungspack aktiviert, vorab aber aktuelle Software installiert. Tabelle mit fixen Zielwerten Anwender überlassen. Überprüfung der gesendeten QC Daten im LIS System, o.K. Analysator funktionsbereit. | Eintrag für fixe Werte vorhanden, jedoch fehlerhaft, Buchstabendreher. Eintrag korrigiert und gespeichert. Neuen Lösungspack aktiviert, vorab aber aktuelle Software installiert. Tabelle mit fixen Zielwerten Anwender überlassen. Überprüfung der gesendeten QC Daten im LIS System, o.K. Analysator funktionsbereit. | SWO-52951 | No | Fixed value entry exists, but incorrect, letter turner. entry corrected and saved. New solution pack activated, but up-to-date software installed in advance. Table with fixed target values. Checking the sent QC data in the LIS system, o.K. analyser is ready for use. | Fixed value entry exists, but incorrect, letter turner. entry corrected and saved. New solution pack activated, but up-to-date software installed in advance. Table with fixed target values. Checking the sent QC data in the LIS system, o.K. analyser is ready for use. | Setup | Fixed assigned values
ABL90 FLEX | Fließwege, Dichtungen, Kassette und LP überprüft und gereinigt, Schlauch Messkammer Oxi ersetzt, Einlaßdichtung ersetzt, LP erneut ersetzt, mehrfach Kal, QC | Fließwege, Dichtungen, Kassette und LP überprüft und gereinigt, Schlauch Messkammer Oxi ersetzt, Einlaßdichtung ersetzt, LP erneut ersetzt, mehrfach Kal, QC | SWO-52986 | No | Flow paths, seals, cassette and LP checked and cleaned, tube measuring chamber Oxi replaced, inlet seal replaced, LP replaced, multiple Kal, QC | Flow paths, seals, cassette and LP checked and cleaned, tube measuring chamber Oxi replaced, inlet seal replaced, LP replaced, multiple Kal, QC | Assisted Replacement | SP | pO2
ABL90 FLEX | Display ausgebaut und erneuert, Scanner umgebaut, Touchscreen kalibriert und Werte gespeichert. Funktion OK. | Display ausgebaut und erneuert, Scanner umgebaut, Touchscreen kalibriert und Werte gespeichert. Funktion OK. | SWO-52837 | No | Display removed and renewed, scanner converted, touchscreen calibrated and values stored. Function OK. | Display removed and renewed, scanner converted, touchscreen calibrated and values stored. Function OK. | Defect | HW | Display
ABL90 | Revisión del equipo. Comprobación del cassette de electrodos, pack de soluciones y sustitución del este último por fallo en el mismo. Revisión y comprobación del equipo con QC´s. | Cambio del SP. | CF874A71EA4C9DD2 | No | Equipment review. Checking the electrode cassette, solution pack and replacing the latter with failure in the latter. Review and check the equipment with QC's. | Switching the SP. | Assisted Replacement | SP | Unknown
ABL90 | Limpeza do circuito da amostra. Reinstalação do cartucho de soluções. Inicialização. Cópia de segurança dos dados e configuração. Actualização de software para a versão 3.1MR6.PT. Verificação do correcto funcionamento. |  | CF878D5AEA091277 | No | Cleaning the sample circuit. Reinstallation of the solution cartridge. Startup. Data backup and configuration. Software update for version 3.1MR6.PT. Checking the correct operation. |  | Assisted Replacement | SC | Wetability | MR6 solved problem
ABL835 FLEX |  |  | CF874C6FEA5D58D4 | No |  |  | Defect | HW | Oxi module | PCB
ABL827 FLEX | neue Einl.nadel eingebaut, Gerät kalibriert, | neue Einl.nadel eingebaut, Gerät kalibriert, | SWO-52909 | No | new single needle built-in, device calibrated, | new single needle built-in, device calibrated, | Defect | HW | Inlet module | Probe bend
ABL835 FLEX | Durchführung einer Proteinentfernung und Kalibrationen. Durchführung einer tHb Kalibration im Service und im Anwendermodus. Es zeigt sich, dass der Küvettenfaktor, durch Anwender zu hoch gemessen wurde. QC wieder zielwertig. Einlassdichtung verschmutzt, erneuert. Analysator messbereit. | Durchführung einer Proteinentfernung und Kalibrationen. Durchführung einer tHb Kalibration im Service und im Anwendermodus. Es zeigt sich, dass der Küvettenfaktor, durch Anwender zu hoch gemessen wurde. QC wieder zielwertig. Einlassdichtung verschmutzt, erneuert. Analysator messbereit. | SWO-52950 | No | Performing protein removal and calibration. Performing a tHb calibration in service and user mode. It turns out that the cuvette factor was measured too high by users. QC again targeted. Inlet seal dirty, renewed. Analyzer ready to measure. | Performing protein removal and calibration. Performing a tHb calibration in service and user mode. It turns out that the cuvette factor was measured too high by users. QC again targeted. Inlet seal dirty, renewed. Analyzer ready to measure. | User Error | HW | Oxi module | tHb cal
ABL825 FLEX | Gerät überprüft, Meßkammer und Elektrode gereinigt, pCO2 nochmals membraniert, | Gerät überprüft, Meßkammer und Elektrode gereinigt, pCO2 nochmals membraniert, | SWO-53069 | No | device checked, measuring chamber and electrode cleaned, pCO2 membraneed again, | device checked, measuring chamber and electrode cleaned, pCO2 membraneed again, | Assisted Replacement | Electrode | Membrane | pCO2
ABL800 BASIC | Checked the input voltage to analyser through UPS which was on the higher side, so asked them to use the voltage stabiliser. Restarted the analyser through voltage stabiliser & checked the touch screen, which was working properly. | Checked the input voltage to analyser through UPS which was on the higher side, so asked them to use the voltage stabiliser. Restarted the analyser through voltage stabiliser & checked the touch screen, which was working properly. | SWO-53456 | No | Checked the input voltage to analyser through UPS which was on the higher side, so asked them to use the voltage stabiliser. Restarted the analyser through voltage stabiliser & checked the touch screen, which was working properly. | Checked the input voltage to analyser through UPS which was on the higher side, so asked them to use the voltage stabiliser. Restarted the analyser through voltage stabiliser & checked the touch screen, which was working properly. | Unknown | HW | DMS_PC | Touch
TCM-COMBIM |  |  | CF877F42EA5CCA84 | No |  |  | Defect | HW | Power Supply | Power Supply
ABL80 OSM | oxy error on QC and blood samples | replaced the hemolyzer and tubings, replaced the pump head | CF887B31603A9A41 | No | oxy error on QC and blood samples | replaced the hemolyzer and tubings, replaced the pump head
ABL90 FLEX | Fehlersuche Überprüfung Datein (Kal, QC) Austausch SC ( da mit einsetzen der bisherigen SC der Fehler auftrat) anschl. Werte i.O. | Fehlersuche Überprüfung Datein (Kal, QC) Austausch SC ( da mit einsetzen der bisherigen SC der Fehler auftrat) anschl. Werte i.O. | SWO-53118 | No | Troubleshooting Check File (Kal, QC) Replacement SC (since with insert the previous SC the error occurred) followed. Values in the | Troubleshooting Check File (Kal, QC) Replacement SC (since with insert the previous SC the error occurred) followed. Values in the | Assisted Replacement | SC | pO2
ABL90 FLEX | Überprüfung Dateiarchiv Aktivitäten und Betriebsmittel Sitz von SC und LP überprüft => Kontakte zur Auslesung der LP-Chipdaten leicht verbogen => Kontakte nachjustiert Durchführung aller Fluid Serviceprogramme => Gerinnsel wird erkannt => Gerät gespült => keine Verstopfung feststellbar Neustart mit aktuell verwendeten LP scheitert kurz nach Start Austausch LP und Startphase i.O. Fehlerursache war schon vor Verwendung aktivierter LP | Überprüfung Dateiarchiv Aktivitäten und Betriebsmittel Sitz von SC und LP überprüft => Kontakte zur Auslesung der LP-Chipdaten leicht verbogen => Kontakte nachjustiert Durchführung aller Fluid Serviceprogramme => Gerinnsel wird erkannt => Gerät gespült => keine Verstopfung feststellbar Neustart mit aktuell verwendeten LP scheitert kurz nach Start Austausch LP und Startphase i.O. Fehlerursache war schon vor Verwendung aktivierter LP | SWO-53103 | No | Checking file archive activities and equipment Seat of SC and LP checked => Contacts for reading of the LP chip data slightly bent => Contacts readjusted Execution of all fluid service programs => Clot is detected => Device flushed => no constipation detectable restart with currently used LP fails shortly after start replacement LP and start phase i.O. Error cause was already before the use of activated LP | Checking file archive activities and equipment Seat of SC and LP checked => Contacts for reading of the LP chip data slightly bent => Contacts readjusted Execution of all fluid service programs => Clot is detected => Device flushed => no constipation detectable restart with currently used LP fails shortly after start replacement LP and start phase i.O. Error cause was already before the use of activated LP | Assisted Replacement | SP | Activated 7 days before install.
ABL90 FLEX | Überprüfung der Verstärkerkanäle mit dem Sensor Simulator, keine Beanstandung. Tausch der SC, nach Konditionierung und Kalibration, alle QC Werte wieder o.K. Analysator funktionsbereit. Reklamationsmeldung Anwender da gelassen. Anakysator funktionsbereit. |  | SWO-53153 | No | Checking the amplifier channels with the sensor simulator, no complaint. Swap the SC, after conditioning and calibration, all QC values again o.K. analyzer ready for use. Complaint notification user left there. Anakysator is ready for use. |  | Assisted Replacement | SC | Glu
AQT90 FLEX | 更换902-924一个正常。 | 更换902-924一个正常。 | SWO-52729 | No | Replace the 902-924 one normal. | Replace the 902-924 one normal. | Defect | HW | Analyzer control Module | Barcode reader, display
ABL825 FLEX | komplettes OXI Modul getauscht, Lamp adju. durchgeführt, Oxi complett test durchgeführt, tHb Kal vermessen, Kal und QC | komplettes OXI Modul getauscht, Lamp adju. durchgeführt, Oxi complett test durchgeführt, tHb Kal vermessen, Kal und QC | SWO-53156 | No | complete OXI module exchanged, Lamp adju. carried out, Oxi complett test performed, tHb Kal measure, Kal and QC | complete OXI module exchanged, Lamp adju. carried out, Oxi complett test performed, tHb Kal measure, Kal and QC | Defect | HW | Oxi module | Hemolyzer
ABL825 FLEX | 更换945-614修复机器 |  | SWO-52814 | No | Replace the 945-614 repair machine |  | Assisted Replacement | Electrode | Electrode | pH
ABL835 FLEX |  |  | CF8D4F7B3825B2D9 | No |  |  | Unknown | -
ABL820 FLEX | Verified errors in data logs. | Replaced hemolyzer, performed all associated adjustments & replaced tube valve. Tube valve was replaced due to original not resetting into housing. | CF88D0436062FD11 | No | Verified errors in data logs. | Replaced hemolyzer, performed all associated adjustments & replaced tube valve. Tube valve was replaced due to original not resetting into housing. | Defect | HW | Oxi module | Hemolyzer
ABL835 FLEX | FLEXQモジュール引き上げ 代替えモジュール設置及び調整 | FLEXQモジュール引き上げ 代替えモジュール設置及び調整 | SWO-53781 | No | FlexQ module pulling replacement module installation and adjustment | FlexQ module pulling replacement module installation and adjustment | Defect | HW | FlexQ Module
ABL835 FLEX | Level 2 QC low outside insert limits. | Cleaned electrode contacts. Checked measuring module valves. Rescanned gas. No change. Replaced measuring chambers and pH seal ring. Customers 6 other analyzers all run below the insert mean. | CF88713B602271B6 | No | Level 2 QC low outside insert limits. | Cleaned electrode contacts. Checked measuring module valves. Rescanned gas. No change. Replaced measuring chambers and pH seal ring. Customers 6 other analyzers all run below the insert mean. | Defect | HW | pH/BG Module | Measuring chamber
ABL835 FLEX | Verified that front scanner is not working. | Reseated cable to barcode scanner at both ends. Also reset draw time in set up. | CF887C2A604E71B9 | No | Verified that front scanner is not working. | Reseated cable to barcode scanner at both ends. Also reset draw time in set up. | Defect | HW | Basic Unit | Cable
ABL835 FLEX | REPLACED ALL TUBING STILL HAVING PROBLEM. | REPLACED HEATER FOR INLET, REPLACED INLET PROBE, REPLACED LIQUID SENSOR, REPLACED INLET NIPPLES. | CF8877AA60071794 | No | REPLACED ALL TUBING STILL HAVING PROBLEM. | REPLACED HEATER FOR INLET, REPLACED INLET PROBE, REPLACED LIQUID SENSOR, REPLACED INLET NIPPLES. | Defect | HW | Inlet module | Spinlet Heater
AQT90 FLEX | Auf Wunsch von TS noch mal fünf Blank Messungen mit Bufferlösung durchgeführt und Datalogs und ein Service Dump gezogen. SD an TS geschickt. Analysator funktionsbereit. | Kalibrierung eines neuen Testkits, keine neue Losnummer vorhanden. Testkit mit neuer Losnummer bestellt für Kunden. Anwender kalibrieren neue Losnummer und führen QC durch. Ist QC nicht mehr grenzwertig, ist kein Service mehr notwendig. Ist QC weiterhin grenzwertig, ist ein Tausch der Optischen Einheit notwendig. Service wird von Kunde angefordert. | SWO-53155 | No | At the request of TS, another five blank measurements were performed with buffer solution and datalogs and a service dump drawn. SD to TS. Analyzer ready for use. | Calibration of a new test kit, no new lot number available. Test kit with new lot number ordered for customers. Users calibrate new lot number and perform QC. If QC is no longer borderline, no service is necessary. If QC continues to be borderline, an exchange of the optical unit is necessary. Service is requested by customer. | Performance | LQC | TnT | New LQC lot
ABL80 FLEX BASIC | 更换新测试卡修复机器 | 更换新测试卡修复机器 | SWO-53469 | No | Replace the new test card repair machine | Replace the new test card repair machine
ABL80 FLEX BASIC | 更换新测试卡修复机器 | 更换新测试卡修复机器 | SWO-53274 | No | Replace the new test card repair machine | Replace the new test card repair machine
AQT90 FLEX | 更换进样针。进样针定标。 | 更换进样针。进样针定标。 | SWO-53256 | No | Replace the sample needle. The sample needle is calibrated. | Replace the sample needle. The sample needle is calibrated. | Defect | HW | Aspiration Module | Bend needle | 4794
AQT90 FLEX | removed the tube and train them how to fix the tube properly so that this will not happen in future. | removed the tube and train them how to fix the tube properly so that this will not happen in future. | SWO-53198 | No | removed the tube and train them how to fix the tube properly so that this will not happen in future. | removed the tube and train them how to fix the tube properly so that this will not happen in future. | User Error | Sample measurement | Aspiration Module | Tube mounting
ABL80 FLEX BASIC | 更换902-972修复机器 | 更换902-972修复机器 | SWO-53242 | No | Replace the 902-972 repair machine | Replace the 902-972 repair machine
ABL80 FLEX BASIC | Analyzer Checked Sc not Installing. ABL80- Standby Analyzer installed now for Temporary solution. | Solution not sucking by solution pump. Thin tube behind SC plucking out Again and Again. Try to fixing but same problem reported. Touch panel opened and cleaned- and re-fixed But same problem persist. Analyzer restarted and try to calibrate so many times but same problem reported. | SWO-53959 | No | Analyzer Checked Sc not Installing. ABL80- Standby Analyzer installed now for Temporary solution. | Solution not sucking by solution pump. Thin tube behind SC plucking out Again and Again. Try to fixing but same problem reported. Touch panel opened and cleaned- and re-fixed But same problem persist. Analyzer restarted and try to calibrate so many times but same problem reported.
ABL80 FLEX BASIC | 更换新测试卡修复机器 | 更换新测试卡修复机器 | SWO-53474 | No | Replace the new test card repair machine | Replace the new test card repair machine
ABL80 FLEX BASIC | 换测试卡正常。 | 换测试卡正常。 | SWO-53995 | No | Swap ping test card OK. | Swap ping test card OK.
ABL80 FLEX BASIC | 换测试卡正常。 | 换测试卡正常。 | SWO-53326 | No | Swap ping test card OK. | Swap ping test card OK.
ABL80 FLEX BASIC | 换测试卡正常。 | 换测试卡正常。 | SWO-53325 | No | Swap ping test card OK. | Swap ping test card OK.
ABL80 FLEX BASIC | Solution not sucking by solution pump. Thin tube behind SC plucking out Again and Again. Try to fixing but same problem reported. Touch panel opened and cleaned- and re-fixed But same problem persist. Analyzer restarted and try to calibrate so many times but same problem reported. | Solution not sucking by solution pump. Thin tube behind SC plucking out Again and Again. Try to fixing but same problem reported. Touch panel opened and cleaned- and re-fixed But same problem persist. Analyzer restarted and try to calibrate so many times but same problem reported. | SWO-53960 | No | Solution not sucking by solution pump. Thin tube behind SC plucking out Again and Again. Try to fixing but same problem reported. Touch panel opened and cleaned- and re-fixed But same problem persist. Analyzer restarted and try to calibrate so many times but same problem reported. | Solution not sucking by solution pump. Thin tube behind SC plucking out Again and Again. Try to fixing but same problem reported. Touch panel opened and cleaned- and re-fixed But same problem persist. Analyzer restarted and try to calibrate so many times but same problem reported.
ABL90 FLEX | ABL überprüft. LP erneuert. Fließwege mehrfach gespült. Gerätestartphase inkl QC's durchgeführt. tHB-Kal durchgeführt (Fällig seit dem 02.09.2015) | ABL überprüft. Beim Eintreffen steht Gerät im Bildschirm Sensorkassette erneuern , da ein Fehler an der Referenzelektrode erkannt wurde. Aktivitätenspeicher auf Hinweise für LP geprüft. Keine Hinweise erkennbar, das das Gerät das LP ersetzt haben will. Sensorkassette aus Kundenbestand erneuert. Gerätestartphase inkl. QC's durchgeführt. Statusausdruck der defekten SC erstellt und an RDEU gesendet. | SWO-53490 | No | ABL. LP renewed. Flow paths rinsed several times. device start phase incl. QC's. tHB-Kal carried out (due on 02.09.2015) | ABL. Upon arrival, the device is on the Renew Sensor Cartridge screen because an error in the reference electrode has been detected. Activity store checked for hints for LP. No indication that the device wants to replace the LP. Sensor cassette from customer base renewed. device start phase incl. QC's. Status expression of the defective SC created and sent to RDEU. | Assisted Replacement | SC | Wetability
ABL90 FLEX | Alle Schläuche erneuert (Schlauch am Quetschventil, Abfallschlauch, Einlasszwischendichtung und Einlassdichtung), System im Service überprüft, System und Funktion i.O. | Alle Schläuche erneuert (Schlauch am Quetschventil, Abfallschlauch, Einlasszwischendichtung und Einlassdichtung), System im Service überprüft, System und Funktion i.O. | SWO-53495 | No | All hoses renewed (hose on the crushing valve, waste hose, inlet intermediate seal and inlet seal), system checked in service, system and function in the following. | All hoses renewed (hose on the crushing valve, waste hose, inlet intermediate seal and inlet seal), system checked in service, system and function in the following. | User Error | Probe positioning
ABL90 | Troubleshot to a clot in the fluid transport system. | Back flushed the analyzer several times. Adjusted the manifold and liquid sensor. Performed a pump calibration. Performed a pressure test and clot removal in service. | CF8C6650C227F524 | No | Troubleshot to a clot in the fluid transport system. | Back flushed the analyzer several times. Adjusted the manifold and liquid sensor. Performed a pump calibration. Performed a pressure test and clot removal in service. | Cleaned | Fluidic pathway
ABL90 FLEX | Sensorkassette erneuert | Sensorkassette erneuert | SWO-53561 | No | Sensor cassette renewed | Sensor cassette renewed | Assisted Replacement | SC | pH
ABL90 FLEX | Gerät überprüft, Lösungspack kontrolliert, Fließwege gespült, Neustart durchgeführt, Einlaßdichtung ersetzt | Gerät überprüft, Lösungspack kontrolliert, Fließwege gespült, Neustart durchgeführt, Einlaßdichtung ersetzt | SWO-53483 | No | Device checked, solution pack checked, flow paths flushed, restart performed, inlet seal replaced | Device checked, solution pack checked, flow paths flushed, restart performed, inlet seal replaced | Cleaned | Fluidic pathway
ABL90 | I transfered the setting, but did not realize the older units had MR3 software. When I rebooted the analyzer Windows was corrupted. | I downloaded MR5 and reinstall, but this did not solve the problem Ordered another software card for tomorrow. 9 15 Came to Fedex, but customer service did not ship me the part 9 18 Matt sent me a USB Windows. I had problems loading it, but finally got it to work Loaded the Analyzer software. | CF8C66A7C25651DF | No | I transfered the setting, but did not realize the older units had MR3 software. When I rebooted the analyzer Windows was corrupted. | I downloaded MR5 and reinstall, but this did not solve the problem Ordered another software card for tomorrow. 9 15 Came to Fedex, but customer service did not ship me the part 9 18 Matt sent me a USB Windows. I had problems loading it, but finally got it to work Loaded the Analyzer software. | Setup | Analyzer settings
ABL90 |  |  | CF8C7518C22C9552 | No |  |  | Cleaned | Fluidic pathway | Clot Probe
AQT90 FLEX |  |  | SWO-53315 | No |  |  | Defect | HW | Aspiration Module | Needle Wash Unit | FTC-13
AQT90 FLEX | 重新定标 正常。 | 重新定标 正常。 | SWO-53351 | No | Re-calibration is OK. | Re-calibration is OK. | Assisted Replacement | Cartridge | TnI | Calibration
AQT90 FLEX |  |  | SWO-53316 | No |  |  | Defect | HW | Process Module | Process Unit sensor | 5784
ABL825 FLEX | Lamp adjust durchgeführt und Lampe justiert, Kal,s und QC durchgeführt | Lamp adjust durchgeführt und Lampe justiert, Kal,s und QC durchgeführt | SWO-53445 | No | Lamp adjust performed and lamp adjusted, Kal,s and QC performed | Lamp adjust performed and lamp adjusted, Kal,s and QC performed | Adjustment | HW | Oxi module | Lamp unit | Classical: They have installed the Analyzer withou adjusting the Lamp after transport
ABL800 BASIC | Gerät überprüft, Einlaß kontrolliert, Netzteil erneuert | Gerät überprüft, Einlaß kontrolliert, Netzteil erneuert | SWO-53537 | No | Device checked, inlet checked, power supply renewed | Device checked, inlet checked, power supply renewed | Defect | HW | Inlet module
ABL825 FLEX |  |  | SWO-53321 | No |  |  | Defect | HW | Oxi module
ABL835 FLEX | Jeg har skiftet Autocheck modulet. Alle ampuller er scannet og der er kørt 4 QCer. Alt er OK | Jeg har skiftet Autocheck modulet. Alle ampuller er scannet og der er kørt 4 QCer. Alt er OK | SWO-53398 | No | I changed the Auto check module. All ampoules have been scanned and 4 QCer has been run. Everything is OK | I changed the Auto check module. All ampoules have been scanned and 4 QCer has been run. Everything is OK | Defect | HW | AutoCheck Module
ABL800 BASIC | Austausch des DMS covers, Kunde reinigt Geräteoberflächen mit Mikrobac Tissues, Einsendung des defekten DMS covers zur Übeprüfung |  | SWO-53334 | No | Replacement of the DMS cover, customer cleans device surfaces with Mikrobac tissues, submission of the defective DMS cover for examination |  | Defect | HW | Printer module | Cover
ABL825 FLEX | checked data logs, sealing rings, measuring chambers. Ran QC with dummy electrodes in ph, ref, and po2 chambers. PCO2 measured lower not higher. | replaced pH BG module PCB for troubleshooting purposes. Will continue to monitor and see if PCO2 trends higher or lower. | CF8C4428C2422881 | No | checked data logs, sealing rings, measuring chambers. Ran QC with dummy electrodes in ph, ref, and po2 chambers. PCO2 measured lower not higher. | replaced pH BG module PCB for troubleshooting purposes. Will continue to monitor and see if PCO2 trends higher or lower. | Defect | HW | pH/BG Module
ABL825 FLEX | Gerät und Elektrode kontrolliert, pH Elektrode (RV-02) erneuert, Meßkammer und Luftauslass gesäubert, Touchscreen kalibriert | Gerät und Elektrode kontrolliert, pH Elektrode (RV-02) erneuert, Meßkammer und Luftauslass gesäubert, Touchscreen kalibriert | SWO-53484 | No | Device and electrode controlled, pH electrode (RV-02) renewed, measuring chamber and air outlet cleaned, touchscreen calibrated | Device and electrode controlled, pH electrode (RV-02) renewed, measuring chamber and air outlet cleaned, touchscreen calibrated | Assisted Replacement | Electrode | Electrode | pH
AQT90 FLEX | LS gereinigt, Pumpe 3 und Pumpe 4 ersetzt, SP ersetzt, Probemessungen durchgeführt, altes SP SK-04 | LS gereinigt, Pumpe 3 und Pumpe 4 ersetzt, SP ersetzt, Probemessungen durchgeführt, altes SP SK-04 | SWO-53480 | No | LS cleaned, pump 3 and pump 4 replaced, SP replaced, test measurements carried out, old SP SK-04 | LS cleaned, pump 3 and pump 4 replaced, SP replaced, test measurements carried out, old SP SK-04 | Assisted Replacement | Consumables | SP | Aspiration Module | FTC15
AQT90 FLEX | Kundengerät aufgestellt Konfiguration überprüft Startphase i.O. Testkassetten aufgefüllt LQC i.O. | Kundengerät aufgestellt Konfiguration überprüft Startphase i.O. Testkassetten aufgefüllt LQC i.O. | SWO-53382 | No | Customer device set up Configuration checks start phase i.O. test cartridges replenished LQC i.O. | Customer device set up Configuration checks start phase i.O. test cartridges replenished LQC i.O. | Defect | HW | Power Supply | Power Supply | 42173
ABL80 FLEX BASIC | 换测试卡正常。 | 换测试卡正常。 | SWO-53773 | No | Swap ping test card OK. | Swap ping test card OK.
ABL80 FLEX BASIC | 换测试卡正常。 | 换测试卡正常。 | SWO-53590 | No | Swap ping test card OK. | Swap ping test card OK.
ABL80 CO-OX | Tried 2 SP's and 2 SC's problem became both pCO2 and pH errors | Replaced SC | CF8DA9A43848E141 | No | Tried 2 SP's and 2 SC's problem became both pCO2 and pH errors | Replaced SC
ABL80 FLEX BASIC | 更换新测试卡 | 更换新测试卡 | SWO-54453 | No | Replace a new test card | Replace a new test card
ABL80 FLEX BASIC | 换测试卡正常。 | 换测试卡正常。 | SWO-53579 | No | Swap ping test card OK. | Swap ping test card OK.
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-54746 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL80 FLEX BASIC | 换测试卡正常。 | 换测试卡正常。 | SWO-53625 | No | Swap ping test card OK. | Swap ping test card OK.
ABL90 |  |  | CF8D71FA38217A21 | No |  |  | Assisted Replacement | SC | Wetability
ABL90 FLEX | ABL überprüft. Beim Eintreffen steht Gerät im Bildschirm Sensorkassette erneuern , da ein Fehler an der Referenzelektrode erkannt wurde. Aktivitätenspeicher auf Hinweise für LP geprüft. Keine Hinweise erkennbar, das das Gerät das LP ersetzt haben will. Sensorkassette aus Kundenbestand erneuert. Gerätestartphase inkl. QC's durchgeführt. Statusausdruck der defekten SC erstellt und an RDEU gesendet. | ABL überprüft. Beim Eintreffen steht Gerät im Bildschirm Sensorkassette erneuern , da ein Fehler an der Referenzelektrode erkannt wurde. Aktivitätenspeicher auf Hinweise für LP geprüft. Keine Hinweise erkennbar, das das Gerät das LP ersetzt haben will. Sensorkassette aus Kundenbestand erneuert. Gerätestartphase inkl. QC's durchgeführt. Statusausdruck der defekten SC erstellt und an RDEU gesendet. | SWO-53737 | No | ABL. Upon arrival, the device is on the Renew Sensor Cartridge screen because an error in the reference electrode has been detected. Activity store checked for hints for LP. No indication that the device wants to replace the LP. Sensor cassette from customer base renewed. device start phase incl. QC's. Status expression of the defective SC created and sent to RDEU. | ABL. Upon arrival, the device is on the Renew Sensor Cartridge screen because an error in the reference electrode has been detected. Activity store checked for hints for LP. No indication that the device wants to replace the LP. Sensor cassette from customer base renewed. device start phase incl. QC's. Status expression of the defective SC created and sent to RDEU. | Defect | Copy/paste?
ABL90 | The analyzer keeps on going into intervention required, requesting to replace a new inlet gasket. | Flushed the fluid transport path. Upon replacing the inlet gasket, I found the inlet probe bent. Replaced the inlet probe and, the inlet gasket connector. Also, found the analyzer missing the inlet clip, replaced a new clip. Found error 1134 The chip information for the solution pack could not be read. Replaced a new solution pack. The new solution pack started leaking qc solution so I had to replace a second solution pack. | CF8DD383384BBD06 | No | The analyzer keeps on going into intervention required, requesting to replace a new inlet gasket. | Flushed the fluid transport path. Upon replacing the inlet gasket, I found the inlet probe bent. Replaced the inlet probe and, the inlet gasket connector. Also, found the analyzer missing the inlet clip, replaced a new clip. Found error 1134 The chip information for the solution pack could not be read. Replaced a new solution pack. The new solution pack started leaking qc solution so I had to replace a second solution pack. | Replaced | Probe | Bend
AQT90 FLEX |  |  | SWO-53583 | No |  |  | Defect | HW | Analyzer control Module | Barcode reader, display | -
AQT90 FLEX | 更换903-171，910-304各一个正常。 | 更换903-171，910-304各一个正常。 | SWO-53378 | No | Replace 903-171, 910-304 each normal. | Replace 903-171, 910-304 each normal. | Defect | HW | Aspiration Module | Bend needle | 4794
AQT90 FLEX |  |  | SWO-53649 | No |  |  | Defect | HW | Analyzer control Module | Main Board, PCB | 4587.5766
AQT90 FLEX |  |  | SWO-53596 | No |  |  | Defect | HW | Process Module | Manifold, transparent | 3412
ABL800 BASIC | Einlaßmodul ersetzt, | Gerät überprüft, Einlaß kontrolliert, Netzteil erneuert | SWO-53823 | No | Inlet module replaced, | Device checked, inlet checked, power supply renewed | Defect | HW | Inlet module |  | Identical to 16162
ABL825 FLEX | Bildschirm gereinigt, Funktionskontrolle ok | Herr D. meldet das es in Nacht öfters vorgekommen ist das Bildschirmeingaben, wenn das FlexQ-Modul bestückt wird, nicht genommen werden. Der Bildschirm und der schwarze Rahmen wurde schon gereinigt. Herr D. wünscht eine Überprüfung des Analysators durch den Radiometer-Techniker. | SWO-53817 | Yes | Screen cleaned, function control ok | Mr. D. reports that it happened more often at night the screen input when the FlexQ module is equipped, not to be taken. The screen and black frame have already been cleaned. Mr D. would like the radiometer technician to check the analyzer. | Re classified RMED | Hotline
ABL820 FLEX | Pco2 bas sur les niveaux de controles | Remplac tube de pompe BG et O'ring PH (Blue) and a new BG modules (leoner from london office) | CF8D643A383C20EE | No | Pco2 low on control levels | Remplac tube de pompe BG et O'ring PH (Blue) and a new BG modules (leoner from london office) | Assisted Replacement | Electrode | Electrode | pCO2
AQT90 FLEX | Logfiles ausgewertet, keine Auffälligkeit erkennbar. Einmal war der Fehler das die geschlossene Cardrige Door nicht erkannt wurde und ein Neustart wurde verlangt. Halbjahreswartung wird durchgeführt um einen sauberen Betrieb aller Sensoren sicher zu stellen | momentan ist das Gerät Messbereit mit Frau Lippel (0895998833) den Halbjahreswartungstermin auf Dienstag 15.09.2015 ab 12.00 Uhr verschoben | SWO-53687 | No | Logfiles evaluated, no noticeable. Once the error was that the closed Cardrige Door was not recognized and a restart was requested. Half-year maintenance is carried out to ensure clean operation of all sensors | currently the device ready to be measured with Ms. Lippel (0895998833) has postponed the half-year maintenance date to Tuesday 15.09.2015 from 12.00 o'clock | Unknown | SW | SW | Power off/on | 4273
AQT90 FLEX | Barcodereader erneuert. Startphase durchgeführt. Mit neuem Barcodeleser erneut kein auslesen möglich. Barcodeleser zurückgetauscht. | Barcodereader erneuert. Startphase durchgeführt. Mit neuem Barcodeleser erneut kein auslesen möglich. Barcodeleser zurückgetauscht. | SWO-53862 | No | Barcode reader renewed. started phase. With new barcode reader again no reading possible. barcode reader. | Barcode reader renewed. started phase. With new barcode reader again no reading possible. barcode reader. | Cleaned | HW | Storage Module | Barcode reader, Storage wheel | -
AQT90 FLEX | Fehlersuche, Kal.- Aktivitätendateien überprüft, Analysator ging direkt bei Einlegen der Kal. Kassette auf Erneuerung, Solutionpack musste getauscht werden, diesen heute erneuert->Messbereit, Kal. Kassette von gestern NT-ProBnp wird nicht mehr akzeptiert->Nachbestellung, sobald Kal.-Kassette da ist wird der Anwender den Parameter Kalibrieren System-und Funktionskontrolle i.O. | Fehlersuche, Kal.- Aktivitätendateien überprüft, Analysator ging direkt bei Einlegen der Kal. Kassette auf Erneuerung, Solutionpack musste getauscht werden, diesen heute erneuert->Messbereit, Kal. Kassette von gestern NT-ProBnp wird nicht mehr akzeptiert->Nachbestellung, sobald Kal.-Kassette da ist wird der Anwender den Parameter Kalibrieren System-und Funktionskontrolle i.O. | SWO-53621 | No | Troubleshooting, Kal.- Activity files checked, Analyzer went directly when inserting the Cal. Cassette on renewal, Solutionpack had to be replaced, this today renewed->Ready to measure, Kal. Cassette of yesterday NT-ProBnp is no longer accepted->reordering as soon as cal.cassette is there the user will calibrate the parameter calibration system and function control i.O. | Troubleshooting, Kal.- Activity files checked, Analyzer went directly when inserting the Cal. Cassette on renewal, Solutionpack had to be replaced, this today renewed->Ready to measure, Kal. Cassette of yesterday NT-ProBnp is no longer accepted->reordering as soon as cal.cassette is there the user will calibrate the parameter calibration system and function control i.O. | User Error | Consumables | SP | Calibration | 42186
ABL80 FLEX BASIC | 更换新测试卡修复机器 | 更换新测试卡修复机器 | SWO-53814 | No | Replace the new test card repair machine | Replace the new test card repair machine
ABL80 FLEX BASIC | 换试剂包正常。 | 换试剂包正常。 | SWO-54004 | No | The reagent pack is OK. | The reagent pack is OK.
ABL80 OSM | HAD ERROR CODES 582, 584, AND 581. FOUND BLOCKAGE IN HEMOLYZER, FLUSHED 10 TIMES AND PUT MAINTENANCE KIT ON ANALYZER | DID COMPLETE MAINTENANCE KIT ON ANALYZER. COMPLETED THB CALIBRATION. | CF8EF883AE562B60 | No | HAD ERROR CODES 582, 584, AND 581. FOUND BLOCKAGE IN HEMOLYZER, FLUSHED 10 TIMES AND PUT MAINTENANCE KIT ON ANALYZER | DID COMPLETE MAINTENANCE KIT ON ANALYZER. COMPLETED THB CALIBRATION.
ABL80 FLEX BASIC | 更换新测试卡修复机器 | 更换新测试卡修复机器 | SWO-53855 | No | Replace the new test card repair machine | Replace the new test card repair machine
ABL80 CO-OX | Verified there is no communication. | Reloaded RALS into the system and checked the addresses. | CF8EDCF2AE3449C1 | No | Verified there is no communication. | Reloaded RALS into the system and checked the addresses.
ABL90 FLEX | LP erneuert Kalibration und QC durchgeführt | LP erneuert Kalibration und QC durchgeführt | SWO-54184 | No | LP renews calibration and QC performed | LP renews calibration and QC performed | Assisted Replacement | SP | Parameter
ABL835 FLEX | Checked data logs. Other parameters are also off the mean values and unstable. | Freed immobilized flat membrane valves in measuring modules. No help. Freed flat membrane valve in fluidic module even though it appeared to be moving. No help. Isolation check OK. | CF8EC946AE34CA55 | No | Checked data logs. Other parameters are also off the mean values and unstable. | Freed immobilized flat membrane valves in measuring modules. No help. Freed flat membrane valve in fluidic module even though it appeared to be moving. No help. Isolation check OK. | Assisted Replacement | Electrode | Electrode | pCO2
ABL835 FLEX | found that they were Getting Neon Voltage to low, AD trigger error. | Did a complete oxi test did not correct the issue. Rebooted the analyzer which corrected the issue. I then reran the complete oxi test to make sure it was ok everything was fine. | CF9DA940363976E3 | No | found that they were Getting Neon Voltage to low, AD trigger error. | Did a complete oxi test did not correct the issue. Rebooted the analyzer which corrected the issue. I then reran the complete oxi test to make sure it was ok everything was fine. | Defect | HW | Basic Unit | Power Supply | missing trigger
AQT90 FLEX | Analysator bei Eintreffen auf 'grün'. Service Dump Auswertung über SD Analysis Tool ohne Beanstandung - 5 Blank Messungen durchgeführt - Auswertung SD Analysis Tool ohne Beanstandung (1261-1397) - Kunde hat keine CRP LQC im Bestand (Bestellung erfolgt) - Rücksprache mit Kunde bezüglich Auftauzeit CRP LQC - Kunde meldet sich bei erneuten Problemen - Rücksprache bezüglich Termin für Jahreswartung erfolgt. | Analysator bei Eintreffen auf 'grün'. Service Dump Auswertung über SD Analysis Tool ohne Beanstandung - 5 Blank Messungen durchgeführt - Auswertung SD Analysis Tool ohne Beanstandung (1261-1397) - Kunde hat keine CRP LQC im Bestand (Bestellung erfolgt) - Rücksprache mit Kunde bezüglich Auftauzeit CRP LQC - Kunde meldet sich bei erneuten Problemen - Rücksprache bezüglich Termin für Jahreswartung erfolgt. | SWO-53925 | No | Analyzer when it arrives on 'green'. Service Dump Evaluation via SD Analysis Tool without complaint - 5 Blank Measurements Performed - Evaluation SD Analysis Tool without Complaint (1261-1397) - Customer has no CRP LQC in stock (order made) - Consultation with customer regarding thawing time CRP LQC - Customer reports in the event of further problems - consultation is given regarding the date for annual maintenance. | Analyzer when it arrives on 'green'. Service Dump Evaluation via SD Analysis Tool without complaint - 5 Blank Measurements Performed - Evaluation SD Analysis Tool without Complaint (1261-1397) - Customer has no CRP LQC in stock (order made) - Consultation with customer regarding thawing time CRP LQC - Customer reports in the event of further problems - consultation is given regarding the date for annual maintenance. | No fault found | Performance | CRP | LQC | Borderline
ABL80 FLEX BASIC | 更换新测试卡修复机器 | 更换新测试卡修复机器 | SWO-54026 | No | Replace the new test card repair machine | Replace the new test card repair machine
ABL80 FLEX BASIC | 已确认为测试卡和试剂包的问题，走质保流程。仪器在11月份会再做实验，到时候换上新的测试卡和试剂包。 | 已确认为测试卡和试剂包的问题，走质保流程。仪器在11月份会再做实验，到时候换上新的测试卡和试剂包。 | SWO-54285 | No | Issues identified as test cards and reagent kits go through the warranty process. The instrument will be tested again in November, when new test cards and reagent kits will be replaced. | Issues identified as test cards and reagent kits go through the warranty process. The instrument will be tested again in November, when new test cards and reagent kits will be replaced.
ABL80 FLEX BASIC | 更换新测试卡 | 更换新测试卡 | SWO-54479 | No | Replace a new test card | Replace a new test card
ABL80 FLEX BASIC |  |  | SWO-54002 | No
ABL90 FLEX | Sensorkassette erneuert, System gespült, optisches Element im kassettenfach gereinigt, Einlassdichtung gereinigt, pO2 compleet Test durchgeführt und Lösungspack erneuert. Nach LP Tausch sind die Kal-Werte wieder ok. | Sensorkassette erneuert, System gespült, optisches Element im kassettenfach gereinigt, Einlassdichtung gereinigt, pO2 compleet Test durchgeführt und Lösungspack erneuert. Nach LP Tausch sind die Kal-Werte wieder ok. | SWO-54194 | No | Sensor cassette renewed, system flushed, optical element cleaned in the cassette compartment, inlet seal cleaned, pO2 compleet test performed and solution pack renewed. After LP exchange, the Cal values are ok again. | Sensor cassette renewed, system flushed, optical element cleaned in the cassette compartment, inlet seal cleaned, pO2 compleet test performed and solution pack renewed. After LP exchange, the Cal values are ok again. | Assisted Replacement | SP | pO2
ABL90 FLEX | Nach LP - Wechsel Gerät wieder messbereit. LP macht in anderem Gerät gleichen Fehler. Gespräch mit Medizintechnik und Stationsleitung geführt. Gerät überprüft. Abfallschlauch, Ventilschlauch, Anschlussdichtung und Einlassdichtung erneuert.Gerät manuell gespült. Oxi complete test ok. Kalibration und QC's durchgeführt. | Nach LP - Wechsel Gerät wieder messbereit. LP macht in anderem Gerät gleichen Fehler. Gespräch mit Medizintechnik und Stationsleitung geführt. Gerät überprüft. Abfallschlauch, Ventilschlauch, Anschlussdichtung und Einlassdichtung erneuert.Gerät manuell gespült. Oxi complete test ok. Kalibration und QC's durchgeführt. | SWO-54196 | No | After LP - change device ready to be measured again. LP makes the same mistake in other devices. conversation with medical technology and station management. device checked. Waste hose, valve hose, connection seal and inlet seal. device manually rinsed. Oxi complete test ok. calibration and QC's. | After LP - change device ready to be measured again. LP makes the same mistake in other devices. conversation with medical technology and station management. device checked. Waste hose, valve hose, connection seal and inlet seal. device manually rinsed. Oxi complete test ok. calibration and QC's. | Assisted Replacement | SP | pCO2
ABL90 FLEX | Netzwerkkarte auf 10MBit halbduplex eingestellt. Powersavings Netzwerkkarte abgeschaltet. Daten auf CF-Karte gesichert. | Netzwerkkarte auf 10MBit halbduplex eingestellt. Powersavings Netzwerkkarte abgeschaltet. Daten auf CF-Karte gesichert. | SWO-54174 | No | NETWORK card set to 10MBit half-duplex. Powersavings network card switched off. Data backed up to CF card. | NETWORK card set to 10MBit half-duplex. Powersavings network card switched off. Data backed up to CF card. | Setup | HIS/LIS
ABL90 FLEX | Sensorkassette aus Kundenbestand erneuert. Einlassdichtung und Einlasssonde erneuert. --Garantie-- | Sensorkassette aus Kundenbestand erneuert. Einlassdichtung und Einlasssonde erneuert. --Garantie-- | SWO-54229 | No | Sensor cassette from customer base renewed. Inlet seal and inlet probe renewed. --Guarantee-- | Sensor cassette from customer base renewed. Inlet seal and inlet probe renewed. --Guarantee-- | Assisted Replacement | SC | Wetability
ABL90 FLEX | Netzwerkdaten überprüft u. eingegeben, Kommunikation eingestellt, Datenübertragung aktiviert, LP ersetzt, Kal u. QC i.O. Nachfrage CobasIT: Daten werden übertragen. | Netzwerkdaten überprüft u. eingegeben, Kommunikation eingestellt, Datenübertragung aktiviert, LP ersetzt, Kal u. QC i.O. Nachfrage CobasIT: Daten werden übertragen. | SWO-54176 | No | Network data checked and entered, communication set, data transfer enabled, LP replaced, Kal & QC i.O. Demand CobasIT: Data is transferred. | Network data checked and entered, communication set, data transfer enabled, LP replaced, Kal & QC i.O. Demand CobasIT: Data is transferred. | Setup | HIS/LIS
ABL90 FLEX | Downgrade auf 3.1MR3 erfolglos, Datenbanken gelöscht erfolglos, CF Karte mit Windows überspielt erfolglos, CF Karte erneuert, Software aufgespielt kein Erfolg, Upgrade auf 3.1MR6 erfolgreich |  | SWO-54232 | No | Downgrade to 3.1MR3 unsuccessful, databases deleted unsuccessfully, CF card with Windows overplayed unsuccessfully, CF card renewed, software played up no success, upgrade to 3.1MR6 successful |  | Defect | SW | Status Code 2
ABL90 FLEX | 换测试卡正常。 | 换测试卡正常。 | SWO-54225 | No | Swap ping test card OK. | Swap ping test card OK. | Assisted Replacement | SC | Reference
AQT90 FLEX | 需要两张定标卡，重开SWO加借一张。 | 重新定标正常。 | SWO-53984 | No | Two bid cards are required to re-open sWO plus one. | Recalibration is normal. | Assisted Replacement | Cartridge | Myo | Calibration | 159
ABL820 FLEX | Checked electrode chamber. Checked all tubings. Checked inlet area | Replaced inlet probe. Ensured leakage cleared. | CF903DEE241A9E18 | No | Checked electrode chamber. Checked all tubings. Checked inlet area | Replaced inlet probe. Ensured leakage cleared. | Defect | HW | Inlet module | Probe | Clot
ABL815 FLEX | Funktion überprüft, Messkammer optisch OK, gereinigt, E1001 membraniert mit Lot R1827 neu membraniert (Lot 1856). O-Ring an der pH Elektrode ersetzt. Intervall der 1 Punkt Kal auf 2 h gesetzt. Funktion z. Zt. OK | Funktion überprüft, Messkammer optisch OK, gereinigt, E1001 membraniert mit Lot R1827 neu membraniert (Lot 1856). O-Ring an der pH Elektrode ersetzt. Intervall der 1 Punkt Kal auf 2 h gesetzt. Funktion z. Zt. OK | SWO-54051 | No | Function checked, measuring chamber optically OK, cleaned, E1001 membraneed with lot R1827 re-membraneed (lot 1856). O-ring on the pH electrode. Interval of 1 point cal set to 2 h. Function currently OK | Function checked, measuring chamber optically OK, cleaned, E1001 membraneed with lot R1827 re-membraneed (lot 1856). O-ring on the pH electrode. Interval of 1 point cal set to 2 h. Function currently OK | Assisted Replacement | Electrode | Membrane | Reference
ABL825 FLEX |  |  | CF9051CE2446BCBB | No |  |  | Defect | HW | Gas module | Pressure Transducer | pCO2
ABL825 FLEX | Option Keys installiert. Update Type Gerät okay | Option Keys installiert. Update Type Gerät okay | SWO-54034 | Yes | Option Keys installed. Update Type device okay | Option Keys installed. Update Type device okay | Re classified RMED | Demo key exp
ABL835 FLEX | Oldatcserék, Protein eltávolítás, Membránozás (Lac, Glu, Ref), Kalibrálások (Elektródák, tHb), QC mérés hitelesítés |  | SWO-54089 | No | Solution exchanges, Protein removal, membrane (LAC, Glu, Ref), calibrations (electrodes, tHb), QC measurement authentication |  | Defect | HW | Power Supply | Neon Lamp intensity low
ABL837 FLEX | Verschiedene Patientenberichte für die unterschiedlichen Profile angelegt. Unterschiedliche Vorgehensweisen mit flexQ-Modul aufgezeigt. Übertragung in EDV überprüft. | Verschiedene Patientenberichte für die unterschiedlichen Profile angelegt. Unterschiedliche Vorgehensweisen mit flexQ-Modul aufgezeigt. Übertragung in EDV überprüft. | SWO-54180 | No | Different patient reports for the different profiles. Different procedures with flexQ module. Transmission in computer checked. | Different patient reports for the different profiles. Different procedures with flexQ module. Transmission in computer checked. | Setup | Report layout
ABL90 | I logged in and the approval button was active in the settings, and I could see it on patient results that had not been already approved. I checked the customers securty profile and proper access was granted. The person that was having the issue was not available. | No defects noted. | CF94C30DFC62443A | No | I logged in and the approval button was active in the settings, and I could see it on patient results that had not been already approved. I checked the customers securty profile and proper access was granted. The person that was having the issue was not available. | No defects noted. | No fault found
ABL90 FLEX | Neuen Akku eingebaut, Gerät neu gestartet | Neuen Akku eingebaut, Gerät neu gestartet | SWO-54387 | No | New battery installed, device restarted | New battery installed, device restarted | Defect | HW | Battery module
ABL90 FLEX | Gerät überprüft seit 24 Stunden kein Fehler Verschlauchung überprüft--> kein Mangel ersichtlich | Gerät überprüft seit 24 Stunden kein Fehler Verschlauchung überprüft--> kein Mangel ersichtlich | SWO-54251 | No | Device checks for 24 hours no error hose checked--> no defect visible | Device checks for 24 hours no error hose checked--> no defect visible | No fault found | Leak current
ABL90 | Unit working upon arrival.Unit displaying more parameters than the unit is supposed to report.Wron SCs in stock. | Ran all levels Qc=passed. Got with Sales about type of SC customer is supposed to have. | CFA182669821BA73 | No | Unit working upon arrival. Unit displaying more parameters than the unit is supposed to report. Wron SCs in stock. | Ran all levels Qc=passed. Got with Sales about type of SC customer is supposed to have. | No fault found | Glu
ABL90 | Unit down. Flushed unit. | Replaced SC and reran all test. | CFA183EA9857BA76 | No | Unit down. Flushed unit. | Replaced SC and reran all test. | Assisted Replacement | SC | Wetability
ABL90 FLEX | Verbindung hergestellt und Überprüft, Kommunikation mit der IT KFH nicht möglich (interne Betriebsversammlung) |  | SWO-54390 | No | Connection established and verified, communication with the IT KFH not possible (internal company meeting) |  | Adjustment | re boot
ABL90 FLEX | 换试剂包正常。 | 换试剂包正常。 | SWO-54260 | No | The reagent pack is OK. | The reagent pack is OK. | Assisted Replacement | SP | FT
ABL90 FLEX | 换测试卡正常。 | 换测试卡正常。 | SWO-54264 | No | Swap ping test card OK. | Swap ping test card OK. | Assisted Replacement | SC | Wetability
ABL90 FLEX | 更换910-342一个正常。 | 更换910-342一个正常。 | SWO-54010 | No | Replace 910-342 one normal. | Replace 910-342 one normal. | Defect | HW | CPU
ABL825 FLEX | Checked data logs for erors, checked valves and liquid sensors, checked inlet probe, checked electrode and electrode chamber | removed inlet module and replaced inlet nipples, inlet tubing, and bent inlet probe, replaced electrode membranes and cleaned electrode chambers, | CF945013863494EB | No | Checked data logs for erors, checked valves and liquid sensors, checked inlet probe, checked electrode and electrode chamber | removed inlet module and replaced inlet nipples, inlet tubing, and bent inlet probe, replaced electrode membranes and cleaned electrode chambers, | Defect | HW | Inlet module | Nipple kit | Nipple kit
ABL825 FLEX | No problems found. | Perform Complete OXI Test. Reseat all cables in OXI Module. | CF9152E29A5F9BA0 | No | No problems found. | Perform Complete OXI Test. Reseat all cables in OXY Module. | Unknown | -
AQT90 FLEX | Probenhalter und Nadel ausgebaut und erneuert, Neustart OK. | Rädchen an Halterung des Einlasses abgebrochen; Pipette verbogen oder abgebrochen; der Anwender wünscht einen Vorort-Service zur Beseitigung der Geräte-Störung; | SWO-54321 | No | Sample holder and needle removed and renewed, restart OK. | Rädchen aborted at bracket of the inlet; pipette bent or aborted; the user wants a suburban service to eliminate the device malfunction; | Defect | HW | Aspiration Module | Bend needle
ABL80 FLEX BASIC | 更换测试卡正常。 | 更换测试卡正常。 | SWO-54431 | No | The replacement test card is OK. | The replacement test card is OK.
ABL80 FLEX BASIC | 更换测试卡正常。 | 更换测试卡正常。 | SWO-54432 | Yes | The replacement test card is OK. | The replacement test card is OK. | Re classified RMED | Double upSWO
ABL80 FLEX BASIC | 更换636-520一个正常。 | 更换636-520一个正常。 | SWO-54236 | No | Replace 636-520 one normal. | Replace 636-520 one normal.
ABL80 FLEX BASIC | 更换636-520一个正常 | 更换636-520一个正常 | SWO-54241 | No | Replace 636-520 one normal | Replace 636-520 one normal
ABL90 FLEX PLUS | LP entnommen und nochmals eingelegt. LP wird korrekt erkannt und installiert. Abfallschlauch und Ventilschlauch und Inlet Probe überprüft - ok. Kal Fehler pCO2, Luc, Glu. SC (GS Bestand Lot 0380) erneuert. Kalibration und QC's vorgenommen. Patientenmessung ok. 1x LP und SC beim Kunden gelassen. | LP entnommen und nochmals eingelegt. LP wird korrekt erkannt und installiert. Abfallschlauch und Ventilschlauch und Inlet Probe überprüft - ok. Kal Fehler pCO2, Luc, Glu. SC (GS Bestand Lot 0380) erneuert. Kalibration und QC's vorgenommen. Patientenmessung ok. 1x LP und SC beim Kunden gelassen. | SWO-54685 | No | LP and inserted again. LP is detected and installed correctly. Waste hose and valve hose and inlet sample checked - ok. Kal Error pCO2, Luc, Glu. SC (GS Stock Lot 0380). calibration and QC's. Patient measurement ok. 1x LP and SC at the customer' . | LP and inserted again. LP is detected and installed correctly. Waste hose and valve hose and inlet sample checked - ok. Kal Error pCO2, Luc, Glu. SC (GS Stock Lot 0380). calibration and QC's. Patient measurement ok. 1x LP and SC at the customer' . | User Error | SP install
AQT90 FLEX | 重新定标正常。 | 重新定标正常。 | SWO-53991 | No | Recalibration is normal. | Recalibration is normal. | Assisted Replacement | Cartridge | Myo | Calibration | 159
ABL825 FLEX | 更换定标液1正常 | 更换定标液1正常 | SWO-54221 | No | Replace calibration fluid 1 OK | Replace calibration fluid 1 OK | Assisted Replacement | Electrode | Electrode | Reference
ABL825 FLEX |  |  | CF93918B8617B683 | No |  |  | Assisted Replacement | Electrode | Membrane | Cl
ABL80 FLEX BASIC | 更换两张新测试卡 | 更换两张新测试卡 | SWO-54456 | No | Replace two new test cards | Replace two new test cards
AQT90 FLEX | cleaned the process wheel | cleaned the process wheel | SWO-55567 | No | cleaned the process wheel | cleaned the process wheel | Cleaned | HW | Process Module | Process wheel | 129
ABL80 FLEX BASIC |  |  | SWO-54473 | No
ABL90 FLEX | Erneuerung Einlassnadel und anschließende Jahreswartung (siehe SWO) da Undichtigkeit an Abfallstrecke festgestellt | Erneuerung Einlassnadel und anschließende Jahreswartung (siehe SWO) da Undichtigkeit an Abfallstrecke festgestellt | SWO-54595 | No | Renewal inlet needle and subsequent annual maintenance (see SWO) detected as leakage at waste line | Renewal inlet needle and subsequent annual maintenance (see SWO) detected as leakage at waste line | Replaced | Probe | Bend
ABL90 FLEX | Einlassdichtung und Nadel ersetzt. | Anscheinend war die Einlassdichtung mit Halter im Einlassmodul nicht richtig eingehackt, mit Anwenderin Einlassmodul ausgebaut und Einlassdichtung spülen lassen Nadel soll gerade sein, danach ist der Einlass wieder leichtgängig | SWO-54497 | No | inlet seal and needle. | Apparently the inlet seal with holder in the inlet module was not hacked properly, with user inlet module removed and inlet seal rinse needle should be straight, after that the inlet is again smooth | Defect | HW | Inlet Gasket
ABL90 |  |  | CF95450DFC52497C | No |  |  | Replaced | Probe | Bend
ABL90 FLEX | LP erneut installiert, Kaltstartphase abgewartet, manuelle Kals durchgeführt, neue SC installiert, Initialisierungsphase abgewartet, Kals und QCs durchgeführt | Der Kunde hat das Gerät neu gebootet und den LP noch einmal installiert, aber ohne Erfolg: LP wird nicht erkannt. Der Kunde bittet um einen Service. | SWO-54731 | No | LP reinstalled, cold start phase waited, manual kals performed, new SC installed, initialization phase awaited, Kals and QCs performed | The customer has rebooted the device and reinstalled the LP, but to no avail: LP is not recognized. The customer asks for a service. | Assisted Replacement | SP | Unknown
ABL90 FLEX | Fehlersuche verbogene Einlassnadel gerichtet Einlaß gereinigt Testmessung i.O. | Fehlersuche verbogene Einlassnadel gerichtet Einlaß gereinigt Testmessung i.O. | SWO-54618 | No | Troubleshooting bent inlet needle directed inlet cleaned test measurement in the order | Troubleshooting bent inlet needle directed inlet cleaned test measurement in the order | Adjustment | Probe | Bend
ABL90 | found that analyzer keeps giving flush analyzer. | tried flushing the analyzer out with hypochlorite but did not work. Replaced the sc with corrected the issue. | CF954CD6FC54EE3E | No | found that analyzer keeps giving flush analyzer. | tried flushing the analyzer out with hypochlorite but did not work. Replaced the sc with corrected the issue. | Assisted Replacement | SC | Wetability
ABL90 FLEX | Fehlersuche, Einlasssonde ok, Einlassdichtung verschmutzt->erneuert, manuelle Spülung durchgeführt->durchgängig->Test wiederholt, keine Besserung->über Visual Flow Check Blut angesaugt->Softwareupdate V3.1MR6 siehe Case Update durchgeführt System-und Funktionskontrolle i.O. | Fehlersuche, Einlasssonde ok, Einlassdichtung verschmutzt->erneuert, manuelle Spülung durchgeführt->durchgängig->Test wiederholt, keine Besserung->über Visual Flow Check Blut angesaugt->Softwareupdate V3.1MR6 siehe Case Update durchgeführt System-und Funktionskontrolle i.O. | SWO-54851 | No | Troubleshooting, inlet probe ok, inlet seal dirty->renewed, manual flushing performed->continuous->Test repeated, no improvement->via Visual Flow Check blood sucked-in->Software update V3.1MR6 see Case Update performed system and Functional control in the | Troubleshooting, inlet probe ok, inlet seal dirty->renewed, manual flushing performed->continuous->Test repeated, no improvement->via Visual Flow Check blood sucked-in->Software update V3.1MR6 see Case Update performed system and Functional control in the | Assisted Replacement | SC | Wetability | MR6 solved problem
ABL90 FLEX | Fehlersuche, laut Anwender wurden bereits alle durchführenden, Einlasssonde verbogen->erneuert, neue Einlassdichtung angebracht, beim manuellen Spülen des Analysators tritt gefärbte Lösung aus dem Einlass aus->Verstopfung gelöst nach mehrmaligem Spülen System-und Funktionskontrolle i.O. | Fehlersuche, laut Anwender wurden bereits alle durchführenden, Einlasssonde verbogen->erneuert, neue Einlassdichtung angebracht, beim manuellen Spülen des Analysators tritt gefärbte Lösung aus dem Einlass aus->Verstopfung gelöst nach mehrmaligem Spülen System-und Funktionskontrolle i.O. | SWO-54854 | No | Troubleshooting, according to the user, all performing, inlet probe bent->renewed, new inlet seal attached, during manual rinsing of the analyzer, colored solution emerges from the inlet->blockage released after repeated rinsing system and Functional control in the | Troubleshooting, according to the user, all performing, inlet probe bent->renewed, new inlet seal attached, during manual rinsing of the analyzer, colored solution emerges from the inlet->blockage released after repeated rinsing system and Functional control in the | Cleaned | Fluidic pathway | Clot
ABL90 | Checked CF card to see if it was seated properly and it was. | Installed new CF card and reloaded windows. | CF950C28FC16640A | No | Checked CF card to see if it was seated properly and it was. | Installed new CF card and reloaded windows. | Defect | HW | CF Card
ABL90 FLEX | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-54809 | No | Repair after replacing a new test card | Repair after replacing a new test card | Assisted Replacement | SC | Reference
ABL90 FLEX | 更换902-980一个后修复 | 更换902-980一个后修复 | SWO-54244 | No | Replace 902-980 one post-fix | Replace 902-980 one post-fix | Defect | HW | Hemolyzer
AQT90 FLEX |  |  | SWO-54524 | No |  |  | Defect | HW | Storage Module | Barcode reader, Storage wheel | -
AQT90 FLEX | 更换847-607一个后修复 | 更换847-607一个后修复 | SWO-54475 | No | Replace 847-607 one post-fix | Replace 847-607 one post-fix | Defect | HW | Process Module | Cup dryer, Air filter | -
AQT90 FLEX |  |  | SWO-54512 | No |  |  | Defect | HW | Process Module | Optical Unit | 131
ABL835 FLEX | Druckeranschlußkabel lose. | Druckeranschlußkabel lose. | SWO-54703 | No | Druckeranschlußkabel lose. | Druckeranschlußkabel lose. | Adjustment | HW | Printer module | Cable | loosened
ABL825 FLEX | Option Key installiert.Update Type okay Hemolysator ersetzt. Justage durchgeführt,tHb Kal | der Anwender hat bereits eine tHb-Kalibrierung durchgeführt (Fcuv = 0,9976); der Anwender wünscht einen Vorort-Service zur Beseitigung der Geräte-Störung; | SWO-54655 | No | Option Key installed. Update Type okay Hemolysator replaced. Adjustment performed,tHb Kal | the user has already performed a tHb calibration (Fcuv = 0.9976); the user wants a suburban service to eliminate the device malfunction; | Defect | HW | Oxi module | Hemolyzer | FCOHb
ABL825 FLEX | Pumpen und Schlauchsystem geprüft, Pumpenkalibration durchgeführt | Pumpen und Schlauchsystem geprüft, Pumpenkalibration durchgeführt | SWO-54617 | No | Pumps and hose system tested, pump calibration carried out | Pumps and hose system tested, pump calibration carried out | Defect | HW | Fluidic manifold | Pump tubing
ABL827 FLEX | Checked data logs. Found Glucose unstable and zero current errors on calibrations. | Replaced Glucose electrode. Performed zero current calibration in service mode. Restart calibrations. | CF95406CFC1A6775 | No | Checked data logs. Found Glucose unstable and zero current errors on calibrations. | Replaced Glucose electrode. Performed zero current calibration in service mode. Restart calibrations. | Assisted Replacement | Electrode | Electrode | Glu
ABL835 FLEX | Elektrode gereinigt, Messkammern gereinigt, Membrane für Air Ventil erneuert, Schlauch zwischen Pumpen Modul und El Met Modul erneuert, Lac mehrfach membraniert, Lac Elektrode Lot RT04 durch Lot SK05 ersetzt,Kal mehrfach durchgeführt . Kal und QC i.O. Probe Messung ohne Beanstandung | Elektrode gereinigt, Messkammern gereinigt, Membrane für Air Ventil erneuert, Schlauch zwischen Pumpen Modul und El Met Modul erneuert, Lac mehrfach membraniert, Lac Elektrode Lot RT04 durch Lot SK05 ersetzt,Kal mehrfach durchgeführt . Kal und QC i.O. Probe Messung ohne Beanstandung | SWO-54694 | No | Electrode cleaned, measuring chambers cleaned, membrane for air valve renewed, hose renewed between pumps module and El Met module, Lac multiple membranes, Lac electrode lot RT04 replaced by lot SK05,Kal performed several times. Kal and QC i.O. sample measurement without complaint | Electrode cleaned, measuring chambers cleaned, membrane for air valve renewed, hose renewed between pumps module and El Met module, Lac multiple membranes, Lac electrode lot RT04 replaced by lot SK05,Kal performed several times. Kal and QC i.O. sample measurement without complaint | Assisted Replacement | Electrode | Electrode | Lac
ABL827 FLEX | Verified the complaint within the log files. Also after performing the service complete oxi tests. | Trouble shot the problem to the Lamp Unit. During troubleshooting confirmed that the OXI Control PCB is working correctly. | CF9509BDFC20F378 | No | Verified the complaint within the log files. Also after performing the service complete oxi tests. | Trouble shot the problem to the Lamp Unit. During troubleshooting confirmed that the OXI Control PCB is working correctly. | Defect | HW | Oxi module | Lamp unit | Neon
AQT90 FLEX | Se sustituyo unidad óptica y process Wheel, tras calibrar de nuevo el Dimero este presenta la misma anomalía al pasar los QC. Queda pendiente de sustitución repuesto(Cavro) |  | CFA076AF22434FA0 | No | Optical drive and process wheel are replaced, after recalibrating the Dimero it has the same anomaly when passing the CMs. Remains replacement spare(Cavro) |  | Defect | HW | Aspiration Module | Cavro pump
AQT90 FLEX | im Service Dark Measurement gestartet, RawPhoto Value bei 255 nach Reinigung des Process Wheels inkl. der beiden Ref. Pads Value bei 80-100 5 Blankmessungen durchgeführt - > in Spec | im Service Dark Measurement gestartet, RawPhoto Value bei 255 nach Reinigung des Process Wheels inkl. der beiden Ref. Pads Value bei 80-100 5 Blankmessungen durchgeführt - > in Spec | SWO-54629 | No | in service dark measurement, RawPhoto Value at 255 after cleaning the process wheel incl. the two ref. Pads Value performed at 80-100 5 blank measurements - > in Spec | in service dark measurement, RawPhoto Value at 255 after cleaning the process wheel incl. the two ref. Pads Value performed at 80-100 5 blank measurements - > in Spec | Defect | HW | Process Module | Optical Unit | 135
ABL80 FLEX CO-OX | LP ausgetauscht Funktionskontrolle i.o | LP ausgetauscht Funktionskontrolle i.o | SWO-54887 | No | LP replaced function control i.o | LP replaced function control i.o
ABL80 FLEX BASIC | 更换新测试卡后pH定标不过 | 更换第四张测试卡后修复 | SWO-54951 | No | PH calibration after replacing a new test card | Repair after replacing the fourth test card
ABL80 FLEX BASIC | 更换第二张测试卡后O2、K定标不过 | 更换第四张测试卡后修复 | SWO-54952 | Yes | After replacing the second test card O2, K calibration but | Repair after replacing the fourth test card | Re classified RMED | Double upSWO
ABL80 FLEX BASIC | 更换第三张测试卡后K定标不过 | 更换第四张测试卡后修复 | SWO-54953 | No | K calibration after replacing the third test card | Repair after replacing the fourth test card
ABL80 FLEX BASIC | 更换第四张测试卡后修复 | 更换第四张测试卡后修复 | SWO-54954 | Yes | Repair after replacing the fourth test card | Repair after replacing the fourth test card | Re classified RMED
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-54715 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-54710 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL80 CO-OX | Customer failed proficiency test for PO2 and PCO2. Customer was using WSLH proficiency testing for an ABL80coox, but the peer data was combined with the ABL700 series so customer will use an alternative manufacturer. Performed zero check procedure for PO2 and PCO2 - PASS. Will run linearity kit to verify PO2 and PCO2. | Ran VK linearity kit. Verified instrument operation. | CF964B2472081A4D | No | Customer failed proficiency test for PO2 and PCO2. Customer was using WSLH proficiency testing for an ABL80coox, but the peer data was combined with the ABL700 series so customer will use an alternative manufacturer. Performed zero check procedure for PO2 and PCO2 - PASS. Will run linearity kit to verify PO2 and PCO2. | Ran VK linearity kit. Verified instrument operation.
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-55406 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL90 FLEX | Einlassdichtung gereinigt, Einlasssonde und Einlassdichtung aus Kundenbestand erneuert, LP erneuert, Kal und QCs durchgeführt. | Einlassdichtung gereinigt, Einlasssonde und Einlassdichtung aus Kundenbestand erneuert, LP erneuert, Kal und QCs durchgeführt. | SWO-54911 | No | Inlet seal cleaned, inlet probe and inlet seal renewed from customer stock, LP renewed, Kal and QCs performed. | Inlet seal cleaned, inlet probe and inlet seal renewed from customer stock, LP renewed, Kal and QCs performed. | Cleaned | Fluidic pathway
ABL90 | After the customer had setup the radiance settings an error came up. The error said an an unexpected error has occurred. Customer rebooted the system couple times. | Unchecked the Connect to radiance box. Tried downgrading the SW to MR3 as the testing was done on MR3. Status code 2 appeared. Tried reloading windows. Could not boot from USB. Possibly needs a new CF card. | CF9DDCDB36508E84 | No | After the customer had setup the radiance settings an error came up. The error said an an unexpected error has occurred. Customer rebooted the system couple times. | Unchecked the Connect to radiance box. Tried downgrading the SW to MR3 as the testing was done on MR3. Status code 2 appeared. Tried reloading windows. Could not boot from USB. Possibly needs a new CF card. | Defect | SW | Status Code 2
ABL800 BASIC |  |  | SWO-54789 | No |  |  | Assisted Replacement | Electrode | Electrode | pH
ABL805 FLEX | Im Service Menü die Autocheckampullen Temperatur ist bei 55°C. Den Abl auf Manuelle QC Messung umgestellt und E- Teile Bestellt. | Temperatur Geber Leitung oxidiert>> PCB und Leitung Ersetzt. | SWO-55048 | No | In the service menu the autocheck ampoules temperature is at 55°C. Changed the Abl to Manual QC Measurement and Ordered E-Parts. | Temperature encoder line oxidized>> PCB and line Replaced. | Revisit required | HW | AutoCheck Module | Temperature sensor | Corrosion
ABL820 FLEX | Nipple broken. | Replace nipples for inlet gasket. | CF966D7D723380EC | No | Nipple broken. | Replace nipples for inlet gasket. | Defect | HW | Inlet module | Leaking | Nipple kit
ABL815 FLEX | pH BG Modul zur Probe ausgewechselt, Kundenelektroden sind im ABL verblieben. Funktion z. Zt. OK. | Der Kunde bittet um einen Service. | SWO-54817 | No | pH BG module for sample, customer electrodes remain in the ABL. Function currently OK. | The customer asks for a service. | Defect | HW | pH/BG Module | ph + El errors
ABL825 FLEX |  |  | SWO-54838 | No |  |  | Defect | HW | pH/BG Module | Measuring chamber
ABL820 FLEX | DATA LOGS SHOW ALL 3 LEVELS OF QC RUNNING ON THE LOW SIDE AND LEVEL 2 OUT ON THE LOW SIDES7345 LOT 320 RUNNING AT 37.6 WITH A RANGE OF 38.1 TO 41.1.VISUAL INSPECTION OF PH CHAMBER SHOWS NO SIGNS OF RINSE OR CARRYOVER.INSPECTION ALSO SHOWS THAT THE PH ELECTRODE HAS THE NEW BLUE SEALING RING INSTALLED. | CLEANED THE PH CHAMBERCLEANED THE PCO2 ELECTRODE.INSTALLED NEW PCO2 AND REFERENCE MEMBRANES.INSTALLED A NEW GAS 2 BOTTLE.INSTALLED ANOTHER NEW BLUE O RING ON THE PH ELECTRODE.PERFORMED MULTIPLE TWO POINT CALIBRATIONS.I ALSO RECIEVED AUTHORIZATION FROM TECH SUPPORT TO ORDER A BOX OF REPLACEMENT MEMBRANES AND REPLACEMENT PCO2 ELECTRODE AS IT IS UNDER WARRANTY.THIS WILL BE SHIPPED DIRECTLY TO THE CUSTOMER TO THE ATTENTION OF CINDY GRANT | CF963D00721163D7 | No | DATA LOGS SHOW ALL 3 LEVELS OF QC RUNNING ON THE LOW SIDE AND LEVEL 2 OUT ON THE LOW SIDES7345 LOT 320 RUNNING AT 37.6 WITH A RANGE OF 38.1 TO 41.1.VISUAL INSPECTION OF PH CHAMBER SHOWS NO SIGNS OF RINSE OR CARRYOVER. INSPECTION ALSO SHOWS THAT THE PH ELECTRODE HAS THE NEW BLUE SEALING RING INSTALLED. | CLEANED THE PH CHAMBERCLEANED THE PCO2 ELECTRODE. INSTALLED NEW PCO2 AND REFERENCE MEMBRANES. INSTALLED A NEW GAS 2 BOTTLE. INSTALLED ANOTHER NEW BLUE O RING ON THE PH ELECTRODE. PERFORMED MULTIPLE TWO POINT CALIBRATIONS. I ALSO RECIEVED AUTHORIZATION FROM TECH SUPPORT TO ORDER A BOX OF REPLACEMENT MEMBRANES AND REPLACEMENT PCO2 ELECTRODE AS IT IS UNDER WARRANTY. THIS WILL BE SHIPPED DIRECTLY TO THE CUSTOMER TO THE ATTENTION OF CINDY GRANT | Assisted Replacement | Electrode | Membrane | pCO2
ABL815 FLEX | Option Key eingelesen | Option Key eingelesen | SWO-54913 | No | Option Key read in | Option Key read in | Setup | Option key exp
ABL825 FLEX | Option Key eingelesen | Option Key eingelesen | SWO-54915 | No | Option Key read in | Option Key read in | Setup | Option key exp
ABL837 FLEX | Multiple Liquid Sensor errors. | Cleaned the Gas valve in the fluidic module. | CF96A8387237BC42 | No | Multiple Liquid Sensor errors. | Cleaned the Gas valve in the fluidic module. | Cleaned | HW | Fluidic manifold | Air vent
ABL825 FLEX |  |  | CF96447B7263FF9A | No |  |  | Defect | HW | Oxi module | Lamp unit | Neon lamp
ABL825 FLEX | QC Modul durch Ersatzmodul aus GS ersetzt. Probelauf okay | QC Modul durch Ersatzmodul aus GS ersetzt. Probelauf okay | SWO-54914 | No | QC module replaced by replacement module from GS. Trial run okay | QC module replaced by replacement module from GS. Trial run okay | Defect | HW | AutoCheck Module | Breaker arm w/o function
ABL835 FLEX | Messkammer und Schlauchsystem geprüft, Ventile V1-V4 demontiert und gereinigt, pCO2 Elektrode neu membranieren, Funktionskontrolle ok | Messkammer und Schlauchsystem geprüft, Ventile V1-V4 demontiert und gereinigt, pCO2 Elektrode neu membranieren, Funktionskontrolle ok | SWO-55036 | No | Measuring chamber and hose system tested, valves V1-V4 dismantled and cleaned, pCO2 electrode re-membrane, function control ok | Measuring chamber and hose system tested, valves V1-V4 dismantled and cleaned, pCO2 electrode re-membrane, function control ok | Assisted Replacement | Electrode | Membrane | pCO2
ABL825 FLEX | Glu. Elektrode geprüft, anderes Membranmaterial (Lot 2241) verwendet, Funktionskontrolle ok | Glu. Elektrode geprüft, anderes Membranmaterial (Lot 2241) verwendet, Funktionskontrolle ok | SWO-55035 | No | Glu. Electrode tested, other membrane material (lot 2241) used, function control ok | Glu. Electrode tested, other membrane material (lot 2241) used, function control ok | Assisted Replacement | Electrode | Membrane | Glu
ABL825 FLEX | checked data logs, performed complete oximetry test, | Found no errors on patient results or autcheck QC. Level 4 and 6 of SO CAP survey are normal. Level 5 of SO CAP survey consistently have error 581 Oxi spectrum mismatch. Conclusions are that it is the CAP material. Results are consistent on all 3 ABL's. Customer will report to CAP their findings and get more material to rerun the survey. | CF96492972386A2C | Yes | checked data logs, performed complete oximetry test, | Found no errors on patient results or autcheck QC. Level 4 and 6 of SO CAP survey are normal. Level 5 of SO CAP survey consistently have error 581 Oxi spectrum mismatch. Conclusions are that it is the CAP material. Results are consistent on all 3 ABL's. Customer will report to CAP their findings and get more material to rerun the survey. | Re classified RMED | CAP Survey method
ABL825 FLEX |  |  | CF96D4AD721DB053 | No |  |  | Cleaned | Fluidic pathway
ABL800 BASIC | Fehlersuche, auf Serviceprogrammebene FlexQ Modul überprüft->Mischer dreht sich nur wenn er von Hand angestoßen wird und stoppt nach kurzer Zeit wieder (Fest), Ersatzteil muss bestellt werden und FlexQ Modul wird in GS repariert, vorübergehend FlexQ Modul SN754R1053N001 installiert, altes Modul wird dann wieder installiert, Blutmessung ok System-und Funktionskontrolle i.O. | Fehlersuche, auf Serviceprogrammebene FlexQ Modul überprüft->Mischer dreht sich nur wenn er von Hand angestoßen wird und stoppt nach kurzer Zeit wieder (Fest), Ersatzteil muss bestellt werden und FlexQ Modul wird in GS repariert, vorübergehend FlexQ Modul SN754R1053N001 installiert, altes Modul wird dann wieder installiert, Blutmessung ok System-und Funktionskontrolle i.O. | SWO-54942 | No | Troubleshooting, checked at service program level FlexQ module->mixer only rotates when it is triggered by hand and stops again after a short time (fixed), spare part must be ordered and FlexQ module is repaired in GS, temporarily FlexQ module SN754R1053N001 installed, old module is then reinstalled, blood measurement ok system and function control i.O. | Troubleshooting, checked at service program level FlexQ module->mixer only rotates when it is triggered by hand and stops again after a short time (fixed), spare part must be ordered and FlexQ module is repaired in GS, temporarily FlexQ module SN754R1053N001 installed, old module is then reinstalled, blood measurement ok system and function control i.O. | Defect | HW | FlexQ Module | No mixing
ABL825 FLEX | Gerät überprüft, Meßkammer und Luftauslass gesäubert, Elektroden kontrolliert, Ref.-Elektrode und Lactat membraniert, defekte Einlaßdichtung erneuert, Vorverstärker kalibriert, Dichtungsring Lactat ersetzt |  | SWO-54940 | No | Device checked, measuring chamber and air outlet cleaned, electrodes checked, ref electrode and lactate membraneed, defective inlet seal renewed, preamplifier calibrated, sealing ring lactate replaced |  | Assisted Replacement | HW | Inlet module | Inlet Gasket
ABL827 FLEX | Tried the setup files from a different analyzer and also the default setup. Radiance is still not connecting. | Reloaded the DMS software and Radiance is not connected. | CF962E5072353158 | No | Tried the setup files from a different analyzer and also the default setup. Radiance is still not connecting. | Reloaded the DMS software and Radiance is not connected. | Unknown | Radiance connection
ABL80 FLEX BASIC |  |  | SWO-54975 | No
ABL80 FLEX BASIC | 更换测试卡正常。 | 更换测试卡正常。 | SWO-55671 | No | The replacement test card is OK. | The replacement test card is OK.
ABL90 | INSPECTED DATA LOGS, COMPLETED SERVICE TEST, COMPLETE THB CAL.ALL TESTS AND CONTROLS PASSED. BLOOD RESULTS STILL HAVE SPECTRUM ERRORS | REMOVED HEMOLYZER AND FLUSHED WITH BLEACH AND DI WATER. | CF97DEC2E855B8C6 | No | INSPECTED DATA LOGS, COMPLETED SERVICE TEST, COMPLETE THB CAL. ALL TESTS AND CONTROLS PASSED. BLOOD RESULTS STILL HAVE SPECTRUM ERRORS | REMOVED HEMOLYZER AND FLUSHED WITH BLEACH AND DI WATER. | Unknown | Revisit required | Hemolyzer
ABL90 |  |  | CF98D0315E028200 | Yes |  |  | Re classified RMED | Logistical upSWO | PER
ABL90 FLEX | Kontakte am Analysator mechanisch justiert, LO neu eingesetzt, Gerät kalibriert, | Kontakte am Analysator mechanisch justiert, LO neu eingesetzt, Gerät kalibriert, | SWO-55465 | No | Contacts on the analyzer mechanically adjusted, LO re-inserted, device calibrated, | Contacts on the analyzer mechanically adjusted, LO re-inserted, device calibrated, | Adjustment | SP PCB
ABL90 |  |  | CF98D17F5E5E8203 | Yes |  |  | Re classified RMED | Logistical upSWO | PER
ABL90 FLEX | Fluid transport alles ausgeführt danach i.o Nach initialisierung Kal. i.o Alle QC oximetrie nicht bestanden ( 1087: Inhomogene QC-3 Lösung) Initialisierung mehrmals probiert Neues Lösungspack Kal. i.o Alle QC fehler wieder reproduziert Oxi pO2 test durchgeführt i.o Spülverfahren des Flüssigkeitstransportsystem durchgeführt i.o Mehrmals initialisiert Solutionpack fluidtransport fail Sensorkassette erneuert Alles i.o |  | SWO-55162 | No | Fluid transport everything executed after i.o After initialization Kal. i.o All QC oximetry failed ( 1087: Inhomogeneous QC-3 Solution) Initialization several times tried New Solution Pack Kal. i.o All QC error reproduced Oxi pO2 test carried out i.o flushing procedure of the liquid transport system carried out i.o Multiple initialized Solutionpack fluidtransport fail sensor cartridge renewed All i.o |  | Assisted Replacement | SC | Wetability
ABL90 |  |  | CF9DAECE36392576 | No |  |  | Assisted Replacement | SP | pO2
ABL90 FLEX | Defekte Kassette, Lot 383-370 erneuert. Reklamation bearbeitet und versendet. Gerät überprüft. Die Kassette mit der Seriennummer 946-009R0378N009 stammt aus dem Bestand des KfH Kassel und konnte als Ersatzteil nicht im CRM angegeben werden, da die Seriennummer nicht akzeptiert wurde. | Defekte Kassette, Lot 383-370 erneuert. Reklamation bearbeitet und versendet. Gerät überprüft. Die Kassette mit der Seriennummer 946-009R0378N009 stammt aus dem Bestand des KfH Kassel und konnte als Ersatzteil nicht im CRM angegeben werden, da die Seriennummer nicht akzeptiert wurde. | SWO-55134 | No | Defective cassette, lot 383-370 renewed. Complaint processed and sent. device checked. The cassette with the serial number 946-009R0378N009 comes from the kfH Kassel inventory and could not be specified as a spare part in the CRM because the serial number was not accepted. | Defective cassette, lot 383-370 renewed. Complaint processed and sent. device checked. The cassette with the serial number 946-009R0378N009 comes from the kfH Kassel inventory and could not be specified as a spare part in the CRM because the serial number was not accepted. | Assisted Replacement | SC | pCO2
ABL90 FLEX | Auf Kundenwunsch den berichtbaren Bereich bei pCO2 auf 20kPa erhöht, Kunden Einstellungen erläutert | Auf Kundenwunsch den berichtbaren Bereich bei pCO2 auf 20kPa erhöht, Kunden Einstellungen erläutert | SWO-55040 | No | At customer request, the reportable area at pCO2 is increased to 20kPa, customer settings explained | At customer request, the reportable area at pCO2 is increased to 20kPa, customer settings explained | Setup | Reportable range
ABL90 FLEX | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-55182 | No | Repair after replacing a new test card | Repair after replacing a new test card | Assisted Replacement | SC | Reference
ABL90 FLEX |  |  | SWO-54987 | No |  |  | Assisted Replacement | SC | Lac
AQT90 FLEX |  |  | SWO-54946 | No |  |  | Defect | HW | Incubator Module | Cup lifter | 5297
ABL805 FLEX | Temperatur Geber Leitung oxidiert>> PCB und Leitung Ersetzt. | Temperatur Geber Leitung oxidiert>> PCB und Leitung Ersetzt. | SWO-55060 | No | Temperature encoder line oxidized>> PCB and line Replaced. | Temperature encoder line oxidized>> PCB and line Replaced. | Defect | HW | AutoCheck Module | Temperature sensor | Corrosion
ABL820 FLEX | CUSTOMER SERVICE ERRONEOUSLY SHIPPED A PO2 ELECTRODE INSTEAD OF A PCO2 ELECTRODE AND DID NOT SHIP THE PCO2 MEMBRANES I WILL SEND BACK THE UNUSED PO2 ELECTRODE ON RGA Y15-99394D AND I REORDERED THE PC02 MEMBRANES AND PC02 ELECTRODE TO ARRIVE TOMORROW 9 24 2015 AND WILL RETURN TO THE CUSTOMER WITH THOSE ITEMS. | .I REPLACED THE RINSE AND AIR VALVES AND THE RUBBER PLUG ON TOP OF THE PH BG MODULE.RAN MULTIPLE CALIBRATIONS.WILL RETURN TOMORROW TO CONTINUE THE CALL. | CF97C9FEE82E5840 | Yes | CUSTOMER SERVICE ERRONEOUSLY SHIPPED A PO2 ELECTRODE INSTEAD OF A PCO2 ELECTRODE AND DID NOT SHIP THE PCO2 MEMBRANES I WILL SEND BACK THE UNUSED PO2 ELECTRODE ON RGA Y15-99394D AND I REORDERED THE PC02 MEMBRANES AND PC02 ELECTRODE TO ARRIVE TOMORROW 9 24 2015 AND WILL RET URN TO THE CUSTOMER WITH THOSE ITEMS. | . I REPLACED THE RINSE AND AIR VALVES AND THE RUBBER PLUG ON TOP OF THE PH BG MODULE. RAN MULTIPLE CALIBRATIONS. WILL RETURN TOMORROW TO CONTINUE THE CALL. | Re classified RMED | Revisit w/o sparepart
ABL825 FLEX | Gerät überprüft, probeweise Lactatelektrode erneuert, Membrancharge gewechselt, Ventilmembran Gasventil ersetzt, Elektroden im EL Met Modul membraniert, Lactatelektrode zurückgetauscht |  | SWO-55261 | No | Device checked, lactate electrode renewed on a trial basis, membrane charge changed, valve diaphragm gas valve replaced, electrodes membraneed in el Met module, lactate electrode replaced |  | Assisted Replacement | Electrode | Membrane | Lac
ABL820 FLEX | Checked logs to verify past performance. Logs show PO2 ssensitivity at the highest when new 7.5 and slowly lowered to 4.9-5.0. | replaced the PO2 electrode, cleaned the measuring chamber and inspected the electrode connectors. Performed all calibrations and the sensitivity recovered to 28.9 with the new electrode. | CF979A56E85CC189 | No | Checked logs to verify past performance. Logs show PO2 ssensitivity at the highest when new 7.5 and slowly lowered to 4.9-5.0. | replaced the PO2 electrode, cleaned the measuring chamber and inspected the electrode connectors. Performed all calibrations and the sensitivity recovered to 28.9 with the new electrode. | Assisted Replacement | Electrode | Electrode | pO2
ABL800 BASIC | pCO2 Elektrode neu membraniert. Kalibrationen und QC's vorgenommen. | pCO2 Elektrode neu membraniert. Kalibrationen und QC's vorgenommen. | SWO-55081 | No | pCO2 electrode. calibrations and QC's. | pCO2 electrode. calibrations and QC's. | Assisted Replacement | Electrode | Membrane | pCO2
AQT90 FLEX | Beim SW-Update auf 8.8.15 astm.ini Datei überschrieben Korrekte astm.ini Datei eingespielt => Sicherung auf CF-Karte Überprüfung Netzwerkdatei i.O. Server lässt sich anpingen => Netzwerkverbindung i.O. Patientenberichte und LQC nachgesendet => Rückruf in Burg => Daten kommen an | Beim SW-Update auf 8.8.15 astm.ini Datei überschrieben Korrekte astm.ini Datei eingespielt => Sicherung auf CF-Karte Überprüfung Netzwerkdatei i.O. Server lässt sich anpingen => Netzwerkverbindung i.O. Patientenberichte und LQC nachgesendet => Rückruf in Burg => Daten kommen an | SWO-55116 | No | During the SW update to 8.8.15 astm.ini file overwritten Correct astm.ini file recorded => Backup to CF card verification network file i.O. server can be pinged => network connection i.O. patient reports and LQC sent back => callback in castle => data coming expresses proximity or touch from the side | During the SW update to 8.8.15 astm.ini file overwritten Correct astm.ini file recorded => Backup to CF card verification network file i.O. server can be pinged => network connection i.O. patient reports and LQC sent back => callback in castle => data coming expresses proximity or touch from the side | Setup | SW | His/lis | ASTM overwrite
ABL80 FLEX CO-OX | Gerät überprüft, Meßkassette gespült, Lösungspack erneuert | der Anwender sieht keine Möglichkeit das Problem selber zu beheben; der Anwender wünscht einen Vorort-Service zur Beseitigung der Geräte-Störung; | SWO-55424 | No | Device checked, measuring cassette rinsed, solution pack renewed | the user sees no way to fix the problem himself; the user wants a suburban service to eliminate the device malfunction;
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-55401 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL80 FLEX BASIC |  |  | SWO-55217 | No
ABL80 FLEX Basic version |  |  | CF9923475E614B91 | No
ABL80 FLEX BASIC |  | 更换信号板903-334一个后测量正常 | SWO-55200 | No |  | Replace the signal plate 903-334 one rear measurement normal
ABL80 FLEX BASIC | Touch Panel replaced. Thin tube replaced. Waist Tubes opened and Cleaned Flushed as Waist not coming out. Non writen valve interchanged. Analyser Sericed and Calibrated. Working fine. Reinstalled. Ms Lata. | Touch Panel replaced. Thin tube replaced. Waist Tubes opened and Cleaned Flushed as Waist not coming out. Non writen valve interchanged. Analyser Sericed and Calibrated. Working fine. Reinstalled. Ms Lata. | SWO-55642 | No | Touch Panel replaced. Thin tube replaced. Waist Tubes opened and Cleaned Flushed as Waist not coming out. Non writen valve interchanged. Analyser Sericed and Calibrated. Working fine. Reinstalled. Ms Lata. | Touch Panel replaced. Thin tube replaced. Waist Tubes opened and Cleaned Flushed as Waist not coming out. Non writen valve interchanged. Analyser Sericed and Calibrated. Working fine. Reinstalled. Ms Lata. |  |  | Process Module | Process wheel | 129
ABL90 | PREVIOUS SWO FOUND FAULTY HEMOLYZER | REPLACED HEMOLYZER AND SPECTROMETER | CF98CA545E39BDB5 | No | PREVIOUS SWO FOUND FAULTY HEMOLYZER | REPLACED HEMOLYZER AND SPECTROMETER | Defect | HW | Hemolyzer
ABL90 |  | Cusytomer sent in Printout. | CF990D445E58AE0E | Yes |  | Cusytomer sent in Printout. | Re classified RMED | Logistical upSWO
ABL90 FLEX | Anbindungsdaten eingegegen |  | SWO-55356 | No | Anbindungsdaten eingegegen |  | Setup | HIS/LIS | Aqure
ABL90 | Error code 2 occured while rebooting, tried installing the windows SW. Tried SW MR3 and got same error code. tried new CF card and old CF card. got same error 2. Tried new CPU unit and then got the same error code. Error showed invalid device. | Replaced wet section PCB. Reinstalled windows and Roma. Radiance still not working. Provided all data logs and service dump to Darcy. | CF9F7100AC44408F | No | Error code 2 occured while rebooting, tried installing the windows SW. Tried SW MR3 and got same error code. tried new CF card and old CF card. got same error 2. Tried new CPU unit and then got the same error code. Error showed invalid device. | Replaced wet section PCB. Reinstalled windows and Roma. Radiance still not working. Provided all data logs and service dump to Darcy. | Defect | SW | Status Code 2
ABL90 | 1343,1324 | replaced SC | CF98ED485E4B7440 | No | 13431324 | replaced SC | Assisted Replacement | SC | Wetability
TCM CombiM |  |  | SWO-55190 | No
ABL90 FLEX | 全チューブ交換、溶液パック交換、ＯＫ | 全チューブ交換、溶液パック交換、ＯＫ | SWO-56351 | No | All tube replacement, solution pack replacement, OK | All tube replacement, solution pack replacement, OK | Assisted Replacement | SP | FT
ABL825 FLEX |  |  | CF98E6D95E0E1E33 | Yes |  |  | Re classified RMED | PM
ABL825 FLEX | CONFIRMED ISSUE BY FSE JAY TREESE | SENT OUT PO2 ELECTRODE FOR TROUBLESHOOTING. FEDEX TRACKING NUMBER 808432231508. | CF9935CE5E2F2FF5 | Yes | CONFIRMED ISSUE BY FSE JAY TREESE | SENT OUT PO2 ELECTRODE FOR TROUBLESHOOTING. FEDEX TRACKING NUMBER 808432231508. | Re classified RMED | Logistical upSWO
ABL825 FLEX |  |  | SWO-55228 | No |  |  | Assisted Replacement | Electrode | Electrode | pO2
ABL820 FLEX | DATA LOGS SHOW QC PLOT SHIFTED ON 9 3 2015 AND THAT THE REFERENCE MEMBRANE WAS ALSO CHANGED ON THAT DAY.I REPLACED THE REFERENCE MEMBRANE 2 DAYS AGO FROM THE SAME LOT OF MEMBRANES | REPLACED PCO2 ELECTRODE AND MEMBRANE.REPLACED PH BG MEASURING CHAMBER AND SEALING RINGS.REPLACED INLET TUBING AND NIPPLE KIT.ONLY AFTER REPLACING THE REFERENCE MEMBRANE FOR THE SECOND TIME WITH A NEW LOT NUMBER DID ALL QC VALUES SHIFT BACK WITHIN THE CUSTOMERS WDC PEER GROUP VALUE.MULTIPLE CALIBRATIONS. | CF98F37F5E4A187C | No | DATA LOGS SHOW QC PLOT SHIFTED ON 9 3 2015 AND THAT THE REFERENCE MEMBRANE WAS ALSO CHANGED ON THAT DAY. I REPLACED THE REFERENCE MEMBRANE 2 DAYS AGO FROM THE SAME LOT OF MEMBRANES | REPLACED PCO2 ELECTRODE AND MEMBRANE. REPLACED PH BG MEASURING CHAMBER AND SEALING RINGS. REPLACED INLET TUBING AND NIPPLE KIT. ONLY AFTER REPLACING THE REFERENCE MEMBRANE FOR THE SECOND TIME WITH A NEW LOT NUMBER DID ALL QC VALUES SHIFT BACK WITHIN THE CUSTOMERS WDC PEER GROUP VALUE. MULTIPLE CALIBRATIONS. | Assisted Replacement | Electrode | Membrane | pCO2
ABL825 FLEX | Error code 314, Oxi wavelength deviation high. Problem only occurs when running QCs. | Replace hemolyzer unit and lamp unit. Perform lamp alignment. Perform Complet OXI Test. Perform service tHb calibration. | CF98D2555E605BC2 | No | Error code 314, Oxi wavelength deviation high. Problem only occurs when running QCs. | Replace hemolyzer unit and lamp unit. Perform lamp alignment. Perform Complet OXI Test. Perform service tHb calibration. | Defect | HW | Oxi module | Hemolyzer
ABL825 FLEX | Hemolysator ersetzt. Justage und Service tHb Kal SO2 wieder nahe ZW | Hemolysator ersetzt. Justage und Service tHb Kal SO2 wieder nahe ZW | SWO-56135 | No | Hemolysator. Adjustment and service tHb Kal SO2 again close to ZW | Hemolysator. Adjustment and service tHb Kal SO2 again close to ZW | Defect | HW | Oxi module | Hemolyzer | COHb
ABL800 BASIC | ヘモライザー交換、OXIバルブ分解・清掃・潤滑 コンプリートテスト確認 CAL、tHb・StHbCAL、QC測定データ確認 | ヘモライザー交換、OXIバルブ分解・清掃・潤滑 コンプリートテスト確認 CAL、tHb・StHbCAL、QC測定データ確認 | SWO-55418 | No | Hemolyzer replacement, OXI valve disassembly, cleaning and lubrication Complete test confirmation CAL, tHb/StHbCAL, QC measurement data confirmation | Hemolyzer replacement, OXI valve disassembly, cleaning and lubrication Complete test confirmation CAL, tHb/StHbCAL, QC measurement data confirmation | Defect | HW | Oxi module | Hemolyzer
ABL825 FLEX | Patientenberichtslayout angepasst Query Abfrage deaktiviert Mit Edv, Labor und Pflegepersonal Übertragung getestet | Patientenberichtslayout angepasst Query Abfrage deaktiviert Mit Edv, Labor und Pflegepersonal Übertragung getestet | SWO-55270 | No | Patient report layout customized query query disabled With Edv, lab and nurse transmission tested | Patient report layout customized query query disabled With Edv, lab and nurse transmission tested | Setup | HIS/LIS
ABL80 FLEX BASIC | 更换新测试卡修复机器 | 更换新测试卡修复机器 | SWO-55480 | No | Replace the new test card repair machine | Replace the new test card repair machine
ABL80 FLEX BASIC | 更换测试卡正常。 | 更换测试卡正常。 | SWO-55484 | No | The replacement test card is OK. | The replacement test card is OK.
ABL90 FLEX | Tausch der Einlassdichtung und der Einlasssonde, da krum, und der grauen Anschlussdichtung. Analysator durchgespült vom Ablassstutzen. Leckage behoben. Kein weiterer Fehler vorhanden. Info an Vertriebskollegen zur Kontaktaufnahme wegen Tiefeneinweisung. Einlassonde und Einkassdichtung Vorort gelassen, da keine am Analysator. Aktuelle Software eingespielt, siehe extra Servicereport. tHb Kalibration fällig, mit eigenem Kalibrator durchgeführt. Analysator messbereit. | Tausch der Einlassdichtung und der Einlasssonde, da krum, und der grauen Anschlussdichtung. Analysator durchgespült vom Ablassstutzen. Leckage behoben. Kein weiterer Fehler vorhanden. Info an Vertriebskollegen zur Kontaktaufnahme wegen Tiefeneinweisung. Einlassonde und Einkassdichtung Vorort gelassen, da keine am Analysator. Aktuelle Software eingespielt, siehe extra Servicereport. tHb Kalibration fällig, mit eigenem Kalibrator durchgeführt. Analysator messbereit. | SWO-55610 | No | Replacement of the inlet seal and the inlet probe, there krum, and the grey connection seal. Analyzer rinsed from drain nozzle. Leakage fixed. No further error exists. Info to sales colleagues to contact for in-depth instruction. Inlet probe and cash-in seal suburb left, as none on the analyzer. Current software, see extra Servicereport. tHb calibration due, carried out with its own calibrator. Analyzer ready to measure. | Replacement of the inlet seal and the inlet probe, there krum, and the grey connection seal. Analyzer rinsed from drain nozzle. Leakage fixed. No further error exists. Info to sales colleagues to contact for in-depth instruction. Inlet probe and cash-in seal suburb left, as none on the analyzer. Current software, see extra Servicereport. tHb calibration due, carried out with its own calibrator. Analyzer ready to measure. | Replaced | Probe | Bend
ABL90 | Troubleshot and found that the Windows time was incorrect. | Corrected Windows time to local time, and this resolved the problem. | CF9A20C5D41F59AB | No | Troubleshot and found that the Windows time was incorrect. | Corrected Windows time to local time, and this resolved the problem. | Setup | Time Setting
ABL90 | Troubleshot and found that the Windows time was incorrect. | Corrected Windows time to local time, and this resolved the problem. | CF9A1327D43A03ED | No | Troubleshot and found that the Windows time was incorrect. | Corrected Windows time to local time, and this resolved the problem. | Setup | Time Setting
ABL90 |  |  | CF9DC5CF36348AF4 | No |  |  | Assisted Replacement | SC | Cl
ABL90 FLEX | Fehlersuche, Pressure Test und Sol.Pack Check durchgefallen, manuelle Spülung mehrmals durchgeführt, Einlasssonde ok, Einlassdichtung mit Halter erneuert, danach alle Service Tests bestanden, LS Justage und Pumpenkal. ok System-und Funktionskontrolle i.O. | Fehlersuche, Pressure Test und Sol.Pack Check durchgefallen, manuelle Spülung mehrmals durchgeführt, Einlasssonde ok, Einlassdichtung mit Halter erneuert, danach alle Service Tests bestanden, LS Justage und Pumpenkal. ok System-und Funktionskontrolle i.O. | SWO-55813 | No | Troubleshooting, Pressure Test and Sol.Pack Check failed, manual flushing performed several times, inlet probe ok, inlet seal with holder renewed, then passed all service tests, LS adjustment and pump scale. ok System and function control i.O. | Troubleshooting, Pressure Test and Sol.Pack Check failed, manual flushing performed several times, inlet probe ok, inlet seal with holder renewed, then passed all service tests, LS adjustment and pump scale. ok System and function control i.O. | Defect | HW | Inlet gasket with holder
ABL90 FLEX | Roma.ini korrigiert, LP zur Aktivierung getauscht, Kal u. QC i.O. |  | SWO-55693 | No | Roma.ini corrected, LP exchanged for activation, Kal & QC i.O. |  | Unknown | Fixed assigned values
ABL90 FLEX | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-55684 | No | Repair after replacing a new test card | Repair after replacing a new test card | Assisted Replacement | SC | Reference
AQT90 FLEX |  |  | SWO-55454 | No |  |  | Defect | HW | Analyzer control Module | Barcode reader, display
AQT90 FLEX |  |  | SWO-55461 | No |  |  | Defect | HW | Analyzer control Module | Barcode reader, display
ABL805 FLEX | checked logs for errors checked inlet gasket, tubes, probe, sensors, | changed out lower inlet and lower ph BG sensor tubes gasket and probe and put in new nipples. | CF9A155DD424C78E | Yes | checked logs for errors checked inlet gasket, tubes, probe, sensors, | changed out lower inlet and lower ph BG sensor tubes gasket and probe and put in new nipples. | Re classified RMED | SenDx OPS
ABL835 FLEX | ミキサートレイコンプリート交換 | ミキサートレイコンプリート交換 | SWO-56352 | No | Mixer Tray Complete Replacement | Mixer Tray Complete Replacement | Defect | HW | FlexQ Module | Barcode Reader
ABL825 FLEX | Error code 314. OXI wavelength deviation high. | Replace Spectrometer unit. Perform lamp alignment. Perfor Complete OXI Test. Perform Service tHb calibration. | CF9A3C6CD4475A57 | No | Error code 314. OXI wavelength deviation high. | Replace Spectrometer unit. Perform lamp alignment. Perfor Complete OXI Test. Perform Service tHb calibration. | Defect | HW | Oxi module | Spectrometer
ABL835 FLEX | FLEXQモジュール修理完了の為、返却。代替えモジュール引き上げ | FLEXQモジュール修理完了の為、返却。代替えモジュール引き上げ | SWO-55682 | No | Returned for flexq module repair completion. Alternative module pull | Returned for flexq module repair completion. Alternative module pull | Defect | HW | FlexQ Module
ABL800 BASIC | プレッシャートランスデューサー交換 |  | SWO-55405 | No | Pressure transducer replacement |  | Defect | HW | Gas module | Pressure Transducer
ABL825 FLEX | SW neu install. bricht ab. Gerät bootet nicht mehr. ABL mit Absprache Stat.Leitung außer Betrieb gesetzt. | SW neu install. bricht ab. Gerät bootet nicht mehr. ABL mit Absprache Stat.Leitung außer Betrieb gesetzt. | SWO-56145 | No | SW new install. aborts. Device no longer boots. ABL with agreement Stat.Management decommissioned. | SW new install. aborts. Device no longer boots. ABL with agreement Stat.Management decommissioned. | Unknown | HW | DMS_PC | unknown
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-55683 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL80 FLEX BASIC | 更换测试卡正常。 | 更换测试卡正常。 | SWO-55663 | No | The replacement test card is OK. | The replacement test card is OK.
ABL80 FLEX BASIC | 更换测试卡正常。 | 更换测试卡正常。 | SWO-55672 | No | The replacement test card is OK. | The replacement test card is OK.
ABL90 FLEX | 更换测试卡正常。 | 更换测试卡正常。 | SWO-55674 | No | The replacement test card is OK. | The replacement test card is OK. | Assisted Replacement | SC | Reference
AQT90 FLEX |  |  | SWO-55670 | No |  |  | Defect | HW | Storage Module |  | BR1
AQT90 FLEX | 更换传感器636-549一个和进样针903-171一个后修复 | 更换传感器636-549一个和进样针903-171一个后修复 | SWO-55473 | No | Replace sensor 636-549 one and sample needle 903-171 one after repair | Replace sensor 636-549 one and sample needle 903-171 one after repair | Defect | HW | Aspiration Module | Needle sensor | 4974
ABL80 FLEX BASIC | 更换测试卡正常。 | 更换测试卡正常。 | SWO-55766 | No | The replacement test card is OK. | The replacement test card is OK.
ABL80 FLEX BASIC | 更换测试卡正常。 | 更换测试卡正常。 | SWO-55741 | No | The replacement test card is OK. | The replacement test card is OK.
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-55952 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL80 FLEX BASIC | 更换测试卡，更换902-972一个正常。 | 更换测试卡，更换902-972一个正常。 | SWO-55762 | No | Replace the test card and replace the 902-972 one normal. | Replace the test card and replace the 902-972 one normal.
ABL80 FLEX BASIC | SC flushed and reinstalled- Parameteres falling again- PER Made- Need to replace- Ms Rose mary | Working fine, Calibrated and Serviced properly. Meating done with ms rose mary about Consumable inventory. | SWO-56550 | No | SC flushed and reinstalled- Parameteres falling again- PER Made- Need to replace- Ms Rose mary | Working fine, Calibrated and Serviced properly. Meating done with ms rose mary about Consumable inventory.
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-57821 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL90 FLEX | Allgemeine Kontrolle ok. Gerät durchgespült ok. Mit zwei verschieden Lösungspack getestet, gleicher Fehler. Sensorkassette ausgetauscht. 55 min Initialisierung -> werde Kunde morgen kontaktieren |  | SWO-56026 | No | General control ok. Device flushed ok. Tested with two different solution packs, same error. Sensor cartridge. 55 min initialization -> will contact customer tomorrow |  | Assisted Replacement | SC | Wetability
ABL90 |  |  | CFA03A552240D35B | No |  |  | Defect | SW | Windows
ABL90 FLEX | neue SC eingebaut, Gerät kalibriert, neue SC R0396 | neue SC eingebaut, Gerät kalibriert, neue SC R0396 | SWO-55868 | No | new SC built-in, device calibrated, new SC R0396 | new SC built-in, device calibrated, new SC R0396 | Assisted Replacement | SC | Unknown
ABL825 FLEX | Flachbandkabel defekt, getauscht, Kal und QC IO | Flachbandkabel defekt, getauscht, Kal und QC IO | SWO-55876 | No | Ribbon cable defective, replaced, Kal and QC IO | Ribbon cable defective, replaced, Kal and QC IO | Defect | HW | AutoCheck Module | FPC Cable
ABL825 FLEX | pCO2 result is 16.5 on the last qc level 3 that ran this morning. | Cleaned the pH electrode and replaced the blue seal ring. Also, cleaned the dummy electrode and replaced the seal ring. Cleaned the pCO2 electrode and, replaced the seal ring. Note: Left Patricia 2 pH seal ring to keep as spares. Show her how to replace the ring. | CF9DD9063651BFEB | No | pCO2 result is 16.5 on the last qc level 3 that ran this morning. | Cleaned the pH electrode and replaced the blue seal ring. Also, cleaned the dummy electrode and replaced the seal ring. Cleaned the pCO2 electrode and, replaced the seal ring. Note: Left Patricia 2 pH seal ring to keep as spares. Show her how to replace the ring. | Assisted Replacement | Electrode | Electrode | pCO2
ABL800 BASIC | Netzteil ausgetauscht, Fehler blieb weiterhin bestehen . In die GS gefahren, Bildschirm und CPU Unit geholt. Zuerst den Bildschirm ausgetauscht,keine Veränderung, Nachdem CPU Austausch ist der Fehler behoben. Kal mehrfach durchgeführt. QC i.O | Netzteil ausgetauscht, Fehler blieb weiterhin bestehen . In die GS gefahren, Bildschirm und CPU Unit geholt. Zuerst den Bildschirm ausgetauscht,keine Veränderung, Nachdem CPU Austausch ist der Fehler behoben. Kal mehrfach durchgeführt. QC i.O | SWO-55883 | No | Power supply replaced, errors remained. Driven into the GS, screen and CPU unit fetched. First the screen is replaced,no change, After CPU exchange the error is fixed. Kal performed several times. QC i.O | Power supply replaced, errors remained. Driven into the GS, screen and CPU unit fetched. First the screen is replaced,no change, After CPU exchange the error is fixed. Kal performed several times. QC i.O | Defect | HW | DMS_PC | DMS_PC
ABL825 FLEX | 597 error. Liquid detected in pH BG module during gas calibration. | Replaced membrane in pH BG module flat membrane valve. Performed refill, pump cal and LS adjustment. | CF9DFBC736134073 | No | 597 error. Liquid detected in pH BG module during gas calibration. | Replaced membrane in pH BG module flat membrane valve. Performed refill, pump cal and LS adjustment. | Defect | HW | pH/BG Module | Valve, air | membrane
ABL825 FLEX | Überprüfung, Ausbau und Reinigung der Measuring Ventile Austausch Verschlauchung Inlet zu pH BG-Modul, da kleine Leckage Spülung AirKanal => Nachfüllung i.O. Bei LS-Justage ist Durchfluss durch Einlassdichtung behindert => EinlassKal ohne Besserung => Austausch Einlassdichtung aus Kundenbestand LS-Justage und PumpenKal i.O. Durchführung Kal1, Kal2 i.O. QC wird nicht richtig angesaugt => Druckabfall -Aufbau in Fließsystem Test mit neuer Einlassnadel und Einlassschlauch => keine Besserung Einbau Einlassmodul für morgen | Überprüfung, Ausbau und Reinigung der Measuring Ventile Austausch Verschlauchung Inlet zu pH BG-Modul, da kleine Leckage Spülung AirKanal => Nachfüllung i.O. Bei LS-Justage ist Durchfluss durch Einlassdichtung behindert => EinlassKal ohne Besserung => Austausch Einlassdichtung aus Kundenbestand LS-Justage und PumpenKal i.O. Durchführung Kal1, Kal2 i.O. QC wird nicht richtig angesaugt => Druckabfall -Aufbau in Fließsystem Test mit neuer Einlassnadel und Einlassschlauch => keine Besserung Einbau Einlassmodul für morgen | SWO-55879 | Yes | Inspection, removal and cleaning of measuring valves replacement hose inlet to pH BG module, since small leakage flushing AirKanal => refilling i.O. With LS-Adjustment, flow is hindered by inlet seal => InletCal without improvement => Replacement inlet seal from customer stock LS-Adjustment and pumpsKal i.O. Execution Kal1, Kal2 i.O. QC is not properly sucked in => Pressure drop -construction in flow system test with new inlet needle and inlet hose => no improvement installation inlet module for tomorrow | Inspection, removal and cleaning of measuring valves replacement hose inlet to pH BG module, since small leakage flushing AirKanal => refilling i.O. With LS-Adjustment, flow is hindered by inlet seal => InletCal without improvement => Replacement inlet seal from customer stock LS-Adjustment and pumpsKal i.O. Execution Kal1, Kal2 i.O. QC is not properly sucked in => Pressure drop -construction in flow system test with new inlet needle and inlet hose => no improvement installation inlet module for tomorrow | Re classified RMED | Double upSWO
ABL815 FLEX | Messkammer gereinigt alle Schläuche BG pH Modul ersetzt. An der Dammyelektrode und der pH Elektrode Dichtringe ersetzt Kalibration und QC gemessen>> QC Verletzung. pCO2 Elektrode probeweise eingebaut. Beide Gasventile probeweise eingebaut. Beide Gasflaschen ersetzt. Nach wie vor QC Verletzung pCO2. AU PCB und BG Modul bestellt. |  | SWO-55995 | No | Measuring chamber cleaned all hoses BG pH module replaced. At the dammy electrode and the pH electrode sealing rings replaces calibration and QC measured>> QC violation. pCO2 electrode installed on a trial basis. Both gas valves installed on a trial basis. Replaced both gas cylinders. Still QC violation pCO2. AU PCB and BG module. |  | Revisit required | HW | pH/BG Module | unknown
ABL825 FLEX | Die Taste Start ist unterlegt, reagiert aber nicht. Nach dem Neustart ist immer das Problem behoben. Ich habe provisorisch die Safe Tipcap Starttaste noch hinzugefügt. | Die Taste Start ist unterlegt, reagiert aber nicht. Nach dem Neustart ist immer das Problem behoben. Ich habe provisorisch die Safe Tipcap Starttaste noch hinzugefügt. | SWO-55881 | No | The Start button is underpinned, but does not respond. After restarting, the problem is always resolved. I temporarily added the Safe Tipcap Start button. | The Start button is underpinned, but does not respond. After restarting, the problem is always resolved. I temporarily added the Safe Tipcap Start button. | Unknown | SW Slow
AQT90 FLEX | Alufolienschnipsel von Lichtsensor entfernt. Systemüberprüfung durchgeführt, in Ordnung. Kalibrierung NT-proBNP gestartet. | Alufolienschnipsel von Lichtsensor entfernt. Systemüberprüfung durchgeführt, in Ordnung. Kalibrierung NT-proBNP gestartet. | SWO-55892 | No | Removed aluminium foil snippets from light sensor. System check performed, all right. Calibration NT-proBNP started. | Removed aluminium foil snippets from light sensor. System check performed, all right. Calibration NT-proBNP started. | Cleaned | HW | Incubator Module | Cup Sensor | 2167
ABL80 FLEX BASIC | 更换新试剂包后修复 | 更换新试剂包后修复 | SWO-56119 | No | Repair after replacing a new kit | Repair after replacing a new kit
ABL80 CO-OX | Sensor cassette connector PCB corroded. Leak from sol pack. | Replace sensor connector PCB. Replace solution pack. Replace inlet tubing. | CF9F3414AC4D4942 | No | Sensor cassette connector PCB corroded. Leak from sol pack. | Replace sensor connector PCB. Replace solution pack. Replace inlet tubing.
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-56028 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL90 FLEX | Kontakte wurden justiert Gerät wieder Messbereit Info Meldung 1405 | Kontakte wurden justiert Gerät wieder Messbereit Info Meldung 1405 | SWO-56078 | No | Contacts have been adjusted device again ready to measure Info Message 1405 | Contacts have been adjusted device again ready to measure Info Message 1405 | Adjustment | HW | SP PCB
ABL90 | Inspection of the unit found no issues mechanically. Lactate is corrected for ion strength, calculated from K+, Na+, Ca2+. | Per info povided from QA (Jim G.) & approval from Cindy Ghandour (RT Dir.) electrolytes were enabled to allow corrections to LAC to pass level 3 external QC. All electrolytes were removed from the customer's patient report layout. | CF9F0B65AC162867 | No | Inspection of the unit found no issues mechanically. Lactate is corrected for ion strength, calculated from K+, Na+, Ca2+. | Per info povided from QA (Jim G.) & approval from Cindy Ghandour (RT Dir.) electrolytes were enabled to allow corrections to LAC to pass level 3 external QC. All electrolytes were removed from the customer's patient report layout. | Setup | Lac QC | Electrolytes
ABL90 FLEX | Fehlersuche, Neustart erfolglos Austausch DisplayUnit ABL hängt sich beim hochfahren auf, Software installiert V3.1MR5 Datenbankfehler, SW erneut installiert, Pointer justiert Konfiguration eingestellt SC 90 erneuert, da Lac Kalibrierfehler Einlaufverfahren |  | SWO-55831 | No | Troubleshooting, restart unsuccessful replacement DisplayUnit ABL hangs on when booting up, software installs V3.1MR5 database error, SW reinstalled; Pointer adjusted configuration set SC 90 renewed, since Lac calibration error inlet method |  | Defect | HW | Touch screen
ABL90 FLEX | Es wurde schon ein neuer LP eingesetzt, aber Gerät bleibt im Menü stehen, Aus Ein bringt keine Änderung, lt. Anwenderin gibt das Gerät keinen Grund dafür an, bei Eintreffen festgestellt das auf dem Bildschirm die Meldung Verfallsdatum überschritten steht, vorherigen LP ( Chipdaten können nicht gelesen werden) installiert, Neustart i.O., Einlassdichtung ersetzt, Kal und QC durchgeführt, QC Level A pCO2 oberhalb der Grenze, SC ersetzt, Kal und QC durchgeführt, Level A pCO2 weiterhin leicht zu hoch, kein LP mehr vor Ort, Anwender wird LP bestellen und ersetzten | Es wurde schon ein neuer LP eingesetzt, aber Gerät bleibt im Menü stehen, Aus Ein bringt keine Änderung, lt. Anwenderin gibt das Gerät keinen Grund dafür an, bei Eintreffen festgestellt das auf dem Bildschirm die Meldung Verfallsdatum überschritten steht, vorherigen LP ( Chipdaten können nicht gelesen werden) installiert, Neustart i.O., Einlassdichtung ersetzt, Kal und QC durchgeführt, QC Level A pCO2 oberhalb der Grenze, SC ersetzt, Kal und QC durchgeführt, Level A pCO2 weiterhin leicht zu hoch, kein LP mehr vor Ort, Anwender wird LP bestellen und ersetzten | SWO-56151 | No | A new LP has already been inserted, but device remains in the menu, From A does not change, according to the user does not give the device a reason for it when it arrives, the message expiry date is exceeded on the screen, previous LP ( chip data can not be read) installed, restart usually replaced, inlet seal replaced, Kal and QC performed, QC level A pCO2 above the limit, SC replaced, Kal and QC performed, level A pCO2 still slightly too high, no more LP on site, user will order and replace LP | A new LP has already been inserted, but device remains in the menu, From A does not change, according to the user does not give the device a reason for it when it arrives, the message expiry date is exceeded on the screen, previous LP ( chip data can not be read) installed, restart usually replaced, inlet seal replaced, Kal and QC performed, QC level A pCO2 above the limit, SC replaced, Kal and QC performed, level A pCO2 still slightly too high, no more LP on site, user will order and replace LP | User Error | SP Expired
ABL90 | found that would not pass c level of internal qc | replaced sp did not fix the issue. replaced the SC on analyzer which corrected the issue. | CF9F68B3AC1A44A8 | No | found that would not pass c level of internal qc | replaced sp did not fix the issue. replaced the SC on analyzer which corrected the issue. | Assisted Replacement | SC | Glu
ABL90 | Came in to gather data logs per Jim Halbleib and ASM. Asking for SP replacement and customer does not have their replacements yet. Don Masters, ASM, was able to get permission to load 3.1 MR6 SW. | Flushed thoroughly. Will send data logs to Jim Halbleib. | CF9F28B6AC0B73E8 | No | Came in to gather data logs per Jim Halbleib and ASM. Asking for SP replacement and customer does not have their replacements yet. Don Masters, ASM, was able to get permission to load 3.1 MR6 SW. | Flushed thoroughly. Will send data logs to Jim Halbleib. | Assisted Replacement | SP | pO2
ABL90 | Upon arrival, analyzer was in hold, wanting a reagent to be changed. Ejected SP and installed different ones . Analyzer turned around ok, but customer explained that the longest the analyzer is going to be functional for is 8 hours. | Started CQC, performed PM, upgraded software to 3.1MR6 as approved by ASM and VP Sales. Captured data logs. | CF9F9A31AC104524 | No | Upon arrival, analyzer was in hold, wanting a reagent to be changed. Ejected SP and installed different ones . Analyzer turned around ok, but customer explained that the longest the analyzer is going to be functional for is 8 hours. | Started CQC, performed PM, upgraded software to 3.1MR6 as approved by ASM and VP Sales. Captured data logs. | Assisted Replacement | SP | FT
ABL90 FLEX | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-58249 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL90 FLEX | 更换新试剂包后修复 | 更换新试剂包后修复 | SWO-56199 | No | Repair after replacing a new kit | Repair after replacing a new kit | Assisted Replacement | SP | FT
AQT90 FLEX | 重新定标正常。 | 重新定标正常。 | SWO-55951 | No | Recalibration is normal. | Recalibration is normal. | Assisted Replacement | Cartridge | TnI | Calibration | 167
ABL825 FLEX | Austausch der schwarzen Verschlauchung und der Pumpenschläuche => Keine Besserung Austausch unteren Dichtring pH, da Messkammer vollgelaufen Austausch Lac, Glu und Cl Membran, da laut Aktivitäten gestern Dekontamination durchgeführt wurde Austausch Einlassmodul aus GS Gerät => Kundenmodul zur Überprüfung in GS Kal i.O. => im QC Level 3 Ca außer Bereich => Ca und Ref neu positioniert => QC Level 3 i.O. QC Level 4 nicht durchführbar da nicht vorrätig | Überprüfung, Ausbau und Reinigung der Measuring Ventile Austausch Verschlauchung Inlet zu pH BG-Modul, da kleine Leckage Spülung AirKanal => Nachfüllung i.O. Bei LS-Justage ist Durchfluss durch Einlassdichtung behindert => EinlassKal ohne Besserung => Austausch Einlassdichtung aus Kundenbestand LS-Justage und PumpenKal i.O. Durchführung Kal1, Kal2 i.O. QC wird nicht richtig angesaugt => Druckabfall -Aufbau in Fließsystem Test mit neuer Einlassnadel und Einlassschlauch => keine Besserung Einbau Einlassmodul für morgen | SWO-56021 | No | Replacement of the black hose and the pump hoses => No improvement replacement lower sealing ring pH, since measuring chamber full exchange Lac, Glu and Cl membrane, since according to activities yesterday decontamination was carried out replacement inlet module from GS device => Customer module for verification in GS Kal i.O. => in QC Level 3 Ca out of range => Ca and Ref repositioned => QC Level 3 i.O. QC Level 4 not feasible as not in stock | Inspection, removal and cleaning of measuring valves replacement hose inlet to pH BG module, since small leakage flushing AirKanal => refilling i.O. With LS-Adjustment, flow is hindered by inlet seal => InletCal without improvement => Replacement inlet seal from customer stock LS-Adjustment and pumpsKal i.O. Execution Kal1, Kal2 i.O. QC is not properly sucked in => Pressure drop -construction in flow system test with new inlet needle and inlet hose => no improvement installation inlet module for tomorrow | Revisit required | HW | Inlet module | unknown
ABL825 FLEX | Gerät überprüft, Vorverstärker kalibriert, pO2 Elektrode (RT04) membraniert und erneuert, Neustart durchgeführt, pO2 Elektrode aus Laborgerät eingesetzt, Kal und QC o.k. | Gerät überprüft, Vorverstärker kalibriert, pO2 Elektrode (RT04) membraniert und erneuert, Neustart durchgeführt, pO2 Elektrode aus Laborgerät eingesetzt, Kal und QC o.k. | SWO-56110 | No | Device checked, preamplifier calibrated, pO2 electrode (RT04) membraneed and renewed, restart performed, pO2 electrode used from laboratory device, Kal and QC ok. | Device checked, preamplifier calibrated, pO2 electrode (RT04) membraneed and renewed, restart performed, pO2 electrode used from laboratory device, Kal and QC ok. | Assisted Replacement | Electrode | Electrode | pO2
ABL825 FLEX | Gerät überprüft, Touch Screen gesäubert und kalibriert Keine Besserung | Gerät überprüft, Touch Screen gesäubert und kalibriert Keine Besserung | SWO-56111 | No | Device checked, touch screen cleaned and calibrated No improvement | Device checked, touch screen cleaned and calibrated No improvement | Unknown | HW | Touch screen | unknown
ABL815 FLEX | Au PCB und pH BG Modul PCB ersetzt. AU Software und Kay installiert. Startphase QC und Kalibration. |  | SWO-55998 | No | Au PCB and pH BG module PCB replaced. AU Software and Kay installed. Start phase QC and calibration. |  | Defect | HW | pH/BG Module | pCO2
ABL827 FLEX | Checked the data logs for a cause and could not find a cause. | Printed the set up and then reset the setup to manufacturers specifications. Entered the setup information manuall. Ran 4 test blood samples to verify that the problem is solved. | CF9F3075AC168256 | No | Checked the data logs for a cause and could not find a cause. | Printed the set up and then reset the setup to manufacturers specifications. Entered the setup information manuall. Ran 4 test blood samples to verify that the problem is solved. | Unknown | SW Slow
TCM/Tosca |  |  | CF9F1CBBAC440784 | No |  |  | Defect | HW | Electrode Switch | Electrode Switch
TCM/Tosca |  |  | CF9F2CB2AC540786 | No |  |  | Defect | HW | PCB, Main CPU | PCB, Main CPU
ABL80 FLEX BASIC | 更换两张测试卡正常。 | 更换两张测试卡正常。 | SWO-56154 | No | Replacing the two test cards is ok. | Replacing the two test cards is ok.
ABL80 FLEX BASIC | SC replaced- Analyser serviced and Calibrated- Working fine- Dr Gupta, Dr jaydeep, Brother vinod | Working fine, Calibrated and Serviced properly. Meating done with ms rose mary about Consumable inventory. | SWO-56565 | No | SC replaced- Analyser serviced and Calibrated- Working fine- Dr Gupta, Dr jaydeep, Brother vinod | Working fine, Calibrated and Serviced properly. Meating done with ms rose mary about Consumable inventory.
ABL80 CO-OX | Joe investigated external CD drives and Radiometer researched and tested the drive before installation. | Installed external Cd Drive on ABL80 | CFA0621322272882 | No | Joe investigated external CD drives and Radiometer researched and tested the drive before installation. | Installed external Cd Drive on ABL80
ABL90 FLEX | Fehlersuche, Test im Service Programm durchlaufen lassen GERÄTENEUSTART | Fehlersuche, Test im Service Programm durchlaufen lassen GERÄTENEUSTART | SWO-56304 | No | Troubleshooting, testing in the service program let device NEUSTART | Troubleshooting, testing in the service program let device NEUSTART | Adjustment | Restart | SP ??
ABL90 | CONFIRMED ERROR WHEN ARRIVED ONSITE. | PROCEED TO UPGRADE SOFTWARE FROM VERSION 3.1 MR3 TO MR6. RENEE ALSO REQUESTED TO ADD HER CONTACTS IN CREDIT REQUEST FORM. | CFA047B322604B28 | No | CONFIRMED ERROR WHEN ARRIVED ONSITE. | PROCEED TO UPGRADE SOFTWARE FROM VERSION 3.1 MR3 TO MR6. RENEE ALSO REQUESTED TO ADD HER CONTACTS IN CREDIT REQUEST FORM. | Assisted Replacement | SC | Wetability | MR6 installed
ABL90 FLEX | Konfiguration von Mehreren Geräten ausprobiert, immer wieder Fehlermeldung. ABL SW mehrmals installiert, Bios Einstellungen auf 8 umgestellt, ohne Erfolg. | Die Auslieferung abgebrochen und an RDEU Service gemeldet. | SWO-56657 | Yes | Configuration of multiple devices tried, repeatedly error message. ABL SW installed several times, Bios settings changed to 8, to no avail. | Delivery aborted and reported to RDEU Service. | Re classified RMED | installation og analyzer | 6 upSWO on 3 analyzer
ABL90 FLEX | Konfiguration von Mehreren Geräten ausprobiert, immer wieder Fehlermeldung. ABL SW mehrmals installiert, Bios Einstellungen auf 8 umgestellt, ohne Erfolg. | Die Auslieferung abgebrochen und an RDEU Service gemeldet. | SWO-56662 | Yes | Configuration of multiple devices tried, repeatedly error message. ABL SW installed several times, Bios settings changed to 8, to no avail. | Delivery aborted and reported to RDEU Service. | Re classified RMED | installation og analyzer | 6 upSWO on 3 analyzer
ABL90 FLEX | Konfiguration von Mehreren Geräten ausprobiert, immer wieder Fehlermeldung. ABL SW mehrmals installiert, Bios Einstellungen auf 8 umgestellt, ohne Erfolg. | Die Auslieferung abgebrochen und an RDEU Service gemeldet. | SWO-56650 | Yes | Configuration of multiple devices tried, repeatedly error message. ABL SW installed several times, Bios settings changed to 8, to no avail. | Delivery aborted and reported to RDEU Service. | Re classified RMED | installation og analyzer | 6 upSWO on 3 analyzer
ABL90 | Shortly after power up units video went blank.Loss of 12 volt supply. | New unit to be ordered. | CFA1EEDB9851BA78 | Yes | Shortly after power up units video went blank. Loss of 12 volt supply. | New unit to be ordered. | Re classified RMED | installation of analyzer
ABL90 FLEX PLUS |  |  | CFA0C5B6224491AC | Yes |  |  | Re classified RMED | Logistical upSWO
AQT90 FLEX |  |  | SWO-56185 | No |  |  | Defect | HW | Process Module | Optical Unit | 131
AQT90 FLEX |  |  | SWO-56186 | No |  |  | Defect | HW | Process Module | Optical Unit | 131
AQT90 FLEX |  |  | CFA05CE122424BA3 | No |  |  | Defect | HW | Incubator Module | Incubator, PCB | 2168
ABL80 FLEX BASIC | Working fine, Calibrated and Serviced properly. Meating done with ms rose mary about Consumable inventory. | Working fine, Calibrated and Serviced properly. Meating done with ms rose mary about Consumable inventory. | SWO-56566 | No | Working fine, Calibrated and Serviced properly. Meating done with ms rose mary about Consumable inventory. | Working fine, Calibrated and Serviced properly. Meating done with ms rose mary about Consumable inventory.
ABL90 FLEX | Kifogyó oldatpakk hibás volt. Oldatpakk csere után megfelelően működik a készülék. Kalibrálás, QC mérés. |  | SWO-56398 | No | The solution pack was defective. After replacing the solution pack, the appliance is working properly. Calibration, QC measurement. |  | Assisted Replacement | SP | Unknown
ABL90 FLEX | SP richtig eingelegt Im Service Fluid Transport überprüft Einlassnadel getauscht | SP richtig eingelegt Im Service Fluid Transport überprüft Einlassnadel getauscht | SWO-56380 | No | SP correctly inserted In service Fluid Transport checked inlet needle exchanged | SP correctly inserted In service Fluid Transport checked inlet needle exchanged | User Error | SP activation
ABL90 FLEX | Einweisung Gerätebetreuer ITS OP am ABL 90 Fr. Witte Hr. Weiß Ausdrucken Credit Claim an ABL 90 der UniKlinikums Jena zwecks Reklamation etc. | Einweisung Gerätebetreuer ITS OP am ABL 90 Fr. Witte Hr. Weiß Ausdrucken Credit Claim an ABL 90 der UniKlinikums Jena zwecks Reklamation etc. | SWO-56515 | Yes | Instruction equipment supervisor ITS OP at ABL 90 Ms. Witte Mr. White Print Credit Claim to ABL 90 of the UniKlinikums Jena for complaint etc. | Instruction equipment supervisor ITS OP at ABL 90 Ms. Witte Mr. White Print Credit Claim to ABL 90 of the UniKlinikums Jena for complaint etc. | Re classified RMED | Logistical upSWO
ABL90 FLEX | Barcodescanner erneuet, Test i.O | Barcodescanner erneuet, Test i.O | SWO-56385 | No | Barcode scanner is new, test i.O | Barcode scanner is new, test i.O | Defect | HW | Barcode Reader
ABL90 FLEX | Die Auslieferung abgebrochen und an RDEU Service gemeldet. | Die Auslieferung abgebrochen und an RDEU Service gemeldet. | SWO-56652 | Yes | Delivery aborted and reported to RDEU Service. | Delivery aborted and reported to RDEU Service. | Re classified RMED | installation og analyzer | 6 upSWO on 3 analyzer
ABL825 FLEX | OXI modul Neon lámpa csere, OXI tesztek, Kalibrálások, QC mérés hitelesítés |  | SWO-56494 | No | OXY module Neon lamp replacement, OXY tests, calibrations, QC measurement authentication |  | Defect | HW | Oxi module | Lamp unit
ABL835 FLEX | Monitor kábel újracsatlakoztatása, Képernyő kalibrálás, QC mérés |  | SWO-56420 | No | Reconnecting Monitor cable, screen calibration, QC measurement |  | Adjustment | HW | DMS_PC | Touch screen
ABL825 FLEX | Proteinentfernung durchgeführt Ventile im Ventiltest überprüft, ok Luftweg im Fluidic Modul gespült, pH Dichtring erneuert LS Justage durchgeführt, Auto Check Modul befüllt Gerät kalibriert und überprüft | Proteinentfernung durchgeführt Ventile im Ventiltest überprüft, ok Luftweg im Fluidic Modul gespült, pH Dichtring erneuert LS Justage durchgeführt, Auto Check Modul befüllt Gerät kalibriert und überprüft | SWO-56519 | No | Protein removal carried out valves in valve test checked, ok airway flushed in fluidic module, pH sealing ring renewed LS adjustment performed, Auto Check module filled device calibrated and checked | Protein removal carried out valves in valve test checked, ok airway flushed in fluidic module, pH sealing ring renewed LS adjustment performed, Auto Check module filled device calibrated and checked | Cleaned | Fluidic pathway | Fluidic manifold | Air vent
ABL825 FLEX | Touch Screen erneuert, Kalibration und Funktionsüberprüfung o.k. | Gerät überprüft, Touch Screen gesäubert und kalibriert Keine Besserung | SWO-56531 | No | Touch screen renewed, calibration and function verification ok. | Device checked, touch screen cleaned and calibrated No improvement | Defect | HW | DMS_PC | Touch screen
ABL835 FLEX |  |  | CFA1C2C29800C637 | No |  |  | Assisted Replacement | Electrode | Membrane | Reference
AQT90 FLEX |  |  | CFA2B8B10E37B10F | No |  |  | Cleaned | HW | Process Module | Process wheel
ABL80 FLEX CO-OX | ASTM.ini ersetzt, Kal u. IQC durchgeführt, Kommunikation überprüft, | ASTM.ini ersetzt, Kal u. IQC durchgeführt, Kommunikation überprüft, | SWO-56632 | No | ASTM.ini, Kal & IQC performed, communication checked, | ASTM.ini, Kal & IQC performed, communication checked,
ABL90 FLEX | Anwender ein Leihgerät aufgestellt, in extra SWO dokumentiert. Es handelt sich um ein ABL90 Flex Plus der weniger Probenvolumen benötigt. Gleichzeitig Test des Analysators. Vorteilhaft für MBU Messungen. Einweisung der Anwender, Ärzte und Hebammen, notwendig. Meldung an Vertrieb erfolgt. Analysator funktionsbereit. | Anwender ein Leihgerät aufgestellt, in extra SWO dokumentiert. Es handelt sich um ein ABL90 Flex Plus der weniger Probenvolumen benötigt. Gleichzeitig Test des Analysators. Vorteilhaft für MBU Messungen. Einweisung der Anwender, Ärzte und Hebammen, notwendig. Meldung an Vertrieb erfolgt. Analysator funktionsbereit. | SWO-56749 | No | a rental device, documented in extra SWO. It is an ABL90 Flex Plus that requires less sample volume. At the same time, test of the analyzer. Advantageous for MBU measurements. Instruction of users, doctors and midwives. message to sales. Analyzer ready for use. | a rental device, documented in extra SWO. It is an ABL90 Flex Plus that requires less sample volume. At the same time, test of the analyzer. Advantageous for MBU measurements. Instruction of users, doctors and midwives. message to sales. Analyzer ready for use. | User Error | pre analytical error | flex plus installed
ABL90 | Performed linearity to evaluate analyzer. | No correction necessary. Analyzer running according to specs. Suggested to customer to look at qc handling protocol. | CFA326AE0E4272DC | No | Performed linearity to evaluate analyzer. | No correction necessary. Analyzer running according to specs. Suggested to customer to look at qc handling protocol. | No fault found | pO2
ABL90 FLEX | Gerät geprüft. Sensorcassette erneuert Diverse Kal und QC durchgeführt. - Neue Sensorcassette bringt Sensor Ansprachefehler mit pCO2. - Weitere Betriebsmittel zum wechseln nicht vorhanden. | Kundin geraten den LP90 zu tauschen und zu reklamieren, Kundin muss sich erst einenLP90 bestellen und wird soweit auf ein anderes Gerät ausweichen, sollte sie weitere Probleme mit dem ABL haben wird sie sich wieder melden. | SWO-56897 | No | device checked. Sensorcassette renewed Diverse Kal and QC performed. - New sensor cassette brings sensor response error with pCO2. - No other equipment to change. | Customer advised to exchange and complain about the LP90, customer must first order an LP90 and will switch to another device if she has further problems with the ABL she will report again.
ABL90 FLEX | Gerät überprüft, System gespült, Sensorkassette erneuert, |  | SWO-56735 | No | device checked, system flushed, sensor cartridge renewed, |  | Assisted Replacement | SC | pH, Glu; Lac
ABL90 FLEX | Fehlersuche, stabile Kal. abgewartet, Kal. + QC i.O. | Fehlersuche; stabile Kal. abgewartet; Kal. + QC i.O. | SWO-56727 | No | Troubleshooting; stable Kal. waited; Kal. + QC in the first place | Troubleshooting; stable Kal. waited; Kal. + QC in the first place | No fault found | Wait for SC startup
ABL90 |  |  | CFA2FD5C0E46F806 | No |  |  | Defect | HW | SC PCB
ABL90 FLEX | Die Auslieferung abgebrochen und an RDEU Service gemeldet. | Die Auslieferung abgebrochen und an RDEU Service gemeldet. | SWO-56658 | Yes | Delivery aborted and reported to RDEU Service. | Delivery aborted and reported to RDEU Service. | Re classified RMED | installation og analyzer | 6 upSWO on 3 analyzer
ABL90 FLEX | Die Auslieferung abgebrochen und an RDEU Service gemeldet. | Die Auslieferung abgebrochen und an RDEU Service gemeldet. | SWO-56664 | Yes | Delivery aborted and reported to RDEU Service. | Delivery aborted and reported to RDEU Service. | Re classified RMED | installation og analyzer | 6 upSWO on 3 analyzer
ABL835 FLEX | Jeg har skiftet blæser i QC modulet. Alt er OK | Jeg har skiftet blæser i QC modulet. Alt er OK | SWO-56693 | No | I changed the fan in the QC module. Everything is OK | I changed the fan in the QC module. Everything is OK | Defect | HW | AutoCheck Module | FAN
ABL835 FLEX | Jeg har skiftet FLEX-Q modulet. Jeg har justeret og testet modulet. Der er kørt 2 punkts kal. Der er kørt QCer. Der er kørt patient prøve. Alt er OK. | Jeg har skiftet FLEX-Q modulet. Jeg har justeret og testet modulet. Der er kørt 2 punkts kal. Der er kørt QCer. Der er kørt patient prøve. Alt er OK. | SWO-56692 | No | I changed the FLEX-Q module. I have adjusted and tested the module. Two-point Kal has been run. QCer has been run. A patient sample has been run. Everything is OK. | I changed the FLEX-Q module. I have adjusted and tested the module. Two-point Kal has been run. QCer has been run. A patient sample has been run. Everything is OK. | Defect | HW | FlexQ Module | Positioning
ABL837 FLEX | QC-Werte überprüft, keine Abweichung feststellbar. Zielwerte an Labor übermittelt. Temperatur angepasst. In Labor kommen abweichende Crea-Werte an. Gesendete Daten per ASTMNet von Gerät abgegriffen. Kunde verwendet für Auswertung an Conworx nicht-Temperaturkorrigierte Crea-Werte. Herrn Braun informiert, in Conworx temperaturkorrigierte Werte zu verwenden. | QC-Werte überprüft, keine Abweichung feststellbar. Zielwerte an Labor übermittelt. Temperatur angepasst. In Labor kommen abweichende Crea-Werte an. Gesendete Daten per ASTMNet von Gerät abgegriffen. Kunde verwendet für Auswertung an Conworx nicht-Temperaturkorrigierte Crea-Werte. Herrn Braun informiert, in Conworx temperaturkorrigierte Werte zu verwenden. | SWO-56701 | No | QC values checked, no deviation detectable. target values to laboratory. Temperature adjusted. Different Crea values arrive in the laboratory. Data sent via ASTMNet intercepted by device. Customer uses non-temperature-corrected Crea values for evaluation of Conworx. Mr Braun is informed to use temperature-corrected values in Conworx. | QC values checked, no deviation detectable. target values to laboratory. Temperature adjusted. Different Crea values arrive in the laboratory. Data sent via ASTMNet intercepted by device. Customer uses non-temperature-corrected Crea values for evaluation of Conworx. Mr Braun is informed to use temperature-corrected values in Conworx. | Setup | HIS/LIS | Conworx | Temperature setup
AQT90 FLEX | Optische Einheit gewechselt alle alten Kalibrierungen gelöscht und die 3 Parameter neu kalibriert 5 Blank Messungen + QC Parameter durchgeführt | Optische Einheit gewechselt alle alten Kalibrierungen gelöscht und die 3 Parameter neu kalibriert 5 Blank Messungen + QC Parameter durchgeführt | SWO-56745 | No | Optical unit changed all old calibrations deleted and the 3 parameters recalibrated 5 blank measurements + QC parameters performed | Optical unit changed all old calibrations deleted and the 3 parameters recalibrated 5 blank measurements + QC parameters performed | Defect | HW | Process Module | Optical Unit | 135
TCM400-4 |  |  | CFA30EC60E610183 | No
ABL80 FLEX BASIC | 更换新测试卡后 | 更换新测试卡后 | SWO-57699 | No | After replacing a new test card | After replacing a new test card
ABL80 FLEX BASIC | 更换测试卡修复机器 | 更换测试卡修复机器 | SWO-56809 | No | Replace the test card repair machine | Replace the test card repair machine
ABL800 BASIC |  |  | SWO-56812 | No |  |  | Defect | HW | El/Met Module | PCB | Leak current
ABL80 OSM | Verified the problem. Checked the internal tubings and found the drain tube kinked, Problem happened again while running QC. | Put in an annual PM kit. Replaced the pump motor by the hemolyzer. | CFA6998D7043382C | No | Verified the problem. Checked the internal tubings and found the drain tube kinked, Problem happened again while running QC. | Put in an annual PM kit. Replaced the pump motor by the hemolyzer.
ABL90 FLEX | Fehlersuche, entgegen der telef. Auskunft wurde bisher nur der Lösungspack getauscht, SC getauscht | Fehlersuche, entgegen der telef. Auskunft wurde bisher nur der Lösungspack getauscht, SC getauscht | SWO-56983 | No | Troubleshooting, contrary to the telef. Information has so far only been exchanged for the solution pack, SC has been exchanged | Troubleshooting, contrary to the telef. Information has so far only been exchanged for the solution pack, SC has been exchanged | Assisted Replacement | SC | pO2
ABL90 FLEX | ABL überprüft. Lösungspack-PCB erneuert. Gerätestartphase inkl. QC's durchgeführt. Das LP wurde einwanfrei beim Neustart erkannt. | ABL überprüft. Lösungspack-PCB erneuert. Gerätestartphase inkl. QC's durchgeführt. Das LP wurde einwanfrei beim Neustart erkannt. | SWO-56944 | No | ABL. Solution pack PCB. device start phase incl. QC's. The LP was detected one wan-free on reboot. | ABL. Solution pack PCB. device start phase incl. QC's. The LP was detected one wan-free on reboot. | Defect | HW | SP PCB
ABL90 FLEX | Neuen Lösungspack angeliefert und installiert. | Kundin geraten den LP90 zu tauschen und zu reklamieren, Kundin muss sich erst einenLP90 bestellen und wird soweit auf ein anderes Gerät ausweichen, sollte sie weitere Probleme mit dem ABL haben wird sie sich wieder melden. | SWO-57121 | No | New solution pack delivered and installed. | Customer advised to exchange and complain about the LP90, customer must first order an LP90 and will switch to another device if she has further problems with the ABL she will report again. | Assisted Replacement | SP | pCO2
ABL90 FLEX | Sensor ersetzt | Sensor ersetzt | SWO-56978 | No | Sensor replaced | Sensor replaced | Assisted Replacement | SC | Lac
ABL90 | Battery not changing charging state. | Warranty replace Battery Pack. | CFA6EBB370569DAD | No | Battery not changing charging state. | Warranty replace Battery Pack. | Defect | HW | Battery module
ABL90 | Pack has flashing LED, will not charge. | Warranty replacement of Battery Pack. | CFA6ED2F700B9DB1 | No | Pack has flashing LED, will not charge. | Warranty replacement of Battery Pack. | Defect | HW | Battery module
ABL90 FLEX | 更换测试卡修复机器 | 更换测试卡修复机器 | SWO-57038 | No | Replace the test card repair machine | Replace the test card repair machine | Assisted Replacement | SC
ABL90 FLEX | 更换910-342修复机器 | 更换910-342修复机器 | SWO-56598 | No | Replace the 910-342 repair machine | Replace the 910-342 repair machine | Defect | HW | CPU
ABL90 FLEX | 更换试剂包正常。 | 更换试剂包正常。 | SWO-58442 | No | The replacement kit is OK. | The replacement kit is OK. | Assisted Replacement | SP | Lac QC high
AQT90 FLEX | ストレージホイールユニットの傾きによる為ストレージホイール引き上げ修理 | ストレージホイールユニットの傾きによる為ストレージホイール引き上げ修理 | SWO-57230 | No | Storage wheel pull-up repair due to tilt of storage wheel unit | Storage wheel pull-up repair due to tilt of storage wheel unit | Unknown | HW | Storage Module | Storage wheel | 5784
ABL825 FLEX | OXIméter panel csere, OXI tesztek,Protein eltávolítás, LS beállítás, Kalibrálások (elektródák, tHb), QC mérés hitelesítés ****** Az legutóbbi alkalommal kicserél Neon lámpa visszakerült a műszerbe Garanciális javítás ****** | OXIméter panel csere, OXI tesztek,Protein eltávolítás, LS beállítás, Kalibrálások (elektródák, tHb), QC mérés hitelesítés ****** Az legutóbbi alkalommal kicserél Neon lámpa visszakerült a műszerbe Garanciális javítás ****** | SWO-56977 | No | Oximeter panel Replacement, OXY tests, Protein removal, LS adjustment, calibrations (electrodes, tHb), QC measurement Authentication * * * * * * The last time you replace Neon light returned to the instrument warranty repair * * * * * * | Oximeter panel Replacement, OXY tests, Protein removal, LS adjustment, calibrations (electrodes, tHb), QC measurement Authentication * * * * * * The last time you replace Neon light returned to the instrument warranty repair * * * * * * | Defect | HW | Oxi module | PCB
ABL800 FLEX | FPC Cable ern., Gerät kalibriert, QC´s durchgef. - Mechanik wieder OK | FPC Cable ern., Gerät kalibriert, QC´s durchgef. - Mechanik wieder OK | SWO-56969 | No | FPC Cable ern., Device calibrated, QC's through. - Mechanics ok again | FPC Cable ern., Device calibrated, QC's through. - Mechanics ok again | Defect | HW | AutoCheck Module | FPC Cable
ABL835 FLEX | PCO 2 Membraniert, Messkammer gereinigt. Mehrfach Kalibriert. | PCO 2 Membraniert, Messkammer gereinigt. Mehrfach Kalibriert. | SWO-57596 | No | PCO 2 Membraneed, measuring chamber cleaned. Multiple calibrated. | PCO 2 Membraneed, measuring chamber cleaned. Multiple calibrated. | Assisted Replacement | Electrode | Membrane | pCO2
ABL825 FLEX | Wiedereinbau folgender Module nach Überprüfung ...BG..Elyt...Fluidic Modul nach Justagen und Kalibrationen QC ok Zur weiteren Beobachtung. Dummy Elektrode abgedichtet. | Wiedereinbau folgender Module nach Überprüfung ...BG..Elyt...Fluidic Modul nach Justagen und Kalibrationen QC ok Zur weiteren Beobachtung. Dummy Elektrode abgedichtet. | SWO-57245 | No | Re-installation of the following modules after verification ... Bg.. Elyt... Fluidic module after adjustments and calibrations QC ok For further observation. Dummy electrode sealed. | Re-installation of the following modules after verification ... Bg.. Elyt... Fluidic module after adjustments and calibrations QC ok For further observation. Dummy electrode sealed. | Assisted Replacement | Electrode | Electrode | Dummy | Sealing ring
ABL835 FLEX | Komplett monitor csere,Tápegység visszacserélése az eredetire, Képernyő kalibrálás, OXI tesztek, Kalibrálások (elektródák, tHB), QC mérés hitelesítés | Monitor kábel újracsatlakoztatása, Képernyő kalibrálás, QC mérés | SWO-56914 | No | Complete monitor replacement, replace power supply to original, screen calibration, OXY tests, calibrations (electrodes, tHB), QC measurement authentication | Reconnecting Monitor cable, screen calibration, QC measurement | Defect | HW | DMS_PC | Touch screen
ABL825 FLEX | Nach Testlauf in GS Rücktausch von Kundenmodul Beim Eintreffen QC Level 4 mit Fehler bei pH, Ca und Na => Überprüfung Elektroden und Neumembranierung Ref Mehrmals Kal bis i.O. Fehlerhafte QC nachgeholt | Nach Testlauf in GS Rücktausch von Kundenmodul Beim Eintreffen QC Level 4 mit Fehler bei pH, Ca und Na => Überprüfung Elektroden und Neumembranierung Ref Mehrmals Kal bis i.O. Fehlerhafte QC nachgeholt | SWO-57178 | No | After test run in GS Return of customer module When arriving QC Level 4 with error at pH, Ca and Na => Inspection electrodes and re-membraneing Ref Multiple times Kal to i.O. Faulty QC | After test run in GS Return of customer module When arriving QC Level 4 with error at pH, Ca and Na => Inspection electrodes and re-membraneing Ref Multiple times Kal to i.O. Faulty QC | Assisted Replacement | Electrode | Membrane | Reference
ABL825 FLEX | Zebra Printer an neue Gegebenheiten (+Glukose und Calcium angepasst), Etikettenformat 55x95mm, Ausdruck zusammen mit Herrn Frank STL überprüft ok System-und Funktionskontrolle i.O. | Zebra Printer an neue Gegebenheiten (+Glukose und Calcium angepasst), Etikettenformat 55x95mm, Ausdruck zusammen mit Herrn Frank STL überprüft ok System-und Funktionskontrolle i.O. | SWO-56975 | Yes | Zebra Printer adapted to new conditions (+glucose and calcium), label format 55x95mm, printout together with Mr. Frank STL checks ok system and function control i.O. | Zebra Printer adapted to new conditions (+glucose and calcium), label format 55x95mm, printout together with Mr. Frank STL checks ok system and function control i.O. | Re classified RMED | Ext printer
ABL825 FLEX | Geräteneustart, Funktionskontrolle | Geräteneustart, Funktionskontrolle | SWO-57004 | No | Geräteneustart, Funktionskontrolle | Geräteneustart, Funktionskontrolle | Unknown | HW | AutoCheck Module | Barcode reader | Restart
ABL827 FLEX | Verifed the problem. Checked the voltages on the power supply and they werte all good. | Replaced the cpu unit. | CFA71F0770315067 | No | Verifed the problem. Checked the voltages on the power supply and they werte all good. | Replaced the cpu unit. | Defect | HW | DMS_PC | CPU
AQT90 FLEX | run the safe position and checked the punch up service program, now working satisfactorily | run the safe position and checked the punch up service program, now working satisfactorily | SWO-57906 | No | run the safe position and checked the punch up service program, now working satisfactorily | run the safe position and checked the punch up service program, now working satisfactorily | Other | SW | Incubator Module | Punchup
ABL80 FLEX BASIC | Gerät überprüft, Sensorkassette erneuert, Öösungspack ausgetauscht, Kal o.k. | Gerät überprüft, Sensorkassette erneuert, Öösungspack ausgetauscht, Kal o.k. | SWO-57240 | No | Device checked, sensor cartridge renewed, oven pack replaced, Kal ok. | Device checked, sensor cartridge renewed, oven pack replaced, Kal ok.
ABL90 |  |  | CFA827E8E65C299E | No |  |  | Assisted Replacement | SP | pO2
ABL90 | Error 1343- unable to pump solution | Flushed unit. Installed 3.1 MR6 software | CFA7FE1FE6101EE8 | No | Error 1343- unable to pump solution | Flushed unit. Installed 3.1 MR6 software | Assisted Replacement | SC | Wetability
ABL90 | 5xx 6xx high level encryption became unselected on ABL. | Enabled 5xx 6xx encryption. | CFA7FD91E6054E5D | No | 5xx 6xx high level encryption became unselected on ABL. | Enabled 5xx 6xx encryption. | Setup | HIS/LIS | Aqure
ABL90 FLEX | Fehler deuten auf Fehler in Betriebsmittel hin. Überprüfung des Sensor boards über Sensor Simulator Kassette. Ergebniss o.K., keine Beanstandung der Verstärkerkanäle. Einlasssonde krum, erneuert und eine weitere Vorort gelassen, da keine am Analysator. Einlassdichtung verschmutzt, erneuert. Fließwege gespült, keine Beanstandung. Analysator messbereit. | Fehler deuten auf Fehler in Betriebsmittel hin. Überprüfung des Sensor boards über Sensor Simulator Kassette. Ergebniss o.K., keine Beanstandung der Verstärkerkanäle. Einlasssonde krum, erneuert und eine weitere Vorort gelassen, da keine am Analysator. Einlassdichtung verschmutzt, erneuert. Fließwege gespült, keine Beanstandung. Analysator messbereit. | SWO-57223 | No | Errors indicate errors in equipment. Checking the sensor board via Sensor Simulator cassette. Result o.K., no complaint of the amplifier channels. Entrance probe crooked, renewed and left another suburb, since none on the analyzer. Inlet seal dirty, renewed. Flow paths rinsed, no complaint. Analyzer ready to measure. | Errors indicate errors in equipment. Checking the sensor board via Sensor Simulator cassette. Result o.K., no complaint of the amplifier channels. Entrance probe crooked, renewed and left another suburb, since none on the analyzer. Inlet seal dirty, renewed. Flow paths rinsed, no complaint. Analyzer ready to measure. | Assisted Replacement | SC | Reference
ABL90 FLEX | Tausch der CF Karte, vorab Sichern der Daten. Einspielung des Betriebssystems und der Aplikationssoftware. Einspielung der alten Datenbank und Konfiguration. Neuer LP und SC sind notwendig durch CF Karten Tausch, vom Anwender genommen. | Tausch der CF Karte, vorab Sichern der Daten. Einspielung des Betriebssystems und der Aplikationssoftware. Einspielung der alten Datenbank und Konfiguration. Neuer LP und SC sind notwendig durch CF Karten Tausch, vom Anwender genommen. | SWO-57224 | No | Exchange of the CF card, back up the data in advance. Recording of the operating system and the application software. Recording of the old database and configuration. New LP and SC are necessary by CF card swap, taken by the user. | Exchange of the CF card, back up the data in advance. Recording of the operating system and the application software. Recording of the old database and configuration. New LP and SC are necessary by CF card swap, taken by the user. | Defect | HW | CF Card | Double printing
ABL90 | EXTERNAL QC WILL NOT PASS | PUT EXTERNAL QC IN ROOM THAT IS 25C. | CFA80E1AE62D864D | Yes | EXTERNAL QC WILL NOT PASS | PUT EXTERNAL QC IN ROOM THAT IS 25C. | Re classified RMED | Double entry
ABL90 | Limpeza do circuito, ajuste dos sensores de liquidos, substituição do SP, teste com água e 2 testes com sangue, tudo OK, passagem do QC. |  | CFA83D45E6330D31 | No | Cleaning the circuit, adjusting the liquid sensors, replacing the SP, testing with water and 2 tests with blood, all OK, pass the QC. |  | Adjustment | LS
ABL90 | Verificação dos registros, limpeza do circuito, substituição da sonda e da entrada, ajustes dos SLs, calibração das bombas, execução do teste de pressão, estabilização e passagem dos QCs. |  | CFA74FABE6396D99 | No | Checking records, cleaning the circuit, replacing the probe and inlet, adjusting the SLs, calibrating the pumps, running the pressure test, stabilizing and passing the QCs. |  | Cleaned | Fluidic pathway
ABL825 FLEX | Verifed the problem by checking the data logs. Passed a complete oxi test. Checked the lamp calibration. After performing the service THB cal the unit would not pass the THB during cals. Performed more complete oxi test and found that the it was failing the neon position. Concluded that the spectrometer is not reading the neon lamp correctly. Ordering a spectrometer and a fiber optic cable. | Performed a complete oxi test. Calibrated the lamp. Replaced the CA membrane. Replaced the lamp assembly. Replaced the spectrometer | CFA7EE24E63A4E30 | No | Verifed the problem by checking the data logs. Passed a complete oxi test. Checked the lamp calibration. After performing the service THB cal the unit would not pass the THB during cals. Performed more complete oxi test and found that the it was failing the neon position. Concluded that the spectrometer is not reading the neon lamp correctly. Ordering a spectrometer and a fiber optic cable. | Performed a complete oxi test. Calibrated the lamp. Replaced the CA membrane. Replaced the lamp assembly. Replaced the spectrometer | Defect | HW | Oxi module | Spectrometer
ABL835 FLEX | Ca. Elektrode neu membraniert, mehrfach kalibriert, Funktionskontrolle ok | Ca. Elektrode neu membraniert, mehrfach kalibriert, Funktionskontrolle ok | SWO-57274 | No | Approx. electrode re-membraneed, multiple calibrated, function control ok | Approx. electrode re-membraneed, multiple calibrated, function control ok | Assisted Replacement | Electrode | Membrane | Ca
ABL825 FLEX | Lites failing cal every night at 4am then after recal is done works again | Replaced REF electrode under warranty | CFA7DD6DE6046B93 | No | Lites failing cal every night at 4am then after recal is done works again | Replaced REF electrode under warranty | Assisted Replacement | Electrode | Electrode | Reference
ABL825 FLEX |  |  | CFA7ED72E65613CD | No |  |  | Defect | HW | Oxi module | Lamp unit
AQT90 FLEX | Notwendige Datei ASTM.ini für Convorx ist bereits eingespielt. IT Einbindung o.K., kein Fehler feststellbar. Übertragungsproblem liegt an der Einstellung im Poct Cellerator von Convorx. Erst wenn QC freigegeben wird werden Daten zu Convorx und dann ins KIS übertragen. Die beiden weiteren angeschlossenen Analysatoren müssen somit nicht angefahren werden. Analysator messbereit. | Notwendige Datei ASTM.ini für Convorx ist bereits eingespielt. IT Einbindung o.K., kein Fehler feststellbar. Übertragungsproblem liegt an der Einstellung im Poct Cellerator von Convorx. Erst wenn QC freigegeben wird werden Daten zu Convorx und dann ins KIS übertragen. Die beiden weiteren angeschlossenen Analysatoren müssen somit nicht angefahren werden. Analysator messbereit. | SWO-57225 | No | Necessary file ASTM.ini for Convorx is already saved. IT integration o.K., no error detectable. Transmission problem is due to the setting in The Poct Cellerator of Convorx. Only when QC is released data about Convorx and then transferred to the KIS. The two other connected analyzers do not have to be approached. Analyzer ready to measure. | Necessary file ASTM.ini for Convorx is already saved. IT integration o.K., no error detectable. Transmission problem is due to the setting in The Poct Cellerator of Convorx. Only when QC is released data about Convorx and then transferred to the KIS. The two other connected analyzers do not have to be approached. Analyzer ready to measure. | Setup | SW | Analyzer control Module | HIS/LIS
AQT90 FLEX |  |  | CFB1112B2061DB38 | No |  |  | Defect | HW | Analyzer control Module | CPU
ABL80 FLEX BASIC | 更换测试卡修复机器 | 更换测试卡修复机器 | SWO-57322 | No | Replace the test card repair machine | Replace the test card repair machine
ABL80 FLEX BASIC | 更换测试卡修复机器 | 更换测试卡修复机器 | SWO-57324 | No | Replace the test card repair machine | Replace the test card repair machine
ABL80 FLEX BASIC | 更换两张测试卡正常。 | 更换两张测试卡正常。 | SWO-57512 | No | Replacing the two test cards is ok. | Replacing the two test cards is ok.
ABL80 CO-OX | Seems to be happening ramdomly. | Customer reploaced the solution pack and it hasn't happened since. I replaced the SP latch. | CFA9652F5C1B042E | No | Seems to be happening ramdomly. | Customer reploaced the solution pack and it hasn't happened since. I replaced the SP latch.
ABL90 |  |  | CFA91BE85C4ED4B4 | No |  |  | Defect | HW | Sensor Bay Unit
ABL90 | Roma.ini to be modified. | Modified ini. | CFB453FB820E3F55 | No | Roma.ini to be modified. | Modified this. | Setup | Roma.ini | Barcode Reader
ABL90 | Codebar to be added to .ini file. | Modified Roma.ini. | CFB45119822D3F52 | No | Codebar to be added to .ini file. | Modified Roma.ini. | Setup | Roma.ini | Barcode Reader
ABL90 | Customer using Codebar bar codes. | Modified Ini. to accept Codebar. | CFB44F46824C3F4E | No | Customer using Codebar bar codes. | Modified Ini. to accept Codebar. | Setup | Roma.ini | Barcode Reader
ABL90 | Roma.ini file to be modified. | Modified .ini to accept bar codes. | CFB4502982223F50 | No | Roma.ini file to be modified. | Modified .ini to accept bar codes. | Setup | Roma.ini | Barcode Reader
ABL90 | EXTERNAL QC LEVEL 1 IS HIGH AND LEVEL 3 IS LOW | FRANK CASTANEDO REPLACED SENSOR CASSETTE | CFAA3AABD236516E | No | EXTERNAL QC LEVEL 1 IS HIGH AND LEVEL 3 IS LOW | FRANK CASTANEDO REPLACED SENSOR CASSETTE | Revisit required | Techline
ABL90 | EXTERNAL QC LEVEL 1 IS HIGH AND LEVEL 3 IS LOW | FRANK CASTANEDO REPLACED SENSOR CASSETTE | CFAA3BB0D21A5170 | No | EXTERNAL QC LEVEL 1 IS HIGH AND LEVEL 3 IS LOW | FRANK CASTANEDO REPLACED SENSOR CASSETTE | Revisit required | Techline
ABL90 | EXTERNAL QC LEVEL 1 IS HIGH AND LEVEL 3 IS LOW | FRANK CASTANEDO REPLACED SENSOR CASSETTE | CFAA3CBED2305172 | No | EXTERNAL QC LEVEL 1 IS HIGH AND LEVEL 3 IS LOW | FRANK CASTANEDO REPLACED SENSOR CASSETTE | Revisit required | Techline
ABL90 | EXTERNAL QC LEVEL 1 IS HIGH AND LEVEL 3 IS LOW | FRANK CASTANEDO REPLACED SENSOR CASSETTE | CFAA3D36D2555174 | No | EXTERNAL QC LEVEL 1 IS HIGH AND LEVEL 3 IS LOW | FRANK CASTANEDO REPLACED SENSOR CASSETTE | Revisit required | Techline
ABL90 FLEX | 更换910-285修复机器. | 更换910-285修复机器. | SWO-56880 | No | Replace the 910-285 repair machine. | Replace the 910-285 repair machine. | Defect | HW | PS
ABL90 FLEX | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-57809 | No | Repair after replacing a new test card | Repair after replacing a new test card | Assisted Replacement | SC | SC | SC | 17
AQT90 FLEX | 更换903-171修复机器 | 更换903-171修复机器 | SWO-57036 | No | Replace the 903-171 repair machine | Replace the 903-171 repair machine | Defect | HW | Aspiration Module | Bend needle | 4794
ABL825 FLEX | Überprüfung Kal und QC Ergebnisse Überprüfung Halogenlampe => Intensität zu hoch => Lampe nachjustiert Oximetrie Complete Test i.O. => Stdev Dark Value bei 1,5 mV Halogenlampenintensität bei ca. 3200 Durchführung Proteinentfernung Service-tHb-Kal i.O. Kal und tHb-Kal i.O. => Fcuv = 0,999 QC Level 1 bleibt fehlerhaft bei FMetHb (zu hoch) Geräteneustart inkl. Trennung von Netzkabel Kal1 und QC beide Level dann i.O. | Überprüfung Kal und QC Ergebnisse Überprüfung Halogenlampe => Intensität zu hoch => Lampe nachjustiert Oximetrie Complete Test i.O. => Stdev Dark Value bei 1,5 mV Halogenlampenintensität bei ca. 3200 Durchführung Proteinentfernung Service-tHb-Kal i.O. Kal und tHb-Kal i.O. => Fcuv = 0,999 QC Level 1 bleibt fehlerhaft bei FMetHb (zu hoch) Geräteneustart inkl. Trennung von Netzkabel Kal1 und QC beide Level dann i.O. | SWO-57331 | No | Review Kal and QC Results Check halogen lamp => Intensity too high => Lamp readjusted oximetry Complete Test i.O. => Stdev Dark Value at 1.5 mV halogen lamp intensity at approx. 3200 performing protein removal service-tHb-Kal i.O. Kal and tHb-Kal i.O. => Fcuv = 0.999 QC Level 1 remains faulty at FMetHb (too high) device restart incl. disconnection of power cable Kal1 and QC both levels then i.O. | Review Kal and QC Results Check halogen lamp => Intensity too high => Lamp readjusted oximetry Complete Test i.O. => Stdev Dark Value at 1.5 mV halogen lamp intensity at approx. 3200 performing protein removal service-tHb-Kal i.O. Kal and tHb-Kal i.O. => Fcuv = 0.999 QC Level 1 remains faulty at FMetHb (too high) device restart incl. disconnection of power cable Kal1 and QC both levels then i.O. | Adjustment | HW | Oxi module | Lamp unit
ABL835 FLEX | Qc is within insert limits and consistently better than previous checks. All 7 of customers analyzers are below the mean value for S7745 lot 619. | Installed another pH seal ring. Installed anran S7745 lot 630 QC. This QC came in closer to the mean value but was still below the mean value by 1.7mmHg. | CFA925945C624C87 | No | Qc is within insert limits and consistently better than previous checks. All 7 of customers analyzers are below the mean value for S7745 lot 619. | Installed another pH seal ring. Installed anran S7745 lot 630 QC. This QC came in closer to the mean value but was still below the mean value by 1.7mmHg. | Assisted Replacement | Electrode | Electrode | pH | Sealing ring
ABL820 FLEX | System gespült | System gespült | SWO-57424 | No | System flushed | System flushed | Cleaned | Fluidic pathway
ABL820 FLEX | Found a small leak under the inlet gasket. The leak is small enough that no fluid is leaking but residue is building up under the inlet gasket. | I changed the inlet gasket, inlet tubing nipples, pump tubes and tubing organizer. | CFA9805E5C20F5D5 | No | Found a small leak under the inlet gasket. The leak is small enough that no fluid is leaking but residue is building up under the inlet gasket. | I changed the inlet gasket, inlet tubing nipples, pump tubes and tubing organizer. | Defect | HW | Inlet module | Leaking | Nipple kit
ABL825 FLEX | Seems to be some left over fluid in pH BG measuring chamber in front of Reference electrode, pH electrode and PCO2 Electrode. | Cleaned measuring chamber with Q tip after removing all electrodes. Replaced pH sealing ring and Reference and PO2 Membrane. Adjusted white membrane on CL electrode. Fingers were slightly past the ring. Ran Calibrations and QC. Green QC was 95.5. Now reads 98.2. | CFA938E45C55D8CE | No | Seems to be some left over fluid in pH BG measuring chamber in front of Reference electrode, pH electrode and PCO2 Electrode. | Cleaned measuring chamber with Q tip after removing all electrodes. Replaced pH sealing ring and Reference and PO2 Membrane. Adjusted white membrane on CL electrode. Fingers were slightly past the ring. Ran Calibrations and QC. Green QC was 95.5. Now reads 98.2. | Assisted Replacement | Electrode | Electrode | pH | Sealing ring
AQT90 FLEX |  |  | CFB1F77C96632775 | No |  |  | Defect | HW | Cover and Chassis | Replenishment pack PCB
AQT90 FLEX | Optische Einheit getauscht, bei Wartungsmassnahme. Die Factory Referenz durchgeführt, gemessener Wert 3101204,3. Die zwei Parameter TnT und D-Dimer neu kalibriert. Die bestehenden Kalibrationen durch vorhergehnde Optischen Einheit gelöscht. Kalibrationen und QC o.K. Service Dump für TS und RMED gezogen. Analysator messbereit. | Optische Einheit getauscht, bei Wartungsmassnahme. Die Factory Referenz durchgeführt, gemessener Wert 3101204,3. Die zwei Parameter TnT und D-Dimer neu kalibriert. Die bestehenden Kalibrationen durch vorhergehnde Optischen Einheit gelöscht. Kalibrationen und QC o.K. Service Dump für TS und RMED gezogen. Analysator messbereit. | SWO-57429 | No | Optical unit exchanged, in case of maintenance measures. The factory reference performed, measured value 3101204,3. The two parameters TnT and D-Dimer are recalibrated. The existing calibrations are deleted by previous optical unit. Calibrations and QC o.K. Service Dump for TS and RMED. Analyzer ready to measure. | Optical unit exchanged, in case of maintenance measures. The factory reference performed, measured value 3101204,3. The two parameters TnT and D-Dimer are recalibrated. The existing calibrations are deleted by previous optical unit. Calibrations and QC o.K. Service Dump for TS and RMED. Analyzer ready to measure. | Performance | QC | TNT | Optical Unit
ABL80 FLEX BASIC | 清洁管路。 | 清洁管路。 | SWO-57888 | No | Clean the lines. | Clean the lines.
ABL80 CO-OX | Found display unit damaged. Will perform pm to fix fluid blockage. Sensor cassette that was shipped with instrument was damaged and needed to be repaired. Solution pack installed in instrument was voided and unable to be installed and used. | Replaced damaged display unit. Performed pm according to manufacturers specifications. Installed new sensor cassette and solution pack. Verified instrument operation. | CFABC4D44803DF3F | No | Found display unit damaged. Will perform pm to fix fluid blockage. Sensor cassette that was shipped with instrument was damaged and needed to be repaired. Solution pack installed in instrument was voided and unable to be installed and used. | Replaced damaged display unit. Performed pm according to manufacturers specifications. Installed new sensor cassette and solution pack. Verified instrument operation.
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-57553 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-57510 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL80 FLEX BASIC |  |  | SWO-57477 | No
ABL80 FLEX CO-OX | tHb Kal. durchgeführt. Brachte jedoch keinen Erfolg. Vergleichsmessungen auch zu ABL 835 durchgeführt. COHB und MetHB Korrekturen in Absprache mit Dr. Mayer vom Labor Enders und Partner eingeführt. Kunde wird selbst noch weitere Vergleichsmessungen durchführen. Weitere Absprache in der nächsten Woche. | tHb Kal. durchgeführt. Brachte jedoch keinen Erfolg. Vergleichsmessungen auch zu ABL 835 durchgeführt. COHB und MetHB Korrekturen in Absprache mit Dr. Mayer vom Labor Enders und Partner eingeführt. Kunde wird selbst noch weitere Vergleichsmessungen durchführen. Weitere Absprache in der nächsten Woche. | SWO-57676 | No | tHb Kal. Performed. But it didn't succeed. comparative measurements were also carried out for ABL 835. COHB and MetHB corrections in consultation with Dr. Mayer from the Enders and Partner laboratory. Customer will carry out further comparative measurements himself. Further agreement next week. | tHb Kal. Performed. But it didn't succeed. comparative measurements were also carried out for ABL 835. COHB and MetHB corrections in consultation with Dr. Mayer from the Enders and Partner laboratory. Customer will carry out further comparative measurements himself. Further agreement next week.
ABL80 FLEX | Replaced SC | Replaced SC | SWO-57971 | No | Replaced SC | Replaced SC
ABL90 FLEX | Durch Ersatzgerät (090R0592N001) ersetzt i.o Allgemeine Kontrolle i.o Sensorkassette ersetzt i.o System i.o Ka., + QC i.o |  | SWO-57562 | No | Replaced by replacement device (090R0592N001) i.o General control i.o sensor cartridge replaces i.o System i.o Ka., + QC i.o |  | Assisted Replacement | SC | Wetability
ABL90 FLEX | Lösungpack ersetzt | Lösungpack ersetzt | SWO-57624 | No | Solution pack replaced | Solution pack replaced | Assisted Replacement | SP | pCO2 | Level A
ABL90 FLEX PLUS |  |  | CFBAFB84D020F8A3 | No |  |  | Unknown | FT
ABL90 FLEX PLUS | Ventilschlauch erneuert und Einlasssonde, da krum. Vmtl. SC nicht o.K. Aus GS ein LP und eine SC mitgebracht. Analysator messbereit nach Software Update. | Ventilschlauch erneuert und Einlasssonde, da krum. Vmtl. SC nicht o.K. Aus GS ein LP und eine SC mitgebracht. Analysator messbereit nach Software Update. | SWO-57648 | No | Valve hose renewed and inlet probe, since krum. Vmtl. SC not o.K. Brought an LP and a SC from GS. Analyzer ready for software update. | Valve hose renewed and inlet probe, since krum. Vmtl. SC not o.K. Brought an LP and a SC from GS. Analyzer ready for software update. | Assisted Replacement | SC | Wetability
ABL90 FLEX PLUS | Onderstaande handelingen heb uitgevoerd: · Laat de analyser verder doorstarten, waardoor je in de instellingen kunt komen · Selecteer Menu – Extra functies – USB functies – Fabrieksinstellingen en in het volgende scherm “Doorgaan” · De analyser start nu op met de Radiometer instellingen | Onderstaande handelingen heb uitgevoerd: · Laat de analyser verder doorstarten, waardoor je in de instellingen kunt komen · Selecteer Menu – Extra functies – USB functies – Fabrieksinstellingen en in het volgende scherm “Doorgaan” · De analyser start nu op met de Radiometer instellingen | SWO-57635 | No | Perform the following actions: · Allow the analyser to be further started, allowing you to enter the settings · Select Menu – Additional Functions – USB functions – Factory settings and in the next screen "continue" · The analyser now starts with the Radio meter settings | Perform the following actions: · Allow the analyser to be further started, allowing you to enter the settings · Select Menu – Additional Functions – USB functions – Factory settings and in the next screen "continue" · The analyser now starts with the Radio meter settings | Defect | SW | Load setup?
ABL90 FLEX PLUS | Het volgende heb uitgevoerd, Druk NIET op de “OK” knop in het error venster · Laat de analyser verder doorstarten, waardoor je in de instellingen kunt komen · Selecteer Menu – Extra functies – USB functies – Fabrieksinstellingen en in het volgende scherm “Doorgaan” · De analyser start nu op met de Radiometer instellingen | Het volgende heb uitgevoerd; Druk NIET op de “OK” knop in het error venster · Laat de analyser verder doorstarten, waardoor je in de instellingen kunt komen · Selecteer Menu – Extra functies – USB functies – Fabrieksinstellingen en in het volgende scherm “Doorgaan” · De analyser start nu op met de Radiometer instellingen | SWO-57636 | No | Have carried out the following; Do NOT Press the "OK" button in the error window · Allow the analyser to be further started, allowing you to enter the settings · Select Menu – Additional Functions – USB functions – Factory settings and in the next screen "continue" · The analyser now starts with the Radio meter settings | Have carried out the following; Do NOT Press the "OK" button in the error window · Allow the analyser to be further started, allowing you to enter the settings · Select Menu – Additional Functions – USB functions – Factory settings and in the next screen "continue" · The analyser now starts with the Radio meter settings | Defect | SW | Load setup?
ABL90 FLEX | 更换910-285一个后修复 | 更换910-285一个后修复 | SWO-57103 | No | Replace 910-285 one post-fix | Replace 910-285 one post-fix | Defect | HW | PS
ABL90 FLEX | 更换测试卡修复机器 | 更换测试卡修复机器 | SWO-57557 | No | Replace the test card repair machine | Replace the test card repair machine | Assisted Replacement | SC | Reference
ABL90 FLEX | 更换试剂包正常。 | 更换试剂包正常。 | SWO-57534 | No | The replacement kit is OK. | The replacement kit is OK. | Assisted Replacement | SP | FT
ABL90 | Limpeza do circuito,substituição do SP sem sucesso, substituição do SC, establização e passagem do QC. Tentativa de ligação à rede sem sucesso, não adquire IP. |  | CFAAAACCD2128D12 | No | Cleaning the circuit, replacing the SP without success, replacing the SC, establishing and passing the QC. Attempt to connect to the network without success, does not acquire IP. |  | Assisted Replacement | SC | pO2
AQT90 FLEX |  |  | SWO-57466 | No |  |  | Defect | HW | Process Module | Optical Unit | 124
AQT90 FLEX | 更换903-198一个后未修复，CPU板和存储板也已损坏，需要再借件 | 更换903-198一个后未修复，CPU板和存储板也已损坏，需要再借件 | SWO-57566 | No | Replacing 903-198 one has not been repaired, the CPU board and storage board have also been damaged and need to be borrowed again | Replacing 903-198 one has not been repaired, the CPU board and storage board have also been damaged and need to be borrowed again | Defect | HW | Analyzer control Module | Main Board, PCB | 2
AQT90 FLEX |  |  | SWO-57511 | No |  |  | Defect | HW | Aspiration Module | Bend needle | 4794
AQT90 FLEX | 更换903-364一个正常。 | 更换903-364一个正常。 | SWO-57037 | No | Replace 903-364 one normal. | Replace 903-364 one normal. | Defect | HW | Process Module | Optical Unit | 131
ABL835 FLEX | Neue pCO2 Elektrode eingebaut. Nach mehrfachen Kalibrieren ist der AQC wert wieder OK. | Neue pCO2 Elektrode eingebaut. Nach mehrfachen Kalibrieren ist der AQC wert wieder OK. | SWO-58033 | No | New pCO2 electrode installed. After multiple calibrations, the AQC is worth OK again. | New pCO2 electrode installed. After multiple calibrations, the AQC is worth OK again. | Assisted Replacement | Electrode | Electrode | pCO2
ABL835 FLEX | Monitoreinheit geprüft, Kabelbruch im Monitorkabel, Ersatzteil bestellt, |  | SWO-57608 | No | Monitor unit checked, cable breakage in the monitor cable, spare part ordered, |  | Defect | HW | DMS_PC | Cable
ABL800 BASIC | pCO2 Elektrode erneuert. Messkammer gereinigt. Kalibrationen und QC's vorgenommen. | pCO2 Elektrode erneuert. Messkammer gereinigt. Kalibrationen und QC's vorgenommen. | SWO-57662 | No | pCO2 electrode. measuring chamber. calibrations and QC's. | pCO2 electrode. measuring chamber. calibrations and QC's. | Assisted Replacement | Electrode | Electrode | pCO2
ABL800 BASIC | Inlet Probe und Einlaßdichtung erneuert. Drei Leistungstests mit tHb Kalibrator durchgeführt - ok. Kalibration und QC's durchgeführt. | Inlet Probe und Einlaßdichtung erneuert. Drei Leistungstests mit tHb Kalibrator durchgeführt - ok. Kalibration und QC's durchgeführt. | SWO-57666 | No | Inlet sample and inlet seal renewed. Three performance tests performed with tHb calibrator - ok. calibration and QC's. | Inlet sample and inlet seal renewed. Three performance tests performed with tHb calibrator - ok. calibration and QC's. | Defect | HW | Inlet module | Probe bend
ABL800 BASIC | Proteinentfernung vorgenommen. Halogen Lamp justiert. Service tHb Kalibration vorgenommen. Glu, Ref und pO2 Elektroden neu membraniert. tHb Kalibration, Kalibrationen und QC's vorgenommen. | Proteinentfernung vorgenommen. Halogen Lamp justiert. Service tHb Kalibration vorgenommen. Glu, Ref und pO2 Elektroden neu membraniert. tHb Kalibration, Kalibrationen und QC's vorgenommen. | SWO-57660 | No | protein removal. Halogen Lamp. service tHb calibration. Glu, Ref and pO2 electrodes. tHb calibration, calibrations and QC's. | protein removal. Halogen Lamp. service tHb calibration. Glu, Ref and pO2 electrodes. tHb calibration, calibrations and QC's. | Assisted Replacement | Electrode | Membrane | pO2
AQT90 FLEX | Se sustituyo needle assembly y se calibro según protocolo del fabricante. Se comprobó correcto funcionamiento mediante QC. Se realizo el mantenimiento preventivo y se comprobó correcto funcionamiento mediante QC: |  | CFC138951E1C9564 | Yes | Needle assembly is replaced and calibro according to manufacturer's protocol. Proper operation was checked by QC. Preventive maintenance was performed and proper operation was found by QC: |  | Re classified RMED | PM
AQT90 FLEX | Barcode Reader (902-924R0060N008 durch 902-924R1007N124 ersetzt)ausgetauscht | Barcode Reader (902-924R0060N008 durch 902-924R1007N124 ersetzt)ausgetauscht | SWO-57645 | No | Barcode Reader (902-924R0060N008 replaced by 902-924R1007N124) | Barcode Reader (902-924R0060N008 replaced by 902-924R1007N124) | Defect | HW | Analyzer control Module | Barcode reader, display
AQT90 FLEX | Kassettenführung justiert, Gerät neu gestartet, QC´s durchgef. | Kassettenführung justiert, Gerät neu gestartet, QC´s durchgef. | SWO-57565 | No | Cassette guide adjusted, device restarted, QC's through. | Cassette guide adjusted, device restarted, QC's through. | Adjustment | HW | Cover and Chassis | Door guide | 5784
ABL80 FLEX BASIC | 更换新试剂包后修复 | 更换新试剂包后修复 | SWO-57723 | No | Repair after replacing a new kit | Repair after replacing a new kit
ABL80 FLEX BASIC | 更换测试卡正常。 | 更换测试卡正常。 | SWO-57889 | No | The replacement test card is OK. | The replacement test card is OK.
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-57811 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL80 FLEX BASIC |  |  | SWO-57700 | No
ABL80 FLEX BASIC |  |  | SWO-57746 | No
ABL80 FLEX BASIC | 更换测试卡修复机器 | 更换测试卡修复机器 | SWO-57779 | No | Replace the test card repair machine | Replace the test card repair machine
ABL80 FLEX BASIC | 更换测试卡正常。 | 更换测试卡正常。 | SWO-58441 | No | The replacement test card is OK. | The replacement test card is OK.
ABL90 FLEX | Fehlersuche, Reinigung Analysator, Installation neue Sensorkassette, Funktionskontrolle | Fehlersuche, Reinigung Analysator, Installation neue Sensorkassette, Funktionskontrolle | SWO-57867 | No | Troubleshooting, cleaning analyzer, installation of new sensor cartridge, function control | Troubleshooting, cleaning analyzer, installation of new sensor cartridge, function control | Assisted Replacement | SC | Reference
ABL90 FLEX | Auf Grund von langen Lieferzeiten Abholung eines Lösungspacks in Praxis Dr. Morgner Freiberg ( Leihweise), Lieferung in KH Freiberg Labor | Auf Grund von langen Lieferzeiten Abholung eines Lösungspacks in Praxis Dr. Morgner Freiberg ( Leihweise), Lieferung in KH Freiberg Labor | SWO-57868 | No | Due to long delivery times pick-up of a solution pack in practice Dr. Morgner Freiberg (loan), delivery in KH Freiberg laboratory | Due to long delivery times pick-up of a solution pack in practice Dr. Morgner Freiberg (loan), delivery in KH Freiberg laboratory | Assisted Replacement | SP | pCO2 | Level A
ABL90 FLEX | Beim ABL 825 die Oxymetrie kalibriert, THB Kal im Service Modus durchgeführt. Am ABL 90 das System gespült und Thb. Kal durchgeführt. Bei den nächsten Vergleichsmessungen lag die max. Abweichung THB bei 0,4 g dl. | Beim ABL 825 die Oxymetrie kalibriert, THB Kal im Service Modus durchgeführt. Am ABL 90 das System gespült und Thb. Kal durchgeführt. Bei den nächsten Vergleichsmessungen lag die max. Abweichung THB bei 0,4 g dl. | SWO-58032 | No | With the ABL 825 the oxymetry calibrates, THB Kal performed in service mode. On the ABL 90 the system rinsed and Thb. Kal performed. For the next comparative measurements, the maximum deviation THB was 0.4 g dl. | With the ABL 825 the oxymetry calibrates, THB Kal performed in service mode. On the ABL 90 the system rinsed and Thb. Kal performed. For the next comparative measurements, the maximum deviation THB was 0.4 g dl. | Cleaned | Fluidic pathway | tHb
ABL90 | 7730 Lot 138 range 145-165 (144) 7740 lot 143 range 97-113 (123) 9050 Lot 380803 4.0-29.5 (73.3 | solution pack and sensor cassette replaced by Sales solution sent to hopsital; on site 10 9 all test are passing | CFABD306481A6AB5 | Yes | 7730 Lot 138 range 145-165 (144) 7740 lot 143 range 97-113 (123) 9050 Lot 380803 4.0-29.5 (73.3 | solution pack and sensor cassette replaced by Sales solution sent to hopsital; on site 10 9 all test are passing | Re classified RMED | Courtesy
ABL90 | came back in to try more trouble shooting | replaced sensor interface PCB not help replaced sensor cassette | CFAC01E148486ABD | No | came back in to try more trouble shooting | replaced sensor interface PCB not help replaced sensor cassette | Defect | HW | SC PCB
ABL90 | 9050 range 4.0-29.5 coming in at 73.3, 9030 range 175-211 coming in at 171 , external QC 7730 range is 145-165 (141) | replaced sensor cassette R0700 today; ran QC all passing then ran blood samples reran blood and QC failed | CFABC2E0481F6AAE | No | 9050 range 4.0-29.5 coming in at 73.3; 9030 range 175-211 coming in at 171 ; external QC 7730 range is 145-165 (141) | replaced sensor cassette R0700 today; ran QC all passing then ran blood samples reran blood and QC failed | Unknown | Revisit required | SC pO2
ABL90 | 9050 range 4.0-29.5 (73.3) 7730 range 145-165 (141) 7740 range 97-113 (123)---- problem is intermittent and very erractic CALLED RADIOMETER FOR INSTRUCTION ON FILLING OUT THE CREDIT FORM NO ANSWER AS OF YET | replaced sensor cassette; ran all levels of QC and they passed; ran blood samples then reran QC and it failed low for pO2 sending analyzer back to Radiometer for testing; QC used for testing | CFABC9BD48026AB1 | No | 9050 range 4.0-29.5 (73.3) 7730 range 145-165 (141) 7740 range 97-113 (123)---- problem is intermittent and very erractic CALLED RADIOMETER FOR INSTRUCTION ON FILLING OUT THE CREDIT FORM NO ANSWER AS OF YET | replaced sensor cassette; ran all levels of QC and they passed; ran blood samples then reran QC and it failed low for pO2 sending analyzer back to Radiometer for testing; QC used for testing | Assisted Replacement | SC | pO2
ABL90 | range level 9030 175-211 coming in at 172, 9050 4.0-29.5 coming in at 62.3, 7730 range is 1450165 coming in at 137, pCO2 failing 7750 range 20.1-25.1 coming in at 14.6 | performed a complete overhaul; replaced sensor cassette and solution pack will complete warranty replacement form; test are passing ran blood then QC and still passing | CFABBD53481C6AA8 | No | range level 9030 175-211 coming in at 172; 9050 4.0-29.5 coming in at 62.3; 7730 range is 1450165 coming in at 137; pCO2 failing 7750 range 20.1-25.1 coming in at 14.6 | performed a complete overhaul; replaced sensor cassette and solution pack will complete warranty replacement form; test are passing ran blood then QC and still passing | Assisted Replacement | SC | pO2
ABL90 | Chequeo de equipo, sustitución de sonda doblada, inlet gasket y conector, sustitución de SP, ejecución de programas de servicio y puesta en marcha, formación de personal para mnto equipo, calibración y pasado ciclos. | Sustitución de SP y sonada, gasket, conector gasket | CFABE01E485C4FA5 | No | Equipment check, bent probe replacement, inlet gasket and connector, SP replacement, service and commissioning program execution, personnel training for equipment, calibration and past cycles. | Sp replacement and sonada, gasket, gasket connector | Assisted Replacement | SP | FT
ABL90 FLEX | 更换903-183一个正常。 | 更换903-183一个正常。 | SWO-57305 | No | Replace 903-183 one normal. | Replace 903-183 one normal. | Defect | HW | SC PCB
ABL90 | Descontaminação do circuito, substituição da entrada, verificação da sonda de entrada, establização e pasasgem do QC. |  | CFABF664480F53FD | No | Decontamination of the circuit, replacement of the inlet, verification of the inlet probe, establishment and ticket of the QC. |  | Cleaned | Fluidic pathway | FT | Inhom QC3
ABL90 | Inspeção do módulo da entrada, ajuste da sonda, substituição da entrada, limpeza do circuito, ajuste dos SLs, calibração da bomba, estabilização e passagem dos QCs. |  | CFAC2FE94838413F | No | Input module inspection, probe adjustment, inlet replacement, circuit cleaning, SLs adjustment, pump calibration, stabilization and QCs passage. |  | Defect | HW | Inlet gasket with holder
AQT90 FLEX | 重新定标正常。 | 重新定标正常。 | SWO-57922 | Yes | Recalibration is normal. | Recalibration is normal. | Re classified RMED | Double
AQT90 FLEX | 更换Hct模块和条码扫描器后，需要再借定标卡用于定标 | 重新定标正常。 | SWO-57895 | No | After replacing the Hct module and barcode scanner, the calibration card needs to be borrowed for calibration | Recalibration is normal. | Defect | HW | Aspiration Module | Needle Wash Unit | 1257
AQT90 FLEX |  |  | SWO-57757 | No |  |  | Defect | HW | Storage Module | Barcode reader, Storage wheel | BR1
AQT90 FLEX |  | 更换903-198一个后未修复，CPU板和存储板也已损坏，需要再借件 | SWO-57731 | No |  | Replacing 903-198 one has not been repaired, the CPU board and storage board have also been damaged and need to be borrowed again | Defect | HW | Analyzer control Module | CPU | 2
AQT90 FLEX |  |  | SWO-57683 | No |  |  | Defect | HW | Aspiration Module | Needle sensor | 4794
ABL825 FLEX |  |  | CFABB2164843E16B | No |  |  | Adjustment | HW | Oxi module | Lamp unit
ABL800 BASIC | 更换902-821一个正常。 | 更换902-821一个正常。 | SWO-57727 | No | Replace the 902-821 one normal. | Replace the 902-821 one normal. | Defect | HW | AutoCheck Module | Barcode reader
ABL825 FLEX |  |  | SWO-57745 | No |  |  | Defect | HW | Printer module | Printer head
ABL820 FLEX | Data logs indicate 704, AutoCheck measurement repeated and 742, AutoCheck carousel failure errors. AutoCheck module was found to have broken glass present, which may have prevented the carousel from operating properly. Site removed much of the broken glass that was present. | Additional carousel cleaning was performed, along with cleaning of the carousel mirror. Verified carousel and inlet probe alignment. AutoCheck ampoules were then successfully run without error. Instrument is within acceptable operating specifications. | CFAB9D0C48514929 | No | Data logs indicate 704, AutoCheck measurement repeated and 742, AutoCheck carousel failure errors. AutoCheck module was found to have broken glass present, which may have prevented the carousel from operating properly. Site removed much of the broken glass that was present. | Additional carousel cleaning was performed, along with cleaning of the carousel mirror. Verified carousel and inlet probe alignment. AutoCheck ampoules were then successfully run without error. Instrument is within acceptable operating specifications. | Cleaned | HW | AutoCheck Module | Ampoule broken
ABL825 FLEX |  |  | SWO-57749 | No |  |  | Defect | HW | DMS_PC | Touch Screen
ABL827 FLEX | Gerät überprüft, Neustart durchgeführt, fällige Elektroden membraniert | Gerät überprüft, Neustart durchgeführt, fällige Elektroden membraniert | SWO-58010 | Yes | Device checked, restart performed, due electrodes membraneed | Device checked, restart performed, due electrodes membraneed | Re classified RMED | Courtesy
ABL837 FLEX |  |  | CFABC7B84849E8F6 | Yes |  |  | Re classified RMED | Installation
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-57923 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL80 FLEX BASIC | 更换测试卡正常。 | 更换测试卡正常。 | SWO-57980 | No | The replacement test card is OK. | The replacement test card is OK.
ABL80 FLEX BASIC | 更换三张新测试卡修复机器 | 更换三张新测试卡修复机器 | SWO-57938 | No | Replace three new test cards to repair the machine | Replace three new test cards to repair the machine
ABL800 BASIC | Checked to find out the flow issues & leakage., checked fluidic & measurement valves, found working OK. Cleaned & reseat pressure transducer for solving gas offset issue. But still issue persists. Need to come again tomorrow for further troubleshooting. | Replaced cable from analyser control board to pressure transducer. Issue resolved. Analyzer ready. | SWO-58429 | Yes | Checked to find out the flow issues & leakage., checked fluidic & measurement valves, found working OK. Cleaned & reseat pressure transducer for solving gas offset issue. But still issue persists. Need to come again tomorrow for further troubleshooting. | Replaced cable from analyser control board to pressure transducer. Issue resolved. Analyzer ready. | Re classified RMED | Double reg w diff date | SWO-5843
ABL800 BASIC | Replaced Calibration 1 solution with a new one, done the calibration, run a sample, machine is working fine. | Replaced Calibration 1 solution with a new one, done the calibration, run a sample, machine is working fine. | SWO-58205 | No | Replaced Calibration 1 solution with a new one, done the calibration, run a sample, machine is working fine. | Replaced Calibration 1 solution with a new one, done the calibration, run a sample, machine is working fine. | Assisted Replacement | Solution bottle | Cal1
ABL80 FLEX BASIC | 更换903-334一个后修复 | 更换903-334一个后修复 | SWO-57947 | No | Replace 903-334 one post-fix | Replace 903-334 one post-fix
ABL90 |  |  | CFBD674CBC5DE9AF | No |  |  | Cleaned | Fluidic pathway
ABL90 |  |  | CFBD6888BC2BE9B1 | No |  |  | Cleaned | Fluidic pathway
AQT90 FLEX | 更换903-127和910-304修复机器 |  | SWO-58040 | No | Replace the 903-127 and 910-304 repair machines |  | Defect | HW | Storage Module | Barcode reader, Storage wheel | 3416
AQT90 FLEX |  |  | SWO-57973 | No |  |  | Defect | HW | Cover and Chassis | Cartridge door | 5784
AQT90 FLEX | 更换903-364一个正常。 | 更换903-364一个正常。 | SWO-57789 | No | Replace 903-364 one normal. | Replace 903-364 one normal. | Defect | HW | Process Module | Optical Unit | 131
ABL800 BASIC |  |  | SWO-57967 | No |  |  | Defect | HW | Analyzer control Module | PCB
ABL800 BASIC | Replaced cable from analyser control board to pressure transducer. Issue resolved. Analyzer ready. | Replaced cable from analyser control board to pressure transducer. Issue resolved. Analyzer ready. | SWO-58430 | No | Replaced cable from analyser control board to pressure transducer. Issue resolved. Analyzer ready. | Replaced cable from analyser control board to pressure transducer. Issue resolved. Analyzer ready. | Defect | HW | Gasmodule | Cable
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-58313 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-58105 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-58246 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL80 CO-OX | Examined display assembly. | Replaced display assembly and cable. Verified operation of ABL 80. | CFAF8711AA385ACE | No | Examined display assembly. | Replaced display assembly and cable. Verified operation of ABL 80.
ABL90 FLEX | Überprüfung im Service Programm Funktionstest i.O. Überprüfung pO2 Kanal- i.O. SP90 und Einlaßdichtung wurde durchKunden erneuert, Gerät meldet aber noch immer inhomogene Lösungen, erneuter Wechsel SP, Fließwege gespült | Überprüfung im Service Programm Funktionstest i.O. Überprüfung pO2 Kanal- i.O. SP90 und Einlaßdichtung wurde durchKunden erneuert, Gerät meldet aber noch immer inhomogene Lösungen; erneuter Wechsel SP, Fließwege gespült | SWO-58203 | No | Verification in the service program function test i.O. check pO2 channel- i.O. SP90 and inlet seal has been renewed bycustomers, but device still reports inhomogeneous solutions; re-change SP, flow paths rinsed | Verification in the service program function test i.O. check pO2 channel- i.O. SP90 and inlet seal has been renewed bycustomers, but device still reports inhomogeneous solutions; re-change SP, flow paths rinsed | Assisted Replacement | SP | FT
ABL90 FLEX | tHb-Kal wiederholt. O.K. | tHb-Kal wiederholt. O.K. | SWO-58170 | No | tHb-Kal wiederholt. O.K. | tHb-Kal wiederholt. O.K. | Adjustment | tHb Cal
ABL90 FLEX | Fehlersuche, bei gewissen Darstellungen (vor allem bei hellen Farben) ist der Bildschirm lila gefärbt, der Analysator ist aber Messbereit, auch ein Neustart brachte keine Besserung, Software V3.1MR6 nochmal installiert->keine Besserung, vermutlich Bildschirm oder CPU defekt, die Ersatzteile müssen bestellt werden, SC erneuert(fällig) Anwender ist informiert |  | SWO-58151 | No | Troubleshooting, with certain displays (especially with bright colors) the screen is colored purple, but the analyzer is ready to measure, even a restart did not bring any improvement, software V3.1MR6 reinstalled->no improvement, presumably screen or CPU defective , the spare parts must be ordered, SC renewed (due) user is informed |  | Defect | HW | Display
TCM400 |  |  | SWO-58027 | No
ABL90 FLEX | 更换测试卡修复机器 | 更换测试卡修复机器 | SWO-58091 | No | Replace the test card repair machine | Replace the test card repair machine | Assisted Replacement | SC | Response | K, Na
AQT90 FLEX | 更换903-210一个正常。 | 更换903-210一个正常。 | SWO-57937 | No | Replace the 903-210 one normal. | Replace the 903-210 one normal. | Defect | HW | Process Module | Cup dryer, PCB | 17
AQT90 FLEX |  |  | SWO-58025 | No |  |  | Assisted Replacement | Cartridge | CKMB | Calibration
AQT90 FLEX | 更换进样针903-171一个后修复 | 更换进样针903-171一个后修复 | SWO-58278 | No | Replace the sample needle 903-171 after repair | Replace the sample needle 903-171 after repair | Defect | HW | Aspiration Module | Bend needle | 4794
ABL800 BASIC | Austausch der Display Unit und des Verbindungskabels, nach Wechsel sind Tastetöne beim Berühren des Displays vorhandenen, diese lassen sich nicht abstellen-> Info an Kunden, evtl. Tausch der CPU Unit nötig, Kunde beobachtet Verhalten des ABL | Austausch der Display Unit und des Verbindungskabels, nach Wechsel sind Tastetöne beim Berühren des Displays vorhandenen, diese lassen sich nicht abstellen-> Info an Kunden, evtl. Tausch der CPU Unit nötig, Kunde beobachtet Verhalten des ABL | SWO-58191 | No | Replacement of the display unit and the connection cable, after change, button tones are present when touching the display, these cannot be turned off-> info to customers, possibly exchange of the CPU unit necessary, customer observes behavior of the ABL | Replacement of the display unit and the connection cable, after change, button tones are present when touching the display, these cannot be turned off-> info to customers, possibly exchange of the CPU unit necessary, customer observes behavior of the ABL | Defect | HW | DMS_PC | Touch Screen
ABL825 FLEX | Messkammer und Kontakte pH und Referenz gereinigt - Erneuerung pH Dichtring - Funktionsüberprüfung etc. | Messkammer und Kontakte pH und Referenz gereinigt - Erneuerung pH Dichtring - Funktionsüberprüfung etc. | SWO-58154 | No | Measuring chamber and contacts pH and reference cleaned - renewal pH sealing ring - function check etc. | Measuring chamber and contacts pH and reference cleaned - renewal pH sealing ring - function check etc. | Defect | Electrode | Electrode | Sealing ring
ABL825 FLEX |  |  | CFAF8118AA15EDD1 | No |  |  | Assisted Replacement | Electrode | Membrane | pCO2
ABL825 FLEX | Gerät überprüft, pH BG Modul 823R2242N003 ausgetauscht, Abgleich durchgeführt, Kalibrationen und QC o.k. | Dichtungsring pH Elektrode ersetzt, pH BG Modul zurückgetauscht 902-823R2242N003 eingebaut, 823R2278N002 zurück im Bestand, Software aktualisiert | SWO-58372 | No | Device checked, pH BG module 823R2242N003 replaced, matching performed, calibrations and QC ok. | Sealing ring pH electrode replaced, pH BG module replaced 902-823R2242N003 built-in, 823R2278N002 back in stock, software updated | Defect | HW | pH/BG Module | unknown
ABL825 FLEX |  |  | CFAF89C9AA5FEDD9 | No |  |  | Assisted Replacement | Electrode | Membrane | pCO2
ABL825 FLEX |  |  | CFAF8576AA12EDD5 | No |  |  | Assisted Replacement | Electrode | Electrode | Reference
ABL825 FLEX | Fehlersuche, Einlasssonde begradigt, Initialisierung abgewartet, Kal u. QC i.O. |  | SWO-58251 | No | Troubleshooting, inlet probe straightened, initialization awaited, Kal & QC i.O. |  | Defect | HW | Inlet module | Probe bend
ABL835 FLEX |  |  | CFAF8951AA469010 | Yes |  |  | Re classified RMED | Logistical upSWO
AQT90 FLEX | Motor getauscht alle Funktionen IO | der Anwender hat bzw. wollte eine neue Testkassette einlegen; die Sorage-Door schließt nicht mehr; der Anwender hat das Gerät schon kurzzeitig aus- und wieder eingeschaltet, allerdings ohne Erfolg; Fehler bleibt; der Anwender wünscht einen Vorort-Service zur Beseitigung der Geräte-Störung; | SWO-58095 | No | Motor swaps all functions IO | the user has or wanted to insert a new test cassette; the Sorage-Door no longer closes; the user has already switched the device off and on again for a short time, but to no avail; Error remains; the user wants a suburban service to eliminate the device malfunction; | Defect | HW | Storage Module | Stepper motor | 1999
AQT90 FLEX | Motor defekt ET muss besorgt werden | der Anwender hat bzw. wollte eine neue Testkassette einlegen; die Sorage-Door schließt nicht mehr; der Anwender hat das Gerät schon kurzzeitig aus- und wieder eingeschaltet, allerdings ohne Erfolg; Fehler bleibt; der Anwender wünscht einen Vorort-Service zur Beseitigung der Geräte-Störung; | SWO-58096 | Yes | Engine defective ET must be taken care of | the user has or wanted to insert a new test cassette; the Sorage-Door no longer closes; the user has already switched the device off and on again for a short time, but to no avail; Error remains; the user wants a suburban service to eliminate the device malfunction; | Re classified RMED | Double entry
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-58233 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL80 FLEX BASIC | 更换测试卡正常。 | 更换测试卡正常。 | SWO-58526 | No | The replacement test card is OK. | The replacement test card is OK.
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-58239 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL80 FLEX BASIC | 更换测试卡正常 | 更换测试卡正常 | SWO-58341 | No | Replace the test card ok | Replace the test card ok
ABL90 FLEX | Inlet module verwijdert, nu is duidelijk te zien dat de naald aan de voorkant krom is, waardoor de inletgasket niet vloeiend terug kan schuiven. Dit is zeer waarschijnlijk ontstaan door het verkeerd verwijderen van een spuit, na aanbieden monster. Derhalve naald en inletgasket vervangen. Hierna werkt inletmodule weer vloeiend. | Inlet module verwijdert, nu is duidelijk te zien dat de naald aan de voorkant krom is, waardoor de inletgasket niet vloeiend terug kan schuiven. Dit is zeer waarschijnlijk ontstaan door het verkeerd verwijderen van een spuit, na aanbieden monster. Derhalve naald en inletgasket vervangen. Hierna werkt inletmodule weer vloeiend. | SWO-58354 | No | Inlet module, now it is clear to see that the needle on the front is crooked, so the inletgasket cannot slide back smoothly. This is very likely caused by the improper removal of a syringe, after offering sample. Thereupon needle and Inletgasket replaced. After this, inletmodule works smoothly again. | Inlet module, now it is clear to see that the needle on the front is crooked, so the inletgasket cannot slide back smoothly. This is very likely caused by the improper removal of a syringe, after offering sample. Thereupon needle and Inletgasket replaced. After this, inletmodule works smoothly again. | Replaced | probe | Bend
ABL90 FLEX | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-58252 | No | Repair after replacing a new test card | Repair after replacing a new test card | Assisted Replacement | SC | pCO2
ABL90 FLEX |  | B | SWO-58227 | No |  |  | Defect | HW | Inlet module
ABL90 | Verificação dos registros, limpezas sucessivas ao circuito, calibração e passagem dos QCs. | B | CFB0B2422030C28C | No | Verification of records, successive cleanings to the circuit, calibration and passage of the QCs. |  | Cleaned | Fluidic pathway | pO2 | QC
ABL825 FLEX | De driver werd opnieuw ingeladen. Omdat deze niet automatisch werd gevonden moest deze volledig handmatig worden geïnstalleerd . (C windows system32) Na het opstarten werd er een schermkalibratie uitgevoerd. Voorlopig afwachten of het probleem hiermee opgelost werd. | De driver werd opnieuw ingeladen. Omdat deze niet automatisch werd gevonden moest deze volledig handmatig worden geïnstalleerd . (C windows system32) Na het opstarten werd er een schermkalibratie uitgevoerd. Voorlopig afwachten of het probleem hiermee opgelost werd. | SWO-58418 | No | The driver was re-loaded. Because it was not found automatically, it had to be installed completely manually. (C windows System32) After booting, a screen calibration was performed. For now, wait and see if this solves the problem. | The driver was re-loaded. Because it was not found automatically, it had to be installed completely manually. (C windows System32) After booting, a screen calibration was performed. For now, wait and see if this solves the problem. | Defect | SW | DMS_PC | Windows | Blue screen
ABL825 FLEX |  | B | CFB102B3201FEEC2 | No |  |  | Cleaned | Fluidic pathway | - | pCO2
ABL825 FLEX | Dichtung erneuert | Dichtung erneuert | SWO-58530 | No | Seal renewed | Seal renewed | Assisted Replacement | Electrode | Electrode | Glu | Sealing ring
ABL825 FLEX | Belüftungsstrecke gespült, LS-Justage, Kal, QC | Belüftungsstrecke gespült, LS-Justage, Kal, QC | SWO-58436 | No | Ventilation line rinsed, LS-Adjustment, Kal, QC | Ventilation line rinsed, LS-Adjustment, Kal, QC | Defect | HW | El/Met Module | Valve
ABL825 FLEX |  | B | CFB0B158202806B3 | No |  |  | Defect | HW | Oxi module | Spectrometer | Neon number high
ABL825 FLEX | Nullpunkt für PO2, Glu. und Lac. neu kalibriert, Funktionskontrolle ok | Nullpunkt für PO2, Glu. und Lac. neu kalibriert, Funktionskontrolle ok | SWO-58494 | No | Zero point for PO2, Glu. and Lac. recalibrated, function control ok | Zero point for PO2, Glu. and Lac. recalibrated, function control ok | Adjustment | HW | pH/BG Module | zero current adjustment
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-58709 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL80 FLEX CO-OX | Reinigung des Gerätes, Installation LP, Funktionskontrolle | Reinigung des Gerätes, Installation LP, Funktionskontrolle | SWO-58680 | No | Cleaning the device, installation LP, function control | Cleaning the device, installation LP, function control
ABL80 CO-OX | CONFIRMED WHEN ARRIVED ONSITE | REPLACED HEMOLYZER. | CFB2388C96412878 | No | CONFIRMED WHEN ARRIVED ONSITE | REPLACED HEMOLYZER.
ABL90 |  | B | CFB215F0964A9329 | No |  |  | Cleaned | Fluidic pathway
ABL90 |  | B | CFBD6B56BC56E9B4 | No |  |  | Defect | HW | Inlet Gasket
ABL90 | Examined event logs | Advised customer to replace solution pack, for troubleshooting purposes and to flush analyzer. Upon arrival, customer had resolve the issue by replacing solution pack. Handed solution pack replacement to customer. Verified operation of ABL 90. | CFB82927E435DA27 | No | Examined event logs | Advised customer to replace solution pack, for troubleshooting purposes and to flush analyzer. Upon arrival, customer had resolve the issue by replacing solution pack. Handed solution pack replacement to customer. Verified operation of ABL 90. | Assisted Replacement | SP | FT
AQT90 FLEX | 更换902-924一个正常。 | 更换902-924一个正常。 | SWO-58279 | No | Replace the 902-924 one normal. | Replace the 902-924 one normal. | Defect | HW | Analyzer control Module | Barcode reader, display | BR2
ABL805 FLEX | checked logs for the errors checked valve in service menu for failurs found over 200 fails on cal 1 valve and over 100 fails on valve 2 and 3 | replaced the 3 valves | CFB1E2D0964A6250 | Yes | checked logs for the errors checked valve in service menu for failurs found over 200 fails on cal 1 valve and over 100 fails on valve 2 and 3 | replaced the 3 valves | Re classified RMED
ABL800 BASIC | 比較電極交換実施。 | 比較電極交換実施。 | SWO-60417 | No | Comparative electrode exchange was implemented. | Comparative electrode exchange was implemented. | Assisted Replacement | Electrode | Electrode | Reference
ABL825 FLEX | Inlet Modul zurück gebaut .Überprüfung des Gerätes ohne Erkenntnisse. Kal und QC ok. | Inlet Modul zurück gebaut .Überprüfung des Gerätes ohne Erkenntnisse. Kal und QC ok. | SWO-58624 | No | Inlet module built back. Checking the device without knowledge. Kal and QC ok. | Inlet module built back. Checking the device without knowledge. Kal and QC ok. | Unknown | Unknown | Unknown | Unknown
ABL827 FLEX | mehrere Kal durchgeführt, Elektroden kontrolliert, QC, Remotevervindung getestet - i.O. | mehrere Kal durchgeführt, Elektroden kontrolliert, QC, Remotevervindung getestet - i.O. | SWO-58596 | No | several cal, electrodes controlled, QC, remote vvination tested - i.O. | several cal, electrodes controlled, QC, remote vvination tested - i.O. | Adjustment | Electrode | Electrode | Crea
ABL835 FLEX | FLXQモジュール引き上げ | FLXQモジュール引き上げ | SWO-58573 | No | FLXQ Module Pull | FLXQ Module Pull | Defect | HW | FlexQ Module | Mixer
ABL800 BASIC | Gerät überprüft, Meßkammer und Luftauslass gesäubert, Elektroden kontrolliert, pCO2 membraniert, E777 (RV-02) erneuert | Gerät überprüft, Meßkammer und Luftauslass gesäubert, Elektroden kontrolliert, pCO2 membraniert, E777 (RV-02) erneuert | SWO-58504 | No | Device checked, measuring chamber and air outlet cleaned, electrodes checked, pCO2 membraneed, E777 (RV-02) renewed | Device checked, measuring chamber and air outlet cleaned, electrodes checked, pCO2 membraneed, E777 (RV-02) renewed | Assisted Replacement | Electrode | Electrode | pH
ABL810 FLEX | Pumpenschlauch BG-Modul erneuert. | Pumpenschlauch BG-Modul erneuert. | SWO-58581 | No | Pump hose BG module renewed. | Pump hose BG module renewed. | Defect | HW | pH/BG Module | Pump tubing
ABL815 FLEX | Beschädigtes Tray ersetzt und sämtliche Funktionen des Moduls im Servicemenü überprüft kalibriert. Dichtringe Cl Dummy sowie pH ersetzt, Dichtigkeit überprüft. Barometrischen Druck korrigiert von 733mmHg auf 753mmHg (aktueller Wert von wetter.com) Kalibration, QC Messung => OK | Beschädigtes Tray ersetzt und sämtliche Funktionen des Moduls im Servicemenü überprüft kalibriert. Dichtringe Cl Dummy sowie pH ersetzt, Dichtigkeit überprüft. Barometrischen Druck korrigiert von 733mmHg auf 753mmHg (aktueller Wert von wetter.com) Kalibration, QC Messung => OK | SWO-58673 | No | Replaced damaged tray and calibrated all functions of the module checked in the service menu. Sealing rings Cl dummy as well as pH replaced, tightness checked. Barometric pressure corrected from 733mmHg to 753mmHg (current value of wetter.com) calibration, QC measurement => OK | Replaced damaged tray and calibrated all functions of the module checked in the service menu. Sealing rings Cl dummy as well as pH replaced, tightness checked. Barometric pressure corrected from 733mmHg to 753mmHg (current value of wetter.com) calibration, QC measurement => OK | Defect | HW | FlexQ Module | Tray
ABL825 FLEX |  | B | CFB846FFE43850C2 | No |  |  | Assisted Replacement | Electrode | Electrode | pCO2
AQT90 FLEX |  | B | CFB228FD96575FE2 | No |  |  | Defect | HW | Process Module | Optical Unit | 135
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-59398 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL80 FLEX BASIC | 更换测试卡正常。 | 更换测试卡正常。 | SWO-59576 | No | The replacement test card is OK. | The replacement test card is OK.
ABL80 FLEX BASIC |  | B | SWO-58705 | No
ABL90 FLEX | Kein Fehler mehr feststellbar | Kein Fehler mehr feststellbar | SWO-59590 | No | No more error detectable | No more error detectable | No fault found | Solved by customer | SP replaced
ABL90 FLEX | div. Einstellungen am Gerät vorgenommen, alles i.O. | div. Einstellungen am Gerät vorgenommen, alles i.O. | SWO-59625 | No | div. Settings made on the device, everything in the | div. Settings made on the device, everything in the | Setup | Fixed assigned values
ABL90 FLEX | Auf Anforderung von RDEU TS Service Dump erstellt und an TS per Mail geschickt | Auf Anforderung von RDEU TS Service Dump erstellt und an TS per Mail geschickt | SWO-59501 | No | Created on request of RDEU TS Service Dump and sent to TS by e-mail | Created on request of RDEU TS Service Dump and sent to TS by e-mail | Unknown | FT | MR6
ABL90 | The analyzer was tring to communicate with Cerner. | I set the analyzer up for no communication and rebooted the analyzer with no errors. The analyzer cannot be set up for communications at this time. | CFB35FC40C17875F | No | The analyzer was tring to communicate with Cerner. | I set the analyzer up for no communication and rebooted the analyzer with no errors. The analyzer cannot be set up for communications at this time. | Defect | SW | OS sybase 16
ABL90 FLEX | 更换新测试卡修复机器 | 更换新测试卡修复机器 | SWO-59407 | No | Replace the new test card repair machine | Replace the new test card repair machine | Assisted Replacement | SC | Glu/Lac
ABL90 FLEX | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-58737 | No | Repair after replacing a new test card | Repair after replacing a new test card | Assisted Replacement | SC | Cl | Cal
AQT90 FLEX | 更换903-236一个正常。 | 更换903-236一个正常。 | SWO-58447 | No | Replace the 903-236 one normal. | Replace the 903-236 one normal. | Defect | HW | Process Module | Process wheel | 124
ABL825 FLEX | PH Elektrode (E 777 RI 01) geprüft und ausgetauscht, Funktionskontrolle ok | B | SWO-59436 | No | PH electrode (E 777 RI 01) tested and replaced, function control ok |  | Assisted Replacement | Electrode | Electrode | pH
ABL835 FLEX | komplettes BG Modul getauscht brachte auch keine Verbesserung. Erneut eine pCO 2 Elektrode getauscht. | komplettes BG Modul getauscht brachte auch keine Verbesserung. Erneut eine pCO 2 Elektrode getauscht. | SWO-59579 | No | complete BG module exchanged also brought no improvement. Again, a pCO 2 electrode is replaced. | complete BG module exchanged also brought no improvement. Again, a pCO 2 electrode is replaced. | Assisted Replacement | Electrode | Electrode | pCO2
ABL825 FLEX | Friction Unit in Einlassmodul ersetzt | Inlet Modul zurück gebaut .Überprüfung des Gerätes ohne Erkenntnisse. Kal und QC ok. | SWO-59482 | No | Friction Unit replaced in inlet module | Inlet module built back. Checking the device without knowledge. Kal and QC ok. | Defect | HW | Inlet module | Friction detector unit
ABL825 FLEX | Messkammer und Schlauchsystem geprüft und gereinigt, mehrfach Proteinentfernerprogramm durchlaufen lassen, Funktionskontrolle ok | Kundin bittet um einen Rückruf und Service. | SWO-59437 | No | Measuring chamber and hose system tested and cleaned, allow multiple protein remover program to pass through, function control ok | Customer asks for a callback and service. | Cleaned | Fluidic pathway | El/Met Module | Clot
ABL825 FLEX |  | B | CFB3739B0C274A77 | No |  |  | Defect | HW | Oxi module | Hemolyzer | COHb
ABL800 BASIC | Switched OFF the machine, removed the display connections, done the touch screen calibration, but still the same problem persists. So the display unit has to be replaced. Ordered the Display unit from Mumbai Office, will receive it on or before Monday. | B | SWO-59807 | No | Switched OFF the machine, removed the display connections, done the touch screen calibration, but still the same problem persists. So the display unit has to be replaced. Ordered the Display unit from Mumbai Office, will receive it on or before Monday. |  | Defect | HW | DMS_PC | Touch Screen
ABL800 BASIC | pCO2 Elektrode, unteren Dichtring von pH und Dummyelektrode erneuert. Kalibrationen und QC's vorgenommen | pCO2 Elektrode, unteren Dichtring von pH und Dummyelektrode erneuert. Kalibrationen und QC's vorgenommen | SWO-59618 | No | pCO2 electrode, lower sealing ring of pH and dummy electrode renewed. Calibrations and QC's made | pCO2 electrode, lower sealing ring of pH and dummy electrode renewed. Calibrations and QC's made | Assisted Replacement | Electrode | Electrode | pCO2
ABL800 BASIC | pCO2 Elektrode neu membraniert. Kalibrationen und QC's durchgeführt. | pCO2 Elektrode neu membraniert. Kalibrationen und QC's durchgeführt. | SWO-59616 | No | pCO2 electrode. calibrations and QC's. | pCO2 electrode. calibrations and QC's. | Assisted Replacement | Electrode | Membrane | pCO2
ABL800 BASIC | pCO2 Elektrode, unteren Dichtring von pH und Dummyelektrode erneuert. Referenzelektrode neu membraniert. Kalibrationen und QC's vorgenommen. | pCO2 Elektrode, unteren Dichtring von pH und Dummyelektrode erneuert. Referenzelektrode neu membraniert. Kalibrationen und QC's vorgenommen. | SWO-59614 | No | pCO2 electrode, lower sealing ring of pH and dummy electrode renewed. reference electrode re-membraneed. calibrations and QC's. | pCO2 electrode, lower sealing ring of pH and dummy electrode renewed. reference electrode re-membraneed. calibrations and QC's. | Assisted Replacement | Electrode | Electrode | pCO2
ABL800 BASIC | Elektrodenspitze der pCO2 Membran gereinigt, E788 mit Membranmaterial anderer Lot Nr. membraniert. Mehrere Kal und QC durchgeführt und Verlauf betrachtet. Verlauf muss weiter beobachtet werden. | Elektrodenspitze der pCO2 Membran gereinigt, E788 mit Membranmaterial anderer Lot Nr. membraniert. Mehrere Kal und QC durchgeführt und Verlauf betrachtet. Verlauf muss weiter beobachtet werden. | SWO-59587 | No | Electrode tip of the pCO2 membrane cleaned, E788 membraneed with membrane material of other solder No. Several Kal and QC performed and considered course. The course must be monitored further. | Electrode tip of the pCO2 membrane cleaned, E788 membraneed with membrane material of other solder No. Several Kal and QC performed and considered course. The course must be monitored further. | Assisted Replacement | Electrode | Membrane | pCO2
AQT90 FLEX | AQT überprüft. Sensoren am StorageWheel und Punsh-Down gereinigt. Cartridge-Halterung am Storage-Wheel aus der Halterung raus. Halterung i.O.! Halterung neu justiert. Gerätestartphase und QC durchgeführt. | AQT überprüft. Sensoren am StorageWheel und Punsh-Down gereinigt. Cartridge-Halterung am Storage-Wheel aus der Halterung raus. Halterung i.O.! Halterung neu justiert. Gerätestartphase und QC durchgeführt. | SWO-59518 | No | AQT. sensors on the StorageWheel and Punsh-Down. Cartridge mount on the storage wheel out of the bracket. Bracket in the end! bracket is re-adjusted. device start phase and QC. | AQT. sensors on the StorageWheel and Punsh-Down. Cartridge mount on the storage wheel out of the bracket. Bracket in the end! bracket is re-adjusted. device start phase and QC. | Cleaned | HW | Storage Module | Storage wheel, Sensor | 1999,6,24
ABL80 FLEX BASIC | 更换新测试卡修复机器 | 更换新测试卡修复机器 | SWO-59652 | No | Replace the new test card repair machine | Replace the new test card repair machine
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-59559 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL80 FLEX CO-OX | Fehlersuche, Reinigung LS, Erneuerung Schlauch, Funktionskontrolle | Fehlersuche, Reinigung LS, Erneuerung Schlauch, Funktionskontrolle | SWO-59657 | No | Troubleshooting, cleaning LS, renewal hose, function control | Troubleshooting, cleaning LS, renewal hose, function control
ABL80 FLEX BASIC | 更换新的感应板后故障消除，再将原感应板替换新装感应板后故障未重现。最终未进行备件更换，机器恢复正常使用 | 更换新的感应板后故障消除，再将原感应板替换新装感应板后故障未重现。最终未进行备件更换，机器恢复正常使用 | SWO-59404 | No | The fault is eliminated after replacing the new sensing plate, and the fault does not recur after replacing the original sensing plate with the new sensing plate. Final lying spare parts were not replaced and the machine was returned to normal use | The fault is eliminated after replacing the new sensing plate, and the fault does not recur after replacing the original sensing plate with the new sensing plate. Final lying spare parts were not replaced and the machine was returned to normal use
ABL80 FLEX BASIC | 更换测试卡正常。 | 更换测试卡正常。 | SWO-59896 | No | The replacement test card is OK. | The replacement test card is OK.
ABL90 FLEX | Fehlersuche, div. Teile gew. u. Serviceprogramme getestet, LP gew., Verstopfung aus Hemolyzer entf., Neustart, Kal. + QC i.O. | Fehlersuche; div. Teile gew. u. Serviceprogramme getestet; LP gew.; Verstopfung aus Hemolyzer entf.; Neustart; Kal. + QC i.O. | SWO-59627 | No | Troubleshooting; various parts. and service programs tested; LP. Constipation from hemolyzer sether;; Restart; Kal. + QC in the first place | Troubleshooting; various parts. and service programs tested; LP. Constipation from hemolyzer sether;; Restart; Kal. + QC in the first place | Cleaned | Fluidic pathway | clot hemolyzer
ABL90 FLEX | Den Netzwerkstecker am ABL gezogen und neu gestartet. Seit dem Neustart kam wieder die Bekannte Meldung und war nicht wegzukriegen (Radiance Häkchen raus>> Neustart Meldung weg. Radiance Häkchen gesetzt >> Meldung wird wieder angezeigt>> Neustart Meldung nach wie vor da.) Nach dem ich wieder die AU-Software neu installiert habe, ist die Meldung weg, diese kommt aber wieder wenn ich den Netzwerkstecker ziehe und den Analysator neu starte. Um die Meldung wegzubekommen >AU-Software neu installiert. Radiance Verbindung getrennt, Abl runtergefahren und auf der Station aufgestellt>>Radiance Verbindung aktiviert >>die Meldung („an unexpected problem has occurred. the application cannot continue executing“ ) wird wieder angezeigt, mit O.K bestätigt>>ABL neu gestartet>>die Meldung wird wieder angezeigt. | Den Netzwerkstecker am ABL gezogen und neu gestartet. Seit dem Neustart kam wieder die Bekannte Meldung und war nicht wegzukriegen (Radiance Häkchen raus>> Neustart Meldung weg. Radiance Häkchen gesetzt >> Meldung wird wieder angezeigt>> Neustart Meldung nach wie vor da.) Nach dem ich wieder die AU-Software neu installiert habe, ist die Meldung weg, diese kommt aber wieder wenn ich den Netzwerkstecker ziehe und den Analysator neu starte. Um die Meldung wegzubekommen >AU-Software neu installiert. Radiance Verbindung getrennt, Abl runtergefahren und auf der Station aufgestellt>>Radiance Verbindung aktiviert >>die Meldung („an unexpected problem has occurred. the application cannot continue executing“ ) wird wieder angezeigt, mit O.K bestätigt>>ABL neu gestartet>>die Meldung wird wieder angezeigt. | SWO-59966 | No | Pulled and restarted the network plug on the ABL. Since the reboot came again the known message and could not get away (Radiance check box out>> Restart message away. Radiance check mark set >> Message is displayed again>> Restart message still there.) After I have reinstalled the AU software, the message is gone, but it comes back when I pull the network plug and restart the analyzer. To get rid of the message >AU software reinstalled. Radiance connection disconnected, Abl shut down and set up on the station>>Radiance connection activated >>the message ("an unexpected problem has occurred. the application cannot continue executing" ) is displayed again, confirmed with O.K>>ABL restarted>>the message will be displayed again. | Pulled and restarted the network plug on the ABL. Since the reboot came again the known message and could not get away (Radiance check box out>> Restart message away. Radiance check mark set >> Message is displayed again>> Restart message still there.) After I have reinstalled the AU software, the message is gone, but it comes back when I pull the network plug and restart the analyzer. To get rid of the message >AU software reinstalled. Radiance connection disconnected, Abl shut down and set up on the station>>Radiance connection activated >>the message ("an unexpected problem has occurred. the application cannot continue executing" ) is displayed again, confirmed with O.K>>ABL restarted>>the message will be displayed again. | Defect | SW | OS sybase 16
ABL90 FLEX | Den Netzwerkstecker am ABL gezogen und neu gestartet. Seit dem Neustart kam wieder die Bekannte Meldung und war nicht wegzukriegen (Radiance Häkchen raus>> Neustart Meldung weg. Radiance Häkchen gesetzt >> Meldung wird wieder angezeigt>> Neustart Meldung nach wie vor da.) Nach dem ich wieder die AU-Software neu installiert habe, ist die Meldung weg, diese kommt aber wieder wenn ich den Netzwerkstecker ziehe und den Analysator neu starte. Um die Meldung wegzubekommen >AU-Software neu installiert. Radiance Verbindung getrennt, Abl runtergefahren und auf der Station aufgestellt>>Radiance Verbindung aktiviert >>die Meldung („an unexpected problem has occurred. the application cannot continue executing“ ) wird wieder angezeigt, mit O.K bestätigt>>ABL neu gestartet>>die Meldung wird wieder angezeigt. | Den Netzwerkstecker am ABL gezogen und neu gestartet. Seit dem Neustart kam wieder die Bekannte Meldung und war nicht wegzukriegen (Radiance Häkchen raus>> Neustart Meldung weg. Radiance Häkchen gesetzt >> Meldung wird wieder angezeigt>> Neustart Meldung nach wie vor da.) Nach dem ich wieder die AU-Software neu installiert habe, ist die Meldung weg, diese kommt aber wieder wenn ich den Netzwerkstecker ziehe und den Analysator neu starte. Um die Meldung wegzubekommen >AU-Software neu installiert. Radiance Verbindung getrennt, Abl runtergefahren und auf der Station aufgestellt>>Radiance Verbindung aktiviert >>die Meldung („an unexpected problem has occurred. the application cannot continue executing“ ) wird wieder angezeigt, mit O.K bestätigt>>ABL neu gestartet>>die Meldung wird wieder angezeigt. | SWO-59973 | No | Pulled and restarted the network plug on the ABL. Since the reboot came again the known message and could not get away (Radiance check box out>> Restart message away. Radiance check mark set >> Message is displayed again>> Restart message still there.) After I have reinstalled the AU software, the message is gone, but it comes back when I pull the network plug and restart the analyzer. To get rid of the message >AU software reinstalled. Radiance connection disconnected, Abl shut down and set up on the station>>Radiance connection activated >>the message ("an unexpected problem has occurred. the application cannot continue executing" ) is displayed again, confirmed with O.K>>ABL restarted>>the message will be displayed again. | Pulled and restarted the network plug on the ABL. Since the reboot came again the known message and could not get away (Radiance check box out>> Restart message away. Radiance check mark set >> Message is displayed again>> Restart message still there.) After I have reinstalled the AU software, the message is gone, but it comes back when I pull the network plug and restart the analyzer. To get rid of the message >AU software reinstalled. Radiance connection disconnected, Abl shut down and set up on the station>>Radiance connection activated >>the message ("an unexpected problem has occurred. the application cannot continue executing" ) is displayed again, confirmed with O.K>>ABL restarted>>the message will be displayed again. | Defect | SW | OS sybase 16
ABL90 | Battery will not take charge. | Battery replaced under Warranty. | CFB48CB082451D75 | No | Battery will not take charge. | Battery replaced under Warranty. | Defect | HW | Battery module
ABL90 | Battery unable to come fully charged. | Replace Battery | CFB48DC0824E1D78 | No | Battery unable to come fully charged. | Replace Battery | Defect | HW | Battery module
ABL90 FLEX | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-59600 | No | Repair after replacing a new test card | Repair after replacing a new test card | Assisted Replacement | SC | Reference | 1361
ABL90 FLEX | 更换试剂包正常。 | 更换试剂包正常。 | SWO-59569 | No | The replacement kit is OK. | The replacement kit is OK. | Assisted Replacement | SP | Glu/Lac | H2??
ABL90 FLEX | 更换新试剂包修复机器 | 更换新试剂包修复机器 | SWO-59672 | No | Replace the new reagent kit repair machine | Replace the new reagent kit repair machine | Assisted Replacement | SP | FT | QC3
AQT90 FLEX | 更换847-606一个正常。 | 更换847-606一个正常。 | SWO-58725 | No | Replace 847-606 one normal. | Replace 847-606 one normal. | Defect | HW | Process Module | Cup dryer, Air filter | 17
ABL837 FLEX | did the cleaning and protein removal, did the calibration now instrument working. | did the cleaning and protein removal, did the calibration now instrument working. | SWO-59857 | No | did the cleaning and protein removal, did the calibration now instrument working. | did the cleaning and protein removal, did the calibration now instrument working. | Revisit required | HW | Oxi module | tHb
ABL825 FLEX |  | B | CFB4739D8217C6F1 | No |  |  | Defect | HW | Oxi module | PCB
ABL825 FLEX | 更换945-614一个正常。 | 更换945-614一个正常。 | SWO-58699 | No | Replace 945-614 one normal. | Replace 945-614 one normal. | Assisted Replacement | Electrode | Electrode | pH
ABL825 FLEX | upon inspection the Glu membrane jacket had no liquid and the Lac was giving zero current errors | replaced Glu membrane performed Lac zero current in serivce | CFB47B6982425644 | No | upon inspection the Glu membrane jacket had no liquid and the Lac was giving zero current errors | replaced Glu membrane performed Lac zero current in serivce | Assisted Replacement | Electrode | Membrane | Glu
ABL825 FLEX | Ca. Elektrode (E733 17-12) geprüft und ausgetauscht, Funktionskontrolle ok | Ca. Elektrode (E733 17-12) geprüft und ausgetauscht, Funktionskontrolle ok | SWO-59679 | No | Approx. electrode (E733 17-12) tested and replaced, function control ok | Approx. electrode (E733 17-12) tested and replaced, function control ok | Assisted Replacement | Electrode | Electrode | Ca
ABL815 FLEX | Fehlersuche, div. Fließwege gereinigt, Einlaßdichtung gew. -> weiter Flüssigkeitsaustritt, Einlaßnadel überprüft = unrund -> gew., danach alles i.O., somit Anwenderfehler | Kunden empfohlen eine LS Justage und Pumpen Kal. zu starten, sollte der Fehler danach nicht behoben sein wird sich der Kunde wieder melden. | SWO-59712 | No | Troubleshooting; div. flow paths cleaned; inlet seal. -> further liquid leakage; Inlet needle checked = unrounded -> wt., then everything in the order, thus user error | Customers recommended an LS adjustment and pumps Kal. if the error is not fixed after that, the customer will report back. | Defect | HW | Inlet module | Probe
ABL825 FLEX | Elyt und BGA Modul Pump Tubing erneuert. PO2 erneut membraniert, wegen Kal. Fehler. | Elyt und BGA Modul Pump Tubing erneuert. PO2 erneut membraniert, wegen Kal. Fehler. | SWO-59866 | No | Elyt and BGA module pump tubing renewed. PO2 again membraneed, due to Cal. Error. | Elyt and BGA module pump tubing renewed. PO2 again membraneed, due to Cal. Error. | Defect | HW | pH/BG Module | Pump tubing
ABL825 FLEX |  | B | CFB4AF678201B227 | No |  |  | Assisted Replacement | Electrode | Electrode | pCO2
ABL825 FLEX | képernyő újra kalibrálása, ideiglenes teszt alaplap (910-321) beépítése a készülékbe, WinXPE és analizátor szoftver újra telepítése | B | SWO-59428 | No | Screen re-calibrating, temporary test motherboard (910-321) incorporating into the device, WinXPE and Analyzer software re-installing |  | Defect | SW | DMS_PC | Windows
AQT90 FLEX | cleaned the cup sensor , now working fine | cleaned the cup sensor , now working fine | SWO-59754 | No | cleaned the cup sensor , now working fine | cleaned the cup sensor , now working fine | Cleaned | HW | Incubator Module | Cup Sensor
AQT90 FLEX |  | B | CFB47B39824AD0A7 | No |  |  | Defect | HW | Process Module | Optical Unit | CRP LQC
ABL80 FLEX BASIC | 更换测试卡正常。 | 更换测试卡正常。 | SWO-60280 | No | The replacement test card is OK. | The replacement test card is OK.
ABL90 FLEX | 更换新测试卡修复仪器 | 更换新测试卡修复仪器 | SWO-59851 | No | Replace the new test card repair instrument | Replace the new test card repair instrument | Assisted Replacement | SC | pCO2 | Cal
ABL825 FLEX | Customer tried troubleshooting over the phone. Found connector seemed to be at fault. | I replaced connector on PH electrode. I fixed the connector on the reference electrode. Once this was done I was able to get calibrations to settle down and come in to range. | CFB85BDCE43CD436 | No | Customer tried troubleshooting over the phone. Found connector seemed to be at fault. | I replaced connector on PH electrode. I fixed the connector on the reference electrode. Once this was done I was able to get calibrations to settle down and come in to range. | Defect | HW | pH/BG Module | Electrode connector
ABL80 FLEX BASIC | 更换两个试剂包正常。 | 更换两个试剂包正常。 | SWO-59913 | No | Replacing both reagent kits is normal. | Replacing both reagent kits is normal.
ABL80 FLEX BASIC | 更换新测试卡修复仪器 | 更换新测试卡修复仪器 | SWO-60485 | No | Replace the new test card repair instrument | Replace the new test card repair instrument
ABL80 FLEX BASIC | 更换测试卡正常 | 更换测试卡正常 | SWO-59992 | No | Replace the test card ok | Replace the test card ok
ABL80 FLEX BASIC | 更换两张新测试卡修复仪器 | 更换两张新测试卡修复仪器 | SWO-59982 | No | Replace two new test card repair instruments | Replace two new test card repair instruments
ABL80 FLEX BASIC | 更换测试卡正常。 | 更换测试卡正常。 | SWO-59915 | No | The replacement test card is OK. | The replacement test card is OK.
ABL80 FLEX BASIC |  | B | SWO-59958 | No
ABL80 FLEX BASIC | 更换测试卡正常。 | 更换测试卡正常。 | SWO-59961 | No | The replacement test card is OK. | The replacement test card is OK.
ABL80 FLEX BASIC | 更换测试卡正常。 | 更换测试卡正常。 | SWO-59960 | No | The replacement test card is OK. | The replacement test card is OK.
ABL90 |  | B | CFB99C065A514194 | No |  |  | Assisted Replacement | SC | Wetability
ABL90 FLEX | Sensor Kassette erneuert Analysator kalibriert und überprüft QC Messung durchgeführt | Sensor Kassette erneuert Analysator kalibriert und überprüft QC Messung durchgeführt | SWO-60144 | No | Sensor cassette renewed analyzer calibrated and checked QC measurement performed | Sensor cassette renewed analyzer calibrated and checked QC measurement performed | No fault found | pCO2 | QC C
ABL90 FLEX | Remplacé le Main Board CPU 945 (défecteux:garantie) e.o Installation d'un nouveau type CPU 945 910-342P1507N0849 e.o Remplacé la sensor de mesure (demande de garantie par le client) e.o Calibration e.o contrôle de qualité e.o | Remplacé le Main Board CPU 945 (défecteux:garantie) e.o Installation d'un nouveau type CPU 945 910-342P1507N0849 e.o Remplacé la sensor de mesure (demande de garantie par le client) e.o Calibration e.o contrôle de qualité e.o | SWO-60122 | No | Replaced the Main Board CPU 945 (defective:guarantee) e.o Installation of a new type CPU 945 910-342P1507N0849 e.o Replaced the sensor measurement (customer warranty request) e.o Calibration e.o quality control e.o | Replaced the Main Board CPU 945 (defective:guarantee) e.o Installation of a new type CPU 945 910-342P1507N0849 e.o Replaced the sensor measurement (customer warranty request) e.o Calibration e.o quality control e.o | Defect | HW | CPU
ABL90 FLEX | Kunde hat aus den Boxen ein paar aussortiert, diese abgeholt und via Rückmeldung technischer Probelem an TS geschickt, thb Kal durchgeführt | Kunde hat aus den Boxen ein paar aussortiert, diese abgeholt und via Rückmeldung technischer Probelem an TS geschickt, thb Kal durchgeführt | SWO-60166 | No | Customer has sorted out a few from the boxes, picked them up and sent them to TS via feedback from technical sampler, thb Kal performed | Customer has sorted out a few from the boxes, picked them up and sent them to TS via feedback from technical sampler, thb Kal performed | Defect | HW | Clot catcher
ABL90 FLEX | Script aufgespielt und gestartet | Script aufgespielt und gestartet | SWO-60223 | Yes | Script played and started | Script played and started | Re classified RMED | Hospital network
ABL90 FLEX | Bildschirm immer noch lila gefärbt(vor allem bei Animationen Eingabefeldern), CPU Unit LOT 342P1516N1078erneuert gegen LOT 342P1504N0605->nach Neustart ist die Bildschirmdarstellung wieder ok System-und Funktionskontrolle i.O. | Bildschirm immer noch lila gefärbt(vor allem bei Animationen Eingabefeldern), CPU Unit LOT 342P1516N1078erneuert gegen LOT 342P1504N0605->nach Neustart ist die Bildschirmdarstellung wieder ok System-und Funktionskontrolle i.O. | SWO-60038 | No | Screen still colored purple (especially for animations input fields), CPU Unit LOT 342P1516N1078renewed against LOT 342P1504N0605->after restart, the screen display is again ok system and function control i.O. | Screen still colored purple (especially for animations input fields), CPU Unit LOT 342P1516N1078renewed against LOT 342P1504N0605->after restart, the screen display is again ok system and function control i.O. | Defect | HW | CPU
ABL90 FLEX | 更换新测试卡修复仪器 | 更换新测试卡修复仪器 | SWO-60194 | No | Replace the new test card repair instrument | Replace the new test card repair instrument | Assisted Replacement | SC | Reference | 1361
AQT90 FLEX |  | B | SWO-59937 | No |  |  | Defect | HW | Process Module | Optical Unit | 124
AQT90 FLEX |  | B | SWO-59938 | No |  |  | Defect | HW | Storage Module | Barcode reader, Storage wheel | BR1
ABL837 FLEX | Replaced the Oxi module for standby purpose, reinstalled the software calibrated instrument performed the QC both results are with in the range. | Replaced the Oxi module for standby purpose, reinstalled the software calibrated instrument performed the QC both results are with in the range. | SWO-60176 | No | Replaced the Oxi module for standby purpose, reinstalled the software calibrated instrument performed the QC both results are with in the range. | Replaced the Oxi module for standby purpose, reinstalled the software calibrated instrument performed the QC both results are with in the range. | Defect | HW | Oxi module | tHB
ABL835 FLEX | Skiftet pH fodpakning Skiftet pco2 membran Kalibrering og QC | Skiftet pH fodpakning Skiftet pco2 membran Kalibrering og QC | SWO-60053 | No | Switched pH Foot gasket switched PCO2 membrane calibration and QC | Switched pH Foot gasket switched PCO2 membrane calibration and QC | Assisted Replacement | Electrode | Membrane | pCO2
ABL825 FLEX |  | B | SWO-59995 | No |  |  | Defect | HW | Oxi module | PCB | 314
ABL825 FLEX | Driftgrenzen erweitert, Meßkammer gesäubert, Ref. Elektrode membraniert, E777 (RV02) ausgetauscht | Dichtungsring pH Elektrode ersetzt, pH BG Modul zurückgetauscht 902-823R2242N003 eingebaut, 823R2278N002 zurück im Bestand, Software aktualisiert | SWO-60226 | No | Drift limits extended, measuring chamber cleaned, ref. Electrode membrane, E777 (RV02) replaced | Sealing ring pH electrode replaced, pH BG module replaced 902-823R2242N003 built-in, 823R2278N002 back in stock, software updated | Assisted Replacement | Electrode | Electrode | pH | Sealing ring
ABL825 FLEX | Sript aufgespielt und ausgeführt | B | SWO-60220 | No | Sript played and executed |  | Setup | HIS/LIS
ABL825 FLEX | Sript aufgespielt und ausgeführt | B | SWO-60225 | No | Sript played and executed |  | Setup | HIS/LIS
ABL825 FLEX | Airventile überarbeitet, zus. Mit HJW Kalibration, QC- Kontrolle | Airventile überarbeitet, zus. Mit HJW Kalibration, QC- Kontrolle | SWO-60168 | No | Air valves revised, too. With HJW calibration, QC control | Air valves revised, too. With HJW calibration, QC control | Defect | HW | pH/BG Module | Valve, air
ABL825 FLEX | Fluidic Board aus Rücknahmegerät vorübergehend eingesetzt. Kundenmodul zur Reparatur nach Willich. | Fluidic Board aus Rücknahmegerät vorübergehend eingesetzt. Kundenmodul zur Reparatur nach Willich. | SWO-60231 | No | Fluidic board from take-back device temporarily inserted. Customer module for repair according to Willich. | Fluidic board from take-back device temporarily inserted. Customer module for repair according to Willich. | Defect | HW | Fluidic manifold | Stepper motor
ABL825 FLEX | Elektroden ok.Kal1Lös. RT 01 scheint fehlerhaft..ersetzt.durch RV 01 Bestand GS. Kal und QC Werte okay | Elektroden ok.Kal1Lös. RT 01 scheint fehlerhaft..ersetzt.durch RV 01 Bestand GS. Kal und QC Werte okay | SWO-60059 | No | Electrodes ok. Kal1Lös. RT 01 seems flawed. replaced by RV 01 stock GS. Kal and QC values okay | Electrodes ok. Kal1Lös. RT 01 seems flawed. replaced by RV 01 stock GS. Kal and QC values okay | Assisted Replacement | Solution bottle | - | Cal 1 | RT1
AQT90 FLEX |  | B | CFB82544E4180AEF | No |  |  | Assisted Replacement | Consumables | Test Kit | TNI
AQT90 FLEX | CLEANED THE SENSOR | CLEANED THE SENSOR | SWO-61186 | No | CLEANED THE SENSOR | CLEANED THE SENSOR | Not Cat | Not Cat | Not Cat | Not Cat
AQT90 FLEX | System geöffnet und Nadel ersetzt, justiert und Tests durchgeführt, mehrere Puffermessungen - i.O, QC | System geöffnet und Nadel ersetzt, justiert und Tests durchgeführt, mehrere Puffermessungen - i.O, QC | SWO-60261 | No | System opened and needle replaced, adjusted and tests performed, multiple buffer measurements - i.O, QC | System opened and needle replaced, adjusted and tests performed, multiple buffer measurements - i.O, QC | Defect | HW | Aspiration Module | Bend needle | 4794
ABL80 FLEX BASIC | 尝试更换试剂包无效，更换新测试卡后修复 | 尝试更换试剂包无效，更换新测试卡后修复 | SWO-60716 | No | Trying to replace the reagent kit is not valid and is repaired after replacing a new test card | Trying to replace the reagent kit is not valid and is repaired after replacing a new test card
ABL80 FLEX BASIC | 更换测试卡正常。 | 更换测试卡正常。 | SWO-60302 | No | The replacement test card is OK. | The replacement test card is OK.
ABL80 CO-OX | Customer has had to change SC SP to correct issues with pCO2 and pO2 at different times. Customer is following the troubleshooting guidance sheet. Customer does not flush SC when errors start. I suggested to flush SC then reinitialize to see if problem is corrected before changing consumables. | Performed zero check and all passed. | CFB99D465A163A2C | No | Customer has had to change SC SP to correct issues with pCO2 and pO2 at different times. Customer is following the troubleshooting guidance sheet. Customer does not flush SC when errors start. I suggested to flush SC then reinitialize to see if problem is corrected before changing consumables. | Performed zero check and all passed.
ABL80 FLEX BASIC | 更换新测试卡修复仪器 | 更换新测试卡修复仪器 | SWO-60482 | No | Replace the new test card repair instrument | Replace the new test card repair instrument
ABL80 FLEX BASIC | 更换新测试卡修复仪器 | 更换新测试卡修复仪器 | SWO-60483 | No | Replace the new test card repair instrument | Replace the new test card repair instrument
ABL80 FLEX BASIC | 更换三张测试卡，试剂包正常。 | 更换三张测试卡，试剂包正常。 | SWO-60286 | No | Replace three test cards and the reagent package is OK. | Replace three test cards and the reagent package is OK.
ABL80 FLEX CO-OX | ABL neu gestartet, tHb Kalibration, Testmessung. | ABL neu gestartet, tHb Kalibration, Testmessung. | SWO-60422 | No | ABL restarted, tHb calibration, test measurement. | ABL restarted, tHb calibration, test measurement.
ABL80 CO-OX | Upon arrival, customer had replaced sensor cassette and solution pack, running out of their last stock. Apparently consumables failed out of box. | Replaced sensor cassette and solution pack for troubleshooting purposes. Verified operation of ABL 80 COOX. Handed solution pack and sensor cassette replacement to customer. | CFB9CE8C5A281A83 | No | Upon arrival, customer had replaced sensor cassette and solution pack, running out of their last stock. Apparently consumables failed out of box. | Replaced sensor cassette and solution pack for troubleshooting purposes. Verified operation of ABL 80 COOX. Handed solution pack and sensor cassette replacement to customer.
ABL80 OSM | Confirmed the problem. Coppied the setup from a ABL80 OSM that is working properly. | Copied the setup. | CFB995635A51CB1F | No | Confirmed the problem. Coppied the setup from a ABL80 OSM that is working properly. | Copied the setup.
ABL90 |  | B | CFC61B71F6316569 | No
ABL90 |  | B | CFC9CF34581DA449 | No
ABL90 |  | B | CFB99AD85A1F4191 | No |  |  | Assisted Replacement | SC | Wetability | MR6 installed
ABL90 | While attempting to duplicate the malfunction, I found out that the customer was trying to leave the entry blank, when it requires a number or a decimal point. When the min or max value was left blank, it induced the program crash. | This is normal operation, no defect noted. 0.0 was used for the PRESC min and max which meets the customers requirements. | CFB9966D5A17ABA9 | No | While attempting to duplicate the malfunction, I found out that the customer was trying to leave the entry blank, when it requires a number or a decimal point. When the min or max value was left blank, it induced the program crash. | This is normal operation, no defect noted. 0.0 was used for the PRESC min and max which meets the customers requirements. | Defect | SW | crash
ABL90 FLEX | Sript aufgespielt und ausgeführt | B | SWO-60504 | Yes | Sript played and executed |  | Re classified RMED | Hospital network
ABL90 FLEX | Sript aufgespielt und ausgeführt | B | SWO-60502 | Yes | Sript played and executed |  | Re classified RMED | Hospital network
ABL90 |  | B | CFC6914AF60898A1 | No |  |  | Assisted Replacement | SC | pCO2
ABL90 |  | B | CFB97F485A35787E | No |  |  | No fault found | HIS/LIS
ABL90 | When running QC level 3 received inhomogeneous rinse solution(LS sensor) and liquid transport error during rinse. | Back flush the analyzer, performe PM,Replaced sensor cassette bay unit and installed new SC and SP. Ran QC level 3 several times no error occured. | CFB9D4055A5A1D8A | No | When running QC level 3 received inhomogeneous rinse solution(LS sensor) and liquid transport error during rinse. | Back flush the analyzer, performe PM,Replaced sensor cassette bay unit and installed new SC and SP. Ran QC level 3 several times no error occured. | Defect | HW | Sensor Bay Unit
ABL90 | Customer replaced the Sensor Cassette, but analyzer failed a different level Po2. Customer replaced the SC a second time and cals were iIN OUT on the PO2 Status between -.6 and 35(Not Stable). | Replaced the Sensor Cassette interface PCB. Recalibrated the Thermistors. If problem persist with the Po2 Status the SC should be replaced. | CFB9E8275A3D37E1 | No | Customer replaced the Sensor Cassette, but analyzer failed a different level Po2. Customer replaced the SC a second time and cals were iIN OUT on the PO2 Status between -.6 and 35(Not Stable). | Replaced the Sensor Cassette interface PCB. Recalibrated the Thermistors. If problem persist with the Po2 Status the SC should be replaced. | Defect | HW | SC PCB
ABL90 |  | B | CFBAB14AD034D727 | Yes |  |  | Re classified RMED | Demo, no info
ABL90 FLEX | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-60497 | No | Repair after replacing a new test card | Repair after replacing a new test card | Assisted Replacement | SC | Reference | 1361
ABL90 FLEX | 更换两张测试卡修复仪器 | 更换两张测试卡修复仪器 | SWO-60484 | No | Replace two test card repair instruments | Replace two test card repair instruments | Assisted Replacement | SC | Lac | Cal
ABL90 FLEX | 更换新测试卡修复仪器 | 更换新测试卡修复仪器 | SWO-60489 | No | Replace the new test card repair instrument | Replace the new test card repair instrument | Assisted Replacement | SC | Lac | Cal
AQT90 FLEX |  | B | SWO-60227 | No |  |  | Defect | HW | Storage Module | Barcode reader, Storage wheel
AQT90 FLEX | 更换进样针903-171一个后修复 | 更换进样针903-171一个后修复 | SWO-60490 | No | Replace the sample needle 903-171 after repair | Replace the sample needle 903-171 after repair | Defect | HW | Aspiration Module | Bend needle | 4794
AQT90 FLEX |  | B | SWO-60300 | No |  |  | Defect | HW | Process Module | Optical Unit | 131
ABL837 FLEX | Cleaned the electrode but problem is remaining same, replaced the cal1, cal2, cleaning solutions, Replaced Crea membrane calibrated the instrument found results are with in the range. | Cleaned the electrode but problem is remaining same, replaced the cal1, cal2, cleaning solutions, Replaced Crea membrane calibrated the instrument found results are with in the range. | SWO-61239 | No | Cleaned the electrode but problem is remaining same, replaced the cal1, cal2, cleaning solutions, Replaced Crea membrane calibrated the instrument found results are with in the range. | Cleaned the electrode but problem is remaining same, replaced the cal1, cal2, cleaning solutions, Replaced Crea membrane calibrated the instrument found results are with in the range. | Assisted Replacement | Electrode | Membrane | Crea
ABL815 FLEX | AC Modul ausgebaut, Funktion im Serviceprogramm überprüft, OK 14 poliges Kabel erneuert, Test QC OK. | B | SWO-60433 | No | AC module removed, function checked in service program, OK 14 pin cable renewed, test QC OK. |  | Defect | HW | AutoCheck Module | FPC Cable
AQT90 FLEX |  | B | CFB9917C5A256ED6 | No |  |  | Defect | HW | Process Module | Optical Unit | 135
ABL80 FLEX BASIC | 更换新测试卡修复仪器 | 更换新测试卡修复仪器 | SWO-60798 | No | Replace the new test card repair instrument | Replace the new test card repair instrument
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-60501 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL80 FLEX BASIC | Replaced the signal data board & installed the sensor cassette checked equipment is functional | Replaced the signal data board & installed the sensor cassette checked equipment is functional | SWO-61162 | No | Replaced the signal data board & installed the sensor cassette checked equipment is functional | Replaced the signal data board & installed the sensor cassette checked equipment is functional
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-60554 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL90 FLEX | Uhrzeit auf Windowsebene eingestellt und gespeichert | Uhrzeit auf Windowsebene eingestellt und gespeichert | SWO-60592 | No | Time set and saved at the Windows level | Time set and saved at the Windows level | Defect | SW | Time setting
ABL90 FLEX | Gerät gespült, SP erneuert und Jahreswartung durchgeführt Einstellungen für Uhrzeit überprüft und korrigiert | Gerät gespült, SP erneuert und Jahreswartung durchgeführt Einstellungen für Uhrzeit überprüft und korrigiert | SWO-60562 | No | Device flushed, SP renewed and year maintenance performed settings for time checked and corrected | Device flushed, SP renewed and year maintenance performed settings for time checked and corrected | Cleaned | Fluidic pathway
ABL90 | Tried many different things to fix. I was able to determine that CF card was not corrupt and Computer board was bad. | Replaced CPU. Loaded customer CF card and calibrated analyzer. Reconnected analyzer to customer network and made sure everything was working properly. | CFBAE2BFD0519FA0 | No | Tried many different things to fix. I was able to determine that CF card was not corrupt and Computer board was bad. | Replaced CPU. Loaded customer CF card and calibrated analyzer. Reconnected analyzer to customer network and made sure everything was working properly. | Defect | HW | CPU
ABL90 FLEX | SC und LP aus Kundenbestand ersetzt, Funktion OK. Rückruf mit Dr. Lysko besprochen. | SC und LP aus Kundenbestand ersetzt, Funktion OK. Rückruf mit Dr. Lysko besprochen. | SWO-60742 | No | SC and LP replaced from customer inventory, function OK. Recall discussed with Dr. Lysko. | SC and LP replaced from customer inventory, function OK. Recall discussed with Dr. Lysko. | Performance | BE to high
ABL90 | After reboot the analyzer will come up like normal, but the error will occur immediately there after. | Reloaded software version MR6, but this did not work. Deleted all roma programs, then reloaded MR6, and this resolved the problem. When I reloaded the customer settings, the analyzer failed again. Manually loaded customer settings, | CFBAF0C3D01A1120 | No | After reboot the analyzer will come up like normal, but the error will occur immediately there after. | Reloaded software version MR6, but this did not work. Deleted all roma programs, then reloaded MR6, and this resolved the problem. When I reloaded the customer settings, the analyzer failed again. Manually loaded customer settings, | Defect | SW | Sybase 16
TCM4 |  | B | SWO-60521 | No
ABL90 FLEX |  | B | SWO-60535 | No |  |  | Defect | HW | CPU
AQT90 FLEX | 需要更換光學單元 | B | SWO-60493 | No | Optical unit needs to be replaced |  | Defect | HW | Process Module | Optical Unit | 124
AQT90 FLEX |  | 更换903-364一个正常。 | SWO-60476 | No |  | Replace 903-364 one normal. | Defect | HW | Process Module | Optical Unit | 131
ABL825 FLEX | LOOKED IN ACTIVITIY LOGS AND FOUND ERROR CODES 316, 258, NEON LAMP INTENSITY TOO LOW. HAPPEN 6 TIMES IN LAST 2 DAYS. | REPLACED LAMP. ADJUSTED LAMP, DID COMPLETE OXIMETRY TEST. DID THB IN SERVICE, IN USER, AND AS SAMPLE. | CFBAECC3D0062457 | No | LOOKED IN ACTIVITIY LOGS AND FOUND ERROR CODES 316, 258; NEON LAMP INTENSITY TOO LOW. HAPPEN 6 TIMES IN LAST 2 DAYS. | REPLACED LAMP. ADJUSTED LAMP, DID COMPLETE OXIMETRY TEST. DID THB IN SERVICE, IN USER, AND AS SAMPLE. | Defect | HW | Oxi module | Lamp unit | 316, 258
ABL825 FLEX | Upon arrival, found the analyzer system messages presenting errors 316 Neon number too low and, 258 Neon lamp intensity too low. | Replaced all listed parts. Adjusted the hemolyzer lamp. Performed the Complete oximetry test, passing within specifications. Performed the service and, customer mode ctHb calibrations, passing within specifications. | CFBAB294D012925B | No | Upon arrival, found the analyzer system messages presenting errors 316 Neon number too low and, 258 Neon lamp intensity too low. | Replaced all listed parts. Adjusted the hemolyzer lamp. Performed the Complete oximetry test, passing within specifications. Performed the service and, customer mode ctHb calibrations, passing within specifications. | Defect | HW | Oxi module | Lamp unit | 316, 258
ABL825 FLEX | Optionkey installiert, jedoch ist der Blood Gas Key nicht mit dabei gewesen. Mit Rolf Lang abgesprochen, er besorgt den Key bis morgen. | Key per Mail an das Labor gesendet und via telefonischer Unterstützung mit einer fähigen Mitarbeiterin des Labors den Blood Gas Key aufgespielt. | SWO-60881 | Yes | Optionkey installed, but the Blood Gas Key was not included. Agreed with Rolf Lang, he will take care of the key until tomorrow. | Key sent by e-mail to the laboratory and played the Blood Gas Key via telephone support with a capable employee of the laboratory. | Re classified RMED | Telephone support
ABL827 FLEX | Reinigungsarbeiten, Fehlersuche, Bildschirm demontiert und trockengelegt, Im unteren Bereich des Bildschirms Geruch von verschmortem Kunststoff, Bildschirm muss ersetzt werden, Lieferung und Montage des neuen Bildschirms am 22.10.2015 Gerät für Anwender gesperrt !!! Neuer Monitor wurde am 22.10.2015 geliefert und montiert ! Mailto: Reiner.schmidt@kliniken-service.de | Reinigungsarbeiten, Fehlersuche, Bildschirm demontiert und trockengelegt, Im unteren Bereich des Bildschirms Geruch von verschmortem Kunststoff, Bildschirm muss ersetzt werden, Lieferung und Montage des neuen Bildschirms am 22.10.2015 Gerät für Anwender gesperrt !!! Neuer Monitor wurde am 22.10.2015 geliefert und montiert ! Mailto: Reiner.schmidt@kliniken-service.de | SWO-60821 | No | Cleaning, troubleshooting, screen dismantled and drained, In the lower part of the screen smell of charred plastic, screen must be replaced, delivery and assembly of the new screen on 22.10.2015 device locked for users !!! New monitor was delivered and mounted on 22.10.2015 ! Mailto: Reiner.schmidt@kliniken-service.de | Cleaning, troubleshooting, screen dismantled and drained, In the lower part of the screen smell of charred plastic, screen must be replaced, delivery and assembly of the new screen on 22.10.2015 device locked for users !!! New monitor was delivered and mounted on 22.10.2015 ! Mailto: Reiner.schmidt@kliniken-service.de | Defect | HW | DMS_PC | Touch Screen
ABL825 FLEX | Überprüfung aller Messmodulenventile Spülung Airkanal Reinigung der Messkammer bei pH, Ref und Elektrolyten Austausch Ref Elektrode Durchführung Proteinentfernung Mehrmals Kal bis i.O. Durchführung alle QC Level | Überprüfung aller Messmodulenventile Spülung Airkanal Reinigung der Messkammer bei pH, Ref und Elektrolyten Austausch Ref Elektrode Durchführung Proteinentfernung Mehrmals Kal bis i.O. Durchführung alle QC Level | SWO-60760 | No | Inspection of all measuring module valves flushing air duct cleaning of the measuring chamber at pH, ref and electrolytes replacement ref electrode carrying out protein removal Several times Kal to i.O. Execution all QC level | Inspection of all measuring module valves flushing air duct cleaning of the measuring chamber at pH, ref and electrolytes replacement ref electrode carrying out protein removal Several times Kal to i.O. Execution all QC level | Cleaned | Fluidic pathway | -
ABL825 FLEX | Überprüfung und Reinigung aller Ventile in den Messmodulen Reinigung der Messkammer bei pH, Ref und Elektrolyten Austausch Ref Elektrode Durchführung Proteinentfernung Mehrmals Kal bis i.O. Durchführung alle QC Level => pH und Elektrolyte nahe zu überall auf Zielwert | Überprüfung und Reinigung aller Ventile in den Messmodulen Reinigung der Messkammer bei pH, Ref und Elektrolyten Austausch Ref Elektrode Durchführung Proteinentfernung Mehrmals Kal bis i.O. Durchführung alle QC Level => pH und Elektrolyte nahe zu überall auf Zielwert | SWO-60757 | No | Checking and cleaning of all valves in the measuring modules Cleaning the measuring chamber at pH, ref and electrolytes exchange ref electrode performing protein removal Several times Kal to i.O. Performing all QC levels => pH and electrolytes close to anywhere at target value | Checking and cleaning of all valves in the measuring modules Cleaning the measuring chamber at pH, ref and electrolytes exchange ref electrode performing protein removal Several times Kal to i.O. Performing all QC levels => pH and electrolytes close to anywhere at target value | Assisted Replacement | Electrode | Electrode | Reference
ABL825 FLEX | Austausch Display Unit im Rahmen der Garantie | Austausch Display Unit im Rahmen der Garantie | SWO-60761 | No | Replacement Display Unit as part of the warranty | Replacement Display Unit as part of the warranty | Defect | HW | DMS_PC | Touch Screen
ABL835 FLEX | Versuch die Software komplett neu zu installieren scheitert neue Festplatte eingebaut, Software installiert und Gerät konfiguriert Gerät kalibriert und überprüft | Versuch die Software komplett neu zu installieren scheitert neue Festplatte eingebaut, Software installiert und Gerät konfiguriert Gerät kalibriert und überprüft | SWO-60573 | No | Attempt to completely reinstall the software fails new hard drive built-in, software installed and device configured device calibrated and checked | Attempt to completely reinstall the software fails new hard drive built-in, software installed and device configured device calibrated and checked | Defect | HW | DMS_PC | HDD
ABL825 FLEX | Sript aufgespielt und ausgeführt | Sript aufgespielt und ausgeführt | SWO-60843 | No | Sript played and executed | Sript played and executed | Setup | HIS/LIS
ABL827 FLEX | Sript aufgespielt und ausgeführt | Sript aufgespielt und ausgeführt | SWO-60841 | No | Sript played and executed | Sript played and executed | Setup | HIS/LIS
ABL827 FLEX | Crea neu membraniert und Lösungen erneuert | Crea neu membraniert und Lösungen erneuert | SWO-60840 | No | Crea re-membraneed and solutions renewed | Crea re-membraneed and solutions renewed | Assisted Replacement | Electrode | Membrane | Crea
ABL825 FLEX | PH Elektrode ersetzt da mechanisch zerstört, Kalibration, QC Kontrolle | PH Elektrode ersetzt da mechanisch zerstört, Kalibration, QC Kontrolle | SWO-60643 | No | PH electrode replaces since mechanically destroyed, calibration, QC control | PH electrode replaces since mechanically destroyed, calibration, QC control | Assisted Replacement | Electrode | Electrode | pH
ABL800 BASIC | pCO2 Elektrode neu membraniert. Kontakte der Elektrode und der Elektrodenkappe gereinigt. Kalibrationen und QC's vorgenommen. | pCO2 Elektrode neu membraniert. Kontakte der Elektrode und der Elektrodenkappe gereinigt. Kalibrationen und QC's vorgenommen. | SWO-60632 | No | pCO2 electrode. contacts of the electrode and the electrode cap. calibrations and QC's. | pCO2 electrode. contacts of the electrode and the electrode cap. calibrations and QC's. | Assisted Replacement | Electrode | Membrane | pCO2
ABL800 BASIC | pCO2 Elektrode mit andrem Lot neu membraniert. Kontakte der Elektrode und der Elektrodenkappe gereinigt. Kalibrationen und QC's vorgenommen. | pCO2 Elektrode mit andrem Lot neu membraniert. Kontakte der Elektrode und der Elektrodenkappe gereinigt. Kalibrationen und QC's vorgenommen. | SWO-60635 | No | pCO2 electrode with other solder re-membrane. contacts of the electrode and the electrode cap. calibrations and QC's. | pCO2 electrode with other solder re-membrane. contacts of the electrode and the electrode cap. calibrations and QC's. | Assisted Replacement | Electrode | Membrane | pCO2
ABL800 BASIC | Unteren Dichtring der pH und Dummyelektrode erneuert. Kontakte der pCO2 Elektrode und der Kontaktkappe gereinigt. Kalibrationen und QC's vorgenommen. Kunde wird Gerät beobachten. | Unteren Dichtring der pH und Dummyelektrode erneuert. Kontakte der pCO2 Elektrode und der Kontaktkappe gereinigt. Kalibrationen und QC's vorgenommen. Kunde wird Gerät beobachten. | SWO-60636 | No | Lower sealing ring of pH and dummy electrode renewed. contacts of the pCO2 electrode and the contact cap. calibrations and QC's. Customer will watch device. | Lower sealing ring of pH and dummy electrode renewed. contacts of the pCO2 electrode and the contact cap. calibrations and QC's. Customer will watch device. | Assisted Replacement | Electrode | Electrode | pH | Sealing ring
ABL835 FLEX |  | B | CFBADA25D019D79D | No |  |  | Defect | HW | Oxi module | Hemolyzer
AQT90 FLEX |  | B | CFBAF1DCD0274D1D | No |  |  | Performance | HW | Process Module | Process wheel | CRP LQC
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-61738 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL80 FLEX BASIC | 更换前面板924-738一个 | 更换前面板924-738一个 | SWO-60744 | No | Replace the front panel 924-738 one | Replace the front panel 924-738 one
ABL80 FLEX BASIC | 更换新测试卡 | 更换新测试卡 | SWO-63069 | No | Replace a new test card | Replace a new test card
ABL80 FLEX BASIC | 更换测试卡正常。 | 更换测试卡正常。 | SWO-60992 | No | The replacement test card is OK. | The replacement test card is OK.
ABL80 |  | B | CFBC0D0A465B5D5C | No
ABL80 FLEX BASIC | 更换被腐蚀的接口板和定标未通过的测试卡后修复 | 更换被腐蚀的接口板和定标未通过的测试卡后修复 | SWO-60734 | No | Repair after replacing corroded interface plates and failed test cards for calibration | Repair after replacing corroded interface plates and failed test cards for calibration
ABL80 FLEX CO-OX | Gerät getauscht gegen ABL 80 Coox 306894 ( geliefert von Zentrale ) Kundenkonfiguration durchgeführt. Glu . Kal. Fehler 1225. Keine weitere SC vorhanden, Bestellung konnte nicht rechtzeitig geliefert werden. | Gerät getauscht gegen ABL 80 Coox 306894 ( geliefert von Zentrale ) Kundenkonfiguration durchgeführt. Glu . Kal. Fehler 1225. Keine weitere SC vorhanden, Bestellung konnte nicht rechtzeitig geliefert werden. | SWO-60960 | No | Device exchanged for ABL 80 Coox 306894 (supplied by Central ) customer configuration carried out. Glu. Kal. Error 1225. No further SC available, order could not be delivered on time. | Device exchanged for ABL 80 Coox 306894 (supplied by Central ) customer configuration carried out. Glu. Kal. Error 1225. No further SC available, order could not be delivered on time.
ABL90 FLEX | SW Update auf V 3.1MR6 durchgeführt, Gerät komplett neu konfiguriert. Testmessung iO. | SW Update auf V 3.1MR6 durchgeführt, Gerät komplett neu konfiguriert. Testmessung iO. | SWO-60919 | No | SW update performed on V 3.1MR6, device completely reconfigured. Test measurement iO. | SW update performed on V 3.1MR6, device completely reconfigured. Test measurement iO. | Unknown | Pump cal err | Off/On
ABL90 FLEX | Gerät gereinigt und überprüft, ABL 90 Jahreswartungskit komplett verbaut, Wartungsmaßnahmen durchgeführt, LP erneuert, tHb Kal, Kals und QCs durchgeführt | Gerät gereinigt und überprüft, ABL 90 Jahreswartungskit komplett verbaut, Wartungsmaßnahmen durchgeführt, LP erneuert, tHb Kal, Kals und QCs durchgeführt | SWO-60955 | Yes | Device cleaned and checked, ABL 90 annual maintenance kit fully installed, maintenance carried out, LP renewed, tHb Kal, Kals and QCs carried out | Device cleaned and checked, ABL 90 annual maintenance kit fully installed, maintenance carried out, LP renewed, tHb Kal, Kals and QCs carried out | Re classified RMED | Yearly maintenance
ABL90 |  | B | CFBC2369461567AF | Yes |  |  | Re classified RMED | Telephone support
ABL90 | Checked all logs to verify past performance. Analyzer was operating fine until staff replaced the SC. GLU and LAC failing internal QC | Staff replaced the sensor cass. and the internal QC is still failing (HIGH) on GLU and LAC. I had staff to Replaced the Solution Pack - cals and qc passed Instructed customer to send for credit on consumable | CFBC31E7462D6FED | Yes | Checked all logs to verify past performance. Analyzer was operating fine until staff replaced the SC. GLU and LAC failing internal QC | Staff replaced the sensor cass. and the internal QC is still failing (HIGH) on GLU and LAC. I had staff to Replaced the Solution Pack - cals and qc passed Instructed customer to send for credit on consumable | Re classified RMED | PER
ABL90 | Troubleshot and found that activating Radiance in settings is causing the problem. | I deselected Radiance in the settings menu, and this resolved the problem. During QC the internal QC failed for pH, Na, K so I replaced the solution pack (Provided by customer) and this resolved the problem. External QC failed for pO2, so I replaced the Sensor Cassette (Provided by customer) and this resolved the problem. | CFBD6DD1BC500C61 | No | Troubleshot and found that activating Radiance in settings is causing the problem. | I deselected Radiance in the settings menu, and this resolved the problem. During QC the internal QC failed for pH, Na, K so I replaced the solution pack (Provided by customer) and this resolved the problem. External QC failed for pO2, so I replaced the Sensor Cassette (Provided by customer) and this resolved the problem. | Defect | SW | Sybase 16
ABL90 FLEX | 更换试剂包正常。 | 更换试剂包正常。 | SWO-61273 | No | The replacement kit is OK. | The replacement kit is OK. | Assisted Replacement | SP | FT |  | 1343
ABL90 FLEX | インレットガスケット、コネクトガスケット、インレットプローブ交換。SC,SP交換。流路洗浄実施。 | インレットガスケット、コネクトガスケット、インレットプローブ交換。SC,SP交換。流路洗浄実施。 | SWO-61602 | No | Inlet gaskets, connect gaskets, inlet probe replacement. SC, SP replacement. Flow path cleaning was carried out. | Inlet gaskets, connect gaskets, inlet probe replacement. SC, SP replacement. Flow path cleaning was carried out. | Assisted Replacement | SC | Wetability
ABL90 | Reinstalação do cartucho, verificação dos regsitros, estabilização e passagem dos QCs. | B | CFBC097746115ECE | No | Reinstallation of the cartridge, verification of the Regsitros, stabilization and passage of the QCs. |  | User error | SP install
ABL825 FLEX |  | B | CFBD4C4EBC0B62D1 | Yes |  |  | Re classified RMED | Training
ABL815 FLEX |  | B | CFBD0166BC1E7344 | No |  |  | Cleaned | Fluidic pathway | - | pH
ABL825 FLEX | Aufgeführte Ventilmembranen des EL MET Moduls ersetzt, da verschlissen, Nachfüllung, Kal, QC | Aufgeführte Ventilmembranen des EL MET Moduls ersetzt, da verschlissen, Nachfüllung, Kal, QC | SWO-61060 | No | Replaced listed valve membranes of the EL MET module, as worn, refill, Kal, QC | Replaced listed valve membranes of the EL MET module, as worn, refill, Kal, QC | Defect | HW | El/Met Module | Valve | Membrane
ABL800 BASIC | 代替モジュール貸出 | インレットニップル交換 高さ調整 インレットモジュール分解・清掃・グリースアップ 動作確認 | SWO-61303 | No | Alternative Module Lending | Inlet nipple replacement Height adjustment Inlet module disassembly, cleaning, grease-up operation confirmation | Defect | HW | Inlet module | Nipple kit | Nipple kit
AQT90 FLEX | Inlet Wheel, Nadel im Service überprüft - needle cal (only washblock) erneuert Pipette und washblock sind sauber Blankmessung und QC durchgeführt | Inlet Wheel, Nadel im Service überprüft - needle cal (only washblock) erneuert Pipette und washblock sind sauber Blankmessung und QC durchgeführt | SWO-60870 | No | Inlet wheel, needle checked in service - needle cal (only washblock) renewed pipette and washblock are clean blank measurement and QC performed | Inlet wheel, needle checked in service - needle cal (only washblock) renewed pipette and washblock are clean blank measurement and QC performed | Defect | SW | Analyzer control Module | Power off/on | 35
ABL80 FLEX BASIC | 更换测试卡正常。 | 更换测试卡正常。 | SWO-61263 | No | The replacement test card is OK. | The replacement test card is OK.
ABL80 FLEX BASIC | 更换902-972一个后修复 | 更换902-972一个后修复 | SWO-60729 | No | Replace 902-972 one post-fix | Replace 902-972 one post-fix
ABL80 FLEX BASIC | 更换测试卡正常。 | 更换测试卡正常。 | SWO-61286 | No | The replacement test card is OK. | The replacement test card is OK.
ABL80 FLEX BASIC | 更换新测试卡修复仪器 | 更换新测试卡修复仪器 | SWO-61054 | No | Replace the new test card repair instrument | Replace the new test card repair instrument
ABL80 FLEX CO-OX | SC wird von Zentrale ersetzt. Defekte SC geht an Zentrale Willich zur Begutachtung und Kenntnisnahme etc.. | SC wird von Zentrale ersetzt. Defekte SC geht an Zentrale Willich zur Begutachtung und Kenntnisnahme etc.. | SWO-61094 | No | SC is replaced by Headquarters. Defective SC goes to Central Willich for inspection and knowledge, etc. | SC is replaced by Headquarters. Defective SC goes to Central Willich for inspection and knowledge, etc.
ABL90 | CONFIRMED ISSUE WHEN ONSITE. | PROCEED TO UPGRADE SOFTWARE FROM VERSION 3.1 MR3 TO MR6. ALSO INSTALLED NEW SENSOR CASSETTE AND SOLUTION PACK BORROWED FROM FAIREVIEW U OF MN (GRAHAM GREGORICH). REPLACED BOTH INLET GASKET HOLDER AND PROBE AS THEY WERE WORN. CLEANED UP BLOOD SPECKS ON SYSTEM. | CFBD3D42BC0B1E67 | No | CONFIRMED ISSUE WHEN ONSITE. | PROCEED TO UPGRADE SOFTWARE FROM VERSION 3.1 MR3 TO MR6. ALSO INSTALLED NEW SENSOR CASSETTE AND SOLUTION PACK BORROWED FROM FAIREVIEW U OF MN (GRAHAM GREGORICH). REPLACED BOTH INLET GASKET HOLDER AND PROBE AS THEY WERE WORN. CLEANED UP BLOOD SPECKS ON SYSTEM. | Assisted Replacement | SC | Wetability | MR6 installed
ABL90 | Site reports that LIS connectivity is periodically dropped. | Site visit with Steve Powers verified LIS connectivity. Verified instrument performance via trace route, pinging of the soft server and gathering of comm logs. Instrument is within acceptable operating specifications. | CFC2581A94490422 | No | Site reports that LIS connectivity is periodically dropped. | Site visit with Steve Powers verified LIS connectivity. Verified instrument performance via trace route, pinging of the soft server and gathering of comm logs. Instrument is within acceptable operating specifications. | No fault found | HIS/LIS
ABL90 FLEX | Credit claim im Klinikum ausgedruckt und an RDeu gesendet | Credit claim im Klinikum ausgedruckt und an RDeu gesendet | SWO-61156 | Yes | Credit claim printed in the hospital and sent to RDeu | Credit claim printed in the hospital and sent to RDeu | Re classified RMED | PER
ABL90 FLEX | Lösungspack entfernt und neu eingelegt Fluid Transport über Service Menü überprüft SC getauscht Analysator kalibriert und überprüft QC Messung durchgeführt | Lösungspack entfernt und neu eingelegt Fluid Transport über Service Menü überprüft SC getauscht Analysator kalibriert und überprüft QC Messung durchgeführt | SWO-61242 | No | Solution pack removed and re-inserted Fluid Transport checked via service menu replaced SC analyzer calibrated and verified QC measurement performed | Solution pack removed and re-inserted Fluid Transport checked via service menu replaced SC analyzer calibrated and verified QC measurement performed | Assisted Replacement | SC | Wetability
ABL90 FLEX | Anwender hatte bereits SC,SP, Einlassdichtung und -nadel erneuert --> ohne Erfolg Flüssigkeitsstrecke gespült, alten SP inst. --> Fluidictransport im Serviceprogramm i.O --> Neustart | Anwender hatte bereits SC,SP, Einlassdichtung und -nadel erneuert --> ohne Erfolg Flüssigkeitsstrecke gespült, alten SP inst. --> Fluidictransport im Serviceprogramm i.O --> Neustart | SWO-61176 | No | User had already renewed SC,SP, inlet seal and needle --> without success liquid section flushed, old SP inst. --> Fluidic transport in the service program i.O --> Restart | User had already renewed SC,SP, inlet seal and needle --> without success liquid section flushed, old SP inst. --> Fluidic transport in the service program i.O --> Restart | Cleaned | Fluidic pathway
ABL90 FLEX | neuer PC wird bestellt | neuer PC wird bestellt | SWO-61143 | No | new PC is ordered | new PC is ordered | Defect | HW | PC Module
ABL5 | Halbjahreswartung durchgeführt. | Halbjahreswartung durchgeführt. | SWO-61764 | No | half-yearly maintenance. | half-yearly maintenance.
ABL90 FLEX | 更换新测试卡修复仪器 | 更换新测试卡修复仪器 | SWO-61049 | No | Replace the new test card repair instrument | Replace the new test card repair instrument | Assisted Replacement | SC | Reference
ABL90 FLEX | 更换新测试卡修复仪器 | 更换新测试卡修复仪器 | SWO-61069 | No | Replace the new test card repair instrument | Replace the new test card repair instrument | Assisted Replacement | SC | pCO2 | Cal | 13221
AQT90 FLEX | 更换902-978一个正常。 | 更换902-978一个正常。 | SWO-60722 | No | Replace 902-978 one normal. | Replace 902-978 one normal. | Defect | HW | Aspiration Module | Cavro Pump | 1218
AQT90 FLEX | 更换636-550一个后修复 | 更换636-550一个后修复 | SWO-60477 | No | Replace 636-550 one post-fix | Replace 636-550 one post-fix | Defect | HW | Storage Module | Cartridge Sensor | 5784
ABL825 FLEX | Finalen Glu-Key eingespielt | Finalen Glu-Key eingespielt | SWO-61259 | No | Final Glu-Key recorded | Final Glu-Key recorded | Setup | Demo key exp | Demo Key Exp | Demo key expired
ABL825 FLEX |  | B | CFBD85E7BC1630BA | No |  |  | Defect | HW | El/Met Module | Measuring chamber | Lac
AQT90 FLEX |  | B | CFBD2854BC62DC86 | Yes |  |  | Re classified RMED | Remote support
AQT90 FLEX | Scanner ausgebaut, alle Spiegel gereinigt, alle Kabel neu befestigt, Scanner getestet, | Scanner ausgebaut, alle Spiegel gereinigt, alle Kabel neu befestigt, Scanner getestet, | SWO-61443 | No | Scanner removed, all mirrors cleaned, all cables reattached, scanner tested, | Scanner removed, all mirrors cleaned, all cables reattached, scanner tested, | Cleaned | HW | Analyzer control Module | Barcode reader, display
ABL80 FLEX BASIC | 连续更换两张测试卡后修复 | 连续更换两张测试卡后修复 | SWO-61216 | No | Repair after replacing two test cards in a row | Repair after replacing two test cards in a row
ABL90 | The user has replaced the solution pack and inlet tubing and flushed the analyzer multiple times, but still has errors. Found pump test, vacuum test and clot test all failing. No obvious leaks found. | Replaced the waste tubing, pinch valve tubing, connector gasket and inlet gasket and resolved the errors. | CFBED5AA320AAA66 | No | The user has replaced the solution pack and inlet tubing and flushed the analyzer multiple times, but still has errors. Found pump test, vacuum test and clot test all failing. No obvious leaks found. | Replaced the waste tubing, pinch valve tubing, connector gasket and inlet gasket and resolved the errors. | Defect | HW | Tubes | Leaking
ABL80 FLEX BASIC | 更换新测试卡修复仪器 | 更换新测试卡修复仪器 | SWO-61271 | No | Replace the new test card repair instrument | Replace the new test card repair instrument
AQT90 FLEX | 更换903-198一个后修复 | 更换903-198一个后修复 | SWO-61079 | No | Replace 903-198 one post-fix | Replace 903-198 one post-fix | Defect | HW | Analyzer control Module | Main Board, PCB | 2132
AQT90 FLEX |  | B | SWO-61248 | No |  |  | Defect | HW | Storage Module | Barcode reader, Storage wheel | BR1
ABL80 FLEX BASIC | 更换新测试卡修复仪器 | 更换新测试卡修复仪器 | SWO-61274 | No | Replace the new test card repair instrument | Replace the new test card repair instrument
ABL80 FLEX BASIC | SP80 erneuert | Medizintechniker Herr Haak hat div. Male Material erneuert. Er besteht auf einen Vorortservice! | SWO-61446 | No | SP80 renewed | Medical technician Mr. Haak has renewed material several times. He insists on a suburban service!
ABL80 FLEX BASIC | 更换两张测试卡 | 更换两张测试卡 | SWO-63070 | No | Replace two test cards | Replace two test cards
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-61521 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL80 FLEX BASIC | 更换测试卡正常。 | 更换测试卡正常。 | SWO-62874 | No | The replacement test card is OK. | The replacement test card is OK.
ABL80 FLEX BASIC | 更换两张新测试卡修复仪器 | 更换两张新测试卡修复仪器 | SWO-61528 | No | Replace two new test card repair instruments | Replace two new test card repair instruments
ABL90 | Found that the solution pack was not seating fully in the opening causing the solution pack ID PCB to not fully contact the chip onthe solution pack itself. | All parts on the measurin module chassis were replaced to try to correct. All original parts returned to analyzer with the exception of the measuring module chassis which when tried resolved the issue. Chassis removed from another unit, chassis is not a new part. | CFC12F121E39A4A9 | No | Found that the solution pack was not seating fully in the opening causing the solution pack ID PCB to not fully contact the chip onthe solution pack itself. | All parts on the measurin module chassis were replaced to try to correct. All original parts returned to analyzer with the exception of the measuring module chassis which when tried resolved the issue. Chassis removed from another unit, chassis is not a new part. | Defect | HW | Measring module chassis
ABL90 FLEX | Remplacé le kit de maintenance e.o Nettoyage de la section fluidique e.o Upgrade software 3.1MR6 e.o Calibration e.o CQ e.o | Remplacé le kit de maintenance e.o Nettoyage de la section fluidique e.o Upgrade software 3.1MR6 e.o Calibration e.o CQ e.o | SWO-61475 | No | Replaced the maintenance kit e.o Cleaning of the fluidic section e.o Upgrade software 3.1MR6 e.o Calibration e.o CQ e.o | Replaced the maintenance kit e.o Cleaning of the fluidic section e.o Upgrade software 3.1MR6 e.o Calibration e.o CQ e.o | Assisted Replacement | SC | Wetability | MR6 installed
ABL90 FLEX | Überprüfung Einlassnadel => Nadel verbogen Austausch Nadel Fluid Programme i.O. Testmessung i.O. | Überprüfung Einlassnadel => Nadel verbogen Austausch Nadel Fluid Programme i.O. Testmessung i.O. | SWO-61456 | No | Check inlet needle => Needle bent exchange needle fluid programs in the form of test measurement in the | Check inlet needle => Needle bent exchange needle fluid programs in the form of test measurement in the | Replaced | Probe bend
ABL90 FLEX | Kunde berichtet, dass nach LP & SC wechsle noch Lac mit Kal & QC Fehler, Kunde führt noch eine paar Kals durch und wiederholt dann die fehlerhaften QC, bei Eintreffen (siehe SWO614290) Gerät grün, nach SW Update pCO2 QC Fehler im LevelA, Gerät gespült, Kal & QC wiederholt, keine Verbesserung LP ersetzt | Kunde berichtet, dass nach LP & SC wechsle noch Lac mit Kal & QC Fehler, Kunde führt noch eine paar Kals durch und wiederholt dann die fehlerhaften QC, bei Eintreffen (siehe SWO614290) Gerät grün, nach SW Update pCO2 QC Fehler im LevelA, Gerät gespült, Kal & QC wiederholt, keine Verbesserung LP ersetzt | SWO-61544 | No | Customer reports that after LP & SC still change Lac with Kal & QC error, customer performs a few more kals and then repeats the faulty QC, upon arrival (see SWO614290) device green, according to SW update pCO2 QC error in LevelA, device flushed, Kal & QC repeated, no repair ung LP replaced | Customer reports that after LP & SC still change Lac with Kal & QC error, customer performs a few more kals and then repeats the faulty QC, upon arrival (see SWO614290) device green, according to SW update pCO2 QC error in LevelA, device flushed, Kal & QC repeated, no repair ung LP replaced | Assisted Replacement | SP | pCO2 | QC-A
ABL90 FLEX | Neues Lösungspack installiert, Gerät kalibriert und QC durchgeführt. | Neues Lösungspack installiert, Gerät kalibriert und QC durchgeführt. | SWO-61816 | No | New solution pack installed, device calibrated and QC performed. | New solution pack installed, device calibrated and QC performed. | Assisted Replacement | SP | FT
ABL90 | Installing analyzer after repair | started up analyzer installed sensor casette and solution pack; during running Vk transport error and 1238; inhomogeneous rinse solution 1081; 521 inhomogeneous sample ; found old style inlet gasket replaced with one from PM kit | CFC13ACE1E027645 | Yes | Installing analyzer after repair | started up analyzer installed sensor casette and solution pack; during running Vk transport error and 1238; inhomogeneous rinse solution 1081; 521 inhomogeneous sample ; found old style inlet gasket replaced with one from PM kit | Re classified RMED | Install after repair
ABL90 FLEX PLUS | Gerät funktioniert einwandfrei, jedoch wird es beim nächsten wechseln der Dichtung Probleme geben. Ersatzteile bestellt um dies zu verhindern. | Gerät funktioniert einwandfrei, jedoch wird es beim nächsten wechseln der Dichtung Probleme geben. Ersatzteile bestellt um dies zu verhindern. | SWO-61441 | No | Device works perfectly, but there will be problems the next time you change the seal. spare parts ordered to prevent this. | Device works perfectly, but there will be problems the next time you change the seal. spare parts ordered to prevent this. | Defect | HW | FLEX PLUS | Inlet
ABL90 FLEX | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-61295 | No | Repair after replacing a new test card | Repair after replacing a new test card | Assisted Replacement | SC | pO2 | Cal | 1317
AQT90 FLEX | 清理处理盘上的反射白块和校准白块 | 清理处理盘上的反射白块和校准白块 | SWO-61313 | No | Clean up reflected white blocks and calibrated white blocks on the processing disc | Clean up reflected white blocks and calibrated white blocks on the processing disc | Cleaned | HW | Process Module | Process wheel | 129
AQT90 FLEX |  | B | SWO-61269 | No |  |  | Defect | HW | Process Module | Cup dryer, Air filter | 17
ABL800 BASIC |  | B | SWO-61278 | No |  |  | Defect | HW | Oxi module |  | 251
ABL835 FLEX | PH Elektrode und Messkammer gereinigt. ABL mehrfach Kalibriert. | PH Elektrode und Messkammer gereinigt. ABL mehrfach Kalibriert. | SWO-61509 | No | PH electrode and measuring chamber. ABL calibrated several times. | PH electrode and measuring chamber. ABL calibrated several times. | Cleaned | Fluidic pathway | pH/BG Module | pH
ABL800 BASIC | Given user training for the Biomedical persons, done the maitanance for both the ABL 800 Machines. Started with Touch screen problem and machine is still down due to software and hardware issue. | B | SWO-61508 | No | Given user training for the Biomedical persons, done the maitanance for both the ABL 800 Machines. Started with Touch screen problem and machine is still down due to software and hardware issue. |  | Defect | HW | DMS_PC | Touch Screen
ABL815 FLEX | Analyzer leaking out the front of the inlet. Customer could not stop it. | Changed inlet gasket and probe. | CFC124341E595D51 | No | Analyzer leaking out the front of the inlet. Customer could not stop it. | Changed inlet gasket and probe. | Assisted Replacement | HW | Inlet module | Inlet gasket
ABL825 FLEX |  | B | CFC0DC6E1E51148E | Yes |  |  | Re classified RMED | PM
ABL825 FLEX |  | B | CFC129DF1E4FAA2D | No |  |  | Unknown | -
ABL827 FLEX | The tHb continues to drift out on cals high. Would not stabalize. Hemolyzer corrected the issue. This was the install. See SWO B9P62PUL4 for install. | Replaced hemolyzer. Adjusted lamp. Ran complete OXI test. | CFC1445D1E5AB39B | Yes | The tHb continues to drift out on cals high. Would not stabalize. Hemolyzer corrected the issue. This was the install. See SWO B9P62PUL4 for install. | Replaced hemolyzer. Adjusted lamp. Ran complete OXI test. | Re classified RMED | HW | Oxi module | Hemolyzer
AQT90 FLEX |  | B | CFC113A91E4CFB8A | No |  |  | Defect | HW | Aspiration Module | Bend needle | Pipette 38
AQT90 FLEX | Reinigung, alle Kassetten kontrolliert, Fehlerhafte Kassette entfernt, SafePositions ok, Neustart, Probelauf ok, QK ok, Patientenmessung -> erneuter Fehler, Fehlersuche, Storagewheel demontiert, losen Plastikclip von Kassette aus AQT entfernt, Neustart, QK CRP ok, Patientenmessung | Fehlersuche; div. Komponenten überprüft = div. cups nicht ausgeworfen; cups entfernt; Neustart; Service-dump gezogen; LQC i.O. | SWO-61515 | No | Cleaning, all cassettes checked, Faulty cassette removed, SafePositions ok, restart, test run ok, QK ok, patient measurement -> re-fault, troubleshooting, storagewheel disassembled, loose plastic clip removed from cassette from AQT, restart, QK CRP ok, Patient measurement | Troubleshooting; div. Components checked = div. cups not ejected; cups removed; Restart; Service dump dragged; LQC i.O. | Defect | HW | Storage Module | Clips | Setsafe_22
ABL80 FLEX BASIC | 更换新试剂包后修复 | 更换新试剂包后修复 | SWO-61763 | No | Repair after replacing a new kit | Repair after replacing a new kit
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-62473 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL80 CO-OX | Perform a cxomplete oxi test and found that the LED failed.After installing the Hemolyzer the ABL80 got a fluid blockage. Tried repositioning the tubes several times. Flushed the tubing and it was still clogged. | Replaced te Hemolyzer. Flushed the new Hemolyzer to solve the problem. | CFC26BC594138CFD | No | Perform a cxomplete oxi test and found that the LED failed. After installing the Hemolyzer the ABL80 got a fluid blockage. Tried repositioning the tubes several times. Flushed the tubing and it was still clogged. | Replaced te Hemolyzer. Flushed the new Hemolyzer to solve the problem.
ABL90 FLEX | letzte SC mit Ref. Elektroden Error SC ersetzt, vor Ort war keine mehr LP zusätzlich gewechselt, da fast leer | letzte SC mit Ref. Elektroden Error SC ersetzt, vor Ort war keine mehr LP zusätzlich gewechselt, da fast leer | SWO-61717 | Yes | last SC with Ref. electrodes Error SC replaced, no additional LP was changed on site, as almost empty | last SC with Ref. electrodes Error SC replaced, no additional LP was changed on site, as almost empty | Re classified RMED | Logistical upSWO
ABL90 | Found blockage in waste connector. | Replace waste connector and install PM kit. Install version 3.1 MR6 software. | CFC25C3894599002 | No | Found blockage in waste connector. | Replace waste connector and install PM kit. Install version 3.1 MR6 software. | Defect | HW | Waste connector
ABL90 | Checked data logs, found multiple QC levels ran and out of range. Internal QC is within range. Reran all 3 levels of external QC. Level 3 PO2 still out. Tried running at different temperatures with same results. Ran PO2 complete test in service. Possible contamination suspected. Ordered PM kit for overnight delivery. | Replaced all tubings, seals, and gaskets. Cleaned and decontaminated analyzer. Flushed wet section during overhaul. | CFC26EC8941F4069 | No | Checked data logs, found multiple QC levels ran and out of range. Internal QC is within range. Reran all 3 levels of external QC. Level 3 PO2 still out. Tried running at different temperatures with same results. Ran PO2 complete test in service. Possible contamination suspected. Ordered PM kit for overnight delivery. | Replaced all tubings, seals, and gaskets. Cleaned and decontaminated analyzer. Flushed wet section during overhaul. | Adjustment | Decontamination | Tubes replaced | pO2
ABL90 | came in yesterday to run VK, experienced error 1084 inhomogeneouse cal 3 solution 1189 FTC aborted LS state , change error 1268 transport error | ran waste blood; ran service test for PO2 finished running VK Kit; Ran 10 samples of QC for all 3 levels ran 15 blood comparisons | CFC237B0943CE1AA | No | came in yesterday to run VK; experienced error 1084 inhomogeneouse cal 3 solution 1189 FTC aborted LS state ; change error 1268 transport error | ran waste blood; ran service test for PO2 finished running VK Kit; Ran 10 samples of QC for all 3 levels ran 15 blood comparisons | Assisted Replacement | SC | Wetability | MTA
TCM CombiM |  | B | SWO-61590 | No
AQT90 FLEX |  | B | SWO-61605 | No |  |  | Not Cat | Not Cat | Not Cat | Not Cat | JMP mail
ABL825 FLEX | インレットガスケット交換 | B | SWO-63029 | No | Inlet gasket replacement |  | Cleaned | Fluidic pathway | Inlet module | Inlet gasket
ABL835 FLEX |  | B | CFC250CA94281173 | No |  |  | Defect | HW | El/Met Module | Electrode connector | Lac
ABL800 BASIC | Austausch der CPU Unit, BIOS Update durchgeführt | Austausch der CPU Unit, BIOS Update durchgeführt | SWO-61653 | No | Replacement of CPU Unit, BIOS Update Performed | Replacement of CPU Unit, BIOS Update Performed | Defect | HW | DMS_PC | CPU
ABL825 FLEX |  | B | CFC2762394115907 | No |  |  | Defect | HW | Oxi module | Spectrometer | 258, 316
ABL825 FLEX | Airventile in beiden Modulen ersetzt da sie nicht synchron laufen, sporadisch kleben und Fehler im Ventiltest machen, pumpenkal. LS justage, Kalibration, QC- Kontrolle QC | Airventile in beiden Modulen ersetzt da sie nicht synchron laufen, sporadisch kleben und Fehler im Ventiltest machen, pumpenkal. LS justage, Kalibration, QC- Kontrolle QC | SWO-61719 | No | Air valves in both modules are replaced as they do not run synchronously, stick sporadically and make errors in the valve test, pump scale. LS adjustment, calibration, QC-control QC | Air valves in both modules are replaced as they do not run synchronously, stick sporadically and make errors in the valve test, pump scale. LS adjustment, calibration, QC-control QC | Defect | HW | pH/BG Module | Valve, air
ABL837 FLEX | オートチェックモジュール差し替え、引き上げ対応 | B | SWO-62924 | No | Auto check module replacement, raise dissonpable |  | Defect | HW | AutoCheck Module | unknown
ABL827 FLEX | ABL neu gestartet, Scanner ok, Crea-Messkammern gereinigt und Luftblasen von Creaelektroden entfernt, Kal1&Kal2-Lösungen ersetzt, mehrere Kal abgewartet, QK ok | B | SWO-61677 | No | ABL restarted, scanner ok, Crea measuring chambers cleaned and air bubbles removed from creaelectrodes, Kal1&Kal2 solutions replaced, several kalwaited, QK ok |  | Cleaned | Fluidic pathway | Met II Module | Crea
ABL835 FLEX | Installation ABL and various flow error and inlet leaking windows crash | Tried to fix various issue but stop customer request a new instument he is afraid to keep that ABL see also SWO B9P14QJAB | CFC268B7941DA4B0 | No | Installation ABL and various flow error and inlet leaking windows crash | Tried to fix various issue but stop customer request a new instument he is afraid to keep that ABL see also SWO B9P14QJAB | Unknown | New analyzer required
ABL825 FLEX | Oximetriemodul erneuert | Oximetriemodul erneuert | SWO-61676 | No | Oximetry module renewed | Oximetry module renewed | Defect | HW | Oxi module | unknown
AQT90 FLEX | Nadel erneuert. | Nadel erneuert. | SWO-61708 | No | Needle renewed. | Needle renewed. | User error | HW | Aspiration Module | Needle assembly | 15
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-62484 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL80 FLEX BASIC | 更换新测试卡修复仪器 | 更换新测试卡修复仪器 | SWO-62523 | No | Replace the new test card repair instrument | Replace the new test card repair instrument
ABL80 FLEX BASIC |  | B | SWO-61745 | No
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-62458 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL80 FLEX CO-OX | aufgrund v. div. QC Fehlern in der Vergangenheit, wurde das bestehende Gerät mit einem anderen ABL80CoOx (id: 308198) testweise ausgetauscht nach ca. 8 Wochen wird die Situation neu bewertet, um zu entscheiden ob das Austauschgerät weniger Fehler verursacht das Kundengerät wird für diese Zeit mitgenommen | aufgrund v. div. QC Fehlern in der Vergangenheit, wurde das bestehende Gerät mit einem anderen ABL80CoOx (id: 308198) testweise ausgetauscht nach ca. 8 Wochen wird die Situation neu bewertet, um zu entscheiden ob das Austauschgerät weniger Fehler verursacht das Kundengerät wird für diese Zeit mitgenommen | SWO-61884 | No | due to v. div. QC errors in the past, the existing device was replaced with another ABL80CoOx (id: 308198) test-tested after about 8 weeks the situation is reassessed to decide whether the replacement device causes less error the customer device will be taken away for this time | due to v. div. QC errors in the past, the existing device was replaced with another ABL80CoOx (id: 308198) test-tested after about 8 weeks the situation is reassessed to decide whether the replacement device causes less error the customer device will be taken away for this time
ABL90 | Performed service solution pack test with no errors. | Flushed analyzer and installed software version 3.1 MR6. | CFC3C01E0A538D32 | No | Performed service solution pack test with no errors. | Flushed analyzer and installed software version 3.1 MR6. | Assisted Replacement | SC | Wetability | MR6 installed
ABL90 FLEX | 更换新的测试卡修复仪器 | 更换新的测试卡修复仪器 | SWO-61782 | No | Replace a new test card repair instrument | Replace a new test card repair instrument | Assisted Replacement | SC | Na+
ABL90 FLEX | 更换测试卡正常。 | 更换测试卡正常。 | SWO-62872 | No | The replacement test card is OK. | The replacement test card is OK. | Assisted Replacement | SC | Reference
ABL90 FLEX | 更换新的测试卡后修复 | 更换新的测试卡后修复 | SWO-61850 | No | Repair after replacing a new test card | Repair after replacing a new test card | Assisted Replacement | SC | CL, pCO2
ABL90 | Comprobar SC y SP, posicionamiento del selector de flujos y motor del selector. Cambiar el detector de posición del selector de flujos. | B | CFC38F790A08C660 | No | Check SC and SP, flow selector positioning and selector motor. Change the position detector of the flow selector. |  | Defect | HW | Flow selector PCB
AQT90 FLEX | 更换条码扫描器910-304一个后修复 | 更换条码扫描器910-304一个后修复 | SWO-61568 | No | Replace barcode scanner 910-304 a post-repair | Replace barcode scanner 910-304 a post-repair | Defect | HW | Storage Module | Barcode reader, Storage wheel | BR1
AQT90 FLEX |  | B | SWO-61759 | No |  |  | Defect | HW | Aspiration Module | Bend needle | 4794
ABL80 FLEX CO-OX | CO-OXﾓｼﾞｭｰﾙのｺﾈｸﾀｰ類の確認・抜き先で再立ち上げで復帰。 | B | SWO-61824 | No | Re-launch ed.
ABL800 BASIC | 比較電極交換実施。 | 比較電極交換実施。 | SWO-61865 | No | Comparative electrode exchange was implemented. | Comparative electrode exchange was implemented. | Assisted Replacement | Electrode | Electrode | Reference
ABL835 FLEX | 28-10-2015: Skiftet pH o-ring, ventiler i pHmodul pumpeslanger i pH og EL MET modul Kal. og QC | Skiftet pH fodpakning Skiftet pco2 membran Kalibrering og QC | SWO-61872 | No | 28-10-2015: Switched pH o-ring, valves in pHmodul pump hoses in pH and EL MET module KAL. and QC | Switched pH Foot gasket switched PCO2 membrane calibration and QC | Assisted Replacement | Electrode | Membrane | pCO2
ABL800 BASIC | Need to change the New PH BG Module PCB, cross checked with another machine Boards installed in the same Hospital. | B | SWO-62615 | No | Need to change the New PH BG Module PCB, cross checked with another machine Boards installed in the same Hospital. |  | Defect | HW | pH/BG Module | PCB
ABL820 FLEX | At random times the unit locks up causing the customer to power down and reboot. | Saved set up and data files. Reloaded software 6.15 and loaded files and set up. | CFC4C7888028559B | No | At random times the unit locks up causing the customer to power down and reboot. | Saved set up and data files. Reloaded software 6.15 and loaded files and set up. | Unknown | SW | Restart analyzer
AQT90 FLEX | Cleaned dust in the barcode encapsulation. Machine shifted to an another convenient position for avoiding dust. | Cleaned dust in the barcode encapsulation. Machine shifted to an another convenient position for avoiding dust. | SWO-61873 | No | Cleaned dust in the barcode encapsulation. Machine shifted to an another convenient position for avoiding dust. | Cleaned dust in the barcode encapsulation. Machine shifted to an another convenient position for avoiding dust. | Cleaned | HW | Storage Module | Barcode reader, Storage wheel
AQT90 FLEX | Aufstellung des Reparaturaustauschgerätes 838R0021N005 Testkassetten für alle Parameter bis auf PCT eingelegt Systemreinigung i.O. Netzwerkverbindung zu Conworx i.O. => Server lässt sich anpingen Neue PCT Kal-Kassette Lot 10934 bzw. neues Test-Kit für Kalibrierung auf Austauschgerät bestellt Kundengerät zur Reparatur mitgenommen | Aufstellung des Reparaturaustauschgerätes 838R0021N005 Testkassetten für alle Parameter bis auf PCT eingelegt Systemreinigung i.O. Netzwerkverbindung zu Conworx i.O. => Server lässt sich anpingen Neue PCT Kal-Kassette Lot 10934 bzw. neues Test-Kit für Kalibrierung auf Austauschgerät bestellt Kundengerät zur Reparatur mitgenommen | SWO-61899 | No | Installation of the repair replacement device 838R0021N005 test cassettes for all parameters except PCT inserted System cleaning according to network connection to Conworx i.O. => Server can be pinged New PCT Kal cassette Lot 10934 or new test kit for calibration on Replacement device orders customer device taken for repair | Installation of the repair replacement device 838R0021N005 test cassettes for all parameters except PCT inserted System cleaning according to network connection to Conworx i.O. => Server can be pinged New PCT Kal cassette Lot 10934 or new test kit for calibration on Replacement device orders customer device taken for repair | Defect | SW | Analyzer control Module | Power off/on | 3264, 4587
ABL80 FLEX BASIC | 更换新测试卡 | 更换新测试卡 | SWO-62463 | No | Replace a new test card | Replace a new test card
ABL80 CO-OX | Performed zero check procedure and found PO2 was high out of range on gains. Found sensor connector board had corroded contacts which could cause PO2 gains to be out of range. | Replaced sensor connector board. Verified instrument operation. | CFC5FB7FF613E2BE | No | Performed zero check procedure and found PO2 was high out of range on gains. Found sensor connector board had corroded contacts which could cause PO2 gains to be out of range. | Replaced sensor connector board. Verified instrument operation.
ABL80 CO-OX | Examined sensor cassete, solution pack, hemolyzer and sensor connector board. Performed zero cal procedure | Replaced multiple solution packs and 2 sensor cassettes for troubleshooting purposes. Performed zero cal procedure. Replaced manifold assembly. Jeff Ostroski replaced sensor connector board with his stock. Verfied operation of ABL80 | CFC7BE136C52D9AB | No | Examined sensor cassete, solution pack, hemolyzer and sensor connector board. Performed zero cal procedure | Replaced multiple solution packs and 2 sensor cassettes for troubleshooting purposes. Performed zero cal procedure. Replaced manifold assembly. Jeff Ostroski replaced sensor connector board with his stock. Verfied operation of ABL80
ABL80 FLEX BASIC | 试剂管路接头松了，将松的地方剪掉再连接后修复 | 试剂管路接头松了，将松的地方剪掉再连接后修复 | SWO-62710 | No | Reagent line connector loose, cut loose place and reconnect after repair | Reagent line connector loose, cut loose place and reconnect after repair
ABL90 | Reviewed data logs, reported issue appears to have begun on 10 25 2015. Noted frequent Calibration sensitivity issue, external QC recovery low of range, internal QC recovering both low and high out of range dependent upon level run. | Ran service diagnostics, no reported issues in diagnostics. Flushed fluid path, flushed sample probe, replace inlet gasket holder and grey inlet gasket connection. Replaced Solution Pack to resolve issue. Replaced Sensor Cassette as precautionary measure, requested replacement for Sensor Cassette for customer via credit claim form scanned and emailed to Radiometer Support. | CFC4EAE1801BDE38 | No | Reviewed data logs, reported issue appears to have begun on 10 25 2015. Noted frequent Calibration sensitivity issue, external QC recovery low of range, internal QC recovering both low and high out of range dependent upon level run. | Ran service diagnostics, no reported issues in diagnostics. Flushed fluid path, flushed sample probe, replace inlet gasket holder and grey inlet gasket connection. Replaced Solution Pack to resolve issue. Replaced Sensor Cassette as precautionary measure, requested replacement for Sensor Cassette for customer via credit claim form scanned and emailed to Radiometer Support. | Assisted Replacement | SC | Lac
ABL90 FLEX | SC ersetzt Startup abgewartet | SC ersetzt Startup abgewartet | SWO-62587 | No | SC replaces startup waited for | SC replaces startup waited for | Assisted Replacement | SC | Lac
ABL90 | Troubleshot and found that the problem only happened on internal QC. pO2 was also out on level one of external QC, but that just had to be ran again because they replaced the sensor cassette last night. The customer had put in an older solution pack to troubleshooting, and its this solution pack that was causing the problem. | Reinstalled the newer solution pack, and this resolved the problem. | CFC4E7BB805B93DE | No | Troubleshot and found that the problem only happened on internal QC. pO2 was also out on level one of external QC, but that just had to be ran again because they replaced the sensor cassette last night. The customer had put in an older solution pack to troubleshooting, and its this solution pack that was causing the problem. | Reinstalled the newer solution pack, and this resolved the problem. | Assisted Replacement | SP | Lac
ABL90 FLEX PLUS | Nach den oben beschriebenen Maßnahmen ist der Analysator wieder Einsatz bereit. | Nach den oben beschriebenen Maßnahmen ist der Analysator wieder Einsatz bereit. | SWO-62713 | No | After the measures described above, the analyzer is ready for use again. | After the measures described above, the analyzer is ready for use again. | Assisted Replacement | SP | FT
AQT90 FLEX |  | B | SWO-62534 | No |  |  | Defect | HW | Process Module | Optical Unit | 131
AQT90 FLEX |  | B | SWO-61911 | No |  |  | Defect | HW | Process Module | Cup dryer, PCB | 17
AQT90 FLEX | 更换903-364一个正常。 | 更换903-364一个正常。 | SWO-62446 | No | Replace 903-364 one normal. | Replace 903-364 one normal. | Defect | HW | Process Module | Optical Unit | 131
ABL835 FLEX |  | B | CFC618FEF658F657 | No |  |  | Assisted Replacement | Electrode | Electrode | Reference
ABL825 FLEX | AC Modul geprüft, Karussell ersetzt, Funktionskontrolle ok | AC Modul geprüft, Karussell ersetzt, Funktionskontrolle ok | SWO-62580 | No | AC module checked, carousel replaced, function control ok | AC module checked, carousel replaced, function control ok | Defect | HW | AutoCheck Module | Carrousel
ABL835 FLEX | Ersatzmonitor 910-320r1113n2758 installiert, Funktionskontrolle ok | Ersatzmonitor 910-320r1113n2758 installiert, Funktionskontrolle ok | SWO-62579 | No | Replacement monitor 910-320r1113n2758 installed, function control ok | Replacement monitor 910-320r1113n2758 installed, function control ok | Defect | HW | DMS_PC | Touch Screen | LEH Mail
ABL820 FLEX | 新品部品を交換ののち症状改善を確認 | B | SWO-62941 | No | Check for symptom improvement after replacing new parts |  | Defect | HW | Oxi module | Lamp unit
ABL827 FLEX |  | B | CFC4F54F80372496 | No |  |  | Assisted Replacement | Electrode | Electrode | pH
ABL827 FLEX | Jim Halbleib, CAS, is doing the training today. He ran a sample and the analyzer started leaking from the inlet gasket. It was also coming from the inlet tube over to the pH BG module. He found a clot under the pH electrode and one of the dummy electrodes of the CREA module. I saw this when I installed the analyzer. I replaced the inlet probe, and inlet nipples to make sure there is no obstruction. | Replaced the inlet probe and inlet nipple kit. Replaced all clear waste tubes. Also replaced the inlet probe tube with non-stock tube. | CFC4DB4B8014C2B8 | No | Jim Halbleib, CAS, is doing the training today. He ran a sample and the analyzer started leaking from the inlet gasket. It was also coming from the inlet tube over to the pH BG module. He found a clot under the pH electrode and one of the dummy electrodes of the CREA module. I saw this when I installed the analyzer. I replaced the inlet probe, and inlet nipples to make sure there is no obstruction. | Replaced the inlet probe and inlet nipple kit. Replaced all clear waste tubes. Also replaced the inlet probe tube with non-stock tube. | Defect | HW | Inlet module | Probe
AQT90 FLEX | Barcodeleser erneuert. | Barcodeleser erneuert. | SWO-62487 | No | Barcode reader renewed. | Barcode reader renewed. | Defect | HW | Aspiration Module | Barcode reader, Inlet
TCM-COMBIM |  | B | CFC4BDA080127546 | No
TCM-COMBIM |  | B | CFC63A13F60AE890 | No
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-62696 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL80 FLEX | check the power supply & main distributor board. need to Replace both the part. System has measure breakdown. | B | SWO-62733 | No | check the power supply & main distributor board. need to Replace both the part. System has measure breakdown.
ABL90 |  | B | CFC61CFCF6552F07 | Yes |  |  | Re classified RMED | Logistical upSWO | PER
ABL90 | Error is intermittent an I was not able to duplicate the error. | Replaced the Inlet Position Detector PCB. | CFC62C72F607039A | No | Error is intermittent an I was not able to duplicate the error. | Replaced the Inlet Position Detector PCB. | Defect | HW | Inlet Pos. PCB
ABL90 FLEX | Hemolyzer unit heb vervangen en gecontroleerd, OK. | Hemolyzer unit heb vervangen en gecontroleerd, OK. | SWO-62835 | No | Hemolyzer unit have replaced and checked, OK. | Hemolyzer unit have replaced and checked, OK. | Defect | HW | Hemolyzer
ABL90 |  | B | CFC689A8F617989D | Yes |  |  | Re classified RMED | PER reg.
ABL90 |  | B | CFC5F2B3F61EC776 | No
ABL80 CO-OX | Examined sensor cassette and solution pack. | Ran 2 system cycles. Verified operation of ABL 80. | CFC5F7FEF633C2B1 | No | Examined sensor cassette and solution pack. | Ran 2 system cycles. Verified operation of ABL 80.
ABL90 FLEX | 更换试剂包 | 更换试剂包 | SWO-62722 | No | Replace the reagent kit | Replace the reagent kit | Assisted Replacement | SP | Glu/Lac | H2
AQT90 FLEX |  | 更换光学单需要重新定标，再借定标卡。 | SWO-62729 | No |  | Replacing the optical sheet requires a recalibration and borrowing the bid card. | Defect | HW | Process Module | Optical Unit | 131
AQT90 FLEX | 更换光学单需要重新定标，再借定标卡。 | 更换光学单需要重新定标，再借定标卡。 | SWO-62508 | No | Replacing the optical sheet requires a recalibration and borrowing the bid card. | Replacing the optical sheet requires a recalibration and borrowing the bid card. | Defect | HW | Process Module | Optical Unit | 131
AQT90 FLEX |  | B | SWO-62662 | No |  |  | Defect | HW | Storage Module | Barcode reader, Storage wheel | -
AQT90 FLEX |  | 需要加借定标卡。 | SWO-62723 | Yes |  | A bid card needs to be borrowed. | Re classified RMED | Double upSWO
AQT90 FLEX | 需要加借定标卡。 | 需要加借定标卡。 | SWO-62697 | No | A bid card needs to be borrowed. | A bid card needs to be borrowed. | Defect | HW | Process Module | Optical Unit | 131
ABL80 | Limpieza de canales, y ajuste del cero. sustitucion de sc y sp. Ciclo de calidad | B | CFC90BD5E2347A63 | No | Channel cleaning, and zero adjustment. substitution of sc and sp. Quality cycle
ABL837 FLEX | removed disk and checked for data, cannot access rome.db | removed DMS and replaced hard drive, attempted to load windows, windows software will not load on to new hard drive. | CFC9AB5A5847DC8D | No | removed disk and checked for data, cannot access rome.db | removed DMS and replaced hard drive, attempted to load windows, windows software will not load on to new hard drive.
ABL825 FLEX |  | B | CFC70B6C6C178A25 | No |  |  | Cleaned | Fluidic pathway
ABL815 FLEX | Keys installiert. | Keys installiert. | SWO-62807 | Yes | Keys installed. | Keys installed. | Re classified RMED | Demo key exp
ABL825 FLEX | Error 793 Assymmetry pump calibration | Found the Gas valve in the Electrolyte module frozen Replaced valve and pump tubing. | CFC60F8CF611B266 | No | Error 793 Assymmetry pump calibration | Found the Gas valve in the Electrolyte module frozen Replaced valve and pump tubing.
AQT90 FLEX |  | B | CFC640C3F612C680 | No |  |  | Defect | HW | Process Module | Cup dryer, PCB | 168
AQT90 FLEX | 更换903-364一个正常。 | 更换903-364一个正常。 | SWO-62870 | No | Replace 903-364 one normal. | Replace 903-364 one normal. | Defect | HW | Process Module | Optical Unit | 131
ABL800 BASIC | cross checked with working Oxi Module, instrument working fine it need to replace with new Module. requesting new Module. | cross checked with working Oxi Module, instrument working fine it need to replace with new Module. requesting new Module. | SWO-63222 | No | cross checked with working Oxi Module, instrument working fine it need to replace with new Module. requesting new Module. | cross checked with working Oxi Module, instrument working fine it need to replace with new Module. requesting new Module. | Defect | HW | Oxi module | Lamp unit | Neon
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-63008 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL90 FLEX | 更换903-165一个正常。 | 更换903-165一个正常。 | SWO-62840 | No | Replace the 903-165 one normal. | Replace the 903-165 one normal. | Defect | HW | Wet section Control PCB
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-63231 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL80 FLEX BASIC | 更换新试剂包后修复 | 更换新试剂包后修复 | SWO-63013 | No | Repair after replacing a new kit | Repair after replacing a new kit
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-62986 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL90 FLEX |  | Fehlerbehebung durch Techniker notwendig | SWO-63496 | Yes |  | Troubleshooting by technicians necessary | Re classified RMED | Telephone support
ABL90 | SP lid had not been depressed completely. | Compressed SP cap completely. | CFCA1F5D583650F2 | No | SP lid had not been depressed completely. | Compressed SP cap completely. | User error | SP Activation
ABL90 | Unit started displaying error after doing upgrade to version MR6. | Unresolved. | CFCA215C582E50F5 | No | Unit started displaying error after doing upgrade to version MR6. | Unresolved. | Defect | SW | MR6
ABL90 FLEX PLUS |  | B | CFC9E75258431558 | Yes |  |  | Re classified RMED | Logistical upSWO
ABL90 FLEX | 更换新试剂包后修复 | 更换新试剂包后修复 | SWO-63467 | No | Repair after replacing a new kit | Repair after replacing a new kit | Assisted Replacement | SP | H2
ABL90 | Verificação dos registros, limpeza do circuito. inspecção da sonda, subsituição da entrada, inspeção dos roletes da bomba, ajuste dos SLs, calibração da bomba, execução do teste de pressão, substituição do pack de soluções, estabilização e passagem dos QCs. | B | CFC9FDC65851BF62 | No | Checking records, cleaning the circuit. Probe inspection, input subsitution, pump roller inspection, SLs adjustment, pump calibration, pressure test execution, Solution pack replacement, QCs stabilization and passage. |  | Assisted Replacement | SP | Glu/Lac
ABL837 FLEX | removed disk and checked for data, cannot access rome.db | removed DMS and replaced hard drive, loaded windows software usning new software disk, loaded ABL800 software ver 6.13, loaded customer setup, | CFCA73D55847DC99 | No | removed disk and checked for data, cannot access rome.db | removed DMS and replaced hard drive, loaded windows software usning new software disk, loaded ABL800 software ver 6.13, loaded customer setup, | Defect | HW | DMS_PC | HDD
ABL835 FLEX | ハロゲンランプ調整 | ハロゲンランプ調整 | SWO-63089 | No | Halogen lamp adjustment | Halogen lamp adjustment | Adjustment | HW | Oxi module | Lamp unit
ABL825 FLEX | The analyzer attempts to run qc, the ampoule is opened but it will not spin the carousel to the aspiration position. | Replaced the autocheck FPC cable and the clamp for the cable. | CFCA20365834616A | No | The analyzer attempts to run qc, the ampoule is opened but it will not spin the carousel to the aspiration position. | Replaced the autocheck FPC cable and the clamp for the cable. | Defect | HW | AutoCheck Module | FPC Cable
ABL825 FLEX | Gerät überprüft, pH BG Modul ausgetauscht (823R0125N003) Siehe HJW | B | SWO-63135 | No | Device checked, pH BG module replaced (823R0125N003) See HJW |  | Defect | HW | pH/BG Module | ph/BG Module
ABL827 FLEX | Gerät überprüft, Sicherungen und Verbindungen kontrolliert, | Gerät überprüft, Neustart durchgeführt, fällige Elektroden membraniert | SWO-63500 | No | device checked, fuses and connections checked, | Device checked, restart performed, due electrodes membraneed | Revisit required | HW | DMS_PC | DMS_PC
ABL827 FLEX | Behälter erneuert, CREA membraniert | B | SWO-63495 | No | Container renewed, CREA membraneed |  | User error | Solution bottle | Not activated
ABL825 FLEX | Gerät aus- und eingeschaltet, Einlassdichtung u. -Sonde ersetzt, Tray kalibriert, Neustart, akz. Kal abgewartet, Kal u. QC i.O. (Einlasssonde u. Dichtung aus Kundenbestand) | Gerät aus- und eingeschaltet, Einlassdichtung u. -Sonde ersetzt, Tray kalibriert, Neustart, akz. Kal abgewartet, Kal u. QC i.O. (Einlasssonde u. Dichtung aus Kundenbestand) | SWO-63190 | No | Device switched off and on, inlet seal and probe replaced, tray calibrated, restart, etc. Kal waited, Kal & QC i.O. (inlet probe and seal from customer stock) | Device switched off and on, inlet seal and probe replaced, tray calibrated, restart, etc. Kal waited, Kal & QC i.O. (inlet probe and seal from customer stock) | Defect | HW | Inlet module | Probe bend
ABL810 FLEX | Kontakt gereinigt. | Kontakt gereinigt. | SWO-63522 | No | Contact cleaned. | Contact cleaned. | Assisted Replacement | Electrode | Electrode | pCO2
TCM CombiM | Geen fout gevonden. | Geen fout gevonden. | SWO-63072 | No | No error found. | No error found.
TCM400-3 | pO2 sens. error during calibration | Replaced membrane. Returning to replace electrode. | CFCA3865585B0975 | No | pO2 sens. error during calibration | Replaced membrane. Returning to replace electrode.
ABL80 FLEX BASIC | 更换新测试卡 | 更换新测试卡 | SWO-63480 | No | Replace a new test card | Replace a new test card
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-63248 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL80 FLEX BASIC | 更换新测试卡 | 更换新测试卡 | SWO-63725 | No | Replace a new test card | Replace a new test card
ABL80 FLEX BASIC | 更换新测试卡 | 更换新测试卡 | SWO-63736 | No | Replace a new test card | Replace a new test card
ABL90 | WHEN THE VK KIT IS RUN THE THB VALUES ARE LOW FOR ALL THREE LEVELS | THE POWER SUPPLY, WET SECTION CONTROL BOARD AND HEMOLYZER WERE REPLACED. THE THB CAL WAS PERFORMED, INTERNAL QC WAS RUN AND EXTERNAL TOO. THE QC'S PASS AND ARE CLOSE TO THE ESTABLISHED MEANS. | CFCB2B2BCE59F980 | No | WHEN THE VK KIT IS RUN THE THB VALUES ARE LOW FOR ALL THREE LEVELS | THE POWER SUPPLY, WET SECTION CONTROL BOARD AND HEMOLYZER WERE REPLACED. THE THB CAL WAS PERFORMED, INTERNAL QC WAS RUN AND EXTERNAL TOO. THE QC'S PASS AND ARE CLOSE TO THE ESTABLISHED MEANS. | Defect | HW | Hemolyzer
ABL90 | The ribbon cable was not plegged in all the way into the connector. | Put the ribbon cable into the connector correctly. | CFCB4A2ACE3FFFE3 | No | The ribbon cable was not plegged in all the way into the connector. | Put the ribbon cable into the connector correctly. | Adjustment | Mixer tray | Cable seating
ABL90 FLEX | Mit Frau Cramer SC gewechselt, den Reklamationsvorgang erläutert und das Retourenformular generiert. Formular an Service@radiometer.de gemailt. Funktion OK. | Mit Frau Cramer SC gewechselt, den Reklamationsvorgang erläutert und das Retourenformular generiert. Formular an Service@radiometer.de gemailt. Funktion OK. | SWO-63323 | No | Changed with Ms. Cramer SC, explained the complaint process and generated the returns form. Form emailed to Service@radiometer.de. Function OK. | Changed with Ms. Cramer SC, explained the complaint process and generated the returns form. Form emailed to Service@radiometer.de. Function OK. | Assisted Replacement | SC | pH
ABL90 FLEX | Monitor ausgebaut, Touchsreen erneuert | Fehlerbehebung durch Techniker notwendig | SWO-63499 | No | Monitor removed, Touchsreen renewed | Troubleshooting by technicians necessary | Defect | HW | Touch screen
ABL90 FLEX | Display neu kalibriert Funktion überprüft Analysator kalibriert und überprüft QC Messung durchgeführt | Display neu kalibriert Funktion überprüft Analysator kalibriert und überprüft QC Messung durchgeführt | SWO-63346 | No | Display recalibrated function checks analyzer calibrated and verified QC measurement performed | Display recalibrated function checks analyzer calibrated and verified QC measurement performed | Adjustment | Touch screen cal.
ABL90 FLEX | Weiterer Neustart erfolglos, Konfiguration von anderer Station installiert, ResetDataBase, Neustart ok, Konfiguration eingespielt, erneuter Fehler, ResetDatabase, Neustart ok, Konfiguration manuell vorgenommen, Inkonsistente Softwaremeldung vom Lösungspack, neuen Lösungspack installiert und ResetDatabase, manuelle Konfiguration des ABL, Neustart ok, Kal abgewartet ok, fehlerhafte QK bei pH in allen Leveln, Sensorkassette ersetzt, Startphase, erneute pH-Fehler in allen QK-Leveln, SP90 erneut ersetzt, Startphase, mehrere Kal abgewartet, QK ok | Weiterer Neustart erfolglos, Konfiguration von anderer Station installiert, ResetDataBase, Neustart ok, Konfiguration eingespielt, erneuter Fehler, ResetDatabase, Neustart ok, Konfiguration manuell vorgenommen, Inkonsistente Softwaremeldung vom Lösungspack, neuen Lösungspack installiert und ResetDatabase, manuelle Konfiguration des ABL, Neustart ok, Kal abgewartet ok, fehlerhafte QK bei pH in allen Leveln, Sensorkassette ersetzt, Startphase, erneute pH-Fehler in allen QK-Leveln, SP90 erneut ersetzt, Startphase, mehrere Kal abgewartet, QK ok | SWO-63340 | No | Further restart unsuccessful, configuration installed from other station, ResetDataBase, restart ok, configuration played in, re-error, ResetDatabase, restart ok, configuration done manually, Inconsistent software message from the solution pack, new Solution pack installed and ResetDatabase, manual configuration of the ABL, restart ok, Kal waited ok, faulty QK at pH in all levels, sensor cartridge replaced, start phase, new pH error in all QC levels, SP90 replaced, start phase, several cal waited, QK ok | Further restart unsuccessful, configuration installed from other station, ResetDataBase, restart ok, configuration played in, re-error, ResetDatabase, restart ok, configuration done manually, Inconsistent software message from the solution pack, new Solution pack installed and ResetDatabase, manual configuration of the ABL, restart ok, Kal waited ok, faulty QK at pH in all levels, sensor cartridge replaced, start phase, new pH error in all QC levels, SP90 replaced, start phase, several cal waited, QK ok | Defect | SW | database
ABL90 | Unit previously at Ver MR5 with Fluid Flow issues. | Retrograde Wetsection software back to MR5. | CFCB47EECE5BE362 | No | Unit previously at Ver MR5 with Fluid Flow issues. | Retrograde Wetsection software back to MR5. | Defect | SW | Communication | MR6
AQT90 FLEX | 更换光学传感器636-550一个后修复 | B | SWO-63238 | No | Replacement of optical sensor 636-550 one post-repair |  | Defect | HW | Process Module | Process Unit sensor | 5257
ABL825 FLEX | 更换910-321一个正常。 | 更换910-321一个正常。 | SWO-63055 | No | Replace 910-321 one normal. | Replace 910-321 one normal. | Defect | HW | DMS_PC | CPU
ABL825 FLEX | Messkammer und Schlauchsystem geprüft, Ca. Elektrode neu membraniert, Funktionskontrolle ok | Messkammer und Schlauchsystem geprüft, Ca. Elektrode neu membraniert, Funktionskontrolle ok | SWO-63404 | No | Measuring chamber and hose system tested, approx. electrode re-membraneed, function control ok | Measuring chamber and hose system tested, approx. electrode re-membraneed, function control ok | Assisted Replacement | Electrode | Membrane | Ca
ABL825 FLEX | Hemolysator ersetzt, Complete Oxymetrietest, Hb-Lampe neu eingestellt , Complete Oxymetrietest i.O., Kal, Service tHb-Kal, Kal, QC | B | SWO-63428 | No | Hemolysator ersetzt, Complete Oxymetrietest, Hb-Lampe neu eingestellt , Complete Oxymetrietest i.O., Kal, Service tHb-Kal, Kal, QC |  | Defect | HW | Oxi module | Hemolyzer | Noise
ABL827 FLEX | Gerät überprüft, probeweise Monitor gewechselt, CPU Unit erneuert, Funktion und Einstellungen kontrolliert | Gerät überprüft, Neustart durchgeführt, fällige Elektroden membraniert | SWO-63501 | No | Device checked, monitor changed on a trial basis, CPU unit renewed, function and settings controlled | Device checked, restart performed, due electrodes membraneed | Unknown | HW | DMS_PC | CPU
ABL827 FLEX | pH Elektrode (RV02) ausgetauscht, Ref, und Lactat membraniert | B | SWO-63497 | No | pH electrode (RV02) replaced, ref, and lactate membraneed |  | Assisted Replacement | Electrode | Electrode | pH
ABL820 FLEX | When the customer performs maintenance and, puts the instrument on hold, the restart button is greyed out. The error presented is Inlet Flaps removed. | Replaced the inlet flap detection PCB. Replaced the inlet probe. explained the customer how the software version 6.15 forces the user to tell the instrument what was done when the analyzer is placed on 'Hold'. Explained the customer how to restart the analyzer if nothing was replaced. Replaced the seal ring for the pH electrodeand, cleaned the pCO2 electrode, found pCO2 running low on qc level 1. | CFCB3204CE4CC343 | No | When the customer performs maintenance and, puts the instrument on hold, the restart button is greyed out. The error presented is Inlet Flaps removed. | Replaced the inlet flap detection PCB. Replaced the inlet probe. explained the customer how the software version 6.15 forces the user to tell the instrument what was done when the analyzer is placed on 'Hold'. Explained the customer how to restart the analyzer if nothing was replaced. Replaced the seal ring for the pH electrodeand, cleaned the pCO2 electrode, found pCO2 running low on qc level 1. | Defect | HW | Inlet module | Inlet Flap Det. Unit
ABL800 BASIC | Replaced the lactate membrane | Replaced the lactate membrane | SWO-63595 | No | Replaced the lactate membrane | Replaced the lactate membrane | Assisted Replacement | Electrode | Membrane | Lac
ABL835 FLEX | Confirmed problem. | Installed AU software. | CFCB651DCE4055E4 | Yes | Confirmed problem. | Installed AU software. | Re classified RMED | Installation
AQT90 FLEX | change the optical unit, now working satisfactorily | change the optical unit, now working satisfactorily | SWO-64141 | No | change the optical unit, now working satisfactorily | change the optical unit, now working satisfactorily | Defect | HW | Process Module | Optical Unit | 129
ABL80 FLEX BASIC |  | B | SWO-63484 | No
ABL80 FLEX BASIC | 更换两张测试卡 | 更换两张测试卡 | SWO-63531 | No | Replace two test cards | Replace two test cards
ABL80 FLEX BASIC |  | B | SWO-63515 | No
ABL80 FLEX BASIC | 未修复， 再借902-972一个 | B | SWO-63281 | No | Unrepaired, borrow another 902-972
ABL80 FLEX CO-OX | Abfallpumpe erneuert. | Abfallpumpe erneuert. | SWO-63634 | No | Waste pump renewed. | Waste pump renewed.
ABL80 FLEX BASIC | 更换新测试卡 | 更换新测试卡 | SWO-63537 | No | Replace a new test card | Replace a new test card
ABL80 FLEX BASIC |  | B | SWO-63479 | No
ABL90 | Unit has been downgraded to a lower Wet Section version.MR6 to MR5 for Sysbase16 issues. | Unable to configure unit to communicate. | CFCD6C85BA56E367 | No | Unit has been downgraded to a lower Wet Section version. MR6 to MR5 for Sysbase16 issues. | Unable to configure unit to communicate. | Defect | SW | Sybase 16
ABL90 FLEX | Sobald der Abl Zwegs Fehler Behebung in die Medizintechnik gebracht wird oder die Kommunikationeinstellungen geändert werden wird immer wieder die Meldung angezeigt, diese Meldung ist nur weg zu bekommen in dem man die Konfiguration wie beschrieben neu aufsetzt. | Sobald der Abl Zwegs Fehler Behebung in die Medizintechnik gebracht wird oder die Kommunikationeinstellungen geändert werden wird immer wieder die Meldung angezeigt, diese Meldung ist nur weg zu bekommen in dem man die Konfiguration wie beschrieben neu aufsetzt. | SWO-63681 | No | As soon as the Abl Zwegs troubleshooting is brought into the medical technology or the communication settings are changed again and again the message is displayed, this message can only be got away by resetting the configuration as described. | As soon as the Abl Zwegs troubleshooting is brought into the medical technology or the communication settings are changed again and again the message is displayed, this message can only be got away by resetting the configuration as described. | Defect | SW | Communication
ABL90 FLEX | Sobald der Abl Zwegs Fehler Behebung in die Medizintechnik gebracht wird oder die Kommunikationeinstellungen geändert werden wird immer wieder die Meldung angezeigt, diese Meldung ist nur weg zu bekommen in dem man die Konfiguration wie beschrieben neu aufsetzt. | Sobald der Abl Zwegs Fehler Behebung in die Medizintechnik gebracht wird oder die Kommunikationeinstellungen geändert werden wird immer wieder die Meldung angezeigt, diese Meldung ist nur weg zu bekommen in dem man die Konfiguration wie beschrieben neu aufsetzt. | SWO-63695 | No | As soon as the Abl Zwegs troubleshooting is brought into the medical technology or the communication settings are changed again and again the message is displayed, this message can only be got away by resetting the configuration as described. | As soon as the Abl Zwegs troubleshooting is brought into the medical technology or the communication settings are changed again and again the message is displayed, this message can only be got away by resetting the configuration as described. | Defect | SW | Communication
AQT90 FLEX |  | B | SWO-63504 | No |  |  | Defect | HW | Process Module | Optical Unit | 131
AQT90 FLEX | 清理孵育盘相关传感器，并用定标卡944-258重新定标后修复 | 清理孵育盘相关传感器，并用定标卡944-258重新定标后修复 | SWO-63721 | No | Clean up the incubation disk-related sensors and repair them after recalibration with the calibration card 944-258 | Clean up the incubation disk-related sensors and repair them after recalibration with the calibration card 944-258 | Cleaned | HW | Incubator Module | Incubator wheel | 1999.61.22
ABL835 FLEX | Oximetriemodul geprüft, Service tHb Kalibration durchgeführt, Funktionskontrolle ok | Oximetriemodul geprüft, Service tHb Kalibration durchgeführt, Funktionskontrolle ok | SWO-63632 | No | Oximetry module tested, service tHb calibration performed, function control ok | Oximetry module tested, service tHb calibration performed, function control ok | Adjustment | HW | Oxi module | tHb cal
ABL825 FLEX | Ventile überprüft auf Sicherungsring, Funktion und Festsitz, keine Auffälligkeiten. V1-Ventil im BG-Modul mit Kupfersicherungsring ersetzt. Pumpenschläuche Messmodule erneuert. | Valve 1 und 2 erneuert. anschließend Halbjahreswartung durchgeführt, siehe SWO Garantie !!!!! | SWO-63679 | No | Valves checked for safety ring, function and fixed seat, no abnormalities. V1 valve in the BG module with copper retaining ring. Pump hoses measuring modules renewed. | Valve 1 and 2 renewed. then carried out half-year maintenance, see SWO warranty !!!!! | Defect | HW | pH/BG Module | Pump tubing
ABL835 FLEX | Demokeys deaktiviert Neustart | Demokeys deaktiviert Neustart | SWO-63654 | Yes | Demokeys disabled restart | Demokeys disabled restart | Re classified RMED | Demo key exp
ABL825 FLEX | Messkammer gereinigt, pH-Elektrode erneuert, Einlassnadel erneuert, System und Funktion i.O. | Messkammer gereinigt, pH-Elektrode erneuert, Einlassnadel erneuert, System und Funktion i.O. | SWO-63693 | No | Measuring chamber cleaned, pH electrode renewed, inlet needle renewed, system and function in the end | Measuring chamber cleaned, pH electrode renewed, inlet needle renewed, system and function in the end | Assisted Replacement | Electrode | Electrode | pH
ABL825 FLEX |  | B | CFCCB3AB441DBB40 | No |  |  | Adjustment | HW | Analyzer control Module | Barometer
ABL825 FLEX | Verifed the problem by running test sample. | Went to the Auto data trans fer in the setup and checked the box for patient sample. | CFCC70C8440FE40C | No | Verifed the problem by running test sample. | Went to the Auto data trans fer in the setup and checked the box for patient sample. | Setup | HIS/LIS
ABL825 FLEX |  | B | CFCCAFA0444BBB3B | No |  |  | Assisted Replacement | Electrode | Electrode | pCO2
ABL825 FLEX |  | B | CFCCB4C9440EBB43 | No |  |  | Adjustment | HW | Analyzer control Module | Barometer
ABL825 FLEX |  | B | CFCCB5B6444EBB45 | No |  |  | Adjustment | HW | Analyzer control Module | Barometer
AQT90 FLEX | Service Dump gezogen und TS geschickt. Analysator messbereit. | Service Dump gezogen und TS geschickt. Analysator messbereit. | SWO-63703 | Yes | Service dump pulled and Sent TS. Analyzer ready to measure. | Service dump pulled and Sent TS. Analyzer ready to measure. | Re classified RMED | Only to retrieve Service Dump
AQT90 FLEX | Mehrere Messungen sind abgebrochen wegen Fehler 1035 AQT zur Reparatur mit genommen Optische Einheit getauscht und kalibriert alte S N: R0197N006 CRP Lot 10712 neu kalibriert | B | SWO-63506 | No | Several measurements are cancelled due to error 1035 AQT taken for repair Optical unit exchanged and calibrated old S N: R0197N006 CRP Lot 10712 recalibrated |  | Defect | HW | Process Module | Optical Unit | 135
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-63742 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL80 FLEX BASIC | 更换测试卡和试剂包 | 更换测试卡和试剂包 | SWO-63795 | No | Replace test cards and reagent kits | Replace test cards and reagent kits
ABL80 FLEX | pCO2 mehrmals probiert zu Initialisieren mit Kunden SC (ohne Erfolg) Neue Sensorkassette Initialisiert (ohne Erfolg) Neues SP Initialisiert i.o QA Zyklus i.o | pCO2 mehrmals probiert zu Initialisieren mit Kunden SC (ohne Erfolg) Neue Sensorkassette Initialisiert (ohne Erfolg) Neues SP Initialisiert i.o QA Zyklus i.o | SWO-63974 | No | pCO2 tried several times to initialize with customer SC (without success) New sensor cartridge initialized (without success) New SP Initialized i.o QA cycle i.o | pCO2 tried several times to initialize with customer SC (without success) New sensor cartridge initialized (without success) New SP Initialized i.o QA cycle i.o
ABL80 FLEX BASIC | 更换两张测试卡 | 更换两张测试卡 | SWO-63960 | No | Replace two test cards | Replace two test cards
ABL80 FLEX BASIC | 更换902-972修复 | 更换902-972修复 | SWO-63569 | No | Replacement 902-972 repair | Replacement 902-972 repair
ABL80 FLEX BASIC | 更换新测试卡修复仪器 | 更换新测试卡修复仪器 | SWO-65551 | No | Replace the new test card repair instrument | Replace the new test card repair instrument
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-63737 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL825 FLEX | BG,EL測定室研磨加工 BG,ELモジュール電極ハウジング固定穴、再タップ加工。 Glu,Lac,pO2、0点CAL実施。 コンタミ除去実施。LS、ポンプCAL実施 pCO2,Ref,Glu,Lac,メンブラン交換 各部調整動作チェックOK。 CAL２チェックOK。QCチェックOK。 | BG,EL測定室研磨加工 BG,ELモジュール電極ハウジング固定穴、再タップ加工。 Glu,Lac,pO2、0点CAL実施。 コンタミ除去実施。LS、ポンプCAL実施 pCO2,Ref,Glu,Lac,メンブラン交換 各部調整動作チェックOK。 CAL２チェックOK。QCチェックOK。 | SWO-66507 | No | BG, EL measurement chamber polishing BG, EL module electrode housing fixed hole, retapping. Glu, Lac, pO2, 0-point CAL implementation. Contamination removal was carried out. LS, pump CAL implementation pCO2, Ref, Glu, Lac, membrane exchange each part adjustment operation check OK. CAL2 check OK. QC check OK. | BG, EL measurement chamber polishing BG, EL module electrode housing fixed hole, retapping. Glu, Lac, pO2, 0-point CAL implementation. Contamination removal was carried out. LS, pump CAL implementation pCO2, Ref, Glu, Lac, membrane exchange each part adjustment operation check OK. CAL2 check OK. QC check OK. | Cleaned | Fluidic pathway | - | pCO2
ABL837 FLEX |  | B | CFCE1C17BA34EC85 | No |  |  | Defect | HW | pH/BG Module | Valve, air | Membrane
ABL825 FLEX | Barometer von 746mmHg auf 765mHg angepasst. Kalibrationen und Qualitätskontrollen in Ordnung. | Barometer von 746mmHg auf 765mHg angepasst. Kalibrationen und Qualitätskontrollen in Ordnung. | SWO-63968 | No | Barometer adjusted from 746mmHg to 765mHg. calibrations and quality controls in order. | Barometer adjusted from 746mmHg to 765mHg. calibrations and quality controls in order. | Assisted Replacement | Electrode | Electrode | pCO2 | Baro adj
ABL825 FLEX | Barometer von 746mmHg auf 765mmHg korrigiert. Kalibrationen und Qualitätskontrollen in Ordnung. | Barometer von 746mmHg auf 765mmHg korrigiert. Kalibrationen und Qualitätskontrollen in Ordnung. | SWO-63970 | No | Barometer corrected from 746mmHg to 765mmHg. calibrations and quality controls in order. | Barometer corrected from 746mmHg to 765mmHg. calibrations and quality controls in order. | Assisted Replacement | Electrode | Electrode | pCO2 | Baro adj
ABL825 FLEX |  | B | CFCDBC22BA3DC11B | No |  |  | Defect | HW | Oxi module | Halogen Lamp
ABL825 FLEX | Analysator Software neu installiert, Konfiguration durchgeführt, tHb-Kal durchgeführt | Analysator Software neu installiert, Konfiguration durchgeführt, tHb-Kal durchgeführt | SWO-63940 | No | Analyzer software reinstalled, configuration performed, tHb-Kal performed | Analyzer software reinstalled, configuration performed, tHb-Kal performed | Defect | SW | Unknown
ABL825 FLEX | Inaktive Verbindung1 deaktiviert. Powermanagement der Netzwerkkarte deaktiviert. TraceLog für Verbindung iSoft aktiviert. | Inaktive Verbindung1 deaktiviert. Powermanagement der Netzwerkkarte deaktiviert. TraceLog für Verbindung iSoft aktiviert. | SWO-63969 | No | Inactive connection1 disabled. Power management of the NIC is disabled. TraceLog enabled for iSoft connection. | Inactive connection1 disabled. Power management of the NIC is disabled. TraceLog enabled for iSoft connection. | Setup | HIS/LIS
AQT90 FLEX | Mit Herrn Brand aus der Technik in Willich und mit Radiometer Kopenhagen Rücksprache gehalten und Frau Lederer (Apotheke) sowie der Notaufnahme Herrn Freiburg mitgeteilt dass das PCT nicht verwendet werden sollte da es zu falschen Ergebnissen bei der Messung kommen könnte. Herrn Freiburg ( Leiter Notaufnahme) Bescheid gegeben. Herr Freiburg wird sofort eine neue Schachtel PCT bestellen. Die Apotheke wird die tiefgefrorene Schachtel PCT laut der Aussage von Frau Lederer entsorgen. | Mit Herrn Brand aus der Technik in Willich und mit Radiometer Kopenhagen Rücksprache gehalten und Frau Lederer (Apotheke) sowie der Notaufnahme Herrn Freiburg mitgeteilt dass das PCT nicht verwendet werden sollte da es zu falschen Ergebnissen bei der Messung kommen könnte. Herrn Freiburg ( Leiter Notaufnahme) Bescheid gegeben. Herr Freiburg wird sofort eine neue Schachtel PCT bestellen. Die Apotheke wird die tiefgefrorene Schachtel PCT laut der Aussage von Frau Lederer entsorgen. | SWO-63901 | Yes | Mr. Brand from the technical department in Willich and radiometer Copenhagen held consultation and informed Mrs Lederer (pharmacy) and the emergency department Mr. Freiburg that the PCT should not be used as there could be incorrect results in the measurement. Mr Freiburg (Head of Emergency Department). Mr. Freiburg will immediately order a new box of PCT. The pharmacy will dispose of the frozen box PCT, according to Ms Lederer. | Mr. Brand from the technical department in Willich and radiometer Copenhagen held consultation and informed Mrs Lederer (pharmacy) and the emergency department Mr. Freiburg that the PCT should not be used as there could be incorrect results in the measurement. Mr Freiburg (Head of Emergency Department). Mr. Freiburg will immediately order a new box of PCT. The pharmacy will dispose of the frozen box PCT, according to Ms Lederer. | Re classified RMED | Customer informed re how to store PCT
ABL80 FLEX CO-OX | Fehlersuche Abgeschaltete Parameter aktiviert Fließwege überprüft, Schlauch war im Gerät abgeknickt, dadurch Ansaugung der Lsg. fehlerhaft Einlaufverfahren durchgeführt | Fehlersuche Abgeschaltete Parameter aktiviert Fließwege überprüft, Schlauch war im Gerät abgeknickt, dadurch Ansaugung der Lsg. fehlerhaft Einlaufverfahren durchgeführt | SWO-64086 | No | Troubleshooting Switched-off parameters activated flow paths checked, hose was bent in the device, thereby suction of the Lsg. incorrectly performed incoming procedure | Troubleshooting Switched-off parameters activated flow paths checked, hose was bent in the device, thereby suction of the Lsg. incorrectly performed incoming procedure
ABL80 FLEX BASIC | 更换两张新测试卡修复仪器 | 更换两张新测试卡修复仪器 | SWO-64166 | No | Replace two new test card repair instruments | Replace two new test card repair instruments
ABL80 FLEX BASIC | SC removed and flushed with distiled Water. SP reset. SP and SC reinstalled. Working fine. Test run. Pre sampling error reported to Incharge. | SC removed and flushed with distiled Water. SP reset. SP and SC reinstalled. Working fine. Test run. Pre sampling error reported to Incharge. | SWO-64211 | No | SC removed and flushed with distiled Water. SP reset. SP and SC reinstalled. Working fine. Test run. Pre sampling error reported to Incharge. | SC removed and flushed with distiled Water. SP reset. SP and SC reinstalled. Working fine. Test run. Pre sampling error reported to Incharge.
ABL90 | THE INSTRUMENT IS STIL FAILING THE THB PARAMETER FOR THE VK KIT | THE SOFTWARE WAS CHECKED FOR ANY OPTIONS THAT COULD HAVE BEN SELECTED TO AFECT THE RUNING OF THE VK KIT | CFCEF4A7301463DD | No | THE INSTRUMENT IS STIL FAILING THE THB PARAMETER FOR THE VK KIT | THE SOFTWARE WAS CHECKED FOR ANY OPTIONS THAT COULD HAVE BEN SELECTED TO AFECT THE RUNING OF THE VK KIT | Unknown | tHb cal
ABL90 FLEX | Hemolyzer getauscht, Schläuche getauscht, Kaltstartphase, tHb Kal, OXI Test, Kal und QC | Hemolyzer getauscht, Schläuche getauscht, Kaltstartphase, tHb Kal, OXI Test, Kal und QC | SWO-64226 | No | Hemolyzer exchanged, hoses exchanged, cold start phase, tHb Kal, OXI Test, Kal and QC | Hemolyzer exchanged, hoses exchanged, cold start phase, tHb Kal, OXI Test, Kal and QC | Defect | HW | Hemolyzer | Spectrum
ABL90 FLEX PLUS |  | B | CFCECFBE30349271 | No |  |  | Assisted Replacement | SP | Pump tube | 1178
ABL90 | Desmontagem do equipamento, realização do kit de tubos, montagem, teste sem sucesso, substituição do Sp sem sucesso, substituição da SC, establização e passagem do QC. Tudo OK. | B | CFCEF8443038FE95 | No | Disassembly of the equipment, completion of the tube kit, assembly, unsuccessful test, replacement of the Sp without success, replacement of SC, establishment and passage of QC. All OK. |  | Assisted Replacement | SC | Wetability
AQT90 FLEX | 更换了903-138修复了仪器 | 更换了903-138修复了仪器 | SWO-63729 | No | Replaced 903-138 repaired instrument | Replaced 903-138 repaired instrument | Defect | HW | Incubator Module | Incubator, PCB | 2168
ABL800 BASIC | Replace the gas 1 cylinder. Run multiple calibration. Calibration pass. | Replace the PCO2 membrane. after replacement calibration done successful. | SWO-64157 | No | Replace the gas 1 cylinder. Run multiple calibration. Calibration pass. | Replace the PCO2 membrane. after replacement calibration done successful. | Assisted Replacement | Electrode | Membrane | pCO2
ABL835 FLEX |  | B | CFCEF5F5305089BD | No |  |  | Defect | HW | Inlet module | Leaking | Nipple kit
ABL825 FLEX | BLACK SCREEN STAYS ON ENTER SET UP AND DOES NOT RECOGNIZE THE KEYBOARD | INSTALL NEW CPU BOARD.PERFORM START UP | CFCECEA3300A74F8 | No | BLACK SCREEN STAYS ON ENTER SET UP AND DOES NOT RECOGNIZE THE KEYBOARD | INSTALL NEW CPU BOARD. PERFORM START UP | Defect | HW | DMS_PC | CPU
ABL825 FLEX | Gerät überprüft, letzte Membranierung am 23.10.2015, glu und Lac Elektrode membraniert, Kal und QC O.K. | Gerät überprüft, letzte Membranierung am 23.10.2015, glu und Lac Elektrode membraniert, Kal und QC O.K. | SWO-64187 | No | Device checked, last membraneing on 23.10.2015, glu and lac electrode membraneed, Kal and QC O.K. | Device checked, last membraneing on 23.10.2015, glu and lac electrode membraneed, Kal and QC O.K. | Assisted Replacement | Electrode | Membrane | Lac
TCM400-3 | During calibration a sensitivity error occurs | Replaced electrode. | CFCF41413050321D | No | During calibration a sensitivity error occurs | Replaced electrode.
TCM-COMBIM |  | B | CFCECDE63032AE10 | No
ABL80 FLEX BASIC | 更换新测试卡修复仪器 | 更换新测试卡修复仪器 | SWO-64176 | No | Replace the new test card repair instrument | Replace the new test card repair instrument
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-64278 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL90 | Power off and on with no help. | Installed new wet section pcb. | CFD01225A6463290 | No | Power off and on with no help. | Installed new wet section pcb. | Defect | HW | Wet section Control PCB
ABL90 FLEX | LP erneuert Kalibration und QC durchgeführt | LP erneuert Kalibration und QC durchgeführt | SWO-64204 | No | LP renews calibration and QC performed | LP renews calibration and QC performed | Assisted Replacement | SP | QC C pO2
ABL80 FLEX BASIC | 更换902-985修复仪器 | 更换902-985修复仪器 | SWO-63981 | No | Replace the 902-985 repair instrument | Replace the 902-985 repair instrument
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-64246 | No | Repair after replacing a new test card | Repair after replacing a new test card
AQT90 FLEX | 更换910-304一个修复仪器 | 更换910-304一个修复仪器 | SWO-64164 | No | Replace 910-304 a repair instrument | Replace 910-304 a repair instrument | Defect | HW | Storage Module | Barcode reader, Storage wheel
ABL825 FLEX | 更换Lac电极膜修复仪器 | 更换Lac电极膜修复仪器 | SWO-64217 | No | Replace Lac electrode membrane repair instrument | Replace Lac electrode membrane repair instrument | Assisted Replacement | Electrode | Membrane | Lac
ABL80 FLEX BASIC | 更换新测试卡 | 更换新测试卡 | SWO-64737 | No | Replace a new test card | Replace a new test card
ABL80 FLEX BASIC | 更换测试卡 | 更换测试卡 | SWO-64252 | No | Replace the test card | Replace the test card
ABL80 FLEX BASIC |  | B | SWO-64262 | No
ABL80 FLEX BASIC |  | B | SWO-64265 | No
ABL80 FLEX CO-OX | Temperaturkorrigierte Werte Konfiguriert | Temperaturkorrigierte Werte Konfiguriert | SWO-64494 | No | Temperature-corrected values Configured | Temperature-corrected values Configured
ABL80 CO-OX |  | B | CFD2CF65924E6C1E | No
ABL80 FLEX CO-OX | Anschlußstutzen des Lösungspacks gereinigt, Lösungspack erneut installiert, Initialisierung abgewartet, Kal und QCs durchgeführt. | Anschlußstutzen des Lösungspacks gereinigt, Lösungspack erneut installiert, Initialisierung abgewartet, Kal und QCs durchgeführt. | SWO-64451 | No | Connection plug of the solution pack cleaned, solution pack reinstalled, initialization awaited, Kal and QCs performed. | Connection plug of the solution pack cleaned, solution pack reinstalled, initialization awaited, Kal and QCs performed.
ABL90 |  | B | CFD7F9076A1AB466 | No |  |  | Assisted Replacement | SC | Wetability
ABL90 |  | B | CFD2C20C92111140 | Yes |  |  | Re classified RMED | Planned SW upgrade
ABL90 | Checked operation of sensor bay door. Slightly open door manually to check for binding. | Exercise sensor door. Sensor cassette needs replace. Replaced for customer. Started calibrations. Lac failing level 2 internal QC. pO2 failing external level 2. Ran multiple calibrations and QC to try and correct. | CFD2CE7B924EC02B | No | Checked operation of sensor bay door. Slightly open door manually to check for binding. | Exercise sensor door. Sensor cassette needs replace. Replaced for customer. Started calibrations. Lac failing level 2 internal QC. pO2 failing external level 2. Ran multiple calibrations and QC to try and correct. | Adjustment | SC Door
ABL90 FLEX | Flüssigkeitsstrecke gespült, Kal und QC i.O Probleme mit Station besprochen. Anteil an Abbrüchen ist gleich geblieben | Flüssigkeitsstrecke gespült, Kal und QC i.O Probleme mit Station besprochen. Anteil an Abbrüchen ist gleich geblieben | SWO-64344 | No | Liquid track flushed, Kal and QC i.O problems discussed with station. Percentage of cancellations has remained the same | Liquid track flushed, Kal and QC i.O problems discussed with station. Percentage of cancellations has remained the same | Cleaned | Fluidic pathway
ABL90 | Analyzer is using software version 3.1MR3 | Upgraded software version to 3.1MR6, and this resolved the problem. | CFD294C7920AD43A | No | Analyzer is using software version 3.1MR3 | Upgraded software version to 3.1MR6, and this resolved the problem. | Assisted Replacement | SC | Wetability | MR6
ABL90 | Inspection of the system shows all calibrations passing and no cal errors on calibration log. | Replaced SP & SC multiple times for different errors. | CFD2DD569212CFAE | No | Inspection of the system shows all calibrations passing and no cal errors on calibration log. | Replaced SP & SC multiple times for different errors. | No fault found | Cl
ABL90 FLEX | 更换测试卡。 | 更换测试卡。 | SWO-64472 | No | Replace the test card. | Replace the test card. | Assisted Replacement | SC | pH/K/Na
ABL90 FLEX |  | B | SWO-64244 | No |  |  | Defect | HW | Hemolyzer | Clot
ABL90 FLEX | 更换新试剂包后修复 | 更换新试剂包后修复 | SWO-64313 | No | Repair after replacing a new kit | Repair after replacing a new kit | Assisted Replacement | SP | H2
ABL90 | Revisión del equipo (entrada en el sistema y verificación de los diferentes programas de servicio), sustitución del SP (se observa fallos en rinse, intensidad de neón y presión de vacio, etc..) y SC, siendo este último el causante de los problemas del equipo. Estabilización de las lecturas ce cal y pasar los QC´s, así como comprobación con muestras. | Cambio del SC. | CFD40FEA08433436 | No | Equipment review (entry into the system and verification of the different service programs), replacement of the SP (errors in rinse, neon intensity and void pressure, etc.) and SC, the latter being the cause of the equipment problems. Stabilization of ce cal readings and passing QC's, as well as sample checking. | SC change. | Assisted Replacement | SC | Wetability
ABL835 FLEX | Met CD-ROM 914-616 windows XPe v 1.3 (foutieve versie) opnieuw geïnstalleerd (formateren HDD) na maken van een backup, uitprinten vervangingsschema en opslaan van de instellingen. Hierna software versie 6.13 geïnstalleerd. Back-up trachten in te lezen, weer status code 6. Nogmaals windows XPe installatie nu versie 2.2 en applicatie 6.13 software. Regionale- en tijdsinstellingen aangepast aan Nederland. Instellingen inlezen vanaf versie 6.15 gaat niet. Instellingen inlezen vanaf versie 6.13 LC1 analyser (28-07-2015) gaat wel. Controles opnieuw ingelezen en schema ingesteld. Vervangingsschema aangepast. Koppeling gestest door een meting uit te voeren, resultaten komen keurig over in Radiance en LIS. | Met CD-ROM 914-616 windows XPe v 1.3 (foutieve versie) opnieuw geïnstalleerd (formateren HDD) na maken van een backup, uitprinten vervangingsschema en opslaan van de instellingen. Hierna software versie 6.13 geïnstalleerd. Back-up trachten in te lezen, weer status code 6. Nogmaals windows XPe installatie nu versie 2.2 en applicatie 6.13 software. Regionale- en tijdsinstellingen aangepast aan Nederland. Instellingen inlezen vanaf versie 6.15 gaat niet. Instellingen inlezen vanaf versie 6.13 LC1 analyser (28-07-2015) gaat wel. Controles opnieuw ingelezen en schema ingesteld. Vervangingsschema aangepast. Koppeling gestest door een meting uit te voeren, resultaten komen keurig over in Radiance en LIS. | SWO-64364 | No | With CD-ROM 914-616 Windows XPe v 1.3 (faulty version) reinstalled (formating HDD) After making a backup, print out replacement schedule and save the settings. Hereafter, software version 6.13 is installed. Backup attempt to read, again status Code 6. Again Windows XPe installation now version 2.2 and application 6.13 software. Regional and time settings adapted to the Netherlands. Reading settings from version 6.15 does not go. Reading settings from version 6.13 LC1 Analyser (28-07-2015) is going to be. Checks re-read and schedule set up. Replacement schedule. Coupling tested by performing a measurement, results are neatly passed in Radiance and LIS. | With CD-ROM 914-616 Windows XPe v 1.3 (faulty version) reinstalled (formating HDD) After making a backup, print out replacement schedule and save the settings. Hereafter, software version 6.13 is installed. Backup attempt to read, again status Code 6. Again Windows XPe installation now version 2.2 and application 6.13 software. Regional and time settings adapted to the Netherlands. Reading settings from version 6.15 does not go. Reading settings from version 6.13 LC1 Analyser (28-07-2015) is going to be. Checks re-read and schedule set up. Replacement schedule. Coupling tested by performing a measurement, results are neatly passed in Radiance and LIS. | Unknown | SW | - | Version 6.15
ABL800 BASIC | Replace the PCO2 membrane. after replacement calibration done successful. | Replace the PCO2 membrane. after replacement calibration done successful. | SWO-65453 | No | Replace the PCO2 membrane. after replacement calibration done successful. | Replace the PCO2 membrane. after replacement calibration done successful. | Assisted Replacement | Electrode | Membrane | pCo2
ABL837 FLEX | Caドリフト許容範囲がデフォルト値より厳しく設定されていた為設定しなおし実施。全デュアルバルブメンブランアンカー交換。別途pCO2のQCレベル4Low発生にて確認、pHモジュール~インレット間チューブより圧漏れ確認、交換にて改善。 | Caドリフト許容範囲がデフォルト値より厳しく設定されていた為設定しなおし実施。全デュアルバルブメンブランアンカー交換。別途pCO2のQCレベル4Low発生にて確認、pHモジュール~インレット間チューブより圧漏れ確認、交換にて改善。 | SWO-64871 | No | Since the Ca drift tolerance was set more severely than the default value, it was re-set. All dual valve membrane anchor replacement. Confirmed by QC level 4Low generation of pCO2 separately, pH module - inlet tube from pressure leakage confirmation, improved by replacement. | Since the Ca drift tolerance was set more severely than the default value, it was re-set. All dual valve membrane anchor replacement. Confirmed by QC level 4Low generation of pCO2 separately, pH module - inlet tube from pressure leakage confirmation, improved by replacement. | Defect | HW | pH/BG Module | Valve
ABL837 FLEX | キャリブレーションドリフト許容範囲調整、メンブランアンカー交換。 | キャリブレーションドリフト許容範囲調整、メンブランアンカー交換。 | SWO-68048 | No | Calibration drift tolerance adjustment, membrane anchor replacement. | Calibration drift tolerance adjustment, membrane anchor replacement. | Defect | HW | El/Met Module | Valve
ABL825 FLEX | pH-Modul erneuert. | pH-Modul erneuert. | SWO-64382 | No | pH module. | pH module. | Defect | HW | pH/BG Module | pCO2
ABL825 FLEX | Found leakage coming from EL MET module. Found NA membrane was on cricked. | Replaced membrane on NA electrode. Performed all overdue calibrations and cleanings. Performed necessary membrane changes that were due. Performed necessary calibrations to bring analyzer to a ready state. | CFD320399208156C | No | Found leakage coming from EL MET module. Found NA membrane was on cricked. | Replaced membrane on NA electrode. Performed all overdue calibrations and cleanings. Performed necessary membrane changes that were due. Performed necessary calibrations to bring analyzer to a ready state. | Assisted Replacement | Electrode | Membrane | Na+
ABL800 BASIC | 修理したモジュール返却。全メンブラン交換。 | 修理したモジュール返却。全メンブラン交換。 | SWO-64437 | No | Return repaired modules. All membrane replacement. | Return repaired modules. All membrane replacement. | Cleaned | Fluidic pathway | - | Clot
ABL825 FLEX | Oxi Modul geprüft, tHb Kalibration durchgeführt, Einlassnadel (aus Kundenbestand) ersetzt, Funktionskontrolle ok | Oxi Modul geprüft, tHb Kalibration durchgeführt, Einlassnadel (aus Kundenbestand) ersetzt, Funktionskontrolle ok | SWO-64387 | No | Oxi module tested, tHb calibration performed, inlet needle (from customer stock) replaced, function control ok | Oxi module tested, tHb calibration performed, inlet needle (from customer stock) replaced, function control ok | Adjustment | HW | Oxi module | tHb cal
ABL835 FLEX | Neue Monitoreinheit ( ) installiert, Funktionskontrolle ok | Neue Monitoreinheit ( ) installiert, Funktionskontrolle ok | SWO-64380 | No | New monitor unit ( ) installed, function control ok | New monitor unit ( ) installed, function control ok | Defect | HW | DMS_PC | Touch Screen
ABL800 BASIC | ディスプレイ交換し症状改善。 | ディスプレイ交換し症状改善。 | SWO-64346 | No | Display replacement and symptom improvement. | Display replacement and symptom improvement. | Defect | HW | DMS_PC | Touch Screen
ABL825 FLEX | Checked fluid pathway. Found waste tubing disconnected below the inlet gasket. Checked nipples, found one to be occluded. Unable to flush out. | Replaced inlet nipples. | CFD29E0892508047 | No | Checked fluid pathway. Found waste tubing disconnected below the inlet gasket. Checked nipples, found one to be occluded. Unable to flush out. | Replaced inlet nipples. | Defect | HW | Inlet module | Leaking | Nipple kit
AQT90 FLEX |  | B | CFD298A29250E7D7 | No |  |  | Defect | HW | SP | SP | Lot SK-18
AQT90 FLEX | Fehlersuche Test Kassette blockiert Storage Wheel, Führungsnase abgebrochen Storage Wheel auseinander gebaut und Kassette entfernt Systemüberprüfung OK Systemreinigung durchgeführt QC Messung durchgeführt | Fehlersuche Test Kassette blockiert Storage Wheel, Führungsnase abgebrochen Storage Wheel auseinander gebaut und Kassette entfernt Systemüberprüfung OK Systemreinigung durchgeführt QC Messung durchgeführt | SWO-64279 | No | Troubleshooting Test Cassette Blocks Storage Wheel, Guide Nose Aborted Storage Wheel Disassembled and Cassette Removed System Check OK System Cleaning Performed QC Measurement Performed | Troubleshooting Test Cassette Blocks Storage Wheel, Guide Nose Aborted Storage Wheel Disassembled and Cassette Removed System Check OK System Cleaning Performed QC Measurement Performed | Defect | Consumables | Test Kit | Cartridge | 5689 (Broken)
AQT90 FLEX | Verificação dos registos. Desligar equipamento, verificação da ligação do do leitor de codigo de barras com o monitor. Inicialização. Verificação do estado dos tubos, montagem, verificação do correcto funcionamento. | B | CFD2C67B9223C062 | No | Checking records. Disconnect equipment, check the connection of the barcode reader with the monitor. Startup. Checking the state of the tubes, mounting, checking the correct operation. |  | Cleaned | HW | Analyzer control Module | Barcode reader, display
ABL80 FLEX BASIC |  | B | SWO-64500 | No
ABL80 FLEX BASIC |  | B | SWO-64537 | No
ABL80 FLEX CO-OX | Hemolyzer + Pinch valve ausgetauscht i.o (Garantie) THb Kalibration justierung i.o QA Zyklus i.o | Hemolyzer + Pinch valve ausgetauscht i.o (Garantie) THb Kalibration justierung i.o QA Zyklus i.o | SWO-64508 | No | Hemolyzer + Pinch valve replaced i.o (guarantee) THb calibration adjustment i.o QA cycle i.o | Hemolyzer + Pinch valve replaced i.o (guarantee) THb calibration adjustment i.o QA cycle i.o
ABL80 FLEX BASIC | 更换测试卡正常。 | 更换测试卡正常。 | SWO-64715 | No | The replacement test card is OK. | The replacement test card is OK.
ABL80 FLEX BASIC |  | B | SWO-64473 | No
ABL90 FLEX | Inlezen van software V3.1 MR3 en laten opstarten met lege database, nu start de ABL gewoon door, terugzetten van de instellingen (copy van IZ) Controle van instellingen via meting bloedgas : OK | Inlezen van software V3.1 MR3 en laten opstarten met lege database, nu start de ABL gewoon door, terugzetten van de instellingen (copy van IZ) Controle van instellingen via meting bloedgas : OK | SWO-64631 | No | Reading software V 3.1 MR3 and letting it boot with empty database, now start the ABL just by, restoring the settings (copy from IZ) checking settings via blood gas measurement: OK | Reading software V 3.1 MR3 and letting it boot with empty database, now start the ABL just by, restoring the settings (copy from IZ) checking settings via blood gas measurement: OK | Defect | SW | Status Code 2
ABL90 FLEX | Fehlersuche, CF-Karte gew. - ohne Erfolg, SW neu aufgespielt, Konfig. v. ITS aufgespielt, roma.ini angepaßt, div. Einstellungen vorgenommen, alles i.O. | Fehlersuche; CF-Karte gew. - ohne Erfolg; SW neu aufgespielt; Konfig. v. ITS aufgespielt; roma.ini angepaßt; div. Einstellungen vorgenommen; alles i.O. | SWO-64674 | No | Troubleshooting; CF card. - to no avail; SW replayed; Config. v. ITS; roma.ini adapted; div. Settings made; everything in the | Troubleshooting; CF card. - to no avail; SW replayed; Config. v. ITS; roma.ini adapted; div. Settings made; everything in the | Defect | SW | Roma.ini
ABL90 FLEX | ABL gegen ein Reparaturersatz Gerät SN. 090R604N008 ausgetauscht. Der ABL bleibt solange beim Kunden stehen bis der Fehler an dem Kundengerät behoben ist. | ABL gegen ein Reparaturersatz Gerät SN. 090R604N008 ausgetauscht. Der ABL bleibt solange beim Kunden stehen bis der Fehler an dem Kundengerät behoben ist. | SWO-64673 | No | ABL against a repair replacement device SN. 090R604N008. The ABL remains with the customer until the error on the customer device is fixed. | ABL against a repair replacement device SN. 090R604N008. The ABL remains with the customer until the error on the customer device is fixed. | Defect | SW | Maybe Sybase 16
ABL90 FLEX | 更换测试卡正常。 | 更换测试卡正常。 | SWO-64479 | No | The replacement test card is OK. | The replacement test card is OK. | Assisted Replacement | SC | Lac
ABL90 FLEX | 连续更换两张测试卡后修复 | 连续更换两张测试卡后修复 | SWO-65221 | No | Repair after replacing two test cards in a row | Repair after replacing two test cards in a row | Assisted Replacement | SC | Glu
AQT90 FLEX |  | B | SWO-64527 | No |  |  | Defect | HW | Aspiration Module | Bend needle | 4794
ABL815 FLEX | Einlaßsystem gereinigt, neue Einl.dichtung eingebaut, Nadel gereinigt und gerichtet, Fließwege gereinigt, Hb-Lampe mechanisch neu eingestellt, | Einlaßsystem gereinigt, neue Einl.dichtung eingebaut, Nadel gereinigt und gerichtet, Fließwege gereinigt, Hb-Lampe mechanisch neu eingestellt, | SWO-64684 | No | Inlet system cleaned, new inlet seal installed, needle cleaned and directed, flow paths cleaned, Hb lamp mechanically re-adjusted, | Inlet system cleaned, new inlet seal installed, needle cleaned and directed, flow paths cleaned, Hb lamp mechanically re-adjusted, | Defect | HW | Inlet module | Probe bend
ABL827 FLEX | SMD Sicherungen überprüft, Kontaktflächen gereinigt. ABL ein kalibriert. Verfallene Lösungen ersetzt. Crea, Ref.,Glu und Lac Membraniert. | SMD Sicherungen überprüft, Kontaktflächen gereinigt. ABL ein kalibriert. Verfallene Lösungen ersetzt. Crea, Ref.,Glu und Lac Membraniert. | SWO-64675 | No | SMD fuses checked, contact surfaces cleaned. ABL calibrated. Replaced dilapidated solutions. Crea, Ref.,Glu and Lac Membraned. | SMD fuses checked, contact surfaces cleaned. ABL calibrated. Replaced dilapidated solutions. Crea, Ref.,Glu and Lac Membraned. | Defect | HW | Basic unit | Analyzer Control PCB | Fuse
ABL815 FLEX | Gerät gereinigt und überprüft, Proteinentfernung durchgeführt, Lufteinlass gespült, Messkammern gereinigt, O2 neu membraniert, Ventile in den Messmodulen überprüft, Ventilmembran im pH BG Modul erneuert, LS Justage, Pumpenkal, Kals und QCs durchgeführt | Herr M. bemängelt, dass der pO2 in der BGA-Analyse um bis zu 30% niedriger gemessen wird. Herr M. hat dieses Ergebnis mit dem Labor verglichen. AQC ist in allen drei Level mittig. Es liegen keine offensichtlichen Störungen vor. Preanalytische Fehler wurden vom Kunden ausgeschlossen. Herr M. bittet um eine Überprüfung vor Ort. | SWO-64697 | No | Device cleaned and checked, protein removal performed, air inlet flushed, measuring chambers cleaned, O2 re-membraneed, valves checked in the measuring modules, valve membrane renewed in pH BG module, LS adjustment, pump scale, kals and QCs performed | Mr M. complains that pO2 is measured up to 30% lower in the BGA analysis. Mr M. compared this result to the laboratory. AQC is central in all three levels. There are no obvious disturbances. Preanalytical errors have been excluded by the customer. Mr M. asks for an on-the-spot check. | Defect | HW | pH/BG Module | Valve
ABL825 FLEX | PH電極交換 | PH電極交換 | SWO-68137 | No | PH electrode replacement | PH electrode replacement | Assisted Replacement | Electrode | Electrode | pH
ABL825 FLEX | When running qc, all oxy values are question marked due to message 269 Hemolyzer phase regulation failed. | Replaced the hemolyzer unit. Adjusted the halogen lamp, performed the complete oximetry test and, the ctHb calibration in service mode, all passing within specifications. Also, performed the customer mode ctHb calibration. | CFD4202F08525864 | No | When running qc, all oxy values are question marked due to message 269 Hemolyzer phase regulation failed. | Replaced the hemolyzer unit. Adjusted the halogen lamp, performed the complete oximetry test and, the ctHb calibration in service mode, all passing within specifications. Also, performed the customer mode ctHb calibration. | Defect | HW | Oxi module | Hemolyzer
ABL835 FLEX | Confirmed problem. | Replaced the AU control PCB. Installed OS and Radiometer software. Left spare inlet probes. | CFD41589084BBD6A | No | Confirmed problem. | Replaced the AU control PCB. Installed OS and Radiometer software. Left spare inlet probes. | Defect | HW | Basic unit | Analyzer Control PCB
ABL827 FLEX | I appears the inlet heater is leaking from the bottom. The is a pool of liquid under the heater. | Replaced inlet heater. | CFD3FB6F08012241 | No | I appears the inlet heater is leaking from the bottom. The is a pool of liquid under the heater. | Replaced inlet heater. | Defect | HW | Inlet module | Spinlet count(>149) | Maybe replace if reoccur
AQT90 FLEX | Probennadel ausgebaut und gereinigt Wash unit gereinigt LQC-Lösung (LQC Multi-Check Level 2) mehrmals gemessen | Probennadel ausgebaut und gereinigt Wash unit gereinigt LQC-Lösung (LQC Multi-Check Level 2) mehrmals gemessen | SWO-64678 | No | Sample needle removed and cleaned Wash unit cleaned LQC solution (LQC Multi-Check Level 2) measured several times | Sample needle removed and cleaned Wash unit cleaned LQC solution (LQC Multi-Check Level 2) measured several times | Cleaned | QC | Aspiration Module | Needle | LQC-TnI
AQT90 FLEX | Auslieferung AQT90 Gerät mit Tests bestückt Testmessung CRP erfolgreich | Auslieferung AQT90 Gerät mit Tests bestückt Testmessung CRP erfolgreich | SWO-64627 | No | Delivery AQT90 device equipped with tests Test measurement CRP successful | Delivery AQT90 device equipped with tests Test measurement CRP successful | Not Cat | Not Cat | Not Cat | Not Cat | 3264
ABL80 FLEX BASIC | 更换新测试卡修复仪器 | 更换新测试卡修复仪器 | SWO-65555 | No | Replace the new test card repair instrument | Replace the new test card repair instrument
ABL80 FLEX BASIC |  | B | SWO-64744 | No
ABL80 FLEX BASIC | 更换3张新测试卡 | 更换3张新测试卡 | SWO-64770 | No | Replace 3 new test cards | Replace 3 new test cards
ABL80 FLEX BASIC |  | B | SWO-64776 | No
ABL90 FLEX | Fehlersuche SW aktualisiert V3.1MR6 Überprüfung im Service Programm SC erneut eingesetzt bei Lac bleibt Fehler bestehen (Reaktionsfehler,Empf, Drift),Glu Drift ,Empf.fehler SC90 gewechselt Einlaufverfahren tHb Kal durchgeführt (Fälligkeit) | Überprüfung Kal der letzten 12 h Fehler durchgängig vorhanden Fließwege überprüft, Kontakte gewartet Austausch Sensor interface PCB- Lac Fehler bleibt trotzdem bestehen erneuter Tausch SC90 nach mehreren Kalibrierungen läuft Lac Empf. langsam ein Reaktionsfehler beseitigt | SWO-64854 | No | Troubleshooting SW updates V3.1MR6 check re-used in service program SC with Lac remains error (reaction error,empf, drift),Glu Drift,Empf.error SC90 changed run-in procedure tHb Kal performed (due date) | Checking Kal of the last 12 h error consistently present flow paths checked, contacts maintained replacement sensor interface PCB- Lac error remains nevertheless insisting re-exchange SC90 after several calibrations running Lac Empf. slowly corrects a reaction error | Unknown | Revisit required
ABL90 | se detecta puerta de entrada partida y aguja doblada. se procede a sustituir puerta, aguja y tubo de conexión. tras esto, se instala clip protector. se calibra equipo y se comprueba correcto funcionamiento mediante QCs. | B | CFE6F191F24CB293 | No | split entry door and bent needle are detected. the door, needle and connecting tube are replaced. after this, protective clip is installed. equipment is calibrated and properly functioned using QCs. |  | User error | Probe bend
ABL90 FLEX | 更换三张测试卡正常。 | 更换三张测试卡正常。 | SWO-64801 | No | Replacing the three test cards is OK. | Replacing the three test cards is OK. | Assisted Replacement | SC | Reference
AQT90 FLEX |  | B | SWO-64708 | No |  |  | Defect | HW | Incubator Module | Incubator, PCB | 2168
ABL800 BASIC |  | B | SWO-64746 | No |  |  | Defect | HW | pH/BG Module | Valve
ABL825 FLEX | Einlasssonde erneuert, Einlassdichtung wurde schon von Anwendern erneuert, Abfallnippel unter der Einlassdichtung erneuert, Pumpenkalibration und LS Justage durchgeführt, Messkammern gereinigt, Elektroden kontrolliert, Kals und QCs durchgeführt | B | SWO-64961 | No | Inlet probe renewed, inlet seal has already been renewed by users, waste nipples under the inlet seal have been renewed, pump calibration and LS adjustment performed, measuring chambers cleaned, electrodes checked, kals and QCs performed |  | User error | Measurement | Inlet module | Capillary broken
ABL835 FLEX | Membraniert mehrfach Kalibriert. | Membraniert mehrfach Kalibriert. | SWO-64933 | No | Membranes multiple calibrated. | Membranes multiple calibrated. | Assisted Replacement | Electrode | Membrane | pCO2
ABL800 FLEX | Kundin hat Fenster geschlossen. Drucker funktionierte wieder. Sollte Fehler nochmals auftreten, wird Kunde sich wieder melden. (Dann evtl. PC Unit tauschen) | Kundin hat Fenster geschlossen. Drucker funktionierte wieder. Sollte Fehler nochmals auftreten, wird Kunde sich wieder melden. (Dann evtl. PC Unit tauschen) | SWO-64804 | No | Customer has closed windows. Printer worked again. If errors occur again, the customer will report back. (Then replace PC Unit if necessary) | Customer has closed windows. Printer worked again. If errors occur again, the customer will report back. (Then replace PC Unit if necessary) | Unknown | HW | Printer module | unknown
ABL825 FLEX |  | B | SWO-64765 | No |  |  | Defect | HW | Fluidic manifold | Fluidic manifold
ABL800 BASIC |  | B | SWO-64717 | No |  |  | Defect | HW | DMS_PC | Touch Screen
ABL825 FLEX | E1001 ern. und den Ref. Connector ern., Reinigung auf alle 24 Stunden verstellt, Gerät kalibriert, QC´s durchgef. | B | SWO-64807 | No | E1001 ern. and the ref. Connector ern., cleaning adjusted to every 24 hours, device calibrated, QC's through. |  | Assisted Replacement | Electrode | Electrode | Reference
ABL827 FLEX |  | B | CFD5378B7E153A40 | Yes |  |  | Re classified RMED | SWU
ABL825 FLEX | Abfallstrecke überprüft und gespült - Erneuerung Einlassnadel mit Einlassschlauch - Kunde hat bereits Einlassdichtung erneuert - LS Justage und Pumpenkalibration - Funktionsüberprüfung etc. (5 Messungen) | Abfallstrecke überprüft und gespült - Erneuerung Einlassnadel mit Einlassschlauch - Kunde hat bereits Einlassdichtung erneuert - LS Justage und Pumpenkalibration - Funktionsüberprüfung etc. (5 Messungen) | SWO-64861 | No | Waste section checked and rinsed - Renewal inlet needle with inlet hose - customer has already renewed inlet seal - LS adjustment and pump calibration - function check etc. (5 measurements) | Waste section checked and rinsed - Renewal inlet needle with inlet hose - customer has already renewed inlet seal - LS adjustment and pump calibration - function check etc. (5 measurements) | Defect | HW | Inlet module | Probe
ABL815 FLEX | Ran complete oxy test and passed. | Replaced the OXY PCB. | CFD52D817E5F412B | No | Ran complete oxy test and passed. | Replaced the OXY PCB. | Defect | HW | Oxi module | PCB
ABL825 FLEX | Netzwerkkabel, Anschlussdose und IP-Adressen mit benachbartem Gerät getauscht. Port zu iSoft unverändert. Verbindung zu iSoft überprüft, nach Treiberneustart in Ordnung. Frau Baechtle wird sich wieder melden, bei welchem Gerät Problem erneut auftritt. | Inaktive Verbindung1 deaktiviert. Powermanagement der Netzwerkkarte deaktiviert. TraceLog für Verbindung iSoft aktiviert. | SWO-64735 | No | Network cable, junction box and IP addresses are exchanged with a neighboring device. Port to iSoft unchanged. Connection to iSoft checked, after driver restart fine. Ms Baechtle will come back to see which device is reoccurring. | Inactive connection1 disabled. Power management of the NIC is disabled. TraceLog enabled for iSoft connection. | Setup | HIS/LIS
AQT90 FLEX | SUSTITUCION DE IMPRESORA POR ESTAR DETERIORADA IMPIDIENDO EL NORMAL FUNCIONAMIENTO. PRUEBAS DE FUNCIONAMIENTO. | B | CFD587037E44F750 | No | PRINTER REPLACEMENT BECAUSE IT IS IMPAIRED PREVENTING NORMAL OPERATION. OPERATION TESTS. |  | Defect | HW | Analyzer control Module | Printer
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-64947 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL80 FLEX BASIC | 更换新测试卡 | 更换新测试卡 | SWO-64957 | No | Replace a new test card | Replace a new test card
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-64960 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-65212 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-64944 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-65449 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL90 FLEX | Überprüfung Kal der letzten 12 h Fehler durchgängig vorhanden Fließwege überprüft, Kontakte gewartet Austausch Sensor interface PCB- Lac Fehler bleibt trotzdem bestehen erneuter Tausch SC90 nach mehreren Kalibrierungen läuft Lac Empf. langsam ein Reaktionsfehler beseitigt | Überprüfung Kal der letzten 12 h Fehler durchgängig vorhanden Fließwege überprüft, Kontakte gewartet Austausch Sensor interface PCB- Lac Fehler bleibt trotzdem bestehen erneuter Tausch SC90 nach mehreren Kalibrierungen läuft Lac Empf. langsam ein Reaktionsfehler beseitigt | SWO-65001 | No | Checking Kal of the last 12 h error consistently present flow paths checked, contacts maintained replacement sensor interface PCB- Lac error remains nevertheless insisting re-exchange SC90 after several calibrations running Lac Empf. slowly corrects a reaction error | Checking Kal of the last 12 h error consistently present flow paths checked, contacts maintained replacement sensor interface PCB- Lac error remains nevertheless insisting re-exchange SC90 after several calibrations running Lac Empf. slowly corrects a reaction error | Defect | HW | Sensor cassette PCB
ABL90 FLEX | Akku erneuert, Gerät neu gestartet | B | SWO-65415 | No | Battery renewed, device restarted |  | Defect | HW | Battery module
ABL90 FLEX | SC 90 Lot R409 installiert. Funktion OK . Reklamationsausdrucke an RDEU gemailt. Fehler trat nach kurzer Zeit wieder auf. Fließwege und Wetsection überprüft, um Service fiel auf, dass die SC Abdeckung sporadisch nicht schließt. Abdeckung justiert. Eventuell Werden Kontakte nicht richtig geschlossen. Funktion z. Zt OK. | Frau Lücke hat die selbe SC heute Morgen ohne Erfolg noch einmal installiert und zwischenzeitlich die SC ersetzt. Der Fehler bleibt bestehen. Frau L. wünscht einen Vorortservice. | SWO-65082 | No | SC 90 Lot R409 installed. Function OK . Complaint printouts to RDEU. Error reappeared after a short time. Flow paths and wetsection checked to service noticed that the SC cover sporadically does not close. Cover adjusted. Contacts may not be closed properly. Function at an ok. | Ms Lücke unsuccessfully reinstalled the same SC this morning and replaced the SC in the meantime. The error persists. Ms. L. wants a suburban service. | Unknown | SC | pCO2
ABL90 FLEX | ABL gegen ein Reparaturersatz Gerät SN. 090R0463N010 ausgetauscht. Der ABL bleibt solange beim Kunden stehen bis der Fehler an dem Kundengerät behoben ist. | ABL gegen ein Reparaturersatz Gerät SN. 090R0463N010 ausgetauscht. Der ABL bleibt solange beim Kunden stehen bis der Fehler an dem Kundengerät behoben ist. | SWO-65132 | No | ABL against a repair replacement device SN. 090R0463N010. The ABL remains with the customer until the error on the customer device is fixed. | ABL against a repair replacement device SN. 090R0463N010. The ABL remains with the customer until the error on the customer device is fixed. | Defect | SW | Connectivity
ABL90 FLEX | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-64943 | No | Repair after replacing a new test card | Repair after replacing a new test card | Assisted Replacement | SC | Reference
ABL90 FLEX | 更换新测试卡 | 更换新测试卡 | SWO-65039 | No | Replace a new test card | Replace a new test card | Assisted Replacement | SC | Reference
AQT90 FLEX |  | B | SWO-65022 | No |  |  | Defect | HW | Process Module | Cup dryer, PCB | 168
ABL825 FLEX |  | B | CFD690F7F4338DB6 | No |  |  | Defect | HW | pH/BG Module | Pump tubing | Assymetri
ABL835 FLEX | Remplacé membrane de reference et téléchargé data pour ouvrir un techline | Remplac membrane de reference | CFD65551F42323C0 | No | Replaced reference membrane and downloaded data to open a techline | Remplac membrane de reference | Assisted Replacement | Electrode | Membrane | Reference
ABL835 FLEX | Pco2 bas sur les 3 niveaux | Remplacé membrane de reference et téléchargé data pour ouvrir un techline | CFD65B05F41A23C4 | No | Pco2 low on 3 levels | Replaced reference membrane and downloaded data to open a techline | Assisted Replacement | Electrode | Membrane | Reference
ABL825 FLEX | Gas flow path suspect. | Replaced misc parts from PM kit. Replaced Inlet probe. Replaced pH Sealing Ring. Replaced Inlet Nipples. | CFD66829F41CBB15 | No | Gas flow path suspect. | Replaced misc parts from PM kit. Replaced Inlet probe. Replaced pH Sealing Ring. Replaced Inlet Nipples. | Defect | HW | Inlet module | Nipple kit | Nipple kit
ABL815 FLEX | FlexQ-Modul von Station 3b mit dem von der Station 1b getauscht, da man dort nur kapillär misst. Gerät überprüft. | FlexQ-Modul von Station 3b mit dem von der Station 1b getauscht, da man dort nur kapillär misst. Gerät überprüft. | SWO-65421 | No | FlexQ module of station 3b is exchanged with that of station 1b, as it is only measured in a capillary manner. device checked. | FlexQ module of station 3b is exchanged with that of station 1b, as it is only measured in a capillary manner. device checked. | Defect | HW | FlexQ Module | Alignment
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-65218 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-65309 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL90 |  | B | CFD7AF866A0BF150 | No |  |  | Defect | HW | Waste tube
ABL90 | Error 0751 was on the screen upon arrival requesting a flush of the system. Found the inlet gasket was covered with blood from top to bottom. | Did a complete flush of the system. Replaced the SC customer stock. The system is now running correctly. | CFD799766A321120 | No | Error 0751 was on the screen upon arrival requesting a flush of the system. Found the inlet gasket was covered with blood from top to bottom. | Did a complete flush of the system. Replaced the SC customer stock. The system is now running correctly. | Cleaned | Fluidic pathway
ABL90 | THIS IS AFTER THE CUSTOMER SHIPPED THE ANALYZER TO COLORADO AND BACK | I REINSTALLED THE SENSOR CASETTE AND SOLUTION PACK WITHOUT ISSUE.INSTRUCTED THE CUSTOMER ON HOW TO PERFORM FLUID TRANSPORT ADJUSTMWNTS IN SERVICE | CFD7C3EB6A025E8F | No | THIS IS AFTER THE CUSTOMER SHIPPED THE ANALYZER TO COLORADO AND BACK | I REINSTALLED THE SENSOR CASETTE AND SOLUTION PACK WITHOUT ISSUE. INSTRUCTED THE CUSTOMER ON HOW TO PERFORM FLUID TRANSPORT ADJUSTMWNTS IN SERVICE | User error | Incorrect power down
ABL90 FLEX | Windows Einstellungen überprüft>I.O PC Modul ausgebaut und Schalter überprüft>I.O Akku Ersetzt. | Windows Einstellungen überprüft>I.O PC Modul ausgebaut und Schalter überprüft>I.O Akku Ersetzt. | SWO-65328 | No | Windows Settings Checked>I.O PC Module Removed and Switch Checked>I.O Battery Replaced. | Windows Settings Checked>I.O PC Module Removed and Switch Checked>I.O Battery Replaced. | Defect | HW | Battery module
ABL90 FLEX | Problem mit SW Update behoben. | Problem mit SW Update behoben. | SWO-65389 | No | Fixed an issue with SW Update. | Fixed an issue with SW Update. | Assisted Replacement | SC | Wetability
ABL90 FLEX | Fehlersuche SW aktualisiert V3.1MR6 Fließwege überprüft Neustart , anschl. i.O. Zählerstände im KH abgelesen | Fehlersuche SW aktualisiert V3.1MR6 Fließwege überprüft Neustart , anschl. i.O. Zählerstände im KH abgelesen | SWO-65394 | No | Troubleshooting SW updates V3.1MR6 flow paths checked restart , see also i.O. meter readings read in the KH | Troubleshooting SW updates V3.1MR6 flow paths checked restart , see also i.O. meter readings read in the KH | Assisted Replacement | SP | FT
ABL90 | Verificação dos registros, limpeza do circuito, limpeza dos contactos da sensor interface com alcool isopropilico, inicialização da cassete, estabilização e passagem dos QCs. | B | CFD84D926A149F9E | No | Checking records, cleaning the circuit, cleaning the contacts of the interface sensor with ISOPROPYIC alcohol, cassette initialization, stabilization and passage of the QCs. |  | Adjustment | Sensor cassette contact
AQT90 FLEX |  | B | SWO-65236 | No |  |  | Defect | HW | Process Module | Cup dryer, PCB | 168
ABL815 FLEX |  | B | CFD9ED895618C0B4 | No |  |  | Assisted Replacement | Electrode | Electrode | Reference
ABL835 FLEX | Checked machine for signs of problems. Checked peer group data to see if customer's qc was running outside of peer group. All was running within specs. | Performed protein remover and service thb calibration. Checked results and and results were in package insert range. Performed 3 levels of qc and compared to peer group. Customer is within 2 sd of peer group range on all 3 levels of qc. | CFDCB5A142147BBE | No | Checked machine for signs of problems. Checked peer group data to see if customer's qc was running outside of peer group. All was running within specs. | Performed protein remover and service thb calibration. Checked results and and results were in package insert range. Performed 3 levels of qc and compared to peer group. Customer is within 2 sd of peer group range on all 3 levels of qc. | Adjustment | HW | Oxi module | tHb cal
ABL825 FLEX | Service done yesterday, possible carryover issue | Replace reference membrane with a lot from my inventory | CFD804956A056841 | No | Service done yesterday; possible carryover issue | Replace reference membrane with a lot from my inventory | Assisted Replacement | Electrode | Membrane | Reference
ABL825 FLEX | SW 6.15 neu installiert, Funktionskontrolle ok | Prof. Dr. Otto bittet um eine Überprüfung vor dem Wochenende. Im Augenblick läuft eine Versuchsreihe, bei der kontinuierlich gemessen werden muss, auch am Wochenende. Kunde bittet um einen Rückruf 05115323890. Bei einen Serviceeinsatz bitte vorher bei Fr. Domberg melden, Sie hat noch etwas für das BGA Gerät. | SWO-65380 | No | SW 6.15 reinstalled, function control ok | Prof. Dr. Otto asks for a review before the weekend. At the moment, a series of experiments, in which continuous measurements must be carried out, also runs on weekends. Customer asks for a callback 05115323890. If you have a service, please report to Ms. Domberg in advance, she still has something for the BGA device. | Unknown | SW | SW | SW
ABL825 FLEX | Abfallstrecke gespült - Einlassschlauch erneuert - Einlassnadel erneuert (Kundenbestand) - LS Justage und Pumpenkalibration - Funktionsüberprüfung etc. | Abfallstrecke gespült - Einlassschlauch erneuert - Einlassnadel erneuert (Kundenbestand) - LS Justage und Pumpenkalibration - Funktionsüberprüfung etc. | SWO-65411 | No | Waste section rinsed - inlet hose renewed - inlet needle renewed (customer stock) - LS adjustment and pump calibration - function check etc. | Waste section rinsed - inlet hose renewed - inlet needle renewed (customer stock) - LS adjustment and pump calibration - function check etc. | Cleaned | Fluidic pathway | Inlet module | Inlet module
ABL825 FLEX | Erneuerung pO2 Elektrode - Kalibrierungen und Funktionsüberprüfung etc. | Erneuerung pO2 Elektrode - Kalibrierungen und Funktionsüberprüfung etc. | SWO-65410 | No | Renewal pO2 electrode - calibrations and function verification etc. | Renewal pO2 electrode - calibrations and function verification etc. | Assisted Replacement | Electrode | Electrode | pO2
ABL827 FLEX | found that they keep getting inlet Lower LS error | replaced the ls to correct the issue | CFD7D5356A2F71E6 | No | found that they keep getting inlet Lower LS error | replaced the ls to correct the issue | Defect | HW | Inlet module | LS
ABL80 FLEX BASIC | 更换新的试剂包后修复 | 更换新的试剂包后修复 | SWO-65465 | No | Repair after replacing a new reagent kit | Repair after replacing a new reagent kit
ABL90 FLEX | 更换测试卡试剂包正常。 | 更换测试卡正常。 | SWO-65499 | No | Replacing the test card kit is OK. | The replacement test card is OK. | Assisted Replacement | SC | pO2
AQT90 FLEX | Inlettürbereich gereinigt - Schalter händlich betätigt - Registrierungsoptik gereinigt - im Service Einlasstür mehrfach geschlossen und geöffnet - Funktionsüberprüfung bei mehreren Messungen ohne Beanstandung ! Bem.: LQC wird von Kundin Fr. Völkert durchgeführt ! | Inlettürbereich gereinigt - Schalter händlich betätigt - Registrierungsoptik gereinigt - im Service Einlasstür mehrfach geschlossen und geöffnet - Funktionsüberprüfung bei mehreren Messungen ohne Beanstandung ! Bem.: LQC wird von Kundin Fr. Völkert durchgeführt ! | SWO-65469 | No | Inlet door area cleaned - switch manually operated - registration optics cleaned - in the service inlet door closed several times and opened - function check for several measurements without complaint ! Bem.: LQC is carried out by customer Fr. Völkert ! | Inlet door area cleaned - switch manually operated - registration optics cleaned - in the service inlet door closed several times and opened - function check for several measurements without complaint ! Bem.: LQC is carried out by customer Fr. Völkert ! | Cleaned | HW | Cover and Chassis | Switch,Inlet Door | 4587
ABL90 FLEX |  | B | SWO-65479 | No |  |  | Defect | HW | Hemolyzer
Tosca-500 | Observed error upon power up and attempt to calibrate sensor. Noticed white material in left sensor port. | Cleaned sensor port, cleaned calibration port assembly. Re initialized monitor, completed successful calibration X2 with no errors reported. Monitor requesting sensor to be remembraned. Customer will remembrane and test unit and report any further issue. | CFDB8D49CC359236 | No | Observed error upon power up and attempt to calibrate sensor. Noticed white material in left sensor port. | Cleaned sensor port, cleaned calibration port assembly. Re initialized monitor, completed successful calibration X2 with no errors reported. Monitor requesting sensor to be remembraned. Customer will remembrane and test unit and report any further issue. | Cleaned | HW | TCM4 module | Sensor Contacts
ABL80 FLEX BASIC | 更换新测试卡修复仪器 | 更换新测试卡修复仪器 | SWO-65558 | No | Replace the new test card repair instrument | Replace the new test card repair instrument
ABL80 FLEX BASIC | 更换新测试卡修复仪器 | 更换新测试卡修复仪器 | SWO-65559 | No | Replace the new test card repair instrument | Replace the new test card repair instrument
ABL80 FLEX BASIC | 更换测试卡正常。 | 更换测试卡正常。 | SWO-66280 | No | The replacement test card is OK. | The replacement test card is OK.
ABL80 FLEX BASIC | 清理灰尘 | 清理灰尘 | SWO-65749 | No | Clean up the dust | Clean up the dust
ABL80 FLEX BASIC | 更换新测试卡 | 更换新测试卡 | SWO-65601 | No | Replace a new test card | Replace a new test card
ABL80 FLEX BASIC | 更换新测试卡 | 更换新测试卡 | SWO-66925 | No | Replace a new test card | Replace a new test card
ABL80 FLEX BASIC | 更换两张新测试卡修复仪器 | 更换两张新测试卡修复仪器 | SWO-66121 | No | Replace two new test card repair instruments | Replace two new test card repair instruments
ABL80 FLEX BASIC | Restarted the analyzer. Checked the SC and Tubings for any leakage. Found that its fine. Reinstalled the solution pack and sensor cassette. | Restarted the analyzer. Checked the SC and Tubings for any leakage. Found that its fine. Reinstalled the solution pack and sensor cassette. | SWO-66686 | No | Restarted the analyzer. Checked the SC and Tubings for any leakage. Found that its fine. Reinstalled the solution pack and sensor cassette. | Restarted the analyzer. Checked the SC and Tubings for any leakage. Found that its fine. Reinstalled the solution pack and sensor cassette.
ABL90 FLEX | LP ST 06 eingesetzt.LP ST02 fehlerhaft. Nach Kal mehrere QC durchgeführt ok | B | SWO-65693 | No | LP ST 06. LP ST02 faulty. After Kal several QC performed ok |  | Assisted Replacement | SP | pCO2 - QC-A high
ABL90 FLEX | Gerät war schon wieder o.k. Hr. Janssen aus dem Labor hatte noch einmal die Wege gespült, danach war wieder alles o.k. nur noch die Kal.und QC abgewartet, | Gerät war schon wieder o.k. Hr. Janssen aus dem Labor hatte noch einmal die Wege gespült, danach war wieder alles o.k. nur noch die Kal.und QC abgewartet, | SWO-65843 | Yes | The device was already o.k. Mr. Janssen from the laboratory had washed the paths again, after that everything was ok again only the Cal.and QC waited, | The device was already o.k. Mr. Janssen from the laboratory had washed the paths again, after that everything was ok again only the Cal.and QC waited, | Re classified RMED | Phone support
ABL90 FLEX | LP erneuert | LP erneuert | SWO-65920 | No | LP renewed | LP renewed | Assisted Replacement | SP | FT
ABL90 FLEX | Lösungspack ersetzt, Neustart, Kalibration, QC- Kontrolle. QC in Ordnung | Lösungspack ersetzt, Neustart, Kalibration, QC- Kontrolle. QC in Ordnung | SWO-65730 | No | Solution pack replaced, restart, calibration, QC control. QC in order | Solution pack replaced, restart, calibration, QC control. QC in order | Assisted Replacement | SP | pCO2 - QC-A high
ABL90 FLEX | Akku getauscht | Akku getauscht | SWO-65681 | No | Battery replaced | Battery replaced | Defect | HW | Battery module
ABL90 FLEX | SC erneuert, Initialisierungsphase abgewartet, Gerät mehrfach manuell kalibriert und QCs manuell wiederholt. | B | SWO-66085 | No | SC renews, waits for initialization phase, manually calibrates the device several times and repeats QCs manually. |  | Assisted Replacement | SC | Glu-Lac
ABL90 |  | B | CFDBA4F7CC074F69 | No |  |  | Adjustment | HW | SC contacts
ABL90 |  | B | CFDBA6EECC2E4F6B | Yes |  |  | Re classified RMED | SWU
ABL90 FLEX |  | B | SWO-65582 | No |  |  | Defect | HW | SP PCB
ABL90 FLEX | 更换新试剂包修复仪器 | 更换新试剂包修复仪器 | SWO-65756 | No | Replace the new reagent kit repair instrument | Replace the new reagent kit repair instrument | User error | SP empty
ABL90 FLEX | 更换两个试剂包修复仪器 | 更换两个试剂包修复仪器 | SWO-65580 | No | Replace two reagent kit repair instruments | Replace two reagent kit repair instruments | Assisted Replacement | SP | Flow selector
AQT90 FLEX |  | B | SWO-65488 | No |  |  | Defect | HW | Process Module | Optical Unit | 129
AQT90 FLEX | 更换903-364，636-550各一个正常。 | 更换903-364，636-550各一个正常。 | SWO-65746 | No | Replace 903-364, 636-550 each normal. | Replace 903-364, 636-550 each normal. | Not Cat | Not Cat | Not Cat | Not Cat
ABL825 FLEX | Fehlersuche: Abfall-Pumpenschläuche bzw. oberer Pumpenschlauch wurde falsch in der Halterung positioniert, dadurch Quetschung mit Blockade und Flüssigkeitsaustritt. Alle Pumpenschläuche sowie die Einlassdichtung ersetzt und den Bereich rund um das Einlass- AutoCheck Modul gereinigt. pH Dichtring ersetzt durch 834-660. AutoCheck Modul demontiert und beide Scanner (Ampullen- und Hauptscanner) gereinigt. Neustart, Kalibrationen, Funktionstest, QC Messungen => Alle Werte OK | Fehlersuche: Abfall-Pumpenschläuche bzw. oberer Pumpenschlauch wurde falsch in der Halterung positioniert, dadurch Quetschung mit Blockade und Flüssigkeitsaustritt. Alle Pumpenschläuche sowie die Einlassdichtung ersetzt und den Bereich rund um das Einlass- AutoCheck Modul gereinigt. pH Dichtring ersetzt durch 834-660. AutoCheck Modul demontiert und beide Scanner (Ampullen- und Hauptscanner) gereinigt. Neustart, Kalibrationen, Funktionstest, QC Messungen => Alle Werte OK | SWO-65640 | No | Troubleshooting: Waste pump hoses or upper pump hose were positioned incorrectly in the bracket, thus crushing with blockage and liquid leakage. Replaced all pump hoses as well as the inlet seal and cleaned the area around the inlet AutoCheck module. pH sealing ring replaced by 834-660. AutoCheck module dismantled and both scanners (ampoule and main scanner) cleaned. Restart, calibrations, function test, QC measurements => All values OK | Troubleshooting: Waste pump hoses or upper pump hose were positioned incorrectly in the bracket, thus crushing with blockage and liquid leakage. Replaced all pump hoses as well as the inlet seal and cleaned the area around the inlet AutoCheck module. pH sealing ring replaced by 834-660. AutoCheck module dismantled and both scanners (ampoule and main scanner) cleaned. Restart, calibrations, function test, QC measurements => All values OK | Assisted Replacement | HW | Fluidic manifold | Waste tubes
ABL835 FLEX | Jeg fandt at den blå luftstuds var stoppet til. Den er nu renset. Der er udført væskesensor justering og pumpekalibrering. Alt er OK igen. | Jeg fandt at den blå luftstuds var stoppet til. Den er nu renset. Der er udført væskesensor justering og pumpekalibrering. Alt er OK igen. | SWO-66194 | No | I found that the blue air nozzle was stopped for. It is now cleaned. Fluid sensor adjustment and pump calibration have been performed. Everything is OK again. | I found that the blue air nozzle was stopped for. It is now cleaned. Fluid sensor adjustment and pump calibration have been performed. Everything is OK again. | Cleaned | Fluidic pathway | Fluidic manifold | Air Vent
ABL825 FLEX | 367 pH BG module not filled and 879 no Cal 1 solution errors on aborted cals. | Removed and cleaned all modules. Replaced pH seal ring. Flushed sample path. | CFDB8AC0CC0620FE | No | 367 pH BG module not filled and 879 no Cal 1 solution errors on aborted cals. | Removed and cleaned all modules. Replaced pH seal ring. Flushed sample path. | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
ABL820 FLEX | ALL 3 LEVELS OF PC02 ARE ON THE LOW SIDE OF PEER GROUP AND LEVEL 1 ALMOST OUT OF PACKAGE INSERT LOW AT62.1 | REPLACED GAS 1 TANK,REPLACED PC02 AND REFERENCE MEMBRANES.LEVEL 1 NOW AT 65.3 | CFDB797CCC5BB5FF | No | ALL 3 LEVELS OF PC02 ARE ON THE LOW SIDE OF PEER GROUP AND LEVEL 1 ALMOST OUT OF PACKAGE INSERT LOW AT62.1 | REPLACED GAS 1 TANK,REPLACED PC02 AND REFERENCE MEMBRANES. LEVEL 1 NOW AT 65.3 | Assisted Replacement | Electrode | Membrane | pCO2
ABL825 FLEX | Oxi Modul geprüft, Service tHb Kalibration durchgeführt, nach Rücksprache mit Frau Domberg die Herstellergrenzen programmiert, Funktionskontrolle ok | die tHb-Kalibrierung ist ok (Fcuv 1,0052), der Anwender hat schon Proteinentfernung durchgeführt, ohne jegliche Besserung; vermutlich Hemolyzer Unit defekt; der Anwender wünscht einen Vorort-Service zur Beseitigung der Geräte-Störung; | SWO-65813 | No | Oxi module tested, service tHb calibration carried out, after consultation with Ms. Domberg the manufacturer limits programmed, function control ok | the tHb calibration is ok (Fcuv 1,0052), the user has already performed protein removal without any improvement; probably Hemolyzer unit defective; the user wants a suburban service to eliminate the device malfunction; | Defect | HW | Oxi module | Hemolyzer
ABL825 FLEX | Found cable that runs from the sensor for the vial cracker arm not connected properly. This was causing vial cracker arm to remain in the up position. | Adjusted connection on cracker arm. Rebooted analyzer. Performed necessary calibrations. | CFDBA6E2CC1116D0 | No | Found cable that runs from the sensor for the vial cracker arm not connected properly. This was causing vial cracker arm to remain in the up position. | Adjusted connection on cracker arm. Rebooted analyzer. Performed necessary calibrations. | Defect | HW | AutoCheck Module | Cable
ABL825 FLEX | Found machine leaking from the PH BG module. Electrodes not seating correctly. | Remembraned the CL electrode. Reseated other electrodes and cleaned the chambers. Performed general cleaning of analyzer. Performed necessary calibrations. | CFDB9421CC2616CD | No | Found machine leaking from the PH BG module. Electrodes not seating correctly. | Remembraned the CL electrode. Reseated other electrodes and cleaned the chambers. Performed general cleaning of analyzer. Performed necessary calibrations. | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
ABL815 FLEX |  | B | CFDDECF5B8448589 | No |  |  | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
AQT90 FLEX | Barcodescanner justiert--> Kassetten werden erkannt | Barcodescanner justiert--> Kassetten werden erkannt | SWO-65682 | No | Barcode scanner adjusted--> cassettes are detected | Barcode scanner adjusted--> cassettes are detected | Adjustment | HW | Storage Module | Barcode reader, Storage wheel
ABL80 FLEX BASIC | 清洁内部管路，重新安装测试卡试剂包 | 清洁内部管路，重新安装测试卡试剂包 | SWO-65777 | No | Clean the internal lines and reinstall the test card kit | Clean the internal lines and reinstall the test card kit
ABL80 FLEX BASIC |  | B | SWO-65767 | No
ABL80 FLEX BASIC | 更换2张新测试卡 | 更换2张新测试卡 | SWO-65761 | No | Replace 2 new test cards | Replace 2 new test cards
ABL80 FLEX BASIC |  | B | SWO-65754 | No
ABL80 FLEX BASIC | 更换新测试卡修复仪器 | 更换新测试卡修复仪器 | SWO-65755 | No | Replace the new test card repair instrument | Replace the new test card repair instrument
ABL80 OSM |  | B | CFE75469F2001060 | No
ABL80 FLEX BASIC | 更换测试卡正常。 | 更换测试卡正常。 | SWO-65757 | No | The replacement test card is OK. | The replacement test card is OK.
ABL80 FLEX BASIC | cleaned SC with cleaning solution and kept in fridge for 1 hour than reinstalled SC. | cleaned SC with cleaning solution and kept in fridge for 1 hour than reinstalled SC. | SWO-66664 | No | cleaned SC with cleaning solution and kept in fridge for 1 hour than reinstalled SC. | cleaned SC with cleaning solution and kept in fridge for 1 hour than reinstalled SC.
ABL90 | Troubleshot to a faulty inlet module | Replaced inlet module, and this resolved the problem. | CFDCEF9842409B02 | No | Troubleshot to a faulty inlet module | Replaced inlet module, and this resolved the problem. | Defect | HW | Inlet module
ABL90 | Activity log indicates 1066 temperature in sensor cassette substrate out of range errors. | Replaced sensor interface PCB and verified proper instrument performance. Also replaced solution pack due to low number of remaining tests. Instrument is within acceptable operating specifications. | CFDCF8F9420C2186 | No | Activity log indicates 1066 temperature in sensor cassette substrate out of range errors. | Replaced sensor interface PCB and verified proper instrument performance. Also replaced solution pack due to low number of remaining tests. Instrument is within acceptable operating specifications. | Defect | HW | Sensor cassette PCB |  | Temp
ABL90 FLEX | Tätigkeiten durch Anwender: System gespült, Einlassdichtung und verbogene Einlassnadel ersetzt, diverse Systemprogramme durchgeführt Überprüfung des Systems, Lösungspack ersetzt, Kalibrationsprogramme der Fließwege durchgeführt Reklamationsmeldungen ausgefüllt Einlaufphase der Sensorkassette und des Lösungspacks abgewartet Kal- und QC --> Fehler pCO2 Level 2 zu niedrig, mehrere Kal- und QC ohne Erfolg Lösungspack ausgetauscht Kal- und QC | Tätigkeiten durch Anwender: System gespült, Einlassdichtung und verbogene Einlassnadel ersetzt, diverse Systemprogramme durchgeführt Überprüfung des Systems, Lösungspack ersetzt, Kalibrationsprogramme der Fließwege durchgeführt Reklamationsmeldungen ausgefüllt Einlaufphase der Sensorkassette und des Lösungspacks abgewartet Kal- und QC --> Fehler pCO2 Level 2 zu niedrig, mehrere Kal- und QC ohne Erfolg Lösungspack ausgetauscht Kal- und QC | SWO-65958 | No | Activities by user: system flushed, inlet seal and bent inlet needle replaced, various system programs carried out inspection of the system, solution pack replaced, calibration programs of flow paths carried out complaint messages completed Inlet phase of the sensor cartridge and the solution pack awaited Kal- and QC --> Error pCO2 Level 2 too low, multiple Cal- and QC without success solution pack replaced Cal- and QC | Activities by user: system flushed, inlet seal and bent inlet needle replaced, various system programs carried out inspection of the system, solution pack replaced, calibration programs of flow paths carried out complaint messages completed Inlet phase of the sensor cartridge and the solution pack awaited Kal- and QC --> Error pCO2 Level 2 too low, multiple Cal- and QC without success solution pack replaced Cal- and QC | Assisted Replacement | SP | Pump cal error
ABL90 | internal printer prints charters not the patient results, prints nL2=n?u0un, external printer not working--- Came on site because of known issues with the MR6 communication issues | loaded MR 7, disconnected exteranl printer and rebooted--internal printer worked right on reboot--printed out results from printing request before the reboot, tried to print on thermal printer and it would not work; reconnected external printer and made sure correct copy was selected External printer works now which is the one the customer wants to use; internal printer is still bad | CFDCBEAA423B11A3 | No | internal printer prints charters not the patient results; prints nL2=n?u0un, external printer not working--- Came on site because of known issues with the MR6 communication issues | loaded MR 7, disconnected exteranl printer and rebooted--internal printer worked right on reboot--printed out results from printing request before the reboot, tried to print on thermal printer and it would not work; reconnected external printer and made sure correct copy was selected External printer works now which is the one the customer wants to use; internal printer is still bad | Setup | Printer Comm
ABL90 FLEX | Apklikation beendet - Roma.ini angepasst - Software Update durchgeführt - siehe TU0134 | Gerät stürzt ab sobald sich jemand anmeldet, Gerät steht in Intervention notwendig, Meldung an unexpected Error has occured, beim bestätigen folgt Geräte Absturz, Rückruf durch KDB erwünscht | SWO-66025 | No | Apklikation ended - Roma.ini adapted - Software update performed - see TU0134 | Device crashes as soon as someone logs on, device is required in intervention, message to unexpected error has occurred, when confirming follows device crash, recall requested by KDB | Defect | SW | Connectivity
ABL90 FLEX PLUS |  | B | CFDCD4B042022198 | Yes |  |  | Re classified RMED | Planned SW upgrade
ABL90 FLEX | 更换试剂包 | 更换试剂包 | SWO-66100 | No | Replace the reagent kit | Replace the reagent kit | Assisted Replacement | SP | FT
ABL90 FLEX | 更换新测试卡修复仪器 | 更换新测试卡修复仪器 | SWO-65800 | No | Replace the new test card repair instrument | Replace the new test card repair instrument | Assisted Replacement | SC | Reference
AQT90 FLEX |  | B | SWO-65753 | No |  |  | Defect | HW | Process Module | Optical Unit | 135
AQT90 FLEX | 更换测试卡正常。 | 更换测试卡正常。 | SWO-66045 | No | The replacement test card is OK. | The replacement test card is OK. | Defect | HW | Test Kit | Cartridge | 2872 (Broken)
ABL825 FLEX | Undichten Dichtring pH-Elektrode erneuert. QCs weiterhin außerhalb Bereich, pCO2-Elektrode erneuert. Alte Elektrode Lot: RT04. Neu: Lot SB01. Kalibrierungen und Qualitätskontrollen in Ordnung. | Undichten Dichtring pH-Elektrode erneuert. QCs weiterhin außerhalb Bereich, pCO2-Elektrode erneuert. Alte Elektrode Lot: RT04. Neu: Lot SB01. Kalibrierungen und Qualitätskontrollen in Ordnung. | SWO-65788 | No | Leaking sealing ring pH electrode renewed. QCs continue to be out of range, pCO2 electrode renewed. Old electrode lot: RT04. New: Lot SB01. Calibrations and quality controls in order. | Leaking sealing ring pH electrode renewed. QCs continue to be out of range, pCO2 electrode renewed. Old electrode lot: RT04. New: Lot SB01. Calibrations and quality controls in order. | Assisted Replacement | Electrode | Electrode | pCo2
ABL800 BASIC | Gerät überprüft und gereinigt, Meßkammer und Luftauslass gesäubert, Dichtungsring pH Elektrode ersetzt, Lufteinschluss in pCO2 Elektrode beseitigt, Endkeys installiert, tHb Kal o.k. | Zusätzlich mit dem Kunden den Barometerstand überprüft. Barometer ist mit 755mmHG korrekt auf den aktuellen Luftdruck eingestellt. Frau S. Wünscht einen Vorortservice. | SWO-65966 | No | Device checked and cleaned, measuring chamber and air outlet cleaned, sealing ring pH electrode replaced, air inclusion eliminated in pCO2 electrode, end keys installed, tHb Kal ok. | In addition, check the barometer status with the customer. With 755mmHG, the barometer is correctly adjusted to the current air pressure. Ms. S. wants a suburban service. | Assisted Replacement | Electrode | Membrane | pCo2
ABL825 FLEX | Gerät überprüft, Meßkammer und Luftauslass gesäubert, Elektroden und Kontakte kontrolliert, pH Dichtungsring ersetzt, Kal und QC o.k. | B | SWO-65965 | No | Device checked, measuring chamber and air outlet cleaned, electrodes and contacts checked, pH sealing ring replaced, Kal and QC ok. |  | Assisted Replacement | Electrode | Electrode | Sealing ring
ABL825 FLEX | Gerät überprüft, pH BG Modul R2242N008 ausgetauscht, R0067N001 eingebaut, Dichtungsring pH Elektrode ersetzt | B | SWO-66052 | No | Device checked, pH BG module R2242N008 replaced, R0067N001 built-in, sealing ring pH electrode replaced |  | Defect | HW | pH/BG Module | ph/BG Module
ABL837 FLEX |  | B | CFDCA7894214A605 | No |  |  | Unknown
ABL815 FLEX | Keine Auffälligkeiten feststellbar, Archiv mit Kunden durchgesprochen und analysiert | Keine Auffälligkeiten feststellbar, Archiv mit Kunden durchgesprochen und analysiert | SWO-66221 | No | No abnormalities detectable, archive discussed with customers and analyzed | No abnormalities detectable, archive discussed with customers and analyzed | No fault found
ABL835 FLEX |  | B | CFDCFF81423C3FF8 | No |  |  | Assisted Replacement | Electrode | Electrode | pCo2
ABL815 FLEX | Friktion Detektor erneuert, Kalibrierung einwandfrei | Friktion Detektor erneuert, Kalibrierung einwandfrei | SWO-66219 | No | Friction detector renewed, calibration flawless | Friction detector renewed, calibration flawless | Defect | HW | Inlet module | Friction detector unit
ABL815 FLEX | Kein pCO Fehler mehr seit einer Woche aufgetreten, Wert konstant unter Zielwert aber konform, nach Rücksprache werden derzeit keine Maßnahmen ergriffen | Kein pCO Fehler mehr seit einer Woche aufgetreten, Wert konstant unter Zielwert aber konform, nach Rücksprache werden derzeit keine Maßnahmen ergriffen | SWO-66220 | No | No pCO error has occurred for a week, value constant below target but compliant, no action is currently taken after consultation | No pCO error has occurred for a week, value constant below target but compliant, no action is currently taken after consultation | No fault found
ABL80 FLEX CO-OX | Manifold und LP überprüft, neue Dichtung an Abfallstutzen angebracht, mehrere Kal, QC | kompletten Wartungssatz verbaut Anschlüsse auf manifold auf korrekten Sitz überprüft, mehrere Kal, QC | SWO-66883 | No | Manifold and LP checked, new seal attached to waste nozzles, several Kal, QC | Complete maintenance kit installed connections on manifold checked for correct fit, several Kal, QC
ABL80 FLEX BASIC |  | B | SWO-66040 | No
ABL80 FLEX BASIC | Sc flushed and SP reset. reinstalled. Clote foung. Sample not aspirating properly. SC Cl parameter faieled in auto cal. tried to reinstall. Not Ok. PER made and SC replaced. Calibrated and working fine. | Sc flushed and SP reset. reinstalled. Clote foung. Sample not aspirating properly. SC Cl parameter faieled in auto cal. tried to reinstall. Not Ok. PER made and SC replaced. Calibrated and working fine. | SWO-66756 | No | Sc flushed and SP reset. reinstalled. Clote foung. Sample not aspirating properly. SC Cl parameter faieled in auto cal. tried to reinstall. Not Ok. PER made and SC replaced. Calibrated and working fine. | Sc flushed and SP reset. reinstalled. Clote foung. Sample not aspirating properly. SC Cl parameter faieled in auto cal. tried to reinstall. Not Ok. PER made and SC replaced. Calibrated and working fine.
ABL90 FLEX | Fehlersuche Kunde hat erst SC und SP erneuert Fließwege gespült SC erneuert Einlaufprozedur SW aktualisiert V3.1Mr6 Handbuch erneuert | Fehlersuche Kunde hat erst SC und SP erneuert Fließwege gespült SC erneuert Einlaufprozedur SW aktualisiert V3.1Mr6 Handbuch erneuert | SWO-66104 | No | Troubleshooting customer has first redesigned SC and SP flow paths rinsed SC renewed inlet procedure SW updated V3.1Mr6 manual renewed | Troubleshooting customer has first redesigned SC and SP flow paths rinsed SC renewed inlet procedure SW updated V3.1Mr6 manual renewed | Assisted Replacement | SC | Lac
ABL90 FLEX | Gerät überprüft und gereinigt, Fließwege kontrolliert und gespült, Abfallschlauch ersetzt, Hemolysator überprüft, Lösungspack erneuert | Gerät überprüft und gereinigt, Fließwege kontrolliert und gespült, Abfallschlauch ersetzt, Hemolysator überprüft, Lösungspack erneuert | SWO-66172 | No | Device checked and cleaned, flow paths checked and rinsed, waste hose replaced, hemolysator checked, solution pack renewed | Device checked and cleaned, flow paths checked and rinsed, waste hose replaced, hemolysator checked, solution pack renewed | Assisted Replacement | SP | Leak current
ABL90 | Several SCs' all having Temp errors. | Replaced Sensor Interface PCB. | CFDE6234B80471C6 | No | Several SCs' all having Temp errors. | Replaced Sensor Interface PCB. | Defect | HW | Sensor cassette PCB |  | temp
ABL90 |  | B | CFDE5081B843202E | No |  |  | Assisted Replacement | SC | Wetability
ABL90 FLEX | 更换新测试卡修复仪器 | 更换新测试卡修复仪器 | SWO-66403 | No | Replace the new test card repair instrument | Replace the new test card repair instrument | Assisted Replacement | SC | Termistor
ABL90 FLEX | 更换903-199一个正常。 | 更换903-199一个正常。 | SWO-65618 | No | Replace 903-199 one normal. | Replace 903-199 one normal. | Defect | HW | Cassette retainer
AQT90 FLEX |  | B | SWO-66041 | No |  |  | Defect | HW | Process Module | Optical Unit | 131
ABL815 FLEX |  | B | SWO-66042 | No |  |  | Defect | HW | pH/BG Module | Measuring chamber
ABL820 FLEX | Gerät überprüft, 14pol. Kommunikationskabel im AQC Modul erneuert. | Gerät überprüft, 14pol. Kommunikationskabel im AQC Modul erneuert. | SWO-66215 | No | Device checked, 14pol. communication cable in the AQC module. | Device checked, 14pol. communication cable in the AQC module. | Defect | HW | AutoCheck Module | FPC Cable
ABL825 FLEX | Examined auto-check and inlet module. | Replaced 14 pole FPC cable. Verified operation of ABL825. | CFDE05DBB811D0C4 | No | Examined auto-check and inlet module. | Replaced 14 pole FPC cable. Verified operation of ABL825. | Defect | HW | AutoCheck Module | FPC Cable
ABL825 FLEX | Neue Elektrode eingebaut | neue Membrane aus einer anderen Charge aufgezogen. Kontakte und Messkammer gereinigt. Mehrfach Kalibriert. | SWO-66204 | No | New electrode installed | new membrane from another batch. contacts and measuring chamber. Multiple calibrated. | Assisted Replacement | Electrode | Electrode | pCO2
ABL825 FLEX | Ampullen werden nicht erkannt. Es sind die QCs doppelt eingelesen gewesen. Alle Spiegel grereinigt. ABL neu gestartet. Alle Ampullen werden wieder eingelesen. Liquid error vorhanden. Lufteinlass gespült. Liquid error verschwunden. Spüllösung erneuert. Pumpenkal und LS Justage durchgeführt. Mehrere kalibrierungen durchgeführt. Ca++ neu membraniert da Probleme bei QC Nochmals Kalibrierungen durchgeführt. QCs für alle vorhandenen Level durchgeführt. ABL wieder Messbereit. | Ampullen werden nicht erkannt. Es sind die QCs doppelt eingelesen gewesen. Alle Spiegel grereinigt. ABL neu gestartet. Alle Ampullen werden wieder eingelesen. Liquid error vorhanden. Lufteinlass gespült. Liquid error verschwunden. Spüllösung erneuert. Pumpenkal und LS Justage durchgeführt. Mehrere kalibrierungen durchgeführt. Ca++ neu membraniert da Probleme bei QC Nochmals Kalibrierungen durchgeführt. QCs für alle vorhandenen Level durchgeführt. ABL wieder Messbereit. | SWO-66193 | No | Ampoules are not detected. The QCs have been read in twice. All mirrors grereinigt. ABL. All ampoules are read in again. Liquid error present. Air inlet rinsed. Liquid error disappeared. Rinsing solution renewed. pump scale and LS adjustment. Several calibrations performed. Ca++ re-membranes as problems with QC performed calibrations again. QCs for all existing levels. ABL ready to measure again. | Ampoules are not detected. The QCs have been read in twice. All mirrors grereinigt. ABL. All ampoules are read in again. Liquid error present. Air inlet rinsed. Liquid error disappeared. Rinsing solution renewed. pump scale and LS adjustment. Several calibrations performed. Ca++ re-membranes as problems with QC performed calibrations again. QCs for all existing levels. ABL ready to measure again. | Cleaned | HW | AutoCheck Module | Barcode reader
ABL825 FLEX | Pumpenschlauch pH BG-Modul getauscht (KD-Bestand) Proteinentferner Programm Ref.Elektrode membraniert | Pumpenschlauch pH BG-Modul getauscht (KD-Bestand) Proteinentferner Programm Ref.Elektrode membraniert | SWO-66143 | No | Pump hose pH BG module exchanged (KD stock) protein remover program Ref.Electrode membraneed | Pump hose pH BG module exchanged (KD stock) protein remover program Ref.Electrode membraneed | Assisted Replacement | Electrode | Membrane | pCo2
ABL815 FLEX | Tried recovering file. Could not. Found huge whole in inlet tube. Inlet probe bent. | Installed windows and Au software. Replaced inlet tube and probe. Cleaned analyzer. Performed necessary calibrations. provided customer with some spare parts to help us troubleshoot when there are problems. PM kit, and pump rotors. | CFDE3437B8431266 | No | Tried recovering file. Could not. Found huge whole in inlet tube. Inlet probe bent. | Installed windows and Au software. Replaced inlet tube and probe. Cleaned analyzer. Performed necessary calibrations. provided customer with some spare parts to help us troubleshoot when there are problems. PM kit, and pump rotors. | Defect | HW | Inlet module | Inlet tube
ABL80 FLEX BASIC | 更换两张测试卡正常。 | 更换两张测试卡正常。 | SWO-66312 | No | Replacing the two test cards is ok. | Replacing the two test cards is ok.
ABL80 FLEX BASIC | 更换新测试卡 | 更换新测试卡 | SWO-66513 | No | Replace a new test card | Replace a new test card
ABL80 FLEX BASIC | cleaned Sc with distilled water and also suggested not to clean SC with hypochloride solution.Run QC solutions and cross checked. results are coming in range. | cleaned Sc with distilled water and also suggested not to clean SC with hypochloride solution.Run QC solutions and cross checked. results are coming in range. | SWO-66669 | No | cleaned Sc with distilled water and also suggested not to clean SC with hypochloride solution. Run QC solutions and cross checked. results are coming in range. | cleaned Sc with distilled water and also suggested not to clean SC with hypochloride solution. Run QC solutions and cross checked. results are coming in range.
ABL80 CO-OX | Ran oxi test in service cavatation error- New hemolyzer on order | B | CFDF799B2E04AB9A | No | Ran oxi test in service cavatation error- New hemolyzer on order
ABL90 |  | B | CFE3265F90080F7F | No |  |  | Assisted Replacement | SP | pO2
ABL90 | Subnet Mask did not have the correct address, and RALS did not have the correct IP address. | Corrected the Subnet Mask and RALS IP adresses. | CFDF15ED2E2C3E3F | No | Subnet Mask did not have the correct address, and RALS did not have the correct IP address. | Corrected the Subnet Mask and RALS IP adresses. | Setup | HIS/LIS
ABL90 FLEX | Eintrag für fixe Zielwerte ergänzt, Eintrag gespeichert mit ewfmgr, Neustart, Lösungspack ersetzt, Probelauf | Eintrag für fixe Zielwerte ergänzt, Eintrag gespeichert mit ewfmgr, Neustart, Lösungspack ersetzt, Probelauf | SWO-66477 | No | Added entry for fixed target values, entry saved with ewfmgr, restart, solution pack replaced, test run | Added entry for fixed target values, entry saved with ewfmgr, restart, solution pack replaced, test run | Setup | Roma.ini | Fixed assigned values
ABL90 | Software upgrade. | Install software upgrade version 3.1 MR7. | CFDF51FD2E2197C8 | Yes | Software upgrade. | Install software upgrade version 3.1 MR7. | Re classified RMED
ABL90 | Software upgrade. | Install software upgrade version 3.1 MR7. | CFDF57B32E3CC8C7 | Yes | Software upgrade. | Install software upgrade version 3.1 MR7. | Re classified RMED
ABL90 FLEX PLUS |  | B | CFDF202E2E049970 | No |  |  | Assisted Replacement | SP | - |  | 377
ABL5 | Elektrode neu bezogen, Gerät kalibriert, | Elektrode neu bezogen, Gerät kalibriert, | SWO-66604 | No | electrode re-covered, device calibrated, | electrode re-covered, device calibrated,
ABL90 FLEX | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-66528 | No | Repair after replacing a new test card | Repair after replacing a new test card | Assisted Replacement | SC | Sensitivity
ABL90 FLEX | 更换测试卡正常。 | 更换测试卡正常。 | SWO-66311 | No | The replacement test card is OK. | The replacement test card is OK. | Assisted Replacement | SC | Reference
ABL835 FLEX |  | B | CFDF4FB72E56E08E | No |  |  | Defect | HW | Oxi module | Spectrometer
ABL820 FLEX | .BONNIE KELLER AND I CHECKED ALL CALIBRATION SETTINGS ANALYZER IS GETTING MANY MESSAGES THAT SAY PH BG CALIBRATION UPGRADED DUE TO DRIFT IBN PH | EXPLAINED TO THE CUSTOMER THAT THE 30 MINUTE PH BG CALIBRATION THAT RUNS EVERY 30 MINUTES USES RINSE SOLUTION.PULLED DATA LOGS TO SEND TO COPENHAGEN TO DETERMINE WHETHER OR NOT TOO MUCH CAL 1 SOLUTION IS BEING USED | CFDF4ACB2E53203B | Yes | . BONNIE KELLER AND I CHECKED ALL CALIBRATION SETTINGS ANALYZER IS GETTING MANY MESSAGES THAT SAY PH BG CALIBRATION UPGRADED DUE TO DRIFT IBN PH | EXPLAINED TO THE CUSTOMER THAT THE 30 MINUTE PH BG CALIBRATION THAT RUNS EVERY 30 MINUTES USES RINSE SOLUTION. PULLED DATA LOGS TO SEND TO COPENHAGEN TO DETERMINE WHETHER OR NOT TOO MUCH CAL 1 SOLUTION IS BEING USED | Re classified RMED
ABL800 BASIC | zum Test einen anderen Bildschirm angeschlossen, keine Veränderung, Gerät läuft im Moment, Daten und Konfig. gezogen um eine neue Festplatte zu bespielen, | Der Kunde bittet um einen Service. | SWO-66605 | No | for testing another screen connected, no change, device running at the moment, data and config. pulled to play a new hard drive, | The customer asks for a service. | Defect | HW | DMS_PC | HDD
ABL825 FLEX | Drucker erneuert. Ausdruck in Ordnung. | Drucker erneuert. Ausdruck in Ordnung. | SWO-66598 | No | Printer renewed. Expression in order. | Printer renewed. Expression in order. | Defect | HW | Printer module | Printer head
ABL835 FLEX | Endkeys installiert | Endkeys installiert | SWO-66521 | Yes | Endkeys installed | Endkeys installed | Re classified RMED
ABL827 FLEX | Mac Adresse ausgelesen, Verbindungseinstellungen geprüft, Meßkammer gesäubert, Ref. Membran ersetzt | Mac Adresse ausgelesen, Verbindungseinstellungen geprüft, Meßkammer gesäubert, Ref. Membran ersetzt | SWO-66522 | No | Mac address read out, connection settings checked, measuring chamber cleaned, ref. Membrane replaced | Mac address read out, connection settings checked, measuring chamber cleaned, ref. Membrane replaced | Setup | HIS/LIS
ABL825 FLEX | Erneuerung Abfallverschlauchung unter Einlassdichtung - Erneuerung Nippel Kit für Einlassmodul - Mehrere Messungen ohne Flüssigkeitsaustritt - Funktionsüberprüfung etc. | Erneuerung Abfallverschlauchung unter Einlassdichtung - Erneuerung Nippel Kit für Einlassmodul - Mehrere Messungen ohne Flüssigkeitsaustritt - Funktionsüberprüfung etc. | SWO-66412 | No | Renewal of waste hoses under inlet seal - Renewal nipple kit for inlet module - Multiple measurements without liquid leakage - Function check etc. | Renewal of waste hoses under inlet seal - Renewal nipple kit for inlet module - Multiple measurements without liquid leakage - Function check etc. | Defect | HW | Inlet module | Leaking | Nipple kit
AQT90 FLEX |  | B | CFE0493DA402BE2B | No |  |  | Assisted Replacement | Consumables | Test kit | Cartridge, TnT | TnT
TCM/Tosca |  | B | CFDF1D832E42278B | No |  |  | Defect | HW | Touch Screen | Touch Screen
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-66568 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-66512 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL80 FLEX BASIC | 更换测试卡正常。 | 更换测试卡正常。 | SWO-66712 | No | The replacement test card is OK. | The replacement test card is OK.
ABL80 FLEX BASIC | 更换新测试卡 | 更换新测试卡 | SWO-66926 | No | Replace a new test card | Replace a new test card
ABL80 FLEX CO-OX | kompletten Wartungssatz verbaut Anschlüsse auf manifold auf korrekten Sitz überprüft, mehrere Kal, QC | kompletten Wartungssatz verbaut Anschlüsse auf manifold auf korrekten Sitz überprüft, mehrere Kal, QC | SWO-66884 | No | Complete maintenance kit installed connections on manifold checked for correct fit, several Kal, QC | Complete maintenance kit installed connections on manifold checked for correct fit, several Kal, QC
ABL90 FLEX | Lösungspack geliefert und installiert. | Lösungspack geliefert und installiert. | SWO-66847 | Yes | solution pack is delivered and installed. | solution pack is delivered and installed. | Re classified RMED | Logistical upSWO
ABL90 FLEX | Remplacé le sensor interface PCB e.o garantie Remplacé le kit an. + l'aiguille d'aspiration e.o garantie Calibration e.o contrôle de qualité e.o | Remplacé le sensor interface PCB e.o garantie Remplacé le kit an. + l'aiguille d'aspiration e.o garantie Calibration e.o contrôle de qualité e.o | SWO-66539 | No | Replaced the sensor interface PCB e.o warranty Replaced the an kit. E.o guaranteed e.o calibration e.o quality control e.o | Replaced the sensor interface PCB e.o warranty Replaced the an kit. E.o guaranteed e.o calibration e.o quality control e.o | Defect | HW | Sensor cassette PCB |  | temp
ABL90 FLEX | Bei erneuter tHb-Kalibrierung kein Fehler feststellbar. Fehler lag vermutlich am Lösungpack. | Bei erneuter tHb-Kalibrierung kein Fehler feststellbar. Fehler lag vermutlich am Lösungpack. | SWO-66613 | No | No error can be detected when tHb calibration is re-calibrated. Errors were probably due to the solution pack. | No error can be detected when tHb calibration is re-calibrated. Errors were probably due to the solution pack. | Adjustment | tHb cal
ABL90 FLEX | Beim Eintreffen steht der ABL messbereit mit Lactat-Fehler. Diverse Problem mit Station besprochen . SC erneuert, da SP mehrfach erneuert wurde Station überwach Einlaufphase | Beim Eintreffen steht der ABL messbereit mit Lactat-Fehler. Diverse Problem mit Station besprochen . SC erneuert, da SP mehrfach erneuert wurde Station überwach Einlaufphase | SWO-66680 | No | Upon arrival, the ABL is ready for measurement with lactate errors. Various problem discussed with station . SC renewed as SP has been renewed several times station overwake run-in phase | Upon arrival, the ABL is ready for measurement with lactate errors. Various problem discussed with station . SC renewed as SP has been renewed several times station overwake run-in phase | Assisted Replacement | SC | Lac
ABL90 | Status on lactate reading above upperlimit of 3000. | Replaced sensor cassette. Installed 3.1 MR7 software. Replaced solution pack to correct S9030 qc readings. | CFE08FF7A44CDC04 | No | Status on lactate reading above upperlimit of 3000. | Replaced sensor cassette. Installed 3.1 MR7 software. Replaced solution pack to correct S9030 qc readings. | Assisted Replacement | SC | Lac
ABL90 FLEX PLUS | J'ai débouché le système avec le seringue et eau chaude. J'ai trouvé un bouchage dans le aiguille : remplacer. Test en service : OK Upgrade software V3.1MR7 Calibrations et CQ : OK J'ai laisser le nouveau pack de solution et le cassette sensor chez le client. | J'ai débouché le système avec le seringue et eau chaude. J'ai trouvé un bouchage dans le aiguille : remplacer. Test en service : OK Upgrade software V3.1MR7 Calibrations et CQ : OK J'ai laisser le nouveau pack de solution et le cassette sensor chez le client. | SWO-66662 | No | I opened the system with the syringe and hot water. I found a plug in the needle: replace. Test in service: OK Upgrade software V3.1MR7 Calibrations and CQ: OK I left the new solution pack and the sensor cassette at the customer's home. | I opened the system with the syringe and hot water. I found a plug in the needle: replace. Test in service: OK Upgrade software V3.1MR7 Calibrations and CQ: OK I left the new solution pack and the sensor cassette at the customer's home. | Assisted Replacement | SC | Wetability
ABL800 BASIC | cleaned the pH electrode, performed the protein removal and did the calibration for PUMP,LS now instrument working. | cleaned the pH electrode, performed the protein removal and did the calibration for PUMP,LS now instrument working. | SWO-66717 | No | cleaned the pH electrode, performed the protein removal and did the calibration for PUMP,LS now instrument working. | cleaned the pH electrode, performed the protein removal and did the calibration for PUMP,LS now instrument working. | Cleaned | Fluidic pathway
ABL825 FLEX | 更换945-619修复 | 更换945-619修复 | SWO-66287 | No | Replacement 945-619 repair | Replacement 945-619 repair | Assisted Replacement | Electrode | Electrode | Lac
ABL835 FLEX |  | B | CFE07AA1A4233CD2 | Yes |  |  | Re classified RMED | Training
ABL815 FLEX | alle Pumpenschläuche ersetzt aus einem Hjw kit ersetzt, | alle Pumpenschläuche ersetzt aus einem Hjw kit ersetzt, | SWO-66684 | No | replaced all pump hoses from one Hjw kit, | replaced all pump hoses from one Hjw kit, | Assisted Replacement | HW | El/Met Module | Pump tubing
ABL837 FLEX |  | B | CFE06372A429745C | No |  |  | Defect | HW | Printer module | Printer head
ABL825 FLEX | neue Membrane aus einer anderen Charge aufgezogen. Kontakte und Messkammer gereinigt. Mehrfach Kalibriert. | neue Membrane aus einer anderen Charge aufgezogen. Kontakte und Messkammer gereinigt. Mehrfach Kalibriert. | SWO-66653 | No | new membrane from another batch. contacts and measuring chamber. Multiple calibrated. | new membrane from another batch. contacts and measuring chamber. Multiple calibrated. | Assisted Replacement | Electrode | Membrane | pCO2
TCM TOSCA | Remambraine required. Tried to calibrate. old mambraine calibrated. A study done of 15 min. Readings prper. Dr's Qustion noted. come back soon with prper answeres. | Remambraine required. Tried to calibrate. old mambraine calibrated. A study done of 15 min. Readings prper. Dr's Qustion noted. come back soon with prper answeres. | SWO-66765 | No | Remambraine required. Tried to calibrate. old mambraine calibrated. A study done of 15 min. Readings prper. Dr's Qustion noted. come back soon with prper answeres. | Remambraine required. Tried to calibrate. old mambraine calibrated. A study done of 15 min. Readings prper. Dr's Qustion noted. come back soon with prper answeres.
ABL80 FLEX BASIC | 更换测试卡正常。 | 更换测试卡正常。 | SWO-66734 | No | The replacement test card is OK. | The replacement test card is OK.
ABL90 FLEX | Fehlersuche unter Zuhilfenahme der Systemprogramme Reparatur mit Einsatz eines Jahreswartungskits durchgeführt Startphase mehrere Kalibrationen nötig (Lac Empfindlichkeit zu gering) Kal- und QC | Fehlersuche unter Zuhilfenahme der Systemprogramme Reparatur mit Einsatz eines Jahreswartungskits durchgeführt Startphase mehrere Kalibrationen nötig (Lac Empfindlichkeit zu gering) Kal- und QC | SWO-66770 | No | Troubleshooting with the help of the system programs Repair with use of an annual maintenance kit carried out start phase several calibrations necessary (Lac sensitivity too low) Cal and QC | Troubleshooting with the help of the system programs Repair with use of an annual maintenance kit carried out start phase several calibrations necessary (Lac sensitivity too low) Cal and QC | Defect | HW | Tubes
ABL90 FLEX | 更换测试卡正常。 | 更换测试卡正常。 | SWO-66766 | No | The replacement test card is OK. | The replacement test card is OK. | Assisted Replacement | SC | Cl | Cal
AQT90 FLEX | 重新定标正常。 | 重新定标正常。 | SWO-66505 | No | Recalibration is normal. | Recalibration is normal. | Assisted Replacement | Consumables | Test kit | Cartridge, CKMB | CKMB - 167
ABL90 FLEX | 更换903-183,636-461各一个正常。 | 更换903-183,636-461各一个正常。 | SWO-66573 | No | Replace 903-183, 636-461 each normal. | Replace 903-183, 636-461 each normal. | Defect | HW | Sensor cassette PCB
AQT90 FLEX | 更换903-236，903-364各一个正常。 | 更换903-236，903-364各一个正常。 | SWO-66746 | No | Replace 903-236, 903-364 each normal. | Replace 903-236, 903-364 each normal. | Defect | HW | Process Module | Optical Unit | 131
ABL825 FLEX | 更换945-603 一个正常。 | 更换945-603 一个正常。 | SWO-66517 | No | Replace 945-603 one normal. | Replace 945-603 one normal. | Assisted Replacement | Electrode | Electrode | Reference
ABL80 FLEX BASIC | 更换两个试剂包正常。 | 更换两个试剂包正常。 | SWO-67572 | No | Replacing both reagent kits is normal. | Replacing both reagent kits is normal.
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-67346 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL80 FLEX BASIC | 更换902-972,903-334修复 | 更换902-972,903-334修复 | SWO-66760 | No | Replacement 902-972,903-334 repair | Replacement 902-972,903-334 repair
ABL80 FLEX BASIC |  | B | SWO-66802 | No
ABL80 FLEX BASIC | 更换测试卡正常。 | 更换测试卡正常。 | SWO-66823 | No | The replacement test card is OK. | The replacement test card is OK.
ABL80 OSM | Some doctor have complained the THB seems to ready low. | Ran vk kit Downloaded the datalog and added to the Techline case. | CFE4A6DE061344A8 | No | Some doctor have complained the THB seems to ready low. | Ran vk kit Downloaded the datalog and added to the Techline case.
ABL80 FLEX BASIC | 更换903-334修复 | 更换903-334修复 | SWO-66785 | No | Replacement 903-334 repair | Replacement 903-334 repair
ABL80 OSM | THB failed during proficiency testing. | Ran vk kit Downloaded datalogs and open techline case. | CFE4A4CC061044A3 | No | THB failed during proficiency testing. | Ran vk kit Downloaded datalogs and open techline case.
ABL80 FLEX CO-OX | Zählerstand ablesen Für Conworks Anbindung vorbereiten, Eintragungen In Windows vorgenommen | Zählerstand ablesen Für Conworks Anbindung vorbereiten, Eintragungen In Windows vorgenommen | SWO-66959 | No | Reading Counter Reading For Conworks Connection Prepare, Entries Made In Windows | Reading Counter Reading For Conworks Connection Prepare, Entries Made In Windows
ABL90 FLEX | Fehlersuche, Kontakte gerichtet, LP erneut eingelegt, Startphase abgewartet, Kal. = div. Fehler, LP gew., Kal. weiter pCO2-Fehler, vorhandene SC neu initialisiert -> Glu-, Lac- u. pCO2-Fehler in Kal., neue SC initialisiert, Kal. i.O., QC = Lac-Fehler Level B, weitere Kal., Kal. + QC i.O. | Fehlersuche; Kontakte gerichtet; LP erneut eingelegt; Startphase abgewartet; Kal. = div. Fehler; LP gew.; Kal. weiter pCO2-Fehler; vorhandene SC neu initialisiert -> Glu-, Lac- u. pCO2-Fehler in Kal.; neue SC initialisiert; Kal. i.O.; QC = Lac-Fehler Level B; weitere Kal.; Kal. + QC i.O. | SWO-66947 | No | Troubleshooting; contacts directed; LP re-inserted; start-up phase awaited; Kal. = div. error; LP. Kal. further pCO2 errors; existing SC reinitialized -> Glu, Lac and pCO2 errors in Cal.; new SC initialized; Kal. i.O.; QC = Lac Error Level B; more Kal.; Kal. + QC in the first place | Troubleshooting; contacts directed; LP re-inserted; start-up phase awaited; Kal. = div. error; LP. Kal. further pCO2 errors; existing SC reinitialized -> Glu, Lac and pCO2 errors in Cal.; new SC initialized; Kal. i.O.; QC = Lac Error Level B; more Kal.; Kal. + QC in the first place | Adjustment | HW | SP PCB
ABL90 FLEX | Beim Eintreffen Leckage erkannt LP lässt sich Problemlos tauschen Mehrmals eingesetzt und ausgeworfen Einlassdichtung getauscht Im Service Menü Fluid Transport durchgegangen OK Analysator kalibriert und überprüft QC Messung durchgeführt | Beim Eintreffen Leckage erkannt LP lässt sich Problemlos tauschen Mehrmals eingesetzt und ausgeworfen Einlassdichtung getauscht Im Service Menü Fluid Transport durchgegangen OK Analysator kalibriert und überprüft QC Messung durchgeführt | SWO-66877 | No | Upon arrival leakage detected LP can be easily exchanged multiple times inserted and ejected inlet seal exchanged In the service menu Fluid Transport passed through OK analyzer calibrated and checked QC measurement performed | Upon arrival leakage detected LP can be easily exchanged multiple times inserted and ejected inlet seal exchanged In the service menu Fluid Transport passed through OK analyzer calibrated and checked QC measurement performed | Unknown | SP Activation
ABL90 | Po2 having inhomogeneous sample errors. All the other parameters are ok. Cals and QC look good for po2 as well. Checked all tubing. | Flushed the analyzer multiple times. Ran all service checks. Cleaned inlet gasket.Upgraded the SW to 3.1 MR7. Ran multiple samples for testing. Ran Both Venous and arterial samples. | CFE45702064A4A00 | No | Po2 having inhomogeneous sample errors. All the other parameters are ok. Cals and QC look good for po2 as well. Checked all tubing. | Flushed the analyzer multiple times. Ran all service checks. Cleaned inlet gasket. Upgraded the SW to 3.1 MR7. Ran multiple samples for testing. Ran Both Venous and arterial samples. | Cleaned | Fluidic pathway
ABL90 |  | B | CFE45A7906178173 | No |  |  | Defect | HW | Wet section Control PCB
ABL90 FLEX PLUS | 3 Gruppen Anwender und Ärzte Nacheinweisung in ABL90Flex Plus | 3 Gruppen Anwender und Ärzte Nacheinweisung in ABL90Flex Plus | SWO-67043 | Yes | 3 Groups Users and Doctors Follow-up in ABL90Flex Plus | 3 Groups Users and Doctors Follow-up in ABL90Flex Plus | Re classified RMED | Training
ABL90 FLEX | 更换试剂包正常。 | 更换试剂包正常。 | SWO-67105 | No | The replacement kit is OK. | The replacement kit is OK. | User error | SP Activation
ABL90 FLEX |  | B | SWO-66830 | No |  |  | Defect | HW | Wet section Control PCB
ABL90 FLEX | 更换新测试卡 | 更换新测试卡 | SWO-66917 | No | Replace a new test card | Replace a new test card | Assisted Replacement | SC | Glu
ABL90 FLEX | 更换试剂包 | 更换试剂包 | SWO-66929 | No | Replace the reagent kit | Replace the reagent kit | Assisted Replacement | SP | Pump cal error
ABL90 FLEX | 更换903-199修复 | 更换903-199修复 | SWO-66779 | No | Replacement 903-199 repair | Replacement 903-199 repair | Defect | HW | Cassette retainer |  | temp
AQT90 FLEX |  | B | SWO-66796 | No |  |  | Defect | HW | Process Module | Optical Unit | 129
AQT90 FLEX | 更换新测试卡修复仪器 | 更换新测试卡修复仪器 | SWO-66937 | No | Replace the new test card repair instrument | Replace the new test card repair instrument | Assisted Replacement | Consumables | SP | SP | 167
AQT90 FLEX | 更换新测试卡修复仪器 | 更换新测试卡修复仪器 | SWO-66939 | Yes | Replace the new test card repair instrument | Replace the new test card repair instrument | Re classified RMED | Double entry
AQT90 FLEX |  | B | SWO-66807 | No |  |  | Defect | HW | Analyzer control Module | CPU
AQT90 FLEX |  | B | SWO-66808 | No |  |  | Defect | HW | Process Module | Optical Unit | 131
ABL825 FLEX |  | B | CFE462590617FDB7 | Yes |  |  | Re classified RMED | Installation
ABL825 FLEX | E1001 ern., Ref. Connector ern., Gerät kalibriert, QC´s durchgef. | B | SWO-66893 | No | E1001 ern., Ref. Connector ern., device calibrated, QC's through. |  | Assisted Replacement | Electrode | Electrode | Reference
ABL810 FLEX | Fehlersuche durchgeführt. Elektrode erneuert. Luftdruck genau eingestellt. Mehrere Kal und QCs durchlaufen lassen. Alle Werte sind zwar weiter von der unteren RiliBÄK Grenze entfernt als vorher, jedoch befinden sich die Werte noch im unteren Drittel. Tendenz ist steigend. | Fehlersuche durchgeführt. Elektrode erneuert. Luftdruck genau eingestellt. Mehrere Kal und QCs durchlaufen lassen. Alle Werte sind zwar weiter von der unteren RiliBÄK Grenze entfernt als vorher, jedoch befinden sich die Werte noch im unteren Drittel. Tendenz ist steigend. | SWO-66986 | No | Troubleshooting. electrode renewed. air pressure is precisely adjusted. Allow to pass through several Kal and QCs. All values are further away from the lower RiliBÄK boundary than before, but the values are still in the lower third. The trend is rising. | Troubleshooting. electrode renewed. air pressure is precisely adjusted. Allow to pass through several Kal and QCs. All values are further away from the lower RiliBÄK boundary than before, but the values are still in the lower third. The trend is rising. | Assisted Replacement | Electrode | Electrode | pCO2
ABL800 BASIC | Fehlersuche, Kal.- QC- und Aktivitätendateien über die letzten Wochen überprüft, Spüllösung am 11.11.15 und 21.11.15 leergelaufen->nach CA 8h erneuert, pO2 und pCO2 neu membraniert und Messkammern gereinigt,(Verschmutzung in pH Messkammer), QC Verlauf überprüft ok, am 22.11.15 wurde keine QC durchgeführt, Karussell neu und korrekt eingelegt->danach ok, Nadel ok, Ventile im Service angesteuert ok, am 21.11.15 Probenabbrüche mit ? markiert->ungenügend Probe Probenfehler bzw. Inhomogene Probe Messung instabil und andere Probenfehler->Einlassdichtung stark verschmutzt erneuert, mehrere Kal., QCs ok, den Anwender auf die Präanalytik die defekte Einlassdichtung und Durchführung der Erneuerungen der Betriebsmittel hingewiesen System-und Funktionskontrolle i.O. | Fehlersuche, Kal.- QC- und Aktivitätendateien über die letzten Wochen überprüft, Spüllösung am 11.11.15 und 21.11.15 leergelaufen->nach CA 8h erneuert, pO2 und pCO2 neu membraniert und Messkammern gereinigt,(Verschmutzung in pH Messkammer), QC Verlauf überprüft ok, am 22.11.15 wurde keine QC durchgeführt, Karussell neu und korrekt eingelegt->danach ok, Nadel ok, Ventile im Service angesteuert ok, am 21.11.15 Probenabbrüche mit ? markiert->ungenügend Probe Probenfehler bzw. Inhomogene Probe Messung instabil und andere Probenfehler->Einlassdichtung stark verschmutzt erneuert, mehrere Kal., QCs ok, den Anwender auf die Präanalytik die defekte Einlassdichtung und Durchführung der Erneuerungen der Betriebsmittel hingewiesen System-und Funktionskontrolle i.O. | SWO-67020 | No | Troubleshooting, Kal.- QC and activity files checked over the last weeks, rinsing solution on 11.11.15 and 21.11.15 empty-run->renewed after CA 8h, pO2 and pCO2 re-logged and measuring chambers cleaned,(contamination in pH measuring chamber), QC course checked ok, on 22.11.15 no QC was performed, carousel new and correct inserted->after that ok, needle ok, valves in the service controlled ok, on 21.11.15 sample aborts with ? marked->insufficient sample sample error or inhomogeneous sample measurement unstable and other sample error->inlet seal strongly polluted renewed, several cal., QCs ok, the user on the pre-analysis the defective inlet seal and carrying out the renewals of the Equipment indicated system and function control as the case | Troubleshooting, Kal.- QC and activity files checked over the last weeks, rinsing solution on 11.11.15 and 21.11.15 empty-run->renewed after CA 8h, pO2 and pCO2 re-logged and measuring chambers cleaned,(contamination in pH measuring chamber), QC course checked ok, on 22.11.15 no QC was performed, carousel new and correct inserted->after that ok, needle ok, valves in the service controlled ok, on 21.11.15 sample aborts with ? marked->insufficient sample sample error or inhomogeneous sample measurement unstable and other sample error->inlet seal strongly polluted renewed, several cal., QCs ok, the user on the pre-analysis the defective inlet seal and carrying out the renewals of the Equipment indicated system and function control as the case | Defect | HW | Inlet module | Inlet gasket
ABL815 FLEX | Einlassnadel gerichtet und Einlassdichtung erneuert | Einlassnadel gerichtet und Einlassdichtung erneuert | SWO-66993 | No | Inlet needle directed and inlet seal renewed | Inlet needle directed and inlet seal renewed | Defect | HW | Inlet module | Probe bend
ABL815 FLEX | Kein Fehler vom System feststellbar, Anwender Aktionen | Kein Fehler vom System feststellbar, Anwender Aktionen | SWO-67140 | No | No error detected by the system, user actions | No error detected by the system, user actions | Unknown
ABL835 FLEX | Attempted to run complete oxi test during service testing of analyzer, and discovered that halogen lamp intensity was below tolerence limits. Attempted to adjust lamp unit 3 times, with each time the lamp being quite high or quite low during the repeated complete oxi test. During testing, also observed pH and Ca sensitivity being above tolerence limit. pH and lac failing QC. Replaced reference and lac membrane, cal 1 and cal2 bottles. | Replaced lamp unit. Contacted RAME TechSupport | CFE452930645F885 | No | Attempted to run complete oxi test during service testing of analyzer, and discovered that halogen lamp intensity was below tolerence limits. Attempted to adjust lamp unit 3 times, with each time the lamp being quite high or quite low during the repeated complete oxi test. During testing, also observed pH and Ca sensitivity being above tolerence limit. pH and lac failing QC. Replaced reference and lac membrane, cal 1 and cal2 bottles. | Replaced lamp unit. Contacted RAME TechSupport | Defect | HW | Oxi module | Lamp unit
TCM/Tosca |  | B | CFE46BA1063CED57 | No
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-67068 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL80 OSM | Performed liquid sensor calibration on diagnostics and value was high out of range. Suspect is blockage in hemolyzer or dirty liquid sensor. | Flushed hemolyzer with protein remover and cleaned liquid sensor, but sensor was shorted from getting wet and a new sensor needs to be ordered. | CFE6930EF256968C | No | Performed liquid sensor calibration on diagnostics and value was high out of range. Suspect is blockage in hemolyzer or dirty liquid sensor. | Flushed hemolyzer with protein remover and cleaned liquid sensor, but sensor was shorted from getting wet and a new sensor needs to be ordered.
ABL80 FLEX BASIC | SC Flushed and calibrated, Calibration time changed 8 hours to 4 hours, Rinse button enabled on main screen. Sample performed for 5 times. And found all readings in range. Analyzer report able. Working fine. | SC Flushed and calibrated, Calibration time changed 8 hours to 4 hours, Rinse button enabled on main screen. Sample performed for 5 times. And found all readings in range. Analyzer report able. Working fine. | SWO-67627 | No | SC Flushed and calibrated, Calibration time changed 8 hours to 4 hours, Rinse button enabled on main screen. Sample performed for 5 times. And found all readings in range. Analyzer report able. Working fine. | SC Flushed and calibrated, Calibration time changed 8 hours to 4 hours, Rinse button enabled on main screen. Sample performed for 5 times. And found all readings in range. Analyzer report able. Working fine.
ABL90 |  | B | CFE596827C4564C8 | No |  |  | Assisted Replacement | SC | pCO2
ABL90 FLEX | Gerät demontiert und Hemolyzer erneuert. - neuer Hemolyzer ist defekt und besteht die THB Kal nicht , meldet Oxi mismatch. - Kunde arbeitet erstmal ohne Oxi Parameter. Ersatzteil wird besorgt! | Gerät demontiert und Hemolyzer erneuert. - neuer Hemolyzer ist defekt und besteht die THB Kal nicht , meldet Oxi mismatch. - Kunde arbeitet erstmal ohne Oxi Parameter. Ersatzteil wird besorgt! | SWO-67534 | No | device and hemolyzer renewed. - new Hemolyzer is defective and does not exist the THB Kal, reports Oxi mismatch. - Customer works without Oxi parameters for the first time. Spare part is worried! | device and hemolyzer renewed. - new Hemolyzer is defective and does not exist the THB Kal, reports Oxi mismatch. - Customer works without Oxi parameters for the first time. Spare part is worried! | Unknown | Revisit required
ABL90 | Instrument was running 3.1 MR3. Will load new 3.1 MR6 software. | Installed system software version 3.1 MR6. Verified instrument operation. | CFE617AB7C389680 | No | Instrument was running 3.1 MR3. Will load new 3.1 MR6 software. | Installed system software version 3.1 MR6. Verified instrument operation. | Defect | SW | MR3
ABL90 FLEX | Fehlersuche am Analysator, unter Service Fluid Transport überprüft, Analysator im Service überprüft ok, Spülung wurde bereits durchgeführt, Einlassdichtung mit Halter defekt->erneuert, Leckagetest wurde bestanden, Starphase ok System-und Funktionskontrolle i.O. | Fehlersuche am Analysator, unter Service Fluid Transport überprüft, Analysator im Service überprüft ok, Spülung wurde bereits durchgeführt, Einlassdichtung mit Halter defekt->erneuert, Leckagetest wurde bestanden, Starphase ok System-und Funktionskontrolle i.O. | SWO-67223 | No | Troubleshooting on the analyzer, checked under Service Fluid Transport, Analyzer checked in service ok, flushing has already been performed, inlet seal with holder defective->renewed, leakage test has been passed, star phase ok system and function control i.O. | Troubleshooting on the analyzer, checked under Service Fluid Transport, Analyzer checked in service ok, flushing has already been performed, inlet seal with holder defective->renewed, leakage test has been passed, star phase ok system and function control i.O. | Defect | HW | Inlet Gasket
ABL90 FLEX | Lösungspack neu installiert, Serviceprogramm PumpenKal, Einlaßbereich geprüft, Startphase abgewartet, Kal u. QK ok | Lösungspack neu installiert, Serviceprogramm PumpenKal, Einlaßbereich geprüft, Startphase abgewartet, Kal u. QK ok | SWO-67277 | No | Solution pack newly installed, service program PumpsKal, intake area checked, start-up phase awaited, Kal & QK ok | Solution pack newly installed, service program PumpsKal, intake area checked, start-up phase awaited, Kal & QK ok | User error | SP Activation
ABL90 | Troubleshot analyzer in service, and the pressure test would not pass. | Replaced the inlet gasket, and this resolved the pressure test issue. | CFE577EC7C186ED1 | No | Troubleshot analyzer in service, and the pressure test would not pass. | Replaced the inlet gasket, and this resolved the pressure test issue. | Defect | HW | Inlet Gasket
ABL90 |  | B | CFE6DB8BF244CD8E | No |  |  | Defect | HW | Inlet Gasket
ABL90 |  | B | CFE5943C7C59B325 | No |  |  | Defect | HW | Sensor cassette Bay
ABL90 | Not connecting to Radiance due to MR6 software. | Loaded MR7 to resolve the problem | CFE5B2EA7C1D7C3C | No | Not connecting to Radiance due to MR6 software. | Loaded MR7 to resolve the problem | Defect | SW | Connectivity | MR6
ABL90 FLEX PLUS |  | B | CFE5A0E77C5ED0E3 | No |  |  | Defect | HW | Sensor Cassette PCB |  | temp
TCM CombiM |  | B | SWO-67119 | No
ABL90 | Verificação do estado da sonda e da entrada, sem problemas, limpeza da entrada e do circuito, establização, erros na gluc e lac, substituição da cassete, establização.e passagem do QC | B | CFE5EBB97C09A5DB | No | Checking the state of the probe and the inlet, without problems, cleaning the inlet and circuit, establishing, errors in Gluc and Lac, replacing the cassette, establishing. and QC Passage |  | Cleaned | Fluidic pathway
AQT90 FLEX |  | B | SWO-67091 | No |  |  | Defect | HW | Process Module | Optical Unit | 131
ABL825 FLEX | El Met målekammer udskiftet, slutafprøvet. | B | SWO-67290 | No | The El Met Measuring chamber replaced, final tested. |  | Defect | HW | El/Met Module | Measuring chamber
ABL800 BASIC | Seit letztem Service kein Neustart im Speicher, Netzteil zur Probe ersetzt, Kal und QC durchgeführt, Anwender informiert, Anwender beobachtet, lässt wenn möglich im Fehlerfall das Gerät an und meldet sich | Seit letztem Service kein Neustart im Speicher, Netzteil zur Probe ersetzt, Kal und QC durchgeführt, Anwender informiert, Anwender beobachtet, lässt wenn möglich im Fehlerfall das Gerät an und meldet sich | SWO-67245 | No | Since last service no restart in memory, power supply replaced to sample, Kal and QC performed, user informed, user observed, if possible in case of error the device on and logs in | Since last service no restart in memory, power supply replaced to sample, Kal and QC performed, user informed, user observed, if possible in case of error the device on and logs in | Defect | HW | Power Supply
ABL800 BASIC | Gerät überprüft, Meßkammer gesäubert, Elektrode abgewaschen und membraniert (0542), K und Ref membraniert, diverse Kal und QC o.k.. Baro kontrolliert | QC bei pCO2 weiter ausserhalb Herstellergrenzen. Kd. bittet um Rückruf durch H. Rotter. | SWO-67188 | No | Device checked, measuring chamber cleaned, electrode washed off and membraned (0542), K and ref membraneed, various Kal and QC o.k.. Baro controlled | QC at pCO2 further outside the manufacturer's limits. Kd. asks for a callback from H. Rotter. | Assisted Replacement | Electrode | Membrane | Reference
ABL800 BASIC | Replaced the Cal-II and Rinse soln bottles, run the 2-point calibration, run sample. Machine is working fine. Meet Naik sir to discuss for maintaining the stock at KEM hospital. | Replaced the Cal-II and Rinse soln bottles, run the 2-point calibration, run sample. Machine is working fine. Meet Naik sir to discuss for maintaining the stock at KEM hospital. | SWO-67910 | No | Replaced the Cal-II and Rinse soln bottles, run the 2-point calibration, run sample. Machine is working fine. Meet Naik sir to discuss for maintaining the stock at KEM hospital. | Replaced the Cal-II and Rinse soln bottles, run the 2-point calibration, run sample. Machine is working fine. Meet Naik sir to discuss for maintaining the stock at KEM hospital. | User error | Solution empty
ABL835 FLEX | Pco2 sensibilité basse 91% meme avec membrane neuve | Remplacé electrode de PCo2 | CFE58BDC7C0CBA1B | No | Pco2 low sensitivity 91% even with new membrane | Replaced electrode from PCo2 | Assisted Replacement | Electrode | Electrode | pCO2
ABL825 FLEX | Tried different lots of Ca++ membranes-No change | Changed all pump tubes and connector tubes Membrane the reference and Ca++ electrodes | CFE595567C520E7D | No | Tried different lots of Ca++ membranes-No change | Changed all pump tubes and connector tubes Membrane the reference and Ca++ electrodes | Assisted Replacement | Electrode | Membrane | Reference
ABL825 FLEX | Main board CPU 910 i.o Garantie Kalibration i.o QC i.o | Main board CPU 910 i.o Garantie Kalibration i.o QC i.o | SWO-67221 | No | Main Board CPU 910 I. A guarantee Kalibration i. o QC i. O | Main Board CPU 910 I. A guarantee Kalibration i. o QC i. O | Defect | HW | DMS_PC | CPU
AQT90 FLEX |  | B | CFE724B2F25EE0B0 | No |  |  | Cleaned | HW | Analyzer control Module | Barcode reader, display
ABL80 FLEX BASIC |  | B | SWO-67385 | No
ABL80 FLEX BASIC | 更换新测试卡修复仪器 | 更换新测试卡修复仪器 | SWO-67582 | No | Replace the new test card repair instrument | Replace the new test card repair instrument
ABL90 FLEX | Selbiges LP installiert, Gleicher Fehler, inzwischen nur noch 313 Zyklen im LP (gestern installiert), aufgef. Verschleißteile ersetzt, Testprogramme im Service durchgeführt, u.a. Pumpenkalibrierfehler, kein weiteres Lösungspack vorrätig, Abbruch, LP wird vom Kunden aus anderem Standort organisiert | Selbiges LP installiert, Gleicher Fehler, inzwischen nur noch 313 Zyklen im LP (gestern installiert), aufgef. Verschleißteile ersetzt, Testprogramme im Service durchgeführt, u.a. Pumpenkalibrierfehler, kein weiteres Lösungspack vorrätig, Abbruch, LP wird vom Kunden aus anderem Standort organisiert | SWO-67561 | No | Same LP installed, Same error, now only 313 cycles in the LP (installed yesterday), on. Replacement of wear parts, test programs carried out in service, among other things pump calibration errors, no further solution pack in stock, demolition, LP is organized by the customer from another location | Same LP installed, Same error, now only 313 cycles in the LP (installed yesterday), on. Replacement of wear parts, test programs carried out in service, among other things pump calibration errors, no further solution pack in stock, demolition, LP is organized by the customer from another location | Assisted Replacement | SP | Pump cal error
ABL90 FLEX | Fehlersuche, Einlaßdichtung erneuert, LP neu installiert, div. Test`s im Serviceprogr. durchgef., Kal. + QC i.O. | Fehlersuche; Einlaßdichtung erneuert; LP neu installiert; div. Test`s im Serviceprogr. durchgef.; Kal. + QC i.O. | SWO-67483 | No | Troubleshooting; Inlet seal renewed; LP reinstalled; div. Test it in the service progr. through it; Kal. + QC in the first place | Troubleshooting; Inlet seal renewed; LP reinstalled; div. Test it in the service progr. through it; Kal. + QC in the first place | Defect | HW | Inlet Gasket
ABL90 FLEX | Im Service Menü Drucküberprüfung und Pumpenkal fehlerhaft System durchgespült Schläuche auf Fließwege erneuert Einlasssonde erneuert Fluid Transport im Service Menü durchgeführt OK SC getauscht da Kal Fehler bei Lac Empfindlichkeit nur bei 84% Analysator kalibriert und überprüft QC Messung durchgeführt | Im Service Menü Drucküberprüfung und Pumpenkal fehlerhaft System durchgespült Schläuche auf Fließwege erneuert Einlasssonde erneuert Fluid Transport im Service Menü durchgeführt OK SC getauscht da Kal Fehler bei Lac Empfindlichkeit nur bei 84% Analysator kalibriert und überprüft QC Messung durchgeführt | SWO-67599 | No | In the service menu pressure check and pump scale faulty system flushed hoses on flow paths renewed inlet probe renewed fluid transport carried out in the service menu OK SC exchanged since Kal error at Lac sensitivity only calibrated at 84% analyzer and checked QC measurement performed | In the service menu pressure check and pump scale faulty system flushed hoses on flow paths renewed inlet probe renewed fluid transport carried out in the service menu OK SC exchanged since Kal error at Lac sensitivity only calibrated at 84% analyzer and checked QC measurement performed | Assisted Replacement | SP | Pump cal error
ABL90 FLEX | Erneut neuen Hemolyzer angeliefert. Gerät meldet immer noch Ableitströme. Fehler muss eine andere Ursache haben. Lösungspack und Sensorcassette überprüft. Inpendancetest mit anderer Sensorcassette im Service durchgeführt. - Fehler bleibt. Ersatzgerät besorgt und aufgestellt. - SN 090R0147N0009 Lösungspack und Sensorcassette installiert. Konfiguration angepasst. Kundengerät zur Überprüfung Reparatur mitgenommen. | Gerät demontiert und Hemolyzer erneuert. - neuer Hemolyzer ist defekt und besteht die THB Kal nicht , meldet Oxi mismatch. - Kunde arbeitet erstmal ohne Oxi Parameter. Ersatzteil wird besorgt! | SWO-67531 | No | New Hemolyzer delivered again. Device still reports discharge currents. Error must have a different cause. solution pack and sensor cassette checked. Inpendance test performed with other sensor cassette in the service. - Error remains. Replacement device is provided and set up. - SN 090R0147N0009 solution pack and sensor cassette installed. configuration. Customer device for verification repair. | device and hemolyzer renewed. - new Hemolyzer is defective and does not exist the THB Kal, reports Oxi mismatch. - Customer works without Oxi parameters for the first time. Spare part is worried! | Defect | Hw | Hemolyzer
AQT90 FLEX |  | B | SWO-67347 | No |  |  | Defect | HW | Process Module | Process Unit sensor | 5257
ABL800 BASIC | Gerät überprüft und gereinigt, Meßkammer und Luftauslass gesäubert, Elektroden membraniert, tHb Kal o.k. Cl Level 4 weiterhin grenzwertig | Gerät überprüft und gereinigt, Meßkammer und Luftauslass gesäubert, Elektroden membraniert, tHb Kal o.k. Cl Level 4 weiterhin grenzwertig | SWO-67565 | No | Device checked and cleaned, measuring chamber and air outlet cleaned, electrodes membraneed, tHb Kal o.k. Cl Level 4 still borderline | Device checked and cleaned, measuring chamber and air outlet cleaned, electrodes membraneed, tHb Kal o.k. Cl Level 4 still borderline | Assisted Replacement | Electrode | Membrane | Cl
ABL800 BASIC | Changed the New Display Cable & New Touch Screen, need to observe for 2 weeks. Under Observation. | B | SWO-67519 | No | Changed the New Display Cable & New Touch Screen, need to observe for 2 weeks. Under Observation. |  | Defect | HW | DMS_PC | Touch Screen
ABL800 BASIC | 本体流路清掃・タンパク除去・エアー回路・Ref&pH電極先端清掃 Cal・QC→OK | B | SWO-68125 | No | Body flow path cleaning, protein removal, air circuit, Ref&pH electrode tip cleaning Cal, QC → OK |  | Cleaned | Fluidic pathway | pH/BG Module | Measuring chamber
ABL810 FLEX | E1001 mit D711 R1824 membraniert, danach keine Restlösung vor der Elektrode. Barometer um 10 mmHg korrigiert. | E1001 mit D711 R1824 membraniert, danach keine Restlösung vor der Elektrode. Barometer um 10 mmHg korrigiert. | SWO-67498 | No | E1001 with D711 R1824 membraneed, then no residual solution before the electrode. barometer corrected by 10 mmHg. | E1001 with D711 R1824 membraneed, then no residual solution before the electrode. barometer corrected by 10 mmHg. | Assisted Replacement | Electrode | Membrane | pCO2
ABL800 BASIC | 更换910-310一个后修复 | 更换910-310一个后修复 | SWO-67066 | No | Replace 910-310 one post-fix | Replace 910-310 one post-fix | Defect | HW | AutoCheck Module | Barcode Scanner
ABL825 FLEX | 更换Lac电极膜942-066一个和电极头胶圈834-618一个后修复 | 更换Lac电极膜942-066一个和电极头胶圈834-618一个后修复 | SWO-67613 | No | Replace the Lac electrode film 942-066 one and the electrode head bond ring 834-618 after one repair | Replace the Lac electrode film 942-066 one and the electrode head bond ring 834-618 after one repair | Assisted Replacement | Electrode | Membrane | Lac
ABL815 FLEX | FlexQ Modul überprüft und Verhalten dem Kunden erklärt, Konfiguration gemäß Kundenwunsch angepasst | FlexQ Modul überprüft und Verhalten dem Kunden erklärt, Konfiguration gemäß Kundenwunsch angepasst | SWO-67719 | Yes | FlexQ module checked and behavior explained to the customer, configuration adapted according to customer request | FlexQ module checked and behavior explained to the customer, configuration adapted according to customer request | Re classified RMED | Setup/Courtesy
ABL80 FLEX BASIC | 更换新测试卡修复仪器 | 更换新测试卡修复仪器 | SWO-67630 | No | Replace the new test card repair instrument | Replace the new test card repair instrument
ABL80 FLEX BASIC |  | B | SWO-67608 | No
ABL80 FLEX BASIC | 更换新测试卡修复仪器 | 更换新测试卡修复仪器 | SWO-67638 | No | Replace the new test card repair instrument | Replace the new test card repair instrument
ABL80 FLEX CO-OX | da in der Vergangenheit pO2 regelmäßig fehlerhaft war, keine Verbrauchsmaterialien eingesetzt Vorschlag Analog Board zum Test tauschen oder Ersatzgerät beim Kunden belassen, da Kunde schon sehr verärgert ist das Ersatzgerät läuft bis jetzt ohne Probleme | da in der Vergangenheit pO2 regelmäßig fehlerhaft war, keine Verbrauchsmaterialien eingesetzt Vorschlag Analog Board zum Test tauschen oder Ersatzgerät beim Kunden belassen, da Kunde schon sehr verärgert ist das Ersatzgerät läuft bis jetzt ohne Probleme | SWO-67746 | No | since in the past pO2 was regularly faulty, no consumables used proposal Analog Board to exchange for testing or leave replacement device at the customer, since customer is already very annoyed the replacement device is running so far without any problems | since in the past pO2 was regularly faulty, no consumables used proposal Analog Board to exchange for testing or leave replacement device at the customer, since customer is already very annoyed the replacement device is running so far without any problems
ABL80 FLEX BASIC | Cleaning performed. Reinstalled sensor cassette. Performed two point calibration. | B | SWO-67793 | No | Cleaning performed. Reinstalled sensor cassette. Performed two point calibration.
ABL80 FLEX CO-OX | Gerät aus, Fließwege u. Gerät gereinigt, Gerät ein, Kal. + QC i.O. | Gerät aus; Fließwege u. Gerät gereinigt; Gerät ein; Kal. + QC i.O. | SWO-67670 | No | device off; Flow paths and device cleaned; device; Kal. + QC in the first place | device off; Flow paths and device cleaned; device; Kal. + QC in the first place
ABL90 FLEX | Kunde hat SP Lot ST03 und ST10 ersetzt bei beiden trat der Fehler 1178 auf. Kunde hatte kein weiteres SP. SP 90 Lot ST08 geliefert und installiert. Funktion OK. Reklamationsprotokolle besprochen. | Kunde hat SP Lot ST03 und ST10 ersetzt bei beiden trat der Fehler 1178 auf. Kunde hatte kein weiteres SP. SP 90 Lot ST08 geliefert und installiert. Funktion OK. Reklamationsprotokolle besprochen. | SWO-67865 | Yes | Customer has replaced SP Lot ST03 and ST10 with both occurred the error 1178. Customer had no further SP. SP 90 Lot ST08 supplied and installed. Function OK. complaint logs. | Customer has replaced SP Lot ST03 and ST10 with both occurred the error 1178. Customer had no further SP. SP 90 Lot ST08 supplied and installed. Function OK. complaint logs. | Re classified RMED | PER
ABL90 FLEX | Diverse slangetjes en de naald heb vervangen, flush procedure heb uitgevoerd, hierna calibratie QC's heb bepaald, OK. | Diverse slangetjes en de naald heb vervangen, flush procedure heb uitgevoerd, hierna calibratie QC's heb bepaald, OK. | SWO-67708 | No | Several tubes and needle have replaced, flush procedure have been performed, hereafter calibration Qcs have determined, OK. | Several tubes and needle have replaced, flush procedure have been performed, hereafter calibration Qcs have determined, OK. | Cleaned | Fluidic pathway
ABL90 FLEX | QC nach Kal i.O., aber grenzwertig um sicher zu gehen, SC gewechselt Bitte auf Umgebungstemperatur des Gerätes achten. Gerät stand bei winterlichen Temperaturen am offenen Fenster, was die Gase negativ beeinflusst. | QC nach Kal i.O., aber grenzwertig um sicher zu gehen, SC gewechselt Bitte auf Umgebungstemperatur des Gerätes achten. Gerät stand bei winterlichen Temperaturen am offenen Fenster, was die Gase negativ beeinflusst. | SWO-67747 | No | QC to Kal i.O., but borderline to be sure, SC changed please pay attention to ambient temperature of the device. Device stood at the open window in wintry temperatures, which negatively affects the gases. | QC to Kal i.O., but borderline to be sure, SC changed please pay attention to ambient temperature of the device. Device stood at the open window in wintry temperatures, which negatively affects the gases. | User error | Location of analyzer, cold
ABL90 FLEX | Nach der Installation der Software Version 3.1 MR7 wurde die Hauskonfiguration ohne Fehlermeldung übernommen und der ABL konnte erfolgreich angebunden werden. | Nach der Installation der Software Version 3.1 MR7 wurde die Hauskonfiguration ohne Fehlermeldung übernommen und der ABL konnte erfolgreich angebunden werden. | SWO-68034 | No | After installing the software version 3.1 MR7, the house configuration was taken over without an error message and the ABL could be successfully connected. | After installing the software version 3.1 MR7, the house configuration was taken over without an error message and the ABL could be successfully connected. | Setup | HIS/LIS | MR7
ABL90 FLEX | Fehlersuche, SW neu aufgespielt -> ohne Erfolg, Datenbank gelöscht, SW erneut aufgespielt -> i.O., Neustart = Austausch d. LP u.d. SC notwendig, Startphase u. Initialisierung abgewartet, Konfig. geladen, Kal. = Lac-Fehler, weitere Kal. durchgef., t-Hb-Kal., Kal. + QC i.O. | Fehlersuche; SW neu aufgespielt -> ohne Erfolg; Datenbank gelöscht; SW erneut aufgespielt -> i.O.; Neustart = Austausch d. LP u.d. SC notwendig; Startphase u. Initialisierung abgewartet; Konfig. geladen; Kal. = Lac-Fehler; weitere Kal. durchgef.; t-Hb-Kal.; Kal. + QC i.O. | SWO-67753 | Yes | Troubleshooting; SW replayed -> to no avail; Database deleted; SW replayed -> i.O.; Restart = replacement of LP and SC necessary; start-up phase and initialization awaited; Config. loaded; Kal. = Lac error; more Kal. through it; t-Hb-Kal.; Kal. + QC in the first place | Troubleshooting; SW replayed -> to no avail; Database deleted; SW replayed -> i.O.; Restart = replacement of LP and SC necessary; start-up phase and initialization awaited; Config. loaded; Kal. = Lac error; more Kal. through it; t-Hb-Kal.; Kal. + QC in the first place | Re classified RMED | OOB
ABL90 FLEX PLUS |  | B | CFE83B9C6833EA1C | Yes |  |  | Re classified RMED
ABL90 FLEX PLUS |  | B | CFE9393ADE33D639 | No |  |  | Defect | HW | Pinch Valve Tube
ABL90 FLEX | 更换新试剂包修复仪器 | 更换新试剂包修复仪器 | SWO-67655 | No | Replace the new reagent kit repair instrument | Replace the new reagent kit repair instrument | Assisted Replacement | SP | Pump cal error
AQT90 FLEX | 更换试剂包 | 更换试剂包 | SWO-67862 | No | Replace the reagent kit | Replace the reagent kit | Defect | Consumables | SP | Solution Pack Chip | 5777
ABL835 FLEX | Zeitweise lag ein Kalibrierfehler vor. Messkammer gereinigt. Kal-Werte stabil. Div. Einstellungen in der Konfiguration vorgenommen. Anwender wollen Elektrolyte vorerst bei Kapillarmessungen mitbestimmen laut Frau Dr. Stelzer. QC-Ergebnisse in allen Leveln im Bereich. Gerät überprüft. | Zeitweise lag ein Kalibrierfehler vor. Messkammer gereinigt. Kal-Werte stabil. Div. Einstellungen in der Konfiguration vorgenommen. Anwender wollen Elektrolyte vorerst bei Kapillarmessungen mitbestimmen laut Frau Dr. Stelzer. QC-Ergebnisse in allen Leveln im Bereich. Gerät überprüft. | SWO-67783 | No | At times there was a calibration error. measuring chamber. Kal values stable. Div. Settings made in the configuration. For the time being, users want to influence electrolytes in capillary measurements, according to Dr. Stelzer. QC results at all levels in the range. device checked. | At times there was a calibration error. measuring chamber. Kal values stable. Div. Settings made in the configuration. For the time being, users want to influence electrolytes in capillary measurements, according to Dr. Stelzer. QC results at all levels in the range. device checked.
ABL800 BASIC | ディスプレイユニット交換 | ディスプレイユニット交換 | SWO-67825 | No | Display unit replacement | Display unit replacement | Defect | HW | DMS_PC | Touch Screen
ABL800 FLEX | Druckereinheit inkl. Druckertreiberplatine erneuert Software neu installiert, Gerät konfiguriert ABL kalibriert und überprüft | Druckereinheit inkl. Druckertreiberplatine erneuert Software neu installiert, Gerät konfiguriert ABL kalibriert und überprüft | SWO-67779 | No | Printer unit incl. printer driver board renews software reinstalled, device configures ABL calibrated and checks | Printer unit incl. printer driver board renews software reinstalled, device configures ABL calibrated and checks | Defect | HW | Printer module | Printer head
ABL835 FLEX | pH modul skiftet | pH modul skiftet | SWO-67754 | No | pH Module Switch | pH Module Switch | Assisted Replacement | Electrode | Electrode | pCO2
ABL825 FLEX | Verbogene Einlassende>> Sonde ersetzt. QC und Kalibration. | Verbogene Einlassende>> Sonde ersetzt. QC und Kalibration. | SWO-67964 | No | Bent inlet>> probe replaced. QC and Calibration. | Bent inlet>> probe replaced. QC and Calibration. | Defect | HW | Inlet module | Probe bend
ABL825 FLEX | Fehlersuche, Schlauch zu Vorwärmer defekt, Schläuche ersetzt, Einlass gereinigt, Neustart, akz. Kal abgewartet, Kal u. QC i.O. Probenmessung | Fehlersuche, Schlauch zu Vorwärmer defekt, Schläuche ersetzt, Einlass gereinigt, Neustart, akz. Kal abgewartet, Kal u. QC i.O. Probenmessung | SWO-67671 | No | Troubleshooting, hose to preheater defective, hoses replaced, inlet cleaned, restart, etc. Kal waited, Kal & QC i.O. sample measurement | Troubleshooting, hose to preheater defective, hoses replaced, inlet cleaned, restart, etc. Kal waited, Kal & QC i.O. sample measurement | Defect | HW | Inlet module | Tubing
ABL80 FLEX BASIC | 更换测试卡正常。 | 更换测试卡正常。 | SWO-67859 | No | The replacement test card is OK. | The replacement test card is OK.
ABL80 FLEX BASIC | 更换新测试卡修复仪器 | 更换新测试卡修复仪器 | SWO-68124 | No | Replace the new test card repair instrument | Replace the new test card repair instrument
ABL90 FLEX | roma.ini korrigiert (Barcodetypen neu eingegeben), Neustart, Kal u. QC i.O. | roma.ini korrigiert (Barcodetypen neu eingegeben), Neustart, Kal u. QC i.O. | SWO-68053 | No | roma.ini corrected (barcode types re-entered), restart, Cal and QC i.O. | roma.ini corrected (barcode types re-entered), restart, Cal and QC i.O. | Setup | Roma.ini
ABL90 |  | B | CFECF2734054D02A | No
ABL90 FLEX | System gespült, div. Maschinenprogramme durchgeführt Kal und QC | System gespült, div. Maschinenprogramme durchgeführt Kal und QC | SWO-67989 | No | System flushed, various machine programs carried out Kal and QC | System flushed, various machine programs carried out Kal and QC | Cleaned | Fluidic pathway
ABL90 FLEX | Vergleichsmessung mit einem anderen ABL90 durchgeführt, Differenz im pH 0,002 und bei den Gasen < 0,5 mmHg. Sonst Funktion OK. SP Reklamation an RDEU gemailt. | Überprüfung der ID PCB Kontakte, o.K. Einlasssonde durch Anwender bereits getauscht und Fließwege durchgespült über Abfallanschlussstutzen. Einlassdichtung auch getauscht, nicht erfolgreich. Bitte um Service durch Techniker. | SWO-68014 | No | Comparative measurement carried out with another ABL90, difference in pH 0.002 and for gases of 0.5 mmHg. Otherwise function OK. SP complaint to RDEU. | Check of the ID PCB contacts, o.K. inlet probe by user already exchanged and flow paths flushed through via waste connection nozzles. Inlet seal also exchanged, not successful. Request for service by technicians. | Assisted Replacement | SP | Aspiration
ABL90 FLEX | Nach der Installation der Software Version 3.1 MR7 wurde die Hauskonfiguration ohne Fehlermeldung übernommen und der ABL konnte erfolgreich angebunden werden. | Nach der Installation der Software Version 3.1 MR7 wurde die Hauskonfiguration ohne Fehlermeldung übernommen und der ABL konnte erfolgreich angebunden werden. | SWO-67973 | No | After installing the software version 3.1 MR7, the house configuration was taken over without an error message and the ABL could be successfully connected. | After installing the software version 3.1 MR7, the house configuration was taken over without an error message and the ABL could be successfully connected. | Setup | HIS/LIS | MR7
ABL90 FLEX | Nach der Installation der Software Version 3.1 MR7 wurde die Hauskonfiguration ohne Fehlermeldung übernommen und der ABL konnte erfolgreich angebunden werden. | Nach der Installation der Software Version 3.1 MR7 wurde die Hauskonfiguration ohne Fehlermeldung übernommen und der ABL konnte erfolgreich angebunden werden. | SWO-68035 | No | After installing the software version 3.1 MR7, the house configuration was taken over without an error message and the ABL could be successfully connected. | After installing the software version 3.1 MR7, the house configuration was taken over without an error message and the ABL could be successfully connected. | Setup | HIS/LIS | MR7
ABL90 FLEX | Abfallschlauch kontrolliert - ok. ABL manuell gespült. Einlassdichtung, Anschlussdichtung, Inletprobe und Ventilschlauch präventiv erneuert. Lösungspack aus Kundenbestand erneuert (LP neu SV02 alt ST01) Kalibration und QC's durchgeführt. | Abfallschlauch kontrolliert - ok. ABL manuell gespült. Einlassdichtung, Anschlussdichtung, Inletprobe und Ventilschlauch präventiv erneuert. Lösungspack aus Kundenbestand erneuert (LP neu SV02 alt ST01) Kalibration und QC's durchgeführt. | SWO-67999 | No | Waste hose controlled - ok. ABL manually rinsed. Preventively renewed inlet seal, connection seal, inlet sample and valve hose. Solution pack renewed from customer base (LP new SV02 old ST01) calibration and QC's performed. | Waste hose controlled - ok. ABL manually rinsed. Preventively renewed inlet seal, connection seal, inlet sample and valve hose. Solution pack renewed from customer base (LP new SV02 old ST01) calibration and QC's performed. | Assisted Replacement | SP | Pump cal error
ABL90 FLEX PLUS |  | B | CFECEFC6400DD021 | No |  |  | Assisted Replacement | SC | SC | Glu
ABL90 FLEX PLUS | Bei Eintreffen pCO2 QC Fehler in allen 3 Level, neuen LP eingesetzt da verwendeter LP Lot ST-02 -> keine Verbesserungen, Austausch der SC Kunde hat zwei SC vom Lot 395-9 eingesetzt bei dem das Gerät mitteilte, das Haltbarkeit im Analysator überschritten ist ( zuerst eingesetzt am 06.10.15), SC waren aber lt. Kundenaussage ungeöffnet | Bei Eintreffen pCO2 QC Fehler in allen 3 Level, neuen LP eingesetzt da verwendeter LP Lot ST-02 -> keine Verbesserungen, Austausch der SC Kunde hat zwei SC vom Lot 395-9 eingesetzt bei dem das Gerät mitteilte, das Haltbarkeit im Analysator überschritten ist ( zuerst eingesetzt am 06.10.15), SC waren aber lt. Kundenaussage ungeöffnet | SWO-67992 | No | Upon arrival pCO2 QC error in all 3 levels, new LP used since used LP Lot ST-02 -> no improvements, replacement of sc customer has used two SC from lot 395-9 where the device announced the durability in the analyzer is exceeded (first used on 06.10.15), SC were unopened according to customer statement | Upon arrival pCO2 QC error in all 3 levels, new LP used since used LP Lot ST-02 -> no improvements, replacement of sc customer has used two SC from lot 395-9 where the device announced the durability in the analyzer is exceeded (first used on 06.10.15), SC were unopened according to customer statement | Assisted Replacement | SC | pO2/pCO2
AQT90 FLEX | 更换636-550修复（Denis 20151127 该机器光学单元上下运动异常，机器报警安全位置-22错误，经检查为检测光学单元上下运动的传感器故障，所以更换了光学单元上下运动传感器，代理商工程师反应的卡盘错误，为光学单元运动过程中会卡住处理盘，所以该故障表现为光学单元上下运动故障，会卡住处理盘，需要更换636-550） | 更换636-550修复（Denis 20151127 该机器光学单元上下运动异常，机器报警安全位置-22错误，经检查为检测光学单元上下运动的传感器故障，所以更换了光学单元上下运动传感器，代理商工程师反应的卡盘错误，为光学单元运动过程中会卡住处理盘，所以该故障表现为光学单元上下运动故障，会卡住处理盘，需要更换636-550） | SWO-67899 | No | Replacement 636-550 repair (Denis 20151127 The machine optical unit up and down abnormal motion, machine alarm safety position-22 error, checked to detect the optical unit up and down motion sensor failure, so replaced the optical unit up and down motion sensor, dealer engineer steaming card disk error, The processing disc will be stuck during the optical unit movement, so the fault is manifested as the optical unit up and down motion failure, the processing disc will get stuck, need to replace 636-550) | Replacement 636-550 repair (Denis 20151127 The machine optical unit up and down abnormal motion, machine alarm safety position-22 error, checked to detect the optical unit up and down motion sensor failure, so replaced the optical unit up and down motion sensor, dealer engineer steaming card disk error, The processing disc will be stuck during the optical unit movement, so the fault is manifested as the optical unit up and down motion failure, the processing disc will get stuck, need to replace 636-550) | Defect | HW | Process Module | Process Unit sensor | 5784
AQT90 FLEX | 更换测试卡正常。 | 更换测试卡正常。 | SWO-67858 | No | The replacement test card is OK. | The replacement test card is OK. | Defect | Consumables | Test Kit | Cartridge | 1997 (Broken)
ABL835 FLEX | Régi készülék elszállítása (SN: 754R1969N002) Számláló állása: 1556. Új készülék üzembe helyezése (SN: 754R2143N0001). Membrán-cserék (Ref, Glu, Lac, Cl,Ca, K, Na). Kalibrálások (Inlet, LS, Pumpa, tHb, Elektródák), QC mérés hitelesítés (Kinyomtatott eredmény mellékelve) | B | SWO-68009 | No | Old Appliance Removal (SN: 754R1969N002) Counter: 1556. Installing a new Appliance (SN: 754R2143N0001). Membrane-Replacements (Ref, Glu, Lac, Cl, Ca, K, Na). Calibrations (Inlet, LS, pump, tHb, electrodes), QC measurement calibration (printed result included) |  | Unknown | Analyzer replaced
ABL825 FLEX | Halbjahreswartung durchgeführt, Meßkammer gesäubert, verbogene Einlaßnadel erneuert | Halbjahreswartung durchgeführt, Meßkammer gesäubert, verbogene Einlaßnadel erneuert | SWO-68108 | No | Half-year maintenance carried out, measuring chamber cleaned, bent inlet needle renewed | Half-year maintenance carried out, measuring chamber cleaned, bent inlet needle renewed | Cleaned | Fluidic pathway | El/Met Module | Glu
ABL835 FLEX |  | B | CFED0B09400C0C60 | No
AQT90 FLEX |  | B | CFE91F73DE3ECB83 | No |  |  | Not Cat | Not Cat | Not Cat | Not Cat
AQT90 FLEX |  | B | SWO-68122 | No |  |  | Defect | HW | Cover and Chassis | Replenishment pack PCB
AQT90 FLEX | 更换定标卡重新校准.定标修复仪器 | 更换定标卡重新校准.定标修复仪器 | SWO-66320 | No | Replacement of calibration card recalibration - Calibration repair instrument | Replacement of calibration card recalibration - Calibration repair instrument | Assisted Replacement | Consumables | Test Kit | Cartridge, D-Dimer | 159
ABL825 FLEX | 再次更换一个新的氧电极膜后修复 | 再次更换一个新的氧电极膜后修复 | SWO-68210 | No | Repairafter after replacing a new oxygen electrode membrane again | Repairafter after replacing a new oxygen electrode membrane again | Assisted Replacement | Electrode | Membrane | pO2
ABL80 FLEX BASIC | 更换新测试卡修复仪器 | 更换新测试卡修复仪器 | SWO-68224 | No | Replace the new test card repair instrument | Replace the new test card repair instrument
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-68420 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL90 FLEX | 更换测试卡正常。 | 更换测试卡正常。 | SWO-68222 | No | The replacement test card is OK. | The replacement test card is OK. | Assisted Replacement | SC | Lac
AQT90 FLEX | 更换903-364修复 | 更换903-364修复 | SWO-68126 | No | Replacement 903-364 repair | Replacement 903-364 repair | Defect | HW | Process Module | Optical Unit | 131
AQT90 FLEX | 更换903-364一个修复仪器 | 更换903-364一个修复仪器 | SWO-68250 | No | Replace 903-364 a repair instrument | Replace 903-364 a repair instrument | Defect | HW | Process Module | Optical Unit | 131
ABL80 FLEX | オンライン接続及びテスト実施。 | オンライン接続及びテスト実施。 | SWO-68415 | No | Online connection and test. | Online connection and test.
ABL80 FLEX BASIC |  | B | SWO-68225 | No
ABL80 FLEX BASIC | 代替設置 持ち帰り修理 | 代替設置 持ち帰り修理 | SWO-68294 | No | Alternative installation take-away repair | Alternative installation take-away repair
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-68422 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL80 FLEX BASIC | 更换两张新测试卡 | 更换两张新测试卡 | SWO-68246 | No | Replace two new test cards | Replace two new test cards
ABL80 FLEX BASIC | 更换测试卡正常。 | 更换测试卡正常。 | SWO-68256 | No | The replacement test card is OK. | The replacement test card is OK.
ABL80 FLEX BASIC | 更换新测试卡修复仪器 | 更换新测试卡修复仪器 | SWO-68184 | No | Replace the new test card repair instrument | Replace the new test card repair instrument
ABL80 OSM | Examined wet section and liquid sensors. Found defective hemolizer liquid sensor. | Replaced liquid sensor. Performed tHB calibration. Verified operation of ABL80 OSM. | CFED087D4034C47A | No | Examined wet section and liquid sensors. Found defective hemolizer liquid sensor. | Replaced liquid sensor. Performed tHB calibration. Verified operation of ABL80 OSM.
ABL80 FLEX CO-OX | Überprüfung Anbindung Conworks und Korrektur Überprüfung durch EDV i.O. | Überprüfung Anbindung Conworks und Korrektur Überprüfung durch EDV i.O. | SWO-68360 | No | Verification connection Conworks and correction check by EDP i.O. | Verification connection Conworks and correction check by EDP i.O.
ABL90 FLEX PLUS | Einlassmodul erneuert. Teilenummer 903-338 | Einlassmodul erneuert. Teilenummer 903-338 | SWO-68497 | No | Inlet module renewed. Part number 903-338 | Inlet module renewed. Part number 903-338 | Defect | HW | Auto Inlet module | Email - return device
TCM CombiM |  | B | SWO-68284 | No
ABL90 FLEX | 更换新测试卡 | 更换新测试卡 | SWO-68207 | No | Replace a new test card | Replace a new test card | Assisted Replacement | SC | Reference
ABL90 FLEX | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-68448 | Yes | Repair after replacing a new test card | Repair after replacing a new test card | Re classified RMED | Double entry
ABL90 FLEX | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-68449 | No | Repair after replacing a new test card | Repair after replacing a new test card | Assisted Replacement | SC | Lac
ABL90 FLEX | 更换新测试卡修复仪器 | 更换新测试卡修复仪器 | SWO-68443 | No | Replace the new test card repair instrument | Replace the new test card repair instrument | Assisted Replacement | SC | Reference
AQT90 FLEX | 更换两张测试卡正常。 | 更换两张测试卡正常。 | SWO-68255 | No | Replacing the two test cards is ok. | Replacing the two test cards is ok. | Assisted Replacement | Consumables | Test Kit | Cartridge | 5784 (Broken or Label)
AQT90 FLEX |  | B | SWO-68195 | No |  |  | Defect | HW | Process Module | Optical Unit | 124
ABL835 FLEX | Sustitucion de los nipples. Limpieza salida de aire y desperdicios. Cal, qc. | B | CFED51BF40442CC7 | No | Substitution of nipples. Cleaning air outlet and waste. Cal, qc. |  | Defect | HW | Inlet module | Leaking | Nipple kit
ABL837 FLEX | The analyzer is leaking from the inlet syringe port at the inlet module. | Replaced all the waste lines from spares. Replaced Lac, Ca, K, cCrea and the reference membranes, due to failure on calibrations and qc. replaced the inlet waste nipple kit. | CFED10374051BF6C | No | The analyzer is leaking from the inlet syringe port at the inlet module. | Replaced all the waste lines from spares. Replaced Lac, Ca, K, cCrea and the reference membranes, due to failure on calibrations and qc. replaced the inlet waste nipple kit. | Defect | HW | Fluidic manifold | Tubings
ABL825 FLEX | checked autocheck in service mode. unable to align carousel | Replaced FPC cable 14-pole 636-428 for diagnostic purposes. | CFECE9A240249DA6 | No | checked autocheck in service mode. unable to align carousel | Replaced FPC cable 14-pole 636-428 for diagnostic purposes. | Defect | HW | AutoCheck Module | FPC Cable
ABL827 FLEX | komplettes Oxi Modul mit neuem Hemolyzer eingesetzt Services tHb Kal, Kal, Anwender tHb Kal durchgeführt | komplettes Oxi Modul mit neuem Hemolyzer eingesetzt Services tHb Kal, Kal, Anwender tHb Kal durchgeführt | SWO-68351 | No | complete Oxi module with new Hemolyzer used services tHb Kal, Kal, user tHb Kal carried out | complete Oxi module with new Hemolyzer used services tHb Kal, Kal, user tHb Kal carried out | Defect | HW | Oxi module | Hemolyzer
ABL815 FLEX | Lampenabgleich durchgeführt, Funktion ,Kal und QC OK. | Lampenabgleich durchgeführt, Funktion ,Kal und QC OK. | SWO-68461 | No | Lamp leveling performed, function ,Kal and QC OK. | Lamp leveling performed, function ,Kal and QC OK. | Adjustment | HW | Oxi module | Lamp unit
ABL815 FLEX | QC Level 3 ausgeführt - OK Nacheinweisung FlexQ durchgeführt | QC Level 3 ausgeführt - OK Nacheinweisung FlexQ durchgeführt | SWO-68454 | No | QC Level 3 executed - OK follow-up FlexQ performed | QC Level 3 executed - OK follow-up FlexQ performed | Assisted Replacement | Electrode | Membrane | pCO2
ABL835 FLEX | Oxi modul halogen lámpa intenzitásának beállítása. | Oxi modul halogen lámpa intenzitásának beállítása. | SWO-68532 | No | To adjust the intensity of the Oxy module halogen lamp. | To adjust the intensity of the Oxy module halogen lamp. | Adjustment | HW | Oxi module | Lamp unit
AQT90 FLEX |  | B | CFED281A405B1A5D | No |  |  | Cleaned | HW | Aspiration Module | Needle Wash Unit | FTC_16
AQT90 FLEX | Training and Demonstration completed. | Training and Demonstration completed. | SWO-68738 | Yes | Training and Demonstration completed. | Training and Demonstration completed. | Re classified RMED | Training only
ABL80 FLEX BASIC | 更换新试剂包修复仪器 | 更换新试剂包修复仪器 | SWO-68445 | No | Replace the new reagent kit repair instrument | Replace the new reagent kit repair instrument
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-68658 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL80 FLEX BASIC | 更换903-381一个后修复 | 更换903-381一个后修复 | SWO-68175 | No | Replace 903-381 one post-fix | Replace 903-381 one post-fix
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-68927 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL80 FLEX BASIC | 连续更换四张测试卡后修复 | 连续更换四张测试卡后修复 | SWO-68928 | No | Repair after four consecutive test cards | Repair after four consecutive test cards
ABL80 OSM | Verfied upon arrival. | Performed extended troubleshooting. Found a blockage in the waste tubing. Performed a back flush, problem resolved. | CFEE7C46B6228AA2 | No | Verfied upon arrival. | Performed extended troubleshooting. Found a blockage in the waste tubing. Performed a back flush, problem resolved.
ABL90 FLEX | Gerät steht in Intervention erforderlich => PumpenKal-Fehler Überprüfung Dateiarchiv Aktivitäten, Betriebsmittel, Kal und QC => aktuelles LP wurde am 30.11. wegen inhomogener Lösung und Kal Backlog-Fehlern eingesetzt und mehrmals reinstalliert => seit erstem Einsetzen Vakuum- und PumpenKal-Fehler Einlassnadel gerichtet, da leicht verbogen => Einlassdichtung sonst i.O. Einlassanschlussdichtung neu eingesetzt Reinigung Quetschventilschlauch Spülung er Fließwege mit und ohne SC => ohne Beanstandung Geräteneustart => Fluid Transport mittels Serviceprogramme überprüft => alles i.O. Aktuellen LP reinstalliert => weiterhin PumpenKal Fehler => Reinstallierung alter LP vom Tausch am 30.11. Startphase, Kal und QC i.O. Uhrzeit gestellt | Gerät steht in Intervention erforderlich => PumpenKal-Fehler Überprüfung Dateiarchiv Aktivitäten, Betriebsmittel, Kal und QC => aktuelles LP wurde am 30.11. wegen inhomogener Lösung und Kal Backlog-Fehlern eingesetzt und mehrmals reinstalliert => seit erstem Einsetzen Vakuum- und PumpenKal-Fehler Einlassnadel gerichtet, da leicht verbogen => Einlassdichtung sonst i.O. Einlassanschlussdichtung neu eingesetzt Reinigung Quetschventilschlauch Spülung er Fließwege mit und ohne SC => ohne Beanstandung Geräteneustart => Fluid Transport mittels Serviceprogramme überprüft => alles i.O. Aktuellen LP reinstalliert => weiterhin PumpenKal Fehler => Reinstallierung alter LP vom Tausch am 30.11. Startphase, Kal und QC i.O. Uhrzeit gestellt | SWO-68636 | No | Device is required in intervention => PumpsKal error checking file archive activities, equipment, cal and QC => current LP was used on 30.11. due to inhomogeneous solution and Kal backlog errors and reinstalled several times => since first insertion Vacuum and pumpsKal error inlet needle directed, as slightly bent => inlet seal otherwise usually inserted inlet connection seal newly inserted cleaning squeegee valve hose flushing he flow paths with and without SC => without complaint device restart => Fluid Transport checked by means of service programs => all i.O. current LP reinstalled => further pumpsKal error => reinstall of old LP from the exchange on 30.11. start phase, Kal and QC i.O. time set | Device is required in intervention => PumpsKal error checking file archive activities, equipment, cal and QC => current LP was used on 30.11. due to inhomogeneous solution and Kal backlog errors and reinstalled several times => since first insertion Vacuum and pumpsKal error inlet needle directed, as slightly bent => inlet seal otherwise usually inserted inlet connection seal newly inserted cleaning squeegee valve hose flushing he flow paths with and without SC => without complaint device restart => Fluid Transport checked by means of service programs => all i.O. current LP reinstalled => further pumpsKal error => reinstall of old LP from the exchange on 30.11. start phase, Kal and QC i.O. time set | Assisted Replacement | SP | Pump cal error
ABL90 |  | B | CFEE8633B65BC401 | No |  |  | Assisted Replacement | SC | Ca
ABL90 | Analyzer exhibits multiple screen errors and lockups and patient sample reppoort errors. Screen shot shows patient ID screen with no entry available, and diagonal lines through screen number entry. sO2 missing on one patient printout. | Installed MR7 software. Customer will monitor and advise if any errors occur. | CFEEAE4DB601F3CF | No | Analyzer exhibits multiple screen errors and lockups and patient sample reppoort errors. Screen shot shows patient ID screen with no entry available, and diagonal lines through screen number entry. sO2 missing on one patient printout. | Installed MR7 software. Customer will monitor and advise if any errors occur. | Defect | SW | Lock up
ABL90 | Analyzer exhibits multiple screen errors and lockups and patient sample report errors. Screen shot shows patient ID screen with no entry available, and diagonal lines through screen number entry. sO2 missing on one patient printout. | Installed MR7 software. Customer will monitor and advise if any errors occur. | CFEEDA57B648F237 | No | Analyzer exhibits multiple screen errors and lockups and patient sample report errors. Screen shot shows patient ID screen with no entry available, and diagonal lines through screen number entry. sO2 missing on one patient printout. | Installed MR7 software. Customer will monitor and advise if any errors occur. | Defect | SW | Lock up
ABL90 FLEX | Kundengerät abgezogen und ABL 5 von der Lungenfunktion aufgestellt | ABL90 aufgestellt Schlauchsystem gespült ABL 5 Leihgerät zurück in die LUFU geliefert | SWO-70036 | No | customer device and ABL 5 set up from lung function | ABL90 installed hose system flushed ABL 5 loaner delivered back to the LUFU | Defect | HW | CPU
TCM CombiM | 更换电池431-018后修复 | 更换电池431-018后修复 | SWO-68258 | No | Repair after replacing battery 431-018 | Repair after replacing battery 431-018
ABL90 FLEX | 更换910-285一个后修复 | 更换910-285一个后修复 | SWO-68265 | No | Replace 910-285 one post-fix | Replace 910-285 one post-fix | Defect | HW | Power Supply
ABL90 FLEX | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-68513 | No | Repair after replacing a new test card | Repair after replacing a new test card | Assisted Replacement | SC | Multi
ABL90 FLEX | 更换新测试卡修复仪器 | 更换新测试卡修复仪器 | SWO-68436 | No | Replace the new test card repair instrument | Replace the new test card repair instrument | Assisted Replacement | SC | Reference
ABL90 FLEX | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-69503 | No | Repair after replacing a new test card | Repair after replacing a new test card | Assisted Replacement | SC | Reference
AQT90 FLEX | 更换910-304修复仪器 | 更换910-304修复仪器 | SWO-68450 | No | Replace the 910-304 repair instrument | Replace the 910-304 repair instrument | Defect | HW | Storage Module | Barcode reader, Storage wheel | BR1
AQT90 FLEX |  | B | SWO-68499 | No |  |  | Defect | HW | Incubator Module | Incubator, PCB | 2168/5658
ABL80 FLEX Basic version |  | B | D003D3998C3F10B9 | No
ABL805 FLEX | autocheck functions correctly except when it gets to the point to break the ampoule it returns to the down position | changed the 14 pole cable | CFEE5B8FB6588B93 | Yes | autocheck functions correctly except when it gets to the point to break the ampoule it returns to the down position | changed the 14 pole cable | Re classified RMED | SenDx Production
ABL827 FLEX |  | B | CFEF9B0F2C100236 | No |  |  | Defect | HW | Gas Module | Pressure Transducer
ABL800 BASIC |  | B | SWO-68433 | No |  |  | Defect | HW | DMS_PC | Touch Screen
ABL825 FLEX | Viewed logs and tHb is low when ran as a sample. | Adjusted hemolyzer lamp level and ran a oxy test. | CFEE5767B647AFA0 | No | Viewed logs and tHb is low when ran as a sample. | Adjusted hemolyzer lamp level and ran a oxy test. | Adjustment | HW | Oxi module | Lamp unit
ABL800 BASIC | did the calibration for pumps, LS ,inlet calibration ,2point calibration and QC all results within the range. instrument kept for observation for root cause investigation. | replaced the cleaning solution, checked 10 samples and QC all results found with in the range. | SWO-69449 | No | did the calibration for pumps, LS ,inlet calibration ,2point calibration and QC all results within the range. instrument kept for observation for root cause investigation. | replaced the cleaning solution, checked 10 samples and QC all results found with in the range. | No fault found | Solution bottle
AQT90 FLEX | Se limpia el cup sensor del analizador se realizan diferentes test y calibraciones por servicio, se introducen muestras de blanco y sangre, funcionando correctamente. | Limpieza del cup sensor. | CFEE4D67B648C91B | No | Different tests and calibrations are cleaned by service, white and blood samples are introduced, working properly. | Cleaning the sensor cup. | Cleaned | HW | Incubator Module | Cup Sensor
AQT90 FLEX | Fehlersuche, Pipette in Pierce Position->Pipette verbogen->erneuert, Safe all Positions bestanden, Startphase und Systemüberprüfung ok, Reinigung gestartet und LQC MC2 ok System-und Funktionskontrolle i.O. | P4 erneuert, Testmessung iO. | SWO-68517 | No | Troubleshooting, pipette in Pierce position->pipette bent->renewed, Safe all positions passed, start phase and system check ok, cleaning started and LQC MC2 ok system and function control i.O. | P4 renewed, test measurement iO. | Defect | HW | Aspiration Module | Bend needle
AQT90 FLEX | Platine der NWU und dieselbe gereinigt und getrocknet, Filter gereinigt, Testmessung leider wieder fehlerhaft. Folgeeinsatz erforderlich. | P4 erneuert, Testmessung iO. | SWO-68799 | No | Board of the NWU and the same cleaned and dried, filter cleaned, test measurement unfortunately faulty again. Follow-up use required. | P4 renewed, test measurement iO. | Not Cat | Not Cat | Not Cat | Not Cat
ABL80 FLEX BASIC | 更换新测试卡修复仪器 | 更换新测试卡修复仪器 | SWO-68667 | No | Replace the new test card repair instrument | Replace the new test card repair instrument
ABL80 FLEX BASIC |  | B | SWO-68772 | No
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-69546 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL80 FLEX BASIC |  | B | SWO-68654 | No
ABL80 FLEX BASIC | 更换三张测试卡后修复 | 更换三张测试卡后修复 | SWO-69501 | No | Repair after replacing three test cards | Repair after replacing three test cards
ABL90 FLEX | Anschlußdichtung u. Nadel gew., neues LP eingesetzt, Startphase abgewartet, Kal. = Fehler, SC gew., Startphase abgewartet, Kal. + QC i.O., ein LP beim Kunden hinterlegt | Anschlußdichtung u. Nadel gew.; neues LP eingesetzt; Startphase abgewartet; Kal. = Fehler; SC gew.; Startphase abgewartet; Kal. + QC i.O.; ein LP beim Kunden hinterlegt | SWO-68828 | No | connection seal and needle; new LP inserted; start-up phase awaited; Kal. = Error; SC. start-up phase awaited; Kal. + QC in the first place; an LP deposited with the customer | connection seal and needle; new LP inserted; start-up phase awaited; Kal. = Error; SC. start-up phase awaited; Kal. + QC in the first place; an LP deposited with the customer | Assisted Replacement | SC | pCO2
ABL90 FLEX | Gerät überprüft, Touchsreen erneuert | Gerät überprüft, Touchsreen erneuert | SWO-68944 | No | Device checked, touchsreen renewed | Device checked, touchsreen renewed | Defect | HW | Touchscreen | Wet
ABL90 FLEX | Gerät aus- eingeschaltet neue SC eingesetzt und Startup abgewartet Reklamationen zur Zentrale geschickt | Gerät aus- eingeschaltet neue SC eingesetzt und Startup abgewartet Reklamationen zur Zentrale geschickt | SWO-68801 | No | Device switched off new SC used and startup awaited complaints sent to the head office | Device switched off new SC used and startup awaited complaints sent to the head office | Assisted Replacement | SC | Temp
ABL90 FLEX | Fogyóanyag tesztelés cseregépen. Ott megfelelő. Bemeneti tömítés cseréje. | B | SWO-68884 | No | Consumables testing on a replacement machine. Right there. Replacing the inlet seal. |  | Defect | HW | Inlet Gasket
ABL90 FLEX | Hemolyzer wieder ausgebaut umgebaut. Sensorcassette erneuert. Gerät beim Kunden wieder messbereit aufgestellt. Ableitströme wurden wegen defekter Sensorcassette angezeigt! | Hemolyzer wieder ausgebaut umgebaut. Sensorcassette erneuert. Gerät beim Kunden wieder messbereit aufgestellt. Ableitströme wurden wegen defekter Sensorcassette angezeigt! | SWO-69245 | No | Hemolyzer rebuilt. Sensorcassette renewed. device at the customer's place ready for measurement again. Discharge currents were indicated due to defective sensor cassette! | Hemolyzer rebuilt. Sensorcassette renewed. device at the customer's place ready for measurement again. Discharge currents were indicated due to defective sensor cassette! | Assisted Replacement | SC | Leak current
ABL90 | Determined SC had recently been replaced. Internal Qc and Cals=passed. | Ran pO2 test, and ran several cals. Cleared error. | D00005322A5EEAAD | No | Determined SC had recently been replaced. Internal Qc and Cals=passed. | Ran pO2 test, and ran several cals. Cleared error. | User error | QC verification ext. | pO2
ABL90 FLEX | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-68926 | No | Repair after replacing a new test card | Repair after replacing a new test card | Assisted Replacement | SC | Multi
AQT90 FLEX | 更换903-171修复仪器 | 更换903-171修复仪器 | SWO-68435 | No | Replace the 903-171 repair instrument | Replace the 903-171 repair instrument | Defect | HW | Aspiration Module | Bend needle | 4794
ABL825 FLEX | 更换Lac电极膜942-066后修复 | 更换Lac电极膜942-066后修复 | SWO-68670 | No | Repair after replacing Lac electrode film 942-066 | Repair after replacing Lac electrode film 942-066 | Assisted Replacement | Electrode | Membrane | Lac
ABL825 FLEX | Found inlet needle bent. Also found small clot in waste line coming from inlet gasket. | Replaced inlet probe with one from customer stock. Cleared waste line with protein remover solution and water. Replaced waste pump tubing with customer stock as a preventitive measure. | CFEFDC7D2C1E1285 | No | Found inlet needle bent. Also found small clot in waste line coming from inlet gasket. | Replaced inlet probe with one from customer stock. Cleared waste line with protein remover solution and water. Replaced waste pump tubing with customer stock as a preventitive measure. | Defect | HW | Inlet module | Probe bend
ABL800 BASIC | 引上げ修理実施。 | B | SWO-70740 | No | Raiserepair carried out. |  | Defect | HW | Gas Module | Flow Regulator, left
ABL825 FLEX | Intensität mit Oxi Complete Test überprüft, Wert >4095, Halogenlampe einjustiert auf 3100, Complete Test ohne Beanstandung, Gerät einkalibriert, Kalibration und QC i.O. | Intensität mit Oxi Complete Test überprüft, Wert >4095, Halogenlampe einjustiert auf 3100, Complete Test ohne Beanstandung, Gerät einkalibriert, Kalibration und QC i.O. | SWO-68965 | No | Intensity checked with Oxi Complete Test, value >4095, halogen lamp adjusted to 3100, Complete test without complaint, device calibrated, calibration and QC i.O. | Intensity checked with Oxi Complete Test, value >4095, halogen lamp adjusted to 3100, Complete test without complaint, device calibrated, calibration and QC i.O. | Adjustment | HW | Oxi module | Lamp unit
ABL835 FLEX | Serial output not sending Data | replaced CPU Pcb from a loaner | CFEFB94D2C164EF2 | No | Serial output not sending Data | replaced CPU Pcb from a loaner | Defect | HW | DMS_PC | CPU
ABL835 FLEX |  | B | D010524D282C8FCD | No
ABL835 FLEX | OXI tesztek,Alvadék eltávolítás, Protein eltávolítás, Műszer tesztelése, Kalibrálások(Inlet, LS, Pumpa, tHb, Elektródák), QC mérés hitelesítés *********** Valószinüleg alvadék került a készülékbe, Nem megfelelő mennyiségi mintát tudott felszívni a műszer. Fontos, hogy 95ul-es kapilláris csövet teljesen megtöltsék és fel legyen keverve a heparinozott kap. csővel. *********** | OXI tesztek,Alvadék eltávolítás, Protein eltávolítás, Műszer tesztelése, Kalibrálások(Inlet, LS, Pumpa, tHb, Elektródák), QC mérés hitelesítés *********** Valószinüleg alvadék került a készülékbe, Nem megfelelő mennyiségi mintát tudott felszívni a műszer. Fontos, hogy 95ul-es kapilláris csövet teljesen megtöltsék és fel legyen keverve a heparinozott kap. csővel. *********** | SWO-68867 | No | OXY tests, curing removal, Protein removal, instrument testing, calibration (Inlet, LS, pump, tHb, electrodes), QC measurement Authentication * * * * * * * * * * * The device has been sufficiently cured, it could absorb an inadequate quantitative sample Instrument. It is important that the 95ul capillary tube is fully filled and mixed with the heparinized cap. Tube. *********** | OXY tests, curing removal, Protein removal, instrument testing, calibration (Inlet, LS, pump, tHb, electrodes), QC measurement Authentication * * * * * * * * * * * The device has been sufficiently cured, it could absorb an inadequate quantitative sample Instrument. It is important that the 95ul capillary tube is fully filled and mixed with the heparinized cap. Tube. *********** | Cleaned | Fluidic pathway
AQT90 FLEX | P4 erneuert, Testmessung iO. | P4 erneuert, Testmessung iO. | SWO-68800 | No | P4 renewed, test measurement iO. | P4 renewed, test measurement iO. | Not Cat | Not Cat | Not Cat | Not Cat
AQT90 FLEX |  | B | CFEF76CD2C2D8A19 | No |  |  | Defect | HW | Process Module | Optical Unit | 131
TCM-COMBIM |  | B | CFEF8A792C328BDA | No
ABL80 FLEX BASIC | 更换测试卡正常。 | 更换测试卡正常。 | SWO-69208 | No | The replacement test card is OK. | The replacement test card is OK.
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-69204 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL80 FLEX BASIC | 更换测试卡正常。 | 更换测试卡正常。 | SWO-68963 | No | The replacement test card is OK. | The replacement test card is OK.
ABL80 FLEX BASIC | Reconnected sensor connector board and done zeroing procedure. then reinstalled sc. | Reconnected sensor connector board and done zeroing procedure. then reinstalled sc. | SWO-69397 | No | Reconnected sensor connector board and done zeroing procedure. then reinstalled sc. | Reconnected sensor connector board and done zeroing procedure. then reinstalled sc.
ABL90 FLEX | Im Service Fließwege gereinigt, Fluid Transport = i.O., pO2- u. Oxitest, Test wiederholen , Kal wiederholt - i.O., QC Level C fehlerhaft ( pO2), LP ersetzt, Kal u. QC i.O. | Im Service Fließwege gereinigt, Fluid Transport = i.O., pO2- u. Oxitest, Test wiederholen , Kal wiederholt - i.O., QC Level C fehlerhaft ( pO2), LP ersetzt, Kal u. QC i.O. | SWO-68970 | No | In service flow paths cleaned, fluid transport = i.O., pO2 and Oxitest, test repeat, kal repeated - i.O., QC Level C faulty ( pO2), LP replaced, Kal & QC i.O. | In service flow paths cleaned, fluid transport = i.O., pO2 and Oxitest, test repeat, kal repeated - i.O., QC Level C faulty ( pO2), LP replaced, Kal & QC i.O. | Assisted Replacement | SP | pO2
ABL90 FLEX | CPU unit UIM945 vervangen. | CPU unit UIM945 vervangen. | SWO-69155 | No | Replace CPU unit UIM945. | Replace CPU unit UIM945. | Defect | HW | CPU
ABL90 | Confirmed observed by review of data logs. Inspected analyzer, no errors reported in diagnostics, no issues observed. | Replaced Sensor cassette to resolve issue. Replace solution pack as precautionary measure. | CFF0DB27A25FC8EA | No | Confirmed observed by review of data logs. Inspected analyzer, no errors reported in diagnostics, no issues observed. | Replaced Sensor cassette to resolve issue. Replace solution pack as precautionary measure. | Assisted Replacement | SC | pCO2
ABL90 | Returned to find pO2 out.Reloaded software, replace Inlet Gasket, Flushed, pO2 still out. | Replaced SC from customer stock. Also, moved Lab analyzer with Stat Lab analyzer at customers request. | D000063E2A1BEAB0 | No | Returned to find pO2 out. Reloaded software, replace Inlet Gasket, Flushed, pO2 still out. | Replaced SC from customer stock. Also, moved Lab analyzer with Stat Lab analyzer at customers request. | Assisted Replacement | SC | pO2
ABL90 | Tech support suggested removing Jumper on battery circuit board to establish icon. | Remove Replace Jumper on battery circuit. | CFF10F88A22ECC25 | No | Tech support suggested removing Jumper on battery circuit board to establish icon. | Remove Replace Jumper on battery circuit. | Unknown | Revisit required | New upSWO next day
ABL90 FLEX PLUS |  | B | CFF0CEBBA2422AE1 | No |  |  | Assisted Replacement | SC | Unknown
ABL90 | Verificação do funcionamento , limpeza do circuito, verificação do estado da agulha e entrada, teste de funcionamento sem sucesso, substituição do pack, teste de funcionamento, establização e passagem do QC. | B | CFF0FCCFA2277C11 | No | Checking the operation, cleaning the circuit, checking the state of the needle and inlet, unsuccessful operation test, replacement of the pack, test of operation, establishment and passage of QC. |  | Assisted Replacement | SP | FT
ABL825 FLEX | Lactat Elektrode aus alt Bestand ersetzt | Lactat Elektrode aus alt Bestand ersetzt | SWO-69028 | No | Lactate electrode replaced from old stock | Lactate electrode replaced from old stock | Assisted Replacement | Electrode | Electrode | Lac
ABL827 FLEX |  | B | CFF2051718113682 | No |  |  | Assisted Replacement | Electrode | Electrode | pCO2
ABL825 FLEX | Tested power supply and it worked find. Deducted hard drive was the problem. | 12-1: Pulled machine apart to replace hard drive and realized I had ordered wrong part. I needed the SATA drive. 12 3: Got new hard drive in and installed and had an out of box failure. However, when I put the original hard drive back in on 12 1, machine started back up and is oporational. Machine has continued to work without problem all week. At this point I have put all original parts back in analyzer and will continue to monitor. | CFF1095BA21358F1 | No | Tested power supply and it worked find. Deducted hard drive was the problem. | 12-1: Pulled machine apart to replace hard drive and realized I had ordered wrong part. I needed the SATA drive. 12 3: Got new hard drive in and installed and had an out of box failure. However, when I put the original hard drive back in on 12 1, machine started back up and is oporational. Machine has continued to work without problem all week. At this point I have put all original parts back in analyzer and will continue to monitor. | Defect | HW | DMS_PC | HDD
ABL835 FLEX |  | B | CFF1F960180465BA | Yes |  |  | Re classified RMED
ABL800 BASIC | Firstly replaced the Lac membrane since there was Lac calibration failure. then shut down the machine for 30 minutes, restarted the machine after 2nd 2-point calibration the same error appears. so again switched off the machine, reset the Analyzer control PCB connections. After restarting, done 2-point calibration, the machine is working fine. Kept under observation. | B | SWO-69442 | No | Firstly replaced the Lac membrane since there was Lac calibration failure. then shut down the machine for 30 minutes, restarted the machine after 2nd 2-point calibration the same error appears. so again switched off the machine, reset the Analyzer control PCB connections. After restarting, done 2-point calibration, the machine is working fine. Kept under observation. |  | Unknown | Restart analyzer
ABL825 FLEX | Performed a complete oxy test and passed and failed lamp test. | Performed a THB in the service mode. Adjsuted lamp to 3210. | CFF0C9AFA20FE33F | No | Performed a complete oxy test and passed and failed lamp test. | Performed a THB in the service mode. Adjsuted lamp to 3210. | Adjustment | HW | Oxi module | Lamp unit
ABL825 FLEX | Ran complete oxy test and failed lamp test. | Adjusted the lamp to 3215 performed a THB calibration in the service mode. | CFF0D746A236E342 | No | Ran complete oxy test and failed lamp test. | Adjusted the lamp to 3215 performed a THB calibration in the service mode. | Adjustment | HW | Oxi module | Lamp unit
AQT90 FLEX | Has to change the optical unit. 902-967 Optical Unit | B | SWO-71655 | No | Has to change the optical unit. 902-967 Optical Unit |  | Defect | HW | Process Module | Optical Unit | 131
AQT90 FLEX |  | B | CFF0CE07A2552B29 | No |  |  | Defect | Consumables | Test Kit | Cartridge | 167 (Broken)
ABL80 FLEX CO-OX | Fehlersuche SP wurde durch Kunden bereits erneuert, Neue SC derzeit nicht vorhanden, SC gespült, Gerinnsel aus SC entfernt Ansaugschlauch war vom Stutzen abgesprungen, Fehler behoben, | Fehlersuche SP wurde durch Kunden bereits erneuert, Neue SC derzeit nicht vorhanden, SC gespült, Gerinnsel aus SC entfernt Ansaugschlauch war vom Stutzen abgesprungen, Fehler behoben, | SWO-69382 | No | Troubleshooting SP has already been renewed by customers, New SC currently not available, SC flushed, clot removed from SC suction hose had bounced off the nozzle, bug fixed, | Troubleshooting SP has already been renewed by customers, New SC currently not available, SC flushed, clot removed from SC suction hose had bounced off the nozzle, bug fixed,
ABL80 FLEX BASIC | Gerät überprüft, pO2 Zerotest durchgeführt, keine Kalibration möglich SC 80 945-776 216369, SP80 944-309 214899 defekt | B | SWO-69402 | No | Device checked, pO2 zero test performed, no calibration possible SC 80 945-776 216369, SP80 944-309 214899 defective
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-69420 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL80 FLEX BASIC | 更换新测试卡修复仪器 | 更换新测试卡修复仪器 | SWO-69802 | No | Replace the new test card repair instrument | Replace the new test card repair instrument
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-69421 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL90 FLEX | Hemolysator erneuert, Einlaßsonde erneuert, tHb Kalibration. | Hemolysator erneuert, Einlaßsonde erneuert, tHb Kalibration. | SWO-69369 | No | Hemolysator renewed, inlet probe renewed, tHb calibration. | Hemolysator renewed, inlet probe renewed, tHb calibration. | Defect | HW | Hemolyzer
ABL90 | Unit failing LS adjust, pump cal and pressure test in fluidic transport programs. Checked unit for signs of leakage. | Performed fludic transport programs. All passed. Left customer spair inlet gasket. | CFF2384D184FE32B | No | Unit failing LS adjust, pump cal and pressure test in fluidic transport programs. Checked unit for signs of leakage. | Performed fludic transport programs. All passed. Left customer spair inlet gasket. | Adjustment | Service pgm
ABL90 | Service tip indicates checking position of micro switches. Micro switches in wrong position for Icon display. | Make both switches. | CFF34B128E58F0EA | No | Service tip indicates checking position of micro switches. Micro switches in wrong position for Icon display. | Make both switches. | Adjustment | Battery | Jumper off/on
TCM CombiM |  | B | SWO-69203 | No
AQT90 FLEX | 更换903-171一个正常。 | 更换903-171一个正常。 | SWO-68918 | No | Replace 903-171 one normal. | Replace 903-171 one normal. | Defect | HW | Aspiration Module | Bend needle | 4794
ABL825 FLEX | Laktat Elektrode (E7077 RP01) geprüft und ausgetauscht, Funktionskontrolle ok | Laktat Elektrode (E7077 RP01) geprüft und ausgetauscht, Funktionskontrolle ok | SWO-69298 | No | Lactate electrode (E7077 RP01) tested and replaced, function control ok | Lactate electrode (E7077 RP01) tested and replaced, function control ok | Assisted Replacement | Electrode | Electrode | Lac
ABL800 BASIC | Membrane need to replacement. Replaced the membrane with new one. Perform 2 point calibration.After replace the po2 membrane, unit is working properly. | Membrane need to replacement. Replaced the membrane with new one. Perform 2 point calibration.After replace the po2 membrane, unit is working properly. | SWO-69366 | No | Membrane need to replacement. Replaced the membrane with new one. Perform 2 point calibration. After replace the po2 membrane, unit is working properly. | Membrane need to replacement. Replaced the membrane with new one. Perform 2 point calibration. After replace the po2 membrane, unit is working properly. | Assisted Replacement | Electrode | Membrane | pO2
ABL800 BASIC | replaced the cleaning solution, checked 10 samples and QC all results found with in the range. | replaced the cleaning solution, checked 10 samples and QC all results found with in the range. | SWO-69468 | Yes | replaced the cleaning solution, checked 10 samples and QC all results found with in the range. | replaced the cleaning solution, checked 10 samples and QC all results found with in the range. | Re classified RMED | Double entry | SWO-69449
ABL837 FLEX | Oxi test indicates Lamp. Neon number too low. | Tried Oxi PCB, unable too eliminate error. | CFF349708E5CF0E7 | No | Oxi test indicates Lamp. Neon number too low. | Tried Oxygen PCB, unable too eliminate error. | Unknown | Revisit required
AQT90 FLEX | Scanner bestellt. | B | SWO-69407 | No | Scanner ordered. |  | Defect | HW | Storage Module | Barcode reader, Storage wheel
AQT90 FLEX | Dismantiled the process wheel and incubator wheel. Cleaned both of them. Cleaned the sensors. Reassembled the process wheel optical unit and incubator wheel. | B | SWO-69410 | No | Dismantiled the process wheel and incubator wheel. Cleaned both of them. Cleaned the sensors. Reassembled the process wheel optical unit and incubator wheel. |  | Cleaned | HW | Incubator Module | Punchup
ABL90 FLEX | 调节接口板的位置后修复 | 调节接口板的位置后修复 | SWO-68975 | No | Repair after adjusting the position of the interface board | Repair after adjusting the position of the interface board | Defect | HW | SC ID PCB
ABL80 FLEX BASIC | 更换新测试卡修复仪器 | 更换新测试卡修复仪器 | SWO-69594 | No | Replace the new test card repair instrument | Replace the new test card repair instrument
ABL80 FLEX BASIC | 更换新测试卡修复仪器 | 更换新测试卡修复仪器 | SWO-69600 | No | Replace the new test card repair instrument | Replace the new test card repair instrument
ABL80 FLEX BASIC | 更换新测试卡修复仪器 | 更换新测试卡修复仪器 | SWO-69572 | No | Replace the new test card repair instrument | Replace the new test card repair instrument
ABL80 OSM |  | B | CFF52AE7043BB979 | No
ABL90 FLEX | 更换新试剂包后修复 | 更换新试剂包后修复 | SWO-69479 | No | Repair after replacing a new kit | Repair after replacing a new kit | Assisted Replacement | SP | FT
ABL800 BASIC | replaced the cleaning solution,checked 10 samples and QC all results found with in the range. | replaced the cleaning solution, checked 10 samples and QC all results found with in the range. | SWO-69463 | Yes | replaced the cleaning solution,checked 10 samples and QC all results found with in the range. | replaced the cleaning solution, checked 10 samples and QC all results found with in the range. | Re classified RMED | Double entry | SWO-69449
ABL80 FLEX BASIC | Neue Sensorkassette und Lösungspack installiert, Fehler bleibt bestehen, ABL zurückgesetzt und neu konfiguriert | Neue Sensorkassette und Lösungspack installiert, Fehler bleibt bestehen, ABL zurückgesetzt und neu konfiguriert | SWO-69627 | No | New sensor cartridge and solution pack installed, error persists, ABL resets and reconfigured | New sensor cartridge and solution pack installed, error persists, ABL resets and reconfigured
ABL80 FLEX BASIC | 更换测试卡正常。 | 更换测试卡正常。 | SWO-71025 | No | The replacement test card is OK. | The replacement test card is OK.
ABL80 FLEX BASIC |  | B | SWO-69555 | No
ABL80 FLEX BASIC |  | 要借用的CF卡仓库无库存，因为不影响医院使用，暂不进行维修 (denis 20151126 该代理商老板要求定期查看每一台机器数据日志，该分析仪数据日志下载失败，该机器从序列号看为945计算机单元，上海没有945计算机单元CF卡，无法修复，由于该机器的故障并不影响使用，暂时将该排工单关闭，等到CF卡有货后再进行修复） | SWO-69510 | No |  | The CF card warehouse to borrow is out of stock because it does not affect hospital use, no maintenance is carried out (denis 20151126 The dealer boss requires that each machine data log be viewed regularly, the analyzer data log download failed, the machine from the serial number is 945 computer unit, Shanghai does not have 945 computer unit CF card, can not be repaired, because the failure of the machine does not affect the use of the temporary closure of the ticket, wait until the CF card is available for repair)
ABL80 FLEX BASIC | 更换新试剂包后修复 | 更换新试剂包后修复 | SWO-69624 | No | Repair after replacing a new kit | Repair after replacing a new kit
ABL80 FLEX BASIC | 更换新测试卡修复仪器 | 更换新测试卡修复仪器 | SWO-70317 | No | Replace the new test card repair instrument | Replace the new test card repair instrument
ABL80 OSM | Found that the 2106 error was truly failing again. | Replaced the Hemolyzer and installed a new Maintenance kit and the problem is still there. Will return tomorrow with more parts. | CFF642807A216B9B | No | Found that the 2106 error was truly failing again. | Replaced the Hemolyzer and installed a new Maintenance kit and the problem is still there. Will return tomorrow with more parts.
ABL80 FLEX BASIC |  | B | SWO-69513 | No
ABL90 |  | B | CFF71B49F049B058 | No |  |  | Replaced | Probe bend
ABL90 FLEX | LP wurde nicht ordnungsgemäß eingesetzt im Service Menü getestet, Fehlermeldungen Fließweg durchgespült, Einlaßdichtung erneuert neuen LP installiert ( Materialaus Kundenbestand) | LP wurde nicht ordnungsgemäß eingesetzt im Service Menü getestet, Fehlermeldungen Fließweg durchgespült, Einlaßdichtung erneuert neuen LP installiert ( Materialaus Kundenbestand) | SWO-69696 | No | LP was not used properly tested in the service menu, error messages flow path flushed through, inlet seal renewed new LP installed (material from customer stock) | LP was not used properly tested in the service menu, error messages flow path flushed through, inlet seal renewed new LP installed (material from customer stock) | Assisted Replacement | SP | Activation
ABL90 FLEX | Gerät überprüft, Einlaß gesäubert, Kal und QC o.k. | Gerät überprüft, Einlaß gesäubert, Kal und QC o.k. | SWO-69760 | No | Device checked, inlet cleaned, Kal and QC ok | Device checked, inlet cleaned, Kal and QC ok | Adjustment | Inlet gasket
ABL90 | Errors on external QC pO2 - low QC is High. Customer has exchanged multiple SC and SP. | Replaced SC and SP's to troubleshoot, replaced inlet probe and gasket. Final replacement to correct issue was 3rd SC. | CFF62BEC7A023806 | No | Errors on external QC pO2 - low QC is High. Customer has exchanged multiple SC and SP. | Replaced SC and SP's to troubleshoot, replaced inlet probe and gasket. Final replacement to correct issue was 3rd SC. | Assisted Replacement | SC | pO2 | Wetability ?
ABL90 |  | B | CFF5EBD57A5A1151 | No |  |  | Defect | HW | Oxi module | Pinch Valve tube
ABL90 FLEX | Drucker erneuert. | Drucker erneuert. | SWO-69677 | No | Printer renewed. | Printer renewed. | Defect | HW | Printer head
ABL90 FLEX | Bei Eintreffen Analysator in Startphase (Messbereit in 40 Minuten) - Kunde hatte noch SC und Einlassnadel erneuert - Danach Meldung : Chipdaten des SP könnten nicht gelesen werden ! Kontakte überprüft, gereinigt und SP nochmals installiert - Funktionsüberprüfung ohne Beanstandung ! | Bei Eintreffen Analysator in Startphase (Messbereit in 40 Minuten) - Kunde hatte noch SC und Einlassnadel erneuert - Danach Meldung : Chipdaten des SP könnten nicht gelesen werden ! Kontakte überprüft, gereinigt und SP nochmals installiert - Funktionsüberprüfung ohne Beanstandung ! | SWO-69682 | No | On arrival Analyzer in start phase (ready to be measured in 40 minutes) - Customer had still replaced SC and inlet needle - After that message : Chip data of the SP could not be read ! Contacts checked, cleaned and SP reinstalled - function check without complaint ! | On arrival Analyzer in start phase (ready to be measured in 40 minutes) - Customer had still replaced SC and inlet needle - After that message : Chip data of the SP could not be read ! Contacts checked, cleaned and SP reinstalled - function check without complaint ! | Replaced | Probe
ABL90 FLEX PLUS | Koffer ABL90Flex Plus geliefert | Koffer ABL90Flex Plus geliefert | SWO-70081 | No | Case ABL90Flex Plus delivered | Case ABL90Flex Plus delivered | Cleaned | Fluidic pathway
ABL90 FLEX PLUS | Lösungspack erneuert und Einlassmodul geprüft. Einlasskal. durchgeführt. Einlasssonde erneuert und Fließwege gespült. | Lösungspack erneuert und Einlassmodul geprüft. Einlasskal. durchgeführt. Einlasssonde erneuert und Fließwege gespült. | SWO-69933 | No | solution pack renewed and inlet module checked. Entrance scale. Performed. Inlet probe renewed and flow paths rinsed. | solution pack renewed and inlet module checked. Entrance scale. Performed. Inlet probe renewed and flow paths rinsed. | Defect | SW | Auto Inlet module | Auto Inlet module
TCM4 |  | B | SWO-69533 | No
ABL90 FLEX |  | 更换902-980修复仪器 | SWO-69523 | No |  | Replace the 902-980 repair instrument | Defect | HW | Hemolyzer
ABL90 FLEX | 更换新试剂包修复仪器 | 更换新试剂包修复仪器 | SWO-69605 | No | Replace the new reagent kit repair instrument | Replace the new reagent kit repair instrument | Assisted Replacement | SP | Pump cal error
AQT90 FLEX |  | B | SWO-69512 | No |  |  | Defect | HW | Storage Module | Barcode reader, Storage wheel
AQT90 FLEX | 更换新测试卡修复仪器 | 更换新测试卡修复仪器 | SWO-69807 | No | Replace the new test card repair instrument | Replace the new test card repair instrument | Defect | Consumables | Test Kit | Cartridge | 5689 (Broken)
ABL825 FLEX |  | B | CFF60CE37A2FAF36 | No |  |  | Assisted Replacement | Electrode | Electrode | pCO2
ABL825 FLEX |  | B | SWO-69541 | No |  |  | Assisted Replacement | Electrode | Electrode | Reference
ABL815 FLEX |  | B | CFF635EE7A56F56D | Yes |  |  | Re classified RMED | Not RMED issue
ABL800 BASIC | ①電極交換②キャリブレーションの安定を確認③ＱＣを確認 | ①電極交換②キャリブレーションの安定を確認③ＱＣを確認 | SWO-69508 | No | Check the stability of the electrode replacement and calibration .3- QC | Check the stability of the electrode replacement and calibration .3- QC | Assisted Replacement | Electrode | Electrode | pH
ABL825 FLEX | Inspected the inlet module, found the inlet probe badly bent. | Replaced the inlet probe.. Sending the customer extra probes and inlet rubber tubing to replace when needed. | CFF5EE937A34DC61 | No | Inspected the inlet module, found the inlet probe badly bent. | Replaced the inlet probe.. Sending the customer extra probes and inlet rubber tubing to replace when needed. | Defect | HW | Inlet module | Probe bend
ABL800 BASIC | Status der Led gesichert, Netzteil wieder zurückgetauscht, E-Teil wird bestellt, Kal und QC durchgeführt, Anwender informiert | Bildschirm wieder schwarz. Kd. bittet lt. Vereinbarung H. Zallmanzig zu informieren. | SWO-69685 | No | Status of led secured, power supply replaced, E-part is ordered, Kal and QC performed, user informed | Screen black again. Kd. asks to inform H. Zallmanzig according to the agreement. | Unknown | Revisit required
ABL835 FLEX | The hemolyzer is making a louder pitch noise thatn the other hemolyzer. I came by this morning and checked it and it was working properly. Customer called back and asked me to replace it. Said it was too loud. | Replaced hemolyzer. Adjusted lamp. Ran complete OXI test. | CFF5E3197A5BD0EF | No | The hemolyzer is making a louder pitch noise thatn the other hemolyzer. I came by this morning and checked it and it was working properly. Customer called back and asked me to replace it. Said it was too loud. | Replaced hemolyzer. Adjusted lamp. Ran complete OXI test. | Defect | HW | Oxi module | Hemolyzer | Noise
ABL825 FLEX | Barometer Druck angepasst. Kalibration mehrfach durchgeführt. | Barometer Druck angepasst. Kalibration mehrfach durchgeführt. | SWO-69763 | No | Barometer pressure adjusted. calibration carried out several times. | Barometer pressure adjusted. calibration carried out several times. | Assisted Replacement | Electrode | Electrode | pCO2
ABL825 FLEX | Messkammer gereinigt. Pump Tubing BG ersetzt. | Messkammer gereinigt. Pump Tubing BG ersetzt. | SWO-70418 | No | measuring chamber. Pump Tubing BG replaced. | measuring chamber. Pump Tubing BG replaced. | Defect | HW | pH/BG Module | Pump tubing
ABL815 FLEX | Beim Eintreffen Gerät ohne Oximetriefehler, dafür mit FlexQ Positionierfehler Überprüfung Kal-Ergebnisse => Kal von 04.12. 15:09 Uhr bis 06.12. 02:10 Uhr mit Oximetriefehler => jetzt wieder i.O. Durchführung FlexQ Positionskalibierung Durchführung Oximetrie Complete Test => i.O. => Halogenlampe bei 3750 => Lampe nachjustiert Mehrmals Kal1 i.O. | Beim Eintreffen Gerät ohne Oximetriefehler, dafür mit FlexQ Positionierfehler Überprüfung Kal-Ergebnisse => Kal von 04.12. 15:09 Uhr bis 06.12. 02:10 Uhr mit Oximetriefehler => jetzt wieder i.O. Durchführung FlexQ Positionskalibierung Durchführung Oximetrie Complete Test => i.O. => Halogenlampe bei 3750 => Lampe nachjustiert Mehrmals Kal1 i.O. | SWO-69666 | No | On arrival device without oximetry error, but with FlexQ positioning error check Cal results => Kal from 04.12. 15:09 o'clock to 06.12. 02:10 o'clock with oximetry error => now again i.O. Performing FlexQ position calibration carrying out Oximetry Complete Test => i.O. => Halogen lamp at 3750 => Lamp adjusted several times Kal1 I.O. | On arrival device without oximetry error, but with FlexQ positioning error check Cal results => Kal from 04.12. 15:09 o'clock to 06.12. 02:10 o'clock with oximetry error => now again i.O. Performing FlexQ position calibration carrying out Oximetry Complete Test => i.O. => Halogen lamp at 3750 => Lamp adjusted several times Kal1 I.O. | Adjustment | HW | Oxi module | Lamp unit
ABL815 FLEX | Trotz Kalibrierung kommt Fehler wieder Austausch FlexQ Modul Inkonsistente Software angepasst Positionskalibierung | Trotz Kalibrierung kommt Fehler wieder Austausch FlexQ Modul Inkonsistente Software angepasst Positionskalibierung | SWO-69722 | Yes | Despite calibration, error comes again replacing FlexQ module inconsistent software adapted position calibration | Despite calibration, error comes again replacing FlexQ module inconsistent software adapted position calibration | Re classified RMED | FLEXQ Install
AQT90 FLEX | Neuen Sacnner eingebaut und justiert. | Neuen Sacnner eingebaut und justiert. | SWO-69716 | No | New Sacnner built-in and adjusted. | New Sacnner built-in and adjusted. | Defect | HW | Storage Module | Barcode reader, Storage wheel
AQT90 FLEX | Optical Unit und Processwheel ausgebaut und Cups entfernt Incubatorwheel ausgebaut und Cups entfernt Feder verriegelung Inletwheel wieder montiert (war ausgehängt) Lösungspack mit Lot SV-05 installiert Reinigung Funktionskontrolle i.o | Optical Unit und Processwheel ausgebaut und Cups entfernt Incubatorwheel ausgebaut und Cups entfernt Feder verriegelung Inletwheel wieder montiert (war ausgehängt) Lösungspack mit Lot SV-05 installiert Reinigung Funktionskontrolle i.o | SWO-69692 | No | Optical Unit and Processwheel removed and cups removed Incubatorwheel removed and cups removed spring locking Inletwheel reassembled (was suspended) solution pack with lot SV-05 installed cleaning function control i.o | Optical Unit and Processwheel removed and cups removed Incubatorwheel removed and cups removed spring locking Inletwheel reassembled (was suspended) solution pack with lot SV-05 installed cleaning function control i.o | Defect | HW | Process Module | Process wheel | 5684
AQT90 FLEX | Analysator geöffnet. TnT Reagenzien (Lot10563) wurde nach PunchOut nicht wieder in Storage Wheel gezogen. Reagenzien entfernt, kein Bruch der Haltenasen ersichtlich. Schieber der Reagenzien ohne Beanstandung. Initialisierung. Nachlegen neuer Reagenzien. Analysator messbereit | Analysator geöffnet. TnT Reagenzien (Lot10563) wurde nach PunchOut nicht wieder in Storage Wheel gezogen. Reagenzien entfernt; kein Bruch der Haltenasen ersichtlich. Schieber der Reagenzien ohne Beanstandung. Initialisierung. Nachlegen neuer Reagenzien. Analysator messbereit | SWO-69596 | No | Analyzer open. TnT Reagents (Lot10563) was not redrawn into Storage Wheel after PunchOut. reagents removed; no break age of the holding asses. Slide of the reagents without complaint. Initialization. Re-energising new reagents. Analyzer ready for measurement | Analyzer open. TnT Reagents (Lot10563) was not redrawn into Storage Wheel after PunchOut. reagents removed; no break age of the holding asses. Slide of the reagents without complaint. Initialization. Re-energising new reagents. Analyzer ready for measurement | Assisted Replacement | Consumables | Test Kit | Cartridge, TnT | TnT
ABL80 FLEX BASIC | 更换测试卡正常。 | 更换测试卡正常。 | SWO-69943 | No | The replacement test card is OK. | The replacement test card is OK.
ABL80 CO-OX | Found machine communicating, but not delivering data to gastrax. | Name of analyzer in communication settings was wrong. Changed name and got a static address from IT. | CFF74CEFF04CAA46 | No | Found machine communicating, but not delivering data to gastrax. | Name of analyzer in communication settings was wrong. Changed name and got a static address from IT.
ABL80 FLEX BASIC | 更换两个新试剂包修复仪器 | 更换两个新试剂包修复仪器 | SWO-69782 | No | Replace two new reagent kit repair instruments | Replace two new reagent kit repair instruments
ABL80 CO-OX | Inspection of the nit revealed the zero check is out of spec. | Set Zero check per procedure from the latest version service manual. | CFF73FA6F0187859 | No | Inspection of the nit revealed the zero check is out of spec. | Set Zero check per procedure from the latest version service manual.
ABL90 FLEX | versucht Hemolyzer zu spülen, aber ohne Erfolg, gleiche Ergebnis Hemolyzer ausgetauscht, Startup abgewartet, tHb Kal durchgeführt, Testmessungen i.O. | versucht Hemolyzer zu spülen, aber ohne Erfolg, gleiche Ergebnis Hemolyzer ausgetauscht, Startup abgewartet, tHb Kal durchgeführt, Testmessungen i.O. | SWO-69845 | No | tries to flush Hemolyzer, but without success, same result Hemolyzer replaced, startup waited, tHb Kal performed, test measurements i.O. | tries to flush Hemolyzer, but without success, same result Hemolyzer replaced, startup waited, tHb Kal performed, test measurements i.O. | Defect | HW | Hemolyzer | 581
ABL90 FLEX | Gerät überprüft, Software upgrade und tHb Kal durchgeführt | Gerät überprüft, Software upgrade und tHb Kal durchgeführt | SWO-70263 | No | Device checked, software upgrade and tHb Kal performed | Device checked, software upgrade and tHb Kal performed | Assisted Replacement | SP | pCO2
ABL90 FLEX | Gerät überprüft: Einlassnadeln leicht verbogen-> gerichtet, Inlet gereinigt, danach Startphase i.O., Einlassnadeln aus Kundenbestand ersetzt SW Version 3.1MR6 installiert | Gerät überprüft: Einlassnadeln leicht verbogen-> gerichtet, Inlet gereinigt, danach Startphase i.O., Einlassnadeln aus Kundenbestand ersetzt SW Version 3.1MR6 installiert | SWO-69972 | No | Device checked: inlet needles slightly bent-> directed, Inlet cleaned, then start phase i.O., inlet needles replaced by customer stock SW version 3.1MR6 installed | Device checked: inlet needles slightly bent-> directed, Inlet cleaned, then start phase i.O., inlet needles replaced by customer stock SW version 3.1MR6 installed | Defect | HW | Inlet Gasket
ABL90 | Peggy called stating the analyzer keeps asking to replace the solution pack. | Back-Flushed the analyzer. Checked and, replaced the inlet gasket with holder, the inlet gasket connector and, the inlet probe. I reinstalled the solution pack and, realized the SP is expired. I checked the customer's stock and, they have 4 expired solution packs. Shady from the North division will send a new SP by tomorrow. | CFF75ECDF02B301E | No | Peggy called stating the analyzer keeps asking to replace the solution pack. | Back-Flushed the analyzer. Checked and, replaced the inlet gasket with holder, the inlet gasket connector and, the inlet probe. I reinstalled the solution pack and, realized the SP is expired. I checked the customer's stock and, they have 4 expired solution packs. Shady from the North division will send a new SP by tomorrow. | User error | SP | Expired
ABL90 FLEX | ABL90 aufgestellt Schlauchsystem gespült ABL 5 Leihgerät zurück in die LUFU geliefert | ABL90 aufgestellt Schlauchsystem gespült ABL 5 Leihgerät zurück in die LUFU geliefert | SWO-70042 | Yes | ABL90 installed hose system flushed ABL 5 loaner delivered back to the LUFU | ABL90 installed hose system flushed ABL 5 loaner delivered back to the LUFU | Re classified RMED | Re install repaired app.
ABL90 FLEX | LP und SC ern., Gerät kalibriert, QC´s durchgef. | B | SWO-69890 | No | LP and SC ern., device calibrated, QC's through. |  | Assisted Replacement | SP | Leak current
AQT90 FLEX |  | B | SWO-69789 | No |  |  | Defect | HW | Process Module | Optical Unit | 131
AQT90 FLEX |  | B | SWO-69781 | No |  |  | Defect | HW | Process Module | Optical Unit | 131
AQT90 FLEX | 更换新处理包后修复 | 更换新处理包后修复 | SWO-69762 | No | Repair after replacing a new processing package | Repair after replacing a new processing package | Assisted Replacement | Consumables | SP | SP | 1999
ABL800 FLEX | Cleaned all the chamber & tubes & performed 2 point calibration QC done values within the range | B | SWO-70696 | No | Cleaned all the chamber & tubes & performed 2 point calibration QC done values within the range |  | Cleaned | Fluidic pathway | pH/BG Module | Measuring chamber
ABL825 FLEX | Found tubing between pHBG and EI Met modules loose. | Replaced loose tubings. | D00007B52A09EAB3 | No | Found tubing between pHBG and EI Met modules loose. | Replaced loose tubings. | Defect | HW | Basic unit | Tube
ABL810 FLEX | Dummy durch Cl- Elektrode ersetzt, Pumpenschlauch erneuert, Verbindungsschlauch Vorwärmer> Messkammer erneuert, pCO2 Messkammer wird weiter überflutet. Rollenkranz demontiert, Dichtung unterm Rollenkranz ist komplett mit Fett verschmiert. Fett entfernt und 824-647 erneuert, Pumpenkal. bei 91 91. Überflutung der pCO2 Messkammer wurde weniger. 2. Kal. wegen pCO2 Driften wiederholt. QCs haben sich um 2mmHg verbessert. pCO2 Membran erneuert, alt R0566 neu R0672, Kal. i.O. QC in Lvl.1&2 nach unterhalb RiliBÄK Bereich. Mit R0647 membraniert, alle Lvl. QC mittelwertig. | Dummy durch Cl- Elektrode ersetzt, Pumpenschlauch erneuert, Verbindungsschlauch Vorwärmer> Messkammer erneuert, pCO2 Messkammer wird weiter überflutet. Rollenkranz demontiert, Dichtung unterm Rollenkranz ist komplett mit Fett verschmiert. Fett entfernt und 824-647 erneuert, Pumpenkal. bei 91 91. Überflutung der pCO2 Messkammer wurde weniger. 2. Kal. wegen pCO2 Driften wiederholt. QCs haben sich um 2mmHg verbessert. pCO2 Membran erneuert, alt R0566 neu R0672, Kal. i.O. QC in Lvl.1&2 nach unterhalb RiliBÄK Bereich. Mit R0647 membraniert, alle Lvl. QC mittelwertig. | SWO-69938 | No | Dummy replaced by Cl-electrode, pump hose renewed, connecting hose preheater> measuring chamber renewed, pCO2 measuring chamber is further flooded. Roller wreath disassembled, seal under the roller ring is completely smeared with grease. Grease removed and 824-647 renewed, pump scale. at 91 91. Flooding of the pCO2 measuring chamber decreased. 2. Kal. due to pCO2 drifts. QCs have improved by 2mmHg. pCO2 membrane renewed, old R0566 new R0672, Kal. i.O. QC in Lvl.1&2 after below RiliBÄK range. Membraneed with R0647, all Lvl. QC inaverage. | Dummy replaced by Cl-electrode, pump hose renewed, connecting hose preheater> measuring chamber renewed, pCO2 measuring chamber is further flooded. Roller wreath disassembled, seal under the roller ring is completely smeared with grease. Grease removed and 824-647 renewed, pump scale. at 91 91. Flooding of the pCO2 measuring chamber decreased. 2. Kal. due to pCO2 drifts. QCs have improved by 2mmHg. pCO2 membrane renewed, old R0566 new R0672, Kal. i.O. QC in Lvl.1&2 after below RiliBÄK range. Membraneed with R0647, all Lvl. QC inaverage. | Assisted Replacement | Electrode | Membrane | pCO2
ABL825 FLEX |  | B | CFF748D1F01069CC | No |  |  | Assisted Replacement | Electrode | Electrode | pCO2
ABL837 FLEX | Replacement of Oxi pcb did not help. nor Lamp Unit. | Replace Spectrometer. | CFF7517FF05363CE | No | Replacement of Oxi pcb did not help. nor Lamp Unit. | Replace Spectrometer. | Defect | HW | Oxi module | Spectrometer
ABL827 FLEX | Performed a complete oximetry test. Lamp test failed for neon position. Range is 0.100 -2.00 had reading of .09. Replaced oxi module PCB, lamp unit and hemolyzer unit. Still had the same problem. | Replaced the spectrometer unit. | CFF72B72F0089ABF | No | Performed a complete oximetry test. Lamp test failed for neon position. Range is 0.100 -2.00 had reading of .09. Replaced oxi module PCB, lamp unit and hemolyzer unit. Still had the same problem. | Replaced the spectrometer unit. | Defect | HW | Oxi module | Spectrometer
ABL825 FLEX | Checked Flex Q for damage and none was found. | Went into the service mode and re calibrated all of Flex Q. | CFF70205F05885D4 | No | Checked Flex Q for damage and none was found. | Went into the service mode and re calibrated all of Flex Q. | Adjustment | HW | FlexQ Module | Barcode
ABL825 FLEX | Error 316 neon number to low .098 | replaced spectrometer full oxi test et thb calibration value .400 | CFF75315F0485A0B | No | Error 316 neon number to low .098 | replaced spectrometer full oxi test et thb calibration value .400 | Defect | HW | Oxi module | Spectrometer
AQT90 FLEX | PCT Test eingelegt Reinigung und QC durchgeführt | PCT Test eingelegt Reinigung und QC durchgeführt | SWO-69828 | No | PCT test inserted cleaning and QC performed | PCT test inserted cleaning and QC performed | Assisted Replacement | Consumables | Test Kit | Cartridge, PCT | PCT - 211
AQT90 FLEX |  | B | CFF728ACF00E90B8 | No |  |  | Defect | HW | Aspiration Module | Tube | 165-167-1218
ABL80 FLEX Basic version |  | B | CFF88BBB6615843D | No
ABL80 FLEX BASIC | 更换测试卡 | 更换测试卡 | SWO-70253 | No | Replace the test card | Replace the test card
ABL80 FLEX BASIC | 更换测试卡正常。 | 更换测试卡正常。 | SWO-70500 | No | The replacement test card is OK. | The replacement test card is OK.
ABL80 FLEX BASIC | 更换新测试卡修复仪器 | 更换新测试卡修复仪器 | SWO-70320 | No | Replace the new test card repair instrument | Replace the new test card repair instrument
ABL80 FLEX BASIC | 更换新测试卡修复仪器 | 更换新测试卡修复仪器 | SWO-70321 | No | Replace the new test card repair instrument | Replace the new test card repair instrument
ABL80 FLEX BASIC | Tried to Clean SC and SC connector board and reinstall. Not installing SC temp not stable is the massage. PER made for SC replacement. | Zerroing cassette install and try to calibrate. O2 value continuously going down. Sensor connector board Doughtful Need to replace. SC and SP reinstalled. Calibrated working fine. Temporarily Analyser OK. Zerroing calibration not pass. Sensor Board Need to replace. | SWO-70194 | No | Tried to Clean SC and SC connector board and reinstall. Not installing SC temp not stable is the massage. PER made for SC replacement. | Zerroing cassette install and try to calibrate. O2 value continuously going down. Sensor connector board Doughtful Need to replace. SC and SP reinstalled. Calibrated working fine. Temporarily Analyser OK. Zerroing calibration not pass. Sensor Board Need to replace.
ABL80 FLEX CO-OX | Gerät zerlegt und Oximodul ausgebaut, LS Sensor mot Flüssigkeit verunreinigt - ausgebaut und ersetzt, Oximodul gereinigt, alle Fließwege gespült, Neustart und SC initialisiert, QC abgewartet | Gerät zerlegt und Oximodul ausgebaut, LS Sensor mot Flüssigkeit verunreinigt - ausgebaut und ersetzt, Oximodul gereinigt, alle Fließwege gespült, Neustart und SC initialisiert, QC abgewartet | SWO-70200 | No | Device disassembled and oximodul removed, LS sensor mot liquid contaminated - removed and replaced, oximodul cleaned, all flow paths flushed, restart and SC initialized, QC waited | Device disassembled and oximodul removed, LS sensor mot liquid contaminated - removed and replaced, oximodul cleaned, all flow paths flushed, restart and SC initialized, QC waited
ABL90 FLEX | Spülung Fließwege, Dichtung mit Halter ersetzt, Startphase - Fehler 1169 (Lösungen konnten nicht gepumpt werden), LP ersetzt (aus Kundenbestand), Startphase, Kal wiederholt, SC90 ersetzt (Kundenbestand), Kal u. QC i.O. | Spülung Fließwege, Dichtung mit Halter ersetzt, Startphase - Fehler 1169 (Lösungen konnten nicht gepumpt werden), LP ersetzt (aus Kundenbestand), Startphase, Kal wiederholt, SC90 ersetzt (Kundenbestand), Kal u. QC i.O. | SWO-70462 | No | Flushing flow paths, seal with holder replaced, start-up phase - error 1169 (solutions could not be pumped), LP replaced (from customer stock), start phase, sc90 replaced (customer stock), Kal & QC i.O. | Flushing flow paths, seal with holder replaced, start-up phase - error 1169 (solutions could not be pumped), LP replaced (from customer stock), start phase, sc90 replaced (customer stock), Kal & QC i.O. | Assisted Replacement | SC | Wettability
ABL90 FLEX PLUS | LP getauscht und einige Tests durchgeführt, | LP getauscht und einige Tests durchgeführt, | SWO-70071 | No | LP and carried out some tests, | LP and carried out some tests, | Assisted Replacement | SP | Leak current
TCM4 |  | B | SWO-69956 | No
ABL90 FLEX | 更换新试剂包后修复 | 更换新试剂包后修复 | SWO-70012 | No | Repair after replacing a new kit | Repair after replacing a new kit | Assisted Replacement | SP | Pump cal error
ABL90 FLEX | 更换测试卡正常。 | 更换测试卡正常。 | SWO-70028 | No | The replacement test card is OK. | The replacement test card is OK. | Assisted Replacement | SC | Reference
AQT90 FLEX | 更换636-550一个正常。(by andrew) | 更换636-550一个正常。(by andrew) | SWO-69978 | No | Replace 636-550 one normal. (by andrew) | Replace 636-550 one normal. (by andrew) | Defect | HW | Process Module | Process Unit sensor | 5784
AQT90 FLEX | 管路松动，重新连接(by andrew) | 管路松动，重新连接(by andrew) | SWO-70246 | No | The line is loose and reconnection (by andrew) | The line is loose and reconnection (by andrew) | Defect | HW | Aspiration Module | Tube
AQT90 FLEX | 更换903-210一个正常。(by andrew) | 更换903-210一个正常。(by andrew) | SWO-69583 | No | Replace the 903-210 one normal. (by andrew) | Replace the 903-210 one normal. (by andrew) | Defect | HW | Process Module | Cup dryer, PCB | 168
ABL815 FLEX | 541 Solution pump error 539 EI Met module pump error 538 Ph BG pump error | Replaced solution pump tube. Completed protein remove. Replaced pCO2 membrane from customer stock as pCO2 was drifting. Replaced Glucose membrane from customer stock as QC was out on level 3. Completed all service tests. | CFF84A3266275D51 | No | 541 Solution pump error 539 EI Met module pump error 538 Ph BG pump error | Replaced solution pump tube. Completed protein remove. Replaced pCO2 membrane from customer stock as pCO2 was drifting. Replaced Glucose membrane from customer stock as QC was out on level 3. Completed all service tests. | Defect | HW | El/Met Module | Pump tubing
ABL825 FLEX | Instrument was not giving results for thb on calibration. Suspect is malfunctioning oxi pcb. | Replaced oxi pcb. Ran thb cal. Verified instrument operation. | CFF974DCDC20A9BF | No | Instrument was not giving results for thb on calibration. Suspect is malfunctioning oxi pcb. | Replaced oxi pcb. Ran thb cal. Verified instrument operation. | Defect | HW | Oxi module | PCB
ABL825 FLEX | Skiftet pumpe slange | Skiftet pumpe slange | SWO-71513 | No | Changed pump HOSE | Changed pump HOSE | Defect | HW | Fluidic manifold | Tubing
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-71562 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL80 FLEX BASIC | 更换测试卡正常。 | 更换测试卡正常。 | SWO-70294 | No | The replacement test card is OK. | The replacement test card is OK.
ABL80 CO-OX |  | B | CFF9AA3EDC49C3B3 | No
ABL80 FLEX BASIC | Zerroing cassette install and try to calibrate. O2 value continuously going down. Sensor connector board Doughtful Need to replace. SC and SP reinstalled. Calibrated working fine. Temporarily Analyser OK. Zerroing calibration not pass. Sensor Board Need to replace. | Zerroing cassette install and try to calibrate. O2 value continuously going down. Sensor connector board Doughtful Need to replace. SC and SP reinstalled. Calibrated working fine. Temporarily Analyser OK. Zerroing calibration not pass. Sensor Board Need to replace. | SWO-70442 | No | Zerroing cassette install and try to calibrate. O2 value continuously going down. Sensor connector board Doughtful Need to replace. SC and SP reinstalled. Calibrated working fine. Temporarily Analyser OK. Zerroing calibration not pass. Sensor Board Need to replace. | Zerroing cassette install and try to calibrate. O2 value continuously going down. Sensor connector board Doughtful Need to replace. SC and SP reinstalled. Calibrated working fine. Temporarily Analyser OK. Zerroing calibration not pass. Sensor Board Need to replace.
ABL80 FLEX BASIC | 更换测试卡正常 | 更换测试卡正常 | SWO-70302 | No | Replace the test card ok | Replace the test card ok
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-70358 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL80 CO-OX | Customer performed setup as per instructions provided. Found all levels set to other . Reinstalled qc levels. | Instructed customer to run linearity in slot slot four. | CFFA0BFADC1E9788 | No | Customer performed setup as per instructions provided. Found all levels set to other . Reinstalled qc levels. | Instructed customer to run linearity in slot slot four.
ABL90 |  | B | CFFEF6E8B406EBDD | No |  |  | Defect | HW | Oxi module | Pinch Valve tube
ABL90 | customer has an external printer installed this is causing the internal printer to malfunction | replaced internal printer printer worked for test print then stopped working disconnected external printer and still not working rebooted and disconnected external printer internal printer worked when external printer connected internal stopped working ; had to open clam shell for external to work | CFF9A911DC43FDAD | No | customer has an external printer installed this is causing the internal printer to malfunction | replaced internal printer printer worked for test print then stopped working disconnected external printer and still not working rebooted and disconnected external printer internal printer worked when external printer connected internal stopped working ; had to open clam shell for external to work | Unknown | Printer setup
ABL90 |  | B | CFF9DE86DC207F5E | No |  |  | Assisted Replacement | SC | Glu
ABL90 | Customer inidcated level 3 External to be out.Customer had moved unit from one local to a different.During move pO2 corrected itself. | Uased SCs' and SPs' for troubleshooting. | D002AD7D1658F04B | No | Customer inidcated level 3 External to be out. Customer had moved unit from one local to a different. During move pO2 corrected itself. | Uased SCs' and SPs' for troubleshooting. | No fault found
ABL90 FLEX | Sensorkassette getauscht. Barometer zurückgesetzt. Laut Auskunft TSP Radiometer GmbH liegt der Fehler an falschen Zielwerten. Ab Lot UB02 ist Fehler behoben. | Sensorkassette getauscht. Barometer zurückgesetzt. Laut Auskunft TSP Radiometer GmbH liegt der Fehler an falschen Zielwerten. Ab Lot UB02 ist Fehler behoben. | SWO-70414 | No | Sensor cartridge. Barometer reset. According to information TSP Radiometer GmbH, the error is due to incorrect target values. As of Lot UB02, bugs are fixed. | Sensor cartridge. Barometer reset. According to information TSP Radiometer GmbH, the error is due to incorrect target values. As of Lot UB02, bugs are fixed. | Assisted Replacement | SP | pCO2
ABL90 FLEX | 更换测试卡，试剂包正常。 | 更换测试卡，试剂包正常。 | SWO-70244 | No | Replace the test card and the reagent package is OK. | Replace the test card and the reagent package is OK. | Assisted Replacement | SC | Na+
ABL90 FLEX | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-70362 | No | Repair after replacing a new test card | Repair after replacing a new test card | Assisted Replacement | SC | Reference
ABL835 FLEX | pH sensitivity low at 48%. | Cleaned electrode contacts. Checked valves. Reseated connectors to pH BG module. No help. Ordered and replaced pH electrode. | CFF9D1AADC439B47 | No | pH sensitivity low at 48%. | Cleaned electrode contacts. Checked valves. Reseated connectors to pH BG module. No help. Ordered and replaced pH electrode. | Assisted Replacement | Electrode | Electrode | pH
ABL825 FLEX | Checked halogen lamp intensity. | Adjusted halogen lamp voltage to specifications. | CFF9A544DC0A73E0 | No | Checked halogen lamp intensity. | Adjusted halogen lamp voltage to specifications. | Adjustment | HW | Oxi module | Lamp unit
ABL825 FLEX | tried selecting sampler 2 with barcode and would not accept. | Calibration of flex Q tray | CFF98B7DDC24ADBC | No | tried selecting sampler 2 with barcode and would not accept. | Calibration of flex Q tray | Adjustment | HW | FlexQ Module | Positioning
ABL815 FLEX | Alle Serviceprogramme für FlexQ Modul durchgeführt => Sampler Sensor Werte waren außer Bereich Spritzenerkennung auf allen Positionen dann i.O. | Alle Serviceprogramme für FlexQ Modul durchgeführt => Sampler Sensor Werte waren außer Bereich Spritzenerkennung auf allen Positionen dann i.O. | SWO-70492 | No | All service programs performed for FlexQ module => Sampler sensor values were except area syringe detection on all positions then i.O. | All service programs performed for FlexQ module => Sampler sensor values were except area syringe detection on all positions then i.O. | Adjustment | HW | FlexQ Module | Positioning
AQT90 FLEX | Guided customer for EDTI tubes.Run the cleaning tube 5 time & run the sample test in EDTI tube. results get without any error massage. | Guided customer for EDTI tubes.Run the cleaning tube 5 time & run the sample test in EDTI tube. results get without any error massage. | SWO-70624 | No | Guided customer for EDTI tubes. Run the cleaning tube 5 time & run the sample test in EDTI tube. results get without any error massage. | Guided customer for EDTI tubes. Run the cleaning tube 5 time & run the sample test in EDTI tube. results get without any error massage. | Cleaned | HW | Aspiration Module | HCT Sensor
AQT90 FLEX | Fehlersuche am Analysator, Temperatur Überprüfung der Optical Unit auf Serviceprogrammebene alle umstehenden Geräte von AQT weggeschoben, Filter überprüft->nicht verstaubt, keine Besserung dadurch erzielt, Optical Unit SN364R0169N013 defekt->erneuert nach Manual gegen SN364R8209N002, alle Parameter neu kalibriert, Reinigung mit Cleaning Solution, LQCs und Systemüberprüfung bestanden, Testmessung NT-proBNP und D-Dimer ok System-und Funktionskontrolle i.O. | Fehlersuche am Analysator, Temperatur Überprüfung der Optical Unit auf Serviceprogrammebene alle umstehenden Geräte von AQT weggeschoben, Filter überprüft->nicht verstaubt, keine Besserung dadurch erzielt, Optical Unit SN364R0169N013 defekt->erneuert nach Manual gegen SN364R8209N002, alle Parameter neu kalibriert, Reinigung mit Cleaning Solution, LQCs und Systemüberprüfung bestanden, Testmessung NT-proBNP und D-Dimer ok System-und Funktionskontrolle i.O. | SWO-70304 | No | Troubleshooting on the analyzer, temperature check of the optical unit at the service program level all surrounding devices pushed away from AQT, filter checked->not dusty, no improvement achieved by it, Optical Unit SN364R0169N013 defective->renewed according to manual against SN364R8209N002, all parameters recalibrated, cleaning passed with Cleaning Solution, LQCs and system verification, test measurement NT-proBNP and D-Dimer ok system and function control i.O. | Troubleshooting on the analyzer, temperature check of the optical unit at the service program level all surrounding devices pushed away from AQT, filter checked->not dusty, no improvement achieved by it, Optical Unit SN364R0169N013 defective->renewed according to manual against SN364R8209N002, all parameters recalibrated, cleaning passed with Cleaning Solution, LQCs and system verification, test measurement NT-proBNP and D-Dimer ok system and function control i.O. | Defect | HW | Process Module | Optical Unit | 135
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-70582 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL80 FLEX BASIC | 清理排线后，仪器恢复(by andrew) | 清理排线后，仪器恢复(by andrew) | SWO-70268 | No | Instrument recovery (by andrew) after cleaning the line | Instrument recovery (by andrew) after cleaning the line
ABL90 |  | B | CFFB090E5228605A | No |  |  | Assisted Replacement | SP | pH
ABL90 FLEX | Den Sachverhalt Alesp mitgeteilt. | Den Sachverhalt Alesp mitgeteilt. | SWO-70651 | No | The facts of the case were communicated to Alesp. | The facts of the case were communicated to Alesp. | Setup | HIS/LIS | Warnings
ABL90 FLEX PLUS |  | B | CFFEAF26B44C5AA9 | Yes |  |  | Re classified RMED | Logistical upSWO
ABL90 FLEX PLUS | Sensor eingesetzt, Kalibration, QC Kontrolle | Sensor eingesetzt, Kalibration, QC Kontrolle | SWO-70626 | No | Sensor used, calibration, QC control | Sensor used, calibration, QC control | No fault found
ABL90 FLEX | 更换测试卡正常。 | 更换测试卡正常。 | SWO-71026 | No | The replacement test card is OK. | The replacement test card is OK. | Assisted Replacement | SC | Glu/lac
ABL90 | Limpeza do circuito, ajuste dos SLs e calibração das bombas, nova calibração e passagem dos QCs, testes com realização de amostras. | B | CFFAE231522953DA | No | Cleaning the circuit, adjusting the SLs and calibrating the pumps, new calibration and passage of the QCs, tests with sample realization. |  | Cleaned | Fluidic pathway
AQT90 FLEX | 更换903-210一个修复仪器(by andrew) | 更换903-210一个修复仪器(by andrew) | SWO-70047 | No | Replace 903-210 a repair instrument (by andrew) | Replace 903-210 a repair instrument (by andrew) | Defect | HW | Process Module | Cup dryer, PCB | 168
ABL825 FLEX | Fellow FSE advised me of this as he tried to troubleshoot over the phone. Performed Power Supply check. AU lead 3 (-12 + - 0.6V) reading -12.83V. | Replaced power supply, cleaned AU fuses. Replaced glucose membrane to stabilize glucose. Inspected sample path and found very small clot under reference electrode, as well as small accumulation of blood in pH well. | CFFAECC6524C612B | No | Fellow FSE advised me of this as he tried to troubleshoot over the phone. Performed Power Supply check. AU lead 3 (-12 + - 0.6V) reading -12.83V. | Replaced power supply, cleaned AU fuses. Replaced glucose membrane to stabilize glucose. Inspected sample path and found very small clot under reference electrode, as well as small accumulation of blood in pH well. | Defect | HW | Power Supply | Off/On
ABL837 FLEX |  | B | CFFAE20D52031E1D | No |  |  | Adjustment | HW | Oxi module | Lamp unit
ABL825 FLEX |  | B | CFFEF8E2B404EBE2 | No |  |  | Defect | HW | FlexQ Module | Tray
ABL825 FLEX | CONFIRMED ISSUE WHEN ONSITE | REPLACED OXI MODULE PCB & NEON LAMP. | CFFAD87C520F9B5E | No | CONFIRMED ISSUE WHEN ONSITE | REPLACED OXYGEN MODULE PCB & NEON LAMP. | Defect | HW | Oxi module | Neon Lamp
ABL825 FLEX |  | B | D0002C5F2A29D7DC | No |  |  | Assisted Replacement | Electrode | Electrode | Lac
ABL825 FLEX | Customers unit and second new unit compare favorably. | Replaced measuring chambers in El Met module. Cleaned electrode contacts. Replaced pH sealing and remembraned reference electrode. Replaced electrode module pump tubes. Lactate and potassium began to drift on cals. Cleaned remaining electrode contacts. | CFFB193F52522192 | No | Customers unit and second new unit compare favorably. | Replaced measuring chambers in El Met module. Cleaned electrode contacts. Replaced pH sealing and remembraned reference electrode. Replaced electrode module pump tubes. Lactate and potassium began to drift on cals. Cleaned remaining electrode contacts. | Defect | HW | El/Met Module | Measuring chamber
AQT90 FLEX |  | B | CFFED37BB408400F | Yes |  |  | Re classified RMED
ABL1 | Installed new SC and SP and provide training for consumable replacement. | Installed new SC and SP and provide training for consumable replacement. | SWO-72038 | No | Installed new SC and SP and provide training for consumable replacement. | Installed new SC and SP and provide training for consumable replacement.
ABL90 | Mutiple errors with customer, including pump error and finally a blank screen. No SP's available as -R- didn't ship. Troubleshoot errors requiring SC and SP's. | Replaced above parts | CFFCA3A5C8627574 | Yes | Mutiple errors with customer, including pump error and finally a blank screen. No SP's available as -R- didn't ship. Troubleshoot errors requiring SC and SP's. | Replaced above parts | Re classified RMED | Logistical upSWO
ABL800 BASIC |  | B | SWO-70709 | No |  |  | Defect | HW | Basic unit | Reagent illumination, right
ABL815 FLEX | Fehlersuche, Neujustage Hb-Lampe, tHb-Kalibrierung, Funktionskontrolle | Fehlersuche, Neujustage Hb-Lampe, tHb-Kalibrierung, Funktionskontrolle | SWO-70735 | No | Troubleshooting, readjustment Hb-lamp, tHb calibration, function control | Troubleshooting, readjustment Hb-lamp, tHb calibration, function control | Adjustment | HW | Oxi module | Lamp unit
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-70761 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL90 FLEX | 进行血红蛋白定标（by Denis） | 进行血红蛋白定标（by Denis） | SWO-71315 | No | Hemoglobin calibration (by Denis) | Hemoglobin calibration (by Denis) | Adjustment | tHb cal
ABL90 FLEX | 更换新测试卡 | 更换新测试卡 | SWO-71333 | No | Replace a new test card | Replace a new test card | Assisted Replacement | SC | Cl
AQT90 FLEX | 更换910-304一个正常。 | 更换910-304一个正常。 | SWO-70531 | No | Replace 910-304 one normal. | Replace 910-304 one normal. | Defect | HW | Storage Module | Barcode reader, Storage wheel | BR1
AQT90 FLEX | 拆备用机上的903-171一个正常。（denis 20151214 该机器报警Too few counts from LS2,Missing liquid for needle wash，进样针弯曲，导致液路故障，更换进样针后故障排除） | 拆备用机上的903-171一个正常。（denis 20151214 该机器报警Too few counts from LS2,Missing liquid for needle wash，进样针弯曲，导致液路故障，更换进样针后故障排除） | SWO-70572 | No | Dismantling the 903-171 on the standby machine is normal. (denis 20151214 This machine alarmToo counts from LS2, Missing liquid for needle wash, the sample needle bends, causing the liquid circuit failure, after replacing the sample needle troubleshooting) | Dismantling the 903-171 on the standby machine is normal. (denis 20151214 This machine alarmToo counts from LS2, Missing liquid for needle wash, the sample needle bends, causing the liquid circuit failure, after replacing the sample needle troubleshooting) | Defect | HW | Aspiration Module | Bend needle | 3412
AQT90 FLEX | 更换636-550一个正常。 | 更换636-550一个正常。 | SWO-70528 | No | Replace 636-550 one normal. | Replace 636-550 one normal. | Defect | HW | Process Module | Process Unit sensor
ABL80 FLEX BASIC | 更换测试卡正常。 | 更换测试卡正常。 | SWO-71029 | No | The replacement test card is OK. | The replacement test card is OK.
ABL90 FLEX | Einstellungen überprüft, kein Fehler festgestellt. Commit2CDDrive.bat ausgeführt, die Uhrzeit bleibt korrekt. Uhrzeit durch weiteren Neustart überprüft, keine Veränderung. | Einstellungen überprüft, kein Fehler festgestellt. Commit2CDDrive.bat ausgeführt, die Uhrzeit bleibt korrekt. Uhrzeit durch weiteren Neustart überprüft, keine Veränderung. | SWO-71018 | No | Settings checked, no error detected. Commit2CDDrive.bat, the time remains correct. Time checked by further restart, no change. | Settings checked, no error detected. Commit2CDDrive.bat, the time remains correct. Time checked by further restart, no change. | Setup | Time Setting
ABL90 | SV-04 SP is giving unable to pump errors | Replaced SP. | CFFF29B0B413DD12 | No | SV-04 SP is giving unable to pump errors | Replaced SP. | Assisted Replacement | SP | FT
ABL90 | Customer had replaced several SPs' with no luck, and no SCs'. | Unit went green, passed Internal QC after running diagnostics. Pressure,Manifold, etc. | CFFEA8EEB40C0316 | No | Customer had replaced several SPs' with no luck, and no SCs'. | Unit went green, passed Internal QC after running diagnostics. Pressure,Manifold, etc. | Unknown
ABL90 FLEX | Inlet connector erneuert LP, SC, LP ersetzt und jeweils Startup abgewartet QC Lvl 3 immer mit pCO2 RiliBÄK Verletzung - um 15:55Uhr wurde die QC bestanden | Inlet connector erneuert LP, SC, LP ersetzt und jeweils Startup abgewartet QC Lvl 3 immer mit pCO2 RiliBÄK Verletzung - um 15:55Uhr wurde die QC bestanden | SWO-70995 | No | Inlet connector renewed LP, SC, LP replaced and each startup awaited QC Lvl 3 always with pCO2 RiliBÄK violation - at 15:55 the QC was passed | Inlet connector renewed LP, SC, LP replaced and each startup awaited QC Lvl 3 always with pCO2 RiliBÄK violation - at 15:55 the QC was passed | Assisted Replacement | SP | Pump cal error
ABL90 | pO2 is out low on Level 1 external QC lot 139, high on Level 2 external QC lot 144. I had customer change sensor cassette with pO2 external QC failures to other analyzer and problem was replicated on the other analyzer. | Customer changes sensor cassette and QC is ok. | CFFECF2DB4394AE3 | No | pO2 is out low on Level 1 external QC lot 139, high on Level 2 external QC lot 144. I had customer change sensor cassette with pO2 external QC failures to other analyzer and problem was replicated on the other analyzer. | Customer changes sensor cassette and QC is ok. | Assisted Replacement | SC | pO2
ABL90 | The display has a crack down the right side. It is coming up with the needing the embedded password changed. Spoke with Mike at techsupport and he has never seen this. Thinks probably computer damage from being dropped. Spoke with ASM about getting replacement. Barb Fouse (lab manager) refusing analyzer. Will open techline case for out of box failure. | Analyzer being returned on RGA Y15- and Fedex tracking # 8050 4333 6006. Replacement ordered. | CFFEB7C1B43363B6 | No | The display has a crack down the right side. It is coming up with the needing the embedded password changed. Spoke with Mike at techsupport and he has never seen this. Thinks probably computer damage from being dropped. Spoke with ASM about getting replacement. Barb Fouse (lab manager) refusing analyzer. Will open techline case for out of box failure. | Analyzer being returned on RGA Y15- and Fedex tracking # 8050 4333 6006. Replacement ordered. | Defect | HW | Display
ABL90 FLEX PLUS | Druckerpapier getauscht. Mehrere Ausdrucke durchgeführt mit Frau Engelmayr. Ausdrücke waren wieder i.O. Glu wurde leserlich ausgedruckt. ABL ist Messbereit Statusampel und alle Parameter auf grün und Ausdruck funktioniert. | Druckerpapier getauscht. Mehrere Ausdrucke durchgeführt mit Frau Engelmayr. Ausdrücke waren wieder i.O. Glu wurde leserlich ausgedruckt. ABL ist Messbereit Statusampel und alle Parameter auf grün und Ausdruck funktioniert. | SWO-71006 | No | Printer paper exchanged. Several prints carried out with Ms. Engelmayr. Expressions were again I.O. Glu was printed legibly. ABL is ready to measure status light and all parameters on green and expression works. | Printer paper exchanged. Several prints carried out with Ms. Engelmayr. Expressions were again I.O. Glu was printed legibly. ABL is ready to measure status light and all parameters on green and expression works. | Replaced | Printer paper
ABL90 FLEX | 更换新测试卡 | 更换新测试卡 | SWO-71085 | No | Replace a new test card | Replace a new test card | Assisted Replacement | SC | Lac
ABL90 FLEX | 更换测试卡正常。 | 更换测试卡正常。 | SWO-71780 | No | The replacement test card is OK. | The replacement test card is OK. | Assisted Replacement | SC | Na+
ABL800 FLEX | Checked with the new pressure transducer.Still Showing the same issue. Gas module need to be replace. | B | SWO-71982 | No | Checked with the new pressure transducer. Still Showing the same issue. Gas module need to be replace. |  | Defect | HW | Gas Module | Gas module
ABL835 FLEX | SW opdateret til ver. 6.16 Nye QC grænser indlæst | SW opdateret til ver. 6.16 Nye QC grænser indlæst | SWO-70991 | No | SW updated to ver. 6.16 New QC limits Loaded | SW updated to ver. 6.16 New QC limits Loaded | Assisted Replacement | Electrode | Electrode | pCO2
ABL825 FLEX |  | B | CFFEF4F7B438821D | No |  |  | Unknown | Regul?
ABL825 FLEX | Fehlersuche, Austausch defektes Teil, Funktionskontrolle | Fehlersuche, Austausch defektes Teil, Funktionskontrolle | SWO-70980 | No | Troubleshooting, replacement of defective part, function control | Troubleshooting, replacement of defective part, function control | Defect | HW | pH/BG Module | Valve, air
ABL825 FLEX | Gerät überprüft, Gasventilmembran EL Met Modul erneuert, Kal und QC o.k. | Gerät überprüft, Gasventilmembran EL Met Modul erneuert, Kal und QC o.k. | SWO-71046 | No | Device checked, gas valve membrane EL Met module renewed, Kal and QC ok | Device checked, gas valve membrane EL Met module renewed, Kal and QC ok | Defect | HW | El/Met Module | Valve, air | Membrane
ABL810 FLEX | complete oxi test verification of error codes | restarted the unit | CFFE952BB4082D5A | No | complete oxi test verification of error codes | restarted the unit | Defect | HW | Oxi module | PCB | missing trigger
ABL825 FLEX | Glucose dérive sur calibration client a remplacé plusieur membrane | Remplacé electrode glucose et lactate installé logiciel 6.16 et installé nouvelle membrane de reference | CFFE8477B4395F73 | No | Glucose drifts on customer calibration replaced more membrane | Replaced electrode glucose and lactate installed software 6.16 and installed new reference membrane | Assisted Replacement | Electrode | Electrode | Glu
ABL835 FLEX | Bemeneti egység tisztítás, Képernyő vizes papirtörlővel való tisztítása. Kalibrálás, QC mérés A köldökből vérgázmérések miatt alvadékszűrő használata javasolt kappilláris mérésekor. | B | SWO-70964 | No | Input unit cleaning, cleaning of the screen with an aqueous paper towel. Calibration, QC measurement The navel is recommended for the use of A clot filter due to blood gas measurements. |  | Defect | HW | DMS_PC | Touch Screen
AQT90 FLEX |  | B | CFFED0C6B415400B | No |  |  | Defect | HW | Aspiration Module | Needle Wash Unit | 1169
ABL80 FLEX BASIC | 更换测试卡正常。 | 更换测试卡正常。 | SWO-71030 | No | The replacement test card is OK. | The replacement test card is OK.
ABL80 FLEX BASIC | 更换测试卡正常。 | 更换测试卡正常。 | SWO-71778 | No | The replacement test card is OK. | The replacement test card is OK.
ABL90 |  | B | D0025B6F16013091 | No |  |  | Assisted Replacement | SP | Pump cal error
ABL90 FLEX | Einstellungen überprüft und auf CF-Karte gespeichert | Einstellungen überprüft und auf CF-Karte gespeichert | SWO-71275 | No | Settings checked and stored on CF card | Settings checked and stored on CF card | Setup | Time Setting
ABL90 |  | B | D00249821663207A | Yes |  |  | Re classified RMED | Replace SP
ABL90 |  | B | CFFFE7052A33B57E | No |  |  | Defect | HW | Sensor interface PCB |  | temp
ABL90 FLEX | Rep.-Ersatzgerät aufgestellt, Konfig. eingespielt, SC und SP inst. --> Neustart Ersatzteil bestellt | Rep.-Ersatzgerät aufgestellt, Konfig. eingespielt, SC und SP inst. --> Neustart Ersatzteil bestellt | SWO-71200 | No | Rep. replacement device set up, Config. recorded, SC and SP inst. --> Restart spare part ordered | Rep. replacement device set up, Config. recorded, SC and SP inst. --> Restart spare part ordered | Unknown | Revisit required
ABL90 | Tube des d coupé | remplacé tube de la trousse de maintenace flush transport system | CFFFE6E02A3047F9 | No | Tube of the coupe | replaced tube of the handtenacious flush transport system | Defect | HW | Waste tube
ABL90 FLEX | LP ersetzt (Lot: UB-01), mehrfach kalibriert, Kal i.O., QC | B | SWO-71288 | No | LP replaced (lot: UB-01), multiple calibrated, Kal i.O., QC |  | Assisted Replacement | SP | pCO2
ABL90 FLEX PLUS | Hemolyzer vervangen. | Hemolyzer vervangen. | SWO-71136 | No | Hemolyzer replaced. | Hemolyzer replaced. | Defect | HW | Hemolyzer
AQT90 FLEX | 仪器校准，重新定标正常。 | 仪器校准，重新定标正常。 | SWO-71059 | No | The instrument is calibrated and re-calibrated normally. | The instrument is calibrated and re-calibrated normally. | Defect | HW | Process Module | Optical Unit
ABL800 FLEX | Replaced Gas Module and Machine is in under observation | B | SWO-71994 | No | Replaced Gas Module and Machine is in under observation |  | Revisit required
ABL835 FLEX | Checked peer data. Confirmed altitude correct. Customer been having problem for months. Checked electrodes for black o-ring. Another machine in hospital having similar problems. | Customer tried membraning electrode. PCO2 still low outside of peer and on the edge of package insert range for both lot 1 and lot 2. I replaced electrode and membrane. That had no effect so I removed them. Upgraded software to version 6.15. Similar numbers for qc are being reported for analyzer in NICU 754R2044N001. I have performed any service on that analyzer. | CFFFFF202A05F07E | No | Checked peer data. Confirmed altitude correct. Customer been having problem for months. Checked electrodes for black o-ring. Another machine in hospital having similar problems. | Customer tried membraning electrode. PCO2 still low outside of peer and on the edge of package insert range for both lot 1 and lot 2. I replaced electrode and membrane. That had no effect so I removed them. Upgraded software to version 6.15. Similar numbers for qc are being reported for analyzer in NICU 754R2044N001. I have performed any service on that analyzer. | Assisted Replacement | Electrode | Electrode | pCO2
ABL800 BASIC |  | B | SWO-71069 | No |  |  | Defect | HW | AutoCheck Module | Barcode scanner
ABL820 FLEX | Checked all logs to verify error. Found external QC failing PH with unstable errors. Calibrations were failing due to sensitivity errors and customer replaced the reference electrode membrane which fixed the sensitivity problem. Found leakage in the PH measuring chamber. | Replaced the PH seal ring and ran calibrations and QC | D0000FB02A0EAAAE | No | Checked all logs to verify error. Found external QC failing PH with unstable errors. Calibrations were failing due to sensitivity errors and customer replaced the reference electrode membrane which fixed the sensitivity problem. Found leakage in the PH measuring chamber. | Replaced the PH seal ring and ran calibrations and QC | Assisted Replacement | Electrode | Electrode | pH Sealing ring
ABL810 FLEX | pCO2 Membran durch Anwender erneuert. Anderen Pumpenschlauch eingesetzt, pCO2 Messkammer weiter voll Spüllösung. Alle Messkammern gereinigt und dekontaminiert, mehrfach kalibriert und alle Lvl. QC mehrfach gemessen, pCO2 QCs alle fast mittelwertig. | pCO2 Membran durch Anwender erneuert. Anderen Pumpenschlauch eingesetzt, pCO2 Messkammer weiter voll Spüllösung. Alle Messkammern gereinigt und dekontaminiert, mehrfach kalibriert und alle Lvl. QC mehrfach gemessen, pCO2 QCs alle fast mittelwertig. | SWO-71290 | No | pCO2 membrane renewed by users. Other pump hose used, pCO2 measuring chamber continues to be full flushing solution. All measuring chambers cleaned and decontaminated, calibrated several times and all Lvl. QC measured multiple times, pCO2 QCs all almost average. | pCO2 membrane renewed by users. Other pump hose used, pCO2 measuring chamber continues to be full flushing solution. All measuring chambers cleaned and decontaminated, calibrated several times and all Lvl. QC measured multiple times, pCO2 QCs all almost average. | Assisted Replacement | Electrode | Membrane | pCO2
ABL837 FLEX | Errors present upon arrival.Inspected valves, tubings, electrodes. | Ran service sensor cal= cleared errors. | D0000BE02A0EEAB7 | No | Errors present upon arrival. Inspected valves, tubings, electrodes. | Ran service sensor cal= cleared errors. | Adjustment | HW | El/Met Module | LS Calibration
ABL835 FLEX | Bemeneti tű csere. Kalibrálás, QC mérés. | B | SWO-71208 | No | Input needle replacement. Calibration, QC measurement. |  | Defect | HW | Inlet module | Probe bend
AQT90 FLEX | Een andere barcode reader gemonteerd. Nu gaat het uitlezen van diverse labels wel goed en snel. | B | SWO-71183 | No | Another barcode reader mounted. Now reading various labels is going to be good and fast. |  | Defect | HW | Aspiration Module | Barcode reader, Inlet
ABL80 FLEX BASIC | 更换测试卡 | 更换测试卡 | SWO-71359 | No | Replace the test card | Replace the test card
ABL80 FLEX BASIC | 更换新测试卡 | 更换新测试卡 | SWO-72089 | No | Replace a new test card | Replace a new test card
ABL80 CO-OX |  | B | D0011C50A0405EF9 | No
ABL80 FLEX BASIC |  | B | SWO-71322 | No
ABL90 FLEX | Einlassdichtung --> Gummi aus Halterung gerutscht. Einlassdichtung war seit 08.12.2015 im Gerät | Einlassdichtung --> Gummi aus Halterung gerutscht. Einlassdichtung war seit 08.12.2015 im Gerät | SWO-71429 | No | Inlet seal --> Rubber slipped out of bracket. Inlet seal had been in the device since 08.12.2015 | Inlet seal --> Rubber slipped out of bracket. Inlet seal had been in the device since 08.12.2015 | Defect | HW | Inlet Gasket
ABL90 | Verified error by running analzyer. Suspect compromised pinch tubing. | Installed annual maintenance kit and inlet probe. Tested analyzer, issues resolved. | D0016627A00D008D | No | Verified error by running analzyer. Suspect compromised pinch tubing. | Installed annual maintenance kit and inlet probe. Tested analyzer, issues resolved. | Defect | HW | Oxi module | Pinch Valve tube
ABL90 | Dépèché par collègue dû aux problèmes survenus lors de sa visite de service hier. Sensibilité du glucose très basse, tests de service démontrant aussi des problèmes au niveau du circuit pour le détecteur de pCO2. | carte d'interface des détecteurs. Problème de stabilité de température lors de la remise en route, calibration des capteurs de température. Sens glu et lac instable. Changé cassette. Sens glu et lac toujour instable. Test service simulateur de capteurs échouée. Changé carte de contrôle module fluidique. Mise à jour 3.1MR7. Test service simulateur de capteurs échouée. Réinstallé réactifs. Changé pack de solutions. Performé test service simulateur de capteurs sur instrument fonctionel, échoué. | D001906FA025A52C | No | Struggled by a colleague due to the problems that occurred during his service visit yesterday. Very low glucose sensitivity, service tests also showing circuit problems for the pCO2 detector. | interface map of the detectors. Temperature stability problem when restarted, calibration of temperature sensors. Sense of glue and unstable lake. Changed cassette. Sense of glue and unstable lake. Test service sensor simulator failed. Changed fluid module control card. Update 3.1MR7. Test service sensor simulator failed. Reinstalled reactive. Changed solution pack. Performed test service sensor simulator on functional instrument, failed. | Unknown
ABL90 FLEX | Spülung des Analysators, SC reinitialisiert - keine Besserung, neue SC90 installiert, akz. Kal abgewartet, mehrfach kalibriert, QC i.O. | B | SWO-71538 | No | Rinsing of the analyzer, SC reinitialized - no improvement, new SC90 installed, akz. Kal waited, multiple calibrated, QC i.O. |  | Assisted Replacement | SC | pCO2
ABL90 FLEX PLUS | Gerät sowie dessen Einstellungen komplett überprüft, keine Unregelmäßigkeiten festgestellt. ServiceDump gezogen, Datensicherung erstellt, kalibriert und tHb-Kal durchgeführt. Drei QC Messungen OK, QC Statistik ebenfalls. Frau Dr. Mach-Völker auf die unbedingte Verwendung unserer 45µL Clinitubes hingewiesen, sie klärt das mit dem Kinderzimmer, da nicht auszuschließen ist, dass von dort mit anderen Probennehmern gemessen wurde! SW V3.1MR7 installiert. | Gerät sowie dessen Einstellungen komplett überprüft, keine Unregelmäßigkeiten festgestellt. ServiceDump gezogen, Datensicherung erstellt, kalibriert und tHb-Kal durchgeführt. Drei QC Messungen OK, QC Statistik ebenfalls. Frau Dr. Mach-Völker auf die unbedingte Verwendung unserer 45µL Clinitubes hingewiesen, sie klärt das mit dem Kinderzimmer, da nicht auszuschließen ist, dass von dort mit anderen Probennehmern gemessen wurde! SW V3.1MR7 installiert. | SWO-71507 | No | device and its settings, no irregularities detected. ServiceDump dragged, data backup created, calibrated and tHb-Kal performed. Three QC measurements OK, QC statistics also. Dr. Mach-Völker pointed out the unconditional use of our 45-L Clinitubes, she clarifies this with the children's room, since it cannot be excluded that from there was measured with other samplers! SW V3.1MR7 installed. | device and its settings, no irregularities detected. ServiceDump dragged, data backup created, calibrated and tHb-Kal performed. Three QC measurements OK, QC statistics also. Dr. Mach-Völker pointed out the unconditional use of our 45-L Clinitubes, she clarifies this with the children's room, since it cannot be excluded that from there was measured with other samplers! SW V3.1MR7 installed. | User error | Pre analytical
ABL80 FLEX BASIC | Patient test done by 3 times for same sample fluctuated values comes in light, not reliable. SC cleaned and Reinstalled along with SP (By reset) Problem solved. Trained them to SC flush reinstall. Recommendation- Ask them to use distal water to wash reused cups, Not to use Capillary mode, and If problem persist. Then purchase QC for Final Call. | Patient test done by 3 times for same sample fluctuated values comes in light, not reliable. SC cleaned and Reinstalled along with SP (By reset) Problem solved. Trained them to SC flush reinstall. Recommendation- Ask them to use distal water to wash reused cups, Not to use Capillary mode, and If problem persist. Then purchase QC for Final Call. | SWO-71654 | No | Patient test done by 3 times for same sample fluctuated values comes in light, not reliable. SC cleaned and Reinstalled along with SP (By reset) Problem solved. Trained them to SC flush reinstall. Recommendation- Ask them to use distal water to wash reused cups, Not to use Capillary mode, and If problem persist. Then purchase QC for Final Call. | Patient test done by 3 times for same sample fluctuated values comes in light, not reliable. SC cleaned and Reinstalled along with SP (By reset) Problem solved. Trained them to SC flush reinstall. Recommendation- Ask them to use distal water to wash reused cups, Not to use Capillary mode, and If problem persist. Then purchase QC for Final Call.
TCM CombiM |  | B | SWO-71387 | No
ABL90 FLEX | 更换试剂包正常。 | 更换试剂包正常。 | SWO-71774 | No | The replacement kit is OK. | The replacement kit is OK. | Assisted Replacement | SP | FT
ABL90 FLEX | 更换新试剂包 | 更换新试剂包 | SWO-71393 | No | Replace a new kit | Replace a new kit | Assisted Replacement | SP | Pump cal error
ABL90 | SE COMPRUEBA EL ESTADO DE FUNCIONAMIENTO DEL EQUIPO, COMPROBACION DE REGISTROS DE DATOS, SE EXTRAE EL SP Y SC, SE LIMPIA CON AGUA DESTILADA TODO EL CIRCUITO NEUMÁTICO, SE COMPRUEBA EL CORRECTO FUNCIONAMIENTO POR LOS PROGRAMAS DE SERVICE, SE INICIALIZA EL EQUIPO. | B | D0016691A015AC8D | No | THE STATE OF OPERATION OF THE EQUIPMENT IS CHECKED, DATA REGISTRATION CHECK, SP AND SC IS EXTRACTED, CLEANING WITH DESTILATED WATER ALL THE TIRE CIRCUIT, THE CORRECT OPERATION BY SERVICE PROGRAMS IS CHECKED, THE EQUIPMENT IS INITIALIZED. |  | Cleaned | Fluidic pathway
AQT90 FLEX | 更换光学单元903-364一个后修复（denis 20151216 该机器报警1024 定标不能通过，清洁无效，更换光学单元） | 更换光学单元903-364一个后修复（denis 20151216 该机器报警1024 定标不能通过，清洁无效，更换光学单元） | SWO-70784 | No | Replace the optical unit 903-364 a post repair (denis 20151216 This machine alarm 1024 calibration can not pass, cleaning invalid, replacement optical unit) | Replace the optical unit 903-364 a post repair (denis 20151216 This machine alarm 1024 calibration can not pass, cleaning invalid, replacement optical unit) | Defect | HW | Process Module | Optical Unit | 124
AQT90 FLEX | 清理处理盘后修复 (by denis) | B | SWO-71570 | No | Clean-up after-disk repair (by denis) |  | Cleaned | HW | Process Module | Process wheel | 129
AQT90 FLEX |  | B | SWO-71356 | No |  |  | Defect | HW | Process Module | Process Unit sensor | 5784
ABL800 BASIC | CPU Unit ersetzt, Kira deinstallieren, nach Sicherung einspielen startet die Kira nicht mehr, Kira erneut installiert, Konfiguration per Hand eingestellt, Kal THB Kal und QC durchgeführt | CPU Unit ersetzt, Kira deinstallieren, nach Sicherung einspielen startet die Kira nicht mehr, Kira erneut installiert, Konfiguration per Hand eingestellt, Kal THB Kal und QC durchgeführt | SWO-71458 | No | CPU unit replaced, Kira uninstall, after backup play in the Kira no longer starts, Kira reinstalled, configuration set by hand, Kal THB Kal and QC performed | CPU unit replaced, Kira uninstall, after backup play in the Kira no longer starts, Kira reinstalled, configuration set by hand, Kal THB Kal and QC performed | Defect | HW | DMS_PC | CPU
ABL825 FLEX | Ca. Elektrode und Messkammer gprüft, Ca. Elektrode neu membrniert, Funktionskontrolle Ok | Ca. Elektrode und Messkammer gprüft, Ca. Elektrode neu membrniert, Funktionskontrolle Ok | SWO-71417 | No | Approx. electrode and measuring chamber tested, approx. electrode re-membrned, function control Ok | Approx. electrode and measuring chamber tested, approx. electrode re-membrned, function control Ok | Assisted Replacement | Electrode | Membrane | Ca
ABL825 FLEX | Luft Ventile BGA modul gereingt, HJW durchgef., Einlassnippel ersetzt. Gerät kalibriert und QC durchgef I.O. | Luft Ventile BGA modul gereingt, HJW durchgef., Einlassnippel ersetzt. Gerät kalibriert und QC durchgef I.O. | SWO-71588 | No | Air valves BGA module rmed, HJW through, inlet nipples replaced. calibrated and QC through I.O. | Air valves BGA module rmed, HJW through, inlet nipples replaced. calibrated and QC through I.O. | Defect | HW | pH/BG Module | Valve, air
ABL825 FLEX | Neue Software V 6.16 installiert. AQC Insert angepasst. | B | SWO-71503 | Yes | New software V 6.16 installed. AQC Insert. |  | Re classified RMED | SWU
ABL825 FLEX | Verbindung zum OSM ist in Ordnung aber die Pat. berichte werden in OSM nicht gesehen. | Verbindung zum OSM ist in Ordnung aber die Pat. berichte werden in OSM nicht gesehen. | SWO-71571 | Yes | Connection to OSM is fine but the Pat. reports are not seen in OSM. | Connection to OSM is fine but the Pat. reports are not seen in OSM. | Re classified RMED | Installation
AQT90 FLEX | AQT überprüft. Einlassnadel verbogen. Nadel erneuert. Nadelcounter zurückgesetzt. Systemüberprüfung und -reinigung durchgeführt. 3x LQC's Level2 (CRP, DDimer und TnI) durchgeführt. | AQT überprüft. Einlassnadel verbogen. Nadel erneuert. Nadelcounter zurückgesetzt. Systemüberprüfung und -reinigung durchgeführt. 3x LQC's Level2 (CRP, DDimer und TnI) durchgeführt. | SWO-71488 | No | AQT. Inlet needle bent. Needle renewed. Needle counter reset. system verification and cleaning. 3x LQC's Level2 (CRP, DDimer and TnI). | AQT. Inlet needle bent. Needle renewed. Needle counter reset. system verification and cleaning. 3x LQC's Level2 (CRP, DDimer and TnI). | Defect | HW | Aspiration Module | Bend needle | 4794
ABL80 FLEX BASIC | 更换测试卡正常。 | 更换测试卡正常。 | SWO-73700 | No | The replacement test card is OK. | The replacement test card is OK.
ABL90 FLEX | Teile im Umfang einer Jahreswartung ersetzt. Oxitubing stark Blutverkrustet, vermutlich dadurch Vakuumfehler. Einlaufverfahren danach sofort erfolgreich. | Teile im Umfang einer Jahreswartung ersetzt. Oxitubing stark Blutverkrustet, vermutlich dadurch Vakuumfehler. Einlaufverfahren danach sofort erfolgreich. | SWO-71763 | No | parts in the scope of an annual maintenance. Oxitubing is heavily blood-encrusted, presumably due to vacuum errors. immediately successful inlet procedure. | parts in the scope of an annual maintenance. Oxitubing is heavily blood-encrusted, presumably due to vacuum errors. immediately successful inlet procedure. | Defect | HW | Oxi module | Pinch Valve tube
ABL90 |  | B | D002C20D1623351F | No |  |  | Assisted Replacement | SC | Lac
ABL90 | Examined data logs. Found all QC levels failing on pO2. All QC levels were ran without temperature correction. | Observed customer run QC levels. Adjusted auto temp to actual room temp. All QC levels came in OK. Verified operation of ABL90. | D00581A7021E5B27 | No | Examined data logs. Found all QC levels failing on pO2. All QC levels were ran without temperature correction. | Observed customer run QC levels. Adjusted auto temp to actual room temp. All QC levels came in OK. Verified operation of ABL90. | User error | Ext QC Temp
ABL90 FLEX PLUS | wegen aktueller LP Problematik auch durch Materialwechsel keine Besserung erreicht, Kunde( Frau Dr. Fischer) über die aktuelle Situation informiert das dass Gerät derzeit mit eingeschränktem pCO2 funktioniert und desshalb bei Verwendung eigenverantwortlich beurteilt werden muß bineine alösung in Sicht ist - ist einverstanden | wegen aktueller LP Problematik auch durch Materialwechsel keine Besserung erreicht, Kunde( Frau Dr. Fischer) über die aktuelle Situation informiert das dass Gerät derzeit mit eingeschränktem pCO2 funktioniert und desshalb bei Verwendung eigenverantwortlich beurteilt werden muß bineine alösung in Sicht ist - ist einverstanden | SWO-71623 | No | due to current LP problems no improvement is achieved due to material change, customer (Dr. Fischer) informs about the current situation that the device currently works with limited pCO2 and is therefore assessed independently when used must be a solution in sight - agrees | due to current LP problems no improvement is achieved due to material change, customer (Dr. Fischer) informs about the current situation that the device currently works with limited pCO2 and is therefore assessed independently when used must be a solution in sight - agrees | Assisted Replacement | SP | pCO2
ABL90 FLEX PLUS | mehrfach LP und SC ersetzt, ohne Erfolg, während der Nachschicht ist auch das gestern nachmittag ersetzte Material wieder fehlerhaft bei pCO2 QC Level3, Datensicherung durchgeführt und Herrn Dr. Schunk über sie aktuelle Situation unterrichtet, er ist einverstanden das Gerät mit besonderer Vorsicht pCO2 Messwerten gegenüber weiter zu betreiben, fordert aber eine schnelle Lösung des Problems | mehrfach LP und SC ersetzt, ohne Erfolg, während der Nachschicht ist auch das gestern nachmittag ersetzte Material wieder fehlerhaft bei pCO2 QC Level3, Datensicherung durchgeführt und Herrn Dr. Schunk über sie aktuelle Situation unterrichtet, er ist einverstanden das Gerät mit besonderer Vorsicht pCO2 Messwerten gegenüber weiter zu betreiben, fordert aber eine schnelle Lösung des Problems | SWO-71710 | No | Repeatedly replaced LP and SC, without success, during the post-shift the material replaced yesterday afternoon is again faulty in pCO2 QC Level3, data backup performed and Dr. Schunk informs them about the current situation, he agrees the device with special care to continue to operate pCO2 readings, but calls for a quick solution to the problem | Repeatedly replaced LP and SC, without success, during the post-shift the material replaced yesterday afternoon is again faulty in pCO2 QC Level3, data backup performed and Dr. Schunk informs them about the current situation, he agrees the device with special care to continue to operate pCO2 readings, but calls for a quick solution to the problem | Assisted Replacement | SP | pCO2
AQT90 FLEX |  | B | SWO-71548 | No |  |  | Defect | HW | Process Module | Cup dryer, PCB | 168
AQT90 FLEX | 更换910-298一个正常。(denis 20151218 该机器与代理商工程师沟通，该机器触摸屏位置漂移，经过校正后依然漂移，同时触摸屏不灵敏，按键无法识别，更换显示屏后机器正常工作） | 更换910-298一个正常。(denis 20151218 该机器与代理商工程师沟通，该机器触摸屏位置漂移，经过校正后依然漂移，同时触摸屏不灵敏，按键无法识别，更换显示屏后机器正常工作） | SWO-71349 | No | Replace 910-298 one normal. (denis 20151218 The machine communicates with dealer engineers, the machine touch screen position drifts, is still drifted after correction, and the touch screen is not sensitive, the keys are not recognized, the machine works properly after changing the display) | Replace 910-298 one normal. (denis 20151218 The machine communicates with dealer engineers, the machine touch screen position drifts, is still drifted after correction, and the touch screen is not sensitive, the keys are not recognized, the machine works properly after changing the display) | Defect | HW | Analyzer Control Module | Display
ABL800 BASIC | 代替機を設置の上でQAキット交換作業実施。 | B | SWO-71989 | Yes | Qa kit replacement work was implemented after installing an alternative machine. |  | Re classified RMED | PM
ABL825 FLEX |  | B | D00A152FDA287908 | No |  |  | Assisted Replacement | Electrode | Electrode | pCO2
ABL815 FLEX |  | B | D002D6271619B82E | Yes |  |  | Re classified RMED | SWU
ABL815 FLEX |  | B | D002D6C2160BB831 | Yes |  |  | Re classified RMED | SWU
AQT90 FLEX | Vorort Analysator funktioniert nach mehrmaligem ein und ausschalten wieder. Solution Pack ST-09 durch Solution Pack SR-15 ausgetauscht da auf beiden Analysatoren dasselbe Problem Zeitgleich aufgetreten ist | Vorort Analysator funktioniert nach mehrmaligem ein und ausschalten wieder. Solution Pack ST-09 durch Solution Pack SR-15 ausgetauscht da auf beiden Analysatoren dasselbe Problem Zeitgleich aufgetreten ist | SWO-71702 | No | Suburb analyser works after several times on and off again. Solution Pack ST-09 replaced by Solution Pack SR-15 because the same problem occurred on both analyzers at the same time | Suburb analyser works after several times on and off again. Solution Pack ST-09 replaced by Solution Pack SR-15 because the same problem occurred on both analyzers at the same time | Assisted Replacement | Consumables | SP | SP
ABL80 FLEX BASIC | 更换新测试卡修复仪器 | 更换新测试卡修复仪器 | SWO-71825 | No | Replace the new test card repair instrument | Replace the new test card repair instrument
ABL80 FLEX BASIC | 更换测试卡正常。 | 更换测试卡正常。 | SWO-72153 | No | The replacement test card is OK. | The replacement test card is OK.
ABL80 FLEX BASIC | 更换新测试卡修复仪器 | 更换新测试卡修复仪器 | SWO-71798 | No | Replace the new test card repair instrument | Replace the new test card repair instrument
ABL90 |  | B | D00A2B1CDA0F9B6F | No |  |  | Assisted Replacement | SP | unknown
ABL90 |  | B | D0075392EE327B30 | Yes |  |  | Re classified RMED | Customer replaced SP twice
ABL90 |  | B | D003AF7C8C0519F7 | Yes |  |  | Re classified RMED | PM
ABL90 | Problem started when customer replaced sensor cassette. | Flushed system thouroughly and tried calibrating. Machine would not come in. Ordered new sensor cassette and solution pack. Replaced sensor cassette and calibrated. Left machine to calibrate. It didnt come in, so customer ended up changing out sensor cassette again. It finally came in and customer ran quality control. Followed up with customer 12 22 to make sure machine was still functioning properly. | D003C6DE8C523850 | No | Problem started when customer replaced sensor cassette. | Flushed system thouroughly and tried calibrating. Machine would not come in. Ordered new sensor cassette and solution pack. Replaced sensor cassette and calibrated. Left machine to calibrate. It didnt come in, so customer ended up changing out sensor cassette again. It finally came in and customer ran quality control. Followed up with customer 12 22 to make sure machine was still functioning properly. | Assisted Replacement | SC | Ca
ABL90 FLEX | nacheinander SC, LP, Einlaß quergetauscht - Fehler besteht weiterhin - beim 2. Gerät werden Zielwerte gemessen Oxi Schlauch gewechselt - stabile pO2 Messung im Lvl C | nacheinander SC, LP, Einlaß quergetauscht - Fehler besteht weiterhin - beim 2. Gerät werden Zielwerte gemessen Oxi Schlauch gewechselt - stabile pO2 Messung im Lvl C | SWO-71911 | No | one after the other SC, LP, inlet cross-exchanged - error still exists - with the 2nd device target values are measured Oxi hose changed - stable pO2 measurement in Lvl C | one after the other SC, LP, inlet cross-exchanged - error still exists - with the 2nd device target values are measured Oxi hose changed - stable pO2 measurement in Lvl C | Defect | HW | Oxi module | Pinch Valve tube
ABL90 FLEX | Hemolyzer getauscht, SP und SC inst. --> Neustart Oxi-complete-test durchgeführt--> i.O | Hemolyzer getauscht, SP und SC inst. --> Neustart Oxi-complete-test durchgeführt--> i.O | SWO-71875 | No | Hemolyzer exchanged, SP and SC inst. --> Restart Oxi-complete-test performed--> i.O | Hemolyzer exchanged, SP and SC inst. --> Restart Oxi-complete-test performed--> i.O | Defect | HW | Hemolyzer
ABL90 FLEX PLUS | Roma.ini Einträge ergänzt. Touchscreen kalibriert. Anweder- und Patientenbarcode werden erkannt. Kalibration und QC's durchgeführt. | Roma.ini Einträge ergänzt. Touchscreen kalibriert. Anweder- und Patientenbarcode werden erkannt. Kalibration und QC's durchgeführt. | SWO-72232 | Yes | Roma.ini entries added. touchscreen calibrated. Ander- and patient barcode are detected. calibration and QC's. | Roma.ini entries added. touchscreen calibrated. Ander- and patient barcode are detected. calibration and QC's. | Re classified RMED | Roma.ini | Barcode Reader
ABL90 FLEX | 更换902-980一个和903-194一个后修复 （by denis 20151218 该机器在维修过程中又出现新报警，泵定标失败，新故障主要因为试剂包泵管问题，需要更换新试剂包，由于距离公司比较远，需要回公司却试剂包，该机器需要进一步维修） | 更换902-980一个和903-194一个后修复 （by denis 20151218 该机器在维修过程中又出现新报警，泵定标失败，新故障主要因为试剂包泵管问题，需要更换新试剂包，由于距离公司比较远，需要回公司却试剂包，该机器需要进一步维修） | SWO-71619 | No | Replacement of 902-980 one and 903-194 a post-repair (by denis 20151218 the machine in the maintenance process and new alarm, pump calibration failure, new failure mainly because of the reagent package pump tube problems, the need to replace the new reagent package, because of the company is far away, the need to return to the company reagent package, The machine requires further repair) | Replacement of 902-980 one and 903-194 a post-repair (by denis 20151218 the machine in the maintenance process and new alarm, pump calibration failure, new failure mainly because of the reagent package pump tube problems, the need to replace the new reagent package, because of the company is far away, the need to return to the company reagent package, The machine requires further repair) | Defect | HW | Hemolyzer
ABL825 FLEX |  | B | D003C6968C194809 | Yes |  |  | Re classified RMED | Installation
ABL825 FLEX | Messkammer und Schlauchsystem geprüft, Lac. Elektrode (E7077 RP01) geprüft und ausgetauscht, Einlassnadel ersetzt, Funktionskontrolle ok | B | SWO-71916 | No | Measuring chamber and hose system tested, Lac. Electrode (E7077 RP01) tested and replaced, inlet needle replaced, function control ok |  | Assisted Replacement | Electrode | Electrode | Lac
ABL835 FLEX | Nach mehreren Neustarts ist die Funktion des Bildschirms gegeben, Computereinheit geprüft und ein Ersatzteil bestellt, Elektroden Cl. Glu. Und Lac. Neu membraniert, Funktionskontrolle ok | Fr. Domberg bittet um einen Rückruf und dringenden Service! | SWO-71915 | No | After several restarts, the function of the screen is given, computer unit is checked and a spare part ordered, electrodes Cl. Glu. And Lac. Newly membraneed, function control ok | Ms. Domberg asks for a callback and urgent service! | Defect | HW | DMS_PC | Touch Screen
ABL825 FLEX | Oxymetrie Einheit ersetzt, Service THb Kal. PCO2 QC Fehler Ref. Elektrode, pCO 2 und Laktat Membraniert. | Oxymetrie Einheit ersetzt, Service THb Kal. PCO2 QC Fehler Ref. Elektrode, pCO 2 und Laktat Membraniert. | SWO-71971 | No | Oxymetry unit replaced, service THb Kal. PCO2 QC Error Ref. electrode, pCO 2 and lactate membranes. | Oxymetry unit replaced, service THb Kal. PCO2 QC Error Ref. electrode, pCO 2 and lactate membranes. | Defect | HW | Oxi module | Phase regulation
ABL827 FLEX |  | B | D01CD308C43A5105 | No
ABL825 FLEX | 1. OXI hardware error, error 579. 2. Neon lamp low, error 316. | 1. Replace hemolyzer unit. 2. Replace spectrometer unit. | D003E64D8C41FEB3 | No | 1. OXI hardware error, error 579. 2. Neon lamp low, error 316. | 1. Replace hemolyzer unit. 2. Replace spectrometer unit. | Defect | HW | Oxi module | Spectrometer
ABL825 FLEX | Spectrometer and lamp unit failed. | Replace Spectrometer and Lamp units. Perform lamp alignment . Perform Complete OXI Test. | D00406C58C32FEB8 | No | Spectrometer and lamp unit failed. | Replace Spectrometer and Lamp units. Perform lamp alignment . Perform Complete OXI Test. | Defect | HW | Oxi module | Spectrometer
AQT90 FLEX | 機器引き上げ | 機器引き上げ | SWO-72896 | Yes | Equipment pulling up | Equipment pulling up | Re classified RMED | Double entry
AQT90 FLEX | Reset the time to the correct time and performed the maintenance after that. | B | SWO-71984 | No | Reset the time to the correct time and performed the maintenance after that. |  | Assisted Replacement | Consumables | Cleaning Solution | Cleaning Solution
ABL80 FLEX BASIC | 更换新测试卡修复仪器 | 更换新测试卡修复仪器 | SWO-72054 | No | Replace the new test card repair instrument | Replace the new test card repair instrument
ABL80 FLEX BASIC | 更换新测试卡修复仪器 | 更换新测试卡修复仪器 | SWO-72053 | No | Replace the new test card repair instrument | Replace the new test card repair instrument
ABL80 FLEX BASIC |  | B | SWO-72031 | No
AQT90 FLEX |  | B | SWO-72016 | No |  |  | Defect | HW | Analyzer Control Module | Main Board, PCB | 2132
ABL80 FLEX BASIC | 更换新测试卡 | 更换新测试卡 | SWO-72192 | No | Replace a new test card | Replace a new test card
ABL80 FLEX BASIC | 更换测试卡正常。 | 更换测试卡正常。 | SWO-72092 | No | The replacement test card is OK. | The replacement test card is OK.
ABL835 FLEX | Tausch des Oxymetrie Moduls. Überprüfung Einstellung und Nachkorrektur. Durchführung der tHb Kalibration im Service und im Anwenderstatus. | Tausch des Oxymetrie Moduls. Überprüfung Einstellung und Nachkorrektur. Durchführung der tHb Kalibration im Service und im Anwenderstatus. | SWO-72071 | No | Replacement of the oxymetry module. Check setting and post-correction. Execution of the tHb calibration in the service and in the user status. | Replacement of the oxymetry module. Check setting and post-correction. Execution of the tHb calibration in the service and in the user status. | Defect | HW | Oxi module | Oxi module
ABL80 FLEX | Fogyóanyagok felhelyezése (szenzor és oldat-kazetta), Zéró és fűtés ellenőrzése, Kalibráció, QC mérés | B | SWO-72267 | No | Attaching consumables (Sensor and solution cassette), zero and heater control, calibration, QC measurement
ABL80 FLEX BASIC | 更换测试卡正常。 | 更换测试卡正常。 | SWO-72181 | No | The replacement test card is OK. | The replacement test card is OK.
ABL80 FLEX BASIC | 更换测试卡，试剂包正常。 | 更换测试卡，试剂包正常。 | SWO-72358 | No | Replace the test card and the reagent package is OK. | Replace the test card and the reagent package is OK.
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-73046 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL80 FLEX BASIC | 更换测试卡正常。 | 更换测试卡正常。 | SWO-72115 | No | The replacement test card is OK. | The replacement test card is OK.
ABL90 FLEX | Alle Wartungsteile erneuert, Schlauche sowie Dichtungen. Neue Software MR6 aufgespielt(Separater Auftrag). Sollte immer noch keine Besserung auftreten muss die Präanalytik vor Ort genauer betrachtet werden. | Alle Wartungsteile erneuert, Schlauche sowie Dichtungen. Neue Software MR6 aufgespielt(Separater Auftrag). Sollte immer noch keine Besserung auftreten muss die Präanalytik vor Ort genauer betrachtet werden. | SWO-72406 | No | All maintenance parts renewed, hoses as well as seals. New software MR6 applied (separate order). If there is still no improvement, the pre-analysis on site needs to be looked at more closely. | All maintenance parts renewed, hoses as well as seals. New software MR6 applied (separate order). If there is still no improvement, the pre-analysis on site needs to be looked at more closely. | Assisted Replacement | SC | Wettability
ABL90 FLEX | SC gewechselt Startup abgewartet | SC gewechselt Startup abgewartet | SWO-72240 | No | SC changed startup waited | SC changed startup waited | Assisted Replacement | SC | Leak current
ABL90 FLEX | Jahreswartung vorgezogen, Waste Connector Unit gespült, Nadel und Fließwege überprüft, Schläuche und Dichtungen installiert, Sensor Kassette kurz rausgenommen und wieder installiert, Kal mehrfach durchgeführt, Kal und QC i.O | Jahreswartung vorgezogen, Waste Connector Unit gespült, Nadel und Fließwege überprüft, Schläuche und Dichtungen installiert, Sensor Kassette kurz rausgenommen und wieder installiert, Kal mehrfach durchgeführt, Kal und QC i.O | SWO-72306 | No | Annual maintenance preferred, Waste Connector Unit flushed, needle and flow ways checked, hoses and seals installed, sensor cassette briefly removed and reinstalled, Kal performed multiple times, Kal and QC i.O | Annual maintenance preferred, Waste Connector Unit flushed, needle and flow ways checked, hoses and seals installed, sensor cassette briefly removed and reinstalled, Kal performed multiple times, Kal and QC i.O | Cleaned | Fluidic pathway
ABL90 |  | B | D007A064EE11F8E4 | Yes |  |  | Re classified RMED | SWU
ABL90 FLEX | Abdeckung für Scanner gewechselt | B | SWO-72376 | No | Cover changed for scanner |  | User error | Scanner window
ABL90 | Service diags indicated Clot detect errors, Pressure errors, Pump cal, etc. | Ran diagnostics until errors cleared. | D007AEBEEE2E6A1F | No | Service diags indicated Clot detect errors, Pressure errors, Pump cal, etc. | Ran diagnostics until errors cleared. | Cleaned | Fluidic pathway
ABL90 FLEX PLUS |  | B | D00A28C7DA5F2AE6 | No |  |  | Assisted Replacement | SC | unknown
ABL90 FLEX | 更换新测试卡 | 更换新测试卡 | SWO-72211 | No | Replace a new test card | Replace a new test card | Assisted Replacement | SC | SC | SC | 1361
AQT90 FLEX |  | B | SWO-72144 | No |  |  | Defect | HW | Process Module | Optical Unit | 131
AQT90 FLEX | 更换903-364一个正常。 | 更换903-364一个正常。 | SWO-72147 | No | Replace 903-364 one normal. | Replace 903-364 one normal. | Defect | HW | Process Module | Optical Unit | 124
AQT90 FLEX |  | B | SWO-72106 | No |  |  | Defect | HW | Process Module | Optical Unit | 131
AQT90 FLEX | FORMATEO Y BORRADO DE TODOS LOS DATOS DEL EQUIPO, INSTALACIÓN DEL WIND.XPE, INSTALACION DE LA APLICACION CON LA V. 8.8.16, CONFIGURACION DEL EQUIPO, CALIBRACION DE PARÁMETROS, VERIFICACION FINAL CON QC Y MUESTRA DE PACIENTE. | B | D00783E1EE5517C4 | Yes | FORFORMATAND AND DELETE OF ALL EQUIPMENT DATA, WIND INSTALLATION. XPE, INSTALLATION OF APPLICATION WITH V. 8.8.16, EQUIPMENT CONFIGURATION, PARAMETER CALIBRATION, FINAL VERIFICATION WITH QC AND PATIENT SAMPLE. |  | Re classified RMED | SWU
ABL825 FLEX | Oxi Modul geprüft und justiert, Service tHb Kalibration, QC Bereich im Level3 nach Rücksprache mit Herrn Bikker angepasst, Cl Elektrode nen membraniert, Funktionskontrolle ok | B | SWO-72290 | No | Oxi module tested and adjusted, service tHb calibration, QC area in level3 adapted after consultation with Mr. Bikker, Cl electrode nen membrane, function control ok |  | Adjustment | HW | Oxi module | tHb cal
ABL835 FLEX | Opstart Skiftet pH O-ring | Opstart Skiftet pH O-ring | SWO-72254 | No | Start-up changed pH O-ring | Start-up changed pH O-ring | Assisted Replacement | Electrode | Electrode | pH | Sealing ring
ABL827 FLEX | ABL überprüft. OXI-Modul erneuert. (Softwarepflege, LS-Justage, Pumpen-Kalibration, OXI-Complete-Check, Service-tHb-Kal, tHb-Kal und Kal-Verifikation durchgeführt) Gerät kalibriert. QC durchgeführt. | ABL überprüft. OXI-Modul erneuert. (Softwarepflege, LS-Justage, Pumpen-Kalibration, OXI-Complete-Check, Service-tHb-Kal, tHb-Kal und Kal-Verifikation durchgeführt) Gerät kalibriert. QC durchgeführt. | SWO-72246 | No | ABL. OXI module renewed. (Software maintenance, LS-Adjustment, pump calibration, OXI-Complete-Check, Service-tHb-Kal, tHb-Kal and Kal-verification performed) device calibrated. QC. | ABL. OXI module renewed. (Software maintenance, LS-Adjustment, pump calibration, OXI-Complete-Check, Service-tHb-Kal, tHb-Kal and Kal-verification performed) device calibrated. QC. | Defect | HW | Oxi module | Spectrometer
ABL835 FLEX | FlexQ eingebaut, Modul Upgrade Software und Abgleich durchgeführt | FlexQ eingebaut, Modul Upgrade Software und Abgleich durchgeführt | SWO-72213 | Yes | FlexQ built-in, module upgrade software and matching performed | FlexQ built-in, module upgrade software and matching performed | Re classified RMED | Upgrade w FLEXQ
ABL825 FLEX | Gerät überprüft, tHb und Service tHb durchgeführt, Kal und QC o.k, | Gerät überprüft, tHb und Service tHb durchgeführt, Kal und QC o.k, | SWO-72258 | No | device checked, tHb and service tHb performed, Kal and QC o.k, | device checked, tHb and service tHb performed, Kal and QC o.k, | Adjustment | HW | Oxi module | tHb cal
ABL800 BASIC | Fehlersuche, Erneuerung Cl-Membran, Funktionskontrolle | Fehlersuche, Erneuerung Cl-Membran, Funktionskontrolle | SWO-72193 | No | Troubleshooting, renewal Cl-membrane, function control | Troubleshooting, renewal Cl-membrane, function control | Assisted Replacement | Electrode | Membrane | Cl
ABL835 FLEX | Bemeneti egység átöblítése csapvizzel, Kapilláris és fecskendő szennyes elszívó ág átnyomása hipós oldattal. Kalibrálás, QC mérés | B | SWO-72229 | No | Flush input unit with tap water, capillary and syringe fume extraction arm with hippy solution. Calibration, QC measurement |  | Cleaned | Fluidic pathway | Inlet module | Inlet gasket
AQT90 FLEX | done the cleaning and make the cartridge door in right position. | done the cleaning and make the cartridge door in right position. | SWO-72904 | No | done the cleaning and make the cartridge door in right position. | done the cleaning and make the cartridge door in right position. | Adjustment | HW | Cover and Chassis | Storage Door
AQT90 FLEX | Incubatorwheel ausgebaut, Wheel und Gehäuse gereinigt. Cartridgescanner eingestellt. Sensoren von Incubatorgehäuse gereinigt. Abfallkassette getauscht. QC Multicheck Level 2 einmal durchgeführt. --> i.O. AQT wieder Messbereit. | Incubatorwheel ausgebaut, Wheel und Gehäuse gereinigt. Cartridgescanner eingestellt. Sensoren von Incubatorgehäuse gereinigt. Abfallkassette getauscht. QC Multicheck Level 2 einmal durchgeführt. --> i.O. AQT wieder Messbereit. | SWO-72373 | No | Incubatorwheel removed, wheel and housing cleaned. cartridge scanner. sensors from incubator housings. Waste cassette. QC Multicheck Level 2 once performed. --> i.O. AQT ready to measure again. | Incubatorwheel removed, wheel and housing cleaned. cartridge scanner. sensors from incubator housings. Waste cassette. QC Multicheck Level 2 once performed. --> i.O. AQT ready to measure again. | Cleaned | HW | Incubator Module | Incubator, Sensor | Setsafe 22
AQT90 FLEX |  | B | D0076B2DEE1EE83C | Yes |  |  | Re classified RMED | SWU
AQT90 FLEX |  | B | D007BB57EE4F5907 | Yes |  |  | Re classified RMED | SWU
ABL80 FLEX BASIC | 更换测试卡正常。 | 更换测试卡正常。 | SWO-72590 | No | The replacement test card is OK. | The replacement test card is OK.
ABL80 FLEX Basic version |  | B | D008D0686457CB9C | No
ABL80 FLEX BASIC | Sensor connector board replaced. Zerroing calibration done. Found all parameteres well with in range. SC and SP reinstalled and calibrated. Analyser working fine. | Sensor connector board replaced. Zerroing calibration done. Found all parameteres well with in range. SC and SP reinstalled and calibrated. Analyser working fine. | SWO-72518 | No | Sensor connector board replaced. Zerroing calibration done. Found all parameteres well with in range. SC and SP reinstalled and calibrated. Analyser working fine. | Sensor connector board replaced. Zerroing calibration done. Found all parameteres well with in range. SC and SP reinstalled and calibrated. Analyser working fine.
ABL90 FLEX PLUS | Analysatorstatus grün. QC im Herstellerbereich. RiliBÄK wurde abgeschaltet. Lösungspack Lot SV03 vor Ort belassen falls mit bestehenden LP Probleme auftreten. QC Überwachung durch Frau Helberger | Analysatorstatus grün. QC im Herstellerbereich. RiliBÄK wurde abgeschaltet. Lösungspack Lot SV03 vor Ort belassen falls mit bestehenden LP Probleme auftreten. QC Überwachung durch Frau Helberger | SWO-72599 | No | Analyzer status green. QC in the manufacturing sector. RiliBÄK has been switched off. Solution pack Lot SV03 to be left on site if problems occur with existing LP. QC monitoring by Ms Helberger | Analyzer status green. QC in the manufacturing sector. RiliBÄK has been switched off. Solution pack Lot SV03 to be left on site if problems occur with existing LP. QC monitoring by Ms Helberger | Assisted Replacement | SP | pCO2 | QC ranges
AQT90 FLEX |  | B | SWO-72338 | No |  |  | Defect | HW | Process Module | Optical Unit | 131
ABL825 FLEX | Wasteschlauch gekürzt um guten Sitz zu gewährleisten. Ventile und Ventilwege gespült. Ventile auf korrekten Sitz überprüft, in Ordnung. Gerät von Blutaustritt gereinigt. LS-Justagen, Kalibrationen und Qualitätskontrollen in Ordnung. | Wasteschlauch gekürzt um guten Sitz zu gewährleisten. Ventile und Ventilwege gespült. Ventile auf korrekten Sitz überprüft, in Ordnung. Gerät von Blutaustritt gereinigt. LS-Justagen, Kalibrationen und Qualitätskontrollen in Ordnung. | SWO-72552 | No | Waste hose shortened to ensure good fit. valves and valve paths. Valves checked for correct fit, fine. device cleaned of blood leakage. LS adjustments, calibrations and quality controls in order. | Waste hose shortened to ensure good fit. valves and valve paths. Valves checked for correct fit, fine. device cleaned of blood leakage. LS adjustments, calibrations and quality controls in order. | Unknown | HW | Fluidic manifold | Waste Tube
ABL825 FLEX |  | B | D008E921642A0294 | No |  |  | Assisted Replacement | Electrode | Electrode | Na+
AQT90 FLEX | 様子見 | 機器引き上げ | SWO-72898 | No | Wait | Equipment pulling up | Not Cat | Not Cat | Not Cat | Not Cat | 1223
AQT90 FLEX | COMPROBACIÓN DEL REGISTRO DE DATOS DEL EQUIPO, COMPROBACION DEL INTERIOR DEL EQUIPO, TUBOS, FILTROS, UNIDAD DE LAVADO, RUEDA DE PROCESOS, SE VERIFICA EL CORRECTO FUNCIONAMIENTO POR LOS PROGRAMAS DE SERVICE, SE REALIZA SERVICE DUMMP DE TODOS LOS DATOS.SE REALIZA LIMPIEZA CON TUBO. | B | D008C5076445B7D6 | No | CHECKING THE TEAM DATA REGISTRATION, CHECKING THE EQUIPMENT INTERIOR, TUBES, FILTERS, WASHING UNIT, PROCESS WHEEL, THE CORRECT OPERATION BY SERVICE PROGRAMS IS VERIFIED, SERVICE DUMMP IS PERFORMED OF ALL DATOS.SE CLEANING WITH TUBE. |  | Cleaned | Performance | Process Module | Cup wash unit
AQT90 FLEX | Analysator geöffnet. CRP KASSETTE (Lot11201) wurde nach PunchOut nicht wieder in Storage Wheel gezogen. Reagenzien entfernt, kein Bruch der Haltenasen ersichtlich. Messingschiene und Cartridge slider (Toothed Bar ) erneuert. Cup Transport Test mit 16 Cups durchgeführt- ohne Beanstandung. Schieber der Reagenzien ohne Beanstandung. Initialisierung. Nachlegen neuer Reagenzien. Analysator messbereit . Eine Test Messung durchgeführt i.O | Analysator geöffnet. CRP KASSETTE (Lot11201) wurde nach PunchOut nicht wieder in Storage Wheel gezogen. Reagenzien entfernt; kein Bruch der Haltenasen ersichtlich. Messingschiene und Cartridge slider (Toothed Bar ) erneuert. Cup Transport Test mit 16 Cups durchgeführt- ohne Beanstandung. Schieber der Reagenzien ohne Beanstandung. Initialisierung. Nachlegen neuer Reagenzien. Analysator messbereit . Eine Test Messung durchgeführt i.O | SWO-72551 | No | Analyzer open. CRP KASSETTE (Lot11201) was not redrawn into Storage Wheel after PunchOut. reagents removed; no break age of the holding asses. Brass rail and cartridge slider (Toothed Bar) renewed. Cup transport test carried out with 16 cups - without complaint. Slide of the reagents without complaint. Initialization. Re-energising new reagents. Analyzer ready to measure . A test measurement performed i.O | Analyzer open. CRP KASSETTE (Lot11201) was not redrawn into Storage Wheel after PunchOut. reagents removed; no break age of the holding asses. Brass rail and cartridge slider (Toothed Bar) renewed. Cup transport test carried out with 16 cups - without complaint. Slide of the reagents without complaint. Initialization. Re-energising new reagents. Analyzer ready to measure . A test measurement performed i.O | Defect | Consumables | Test Kit | Cartridge, CRP | CRP
ABL80 FLEX BASIC | 连续更换两张测试卡后修复 | 连续更换两张测试卡后修复 | SWO-72780 | No | Repair after replacing two test cards in a row | Repair after replacing two test cards in a row
ABL80 FLEX BASIC | 更换新测试卡修复 | 更换新测试卡修复仪器 | SWO-73015 | No | Replace a new test card repair | Replace the new test card repair instrument
ABL80 FLEX BASIC | 更换新测试卡修复仪器 | 更换新测试卡修复仪器 | SWO-73016 | Yes | Replace the new test card repair instrument | Replace the new test card repair instrument | Re classified RMED | Double entry
ABL80 FLEX BASIC | 更换903-325一个好 | 更换903-325一个好 | SWO-72602 | No | Replace 903-325 a good | Replace 903-325 a good
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-73054 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL80 FLEX BASIC | 更换测试卡正常。 | 更换测试卡正常。 | SWO-72567 | No | The replacement test card is OK. | The replacement test card is OK.
ABL80 FLEX BASIC | 更换新测试卡修复仪器 | 更换新测试卡修复仪器 | SWO-72792 | No | Replace the new test card repair instrument | Replace the new test card repair instrument
ABL90 FLEX | Hemolyzer gewechselt Oxi Complete Test i.O. Kal, tHb Kal durchgeführt gr. Einlassdichtung gewechselt, da sie nicht richtig saß | Hemolyzer gewechselt Oxi Complete Test i.O. Kal, tHb Kal durchgeführt gr. Einlassdichtung gewechselt, da sie nicht richtig saß | SWO-72665 | No | Hemolyzer changed Oxi Complete Test i.O. Kal, tHb Kal performed gr. Inlet seal changed because she was not sitting properly | Hemolyzer changed Oxi Complete Test i.O. Kal, tHb Kal performed gr. Inlet seal changed because she was not sitting properly | Defect | HW | Hemolyzer | 581
ABL90 FLEX | Flüssigkeitsstrecke gespült, SP nochmals getauscht, Einlassdichtung erneuert. Neustart | Flüssigkeitsstrecke gespült, SP nochmals getauscht, Einlassdichtung erneuert. Neustart | SWO-72815 | No | Liquid section rinsed, SP replaced again, inlet seal renewed. Restart | Liquid section rinsed, SP replaced again, inlet seal renewed. Restart | Cleaned | Fluidic pathway
ABL90 FLEX | Telefonat von Hr. Langner mit Kundin: Kunde hat zuerst LP ohne Erfolg erneuert. Danach wurde SC erneuert. Kalibrationen danach ok. Mit Kundin tHB Kalibration telefonisch durchgeführt - ok. QC Fehler vorhanden (auch nach mehrmaliger Kalibration. Telefonat von Hr. geldhauser und Kundin: QC Fehler noch immer vorhanden. Bei Eintreffen war Gerät wieder Messbereit. | Telefonat von Hr. Langner mit Kundin: Kunde hat zuerst LP ohne Erfolg erneuert. Danach wurde SC erneuert. Kalibrationen danach ok. Mit Kundin tHB Kalibration telefonisch durchgeführt - ok. QC Fehler vorhanden (auch nach mehrmaliger Kalibration. Telefonat von Hr. geldhauser und Kundin: QC Fehler noch immer vorhanden. Bei Eintreffen war Gerät wieder Messbereit. | SWO-72710 | Yes | Phone call from Mr. Langner with customer: Customer first renewed LP without success. After that, SC was renewed. Calibrations after that ok. Performed with customer tHB calibration by phone - ok. QC error present (even after multiple calibration. Phone call from Mr. geldhauser and customer: QC error still present. When it arrived, the device was ready for measurement again. | Phone call from Mr. Langner with customer: Customer first renewed LP without success. After that, SC was renewed. Calibrations after that ok. Performed with customer tHB calibration by phone - ok. QC error present (even after multiple calibration. Phone call from Mr. geldhauser and customer: QC error still present. When it arrived, the device was ready for measurement again. | Re classified RMED | Phone support + visit | Courtesy
ABL90 FLEX | 更换新试剂包后修复 | 更换新试剂包后修复 | SWO-72751 | No | Repair after replacing a new kit | Repair after replacing a new kit | Assisted Replacement | SP | pO2 QC
AQT90 FLEX | 清洁进样针传感器上的血凝块，更换新进样针903-171一个，给进样针重新定标后修复 | 清洁进样针传感器上的血凝块，更换新进样针903-171一个，给进样针重新定标后修复 | SWO-72859 | No | Clean the blood clots on the sample needle sensor, replace the new sample needle 903-171 one, repair the sample needle after re-calibration | Clean the blood clots on the sample needle sensor, replace the new sample needle 903-171 one, repair the sample needle after re-calibration | Defect | HW | Aspiration Module | Bend needle
ABL835 FLEX | Computereinheit geprüft, CPU ausgetauscht, Funktionskontrolle ok | Computereinheit geprüft, CPU ausgetauscht, Funktionskontrolle ok | SWO-72650 | No | Computer unit checked, CPU replaced, function control ok | Computer unit checked, CPU replaced, function control ok | Defect | HW | DMS_PC | CPU
ABL825 FLEX | I was unable to duplicate these problems. Customer did show me where a sample was analyzed on 12 23 and the draw date showed 12 22. | Customer will be monitoring. | D00CE5EFC6004A77 | No | I was unable to duplicate these problems. Customer did show me where a sample was analyzed on 12 23 and the draw date showed 12 22. | Customer will be monitoring. | Defect | SW | DMS_PC | Draw Date
ABL800 BASIC | Taken the membranes from the distributor, replced them, done the 2-point calibration, cleaned the inlet casket since some readings are showing 'sample problem', after successful 2-point calibration, run few samples, machine is working fine. | Taken the membranes from the distributor, replced them, done the 2-point calibration, cleaned the inlet casket since some readings are showing 'sample problem', after successful 2-point calibration, run few samples, machine is working fine. | SWO-72916 | No | Taken the membranes from the distributor, replced them, done the 2-point calibration, cleaned the inlet casket since some readings are showing 'sample problem', after successful 2-point calibration, run few samples, machine is working fine. | Taken the membranes from the distributor, replced them, done the 2-point calibration, cleaned the inlet casket since some readings are showing 'sample problem', after successful 2-point calibration, run few samples, machine is working fine. | Assisted Replacement | Electrode | Membrane | Lac
ABL837 FLEX |  | B | D0116DE99E4552A0 | No |  |  | Assisted Replacement | HW | Inlet module | Inlet gasket
ABL827 FLEX | Performed complete OXI service test with no errors. | Flushed hemolyzer with hypochlorite solution and water. | D00A0BBFDA435B74 | No | Performed Complete OXI service test with no errors. | Flushed hemolyzer with hypochlorite solution and water. | Cleaned | HW | Oxi module | Hemolyzer
ABL825 FLEX | Found F2 Fuse on Au board blown. This is the fuse that controls the inactive modules. New fuses blow immediatley on install. | Pulled machine apart. Troubleshot down to find a cable that was kinked. Replaced cable from customer's turned off trade in analyzer. Replaced the fuse. Replaced reference and Ca membrane. Calcium was acting up and reference was due to be replaced in a couple of days. Performed all necessary calibrations to bring analyzer in to service. | D00A4798DA05894C | No | Found F2 Fuse on Au board blown. This is the fuse that controls the inactive modules. New fuses blow immediatley on install. | Pulled machine apart. Troubleshot down to find a cable that was kinked. Replaced cable from customer's turned off trade in analyzer. Replaced the fuse. Replaced reference and Ca membrane. Calcium was acting up and reference was due to be replaced in a couple of days. Performed all necessary calibrations to bring analyzer in to service. | Defect | HW | Basic unit | Cable | Kinked
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-73608 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL80 FLEX BASIC | cleaned the pins and perform the zero check and cleaned the waste path. now working fine. | cleaned the pins and perform the zero check and cleaned the waste path. now working fine. | SWO-72902 | No | cleaned the pins and perform the zero check and cleaned the waste path. now working fine. | cleaned the pins and perform the zero check and cleaned the waste path. now working fine.
ABL90 FLEX | カセット交換 | B | SWO-72873 | No | Cassette replacement |  | Assisted Replacement | SC | Glu
ABL90 FLEX | Einlassdichtung erneuert Kalibration und QC durchgeführt | Einlassdichtung erneuert Kalibration und QC durchgeführt | SWO-72988 | No | Inlet seal renews calibration and QC performed | Inlet seal renews calibration and QC performed | Defect | HW | Inlet Gasket
ABL90 FLEX PLUS |  | B | D00B3469502A9991 | No |  |  | Assisted Replacement | SC | K+
ABL90 FLEX | 更换新试剂包修复仪器 | 更换新试剂包修复仪器 | SWO-72784 | No | Replace the new reagent kit repair instrument | Replace the new reagent kit repair instrument | Assisted Replacement | SP | FT | 183
AQT90 FLEX | 更换580-098和636-550修复了仪器 | 更换580-098和636-550修复了仪器 | SWO-72562 | No | Replacement 580-098 and 636-550 repair instruments | Replacement 580-098 and 636-550 repair instruments | Defect | HW | Process Module | Process Unit sensor | 5784
ABL825 FLEX |  | B | D00B2300505129BB | No |  |  | Defect | HW | Inlet module | Probe bend
ABL800 BASIC | Clean the chamber .Clean the ph electrode. Run protein removal multiple time. Run cleaning multiple time. Replace the Ref membrane. Run sample .Query resolved. calibration pass. | Clean the chamber .Clean the ph electrode. Run protein removal multiple time. Run cleaning multiple time. Replace the Ref membrane. Run sample .Query resolved. calibration pass. | SWO-72887 | No | Clean the chamber . Clean the ph electrode. Run protein removal multiple time. Run cleaning multiple time. Replace the Ref membrane. Run sample . Query resolved. calibration pass. | Clean the chamber . Clean the ph electrode. Run protein removal multiple time. Run cleaning multiple time. Replace the Ref membrane. Run sample . Query resolved. calibration pass. | Cleaned | Fluidic pathway | pH/BG Module | pH
ABL800 BASIC | Perform 2 point calibration calibration.After replace the cal 1 solution, unit is working properly. | Perform 2 point calibration calibration.After replace the cal 1 solution, unit is working properly. | SWO-72882 | No | Perform 2 point calibration calibration. After replace the cal 1 solution, unit is working properly. | Perform 2 point calibration calibration. After replace the cal 1 solution, unit is working properly. | Assisted Replacement | Solution bottle | Solution bottle | Cal 1
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-72907 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL80 FLEX BASIC | 更换测试卡正常。 | 更换测试卡正常。 | SWO-72864 | No | The replacement test card is OK. | The replacement test card is OK.
ABL80 FLEX BASIC | 更换903-001一个正常。 | 更换903-001一个正常。 | SWO-72631 | No | Replace 903-001 one normal. | Replace 903-001 one normal.
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-72901 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL90 FLEX | Sensor justage und Pumpenkal in Serviceprogramm. Lösungspack und Sensorkassette aus dem Kundenbestand ausgetauscht. Alle Schläuche ersetzt. Schlauchsystem war undicht. | Sensor justage und Pumpenkal in Serviceprogramm. Lösungspack und Sensorkassette aus dem Kundenbestand ausgetauscht. Alle Schläuche ersetzt. Schlauchsystem war undicht. | SWO-72978 | No | Sensor adjustment and pump scale in service program. Solution pack and sensor cartridge from the customer's inventory. Replaced all hoses. Hose system was leaking. | Sensor adjustment and pump scale in service program. Solution pack and sensor cartridge from the customer's inventory. Replaced all hoses. Hose system was leaking. | Defect | HW | Waste tube
ABL90 FLEX | 更换试剂包正常。 | 更换试剂包正常。 | SWO-72862 | No | The replacement kit is OK. | The replacement kit is OK. | Assisted Replacement | SP | Inhomogen rinse | 181
ABL90 FLEX | 更换试剂包 | 更换试剂包 | SWO-72857 | No | Replace the reagent kit | Replace the reagent kit | Assisted Replacement | SP | Pump cal error | 1178
ABL90 FLEX | 更换测试卡和试剂包 | 更换测试卡和试剂包 | SWO-72863 | No | Replace test cards and reagent kits | Replace test cards and reagent kits | Assisted Replacement | SC | Reference
AQT90 FLEX | 更换光学单元903-364一个后修复 | 更换光学单元903-364一个后修复 | SWO-72555 | No | Replacement of optical unit 903-364 one post-repair | Replacement of optical unit 903-364 one post-repair | Defect | HW | Process Module | Optical Unit | 131
AQT90 FLEX | 更换903-364、903-119、903-236各一个后修复 | 更换903-364、903-119、903-236各一个后修复 | SWO-72583 | No | Replace 903-364, 903-119, 903-236 each after repair | Replace 903-364, 903-119, 903-236 each after repair | Defect | HW | Process Module | Optical Unit | 131
AQT90 FLEX | 更换光学单元903-364一个，并用定标卡重新定标后修复 | 更换光学单元903-364一个，并用定标卡重新定标后修复 | SWO-72800 | No | Replace the optical unit 903-364 and repair it after re-calibration with the calibration card | Replace the optical unit 903-364 and repair it after re-calibration with the calibration card | Defect | HW | Process Module | Optical Unit | 131
ABL837 FLEX | CPUユニット交換 | B | SWO-73420 | No | CPU unit replacement |  | Defect | HW | DMS_PC | CPU
ABL825 FLEX | Verschmutztes OXI Ventil festgestellt, Ventilanker stark verschmutzt und schwergängig, OXI Ventil gereinigt, Fehler weiter vorhanden. Hemolyzer gereinigt, Lampeneinheit nachgestellt, QC Messungen wieder i.O. | Verschmutztes OXI Ventil festgestellt, Ventilanker stark verschmutzt und schwergängig, OXI Ventil gereinigt, Fehler weiter vorhanden. Hemolyzer gereinigt, Lampeneinheit nachgestellt, QC Messungen wieder i.O. | SWO-72892 | No | Contaminated OXI valve detected, valve anchor heavily polluted and heavy-duty, OXI valve cleaned, fault still present. Hemolyzer cleaned, lamp unit re-enacted, QC measurements again i.O. | Contaminated OXI valve detected, valve anchor heavily polluted and heavy-duty, OXI valve cleaned, fault still present. Hemolyzer cleaned, lamp unit re-enacted, QC measurements again i.O. | Defect | HW | Oxi module | Valve
ABL80 FLEX BASIC | 更换新试剂包后修复 | 更换新试剂包后修复 | SWO-72945 | No | Repair after replacing a new kit | Repair after replacing a new kit
ABL90 FLEX | Reparaturüberbrückungsgerät (090R0733N003) aufgestellt, System und Funktion i.O. | Reparaturüberbrückungsgerät (090R0733N003) aufgestellt, System und Funktion i.O. | SWO-73002 | No | Repair bridging device (090R0733N003) installed, system and function in the | Repair bridging device (090R0733N003) installed, system and function in the | Unknown | Analyzer replaced | CPU
ABL90 FLEX PLUS |  | B | D018F0996243EDE4 | No |  |  | Assisted Replacement | SC | SC | Unknown
AQT90 FLEX | 更换F102 | 更换F102 | SWO-72950 | No | Replace F102 | Replace F102 | Defect | HW | Analyzer Control Module | Fuse | 1999
AQT90 FLEX | 更换636-550一个后修复 | 更换636-550一个后修复 | SWO-72761 | No | Replace 636-550 one post-fix | Replace 636-550 one post-fix | Defect | HW | Process Module | Process Unit sensor | 5784
ABL825 FLEX | Gebrauchte Einlaßnadel eingesetzt, Inlet-Kal und Kal und QC ok | Gebrauchte Einlaßnadel eingesetzt, Inlet-Kal und Kal und QC ok | SWO-72962 | No | Used inlet needle inserted, Inlet-Kal and Kal and QC ok | Used inlet needle inserted, Inlet-Kal and Kal and QC ok | Defect | HW | Inlet module | Probe bend
AQT90 FLEX | E-Teil nicht im Lagerbestand. Leihgerät (838R0050N006) geliefert und konfiguriert. LQC durchgeführt. | E-Teil nicht im Lagerbestand. Leihgerät (838R0050N006) geliefert und konfiguriert. LQC durchgeführt. | SWO-73118 | No | E part not in inventory. (838R0050N006) and configured. LQC. | E part not in inventory. (838R0050N006) and configured. LQC. | Not Cat | Not Cat | Not Cat | Not Cat
ABL90 FLEX | 连续更换4个试剂包后修复 | 连续更换4个试剂包后修复 | SWO-73169 | No | Repair after 4 reagent kits in a row | Repair after 4 reagent kits in a row | Assisted Replacement | SP | Pump cal error
ABL90 FLEX | バッテリー取り付け CPU945へ変更。 | バッテリー取り付け CPU945へ変更。 | SWO-72963 | No | Changed to battery installation CPU945. | Changed to battery installation CPU945. | Defect | HW | CPU | Battery
AQT90 FLEX | 更换光学单元903-364一个，并重新定标后修复 | 更换光学单元903-364一个，并重新定标后修复 | SWO-72807 | No | Replace the optical unit 903-364 one and re-calibrate the repair | Replace the optical unit 903-364 one and re-calibrate the repair | Defect | HW | Process Module | Optical Unit | 131
AQT90 FLEX |  | B | SWO-72984 | No |  |  | Defect | HW | Aspiration Module | Bend needle | 4794
AQT90 FLEX |  | B | SWO-72979 | No |  |  | Defect | HW | Process Module | Process Unit sensor | 5784
ABL815 FLEX |  | B | D00F28DFB20411FD | Yes |  |  | Re classified RMED | Logistical upSWO
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-73126 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-73171 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL80 FLEX CO-OX | Fehlersuche am Analysator, immer wieder andere Parameter bei LQC fehlerhaft (Kalium, dann pO2 etc.). trotz Tausch des LP der SC, auch immer wieder Probleme mit Undichtigkeiten des LP->an Schlauchabzweigung leichte Ablagerungen festgestellt, Schlauchwege komplett erneuert, Anschlussstutzen des Analysators (für SC und LP) gereinigt, SC gespült, Neustart->LP aus Kundenbestand erneuert, Kal. und LQC wieder ok, Testmessungen ok Mit dem Anwender Lagerungstemp. der Betriebsmittel besprochen System-und Funktionskontrolle i.O. | Fehlersuche am Analysator, immer wieder andere Parameter bei LQC fehlerhaft (Kalium, dann pO2 etc.). trotz Tausch des LP der SC, auch immer wieder Probleme mit Undichtigkeiten des LP->an Schlauchabzweigung leichte Ablagerungen festgestellt, Schlauchwege komplett erneuert, Anschlussstutzen des Analysators (für SC und LP) gereinigt, SC gespült, Neustart->LP aus Kundenbestand erneuert, Kal. und LQC wieder ok, Testmessungen ok Mit dem Anwender Lagerungstemp. der Betriebsmittel besprochen System-und Funktionskontrolle i.O. | SWO-73143 | No | Troubleshooting on the analyzer, again and again other parameters in LQC faulty (potassium, then pO2 etc.). despite the replacement of the LP of the SC, also again and again problems with leaks of the LP->at the hose branch light deposits detected, hose paths completely renewed, connection socket of the analyzer (for SC and LP) cleaned, SC flushed, restart>LP from Customer stock renewed, Kal. and LQC ok again, test measurements ok With the user storage temp. of the equipment discussed system and function control as the | Troubleshooting on the analyzer, again and again other parameters in LQC faulty (potassium, then pO2 etc.). despite the replacement of the LP of the SC, also again and again problems with leaks of the LP->at the hose branch light deposits detected, hose paths completely renewed, connection socket of the analyzer (for SC and LP) cleaned, SC flushed, restart>LP from Customer stock renewed, Kal. and LQC ok again, test measurements ok With the user storage temp. of the equipment discussed system and function control as the
ABL90 FLEX | Hemolysator und Fließwege im Serviceprogramm gereinigt, Kal und QC durchgeführt, OK | Funktion Im Service überprüft, Pumpenkalwert grenzwertig, Vakuum Value zu hoch. SC 90 aus Kundenbestand eingebaut, Fehler tritt weiterhin auf. LP90 SV03 aus Kundenbestand ersetzt. Funktion OK. Reklamation an RDEU gemailt. | SWO-73062 | Yes | Hemolysator and flow paths cleaned in service program, Kal and QC performed, OK | Function Checked in service, pump calorific value borderline, vacuum value too high. SC 90 built from customer inventory, error still occurs. LP90 SV03 replaced from customer base. Function OK. Complaint to RDEU. | Re classified RMED | Double entry
ABL90 FLEX | System gespült, Einlassdichtung erneuert Station hat keine Dichtung vor Ort Neustart | System gespült, Einlassdichtung erneuert Station hat keine Dichtung vor Ort Neustart | SWO-73147 | No | System flushed, inlet seal renewed station has no seal on site restart | System flushed, inlet seal renewed station has no seal on site restart | Cleaned | Fluidic pathway
ABL90 FLEX PLUS | Software MR7 opnieuw geinstalleerd, met een lege database, en nieuwe .ini file. Backup en instellingen (van een goedwerkende ABL) teruggezet. | Software MR7 opnieuw geinstalleerd, met een lege database, en nieuwe .ini file. Backup en instellingen (van een goedwerkende ABL) teruggezet. | SWO-73127 | No | Software MR7 re-installed, with a blank database, and new. ini file. Backup and settings (of a properly working ABL) restored. | Software MR7 re-installed, with a blank database, and new. ini file. Backup and settings (of a properly working ABL) restored. | Defect | SW | Frequent calibration
ABL90 FLEX PLUS | Das Einlassmodul wieder richtig eingesetzt und im Service Menü kalibriert. | Das Einlassmodul wieder richtig eingesetzt und im Service Menü kalibriert. | SWO-73131 | No | The inlet module is used correctly again and calibrated in the service menu. | The inlet module is used correctly again and calibrated in the service menu. | USER ERROR | Re-Install Auto inlet
ABL90 FLEX | 连续更换两个试剂包后修复 | 连续更换两个试剂包后修复 | SWO-73177 | No | Repair after replacing two reagent kits in a row | Repair after replacing two reagent kits in a row | Assisted Replacement | SP | Pump cal error
ABL90 FLEX | 更换新试剂包后修复 | 更换新试剂包后修复 | SWO-73183 | No | Repair after replacing a new kit | Repair after replacing a new kit | Assisted Replacement | SP | Pump cal error
AQT90 FLEX | 更换Hct模块903-329一个后修复(by denis) | 更换Hct模块903-329一个后修复(by denis) | SWO-73082 | No | Replace Hct module 903-329 one post-fix (by denis) | Replace Hct module 903-329 one post-fix (by denis) | Defect | HW | Aspiration Module | Needle Wash Unit | 1267
ABL800 FLEX | Comprobación neumática, Limpieza cámara de medida y sustitución de los tubos de bombas. Finalmente se observa que el sensor de liquidos inferior del módulo de pH no funciona. Sustitución del mismo. Calibración bombas y SL. Tests ST. Cal 2 ptos , QC y muestras OK. | Comprobación neumática, Limpieza cámara de medida y sustitución de los tubos de bombas. Finalmente se observa que el sensor de liquidos inferior del módulo de pH no funciona. Sustitución del mismo | D01043B22834D537 | No | Pneumatic check, Cleaning measuring chamber and replacement of pump tubes. Finally it is observed that the lower liquid sensor of the pH module does not work. Substitution of the same. Calibration pumps and SL. ST Tests. Cal 2 pts , QC and samples OK. | Pneumatic check, Cleaning measuring chamber and replacement of pump tubes. Finally it is observed that the lower liquid sensor of the pH module does not work. Replacing it | Defect | HW | pH/BG Module | LS sensor
ABL820 FLEX | 更换945-614修复仪器 (by denis) | B | SWO-72763 | No | Replace the 945-614 repair instrument (by denis) |  | Assisted Replacement | Electrode | Electrode | pH
ABL827 FLEX | Call back. Investigation performed on service work order B9QV43YV7. | Replaced hemolyzer unit for diagnostic reasons. | D01063452825E2F9 | No | Call back. Investigation performed on service work order B9QV43YV7. | Replaced hemolyzer unit for diagnostic reasons. | Defect | HW | Oxi module | Hemolyzer
ABL825 FLEX | Einstellungen der Firewall korrigiert, Korrektur S7745 Lot 613, Neustart des Gerätes, Korrekte QK aller Level abgewartet | Einstellungen der Firewall korrigiert, Korrektur S7745 Lot 613, Neustart des Gerätes, Korrekte QK aller Level abgewartet | SWO-73048 | No | Fixed firewall settings, correction S7745 lot 613, restart of the device, correct QK of all levels awaited | Fixed firewall settings, correction S7745 lot 613, restart of the device, correct QK of all levels awaited | Setup | SW | His/lis | Radiance
ABL80 FLEX BASIC | 更换新测试卡修复仪器 | 更换新测试卡修复仪器 | SWO-73435 | No | Replace the new test card repair instrument | Replace the new test card repair instrument
ABL80 FLEX BASIC | 更换新测试卡修复仪器 | 更换新测试卡修复仪器 | SWO-73369 | No | Replace the new test card repair instrument | Replace the new test card repair instrument
ABL80 FLEX BASIC | 更换新测试卡修复仪器 | 更换新测试卡修复仪器 | SWO-73364 | No | Replace the new test card repair instrument | Replace the new test card repair instrument
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-73221 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL80 FLEX BASIC | 更换新测试卡修复仪器，同事清洁了被测试卡漏液浸泡了的接口板，修复了仪器(denis 20151230 代理商工程师清洁该机器接口板后机器故障消失） | 更换新测试卡修复仪器，同事清洁了被测试卡漏液浸泡了的接口板，修复了仪器(denis 20151230 代理商工程师清洁该机器接口板后机器故障消失） | SWO-73437 | No | Replace the new test card repair instrument, colleagues clean the interface board soaked by the test card leak, repair the instrument (denis 20151230 dealer engineer cleans the machine interface board after the machine fault disappears) | Replace the new test card repair instrument, colleagues clean the interface board soaked by the test card leak, repair the instrument (denis 20151230 dealer engineer cleans the machine interface board after the machine fault disappears)
ABL80 FLEX BASIC | 更换新测试卡修复仪器 | 更换新测试卡修复仪器 | SWO-73410 | No | Replace the new test card repair instrument | Replace the new test card repair instrument
ABL80 FLEX BASIC | Kept SC in fridge and also reinstalled solution pack. Then reinstalled Sensor cassette. Now unit is working fine. | Kept SC in fridge and also reinstalled solution pack. Then reinstalled Sensor cassette. Now unit is working fine. | SWO-73535 | No | Kept SC in fridge and also reinstalled solution pack. Then reinstalled Sensor cassette. Now unit is working fine. | Kept SC in fridge and also reinstalled solution pack. Then reinstalled Sensor cassette. Now unit is working fine.
ABL90 FLEX | Funktion Im Service überprüft, Pumpenkalwert grenzwertig, Vakuum Value zu hoch. SC 90 aus Kundenbestand eingebaut, Fehler tritt weiterhin auf. LP90 SV03 aus Kundenbestand ersetzt. Funktion OK. Reklamation an RDEU gemailt. | Funktion Im Service überprüft, Pumpenkalwert grenzwertig, Vakuum Value zu hoch. SC 90 aus Kundenbestand eingebaut, Fehler tritt weiterhin auf. LP90 SV03 aus Kundenbestand ersetzt. Funktion OK. Reklamation an RDEU gemailt. | SWO-73353 | No | Function Checked in service, pump calorific value borderline, vacuum value too high. SC 90 built from customer inventory, error still occurs. LP90 SV03 replaced from customer base. Function OK. Complaint to RDEU. | Function Checked in service, pump calorific value borderline, vacuum value too high. SC 90 built from customer inventory, error still occurs. LP90 SV03 replaced from customer base. Function OK. Complaint to RDEU. | Assisted Replacement | SP | Vacuum error
ABL90 |  | B | D01172589E30CBE4 | No |  |  | Defect | HW | Oxi module | Pinch Valve tube
ABL90 FLEX | Abklärung mit TS. Änderung Probentyp Filter nicht möglich. Zusätzliches Feld (Bemerkung) mit Auswahlfilters war nicht gewünscht. Kalibration und QC's vorgenommen. | Abklärung mit TS. Änderung Probentyp Filter nicht möglich. Zusätzliches Feld (Bemerkung) mit Auswahlfilters war nicht gewünscht. Kalibration und QC's vorgenommen. | SWO-73423 | No | Clarification with TS. Modification sample type filter not possible. Additional field (note) with selection filter was not desired. calibration and QC's. | Clarification with TS. Modification sample type filter not possible. Additional field (note) with selection filter was not desired. calibration and QC's. | Defect | SW | Windows??
ABL90 | pO2 is high on low pO2 and low on high pO2 internal QC. Backflush has not resolved issue. Customer has changed multiple SC and SP since 12 23 on this issue. | Performed complete annual maintenance replacing above parts. pO2 issue still present. Replaced SC, (both were lot 0729) pO2 now checking OK. | D011948D9E551421 | No | pO2 is high on low pO2 and low on high pO2 internal QC. Backflush has not resolved issue. Customer has changed multiple SC and SP since 12 23 on this issue. | Performed complete annual maintenance replacing above parts. pO2 issue still present. Replaced SC, (both were lot 0729) pO2 now checking OK. | Assisted Replacement | SC | pO2 QC
ABL90 | PERFORMED FLIUID TRANSPORT ADJUSTMENTS.PRESSURE TEST SHOWS FAILING VALVE CHECK | REPLACED INLET GASKET HOLDER AND PROBE.PERFRMED FLUID TRANSORT ADJUSTMENT WITH NO FAILURES ON PRESSURE TEST.PERFORMED START UP.PERFORMED CALIBRATIONS. | D0118F329E0167AE | No | PERFORMED FLIUID TRANSPORT ADJUSTMENTS. PRESSURE TEST SHOWS FAILING VALVE CHECK | REPLACED INLET GASKET HOLDER AND PROBE. PERFRMED FLUID TRANSORT ADJUSTMENT WITH NO FAILURES ON PRESSURE TEST. PERFORMED START UP. PERFORMED CALIBRATIONS. | Defect | HW | Inlet Gasket
ABL90 FLEX PLUS |  | B | D0117F929E075290 | No |  |  | Defect | HW | Auto Inlet module | Inlet Gasket
ABL90 FLEX PLUS |  | B | D018F2CC622BEDE7 | No |  |  | Assisted Replacement | SC | SC | Unknown
ABL90 FLEX PLUS | Diverse tubings vervangen, en Rp-vervangen, maar geeft vooralsnog geen oplossing. | B | SWO-73526 | No | Various tubings replaced, and Rp-replaced, but does not give a solution yet. |  | Unknown | Revisit required | pCO2
ABL90 FLEX | 更换新试剂包后修复 | 更换新试剂包后修复 | SWO-73372 | No | Repair after replacing a new kit | Repair after replacing a new kit | Assisted Replacement | SP | Vacuum error
ABL90 | Comprobación circuito. Sc ya caducada (el problema dió dia y medio antes de la expiración del consumible). Hemolizador OK. SC aparece atascado. Limpieza del circuito, chequeo visual pero aun con problemas de aspirado. Cambio SP pero igual. Sustitución aguja de entrada y goma interconexión. Chequeos neumaticos OK. Arranque, cal, cal thb y QC. | Comprobación circuito. Sc ya caducada (el problema dió dia y medio antes de la expiración del consumible). Hemolizador OK. SC aparece atascado. Limpieza del circuito, chequeo visual pero aun con problemas de aspirado. Cambio SP pero igual. Sustitución aguja de entrada y goma interconexión. | D01195E19E5FD305 | No | Circuit check. Sc has already expired (the problem day and a half before the consumable expires). Hemolizer OK. SC appears stuck. Cleaning the circuit, visual check, but still with vacuuming problems. Change SP but still. Replacement input needle and rubber interconnection. Pneumcheck checks OK. Start, lime, thb lime and QC. | Circuit check. Sc has already expired (the problem day and a half before the consumable expires). Hemolizer OK. SC appears stuck. Cleaning the circuit, visual check, but still with vacuuming problems. Change SP but still. Replacement input needle and rubber interconnection. | Defect | HW | Inlet Gasket
AQT90 FLEX | 更换进样针903-171一个后修复 (by denis) | 更换进样针903-171一个后修复 (by denis) | SWO-73087 | No | Replace the sample needle 903-171 after repair (by denis) | Replace the sample needle 903-171 after repair (by denis) | Defect | HW | Aspiration Module | Bend needle | 4794
AQT90 FLEX | 更换光学单元，仪器恢复正常。 | 更换光学单元，仪器恢复正常。 | SWO-73217 | No | Replace the optical unit and the instrument will return to normal. | Replace the optical unit and the instrument will return to normal. | Defect | HW | Process Module | Optical Unit | 131
ABL825 FLEX | checked ip settings all was configured correctly noticed the connection was cycling on and off tried a new cable tried a power cycle still same results. plugged the analyzer into the wall jack and not the switch. analyzer connected to radiance. reset the switch again and reconnected | reset the switch re imput all the settings analyzer is connected re sent all data from analyzer to radiance that was missing | D01179F39E14C1D7 | No | checked ip settings all was configured correctly noticed the connection was cycling on and off tried a new cable tried a power cycle still same results. plugged the analyzer into the wall jack and not the switch. analyzer connected to radiance. reset the switch again and reconnected | reset the switch re imput all the settings analyzer is connected re sent all data from analyzer to radiance that was missing | Setup | SW | His/lis | Radiance
ABL825 FLEX | SWV 6.16 installiert. Einweisung für einen Person durchgeführt. | SWV 6.16 installiert. Einweisung für einen Person durchgeführt. | SWO-73327 | No | SWV 6.16 installed. instruction for one person. | SWV 6.16 installed. instruction for one person. | Defect | HW | DMS_PC | Touch Screen
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-73607 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL90 FLEX | Analysator geprüft und gereinigt. Mehrere Blutmessungen mit 85ul Kapillaren durchgeführt, kein Fehler aufgetreten. Anwender informiert, dass wenn die Fehler erneut auftreten, dann das LP90 erneuert werden soll. Kein LP90 auf Lager, St.5 Bestand geprüft, 1 Ersatz LP ist im Kühlschrank. | Kunde bittet um einen Rückruf und Service | SWO-73564 | No | Analyzer tested and cleaned. Several blood measurements performed with 85ul capillaries, no error occurred. Informs the user that if the errors occur again, then the LP90 should be renewed. No LP90 in stock, St.5 stock checked, 1 replacement LP is in the refrigerator. | Customer asks for a callback and service | Cleaned | Fluidic pathway
ABL90 FLEX | Erneut Fehler bei Patientenmessungen--> Trübung zu hoch, Oxi spectrum mismatch Oxicompletetest: Cavitation absorb fehlerhaft ( 0,5mabs) --> Hemolyzer erneuert Oxicompletetest durchg führt --> cavitation absorb bei 12,23mabs | Erneut Fehler bei Patientenmessungen--> Trübung zu hoch, Oxi spectrum mismatch Oxicompletetest: Cavitation absorb fehlerhaft ( 0,5mabs) --> Hemolyzer erneuert Oxicompletetest durchg führt --> cavitation absorb bei 12,23mabs | SWO-73491 | No | Erneut Fehler bei Patientenmessungen--> Trübung zu hoch, Oxi spectrum mismatch Oxicompletetest: Cavitation absorb fehlerhaft ( 0,5mabs) --> Hemolyzer erneuert Oxicompletetest durchg führt --> cavitation absorb bei 12,23mabs | Erneut Fehler bei Patientenmessungen--> Trübung zu hoch, Oxi spectrum mismatch Oxicompletetest: Cavitation absorb fehlerhaft ( 0,5mabs) --> Hemolyzer erneuert Oxicompletetest durchg führt --> cavitation absorb bei 12,23mabs
ABL90 FLEX | CreditClaimForms ausgedruckt | CreditClaimForms ausgedruckt | SWO-73418 | No | CreditClaimForms printed | CreditClaimForms printed | No fault found
ABL90 FLEX PLUS |  | B | SWO-73534 | No |  |  | Unknown | Removed by RSUB
ABL90 | Verificação dos registros, limpeza do circuito, subsituição da sonda, actualização do sw para a versão 3.1MR7, estabilização e passagem dos QCs. | B | D012BFC114005D80 | No | Checking the records, cleaning the circuit, subsitution of the probe, updating the SW to version 3.1 MR7, stabilization and passage of the QCs. |  | Cleaned | Fluidic pathway
AQT90 FLEX | 更换903-198一个后修复(denis 20151231 该故障经与陈传伟确认，代理商到现场后发现该机器不能正常启动，经检查发现该机器AU板上保险丝烧毁，更换保险丝后机器依然无法正常启动，然后代理商工程师更换AU板，重新启动计算机，机器恢复正常） | 更换903-198一个后修复 | SWO-73585 | No | Replacement 903-198 a post-fix (denis 20151231 the fault was confirmed with Chen Chuanwei, the dealer to the scene found that the machine can not start normally, after inspection found that the machine AU plate fuse burned, after replacing the fuse machine still can not start normally, and then the agent engineer to replace the AU board, Restart the computer and the machine is back to normal) | Replace 903-198 one post-fix | Defect | HW | Analyzer Control Module | Main Board, PCB | 1999
AQT90 FLEX | 更换903-171一个后修复，903-329未更换。(by denis) | 更换903-171一个后修复，903-329未更换。(by denis) | SWO-73190 | No | Replace 903-171 one post repair, 903-329 not replaced. (by denis) | Replace 903-171 one post repair, 903-329 not replaced. (by denis) | Defect | HW | Aspiration Module | Bend needle | 4794
AQT90 FLEX | 更换新进样针修复仪器（by denis） | 更换新进样针修复仪器（by denis） | SWO-73575 | No | Replace the new sample needle repair instrument (by denis) | Replace the new sample needle repair instrument (by denis) | Defect | HW | Aspiration Module | Bend needle
ABL825 FLEX |  | B | D012BBDB14359CD3 | No |  |  | Defect | HW | AutoCheck Module | PCB
ABL835 FLEX |  | B | D012BA2A1444DD37 | Yes |  |  | Re classified RMED | SWU
ABL825 FLEX |  | B | D012BAE0140DDD3B | Yes |  |  | Re classified RMED | SWU
ABL820 FLEX | Viewed logs for any possible errors - none. | Replaced the HDD, power supply and CPU unit. | D012E2731443DF89 | No | Viewed logs for any possible errors - none. | Replaced the HDD, power supply and CPU unit. | Defect | HW | DMS_PC | HDD
ABL825 FLEX | found that the breaker arm was not retracting all the way. | replaced breaker motor on analyzer to correct the issue. | D012D7B11434984E | No | found that the breaker arm was not retracting all the way. | replaced breaker motor on analyzer to correct the issue. | Defect | HW | AutoCheck Module | Breaker arm w/o function
ABL825 FLEX | Messkammer CO2 gereinigt, pCO2-Elektrode (alt: ST05) ersetzt, Mehrere Kal abgewartet, QK ok | Messkammer CO2 gereinigt, pCO2-Elektrode (alt: ST05) ersetzt, Mehrere Kal abgewartet, QK ok | SWO-73508 | No | Measuring chamber CO2 cleaned, pCO2 electrode replaced (old: ST05), Several kal waited, QK ok | Measuring chamber CO2 cleaned, pCO2 electrode replaced (old: ST05), Several kal waited, QK ok | Assisted Replacement | Electrode | Electrode | pCO2
ABL825 FLEX | Reinigung und Kalibration Service Thb Kal durchgeführt. Lac. Membraniert. Patientenlayout angepasst. | Reinigung und Kalibration Service Thb Kal durchgeführt. Lac. Membraniert. Patientenlayout angepasst. | SWO-73522 | No | Cleaning and calibration service Thb Kal performed. Lac. Membraneed. patient layout. | Cleaning and calibration service Thb Kal performed. Lac. Membraneed. patient layout. | Adjustment | HW | Oxi module | tHb cal
ABL80 FLEX CO-OX | Überprüfung Kal und QC Ergebnisse => letzte tHb am 15.10. Service Test durchgeführt => Oximetrie ohne Fehler Kalibrierung Flüssigkeitssensor Durchführung tHb-Kal => Fcuv = 0,845 QC bleibt fehlerhaft mit Fehler 290, 581 und 688 Gerät demontiert und Verschlauchung am Hemolyzer überprüft => alles i.O. Hemolyzer gespült und LS gereinigt => alles i.O. Gerät neugestartet und nochmaliger Service Test => weiterhin keine Fehler Nochmalige tHb-Kal => Fcuv jetzt bei 0,982 Uhrzeit gestellt | Überprüfung Kal und QC Ergebnisse => letzte tHb am 15.10. Service Test durchgeführt => Oximetrie ohne Fehler Kalibrierung Flüssigkeitssensor Durchführung tHb-Kal => Fcuv = 0,845 QC bleibt fehlerhaft mit Fehler 290, 581 und 688 Gerät demontiert und Verschlauchung am Hemolyzer überprüft => alles i.O. Hemolyzer gespült und LS gereinigt => alles i.O. Gerät neugestartet und nochmaliger Service Test => weiterhin keine Fehler Nochmalige tHb-Kal => Fcuv jetzt bei 0,982 Uhrzeit gestellt | SWO-73599 | No | Check Kal and QC results => last tHb performed on 15.10. Service test => Oximetry without error Calibration liquid sensor carrying out tHb-Kal => Fcuv = 0.845 QC remains faulty with error 290, 581 and 688 device dismantled and hose on hemolyzer checked => all i.O. Hemolyzer rinsed and LS cleaned => all i.O. device restarted and re-service test => still no errors again tHb-Kal => Fcuv now set at 0.982 time | Check Kal and QC results => last tHb performed on 15.10. Service test => Oximetry without error Calibration liquid sensor carrying out tHb-Kal => Fcuv = 0.845 QC remains faulty with error 290, 581 and 688 device dismantled and hose on hemolyzer checked => all i.O. Hemolyzer rinsed and LS cleaned => all i.O. device restarted and re-service test => still no errors again tHb-Kal => Fcuv now set at 0.982 time
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-73699 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL90 FLEX | 更换新试剂包后修复 | 更换新试剂包后修复 | SWO-73591 | No | Repair after replacing a new kit | Repair after replacing a new kit | Assisted Replacement | SP | Pump cal error
AQT90 FLEX | 更换902-924修复仪器(by denis) | 更换902-924修复仪器(by denis) | SWO-73411 | No | Replace the 902-924 repair instrument (by denis) | Replace the 902-924 repair instrument (by denis) | Defect | HW | Analyzer Control Module | Barcode reader, display | BR2
AQT90 FLEX | 更换903-113一个后修复(denis 20151231 该机器表现症状为进样针弯曲，经过测试发现该机器的加样模块不受控，报错，在维修菜单中无法调整加样模块到正确的位置并进行旋转，工程师更换903-113后该机器恢复正常工作） | 更换903-113一个后修复(denis 20151231 该机器表现症状为进样针弯曲，经过测试发现该机器的加样模块不受控，报错，在维修菜单中无法调整加样模块到正确的位置并进行旋转，工程师更换903-113后该机器恢复正常工作） | SWO-73385 | No | Replace 903-113 a post-repair (denis 20151231 The machine shows symptoms of the sample needle bending, after testing found that the machine's dosing module is not controlled, error, in the service menu can not adjust the sample module to the correct position and rotate, The machine returns to normal operation after the engineer replaces the 903-113) | Replace 903-113 a post-repair (denis 20151231 The machine shows symptoms of the sample needle bending, after testing found that the machine's dosing module is not controlled, error, in the service menu can not adjust the sample module to the correct position and rotate, The machine returns to normal operation after the engineer replaces the 903-113) | Defect | HW | Aspiration Module | Pipette PCB | 4794
AQT90 FLEX | 更换CF卡，重新定标正常。by Denis 该机器不能正常启动，更换 cf 卡后重新安装软件并定标后机器正常工作 | 更换CF卡，重新定标正常。by Denis 该机器不能正常启动，更换 cf 卡后重新安装软件并定标后机器正常工作 | SWO-73282 | No | Replace the CF card and re-mark it normally. By Denis The machine does not start properly, the machine works properly after replacing the cf card and reinstalling the software and marking | Replace the CF card and re-mark it normally. By Denis The machine does not start properly, the machine works properly after replacing the cf card and reinstalling the software and marking | Not Cat | Not Cat | Not Cat | Not Cat
ABL825 FLEX | Einlassnadel und Dichtung erneuert, jedoch ohne Erfolg. Einlass im Serviceprogramme kalibriert, Fehler blieb bestehen. Ersatzmodul aus der Geschäftsstelle besorgt und eingebaut. Gerät kalibriert und überprüft. Der Service war nach Aussage von Dr. Sebastian Rojas dringend erforderlich. Ich habe auf die kostenpflichtigen Reparaturkosten in doppelter Höhe aufmerksam gemacht und sie sind von Dr. Rojas akzeptiert worden. Defektes Modul 902-783R2188N001 Austausch: R1478N004 | B | SWO-73651 | No | Inlet needle and seal renewed, but to no avail. Inlet calibrated in the service program, errors remained. Replacement module from the office is provided and installed. calibrated and checked. According to Dr. Sebastian Rojas, the service was urgently needed. I have drawn attention to the double-cost repair costs and they have been accepted by Dr. Rojas. Defective module 902-783R2188N001 Replacement: R1478N004 |  | Defect | HW | Inlet module | Friction detector unit
ABL90 FLEX | 更换测试卡正常。 | 更换测试卡正常。 | SWO-74867 | No | The replacement test card is OK. | The replacement test card is OK. | Assisted Replacement | SC | SC | Reference
ABL825 FLEX | 更换902-605一个正常 | 更换902-605一个正常 | SWO-73365 | No | Replace 902-605 one normal | Replace 902-605 one normal | Defect | HW | Gas Module | Pressure Transducer | Analyzer ctrl PCB?
ABL90 FLEX | Bemeneti tűtartó csere, Felszívó tű csere, Kalibrálás, QC mérés | B | SWO-73666 | No | Inlet needle holder replacement, suction needle replacement, calibration, QC measurement |  | User error | Inlet Gasket broken | Inlet Gasket broken | Inlet Gasket broken
ABL800 BASIC | なし。 | なし。 | SWO-73655 | No | Without. | Without. | Unknown | Revisit required | Revisit required | Revisit required
ABL90 FLEX | 更换903-165一个正常。(by andrew) | 更换903-165一个正常。(by andrew) | SWO-73675 | No | Replace the 903-165 one normal. (by andrew) | Replace the 903-165 one normal. (by andrew) | Defect | Hw | Basic Unit | Wet Sec PCB
AQT90 FLEX | flushed the tubing's, replaced HCT cell. | flushed the tubing's, replaced HCT cell. | SWO-75345 | No | flushed the tubing's, replaced HCT cell. | flushed the tubing's, replaced HCT cell. | Defect | HW | Aspiration Module | Needle Wash Unit
ABL80 FLEX BASIC | 更换新试剂包后修复 | 更换新试剂包后修复 | SWO-74963 | No | Repair after replacing a new kit | Repair after replacing a new kit
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-74618 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-73804 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL80 FLEX BASIC | 换了2张测试卡，，第一张pO2 QC错误，第二张Lac定标错误 | 换了2张测试卡，，第一张pO2 QC错误，第二张Lac定标错误 | SWO-75098 | No | Changed 2 test cards, the first pO2 QC error, the second Lac calibration error | Changed 2 test cards, the first pO2 QC error, the second Lac calibration error
ABL80 CO-OX | Troubleshot with customer over phone. Customer tried different solution pack and sensor cassette. Problem persists. | Performed overhaul on analyzer. Replaced faulty liquid sensor. Performed necessary calibrations to bring analyzer back to working status. | D0193545624E6174 | No | Troubleshot with customer over phone. Customer tried different solution pack and sensor cassette. Problem persists. | Performed overhaul on analyzer. Replaced faulty liquid sensor. Performed necessary calibrations to bring analyzer back to working status.
ABL80 FLEX BASIC | 更换测试卡 | 更换测试卡 | SWO-74961 | No | Replace the test card | Replace the test card
ABL80 FLEX BASIC | 更换测试卡 | 更换测试卡 | SWO-74893 | No | Replace the test card | Replace the test card
ABL80 FLEX CO-OX | Fehlersuche, LP wurde bereits erneuert, Schlauchwege ok, mit Zeroing Kassette im Service Werte überprüft->alle im Bereich!, SC manuell gespült, Flüssigkeitssensorkal. ok, Neustart ok, SC erneuert->leichte Blutablagerung in SC danach Kal. und IQC ok System-und Funktionskontrolle i.O. | Fehlersuche, LP wurde bereits erneuert, Schlauchwege ok, mit Zeroing Kassette im Service Werte überprüft->alle im Bereich!, SC manuell gespült, Flüssigkeitssensorkal. ok, Neustart ok, SC erneuert->leichte Blutablagerung in SC danach Kal. und IQC ok System-und Funktionskontrolle i.O. | SWO-73827 | No | Troubleshooting, LP has already been renewed, hoseways ok, checked with zeroing cassette in service values->all in the range!, SC manually rinsed, liquid sensor scale. ok, restart ok, SC renews->light blood deposit in SC after Kal. and IQC ok system and function control in the | Troubleshooting, LP has already been renewed, hoseways ok, checked with zeroing cassette in service values->all in the range!, SC manually rinsed, liquid sensor scale. ok, restart ok, SC renews->light blood deposit in SC after Kal. and IQC ok system and function control in the
ABL90 FLEX PLUS |  | B | D01AA7A8D8585894 | Yes |  |  | Re classified RMED | installation/training
ABL90 FLEX | 更换新试剂包后修复 | 更换新试剂包后修复 | SWO-74612 | No | Repair after replacing a new kit | Repair after replacing a new kit | Assisted Replacement | SP | SP | Pump cal error
ABL835 FLEX | Inlet coule | Remplac probe holder | D0190DF1625F7437 | No | Inlet sinks | Remplac probe holder | Defect | HW | Inlet module | Probe holder
ABL800 BASIC | Inlet test pass. needle is fine & positioning proper. clean the inlet gasket. machine under observation. | B | SWO-75534 | No | Inlet test pass. needle is fine & positioning proper. clean the inlet gasket. machine under observation. |  | Unknown | Unknown | Unknown | Unknown
ABL800 BASIC | Rinse solution need replacement. Replace the rinse solution. After replace rinse solution unit is working properly. | Rinse solution need replacement. Replace the rinse solution. After replace rinse solution unit is working properly. | SWO-75525 | No | Rinse solution need replacement. Replace the rinse solution. After replace rinse solution unit is working properly. | Rinse solution need replacement. Replace the rinse solution. After replace rinse solution unit is working properly. | Assisted Replacement | Solution bottle | Solution bottle | Rinse
ABL817 FLEX | R-Vorgabe Konfig. geladen, Konfig. neu eingestellt | R-Vorgabe Konfig. geladen, Konfig. neu eingestellt | SWO-73796 | No | R-default Config. loaded, Config. newly discontinued | R-default Config. loaded, Config. newly discontinued | Setup | HIS/LIS | His/lis | His/lis
ABL825 FLEX | Confirmed issue when onsite. | PROCEED TO REPLACE LAMP UNIT AS LAST ATTEMPT TO REPLACE NEON LAMP DID NOT RESOLVE ERROR COMPLETELY. | D019159E6208DCAE | No | Confirmed issue when onsite. | PROCEED TO REPLACE LAMP UNIT AS LAST ATTEMPT TO REPLACE NEON LAMP DID NOT RESOLVE ERROR COMPLETELY. | Defect | HW | Oxi module | Lamp unit
ABL825 FLEX | found that they were getting 0328 | checked and cleaned all the wells. I checked all tubing and inlet probe. preformed inlet calibration and pump calibrations and LS adjust to correct the issue. | D0192C7D622DCFAB | No | found that they were getting 0328 | checked and cleaned all the wells. I checked all tubing and inlet probe. preformed inlet calibration and pump calibrations and LS adjust to correct the issue. | Cleaned | Fluidic pathway | pH/BG Module | Measuring chamber
ABL837 FLEX | THE INSTRUMENT WAS LEAKING | THE WASTE TUBING, THE WASTE CONNECTORS AND THE VALVE WERE REPLACED. ALSO THE WASTE BOTTLE CONNECTOR WAS ALSO REPLACED. 6 PATIENT SAMPLES WERE RUN WITH NO ISUES TO REPORT | D019A488625BBF98 | No | THE INSTRUMENT WAS LEAKING | THE WASTE TUBING, THE WASTE CONNECTORS AND THE VALVE WERE REPLACED. ALSO THE WASTE BOTTLE CONNECTOR WAS ALSO REPLACED. 6 PATIENT SAMPLES WERE RUN WITH NO ISUES TO REPORT | Defect | HW | Fluidic manifold | Pump rotor | Waste
ABL825 FLEX | Erneut ein Neues Oxy Modul eingebaut, Service THb Kal. | Erneut ein Neues Oxy Modul eingebaut, Service THb Kal. | SWO-73880 | No | Once again a new Oxy module installed, Service THb Kal. | Once again a new Oxy module installed, Service THb Kal. | Defect | HW | Oxi module | =AW9797
AQT90 FLEX | Fehlersuche am Analysator, Analysator geöffnet, auf Serviceprogrammebene verbogene Pipette erneuert, Safe Postions ok, Neustart ok, Systemüberprüfung ok, TnT LQC ok System-und Funktionskontrolle i.O. | Fehlersuche am Analysator, Analysator geöffnet, auf Serviceprogrammebene verbogene Pipette erneuert, Safe Postions ok, Neustart ok, Systemüberprüfung ok, TnT LQC ok System-und Funktionskontrolle i.O. | SWO-74699 | No | Troubleshooting on the analyzer, analyzer open, pipette bent at service program level renewed, Safe Postions ok, restart ok, system check ok, TnT LQC ok system and function control i.O. | Troubleshooting on the analyzer, analyzer open, pipette bent at service program level renewed, Safe Postions ok, restart ok, system check ok, TnT LQC ok system and function control i.O. | Defect | HW | Aspiration Module | Bend needle
AQT90 FLEX | Checkouts uitgevoerd. | Checkouts uitgevoerd. | SWO-73861 | No | Checkouts uitgevoerd. | Checkouts uitgevoerd. | Unknown | HW | Storage Module | Storage wheel | 5689
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-78285 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL80 FLEX BASIC | 更换废液泵902-974一个后修复(by andrew) | 更换废液泵902-974一个后修复(by andrew) | SWO-73727 | No | Replace the waste pump 902-974 with a post-repair (by-andrew) | Replace the waste pump 902-974 with a post-repair (by-andrew)
ABL80 FLEX BASIC | 更换电源902-970一个后修复(by andrew) | 更换电源902-970一个后修复(by andrew) | SWO-73723 | No | Replace the power supply 902-970 with one post-fix (by andrew) | Replace the power supply 902-970 with one post-fix (by andrew)
ABL90 FLEX | Gerät im Service geprüft. Sämtliche Serviceprogramme aufgerufen.- Meldung Vakuumfehler. Fliesswege mehrmals gespült. Betriebsmittel neu installiert. | Gerät im Service geprüft. Sämtliche Serviceprogramme aufgerufen.- Meldung Vakuumfehler. Fliesswege mehrmals gespült. Betriebsmittel neu installiert. | SWO-75081 | No | device in service. All service programs called.- Message Vacuum error. Flow paths rinsed several times. Reinstalled equipment. | device in service. All service programs called.- Message Vacuum error. Flow paths rinsed several times. Reinstalled equipment. | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
ABL90 FLEX | neues LP mit LOT UB-03 eingesetzt, Kal + QC i.O. | neues LP mit LOT UB-03 eingesetzt; Kal + QC i.O. | SWO-74702 | No | new LP with LOT UB-03; Kal + QC i.O. | new LP with LOT UB-03; Kal + QC i.O. | Assisted Replacement | SP | SP | Logistical upSWO | pCO2
ABL90 FLEX | Kunde hat keine SC vorrätig, da ihre aufgrund Fehler 1346 nicht funktionierte, SC installiert, Kal, Interne QC | Kunde hat keine SC vorrätig, da ihre aufgrund Fehler 1346 nicht funktionierte, SC installiert, Kal, Interne QC | SWO-74772 | No | Customer does not have a SC in stock as their due to error 1346 did not work, SC installed, Kal, Internal QC | Customer does not have a SC in stock as their due to error 1346 did not work, SC installed, Kal, Internal QC | Assisted Replacement | SC | SC | Reference
ABL90 FLEX PLUS |  | B | D01CB4FBC4435902 | No |  |  | Defect | HW | Printer | Printer
ABL90 FLEX | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-74898 | No | Repair after replacing a new test card | Repair after replacing a new test card | Assisted Replacement | SC | SC | pCO2
AQT90 FLEX | 重新定标(by andrew) | 重新定标(by andrew) | SWO-74695 | No | Re-calibration (by andrew) | Re-calibration (by andrew) | Assisted Replacement | Consumables | Test Kit | NT-proBNP | 124
ABL800 BASIC | Moved the both old analyzers to ABG Lab, Cleaned the analyzers one analyzer was totally it was bad condition. ICU analyzer was OK, need to replace the all the membranes and calibrate the analyzer. | Found and replaced the Oxi module PCB from standby analyzer performed the complete oxi module test found it was all the results are OK, performed the 2p calibration and performed the tHb calibration found instur4ment was OK. | SWO-75543 | Yes | Moved the both old analyzers to ABG Lab, Cleaned the analyzers one analyzer was totally it was bad condition. ICU analyzer was OK, need to replace the all the membranes and calibrate the analyzer. | Found and replaced the Oxi module PCB from standby analyzer performed the complete oxi module test found it was all the results are OK, performed the 2p calibration and performed the tHb calibration found instur4ment was OK. | Re classified RMED | Installation
ABL800Basic |  | B | D01A2D71D85471E8 | No |  |  | Assisted Replacement | Electrode | Electrode | Reference
ABL827 FLEX | pH-Elektrode gereinigt, Dichtung der pH-Elektrode ersetzt, Kalibrationen, Kal2-Lösung ersetzt, pH-Elektrode (alt:SB04) ersetzt, Kalibrationen abgewartet, QK ok | pH-Elektrode gereinigt, Dichtung der pH-Elektrode ersetzt, Kalibrationen, Kal2-Lösung ersetzt, pH-Elektrode (alt:SB04) ersetzt, Kalibrationen abgewartet, QK ok | SWO-74755 | No | pH electrode cleaned, seal of pH electrode replaced, calibrations, Kal2 solution replaced, pH electrode (old:SB04) replaced, calibrations awaited, QK ok | pH electrode cleaned, seal of pH electrode replaced, calibrations, Kal2 solution replaced, pH electrode (old:SB04) replaced, calibrations awaited, QK ok | Assisted Replacement | Electrode | Electrode | pH
ABL800 BASIC | Ventilgehäuse der Messmodule gereinigt. Luftventil im pH BG Modul erneuert. Lösungspumpenschlauch, Pumpenschlauch (beider Messmodule), Einlassdichtung, Inletschlauch und Tube rinse heater erneuert. Luftweg über V7 gespült. Kalibrationen und QC's vorgenommen. | Ventilgehäuse der Messmodule gereinigt. Luftventil im pH BG Modul erneuert. Lösungspumpenschlauch, Pumpenschlauch (beider Messmodule), Einlassdichtung, Inletschlauch und Tube rinse heater erneuert. Luftweg über V7 gespült. Kalibrationen und QC's vorgenommen. | SWO-74722 | No | valve housing of the measuring modules. Air valve in pH BG module renewed. Solution pump hose, pump hose (both measuring modules), inlet seal, inlet hose and tube rinse heater renewed. Airway via V7. calibrations and QC's. | valve housing of the measuring modules. Air valve in pH BG module renewed. Solution pump hose, pump hose (both measuring modules), inlet seal, inlet hose and tube rinse heater renewed. Airway via V7. calibrations and QC's. | Defect | HW | pH/BG Module | Valve, air
ABL825 FLEX | Überpüfung Gasventile im Service - Membranierung pCO2 und Referenz etc. - Erneuerung Dichtringe Dummy (Cl) und pH Elektrode (blau) - Funktionsüberprüfung etc. ohne Beanstandung | Überpüfung Gasventile im Service - Membranierung pCO2 und Referenz etc. - Erneuerung Dichtringe Dummy (Cl) und pH Elektrode (blau) - Funktionsüberprüfung etc. ohne Beanstandung | SWO-74767 | No | Overtesting gas valves in service - membraneing pCO2 and reference etc. - Renewal sealing rings dummy (Cl) and pH electrode (blue) - function check etc. without complaint | Overtesting gas valves in service - membraneing pCO2 and reference etc. - Renewal sealing rings dummy (Cl) and pH electrode (blue) - function check etc. without complaint | Assisted Replacement | Electrode | Membrane | pCO2
ABL817 FLEX | Einlassnadel ern., Gerät kalibriert, QC durchgef. | Einlassnadel ern., Gerät kalibriert, QC durchgef. | SWO-74774 | No | Inlet needle ern., device calibrated, QC through | Inlet needle ern., device calibrated, QC through | Defect | HW | Inlet module | Probe bend | Probe bend
ABL825 FLEX | CPU- Einheit ausgetauscht. Glu Membran nicht vorhanden. Glu Driftet stark. | Display ausgetauscht. | SWO-74908 | No | CPU unit. Glu membrane not present. Glu Driftes strongly. | Display replaced. | Defect | HW | DMS_PC | DMS_PC
AQT90 FLEX |  | B | D01A914FD8191E82 | No |  |  | Defect | HW | Aspiration Module | Bend needle | Problem : Bent needle. 2 tubes were fallen inside the AQT. RC : We don’t really know what happened. CM : Replacement of needle. Replacement of the inlet wheel tube holders.
ABL80 FLEX BASIC | 更换触摸屏552-039一个后修复(by andrew) | 更换触摸屏552-039一个后修复(by andrew) | SWO-73908 | No | Replace touch screen 552-039 a post repair (by andrew) | Replace touch screen 552-039 a post repair (by andrew)
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-75101 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-74907 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL80 FLEX CO-OX | Kalibrierung abgewartet, QK bis auf pH&CO2 ok, keine Ersatzkassette vorrätig (ist bestellt), Konfiguration lt. Kundenwunsch korrigiert, Sensorkassette wird durch Anwender ersetzt | Kalibrierung abgewartet, QK bis auf pH&CO2 ok, keine Ersatzkassette vorrätig (ist bestellt), Konfiguration lt. Kundenwunsch korrigiert, Sensorkassette wird durch Anwender ersetzt | SWO-75020 | No | Calibration waited, QK up to pH&CO2 ok, no replacement cartridge in stock (is ordered), configuration corrected according to customer request, sensor cassette is replaced by user | Calibration waited, QK up to pH&CO2 ok, no replacement cartridge in stock (is ordered), configuration corrected according to customer request, sensor cassette is replaced by user
ABL80 FLEX BASIC | 更换测试卡 | 更换测试卡 | SWO-74885 | No | Replace the test card | Replace the test card
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-75063 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL80 FLEX BASIC | Removed old solution pack and installed new solution pack, Then reinstalled solution pack. Still problem is there.Need to replace SC. | Removed old solution pack and installed new solution pack, Then reinstalled solution pack. Still problem is there.Need to replace SC. | SWO-75498 | No | Removed old solution pack and installed new solution pack, Then reinstalled solution pack. Still problem is there. Need to replace SC. | Removed old solution pack and installed new solution pack, Then reinstalled solution pack. Still problem is there. Need to replace SC.
ABL90 FLEX | Analysator überprüft, Software MR7 eingespielt, Kal, Internes QC | Analysator überprüft, Software MR7 eingespielt, Kal, Internes QC | SWO-75049 | No | Analyzer checked, software MR7 inplayed, Kal, Internal QC | Analyzer checked, software MR7 inplayed, Kal, Internal QC | Assisted Replacement | SP | SP | Unknown
ABL90 | Inhomogeneous Rinse and QC solution. Ran diagnostics found no errors. Valve operating OK. | Replaced Sensor Cassette. | D01BC8F64E43F83F | No | Inhomogeneous Rinse and QC solution. Ran diagnostics found no errors. Valve operating OK. | Replaced Sensor Cassette. | Assisted Replacement | SC | SC | Wetability | pO2
ABL90 FLEX | Gerät überprüft, Lösungspack (SV) erneuert. UB 01 eingebaut, Kal und QC o.k. | Gerät überprüft, Lösungspack (SV) erneuert. UB 01 eingebaut, Kal und QC o.k. | SWO-75073 | No | device checked, solution pack (SV) renewed. UB 01 built-in, Kal and QC ok | device checked, solution pack (SV) renewed. UB 01 built-in, Kal and QC ok | Assisted Replacement | SP | SP | pCO2
ABL90 | Upon arrival found other parameters out on cals.After Sensor PCB replacement unit wanted replacement of SC and SP. | Replaced SensorPCB, and SC and SP. | D02801ABEA54BA0B | No | Upon arrival found other parameters out on cals. After Sensor PCB replacement unit wanted replacement of SC and SP. | Replaced SensorPCB, and SC and SP. | Defect | HW | SC PCB | Lac
ABL90 | Sensor cassette was failing for same error on adjacent analyzer. I had customer swap the sensor cassette from adjacent analyzer to see if problem followed the sensor cassette or stayed with the analyzer. The problem followed the sensor cassette with lot # 729-53. | Replaced sensor cassette from customers stock with lot # 428-218 and printed credit form for customer to submit. pO2 for Level 2 Lot 144 is now @ 102 with a range of 93.0-109. Saved data logs from both analyzers for submission to RMED if needed. | D01B84CD4E4B436B | No | Sensor cassette was failing for same error on adjacent analyzer. I had customer swap the sensor cassette from adjacent analyzer to see if problem followed the sensor cassette or stayed with the analyzer. The problem followed the sensor cassette with lot # 729-53. | Replaced sensor cassette from customers stock with lot # 428-218 and printed credit form for customer to submit. pO2 for Level 2 Lot 144 is now @ 102 with a range of 93.0-109. Saved data logs from both analyzers for submission to RMED if needed. | Assisted Replacement | SC | SC | Wetability | pO2
AQT90 FLEX | 更换光学单元后修复(by andrew) | 更换光学单元后修复(by andrew) | SWO-74667 | No | Repair after replacement of optical unit (by andrew) | Repair after replacement of optical unit (by andrew) | Defect | HW | Process Module | Optical Unit | 129
ABL800 BASIC | Replaced the all the membranes Ref, pCO2, pO2, Na, K, Ca, Cl, Glu, Lac, calibrated analyzer, found analyzer was showing error with Na, replaced the Na electrode and calibrated found instrument was working fine. | Found and replaced the Oxi module PCB from standby analyzer performed the complete oxi module test found it was all the results are OK, performed the 2p calibration and performed the tHb calibration found instur4ment was OK. | SWO-75544 | Yes | Replaced the all the membranes Ref, pCO2, pO2, Na, K, Ca, Cl, Glu, Lac, calibrated analyzer, found analyzer was showing error with Na, replaced the Na electrode and calibrated found instrument was working fine. | Found and replaced the Oxi module PCB from standby analyzer performed the complete oxi module test found it was all the results are OK, performed the 2p calibration and performed the tHb calibration found instur4ment was OK. | Re classified RMED | Installation
ABL827 FLEX | Gerät runtergefahren und ausgeschaltet, alle Steckverbindungen am OxiModul kontrolliert, Neustart, korrekte Kal abgewartet, QK ok | Gerät runtergefahren und ausgeschaltet, alle Steckverbindungen am OxiModul kontrolliert, Neustart, korrekte Kal abgewartet, QK ok | SWO-74979 | No | Device shut down and off, all connectors controlled on OxiModule, restart, correct scale waited, QK ok | Device shut down and off, all connectors controlled on OxiModule, restart, correct scale waited, QK ok | Adjustment | HW | Oxi module | Oxi Complete
ABL825 FLEX | Gerät überprüft, Fluidicchannel ausgebaut, Ventile 6,7,8 gesäubert und Ventilmembranen ersetzt | Gerät überprüft, Fluidicchannel ausgebaut, Ventile 6,7,8 gesäubert und Ventilmembranen ersetzt | SWO-75072 | No | Device inspected, Fluidicchannel removed, valves 6,7,8 cleaned and valve membranes replaced | Device inspected, Fluidicchannel removed, valves 6,7,8 cleaned and valve membranes replaced | Defect | Hw | Fluidic manifold | Valve, membrane
ABL800 BASIC | 更换903-126一个后修复(by andrew) | 更换903-126一个后修复(by andrew) | SWO-73732 | No | Replace 903-126 a post-fix (by-andrew) | Replace 903-126 a post-fix (by-andrew) | Defect | HW | AutoCheck Module | PCB
ABL835 FLEX |  | B | D01BAB234E1F2465 | No |  |  | Assisted Replacement | Electrode | Electrode | pCO2
ABL800 BASIC | tHB calibration done values within the range & QC done | tHB calibration done values within the range & QC done | SWO-75492 | No | tHB calibration done values within the range & QC done | tHB calibration done values within the range & QC done | Adjustment | HW | Oxi module | tHb cal
ABL825 FLEX | found that they are getting no leading air segment on qc samples. | found what looks like gel material from venus tube. in the system. cleaned out the system to correct the issue. and replaced all the tubings from customer stock. | D01B85AF4E448492 | No | found that they are getting no leading air segment on qc samples. | found what looks like gel material from venus tube. in the system. cleaned out the system to correct the issue. and replaced all the tubings from customer stock. | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
AQT90 FLEX | Main Board erneuert, Software aufgespielt, Gerät neu gestartet, Start-und Inkubationszeiten abgewartet, Gerät mit Testkassetten bestückt, Systemreinigung und QC durchgeführt. | Main Board erneuert, Software aufgespielt, Gerät neu gestartet, Start-und Inkubationszeiten abgewartet, Gerät mit Testkassetten bestückt, Systemreinigung und QC durchgeführt. | SWO-75131 | No | Main board renewed, software played up, device restarted, start and incubation times awaited, device equipped with test cassettes, system cleaning and QC performed. | Main board renewed, software played up, device restarted, start and incubation times awaited, device equipped with test cassettes, system cleaning and QC performed. | Defect | HW | Analyzer Control Module | PCB, Main board
AQT90 FLEX |  | B | D01BDD274E1BCF2F | Yes |  |  | Re classified RMED | FAN / TU
ABL80 FLEX BASIC | 更换测试卡 | 更换测试卡 | SWO-75080 | No | Replace the test card | Replace the test card
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-75066 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-75358 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL80 FLEX BASIC | 连续更换两张测试卡后修复 | 连续更换两张测试卡后修复 | SWO-75383 | No | Repair after replacing two test cards in a row | Repair after replacing two test cards in a row
ABL80 FLEX BASIC | 从库里拿了一个换上了(by andrew) | 从库里拿了一个换上了(by andrew) | SWO-75148 | No | Take a replacement from Curry (by andrew) | Take a replacement from Curry (by andrew)
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-75087 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL90 FLEX | Bzgl. Uhrzeitfehler Einstellungen in Systemsteuerung geändert, Gerät durchgespült, Einlassnadeln verbogen -> ersetzt, Einlassdichtung mit Halter ersetzt, danach Meldung Lösung können nicht gepumpt werden u. Leckage erkannt -> SC ersetzt, nach Startphase pCO2 Kal immer noch sehr niedrig SC Lot 279, Kunde besorgt andere SC und setzt diese selbst ein | Bzgl. Uhrzeitfehler Einstellungen in Systemsteuerung geändert, Gerät durchgespült, Einlassnadeln verbogen -> ersetzt, Einlassdichtung mit Halter ersetzt, danach Meldung Lösung können nicht gepumpt werden u. Leckage erkannt -> SC ersetzt, nach Startphase pCO2 Kal immer noch sehr niedrig SC Lot 279, Kunde besorgt andere SC und setzt diese selbst ein | SWO-75307 | No | Regarding time error settings changed in control panel, device flushed through, inlet needles bent -> replaced, inlet seal replaced with holder, after that message solution can not be pumped and leakage detected -> SC replaced, after start phase pCO2 Kal still very low SC Lot 279, customer takes care of other SC and uses them himself | Regarding time error settings changed in control panel, device flushed through, inlet needles bent -> replaced, inlet seal replaced with holder, after that message solution can not be pumped and leakage detected -> SC replaced, after start phase pCO2 Kal still very low SC Lot 279, customer takes care of other SC and uses them himself | Setup | SW | Time setting | Time setting
ABL90 FLEX | Lösungspack richtig aktiviert, Sensorkassette erneuert, Initialisierungsphase abgewartet, Kals und QCs durchgeführt | Der Kunde verlangt einen Vorort-Service. | SWO-75325 | No | Solution pack properly activated, sensor cartridge renewed, initialization phase awaited, kals and QCs performed | The customer requires a suburban service. | User error | SP Activation | SP Activation | SP Activation
ABL90 FLEX | Lösungspack erneuert, Kal wiederholt, QCs wiederholt i.O. Der Kunde braucht dringend eine Schulung. | Lösungspack erneuert, Kal wiederholt, QCs wiederholt i.O. Der Kunde braucht dringend eine Schulung. | SWO-75266 | No | Solution pack renewed, Kal repeated, QCs repeated i.O. The customer urgently needs training. | Solution pack renewed, Kal repeated, QCs repeated i.O. The customer urgently needs training. | Assisted Replacement | SP | SP | H2
ABL90 FLEX | Nach Neustart des ABL90 am 6.1. ca. 19:00Uhr, angeleitet durch Hotline ,befand sich das Gerät im Interventionsmodus. Medizintechnik startete am 07.1.2015 um 9:00Uhr eine manuelle Kal., danach alles o.k. Lösungspack überprüft, sowie Einlassdichtung, Pumpenrolle und Einlasssonde. Solutionpack PCB Kontakte gereinigt. LP wieder eingesetzt, kein Fehlverhalten. Weiterhin alles o.k. Einlasssonde als Ersatzteil bereitgestellt. Software V 3.1MR7 geladen, siehe ext. Bericht | Nach Neustart des ABL90 am 6.1. ca. 19:00Uhr, angeleitet durch Hotline ,befand sich das Gerät im Interventionsmodus. Medizintechnik startete am 07.1.2015 um 9:00Uhr eine manuelle Kal., danach alles o.k. Lösungspack überprüft, sowie Einlassdichtung, Pumpenrolle und Einlasssonde. Solutionpack PCB Kontakte gereinigt. LP wieder eingesetzt, kein Fehlverhalten. Weiterhin alles o.k. Einlasssonde als Ersatzteil bereitgestellt. Software V 3.1MR7 geladen, siehe ext. Bericht | SWO-75275 | No | After restarting the ABL90 on 6.1. approx. 19:00 o'clock, guided by hotline, the device was in intervention mode. Medical technology started on 07.1.2015 at 9:00 a.m. a manual cal., then all o.k. solution pack checked, as well as inlet seal, pump roller and inlet probe. Solutionpack PCB contacts cleaned. LP reinstated, no wrongdoing. Furthermore, all o.k. inlet probe provided as spare part. Software V 3.1MR7, see ext. Report | After restarting the ABL90 on 6.1. approx. 19:00 o'clock, guided by hotline, the device was in intervention mode. Medical technology started on 07.1.2015 at 9:00 a.m. a manual cal., then all o.k. solution pack checked, as well as inlet seal, pump roller and inlet probe. Solutionpack PCB contacts cleaned. LP reinstated, no wrongdoing. Furthermore, all o.k. inlet probe provided as spare part. Software V 3.1MR7, see ext. Report | No fault found | FT | FT | FT
ABL90 |  | B | D01CE859C4149B2C | No |  |  | Defect | HW | SC PCB | Ca & K
ABL90 |  | B | D025AB27FE19A37B | No |  |  | User error | Forgot to replace SP | Forgot to replace SP | Forgot to replace SP
TCM CombiM | 更换功能模块903-111一个后修复 | 更换功能模块903-111一个后修复 | SWO-74834 | No | Replace function module 903-111 one post-repair | Replace function module 903-111 one post-repair
ABL90 FLEX | 更换试剂包 | 更换试剂包 | SWO-75107 | No | Replace the reagent kit | Replace the reagent kit | Assisted Replacement | SP | SP | Pump cal error
AQT90 FLEX | 清洁存储盘，用借用的定标卡重新定标后仪器修复(by andrew) | 清洁存储盘，用借用的定标卡重新定标后仪器修复(by andrew) | SWO-74926 | No | Clean the storage tray and re-order the instrument with a borrowed calibration card (by andrew) | Clean the storage tray and re-order the instrument with a borrowed calibration card (by andrew) | Cleaned | HW | Storage Module | Storage wheel, Sensor | 5784
AQT90 FLEX | 更换910-304一个后修复(by andrew) | 更换910-304一个后修复(by andrew) | SWO-74899 | No | Replace 910-304 a post-fix (by-andrew) | Replace 910-304 a post-fix (by-andrew) | Defect | HW | Storage Module | Barcode reader, Storage wheel | BR1
ABL800 BASIC | Check the analyzer found that all the samples results are with blank for tHb & SO2 parameters. Replaced the Oxi module from their old analyzer, done the complete Oxi module test found neon lamp voltage are low, replaced the lamp it was under observation. | Found and replaced the Oxi module PCB from standby analyzer performed the complete oxi module test found it was all the results are OK, performed the 2p calibration and performed the tHb calibration found instur4ment was OK. | SWO-75546 | Yes | Check the analyzer found that all the samples results are with blank for tHb & SO2 parameters. Replaced the Oxi module from their old analyzer, done the complete Oxi module test found neon lamp voltage are low, replaced the lamp it was under observation. | Found and replaced the Oxi module PCB from standby analyzer performed the complete oxi module test found it was all the results are OK, performed the 2p calibration and performed the tHb calibration found instur4ment was OK. | Re classified RMED | Installation
ABL825 FLEX | Ｐｃｏ２電極先端を電極クリーナーで洗浄 Ｃａｌ２ ＱＣ ＯＫ | Ｐｃｏ２電極先端を電極クリーナーで洗浄 Ｃａｌ２ ＱＣ ＯＫ | SWO-75285 | No | Pco2 electrode tip washed with electrode cleaner Cal2 QC OK | Pco2 electrode tip washed with electrode cleaner Cal2 QC OK | Cleaned | Electrode | Electrode | pCO2
ABL825 FLEX | Oxi Modul geprüft und gereinigt, Funktionskontrolle ok | Oxi Modul geprüft und gereinigt, Funktionskontrolle ok | SWO-75210 | No | Oxi module tested and cleaned, function control ok | Oxi module tested and cleaned, function control ok | Cleaned | Fluidic pathway | Oxi module | Oxi module
ABL825 FLEX | Customer ran calibration to resolve but continues to come back | Replaced OXI PCB. Loaded AU software. Ran tHB cal in service and customer modes | D01CEC1FC42724CC | No | Customer ran calibration to resolve but continues to come back | Replaced OXI PCB. Loaded AU software. Ran tHB cal in service and customer modes | Defect | HW | Oxi module | PCB
ABL837 FLEX | ERROR 258 NEON NUMBER TOO LOW ERROR 516 FNEON NUMBER TOO LOW FULL OXI TEST RESULTED IN FAILED LAMP UNIT | REPLACED LAMP UNIT, ADJUSTED, ERRORS RESOLVED | D01CC63BC457E1B6 | No | ERROR 258 NEON NUMBER TOO LOW ERROR 516 FNEON NUMBER TOO LOW FULL OXI TEST RESULTED IN FAILED LAMP UNIT | REPLACED LAMP UNIT, ADJUSTED, ERRORS RESOLVED | Defect | HW | Oxi module | Lamp unit
AQT90 FLEX | Optische Einheit getauscht. | Optische Einheit getauscht. | SWO-75170 | No | Optical unit exchanged. | Optical unit exchanged. | Defect | HW | Process Module | Optical Unit | LQC
ABL80 FLEX BASIC | 更换新测试卡修复仪器 | 更换新测试卡修复仪器 | SWO-76082 | No | Replace the new test card repair instrument | Replace the new test card repair instrument
ABL80 FLEX BASIC | 更换新测试卡修复仪器 | 更换新测试卡修复仪器 | SWO-76083 | No | Replace the new test card repair instrument | Replace the new test card repair instrument
ABL80 FLEX BASIC | 更换测试卡正常。 | 更换测试卡正常。 | SWO-75332 | No | The replacement test card is OK. | The replacement test card is OK.
ABL80 FLEX BASIC | 更换测试卡正常。 | 更换测试卡正常。 | SWO-75309 | No | The replacement test card is OK. | The replacement test card is OK.
ABL80 FLEX | Run rinse multiple time. Flush sc80 with distilled water. Reinstalled SC. QC & sample Repeated. reading getting proper. preanalytical error. | B | SWO-75502 | No | Run rinse multiple time. Flush sc80 with distilled water. Reinstalled SC. QC & sample Repeated. reading getting proper. preanalytical error.
ABL80 FLEX BASIC | 更换新测试卡和试剂包后修复 | B | SWO-76029 | No | Repair after replacing new test cards and reagent kits
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-76813 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL80 FLEX Basic version |  | B | D020FB64262265EF | No
ABL80 FLEX CO-OX | Reinigung Fließwege, Reinigungsverfahren, | Frau Plate meldet das ihr ABL80coox Luftblasen erkennt. Frau P. hat das LP und die SC ohne Erfolg gewechselt. Ein Vorortservice wird gewünscht. | SWO-75421 | No | Cleaning flow paths, cleaning procedures, | Ms Plate reports that her ABL80coox detects air bubbles. Ms. P. changed the LP and the SC without success. A suburban service is desired.
ABL80 FLEX BASIC | 更换打印机控制板902-985一个后修复(by andrew) | 更换打印机控制板902-985一个后修复(by andrew) | SWO-74894 | No | Replace the printer dashboard 902-985 with a post-repair (by-andrew) | Replace the printer dashboard 902-985 with a post-repair (by-andrew)
ABL80 FLEX CO-OX | Sensorkassette und Lösungspack erneuert (alte Sensorkassette und Lösungspack werden reklamiert) | Sensorkassette und Lösungspack erneuert (alte Sensorkassette und Lösungspack werden reklamiert) | SWO-75504 | No | Sensor cassette and solution pack renewed (old sensor cassette and solution pack are claimed) | Sensor cassette and solution pack renewed (old sensor cassette and solution pack are claimed)
ABL80 FLEX BASIC | リセット、再取付 | B | SWO-76527 | No | Reset, re-install
ABL90 FLEX | neue LP geliefert da im Hause an anderen ABL tätig, Reklamationsausdruck erstellt, Zähler aufgenommen | B | SWO-75439 | No | new LP delivered since working in-house at other ABL, complaint expression created, counter included |  | Assisted Replacement | SP | SP | Logistical upSWO | pCO2
ABL90 FLEX | Temporäre Dateien gelöscht, danach war Speicherproblem beseitigt. Freier Speicher wieder 240 MB. Neue Patres aufgespielt - Ausdruck lt. Frau Guice in Ordnung. Postfilter hinzugefügt. tHb Kalibration, Kalibration und QC´s durchgeführt. | Temporäre Dateien gelöscht, danach war Speicherproblem beseitigt. Freier Speicher wieder 240 MB. Neue Patres aufgespielt - Ausdruck lt. Frau Guice in Ordnung. Postfilter hinzugefügt. tHb Kalibration, Kalibration und QC´s durchgeführt. | SWO-75372 | No | Temporary files deleted, after that memory problem was fixed. Free memory again 240 MB. New Fathers played - expression according to Mrs. Guice in order. Post filter added. tHb calibration, calibration and QC's. | Temporary files deleted, after that memory problem was fixed. Free memory again 240 MB. New Fathers played - expression according to Mrs. Guice in order. Post filter added. tHb calibration, calibration and QC's. | Defect | SW | Disc space / Temp files | Disc space / Temp files
ABL90 FLEX | Flüssigkeitsstrecke gespült, Einlassnadel ersetzt und erneut gespült. Sämtliche flüssigkeitsbezogenen Serviceprogramme durchgeführt, alle ohne Fehler. Neustart, tHb-Kal, Funktionstest. Testmessung mit Blut: OK | Flüssigkeitsstrecke gespült, Einlassnadel ersetzt und erneut gespült. Sämtliche flüssigkeitsbezogenen Serviceprogramme durchgeführt, alle ohne Fehler. Neustart, tHb-Kal, Funktionstest. Testmessung mit Blut: OK | SWO-75393 | No | Liquid section rinsed, inlet needle replaced and rinsed again. All liquid-related service programs performed, all without errors. Restart, tHb-Kal, function test. Test measurement with blood: OK | Liquid section rinsed, inlet needle replaced and rinsed again. All liquid-related service programs performed, all without errors. Restart, tHb-Kal, function test. Test measurement with blood: OK | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
ABL90 | Confirmed issue. S7730 QC out of range low w error 521 Inhomogeneous sample, S7740 out of range high, S7750 out of range high. Calibration passing within specifications. | Performed visual inspection and ran service diagnostics, no issues noted or observed. Replace sensor cassette to resolve issue. | D01E17D83A54F8B9 | No | Confirmed issue. S7730 QC out of range low w error 521 Inhomogeneous sample, S7740 out of range high, S7750 out of range high. Calibration passing within specifications. | Performed visual inspection and ran service diagnostics, no issues noted or observed. Replace sensor cassette to resolve issue. | Assisted Replacement | SC | SC | QC ext | pO2
TCM CombiM |  | B | SWO-75348 | No
AQT90 FLEX | 更换定标卡，用借用的定标卡重新定标后修复(by andrew) | 更换定标卡，用借用的定标卡重新定标后修复(by andrew) | SWO-75282 | No | Replace the calibration card and repair it after re-calibration with a borrowed calibration card (by andrew) | Replace the calibration card and repair it after re-calibration with a borrowed calibration card (by andrew) | Assisted Replacement | Consumables | Test Kit | D-dimer
AQT90 FLEX | 更换903-364一个正常。(by andrew) | 更换903-364一个正常。(by andrew) | SWO-74947 | No | Replace 903-364 one normal. (by andrew) | Replace 903-364 one normal. (by andrew) | Defect | HW | Process Module | Optical Unit | 131
AQT90 FLEX | 更换光学单元903-364一个后修复(by andrew) | 更换光学单元903-364一个后修复(by andrew) | SWO-74919 | No | Replace the optical unit 903-364 with a post-repair (by andrew) | Replace the optical unit 903-364 with a post-repair (by andrew) | Defect | HW | Process Module | Optical Unit | 131
ABL800 BASIC | Found and replaced the Oxi module PCB from standby analyzer performed the complete oxi module test found it was all the results are OK, performed the 2p calibration and performed the tHb calibration found instur4ment was OK. | Found and replaced the Oxi module PCB from standby analyzer performed the complete oxi module test found it was all the results are OK, performed the 2p calibration and performed the tHb calibration found instur4ment was OK. | SWO-75548 | Yes | Found and replaced the Oxi module PCB from standby analyzer performed the complete oxi module test found it was all the results are OK, performed the 2p calibration and performed the tHb calibration found instur4ment was OK. | Found and replaced the Oxi module PCB from standby analyzer performed the complete oxi module test found it was all the results are OK, performed the 2p calibration and performed the tHb calibration found instur4ment was OK. | Re classified RMED
ABL835 FLEX | 882 mixer tray slip errors will not clear. | Replaced and aligned Flex Q module. Got ACE.dll not found error when attempting to load AU software. Reinstalled application software and continued with alignment. Performed maintenance due including pump tubing change. | D01E00C63A0ADC86 | No | 882 mixer tray slip errors will not clear. | Replaced and aligned Flex Q module. Got ACE.dll not found error when attempting to load AU software. Reinstalled application software and continued with alignment. Performed maintenance due including pump tubing change. | Defect | HW | FlexQ Module | Unknown
ABL835 FLEX | pH sensitivity bordering on low limit. Out of limit on arrival. Lactate unstable on QC level 2. | Replaced pH electrode. Performed maintenance due including pump tubing change. Lactate still unstable. Remembraned. | D01E181A3A20DC8A | No | pH sensitivity bordering on low limit. Out of limit on arrival. Lactate unstable on QC level 2. | Replaced pH electrode. Performed maintenance due including pump tubing change. Lactate still unstable. Remembraned. | Assisted Replacement | Electrode | Electrode | pH
ABL800 BASIC | Restarted the analyzer. Performed touch screen calibration. Now working fine. | B | SWO-75513 | No | Restarted the analyzer. Performed touch screen calibration. Now working fine. |  | Adjustment | HW | DMS_PC | Touch Screen
ABL825 FLEX | Allgemeine Überprüfung - Erneuerung Einlassdichtung, Verschlauchung unter Einlassdichtung (Abfallstrecke) - Nippel Kit - Funktionsüberprüfung etc. Bem.: Benutzerverwaltung bereits in Bearbeitung durch Radiometer Mitarbeiter Hr. Rösch | Allgemeine Überprüfung - Erneuerung Einlassdichtung, Verschlauchung unter Einlassdichtung (Abfallstrecke) - Nippel Kit - Funktionsüberprüfung etc. Bem.: Benutzerverwaltung bereits in Bearbeitung durch Radiometer Mitarbeiter Hr. Rösch | SWO-75451 | No | General inspection - Renewal inlet seal, hose under inlet seal (waste section) - Nipple Kit - Function check etc. Bem.: User administration already in progress by Radiometer employee Mr. Rösch | General inspection - Renewal inlet seal, hose under inlet seal (waste section) - Nipple Kit - Function check etc. Bem.: User administration already in progress by Radiometer employee Mr. Rösch | Defect | HW | Inlet module | Leaking | Nipple kit
ABL815 FLEX | Lampenspannung im Serviceprogramm überprüft und justiert. Funktion OK. | Mit Sw. Tanja das Gerät neu gestartet, 1.Kal. durchgeführt, Fehler weiter vorhanden. Service Einsatz notwendig, der Kunde bittet um Terminabsprache | SWO-75609 | No | Lamp voltage in the service program is checked and adjusted. Function OK. | With Sw. Tanja the device is restarted, 1.Kal. performed, errors still present. Service service required, the customer asks for appointment | Adjustment | HW | Oxi module | Lamp unit
ABL825 FLEX | Display ausgetauscht. | Display ausgetauscht. | SWO-75608 | No | Display replaced. | Display replaced. | Defect | HW | DMS_PC | Touch Screen
AQT90 FLEX | routine visit. machine working fine. | Guided customer for EDTI tubes.Run the cleaning tube 5 time & run the sample test in EDTI tube. results get without any error massage. | SWO-75521 | No | routine visit. machine working fine. | Guided customer for EDTI tubes. Run the cleaning tube 5 time & run the sample test in EDTI tube. results get without any error massage. | Cleaned | HW | Aspiration Module | Needle Wash Unit
TCM-COMBIM |  | B | D01E29C43A3F0E46 | No
ABL90 FLEX | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-75547 | No | Repair after replacing a new test card | Repair after replacing a new test card | Assisted Replacement | SC | SC | Reference
ABL90 FLEX | 装置分解清掃 部品交換 パック交換 | B | SWO-75814 | No | Equipment Disassembly Cleaning Parts Replacement Pack Replacement |  | Defect | HW | PM Kit | PM Kit
AQT90 FLEX | 做了整机的保养，重点保养了处理盘，然后用定标卡重新定标后修复(by andrew) | 做了整机的保养，重点保养了处理盘，然后用定标卡重新定标后修复(by andrew) | SWO-75304 | No | Did the maintenance of the whole machine, focus on the maintenance of the processing disk, and then with the calibration card re-calibration after repair (by andrew) | Did the maintenance of the whole machine, focus on the maintenance of the processing disk, and then with the calibration card re-calibration after repair (by andrew) | Cleaned | HW | Process Module | Process wheel | 161
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-75604 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL80 FLEX BASIC | 更换测试卡 | 更换测试卡 | SWO-75829 | No | Replace the test card | Replace the test card
ABL80 FLEX BASIC | 连续更换两张测试卡后修复 | 连续更换两张测试卡后修复 | SWO-75801 | No | Repair after replacing two test cards in a row | Repair after replacing two test cards in a row
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-75610 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL80 OSM | analyzer will not interface no lights on the connection when plugged in | contacted IT ( Richard Turman) walked me through trying to check windows and configuration stil not working; replaced the computer unit; pluged in connection Radiance works; they are still using a test IP address | D023116C1248CACE | No | analyzer will not interface no lights on the connection when plugged in | contacted IT ( Richard Turman) walked me through trying to check windows and configuration stil not working; replaced the computer unit; pluged in connection Radiance works; they are still using a test IP address
ABL90 | Ran Complete Oxi=OK. Check Valve operation.Suspect Sensor Cassette. | Suggested to customer to replace SC. | D021EAE19C4BC622 | No | Ran Complete oxygen = OK. Check Valve operation. Suspect Sensor Cassette. | Suggested to customer to replace SC. | Assisted Replacement | SC | SC | Wetability
ABL90 |  | B | D02EAA3538577C90 | No |  |  | Assisted Replacement | SC | SC | sensitivity | Glu
ABL90 |  | B | D021EA2E9C0A2ACE | No |  |  | Assisted Replacement | SC | SC | backlog error
ABL90 | Temperature corrected value not showing on Printout. | Turned on Enable parameters out of Spec in Miscellaneous setup to give results irregardless of measured values. | D02217489C1BC629 | No | Temperature corrected value not showing on Printout. | Turned on Enable parameters out of Spec in Miscellaneous setup to give results irregardless of measured values. | Setup | results outside reportable | results outside reportable | results outside reportable
ABL90 FLEX PLUS | Analysator im Service etc. überprüft - Erneuerung Einlassnadel (Biegung) und Einlassdichtung - SP nochmals installiert - Erneuerung Sensorkassette notwendig (aus ZNA geholt) - Funktionsüberprüfung etc. ohne Beanstandung ! | Analysator im Service etc. überprüft - Erneuerung Einlassnadel (Biegung) und Einlassdichtung - SP nochmals installiert - Erneuerung Sensorkassette notwendig (aus ZNA geholt) - Funktionsüberprüfung etc. ohne Beanstandung ! | SWO-75776 | No | Analyzer in service etc. checked - renewal inlet needle (bending) and inlet seal - SP reinstalled - renewal sensor cassette necessary (taken from ZNA) - function check etc. without complaint ! | Analyzer in service etc. checked - renewal inlet needle (bending) and inlet seal - SP reinstalled - renewal sensor cassette necessary (taken from ZNA) - function check etc. without complaint ! | Assisted Replacement | SC | SC | QC Level C | Lac
ABL90 FLEX PLUS | Fehlersuche, div. Programme im Service durchgef., Einlaß entfernt, Gerät aus, Einlaß montiert, Gerät ein, Startphase abgewartet, Kal. + QC i.O. | Fehlersuche; div. Programme im Service durchgef.; Einlaß entfernt; Gerät aus; Einlaß montiert; Gerät ein; Startphase abgewartet; Kal. + QC i.O. | SWO-75689 | No | Troubleshooting; various programs in the service; Inlet removed; device off; Entrance mounted; device; start-up phase awaited; Kal. + QC in the first place | Troubleshooting; various programs in the service; Inlet removed; device off; Entrance mounted; device; start-up phase awaited; Kal. + QC in the first place | Adjustment | HW | Auto Inlet module | Auto Inlet module
ABL90 FLEX | 更换新试剂包后修复 | 更换新试剂包后修复 | SWO-75637 | No | Repair after replacing a new kit | Repair after replacing a new kit | Assisted Replacement | SP | SP | Leak current
ABL90 | SE SOLUCIONAN DIFERENTES PROBLEMAS NEUMATICOS, CON LIMPIEZA Y SUSTITUCION DE PIEZAS DE ENTRADA, ASI COMO LECTOR DE CHIP POR NO FUNCIONAR CORRECTAMENTE. CALIBRADO Y VERIFICADO CON Q.C. | B | D02234029C59D5D5 | No | DIFFERENT PNEUMATIC PROBLEMS ARE SOLVED, WITH CLEANING AND SUBSTITUTE OF INPUT PARTS, AS A CHIP READER FOR NOT WORKING CORRECTLY. CALIBRATED AND VERIFIED WITH Q.C. |  | Defect | HW | Basic Unit | Sensor Interface PCB
AQT90 FLEX |  | B | SWO-75576 | No |  |  | Defect | HW | Cover and Chassis | PCB, Cabinet | 2995
ABL825 FLEX | 连续更换两个Lac电极膜后恢复 | 连续更换两个Lac电极膜后恢复 | SWO-76012 | No | Recovery after two consecutive Replaces of Lac Electrode Films | Recovery after two consecutive Replaces of Lac Electrode Films | Assisted Replacement | Electrode | Membrane | Lac
ABL825 FLEX | Carosel alignment out of spec.Also Probe bent. | Align AutoCheck.Probe replaced by customer. | D021EA209C45C621 | No | Carosel alignment out of spec. Also Probe bent. | Align AutoCheck.Probe replaced by customer. | Adjustment | HW | AutoCheck Module | Alignment
ABL837 FLEX |  | B | D021F3C59C41E9BE | No |  |  | Defect | HW | Inlet module | Probe bend | Curve needle
ABL825 FLEX | Analysator heruntergefahren - Erneuerung Waste Detector - Funktionsüberprüfung etc. | Analysator heruntergefahren - Erneuerung Waste Detector - Funktionsüberprüfung etc. | SWO-75703 | No | Analyzer shut down - Renewal Waste Detector - Function check etc. | Analyzer shut down - Renewal Waste Detector - Function check etc. | Defect | HW | Basic unit | Waste detector
ABL825 FLEX | Customer ran calibration but error came back | Replaced OXI PCB. Adjusted lamp voltage. During calibration, error code 0579 OXI hardware error posted. Replaced another OXI PCB. Adjusted lamp voltage. Ran tHB cal in service and customer mode | D021C5B19C1D266E | No | Customer ran calibration but error came back | Replaced OXI PCB. Adjusted lamp voltage. During calibration, error code 0579 OXI hardware error posted. Replaced another OXI PCB. Adjusted lamp voltage. Ran tHB cal in service and customer mode | Defect | HW | Oxi module | PCB
ABL825 FLEX | Tried new lot of membranes and swapped electrodes between 2 instruments-No change | Replaced measuring chamber, reference and Ca++ membranes | D021EB5D9C3089C6 | No | Tried new lot of membranes and swapped electrodes between 2 instruments-No change | Replaced measuring chamber, reference and Ca++ membranes | Defect | HW | El/Met Module | Measuring chamber
ABL820 FLEX | user stated that they have to reboot analyzer to fix it from being frozen | restored radiometer defaults and reloaded setup to correct the issue with analyzer per Matt | D021FBED9C501202 | No | user stated that they have to reboot analyzer to fix it from being frozen | restored radiometer defaults and reloaded setup to correct the issue with analyzer per Matt | Defect | SW | SW | SW
ABL825 FLEX | Lamp Unit von 3900 auf 3250 justiert. Einlassonde und Einlassdichtung erneuert. Dichtungsring an der pH Elektrode erneuert, pCO2 Elektrode membraniert, Gerät kalibriert. | Lamp Unit von 3900 auf 3250 justiert. Einlassonde und Einlassdichtung erneuert. Dichtungsring an der pH Elektrode erneuert, pCO2 Elektrode membraniert, Gerät kalibriert. | SWO-75855 | No | Lamp Unit from 3900 to 3250. Inlet probe and inlet seal renewed. Sealing ring on the pH electrode renewed, pCO2 electrode membraned, device calibrated. | Lamp Unit from 3900 to 3250. Inlet probe and inlet seal renewed. Sealing ring on the pH electrode renewed, pCO2 electrode membraned, device calibrated. | Adjustment | HW | Oxi module | Lamp unit
ABL825 FLEX | pCO2 Elektrode ern., Gerät kalibriert, QCs durchgeführt | B | SWO-75664 | No | pCO2 electrode ern., device calibrated, QCs performed |  | Assisted Replacement | Electrode | Electrode | pCO2
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-75834 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-75840 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-76823 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL80 FLEX CO-OX | Den Analysator gereinigt. Ventile gespült. LP und SC wieder eingesetzt. KAL und QC ok. | Den Analysator gereinigt. Ventile gespült. LP und SC wieder eingesetzt. KAL und QC ok. | SWO-75898 | No | Clean the analyzer. Valves rinsed. LP and SC. KAL and QC ok. | Clean the analyzer. Valves rinsed. LP and SC. KAL and QC ok.
ABL80 FLEX BASIC | Solution pack seating is not proper and the O ring binding that might have broken. Wrapped i with tight insulation., Has to replace the spare., Lactate SC was replaced and installed a new one . The calibration was ok | B | SWO-76746 | No | Solution pack seating is not proper and the O ring binding that might have broken. Wrapped i with tight insulation., Has to replace the spare., Lactate SC was replaced and installed a new one . The calibration was ok
ABL80 FLEX CO-OX | Drucker gereinigt und Papierreste entfernt | Drucker gereinigt und Papierreste entfernt | SWO-75994 | No | Printer cleaned and paper leftovers removed | Printer cleaned and paper leftovers removed
ABL80 FLEX BASIC | 先週間違えて溶液ﾊﾟｯｸを交換してしまったので、元の溶液ﾊﾟｯｸをリセットして使用していた為、機器ではまだｴﾗｰは出ていなかったが、実際の溶液残量がなくなってしまい、その状態で検体測定をし、洗浄するための溶液がない為、そのままｶｾｯﾄ内で血液詰まってしまった。ｾﾝｻｰｶｾｯﾄ交換、溶液ﾊﾟｯｸ交換で改善。 | 先週間違えて溶液ﾊﾟｯｸを交換してしまったので、元の溶液ﾊﾟｯｸをリセットして使用していた為、機器ではまだｴﾗｰは出ていなかったが、実際の溶液残量がなくなってしまい、その状態で検体測定をし、洗浄するための溶液がない為、そのままｶｾｯﾄ内で血液詰まってしまった。ｾﾝｻｰｶｾｯﾄ交換、溶液ﾊﾟｯｸ交換で改善。 | SWO-77810 | No | Because I had replaced the solution pack by mistake last week, because I had used to reset the original solution pack, but the equipment had not yet issued an error, but there is no actual solution level, and the sample measurement in that state, because there is no solution for washing, The blood has become blocked in the cassette as it is. Improved by sensor cassette replacement and solution pack replacement. | Because I had replaced the solution pack by mistake last week, because I had used to reset the original solution pack, but the equipment had not yet issued an error, but there is no actual solution level, and the sample measurement in that state, because there is no solution for washing, The blood has become blocked in the cassette as it is. Improved by sensor cassette replacement and solution pack replacement.
ABL90 FLEX | 流路洗浄。 | 流路洗浄。 | SWO-78423 | No | Flow path cleaning. | Flow path cleaning. | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
ABL90 FLEX | Austausch der Batterie Option alter Akku 910-311P0065N010 | Austausch der Batterie Option alter Akku 910-311P0065N010 | SWO-75977 | No | Battery Replacement Option old battery 910-311P0065N010 | Battery Replacement Option old battery 910-311P0065N010 | Defect | HW | Battery module | Battery
ABL90 FLEX | Austausch der Batterie Option alte 910-311P0067N012, Uhrzeiteinstellungen lt. RDEU Vorgaben geprüft | Austausch der Batterie Option alte 910-311P0067N012, Uhrzeiteinstellungen lt. RDEU Vorgaben geprüft | SWO-75978 | No | Replacement of battery option old 910-311P0067N012, time settings checked according to RDEU specifications | Replacement of battery option old 910-311P0067N012, time settings checked according to RDEU specifications | Defect | HW | Battery module | Battery
ABL90 | CONFIRMED ISSUE WHEN ONSITE. | FLUSHED SYSTEM W HOT WATER. REPLACED PROBE AND HOLDER. | D0235DF91244F7EB | No | CONFIRMED ISSUE WHEN ONSITE. | FLUSHED SYSTEM W HOT WATER. REPLACED PROBE AND HOLDER. | Cleaned | Fluidic pathway | Replaced probe + gasket | Replaced probe + gasket
ABL90 FLEX | Gerät gespült.Flüss.strecke durchgängig Gebrauchten.SP wieder eingesetzt. Kal und QC abgewartet | der Anwender hat schon mehrereSC90 LösungsPacks und mehrere SC90 Sensorkassetten erneuert, allerdings ohne Erfolg; die Flüssigkeitsstrecke mit dem Spülwerkzeug mit Wasser und Luft durchgespült, (es tröpfelt nur aus der Einlasssonde); der Anwender hat kein weiteres Material mehr vorrätig und wünscht einen Vorort-Service zur Beseitigung der Geräte-Störung; | SWO-76304 | No | device rinsed. Flüss.strecke used throughout used.SP reinstated. Kal and QC waited | the user has already renewed severalSC90 solution packs and several SC90 sensor cartridges, but to no avail; the liquid section with the rinsing tool rinsed with water and air, (it drips only from the inlet probe); the user no longer has any other material in stock and wants a suburban service to eliminate the equipment malfunction; | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
ABL90 FLEX PLUS |  | B | D0247DF18812A908 | No |  |  | Setup | SW | Time setting | Time setting
ABL80 CO-OX |  | B | D02304E712220B8D | Yes |  |  | Re classified RMED | not upSWO
ABL90 FLEX | 更换新试剂包修复仪器 | 更换新试剂包修复仪器 | SWO-76147 | No | Replace the new reagent kit repair instrument | Replace the new reagent kit repair instrument | Assisted Replacement | SP | SP | Pump cal error
ABL90 | COMPROBACIÓN DEL ESTADO DE FUNCIONAMIENTO DEL EQUIPO, LIMPIEZA DEL CIRCUITO NEUMÁTICO, SUSTITUCION DE TODOS LOS TUBOS DEL EQUIPO, SUSTITUCIÓN DEL INLET PROBE Y GASKET, COMPROBACION FINAL DE FUNCIONAMIENTO CON QC. | B | D02302BA1256665F | No | CHECKING THE STATE OF OPERATION OF THE EQUIPMENT, CLEANING THE TIRE CIRCUIT, SUBSTITUTE OF ALL EQUIPMENT TUBES, SUBSTITUTE OF THE PROBE AND GASKET INSTITUTION, FINAL OPERATION CHECK WITH QC. |  | Defect | HW | PM Kit | PM Kit
ABL90 | Verificação dos registros, troca da entrada, mantém-se erro, troca de sonda, mantém-se erro, execução do kit de tubos, mantém-se erro, troca de pack de soluções, funcionamento correcto. | B | D02363CE1211112B | No | Checking records, exchanging input, keeping error, changing probe, keeping error, running the tube kit, keeping error, changing solution Pack, correct operation. |  | Assisted Replacement | SP | SP | FT
AQT90 FLEX |  | B | SWO-75856 | No |  |  | Defect | HW | Incubator modul | Rubber disk | 1999.61.22
ABL815 FLEX |  | B | D0231A7412104206 | No |  |  | Defect | HW | AutoCheck Module | Barcode reader
ABL825 FLEX | Einweisung durchgeführt. Kalibrierungen durchgeführt. QCs für alle Level durchgeführt. ABL Messbereit. | Einweisung durchgeführt. Kalibrierungen durchgeführt. QCs für alle Level durchgeführt. ABL Messbereit. | SWO-76059 | No | instruction carried out. calibrations. QCs for all levels. ABL Ready to measure. | instruction carried out. calibrations. QCs for all levels. ABL Ready to measure. | User Error | No error | No error | No error
ABL837 FLEX | Elektroden neu bespannt. Empfindlichkeit 3 bei 74% Kalibrationen und Qualitätskontrollen in Ordnung. Noch vereinzelt Driftfehler Crea. | Elektroden neu bespannt. Empfindlichkeit 3 bei 74% Kalibrationen und Qualitätskontrollen in Ordnung. Noch vereinzelt Driftfehler Crea. | SWO-75987 | No | electrodes. Sensitivity 3 at 74% calibrations and quality controls fine. Still sporadic drift error Crea. | electrodes. Sensitivity 3 at 74% calibrations and quality controls fine. Still sporadic drift error Crea. | Assisted Replacement | Electrode | Membrane | Crea
ABL820 FLEX | Messkammern und Elektroden Kontakte gereinigt, pH Dichtung ersetzt. Die als defekt markierte Spüllösung erneut eingescannt, Füllstand bei 81%. | Messkammern und Elektroden Kontakte gereinigt, pH Dichtung ersetzt. Die als defekt markierte Spüllösung erneut eingescannt, Füllstand bei 81%. | SWO-76093 | No | Measuring chambers and electrodes contacts cleaned, pH seal replaced. The rinsing solution marked as defective is rescanned, level at 81%. | Measuring chambers and electrodes contacts cleaned, pH seal replaced. The rinsing solution marked as defective is rescanned, level at 81%. | Assisted Replacement | Electrode | Electrode | pH | Sealing ring
ABL825 FLEX |  | B | D024594C880136E7 | No |  |  | Defect | SW | - | Re load SW
ABL825 FLEX |  | B | D024564D884336E4 | No |  |  | Defect | SW | - | Re load SW
ABL825 FLEX | found that it is getting no leading air segment. found that the gel at the bottom of venus tube got into the system | replaced all the parts listed to clear the gel out of the system. | D02348461240EEA7 | No | found that it is getting no leading air segment. found that the gel at the bottom of venus tube got into the system | replaced all the parts listed to clear the gel out of the system. | Cleaned | Fluidic pathway | pH/BG Module | Measuring chamber
ABL80 FLEX BASIC | 更换新测试卡修复仪器 | 更换新测试卡修复仪器 | SWO-76148 | No | Replace the new test card repair instrument | Replace the new test card repair instrument
ABL80 FLEX BASIC | 连续更换两张测试卡后修复 | 连续更换两张测试卡后修复 | SWO-76259 | No | Repair after replacing two test cards in a row | Repair after replacing two test cards in a row
ABL80 FLEX CO-OX | SC gespült, IQC wiederholt, LP ersetzt, Kal u. IQC i.O. | SC gespült, IQC wiederholt, LP ersetzt, Kal u. IQC i.O. | SWO-76248 | No | SC rinsed, IQC repeated, LP replaced, Kal & IQC i.O. | SC rinsed, IQC repeated, LP replaced, Kal & IQC i.O.
ABL80 FLEX BASIC | After consulting this issue with Mr. Subba Rao sir, he recommended me to order the CF card, so placed an order for CF card. Installed the old analyzer as a temporary solution. | After consulting this issue with Mr. Subba Rao sir, he recommended me to order the CF card, so placed an order for CF card. Installed the old analyzer as a temporary solution. | SWO-76736 | No | After consulting this issue with Mr. Subba Rao sir, he recommended me to order the CF card, so placed an order for CF card. Installed the old analyzer as a temporary solution. | After consulting this issue with Mr. Subba Rao sir, he recommended me to order the CF card, so placed an order for CF card. Installed the old analyzer as a temporary solution.
ABL90 | Inspected tubings and customer performed several flushes. | I retubed instrument. | D024485388307316 | No | Inspected tubings and customer performed several flushes. | I retubed instrument. | Defect | HW | PM Kit | PM Kit
ABL90 FLEX | Nettoyage de la section fluidique e.o Remplacé le kit de maintenance annuel e.o Calibration e.o contrôle de qualité e.o | Nettoyage de la section fluidique e.o Remplacé le kit de maintenance annuel e.o Calibration e.o contrôle de qualité e.o | SWO-76193 | No | Cleaning the fluidic section e.o Replaced the annual maintenance kit e.o Calibration e.o quality control e.o | Cleaning the fluidic section e.o Replaced the annual maintenance kit e.o Calibration e.o quality control e.o | Cleaned | Fluidic pathway | replaced yearly kit | replaced yearly kit
ABL90 FLEX | Fixe Zielwerte überprüft, aktuelle Grenzwert Tabelle an Kunden übergeben. Lösungspack ersetzt. | Fixe Zielwerte überprüft, aktuelle Grenzwert Tabelle an Kunden übergeben. Lösungspack ersetzt. | SWO-76301 | No | Fixed target values checked, current limit table passed to customers. solution pack replaced. | Fixed target values checked, current limit table passed to customers. solution pack replaced. | Assisted Replacement | SP | SP | pCO2 C low
ABL90 FLEX | 更换新试剂包后修复 | 更换新试剂包后修复 | SWO-76256 | No | Repair after replacing a new kit | Repair after replacing a new kit | Assisted Replacement | SP | SP | Pump cal error
AQT90 FLEX | サンプル検知センサーの汚れ（酸化鉄：薬品の影響らしい）があり清掃。定期的に清掃を行う必要がある。クリーニングの実施。 | サンプル検知センサーの汚れ（酸化鉄：薬品の影響らしい）があり清掃。定期的に清掃を行う必要がある。クリーニングの実施。 | SWO-76517 | No | There is a dirt of the sample detection sensor (iron oxide: it seems to be the influence of the chemical) and cleaning. It is necessary to clean it regularly. Cleaning is carried out. | There is a dirt of the sample detection sensor (iron oxide: it seems to be the influence of the chemical) and cleaning. It is necessary to clean it regularly. Cleaning is carried out. | Cleaned | HW | Aspiration Module | Barcode reader, Inlet
ABL800 BASIC | Found UPS was failed need to replace UPS Battery’s, performed the Complete oximetry test, adjusted the lamp unit performed the 2p calibration found instrument is functional. | Found UPS was failed need to replace UPS Battery’s, performed the Complete oximetry test, adjusted the lamp unit performed the 2p calibration found instrument is functional. | SWO-76222 | No | Found UPS was failed need to replace UPS Battery’s, performed the Complete oximetry test, adjusted the lamp unit performed the 2p calibration found instrument is functional. | Found UPS was failed need to replace UPS Battery’s, performed the Complete oximetry test, adjusted the lamp unit performed the 2p calibration found instrument is functional. | Defect | HW | Power Supply | Oxi error
ABL815 FLEX |  | B | SWO-76031 | No |  |  | Unknown | HW | Oxi module | Oxi module
ABL825 FLEX | Se reinicia equipo. Se comprueba fusibles de analyzer control PCB. Se reinstala software. El fallo de la analyzer controler PCB persiste. | B | D02E809D38639C51 | No | The computer restarts. PCB control analyzer fuses are checked. Software is reinstalled. Pcb controlanalyzer failure persists. |  | Unknown | SW | His/lis | HIS/LIS
ABL825 FLEX | Oxymetrie Einheit gewechselt, Thb Kal. durchgeführt. Mehrfach Kalibriert. | Oxymetrie Einheit gewechselt, Thb Kal. durchgeführt. Mehrfach Kalibriert. | SWO-76299 | No | Oxymetry unit changed, Thb Kal. Performed. Multiple calibrated. | Oxymetry unit changed, Thb Kal. Performed. Multiple calibrated. | Defect | HW | Oxi module | Oxi module | Err. 579
ABL825 FLEX | Customer calibrates to resolve | Returned to site and replaced power supply, OXI PCB,lamp umit and fiber optic cable. Adjusted lamp voltage. Ran tHB cal in service and customer modes | D02451C788342BCF | No | Customer calibrates to resolve | Returned to site and replaced power supply, OXI PCB,lamp umit and fiber optic cable. Adjusted lamp voltage. Ran tHB cal in service and customer modes | Defect | HW | Power Supply | Oxi module | 579
ABL825 FLEX | Examined data logs. Found multiple Oxi errors and sample errors in patient sample logs.Found tHb value at 468 on calibrations. | Adjusted halogen lamp gain to 3200. Replaced inlet and pump tubing. Performed tHb calibration. Verified operation of ABL825. | D025F169FE278696 | No | Examined data logs. Found multiple Oxi errors and sample errors in patient sample logs. Found tHb value at 468 on calibrations. | Adjusted halogen lamp gain to 3200. Replaced inlet and pump tubing. Performed tHb calibration. Verified operation of ABL825. | Adjustment | HW | Oxi module | Lamp unit
ABL825 FLEX | Examined data logs. Found multiple oximetry mismatch errors. Examined Oxi moduel. Found tHB value at 500 on calibration. | Adjusted halogen lamp gain to 3200. Performed tHB calibration. Cleaned electrode wells. Replaced inlet gasket and tubing. | D025EB39FE0C8691 | No | Examined data logs. Found multiple oximetry mismatch errors. Examined Oxi moduel. Found tHB value at 500 on calibration. | Adjusted halogen lamp gain to 3200. Performed tHB calibration. Cleaned electrode wells. Replaced inlet gasket and tubing. | Adjustment | HW | Oxi module | Lamp unit
ABL825 FLEX | Examined data logs. Found multiple oxi mismatch errors. Examined Oxi module. Found tHb value reading high on 2-point calibrations. | Adjusted halogen lamp gain. Ran tHb calibration. Verified operation of ABL825. | D027709174260A7D | No | Examined data logs. Found multiple oxi mismatch errors. Examined Oxi module. Found tHb value reading high on 2-point calibrations. | Adjusted halogen lamp gain. Ran tHb calibration. Verified operation of ABL825. | Adjustment | HW | Oxi module | Lamp unit
ABL825 FLEX | Cl.- Elektrode mit Dummy ersetzt. Kalibrationen mehrfach wiederholt. Rilibäkgrenzen eingestellt. Ampelfunktion verändert. | Cl.- Elektrode mit Dummy ersetzt. Kalibrationen mehrfach wiederholt. Rilibäkgrenzen eingestellt. Ampelfunktion verändert. | SWO-76637 | No | Cl.- Electrode replaced with dummy. calibrations repeated several times. Rilibäk grenzen. traffic light function changed. | Cl.- Electrode replaced with dummy. calibrations repeated several times. Rilibäk grenzen. traffic light function changed. | Defect | Electrode | Electrode | pCO2
ABL825 FLEX | System was green when FSE arrived onsite. STAYED BEHIND FOR CUSTOMERS TO RUN SOME SAMPLES TO ENSURE SYSTEM IS STABLE BEFORE LEAVING. | Performed lamp adjustment to decrease light count down from 3670 to 3400 to correct Error 688. | D0245918883DF7EE | No | System was green when FSE arrived onsite. STAYED BEHIND FOR CUSTOMERS TO RUN SOME SAMPLES TO ENSURE SYSTEM IS STABLE BEFORE LEAVING. | Performed lamp adjustment to decrease light count down from 3670 to 3400 to correct Error 688. | Adjustment | HW | Oxi module | Lamp unit
ABL825 FLEX | Datensicherung, Software v6.16 installiert, Datenrücksicherung | Messkammer CO2 gereinigt, pCO2-Elektrode (alt: ST05) ersetzt, Mehrere Kal abgewartet, QK ok | SWO-76360 | No | Data backup, software v6.16 installed, data backup | Measuring chamber CO2 cleaned, pCO2 electrode replaced (old: ST05), Several kal waited, QK ok | Assisted Replacement | Electrode | Electrode | pCO2
AQT90 FLEX |  | B | D02475908836BDBD | No |  |  | Performance | Discrepant result | Test Kit | False positive TnT
ABL80 FLEX BASIC |  | 打印机控制电路板已经更换三块，每次更换完成，都能正常工作几天，但是故障再次出现，为了确保维修好仪器，需要再增加借用的备件 | SWO-76320 | No |  | The printer control board has been replaced three pieces, each replacement completed, can work for several days, but the fault reappeared, in order to ensure that the equipment is repaired, the need to add borrowed spare parts
ABL80 FLEX BASIC | 打印机控制电路板已经更换三块，每次更换完成，都能正常工作几天，但是故障再次出现，为了确保维修好仪器，需要再增加借用的备件 | 打印机控制电路板已经更换三块，每次更换完成，都能正常工作几天，但是故障再次出现，为了确保维修好仪器，需要再增加借用的备件 | SWO-76282 | No | The printer control board has been replaced three pieces, each replacement completed, can work for several days, but the fault reappeared, in order to ensure that the equipment is repaired, the need to add borrowed spare parts | The printer control board has been replaced three pieces, each replacement completed, can work for several days, but the fault reappeared, in order to ensure that the equipment is repaired, the need to add borrowed spare parts
ABL80 FLEX BASIC | Weitere Messungen im Vergleich mit Intensivgerät geplant. | Weitere Messungen im Vergleich mit Intensivgerät geplant. | SWO-76844 | No | Further measurements planned compared to intensive equipment. | Further measurements planned compared to intensive equipment.
ABL80 CO-OX | CONFIRMED LEAKAGE WHEN ONSITE. | REMOVED SOLUTION PACK AND CLEANED UP SOLUTION PACK BAY, WASTE VALVE AND PUMP. PUMPED SOLUTION 20 TIMES AND REMOVED SOLUTION PACK, NO LEAKS SEEN. RETURNED SYSTEM TO CUSTOMER FOR ROUTINE TESTING. | D0259D2FFE26F7F5 | No | CONFIRMED LEAKAGE WHEN ONSITE. | REMOVED SOLUTION PACK AND CLEANED UP SOLUTION PACK BAY, WASTE VALVE AND PUMP. PUMPED SOLUTION 20 TIMES AND REMOVED SOLUTION PACK, NO LEAKS SEEN. RETURNED SYSTEM TO CUSTOMER FOR ROUTINE TESTING.
ABL80 FLEX BASIC | 更换新测试卡修复仪器 | 更换新测试卡修复仪器 | SWO-76289 | No | Replace the new test card repair instrument | Replace the new test card repair instrument
ABL80 FLEX BASIC | 更换新测试卡修复仪器 | 更换新测试卡修复仪器 | SWO-76293 | No | Replace the new test card repair instrument | Replace the new test card repair instrument
ABL80 FLEX CO-OX | Gerät kalibriert - ohne Erfolg, mehrfach gespült und Oxitestprogramm durchgeführt - Fehler bei Lampenwerten und Hämolysator - defekt, Gerät abgezogen | erneut 2 Punkt Kal. durchgeführt, keine Veränderung, tHb Kal. versucht, bricht aber tHb Kal ab,da vermutl. Nullpunktkal. nicht erfolgreich | SWO-76626 | No | Device calibrated - without success, repeatedly rinsed and oxitest program performed - error in lamp values and hemolyser - defective, device removed | 2 point Kal again. carried out, no change, tHb Kal. tried, but cancels tHb Kal,there presumed. Zero-point cal. not successful
ABL80 FLEX BASIC | Cleaned sensor cassette with distilled water and removed clot. Then reinstalled SC. Now unit is working fine. | Cleaned sensor cassette with distilled water and removed clot. Then reinstalled SC. Now unit is working fine. | SWO-76721 | No | Cleaned sensor cassette with distilled water and removed clot. Then reinstalled SC. Now unit is working fine. | Cleaned sensor cassette with distilled water and removed clot. Then reinstalled SC. Now unit is working fine.
ABL80 CO-OX | Tried resetting cpu and cable - printer inoperable. | Replaced the printer. | D025BBBDFE234228 | No | Tried resetting cpu and cable - printer inoperable. | Replaced the printer.
ABL90 FLEX | Gerät rückgespült, Nadel (leicht verbogen) sowie Einlassdichtung ersetzt und Absolvierung aller relevanten Serviceprogramme => weiterhin wiederholbare Fehler bei Clot Detection und Pumpenkal. Visueller Fließ-Check OK aber weiterhin Fehler. Alten LP (147 Rest-Aktivitäten) erneut eingesetzt, alle Serviceprogramme OK => Neustart des Analysators, Startphase 1h, Kalibrierung, QC Messungen, Alle OK | Gerät rückgespült, Nadel (leicht verbogen) sowie Einlassdichtung ersetzt und Absolvierung aller relevanten Serviceprogramme => weiterhin wiederholbare Fehler bei Clot Detection und Pumpenkal. Visueller Fließ-Check OK aber weiterhin Fehler. Alten LP (147 Rest-Aktivitäten) erneut eingesetzt, alle Serviceprogramme OK => Neustart des Analysators, Startphase 1h, Kalibrierung, QC Messungen, Alle OK | SWO-76504 | No | Device rinsed back, needle (slightly bent) as well as inlet seal replaced and completion of all relevant service programs => still repeatable errors in Clot Detection and Pump Scale. Visual flow check OK but still errors. Old LP (147 residual activities) re-inserted, all service programs OK => Restart of the analyzer, start phase 1h, calibration, QC measurements, All OK | Device rinsed back, needle (slightly bent) as well as inlet seal replaced and completion of all relevant service programs => still repeatable errors in Clot Detection and Pump Scale. Visual flow check OK but still errors. Old LP (147 residual activities) re-inserted, all service programs OK => Restart of the analyzer, start phase 1h, calibration, QC measurements, All OK | Assisted Replacement | SP | SP | Pump cal error
ABL90 | CONFIRMED ISSUE WHEN ONSITE. CUSTOMER HAS TRIED REPLACING MULTIPLE SPS SCS PER ADVICE OF TECH SUPPORT TO NO AVAIL. | PROCEED TO REPLACE PINCH VALVE AND PARTS IN PM KIT. | D025DE34FE0091F9 | No | CONFIRMED ISSUE WHEN ONSITE. CUSTOMER HAS TRIED REPLACING MULTIPLE SPS SCS PER ADVICE OF TECH SUPPORT TO NO AVAIL. | PROCEED TO REPLACE PINCH VALVE AND PARTS IN PM KIT. | Defect | HW | PM Kit | PM Kit
ABL90 | Verified connection drops in the activity log. Found that the power management was set to allow the computer to turn off this device to save power in the properties of the WiFi adapter. | Deselected the option for the computer to turn off this device to save power. | D026179AFE0E438F | No | Verified connection drops in the activity log. Found that the power management was set to allow the computer to turn off this device to save power in the properties of the WiFi adapter. | Deselected the option for the computer to turn off this device to save power. | Setup | Wifi | Wifi | Wifi
ABL90 FLEX PLUS |  | B | D025B1D1FE15DD80 | No |  |  | Assisted Replacement | unknown | unknown | unknown | PER ?
ABL90 FLEX PLUS |  | B | D025A4B6FE5E3664 | Yes |  |  | Re classified RMED | Logistical upSWO
ABL90 FLEX PLUS | SP und SC erneut getauscht Sensor Interface Board getauscht Analysator kalibriert und überprüft QC Messung durchgeführt Mit Kunden noch mal die Kapillaren besprochen und dabei festgestellt das aktuell Siemens Kapillare verwendet werden mit 130-200 IU ml Heparin Einheiten... | der Anwender ist nicht am Gerät und kann auch keien detailierten Fehlerangaben machen --> Kal-Fehler ??, oder QC-Verletzungen ??; der Anwender wünscht einen Vorort-Service zur dauerhaften Beseitigung der Geräte-Störung; | SWO-76449 | No | SP and SC again swapped Sensor Interface Board replaced Analyzer calibrated and checked QC measurement performed With customers again discussed the capillaries and determined that currently Siemens capillaries are used with 130-200 IU heparin units... | the user is not on the device and can also make keien detailed error information --> Kal error ??, or QC violations ??; the user wants a suburban service for the permanent elimination of the device malfunction; | Defect | HW | SC PCB | SC PCB
ABL90 FLEX | 更换新测试卡修复仪器 | 更换新测试卡修复仪器 | SWO-76532 | No | Replace the new test card repair instrument | Replace the new test card repair instrument | Assisted Replacement | SC | SC | pCO2
AQT90 FLEX | 更换944-212修复了仪器 | 更换944-212修复了仪器 | SWO-76122 | No | Replacement 944-212 repaired instrument | Replacement 944-212 repaired instrument | Assisted Replacement | Consumables | Test Kit | TnI | 124
ABL827 FLEX | System overall is testing ok. The lock up happens once or twice a day or sometimes 2 times a week. | Provided the customer with the service tip to correct the problem without the need to restart the analyzer. Will continue to monitor the problem and should it occur use the service tip procedure to fix the problem. Problem has not came back closing this case until further notice. | D025A553FE48417A | No | System overall is testing ok. The lock up happens once or twice a day or sometimes 2 times a week. | Provided the customer with the service tip to correct the problem without the need to restart the analyzer. Will continue to monitor the problem and should it occur use the service tip procedure to fix the problem. Problem has not came back closing this case until further notice. | Unknown | SW | DMS_PC | Analyzer Freeze
ABL837 FLEX |  | B | D025B7B6FE3BF087 | No |  |  | Adjustment | HW | Oxi module | Lamp unit
ABL80 FLEX BASIC | 更换新测试卡修复仪器 | 更换新测试卡修复仪器 | SWO-76538 | No | Replace the new test card repair instrument | Replace the new test card repair instrument
ABL80 FLEX BASIC | 连续更换两张测试卡后修复 | 连续更换两张测试卡后修复 | SWO-76577 | No | Repair after replacing two test cards in a row | Repair after replacing two test cards in a row
ABL80 FLEX BASIC | 更换了903-334和902-972修复仪器, | 更换了903-334和902-972修复仪器, | SWO-76269 | No | The 903-334 and 902-972 repair instruments have been replaced. | The 903-334 and 902-972 repair instruments have been replaced.
ABL90 FLEX | Sensor Interface Board getauscht und zum Sensorfach justiert neue SC eingesetzt, Temp. im Service jetzt i.O. Startup abgewartet | Sensor Interface Board getauscht und zum Sensorfach justiert neue SC eingesetzt, Temp. im Service jetzt i.O. Startup abgewartet | SWO-76664 | No | Sensor interface board exchanged and adjusted to the sensor compartment used new SC, Temp. in the service now i.O. Startup waited | Sensor interface board exchanged and adjusted to the sensor compartment used new SC, Temp. in the service now i.O. Startup waited | Defect | SC PCB | SC PCB | temp | temp
ABL90 | Tried several SCs' and SPs'. | Replaced Hemolyzer. | D026F15A7431BA00 | No | Tried several SCs' and SPs'. | Replaced Hemolyzer. | Defect | HW | Hemolyzer | FT
ABL90 FLEX | Fehlersuche am Analysator, Kal.- QC- und Aktivitätendateien überprüft, Leckage erkannt Vakuum Fehler, Einlassdichtung mit Halter+Inlet Probe+Anschlussdichtung überprüft->Anschlussdichtung korrekt angebracht, manuelle Spülung, Quetschventilschlauch ok->keine Besserung, Fehler trat nach SC LP Wechsel auf, SC erneuert, dann wurde die Startphase bestanden System-und Funktionskontrolle i.O. | Fehlersuche am Analysator, Kal.- QC- und Aktivitätendateien überprüft, Leckage erkannt Vakuum Fehler, Einlassdichtung mit Halter+Inlet Probe+Anschlussdichtung überprüft->Anschlussdichtung korrekt angebracht, manuelle Spülung, Quetschventilschlauch ok->keine Besserung, Fehler trat nach SC LP Wechsel auf, SC erneuert, dann wurde die Startphase bestanden System-und Funktionskontrolle i.O. | SWO-76676 | No | Troubleshooting on the analyzer, Kal.- QC and activity files checked, leakage detected vacuum fault, inlet seal with holder+Inlet probe checked->connection seal correctly attached, manual flushing, crushing valve hose ok->none Improvement, error occurred after SC LP change, SC renewed, then the start phase passed system and function control i.O. | Troubleshooting on the analyzer, Kal.- QC and activity files checked, leakage detected vacuum fault, inlet seal with holder+Inlet probe checked->connection seal correctly attached, manual flushing, crushing valve hose ok->none Improvement, error occurred after SC LP change, SC renewed, then the start phase passed system and function control i.O. | Assisted Replacement | SC | SC | Wetability
ABL90 FLEX | Sehr oft Fehlermeldung 1352 ( Verdacht auf Gerinsel im Oxi-Modul ) Flüssigkeitsstrecke gespült--> lässt sich schwer spülen--> Hemolyzer und JW-kit erneuert SC erneuert | Sehr oft Fehlermeldung 1352 ( Verdacht auf Gerinsel im Oxi-Modul ) Flüssigkeitsstrecke gespült--> lässt sich schwer spülen--> Hemolyzer und JW-kit erneuert SC erneuert | SWO-76604 | No | Very often error message 1352 ( Suspicion of Gerinsel in the Oxi-Module ) Liquid section flushed--> it is difficult to rinse--> Hemolyzer and JW-kit renewed SC renewed | Very often error message 1352 ( Suspicion of Gerinsel in the Oxi-Module ) Liquid section flushed--> it is difficult to rinse--> Hemolyzer and JW-kit renewed SC renewed | Defect | HW | PM Kit | PM Kit
ABL90 |  | FOLLOWED SIT'S SUGGESTION TO UPGRADE CURRENT SW FROM 3.1 MR6 TO MR7. CUSTOMER LETTER STATES IN THE CORRECTION SECTION Unable to connect to Radiance, AQURE, or HIS LIS downstream systems. Software has been changed to prevent this from occurring. | D026E6AC744191FF | No |  | FOLLOWED SIT'S SUGGESTION TO UPGRADE CURRENT SW FROM 3.1 MR6 TO MR7. CUSTOMER LETTER STATES IN THE CORRECTION SECTION Unable to connect to Radiance, AQURE, or HIS LIS downstream systems. Software has been changed to prevent this from occurring. | Defect | SW | Unable to connect to radiance | Unable to connect to radiance
AQT90 FLEX | 更换847-607修复仪器 | 更换847-607修复仪器 | SWO-76353 | No | Replacement of 847-607 repair instruments | Replacement of 847-607 repair instruments | Defect | HW | Process Module | Cup dryer, Air filter | 17
ABL825 FLEX | 更换打印机910-275一个后修复 | 更换打印机910-275一个后修复 | SWO-76081 | No | Replace the printer 910-275 after one repair | Replace the printer 910-275 after one repair | Defect | HW | DMS_PC | Printer module
ABL825 FLEX | Lac Elektrode aus Kundenbestand eingebaut. Kontakte Justiert..Lac nachgearbeitet Empf 599pA | Lac Elektrode aus Kundenbestand eingebaut. Kontakte Justiert..Lac nachgearbeitet Empf 599pA | SWO-76924 | No | Lac electrode from customer base installed. Contacts Adjusted.. Lac reworked Empf 599pA | Lac electrode from customer base installed. Contacts Adjusted.. Lac reworked Empf 599pA | Assisted Replacement | Electrode | Electrode | Lac
ABL825 FLEX | Oxi Modul geprüft, Ersatz Modul (902-833r1943n009) installiert, Software installiert, Funktionskontrolle ok | Oxi Modul geprüft, Ersatz Modul (902-833r1943n009) installiert, Software installiert, Funktionskontrolle ok | SWO-77032 | No | Oxi module tested, replacement module (902-833r1943n009) installed, software installed, function control ok | Oxi module tested, replacement module (902-833r1943n009) installed, software installed, function control ok | Defect | HW | Oxi module | Oxi module
ABL825 FLEX | Gerät geprüft, pO2 und Lac. Elektrode neu membraniert, Nullpunkt Kalibration durchgeführt, Überdruckkanal gespült, AC Karussell ersetzt, Funktionskontrolle ok | Gerät geprüft, pO2 und Lac. Elektrode neu membraniert, Nullpunkt Kalibration durchgeführt, Überdruckkanal gespült, AC Karussell ersetzt, Funktionskontrolle ok | SWO-76706 | No | device tested, pO2 and Lac. Electrode re-membraneed, zero-point calibration performed, overpressure channel flushed, AC carousel replaced, function control ok | device tested, pO2 and Lac. Electrode re-membraneed, zero-point calibration performed, overpressure channel flushed, AC carousel replaced, function control ok | Adjustment | Electrode | Electrode | pO2
ABL835 FLEX | Baro korr. auf +19 mmHg Kal und QC ok | Baro korr. auf +19 mmHg Kal und QC ok | SWO-77001 | No | Baro korr. on +19 mmHg Kal and QC ok | Baro korr. on +19 mmHg Kal and QC ok | Adjustment | HW | Basic unit | Barometer
ABL825 FLEX | Checked halogen lamp intensity. Lamp intensity reading 4095. | Aligned the halogen lamp. | D026DCF674187232 | No | Checked halogen lamp intensity. Lamp intensity reading 4095. | Aligned the halogen lamp. | Adjustment | HW | Oxi module | Lamp unit
ABL80 FLEX BASIC | 更换测试卡正常。 | 更换测试卡正常。 | SWO-77130 | No | The replacement test card is OK. | The replacement test card is OK.
ABL90 FLEX PLUS |  | B | D03A31E05E626CAF | No |  |  | Adjustment | Auto Inlet | Gasket | Gasket
ABL90 FLEX | 更换了新的试剂包，清洗了进样针和管路修复仪器 | 更换了新的试剂包，清洗了进样针和管路修复仪器 | SWO-77331 | No | New reagent kit replaced, feeding needle and pipe repair instrument | New reagent kit replaced, feeding needle and pipe repair instrument | Cleaned | Fluidic pathway | replaced yearly kit | replaced yearly kit
AQT90 FLEX | 更换636-549修复仪器 | 更换636-549修复仪器 | SWO-76560 | No | Replace 636-549 repair instruments | Replace 636-549 repair instruments | Defect | HW | Aspiration Module | Needle Sensor | 4794
ABL800 BASIC | lactate membrane replace | Lactate membrane replace & run 2-point calibration .calibration pass. | SWO-76774 | No | lactate membrane replace | Lactate membrane replace & run 2-point calibration .calibration pass. | Assisted Replacement | Electrode | Membrane | Lac
AQT90 FLEX |  | B | D027F0BFEA351627 | No |  |  | Defect | HW | Process Module | Motor P4
ABL80 FLEX BASIC | 更换测试卡正常。 | 更换测试卡正常。 | SWO-78168 | No | The replacement test card is OK. | The replacement test card is OK.
ABL80 FLEX BASIC | 更换新测试卡修复仪器 | 更换新测试卡修复仪器 | SWO-76858 | No | Replace the new test card repair instrument | Replace the new test card repair instrument
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-77391 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL80 FLEX BASIC | 更换测试卡正常。 | 更换测试卡正常。 | SWO-77128 | No | The replacement test card is OK. | The replacement test card is OK.
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-77048 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-77050 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL80 FLEX BASIC | 更换测试卡正常。 | 更换测试卡正常。 | SWO-78087 | No | The replacement test card is OK. | The replacement test card is OK.
ABL80 FLEX CO-OX | Gerät wieder aufgestellt | Gerät wieder aufgestellt | SWO-76983 | No | Device repositioned | Device repositioned
ABL80 FLEX BASIC | 更换测试卡正常。 | 更换测试卡正常。 | SWO-77876 | No | The replacement test card is OK. | The replacement test card is OK.
ABL80 FLEX BASIC | 更换测试卡正常。 | 更换测试卡正常。 | SWO-78254 | No | The replacement test card is OK. | The replacement test card is OK.
ABL90 FLEX | Im Service Testprogramme überprüft, selbiges LP nochmals installiert, Kal, QC | Im Service Testprogramme überprüft, selbiges LP nochmals installiert, Kal, QC | SWO-76938 | No | Checked in the service test program, same LP reinstalled, Kal, QC | Checked in the service test program, same LP reinstalled, Kal, QC | User error | Incorrect SP installation | Incorrect SP installation | Incorrect SP installation
ABL90 FLEX | Gerät geprüft. Einlassdichtung ersetzt. Angelieferten Lösungspack installiert. | Gerät geprüft. Einlassdichtung ersetzt. Angelieferten Lösungspack installiert. | SWO-77373 | No | device checked. replaced inlet seal. delivered solution pack. | device checked. replaced inlet seal. delivered solution pack. | Defect | SP | Pump cal error | Pump cal error
ABL90 FLEX | SC überprüft = Verfalldatum überschritten, neue SC geliefert + installiert, tHb-Kal. wegen Fälligkeit durchgef., Kal. + QC außer Lac i.O., Anwender informiert | SC überprüft = Verfalldatum überschritten; neue SC geliefert + installiert; tHb-Kal. wegen Fälligkeit durchgef.; Kal. + QC außer Lac i.O.; Anwender informiert | SWO-76903 | No | SC checked = expiry date exceeded; new SC delivered + installed; tHb-Kal. due to due date; Kal. + QC except Lac i.O.; Users informed | SC checked = expiry date exceeded; new SC delivered + installed; tHb-Kal. due to due date; Kal. + QC except Lac i.O.; Users informed | Assisted Replacement | SC | SC | temp
ABL90 FLEX | ASTM Protokoll mitgeschnitten. Übertragung o.k. | ASTM Protokoll mitgeschnitten. Übertragung o.k. | SWO-76981 | No | ASTM protocol. Transmission ok | ASTM protocol. Transmission ok | Setup | HIS/LIS | His/lis | His/lis
ABL90 FLEX | Kal.- QC- und Aktivitätendateien überprüft, zum Zeitpunkt der unplausiblen Messungen ABE->IQC Fehler bei pCO2 vorhanden, Sachverhalt und Fehlerbehebung mit den Anwendern besprochen+Kontrolle des Ampelstatus, Präanalytik besprochen System-und Funktionskontrolle i.O. | Kal.- QC- und Aktivitätendateien überprüft, zum Zeitpunkt der unplausiblen Messungen ABE->IQC Fehler bei pCO2 vorhanden, Sachverhalt und Fehlerbehebung mit den Anwendern besprochen+Kontrolle des Ampelstatus, Präanalytik besprochen System-und Funktionskontrolle i.O. | SWO-76998 | No | Kal.- QC and activity files checked, at the time of implausible measurements ABE->IQC error in pCO2 present, fact and troubleshooting discussed with the users+control of traffic light status, pre-analysis discussed system and function control i.O. | Kal.- QC and activity files checked, at the time of implausible measurements ABE->IQC error in pCO2 present, fact and troubleshooting discussed with the users+control of traffic light status, pre-analysis discussed system and function control i.O. | No fault found | potential user error | potential user error | potential user error
ABL90 FLEX | Im Controlpaneel deutsche Tastatur als Standart gewählt und auf CFKarte gesichert, Funktionskontrolle | Im Controlpaneel deutsche Tastatur als Standart gewählt und auf CFKarte gesichert, Funktionskontrolle | SWO-77026 | No | In the control panel German keyboard selected as standard and secured on CF card, function control | In the control panel German keyboard selected as standard and secured on CF card, function control | Setup | Keyboard language | Keyboard language | Keyboard language
ABL90 FLEX | Deutsche Tastatur als Standart festgelegt und auf CFKarte gesichert | Deutsche Tastatur als Standart festgelegt und auf CFKarte gesichert | SWO-77028 | No | German keyboard set as standard and saved to CFcard | German keyboard set as standard and saved to CFcard | Setup | Keyboard language | Keyboard language | Keyboard language
ABL90 FLEX | Dr. Tastatur als Standart eingestellt & auf CFKarte gesichert ,Funktionskontrolle | Dr. Tastatur als Standart eingestellt & auf CFKarte gesichert ,Funktionskontrolle | SWO-77027 | No | Dr. Keyboard set as standard & secured to CFcard ,Function control | Dr. Keyboard set as standard & secured to CFcard ,Function control | Setup | Keyboard language | Keyboard language | Keyboard language
ABL90 FLEX PLUS | ABL gespoeld. | ABL gespoeld. | SWO-76900 | No | ABL spoofed. | ABL spoofed. | Cleaned | Fluidic pathway | Oxi module | turbidity
ABL90 FLEX | 更换新测试卡修复仪器 | 更换新测试卡修复仪器 | SWO-76899 | No | Replace the new test card repair instrument | Replace the new test card repair instrument | Assisted Replacement | SC | SC | pCO2 Cl | pCO2 & Cl
ABL90 FLEX | 更换新试剂包修复仪器 | 更换新试剂包修复仪器 | SWO-76875 | No | Replace the new reagent kit repair instrument | Replace the new reagent kit repair instrument | Assisted Replacement | SP | SP | Pump cal error
ABL90 FLEX | 更换新测试卡修复仪器 | 更换新测试卡修复仪器 | SWO-76850 | No | Replace the new test card repair instrument | Replace the new test card repair instrument | Assisted Replacement | sc | sc | Lac QC
ABL837 FLEX | Need to replace printer head | B | SWO-77190 | No | Need to replace printer head |  | Defect | HW | DMS_PC | Printer head
ABL825 FLEX | Se cambia CPU. Se recarga software. El fallo persiste. Se intenta reinstalar windows pero da errores. Equipo no operativo. | B | D02E82AD380A9C54 | No | CPU is changed. Software is reloaded. The fault persists. An attempt is made to reinstall windows but gives errors. Non-operational computer. |  | Unknown | SW | His/lis | ?
ABL825 FLEX |  | B | SWO-76804 | No |  |  | Defect | HW | Basic unit | Barcode scanner
ABL825 FLEX | Halogen Lampe neu justiert. Mehrfach Kalibriert, Service und Anwender Thb Kal durchgeführt. Alle AQC i.O. Lampenwert bleibt stabil. | Halogen Lampe neu justiert. Mehrfach Kalibriert, Service und Anwender Thb Kal durchgeführt. Alle AQC i.O. Lampenwert bleibt stabil. | SWO-76947 | No | Halogen lamp re-adjusted. Multiple calibrated, service and user Thb Kal performed. All AQC i.O. lamp value remains stable. | Halogen lamp re-adjusted. Multiple calibrated, service and user Thb Kal performed. All AQC i.O. lamp value remains stable. | Adjustment | HW | Oxi module | Lamp unit
ABL835 FLEX | troubleshot with customer extensivly on telephone over weekend. On-site found bubble in Lac electrode and low solution in PCO2 electrode. | Cleaned and remembraned both PCO2 and LAC electrodes membranes. Replaced PCO2 electrode under warranty. Performed protein and cleaning procedure on machine. Performed protien remover bomb. Replaced CA membrane. Performed multiple calibrations to bring calibration back to green status. | D02AD65BD62302A1 | No | troubleshot with customer extensivly on telephone over weekend. On-site found bubble in Lac electrode and low solution in PCO2 electrode. | Cleaned and remembraned both PCO2 and LAC electrodes membranes. Replaced PCO2 electrode under warranty. Performed protein and cleaning procedure on machine. Performed protien remover bomb. Replaced CA membrane. Performed multiple calibrations to bring calibration back to green status. | Assisted Replacement | Electrode | Electrode | pCO2
ABL837 FLEX |  | B | D02D6B50C248726F | No |  |  | Unknown | HW | DMS_PC | CPU unit
ABL800 BASIC | cleaned the touch screen, done calibration for touch screen presently working. | cleaned the touch screen, done calibration for touch screen presently working. | SWO-78172 | No | cleaned the touch screen, done calibration for touch screen presently working. | cleaned the touch screen, done calibration for touch screen presently working. | Adjustment | HW | DMS_PC | Touch Screen
ABL815 FLEX | QC Chargen gescannt, Kurzeinweisung, Probennehmer abgegeben, Konfiguration angepasst..... | QC Chargen gescannt, Kurzeinweisung, Probennehmer abgegeben, Konfiguration angepasst..... | SWO-76982 | No | QC batches scanned, short instruction, sampler delivered, configuration adapted..... | QC batches scanned, short instruction, sampler delivered, configuration adapted..... | Unknown | Unknown | Unknown | Unknown
TCM-COMBIM |  | B | D02AD6E1D65F0C3C | No
TCM/Tosca |  | B | D02AF7AAD6148991 | No
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-77382 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL80 FLEX CO-OX | Gerät demontiert und LS-Sensor überprüft => Sensor und Sitz vom Sensor verschmutzt => QC Lösung ausgetreten Reinigung von LS und Sitz vom Sensor Geräteneustart => Flüssigkeitssensor-Kal und Service Test i.O. Uhrzeiteinstellung korrigiert SC wird als verfallen angezeigt => Reinstallierung nicht möglich => LP zu wenig Zyklen => Austausch LP aus Kundenbestand Installierung neuer SC aus Kundenbestand | Gerät demontiert und LS-Sensor überprüft => Sensor und Sitz vom Sensor verschmutzt => QC Lösung ausgetreten Reinigung von LS und Sitz vom Sensor Geräteneustart => Flüssigkeitssensor-Kal und Service Test i.O. Uhrzeiteinstellung korrigiert SC wird als verfallen angezeigt => Reinstallierung nicht möglich => LP zu wenig Zyklen => Austausch LP aus Kundenbestand Installierung neuer SC aus Kundenbestand | SWO-77144 | No | Device disassembled and LS sensor checked => Sensor and seat polluted by sensor => QC solution leaked cleaning of LS and seat from sensor device restart => Liquid sensor cal and service test corrected time setting SC is displayed as forfeited => Reinstall not possible => LP too few cycles => Replacement LP from customer stock Installation of new SC from customer stock | Device disassembled and LS sensor checked => Sensor and seat polluted by sensor => QC solution leaked cleaning of LS and seat from sensor device restart => Liquid sensor cal and service test corrected time setting SC is displayed as forfeited => Reinstall not possible => LP too few cycles => Replacement LP from customer stock Installation of new SC from customer stock
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-77061 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL80 FLEX BASIC | 更换测试卡正常。 | 更换测试卡正常。 | SWO-77049 | No | The replacement test card is OK. | The replacement test card is OK.
ABL90 FLEX | Datensicherung, defekten Touchscreen ersetzt, Startphase, Kal & QK ok | Datensicherung, defekten Touchscreen ersetzt, Startphase, Kal & QK ok | SWO-77152 | No | Data backup, defective touchscreen replaced, start phase, Kal & QK ok | Data backup, defective touchscreen replaced, start phase, Kal & QK ok | Defect | HW | DMS_PC | Touch screen
ABL90 | Examined PC connections. Examined IP settings and comm settings. Found mismatch on IP addresses between analyzer and NetOp client. Instrument needs to be upgraded to MR6. | Rebooted analyzer. Entered correct IP address to match analyzer on NetOp client. Upgraded instrument to MR 6 software. Verified operation of ABL 90. | D02BE28C4C5EC3C3 | No | Examined PC connections. Examined IP settings and comm settings. Found mismatch on IP addresses between analyzer and NetOp client. Instrument needs to be upgraded to MR6. | Rebooted analyzer. Entered correct IP address to match analyzer on NetOp client. Upgraded instrument to MR 6 software. Verified operation of ABL 90. | Setup | IP address NetOp | IP address NetOp | IP address NetOp
ABL90 | Error 1152-Solution pack chip data could not be read. Sensor cassette exceeds life time in analyzer. | Re installed solution pack. Un able to installed cassette, none available. | D02BEB1C4C133CD6 | No | Error 1152-Solution pack chip data could not be read. Sensor cassette exceeds life time in analyzer. | Re installed solution pack. Un able to installed cassette, none available. | No fault found
ABL90 | Confirmed problem. | Replaced the tubing, gasket, probe and connector gasket using a yearly maintenance kit. | D02BDE9E4C46C3AE | No | Confirmed problem. | Replaced the tubing, gasket, probe and connector gasket using a yearly maintenance kit. | Defect | HW | PM Kit | PM Kit
ABL90 FLEX | Gerät überprüft und auf Undichtigkeiten kontrolliert, Fließwege gespült, Sensorkassette (416-57) erneuert, tHb Kal durchgeführt | Gerät überprüft und auf Undichtigkeiten kontrolliert, Fließwege gespült, Sensorkassette (416-57) erneuert, tHb Kal durchgeführt | SWO-77471 | No | Device checked and checked for leaks, flow paths flushed, sensor cassette (416-57) renewed, tHb Kal performed | Device checked and checked for leaks, flow paths flushed, sensor cassette (416-57) renewed, tHb Kal performed | Assisted Replacement | SP | SP | Leak current
ABL90 FLEX PLUS | Die SC war nicht richtig kalibriert, auch die QC's waren nicht o.k., Messungen waren deshalb auch nicht o.k., SC erneuert und QC durchgeführt, | Die SC war nicht richtig kalibriert, auch die QC's waren nicht o.k., Messungen waren deshalb auch nicht o.k., SC erneuert und QC durchgeführt, | SWO-77184 | No | The SC was not properly calibrated, even the QC's were not OK, measurements were therefore not OK, SC renewed and QC performed, | The SC was not properly calibrated, even the QC's were not OK, measurements were therefore not OK, SC renewed and QC performed, | Assisted Replacement | SC | SC | pCO2
ABL90 FLEX | 更换试剂包正常。 | 更换试剂包正常。 | SWO-77052 | No | The replacement kit is OK. | The replacement kit is OK. | Assisted Replacement | SP | SP | Pump cal error
AQT90 FLEX | 重新安装系统后，故障依旧 。更换910-343一个正常 | 重新安装系统后，故障依旧 。更换910-343一个正常 | SWO-77086 | No | After reinstalling the system, the problem remains. Replace 910-343 one normal | After reinstalling the system, the problem remains. Replace 910-343 one normal | Defect | HW | Analyzer Control Module | CPU
ABL835 FLEX | Se sustituye la tarjeta del moddulo oxi, se realiizan test y se chequea con calibración y controles de calidad. | Cambio de la placa de oxi | D02FA821AE280194 | No | The oxy moddulo card is replaced, tested and checked with calibration and quality controls. | Changing the oxidboard | Defect | HW | Oxi module | PCB | refurbished
ABL805 FLEX | checked all lines checked tubes foud leak on inlet nipples checked all valves found air valve with a lot of failurs | replaced the air valve 6 and nipple kit | D02BEA4B4C11CC68 | Yes | checked all lines checked tubes foud leak on inlet nipples checked all valves found air valve with a lot of failurs | replaced the air valve 6 and nipple kit | Re classified RMED | Production analyzer
ABL825 FLEX | Se reinstala todo el sistema operativo y se restauran datos y parámetros del cliente. Al restaurar datos se reproduce el error. Se deja analyzer control y cpu unit originales. Se reinstala de nuevo el sistema operativo y aplicación sin restaurar base de datos. | B | D02E85B238469C58 | No | The entire operating system is reinstalled and client data and parameters are restored. Restoring data reproduces the error. Allows original analyzer control and cpu unit. The operating system and application are reinstalled again without restoring database. |  | Unknown | SW | His/lis | HIS/LIS
ABL837 FLEX |  | B | D02D6CABC23B7270 | No |  |  | Unknown | HW | DMS_PC | CPU unit
ABL825 FLEX | Gaseinheit kontrolliert. Gas-Ventil getauscht. pCO2 Elektrode getauscht. PH- Dichtungsring kontrolliert und Messkamer gereigt. | Gaseinheit kontrolliert. Gas-Ventil getauscht. pCO2 Elektrode getauscht. PH- Dichtungsring kontrolliert und Messkamer gereigt. | SWO-77297 | No | gas unit. Gas valve. pCO2 electrode. PH sealing ring is checked and measuring camer is re-energised. | gas unit. Gas valve. pCO2 electrode. PH sealing ring is checked and measuring camer is re-energised. | Assisted Replacement | Electrode | Electrode | pCO2
ABL825 FLEX | user stated that screen randomly freezes | reset the setup on the analyzer and than reloaded it. | D02BE3304C55591B | No | user stated that screen randomly freezes | reset the setup on the analyzer and than reloaded it. | No fault found | Off/On | Off/On | Off/On
AQT90 FLEX | Kabel vom Bildschirm überprüft Keine Hinweise in den Aktivitäten Auf verdacht die DMS Software neu aufgespielt Systemüberprüfung OK Systemreinigung durchgeführt QC Messung durchgeführt Kunde fotografiert die Fehlermeldung falls diese wieder auftreten sollte und meldet sich dann wieder bei der Hotline | Kabel vom Bildschirm überprüft Keine Hinweise in den Aktivitäten Auf verdacht die DMS Software neu aufgespielt Systemüberprüfung OK Systemreinigung durchgeführt QC Messung durchgeführt Kunde fotografiert die Fehlermeldung falls diese wieder auftreten sollte und meldet sich dann wieder bei der Hotline | SWO-77287 | No | Cable from the screen checked No clues in the activities Suspicion the DMS software replayed system check OK system cleaning performed QC measurement driver photographs the error message in case this should occur again and reports itself then back to the hotline | Cable from the screen checked No clues in the activities Suspicion the DMS software replayed system check OK system cleaning performed QC measurement driver photographs the error message in case this should occur again and reports itself then back to the hotline | Defect | SW re-install | Analyzer Control Module | Display
AQT90 FLEX | Fehlersuche Storage Wheel zerlegt und sauber gemacht, Fehler bleibt bestehen, mechanische Geräusch beim einlegen bis hin zu Fehler und Neustart des Geräts Ersatzteile bestellt Storage PCB 903-127 Erneuerter Kundenbesuch notwendig Systemüberprüfung ok Systemreinigung durchgeführt QC Messung durchgeführt | Fehlersuche Storage Wheel zerlegt und sauber gemacht, Fehler bleibt bestehen, mechanische Geräusch beim einlegen bis hin zu Fehler und Neustart des Geräts Ersatzteile bestellt Storage PCB 903-127 Erneuerter Kundenbesuch notwendig Systemüberprüfung ok Systemreinigung durchgeführt QC Messung durchgeführt | SWO-77109 | No | Troubleshooting Storage Wheel disassembled and cleaned, errors persist, mechanical noise when inserting up to error and restart of the device spare parts ordered Storage PCB 903-127 Renewed customer visit necessary system review ok system cleaning performed QC measurement performed | Troubleshooting Storage Wheel disassembled and cleaned, errors persist, mechanical noise when inserting up to error and restart of the device spare parts ordered Storage PCB 903-127 Renewed customer visit necessary system review ok system cleaning performed QC measurement performed | Defect | HW | Storage Module | PCB, Storage
TCM400 |  | B | SWO-77235 | No
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-77395 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-78506 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL80 FLEX BASIC | 更换902-972一个后修复，903-334未更换。 | 更换902-972一个后修复，903-334未更换。 | SWO-77042 | No | Replace 902-972 one post repair, 903-334 not replaced. | Replace 902-972 one post repair, 903-334 not replaced.
ABL80 FLEX BASIC | 零点校正卡校正机器零点，零点电压偏高，医院重新改造线路，添加接地线正常。 | 零点校正卡校正机器零点，零点电压偏高，医院重新改造线路，添加接地线正常。 | SWO-76849 | No | Zero-point correction card correction machine zero point, zero voltage is high, the hospital remodeled the line, add edding wire normal. | Zero-point correction card correction machine zero point, zero voltage is high, the hospital remodeled the line, add edding wire normal.
ABL90 |  | B | D02D52FCC202DEBB | No |  |  | Assisted Replacement | SP | SP | FT
ABL90 FLEX | Kunden neues LP gebracht | Kunden neues LP gebracht | SWO-77662 | No | Customers brought new LP | Customers brought new LP | Assisted Replacement | SP | SP | pO2 QC
ABL90 FLEX | Software neu aufgespielt ohne Erfolg Datenbank defekt Neue CF Karte eingebaut Software neu installiert Analysator konfiguriert Analysator kalibriert und überprüft Fcuv 100,0% QC Messung durchgeführt Mit EDV gesprochen zwecks USV, ein Angebot für einen Akku wird an WKG gestellt | Software neu aufgespielt ohne Erfolg Datenbank defekt Neue CF Karte eingebaut Software neu installiert Analysator konfiguriert Analysator kalibriert und überprüft Fcuv 100,0% QC Messung durchgeführt Mit EDV gesprochen zwecks USV, ein Angebot für einen Akku wird an WKG gestellt | SWO-77367 | No | Software refiled without success database defective New CF card installed software reinstalled analyzer configured analyzer calibrated and verified Fcuv 100.0% QC measurement performed With COMPUTER spoken for UPS, an offer for a battery is made to WKG | Software refiled without success database defective New CF card installed software reinstalled analyzer configured analyzer calibrated and verified Fcuv 100.0% QC measurement performed With COMPUTER spoken for UPS, an offer for a battery is made to WKG | Defect | HW | CF card | CF card
ABL90 FLEX PLUS |  | B | D02D8085C21ACF9B | No |  |  | Assisted Replacement | SC | SC | pCO2 | PER ?
ABL90 FLEX PLUS |  | B | D02D539AC20328B3 | No |  |  | Assisted Replacement | SC | SC | SC | PER ?
TCM CombiM |  | B | SWO-77330 | No
AQT90 FLEX | 更换903-210一个后修复，905-908未更换。 | 更换903-210一个后修复，905-908未更换。 | SWO-76873 | No | Replace the 903-210 one after repair, 905-908 did not replace. | Replace the 903-210 one after repair, 905-908 did not replace. | Defect | HW | Process Module | PCB, Cup dryer | 168
ABL825 FLEX |  | B | D02D3B49C260E8AF | Yes |  |  | Re classified RMED | Logistical upSWO
ABL827 FLEX | Fehlersuche, div. Ventile überprüft -> i.O., Streßtest Kal. & Spülungen i.O., QC i.O. | Fehlersuche; div. Ventile überprüft -> i.O.; Streßtest Kal. & Spülungen i.O.; QC i.O. | SWO-77519 | No | Troubleshooting; div. Valves checked -> i.O.; Stress Test Kal. & rinses i.O.; QC in the first | Troubleshooting; div. Valves checked -> i.O.; Stress Test Kal. & rinses i.O.; QC in the first | Assisted Replacement | Solution bottle | Solution bottle | Crea
ABL825 FLEX |  | B | D02D746FC21D023D | No |  |  | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
ABL825 FLEX |  | B | D02D6FE5C2630238 | No |  |  | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
ABL800 BASIC | Reinigung, Proteinentfernung und mehrere 1.Kals. durchgeführt, sO2 i.O. Glu. Fehler weiter vorhanden, Membran erneuert, wieder mehrere 1.Kals., Glu. auch i.O. | Reinigung, Proteinentfernung und mehrere 1.Kals. durchgeführt, sO2 i.O. Glu. Fehler weiter vorhanden, Membran erneuert, wieder mehrere 1.Kals., Glu. auch i.O. | SWO-77730 | No | Cleaning, protein removal and several 1.Kals. sO2 i.O. Glu. Error still present, membrane renewed, again several 1.Kals., Glu. also in the | Cleaning, protein removal and several 1.Kals. sO2 i.O. Glu. Error still present, membrane renewed, again several 1.Kals., Glu. also in the | Assisted Replacement | Electrode | Membrane | Glu | Rilibäk
ABL837 FLEX |  | B | D02D6F90C20B7271 | Yes |  |  | Re classified RMED | Install Replacement
ABL835 FLEX | Essayé de réinstaller Windows et mise- à- jour de logiciel en utilisant lecteur CD, CD, clé USB différente. Lors de la mise- à- jour de logiciel, remarqué que ancienne version de sysbase était zéro au lieu d'avoir un chifre. | Réinstallé Windows et effectué mise- à- jour de logiciel version 6.16. Reconfiguré appereil et installé nouveaux lots de contrôles. | D02D605BC232D4AE | No | Tried to reinstall Windows and software update using CD player, CD, different USB stick. When updating software, noticed that old version of sysbase was zero instead of having a chifre. | Reinstalled Windows and carried out software update version 6.16. Reconfigured appereil and installed new batches of controls. | Defect | SW | Win | =AW9871
ABL827 FLEX | neue Ca-Membran mit Spüllsg. vorbehandelt, Ca neu membraniert, Kal. i.O., QC fehlerhaft, neue Ca-Elektrode in gl. Membran getestet -> Fehler, neue Ca-Charge(alt: R0578, neu: R0533) mit alter Elektrode getestet, Ca weiter Fehler, Abbruch d. Arbeiten | B | SWO-77544 | No | new Ca-membrane with rinse sg. pre-treated; Ca re-membraneed; Kal. i.O.; QC erroneous; new Ca electrode in gl. Membrane tested -> error; new Ca-Charge (old: R0578; new: R0533) tested with old electrode; Ca continued error; Demolition d. Work |  | Assisted Replacement | Electrode | Electrode | Ca
ABL820 FLEX | Viewed logs for any possible errors - none. | Replaced the inlet module. | D02E217D383E29EA | No | Viewed logs for any possible errors - none. | Replaced the inlet module. | Defect | HW | Inlet module | Inlet module
ABL825 FLEX | Baro Korrektur + 19 mmHg Kal und QC ok | Baro Korrektur + 19 mmHg Kal und QC ok | SWO-77481 | No | Baro Correction + 19 mmHg Kal and QC ok | Baro Correction + 19 mmHg Kal and QC ok | Adjustment | HW | Basic unit | Barometer | pCO2
ABL825 FLEX | Baro Korrektur +19mmHg bis jetzt kein Erfolg feststellbar. | Baro Korrektur +19mmHg bis jetzt kein Erfolg feststellbar. | SWO-77555 | No | Baro correction +19mmHg so far no success can be detected. | Baro correction +19mmHg so far no success can be detected. | Adjustment | HW | Basic unit | Barometer | pCO2
ABL80 FLEX BASIC | Validierung mit ABL 825 FU Intensiv . Verschiedene Blutprobenmessungen ergeben im Mittel eine Differenz von -10 mmHg PO2 Korrektur an ABL80 -10mmHg PO2 Versatz mit Blutmessung bestätigt. | Weitere Messungen im Vergleich mit Intensivgerät geplant. | SWO-77952 | No | Validation with ABL 825 FU Intensive . Various blood sample measurements show on average a difference of -10 mmHg PO2 correction to ABL80 -10mmHg PO2 offset confirmed with blood measurement. | Further measurements planned compared to intensive equipment.
ABL90 FLEX | Heb wastetubing, pinchvalvetubing en connectorgasket vervangen.Na opstarten 50 minuten verschijnt foutmelding 1343 (niet gelukt vloeistof te pompen0 weer, inhomogeen CAL3 en QC3. Via service menu bloed laten opzuigen. Service -fluid transport-testen nu goed. Ook andere servicetesten OK. Calibraties en QC fout, reagenspack (UB-11) vervangen voor UK-02. Nu problemen verholpen. Met gebruiker een tHb calibratie uitgevoerd. | Heb wastetubing, pinchvalvetubing en connectorgasket vervangen.Na opstarten 50 minuten verschijnt foutmelding 1343 (niet gelukt vloeistof te pompen0 weer, inhomogeen CAL3 en QC3. Via service menu bloed laten opzuigen. Service -fluid transport-testen nu goed. Ook andere servicetesten OK. Calibraties en QC fout, reagenspack (UB-11) vervangen voor UK-02. Nu problemen verholpen. Met gebruiker een tHb calibratie uitgevoerd. | SWO-77686 | No | Have wastetubing, pinchvalvetubing and Connectorgasket replaced. After booting 50 minutes pops up error 1343 (failed to pompen0 liquid again, Inhomogen CAL3 and QC3. Suck up Blood Via service menu. Service-Fluid transport-testing now well. Also test other service OK. Calibrations and QC Error, reagent pack (UB-11) replaced for UK-02. Fixed issues now. With User a tHb calibration performed. | Have wastetubing, pinchvalvetubing and Connectorgasket replaced. After booting 50 minutes pops up error 1343 (failed to pompen0 liquid again, Inhomogen CAL3 and QC3. Suck up Blood Via service menu. Service-Fluid transport-testing now well. Also test other service OK. Calibrations and QC Error, reagent pack (UB-11) replaced for UK-02. Fixed issues now. With User a tHb calibration performed. | Assisted Replacement | SP | SP | Pump cal error
ABL90 | Check power management properties in Control panel. Both WiFi and wired network connection were set to allow windows to shut down connection to save power. | Uncheck powermanagement setting for wired and WiFi connection. | D02E4E95385535F9 | No | Check power management properties in Control panel. Both WiFi and wired network connection were set to allow windows to shut down connection to save power. | Uncheck powermanagement setting for wired and WiFi connection. | Setup | Wifi | power | power
ABL90 FLEX | Gerät hat zum Problem lösen nach SP austauschen verlangt und das wurde nicht gemacht. SP ersetzt alle test, kalibration und QC alle 3 level durchgef. I.O. | Gerät hat zum Problem lösen nach SP austauschen verlangt und das wurde nicht gemacht. SP ersetzt alle test, kalibration und QC alle 3 level durchgef. I.O. | SWO-77862 | No | Device has asked to solve problem after SP exchange and that was not done. SP replaces all test, calibration and QC all 3 levels. Ok. | Device has asked to solve problem after SP exchange and that was not done. SP replaces all test, calibration and QC all 3 levels. Ok. | Assisted Replacement | SP | SP | Oxi mismatch | QC C
ABL90 FLEX | Gerät überprüft und gereinigt, Fließwege durchgspült, Einlaß gesäubert, Lösungspack erneuert, Sensorkassette ausgetauscht | B | SWO-77720 | No | Device checked and cleaned, flow paths flushed through, inlet cleaned, solution pack renewed, sensor cassette replaced |  | Assisted Replacement | SC | SC | pCO2
ABL90 | Pressure test vacuum failing Error 1063, Leak detected error. Analyzer in UIR | Checked probe, gasket and oxi valve tubing. Flushed the analyzer and Replaced Inlet gasket with holder. Ran waste blood samples. | D02E984538065008 | No | Pressure test vacuum failing Error 1063, Leak detected error. Analyzer in UIR | Checked probe, gasket and oxi valve tubing. Flushed the analyzer and Replaced Inlet gasket with holder. Ran waste blood samples. | Defect | HW | Inlet Gasket | & cleaned
ABL90 FLEX | Lac-Pin händisch heruntergezogen und ET bestellt. | Lac-Pin erneuert. | SWO-78295 | No | Lac-Pin manually pulled down and ordered ET. | Lac-Pin renewed. | Adjustment | Contact pin | Lac | Lac
ABL90 FLEX PLUS |  | B | D02E5BCA38004DAC | Yes |  |  | Re classified RMED | Logistical upSWO | pO2 | pO2 | PER ?
ABL90 FLEX | 更换903-165 一个正常。 | 更换903-165 一个正常。 | SWO-77337 | No | Replace 903-165 one normal. | Replace 903-165 one normal. | Assisted Replacement | SP | SP | Pump cal error
ABL90 FLEX | 更换新试剂包后修复 | 更换新试剂包后修复 | SWO-78278 | No | Repair after replacing a new kit | Repair after replacing a new kit | Assisted Replacement | SP | SP | Pump cal error
ABL90 FLEX |  | B | SWO-77628 | No |  |  | Defect | HW | Printer | Printer
AQT90 FLEX |  | B | SWO-77626 | No |  |  | Defect | HW | Power Supply | Power supply
AQT90 FLEX |  | B | SWO-77587 | No |  |  | Defect | HW | Incubator module | PCB, Incubator | 3564
AQT90 FLEX | 更换新测试卡修复仪器 | 更换新测试卡修复仪器 | SWO-77668 | No | Replace the new test card repair instrument | Replace the new test card repair instrument | Defect | Consumables | Test Kit | Cartridge | 5689 (Broken)
AQT90 FLEX |  | B | SWO-77635 | No |  |  | Defect | HW | Process Module | Process Unit sensor
AQT90 FLEX | 更换903-364一个正常。 | 更换903-364一个正常。 | SWO-77354 | No | Replace 903-364 one normal. | Replace 903-364 one normal. | Defect | HW | Process Module | Optical Unit | 131
AQT90 FLEX |  | B | SWO-77620 | No |  |  | Defect | HW | Process Module | Optical Unit | 129
ABL835 FLEX | Proteinentfernung durchgeführt - Messkammer gereinigt pH Dichtring erneuert, Halbjahreskit eingebaut Gerät kalibriert und überprüft | Proteinentfernung durchgeführt - Messkammer gereinigt pH Dichtring erneuert, Halbjahreskit eingebaut Gerät kalibriert und überprüft | SWO-77694 | No | Protein removal performed - measuring chamber cleaned pH sealing ring renewed, half-year kit built-in device calibrated and checked | Protein removal performed - measuring chamber cleaned pH sealing ring renewed, half-year kit built-in device calibrated and checked | Defect | HW | Basic unit | Pump tubing
ABL800 BASIC | Fehlersuche, sporadische Probenfehler, Einlassdichtung und Einlasssonde neu, Ref+Elektrolytelektroden neu membraniert, Neustart, mehrere Probenmessungen ok, Kal. und QC ok System-und Funktionskontrolle i.O. | Fehlersuche, sporadische Probenfehler, Einlassdichtung und Einlasssonde neu, Ref+Elektrolytelektroden neu membraniert, Neustart, mehrere Probenmessungen ok, Kal. und QC ok System-und Funktionskontrolle i.O. | SWO-77714 | No | Troubleshooting, sporadic sample errors, inlet seal and inlet probe new, ref+electrolyte electrodes re-membrane, restart, several sample measurements ok, Kal. and QC ok System and function control in the | Troubleshooting, sporadic sample errors, inlet seal and inlet probe new, ref+electrolyte electrodes re-membrane, restart, several sample measurements ok, Kal. and QC ok System and function control in the | Assisted Replacement | Electrode | Membrane | Reference
ABL835 FLEX | Upgrade suggested by tech support through a techline case. | Performed software upgrade to version 6.16. Performed all necessary calibrations. Loaded new qc ranges for the new software. | D02EB5EE38217406 | Yes | Upgrade suggested by tech support through a techline case. | Performed software upgrade to version 6.16. Performed all necessary calibrations. Loaded new qc ranges for the new software. | Re classified RMED | SWU
ABL835 FLEX | Techline suggested installation of this version to fix PCO2 issues. | Performed software upgrade to version 6.16. Scanned new insert sheets according to the software version. Performed all necessary calibrations. | D02E970D384573FC | Yes | Techline suggested installation of this version to fix PCO2 issues. | Performed software upgrade to version 6.16. Scanned new insert sheets according to the software version. Performed all necessary calibrations. | Re classified RMED | SWU
ABL825 FLEX | Actions suggested by open techline case concerning this problem. | Performed software upgrade to version 6.16. Performed all necessary calibrations. Loaded new updated qc inserts for new software version. | D02EAE5A38607403 | Yes | Actions suggested by open techline case concerning this problem. | Performed software upgrade to version 6.16. Performed all necessary calibrations. Loaded new updated qc inserts for new software version. | Re classified RMED | SWU
ABL800 BASIC | 更换Lac电极膜后修复 | 更换Lac电极膜后修复 | SWO-77812 | No | Repair after replacing Lac electrode film | Repair after replacing Lac electrode film | Assisted Replacement | Electrode | Membrane | Lac
ABL825 FLEX |  | B | D02E670F381870A0 | No |  |  | Defect | HW | Gas module | Pressure Transducer
ABL827 FLEX | neue Elektrodenkappe u. Ref.-Elektrode getestet -> Fehler, Gerät aus, alle Ventile d. Fluidic-Moduls ausgebaut u. gereinigt, Gerät ein -> Fehler, Kal.1 + Reinigngslsg. gew. -> Fehler, Abbruch d. Arbeiten | neue Elektrodenkappe u. Ref.-Elektrode getestet -> Fehler; Gerät aus; alle Ventile d. Fluidic-Moduls ausgebaut u. gereinigt; Gerät ein -> Fehler; Kal.1 + Reinigngslsg. gew. -> Fehler; Abbruch d. Arbeiten | SWO-77775 | No | new electrode cap and ref electrode tested -> error; device off; all valves of the fluidic module removed and cleaned; device a -> error; Kal.1 + Reinigngslsg. Wt. -> Error; Demolition d. Work | new electrode cap and ref electrode tested -> error; device off; all valves of the fluidic module removed and cleaned; device a -> error; Kal.1 + Reinigngslsg. Wt. -> Error; Demolition d. Work | Unknown | Revisit required | Revisit required | Revisit required
ABL835 FLEX | Pco2 controles bas sur niveau 1,2 | installé logiciel calibration et scanner les nouvelle valeurs de Pco2 | D02E6DF53844B78A | No | Pco2 controls low on level 1.2 | installed software calibration and scan the new Pco2 values | Assisted Replacement | Electrode | Electrode | pCO2 | SWU
ABL835 FLEX | Pco2 valeur basse sur niveaux 1, 2 | Install logiciel et scanner nouvelles valeur de pco2 | D02E5253385EB784 | No | Pco2 low value on levels 1, 2 | Install software and scanner new value of pco2 | Assisted Replacement | Electrode | Electrode | pCO2 | SWU
AQT90 FLEX | 更换910-304一个正常。 | 更换910-304一个正常。 | SWO-77077 | No | Replace 910-304 one normal. | Replace 910-304 one normal. | Defect | HW | Storage Module | Barcode reader, Storage wheel | BR2
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-78522 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL80 FLEX BASIC | 更换新测试卡修复仪器 | 更换新测试卡修复仪器 | SWO-77898 | No | Replace the new test card repair instrument | Replace the new test card repair instrument
ABL80 FLEX BASIC | 更换新试剂包安装后解决 | 更换新试剂包安装后解决 | SWO-78242 | No | Post-replacement solution for new reagent kit installation | Post-replacement solution for new reagent kit installation
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-78200 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL80 FLEX BASIC | 连续更换两张测试卡后修复 | 连续更换两张测试卡后修复 | SWO-77856 | No | Repair after replacing two test cards in a row | Repair after replacing two test cards in a row
ABL80 FLEX BASIC | 更换新试剂包后修复 | 更换新试剂包后修复 | SWO-78226 | No | Repair after replacing a new kit | Repair after replacing a new kit
ABL80 FLEX BASIC | 更换两张测试卡修复仪器 | 更换两张测试卡修复仪器 | SWO-77855 | No | Replace two test card repair instruments | Replace two test card repair instruments
ABL80 CO-OX | observed errors in data logs. checked tubings, no blockage observed. | Ran cleaning pack, fluid blockage still occurring. Replaced pumps and tubings for troublesehooting purposes. | D02F9CB3AE498D7D | No | observed errors in data logs. checked tubings, no blockage observed. | Ran cleaning pack, fluid blockage still occurring. Replaced pumps and tubings for troublesehooting purposes.
ABL80 FLEX CO-OX | aktive Parameter und andere Einstellungen überprüft, Systemeinstellungen überprüft, Kassette neu initialisiert - nach wie vor pO2 rot, Gerät abgeschalten und ohne SC gestartet pO2 grün, SC installiert - Neustart und korrekte QC abgewartet, i. O. aktuelle Version RiliBÄK Tool installiert iund erklärt | B | SWO-78261 | No | active parameters and other settings checked, system settings checked, cassette reinitialized - still pO2 red, device switched off and started without SC pO2 green, SC installed - restart and correct QC awaited, i.o. current version RiliBÄK tool installed iand explained
ABL90 FLEX | Im Rahmen der Wartung festgestellt, dass der Bügel des Hämolysators zu hoch saß, Position korrigiert | Fr. Räuber wartet seit gestern Nachmittag auf einen Rückruf und einen entsprechenden Service. Auftrag ist gestern um 14.00 Uhr per Mail rausgegangen, Lesebestätigung liegt auch vor. Hat der Kunde die Info intern nicht weitergegeben? Bitte dringend anrufen. | SWO-78074 | No | During maintenance it was determined that the handle of the hemolyser was sitting too high, position corrected | Ms. Räuber has been waiting for a recall and service since yesterday afternoon. Order went out yesterday at 2 p.m. by e-mail, reading confirmation is also available. Has the customer not passed on the information internally? Please call urgently. | Adjustment | Hemolyzer strap? | Hemolyzer strap? | Hemolyzer strap?
ABL90 |  | B | D033BF3E1004D6D2 | No |  |  | Cleaned | Fluidic pathway | Perhaps training | Perhaps training
ABL90 FLEX PLUS |  | B | D02F992DAE5FDC97 | Yes |  |  | Re classified RMED | Logistical upSWO | pO2 | pO2 | PER ?
ABL90 FLEX PLUS | Schlauch des Quetschventils und Abfallstrecke erneuert. Wet Section PCB ausgetauscht. Erneute installation der Software Version 3.1 MR7, inkonsistente Software. Kalibration des Einlaßmoduls. Austausch des LP UB-05. Initialisierungszeit (Fehler 1031)der Oximetrie aufgrund Austausch Wet Section PCB, 50Min. Kalibrationen durchgeführt. tHb Kal. QC ohne Beanstandung Rücksprache mit Personal und Frau Helberger | Schlauch des Quetschventils und Abfallstrecke erneuert. Wet Section PCB ausgetauscht. Erneute installation der Software Version 3.1 MR7, inkonsistente Software. Kalibration des Einlaßmoduls. Austausch des LP UB-05. Initialisierungszeit (Fehler 1031)der Oximetrie aufgrund Austausch Wet Section PCB, 50Min. Kalibrationen durchgeführt. tHb Kal. QC ohne Beanstandung Rücksprache mit Personal und Frau Helberger | SWO-77962 | No | hose of the crushing valve and waste section. Wet Section PCB. Reinstall the software version 3.1 MR7, inconsistent software. Calibration of the inlet module. Replacement of the LP UB-05. Initialization time (error 1031) of oximetry due to exchange Wet Section PCB, 50min calibrations performed. tHb Kal. QC without complaint Consultation with staff and Ms Helberger | hose of the crushing valve and waste section. Wet Section PCB. Reinstall the software version 3.1 MR7, inconsistent software. Calibration of the inlet module. Replacement of the LP UB-05. Initialization time (error 1031) of oximetry due to exchange Wet Section PCB, 50min calibrations performed. tHb Kal. QC without complaint Consultation with staff and Ms Helberger | Defect | HW | Basic Unit | Wet sec PCB
ABL800 BASIC | replaced the printer now printer working | replaced the printer now printer working | SWO-78181 | No | replaced the printer now printer working | replaced the printer now printer working | Defect | HW | Printer module | Printer module
ABL827 FLEX | Fehlersuche, Gerät aus, alle Meßventile ausgebaut u. gereinigt, Pumpenschläuche u. Einlaßdichtung gew., MK` d. EL MET-Module gereinigt, Gerät ein, div. Progr. im Service durchgef., Kal. + QC i.O. | Fehlersuche; Gerät aus; alle Meßventile ausgebaut u. gereinigt; Pumpenschläuche u. Einlaßdichtung gew.; MK` d. EL MET-Module gereinigt; Gerät ein; div. Progr. im Service durchgef.; Kal. + QC i.O. | SWO-78017 | No | Troubleshooting; device off; all measuring valves removed and cleaned; pump hoses and inlet seal; MK' d. EL MET modules cleaned; device; div. Progr. in the service; Kal. + QC in the first place | Troubleshooting; device off; all measuring valves removed and cleaned; pump hoses and inlet seal; MK' d. EL MET modules cleaned; device; div. Progr. in the service; Kal. + QC in the first place | Defect | HW | Met II Module | Pump tubing
ABL837 FLEX |  | B | D02FCC6AAE19C1F3 | No |  |  | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
ABL825 FLEX | Checked cal log as of visit schudule is in track. Did notice that the unit did perform several 1 pt. cals instead of pH BG cals, after checking activity log pH BG cal had been upgrade to a 1 pt. cal. Also noticed there were several unschuduled calibration. | Explained to customer what occurs during different calibrations. Also will follow-up with customer Monday. | D0312EFB2429BCF8 | No | Checked cal log as of visit schudule is in track. Did notice that the unit did perform several 1 pt. cals instead of pH BG cals, after checking activity log pH BG cal had been upgrade to a 1 pt. cal. Also noticed there were several unschuduled calibration. | Explained to customer what occurs during different calibrations. Also will follow-up with customer Monday. | Unknown | Electrode | Electrode | pH
ABL820 FLEX | user called stating that the interface between analyzer and sunquest not working | found that the roma.ini was missing some information in it from the software upgrade. Put the information back in it which corrected the issue. | D0334F951015D891 | No | user called stating that the interface between analyzer and sunquest not working | found that the roma.ini was missing some information in it from the software upgrade. Put the information back in it which corrected the issue. | Setup | HIS/LIS | His/lis | His/lis
ABL825 FLEX | found printer is not printing the whole print out | replaced printer which corrected the issue. | D02FAF3CAE449A43 | No | found printer is not printing the whole print out | replaced printer which corrected the issue. | Defect | HW | Printer module | Printer module
ABL90 FLEX | 更换新试剂包后修复 | 更换新试剂包后修复 | SWO-79066 | No | Repair after replacing a new kit | Repair after replacing a new kit | Assisted Replacement | SP | SP | Pump cal error
AQT90 FLEX | 重新定标正常。 | 重新定标正常。 | SWO-77836 | No | Recalibration is normal. | Recalibration is normal. | Defect | HW | Process Module | Optical Unit | 129
ABL825 FLEX | 更换了一个Glu电极膜和两个Lac电极膜后修复 | 更换了一个Glu电极膜和两个Lac电极膜后修复 | SWO-78175 | No | Repaired after replacing one Glu electrode film and two Lac electrode membranes | Repaired after replacing one Glu electrode film and two Lac electrode membranes | Assisted Replacement | Electrode | Membrane | Lac
ABL800 BASIC |  | B | SWO-78129 | No |  |  | Assisted Replacement | Electrode | Electrode | pH
ABL80 FLEX BASIC | 连续更换两张测试卡后修复 | 连续更换两张测试卡后修复 | SWO-78509 | No | Repair after replacing two test cards in a row | Repair after replacing two test cards in a row
ABL80 FLEX BASIC | 更换测试卡正常。 | 更换测试卡正常。 | SWO-78236 | No | The replacement test card is OK. | The replacement test card is OK.
ABL80 FLEX BASIC | 更换前面板924-738一个后修复(by andrew) | 更换前面板924-738一个后修复(by andrew) | SWO-78259 | No | Replace the front panel 924-738 with a post-repair (by-andrew) | Replace the front panel 924-738 with a post-repair (by-andrew)
ABL80 FLEX BASIC |  | B | SWO-78284 | No
ABL80 FLEX BASIC |  | B | SWO-78194 | No
ABL80 CO-OX | Verified complaint upon arrival. | Extra parts needed to be ordered. Replaced the CPU, Printer Unit. Printer Control Board and the Printer cable. The printer is now operational again. | D033B60E103BB403 | No | Verified complaint upon arrival. | Extra parts needed to be ordered. Replaced the CPU, Printer Unit. Printer Control Board and the Printer cable. The printer is now operational again.
ABL80 FLEX BASIC | There was micro clot inside Sc so cleaned it with distilled water and kept in fridge for 30 minutes. Then reinstalled SC. Now unit is working fine. | There was micro clot inside Sc so cleaned it with distilled water and kept in fridge for 30 minutes. Then reinstalled SC. Now unit is working fine. | SWO-79408 | No | There was micro clot inside Sc so cleaned it with distilled water and kept in fridge for 30 minutes. Then reinstalled SC. Now unit is working fine. | There was micro clot inside Sc so cleaned it with distilled water and kept in fridge for 30 minutes. Then reinstalled SC. Now unit is working fine.
ABL80 OSM | Ran service test and passed. | Replaced hemolyzer. | D0336AE21019D243 | No | Ran service test and passed. | Replaced hemolyzer.
ABL90 |  | B | D033999D104AF66A | No |  |  | Unknown | Revisit required | Oxi QC | Oxi QC
ABL90 FLEX | Lac-Pin erneuert. | Lac-Pin erneuert. | SWO-78296 | No | Lac-Pin renewed. | Lac-Pin renewed. | Defect | HW | SC contact pin | Lac
ABL90 FLEX | SC gestern ersetzt Spülung seit dem unterbrochen, PCO 2 Empf bei 7% , Oxi Schlauch ersetzt, Einlassnadel verbogen, Einlassnadel aus Kundenbestand ersetzt, SC ersetzt, Kal und QC durchgeführt | SC gestern ersetzt Spülung seit dem unterbrochen, PCO 2 Empf bei 7% , Oxi Schlauch ersetzt, Einlassnadel verbogen, Einlassnadel aus Kundenbestand ersetzt, SC ersetzt, Kal und QC durchgeführt | SWO-78479 | No | SC yesterday replaced flushing since the interrupted, PCO 2 Empf at 7%, Oxi hose replaced, inlet needle bent, inlet needle replaced from customer stock, SC replaced, Kal and QC performed | SC yesterday replaced flushing since the interrupted, PCO 2 Empf at 7%, Oxi hose replaced, inlet needle bent, inlet needle replaced from customer stock, SC replaced, Kal and QC performed | Replaced | Probe | Bend | Bend
ABL90 | Review data logs, confirmed issue by running external QC, duplicated reported issue. Flushed fluidics, ran primer, reran QC, no change. Installed Maintenance kit, tested pinch valve, re-testing external QC, QC out of range. | Replaced SC, system calibratin passed. Internal and external QC passed. No issue observed. | D033D5641009A4B6 | No | Review data logs, confirmed issue by running external QC, duplicated reported issue. Flushed fluidics, ran primer, reran QC, no change. Installed Maintenance kit, tested pinch valve, re-testing external QC, QC out of range. | Replaced SC, system calibratin passed. Internal and external QC passed. No issue observed. | Defect | SC | pO2 QC | pO2 QC
ABL90 FLEX | 连续更换两张测试卡后解决 | 连续更换两张测试卡后解决 | SWO-78694 | No | Resolve after replacing two test cards in a row | Resolve after replacing two test cards in a row | Assisted Replacement | SP | SP | Glu/Lac
ABL90 FLEX | 更换新试剂包修复仪器 | 更换新试剂包修复仪器 | SWO-78505 | No | Replace the new reagent kit repair instrument | Replace the new reagent kit repair instrument | Assisted Replacement | SP | SP | FT | 1169
ABL90 FLEX | 更换新试剂包后修复 | 更换新试剂包后修复 | SWO-79248 | No | Repair after replacing a new kit | Repair after replacing a new kit | Assisted Replacement | SP | SP | Pump cal error
ABL800 BASIC | Seit letztem Einsatz keine entsprechenden Fehlermeldungen mehr im Speicher, Anwender hat ebenfalls keine Fehler mehr beobachtet | CPU Unit ersetzt, Kira deinstallieren, nach Sicherung einspielen startet die Kira nicht mehr, Kira erneut installiert, Konfiguration per Hand eingestellt, Kal THB Kal und QC durchgeführt | SWO-78307 | No | Since last use no corresponding error messages in memory, user has also not observed any errors | CPU unit replaced, Kira uninstall, after backup play in the Kira no longer starts, Kira reinstalled, configuration set by hand, Kal THB Kal and QC performed | Defect | HW | DMS_PC | CPU unit
ABL827 FLEX | FlexQ-Modul kontrolliert und justiert, Einlassnadel ersetzt, Probelauf ok, Messungen mit Blut ok | FlexQ-Modul kontrolliert und justiert, Einlassnadel ersetzt, Probelauf ok, Messungen mit Blut ok | SWO-78333 | No | FlexQ module controls and adjusts, inlet needle replaced, test run ok, measurements with blood ok | FlexQ module controls and adjusts, inlet needle replaced, test run ok, measurements with blood ok | Defect | HW | Inlet module | Probe bend | Einlassnadel verbogen, FLEXQ kontroliiert und justiert
ABL800 BASIC | 更换Lac电极膜后修复 | 更换Lac电极膜后修复 | SWO-79266 | No | Repair after replacing Lac electrode film | Repair after replacing Lac electrode film | Assisted Replacement | Electrode | Membrane | Lac
ABL827 FLEX | Modulpumpenschläuche ersetzt, alle Messmodulventile demontiert sowie gereinigt und wieder montiert, Probelauf, Na- & K-Elektrode membraniert, Kal ok, QK S7645 ok ( jedoch niedrig) | Modulpumpenschläuche ersetzt, alle Messmodulventile demontiert sowie gereinigt und wieder montiert, Probelauf, Na- & K-Elektrode membraniert, Kal ok, QK S7645 ok ( jedoch niedrig) | SWO-78332 | No | Module pump hoses replaced, all measuring module valves dismantled and cleaned and reassembled, test run, Na& K electrode membraneed, Kal ok, QK S7645 ok (but low) | Module pump hoses replaced, all measuring module valves dismantled and cleaned and reassembled, test run, Na& K electrode membraneed, Kal ok, QK S7645 ok (but low) | Defect | HW | El/Met Module | Pump tubing
ABL825 FLEX | Checked data logs and saw errors Instrument ran out of Cal 1 and none on site Went to St Joe's and got 2 bottles | Installed a Cal 1 bottle Reboot analyzer to reset Neon lamp Test in service-back in range Ca++ working fine after new Cal 1 put on | D0339E09102722B7 | No | Checked data logs and saw errors Instrument ran out of Cal 1 and none on site Went to St Joe's and got 2 bottles | Installed a Cal 1 bottle Reboot analyzer to reset Neon lamp Test in service-back in range Ca++ working fine after new Cal 1 put on | User Error | Consumables | Solution bottle | Solutions bottle
AQT90 FLEX |  | B | D03376D4100DE939 | No |  |  | Defect | SW | Analyzer Control Module | Off/On
AQT90 FLEX | Storage PCB getauscht alte S N R1002N0449 Systemüberprüfung OK QC Messung durchgeführt | Fehlersuche Storage Wheel zerlegt und sauber gemacht, Fehler bleibt bestehen, mechanische Geräusch beim einlegen bis hin zu Fehler und Neustart des Geräts Ersatzteile bestellt Storage PCB 903-127 Erneuerter Kundenbesuch notwendig Systemüberprüfung ok Systemreinigung durchgeführt QC Messung durchgeführt | SWO-78310 | No | Storage PCB swapped old S N R1002N0449 system check OK QC measurement performed | Troubleshooting Storage Wheel disassembled and cleaned, errors persist, mechanical noise when inserting up to error and restart of the device spare parts ordered Storage PCB 903-127 Renewed customer visit necessary system review ok system cleaning performed QC measurement performed | Revisit required | HW | Storage Module | PCB, Storage
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-78706 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-79221 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL80 FLEX BASIC | 更换新测试卡修复仪器 | 更换新测试卡修复仪器 | SWO-78497 | No | Replace the new test card repair instrument | Replace the new test card repair instrument
ABL80 OSM | Susana is concerned because when she runs external qc, she has to select the qc level she is running. | Explained to Susana that in order for the instruments to auto-detect the qc level, all the parameters in the analyzer have to be 'Active' so that the analyzer can check that all is ok and calculate for the correct qc level automatically. For all the instruments, I activated all the parameters and modified her patient report only to show ctHb and, FO2Hb as she requires. | D034F0E48644EB57 | No | Susana is concerned because when she runs external qc, she has to select the qc level she is running. | Explained to Susana that in order for the instruments to auto-detect the qc level, all the parameters in the analyzer have to be 'Active' so that the analyzer can check that all is ok and calculate for the correct qc level automatically. For all the instruments, I activated all the parameters and modified her patient report only to show ctHb and, FO2Hb as she requires.
ABL80 FLEX BASIC |  | B | SWO-78496 | No
ABL80 CO-OX | When they scan in the patient accession number it is putting a digit at the end of the ID. This is causing the analyzer not to query the demographics. Jim Halbleib sent me the barcodes to program the barcode reader. I worked with Jim and tech support and could not get rid of the extra digit. When they enter the accession manually it does pull overt eh demographics. We have a call in to Sendx for help. | B | D034C543864E7368 | No | When they scan in the patient accession number it is putting a digit at the end of the ID. This is causing the analyzer not to query the demographics. Jim Halbleib sent me the barcodes to program the barcode reader. I worked with Jim and tech support and could not get rid of the extra digit. When they enter the accession manually it does pull overt eh demographics. We have a call in to Sendx for help.
ABL90 FLEX | Bij aankomst zie dat het RP niet correct is geinstalleerd. Zodoende flush procedure heb uitgevoerd, inletgasket, connector en oxivalve tubing heb vervangen. Het RP opnieuw heb geinstalleerd, na 50 min. te lage PCO2 gevoeligheid. Andere SC lot. nr. R0266 heb geinstalleerd, na 50 min. opstarten, OK. Logs van defecte SC lot. R0725N176 en RP lot. UK-03 heb gekopieerd voor CCN. | Bij aankomst zie dat het RP niet correct is geinstalleerd. Zodoende flush procedure heb uitgevoerd, inletgasket, connector en oxivalve tubing heb vervangen. Het RP opnieuw heb geinstalleerd, na 50 min. te lage PCO2 gevoeligheid. Andere SC lot. nr. R0266 heb geinstalleerd, na 50 min. opstarten, OK. Logs van defecte SC lot. R0725N176 en RP lot. UK-03 heb gekopieerd voor CCN. | SWO-78610 | No | On arrival see that the RP is not installed correctly. So flush procedure, Inletgasket, connector and oxivalve tubing have been replaced. Re-installed the RP, after 50 min. too low PCO2 sensitivity. Other SC lot. No. R0266 have installed, after 50 min. booting, OK. Logs of defective SC lot. R0725N176 and RP lot. UK-03 have copied for CCN. | On arrival see that the RP is not installed correctly. So flush procedure, Inletgasket, connector and oxivalve tubing have been replaced. Re-installed the RP, after 50 min. too low PCO2 sensitivity. Other SC lot. No. R0266 have installed, after 50 min. booting, OK. Logs of defective SC lot. R0725N176 and RP lot. UK-03 have copied for CCN. | User error | SP Installation | SP Installation | SP Installation
ABL90 | Confirmed problems. | Replaced the inlet gasket. | D034ACB7862594F9 | No | Confirmed problems. | Replaced the inlet gasket. | Defect | HW | Inlet Gasket | Inlet gasket
ABL90 |  | B | D034AF7B860C04FC | No |  |  | Defect | HW | DMS_PC | Barcode Reader
ABL90 | Customer using terminal servers and NW7 bar code labels | configured LIS comunication for serial connection. Added NW7 code in INI file. | D034F3E6862E37E4 | No | Customer using terminal servers and NW7 bar code labels | configured LIS comunication for serial connection. Added NW7 code in INI file. | Setup | HIS/LIS | DMS_PC | Barcode Reader
ABL90 FLEX | új Tűtartó csere, Kalibrációk, QC mérések | B | SWO-78600 | No | New needle holder replacement, calibrations, QC measurements |  | Defect | HW | Inlet Gasket | Inlet gasket
ABL90 FLEX | 更换新试剂包后修复 | 更换新试剂包后修复 | SWO-78701 | No | Repair after replacing a new kit | Repair after replacing a new kit | Assisted Replacement | SP | SP | Pump cal error
ABL90 FLEX |  | B | SWO-78528 | No |  |  | Defect | HW | Power Supply | Power Supply
ABL90 FLEX | 更换新的测试卡和试剂包后修复 | 更换新的测试卡和试剂包后修复 | SWO-78703 | No | Repair after replacing new test cards and reagent kits | Repair after replacing new test cards and reagent kits
ABL90 | Verificação dos registros, remoção de coágulo do circuito, subistuição da entrada, estabilização e passagem dos QCs. | B | D0353E3E864399CA | No | Verification of records, removal of clot from the circuit, ascent of the entry, stabilization and passage of the QCs. |  | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
ABL825 FLEX | Luftansaugstutzen sehr stark auskristallisiert - gereinigt und gespült, LS Justage und Pumpenkal, pH und Refmesskammer gereinigt, Elektroden gereinigt, mehrere Kal, FlexQ Modul funktionier nicht - Modulsensoren kalibriert, Positionen eingestellt, QC | Luftansaugstutzen sehr stark auskristallisiert - gereinigt und gespült, LS Justage und Pumpenkal, pH und Refmesskammer gereinigt, Elektroden gereinigt, mehrere Kal, FlexQ Modul funktionier nicht - Modulsensoren kalibriert, Positionen eingestellt, QC | SWO-78717 | No | Air intake nozzle very highly crystallized - cleaned and rinsed, LS adjustment and pump scale cleaned, pH and refmeter, electrodes cleaned, multiple cal, FlexQ module not functional - module sensors calibrated, positions set, QC | Air intake nozzle very highly crystallized - cleaned and rinsed, LS adjustment and pump scale cleaned, pH and refmeter, electrodes cleaned, multiple cal, FlexQ module not functional - module sensors calibrated, positions set, QC | Cleaned | Fluidic pathway | Fluidic manifold | Air vent
ABL825 FLEX | Demo Option Key erneut Installiert. | Demo Option Key erneut Installiert. | SWO-78675 | No | Demo Option Key Reinstalled. | Demo Option Key Reinstalled. | Setup | Demo key exp | Demo Key Exp | Demo key expired
ABL825 FLEX | PH- Dichtungsring ausgetauscht. AC Pcb ausgetauscht. Au SW erneut installiert. Das Gerät müsste in Werkstatt überprüft werden | PH- Dichtungsring ausgetauscht. AC Pcb ausgetauscht. Au SW erneut installiert. Das Gerät müsste in Werkstatt überprüft werden | SWO-78719 | No | PH sealing ring. AC Pcb. Au SW reinstalled. The device would have to be checked in the workshop | PH sealing ring. AC Pcb. Au SW reinstalled. The device would have to be checked in the workshop | Unknown | Revisit required | Revisit required | Revisit required
ABL815 FLEX | Kein Fehler feststellbar Baro + 15mmHg | Kein Fehler feststellbar Baro + 15mmHg | SWO-78671 | No | No error detectable Baro + 15mmHg | No error detectable Baro + 15mmHg | Assisted Replacement | Electrode | Electrode | pCO2 | Baro adj
ABL815 FLEX | Baro Korrektur + 15 mmHg | Baro Korrektur + 15 mmHg | SWO-78672 | No | Baro Correction + 15 mmHg | Baro Correction + 15 mmHg | Assisted Replacement | Electrode | Electrode | pCO2
AQT90 FLEX | Verklemmte Kalibrationskassette TnI (Lot: 11242) entfernt TnI Kalibrierung QC-Messungen TnI und D-Dimer mit jeweils LQC Level 1 | Verklemmte Kalibrationskassette TnI (Lot: 11242) entfernt TnI Kalibrierung QC-Messungen TnI und D-Dimer mit jeweils LQC Level 1 | SWO-78668 | No | Clamped calibration cassette TnI (lot: 11242) removes TnI calibration QC measurements TnI and D-Dimer with LQC Level 1 each | Clamped calibration cassette TnI (lot: 11242) removes TnI calibration QC measurements TnI and D-Dimer with LQC Level 1 each | User Error | Consumables | Test Kit | Cartridge, TnI
ABL80 FLEX BASIC | 连续更换两张测试卡后修复 | 连续更换两张测试卡后修复 | SWO-79021 | No | Repair after replacing two test cards in a row | Repair after replacing two test cards in a row
ABL80 FLEX CO-OX OSM | Neuen Hemolyzer eingebaut. | Neuen Hemolyzer eingebaut. | SWO-79303 | No | New Hemolyzer installed. | New Hemolyzer installed.
ABL80 OSM | Ran blood sample to see if I can reproduce error and I did. | I replaced the spectrometer. | D035E3E5FC4AA498 | No | Ran blood sample to see if I can reproduce error and I did. | I replaced the spectrometer.
ABL90 FLEX | Geräteüberprüfung SW nochmals installiert ABL zur Reparatur in GS | Fehlersuche Windows und ABL SW neu installiert (3.1MR6) Geräteneustart Einstellung Konfiguration und Anbindung | SWO-78889 | No | Device inspection SW reinstalled ABL for repair in GS | Troubleshooting Windows and ABL SW reinstalled (3.1MR6) Device restart setting configuration and connection | Defect | SW | MR6 | MR6
ABL90 | Unit failing Pump cal., liquid sensor adjust, and pressure test during fluid transport programs. | Checked for leaks. Replaced valve tubing, checked operation of pinch valve. Re installed solution pack. | D036097CFC5E95BF | No | Unit failing Pump cal., liquid sensor adjust, and pressure test during fluid transport programs. | Checked for leaks. Replaced valve tubing, checked operation of pinch valve. Re installed solution pack. | Defect | HW | Oxi module | Pinch Valve tube
ABL90 FLEX | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-78947 | No | Repair after replacing a new test card | Repair after replacing a new test card | Assisted Replacement | SC | SC | pH
ABL90 FLEX | 更换试剂包正常。 | 更换试剂包正常。 | SWO-78975 | No | The replacement kit is OK. | The replacement kit is OK. | Assisted Replacement | SP | SP | Pump cal error
ABL90 | Verificação dos registros, verificação da calibração da bomba a zero, remoção de coágulo, testes OK, estabilização e passagem dos QCs, realização de uma gasimetria. | B | D035F1B1FC5399CE | No | Checking the records, checking the calibration of the pump to zero, removal of clot, tests OK, stabilization and passage of the QCs, performing a gasimetry. |  | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
ABL835 FLEX |  | B | D0361969FC4CC1D4 | Yes |  |  | Re classified RMED | PM
ABL825 FLEX | 更换LAC电极膜正常。 | 更换LAC电极膜正常。 | SWO-79030 | No | Replacing the LAC electrode film is normal. | Replacing the LAC electrode film is normal. | Assisted Replacement | Electrode | Membrane | Lac
ABL825 FLEX | Ref. Elektrode Lot: 17-12 ersetzt durch Lot: 9707 | Ref. Elektrode Lot: 17-12 ersetzt durch Lot: 9707. Kal. und QCs i.O. | SWO-78838 | No | Ref. Electrode Lot: 17-12 Replaced by Lot: 9707 | Ref. Electrode Lot: 17-12 Replaced by Lot: 9707. Kal. and QCs i.O. | Assisted Replacement | Electrode | Membrane | Reference
ABL825 FLEX | Ref. Elektrode Lot: 17-12 ersetzt durch Lot: 9707. Kal. und QCs i.O. | Ref. Elektrode Lot: 17-12 ersetzt durch Lot: 9707. Kal. und QCs i.O. | SWO-78839 | Yes | Ref. Electrode Lot: 17-12 Replaced by Lot: 9707. Kal. and QCs i.O. | Ref. Electrode Lot: 17-12 Replaced by Lot: 9707. Kal. and QCs i.O. | Re classified RMED | dublicate
ABL837 FLEX | Customer had been keeping unit in Standby due to lack of consumables. Order not being filled. pO2 failing calibration with Zero Current. | Ordered in Electrode. | D036158FFC205A2C | No | Customer had been keeping unit in Standby due to lack of consumables. Order not being filled. pO2 failing calibration with Zero Current. | Ordered in Electrode. | User Error | Standby | Standby | Standby
AQT90 FLEX | Scanner ausgebaut gereinigt und justiert, keine Verbesserung. Scanner bestellt. | Der Kunde bittet um einen Service und Rückruf. | SWO-78978 | No | Scanner removed cleaned and adjusted, no improvement. Scanner ordered. | The customer asks for a service and callback. | Adjustment | HW | Storage Module | Barcode reader, Storage wheel
AQT90 FLEX | Scannerschutz vom Staub befreit und Scanner justiert | Scannerschutz vom Staub befreit und Scanner justiert | SWO-78913 | No | Scanner protection freed from dust and scanner adjusted | Scanner protection freed from dust and scanner adjusted | Adjustment | HW | Storage Module | Barcode reader, Storage wheel
AQT90 FLEX |  | B | D03648D6FC59FAB4 | No |  |  | Not Cat | Not Cat | Not Cat | Not Cat | 5784
ABL80 CO-OX |  | B | D0376C3B724DF0E2 | No
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-79025 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL80 FLEX BASIC | 更换新测试卡修复仪器 | 更换新测试卡修复仪器 | SWO-79211 | No | Replace the new test card repair instrument | Replace the new test card repair instrument
ABL80 FLEX BASIC | K and Na are still red. Sc need replacement. Replace the SC with new one. After replace the SC unit is working properly. | K and Na are still red. Sc need replacement. Replace the SC with new one. After replace the SC unit is working properly. | SWO-79416 | No | K and Na are still red. Sc need replacement. Replace the SC with new one. After replace the SC unit is working properly. | K and Na are still red. Sc need replacement. Replace the SC with new one. After replace the SC unit is working properly.
ABL80 FLEX BASIC | Check the instrument found calibration was ok, found problem with solution pack, suggested customer to replace the solution pack. Conducted preanalytical training for all the Nurses. | Check the instrument found calibration was ok, found problem with solution pack, suggested customer to replace the solution pack. Conducted preanalytical training for all the Nurses. | SWO-79441 | No | Check the instrument found calibration was ok, found problem with solution pack, suggested customer to replace the solution pack. Conducted preanalytical training for all the Nurses. | Check the instrument found calibration was ok, found problem with solution pack, suggested customer to replace the solution pack. Conducted preanalytical training for all the Nurses.
ABL80 OSM | No investigation needed due to parts already replaced . | Replaced the CO-OX control board. | D03727AA721A519B | No | No investigation needed due to parts already replaced . | Replaced the CO-OX control board.
ABL90 |  | B | D037475A7200EC3D | No |  |  | Defect | HW | Spectrometer | Spectrometer
ABL90 FLEX | Jahreswartung durchgeführt - siehe entsprechenden SWO | Jahreswartung durchgeführt - siehe entsprechenden SWO | SWO-79154 | No | Annual maintenance carried out - see corresponding SWO | Annual maintenance carried out - see corresponding SWO | Defect | HW | PM Kit | PM Kit
ABL90 FLEX | ABL auf ein anderen Platz gestellt und an eine Einzelsteckdose angeschlossen. | ABL auf ein anderen Platz gestellt und an eine Einzelsteckdose angeschlossen. | SWO-79096 | No | ABL is placed in a different place and connected to a single socket. | ABL is placed in a different place and connected to a single socket. | User error | Analyzer switched off | Analyzer switched off | Analyzer switched off
AQT90 FLEX |  | B | SWO-78958 | No |  |  | Defect | HW | Process Module | Optical Unit | 124
ABL80 FLEX CO-OX | 清掃、センサールアーチューブ取付、キャル確認 | 清掃、センサールアーチューブ取付、キャル確認 | SWO-79032 | No | Cleaning, sensor lure tube mounting, care check | Cleaning, sensor lure tube mounting, care check
ABL800 FLEX | Cleaned and removed the clot in the pH measuring chamber. Replaced pH electrode with new as a proactive preventive action. | Cleaned and removed the clot in the pH measuring chamber. Replaced pH electrode with new as a proactive preventive action. | SWO-79378 | No | Cleaned and removed the clot in the pH measuring chamber. Replaced pH electrode with new as a proactive preventive action. | Cleaned and removed the clot in the pH measuring chamber. Replaced pH electrode with new as a proactive preventive action. | Cleaned | Fluidic pathway | pH/BG Module | Measuring chamber
ABL817 FLEX | Ventilmembran von V2 (pH BG Modul) erneuert | Ventilmembran von V2 (pH BG Modul) erneuert | SWO-79170 | No | Valve membrane of V2 (pH BG module) renewed | Valve membrane of V2 (pH BG module) renewed | Defect | HW | pH/BG Module | Valve, air
ABL825 FLEX | Allgemeine Kontrolle ok. Hemolyzer ersetzt ok. Alle Einstellungen vorgenommen ok. System i.o. Kal. + QC + Thb i.o. | B | SWO-79132 | No | General control ok. Hemolyzer replaces ok. All settings made ok. System i.o. Kal. + QC + Thb i.o. |  | Defect | HW | Oxi module | Hemolyzer
ABL835 FLEX | Analysator teilweise demontiert - Elektrodendeckel Lactat und Glucose demontiert - Masseblech justiert (Lactat verbogen) - Deckel (Connector) Kontakte und Elektrodenkontakte gebürstet (Glasfaserstift) - Rückmontage EL MET Modul - im Service (Nullstromkalibrierung) Kontakte überprüft - Nullstromkalibrierung - Proteinentfernung und Reinigung durchgeführt - bei Probemessung Lactat instabil - Überprüfung Messventile - Erneuerung Airventil EL MET Modul (mehrere Fehlermeldungen im Service) - Funktionsüberprüfung etc. | Analysator teilweise demontiert - Elektrodendeckel Lactat und Glucose demontiert - Masseblech justiert (Lactat verbogen) - Deckel (Connector) Kontakte und Elektrodenkontakte gebürstet (Glasfaserstift) - Rückmontage EL MET Modul - im Service (Nullstromkalibrierung) Kontakte überprüft - Nullstromkalibrierung - Proteinentfernung und Reinigung durchgeführt - bei Probemessung Lactat instabil - Überprüfung Messventile - Erneuerung Airventil EL MET Modul (mehrere Fehlermeldungen im Service) - Funktionsüberprüfung etc. | SWO-79137 | No | Analyzer partially dismantled - electrode lid lactate and glucose disassembled - ground plate adjusted (lactate bent) - lid (connector) contacts and electrode contacts brushed (glass fiber pin) - back-mount EL MET module - in service (zero current calibration) Contacts checked - zero current calibration - protein removal and cleaning performed - in case of sample measurement lactate unstable - inspection measuring valves - renewal air valve EL MET module (several error messages in service) - function check etc. | Analyzer partially dismantled - electrode lid lactate and glucose disassembled - ground plate adjusted (lactate bent) - lid (connector) contacts and electrode contacts brushed (glass fiber pin) - back-mount EL MET module - in service (zero current calibration) Contacts checked - zero current calibration - protein removal and cleaning performed - in case of sample measurement lactate unstable - inspection measuring valves - renewal air valve EL MET module (several error messages in service) - function check etc. | Defect | HW | El/Met Module | Valve, air
ABL837 FLEX | Entlüftungsanschluss gespült, LS-Justagen in Ordnung. | Entlüftungsanschluss gespült, LS-Justagen in Ordnung. | SWO-79084 | No | Vent connection flushed, LS-justagen sit. | Vent connection flushed, LS-justagen sit. | Cleaned | Fluidic pathway | Fluidic manifold | Air vent
ABL837 FLEX | pO2 electrode at fault. several failed cals with Zero Current error. | Replace Electrode. | D037301172427CC7 | No | pO2 electrode at fault. several failed cals with Zero Current error. | Replace Electrode. | Assisted Replacement | Electrode | Electrode | pO2
ABL825 FLEX | Errors 0210 and 0579. tHb Cal has 634PA. Calibrations that fails have 0251, 0579 and 0376 errors. Lamp unit Value is at 3.846, too high. | Perform Lamp unit calibration. Perform Complete Oxi Test. All ok. tHb on cals read 556.40PA. | D03712E6720F860E | No | Errors 0210 and 0579. tHb Cal has 634PA. Calibrations that fails have 0251, 0579 and 0376 errors. Lamp unit Value is at 3.846, too high. | Perform Lamp unit calibration. Perform Complete Oxi Test. All ok. tHb on cals read 556.40PA. | Adjustment | HW | Oxi module | Lamp unit
ABL80 FLEX CO-OX | Sensorkassette aus Kundenbestand erneuert. Kalibrationen und Qualitätskontrollen in Ordnung. | Sensorkassette aus Kundenbestand erneuert. Kalibrationen und Qualitätskontrollen in Ordnung. | SWO-79412 | No | Sensor cassette from customer base renewed. calibrations and quality controls in order. | Sensor cassette from customer base renewed. calibrations and quality controls in order.
ABL80 FLEX BASIC | 更换新测试卡修复仪器 | 更换新测试卡修复仪器 | SWO-79273 | No | Replace the new test card repair instrument | Replace the new test card repair instrument
ABL80 FLEX BASIC | 更换新测试卡修复仪器 | 更换新测试卡修复仪器 | SWO-79267 | No | Replace the new test card repair instrument | Replace the new test card repair instrument
ABL80 FLEX BASIC | Reset solution pack and reinstalled again. Removed SC and then kept in fridge.. Reinstalled again . Unit is working fine. | Reset solution pack and reinstalled again. Removed SC and then kept in fridge.. Reinstalled again . Unit is working fine. | SWO-79410 | No | Reset solution pack and reinstalled again. Removed SC and then kept in fridge.. Reinstalled again . Unit is working fine. | Reset solution pack and reinstalled again. Removed SC and then kept in fridge.. Reinstalled again . Unit is working fine.
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-79544 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL90 FLEX | Kommunikation überprüft: Ping an 10.135.32.13 iO., Aktivitäten-Datei überprüft: ABL verliert und verbindet sich ca. alle 2 Std. Nachfrage bei IT (Herr Winterfeldt): keine Kommunikationseinschränkung bei Verbindung bzw. Firewall Netzwerkkarte konfiguriert: 100Mbit Half-Duplex, kein Stromsparmodus SC90 installiert (Verfallsdatum erreicht),: Kal-Fehler bei Ca, pCO2-QC-Fehler, neue SC eingesetzt, | B | SWO-79323 | No | Communication checked: Ping to 10.135.32.13 iO., Activity file checked: ABL loses and connects approximately every 2 hours demand from IT (Mr. Winterfeldt): no communication restriction configured on connection or firewall network card: 100Mbit Half duplex, no power saving mode SC90 installed (expiration date reached),: Kal error at Ca, pCO2 QC error, new SC inserted, |  | Setup | His/lis
ABL90 FLEX | Stenose in Hemolysator entfernt. Mehrfaches Manuelles Spülen erfolgreich. Hemol.wieder eingebaut. Kunden Spülversion mit Service Kit erklärt. Techn. Einweisung. | Stenose in Hemolysator entfernt. Mehrfaches Manuelles Spülen erfolgreich. Hemol.wieder eingebaut. Kunden Spülversion mit Service Kit erklärt. Techn. Einweisung. | SWO-79337 | No | Stenosis in Hemolysator. Multiple manual rinsing successful. Hemol.re-installed. Customer rinse version explained with service kit. Techn. Instruction. | Stenosis in Hemolysator. Multiple manual rinsing successful. Hemol.re-installed. Customer rinse version explained with service kit. Techn. Instruction. | Cleaned | Fluidic pathway | Hemolyzer | Hemolyzer
ABL90 FLEX | Messbereichsgrenzen angepasst. Probemessung erfolgreich. Präanalytik Einweisung Personal. | Messbereichsgrenzen angepasst. Probemessung erfolgreich. Präanalytik Einweisung Personal. | SWO-79339 | No | measurement range limits. Test measurement successful. Pre-analytical instruction personnel. | measurement range limits. Test measurement successful. Pre-analytical instruction personnel. | Setup | User Ranges | User Ranges | User Ranges
ABL90 FLEX PLUS | Touchscreen erneuert und kalibriert | Touchscreen erneuert und kalibriert | SWO-79373 | No | Touchscreen renewed and calibrated | Touchscreen renewed and calibrated | Defect | HW | DMS_PC | Touch screen
TCM CombiM | 更换电池431-018C一个后修复(by andrew) | 更换电池431-018C一个后修复(by andrew) | SWO-78957 | No | Replace the battery 431-018C one post repair (by andrew) | Replace the battery 431-018C one post repair (by andrew) | Defect | HW | Battery module | Battery module
ABL90 FLEX | 更换新测试卡修复仪器 | 更换新测试卡修复仪器 | SWO-79481 | No | Replace the new test card repair instrument | Replace the new test card repair instrument | Assisted Replacement | SC | SC | Chip error
AQT90 FLEX |  | B | SWO-79218 | No |  |  | Assisted Replacement | Consumables | Test Kit | NT-proBNP | 167
ABL800 BASIC | Reset the connections. Run oxi test.run protein removal multiple time. check the lamp intensity. machine under observation. | B | SWO-79431 | No | Reset the connections. Run oxi test.run protein removal multiple time. check the lamp intensity. machine under observation. |  | Cleaned | Fluidic pathway | Oxi module | Hemolyzer
ABL825 FLEX | 更换910-275一个正常。(by andrew) | 更换910-275一个正常。(by andrew) | SWO-79023 | No | Replace 910-275 one normal. (by andrew) | Replace 910-275 one normal. (by andrew) | Defect | HW | Printer module | Printer module
ABL827 FLEX | Allgemeine Überprüfung Analysator - Messkammern und Kontakte überprüft und gereinigt etc. - Erneuerung pH Dichtring (Blut unter Dichtring) - Erneuerung Glucose Dichtring (nicht vorhanden) - Referenz, Na, Cl, Ca und Glucose Elektroden membraniert - Funktionsüberprüfung etc. | Allgemeine Überprüfung Analysator - Messkammern und Kontakte überprüft und gereinigt etc. - Erneuerung pH Dichtring (Blut unter Dichtring) - Erneuerung Glucose Dichtring (nicht vorhanden) - Referenz, Na, Cl, Ca und Glucose Elektroden membraniert - Funktionsüberprüfung etc. | SWO-79372 | No | General inspection analyzer - measuring chambers and contacts checked and cleaned etc. - Renewal pH sealing ring (blood under sealing ring) - renewal glucose sealing ring (not present) - reference, Na, Cl, Cl and glucose electrodes membraneed - functional check etc. | General inspection analyzer - measuring chambers and contacts checked and cleaned etc. - Renewal pH sealing ring (blood under sealing ring) - renewal glucose sealing ring (not present) - reference, Na, Cl, Cl and glucose electrodes membraneed - functional check etc. | Cleaned | HW | pH/BG Module | Reference | Contact
ABL825 FLEX | 更换密封圈，做整机保养后解决 | 更换密封圈，做整机保养后解决 | SWO-79512 | No | Replace the seal ring and do the whole machine after maintenance to solve | Replace the seal ring and do the whole machine after maintenance to solve | Defect | Electrode | Electrode | Lac | Sealing ring
ABL815 FLEX | Gerätekonfiguration mit dem Kunden angepasst | Gerätekonfiguration mit dem Kunden angepasst | SWO-79389 | Yes | Device configuration customized with customer | Device configuration customized with customer | Re classified RMED | setup after 12 days
ABL825 FLEX | Anpassung der Gerätekonfiguration in Abstimmung mit dem Kunden vorgenommen | Anpassung der Gerätekonfiguration in Abstimmung mit dem Kunden vorgenommen | SWO-79388 | Yes | Adjustment of the device configuration in coordination with the customer | Adjustment of the device configuration in coordination with the customer | Re classified RMED | setup after 12 days
ABL825 FLEX | SW 6.16 Installiert nach Überprüfung aller Daten, keine andere Möglichkeit. Baro Korr auf +5 mmHg zurückgenommen zur weiteren Beobachtung | SW 6.16 Installiert nach Überprüfung aller Daten, keine andere Möglichkeit. Baro Korr auf +5 mmHg zurückgenommen zur weiteren Beobachtung | SWO-79345 | No | SW 6.16 Installed after checking all data, no other option. Baro Korr to +5 mmHg withdrawn for further observation | SW 6.16 Installed after checking all data, no other option. Baro Korr to +5 mmHg withdrawn for further observation | Assisted Replacement | Electrode | Electrode | pCO2 | 6.16 installed
AQT90 FLEX | Felszívó tű elgörbült. Mechanikailag sérült az inlet rész. Készülék további javításra elszállítva. Cserekészülék kihelyezve. | B | SWO-79400 | No | Absorbent needle bent. The part of the inlet is mechanically damaged. Device is shipped for further repair. Replacement. |  | Defect | HW | Aspiration Module | Bend needle
ABL80 FLEX BASIC | 更换新测试卡修复仪器 | 更换新测试卡修复仪器 | SWO-79527 | No | Replace the new test card repair instrument | Replace the new test card repair instrument
ABL90 FLEX | 更换910-285修复仪器(by andrew) | 更换910-285修复仪器(by andrew) | SWO-79234 | No | Replace the 910-285 repair instrument (by andrew) | Replace the 910-285 repair instrument (by andrew) | Defect | HW | Power Supply | Power Supply
ABL80 FLEX BASIC | 更换新测试卡修复仪器 | 更换新测试卡修复仪器 | SWO-79835 | No | Replace the new test card repair instrument | Replace the new test card repair instrument
ABL80 FLEX BASIC | 更换新测试卡修复仪器 | 更换新测试卡修复仪器 | SWO-80078 | No | Replace the new test card repair instrument | Replace the new test card repair instrument
ABL80 FLEX BASIC |  | B | SWO-79480 | No
ABL90 FLEX | 更换新的试剂包修复仪器 | 更换新的试剂包修复仪器 | SWO-79855 | No | Replace the new reagent kit repair instrument | Replace the new reagent kit repair instrument | Assisted Replacement | SP | SP | Pump cal error
ABL90 FLEX | 更换新测试卡修复仪器 | 更换新测试卡修复仪器 | SWO-79494 | No | Replace the new test card repair instrument | Replace the new test card repair instrument | Assisted Replacement | SC | SC | Lac
ABL80 FLEX BASIC | 更换新测试卡修复仪器 | 更换新测试卡修复仪器 | SWO-79836 | No | Replace the new test card repair instrument | Replace the new test card repair instrument
ABL80 FLEX BASIC | 更换新测试卡修复仪器 | 更换新测试卡修复仪器 | SWO-79652 | No | Replace the new test card repair instrument | Replace the new test card repair instrument
ABL80 FLEX BASIC | 更换新测试卡修复仪器 | 更换新测试卡修复仪器 | SWO-79868 | No | Replace the new test card repair instrument | Replace the new test card repair instrument
ABL80 FLEX BASIC | 更换新测试卡修复仪器 | 更换新测试卡修复仪器 | SWO-79870 | No | Replace the new test card repair instrument | Replace the new test card repair instrument
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-79791 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL80 FLEX BASIC |  | B | SWO-79536 | No
ABL80 FLEX BASIC | 连续更换两张测试卡解决 | 连续更换两张测试卡解决 | SWO-79528 | No | Replace two test cards in a row | Replace two test cards in a row
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-79801 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL80 FLEX BASIC | 更换测试卡正常。 | 更换测试卡正常。 | SWO-80031 | No | The replacement test card is OK. | The replacement test card is OK.
ABL80 FLEX BASIC | 更换测试卡正常。 | 更换测试卡正常。 | SWO-79808 | No | The replacement test card is OK. | The replacement test card is OK.
ABL90 FLEX | Leere Datenbank installiert Gerät konfiguriert und vernetzt Lösungspack aus Kundenbestand erneuert (UB-08) ABL kalibriert und überprüft | Leere Datenbank installiert Gerät konfiguriert und vernetzt Lösungspack aus Kundenbestand erneuert (UB-08) ABL kalibriert und überprüft | SWO-79734 | No | Empty database installs device configured and networked solution pack from customer stock renewed (UB-08) ABL calibrated and verified | Empty database installs device configured and networked solution pack from customer stock renewed (UB-08) ABL calibrated and verified | Defect | SW | Data base | Error code 2
ABL90 | Techline has suggested replacement of Hemolyzer. | Replaced Hemolyzer. | D03C69CB4A158F1B | No | Techline has suggested replacement of Hemolyzer. | Replaced Hemolyzer. | Defect | HW | Hemolyzer | Hemolyzer
ABL90 | Green light blinking on Battery Pack. | Replace defective battery. | D03C6AA64A2F8F1F | No | Green light blinking on Battery Pack. | Replace defective battery. | Defect | HW | Battery module | Battery
ABL90 | Examined software version on instrument where setup was loaded from. Found version mismatch on analyzers. | Reloaded setup with compatible software version. Upgraded instrument to MR7 software, since customer will be interfacing with soft. | D03C6CDA4A3362B3 | No | Examined software version on instrument where setup was loaded from. Found version mismatch on analyzers. | Reloaded setup with compatible software version. Upgraded instrument to MR7 software, since customer will be interfacing with soft. | Defect | SW | Error 145 | Data base
ABL90 FLEX | Probengewinnung angeschaut, darauf hin Kundenschulung durchgeführt - Das Problem lag beim nicht korrekten Mischen der Probe - der Kunde verwendet jetzt safe pico´s und die Mischvorrichtung am ABL | Probengewinnung angeschaut, darauf hin Kundenschulung durchgeführt - Das Problem lag beim nicht korrekten Mischen der Probe - der Kunde verwendet jetzt safe pico´s und die Mischvorrichtung am ABL | SWO-79755 | No | Sample collection looked at, therefore customer training carried out - The problem was with incorrect mixing of the sample - the customer now uses safe pico's and the mixing device on the ABL | Sample collection looked at, therefore customer training carried out - The problem was with incorrect mixing of the sample - the customer now uses safe pico's and the mixing device on the ABL | User error | Pre analytical | Mixing | tHb
ABL90 FLEX | 更换测试卡，试剂包正常。 | 更换测试卡，试剂包正常。 | SWO-79886 | No | Replace the test card and the reagent package is OK. | Replace the test card and the reagent package is OK. | Assisted Replacement | SP | SP Installation | SP Installation | 1178
AQT90 FLEX |  | B | SWO-79534 | No |  |  | Defect | HW | Process Module | Optical Unit | 131
ABL827 FLEX | The Glu, lac Crea and pO2 were out on cals. When I arrived the pO2 was failing the zero. The pO2 zero will not come I. Cleaned chamber and checked for bubbles. The electrode is under warranty. I will order a replacement electrode and return. Customer ran QC and will use everything except pO2. 2-2-16 Went to Fedex to pick up pO2 electrode. When I arrived the pO2 was in. They had run all QC and it was in. Customer had ordered their pump tubes and I walked them thru replacing them today. They will run QC when the cals pass. | Debubbled the pO2 and ran zero calibration. Ordered pO2 electrode. 2-2-16 No correction needed. Analyzer was up and running. | D03C45054A621F83 | No | The Glu, lac Crea and pO2 were out on cals. When I arrived the pO2 was failing the zero. The pO2 zero will not come I. Cleaned chamber and checked for bubbles. The electrode is under warranty. I will order a replacement electrode and return. Customer ran QC and will use everything except pO2. 2-2-16 Went to Fedex to pick up pO2 electrode. When I arrived the pO2 was in. They had run all QC and it was in. Customer had ordered their pump tubes and I walked them thru replacing them today. They will run QC when the cals pass. | Debubbled the pO2 and ran zero calibration. Ordered pO2 electrode. 2-2-16 No correction needed. Analyzer was up and running. | Cleaned | Fluidic pathway | pH/BG Module | pO2
ABL825 FLEX | Messkammer und Kontakte gereinigt, PCO 2 Membrane einer anderen Lot aufgezogen. Mehrfach Kalibriert und QC gemessen. Es ist geplant die neue Software V 6.16 zu Installieren. | Messkammer und Kontakte gereinigt, PCO 2 Membrane einer anderen Lot aufgezogen. Mehrfach Kalibriert und QC gemessen. Es ist geplant die neue Software V 6.16 zu Installieren. | SWO-79821 | No | Measuring chamber and contacts cleaned, PCO 2 membrane of another solder. Multiple calibrated and QC measured. It is planned to install the new software V 6.16. | Measuring chamber and contacts cleaned, PCO 2 membrane of another solder. Multiple calibrated and QC measured. It is planned to install the new software V 6.16. | Assisted Replacement | Electrode | Electrode | pCO2
ABL825 FLEX | Oxi Complete Test durchgeführt --> Halogene lamp intensity >3900, Lampe neu justiert, Oxi Complete Test i.O. (Intensität ~3200), Gerät einkalibriert, tHb neu kalibriert, i.O. | Oxi Complete Test durchgeführt --> Halogene lamp intensity >3900; Lampe neu justiert, Oxi Complete Test i.O. (Intensität ~3200), Gerät einkalibriert, tHb neu kalibriert, i.O. | SWO-79781 | No | Oxi Complete Test performed --> Halogens lamp intensity >3900; Lamp recalibrated, Oxi Complete Test i.O. (intensity 3200), device calibrated, tHb recalibrated, i.O. | Oxi Complete Test performed --> Halogens lamp intensity >3900; Lamp recalibrated, Oxi Complete Test i.O. (intensity 3200), device calibrated, tHb recalibrated, i.O. | Adjustment | HW | Oxi module | Lamp unit
ABL825 FLEX | Levegőző szelep kitisztítása, LS-Pumpa-, Elektróda- kalibrálás, QC mérés | Levegőző szelep kitisztítása, LS-Pumpa-, Elektróda- kalibrálás, QC mérés | SWO-79655 | No | Cleaning the air valve, LS-pump, electrode calibration, QC measurement | Cleaning the air valve, LS-pump, electrode calibration, QC measurement | Cleaned | Fluidic pathway | Fluidic manifold | Air vent
ABL827 FLEX | 1-29-16 Beth called in with rinse error and leaking from inlet gasket. I had her replace the rinse bottle. she did not have an inlet gasket. I had her clean the inlet gasket and then she got inlet cal error. 2-1-16 When I arrived I found the probe bent. I did not have an inlet gasket so I put on a pm kit. They said they had ha several rine errors. I replaced the pH BG valves and rinse, air and waste valves on the fluidic module. I did not see any other problem. Amy is in a meeting so I will sing verbal per Amy. | I put on a PM kit, inlet probe, pH BG air and rinse valves and the fluidic module air, rinse and waste valves. Ran pump cal, inlet liquid sensor cal and inlet cal. I am going to open a techline case. They have had different issue since I installed the analyzer. | D03C27C84A2F38F2 | No | 1-29-16 Beth called in with rinse error and leaking from inlet gasket. I had her replace the rinse bottle. she did not have an inlet gasket. I had her clean the inlet gasket and then she got inlet cal error. 2-1-16 When I arrived I found the probe bent. I did not have an inlet gasket so I put on a pm kit. They said they had ha several rine errors. I replaced the pH BG valves and rinse, air and waste valves on the fluidic module. I did not see any other problem. Amy is in a meeting so I will sing verbal per Amy. | I put on a PM kit, inlet probe, pH BG air and rinse valves and the fluidic module air, rinse and waste valves. Ran pump cal, inlet liquid sensor cal and inlet cal. I am going to open a techline case. They have had different issue since I installed the analyzer. | Defect | HW | Inlet module | Probe bend
AQT90 FLEX | Austausch Drucker Restart des Systems Funktionskontrolle ok | Austausch Drucker Restart des Systems Funktionskontrolle ok | SWO-79747 | No | Replacement printer restart of system function control ok | Replacement printer restart of system function control ok | Defect | HW | Analyzer Control Module | Printer
ABL80 FLEX BASIC | 更换新测试卡修复仪器 | 更换新测试卡修复仪器 | SWO-79871 | No | Replace the new test card repair instrument | Replace the new test card repair instrument
ABL90 FLEX | Gerät überprüft und danach eine Wartung in separaten Auftrag durchgeführt. | Gerät überprüft und danach eine Wartung in separaten Auftrag durchgeführt. | SWO-79904 | No | device and then performed maintenance in a separate order. | device and then performed maintenance in a separate order. | Cleaned | Fluidic pathway | Fluidic manifold | Fluidic pathway
ABL90 FLEX | csere oldat leszállítása a készülékhez, oldat és új SC installálás, kalibráció, QC mérés | B | SWO-79930 | No | Delivery of the replacement solution to the device, solution and new SC installation, calibration, QC |  | Assisted Replacement | SC | SC | Electrolytes
ABL90 FLEX | Gerät gespült, kalibriert und QC wiederholt | Gerät gespült, kalibriert und QC wiederholt | SWO-79964 | No | Device rinsed, calibrated and QC repeated | Device rinsed, calibrated and QC repeated | Cleaned | Fluidic pathway | Fluidic pathway | pH
ABL90 FLEX | Justagen, Kalibrationen und Qualitätskontrollen durchgeführt | Justagen, Kalibrationen und Qualitätskontrollen durchgeführt | SWO-79881 | No | Adjustments, calibrations and quality controls carried out | Adjustments, calibrations and quality controls carried out | Unknown | Inlet | Inlet | pO2
ABL90 FLEX PLUS |  | B | D03D7342C04AEBDC | No |  |  | Defect | HW | PM Kit | PM Kit
ABL90 FLEX PLUS | Konfiguration bedingt ein Anmelden um das ID Feld nachträglich zu editieren. Kinderlabor entspr. informiert. | Konfiguration bedingt ein Anmelden um das ID Feld nachträglich zu editieren. Kinderlabor entspr. informiert. | SWO-80003 | Yes | Configuration requires a login to edit the ID field afterwards. Children's laboratory Informed. | Configuration requires a login to edit the ID field afterwards. Children's laboratory Informed. | Re classified RMED | Layout | Patient ID | Patient ID
ABL90 FLEX | 更换新试剂包后解决。共更换一张测试卡和一个试剂包。 | 更换新试剂包后解决。共更换一张测试卡和一个试剂包。 | SWO-79874 | No | Resolve after replacing the new kit. Replace a test card and a reagent kit. | Resolve after replacing the new kit. Replace a test card and a reagent kit. | Assisted Replacement | SC | SC | Lac
ABL90 | Cambio del SC y SP. Verificación del equipo con QC´s y configuración del puerto de comunicaciones. | Cambio del SC y SP. Verificación del equipo con QC´s y configuración del puerto de comunicaciones. | D03EB517361E643F | No | SC and SP change. Verification of the equipment with QC's and configuration of the communications port. | SC and SP change. Verification of the equipment with QC's and configuration of the communications port. | Assisted Replacement | SC | SC | Reference
AQT90 FLEX | 更换光学单元903-364一个后修复 | 更换光学单元903-364一个后修复 | SWO-79811 | No | Replacement of optical unit 903-364 one post-repair | Replacement of optical unit 903-364 one post-repair | Defect | HW | Process Module | Optical Unit
ABL837 FLEX |  | B | D03D6645C04F3228 | No |  |  | Assisted Replacement | Electrode | Electrode | pO2
ABL835 FLEX | Čišćenje modula Autocheck. Provjera --> OK. | B | SWO-79922 | No | Cleaning the Autocheck module. Check--> OK. |  | Cleaned | HW | AutoCheck Module | Barcode reader
ABL825 FLEX | checked logs and meditech graphs- shark tooth compared to other analyzer Ranges also too tight .04 at low levels .08 at high. insert ranges are 0.2 | Removed measuring chamber and inspected- fine, removed air and gas valves -fine , checked amplifier cap Ica and ref- fine. Replaced Ei met pcb , | D03D90A6C0202D2E | No | checked logs and meditech graphs- shark tooth compared to other analyzer Ranges also too tight .04 at low levels .08 at high. insert ranges are 0.2 | Removed measuring chamber and inspected- fine, removed air and gas valves -fine , checked amplifier cap Ica and ref- fine. Replaced Ei met pcb , | Defect | HW | El/Met Module | PCB | Ca
ABL815 FLEX | AutoQC Modul demontiert, Barcode Laser gereinigt, Gerät kalibriert und QC durchgeführt. | AutoQC Modul demontiert, Barcode Laser gereinigt, Gerät kalibriert und QC durchgeführt. | SWO-80313 | No | AutoQC module dismantled, barcode laser cleaned, device calibrated and QC performed. | AutoQC module dismantled, barcode laser cleaned, device calibrated and QC performed. | Cleaned | HW | AutoCheck Module | Barcode Reader
ABL825 FLEX | Kontakte etc. gereinigt - Nullstromkalibrierung - mehrfache Kalibrationen durchgeführt - Funktionsüberprüfung, Deaktivierung Demokeys und Aktivierung Option Keys (Update Type) | Kontakte etc. gereinigt - Nullstromkalibrierung - mehrfache Kalibrationen durchgeführt - Funktionsüberprüfung, Deaktivierung Demokeys und Aktivierung Option Keys (Update Type) | SWO-79979 | No | Contacts etc. cleaned - zero current calibration - multiple calibrations performed - function check, deactivation demokeys and activation option keys (Update Type) | Contacts etc. cleaned - zero current calibration - multiple calibrations performed - function check, deactivation demokeys and activation option keys (Update Type) | Unknown | Electrode | Electrode | pO2
ABL825 FLEX | Fehlersuche.ohne Erfolg. Rücknahme Baro Korrektur +19mmHg nötig nach SW Upgrade 6.16 Baro Korrektur + 5mmHg belassen . | Fehlersuche.ohne Erfolg. Rücknahme Baro Korrektur +19mmHg nötig nach SW Upgrade 6.16 Baro Korrektur + 5mmHg belassen . | SWO-80289 | No | Troubleshooting.without success. Withdrawal Baro correction +19mmHg necessary after SW upgrade 6.16 Baro correction + 5mmHg leave. | Troubleshooting.without success. Withdrawal Baro correction +19mmHg necessary after SW upgrade 6.16 Baro correction + 5mmHg leave. | Assisted Replacement | Electrode | Electrode | pCO2
ABL835 FLEX | 0376: Calibration drift, Lactate drifting on Cal. Tried a new Lactate electrode but Lactate is drifting on calibartion and causing 379 errors. | Replaced EI Met module PCB. Completed all servcie tests. Took sample counter. | D03DCCB5C001DC75 | No | 0376: Calibration drift, Lactate drifting on Cal. Tried a new Lactate electrode but Lactate is drifting on calibartion and causing 379 errors. | Replaced EI Met module PCB. Completed all servcie tests. Took sample counter. | Defect | HW | El/Met Module | PCB
AQT90 FLEX | SERVICE DUMP AND SUGGESTED THAT WINDOWS BE RELOADED | REINSTALLED WINDOWS AND MAIN SOFTWARE. INSTALLED ALL SETTINGS. | D03DDAE9C03609A5 | Yes | SERVICE DUMP AND SUGGESTED THAT WINDOWS BE RELOADED | REINSTALLED WINDOWS AND MAIN SOFTWARE. INSTALLED ALL SETTINGS. | Re classified RMED | Internal setup / RAME |  |  | 124
ABL80 FLEX BASIC | 更换测试卡正常。 | 更换测试卡正常。 | SWO-80044 | No | The replacement test card is OK. | The replacement test card is OK.
ABL80 FLEX BASIC | 更换打印机主板后修复 | 更换打印机主板后修复 | SWO-80271 | No | Repair after replacing the printer board | Repair after replacing the printer board
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-80034 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL80 CO-OX | Found that the waste port for the solution pack blocked upon initial installation. | Installed a new solution pack and installed the annual maintenance kit components. | D03ED2223652E809 | No | Found that the waste port for the solution pack blocked upon initial installation. | Installed a new solution pack and installed the annual maintenance kit components.
ABL80 OSM | No investigation needed it is a on going problem. | Replaced the pinch valve, annual maintenance manifold. | D03EBCA836408CA6 | No | No investigation needed it is a on going problem. | Replaced the pinch valve, annual maintenance manifold.
ABL90 | Noticed Flow Selector, Actuator in Red in Service. | Ran Service diagnostics and calibrations. | D03F0AB5364FFFD1 | No | Noticed Flow Selector, Actuator in Red in Service. | Ran Service diagnostics and calibrations. | Adjustment | Service pgm | Flow Selector | Actuator
ABL90 FLEX | Konfiguration geprüft, Netzwerkeinstellungen iO, Analysator am Kreißsaal Anschluss getestet ok | Konfiguration geprüft, Netzwerkeinstellungen iO, Analysator am Kreißsaal Anschluss getestet ok | SWO-80647 | No | Configuration checked, network settings iO, analyzer tested at the delivery room connection ok | Configuration checked, network settings iO, analyzer tested at the delivery room connection ok | Setup | His/lis | Radiance connection | Radiance connection
TCM CombiM | 更换电池431-018C一个后修复 | 更换电池431-018C一个后修复 | SWO-79578 | No | Replace the battery 431-018C one after repair | Replace the battery 431-018C one after repair
AQT90 FLEX | 更换903-364修复仪器 | 更换903-364修复仪器 | SWO-80026 | No | Replace the 903-364 repair instrument | Replace the 903-364 repair instrument | Defect | HW | Process Module | Optical Unit
ABL825 FLEX | 做血红蛋白定标，暂时修复，由于备件质量问题，未完全修复，需申请新的902-833一个 | 做血红蛋白定标，暂时修复，由于备件质量问题，未完全修复，需申请新的902-833一个 | SWO-79806 | No | Do hemoglobin calibration, temporary repair, due to spare parts quality problems, not completely repaired, need to apply for a new 902-833 one | Do hemoglobin calibration, temporary repair, due to spare parts quality problems, not completely repaired, need to apply for a new 902-833 one | Revisit required | HW | Oxi module | Oxi module
ABL837 FLEX |  | B | D03EC7243634322B | No |  |  | Unknown | SW | His/lis | HIS/LIS
ABL815 FLEX | Überprüfung Oxi Modul => Halogen Lampen Intensität zu hoch Austausch Halogenlampe und Justage Complete Oximetry Test i.O. => Intensität bei ca. 3350 Kal2, tHb-Kal i.O. => Fcuv = 0,9803 | Überprüfung Oxi Modul => Halogen Lampen Intensität zu hoch Austausch Halogenlampe und Justage Complete Oximetry Test i.O. => Intensität bei ca. 3350 Kal2, tHb-Kal i.O. => Fcuv = 0,9803 | SWO-80109 | No | Check Oxi Module => Halogen lamps Intensity too high Replacement halogen lamp and Adjustment Complete Oximetry Test i.O. => Intensity at approx. 3350 Kal2, tHb-Kal i.O. => Fcuv = 0.9803 | Check Oxi Module => Halogen lamps Intensity too high Replacement halogen lamp and Adjustment Complete Oximetry Test i.O. => Intensity at approx. 3350 Kal2, tHb-Kal i.O. => Fcuv = 0.9803 | Adjustment | HW | Oxi module | Lamp unit
AQT90 FLEX | Checked process wheel & optical unit set up positions, no issues detected there. Finally after thorough check up found out that the support steel rod connected to wet section plate has been broken during transportation.Need to replace the same along with vision barcode flat cable, which is also found with broken connection. | Managed somehow with extra washer & nuts to create enough tension to hold up the optical unit system. Now process wheel is moving freely. Also cup transport test is passing. But vision barcode flat cable need to be replaced. | SWO-80736 | No | Checked process wheel & optical unit set up positions, no issues detected there. Finally after thorough check up found out that the support steel rod connected to wet section plate has been broken during transportation. Need to replace the same along with vision barcode flat cable, which is also found with broken connection. | Managed somehow with extra washer & nuts to create enough tension to hold up the optical unit system. Now process wheel is moving freely. Also cup transport test is passing. But vision barcode flat cable need to be replaced. | Other | HW | Transportation | Transportation
AQT90 FLEX | mit Herrn Sieburg von Convorx Rücksprache gehalten und festgestellt da keine Pat.ID mit 61... vom KH hinterlegt ist, daher wartet unser AQT sehr lange auf eine Antwort von Convorx mit den Namen des Pat. Herr Sieburg wird Frau Spiegel kontaktieren ob diese eingepfelgr werden, bis dahin die automatische Pat.ID Abfrage aus der Konfig. entfernt. | mit Herrn Sieburg von Convorx Rücksprache gehalten und festgestellt da keine Pat.ID mit 61... vom KH hinterlegt ist, daher wartet unser AQT sehr lange auf eine Antwort von Convorx mit den Namen des Pat. Herr Sieburg wird Frau Spiegel kontaktieren ob diese eingepfelgr werden, bis dahin die automatische Pat.ID Abfrage aus der Konfig. entfernt. | SWO-80253 | No | consultation with Mr. Sieburg of Convorx and found since no Pat.ID with 61... is deposited by the KH, therefore our AQT waits very long for a response from Convorx with the name of the Pat. Mr. Sieburg will contact Ms. Spiegel if they will be splendid, until then the automatic Pat.ID query from the config. away. | consultation with Mr. Sieburg of Convorx and found since no Pat.ID with 61... is deposited by the KH, therefore our AQT waits very long for a response from Convorx with the name of the Pat. Mr. Sieburg will contact Ms. Spiegel if they will be splendid, until then the automatic Pat.ID query from the config. away. | Defect | SW | Analyzer Control Module | Analyzer logic
TCM CombiM | Akku getauscht. | B | SWO-80181 | No | Battery replaced.
ABL80 FLEX BASIC | 更换新测试卡修复仪器 | 更换新测试卡修复仪器 | SWO-80308 | No | Replace the new test card repair instrument | Replace the new test card repair instrument
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-80531 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL80 FLEX BASIC | 将固定接口板的定位销调整，把固定的黑色销磨平一些后故障消失。借用的接口板未更换。 | 将固定接口板的定位销调整，把固定的黑色销磨平一些后故障消失。借用的接口板未更换。 | SWO-79873 | No | Adjust the anchor pin of the fixed interface plate and grind the fixed black pin flat out some of the faults to disappear. The borrowed interface board has not been replaced. | Adjust the anchor pin of the fixed interface plate and grind the fixed black pin flat out some of the faults to disappear. The borrowed interface board has not been replaced.
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-80274 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL80 FLEX BASIC | Cleaned Sc and reinstalled again. Given Rinse 3-4 times and then run QC.Now result are matching. | Cleaned Sc and reinstalled again. Given Rinse 3-4 times and then run QC.Now result are matching. | SWO-80733 | No | Cleaned Sc and reinstalled again. Given Rinse 3-4 times and then run QC. Now result are matching. | Cleaned Sc and reinstalled again. Given Rinse 3-4 times and then run QC. Now result are matching.
ABL90 FLEX | Datensicherung, SW-Update Vers. 3.1MR7, Konfiguration überprüft - | WXPE neu installiert, Applikations-SW installiert, Konfiguration angepasst, WXPE konfiguriert, Kal u. QC i.O. | SWO-80356 | No | Data backup, SW update verse. 3.1MR7, configuration checked - | WXPE reinstalled, application SW installed, configuration adapted, WXPE configured, Kal & QC i.O. | Defect | SW | Print multible Reports | Print multible Reports
ABL90 FLEX | WXPE neu installiert, Applikations-SW installiert, Konfiguration angepasst, WXPE konfiguriert, Kal u. QC i.O. | WXPE neu installiert, Applikations-SW installiert, Konfiguration angepasst, WXPE konfiguriert, Kal u. QC i.O. | SWO-80543 | Yes | WXPE reinstalled, application SW installed, configuration adapted, WXPE configured, Kal & QC i.O. | WXPE reinstalled, application SW installed, configuration adapted, WXPE configured, Kal & QC i.O. | Re classified RMED | Double entry
ABL90 FLEX | Complete oxitest : hier wordt er geen probleem gedetecteerd. De hemolyser werd vervangen : complete oxitest + tHb kalibratie OK Kalibraties en QC's : OK Er werden eveneens stalen gemeten: OK | B | SWO-80368 | No | Complete oxitest: No problem is detected here. The Hemolyser was replaced: complete oxitest + tHb calibration OK calibrations and Qcs: OK samples were also measured: OK |  | Defect | HW | Hemolyzer | Hemolyzer
ABL90 | Confirmed problem. | Replaced the tubing and gaskets with a maintenance kit. Also put a o-ring in the inlet. | D0409256AC569962 | No | Confirmed problem. | Replaced the tubing and gaskets with a maintenance kit. Also put a o-ring in the inlet. | Defect | HW | PM Kit | pO2
ABL90 FLEX | Einlassdichtung erneuert, Pressure Test durchgeführt i.O. Gerät in der Start Phase. Kal und QC i.O | Einlassdichtung erneuert, Pressure Test durchgeführt i.O. Gerät in der Start Phase. Kal und QC i.O | SWO-80511 | No | Inlet seal renewed, Pressure test carried out usually device in the start phase. Kal and QC i.O | Inlet seal renewed, Pressure test carried out usually device in the start phase. Kal and QC i.O | Defect | HW | Inlet Gasket | Inlet gasket
ABL90 FLEX | Service Dump & Data Logs erstellt und an TS geschickt Kunde ist über ständige Probleme & häufigen LP SC Wechsel sehr unzufrieden, Credit Claim ausgedruckt 4 Stück in einem Monat Kunden über Sachverhalt informiert | Service Dump & Data Logs erstellt und an TS geschickt Kunde ist über ständige Probleme & häufigen LP SC Wechsel sehr unzufrieden, Credit Claim ausgedruckt 4 Stück in einem Monat Kunden über Sachverhalt informiert | SWO-80448 | No | Service Dump & Data Logs created and sent to TS customer is very dissatisfied about constant problems & frequent LP SC bills of exchange, Credit Claim printed 4 pieces in a month customer informed about facts | Service Dump & Data Logs created and sent to TS customer is very dissatisfied about constant problems & frequent LP SC bills of exchange, Credit Claim printed 4 pieces in a month customer informed about facts | Assisted Replacement | SP | SP | pCO2
ABL90 FLEX | Fehlersuche | Fehlersuche, Austausch Sensorbay, Ersatz Lösungspack, Funktionskontrolle | SWO-83765 | No | Troubleshooting | Troubleshooting, Replacement Sensorbay, Replacement Solution Pack, Function Control | Assisted Replacement | SP | SP | Suspected Clot in Inlet
ABL90 | Rechecked all power save settings, all still disabled. Enabled trace log, waited 2 hours but trace log did not start. Edited Roma.ini file with Trace=Yes and saved changes. Also added a blank line between [NetOp] and [Connection 1]. Waited another 2 hours and copied Trace log for further investigation. | B | D04091A1AC30F56D | No | Rechecked all power save settings, all still disabled. Enabled trace log, waited 2 hours but trace log did not start. Edited Roma.ini file with Trace=Yes and saved changes. Also added a blank line between [NetOp] and [Connection 1]. Waited another 2 hours and copied Trace log for further investigation. |  | Setup | WIFI | Power saving | Power saving
ABL90 FLEX PLUS | Date and Time Einstellung vorgenommen - Funktionsüberprüfung etc. | Date and Time Einstellung vorgenommen - Funktionsüberprüfung etc. | SWO-80362 | No | Date and Time setting made - function check etc. | Date and Time setting made - function check etc. | Setup | SW | Time setting | Time setting
ABL800 BASIC | Da Vor Ort, mit Anwender zusammen Konfig. überprüft und Kal. Zeiten angepasst, Ref. und Natrium Elektrode neu membraniert, da Natrium QC Fehler, externen Scanner freigeschalten System-und Funktionskontrolle i.O. | Da Vor Ort, mit Anwender zusammen Konfig. überprüft und Kal. Zeiten angepasst, Ref. und Natrium Elektrode neu membraniert, da Natrium QC Fehler, externen Scanner freigeschalten System-und Funktionskontrolle i.O. | SWO-80414 | No | Since on site, with users together config. checked and Kal. Times adjusted, ref. and sodium electrode re-membraneed since sodium QC error, external scanner unlock system and function control i.O. | Since on site, with users together config. checked and Kal. Times adjusted, ref. and sodium electrode re-membraneed since sodium QC error, external scanner unlock system and function control i.O. | Assisted Replacement | Electrode | Membrane | Na+
ABL827 FLEX | checked data logs for errors and QC Log. checked valves and Gas supply, checked electrodes and membranes, | removed all electrodes from pH BG module and cleaned electrode chamber, replaced membranes on all electrodes, | D0409053AC303C76 | No | checked data logs for errors and QC Log. checked valves and Gas supply, checked electrodes and membranes, | removed all electrodes from pH BG module and cleaned electrode chamber, replaced membranes on all electrodes, | Assisted Replacement | Electrode | Membrane | pCO2
ABL825 FLEX | checked data logs for errors and QC Log. checked valves and Gas supply, checked electrodes and membranes, | removed all electrodes from pH BG module and cleaned electrode chamber, replaced membranes on all electrodes, | D0408E9EAC443C6F | No | checked data logs for errors and QC Log. checked valves and Gas supply, checked electrodes and membranes, | removed all electrodes from pH BG module and cleaned electrode chamber, replaced membranes on all electrodes, | Assisted Replacement | Electrode | Membrane | pCO2
ABL835 FLEX | checked data logs for errors and QC Log. checked valves and Gas supply, checked electrodes and membranes, | removed all electrodes from pH BG module and cleaned electrode chamber, replaced membranes on all electrodes, | D0408F59AC013C73 | No | checked data logs for errors and QC Log. checked valves and Gas supply, checked electrodes and membranes, | removed all electrodes from pH BG module and cleaned electrode chamber, replaced membranes on all electrodes, | Assisted Replacement | Electrode | Membrane | pCO2
ABL827 FLEX | Erneuerung Referenz Elektrode - Funktionsüberprüfung etc. | Erneuerung Referenz Elektrode - Funktionsüberprüfung etc. | SWO-80482 | No | Renewal Reference Electrode - Function Verification etc. | Renewal Reference Electrode - Function Verification etc. | Assisted Replacement | Electrode | Electrode | Reference
ABL820 FLEX | CO2 Elektrode überprüft, CO2 Elektrode ersetzt. | CO2 Elektrode überprüft, CO2 Elektrode ersetzt. | SWO-80499 | No | CO2 electrode checked, CO2 electrode replaced. | CO2 electrode checked, CO2 electrode replaced. | Assisted Replacement | Electrode | Electrode | pCO2
ABL825 FLEX |  | 做血红蛋白定标，暂时修复，由于备件质量问题，未完全修复，需申请新的902-833一个 | SWO-80269 | No |  | Do hemoglobin calibration, temporary repair, due to spare parts quality problems, not completely repaired, need to apply for a new 902-833 one | Defect | HW | Oxi module | Oxi module
ABL825 FLEX | Fehlersuche, Reinigung Fließwege, LSjustage, Funktionskontrolle | Fehlersuche, Reinigung Fließwege, LSjustage, Funktionskontrolle | SWO-83764 | No | Troubleshooting, cleaning flow paths, LSjustage, function control | Troubleshooting, cleaning flow paths, LSjustage, function control | Cleaned | Fluidic pathway | - | -
ABL800 BASIC | cleaned the touch screen and calibrated. presently working but it need to replace with new. | cleaned the touch screen and calibrated. presently working but it need to replace with new. | SWO-80638 | No | cleaned the touch screen and calibrated. presently working but it need to replace with new. | cleaned the touch screen and calibrated. presently working but it need to replace with new. | Defect | HW | DMS_PC | Touch Screen
ABL800 BASIC | Cleaned the deposit and calibrated. | Cleaned the deposit and calibrated. | SWO-80664 | No | Cleaned the deposit and calibrated. | Cleaned the deposit and calibrated. | Cleaned | Fluidic pathway | pH/BG Module | Measuring chamber
ABL825 FLEX | found that user is getting 796 error on analyzer. | replaced solution pump tubbing on analyzer from customer stock. | D03FF999AC33DB2A | No | found that user is getting 796 error on analyzer. | replaced solution pump tubbing on analyzer from customer stock. | Defect | HW | Fluidic manifold | Solution Tube
ABL815 FLEX | Fehlersuche.ohne Erfolg. Rücknahme Baro Korrektur +19mmHg nötig nach SW Upgrade 6.16 Baro Korrektur + 5mmHg belassen . | Fehlersuche.ohne Erfolg. Rücknahme Baro Korrektur +19mmHg nötig nach SW Upgrade 6.16 Baro Korrektur + 5mmHg belassen . | SWO-80288 | No | Troubleshooting.without success. Withdrawal Baro correction +19mmHg necessary after SW upgrade 6.16 Baro correction + 5mmHg leave. | Troubleshooting.without success. Withdrawal Baro correction +19mmHg necessary after SW upgrade 6.16 Baro correction + 5mmHg leave. | Assisted Replacement | Electrode | Electrode | pCO2
ABL815 FLEX | Waste Connector ersetzt. Wege wieder frei | Waste Connector ersetzt. Wege wieder frei | SWO-80494 | No | Waste Connector replaced. Ways clear again | Waste Connector replaced. Ways clear again | Defect | HW | Basic unit | Waste connector
ABL835 FLEX | Der er udført Lamp Adjustment. Der er kørt Complete Oximetry Test. Der er kørt service og bruger tHb. Der er kørt 2 og 1 punkts kal. Der er kørt QCer. Alt er OK | Der er udført Lamp Adjustment. Der er kørt Complete Oximetry Test. Der er kørt service og bruger tHb. Der er kørt 2 og 1 punkts kal. Der er kørt QCer. Alt er OK | SWO-80587 | No | Lamp Adjustment has been performed. Complete Oximetry Test has been run. Service and user tHb are provided. 2 and 1 point Kal have been run. QCer has been run. Everything is OK | Lamp Adjustment has been performed. Complete Oximetry Test has been run. Service and user tHb are provided. 2 and 1 point Kal have been run. QCer has been run. Everything is OK | Adjustment | HW | Oxi module | Lamp unit
ABL815 FLEX | Fehlersuche.ohne Erfolg. Rücknahme Baro Korrektur +19mmHg nötig nach SW Upgrade 6.16 Baro Korrektur + 5mmHg belassen . | Baro Korrektur + 15 mmHg | SWO-80290 | No | Troubleshooting.without success. Withdrawal Baro correction +19mmHg necessary after SW upgrade 6.16 Baro correction + 5mmHg leave. | Baro Correction + 15 mmHg | Assisted Replacement | Electrode | Electrode | pCO2
AQT90 FLEX | Bestellten Scanner ausgetauscht, Kalibrationsbarcode wird jetzt gescannt, Funktion OK. | Bestellten Scanner ausgetauscht, Kalibrationsbarcode wird jetzt gescannt, Funktion OK. | SWO-80476 | No | Ordered scanner replaced, calibration bar code is now scanned, function OK. | Ordered scanner replaced, calibration bar code is now scanned, function OK. | Defect | HW | Analyzer Control Module | Barcode reader, display
AQT90 FLEX | Managed somehow with extra washer & nuts to create enough tension to hold up the optical unit system. Now process wheel is moving freely. Also cup transport test is passing. But vision barcode flat cable need to be replaced. | Managed somehow with extra washer & nuts to create enough tension to hold up the optical unit system. Now process wheel is moving freely. Also cup transport test is passing. But vision barcode flat cable need to be replaced. | SWO-80737 | Yes | Managed somehow with extra washer & nuts to create enough tension to hold up the optical unit system. Now process wheel is moving freely. Also cup transport test is passing. But vision barcode flat cable need to be replaced. | Managed somehow with extra washer & nuts to create enough tension to hold up the optical unit system. Now process wheel is moving freely. Also cup transport test is passing. But vision barcode flat cable need to be replaced. | Re classified RMED | Double entry
AQT90 FLEX | Scanner gereinigt und justiert Funktionskontrolle ok | Scanner gereinigt und justiert Funktionskontrolle ok | SWO-80309 | No | Scanner cleaned and adjusted function control ok | Scanner cleaned and adjusted function control ok | Adjustment | HW | Storage Module | Barcode reader, Storage wheel
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-81860 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL80 FLEX BASIC | 更换新试剂包修复仪器 | 更换新试剂包修复仪器 | SWO-80552 | No | Replace the new reagent kit repair instrument | Replace the new reagent kit repair instrument
ABL80 FLEX BASIC | 更换两个新试剂包和一个新测试卡修复仪器 | 更换两个新试剂包和一个新测试卡修复仪器 | SWO-80593 | No | Replace two new kits and a new test card repair instrument | Replace two new kits and a new test card repair instrument
ABL90 FLEX | Gerät manuell gespült Jahreswartung durchgeführt Einlasssonde aus Kundenbestand erneuert Blutmessung durchgeführt Gerät kalibriert und überprüft | Gerät manuell gespült Jahreswartung durchgeführt Einlasssonde aus Kundenbestand erneuert Blutmessung durchgeführt Gerät kalibriert und überprüft | SWO-80584 | No | Device manually rinsed annual maintenance carried out inlet probe from customer stock renewed blood measurement carried out device calibrated and checked | Device manually rinsed annual maintenance carried out inlet probe from customer stock renewed blood measurement carried out device calibrated and checked | Cleaned | Fluidic pathway | Fluidic pathway | PM Kit
ABL90 FLEX | Gerät manuell gespült Lösungspack aus Kundenbestand erneuert Gerät kalibriert und überprüft | Gerät manuell gespült Lösungspack aus Kundenbestand erneuert Gerät kalibriert und überprüft | SWO-80651 | No | Device manually rinsed solution pack from customer stock renewed device calibrated and checked | Device manually rinsed solution pack from customer stock renewed device calibrated and checked | Assisted Replacement | SP | SP | Cal3
ABL90 FLEX | Ersatzgerät aufgestellt und konfiguriert, defektes Gerät mitgenommen zwecks Reparatur, Teil nicht zeitnah verfügbar | Fehlersuche, Austausch Sensorbay, Ersatz Lösungspack, Funktionskontrolle | SWO-83774 | Yes | Replacement device set up and configured, defective device taken with him for repair, part not available in a timely manner | Troubleshooting, Replacement Sensorbay, Replacement Solution Pack, Function Control | Re classified RMED | Double entry
ABL90 FLEX | 更换新测试卡修复仪器 | 更换新测试卡修复仪器 | SWO-81933 | No | Replace the new test card repair instrument | Replace the new test card repair instrument | Assisted Replacement | SC | SC | Lac
ABL825 FLEX | 更换LAC电极膜正常。 | 更换LAC电极膜正常。 | SWO-82110 | No | Replacing the LAC electrode film is normal. | Replacing the LAC electrode film is normal. | Assisted Replacement | Electrode | Membrane | Lac
ABL825 FLEX | 更换新测试卡修复仪器 | 更换新测试卡修复仪器 | SWO-80571 | Yes | Replace the new test card repair instrument | Replace the new test card repair instrument | Re classified RMED
ABL835 FLEX | Gerät geprüft, pCO2 Elektrode (E788 RN04) geprüft und ausgetauscht, Funktionskontrolle ok | Gerät geprüft, pCO2 Elektrode (E788 RN04) geprüft und ausgetauscht, Funktionskontrolle ok | SWO-80646 | No | Device tested, pCO2 electrode (E788 RN04) tested and replaced, function control ok | Device tested, pCO2 electrode (E788 RN04) tested and replaced, function control ok | Assisted Replacement | Electrode | Electrode | pCO2
ABL827 FLEX |  | B | D041891A2263A629 | No |  |  | Assisted Replacement | Electrode | Electrode | Crea
ABL827 FLEX | QC wiederholt i.O. | QC wiederholt i.O. | SWO-80687 | No | QC wiederholt i.O. | QC wiederholt i.O. | No fault found | HW | pH/BG Module | pH
ABL827 FLEX | El Met module Rinse Valve doesn't hardly any fluid in it. Need to change Rinse Valve Membranes in all Modules. | Replaced all Rinse Valve Membranes in all Modules. Also replaced Air Valve Membrane in pH Bg module. | D04163A12223AED9 | No | El Met module Rinse Valve doesn't hardly any fluid in it. Need to change Rinse Valve Membranes in all Modules. | Replaced all Rinse Valve Membranes in all Modules. Also replaced Air Valve Membrane in pH Bg module. | Defect | HW | El/Met Module | Valves | 731
AQT90 FLEX |  | B | D04180DA2241B1AC | No |  |  | Revisit required | Revisit required | Revisit required | Revisit required | 124
ABL80 FLEX BASIC | 更换新测试卡修复仪器 | 更换新测试卡修复仪器 | SWO-81941 | No | Replace the new test card repair instrument | Replace the new test card repair instrument
AQT90 FLEX | Given User Training to staff of ICU and Pathology and Biomedical Department. | Given User Training to staff of ICU and Pathology and Biomedical Department. | SWO-81128 | Yes | Given User Training to staff of ICU and Pathology and Biomedical Department. | Given User Training to staff of ICU and Pathology and Biomedical Department. | Re classified RMED | Training
ABL80 FLEX BASIC | 更换测试卡正常。 | 更换测试卡正常。 | SWO-82135 | No | The replacement test card is OK. | The replacement test card is OK.
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-80756 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL90 | Verificação dos registros, troca da cassete, reinicialização da cassete antiga, estabilização, teste com gasimetrias | B | D043D7490E306D40 | No | Checking records, replacing the cassette, restarting the old cassette, stabilizing, testing with gasimetries |  | Assisted Replacement | SP | SP | Glu | Sealing ring
ABL90 FLEX | 更换新试剂包后修复 | 更换新试剂包后修复 | SWO-82179 | No | Repair after replacing a new kit | Repair after replacing a new kit | Assisted Replacement | SP | SP | Pump cal error
ABL800 BASIC | Reinstalled the ifcom thermal printer. Checked driver & port for printer. Checked with new automatic & manual printing. Printer working OK. | Reinstalled the ifcom thermal printer. Checked driver & port for printer. Checked with new automatic & manual printing. Printer working OK. | SWO-80759 | No | Reinstalled the ifcom thermal printer. Checked driver & port for printer. Checked with new automatic & manual printing. Printer working OK. | Reinstalled the ifcom thermal printer. Checked driver & port for printer. Checked with new automatic & manual printing. Printer working OK. | Defect | SW | DMS_PC | Printer Driver
ABL80 FLEX Basic version |  | B | D04549138423DDA7 | No
ABL80 OSM | Examined data logs. Found expired sensor cassette and solution pack. Found tHB calibration expired. | Replaced sensor cassette and solution pack. Ran linearity VK kit. Performed tHB calibration. Verified operation of ABL80 OSM. | D04524A5844B3914 | No | Examined data logs. Found expired sensor cassette and solution pack. Found tHB calibration expired. | Replaced sensor cassette and solution pack. Ran linearity VK kit. Performed tHB calibration. Verified operation of ABL80 OSM.
ABL80 CO-OX | When they scan in the patient accession number it is putting a digit at the end of the ID. This is causing the analyzer not to query the demographics. Jim Halbleib sent me the barcodes to program the barcode reader. I worked with Jim and tech support and could not get rid of the extra digit. When they enter the accession manually it does pull overt eh demographics. We have a call in to Sendx for help. Received instructions from Sendx. | Scanned in the correct barcodes to reprogram the barcode reader to leave off the extra digit of the patient sample. | D0452E238432A961 | No | When they scan in the patient accession number it is putting a digit at the end of the ID. This is causing the analyzer not to query the demographics. Jim Halbleib sent me the barcodes to program the barcode reader. I worked with Jim and tech support and could not get rid of the extra digit. When they enter the accession manually it does pull overt eh demographics. We have a call in to Sendx for help. Received instructions from Sendx. | Scanned in the correct barcodes to reprogram the barcode reader to leave off the extra digit of the patient sample.
ABL90 FLEX | Vervanging van Sensor Cassette corrigeert de clothdetectie in oxi module. Opstart zonder problemen, heb wel geen goedgekeurde calibratie afgewacht. | Vervanging van Sensor Cassette corrigeert de clothdetectie in oxi module. Opstart zonder problemen, heb wel geen goedgekeurde calibratie afgewacht. | SWO-80913 | No | Replacement of Sensor Cassette corrects the cloth detection in Oxi module. Start without problems, did not wait for approved calibration. | Replacement of Sensor Cassette corrects the cloth detection in Oxi module. Start without problems, did not wait for approved calibration. | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
ABL90 FLEX | Einlasssonde und Einlassdichtung getauscht Analysator kalibriert und überprüft QC Messung durchgeführt | Einlasssonde und Einlassdichtung getauscht Analysator kalibriert und überprüft QC Messung durchgeführt | SWO-80896 | No | Inlet probe and inlet seal exchanged analyzer calibrated and checked QC measurement performed | Inlet probe and inlet seal exchanged analyzer calibrated and checked QC measurement performed | Defect | HW | Inlet Gasket | Inlet gasket
ABL90 FLEX | Bei Eintreffen eintreffen pCO2 gelb bei QC im Level , Level C wurde vom Kunden aus QC Plan entfernt, Abschaltung der fixen ZW-> LP ersetzt-> QC immer noch fehlerhaft im Level C, SC Ersatz heute fällig -> SC ersetzt -> keine Verbesserung, fixe ZW wieder zugeschaltet-> LP wieder getauscht um fixe ZW wieder zu aktivieren | Bei Eintreffen eintreffen pCO2 gelb bei QC im Level , Level C wurde vom Kunden aus QC Plan entfernt, Abschaltung der fixen ZW-> LP ersetzt-> QC immer noch fehlerhaft im Level C, SC Ersatz heute fällig -> SC ersetzt -> keine Verbesserung, fixe ZW wieder zugeschaltet-> LP wieder getauscht um fixe ZW wieder zu aktivieren | SWO-80794 | No | Upon arrival pCO2 yellow at QC in level , Level C was removed by the customer from QC plan, shutdown of the fixed ZW-> LP replaced-> QC still faulty in level C, SC replacement due today -> SC replaced -> no improvement, fixed ZW switched back on-changed-> LP replaced u m fixed ZW to be reactivated | Upon arrival pCO2 yellow at QC in level , Level C was removed by the customer from QC plan, shutdown of the fixed ZW-> LP replaced-> QC still faulty in level C, SC replacement due today -> SC replaced -> no improvement, fixed ZW switched back on-changed-> LP replaced u m fixed ZW to be reactivated | Assisted Replacement | SP | SP | QC Level C | pCO2
ABL90 | Reviewed data logs, confirm QC and calibration issue occurred on 2 6 2016 at 10:02AM. Also noted that the ABL90 ran another calibration at 10:05AM on the same day day and the problem was resolved, calibration passed, but the operator running QC on Saturday did not returned to check on the status of the QC (?s) after running it at 10:02am. Calibration cycles continued to run and pass without further incident after the issue at 10:02 on 2 6 2016. The operator running QC on Saturday did not call for assistance. | Suzie, called for assistance on Sunday. The on call FSE asked Suzie to rerun a calibration cycle and rerun level 3, all parameters now within range with no errors reported. No additional issues observed or noted while onsite 2 8 2016. | D04515B78443C079 | Yes | Reviewed data logs, confirm QC and calibration issue occurred on 2 6 2016 at 10:02AM. Also noted that the ABL90 ran another calibration at 10:05AM on the same day day and the problem was resolved, calibration passed, but the operator running QC on Saturday did not returned to check on the status of the QC (?s) after running it at 10:02am. Calibration cycles continued to run and pass without further incident after the issue at 10:02 on 2 6 2016. The operator running QC on Saturday did not call for assistance. | Suzie, called for assistance on Sunday. The on call FSE asked Suzie to rerun a calibration cycle and rerun level 3, all parameters now within range with no errors reported. No additional issues observed or noted while onsite 2 8 2016. | Re classified RMED | Tech Support
AQT90 FLEX |  | B | SWO-80793 | No |  |  | Defect | HW | Process Module | Process Unit sensor | 5784
ABL800 BASIC | cleaned the gasket and perform the protein removal. | cleaned the gasket and perform the protein removal. | SWO-81560 | No | cleaned the gasket and perform the protein removal. | cleaned the gasket and perform the protein removal. | Cleaned | Fluidic pathway | Inlet module | Gasket
ABL825 FLEX | 装置引上げの為、代替機設置・動作確認 | 装置分解・内部点検 各バルブ分解・清掃・潤滑 測定チャンバー交換・接続間チューブ交換・流体回路消毒実施・全メンブラン交換 CAL・QC測定データ確認 | SWO-83433 | No | Installation and operation confirmation of alternative machines for equipment pulling up | Device decomposition and internal inspection Each valve disassembly, cleaning, lubrication Measurement chamber replacement, connection-to-tube exchange, fluid circuit disinfection implementation, all membrane replacement CAL and QC measurement data confirmation | Cleaned | Fluidic pathway | - | -
ABL825 FLEX | Installed inlet module in loaner instrument s n 7 54R1099N007 for testing. | Installed inlet module in loaner instrument and ran several power cycles, service inlet calibrations, regular calibrations, QC. | D044F4B4844620C6 | No | Installed inlet module in loaner instrument s n 7 54R1099N007 for testing. | Installed inlet module in loaner instrument and ran several power cycles, service inlet calibrations, regular calibrations, QC. | Defect | HW | Inlet module | Inlet module
ABL825 FLEX | checked flow of solutions, checked fluidic module for leaks, checked tubing for leaks | removed fluidic moduel and cleaned salt deposits, removed and cleaned valves, replaced waste pump tubiong | D045A54F84453C89 | No | checked flow of solutions, checked fluidic module for leaks, checked tubing for leaks | removed fluidic moduel and cleaned salt deposits, removed and cleaned valves, replaced waste pump tubiong | Cleaned | Fluidic pathway | Fluidic manifold | Valve | High risc of Revisit, Clot in Waste tube
ABL820 FLEX | Pco2 bas surles 3 niveaux de controles | Sauvegarder les donn pour le techline remplacé pump tubing BG,tubing pre heater to top measuring chamber,new membrane | D04532D2840999E1 | No | Pco2 low on 3 control levels | Sauvegarder les donn pour le techline remplacé pump tubing BG,tubing pre heater to top measuring chamber,new membrane | Assisted Replacement | Electrode | Membrane | pCO2
ABL825 FLEX |  | B | D04527918449E02C | No |  |  | Assisted Replacement | Electrode | Membrane | Lac | Logistical upSWO
ABL827 FLEX | The solution would not flow through the analyzer and i found the OXI pinch valve to be stuck | I released the pinch valve and tested it and it was ok. | D045132D844C3A9A | No | The solution would not flow through the analyzer and i found the OXI pinch valve to be stuck | I released the pinch valve and tested it and it was ok. | Defect | HW | Oxi module | Pinch Valve
ABL825 FLEX | checked data logs for errors, checked QC logs, checked valves and liquid sensors | removed all electrodes in pH BG module and repalced all membranes, cleaned electrode chamber | D0459C1484613C7E | No | checked data logs for errors, checked QC logs, checked valves and liquid sensors | removed all electrodes in pH BG module and repalced all membranes, cleaned electrode chamber | Assisted Replacement | Electrode | Membrane | pCO2
ABL825 FLEX | Checked halogen lamp voltage found it to be reading high at 3820mV. | Adjusted halogen lamp voltage. | D0452FA0843550A0 | No | Checked halogen lamp voltage found it to be reading high at 3820mV. | Adjusted halogen lamp voltage. | Adjustment | HW | Oxi module | Lamp unit
ABL835 FLEX | Customer requested service visit. Inspected sample path, tested valves and performed isolation check. | Changed potassium membrane | D045516F844220C8 | No | Customer requested service visit. Inspected sample path, tested valves and performed isolation check. | Changed potassium membrane | Assisted Replacement | Electrode | Membrane | K+ | Kinked
ABL835 FLEX | module leaking from syringe port. All tubing correctly connected. Tried new waste pump tubing, noticed inlet probe was slightly bent and inlet gasket potentially clogged. Inspected sample path and tested valves. | Cleaned inlet nipples, replaced inlet probe and inlet gasket. | D04556EE842920CB | No | module leaking from syringe port. All tubing correctly connected. Tried new waste pump tubing, noticed inlet probe was slightly bent and inlet gasket potentially clogged. Inspected sample path and tested valves. | Cleaned inlet nipples, replaced inlet probe and inlet gasket. | Defect | HW | Inlet module | Inlet gasket
AQT90 FLEX | Need to replace spares. Barcode reader for display unit Inlet door, left Inlet door, right Stepper motor, Peristaltic pump Cable clip, self-adhesive Tube, Viton, ø0.9x2.7 90 mm (B B) Tube, Viton, ø0.9x2.7 120 mm (A A) Tubing connector, black Peristaltic Pump Head Valve 3 2 way with washer (w. screw and securing angle) Manifold Unit | B | SWO-81437 | No | Need to replace spares. Barcode reader for display unit Inlet door, left Inlet door, right Stepper motor, Peristaltic pump Cable clip, self-adhesive Tube, Viton, ø0.9x2.7 90 mm (B B) Tube, Viton, ø0.9x2.7 120 mm (A A) Tubing connector, black Peristaltic Pump Head Valve 3 2 way with washer (w. screw and securing angle) Manifold Unit |  | Defect | HW | Power Supply | Power Supply
AQT90 FLEX | Revision de mantenimiento. Limpieza interna del equipo. Limpieza de needle wash con hipoclorito. Sustiitucion de la placa incubator board. Realizacion de test de servicio. Qc y muestras. | B | D0451630844337BE | No | Maintenance review. Internal cleaning of the equipment. Cleaning of wash needle with hypochlorite. Replacement of the incubator board plate. Performing service tests. Qc and samples. |  | Defect | HW | Incubator module | PCB, Incubator
AQT90 FLEX |  | B | D04527A984348030 | No |  |  | Defect | HW | Incubator module | PCB, Incubator
AQT90 FLEX | Programmed LQC setup, run QC sample, results found ok. | Programmed LQC setup, run QC sample, results found ok. | SWO-81120 | Yes | Programmed LQC setup, run QC sample, results found ok. | Programmed LQC setup, run QC sample, results found ok. | Re classified RMED | Set Up LQC
ABL80 FLEX BASIC | SC SP交換 | SC SP交換 | SWO-81081 | No | SC SP Exchange | SC SP Exchange
ABL80 FLEX BASIC | Nullpunktkalibrierung mit Testkassette durchgeführt, Werte i.O., QC innerhalb des Bereichs aber erhöht, Kunde beobachtet weiteres Verhalten | B | SWO-81022 | No | Zero-point calibration performed with test cartridge, values i.O., QC within the range but increased, customer observes further behavior
ABL80 FLEX CO-OX | Einweisung in die tHb Kal. Durchführung durchgeführt. | Einweisung in die tHb Kal. Durchführung durchgeführt. | SWO-81089 | No | Instruction in the tHb Kal. carried out. | Instruction in the tHb Kal. carried out.
ABL90 FLEX | Sensorkassette erneuert Kalibration und QC durchgeführt | Sensorkassette erneuert Kalibration und QC durchgeführt | SWO-80970 | No | Sensor cassette renews calibration and QC performed | Sensor cassette renews calibration and QC performed | Assisted Replacement | SC | SC | Lac
ABL90 FLEX | Sensor wurde ersetzt, pH Fehler damit behoben. Gerät hat nun sporadisch pCO2 QC Verletzung, Ursache Zielwert der QC Lösung bzw. Temperatur, Kunden entsprechend informiert. Zur Zeit 18,2 mmHg | Sensor wurde ersetzt, pH Fehler damit behoben. Gerät hat nun sporadisch pCO2 QC Verletzung, Ursache Zielwert der QC Lösung bzw. Temperatur, Kunden entsprechend informiert. Zur Zeit 18,2 mmHg | SWO-80985 | No | Sensor has been replaced, pH error has been fixed with it. Device has now sporadically informed pCO2 QC violation, cause target value of QC solution or temperature, customers accordingly. Currently 18.2 mmHg | Sensor has been replaced, pH error has been fixed with it. Device has now sporadically informed pCO2 QC violation, cause target value of QC solution or temperature, customers accordingly. Currently 18.2 mmHg | Assisted Replacement | SC | SC | pH
ABL90 FLEX PLUS | Überprüfung Einlassdichtung => Dichtungsrand ausgefranzt (keine neue Dichtung!) => Austausch Dichtung Spülung des Systems und Überprüfung LP => kein Auffälligkeiten Alle Serviceprogramme der Fließwege ohne Beanstandung Startphase i.O. pCO2 Empfindlichkeit sehr tief => SC war seit 08.02. 0:00 Uhr ohne Durchfluss => Empfindlichkeit steigt langsam Empfindlichkeit momentan bei 24,6 % => Gerät einlaufen lassen, ggf. SC tauschen SC und LP beim Kunden belassen, da kein Material vor Ort | Überprüfung Einlassdichtung => Dichtungsrand ausgefranzt (keine neue Dichtung!) => Austausch Dichtung Spülung des Systems und Überprüfung LP => kein Auffälligkeiten Alle Serviceprogramme der Fließwege ohne Beanstandung Startphase i.O. pCO2 Empfindlichkeit sehr tief => SC war seit 08.02. 0:00 Uhr ohne Durchfluss => Empfindlichkeit steigt langsam Empfindlichkeit momentan bei 24,6 % => Gerät einlaufen lassen, ggf. SC tauschen SC und LP beim Kunden belassen, da kein Material vor Ort | SWO-80973 | No | Check inlet seal => Sealing edge flogged (no new seal!) => Replacement seal flushing of the system and check LP => no abnormalities All service programs of the flow paths without complaint Start phase i.O. pCO2 Sensitivity very deep => SC was since 08.02. 0:00 o'clock without flow => Sensitivity slowly increases sensitivity currently at 24.6% => Allow device to enter, possibly sc exchange SC and LP at the customer, as no material on site | Check inlet seal => Sealing edge flogged (no new seal!) => Replacement seal flushing of the system and check LP => no abnormalities All service programs of the flow paths without complaint Start phase i.O. pCO2 Sensitivity very deep => SC was since 08.02. 0:00 o'clock without flow => Sensitivity slowly increases sensitivity currently at 24.6% => Allow device to enter, possibly sc exchange SC and LP at the customer, as no material on site | Defect | HW | Auto Inlet module | Inlet Gasket
ABL90 FLEX PLUS | Einlasssonde leicht verbogen. Einlasssonde erneuert. Zuvor eingesetzten Lösungspack verwendet und Einlaufverfahren gestartet. Alle Funktionen und QC o.k. Eine Einlasssonde zur Reserve an Med. Technik übergeben. | Einlasssonde leicht verbogen. Einlasssonde erneuert. Zuvor eingesetzten Lösungspack verwendet und Einlaufverfahren gestartet. Alle Funktionen und QC o.k. Eine Einlasssonde zur Reserve an Med. Technik übergeben. | SWO-81090 | No | Inlet probe slightly bent. Inlet probe renewed. Previously used solution pack and started inlet procedure. All functions and QC ok. An inlet probe for the reserve is handed over to Med. Technik. | Inlet probe slightly bent. Inlet probe renewed. Previously used solution pack and started inlet procedure. All functions and QC ok. An inlet probe for the reserve is handed over to Med. Technik. | Replaced | Probe | Bend | Bend
ABL837 FLEX | Check the instrument found printer head having problem, Cleaned the instrument and printer head but problem is remaining same. Need to replace with new waiting for the part. | Check the instrument found printer head having problem, Cleaned the instrument and printer head but problem is remaining same. Need to replace with new waiting for the part. | SWO-81313 | No | Check the instrument found printer head having problem, Cleaned the instrument and printer head but problem is remaining same. Need to replace with new waiting for the part. | Check the instrument found printer head having problem, Cleaned the instrument and printer head but problem is remaining same. Need to replace with new waiting for the part. | Defect | HW | Printer module | Printer head
ABL825 FLEX | Oxi Modul demontiert und geprüft, Hemolyzer justiert, Funktionskontrolle ok | Oxi Modul demontiert und geprüft, Hemolyzer justiert, Funktionskontrolle ok | SWO-81015 | No | Oxi module disassembled and tested, hemolyzer adjusted, function control ok | Oxi module disassembled and tested, hemolyzer adjusted, function control ok | Revisit required | HW | Oxi module | Hemolyzer
ABL825 FLEX |  | B | D04A350F5C0DB7A0 | No |  |  | Assisted Replacement | Electrode | Electrode | Reference
ABL835 FLEX | エアーライン詰まりの為、洗浄。LS調整実施。 | エアーライン詰まりの為、洗浄。LS調整実施。 | SWO-80875 | No | Cleaning because of air line clogging. Ls adjustment scan. | Cleaning because of air line clogging. Ls adjustment scan. | Cleaned | Fluidic pathway | Fluidic manifold | Air vent
ABL835 FLEX | PH BGモジュールの測定チャンバー洗浄。PCO2電辱洗浄、PCO2メンブラン交換。 | PH BGモジュールの測定チャンバー洗浄。PCO2電辱洗浄、PCO2メンブラン交換。 | SWO-81078 | No | Measurement chamber cleaning of PH BG module. PCO2 densifyment cleaning, PCO2 membrane replacement. | Measurement chamber cleaning of PH BG module. PCO2 densifyment cleaning, PCO2 membrane replacement. | Cleaned | Fluidic pathway | pH/BG Module | ph/BG Module
ABL800 BASIC | replace inlet gasket | B | SWO-81765 | No | replace inlet gasket |  | Assisted Replacement | HW | Inlet module | Inlet gasket
ABL825 FLEX | Den Kunden im Kenntnissgesetzt dass die Abdeckung des Karussels nicht geöffnet werden soll um die Anzahl der Ampullen zu kontrollieren. pCO2 Elektrode und Dichtring der PH- Elektrode ausgetauscht. Neustart. | Angeblich sind die Ampullen wie vorgeschrieben ersetzt worden und nicht einfach die Kappe hoch gezogen. QC wiederholt alles wieder in Ordnung | SWO-81026 | No | The customer is aware that the cover of the carousel should not be opened in order to control the number of ampoules. pCO2 electrode and sealing ring of the PH electrode. Restart. | Allegedly, the ampoules have been replaced as prescribed and not simply pulled up the cap. QC repeats everything in order | Assisted Replacement | Electrode | Electrode | pCO2
ABL825 FLEX |  | Ich habe Sie die Umgebungsdruck um 19 mbar erhöhen lassen und die QC wiederholt. Alles wieder in Ordnung | SWO-81060 | No |  | I have had you increase the ambient pressure by 19 mbar and repeated the QC. All right again | Assisted Replacement | Electrode | Electrode | pCO2
ABL800 BASIC | cleaned the pH electrode,done cleaning and 2 point calibration all results found with in the range. | cleaned the pH electrode,done cleaning and 2 point calibration all results found with in the range. | SWO-81776 | No | cleaned the pH electrode,done cleaning and 2 point calibration all results found with in the range. | cleaned the pH electrode,done cleaning and 2 point calibration all results found with in the range. | Cleaned | Fluidic pathway | pH/BG Module | pH
ABL800 BASIC | Calibrated touch screen temporarily working | Calibrated touch screen temporarily working | SWO-81769 | No | Calibrated touch screen temporarily working | Calibrated touch screen temporarily working | Revisit required | HW | DMS_PC | Touch Screen
ABL80 FLEX CO-OX | Austausch CO-OX Control Board Einlauf- Und Initialisierungsprozedur | Austausch CO-OX Control Board Einlauf- Und Initialisierungsprozedur | SWO-81221 | No | Replacement CO-OX Control Board inlet and initialization procedure | Replacement CO-OX Control Board inlet and initialization procedure
ABL80 FLEX BASIC | 连续更换三张测试卡后解决 | 连续更换三张测试卡后解决 | SWO-81094 | No | Resolve after replacing three test cards in a row | Resolve after replacing three test cards in a row
ABL80 FLEX BASIC | 更换新测试卡后解决 | 更换新测试卡后解决 | SWO-81310 | No | Resolve after replacing a new test card | Resolve after replacing a new test card
ABL80 FLEX BASIC | 连续更换三张测试卡后修复 | 连续更换三张测试卡后修复 | SWO-82584 | No | Repair after replacing three test cards in a row | Repair after replacing three test cards in a row
ABL80 FLEX | Na resetten van de SP gegevens en het weghalen van de O-ring kan deze opnieuw geplaatst worden zonder problemen. De defecte SP (lotnr 219092) werd meegenomen om te claimen naar Radiometer. | Na resetten van de SP gegevens en het weghalen van de O-ring kan deze opnieuw geplaatst worden zonder problemen. De defecte SP (lotnr 219092) werd meegenomen om te claimen naar Radiometer. | SWO-81243 | No | After resetting the SP data and removing the O-ring, it can be repositioned without any problems. The defective SP (Lotnr 219092) was taken to claim to radiometer. | After resetting the SP data and removing the O-ring, it can be repositioned without any problems. The defective SP (Lotnr 219092) was taken to claim to radiometer.
ABL80 FLEX BASIC | re connected the Sensor luer tubing and installed the sensor cassette now instrument working. | re connected the Sensor luer tubing and installed the sensor cassette now instrument working. | SWO-81771 | No | re connected the Sensor luer tubing and installed the sensor cassette now instrument working. | re connected the Sensor luer tubing and installed the sensor cassette now instrument working.
ABL90 |  | B | D0477C27705BD5E3 | Yes |  |  | Re classified RMED | Dummy SWO
ABL90 | S9050 pO2 59.3 (4.0-29.5). SP SV-20 and SC 743-202 | Replaced the inlet probe. Internal QC-ok. External QC-out. I flushed the analyzer 6 times. Internal QC-out. Replaced the sensor bay. Internal QC-in. External QC-out. Swapped the SC with the mobile unit. Internal QC-in External QC-in. | D047CDF5703F1C8C | No | S9050 pO2 59.3 (4.0-29.5). SP SV-20 and SC 743-202 | Replaced the inlet probe. Internal QC-ok. External QC-out. I flushed the analyzer 6 times. Internal QC-out. Replaced the sensor bay. Internal QC-in. External QC-out. Swapped the SC with the mobile unit. Internal QC-in External QC-in. | Defect | SC | wetability | wetability
ABL90 | Vacuum pressure error reported | FSE replaced tubing, swapped Solution Pack no change. Swapped Service Cassette, fluid flow issue resolved. | D04F3DDB3407053C | No | Vacuum pressure error reported | FSE replaced tubing, swapped Solution Pack no change. Swapped Service Cassette, fluid flow issue resolved. | Assisted Replacement | SC | SC | Wetability | 131
ABL825 FLEX |  | B | D047918D7018BCEC | No |  |  | Adjustment | HW | AutoCheck Module | Carroussel
ABL825 FLEX | checked flow of solutions, checked fluidic module for leaks, checked tubing for leaks | removed and replaced fluidic channel unit, | D0482443701A3C9A | No | checked flow of solutions, checked fluidic module for leaks, checked tubing for leaks | removed and replaced fluidic channel unit, | Defect | HW | Fluidic manifold | Fluidic manifold
ABL800 BASIC | cleaned the po2 chamber & performed 2 point calibration | B | SWO-81746 | No | cleaned the po2 chamber & performed 2 point calibration |  | Cleaned | Fluidic pathway | pH/BG Module | pO2
ABL827 FLEX | Customer tried to attach tubes, no luck, came off again | Installed new pump tubes and y connector | D047C28270207C31 | No | Customer tried to attach tubes, no luck, came off again | Installed new pump tubes and y connector | Defect | HW | Fluidic manifold | Waste tubings
ABL835 FLEX | Performed oxi test. Halogen value at 4095, cavitation absorbance and drive current failing. | Performed lamp adjustment. complete oxi test. | D047F71270362DA9 | No | Performed oxi test. Halogen value at 4095, cavitation absorbance and drive current failing. | Performed lamp adjustment. Complete oxy test. | Adjustment | HW | Oxi module | Lamp unit
ABL80 FLEX CO-OX | Anbindung Conworks entsprechend eingerichtet Patienten-Barcode wird nicht erkannt | Anbindung Conworks entsprechend eingerichtet Patienten-Barcode wird nicht erkannt | SWO-81439 | No | Connection Conworks set up accordingly Patient barcode is not recognized | Connection Conworks set up accordingly Patient barcode is not recognized
ABL90 |  | B | D04CCD4F4804015B | No |  |  | Assisted Replacement | SP | SP | SP
ABL90 FLEX | Lösungspack erneuert, Initialisierungsphase abgewartet, Kals und QCs durchgeführt. | B | SWO-81559 | No | Solution pack renewed, initialization phase awaited, Kals and QCs performed. |  | Assisted Replacement | SP | SP | Leak Current
ABL90 FLEX | Gerät neugestartet, Bootfehler tritt nach Geräte neustart nicht mehr auf. Kunde wird sich bei weiteren Problemen melden. Einlassdichtung durch Kundeneigene Einlassdichtung erneuert, Sensor Kassette aufgrund von pCO2 Kal Fehler durch Kundeneigene SC erneuert. | Gerät neugestartet, Bootfehler tritt nach Geräte neustart nicht mehr auf. Kunde wird sich bei weiteren Problemen melden. Einlassdichtung durch Kundeneigene Einlassdichtung erneuert, Sensor Kassette aufgrund von pCO2 Kal Fehler durch Kundeneigene SC erneuert. | SWO-81541 | No | Device restarts, boot error no longer occurs after device restart. Customer will contact you if you have any further problems. Inlet seal renewed by customer's own inlet seal, sensor cassette renewed due to pCO2 Kal fault by customer's own SC. | Device restarts, boot error no longer occurs after device restart. Customer will contact you if you have any further problems. Inlet seal renewed by customer's own inlet seal, sensor cassette renewed due to pCO2 Kal fault by customer's own SC. | Defect | SW | Reboot failure | Reboot failure
ABL90 FLEX | Telefonische Fehlerbehebung während Anfahrt. Lösungspack war nicht korrekt aktiviert und klemmte. Frau Daniels hat es soweit wieder hinbekommen | Telefonische Fehlerbehebung während Anfahrt. Lösungspack war nicht korrekt aktiviert und klemmte. Frau Daniels hat es soweit wieder hinbekommen | SWO-81445 | No | Telephone troubleshooting while driving. Solution pack was not activated correctly and jammed. Ms. Daniels got it back | Telephone troubleshooting while driving. Solution pack was not activated correctly and jammed. Ms. Daniels got it back | User error | SP Activation | SP Activation | SP Activation
ABL90 FLEX | Ventilschlauch erneuert, brachte keine Besserung. SC erneuert. Keine Besserung. Ersatzteil (Interface Board) nicht im Bestand. Herr Pfister übernimmt Austausch am 12.02.2016. | Ventilschlauch erneuert, brachte keine Besserung. SC erneuert. Keine Besserung. Ersatzteil (Interface Board) nicht im Bestand. Herr Pfister übernimmt Austausch am 12.02.2016. | SWO-81526 | No | valve hose, brought no improvement. SC renewed. No improvement. spare part (interface board) is not in stock. Mr. Pfister takes over exchange on 12.02.2016. | valve hose, brought no improvement. SC renewed. No improvement. spare part (interface board) is not in stock. Mr. Pfister takes over exchange on 12.02.2016. | Unknown | Revisit required
ABL90 | Desligar e ligar o equipamento sem sucesso, resinstalação do software com a versão 3.1MR6.PT e carregar configuração, calibração e passagem do QC. | B | D04F643C3458461F | Yes | Disconnect and connect the equipment without success, resinstall the software with the 3.1MR6.PT version and load configuration, calibration and passage of QC. |  | Re classified RMED | SWU
AQT90 FLEX | Verificação dos registros, teste da fluidica, desmontagem do analisador, verificação dos tubos, detecção de tubo trilhado, estabilização e passagem do QC | B | D048F8BFE607EF41 | No | Check records, Fluidic test, analyzer disassembly, check tubes, tube detection, stabilization and QC passage |  | Cleaned | HW | Aspiration Module | Tubes
ABL825 FLEX | Oxi Modul geprüft, Ersatzmodul installiert, tHb Kalibration, Funktionskontrolle ok | Hr. Küster bittet um einen Rückruf und Service. | SWO-81523 | No | Oxi module tested, replacement module installed, tHb calibration, function control ok | Mr. Küster is asking for a recall and service. | Defect | HW | Oxi module | Hemolyzer
ABL827 FLEX | CreaModul ersetzt (alt: 902-842R200N004 <-> neu: 902-842R273N003), Software installiert, alle Elektroden membraniert, Startphase, Probelauf, Kal & QK ok, KundenCreaModul zur Werkstattreparatur mitgenommen | B | SWO-81890 | No | CreaModule replaces (old: 902-842R200N004 new: 902-842R273N003), software installed, all electrodes membraneed, start phase, test run, Kal & QK ok, customerCreamodule for workshop repair |  | Defect | HW | Met II Module | Met II module
ABL825 FLEX | Einstellungen für Uhrzeit im Control Panel überprüft Timeserver Einstellung ok, Haken für automatische Sommer-Winterzeitumstellung nicht gesetzt Gerät richtig konfiguriert und überprüft | Einstellungen für Uhrzeit im Control Panel überprüft Timeserver Einstellung ok, Haken für automatische Sommer-Winterzeitumstellung nicht gesetzt Gerät richtig konfiguriert und überprüft | SWO-81446 | No | Settings for time in the control panel checks time server setting ok, hook for automatic summer winter time change not set device properly configured and checked | Settings for time in the control panel checks time server setting ok, hook for automatic summer winter time change not set device properly configured and checked | Setup | SW | System clock | System clock
ABL825 FLEX | Sämtliche Pumpenschläuche ersetzt. Messkammer überprüft und ersetzt, jedoch ohne Erfolg. (Original wieder eingebaut) Alle Pumpenrotoren ersetzt. Mehrere Messungen im Kapillarmodus durchgeführt, alle i.o. Kal., QC und System i.o. Funktionskontrolle i.o. | Sämtliche Pumpenschläuche ersetzt. Messkammer überprüft und ersetzt, jedoch ohne Erfolg. (Original wieder eingebaut) Alle Pumpenrotoren ersetzt. Mehrere Messungen im Kapillarmodus durchgeführt, alle i.o. Kal., QC und System i.o. Funktionskontrolle i.o. | SWO-81564 | No | Replaced all pump hoses. Measuring chamber checked and replaced, but to no avail. (original reinstalled) Replaced all pump rotors. Several measurements performed in capillary mode, all i.o. Cal., QC and System i.o. Function control i.o. | Replaced all pump hoses. Measuring chamber checked and replaced, but to no avail. (original reinstalled) Replaced all pump rotors. Several measurements performed in capillary mode, all i.o. Cal., QC and System i.o. Function control i.o. | Defect | HW | El/Met Module | Pump tubing
ABL827 FLEX | Pumpenschlauch am MET II-Modul erneuert, System und Funktion i.O. | Pumpenschlauch am MET II-Modul erneuert, System und Funktion i.O. | SWO-81461 | No | Pump hose on met II module renewed, system and function in the same way | Pump hose on met II module renewed, system and function in the same way | Defect | HW | Met II Module | Pump tubing
ABL815 FLEX | Lampe abgeglichen, Gerät kalibriert, Service tHb Kal durchgeführt, tHb Kal durchgef., Leistungstest durchgeführt, QCs durchgef. | Der Kunde bittet um einen Service und Rückruf. | SWO-81907 | No | Lamp matched, device calibrated, service tHb Kal performed, tHb Kal carried out, performance test performed, QCs carried out. | The customer asks for a service and callback. | Adjustment | HW | Oxi module | Lamp unit
ABL80 FLEX CO-OX OSM | Fehlersuche, tHb Kal. fehlgeschlagen, Schläuche zum Teil defekt eingeklemmt->Jahreswartungssatz verbaut, Hemolyzer SN902-R0126N006 defekt->erneuert gegen 902-R0613N003, Neustart und Oxi Complete Test ok, tHb Kal. und QA Zyklus ok, alle Parameter wieder aktiv System-und Funktionskontrolle i.O. | Fehlersuche, tHb Kal. fehlgeschlagen, Schläuche zum Teil defekt eingeklemmt->Jahreswartungssatz verbaut, Hemolyzer SN902-R0126N006 defekt->erneuert gegen 902-R0613N003, Neustart und Oxi Complete Test ok, tHb Kal. und QA Zyklus ok, alle Parameter wieder aktiv System-und Funktionskontrolle i.O. | SWO-81724 | No | Troubleshooting, tHb Kal. failed, hoses partially defectively clamped->annual maintenance kit installed, Hemolyzer SN902-R0126N006 defective->renewed against 902-R0613N003, restart and Oxi Complete Test ok, tHb Kal. and QA cycle ok, all parameters again active system and function control i.O. | Troubleshooting, tHb Kal. failed, hoses partially defectively clamped->annual maintenance kit installed, Hemolyzer SN902-R0126N006 defective->renewed against 902-R0613N003, restart and Oxi Complete Test ok, tHb Kal. and QA cycle ok, all parameters again active system and function control i.O.
ABL80 FLEX CO-OX | Gerät abgeholt und Leihgerät wieder aufgebaut, 306354Sr.Nr. | Kundin möchte keine weitere Hilfe durch die Hotline Der zuständige Techniker möge sich doch bitte bei ihr melden | SWO-81660 | No | device and borrower rebuilt, 306354Sr.Nr. | Customer does not want any further help from the hotline The responsible technician should please contact her
ABL80 CO-OX | Verified the problem. The sensor cassette would not pass GLU on the initialization. | Replaced the tubing and pumps in the unit, the sensor cassette, and the solution pack. Ran a linearity kit to prove the unit is working properly. Allow GLU to be turned off so that the sensor cassette would be initialized. | D04A52655C01B5C0 | No | Verified the problem. The sensor cassette would not pass GLU on the initialization. | Replaced the tubing and pumps in the unit, the sensor cassette, and the solution pack. Ran a linearity kit to prove the unit is working properly. Allow GLU to be turned off so that the sensor cassette would be initialized.
ABL80 OSM | tried reinstalling sensor, tried flushing through waste tried to reboot no help, inspected tubings and found clots and build up in tubings | replaced sensor cassette and solution pack; replaced annual maintenance kit; ANNAUL MAINTENANCE COMPLETE WILL UPDATE BAS PM DATE | D04A1B295C02371F | No | tried reinstalling sensor; tried flushing through waste tried to reboot no help; inspected tubings and found clots and build up in tubings | replaced sensor cassette and solution pack; replaced annual maintenance kit; ANNAUL MAINTENANCE COMPLETE WILL UPDATE BAS PM DATE
ABL80 OSM | No investigation needed due to on going problem. | Inspected clamp on waste tubing for correct position. | D049F6055C0E9EF9 | No | No investigation needed due to on going problem. | Inspected clamp on waste tubing for correct position.
ABL90 FLEX | PC Modul für Montag 8:00 bestellt Gerät mit in die GS genommen - alles mit Frau Schöpcke besprochen Gerät sollte Montag nachmittag zurückkommen | PC getauscht Gerät temperiert und Startup abgewartet | SWO-81678 | No | PC module for Monday 8:00 ordered device taken into the GS - all discussed with Ms. Schöpcke device should come back Monday afternoon | PC swaps device tempered and startup waited | Defect | HW | CPU | CPU
ABL90 | Pump calibration error. Change in SP and SC does not eliminate error. | Cleaned all tubing - Pump cal is good. | D04A462C5C4971B3 | No | Pump calibration error. Change in SP and SC does not eliminate error. | Cleaned all tubing - Pump cal is good. | Assisted Replacement | SP | SP | Pump cal error
ABL90 FLEX | Sensor Interface PCB getauscht weiterhin Level3 laut RiliBÄK zu niedrig mit Frau Messerschmidt besprochen Level3 vorübergehend abzuschalten Gerät kalibriert und überprüft | Sensor Interface PCB getauscht weiterhin Level3 laut RiliBÄK zu niedrig mit Frau Messerschmidt besprochen Level3 vorübergehend abzuschalten Gerät kalibriert und überprüft | SWO-81663 | No | Sensor Interface PCB continues to swap Level3 according to RiliBÄK too low with Ms. Messerschmidt discussed Level3 temporarily switching off device calibrated and checked | Sensor Interface PCB continues to swap Level3 according to RiliBÄK too low with Ms. Messerschmidt discussed Level3 temporarily switching off device calibrated and checked | Defect | HW | Basic Unit | Sensor Interface PCB
ABL90 | Lactate slope on the linearity calibration is just above Customer acceptable range of .9 - 1.1. I came in on 2 11 the SC was in for more then 24 hours values did not change Replaced with another SC SP ran some waste blood Calibrated for about 2 hours Values were just about the same for Lactate. Calibrations Internal External QC are just fine. SC are different lot number R0750-120 750-172 750-186, but the third SP was a different number, but it failed with Flow selector error . | 2 12 Ran the Sensor Simulator The test passed on all channels including Lactate. Open Techline case to investigate further. | D04A1B765C03FA2C | No | Lactate slope on the linearity calibration is just above Customer acceptable range of .9 - 1.1. I came in on 2 11 the SC was in for more then 24 hours values did not change Replaced with another SC SP ran some waste blood Calibrated for about 2 hours Values were just about the same for Lactate. Calibrations Internal External QC are just fine. SC are different lot number R0750-120 750-172 750-186, but the third SP was a different number, but it failed with Flow selector error . | 2 12 Ran the Sensor Simulator The test passed on all channels including Lactate. Open Techline case to investigate further. | Unknown | SC | Lac | Lac
ABL90 FLEX PLUS | Toch Screen Kalibration nicht möglich Toch Screen lässt sich im oberen Bereich nicht bedienen Toch Screen erneuert und kalibriert Gerät überprüft | Toch Screen Kalibration nicht möglich Toch Screen lässt sich im oberen Bereich nicht bedienen Toch Screen erneuert und kalibriert Gerät überprüft | SWO-81694 | No | Toch Screen Calibration Not Possible Toch Screen can not be operated in the upper range Toch Screen renewed and calibrated device checked | Toch Screen Calibration Not Possible Toch Screen can not be operated in the upper range Toch Screen renewed and calibrated device checked | Defect | HW | DMS_PC | Touch Screen
ABL835 FLEX | Changed inlet nippel | Changed inlet nippel | SWO-81677 | No | Changed inlet nipples | Changed inlet nipples | Defect | HW | Inlet module | Leaking | Nipple kit
ABL825 FLEX | Allgemeine Kontrolle ok. Diverse Schläuche wieder getauscht ok. Mit Dual valve housing getestet ok. Dual Valve Housing inkl. Adhesive Pad ersetzt ok. Diverse QC und Blutmessungen durchgeführt ok. Blutmessungen ok, ab und zu wieder Lac Instabil während QC Messung. *** Leihweise für 2 Wochen Lac Elektrode zusenden *** | B | SWO-81928 | No | General control ok. Various hoses replaced ok. Tested with dual valve housing ok. Dual Valve Housing incl. Adhesive Pad replaced ok. Various QC and blood measurements performed ok. Blood measurements ok, from time to time Lac unstable during QC measurement. Send Lac Electrode on loan for 2 weeks *** |  | Unknown | Electrode | Membrane | Lac
ABL825 FLEX | Fehlende untere Lac-Elektrodenkappe ersetzt, aufgef. Hb-Lampe ersetzt, Kal, QC | Fehlende untere Lac-Elektrodenkappe ersetzt, aufgef. Hb-Lampe ersetzt, Kal, QC | SWO-81740 | No | Missing lower lac electrode cap replaced, re-enactd. Hb lamp replaced, Kal, QC | Missing lower lac electrode cap replaced, re-enactd. Hb lamp replaced, Kal, QC | Assisted Replacement | Electrode | Membrane | Lac
ABL835 FLEX | Reboot and swap carrousels-No change Changes 14 pole cable-No change, carrousel won't spin Order parts and return tomorrow | Replaced carroust spin motor and PCB Program AU software | D04A11C75C3C37A0 | No | Reboot and swap carousels-No change Changes 14 pole cable-No change, carousel won't spin Order parts and return tomorrow | Replaced carroust spin motor and PCB Program AU software | Revisit required | HW | AutoCheck Module | FPC Cable
TCM-COMBIM | Unable to reproduce the reported problem. | B | D049E1D85C6074D6 | No | Unable to reproduce the reported problem.
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-81770 | No | Repair after replacing a new test card | Repair after replacing a new test card
AQT90 FLEX |  | B | SWO-81773 | No |  |  | Defect | HW | Process Module | Process Unit sensor | 5298
AQT90 FLEX |  | B | SWO-81757 | No |  |  | Defect | HW | Storage Module | PCB, Storage | 5691
AQT90 FLEX |  | B | SWO-81753 | No |  |  | Assisted Replacement | Consumables | Test Kit | NT-PROBNP
AQT90 FLEX | Nadel ersetzt und kalibriert, Fehler während der Nadelkalibrierung im ProcessWheel, Sensor defekt, Sensor für Positionserkennung unten oben ersetzt, Neustart, Nadel erneut kalibriert, Neustart abgewartet, Puffemessungen, QC | Nadel ersetzt und kalibriert, Fehler während der Nadelkalibrierung im ProcessWheel, Sensor defekt, Sensor für Positionserkennung unten oben ersetzt, Neustart, Nadel erneut kalibriert, Neustart abgewartet, Puffemessungen, QC | SWO-81848 | No | Needle replaced and calibrated, error during needle calibration in ProcessWheel, sensor defective, sensor for position detection top replaced, restart, needle recalibrated, restart waited, puff measurements, QC | Needle replaced and calibrated, error during needle calibration in ProcessWheel, sensor defective, sensor for position detection top replaced, restart, needle recalibrated, restart waited, puff measurements, QC | Defect | HW | Aspiration Module | Bend needle | Pipette_38
ABL80 FLEX BASIC | 更换新测试卡修复仪器 | 更换新测试卡修复仪器 | SWO-81826 | No | Replace the new test card repair instrument | Replace the new test card repair instrument
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-81869 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL80 FLEX BASIC | 连续更换两张测试卡后修复 | 连续更换两张测试卡后修复 | SWO-82108 | No | Repair after replacing two test cards in a row | Repair after replacing two test cards in a row
ABL90 FLEX | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-81898 | No | Repair after replacing a new test card | Repair after replacing a new test card | Assisted Replacement | SC | SC | Lac
ABL80 FLEX CO-OX | SC (218468) eingebaut und (217903) ausgebaut, System und Funktion i.O. | Parameter können nicht aktiviert werden, die Felder sind inaktiv. Fr. Egelsehr führt einen Neustart durch, wenn notwendig wird sie die Parameter aktivieren. Bei weiteren Fragen wird sie sich wieder melden. | SWO-82112 | No | SC (218468) built-in and (217903) removed, system and function in the | Parameters cannot be activated, the fields are inactive. Ms. Egelsehr performs a reboot if necessary, she will activate the parameters. She will contact me again if she has any further questions.
ABL80 FLEX CO-OX | Fehlersuche, Austausch Flüssigkeitssensor, Funktionskontrolle | Fehlersuche, Austausch Flüssigkeitssensor, Funktionskontrolle | SWO-83784 | No | Troubleshooting, replacement of liquid sensor, function control | Troubleshooting, replacement of liquid sensor, function control
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-82562 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL80 FLEX BASIC |  | 更换打印机主板后修复 | SWO-81861 | No |  | Repair after replacing the printer board
ABL80 CO-OX | Trouble shot the ABL80 and found that the tubing leading to the waste pump was kinked. | Adjusted the tubing to allow proper flow and flushed it manuallyt to make sure that it flows properly. | D04DD24DBE1BF362 | No | Trouble shot the ABL80 and found that the tubing leading to the waste pump was kinked. | Adjusted the tubing to allow proper flow and flushed it manuallyt to make sure that it flows properly.
ABL80 FLEX | Nazicht van de SP en SC : niets abnormaals kunnen vaststellen. Omdat bij deze fout er een probleem is met een van beiden werd de SC opnieuw geïnstalleerd : de verificatie duurde redelijk lang en er werd vaak gevraagd om de luchtbellen in de tubing te verwijderen. Hierna gaf hij terug de bovenstaande fout. Hierop besloten om een nieuwe SC te installeren : deze starte op zonder problemen. Waarschijnlijk zit er een verstopping in de originele SC : deze werd meegenomen. Systeemcontrole en kalibratie : OK | B | SWO-81948 | No | Review of the SP and SC: Nothing can determine abnormal. Because with this error there is a problem with either the SC was reinstalled: The verification took quite a long time and it was often asked to remove the air bubbles in the tubing. After this he gave back the above error. This decided to install a new SC: This started on without any problems. Probably there is a blockage in the original SC: it was taken. System Check and calibration: OK
ABL80 CO-OX | Verified the problem. | Replaced the printer, printer control board and the printer cable. Printer is now functional. Customer agreeded to monitor this overnight to see if the problem returns. | D04E4486BE10E2E8 | No | Verified the problem. | Replaced the printer, printer control board and the printer cable. Printer is now functional. Customer agreeded to monitor this overnight to see if the problem returns.
ABL80 FLEX CO-OX | iButton vom LP konnte nicht ausgelesen werden, deshalb keine 2Pkt. Kal, daraufhin wurden die tHb Parameter deaktiviert Parameter wieder aktiviert Gerät resetet und Oxi Test durchgeführt | iButton vom LP konnte nicht ausgelesen werden, deshalb keine 2Pkt. Kal, daraufhin wurden die tHb Parameter deaktiviert Parameter wieder aktiviert Gerät resetet und Oxi Test durchgeführt | SWO-81991 | No | iButton from the LP could not be read, therefore no 2pkt. Kal, then the tHb parameters were deactivated parameters reactivated device resets and Oxi test performed | iButton from the LP could not be read, therefore no 2pkt. Kal, then the tHb parameters were deactivated parameters reactivated device resets and Oxi test performed
ABL90 |  | B | D05BF58ED03ABD27 | No |  |  | Assisted Replacement | SC | SC | Reference
ABL90 FLEX | PC getauscht Gerät temperiert und Startup abgewartet | PC getauscht Gerät temperiert und Startup abgewartet | SWO-81959 | No | PC swaps device tempered and startup waited | PC swaps device tempered and startup waited | Defect | HW | CPU | CPU
ABL90 FLEX | SC 90 erneuert ABL nach Reparatur in GS beim Kunden zurück getauscht Einlaufverfahren | der Anwender spült die Flüssigkeitsstrecke jeden Tag durch, allerdings funktioniert das Gerät nicht dauerhaft; es tropft nur aus dem Einlass, kein kontinuierlicher Wasserstrahl; der Anwender wünscht einen Vorort-Service zur Beseitigung der Geräte-Störung; | SWO-81909 | No | SC 90 renews ABL after repair in GS at customer exchanged inlet procedure | the user rinses the liquid section every day, but the device does not work permanently; it drips only from the inlet, no continuous water jet; the user wants a suburban service to eliminate the device malfunction; | Cleaned | Fluidic pathway | Inlet Gasket | Inlet gasket
ABL90 |  | B | D04E14EBBE626D6A | No |  |  | Defect | HW | Basic Unit | Wet sec PCB
AQT90 FLEX |  | B | SWO-81858 | No |  |  | Defect | HW | Process Module | Optical Unit
AQT90 FLEX |  | B | SWO-81954 | No |  |  | Defect | HW | Process Module | Process Unit sensor
AQT90 FLEX |  | B | SWO-81863 | No |  |  | Defect | HW | Incubator module | Sensor, Incubator
ABL835 FLEX |  | B | D04F9EC7340E6425 | No |  |  | Defect | HW | Inlet module | Leaking | Nipple kit
ABL837 FLEX | Čišćenje linije otpada i zamjena crijeva | B | SWO-82041 | No | Cleaning the line of waste and replacing the hose |  | Cleaned | Fluidic pathway | Inlet Module | Tubes
ABL815 FLEX | Mehrfache Kalibrationen waren nötig bist die Parameter wieder in Ordnung waren | B | SWO-82071 | No | Multiple calibrations were necessary the parameters were back in order |  | Assisted Replacement | Electrode | Electrode | pCO2
ABL825 FLEX | Mehrere Kalibrationen durchgeführt. | Mehrere Kalibrationen durchgeführt. | SWO-82070 | No | Several calibrations performed. | Several calibrations performed. | Assisted Replacement | Electrode | Membrane | Lac
ABL815 FLEX | Funktion überprüft, Luftdruck OK, Tropfenbildung vor der Referenzmembran. E1001 mit D711 Lot 1900 membraniert, E788 gereinigt und membraniert. Software V6.16 installiert, QC Einstellungen der Software entsprechend geändert. Funktion z.Zt. OK. | Funktion überprüft, Luftdruck OK, Tropfenbildung vor der Referenzmembran. E1001 mit D711 Lot 1900 membraniert, E788 gereinigt und membraniert. Software V6.16 installiert, QC Einstellungen der Software entsprechend geändert. Funktion z.Zt. OK. | SWO-82113 | No | Function checked, air pressure OK, drop formation in front of the reference membrane. E1001 with D711 Lot 1900 membrane, E788 cleaned and membraneed. Software V6.16 installed, QC settings of the software changed accordingly. Function currently OK. | Function checked, air pressure OK, drop formation in front of the reference membrane. E1001 with D711 Lot 1900 membrane, E788 cleaned and membraneed. Software V6.16 installed, QC settings of the software changed accordingly. Function currently OK. | Assisted Replacement | Electrode | Electrode | pCO2
AQT90 FLEX | Storage PCB und Lichtschranke ersetzt>> Fehler besteht nach wie vor. Schrittmotor ersetzt. Startphase QC und Kalibration. | Storage PCB und Lichtschranke ersetzt>> Fehler besteht nach wie vor. Schrittmotor ersetzt. Startphase QC und Kalibration. | SWO-81971 | No | Replace storage PCB and light barrier>> Error still exists. stepper motor replaced. Start phase QC and calibration. | Replace storage PCB and light barrier>> Error still exists. stepper motor replaced. Start phase QC and calibration. | Defect | HW | Storage Module | Stepper motor
ABL80 FLEX BASIC | 更换新测试卡修复仪器 | 更换新测试卡修复仪器 | SWO-83177 | No | Replace the new test card repair instrument | Replace the new test card repair instrument
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-82600 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-83149 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL80 FLEX BASIC | 连续更换两张测试卡后修复 | 连续更换两张测试卡后修复 | SWO-82283 | No | Repair after replacing two test cards in a row | Repair after replacing two test cards in a row
ABL80 FLEX BASIC |  | B | SWO-82151 | No
ABL80 OSM | Trouble shot the unit and found that the tube on top of he hemolyzer was bent too much. | Adjusted the plastic shield on top of the pump to allow the fluid to flow properly. | D04F1718340800FB | No | Trouble shot the unit and found that the tube on top of he hemolyzer was bent too much. | Adjusted the plastic shield on top of the pump to allow the fluid to flow properly.
ABL80 OSM | Trouble shot the unit and found the the waste connector nipple was not screwed in allthe way. | Tightend the nipple in the waste connector. | D04F19E434270100 | No | Trouble shot the unit and found the the waste connector nipple was not screwed in allthe way. | Tightend the nipple in the waste connector.
ABL80 FLEX BASIC | 更换测试卡正常。 | 更换测试卡正常。 | SWO-82106 | No | The replacement test card is OK. | The replacement test card is OK.
ABL90 | Screen locked up. Unit failing pump cal and pressure value test. | Re booted analyzer to unlock display. Replaced listed items. Performed transport tests in service. | D04F4BD63446948C | No | Screen locked up. Unit failing pump cal and pressure value test. | Re booted analyzer to unlock display. Replaced listed items. Performed transport tests in service. | Defect | HW | Inlet Gasket | Inelt Gasket
ABL90 FLEX PLUS | Toush screen erneuert, Gerät neu gestartet, Zeit eingestellt, Kal und QCs durchgeführt. | B | SWO-82257 | No | Toush screen renewed, device restarted, time set, Kal and QCs performed. |  | Defect | HW | DMS_PC | Touch Screen
ABL90 FLEX | 更换测试卡，试剂包正常。 | 更换测试卡，试剂包正常。 | SWO-82282 | No | Replace the test card and the reagent package is OK. | Replace the test card and the reagent package is OK. | Assisted Replacement | SC | SC | pCO2
ABL90 FLEX | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-82573 | No | Repair after replacing a new test card | Repair after replacing a new test card | Assisted Replacement | SC | SC | pCO2 | 1316
AQT90 FLEX |  | B | SWO-82127 | No |  |  | Defect | HW | Aspiration Module | Needle Wash Unit | 1223
AQT90 FLEX |  | B | SWO-82103 | No |  |  | Defect | HW | Storage Module | Barcode reader, Storage wheel | BR1
AQT90 FLEX |  | B | SWO-82137 | No |  |  | Defect | HW | Process Module | Process Unit sensor | 5784
ABL800 BASIC | Found calibration results are ok, only sample measurment found pH parameter was query, clenaed the pH electrode with protein removal calibrated and performed the sample measurment but problem is reamining same need to replace with new. | Found calibration results are ok, only sample measurment found pH parameter was query, clenaed the pH electrode with protein removal calibrated and performed the sample measurment but problem is reamining same need to replace with new. | SWO-82444 | No | Found calibration results are ok, only sample measurment found pH parameter was query, clenaed the pH electrode with protein removal calibrated and performed the sample measurment but problem is reamining same need to replace with new. | Found calibration results are ok, only sample measurment found pH parameter was query, clenaed the pH electrode with protein removal calibrated and performed the sample measurment but problem is reamining same need to replace with new. | Revisit required | Electrode | Electrode | pH
ABL825 FLEX | Gerät demontiert, Festplatte ersetzt, Software neu installiert, Konfiguration neu geladen, Funktionskontrolle ok | Gerät demontiert, Festplatte ersetzt, Software neu installiert, Konfiguration neu geladen, Funktionskontrolle ok | SWO-82422 | No | Device dismantled, hard drive replaced, software reinstalled, configuration reloaded, function control ok | Device dismantled, hard drive replaced, software reinstalled, configuration reloaded, function control ok | Defect | HW | DMS_PC | HDD
ABL825 FLEX | 装置分解・内部点検 各バルブ分解・清掃・潤滑 測定チャンバー交換・接続間チューブ交換・流体回路消毒実施・全メンブラン交換 CAL・QC測定データ確認 | 装置分解・内部点検 各バルブ分解・清掃・潤滑 測定チャンバー交換・接続間チューブ交換・流体回路消毒実施・全メンブラン交換 CAL・QC測定データ確認 | SWO-83435 | No | Device decomposition and internal inspection Each valve disassembly, cleaning, lubrication Measurement chamber replacement, connection-to-tube exchange, fluid circuit disinfection implementation, all membrane replacement CAL and QC measurement data confirmation | Device decomposition and internal inspection Each valve disassembly, cleaning, lubrication Measurement chamber replacement, connection-to-tube exchange, fluid circuit disinfection implementation, all membrane replacement CAL and QC measurement data confirmation | Defect | HW | pH/BG Module | Measuring chamber
ABL837 FLEX |  | B | D04F6FD334056BBF | No |  |  | Defect | HW | Inlet module | Probe bend
ABL835 FLEX | Gerät überprüft, Meßkammer und Luftauslass gesäubert, pCO2 und diverse Elektroden membraniert, Kal und QC o.k. | B | SWO-82309 | No | Device checked, measuring chamber and air outlet cleaned, pCO2 and various electrodes membraneed, Kal and QC ok. |  | Assisted Replacement | Electrode | Membrane | pCO2
ABL825 FLEX | Im pH BG und El Met Modulen die Ventile ausgebaut und leicht geölt. Mehrere Kalibrationen durchgeführt>> I.O | Im pH BG und El Met Modulen die Ventile ausgebaut und leicht geölt. Mehrere Kalibrationen durchgeführt>> I.O | SWO-82462 | No | In the pH BG and El Met modules, the valves are removed and lightly oiled. Several calibrations performed>> I.O | In the pH BG and El Met modules, the valves are removed and lightly oiled. Several calibrations performed>> I.O | Defect | HW | pH/BG Module | Valve, air
AQT90 FLEX |  | B | D04F34E73406EEF6 | No |  |  | Unknown | Analyzer replaced | Analyzer replaced | Analyzer replaced
ABL80 FLEX BASIC |  | B | SWO-82280 | No
ABL80 |  | B | D0509F8DAA526BC7 | No
ABL80 FLEX BASIC | Found that Sensor cassette got clotted with blood, Cleaned the Sensor cassette reinstalled the SC found installation was OK instrument is functional. | Found that Sensor cassette got clotted with blood, Cleaned the Sensor cassette reinstalled the SC found installation was OK instrument is functional. | SWO-82446 | No | Found that Sensor cassette got clotted with blood, Cleaned the Sensor cassette reinstalled the SC found installation was OK instrument is functional. | Found that Sensor cassette got clotted with blood, Cleaned the Sensor cassette reinstalled the SC found installation was OK instrument is functional.
ABL90 FLEX | Konfiguration angepasst nach Kundenwunsch | Konfiguration angepasst nach Kundenwunsch | SWO-82378 | Yes | Configuration adapted to customer requirements | Configuration adapted to customer requirements | Re classified RMED | Setup by Customer request
ABL90 | Suspect silicone contamination. | Flush system with hot water. Discuss silicone issue with customer. | D0506CAEAA277D61 | No | Suspect silicone contamination. | Flush system with hot water. Discuss silicone issue with customer. | Cleaned | Fluidic pathway | Sampler/Silicone | Sampler/Silicone
ABL90 |  | B | D0508254AA5E6BC5 | No |  |  | Defect | HW | Inlet Gasket | Inelt Gasket
AQT90 FLEX | 清理进样盘，重新安装进样盘，重新安装信号排线 | 清理进样盘，重新安装进样盘，重新安装信号排线 | SWO-82331 | No | Clean the sample disk, reinstall the sample disk, reinstall the signal line | Clean the sample disk, reinstall the sample disk, reinstall the signal line | Defect | HW | Aspiration Module | Inlet Wheel | 436
AQT90 FLEX | 更换新试剂包后修复 | 更换新试剂包后修复 | SWO-82369 | No | Repair after replacing a new kit | Repair after replacing a new kit | Defect | HW | SP | SP | 1999
AQT90 FLEX | Comprobar rueda de procesos y mecanismos. Reiniciar el equipo y pasar una muestra. | No encontrado problema. | D05051EBAA59E1E8 | No | Check process wheel and mechanisms. Restart the computer and pass a sample. | No problem found. | Power off/on | HW | Process Module | Process wheel | 5684
ABL815 FLEX | Gasversorgung im Serviceprogramm überprüft, leere Gasflasche durch Anwender am Vortag ersetzt ausgetauscht. Software 6.16 aufgrund von merkwürdigem Verhalten installiert, AU Sybase wird nicht installiert.AU Board ersetzt, Keys installiert. QC gemäß Software 6.16 konfiguriert Funktion OK. Nachtrag: Nach ca 1 h meldet sich der Kunde erneut. ABL trennt seit dem 15.2. ca. 17:00 Uhr sporadisch DMS von AU. mit dem Anwender Aufstellung eines Leihgerätes und weitere Überprüfung in der Werkstatt besprochen. | Kunde meldet sich heute erneut, aber mit einem anderen Fehler. (siehe SWO von 10:42Uhr) Das Gas2 wurde mehrfach ersetzt und wird nicht als voll angezeigt. Der Kunde versichert, dass es sich um eine neue Gas2-Flasche handelt. Er wünscht einen Vorortservice. | SWO-82571 | No | Gas supply checked in the service program, empty gas bottle replaced by users replaced the day before. Software 6.16 installed due to strange behavior, AU Sybase will not be installed. AU board replaced, keys installed. QC according to software 6.16 configured function OK. Addendum: After approx. 1 h, the customer reports again. ABL has been separating DMS from AU sporadically since 15.2. approx. 17:00. discussed with the user installation of a loaner and further inspection in the workshop. | Customer reports again today, but with a different error. (see SWO from 10:42 a.m.) The gas2 has been replaced several times and is not displayed as full. The customer assures that this is a new Gas2 bottle. He wants a suburban service. | Unknown | HW | Gas module | Gas 2
ABL825 FLEX | Screen intermittently freezes during sample introduction. Customer reboots analyzer to restore screen. | Installed 6.16 software. Follow-up with customer on status. | D0508ED1AA0CE5F3 | No | Screen intermittently freezes during sample introduction. Customer reboots analyzer to restore screen. | Installed 6.16 software. Follow-up with customer on status. | Unknown | His/Lis
ABL825 FLEX | Screen intermittently freezes during sample introduction. Customer re boots analyzer in order to restore display. | Installed 6.16 software. Follow-up with customer on status | D05083A9AA16E5F0 | No | Screen intermittently freezes during sample introduction. Customer re boots analyzer in order to restore display. | Installed 6.16 software. Follow-up with customer on status | Unknown | His/Lis
ABL825 FLEX | Friction detector skiftet | Friction detector skiftet | SWO-82715 | No | Friction detector skiftet | Friction detector skiftet | Defect | HW | Inlet module | Friction detector unit
AQT90 FLEX | ﾆｰﾄﾞﾙｳｫｯｼｭﾕﾆｯﾄ交換 | B | SWO-83106 | No | Needle wash unit replacement |  | Defect | HW | Aspiration Module | Needle Wash Unit
AQT90 FLEX | Comprobación del estado del analizador. Se observa buen funcionamiento de la Power Supply, se procede a colocar teclado al equipo sin respuesta del mismo, se reinicia equipo sin CF pero el equipo siempre se queda en la misma pantalla. Se pide material. Se deja el equipo parado. | Se determina fallo de la CPU | D051B384202146B7 | No | Checking the status of the scanner. The Power Supply is properly functioning, the keyboard is placed without a response from it, the computer is restarted without CF but the computer always stays on the same screen. Material is requested. He's leaving the equipment standing. | CPU failure is determined | Revisit required | HW | Analyzer Control Module | CPU
AQT90 FLEX |  | B | D050D255AA405D8A | No |  |  | Defect | HW | Cover and Chassis | Cover
ABL80 FLEX BASIC | 更换新测试卡修复仪器 | 更换新测试卡修复仪器 | SWO-82583 | No | Replace the new test card repair instrument | Replace the new test card repair instrument
ABL80 FLEX BASIC | 更换新测试卡修复仪器 | 更换新测试卡修复仪器 | SWO-82615 | No | Replace the new test card repair instrument | Replace the new test card repair instrument
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-82668 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL80 FLEX BASIC | 更换试剂包正常。 | 更换试剂包正常。 | SWO-83062 | No | The replacement kit is OK. | The replacement kit is OK.
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-82789 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL80 FLEX BASIC | 更换测试卡正常 。 | 更换测试卡正常 。 | SWO-82936 | No | Replacing the test card is OK. | Replacing the test card is OK.
ABL80 FLEX BASIC |  | B | SWO-82643 | No
ABL80 FLEX BASIC | 更换新测试卡修复仪器 | 更换新测试卡修复仪器 | SWO-82612 | No | Replace the new test card repair instrument | Replace the new test card repair instrument
ABL80 FLEX BASIC | done setting in calibration schedule | done setting in calibration schedule | SWO-83078 | No | done setting in calibration schedule | done setting in calibration schedule
ABL90 FLEX | Garantie CF-Card getauscht, Fehler bleibt bestehen--> CPU-Unit erneuert Neustart --> i.O | der Anwender hat den Analysator schon aus- und wieder eingeschaltet, neu gestartet, Fehler bleibt bestehen; vermutlich CF Card oder CPU Unit defekt; der Anwender wünscht einen Vorort-Service zur Beseitigung der Geräte-Störung; | SWO-82635 | No | Guarantee CF-Card exchanged, error remains--> CPU unit renewed Restart --> i.O | the user has already switched the analyzer off and on again, restarted, error remains; probably CF Card or CPU Unit defective; the user wants a suburban service to eliminate the device malfunction; | Defect | HW | CPU | CPU
ABL90 FLEX | 更换试剂包正常。 | 更换试剂包正常。 | SWO-83142 | No | The replacement kit is OK. | The replacement kit is OK. | Assisted Replacement | SP | SP | SP
AQT90 FLEX |  | B | SWO-82574 | No |  |  | Defect | HW | Analyzer Control Module | Barcode reader, display
AQT90 FLEX |  | B | SWO-82593 | No |  |  | Defect | HW | Process Module | Optical Unit
AQT90 FLEX |  | B | SWO-82559 | No |  |  | Defect | HW | Process Module | PCB, Cup dryer
ABL835 FLEX | Comprobar disco duro y fuente de alimentación, cambiar la fuente de alimentación, calibrar y probar con controles. Revisión normal del equipo con limpieza de la cámara de medida y salida del desperdicio, realización de operaciones de mantenimiento y prueba con controles. | Cambiar la fuente de alimentación. | D051AE31202F036F | No | Check hard drive and power supply, change power supply, calibrate and test with controls. Normal equipment check with measuring chamber cleaning and waste output, maintenance operations and testing with controls. | Change the power supply. | Defect | HW | Power Supply | Power Supply
ABL825 FLEX | Defekten Printer ersetzt, Kal und QK ok | Defekten Printer ersetzt, Kal und QK ok | SWO-82854 | No | Defective Printer Replaced, Kal and QK ok | Defective Printer Replaced, Kal and QK ok | Defect | HW | DMS_PC | Printer head
ABL825 FLEX | FlexQ überprüft - keine Fehler festgestellt, Kira gelöscht u. Vers. 6.13 installiert, Daten eingegeben, FlexQ kalibriert, tHb-Kal durchgeführt, Kal u. QC i.O. Probenmessung i.O. | FlexQ überprüft - keine Fehler festgestellt, Kira gelöscht u. Vers. 6.13 installiert, Daten eingegeben, FlexQ kalibriert, tHb-Kal durchgeführt, Kal u. QC i.O. Probenmessung i.O. | SWO-82762 | No | FlexQ checked - no errors detected, Kira deleted & verse. 6.13 installed, data entered, FlexQ calibrated, tHb-Kal performed, Kal and QC i.O. sample measurement in the order of | FlexQ checked - no errors detected, Kira deleted & verse. 6.13 installed, data entered, FlexQ calibrated, tHb-Kal performed, Kal and QC i.O. sample measurement in the order of | Adjustment | HW | Oxi module | tHb cal
ABL815 FLEX | ABL800 754R1147N001 beim Kunden meßbereit aufgestellt, Kundengerät zur weiteren Überprüfung mitgenommen. LIS KIS Key muss nachinstalliert werden. Funktion OK. | Kunde meldet sich heute erneut, aber mit einem anderen Fehler. (siehe SWO von 10:42Uhr) Das Gas2 wurde mehrfach ersetzt und wird nicht als voll angezeigt. Der Kunde versichert, dass es sich um eine neue Gas2-Flasche handelt. Er wünscht einen Vorortservice. | SWO-82851 | No | ABL800 754R1147N001 set up at the customer's place ready for measurement, customer device taken with you for further inspection. LIS KIS Key must be reinstalled. Function OK. | Customer reports again today, but with a different error. (see SWO from 10:42 a.m.) The gas2 has been replaced several times and is not displayed as full. The customer assures that this is a new Gas2 bottle. He wants a suburban service. | Unknown | HW | Gas module | Gas 2
ABL825 FLEX | Sonde ersetzt. Paar Ampullen auf einer anderen Station ausgeliehen, die Lieferung der AQC Ampullen Level 2 soll heute noch erfolgen. Startphase QC und Kalibration. | Telefonat mit Frau Barnick: Die Anwender seien extrem genervt, weil der ABL seit Lieferung immer wieder und auch in letzter Zeit fortwährend Probleme bereite. Anwenderseitig wird man keine Aktionen mehr durchführen und erwartet einen Service. Ich habe Barnick informiert, dass Bahman krank ist, daher ein Service heute unwahrscheinlich ist und wahrscheinlich morgen erst jemand kommen kann. | SWO-82712 | No | probe replaced. Pair of ampoules borrowed from another station, the delivery of the AQC Ampoules Level 2 is to take place today. Start phase QC and calibration. | Phone call with Ms. Barnick: The users are extremely annoyed because the ABL has been causing problems again and again since delivery and also recently. On the user side, you will no longer perform any actions and expect a service. I have informed Barnick that Bahman is ill, so a service is unlikely today and probably someone will not be able to come tomorrow. | Defect | HW | Inlet module | Probe bend
ABL800 BASIC | Keine Maßnahmen erforderlich. | Keine Maßnahmen erforderlich. | SWO-82677 | Yes | No action required. | No action required. | Re classified RMED | Demo key exp
ABL825 FLEX | checked data logs for errors, checked electrodes and electrode chamber, pH chamber filled with blood, checked valves and liquid sensors | removed electrodes for pH module, cleaned pH electrode and electrode chamber, replaced pH seal ring, replaced electrode membranes, removed modules and cleaned valves, replaced tubing as needed, | D05240B6202F3CDC | No | checked data logs for errors, checked electrodes and electrode chamber, pH chamber filled with blood, checked valves and liquid sensors | removed electrodes for pH module, cleaned pH electrode and electrode chamber, replaced pH seal ring, replaced electrode membranes, removed modules and cleaned valves, replaced tubing as needed, | Defect | Electrode | Electrode | pH
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-83974 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL80 FLEX BASIC | 更换测试卡正常。 | 更换测试卡正常。 | SWO-83063 | No | The replacement test card is OK. | The replacement test card is OK.
ABL90 FLEX | Neues Power Supply eingebaut Wartungsarbeiten durchgeführt | Neues Power Supply eingebaut Wartungsarbeiten durchgeführt | SWO-82924 | No | New Power Supply built-in maintenance work carried out | New Power Supply built-in maintenance work carried out | Defect | HW | Power Supply | Power Supply
ABL90 | Failing in diags.Cleared erros in Diags. Unit still unstable in Calibrations or QC. | Perform MR6 7 upgrade. | D052C0F69611894A | No | Failing in diags. Cleared erros in Diags. Unit still unstable in Calibrations or QC. | Perform MR6 7 upgrade. | Assisted Replacement | SP | SP | Pump cal error
ABL90 FLEX PLUS | Controle van het aantal kalibraties op ABL S4, S5 en S6 vertonen geen afwijkingen. Zodoende geen andere actie uitgevoerd. | Controle van het aantal kalibraties op ABL S4, S5 en S6 vertonen geen afwijkingen. Zodoende geen andere actie uitgevoerd. | SWO-82989 | No | Checking the number of calibrations on ABL S4, S5 and S6 do not show any deviations. No other action performed. | Checking the number of calibrations on ABL S4, S5 and S6 do not show any deviations. No other action performed. | No fault found | -
ABL90 FLEX PLUS |  | B | D05359FC96458C15 | No |  |  | Cleaned | Fluidic pathway
AQT90 FLEX |  | B | SWO-82864 | No |  |  | Defect | HW | Storage Module | Storage wheel
AQT90 FLEX | 更换光学单元903-364修复(by andrew) | 更换光学单元903-364修复(by andrew) | SWO-82353 | No | Replacement optical unit 903-364 repair (by andrew) | Replacement optical unit 903-364 repair (by andrew) | Defect | HW | Process Module | Optical Unit | 129
AQT90 FLEX |  | B | SWO-82835 | No |  |  | Defect | HW | Process Module | Optical Unit | 131
ABL825 FLEX | pCO2 elektrode met electrode cleaner behandeld. | pCO2 elektrode met electrode cleaner behandeld. | SWO-82986 | No | pCO2 probe with electrode cleaner treated. | pCO2 probe with electrode cleaner treated. | Assisted Replacement | Electrode | Electrode | pCO2
ABL815 FLEX | ABL ans Zentrallager geliefert. | Gasversorgung im Serviceprogramm überprüft, leere Gasflasche durch Anwender am Vortag ersetzt ausgetauscht. Software 6.16 aufgrund von merkwürdigem Verhalten installiert, AU Sybase wird nicht installiert.AU Board ersetzt, Keys installiert. QC gemäß Software 6.16 konfiguriert Funktion OK. Nachtrag: Nach ca 1 h meldet sich der Kunde erneut. ABL trennt seit dem 15.2. ca. 17:00 Uhr sporadisch DMS von AU. mit dem Anwender Aufstellung eines Leihgerätes und weitere Überprüfung in der Werkstatt besprochen. | SWO-82959 | No | ABL to the central warehouse. | Gas supply checked in the service program, empty gas bottle replaced by users replaced the day before. Software 6.16 installed due to strange behavior, AU Sybase will not be installed. AU board replaced, keys installed. QC according to software 6.16 configured function OK. Addendum: After approx. 1 h, the customer reports again. ABL has been separating DMS from AU sporadically since 15.2. approx. 17:00. discussed with the user installation of a loaner and further inspection in the workshop. | Unknown | HW | Gas module | ?
ABL800 BASIC | Reinigung pCO2 membraniert | Reinigung pCO2 membraniert | SWO-82964 | No | Cleaning pCO2 membraneed | Cleaning pCO2 membraneed | Assisted Replacement | Electrode | Membrane | pCO2
ABL825 FLEX | Kal 1--2 gleiches Lot ersetzt. Mehrfach Kal. QC wieder ok. Betroffenes Lot UK -01 | Kal 1--2 gleiches Lot ersetzt. Mehrfach Kal. QC wieder ok. Betroffenes Lot UK -01 | SWO-83006 | No | Kal 1--2 equal lot replaced. Multiple Kal. QC ok again. Affected Lot UK -01 | Kal 1--2 equal lot replaced. Multiple Kal. QC ok again. Affected Lot UK -01 | Assisted Replacement | Solution bottle | Kal1 & Kal2 | Kal1 & Kal2
ABL825 FLEX | Inlet skiftet | Inlet skiftet | SWO-83163 | No | Inlet shift | Inlet shift | Defect | HW | Inlet module | Inlet module
ABL825 FLEX | Neue Glu. Elektrode eingebaut , Lac. Und Glu. Membraniert. Mehrfach Kalibriert. | Neue Glu. Elektrode eingebaut , Lac. Und Glu. Membraniert. Mehrfach Kalibriert. | SWO-82941 | No | New Glu. Electrode built-in, Lac. And Glu. Membraneed. Multiple calibrated. | New Glu. Electrode built-in, Lac. And Glu. Membraneed. Multiple calibrated. | Assisted Replacement | Electrode | Electrode | Glu
ABL837 FLEX |  | B | D0531582962D210A | No |  |  | Assisted Replacement | HW | El/Met Module | Pump tubing
ABL815 FLEX | Fließwege gereinigt, Messkammer gereinigt, Elektrodenspitze gereinigt und Elektrode neu bezogen, QC und Probemessungen durchgeführt, | Fließwege gereinigt, Messkammer gereinigt, Elektrodenspitze gereinigt und Elektrode neu bezogen, QC und Probemessungen durchgeführt, | SWO-83011 | No | Flow paths cleaned, measuring chamber cleaned, electrode tip cleaned and electrode re-sourced, QC and sample measurements carried out, | Flow paths cleaned, measuring chamber cleaned, electrode tip cleaned and electrode re-sourced, QC and sample measurements carried out, | Assisted Replacement | Electrode | Membrane | pCO2
ABL80 FLEX BASIC | 更换新测试卡修复仪器 | 更换新测试卡修复仪器 | SWO-83181 | No | Replace the new test card repair instrument | Replace the new test card repair instrument
AQT90 FLEX | 更换903-119，636-550各一个正常。(by zhenjie) | 更换903-119，636-550各一个正常。(by zhenjie) | SWO-83138 | No | Replace the 903-119, 636-550 each normal. (by zhenjie) | Replace the 903-119, 636-550 each normal. (by zhenjie) | Defect | HW | Storage Module | Process Unit sensor | 5257
ABL90 FLEX | 更换903-129修复仪器(by zhenjie) | 更换903-129修复仪器(by zhenjie) | SWO-83084 | No | Replace the 903-129 repair instrument (by zhenjie) | Replace the 903-129 repair instrument (by zhenjie) | Defect | HW | Oxi module | Pinch valve
ABL90 FLEX | Gerät überprüft, Dichtungen und Lösungspack kontrolliert, Fließwege gespült, Kal und QC o.k., tHb Kal durchgeführt | B | SWO-83364 | No | Device checked, seals and solution pack checked, flow paths rinsed, Kal and QC ok, tHb Kal performed |  | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
ABL90 FLEX | durch Sichtkontrolle defekten Oxi Schlauch entdeckt Jahreswartung durchgeführt - siehe nachfolgenden SWO | durch Sichtkontrolle defekten Oxi Schlauch entdeckt Jahreswartung durchgeführt - siehe nachfolgenden SWO | SWO-83247 | No | due to visual inspection defective Oxi hose detected annual maintenance carried out - see following SWO | due to visual inspection defective Oxi hose detected annual maintenance carried out - see following SWO | Defect | HW | Oxi module | Mismatch err
ABL90 | Failing Pressure Test Vacuum too high value.Techline suggested Inlet Module for failing Pressure test. | Installed Maintenance kit and Pinch valve.Also performed software upgrade to Ver. MR7. | D056B638F863D3B8 | No | Failing Pressure Test Vacuum too high value. Techline suggested Inlet Module for failing Pressure test. | Installed Maintenance kit and Pinch valve. Also performed software upgrade to Ver. MR7. | Defect | HW | Oxi module | Pinch valve
ABL90 FLEX | LP zur überbrückung geliefert, | LP zur überbrückung geliefert, | SWO-83527 | Yes | LP supplied for bridging, | LP supplied for bridging, | Re classified RMED | Logistical upSWO
ABL90 | Examined activity log. Examined solution pack. | Replaced solution pack, as part of troubleshooting. New solution pack soved the issue. Verified operation of ABL90. | D0570F2FF8076005 | No | Examined activity log. Examined solution pack. | Replaced solution pack, as part of troubleshooting. New solution pack soved the issue. Verified operation of ABL90. | Assisted Replacement | SP | SP | Flow selector cal
ABL90 FLEX PLUS |  | B | D056CF33F811A707 | No |  |  | Adjustment | HW | Oxi module | pinch valve tube
ABL90 FLEX PLUS | Einlass demontiert und überprüft Einlass montiert, justiert Sensorkassette und Lösungspack installiert Kal- und QC | Einlass demontiert und überprüft Einlass montiert, justiert Sensorkassette und Lösungspack installiert Kal- und QC | SWO-83347 | No | Inlet disassembled and checked inlet mounted, adjusted sensor cassette and solution pack installed Cal and QC | Inlet disassembled and checked inlet mounted, adjusted sensor cassette and solution pack installed Cal and QC | Adjustment | HW | SC Door | SC Door
ABL90 FLEX | 更换两个试剂包修复仪器 | 更换两个试剂包修复仪器 | SWO-83275 | No | Replace two reagent kit repair instruments | Replace two reagent kit repair instruments | Assisted Replacement | SP | SP | Flow selector cal
AQT90 FLEX | 更换903-210一个正常。(By zhenjie) | 更换903-210一个正常。(By zhenjie) | SWO-83086 | No | Replace the 903-210 one normal. (By zhenjie) | Replace the 903-210 one normal. (By zhenjie) | Defect | HW | Storage Module | PCB, Cup dryer | LQC
AQT90 FLEX | 更换903-171修复仪器(by zhenjie )error code 4794 | 更换903-171修复仪器(by zhenjie )error code 4794 | SWO-83221 | No | Replace the 903-171 repair instrument (by-zhenjie) error code 4794 | Replace the 903-171 repair instrument (by-zhenjie) error code 4794 | Defect | HW | Aspiration Module | Bend needle | 4794
ABL825 FLEX |  | B | SWO-83190 | No |  |  | Assisted Replacement | Electrode | Electrode | pH
ABL825 FLEX | Checked all logs to verify past performance and found PH errors. Analyzer has been sitting idle for the past several months and customer is ready to go live. | Checked the electrode and cleaned the measuring chamber and verified the PH errors still exist. Replaced the PH Electrode and ran calibrations to verify. Analyzer required several calibrations to pass drift. CL out level three QC, remembraned electrode and retested. Electrode is covered under warranty. | D056AAD7F847DEE0 | No | Checked all logs to verify past performance and found PH errors. Analyzer has been sitting idle for the past several months and customer is ready to go live. | Checked the electrode and cleaned the measuring chamber and verified the PH errors still exist. Replaced the PH Electrode and ran calibrations to verify. Analyzer required several calibrations to pass drift. CL out level three QC, remembraned electrode and retested. Electrode is covered under warranty. | User error | Shut down
ABL815 FLEX | QC bei so2 Level 1 knapp unterhalb des Grenzwertes, Kal i.o bei THB 580, tHb Kal bei 0,9766, Oxi complete Test i.O., Service THB Kal - tHb Kal durchgeführt, Kal und QC durchgeführt, QC bei so2 mittig, Anwender beobachtet und meldet sich bei weitern Problemen | QC bei so2 Level 1 knapp unterhalb des Grenzwertes, Kal i.o bei THB 580, tHb Kal bei 0,9766, Oxi complete Test i.O., Service THB Kal - tHb Kal durchgeführt, Kal und QC durchgeführt, QC bei so2 mittig, Anwender beobachtet und meldet sich bei weitern Problemen | SWO-83355 | No | QC at so2 Level 1 just below the limit value, Kal i.o at THB 580, tHb Kal at 0.9766, Oxi complete Test i.O., Service THB Kal - tHb Kal performed, Kal and QC performed at so2 in the middle, users observed and reports further problems | QC at so2 Level 1 just below the limit value, Kal i.o at THB 580, tHb Kal at 0.9766, Oxi complete Test i.O., Service THB Kal - tHb Kal performed, Kal and QC performed at so2 in the middle, users observed and reports further problems | Adjustment | HW | Oxi module | tHb cal
ABL835 FLEX |  | B | D056C5E1F85DD627 | No |  |  | Assisted Replacement | HW | Inlet module | Inlet gasket
AQT90 FLEX |  | B | D056A6EDF8597ABE | No |  |  | Defect | HW | Analyzer Control Module | Barcode reader, display
TCM TOSCA | New Membrane replaced. and study done on myself. pCO2 values increasing continually. Not acceptable. Reports taken for further discution and study. Tomorow will come with solution. | New Membrane replaced. and study done on myself. pCO2 values increasing continually. Not acceptable. Reports taken for further discution and study. Tomorow will come with solution. | SWO-83335 | No | New Membrane replaced. and study done on myself. pCO2 values increasing continually. Not acceptable. Reports taken for further discution and study. Tomorow will come with solution. | New Membrane replaced. and study done on myself. pCO2 values increasing continually. Not acceptable. Reports taken for further discution and study. Tomorow will come with solution.
ABL80 FLEX BASIC | 更换测试卡正常。 | 更换测试卡正常。 | SWO-83467 | No | The replacement test card is OK. | The replacement test card is OK.
ABL80 FLEX BASIC | 更换测试卡正常。 | 更换测试卡正常。 | SWO-83463 | No | The replacement test card is OK. | The replacement test card is OK.
ABL80 FLEX BASIC | 更换测试卡正常。 | 更换测试卡正常。 | SWO-83422 | No | The replacement test card is OK. | The replacement test card is OK.
ABL80 CO-OX | CONFIRMED ISSUE WHEN ONSITE. PCO2 LEVEL 2 SLUMPED -15MMHG STARTED ON FEB 1 WITH SP LOT 217988 & SC LOT 1345625. | REPLACED PARTS IN PM KIT AND CHANGED SOLUTION PACK FROM LOT 219161 TO 220047 FROM CUSTOMER STOCK. | D05801D66E10AC31 | No | CONFIRMED ISSUE WHEN ONSITE. PCO2 LEVEL 2 SLUMPED -15MMHG STARTED ON FEB 1 WITH SP LOT 217988 & SC LOT 1345625. | REPLACED PARTS IN PM KIT AND CHANGED SOLUTION PACK FROM LOT 219161 TO 220047 FROM CUSTOMER STOCK.
ABL90 FLEX PLUS |  | B | D05A7D935A5C9FDB | Yes |  |  | Re classified RMED | Double entry |  |  | PER ?
ABL90 FLEX PLUS |  | B | D057EA816E367F44 | Yes |  |  | Re classified RMED | Logistical upSWO
ABL90 FLEX |  | B | SWO-83469 | No |  |  | Assisted Replacement | SP | SP | Pump cal error
ABL90 FLEX | 更换试剂包正常。 | 更换试剂包正常。 | SWO-83446 | No | The replacement kit is OK. | The replacement kit is OK. | Assisted Replacement | SP | SP | Pump cal error | 1178
AQT90 FLEX |  | B | SWO-83481 | No |  |  | Defect | HW | Process Module | Optical Unit | 131
AQT90 FLEX |  | B | SWO-83489 | No |  |  | Defect | HW | Incubator module | Sensor, Cup | 2167
ABL825 FLEX | Software deinstalliert und erneut installiert. Alte Datenbank vom 21.02. eingespielt. Zählerstände überprüft, in Ordnung. | Software deinstalliert und erneut installiert. Alte Datenbank vom 21.02. eingespielt. Zählerstände überprüft, in Ordnung. | SWO-83684 | No | Software uninstalled and reinstalled. Old database from 21.02. recorded. Meter readings checked, all right. | Software uninstalled and reinstalled. Old database from 21.02. recorded. Meter readings checked, all right. | Defect | SW | Upgrade | V6.16
ABL800 BASIC |  | B | SWO-83385 | No |  |  | Defect | HW | AutoCheck Module | AutoCheck Module
ABL837 FLEX |  | B | D0588F6D6E388C1A | Yes |  |  | Re classified RMED | Re install after repair
ABL800 BASIC | diverse Einstellungen am Gerät durchgeführt Ventile überprüft und gereinigt Kal- und QC | diverse Einstellungen am Gerät durchgeführt Ventile überprüft und gereinigt Kal- und QC | SWO-83536 | No | Various settings performed on the device Valves checked and cleaned Cal- and QC | Various settings performed on the device Valves checked and cleaned Cal- and QC | Setup | SW | Solution bottle | Level Graphics
ABL827 FLEX | Not rinsing properly Checked valves in service-OK Found pH sealing ring defective | Replaced sealing o-ring Check rinse cycle-OK Replaced cal 1 solution-empty | D057DABB6E293C9D | No | Not rinsing properly Checked valves in service-OK Found pH sealing ring defective | Replaced sealing o-ring Check rinse cycle-OK Replaced cal 1 solution-empty | Defect | Electrode | Electrode | pH
ABL835 FLEX |  | B | D05851EE6E178C19 | Yes |  |  | Re classified RMED | Installation
ABL815 FLEX | Lampeneinheit mechanisch neu abgeglichen, tHb-Kal. durchgeführt, QC durchgeführt, | Lampeneinheit mechanisch neu abgeglichen, tHb-Kal. durchgeführt, QC durchgeführt, | SWO-83768 | No | Lamp unit mechanically re-matched, tHb-Kal. carried out, QC carried out, | Lamp unit mechanically re-matched, tHb-Kal. carried out, QC carried out, | Adjustment | HW | Oxi module | Lamp unit
AQT90 FLEX |  | B | D057EB816E0279DA | No |  |  | Not Cat | Not cat. | Not cat. | Not cat. | 7664
TCM TOSCA | Membrane Replaced again and perform a study but Same continues increasing values coming. Doubt on Electrode Test Cable. TO will sending. | Membrane Replaced again and perform a study but Same continues increasing values coming. Doubt on Electrode Test Cable. TO will sending. | SWO-83698 | No | Membrane Replaced again and perform a study but Same continues increasing values coming. Doubt on Electrode Test Cable. TO will sending. | Membrane Replaced again and perform a study but Same continues increasing values coming. Doubt on Electrode Test Cable. TO will sending.
ABL80 FLEX BASIC |  | B | SWO-83683 | No
ABL80 FLEX CO-OX | Fehlersuche Überprüfung und Korrektur der entsprechenden Eingaben Barcodescanner geändert | Fehlersuche Überprüfung und Korrektur der entsprechenden Eingaben Barcodescanner geändert | SWO-83894 | No | Troubleshooting Check and correction of the corresponding entries Barcode scanner changed | Troubleshooting Check and correction of the corresponding entries Barcode scanner changed
ABL90 FLEX | Anwender hat bereits SP und Einlassdichtung erneuert --> ABL geht immer wieder in geführte Fehlerbehebung Bei. Eintreffen steht der ABL im Halt --> Einlassdichtung nochmals erneuert --> kein Vakuumfehler--> Neustart | Anwender hat bereits SP und Einlassdichtung erneuert --> ABL geht immer wieder in geführte Fehlerbehebung Bei. Eintreffen steht der ABL im Halt --> Einlassdichtung nochmals erneuert --> kein Vakuumfehler--> Neustart | SWO-83777 | No | User has already renewed SP and inlet seal --> ABL goes into guided troubleshooting again and again. Arrival is the ABL in stop --> Inlet seal renewed again --> no vacuum error--> Restart | User has already renewed SP and inlet seal --> ABL goes into guided troubleshooting again and again. Arrival is the ABL in stop --> Inlet seal renewed again --> no vacuum error--> Restart | Defect | HW | Inlet Gasket | Inlet Gasket
ABL90 | Had customer change the solution pack which resolved pO2 issue. After changing pack Ca++ failing on level 1 and 3 and Lactate failing on level 3. | Reran all levels of external qc. Lactate was failing due errors with Ca++. | D0593A8DE41638DF | No | Had customer change the solution pack which resolved pO2 issue. After changing pack Ca++ failing on level 1 and 3 and Lactate failing on level 3. | Reran all levels of external qc. Lactate was failing due errors with Ca++. | Assisted Replacement | SC | SC | pO2
ABL90 FLEX | Fehlersuche, Fehler LP wurde gestern nicht angenommen, inhomogene Lösung->Einlassdichtung überprüft und gereinigt, Anschlussstutzen am LP gereinigt, gleichen LP nochmal eingesetzt->LP richtig aktiviert, beim erneuten Einsetzen konnte kein Problem festgestellt werden, LP lies sich leicht installieren, Kal. und QC ok, Startphase bestanden System-und Funktionskontrolle i.O. | Fehlersuche, Fehler LP wurde gestern nicht angenommen, inhomogene Lösung->Einlassdichtung überprüft und gereinigt, Anschlussstutzen am LP gereinigt, gleichen LP nochmal eingesetzt->LP richtig aktiviert, beim erneuten Einsetzen konnte kein Problem festgestellt werden, LP lies sich leicht installieren, Kal. und QC ok, Startphase bestanden System-und Funktionskontrolle i.O. | SWO-83754 | No | Troubleshooting, error LP was not accepted yesterday, inhomogeneous solution->inlet seal checked and cleaned, connection plug cleaned on the LP, same LP re-inserted->LP correctly activated, no problem could be detected when re-inserting, LP easy to install, Kal. and QC ok, start phase passed system and function control in the | Troubleshooting, error LP was not accepted yesterday, inhomogeneous solution->inlet seal checked and cleaned, connection plug cleaned on the LP, same LP re-inserted->LP correctly activated, no problem could be detected when re-inserting, LP easy to install, Kal. and QC ok, start phase passed system and function control in the | User error | SP Activation | SP Activation | SP Activation
ABL90 FLEX | Fixe Zielwerte deaktiviert und neuen Lösungspack eingesetzt, QC nun perfekt | Fixe Zielwerte deaktiviert und neuen Lösungspack eingesetzt, QC nun perfekt | SWO-83832 | No | Fixed target values deactivated and new solution pack used, QC now perfect | Fixed target values deactivated and new solution pack used, QC now perfect | Assisted Replacement | SP | SP | pCO2
ABL835 FLEX |  | B | D059283CE434A43B | No |  |  | Unknown | SW | Upgrade | V6.16
ABL835 FLEX |  | B | D0593A09E4178C1C | Yes |  |  | Re classified RMED | Installation
ABL80 FLEX BASIC | 更换测试卡正常。 | 更换测试卡正常。 | SWO-83939 | No | The replacement test card is OK. | The replacement test card is OK.
ABL80 FLEX BASIC | 更换新测试卡修复仪器 | 更换新测试卡修复仪器 | SWO-84050 | No | Replace the new test card repair instrument | Replace the new test card repair instrument
ABL80 FLEX | Omdat dit probleem blijkbaar gekend is door SendX werden alle data uit de analyser verzameld om opgestuurd te worden naar de fabrikant voor verder onderzoek. Blijkbaar zijn er problemen met enkele reagentia waardoor de pO2 meting niet altijd correct kan gemeten worden. Momenteel liggen de QC resultaten terug goed waarop besloten werd om niet opnieuw reagentia te vervangen. Op korte termijn zal er een ABL90 worden geïnstalleerd tot dat het | Omdat dit probleem blijkbaar gekend is door SendX werden alle data uit de analyser verzameld om opgestuurd te worden naar de fabrikant voor verder onderzoek. Blijkbaar zijn er problemen met enkele reagentia waardoor de pO2 meting niet altijd correct kan gemeten worden. Momenteel liggen de QC resultaten terug goed waarop besloten werd om niet opnieuw reagentia te vervangen. Op korte termijn zal er een ABL90 worden geïnstalleerd tot dat het | SWO-84089 | No | Because this problem is apparently known by SendX, all data from the analyser was collected to be sent to the manufacturer for further investigation. Apparently there are problems with some reagents so that the pO2 measurement cannot always be measured correctly. Currently, the QC results are back well which decided not to replace reagents again. In the short term, a ABL90 will be installed until the | Because this problem is apparently known by SendX, all data from the analyser was collected to be sent to the manufacturer for further investigation. Apparently there are problems with some reagents so that the pO2 measurement cannot always be measured correctly. Currently, the QC results are back well which decided not to replace reagents again. In the short term, a ABL90 will be installed until the
ABL80 CO-OX | Printer is not advancing paper. Control panel unable to test print but can clear print jobs. | Replaced printer & control PCB. Tested print and print functions normally. Calibrated and ran QC. Instrument operates normally. | D05A7BE65A2FE9FE | No | Printer is not advancing paper. Control panel unable to test print but can clear print jobs. | Replaced printer & control PCB. Tested print and print functions normally. Calibrated and ran QC. Instrument operates normally.
ABL90 FLEX | 更换新试剂包修复仪器 | 更换新试剂包修复仪器 | SWO-84044 | No | Replace the new reagent kit repair instrument | Replace the new reagent kit repair instrument | Assisted Replacement | SP | SP | Leak current
ABL800 FLEX | Compleat Oxy test run- All well. tHb cal done for Service and Normal by 4 times. Failed every try. Need to Check by replacing Hemolyzer unit. | Compleat Oxy test run- All well. tHb cal done for Service and Normal by 4 times. Failed every try. Need to Check by replacing Hemolyzer unit. | SWO-84176 | No | Compleat Oxy test run- All well. tHb cal done for Service and Normal by 4 times. Failed every try. Need to Check by replacing Hemolyzer unit. | Compleat Oxy test run- All well. tHb cal done for Service and Normal by 4 times. Failed every try. Need to Check by replacing Hemolyzer unit. | Defect | HW | Oxi module | Hemolyzer
ABL800 BASIC | 更换945-614一个正常。(by zhenjie) | 更换945-614一个正常。(by zhenjie) | SWO-83993 | No | Replace 945-614 one normal. (by zhenjie) | Replace 945-614 one normal. (by zhenjie) | Assisted Replacement | Electrode | Electrode | pH
ABL825 FLEX | AC Modul demontiert, Lüfter ausgetauscht, Funktionskontrolle ok | AC Modul demontiert, Lüfter ausgetauscht, Funktionskontrolle ok | SWO-84094 | No | AC module disassembled, fan replaced, function control ok | AC module disassembled, fan replaced, function control ok | Defect | HW | AutoCheck Module | Fan
ABL825 FLEX |  | B | D05A96295A18AE80 | No |  |  | Defect | HW | El/Met Module | PCB
ABL825 FLEX | Ventile Elektrodenmodule überprüft. Dichtungen und Sicherungsringe überprüft und in Ordnung. V2-Ventil in Fluidicmodul quergetauscht. Fehler war erst nach Erneuerung der Modul-Pumpenschläuche behoben. Ähnliches Verhalten würde aber bereits vor der Halbjahreswartung festgestellt. | Ventile Elektrodenmodule überprüft. Dichtungen und Sicherungsringe überprüft und in Ordnung. V2-Ventil in Fluidicmodul quergetauscht. Fehler war erst nach Erneuerung der Modul-Pumpenschläuche behoben. Ähnliches Verhalten würde aber bereits vor der Halbjahreswartung festgestellt. | SWO-84308 | No | Valves electrode modules checked. Seals and retaining rings checked and in order. V2 valve cross-exchanged into fluidic module. Error was only fixed after renewal of the module pump hoses. However, similar behaviour would be noted even before the half-year lye. | Valves electrode modules checked. Seals and retaining rings checked and in order. V2 valve cross-exchanged into fluidic module. Error was only fixed after renewal of the module pump hoses. However, similar behaviour would be noted even before the half-year lye. | Defect | HW | pH/BG Module | Pump tubing
ABL800 BASIC | Gehäuse des Pressure Transducers gereinigt. Pressure Transducer ausgetauscht. Luftventil des pH BG Moduls erneuert. Unteren Dichtring der pH Elektrode erneuert. Kalibrationen und QC's vorgenommen. | Gehäuse des Pressure Transducers gereinigt. Pressure Transducer ausgetauscht. Luftventil des pH BG Moduls erneuert. Unteren Dichtring der pH Elektrode erneuert. Kalibrationen und QC's vorgenommen. | SWO-84115 | No | housing of the Pressure Transducer. Pressure Transducer. air valve of the pH BG module. Lower sealing ring of the pH electrode renewed. calibrations and QC's. | housing of the Pressure Transducer. Pressure Transducer. air valve of the pH BG module. Lower sealing ring of the pH electrode renewed. calibrations and QC's. | Defect | HW | pH/BG Module | Valve, air
ABL835 FLEX | Lampe spænding justeret Service tHb kal. | Lampe spænding justeret Service tHb kal. | SWO-84105 | No | Lampe spænding justeret Service tHb kal. | Lampe spænding justeret Service tHb kal. | Adjustment | HW | Oxi module | Lamp unit
AQT90 FLEX | cleaned the cup-sensor and calibrate the incubator wheel. also calibrate the D-Dimer new lot, now working satisfactorily. | cleaned the cup-sensor and calibrate the incubator wheel. also calibrate the D-Dimer new lot, now working satisfactorily. | SWO-84400 | No | cleaned the cup-sensor and calibrate the incubator wheel. also calibrate the D-Dimer new lot, now working satisfactorily. | cleaned the cup-sensor and calibrate the incubator wheel. also calibrate the D-Dimer new lot, now working satisfactorily. | Cleaned | HW | Incubator module | Sensor, Cup | 2167
AQT90 FLEX |  | B | D05AE2CF5A0F7922 | No |  |  | Cleaned | HW | Aspiration Module | PCB, Vacu vision
ABL80 FLEX BASIC | 更换新测试卡修复仪器 | 更换新测试卡修复仪器 | SWO-84210 | No | Replace the new test card repair instrument | Replace the new test card repair instrument
ABL80 FLEX BASIC | 更换测试卡正常。 | 更换测试卡正常。 | SWO-84181 | No | The replacement test card is OK. | The replacement test card is OK.
ABL80 FLEX BASIC | 更换新测试卡修复仪器 | 更换新测试卡修复仪器 | SWO-84184 | No | Replace the new test card repair instrument | Replace the new test card repair instrument
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-84469 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL80 OSM | Spectrum error 581 on blood samples | Replaced the hemolyzer Peformed THB calibration | D05E184CBC5A3DE1 | No | Spectrum error 581 on blood samples | Replaced the hemolyzer Peformed THB calibration
ABL90 FLEX | Oxi Complete Test ok. Ventilschlauch erneuert (Flüssigkeitsaustritt) Kalibrationen und QC's vorgenommen. | Oxi Complete Test ok. Ventilschlauch erneuert (Flüssigkeitsaustritt) Kalibrationen und QC's vorgenommen. | SWO-84329 | No | Oxi Complete Test ok. Valve hose renewed (liquid outlet) calibrations and QC's made. | Oxi Complete Test ok. Valve hose renewed (liquid outlet) calibrations and QC's made. | Defect | HW | Oxi module | pinch valve tube
ABL90 FLEX PLUS |  | B | D05BAAD4D03D228F | No |  |  | Defect | HW | Oxi module | pinch valve tube | 2995
ABL90 FLEX PLUS | Telefonische Rücksprache mit Labor und Stationsleitung Abholung des Analysators beim Kunden. Lieferung in GS und erstellen der Berichte und Informationen an TS. | Telefonische Rücksprache mit Labor und Stationsleitung Abholung des Analysators beim Kunden. Lieferung in GS und erstellen der Berichte und Informationen an TS. | SWO-84339 | No | Telephone consultation with laboratory and station management Pick-up of the analyzer at the customer. Delivery in GS and create the reports and information to TS. | Telephone consultation with laboratory and station management Pick-up of the analyzer at the customer. Delivery in GS and create the reports and information to TS. | Unknown | Removed by RSUB
AQT90 FLEX | 取出断角的测试卡正常。 | 取出断角的测试卡正常。 | SWO-84172 | No | The test card to remove the break is OK. | The test card to remove the break is OK. | Defect | Consumables | Test Kit | Cartridge | 5784 (Broken or Label)
ABL800 BASIC | Ablass-Sensor erneuert | Ablass-Sensor erneuert | SWO-84337 | No | Drain sensor renewed | Drain sensor renewed | Defect | HW | Basic unit | Waste detector
ABL825 FLEX | Probelock eingesetzt. Gerinnsel entfernt. Kalibrationen und Qualitätskontrollen in Ordnung. | Probelock eingesetzt. Gerinnsel entfernt. Kalibrationen und Qualitätskontrollen in Ordnung. | SWO-84348 | No | probe lock. Clots removed. calibrations and quality controls in order. | probe lock. Clots removed. calibrations and quality controls in order. | Defect | HW | Inlet module | Probe Lock
ABL800 BASIC | 更换902-833一个正常.(by zhenjie) | 更换902-833一个正常.(by zhenjie) | SWO-83936 | No | Replace 902-833 one normal. (by zhenjie) | Replace 902-833 one normal. (by zhenjie) | Defect | HW | Oxi module | Oxi module
ABL825 FLEX |  | B | SWO-84216 | No |  |  | Defect | HW | pH/BG Module | PCB
AQT90 FLEX |  | B | D05BBD42D03E6BB5 | No |  |  | Defect | HW | Process Module | Optical Unit
TCM-COMBIM | Confirmed the reported issue. | Determined that the problem was caused by tc sensor. Confirmed that the sensor was under warranty. Replaced the sensor. | D05BA150D0112691 | No | Confirmed the reported issue. | Determined that the problem was caused by tc sensor. Confirmed that the sensor was under warranty. Replaced the sensor. | Assisted Replacement | Electrode | Electrode | pCO2
ABL80 FLEX BASIC | 更换测试卡正常。 | 更换测试卡正常。 | SWO-84467 | No | The replacement test card is OK. | The replacement test card is OK.
ABL80 FLEX BASIC | 更换新测试卡修复仪器 | 更换新测试卡修复仪器 | SWO-84449 | No | Replace the new test card repair instrument | Replace the new test card repair instrument
ABL80 FLEX BASIC | 更换新测试卡修复仪器 | 更换新测试卡修复仪器 | SWO-84422 | No | Replace the new test card repair instrument | Replace the new test card repair instrument
AQT90 FLEX |  | B | SWO-84419 | No |  |  | Defect | HW | Process Module | Optical Unit | 124
AQT90 FLEX | 重新连接电路板线路,修复仪器 | 重新连接电路板线路,修复仪器 | SWO-83971 | No | Reconnect the board line, repair the instrument | Reconnect the board line, repair the instrument | Adjustment | HW | Cover and Chassis | Reconnect Cable | 4587
ABL800 BASIC | Need to replace touch screen | B | SWO-84435 | No | Need to replace touch screen |  | Defect | HW | DMS_PC | Touch Screen
ABL90 FLEX | 更换废液管修复仪器 | 更换废液管修复仪器 | SWO-84248 | No | Replacement of waste pipe repair instruments | Replacement of waste pipe repair instruments | Defect | HW | LS error | LS | 58
ABL90 FLEX | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-84501 | No | Repair after replacing a new test card | Repair after replacing a new test card | Assisted Replacement | SC | SC | Reference | 1361
ABL815 FLEX | Error code 316 neon number too low. Performed complete oxi test and failed for fneon too low. Suspect is malfunctioning spectrometer. | Replaced spectromer. Performed complete oxi test - PASS. Performed thb cal - PASS. Verified instrument operation. | D05EC1ACBC170FDC | No | Error code 316 neon number too low. Performed complete oxi test and failed for fneon too low. Suspect is malfunctioning spectrometer. | Replaced spectromer. Performed complete oxi test - PASS. Performed thb cal - PASS. Verified instrument operation. | Defect | HW | Oxi module | Spectrometer | 316
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-84518 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL80 FLEX BASIC | 重新装包，卡，估仪器清洁正常。 | 重新装包，卡，估仪器清洁正常。 | SWO-84446 | No | Repack, card, evaluate the instrument clean normal. | Repack, card, evaluate the instrument clean normal.
ABL80 FLEX BASIC | 更换新测试卡后解决 | 更换新测试卡后解决 | SWO-84536 | No | Resolve after replacing a new test card | Resolve after replacing a new test card
ABL90 FLEX | Reinigungsarbeiten, Aktivierung Lösungspack kontrolliert, Serviceprogramme geprüft, Lösungspack geliefert, je eine Sensorkassette sowie Lösungspack wurden vom Kunden reklamiert, Startphase, | pO2 ist nach Reparatur immer noch vorhanden Kunde wünscht Service . | SWO-84947 | No | Cleaning work, activation solution pack controlled, service programs checked, solution pack delivered, one sensor cartridge and solution pack were claimed by the customer, start phase, | pO2 is still present after repair customer wishes service . | Cleaned | Fluidic pathway | pO2 | pO2
ABL90 | Analyzer is running 3.1 MR6 | I reloaded the software and cleaned the inlet gasket. No cleaning | D05F94DF3227155A | No | Analyzer is running 3.1 MR6 | I reloaded the software and cleaned the inlet gasket. No cleaning | Cleaned | HW | Inlet Gasket | Inlet Gasket
ABL90 FLEX | Funktion im Service überprüft, OK. System neu gestartet. OK. Dem Anwender SC Wechsel empfohlen wenn der Fehler nochmal Auftritt. | Sie hat die Einlassüberprüft und gereinigt. Test wiederholt. Sie meldet sich wenn es noch Probleme gibt. | SWO-84711 | No | Function checked in service, OK. System restarted. Ok. The user SC change recommended if the error occurs again. | She checked and cleaned the inlet. Test repeated. She will come forward if there are still problems. | No fault found | No fault found | No fault found | No fault found | 1221
ABL90 | CPU is not powering up and in need of replacement. CPU may have been affected by a power surge or power loss, resulting in a failed CPU. | Replaced CPU and verified that all customer settings are present. Instrument has also been placed on a battery backup UPS. Verified proper analyzer function. Analyzer is within acceptable operating specifications. | D05F7EE6325B8F29 | No | CPU is not powering up and in need of replacement. CPU may have been affected by a power surge or power loss, resulting in a failed CPU. | Replaced CPU and verified that all customer settings are present. Instrument has also been placed on a battery backup UPS. Verified proper analyzer function. Analyzer is within acceptable operating specifications. | Defect | HW | CPU | CPU
ABL90 FLEX |  | 更换废液管修复仪器 | SWO-84488 | Yes |  | Replacement of waste pipe repair instruments | Re classified RMED | Double entry
AQT90 FLEX | 更换新测试卡修复仪器 | 更换新测试卡修复仪器 | SWO-84563 | No | Replace the new test card repair instrument | Replace the new test card repair instrument | Defect | Consumables | Test Kit | Cartridge | 5784 (Broken or Label)
AQT90 FLEX | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-85090 | Yes | Repair after replacing a new test card | Repair after replacing a new test card | Re classified RMED | Contents is ABL9 SWO
AQT90 FLEX | 更换了借用的备件，仪器故障未解决。 给仪器做了清洁并校准后故障消失。借用的备件未更换使用。 | 更换了借用的备件，仪器故障未解决。 给仪器做了清洁并校准后故障消失。借用的备件未更换使用。 | SWO-84434 | No | Borrowed spare parts have been replaced and the instrument failure has not been resolved. The fault disappeared after cleaning and calibrating the instrument. Borrowed spare parts have not been replaced. | Borrowed spare parts have been replaced and the instrument failure has not been resolved. The fault disappeared after cleaning and calibrating the instrument. Borrowed spare parts have not been replaced. | Cleaned | HW | Aspiration Module | Barcode reader, Inlet
AQT90 FLEX | 更换了新测试卡后修复 | 更换了新测试卡后修复 | SWO-84533 | No | Repaired after replacing a new test card | Repaired after replacing a new test card | Defect | Consumables | Test Kit | Cartridge | 5784 (Broken or Label)
ABL837 FLEX | CPUボード交換実施。 その後ping確認・・・エラー無し | B | SWO-84704 | No | Cpu board replacement was implemented. Then ping check ... no error |  | Defect | HW | DMS_PC | CPU
ABL825 FLEX | Aufgef. Oxymodul ersetzt (alt R0460N008), Kal, Complete Oxymetrietest, Kal, Service tHb-Kal, Kal, Anwender tHb-Kal, Kontrolle | Aufgef. Oxymodul ersetzt (alt R0460N008), Kal, Complete Oxymetrietest, Kal, Service tHb-Kal, Kal, Anwender tHb-Kal, Kontrolle | SWO-84567 | No | Aufgef. Oxymodul replaced (old R0460N008), Kal, Complete Oxymetry Test, Kal, Service tHb-Kal, Kal, User tHb-Kal, Control | Aufgef. Oxymodul replaced (old R0460N008), Kal, Complete Oxymetry Test, Kal, Service tHb-Kal, Kal, User tHb-Kal, Control | Defect | HW | Oxi module | Neon Lamp, high
ABL825 FLEX | checked data logs for errors, checked electrodes and electrode membranes, checked electrode measuring chamber, | backed up customer data and updated software to ver 6.16, removed and replaced electrode chamber, replaced electrodes and membranes, | D060008C3238F14D | No | checked data logs for errors, checked electrodes and electrode membranes, checked electrode measuring chamber, | backed up customer data and updated software to ver 6.16, removed and replaced electrode chamber, replaced electrodes and membranes, | Defect | HW | pH/BG Module | Measuring chamber
ABL815 FLEX | Lampeneinheit versucht ein zu stellen, Compl.Oximodul ersetzt, Gerät kalibriert, tHb kal.durchgeführt | Lampeneinheit versucht ein zu stellen, Compl.Oximodul ersetzt, Gerät kalibriert, tHb kal.durchgeführt | SWO-84690 | No | Lamp unit tries to adjust, Compl.Oximodul replaces, device calibrated, tHb kal.performed | Lamp unit tries to adjust, Compl.Oximodul replaces, device calibrated, tHb kal.performed | Defect | HW | Oxi module | Oxi module
ABL825 FLEX |  | B | D05FC6EB3228AEE2 | No |  |  | Cleaned | Fluidic pathway | Inlet module | Inlet gasket | Glass broken
AQT90 FLEX | Scanner für Testkassetten funktioniert nicht mehr. Neuen Scanner und Kabel aus GS geholt zum Test. Neue Sicherung verbaut. Scanner funktioniert wieder. Systemüberprüfung durchgeführt -->i.O. AQT wieder Messbereit | Scanner für Testkassetten funktioniert nicht mehr. Neuen Scanner und Kabel aus GS geholt zum Test. Neue Sicherung verbaut. Scanner funktioniert wieder. Systemüberprüfung durchgeführt -->i.O. AQT wieder Messbereit | SWO-84696 | No | Scanner for test cartridges no longer works. New scanner and cable from GS brought for testing. New fuse installed. Scanner works again. System check performed -->i.O. AQT ready for measurement again | Scanner for test cartridges no longer works. New scanner and cable from GS brought for testing. New fuse installed. Scanner works again. System check performed -->i.O. AQT ready for measurement again | Defect | HW | Storage Module | Barcode reader, Storage wheel
AQT90 FLEX |  | B | D05FE539320F1AFB | No |  |  | Defect | HW | Process Module | Process Unit sensor | SetSafe 22
AQT90 FLEX |  | B | D06816B46C1D114A | No |  |  | Revisit required | HW | Analyzer Control Module | Cable (CPU-Cover)
ABL80 FLEX BASIC | 更换试剂包正常。 | 更换试剂包正常。 | SWO-85597 | No | The replacement kit is OK. | The replacement kit is OK.
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-84800 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-84826 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL80 FLEX BASIC | 更换测试卡正常。 | 更换测试卡正常。 | SWO-85455 | No | The replacement test card is OK. | The replacement test card is OK.
ABL90 | Analyzer having software errors after wireless changes made. | Installed new version software. Worked with CHMC IT to complete wireless install. | D061488BA82002C7 | No | Analyzer having software errors after wireless changes made. | Installed new version software. Worked with CHMC IT to complete wireless install. | Defect | SW | Wireless connection | Inconsistent SW
ABL90 |  | B | D060F8E8A81A91BC | No |  |  | Defect | HW | PM Kit | PM Kit
ABL90 | Continued troubleshooting of wireless platform. Wireless will not connect with dongle in top USB. | Replaced CPU unit, re-installed software and wireless. CHMC installed certificates and setup for wireless. | D0614614A81B02C2 | No | Continued troubleshooting of wireless platform. Wireless will not connect with dongle in top USB. | Replaced CPU unit, re-installed software and wireless. CHMC installed certificates and setup for wireless. | Defect | HW | CPU | Wireless connection
ABL90 |  | B | D061FC801E192A94 | No |  |  | User error | SP Installation | Paper towel | Paper towel
ABL90 FLEX | 更换新试剂包修复仪器 | 更换新试剂包修复仪器 | SWO-84786 | No | Replace the new reagent kit repair instrument | Replace the new reagent kit repair instrument | Assisted Replacement | SP | SP | Chip error
ABL90 FLEX | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-85093 | No | Repair after replacing a new test card | Repair after replacing a new test card | Assisted Replacement | SC | SC | pCO2 | 1318
ABL90 FLEX | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-84911 | No | Repair after replacing a new test card | Repair after replacing a new test card | Assisted Replacement | SC | SC | Reference
AQT90 FLEX |  | 工程师来电需要增加借件，故关单 | SWO-84787 | No |  | Engineer calls need to add borrowings, so close the bill | Defect | HW | Analyzer Control Module | PCB, Main board
AQT90 FLEX | 工程师来电需要增加借件，故关单 | 工程师来电需要增加借件，故关单 | SWO-84549 | Yes | Engineer calls need to add borrowings, so close the bill | Engineer calls need to add borrowings, so close the bill | Re classified RMED | Double upSWO
ABL835 FLEX | Refメンブラン交換。測定チャンバー洗浄。 | Refメンブラン交換。測定チャンバー洗浄。 | SWO-85104 | No | Ref Membrane replacement. Measuring chamber cleaning. | Ref Membrane replacement. Measuring chamber cleaning. | Assisted Replacement | Electrode | Membrane | Reference
ABL835 FLEX | FLEXQ修理完了の為、返却。FLEXQポジション調整実施。代替えモジュール引き上げ。 | FLEXQ修理完了の為、返却。FLEXQポジション調整実施。代替えモジュール引き上げ。 | SWO-85108 | No | Return ed to flexq repair complete. FlexQ position adjustment implemented. Alternative module raise. | Return ed to flexq repair complete. FlexQ position adjustment implemented. Alternative module raise. | Adjustment | HW | FlexQ Module | Positioning
ABL835 FLEX | 関連部品交換 | B | SWO-91092 | No | Replacement of related parts |  | Unknown | Electrode | Electrode | Electrolyte
ABL827 FLEX |  | B | D060A47CA80F3E3F | Yes |  |  | Re classified RMED | SWU
ABL800 BASIC | Überprüfung Daten Reinigung K membraniert, tHb Kal | Überprüfung Daten Reinigung K membraniert, tHb Kal | SWO-85031 | No | Verification data cleaning K membraneed, tHb Kal | Verification data cleaning K membraneed, tHb Kal | Assisted Replacement | Electrode | Membrane | K+
ABL825 FLEX | CAS tried changing membrane | Replaced all pump tubing. Replaced all tubing from inlet to hemolyzer. Checked all fluidic and measuring valves. Checked measuring chamber for cracks. | D060CE4CA8294A7C | No | CAS tried changing membrane | Replaced all pump tubing. Replaced all tubing from inlet to hemolyzer. Checked all fluidic and measuring valves. Checked measuring chamber for cracks. | Defect | HW | Fluidic manifold | Pump tubing
ABL825 FLEX | cals and run every half hour, they start as 1 point ph bg 20 seconds later they end as 1 point cal , checked cal bottle pn 944-127 | B | D060EA49A8531930 | No | cals and run every half hour, they start as 1 point ph bg 20 seconds later they end as 1 point cal , checked cal bottle pn 944-127 |  | Unknown | Electrode | Electrode | pH
ABL825 FLEX |  | B | D060D082A85C88A7 | No |  |  | Defect | HW | Oxi module | Lamp unit
AQT90 FLEX | 機器返却 （ｲﾝﾚｯﾄPCB、ｹｰﾌﾞﾙFPC、ﾆｰﾄﾞﾙｳｫｯｼｭﾌﾞﾛｯｸ交換） | 機器返却 （ｲﾝﾚｯﾄPCB、ｹｰﾌﾞﾙFPC、ﾆｰﾄﾞﾙｳｫｯｼｭﾌﾞﾛｯｸ交換） | SWO-88022 | No | Return equipment (inlet PCB, cable FPC, needle wash block replacement) | Return equipment (inlet PCB, cable FPC, needle wash block replacement) | Defect | HW | Aspiration Module | Needle Wash Unit | 1223
ABL80 FLEX BASIC | 更换新测试卡修复仪器 | 更换新测试卡修复仪器 | SWO-85135 | No | Replace the new test card repair instrument | Replace the new test card repair instrument
ABL80 FLEX BASIC | 更换新试剂包修复仪器 | 更换新试剂包修复仪器 | SWO-86076 | No | Replace the new reagent kit repair instrument | Replace the new reagent kit repair instrument
ABL80 FLEX BASIC | 更换测试卡正常。 | 更换测试卡正常。 | SWO-85118 | No | The replacement test card is OK. | The replacement test card is OK.
ABL80 FLEX BASIC | 更换924-738和902-972后修复 | 更换924-738和902-972后修复 | SWO-84782 | No | Repair after replacement of 924-738 and 902-972 | Repair after replacement of 924-738 and 902-972
ABL80 FLEX BASIC | 更换五张测试卡正常。 | 更换五张测试卡正常。 | SWO-85454 | No | Replacing five test cards is OK. | Replacing five test cards is OK.
ABL90 FLEX | Klemmschlauch und Einlasssonde erneuert, Drucktest i.O. Reklamation besprochen. Häufige LP und SC Fehler in letzter Zeit. | Klemmschlauch und Einlasssonde erneuert, Drucktest i.O. Reklamation besprochen. Häufige LP und SC Fehler in letzter Zeit. | SWO-85361 | No | Clamping hose and inlet probe renewed, pressure test in the current complaint discussed. Frequent LP and SC errors lately. | Clamping hose and inlet probe renewed, pressure test in the current complaint discussed. Frequent LP and SC errors lately. | Defect | HW | Oxi module | Pinch Valve
ABL90 | Confirmed. The analyzer will not even boot. | Tried to access the CF card with a USB stick, but I was unable to. | D06270411E46438F | No | Confirmed. The analyzer will not even boot. | Tried to access the CF card with a USB stick, but I was unable to. | Unknown | SW | CF Card | Memory
ABL90 |  | B | D0626B941E3101A8 | No |  |  | Defect | HW | PM Kit | PM Kit | 1254
ABL90 |  | B | D06269711E4B01A3 | Yes |  |  | Re classified RMED | Installation
ABL90 |  | B | D0626A0F1E4F01A4 | Yes |  |  | Re classified RMED | Installation
ABL90 FLEX PLUS | Analysator überprüft - Erneuerung Schlauch an Quetschventil (leichte Undichtigkeit) - Funktionsüberprüfung und Qualitätskontrollen ohne Beanstandung ! Eventueller Pre-Analytik-Fehler in kapilärer Messung vorhanden ! | Analysator überprüft - Erneuerung Schlauch an Quetschventil (leichte Undichtigkeit) - Funktionsüberprüfung und Qualitätskontrollen ohne Beanstandung ! Eventueller Pre-Analytik-Fehler in kapilärer Messung vorhanden ! | SWO-85344 | No | Analyzer checked - renewal hose to squeezing valve (slight leakage) - functional check and quality controls without complaint ! Possible pre-analysis error in capilary measurement present ! | Analyzer checked - renewal hose to squeezing valve (slight leakage) - functional check and quality controls without complaint ! Possible pre-analysis error in capilary measurement present ! | Defect | HW | Oxi module | Pinch Valve
ABL90 FLEX PLUS | Kunden über Sachverhalt informiert, Kundenanschreiben per Mail zugesandt, CCF erstellt und an TS geschickt | Kunden über Sachverhalt informiert, Kundenanschreiben per Mail zugesandt, CCF erstellt und an TS geschickt | SWO-85253 | No | informed customers of facts, sent customer letters by e-mail, created CCF and sent to TS | informed customers of facts, sent customer letters by e-mail, created CCF and sent to TS | Assisted Replacement | SP | SP | pCO2
AQT90 FLEX | 清理光学单元，更换847-607一个正常。 | 清理光学单元，更换847-607一个正常。 | SWO-85383 | No | Clean the optical unit and replace the 847-607 with a normal one. | Clean the optical unit and replace the 847-607 with a normal one. | Cleaned | HW | Process Module | PCB, Cup dryer | 17
ABL810 FLEX | checked patient logs saw that Po2 and PCo2 temp corrected values not reporting when pco2 or po2 were above or below reportable range. blood gas values were always reporting correctly. Blood Gas values were on the bottom of the print outs temp corrected were on the top | changed the Roma ini file to show values out of reportable range. a > or < sign will be shown with the high or low reportable range value. The actual value was reporting in the Temp corrected value and was reporting on the mckesson side to fix i added the ph pco2 and po2 reportable ranges in those temp corrected parameters. | D06250A51E3AFBB1 | Yes | checked patient logs saw that Po2 and PCo2 temp corrected values not reporting when pco2 or po2 were above or below reportable range. blood gas values were always reporting correctly. Blood Gas values were on the bottom of the print outs temp corrected were on the top | changed the Roma ini file to show values out of reportable range. a > or < sign will be shown with the high or low reportable range value. The actual value was reporting in the Temp corrected value and was reporting on the mckesson side to fix i added the ph pco2 and po2 reportable ranges in those temp corrected parameters. | Re classified RMED
ABL827 FLEX | Could not duplicate problem. Checked pumps, tubing, liquid sensors and valves. All checked okay. | Cleaned measuring chambers and replaced pump tubings. | D062045B1E326D05 | No | Could not duplicate problem. Checked pumps, tubing, liquid sensors and valves. All checked okay. | Cleaned measuring chambers and replaced pump tubings. | Defect | HW | Met II Module | Pump tubing
ABL825 FLEX | checked patient logs saw that Po2 and PCo2 temp corrected values not reporting when pco2 or po2 were above or below reportable range. blood gas values were always reporting correctly. Blood Gas values were on the bottom of the print outs temp corrected were on the top. | changed the Roma ini file to show values out of reportable range. a > or < sign will be shown with the high or low reportable range value. The actual value was reporting in the Temp corrected value and was reporting on the mckesson side to fix i added the ph pco2 and po2 reportable ranges in those temp corrected parameters. | D06255F01E23FBB7 | Yes | checked patient logs saw that Po2 and PCo2 temp corrected values not reporting when pco2 or po2 were above or below reportable range. blood gas values were always reporting correctly. Blood Gas values were on the bottom of the print outs temp corrected were on the top. | changed the Roma ini file to show values out of reportable range. a > or < sign will be shown with the high or low reportable range value. The actual value was reporting in the Temp corrected value and was reporting on the mckesson side to fix i added the ph pco2 and po2 reportable ranges in those temp corrected parameters. | Re classified RMED
ABL825 FLEX | checked patient logs saw that Po2 and PCo2 temp corrected values not reporting when pco2 or po2 were above or below reportable range. blood gas values were always reporting correctly. Blood Gas values were on the bottom of the print outs temp corrected were on the top. | changed the Roma ini file to show values out of reportable range. a > or < sign will be shown with the high or low reportable range value. The actual value was reporting in the Temp corrected value and was reporting on the mckesson side to fix i added the ph pco2 and po2 reportable ranges in those temp corrected parameters. | D06254FC1E56FBB4 | Yes | checked patient logs saw that Po2 and PCo2 temp corrected values not reporting when pco2 or po2 were above or below reportable range. blood gas values were always reporting correctly. Blood Gas values were on the bottom of the print outs temp corrected were on the top. | changed the Roma ini file to show values out of reportable range. a > or < sign will be shown with the high or low reportable range value. The actual value was reporting in the Temp corrected value and was reporting on the mckesson side to fix i added the ph pco2 and po2 reportable ranges in those temp corrected parameters. | Re classified RMED
ABL825 FLEX | ran complete oxy all passing, checked lamp and it drifted up to 3800, on install it was at 3400 | adjusted lamp to 3300 | D06243751E5863E1 | No | ran complete oxy all passing; checked lamp and it drifted up to 3800; on install it was at 3400 | adjusted lamp to 3300 | Adjustment | HW | Oxi module | Lamp unit
AQT90 FLEX | Fehlersuche, im Service kann Process Unit und Process Wheel nicht richtig bewegt werden, Überprüfung der Lichtschranke am Process Unit-> zeigt im Service immer 1 egal bei welcher Position, Lichtschranke defekt->bestellt, wird am Folgetag um 7.30Uhr getauscht | Fehlersuche, im Service kann Process Unit und Process Wheel nicht richtig bewegt werden, Überprüfung der Lichtschranke am Process Unit-> zeigt im Service immer 1 egal bei welcher Position, Lichtschranke defekt->bestellt, wird am Folgetag um 7.30Uhr getauscht | SWO-85636 | No | Troubleshooting, in the service process unit and process wheel can not be moved correctly, inspection of the light barrier at the Process Unit-> always shows in the service 1 no matter at which position, light barrier defective->ordered, is exchanged the following day at 7.30 am | Troubleshooting, in the service process unit and process wheel can not be moved correctly, inspection of the light barrier at the Process Unit-> always shows in the service 1 no matter at which position, light barrier defective->ordered, is exchanged the following day at 7.30 am | Revisit required | HW | Process Module | Process Unit sensor | SetSafe 22
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-85846 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL80 FLEX BASIC | 更换测试卡修复仪器 | 更换测试卡修复仪器 | SWO-85591 | No | Replace the test card repair instrument | Replace the test card repair instrument
ABL80 FLEX BASIC |  | B | SWO-85439 | No
ABL80 FLEX BASIC | 更换测试卡正常。 | 更换测试卡正常。 | SWO-86265 | No | The replacement test card is OK. | The replacement test card is OK.
ABL80 FLEX BASIC | 连续更换了两张测试卡后修复 | 连续更换了两张测试卡后修复 | SWO-85363 | No | Repairafter after two consecutive test cards replaced | Repairafter after two consecutive test cards replaced
ABL80 FLEX Basic version |  | B | D0636F24944B3272 | No
ABL80 FLEX CO-OX | Gerät überprüft, Sensorkassette ausgetauscht, pO2 Level 1 zu niedrig, Sensorkassette neu installiert, Kal und QC o.k. | Kunde hat sowohl LP als auch Sensor getauscht . QC Fehler bei pCo2 und PH bleibt bestehen . Kunde hat kein Material mehr vor Ort. Kunde bittet um Service . | SWO-85695 | No | Device checked, sensor cartridge replaced, pO2 Level 1 too low, sensor cartridge reinstalled, Kal and QC ok. | Customer has exchanged both LP and sensor. QC error with pCo2 and PH remains. Customer no longer has any material on site. Customer asks for service .
ABL90 FLEX | Erneut Kalibration ausgelöst. Qualitätskontrolle Messung anschließend o. K. Überprüfung Analysator im Service, kein Fehler feststellbar. Einlasssonde krum, erneuert. Einlassdichtung verschmutzt, erneuert. Auf Tastatur sind Buchstaben vertauscht. Englische Tastatur eingestellt. Korrektur durchgeführt. | Erneut Kalibration ausgelöst. Qualitätskontrolle Messung anschließend o. K. Überprüfung Analysator im Service, kein Fehler feststellbar. Einlasssonde krum, erneuert. Einlassdichtung verschmutzt, erneuert. Auf Tastatur sind Buchstaben vertauscht. Englische Tastatur eingestellt. Korrektur durchgeführt. | SWO-85559 | No | Re-calibration triggered. Quality control measurement then o.K. check analyzer in service, no error detectable. Inlet probe krum, renewed. Inlet seal dirty, renewed. On keyboard, letters are swapped. English keyboard set. correction has been carried out. | Re-calibration triggered. Quality control measurement then o.K. check analyzer in service, no error detectable. Inlet probe krum, renewed. Inlet seal dirty, renewed. On keyboard, letters are swapped. English keyboard set. correction has been carried out. | Assisted Replacement | SC | SC | Metabolite
ABL90 | Analyzer indicates the battery is 100% charged, but the analyzer will only run a few seconds on battery power before it shuts down. | Replaced the battery module. | D063C0B994583EDA | No | Analyzer indicates the battery is 100% charged, but the analyzer will only run a few seconds on battery power before it shuts down. | Replaced the battery module. | Defect | HW | Battery module | Battery
ABL90 | Analyzer indicates 0% charge at all times. | Replace the battery module. | D063C28E941A3EDF | No | Analyzer indicates 0% charge at all times. | Replace the battery module. | Defect | HW | Battery module | Battery module
ABL90 | Upon loading wireless capability drivers, it was discovered that there was a corrupt file on Windows XPE causing the drivers not to load properties. | Installed new CPU with Windows. Reinstalled all Application software and all customer setup and Data. Installed Wireless driver in unit. | D063412894366775 | No | Upon loading wireless capability drivers, it was discovered that there was a corrupt file on Windows XPE causing the drivers not to load properties. | Installed new CPU with Windows. Reinstalled all Application software and all customer setup and Data. Installed Wireless driver in unit. | Defect | SW | Wireless connection | Wireless connection
ABL90 FLEX PLUS |  | B | D0635B139455BD3B | No |  |  | Defect | HW | Sensor Stepper Motor | Sensor Stepper Motor
ABL90 FLEX | 更换测试卡修复仪器 | 更换测试卡修复仪器 | SWO-85440 | No | Replace the test card repair instrument | Replace the test card repair instrument | Assisted Replacement | SC | SC | Reference | 1361
AQT90 FLEX | 使用BNP定标卡944-258给仪器重新定标后恢复 | 使用BNP定标卡944-258给仪器重新定标后恢复 | SWO-85141 | No | Restored after re-calibration of the instrument using BNP calibration card 944-258 | Restored after re-calibration of the instrument using BNP calibration card 944-258 | Unknown | Mail | Test Kit | Test Kit | 129
ABL835 FLEX | 比較電極、比較メンブラン交換。 | 比較電極、比較メンブラン交換。 | SWO-85835 | No | Comparison electrodes, comparison membrane replacement. | Comparison electrodes, comparison membrane replacement. | Assisted Replacement | Electrode | Electrode | Na+
ABL800 FLEX | Elektrode getauscht | Elektrode getauscht | SWO-85605 | No | Electrode exchanged | Electrode exchanged | Assisted Replacement | Electrode | Electrode | pCO2
ABL825 FLEX | Messkammern und Elektroden gereinigt, Messkappenkontakte gereinigt, Ref. und Glukose neu membraniert(fällig), mehrere Kal. und QCs System-und Funktionskontrolle i.O. | Messkammern und Elektroden gereinigt, Messkappenkontakte gereinigt, Ref. und Glukose neu membraniert(fällig), mehrere Kal. und QCs System-und Funktionskontrolle i.O. | SWO-85637 | No | Measuring chambers and electrodes cleaned, measuring cap contacts cleaned, ref. and glucose re-membraneized (due), several cal. and QC's system and function control i.O. | Measuring chambers and electrodes cleaned, measuring cap contacts cleaned, ref. and glucose re-membraneized (due), several cal. and QC's system and function control i.O. | Assisted Replacement | Electrode | Electrode | pH
ABL825 FLEX | Found DB to be corrupt. Customer had reported problems with battery surge protector. | Removed and reloaded ABL 6.16 software. Performed all necessary adjusments, setups and calibrations. | D06366A0942D889E | No | Found DB to be corrupt. Customer had reported problems with battery surge protector. | Removed and reloaded ABL 6.16 software. Performed all necessary adjusments, setups and calibrations. | Unknown | HW | DMS_PC | HDD | Power spike
ABL825 FLEX | CUSTOMER REPORTED OCCASIONAL GRINDING TICKING SOUND FROM AUTCHECK AREA. | FSE VERIFIED MECHANICS AND FUNTION OF AUTOCHECK, VERIFIED FANS UNDERNEATH, CLEANED SMALL AMOUNT OF DEBRIS FROM UNDER CAROUSEL, TEST QC 1 FOR MECHANICAL FUNCTION NO ISSUE FOUND. SUSPECT ISSUE AROSE WITH RELOADING OF CAROUSEL. | D06345C19452DD1C | No | CUSTOMER REPORTED OCCASIONAL GRINDING TICKING SOUND FROM AUTCHECK AREA. | FSE VERIFIED MECHANICS AND FUNTION OF AUTOCHECK, VERIFIED FANS UNDERNEATH, CLEANED SMALL AMOUNT OF DEBRIS FROM UNDER CAROUSEL, TEST QC 1 FOR MECHANICAL FUNCTION NO ISSUE FOUND. SUSPECT ISSUE AROSE WITH RELOADING OF CAROUSEL. | Cleaned | HW | AutoCheck Module | AutoCheck Module
ABL827 FLEX | Gas Einheit geprüft, Offset bei 0,196. Druckaufnehmer überprüft, Lösungsreste in der Kammer festgestellt. Zuleitung freigepustet und getrocknet. Halogenlampe von 3800 auf 3200 runter geregelt. Bei 1. Kal. die Halogenspannung zu niedrig, andere Halogenlampe eingesetzt, danach Spannung zu hoch, diverse Maßnahmen später sind Spannung und Halogenintensität im Bereich. Gasdruck Offset Spannung erneut aus dem Bereich. Austausch der Gaseinheit geplant. | Gas Einheit geprüft, Offset bei 0,196. Druckaufnehmer überprüft, Lösungsreste in der Kammer festgestellt. Zuleitung freigepustet und getrocknet. Halogenlampe von 3800 auf 3200 runter geregelt. Bei 1. Kal. die Halogenspannung zu niedrig, andere Halogenlampe eingesetzt, danach Spannung zu hoch, diverse Maßnahmen später sind Spannung und Halogenintensität im Bereich. Gasdruck Offset Spannung erneut aus dem Bereich. Austausch der Gaseinheit geplant. | SWO-86120 | No | Gas unit tested, offset at 0.196. Pressure transducer checked, solution residues found in the chamber. Supply free-blown and dried. Halogen lamp from 3800 to 3200 down. At 1. the halogen voltage too low, other halogen lamp used, then voltage too high, various measures later are voltage and halogen intensity in the range. Gas pressure offset voltage again from the range. replacement of the gas unit. | Gas unit tested, offset at 0.196. Pressure transducer checked, solution residues found in the chamber. Supply free-blown and dried. Halogen lamp from 3800 to 3200 down. At 1. the halogen voltage too low, other halogen lamp used, then voltage too high, various measures later are voltage and halogen intensity in the range. Gas pressure offset voltage again from the range. replacement of the gas unit. | Cleaned | Fluidic pathway | Gas module | Pressure Transducer | 689
AQT90 FLEX | Bestellte Lichtschranke erneuert, Neustart ok, Safe all Positions ok, LQC System-und Funktionskontrolle i.O. | Fehlersuche, im Service kann Process Unit und Process Wheel nicht richtig bewegt werden, Überprüfung der Lichtschranke am Process Unit-> zeigt im Service immer 1 egal bei welcher Position, Lichtschranke defekt->bestellt, wird am Folgetag um 7.30Uhr getauscht | SWO-85635 | No | Ordered light barrier renewed, restart ok, safe all positions ok, LQC system and function control i.O. | Troubleshooting, in the service process unit and process wheel can not be moved correctly, inspection of the light barrier at the Process Unit-> always shows in the service 1 no matter at which position, light barrier defective->ordered, is exchanged the following day at 7.30 am | Defect | HW | Process Module | Process Unit sensor | SetSafe 22
AQT90 FLEX | 機器引き上げし、代替え器設置 | ｲﾝｷｭﾍﾞｰﾀｰPCB交換 | SWO-87225 | No | Equipment pulling up and installation of replacement equipment | Incubator PCB Replacement | Revisit required | HW | Incubator module | PCB, Incubator | 2168
AQT90 FLEX | Mit Kaltgeräte Kabel direkt ans Hausnetz angeschlossen und Sicherungen erneuert, ohne Erfolg. Netzteil defekt. Powersupply 903-175P0005N726 erneuert durch P0002N469. Startphase i.O. Systemreinigung, Blanktest und Lvl.2 LQCs aller Parameter im Bereich. Pat. Messung gestartet, Barcode Erkennung dauert lange, Datenpflege durchgeführt. | Mit Kaltgeräte Kabel direkt ans Hausnetz angeschlossen und Sicherungen erneuert, ohne Erfolg. Netzteil defekt. Powersupply 903-175P0005N726 erneuert durch P0002N469. Startphase i.O. Systemreinigung, Blanktest und Lvl.2 LQCs aller Parameter im Bereich. Pat. Messung gestartet, Barcode Erkennung dauert lange, Datenpflege durchgeführt. | SWO-86124 | No | With cold equipment cables connected directly to the house network and backups renewed, to no avail. Power supply defective. Powersupply 903-175P0005N726 renewed by P0002N469. Start phase in line with system cleaning, blank test and Lvl.2 LQCs of all parameters in the range. Pat. Measurement started, barcode recognition takes a long time, data maintenance performed. | With cold equipment cables connected directly to the house network and backups renewed, to no avail. Power supply defective. Powersupply 903-175P0005N726 renewed by P0002N469. Start phase in line with system cleaning, blank test and Lvl.2 LQCs of all parameters in the range. Pat. Measurement started, barcode recognition takes a long time, data maintenance performed. | Defect | HW | Power Supply | Power Supply
ABL80 FLEX BASIC | 更换测试卡正常。 | 更换试剂包正常。 | SWO-85596 | No | The replacement test card is OK. | The replacement kit is OK.
ABL80 FLEX BASIC | 清理废液槽和废液管路 | 清理废液槽和废液管路 | SWO-85642 | No | Clean up waste tanks and waste lines | Clean up waste tanks and waste lines
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-85914 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-85583 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL80 FLEX BASIC | 更换电源902-970一个后修复 | 更换电源902-970一个后修复 | SWO-85111 | No | Replace the power supply 902-970 after one repair | Replace the power supply 902-970 after one repair
ABL90 FLEX | Barcode Einstellung angepasst | Barcode Einstellung angepasst | SWO-85821 | Yes | Barcode setting adjusted | Barcode setting adjusted | Re classified RMED | Planned TU
ABL90 |  | B | D0649E930A5520EE | No |  |  | Assisted Replacement | SC | SC | Reference
TCM4 |  | B | SWO-85657 | No
ABL90 FLEX | 更换新测试卡修复仪器 | 更换新测试卡修复仪器 | SWO-85661 | No | Replace the new test card repair instrument | Replace the new test card repair instrument | Assisted Replacement | SC | SC | Ca Cal
AQT90 FLEX | 更换了进样针903-171和AU板903-198后修复仪器 | 更换了进样针903-171和AU板903-198后修复仪器 | SWO-85154 | No | Replacement of the feeding needle 903-171 and AU plate 903-198 after repair instrument | Replacement of the feeding needle 903-171 and AU plate 903-198 after repair instrument | Defect | HW | Analyzer Control Module | PCB, Main board | 6732
ABL800 BASIC |  | B | SWO-85610 | No |  |  | Assisted Replacement | Electrode | Electrode | pH
ABL825 FLEX | Erreur 258,316 oxi test paase pas avec la lampe | Remplac la lampe unit | D06478780A28C7AC | No | Error 258,316 oxi test paase not with lamp | Replace the lamp unites | Defect | HW | Oxi module | Lamp unit | 258
AQT90 FLEX |  | B | D06814C06C341148 | No |  |  | Revisit required | HW | Analyzer Control Module | Cable (CPU-Cover)
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-86062 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL80 FLEX BASIC | 更换测试卡。 | 更换测试卡。 | SWO-86117 | No | Replace the test card. | Replace the test card.
ABL815 FLEX | Abgebrochenen Einlassnippel ersetzt. Fehlersuche: Fließwege überprüft, pH Elektrode mit Reiniger behandelt, Dichtheitstest, Ventile der Messmodule ausgebaut und überprüft. Neustart, Kalibrationen, Funktionstest. Endkontrolle: AC5+ Messung => Alle Werte OK | Abgebrochenen Einlassnippel ersetzt. Fehlersuche: Fließwege überprüft, pH Elektrode mit Reiniger behandelt, Dichtheitstest, Ventile der Messmodule ausgebaut und überprüft. Neustart, Kalibrationen, Funktionstest. Endkontrolle: AC5+ Messung => Alle Werte OK | SWO-85827 | No | Replaced the broken inlet nipple. Troubleshooting: Flow paths checked, pH electrode treated with cleaner, leak test, valves of the measuring modules removed and checked. Restart, calibrations, function test. Final control: AC5+ measurement => All values OK | Replaced the broken inlet nipple. Troubleshooting: Flow paths checked, pH electrode treated with cleaner, leak test, valves of the measuring modules removed and checked. Restart, calibrations, function test. Final control: AC5+ measurement => All values OK | Defect | HW | Inlet module | Nipple kit | 328
ABL810 FLEX | Checked status indicators, determine Rinse was low. | Had operator replace rinse solution and restart analzyer. Start button still not active, verified no other conditions exisited to impact start function. Had operator power down and restart analzyer, problem is now resolved. | D067A060F62490B3 | Yes | Checked status indicators, determine Rinse was low. | Had operator replace rinse solution and restart analzyer. Start button still not active, verified no other conditions exisited to impact start function. Had operator power down and restart analzyer, problem is now resolved. | Re classified RMED | Phone support
ABL80 FLEX BASIC | 更换测试卡正常。 | 更换测试卡正常。 | SWO-85860 | No | The replacement test card is OK. | The replacement test card is OK.
ABL80 FLEX BASIC | 更换试剂包正常。 | 更换试剂包正常。 | SWO-85965 | No | The replacement kit is OK. | The replacement kit is OK.
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-86079 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL80 FLEX BASIC | 更换测试卡正常。 | 更换测试卡正常。 | SWO-86920 | No | The replacement test card is OK. | The replacement test card is OK.
ABL80 OSM | Inspected waste manifold for cracks and none was found. | Replaced waste manifold. | D0684F306C53F68E | No | Inspected waste manifold for cracks and none was found. | Replaced waste manifold.
ABL90 FLEX | SC300交換。交換後pCO2ｴﾗｰで再交換必要に。 | B | SWO-85849 | No | SC300 replacement. Need to re-replace with pCO2 error after replacement. |  | Assisted Replacement | SC | SC | Wetability
ABL90 FLEX | SC交換 | B | SWO-86164 | Yes | SC Switch |  | Re classified RMED | Double entry
ABL90 FLEX | bei Ankunft Gerät wieder in voller Funktion, Kunde hat den Fehler selber beseitigt und das System durchgespült | bei Ankunft Gerät wieder in voller Funktion, Kunde hat den Fehler selber beseitigt und das System durchgespült | SWO-86093 | No | on arrival device back in full function, customer has fixed the error himself and flushed the system | on arrival device back in full function, customer has fixed the error himself and flushed the system | Cleaned | Fluidic pathway
ABL90 FLEX PLUS | Druckbild geprüft --> Druckerbild teilweise schlecht ( Striche verwischt ) | Druckbild geprüft --> Druckerbild teilweise schlecht ( Striche verwischt ) | SWO-86011 | No | Print image checked --> Printer image partly bad ( strokes blurred ) | Print image checked --> Printer image partly bad ( strokes blurred ) | User error | Barcode Quality
ABL90 FLEX PLUS | Flatcable vervangen tussen Inlet detector PCB en Wet Section Control PCB. | B | SWO-86029 | No | Flatcable vervangen tussen Inlet detector PCB en Wet Section Control PCB. |  | Defect | SW | Auto Inlet module | Inlet Position Cal. | Replace Inlet Gasket?
ABL90 FLEX PLUS |  | B | D06A18B7E22E9A7F | Yes |  |  | Re classified RMED | User training
TCM4 |  | B | SWO-85875 | No
ABL90 | Se comprueba el SP y el SC nuevo instalado por el cliente siendo este último el origen del problema y el cual se ha sustituído de nuevo (946-008 600 30) para dejar operativo el equipo. Calibración y pasar QC´s. | Sustituir (946-008 600 30) para dejar operativo el equipo. | D0849CDB0622DDA0 | No | The SP and new SC installed by the client are checked, the latter being the source of the problem and replaced again (946-008 600 30) to leave the computer operational. Calibration and pass QC's. | Replace (946-008 600 30) to leave the computer operational. | Assisted Replacement | SC | SC | Lac
ABL90 | Lavagem do circuito com hipoclorito e água, substituição da entrada, sonda e do pack de soluções e verificação do funcionamento da válvula, tudo sem sucesso, substituição do da SC, inicialização e verificação do funcionamento com QC. Tudo Ok. | B | D068E4226C597C9E | No | Washing the circuit with hypochlorite and water, replacing the inlet, probe and the solution pack and checking the operation of the valve, all without success, replacing the SC, booting and checking the operation with QC. Okay, all right. |  | Assisted Replacement | SC | SC | Wetability
ABL835 FLEX | Unable to activate breaker in service program. | Replaced breaker motor. Realigned module to correct laser setting to 192. | D0684A916C377A26 | No | Unable to activate breaker in service program. | Replaced breaker motor. Realigned module to correct laser setting to 192. | Defect | HW | AutoCheck Module | Stepper motor, breaker arm
ABL827 FLEX | Verified failure once on site. Analyzer was failing for creatine A for level 2 QC. | Replaced all solutions. Performed 2 point calibration and calibration the problem is now resolved. | D06927FDE25BDE5C | No | Verified failure once on site. Analyzer was failing for creatine A for level 2 QC. | Replaced all solutions. Performed 2 point calibration and calibration the problem is now resolved. | Assisted Replacement | Consumables | Solution bottle | Solutions bottle
AQT90 FLEX | Dismantiled the inlet wheel and reconnected. Cleaned all sensors. Cleaned process wheel. | B | SWO-86918 | No | Dismantiled the inlet wheel and reconnected. Cleaned all sensors. Cleaned process wheel. |  | Cleaned | HW | Aspiration Module | Sensor
AQT90 FLEX | ｲﾝｷｭﾍﾞｰﾀｰPCB交換 | ｲﾝｷｭﾍﾞｰﾀｰPCB交換 | SWO-87227 | No | Incubator PCB Replacement | Incubator PCB Replacement | Defect | HW | Incubator module | PCB, Incubator | 2168
AQT90 FLEX |  | B | D06A9CCD5827AD73 | No |  |  | Not Cat | Not cat. | Not cat. | Not cat.
ABL80 OSM |  | B | D0698783E25456EE | No
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-86287 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL80 CO-OX | CONFIRMED ISSUE WHEN ONSITE. SUSPECT POWER SURGE HAS TAKEN UNIT DOWN. | REPLACED POWER SUPPLY AND LED IS NOW VISIBLE. HOWEVER, STILL NO VISUAL ON SCREEN. SWAP CPU UNIT BETWEEN UNITS AND WORKS - CONFIRMED CPU IS DOWN. PARTS ON ORDER AND FSE HUPCHEW WILL BRING SYSTEM WITH HIM FOR FIXING AND MAIL IT BACK TO CUSTOMER WHEN DONE. | D069AE81E21FC344 | No | CONFIRMED ISSUE WHEN ONSITE. SUSPECT POWER SURGE HAS TAKEN UNIT DOWN. | REPLACED POWER SUPPLY AND LED IS NOW VISIBLE. HOWEVER, STILL NO VISUAL ON SCREEN. SWAP CPU UNIT BETWEEN UNITS AND WORKS - CONFIRMED CPU IS DOWN. PARTS ON ORDER AND FSE HUPCHEW WILL BRING SYSTEM WITH HIM FOR FIXING AND MAIL IT BACK TO CUSTOMER WHEN DONE.
ABL80 CO-OX | This case will be closed due to lack of investigational data. Without FSE feedback, it cannot be determined that the waste line clamp was positioned incorrectly. | B | D069CC97E21C706F | No | This case will be closed due to lack of investigational data. Without FSE feedback, it cannot be determined that the waste line clamp was positioned incorrectly.
AQT90 FLEX | 更换636-550一个后修复(by andrew) | 更换636-550一个后修复(by andrew) | SWO-85880 | No | Replace 636-550 one post-fix (by-andrew) | Replace 636-550 one post-fix (by-andrew) | Defect | HW | Process Module | Process Unit sensor | 5784
AQT90 FLEX |  | B | SWO-86070 | No |  |  | Defect | HW | Process Module | Process Unit sensor | 5784
ABL837 FLEX |  | B | SWO-86182 | No |  |  | Defect | HW | FlexQ Module | FlexQ Module
ABL800 BASIC | 更换Lac电极膜修复仪器(by andrew) | 更换Lac电极膜修复仪器(by andrew) | SWO-86315 | No | Replace Lac electrode membrane repair instrument (by andrew) | Replace Lac electrode membrane repair instrument (by andrew) | Assisted Replacement | Electrode | Membrane | Lac | Calibration
ABL800 BASIC | 更换电极和电极膜修复仪器 by jesson | 更换电极和电极膜修复仪器 by jesson | SWO-87544 | Yes | Replace electrode and electrode membrane repair instrument by jesson | Replace electrode and electrode membrane repair instrument by jesson | Re classified RMED | Same day upSWO
ABL800 BASIC | replaced with new touch screen, done calibration for touch screen. | replaced with new touch screen, done calibration for touch screen. | SWO-86930 | No | replaced with new touch screen, done calibration for touch screen. | replaced with new touch screen, done calibration for touch screen. | Defect | HW | DMS_PC | Touch Screen | email
ABL825 FLEX |  | B | D06C587ACE00207D | No |  |  | Assisted Replacement | Electrode | Electrode | Lac | Unknown
ABL815 FLEX | Kontrolle Bildschirmkabelanschluss, Prüfung LW C:, Kontrolle und Korrektur Netzwerkkartenkonfiguration, Probelauf -> Beobachtet durch Kunden, Kunde meldet sich | B | SWO-86573 | No | Control screen cable connection, check LW C:, control and correction of NETWORK card configuration, trial run -> Observed by customer, customer reports |  | Cleaned | HW | DMS_PC | Touch Screen
ABL820 FLEX | Probleme mit oxymetrie nach neugeräteaufstellung. Altes oxymetriemodul ausgebaut.(R1969N001) Neues Oxymetrimodul eingebaut. (R0081N005) Complete Oxymetrietest durchgeführt --> i.O. Mehrere Kalibrierungen durchgeführt --> i.O. Service tHb Kal durchgeführt. THb Kal durchgeführt --> i.O. Kal Verifizierung duchführen --> i.O. QCs für alle Level durchgeführt --> i.O. ABL Messbereit. | Probleme mit oxymetrie nach neugeräteaufstellung. Altes oxymetriemodul ausgebaut.(R1969N001) Neues Oxymetrimodul eingebaut. (R0081N005) Complete Oxymetrietest durchgeführt --> i.O. Mehrere Kalibrierungen durchgeführt --> i.O. Service tHb Kal durchgeführt. THb Kal durchgeführt --> i.O. Kal Verifizierung duchführen --> i.O. QCs für alle Level durchgeführt --> i.O. ABL Messbereit. | SWO-86259 | Yes | Problems with oxymetry after repositioning. Old oxymetry module removed. (R1969N001) New oxymetrimodul installed. (R0081N005) Complete oxymetry test performed --> i.O. Several calibrations performed --> i.O. Service tHb Kal performed. THb Kal performed --> i.O. Kal verification by conducting --> i.O. QCs performed for all levels --> i.O. ABL Ready to measure. | Problems with oxymetry after repositioning. Old oxymetry module removed. (R1969N001) New oxymetrimodul installed. (R0081N005) Complete oxymetry test performed --> i.O. Several calibrations performed --> i.O. Service tHb Kal performed. THb Kal performed --> i.O. Kal verification by conducting --> i.O. QCs performed for all levels --> i.O. ABL Ready to measure. | Re classified RMED | Installation, OOB
AQT90 FLEX | Cleaned the barcode encapsulation. Performed new lot calibration. Checked the barcode reader using the service program | B | SWO-86919 | No | Cleaned the barcode encapsulation. Performed new lot calibration. Checked the barcode reader using the service program |  | Cleaned | HW | Storage Module | Barcode reader, Storage wheel
ABL80 FLEX BASIC | 更换新的测试卡修复仪器 | 更换新的测试卡修复仪器 | SWO-87246 | No | Replace a new test card repair instrument | Replace a new test card repair instrument
ABL80 FLEX BASIC | 更换新的测试卡修复仪器 | 更换新的测试卡修复仪器 | SWO-86517 | No | Replace a new test card repair instrument | Replace a new test card repair instrument
ABL80 FLEX BASIC |  | B | SWO-86293 | No
ABL80 FLEX BASIC |  | B | SWO-86274 | No
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-87838 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL80 FLEX BASIC | Reconnected Battery and SMPS connection. Then checked with voltages. Instrument is working fine. | Reconnected Battery and SMPS connection. Then checked with voltages. Instrument is working fine. | SWO-86899 | No | Reconnected Battery and SMPS connection. Then checked with voltages. Instrument is working fine. | Reconnected Battery and SMPS connection. Then checked with voltages. Instrument is working fine.
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-87013 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL90 | This unit appears to be performing SC maintenance more often than other units. | Called Techline for support. | D06B16EA581F476B | No | This unit appears to be performing SC maintenance more often than other units. | Called Techline for support. | Unknown | Revisit required |  |  | 1228
ABL90 FLEX | ABL manuell gespült. Einlasswieder zusammengebaut und angebaut. Einlassdichtung gereinigt. Nach Startphase noch Kalibrierfehler pCO2. Sensorkassette aus GS geholt und eingesetzt, da keine mehr lagernd Vorort. Kalibration und QC's durchgeführt. | ABL manuell gespült. Einlasswieder zusammengebaut und angebaut. Einlassdichtung gereinigt. Nach Startphase noch Kalibrierfehler pCO2. Sensorkassette aus GS geholt und eingesetzt, da keine mehr lagernd Vorort. Kalibration und QC's durchgeführt. | SWO-86437 | No | ABL manually rinsed. Inlet reassembled and assembled. Inlet seal cleaned. After start-up phase still calibration error pCO2. Sensor cassette fetched from GS and used as no longer in stock suburb. calibration and QC's. | ABL manually rinsed. Inlet reassembled and assembled. Inlet seal cleaned. After start-up phase still calibration error pCO2. Sensor cassette fetched from GS and used as no longer in stock suburb. calibration and QC's. | Cleaned | Fluidic pathway
ABL90 FLEX | 更换新试剂包修复仪器 | 更换新试剂包修复仪器 | SWO-86515 | No | Replace the new reagent kit repair instrument | Replace the new reagent kit repair instrument | Assisted Replacement | SP | FT | Unable to pump | 1343
AQT90 FLEX |  | B | SWO-86322 | No |  |  | Defect | HW | Process Module | Optical Unit | 131
ABL800 FLEX | Oxymetry Module Replaced and software reinstalled. All old data and settings removed. Re-insert all settings and parameters. All working well. Oxy module response satisfactory. Faulty Oxymetry module collected and PER made for replacement. | Oxymetry Module Replaced and software reinstalled. All old data and settings removed. Re-insert all settings and parameters. All working well. Oxy module response satisfactory. Faulty Oxymetry module collected and PER made for replacement. | SWO-86622 | No | Oxymetry Module Replaced and software reinstalled. All old data and settings removed. Re-insert all settings and parameters. All working well. Oxy module response satisfactory. Faulty Oxymetry module collected and PER made for replacement. | Oxymetry Module Replaced and software reinstalled. All old data and settings removed. Re-insert all settings and parameters. All working well. Oxy module response satisfactory. Faulty Oxymetry module collected and PER made for replacement. | Defect | HW | Oxi module | Oxi module
ABL815 FLEX | Kundengerät meßbereit aufgestellt, Leihgerät in die GS mitgenommen. | Nach Rücksprache mit KDB UWEBA holt dieser das Gerät wieder ab und stellt es beim Kunden wieder auf. | SWO-86653 | Yes | Customer device set up ready for measurement, borrower taken to the GS. | After consultation with KDB UWEBA, the customer picks up the device and re-establishes it at the customer's place. | Re classified RMED | Analyzer re installed
ABL825 FLEX | Reviewed data logs for any issues. | Installed software version 6.16 for diagnostic reasons. | D06AED725829D23F | No | Reviewed data logs for any issues. | Installed software version 6.16 for diagnostic reasons. | Unknown | SW | Analyzer freeze | Analyzer lock up
AQT90 FLEX | changed the optical unit, now working fine | changed the optical unit, now working fine | SWO-90858 | No | changed the optical unit, now working fine | changed the optical unit, now working fine | Defect | HW | Process Module | Optical Unit
ABL80 FLEX BASIC | 更换新的测试卡修复仪器 | 更换新的测试卡修复仪器 | SWO-87278 | No | Replace a new test card repair instrument | Replace a new test card repair instrument
ABL80 FLEX BASIC | 更换新的测试卡修复仪器 | 更换新的测试卡修复仪器 | SWO-86789 | No | Replace a new test card repair instrument | Replace a new test card repair instrument
ABL80 FLEX BASIC | 更换新的测试卡修复仪器 | 更换新的测试卡修复仪器 | SWO-86708 | No | Replace a new test card repair instrument | Replace a new test card repair instrument
ABL80 FLEX BASIC | 更换新的测试卡修复仪器 | 更换新的测试卡修复仪器 | SWO-86712 | No | Replace a new test card repair instrument | Replace a new test card repair instrument
ABL80 FLEX CO-OX | Gerät zerlegt. Stutzen an der Manifold Unit defekt. Flüssigkeit ist aufgrund von Undichtigkeiten am Stutzen ausgelaufen. | Gerät zerlegt. Stutzen an der Manifold Unit defekt. Flüssigkeit ist aufgrund von Undichtigkeiten am Stutzen ausgelaufen. | SWO-86854 | No | device disassembled. Nozzle on the Manifold Unit is defective. Liquid has leaked due to leaks on the nozzle. | device disassembled. Nozzle on the Manifold Unit is defective. Liquid has leaked due to leaks on the nozzle.
ABL90 FLEX | Für den TSP den SD und Data Log gezogen | Für den TSP den SD und Data Log gezogen | SWO-87045 | Yes | For the TSP, the SD and Data Log are drawn | For the TSP, the SD and Data Log are drawn | Re classified RMED | Only to collect Data
ABL90 FLEX | SC 90 demontiert, Messkammer gereinigt und geprüft, neue SC 90 installiert, Funktionskontrolle ok | Dem Kunden geraten die SC zu erneuern. Der Kunde bittet um einen Service, da der Techniker bereits geplant habe ins Haus zu kommen. | SWO-86755 | No | SC 90 dismantled, measuring chamber cleaned and tested, new SC 90 installed, function control ok | Customers are advised to renew the SC. The customer asks for a service, as the technician has already planned to come into the house. | Assisted Replacement | SC | SC | Leaking REF
ABL90 |  | B | D06C531BCE0BF5E6 | No |  |  | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
ABL90 FLEX | 更换新的测试卡修复仪器 | 更换新的测试卡修复仪器 | SWO-86711 | No | Replace a new test card repair instrument | Replace a new test card repair instrument | Assisted Replacement | SC | SC | Lac
ABL90 FLEX | 更换新测试卡修复仪器 | 更换新测试卡修复仪器 | SWO-86500 | No | Replace the new test card repair instrument | Replace the new test card repair instrument | Assisted Replacement | SC | SC | pCO2
ABL825 FLEX |  | B | D06C2FBCCE1EAE72 | No |  |  | Defect | HW | pH/BG Module | Valve
ABL800 FLEX | Reference mam-brain replaced and calibrate. K and Ca mambrain replaced and check. Analyser left on observation. for auto cal. | Reference electrode along with mambrain replaced by a 6 month old electrode. but same problem reported. Solutions replaced- same problem reported. Go thrugh service pprogram everything fine found. Need to test with a new reference electrode. TO sending. tHb calibration Ampules also required as New Oxy Module Relaced from Mumbai Stoke Saprate TO sending for Mumbai Stok Oxymetry Module and tHb calibration ampules. | SWO-86941 | No | Reference mam-brain replaced and calibrate. K and Ca mambrain replaced and check. Analyser left on observation. for auto cal. | Reference electrode along with mambrain replaced by a 6 month old electrode. but same problem reported. Solutions replaced- same problem reported. Go thrugh service pprogram everything fine found. Need to test with a new reference electrode. TO sending. tHb calibration Ampules also required as New Oxy Module Relaced from Mumbai Stoke Saprate TO sending for Mumbai Stok Oxymetry Module and tHb calibration ampules. | Revisit required | HW | Oxi module | Oxi module
ABL827 FLEX | Failure verified once on site. | Performed troubleshooting to find the root cause. Level 2 QC was failing for CREA A. Re-membraned both A and B Electrodes, cleaned both electrodes also. Level 2 QC came in with no problem. Verified that the CREA was still good at 11:20pm. | D06D08BC444E2916 | No | Failure verified once on site. | Performed troubleshooting to find the root cause. Level 2 QC was failing for CREA A. Re-membraned both A and B Electrodes, cleaned both electrodes also. Level 2 QC came in with no problem. Verified that the CREA was still good at 11:20pm. | Assisted Replacement | Electrode | Membrane | Crea
ABL810 FLEX | Analysator überprüft, pO2 und pCO2 Membranen sehen undicht aus, um den Dichtring sieht die Membran rissig aus. PCO2 Membranen Lot:729 laut Daten schlecht in QCs. pO2 und pCO2 Membranen erneuert. QC Barcode >V6.16 für S7345-331 eingelesen. | Analysator überprüft, pO2 und pCO2 Membranen sehen undicht aus, um den Dichtring sieht die Membran rissig aus. PCO2 Membranen Lot:729 laut Daten schlecht in QCs. pO2 und pCO2 Membranen erneuert. QC Barcode >V6.16 für S7345-331 eingelesen. | SWO-86657 | No | Analyzer checked, pO2 and pCO2 membranes look leaking, around the sealing ring the membrane looks cracked. PCO2 Membranes Lot:729 according to data badly renewed in QCs. pO2 and pCO2 membranes. QC Barcode >V6.16 for S7345-331. | Analyzer checked, pO2 and pCO2 membranes look leaking, around the sealing ring the membrane looks cracked. PCO2 Membranes Lot:729 according to data badly renewed in QCs. pO2 and pCO2 membranes. QC Barcode >V6.16 for S7345-331. | Assisted Replacement | Electrode | Membrane | pCO2
ABL825 FLEX | Customer remembraned the pO2 electrode still had the same problem. | Customer replaced the pO2 electrode. | D06C88AACE466385 | No | Customer remembraned the pO2 electrode still had the same problem. | Customer replaced the pO2 electrode. | Assisted Replacement | Electrode | Electrode | pO2
AQT90 FLEX | Ausbau Scanner Unit => Reinigung der internen Umlenkspiegel und des Sichtfensters Scanner Unit neu justiert QC Barcodes für CRP (Lot 11272) und D-Dimer (Lot 10999) eingelesen | Ausbau Scanner Unit => Reinigung der internen Umlenkspiegel und des Sichtfensters Scanner Unit neu justiert QC Barcodes für CRP (Lot 11272) und D-Dimer (Lot 10999) eingelesen | SWO-86645 | No | Removal Scanner Unit => Cleaning of the internal deflection mirrors and the viewing window Scanner Unit re-adjusted QC barcodes for CRP (Lot 11272) and D-Dimer (Lot 10999) read in | Removal Scanner Unit => Cleaning of the internal deflection mirrors and the viewing window Scanner Unit re-adjusted QC barcodes for CRP (Lot 11272) and D-Dimer (Lot 10999) read in | Cleaned | HW | Aspiration Module | Barcode reader, Inlet
AQT90 FLEX | Processwheel ausgebaut und versucht Motor zu trocknen, um das Gerät messbereit zu bekommen, bis der Motor ausgetauscht wird. Ohne Erfolg. Motor ist bestellt und muss am 11.03.16 erneuert werden. | Processwheel ausgebaut und versucht Motor zu trocknen, um das Gerät messbereit zu bekommen, bis der Motor ausgetauscht wird. Ohne Erfolg. Motor ist bestellt und muss am 11.03.16 erneuert werden. | SWO-86774 | No | Processwheel removed and tries to dry the motor to get the device ready for measurement until the motor is replaced. Without success. Engine is ordered and must be renewed on 11.03.16. | Processwheel removed and tries to dry the motor to get the device ready for measurement until the motor is replaced. Without success. Engine is ordered and must be renewed on 11.03.16. | Revisit required | HW | Process Module | Process wheel, motor
AQT90 FLEX | Da Pumpenkopf Geräusche macht, erneuert. Angeschlossene Schläuche, Nr. 6 und 3 Durchgängigkeit geprüft mit Spritze, o. K. Keine Behinderung. Wartungsmassnahme fällig, gleich durchgeführt, siehe extra SWO. Vmtl. Präanalytischer Fehler. Beim Anschluss eines Sarstedt Multiadapters an ein EDTA Röhrchen ist das Heparin in den Adapter gedrückt worden, oder geflossen. Damit hat das EDTA Röhrchen zu wenig Heparin und eine Gerinnsel wurde eingesaugt und Probe abgebrochen. Reinigungsmaßnahme notwendig durchgeführt. Systemüberprüfung o.K., Qualitätskontrolle o. K. | Da Pumpenkopf Geräusche macht, erneuert. Angeschlossene Schläuche, Nr. 6 und 3 Durchgängigkeit geprüft mit Spritze, o. K. Keine Behinderung. Wartungsmassnahme fällig, gleich durchgeführt, siehe extra SWO. Vmtl. Präanalytischer Fehler. Beim Anschluss eines Sarstedt Multiadapters an ein EDTA Röhrchen ist das Heparin in den Adapter gedrückt worden, oder geflossen. Damit hat das EDTA Röhrchen zu wenig Heparin und eine Gerinnsel wurde eingesaugt und Probe abgebrochen. Reinigungsmaßnahme notwendig durchgeführt. Systemüberprüfung o.K., Qualitätskontrolle o. K. | SWO-86687 | No | Since pump head makes noises, renewed. Connected hoses, No. 6 and 3 Continuity tested with syringe, o.k. No obstruction. Maintenance measure due, immediately carried out, see extra SWO. Vmtl. Pre-analytical error. When connecting a Sarstedt multiadapter to an EDTA tube, the heparin has been pressed into the adapter, or flowed. Thus, the EDTA tube has too little heparin and a clot was sucked in and sample aborted. Cleaning action necessary. System check o.K., quality control o.k. | Since pump head makes noises, renewed. Connected hoses, No. 6 and 3 Continuity tested with syringe, o.k. No obstruction. Maintenance measure due, immediately carried out, see extra SWO. Vmtl. Pre-analytical error. When connecting a Sarstedt multiadapter to an EDTA tube, the heparin has been pressed into the adapter, or flowed. Thus, the EDTA tube has too little heparin and a clot was sucked in and sample aborted. Cleaning action necessary. System check o.K., quality control o.k. | Defect | SW | Analyzer Control Module | Corrupt XML | 35
AQT90 FLEX |  | B | D079A655E0043F7B | No |  |  | Revisit required | HW | Analyzer Control Module | Cable (CPU-Cover)
AQT90 FLEX | Nach Fehlersuche. Herausgefunden,dass die PCB für das Incubator Wheel defekt ist. Frau Behnke über den Zustand des Analysator informiert und das Ersatzteil für morgenfrüh bestellt. | Nach Fehlersuche. Herausgefunden,dass die PCB für das Incubator Wheel defekt ist. Frau Behnke über den Zustand des Analysator informiert und das Ersatzteil für morgenfrüh bestellt. | SWO-86613 | No | After troubleshooting. Found that the PCB for the Incubator Wheel is defective. Ms. Behnke informed about the condition of the analyzer and ordered the spare part for tomorrow morning. | After troubleshooting. Found that the PCB for the Incubator Wheel is defective. Ms. Behnke informed about the condition of the analyzer and ordered the spare part for tomorrow morning. | Revisit required | HW | Incubator module | PCB | 2168
ABL80 FLEX BASIC | 更换测试卡故障依旧，更换接口板902-972后，机器正常。By Jesson | 更换测试卡故障依旧，更换接口板902-972后，机器正常。By Jesson | SWO-88111 | No | The replacement test card fault is still the same, after replacing the interface plate 902-972, the machine is normal. By Jesson | The replacement test card fault is still the same, after replacing the interface plate 902-972, the machine is normal. By Jesson
ABL80 FLEX BASIC | 更换测试卡正常。 | 更换测试卡正常。 | SWO-86921 | No | The replacement test card is OK. | The replacement test card is OK.
ABL80 CO-OX |  | INSTALLED CPU UNIT AND PERFORM VK TESTING PER CAS TIM RUFFIN'S ADVICE. | D06D66D4441EC366 | No |  | INSTALLED CPU UNIT AND PERFORM VK TESTING PER CAS TIM RUFFIN'S ADVICE.
ABL90 | Unit displaying errors 1228, 1340,1070 for parameters missing values after measurement. SC maintenance by Analyzer is occuring more often on this particulat Unit. No indication or any message is being displayed on screen while the SC maintenance is occuring. | Replace Sensor Interface Bd. | D06D85FA44539BEB | No | Unit displaying errors 1228, 1340,1070 for parameters missing values after measurement. SC maintenance by Analyzer is occuring more often on this particulat Unit. No indication or any message is being displayed on screen while the SC maintenance is occuring. | Replace Sensor Interface Bd. | Defect | HW | Sensor cassette PCB | Ca, Na
ABL90 FLEX | Einstellungen in Controlpanel überprüft und auf CF Karte gespeichert. tHb Kalibration, Kalibration und QC's durchgeführt. | Einstellungen in Controlpanel überprüft und auf CF Karte gespeichert. tHb Kalibration, Kalibration und QC's durchgeführt. | SWO-86850 | No | Settings checked in Controlpanel and stored on CF card. tHb calibration, calibration and QC's. | Settings checked in Controlpanel and stored on CF card. tHb calibration, calibration and QC's. | User error | tHb Cal exp | tHb Cal exp | tHb Cal exp
ABL90 FLEX | Kein weiterer LP mit anderer Lot Nummer auf Lager. Verwendung einer neues Sensorkassette. Qualitätskontroll Wert wieder grenzwertig, aber keine QC Verletzung. Anwender informiert, ein anderen LP mit anderer Los ynummer zu bestellen. Kalibrationen und Qualitätskontrolle o. K. | Kein weiterer LP mit anderer Lot Nummer auf Lager. Verwendung einer neues Sensorkassette. Qualitätskontroll Wert wieder grenzwertig, aber keine QC Verletzung. Anwender informiert, ein anderen LP mit anderer Los ynummer zu bestellen. Kalibrationen und Qualitätskontrolle o. K. | SWO-86896 | No | No other LP with another lot number in stock. Use of a new sensor cartridge. Quality control value again borderline, but no QC violation. informs users to order another LP with a different lot y number. Calibrations and quality control o.K. | No other LP with another lot number in stock. Use of a new sensor cartridge. Quality control value again borderline, but no QC violation. informs users to order another LP with a different lot y number. Calibrations and quality control o.K. | Assisted Replacement | SC | SC | Glu QC
ABL90 FLEX | Wet Section module gedemonteerd en alle connectoren nagekeken : geen afwijkingen vastgesteld. Na montage en werd correct opgestart, thermostatisering verloopt normaal. Calibratie met fouten voor Ca, Cl, Glu, Lact. SC vervangen met een van goed werkende ABL. Probleem met temperatuurfout blijft, vesloten om ABL mee te nemen voor herstelling op bedrijf. | B | SWO-86844 | No | Wet Section module disassembled and checked all connectors: no deviations established. After assembly and started correctly, Thermostatisation is normal. Calibration with errors for Ca, Cl, Glu, Lact. SC replaced with one of well working ABL. Problem with temperature error remains, to take ABL with you for repairs on the farm. |  | Unknown | Removed by RSUB
ABL90 |  | B | D06D94814430A1AE | No |  |  | Defect | HW | Printer module | Printer module
ABL90 FLEX | Sensor Flow Selector gereinigt. tHb Kalibration, Kalibration und QC's vorgenommen. | Sensor Flow Selector gereinigt. tHb Kalibration, Kalibration und QC's vorgenommen. | SWO-86849 | No | Sensor Flow Selector. tHb calibration, calibration and QC's. | Sensor Flow Selector. tHb calibration, calibration and QC's. | Adjustment | HW | Flow selector | Clean
ABL90 FLEX PLUS |  | B | D079E65FE01D99F8 | Yes |  |  | Re classified RMED | Logistical upSWO
ABL90 FLEX PLUS | Einlasskunststoffteil ersetzt, Kontakte über der SC gereinigt, Kal, QC, gleicher Fehler, Wechsel LP, Kal, gleicher Glukose Fehler im QC, nach Rücksprache mit Dr. soll Parameter nicht deaktiviert werden, der Hinweis auf Fehler ist auf dem Ausdruck sichtbar, sofern Ersatzteil verfügbar, dann Fortsetzung der Reparatur | Einlasskunststoffteil ersetzt, Kontakte über der SC gereinigt, Kal, QC, gleicher Fehler, Wechsel LP, Kal, gleicher Glukose Fehler im QC, nach Rücksprache mit Dr. soll Parameter nicht deaktiviert werden, der Hinweis auf Fehler ist auf dem Ausdruck sichtbar, sofern Ersatzteil verfügbar, dann Fortsetzung der Reparatur | SWO-86962 | No | Inlet plastic part replaced, contacts cleaned over the SC, Kal, QC, same error, change LP, kal, same glucose error in the QC, after consultation with Dr. should not deactivate parameters, the indication of errors is visible on the printout, provided spare part available, then continue the repair | Inlet plastic part replaced, contacts cleaned over the SC, Kal, QC, same error, change LP, kal, same glucose error in the QC, after consultation with Dr. should not deactivate parameters, the indication of errors is visible on the printout, provided spare part available, then continue the repair | Unknown | Revisit required | SC | SC
AQT90 FLEX | 更换925-016修复仪器(by andrew) | B | SWO-86578 | No | Replace the 925-016 repair instrument (by andrew) |  | Defect | HW | Cover and Chassis | Storage Door
ABL800 FLEX | Auto Check barcode reader need to replace | B | SWO-86905 | No | Auto Check barcode reader need to replace |  | Defect | HW | AutoCheck Module | Barcode reader
ABL825 FLEX | Austausch CPU Unit | Austausch CPU Unit | SWO-86796 | No | Replacement CPU Unit | Replacement CPU Unit | User error | HW | DMS_PC | Main board PCB
AQT90 FLEX | Stepper Motor Process Wheel erneuert | Stepper Motor Process Wheel erneuert | SWO-86835 | No | Stepper Motor Process Wheel renewed | Stepper Motor Process Wheel renewed | Defect | HW | Process Module | Process wheel, motor
AQT90 FLEX | Schläuche der Cup Wash Unit überprüft und getauscht, Probelauf ok LQC aller 3 Parameter ok | Schläuche der Cup Wash Unit überprüft und getauscht, Probelauf ok LQC aller 3 Parameter ok | SWO-86898 | Yes | Hoses of the Cup Wash Unit checked and exchanged, test run ok LQC all 3 parameters ok | Hoses of the Cup Wash Unit checked and exchanged, test run ok LQC all 3 parameters ok | Re classified RMED | Same day upSWO
AQT90 FLEX | Inkubator-PCB erneuert. | Inkubator-PCB erneuert. | SWO-86815 | No | Incubator PCB. | Incubator PCB. | Defect | HW | Incubator module | PCB | 2168
ABL80 FLEX BASIC | 更换新测试卡修复仪器 | 更换新测试卡修复仪器 | SWO-88394 | No | Replace the new test card repair instrument | Replace the new test card repair instrument
ABL80 FLEX BASIC | 连续更换两张测试卡后解决 | 连续更换两张测试卡后解决 | SWO-87012 | No | Resolve after replacing two test cards in a row | Resolve after replacing two test cards in a row
ABL80 FLEX BASIC | 溶液パック交換 | B | SWO-86946 | No | Solution pack replacement
ABL80 FLEX BASIC | 检查接口板902-972发现接口有腐蚀情况，更换902-972后，做零点校准。重新做质控，质控结果仍然偏高，更换903-334后，并进行零点校准后，结果正常。by Jesson | 检查接口板902-972发现接口有腐蚀情况，更换902-972后，做零点校准。重新做质控，质控结果仍然偏高，更换903-334后，并进行零点校准后，结果正常。by Jesson | SWO-86979 | No | Check the interface board 902-972 to find corrosion of the interface, replace 902-972, do zero calibration. Redo the quality control, the quality control results are still high, after replacing 903-334, and zero-point calibration, the results are normal. by Jesson | Check the interface board 902-972 to find corrosion of the interface, replace 902-972, do zero calibration. Redo the quality control, the quality control results are still high, after replacing 903-334, and zero-point calibration, the results are normal. by Jesson
ABL80 FLEX BASIC | Replaced power supply. Machine is working perfectly without any issue. | Checked with another power point & power cable. Checked LED of the power supply. Analyzer not starting. Need to check with another power supply. | SWO-88247 | No | Replaced power supply. Machine is working perfectly without any issue. | Checked with another power point & power cable. Checked LED of the power supply. Analyzer not starting. Need to check with another power supply.
ABL90 FLEX | 更换新的试剂包 | 更换新的试剂包 | SWO-86989 | No | Replace a new reagent kit | Replace a new reagent kit
AQT90 FLEX |  | B | SWO-86947 | No |  |  | Defect | HW | Aspiration Module | Inlet lock | SetSafe 22
ABL80 FLEX BASIC | 更换新的测试卡修复仪器 | 更换新的测试卡修复仪器 | SWO-87306 | No | Replace a new test card repair instrument | Replace a new test card repair instrument
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-88343 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL80 FLEX BASIC |  | B | SWO-87021 | No
ABL80 FLEX BASIC | 更换测试卡修复仪器 | 更换测试卡修复仪器 | SWO-87243 | No | Replace the test card repair instrument | Replace the test card repair instrument
ABL80 FLEX BASIC |  | B | SWO-86976 | No
ABL80 FLEX BASIC | Sensor cassette failed with V C failed, found Sensor Board connector problem need to replace with new. | Sensor cassette installation got failure, due to Sensor cassette got leakage of solution, instrument was showing Temp-A error, Temp-B, Suggested to customer need to replace with new SC. | SWO-88272 | No | Sensor cassette failed with V C failed; found Sensor Board connector problem need to replace with new. | Sensor cassette installation got failure, due to Sensor cassette got leakage of solution, instrument was showing Temp-A error, Temp-B, Suggested to customer need to replace with new SC.
ABL80 FLEX BASIC |  | B | SWO-86991 | No
ABL80 FLEX Basic version |  | B | D0714285A634F4E6 | Yes |  |  | Re classified RMED | not upSWO
ABL90 FLEX | Gerät überprüft, Kontakte und Sensorkassette kontrolliert, Sensorkassette noch mal erneuert | Gerät überprüft, Kontakte und Sensorkassette kontrolliert, Sensorkassette noch mal erneuert | SWO-87178 | Yes | Device checked, contacts and sensor cassette checked, sensor cassette renewed again | Device checked, contacts and sensor cassette checked, sensor cassette renewed again | Re classified RMED | Same day upSWO
ABL90 FLEX | Komplette Geräteeinstellungen auf der Windowsebene inkl. Administratoroberfläche durchgeführt. Und auf der Speicherkarte mehrmals gespeichert (ewfmgr c: -commit). Analysatorsoftware neu installiert. Eine Umstellung auf die deutsche Tastatur ist möglich, aber nicht die Speicherung. | Komplette Geräteeinstellungen auf der Windowsebene inkl. Administratoroberfläche durchgeführt. Und auf der Speicherkarte mehrmals gespeichert (ewfmgr c: -commit). Analysatorsoftware neu installiert. Eine Umstellung auf die deutsche Tastatur ist möglich, aber nicht die Speicherung. | SWO-87285 | Yes | Complete device settings performed at the Windows level including the administrator interface. And stored several times on the memory card (ewfmgr c: -commit). Reinstalled analyzer software. A changeover to the German keyboard is possible, but not storage. | Complete device settings performed at the Windows level including the administrator interface. And stored several times on the memory card (ewfmgr c: -commit). Reinstalled analyzer software. A changeover to the German keyboard is possible, but not storage. | Re classified RMED | Same day upSWO
ABL90 FLEX | Kunde meldet sich am 15.03.2016 zwischen 07:30 und 08:00 Uhr, wenn der Analysator Kalibration während der Messung zu Optimierung immer noch durchführt. Dieses wird am 15.03.2016 durch SönLe an den Kunden kommuniziert. | Kunde meldet sich am 15.03.2016 zwischen 07:30 und 08:00 Uhr, wenn der Analysator Kalibration während der Messung zu Optimierung immer noch durchführt. Dieses wird am 15.03.2016 durch SönLe an den Kunden kommuniziert. | SWO-87289 | No | Customer will report between 07:30 and 08:00 on March 15, 2016, when the analyzer is still performing calibration during the optimization measurement. This will be communicated to the customer by SönLe on 15.03.2016. | Customer will report between 07:30 and 08:00 on March 15, 2016, when the analyzer is still performing calibration during the optimization measurement. This will be communicated to the customer by SönLe on 15.03.2016. | Defect | SW | Keyboard Language | Keyboard Language
ABL90 FLEX PLUS | Lösungspack 944-197 UN-01 erneuert. | Lösungspack 944-197 UN-01 erneuert. | SWO-87260 | Yes | Solution pack 944-197 UN-01 renewed. | Solution pack 944-197 UN-01 renewed. | Re classified RMED | OOB, Demo
ABL90 FLEX | 工程师报错机器序列号，取消。 | 工程师报错机器序列号，取消。 | SWO-86988 | No | The engineer incorrectly reported the machine serial number and cancelled it. | The engineer incorrectly reported the machine serial number and cancelled it. | Defect | HW | Sensor Detector Plate | Sensor Detector Plate
ABL90 FLEX | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-87242 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL825 FLEX | Fehlersuche, Reinigung Elektrodenkontakte, Neumembranierung, Erneuerung Elektrode, Funktionskontrolle | Fehlersuche, Reinigung Elektrodenkontakte, Neumembranierung, Erneuerung Elektrode, Funktionskontrolle | SWO-87130 | No | Troubleshooting, cleaning electrode contacts, re-membraneing, renewal electrode, function control | Troubleshooting, cleaning electrode contacts, re-membraneing, renewal electrode, function control | Assisted Replacement | Electrode | Electrode | pCO2
ABL825 FLEX | 更换新的测试卡修复仪器 | 更换新的测试卡修复仪器 | SWO-87253 | No | Replace a new test card repair instrument | Replace a new test card repair instrument
ABL825 FLEX |  | B | D07A7128E00BD061 | No |  |  | Assisted Replacement | Electrode | Electrode | pH | Sealing ring
ABL800 BASIC | Run the pump calibration . Run the ls test. Ls fail. check the ph bg lower LS.clean the chamber & tubing.run protein removal multiple time. clot the ph bg lower tubing. causing rinse error. | B | SWO-87256 | No | Run the pump calibration . Run the ls test. Ls fail. check the ph bg lower LS.clean the chamber & tubing.run protein removal multiple time. clot the ph bg lower tubing. causing rinse error. |  | Cleaned | Fluidic pathway | pH/BG Module | Clot
ABL815 FLEX | Gerät lief schon wieder, keinen Fehler festgestellt, 7 Kalibrationen mit dem Oxi.fehler, Oxi.kompl.test durchgeführt alles o.k., | Gerät lief schon wieder, keinen Fehler festgestellt, 7 Kalibrationen mit dem Oxi.fehler, Oxi.kompl.test durchgeführt alles o.k., | SWO-87194 | No | Device was already running again, no error detected, 7 calibrations with the Oxi.error, Oxi.kompl.test performed all ok, | Device was already running again, no error detected, 7 calibrations with the Oxi.error, Oxi.kompl.test performed all ok, | No fault found | HW | Oxi module | Oxi module
ABL80 FLEX BASIC | 更换新的测试卡修复仪器 | 更换新的测试卡修复仪器 | SWO-87305 | No | Replace a new test card repair instrument | Replace a new test card repair instrument
ABL80 FLEX BASIC | Received the Sensor Board connector from Standby instrument Medisun, replaced the SBC done the zeroing reinstalled the SP SC found both consumables are working fine. | Sensor cassette installation got failure, due to Sensor cassette got leakage of solution, instrument was showing Temp-A error, Temp-B, Suggested to customer need to replace with new SC. | SWO-88273 | No | Received the Sensor Board connector from Standby instrument Medisun, replaced the SBC done the zeroing reinstalled the SP SC found both consumables are working fine. | Sensor cassette installation got failure, due to Sensor cassette got leakage of solution, instrument was showing Temp-A error, Temp-B, Suggested to customer need to replace with new SC.
ABL80 FLEX BASIC | 更换新测试卡修复仪器 | 更换新测试卡修复仪器 | SWO-90299 | No | Replace the new test card repair instrument | Replace the new test card repair instrument
ABL80 FLEX BASIC | 更换新的测试卡修复仪器 | 更换新的测试卡修复仪器 | SWO-87520 | No | Replace a new test card repair instrument | Replace a new test card repair instrument
ABL90 FLEX | Lösungspack erneuert, Kals und QCs durchgeführt | Gerät überprüft, Kontakte und Sensorkassette kontrolliert, Sensorkassette noch mal erneuert | SWO-87446 | No | Solution pack renewed, Kals and QCs performed | Device checked, contacts and sensor cassette checked, sensor cassette renewed again | Assisted Replacement | SC | SC | pH
ABL90 | Analyzers Roam. Obtained Wireless information | Enabled Trace Log on analyzer and rebooted. Obtained Event logs for investigation Sifted through logs and worked with Sam to get Netop Rmeote access to work to analyzer. Investigated logs and will obtain trace tomorrow if anlayzers drop connectivity. Found IPV6 and 4 trying to obtain Addresses. Reached out to wireless specialist to get more info on IPV6. Also believe WZC is causing a conflict by not jumping to the next known good AP as the SSID is not brodcasted. Will investigate further tomorrow. | D072E67E1C4D8768 | No | Analyzers Roam. Obtained Wireless information | Enabled Trace Log on analyzer and rebooted. Obtained Event logs for investigation Sifted through logs and worked with Sam to get Netop Rmeote access to work to analyzer. Investigated logs and will obtain trace tomorrow if anlayzers drop connectivity. Found IPV6 and 4 trying to obtain Addresses. Reached out to wireless specialist to get more info on IPV6. Also believe WZC is causing a conflict by not jumping to the next known good AP as the SSID is not brodcasted. Will investigate further tomorrow. | Unknown | Revisit required | SW | Wireless connection | TU was taken into account/Rob
ABL90 | Unable to get SC to calibrate. | Replace suspect SC. | D07250421C200A0C | No | Unable to get SC to calibrate. | Replace suspect SC. | Assisted Replacement | SC | SC | Lac
ABL90 FLEX PLUS | Heute früh wurde die SC ersetzt und seitdem ist Glukose wieder OK bei QC Level C, Kal, QC | Heute früh wurde die SC ersetzt und seitdem ist Glukose wieder OK bei QC Level C, Kal, QC | SWO-87381 | No | This morning the SC was replaced and since then glucose is OK again at QC Level C, Kal, QC | This morning the SC was replaced and since then glucose is OK again at QC Level C, Kal, QC | Assisted Replacement | SC | SC | Glu
ABL90 FLEX PLUS |  | B | D072919E1C522234 | Yes |  |  | Re classified RMED | Logistical upSWO
ABL90 FLEX | 更换924-609，580-083后修复机器 by jesson | 更换924-609，580-083后修复机器 by jesson | SWO-86994 | No | Repair machine by jesson after replacing 924-609,580-083 | Repair machine by jesson after replacing 924-609,580-083 | Defect | HW | Sensor Door Stepper motor | Sensor Door Stepper motor
ABL90 FLEX | 更换新试剂包后修复 | 更换新试剂包后修复 | SWO-87726 | No | Repair after replacing a new kit | Repair after replacing a new kit | Assisted Replacement | SP | SP Installation | FT
ABL90 | Veriicação dos registros, reinstalação do pack, estabilização e passagem dos QCs. | B | D07282111C29C212 | No | Verification of the records, reinstallation of the pack, stabilization and passage of the QCs. |  | User error | SP Installation | SP Installation | SP Installation
ABL90 FLEX | 更换新试剂包后修复 | 更换新试剂包后修复 | SWO-87773 | No | Repair after replacing a new kit | Repair after replacing a new kit
ABL837 FLEX | Cleaned the analyzer replaced the Printer Head restarted instrument found printer was working satisfactorily. | Cleaned the analyzer replaced the Printer Head restarted instrument found printer was working satisfactorily. | SWO-88269 | No | Cleaned the analyzer replaced the Printer Head restarted instrument found printer was working satisfactorily. | Cleaned the analyzer replaced the Printer Head restarted instrument found printer was working satisfactorily. | Defect | HW | Printer module | Printer head
ABL825 FLEX |  | B | D072B4861C412F53 | Yes |  |  | Re classified RMED | SWU
ABL837 FLEX | found that the waste tubes had holes in them | replaced waste tubes to correct the issue | D07272D01C3EFC9D | No | found that the waste tubes had holes in them | replaced waste tubes to correct the issue | Defect | HW | Fluidic pathway | Pump tubing | Waste
ABL800 BASIC | Checked calibration results and sensitivity values its found ok. Checked with QC that also showing very low value for calcium than the actual range.. All other parameters are within range. Changed the Calcium membrane. Verified Qc Values. | B | SWO-88227 | No | Checked calibration results and sensitivity values its found ok. Checked with QC that also showing very low value for calcium than the actual range.. All other parameters are within range. Changed the Calcium membrane. Verified Qc Values. |  | Assisted Replacement | Electrode | Membrane | Ca
ABL827 FLEX | 910-237 Druckaufnehmer erneuert, Lampeneinheit überholt. Gastest und OXI Complete Test i.O. | 910-237 Druckaufnehmer erneuert, Lampeneinheit überholt. Gastest und OXI Complete Test i.O. | SWO-87583 | No | 910-237 Pressure transducer renewed, lamp unit overhauled. Gastest and OXI Complete Test i.O. | 910-237 Pressure transducer renewed, lamp unit overhauled. Gastest and OXI Complete Test i.O. | Defect | HW | Gas module | Pressure Transducer | 689
ABL825 FLEX | Alle Daten gesichert, Gerät vom Netzwerk getrennt, über Control Panel die 80 verbliebenen Druckaufträge gelöscht, Gerät wieder mit Netzwerk verbunden, mehrere Testausdrücke erstellt, kein Fehler mehr aufgetreten, Abteilung beobachtet und meldet sich bei weiteren Problemen | Alle Daten gesichert, Gerät vom Netzwerk getrennt, über Control Panel die 80 verbliebenen Druckaufträge gelöscht, Gerät wieder mit Netzwerk verbunden, mehrere Testausdrücke erstellt, kein Fehler mehr aufgetreten, Abteilung beobachtet und meldet sich bei weiteren Problemen | SWO-87369 | No | All data backed up, device disconnected from the network, 80 remaining print jobs deleted via Control Panel, device reconnected to network, multiple test expressions created, no more errors occurred, department monitors and reports further problems | All data backed up, device disconnected from the network, 80 remaining print jobs deleted via Control Panel, device reconnected to network, multiple test expressions created, no more errors occurred, department monitors and reports further problems | Defect | SW | Printer module | Printer driver
ABL825 FLEX | System gespült und diverse Schläuche getauscht Kal- und QC | System gespült und diverse Schläuche getauscht Kal- und QC | SWO-87433 | Yes | System flushed and various hoses replaced Cal- and QC | System flushed and various hoses replaced Cal- and QC | Re classified RMED | OOB
AQT90 FLEX | Adjusted position of metal bar used to open waste door on reagent pack, so that door is fully opened. | Adjusted position of metal bar used to open waste door on reagent pack, so that door is fully opened. | SWO-87351 | Yes | Adjusted position of metal bar used to open waste door on reagent pack, so that door is fully opened. | Adjusted position of metal bar used to open waste door on reagent pack, so that door is fully opened. | Re classified RMED | Not in PPM
ABL80 FLEX BASIC | 更换新的测试卡修复仪器 | 更换新的测试卡修复仪器 | SWO-88390 | No | Replace a new test card repair instrument | Replace a new test card repair instrument
ABL80 FLEX BASIC | 更换新的测试卡修复仪器 | 更换新的测试卡修复仪器 | SWO-88395 | No | Replace a new test card repair instrument | Replace a new test card repair instrument
ABL80 FLEX BASIC | 更换新测试卡修复仪器 | 更换新测试卡修复仪器 | SWO-88725 | No | Replace the new test card repair instrument | Replace the new test card repair instrument
ABL80 FLEX BASIC |  | B | SWO-87482 | No
ABL80 FLEX BASIC | 更换新的测试卡修复仪器 | 更换新的测试卡修复仪器 | SWO-87521 | No | Replace a new test card repair instrument | Replace a new test card repair instrument
ABL90 | Examined solution pack and sensor cassette. | Performed pump cal, and pressure test. Replaced solution pack, with customer stock. Verified operation of ABL90. | D073908D923EB5A7 | No | Examined solution pack and sensor cassette. | Performed pump cal, and pressure test. Replaced solution pack, with customer stock. Verified operation of ABL90. | Assisted Replacement | SP | SP | Pump cal error
ABL90 FLEX | Gerät überprüft und gereinigt, Lösungspack erneuert, Meßkassette ausgetauscht | Gerät überprüft und gereinigt, Lösungspack erneuert, Meßkassette ausgetauscht | SWO-87615 | No | Device checked and cleaned, solution pack renewed, measuring cassette replaced | Device checked and cleaned, solution pack renewed, measuring cassette replaced | Assisted Replacement | SC | SC | Glu
ABL90 |  | B | D079E5BDE00A0EF1 | No |  |  | Assisted Replacement | SP | SP | QC Level | Glu/Lac
ABL90 FLEX | Analyzer terug bij klant geplaatst na herstelling in workshop. | Analyzer terug bij klant geplaatst na herstelling in workshop. | SWO-87667 | No | Analyzer placed back at customer after repair in workshop. | Analyzer placed back at customer after repair in workshop. | Defect | HW | Temperature | Temperature
ABL90 | Service performance tests did not reveal any problems. | Replace hemolyzer unit. Perform all service tests. Install 3.1 MR8 software. | D073C30B925DBDCB | No | Service performance tests did not reveal any problems. | Replace hemolyzer unit. Perform all service tests. Install 3.1 MR8 software. | Defect | HW | Oxi module | Hemolyzer | 581
ABL90 | Check data logs. | Perform all service tests in service program. All passed. Load MR8 software upgrade. | D07A15FDE0127332 | No | Check data logs. | Perform all service tests in service program. All passed. Load MR8 software upgrade. | No fault found | No fault found | No fault found | No fault found
ABL90 FLEX | Einlassdichtung geliefert und erneuert, Startphase 2h. Auffällig ist der hohe BG LS wert, weitere Untersuchungen folgen. Bestellformular anfordern lassen und Bestellung ausgelöst. | B | SWO-88297 | No | Inlet seal supplied and renewed, start phase 2h. The high BG LS is worth the high BG LS, further investigations follow. request an order form and trigger an order. |  | Defect | HW | Inlet Gasket | Inlet gasket
ABL90 |  | B | D07B4172563C3709 | Yes |  |  | Re classified RMED | Installation
ABL90 | Comprobación del estado del analizador. A mi llegada se observa equipo con luces de standby encendidas. Se reinicializa el equipo sin respuesta, se procede a desmontar carcasa y comprobar tensiones F.A. sólo se encuentra tensión 5v standby. Se pide material. Equipo parado. | Se determina fallo en F.A | D073965C9225307C | No | Checking the status of the scanner. On arrival i ateam with standby lights on. Unanswered equipment is reinitialized, housing is disassembled and F.A. voltages are checked only 5v standby voltage. Material is requested. Team stopped. | F.A. failure determined | Defect | HW | DMS_PC | Display
ABL90 FLEX |  | B | SWO-87553 | No |  |  | Assisted Replacement | SC | SC | pCO2
AQT90 FLEX |  | B | SWO-87524 | No |  |  | Defect | HW | Process Module | Process Unit sensor | 5784
AQT90 FLEX |  | B | SWO-87555 | No |  |  | Defect | HW | Process Module | Optical Unit | 131
AQT90 FLEX | 重启修复仪器 by jesson | 重启修复仪器 by jesson | SWO-87534 | No | Restart repair instrument by jesson | Restart repair instrument by jesson | Defect | SW | Analyzer Control Module | CPU
ABL825 FLEX | 检查机器，发现打印机传动机构松动，重新紧固修复 By Jesson | 检查机器，发现打印机传动机构松动，重新紧固修复 By Jesson | SWO-87514 | No | Check the machine and find that the printer drive mechanism is loose and retighten repair By Jesson | Check the machine and find that the printer drive mechanism is loose and retighten repair By Jesson | Defect | HW | Printer module | Printer module
ABL825 FLEX | Customer had a leak and Cl- was failing cals. | Customer had stopped leak by the time FSE got to site. Replaced measuring chamber of BG module as it had a clot. Found bent Cl- electrode connector and straightened. Replaced CL- membrane and Ref membrane from customer stock. Replaced electrode connector of Chlorine electode with no good QC results for level 3. Replaced Chlorine electrode. Replaced Na, pO2 and pCO2 membranes from customer stock. Completed all service tests. Went over QC ranges with customer. | D0737B44925C3157 | No | Customer had a leak and Cl- was failing cals. | Customer had stopped leak by the time FSE got to site. Replaced measuring chamber of BG module as it had a clot. Found bent Cl- electrode connector and straightened. Replaced CL- membrane and Ref membrane from customer stock. Replaced electrode connector of Chlorine electode with no good QC results for level 3. Replaced Chlorine electrode. Replaced Na, pO2 and pCO2 membranes from customer stock. Completed all service tests. Went over QC ranges with customer. | Assisted Replacement | Electrode | Electrode | Cl
ABL825 FLEX |  | B | D074BC66083D0CD3 | No |  |  | Defect | HW | pH/BG Module | Pump tubing
ABL825 FLEX | Excessive pH cal upgrades due to drift during a pH BG cal. Checked activity log. pH cal upgrade ocurring 6-8 times day. | Replaced pH electrode. Customer will monitor | D073BC4B9222A700 | No | Excessive pH cal upgrades due to drift during a pH BG cal. Checked activity log. pH cal upgrade ocurring 6-8 times day. | Replaced pH electrode. Customer will monitor | Assisted Replacement | Electrode | Electrode | pH
ABL825 FLEX | Modul installiert, Einlaufverfahren durchgeführt. Verbogene Einlasssonde und Einlassdichtung erneuert. | Modul installiert, Einlaufverfahren durchgeführt. Verbogene Einlasssonde und Einlassdichtung erneuert. | SWO-87974 | No | module installed, inlet procedure performed. Bent inlet probe and inlet seal renewed. | module installed, inlet procedure performed. Bent inlet probe and inlet seal renewed. | Defect | HW | Inlet module | Probe bend | Flüssigkeit tritt aus Einlass -Fluidc Modul zurückgetauscht // Modul installiert, Einlaufverfahren durchgeführt. Verbogene Einlasssonde und Einlassdichtung erneuert.
ABL837 FLEX | found in the data logs where it keeps happening | replaced tubing and membranes to correct the issue | D073C6739204FCA0 | No | found in the data logs where it keeps happening | replaced tubing and membranes to correct the issue | Defect | HW | Inlet module | Tubings
ABL800 BASIC | OSよりクリーンインストール実施 | B | SWO-88686 | No | Clean installation from OS |  | Defect | HW | DMS_PC | DMS_PC
ABL815 FLEX | Customer thought Lactate for patients was too high. Will try another Lactate electrode and rerun linearity. | Installed new lactate electrode. Ran VK linearity kit. Linearity values were very close. The only difference was external level 3 qc went down from 10.5 to 9.7. Level 1 and Level 2 qc only deviated 0.1. Verified instrument operation. | D074F9C30853ADDB | No | Customer thought Lactate for patients was too high. Will try another Lactate electrode and rerun linearity. | Installed new lactate electrode. Ran VK linearity kit. Linearity values were very close. The only difference was external level 3 qc went down from 10.5 to 9.7. Level 1 and Level 2 qc only deviated 0.1. Verified instrument operation. | Assisted Replacement | Electrode | Electrode | Lac | QC High
ABL815 FLEX | Lamp unit mechanisch eingestellt, | Lamp unit mechanisch eingestellt, | SWO-87881 | No | Lamp unit mechanically adjusted, | Lamp unit mechanically adjusted, | Adjustment | HW | Oxi module | Lamp unit
ABL815 FLEX | O-Ring der pH-Elektrode, Schlauch zw. BG-Modul und Einlass und Einlassdichtung erneuert, Gerät kalibriert und mit QCs und div. Messungen überprüft. OK | O-Ring der pH-Elektrode, Schlauch zw. BG-Modul und Einlass und Einlassdichtung erneuert, Gerät kalibriert und mit QCs und div. Messungen überprüft. OK | SWO-87645 | No | O-ring of the pH electrode, hose between BG module and inlet and inlet seal renewed, device calibrated and checked with QCs and various measurements. Ok | O-ring of the pH electrode, hose between BG module and inlet and inlet seal renewed, device calibrated and checked with QCs and various measurements. Ok | Defect | HW | Inlet module | Tubings
AQT90 FLEX | Cleaned the machine from the inside, cleaned all the Storage wheel sensors using cotton swab, reset all the storage wheel connections, restarted the machine, machine is working fine. No sample is available to run on the machine. | Cleaned the machine from the inside, cleaned all the Storage wheel sensors using cotton swab, reset all the storage wheel connections, restarted the machine, machine is working fine. No sample is available to run on the machine. | SWO-88218 | No | Cleaned the machine from the inside, cleaned all the Storage wheel sensors using cotton swab, reset all the storage wheel connections, restarted the machine, machine is working fine. No sample is available to run on the machine. | Cleaned the machine from the inside, cleaned all the Storage wheel sensors using cotton swab, reset all the storage wheel connections, restarted the machine, machine is working fine. No sample is available to run on the machine. | Cleaned | HW | Storage Module | Storage wheel, Sensor | Rotatestoragewheel_24
AQT90 FLEX | Einstellungen an AQT überprüft - ok. Kontakt mit OSM Hotline aufgenommen (längere Wartezeit auf OSM Rückruf). Herr Werner (OSM) hat Query überprüft. Diese wurde von AQT gesendet und von OSM gesendet. Herr Werner hat Änderung an der Übertragung von OSM vorgenommen. Query funktioniert. Messung durchgeführt. Übertragung ok. | Einstellungen an AQT überprüft - ok. Kontakt mit OSM Hotline aufgenommen (längere Wartezeit auf OSM Rückruf). Herr Werner (OSM) hat Query überprüft. Diese wurde von AQT gesendet und von OSM gesendet. Herr Werner hat Änderung an der Übertragung von OSM vorgenommen. Query funktioniert. Messung durchgeführt. Übertragung ok. | SWO-87588 | No | Settings checked to AQT - ok. Contact osM hotline (longer waiting time for OSM callback). Mr Werner (OSM) checked Query. This was sent by AQT and sent by OSM. Mr Werner has made a change to the transfer of OSM. Query works. measurement. Transmission ok. | Settings checked to AQT - ok. Contact osM hotline (longer waiting time for OSM callback). Mr Werner (OSM) checked Query. This was sent by AQT and sent by OSM. Mr Werner has made a change to the transfer of OSM. Query works. measurement. Transmission ok. | Setup | SW | His/lis | His/Lis
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-87788 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL80 FLEX BASIC | 更换测试卡修复仪器 | 更换测试卡修复仪器 | SWO-88087 | No | Replace the test card repair instrument | Replace the test card repair instrument
ABL80 CO-OX | customer reported error after sample, they had replaced SC and Sc but could not lear error | Flushed system, ran LS test and service tests passed reinitilzed system | D074D930084F999B | No | customer reported error after sample, they had replaced SC and Sc but could not lear error | Flushed system, ran LS test and service tests passed reinitilzed system
ABL90 FLEX | Run liquid sensor test, passed. Run pump calibration, passed. Run clot detect remove, passed. Exit service analyser went through start up and calibrations, calibration sensitivity errors on pCo2, Glu and Lac. | Run liquid sensor test, passed. Run pump calibration, passed. Run clot detect remove, passed. Exit service analyser went through start up and calibrations, calibration sensitivity errors on pCo2, Glu and Lac. | SWO-87894 | Yes | Run liquid sensor test, passed. Run pump calibration, passed. Run clot detect remove, passed. Exit service analyser went through start up and calibrations, calibration sensitivity errors on pCo2, Glu and Lac. | Run liquid sensor test, passed. Run pump calibration, passed. Run clot detect remove, passed. Exit service analyser went through start up and calibrations, calibration sensitivity errors on pCo2, Glu and Lac. | Re classified RMED | New in CRM
ABL90 | Confirmed error by phone. | Asked customer to replaced the sensor cassette and request credit via the printed credit form. | D0750BFB084A822D | Yes | Confirmed error by phone. | Asked customer to replaced the sensor cassette and request credit via the printed credit form. | Re classified RMED | Phone support
ABL90 FLEX PLUS |  | B | D074DB1608106263 | No |  |  | Defect | HW | Auto Inlet module | Unknown | Not returned, scraped
ABL90 FLEX | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-90099 | No | Repair after replacing a new test card | Repair after replacing a new test card
AQT90 FLEX |  | B | SWO-87810 | No |  |  | Defect | HW | Power Supply | Power Supply
ABL800 FLEX | Reference electrode replaced. reference membrane replaced. K membrane replaced. Ca membrane replaced. Calibration not passing. Again reference membrane replaced. Calibration passed. QC failed for pH, Electrolyte and Cret. Again reference membrane and Cret A and B membrane replaced and calibration done. Calibration successful. QC passed. Analyser working fine. LIS HIS setting shown to IT department. | Reference electrode replaced. reference membrane replaced. K membrane replaced. Ca membrane replaced. Calibration not passing. Again reference membrane replaced. Calibration passed. QC failed for pH, Electrolyte and Cret. Again reference membrane and Cret A and B membrane replaced and calibration done. Calibration successful. QC passed. Analyser working fine. LIS HIS setting shown to IT department. | SWO-87864 | No | Reference electrode replaced. reference membrane replaced. K membrane replaced. Ca membrane replaced. Calibration not passing. Again reference membrane replaced. Calibration passed. QC failed for pH, Electrolyte and Cret. Again reference membrane and Cret A and B membrane replaced and calibration done. Calibration successful. QC passed. Analyser working fine. LIS HIS setting shown to IT department. | Reference electrode replaced. reference membrane replaced. K membrane replaced. Ca membrane replaced. Calibration not passing. Again reference membrane replaced. Calibration passed. QC failed for pH, Electrolyte and Cret. Again reference membrane and Cret A and B membrane replaced and calibration done. Calibration successful. QC passed. Analyser working fine. LIS HIS setting shown to IT department. | Assisted Replacement | Electrode | Membrane | Reference
ABL835 FLEX | 3 17 データ確認と謝罪 | B | SWO-87746 | No | 3 17 Data confirmation and apology |  | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
AQT90 FLEX | Neuer Tube holder eingebaut und alten entfernt, auch Röllchen. Kleine Metallachse des Röhrchens nicht gefunden. AQT geöffnet, Teile ausgebaut und überprüft wo Metallachse ist, nicht gefunden. Überprüfung Pipette, nicht krum, o. K. Systemüberprüfung und QC o. K. | Neuer Tube holder eingebaut und alten entfernt, auch Röllchen. Kleine Metallachse des Röhrchens nicht gefunden. AQT geöffnet, Teile ausgebaut und überprüft wo Metallachse ist, nicht gefunden. Überprüfung Pipette, nicht krum, o. K. Systemüberprüfung und QC o. K. | SWO-87970 | No | New tube holder built-in and old removed, even rolls. Small metal axis of the tube not found. AQT opened, parts removed and checked where metal axis is, not found. Check pipette, not krum, o.K. system check and QC o. K. | New tube holder built-in and old removed, even rolls. Small metal axis of the tube not found. AQT opened, parts removed and checked where metal axis is, not found. Check pipette, not krum, o.K. system check and QC o. K. | Defect | HW | Aspiration Module | Tube Holder
AQT90 FLEX | 持ち帰り調整 | B | SWO-88096 | No | Take-away adjustment |  | Revisit required | HW | Cover and Chassis | Door?
AQT90 FLEX |  | B | D0750A6808197DB2 | No |  |  | Adjustment | HW | Process Module | Process wheel | 5684
TCM-COMBIM | Insured that the sensor head was in the holder and the electrode head micro switch is good. | Located and replaced the Main Computer PCB and the CF card. | D074C62E08481BCE | No | Insured that the sensor head was in the holder and the electrode head micro switch is good. | Located and replaced the Main Computer PCB and the CF card. | Defect | HW | PCB, Main CPU | PCB, Main CPU
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-89931 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-88344 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL80 FLEX BASIC | 更换一张测试卡和一个试剂包修复仪器 | 更换一张测试卡和一个试剂包修复仪器 | SWO-88099 | No | Replace a test card and a reagent kit repair instrument | Replace a test card and a reagent kit repair instrument
ABL80 FLEX Basic version |  | B | D07616F47E39EBB9 | No
ABL90 | The customer had tried two sensor cassettes, and one solution pack, but this did not help. Error code 1361 was displayed on the intervention screen. Troubleshot to a bad Sensor Interface PCB. | Replaced the Sensor Interface PCB, but error code 1361 (Internal reference electrode error in sensor cassette) is still active. A new cassette is needed but the customer has no more. | D07646EB7E126295 | No | The customer had tried two sensor cassettes, and one solution pack, but this did not help. Error code 1361 was displayed on the intervention screen. Troubleshot to a bad Sensor Interface PCB. | Replaced the Sensor Interface PCB, but error code 1361 (Internal reference electrode error in sensor cassette) is still active. A new cassette is needed but the customer has no more. | Defect | HW | Sensor cassette PCB | Sensor cassette PCB
ABL90 FLEX | Herr Haller rief mich an, dass er den SP tauschen musste. Den ausgebauten, kontaminierten SP ( wurde in der Originalverpackung transportiert ) wollte er auf eine andere Station zum testen bringen. Dabei bemerkte er, dass der SP nicht dicht ist. Rote Flüssigkeit lief auf seine Arbeitskleidung. Wo es genau ausläuft konnte er nicht sagen. Der Pappkarton wurde nicht noch einmal geöffnet. Der Transport des SP erfolgte wie immer. Ich fragte ihn nach der Losnummer des SP . Es ist Lot UN-01. Fotos wurden keine gemacht. Der SP befand sich in dem Pappkarton im Müll. Ich habe den SP nicht aus dem Müll geholt um Fotos zu machen. Kunde ist bzgl. des SP sehr unzufrieden. | Herr Haller rief mich an, dass er den SP tauschen musste. Den ausgebauten, kontaminierten SP ( wurde in der Originalverpackung transportiert ) wollte er auf eine andere Station zum testen bringen. Dabei bemerkte er, dass der SP nicht dicht ist. Rote Flüssigkeit lief auf seine Arbeitskleidung. Wo es genau ausläuft konnte er nicht sagen. Der Pappkarton wurde nicht noch einmal geöffnet. Der Transport des SP erfolgte wie immer. Ich fragte ihn nach der Losnummer des SP . Es ist Lot UN-01. Fotos wurden keine gemacht. Der SP befand sich in dem Pappkarton im Müll. Ich habe den SP nicht aus dem Müll geholt um Fotos zu machen. Kunde ist bzgl. des SP sehr unzufrieden. | SWO-88164 | Yes | Mr. Haller called me that he had to swap the SP. He wanted to take the removed, contaminated SP (was transported in the original packaging) to another station for testing. He noticed that the SP is not dense. Red liquid ran on to his work clothes. He couldn't say exactly where it would end. The cardboard box was not reopened. The transport of the SP took place as usual. I asked him for the lot number of the SP . It's Lot UN-01. No photos were taken. The SP was in the cardboard box in the garbage. I didn't get the SP out of the garbage to take photos. Customer is very dissatisfied with the SP. | Mr. Haller called me that he had to swap the SP. He wanted to take the removed, contaminated SP (was transported in the original packaging) to another station for testing. He noticed that the SP is not dense. Red liquid ran on to his work clothes. He couldn't say exactly where it would end. The cardboard box was not reopened. The transport of the SP took place as usual. I asked him for the lot number of the SP . It's Lot UN-01. No photos were taken. The SP was in the cardboard box in the garbage. I didn't get the SP out of the garbage to take photos. Customer is very dissatisfied with the SP. | Re classified RMED | Phone support
ABL90 FLEX | Flushed sample pathway to clear blood clot. | Flushed sample pathway to clear blood clot. | SWO-88231 | Yes | Flushed sample pathway to clear blood clot. | Flushed sample pathway to clear blood clot. | Re classified RMED | RLTD new in CRM
ABL90 | Analyzer needs to be put on network and connected to Radiance. Also needs MR8 software. | Upgrade to MR8. Load MR8 setup from CVICU analyzer without users. Configure NetOp. Configure Radiance connection. Restart analyzer. Test remote support from Radiance server. Also pint PFT analyzer to Cerner transmission profile. Document Analyzer configuration in Avenue. | D075EB0B7E6161DA | Yes | Analyzer needs to be put on network and connected to Radiance. Also needs MR8 software. | Upgrade to MR8. Load MR8 setup from CVICU analyzer without users. Configure NetOp. Configure Radiance connection. Restart analyzer. Test remote support from Radiance server. Also pint PFT analyzer to Cerner transmission profile. Document Analyzer configuration in Avenue. | Re classified RMED | Setup
ABL90 FLEX PLUS |  | B | D079EE2FE03B5773 | Yes |  |  | Re classified RMED | Logistical upSWO
ABL90 FLEX PLUS |  | B | D0761A8B7E4180A0 | No |  |  | Defect | HW | Connector gasket | Connector Gasket
ABL90 FLEX PLUS | SE HACE KIT DE TUBOS SE HACE LIMPIEZA DEL SISTEMA, TEST DE SERVICIO Y SE COMPRUEBA FUNCIONBAMIENTO CON CAL Y QC | SE HACE KIT DE TUBOS SE HACE LIMPIEZA DEL SISTEMA, TEST DE SERVICIO Y SE COMPRUEBA FUNCIONBAMIENTO CON CAL Y QC | D07632237E390A81 | No | PIPE KIT IS DONE SYSTEM CLEANING, SERVICE TEST AND FUNCTIONING WITH CAL AND QC IS CHECKED | PIPE KIT IS DONE SYSTEM CLEANING, SERVICE TEST AND FUNCTIONING WITH CAL AND QC IS CHECKED | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
AQT90 FLEX |  | B | SWO-88055 | No |  |  | Defect | HW | Process Module | Process Unit sensor | 5784
ABL835 FLEX | Se sustituye la membrana de referencia y se realiza la revisión preventiva según procedimiento, ajuste de sensores de líquidos, chequeo de válvulas limpieza de válvulas V7 y V8 entrada de aire y salida de desperdicios, limpieza filtro de aire, limpieza y descontaminación interna, test controles de calidad, calibración general, funcionando correctamente. | Explicación de cambio de membrana de referencia. | D0763C017E474168 | Yes | The reference membrane is replaced and preventive review is performed according to procedure, adjustment of liquid sensors, valve check cleaning of valves V7 and V8 air inlet and waste outlet, air filter cleaning, cleaning and decontamination internal, test quality controls, general calibration, working properly. | Reference membrane change explanation. | Re classified RMED | Preventive Maintenance
ABL835 FLEX | データ確認、新品ｐH電極交換実施 | B | SWO-88892 | No | Data verification and replacement of new pH electrodes |  | Assisted Replacement | Electrode | Electrode | pH
ABL825 FLEX |  | B | D07606947E46ED31 | Yes |  |  | Re classified RMED | Logistical upSWO
ABL825 FLEX | Beide Messkammern überprüft. An der pH Elektrode den unteren Dichtring ersetzt. An der Glu. Den unteren Dichtring eingebaut. Warte Ventil im El Met. Ersetzt. | Hr. Schulz besteht auf einen Rückruf und Service. | SWO-88293 | No | Both measuring chambers are checked. Replaced the lower sealing ring on the pH electrode. At the Glu. The lower sealing ring is installed. Waiting valve in El Met. Replaced. | Mr. Schulz insists on a recall and service. | Assisted Replacement | Electrode | Electrode | pH | Sealing ring
ABL800 BASIC |  | B | SWO-88039 | No |  |  | Assisted Replacement | Electrode | Electrode | Lac | Calibration
ABL835 FLEX | PCO2電極及びメンブラン交換。PH BGモジュールの測定チャンバー交換。PH BG,WL Metモジュールの各バルブ整備。エアーライン洗浄、LS調整実施。 | PCO2電極及びメンブラン交換。PH BGモジュールの測定チャンバー交換。PH BG,WL Metモジュールの各バルブ整備。エアーライン洗浄、LS調整実施。 | SWO-88089 | No | PCO2 electrode and membrane replacement. Measurement chamber replacement of PH BG module. Valve maintenance of PH BG and WL Met modules. Air line cleaning and LS adjustment. | PCO2 electrode and membrane replacement. Measurement chamber replacement of PH BG module. Valve maintenance of PH BG and WL Met modules. Air line cleaning and LS adjustment. | Assisted Replacement | Electrode | Electrode | pCO2
ABL80 FLEX BASIC | 更换新测试卡后解决 | 更换新测试卡后解决 | SWO-90249 | No | Resolve after replacing a new test card | Resolve after replacing a new test card
AQT90 FLEX | Replaced bar code unit. Checked customer log on again.Now working first time. Analyser now working ok | new barcode reader ordered. arranged to visit site on Monday 21st March | SWO-88309 | Yes | Replaced bar code unit. Checked customer log on again. Now working first time. Analyser now working ok | new barcode reader ordered. arranged to visit site on Monday 21st March | Re classified RMED | Not in PPM
ABL80 FLEX BASIC | 更换新测试卡修复仪器 | 更换新测试卡修复仪器 | SWO-88319 | No | Replace the new test card repair instrument | Replace the new test card repair instrument
ABL80 FLEX BASIC |  | B | SWO-88311 | No
AQT90 FLEX |  | B | SWO-88310 | No |  |  | Defect | HW | Process Module | Motor | 5679
ABL800 BASIC | Cleaned Touch panel, perform touch calibration, working ok. | Cleaned Touch panel, perform touch calibration, working ok. | SWO-90181 | No | Cleaned Touch panel, perform touch calibration, working ok. | Cleaned Touch panel, perform touch calibration, working ok. | Revisit required | HW | DMS_PC | Touch Screen
ABL80 FLEX BASIC | 更换新的测试卡修复仪器 | 更换新的测试卡修复仪器 | SWO-88469 | No | Replace a new test card repair instrument | Replace a new test card repair instrument
ABL80 FLEX BASIC | 更换新测试卡修复仪器 | 更换新测试卡修复仪器 | SWO-88507 | No | Replace the new test card repair instrument | Replace the new test card repair instrument
ABL80 FLEX BASIC | 跟换试剂包修复仪器 | 更换测试卡修复仪器 | SWO-88461 | No | Repair instruments with replacement kits | Replace the test card repair instrument
ABL80 FLEX BASIC | 更换测试卡修复仪器 | 更换测试卡修复仪器 | SWO-88462 | No | Replace the test card repair instrument | Replace the test card repair instrument
ABL90 |  | B | D079D560E030624F | No |  |  | Defect | HW | Oxi module | Pinch Valve tube
ABL90 | Internal qc for level 1 reading low, outside control range. | Replaced solution pack, several times. Changed sensor cassette. Replaced connectors and tubings. Replaced sensor interface. | D079FB6FE037D399 | No | Internal qc for level 1 reading low, outside control range. | Replaced solution pack, several times. Changed sensor cassette. Replaced connectors and tubings. Replaced sensor interface. | Defect | HW | Basic Unit | Sensor Interface PCB | Wetability ?
ABL90 FLEX | Fehlersuche, Einlaßdichtung gew., vorhandene SC erneut eingelegt -> weiter Lac-Fehler, neue SC installiert, Initialisierung abgewartet, Kal. = Fehler, Gerät aus- u. eingeschaltet, Startphase abgewartet, Scanner i.O., Kal. + QC i.O. | Fehlersuche; Einlaßdichtung gew.; vorhandene SC erneut eingelegt -> weiter Lac-Fehler; neue SC installiert; Initialisierung abgewartet; Kal. = Fehler; Gerät aus- u. eingeschaltet; Startphase abgewartet; Scanner i.O.; Kal. + QC i.O. | SWO-88584 | No | Troubleshooting; inlet seal. existing SC re-inserted -> further Lac error; new SC installed; initialization awaited; Kal. = Error; Device off and on; start-up phase awaited; Scanner i.O.; Kal. + QC in the first place | Troubleshooting; inlet seal. existing SC re-inserted -> further Lac error; new SC installed; initialization awaited; Kal. = Error; Device off and on; start-up phase awaited; Scanner i.O.; Kal. + QC in the first place | Assisted Replacement | SC | SC | Lac
ABL90 | Deleted all archived files and error still present. | Replaced CF card. | D079BA1AE0253859 | No | Deleted all archived files and error still present. | Replaced CF card. | Defect | HW | CF card | CF card
ABL90 FLEX | Bei Eintreffen stand ABL in Fehlerbehebung. Nadel war nicht richtig in Einlassdichtung eingeführt. Mehrere Male gespült Einlassdichtung und Lösungspack erneuert. (Wurde von Techniker mitgebracht da beides nicht mehr vor Ort.) Material vor Ort wurde von Intensivpersonal bei Versuch Fehler zu beheben verbraucht. Nach Startphase war Analysator wieder Messbereit. Kalibrierungen und QCs für alle Level wurden während Startphase durchgeführt. ABL wieder Messbereit. KEINE EINLASSDICHTUNGEN UND KEIN LÖSUNGSPACK MEHR VOR ORT! Bitte bestellen. | Bei Eintreffen stand ABL in Fehlerbehebung. Nadel war nicht richtig in Einlassdichtung eingeführt. Mehrere Male gespült Einlassdichtung und Lösungspack erneuert. (Wurde von Techniker mitgebracht da beides nicht mehr vor Ort.) Material vor Ort wurde von Intensivpersonal bei Versuch Fehler zu beheben verbraucht. Nach Startphase war Analysator wieder Messbereit. Kalibrierungen und QCs für alle Level wurden während Startphase durchgeführt. ABL wieder Messbereit. KEINE EINLASSDICHTUNGEN UND KEIN LÖSUNGSPACK MEHR VOR ORT! Bitte bestellen. | SWO-88640 | No | Upon arrival, ABL was in troubleshooting. Needle was not properly inserted into inlet seal. Several times rinsed inlet seal and solution pack renewed. (Was brought by technicians since both are no longer on site.) Material on site was consumed by intensive care personnel trying to fix errors. After the start phase, the analyzer was ready for measurement again. Calibrations and QCs for all levels were performed during the start phase. ABL ready to measure again. NO INTRODUCTIONS AND NO SOLUTION PACK MORE BEFORE PLACE! Please order. | Upon arrival, ABL was in troubleshooting. Needle was not properly inserted into inlet seal. Several times rinsed inlet seal and solution pack renewed. (Was brought by technicians since both are no longer on site.) Material on site was consumed by intensive care personnel trying to fix errors. After the start phase, the analyzer was ready for measurement again. Calibrations and QCs for all levels were performed during the start phase. ABL ready to measure again. NO INTRODUCTIONS AND NO SOLUTION PACK MORE BEFORE PLACE! Please order. | User error | HW | Inlet module | Not closed
ABL90 |  | B | D082B83A1A0AEE73 | No |  |  | Defect | HW | Inlet Gasket | Inlet gasket
ABL90 FLEX | Készülék átmosása, Eredeti tűtartó tisztítása, Kalibrációk, QC mérés | B | SWO-88582 | No | Machine wash, original needle holder cleaning, calibration, QC measurement |  | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
AQT90 FLEX |  | B | SWO-88333 | No |  |  | Defect | HW | Storage Module | PCB, Cup dryer | 168
ABL825 FLEX | Festplatte gewechselt. Mit Herrn Rahner überprüft ob ABL in Radiance richtig erscheint Mehrere Kalibrierungen durchgeführt. THb Kal durchgeführt. Fcuv: 0,9724 Kalverifizierung durchgeführt --> i.O. QCs für alle Level durchgeführt. ABL wieder Messbereit. | Festplatte gewechselt. Mit Herrn Rahner überprüft ob ABL in Radiance richtig erscheint Mehrere Kalibrierungen durchgeführt. THb Kal durchgeführt. Fcuv: 0,9724 Kalverifizierung durchgeführt --> i.O. QCs für alle Level durchgeführt. ABL wieder Messbereit. | SWO-88639 | No | Hard drive changed. With Mr. Rahner checking whether ABL appears correct in Radiance Several calibrations performed. THb Kal. Fcuv: 0.9724 Calverification performed --> i.O. QCs performed for all levels. ABL ready to measure again. | Hard drive changed. With Mr. Rahner checking whether ABL appears correct in Radiance Several calibrations performed. THb Kal. Fcuv: 0.9724 Calverification performed --> i.O. QCs performed for all levels. ABL ready to measure again. | Defect | HW | DMS_PC | HDD
ABL800 BASIC | Fehlersuche, Referenzelektrode wurde bereits membraniert, Messkammern pH und Referenz nochmal gereinigt, Elektroden- und Meskappenkontakte mit Glasfaserpinsel gereinigt, Proteinentfernung, Pumpenkal. und LS Justage, Kal., QC und Testmessung wieder ok System-und Funktionskontrolle i.O. | Fehlersuche, Referenzelektrode wurde bereits membraniert, Messkammern pH und Referenz nochmal gereinigt, Elektroden- und Meskappenkontakte mit Glasfaserpinsel gereinigt, Proteinentfernung, Pumpenkal. und LS Justage, Kal., QC und Testmessung wieder ok System-und Funktionskontrolle i.O. | SWO-88479 | No | Troubleshooting, reference electrode has already been membranized, measuring chambers pH and reference have been cleaned again, electrode and measuring cap contacts cleaned with glass fiber brush, protein removal, pump scale. and LS Adjustment, Cal., QC and test measurement again ok system and function control i.O. | Troubleshooting, reference electrode has already been membranized, measuring chambers pH and reference have been cleaned again, electrode and measuring cap contacts cleaned with glass fiber brush, protein removal, pump scale. and LS Adjustment, Cal., QC and test measurement again ok system and function control i.O. | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
ABL827 FLEX | Ran complete oxi test in service-Failed lamp test Tried lamp alignment-No change | Replaced oxi control PCB Alignment and tHB calibration in service | D079DB95E05CF821 | No | Ran complete oxi test in service-Failed lamp test Tried lamp alignment-No change | Replaced Oxygen Control PCB Alignment and tHB calibration in service | Defect | HW | Oxi module | PCB
ABL825 FLEX | Confirmed problem. | Replaced the Cl membrane and the pH BG and El Met module pump tubing. | D07A0B61E05D5EAD | No | Confirmed problem. | Replaced the Cl membrane and the pH BG and El Met module pump tubing. | Defect | HW | pH/BG Module | Pump tubing
ABL810 FLEX | Fließwege überprüft und Pumpenschlauch BG Modul ersetzt. pCO 2 Membraniert und ein Kalibriert. | B | SWO-88610 | Yes | Flow paths checked and pump hose BG module replaced. pCO 2 membranes and a calibrated. |  | Re classified RMED | Workshop setup
ABL810 FLEX | Neues Oxymetriemodul eingebaut. Service THb Kal. Durchgeführt. | Neues Oxymetriemodul eingebaut. Service THb Kal. Durchgeführt. | SWO-88611 | Yes | New oxymetry module built-in. Service THb Kal. Performed. | New oxymetry module built-in. Service THb Kal. Performed. | Re classified RMED | Workshop setup
ABL825 FLEX | Halogenlampe justiert StHb Kal-, Kal-, tHb Kal- und QC Oxi Completetest: ok Halogenintensity:3299 | Halogenlampe justiert StHb Kal-, Kal-, tHb Kal- und QC Oxi Completetest: ok Halogenintensity:3299 | SWO-88474 | No | Halogen lamp adjusted StHb Cal-, Kal-, tHb Kal- and QC Oxi Completetest: ok Halogenintensity:3299 | Halogen lamp adjusted StHb Cal-, Kal-, tHb Kal- and QC Oxi Completetest: ok Halogenintensity:3299 | Adjustment | HW | Oxi module | Lamp unit
ABL825 FLEX | Schrauben festgezogen und FlexQ justiert | Schrauben festgezogen und FlexQ justiert | SWO-88473 | No | Screws tightened and FlexQ adjusted | Screws tightened and FlexQ adjusted | Defect | HW | FlexQ Module | Screws adjusted
AQT90 FLEX |  | B | D07A0B62E062BC73 | No |  |  | Cleaned | HW | Aspiration Module | Barcode reader, Inlet
AQT90 FLEX | ASTM.ini vom Alten AQT 838R0128N003 nach neuen AQT 838 R0402N005 Kopiert. Commit Befehl eingespielt. Konfiguration überprüft. Verbindung hergestellt. Eine Probe gemessen um zu testen ob die Daten im System ankommen. | ASTM.ini vom Alten AQT 838R0128N003 nach neuen AQT 838 R0402N005 Kopiert. Commit Befehl eingespielt. Konfiguration überprüft. Verbindung hergestellt. Eine Probe gemessen um zu testen ob die Daten im System ankommen. | SWO-88548 | No | ASTM.ini from the old AQT 838R0128N003 copied to new AQT 838 R0402N005. Commit command. Configuration checked. Connection established. A sample is measured to test whether the data arrives in the system. | ASTM.ini from the old AQT 838R0128N003 copied to new AQT 838 R0402N005. Commit command. Configuration checked. Connection established. A sample is measured to test whether the data arrives in the system. | Setup | SW | His/lis | His/Lis
ABL80 FLEX BASIC |  | B | SWO-88688 | No
ABL80 FLEX CO-OX | Hemolyser erneuert Testmessungen ,Kalibration und QC durchgeführt | Hemolyser erneuert Testmessungen ,Kalibration und QC durchgeführt | SWO-88749 | No | Hemolyser renews test measurements ,calibration and QC performed | Hemolyser renews test measurements ,calibration and QC performed
ABL80 CO-OX | Customer tried 2 SC's, same error. They put first SC on other analyzer and it was fine. SC connector board is suspect. | Replaced Sensor Connector Board. Replaced parts from PM kit. | D07B135D563E92B4 | No | Customer tried 2 SC's, same error. They put first SC on other analyzer and it was fine. SC connector board is suspect. | Replaced Sensor Connector Board. Replaced parts from PM kit.
ABL80 FLEX CO-OX | Gerät geöffnet und auf Undichtigkeiten überprüft und OK. Am ABL810 im Labor eine BGA gemessen und diese Probe im Kreißsaal am ABL80 CO-OX gemessen. Die Ergebnisse waren vergleichbar und OK. Probe weiterhin mehrfach gemessen. Die Ergebnisse waren stabil und OK. | Gerät geöffnet und auf Undichtigkeiten überprüft und OK. Am ABL810 im Labor eine BGA gemessen und diese Probe im Kreißsaal am ABL80 CO-OX gemessen. Die Ergebnisse waren vergleichbar und OK. Probe weiterhin mehrfach gemessen. Die Ergebnisse waren stabil und OK. | SWO-89703 | No | device opened and checked for leaks and OK. Measured on the ABL810 in the laboratory a BGA and measured this sample in the delivery room at the ABL80 CO-OX. The results were comparable and OK. sample continues to be measured several times. The results were stable and OK. | device opened and checked for leaks and OK. Measured on the ABL810 in the laboratory a BGA and measured this sample in the delivery room at the ABL80 CO-OX. The results were comparable and OK. sample continues to be measured several times. The results were stable and OK.
ABL90 FLEX | Dem Kunden 2 Ersatz Sensoren kostenlos gegeben. 946-009 Lot R0443 | Dem Kunden 2 Ersatz Sensoren kostenlos gegeben. 946-009 Lot R0443 | SWO-88797 | Yes | The customer is given 2 replacement sensors free of charge. 946-009 Lot R0443 | The customer is given 2 replacement sensors free of charge. 946-009 Lot R0443 | Re classified RMED | PER reporting
ABL90 | Upon arrival, found PO2 on internal QC running low out of range. Upon further investigation, found Date on unit set correct day, but 3 years prior 2013. | Changed and put in correct Date. Upon doing this, all Consumables became Expired. Replaced Solution Pack and Sensor cassette. | D07B08775606252B | No | Upon arrival, found PO2 on internal QC running low out of range. Upon further investigation, found Date on unit set correct day, but 3 years prior 2013. | Changed and put in correct Date. Upon doing this, all Consumables became Expired. Replaced Solution Pack and Sensor cassette. | Assisted Replacement | SC | SC | pO2
ABL90 |  | B | D07B42815638370A | Yes |  |  | Re classified RMED | Training of customer
ABL825 FLEX | Checked upper fluid valves ok, found upper electrode pump tubes were stretched out slowing the sample aspiration on blood. | Replaced electrode tube with client stock | D07B43BC564D851B | No | Checked upper fluid valves ok, found upper electrode pump tubes were stretched out slowing the sample aspiration on blood. | Replaced electrode tube with client stock | Defect | HW | pH/BG Module | Pump tubing
ABL825 FLEX | C HECKED UPPER FLUID VALVES OK , FOUND UPPER ELECTRODE PUMP TUBES WERE STRETCHED OUT SLOWING THE SAMPLE ASPIRATION ON BLOOD . | R EPLACED ELECTRODE TUBE WITH CLIENT STOCK | D07B521D5637851E | No | C HECKED UPPER FLUID VALVES OK , FOUND UPPER ELECTRODE PUMP TUBES WERE STRETCHED OUT SLOWING THE SAMPLE ASPIRATION ON BLOOD . | R EPLACED ELECTRODE TUBE WITH CLIENT STOCK | Defect | HW | pH/BG Module | Pump tubing
ABL825 FLEX | Found waste tube not completely seated and the sample probe was bent. Also upper electrode pump tubes were stretched out slowing the sample asperation on blood. | Adjusted waste tube, replaced electrode tube with client stock, and straightened probe. | D07B2D6356088515 | No | Found waste tube not completely seated and the sample probe was bent. Also upper electrode pump tubes were stretched out slowing the sample asperation on blood. | Adjusted waste tube, replaced electrode tube with client stock, and straightened probe. | Defect | HW | pH/BG Module | Pump tubing
AQT90 FLEX |  | B | D07B3F0C56288711 | No |  |  | Defect | HW | Analyzer Control Module | CPU
AQT90 FLEX |  | B | D07B88F7562CDAFC | No |  |  | Not Cat | Not cat. | Not cat. | Not cat.
AQT90 FLEX | Überprüfungs- und Reinigungsprozess laut Anweisung nach Falsch positivem TNT Ergebnis - Funktionsüberprüfung etc. ohne Beanstandung ! | Überprüfungs- und Reinigungsprozess laut Anweisung nach Falsch positivem TNT Ergebnis - Funktionsüberprüfung etc. ohne Beanstandung ! | SWO-88921 | No | Review and cleaning process according to instruction after False positive TNT result - functional check etc. without complaint ! | Review and cleaning process according to instruction after False positive TNT result - functional check etc. without complaint ! | Performance | Discrepant Result | Discrepant Result | Discrepant Result
AQT90 FLEX |  | B | D07B23A5560F7634 | No |  |  | Defect | HW | Process Module | Optical Unit
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-90246 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL80 FLEX CO-OX | Die Daten der Netzwerkanbindung korrigiert, QC und Pat Messung durchgeführt | Die Daten der Netzwerkanbindung korrigiert, QC und Pat Messung durchgeführt | SWO-89771 | No | Corrected the data of the network connection, QC and Pat measurement performed | Corrected the data of the network connection, QC and Pat measurement performed
ABL90 FLEX | Automatisches Datensenden für die Poctopusverbindung eingeschaltet, das Benutzermanagment eingeschaltet, Pat und QC Messung durchgeführt und mit Herrn Michael abgeglichen ob die Daten im Poctopus ankommen | Automatisches Datensenden für die Poctopusverbindung eingeschaltet, das Benutzermanagment eingeschaltet, Pat und QC Messung durchgeführt und mit Herrn Michael abgeglichen ob die Daten im Poctopus ankommen | SWO-89772 | No | Automatic data sending switched on for the Poctopus connection, user management switched on, Pat and QC measurement performed and matched with Mr. Michael whether the data arrives in the Poctopus | Automatic data sending switched on for the Poctopus connection, user management switched on, Pat and QC measurement performed and matched with Mr. Michael whether the data arrives in the Poctopus | Setup | SW | His/lis | His/Lis
ABL90 | Found analyzer to be DHCP Found analyzer set to 24 hour Lease Found WAP's to possible be configured to not roam or allow roaming of Mobile Devices | Obtained logs and investigated. Will pull email together tomorrow and send to team on update. | D07CCF9ECC564FF6 | No | Found analyzer to be DHCP Found analyzer set to 24 hour Lease Found WAP's to possible be configured to not roam or allow roaming of Mobile Devices | Obtained logs and investigated. Will pull email together tomorrow and send to team on update. | Unknown | SW | Wireless | WAP
ABL90 | Customer will be performing an ECMO procedure while getting the error 1228, and Ca, Na, K not reporting values. | Techline indicates this is normal operation of unit, but unable to explain why values are missing. | D07D7030425B3789 | No | Customer will be performing an ECMO procedure while getting the error 1228, and Ca, Na, K not reporting values. | Techline indicates this is normal operation of unit, but unable to explain why values are missing. | Assisted Replacement | SC | SC | K, Ca, Na
ABL90 FLEX PLUS |  | B | D07C89A9CC1574A1 | Yes |  |  | Re classified RMED | Follwo up visit to 23/3-216
TCM4 |  | B | SWO-88915 | No
ABL90 FLEX | 更换新的测试卡修复仪器 | 更换新的测试卡修复仪器 | SWO-89829 | No | Replace a new test card repair instrument | Replace a new test card repair instrument
ABL90 FLEX | 更换测试卡修复仪器 | 更换测试卡修复仪器 | SWO-89934 | No | Replace the test card repair instrument | Replace the test card repair instrument
ABL90 | Verificação do SPack não totalmente activado, activação do mesmo e explicação ao enf. Nuno, montagem do Spack, establização e passagem do QC. | B | D07CA18BCC0D018A | No | SPack Check not fully activated, same activation and enf explanation. Nuno, installation of Spack, Establization and QC passage. |  | User error | SP Activation | SP Activation | SP Activation
ABL827 FLEX | Cleaned the lactate electrode, chamber & electrode connector. The electrode connector did have deposits on it which would have resulted in a poor connection to the electrode. Replaced the EL MET pump roller. Verification Visit performed. Software upgraded to v6.16 | Cleaned the lactate electrode, chamber & electrode connector. The electrode connector did have deposits on it which would have resulted in a poor connection to the electrode. Replaced the EL MET pump roller. Verification Visit performed. Software upgraded to v6.16 | SWO-89815 | No | Cleaned the lactate electrode, chamber & electrode connector. The electrode connector did have deposits on it which would have resulted in a poor connection to the electrode. Replaced the EL MET pump roller. Verification Visit performed. Software upgraded to v6.16 | Cleaned the lactate electrode, chamber & electrode connector. The electrode connector did have deposits on it which would have resulted in a poor connection to the electrode. Replaced the EL MET pump roller. Verification Visit performed. Software upgraded to v6.16 | Defect | HW | El/Met Module | Pump carroussel
ABL827 FLEX | Software upgraded to v6.16 (the fluid transport has been optimised in this version) Replaced the EL MET module pump roller. Performed the Verification Visit that was due in April. Cleaned the contacts in the Electrode Connector & on the electrode. Replaced the electrode connector because the QC measurements were unstable. | Software upgraded to v6.16 (the fluid transport has been optimised in this version) Replaced the EL MET module pump roller. Performed the Verification Visit that was due in April. Cleaned the contacts in the Electrode Connector & on the electrode. Replaced the electrode connector because the QC measurements were unstable. | SWO-89813 | No | Software upgraded to v6.16 (the fluid transport has been optimised in this version) Replaced the EL MET module pump roller. Performed the Verification Visit that was due in April. Cleaned the contacts in the Electrode Connector & on the electrode. Replaced the electrode connector because the QC measurements were unstable. | Software upgraded to v6.16 (the fluid transport has been optimised in this version) Replaced the EL MET module pump roller. Performed the Verification Visit that was due in April. Cleaned the contacts in the Electrode Connector & on the electrode. Replaced the electrode connector because the QC measurements were unstable. | Defect | HW | El/Met Module | Electrode Connector | Lac
ABL825 FLEX | Checked data logs, pH low on sensitivity at 91%. Calibrations will not pass. No other errors present. | Flushed air line through V7. Replaced pH electrode. Electrode covered under warranty, not billable. Restart calibrations. | D07C8A0BCC3556AB | No | Checked data logs, pH low on sensitivity at 91%. Calibrations will not pass. No other errors present. | Flushed air line through V7. Replaced pH electrode. Electrode covered under warranty, not billable. Restart calibrations. | Assisted Replacement | Electrode | Electrode | pH
ABL825 FLEX |  | B | D07C829BCC564EDD | Yes |  |  | Re classified RMED
ABL835 FLEX | confirmed problem | Added ETXETBClassic=False line to the roma.ini file. | D07C835BCC4FE751 | No | confirmed problem | Added ETXETBClassic=False line to the roma.ini file. | Setup | SW | Roma.ini | Roma.ini
ABL815 FLEX | Reinigungsarbeiten, Fehlersuche, Datensicherung, Neustart, Softwareupdate auf v6.16, Datenrücksicherung Siehe auch SWO zu TU135 | Kontrolle Bildschirmkabelanschluss, Prüfung LW C:, Kontrolle und Korrektur Netzwerkkartenkonfiguration, Probelauf -> Beobachtet durch Kunden, Kunde meldet sich | SWO-89681 | No | Cleaning, troubleshooting, data backup, restart, software update to v6.16, data backup See also SWO to TU135 | Control screen cable connection, check LW C:, control and correction of NETWORK card configuration, trial run -> Observed by customer, customer reports | Setup | SW | His/lis | His/Lis
ABL815 FLEX | Modul-Ventile ausgebaut und gängig gemacht, Probelauf | Modul-Ventile ausgebaut und gängig gemacht, Probelauf | SWO-89733 | No | Module valves removed and made common, test run | Module valves removed and made common, test run | Defect | HW | El/Met Module | Valve
AQT90 FLEX | Reinstalled the Application software. | B | SWO-90182 | No | Reinstalled the Application software. |  | Defect | SW | Analyzer Control Module | Re installed SW
AQT90 FLEX | Replaced needle. During initialisation analysers failing on ftc_106. Checked all tubing related to the fluid pathway. Removed wash unit and found fluid has been leaking. New wash unit required. Parts ordered. This call has been passed to Colin due to myself being on Holiday. Colin aware. John to inform Colin when parts are in stock. | Cleaned needle assembly, calibrated height and in to wash block. Repositioned the wash block to inlet PCB cable. Ran system clean and 5 LQC's. | SWO-89793 | Yes | Replaced needle. During initialisation analysers failing on ftc_106. Checked all tubing related to the fluid pathway. Removed wash unit and found fluid has been leaking. New wash unit required. Parts ordered. This call has been passed to Colin due to myself being on Holiday. Colin aware. John to inform Colin when parts are in stock. | Cleaned needle assembly, calibrated height and in to wash block. Repositioned the wash block to inlet PCB cable. Ran system clean and 5 LQC's. | Re classified RMED | Not in PPM
ABL800 BASIC | Cleaned touch panel by tissue paper, perform touch calibration, problem remains intermediately. | B | SWO-90267 | No | Cleaned touch panel by tissue paper, perform touch calibration, problem remains intermediately. |  | Revisit required | HW | DMS_PC | Touch Screen
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-90078 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-90079 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-89846 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL90 FLEX | Gerät überprüft, Netzteil und Sicherungen durchgemessen, Netzteil und Akku erneuert | Gerät überprüft, Netzteil und Sicherungen durchgemessen, Netzteil und Akku erneuert | SWO-90076 | No | Device checked, power supply and fuses measured, power supply and battery renewed | Device checked, power supply and fuses measured, power supply and battery renewed | Defect | HW | Battery module | Battery
ABL90 | Rebooted analyzer still had the same problem. Found pO2 reading high on S9030 and S9050. | Removed the CPU unit and reseated all connections. Flushed analyzer and replaced the sensor cassette. | D07DA6B642040ABC | No | Rebooted analyzer still had the same problem. Found pO2 reading high on S9030 and S9050. | Removed the CPU unit and reseated all connections. Flushed analyzer and replaced the sensor cassette. | Defect | HW | CPU | CPU
ABL90 FLEX | None needed | None needed | SWO-89798 | No | None needed | None needed | Unknown | Unknown | Unknown | Unknown
ABL820 FLEX | siehe SWO89989 | siehe SWO89989 | SWO-89992 | No | see SWO89989 | see SWO89989 | Assisted Replacement | Electrode | Electrode | pH | Sealing ring
ABL827 FLEX | Gerät geprüft , Einlass-Sonde erneuert , CREA neu membraniert , Kal und QCs durchgeführt | B | SWO-90445 | No | Device tested, inlet probe renewed, CREA re-membraneed, Kal and QCs performed |  | Defect | HW | Inlet module | Probe
ABL827 FLEX | Unteren Dichtungsring der Crea A-Elektrode wieder richtig aufgesetzt, Ls-Justage und Pumpenkal, System und Funktion i.O. | Unteren Dichtungsring der Crea A-Elektrode wieder richtig aufgesetzt, Ls-Justage und Pumpenkal, System und Funktion i.O. | SWO-89981 | No | Lower sealing ring of the Crea A electrode reattached correctly, Ls-adjustment and pump scale, system and function in the entire | Lower sealing ring of the Crea A electrode reattached correctly, Ls-adjustment and pump scale, system and function in the entire | User error | Electrode | Electrode | Crea | Sealing ring
ABL825 FLEX | Oximetrie Modul justiert, Service, tHB Kal. Kal und QC und OK. | Oximetrie Modul justiert, Service, tHB Kal. Kal und QC und OK. | SWO-90063 | No | Oximetry module adjusted, service, tHB Kal. Kal and QC and OK. | Oximetry module adjusted, service, tHB Kal. Kal and QC and OK. | Adjustment | HW | Oxi module | Lamp unit
AQT90 FLEX |  | B | D07DD687425DBA2D | No |  |  | Defect | HW | Storage Module | Barcode reader, Storage wheel
AQT90 FLEX | Performed the system cleaning, up graded the software to 8.8.16V, found sample door was not working need to replace doors and check. | Performed the system cleaning, up graded the software to 8.8.16V, found sample door was not working need to replace doors and check. | SWO-90175 | No | Performed the system cleaning, up graded the software to 8.8.16V, found sample door was not working need to replace doors and check. | Performed the system cleaning, up graded the software to 8.8.16V, found sample door was not working need to replace doors and check. | Defect | HW | Cover and Chassis | Inlet Door
AQT90 FLEX | Cleaned the Barcode Scanner, checked the sensor and inlet door movement found intermediate working of our Inlet door in AQT 90, under observation. | B | SWO-90548 | No | Cleaned the Barcode Scanner, checked the sensor and inlet door movement found intermediate working of our Inlet door in AQT 90, under observation. |  | Cleaned | HW | Aspiration Module | Barcode reader, Inlet
ABL90 FLEX PLUS | Vervolgbezoek brengen om de hemolyzer te vervangen. | Vervolgbezoek brengen om de hemolyzer te vervangen. | SWO-90216 | No | Follow-up visit to replace the Hemolyzer. | Follow-up visit to replace the Hemolyzer. | Unknown | Revisit required | Revisit required | Revisit required
AQT90 FLEX | 更换了636-561一个和903-329一个后解决(by zhenjie) | 更换了636-561一个和903-329一个后解决(by zhenjie) | SWO-89838 | No | Replaced 636-561 and 903-329 one after resolution (by zhenjie) | Replaced 636-561 and 903-329 one after resolution (by zhenjie) | Defect | HW | Aspiration Module | Needle Wash Unit | 3414
ABL835 FLEX | Flow is not proper. Calibrations aborting on arrival. Intermittent 0579 errors on calibrations. | Replaced cal 1 bottle. Cleaned air vent port. Replaced missing seal ring on glucose electrode. Performed complete oxi test. Lamp unit failed neon position test. Replacing Oxi control no help. Resoldered splice on neon lamp wire. Tried to reposition. | D07ED906B819E5F0 | No | Flow is not proper. Calibrations aborting on arrival. Intermittent 0579 errors on calibrations. | Replaced cal 1 bottle. Cleaned air vent port. Replaced missing seal ring on glucose electrode. Performed complete oxi test. Lamp unit failed neon position test. Replacing Oxi control no help. Resoldered splice on neon lamp wire. Tried to reposition. | Assisted Replacement | Electrode | Electrode | Glu | Sealing ring
ABL80 FLEX BASIC | 更换了试剂包卡座903-363后修复(by zhenjie) | 更换了试剂包卡座903-363后修复(by zhenjie) | SWO-90109 | No | Repair after replacement of reagent pack holder 903-363 (by zhenjie) | Repair after replacement of reagent pack holder 903-363 (by zhenjie)
ABL80 FLEX BASIC | 更换新测试卡安装后解决 | 更换新测试卡安装后解决 | SWO-90183 | No | Resolve after replacing new test card installation | Resolve after replacing new test card installation
ABL80 CO-OX | Found that the ribbon cable connector was crimping waste line. | Repositioned waste line. | D082D96D1A181E39 | No | Found that the ribbon cable connector was crimping waste line. | Repositioned waste line.
ABL90 FLEX PLUS | インレットガスケット及び流路洗浄。インレットプローブがインレットガスケットより外れていた為、取り付け。マニホールドアジャスト、プレッシャーテスト、LS調整、ポンプキャリブレーション実施。 | インレットガスケット及び流路洗浄。インレットプローブがインレットガスケットより外れていた為、取り付け。マニホールドアジャスト、プレッシャーテスト、LS調整、ポンプキャリブレーション実施。 | SWO-90337 | No | Inlet gasket and flow path cleaning. The inlet probe was out of the inlet gasket, so it was installed. Manihole door just, pressure test, LS adjustment, pump calibration. | Inlet gasket and flow path cleaning. The inlet probe was out of the inlet gasket, so it was installed. Manihole door just, pressure test, LS adjustment, pump calibration. | User error | HW | Auto Inlet module | Inlet Gasket
AQT90 FLEX | 更换光学单元传感器636-050一个后修复 | 更换光学单元传感器636-050一个后修复 | SWO-90284 | No | Replacement of optical unit sensor 636-050 one post-repair | Replacement of optical unit sensor 636-050 one post-repair | Defect | HW | Process Module | Process Unit sensor | 5784
ABL825 FLEX | Customer reseated carousel | Found broken fan on bottom of assembly blocking carousel from going down. Replaced autocheck fan | D08283C51A0A2A6F | No | Customer reseated carousel | Found broken fan on bottom of assembly blocking carousel from going down. Replaced autocheck fan | Defect | HW | AutoCheck Module | FAN
ABL825 FLEX | Inspected all tubings, inlet gasket, inlet probe and valves and all look good. | Installed another PH electrode and error went away, installed original and error gone. I cleaned out the ph bg measuring chamber and cleaned the bottom of the ph electrode | D0828E751A4420A9 | No | Inspected all tubings, inlet gasket, inlet probe and valves and all look good. | Installed another PH electrode and error went away, installed original and error gone. I cleaned out the ph bg measuring chamber and cleaned the bottom of the ph electrode | Assisted Replacement | Electrode | Electrode | pH
ABL80 FLEX BASIC | 部品交換作業 | 部品交換作業 | SWO-97277 | No | Part replacement work | Part replacement work
ABL80 FLEX BASIC | re installed the SC 80 now its working fine | re installed the SC 80 now its working fine | SWO-91072 | No | re installed the SC 80 now its working fine | re installed the SC 80 now its working fine
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-92571 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL1 | Cleaned Connector, try to reinstalled SC, Sensor connector board need to replace with new one. | B | SWO-91322 | No | Cleaned Connector, try to reinstalled SC, Sensor connector board need to replace with new one.
ABL90 FLEX | Gerät steht seit 25.03.16 in Intervention erforderlich-> Test wiederholt, Startphase abgewartete, div. Anwenderwartungsmassnahmen fällig: Anschlußdichtung ersetzt, thb Kal durchgeführt, nach Startphase pCO2, Glu,Lac Kal fehlerhaft ( Gerät stand eingeschaltet ohne Spülung der SC seit 27.03. lt. Aktivitätendatenbank), Kal mehrfach wiederholt-> keine Verbesserung Austausch der SC nötig -> keine vorrätig, Labor informiert-> SC ersetzt, Installation wird vom Labor überwacht | Gerät steht seit 25.03.16 in Intervention erforderlich-> Test wiederholt, Startphase abgewartete, div. Anwenderwartungsmassnahmen fällig: Anschlußdichtung ersetzt, thb Kal durchgeführt, nach Startphase pCO2, Glu,Lac Kal fehlerhaft ( Gerät stand eingeschaltet ohne Spülung der SC seit 27.03. lt. Aktivitätendatenbank), Kal mehrfach wiederholt-> keine Verbesserung Austausch der SC nötig -> keine vorrätig, Labor informiert-> SC ersetzt, Installation wird vom Labor überwacht | SWO-90452 | No | Device is required since 25.03.16 in intervention-> Test repeated, start phase waited, various user maintenance measures due: connection seal replaced, thb Kal carried out, after start phase pCO2, Glu,Lac Kal faulty (device was switched on without Rinsing of the SC since 27.03. according to activity database), Kal repeated several times-> no improvement replacement of the SC necessary -> no in stock, laboratory informed-> SC replaced, installation is monitored by the laboratory | Device is required since 25.03.16 in intervention-> Test repeated, start phase waited, various user maintenance measures due: connection seal replaced, thb Kal carried out, after start phase pCO2, Glu,Lac Kal faulty (device was switched on without Rinsing of the SC since 27.03. according to activity database), Kal repeated several times-> no improvement replacement of the SC necessary -> no in stock, laboratory informed-> SC replaced, installation is monitored by the laboratory | Defect | HW | Inlet Gasket | Inlet gasket
ABL90 FLEX | Arrived on site and attended machine to confirm customer reported fault. Pluuged in keyboard to analyser and navigated to C: and D: drive of the machine. Found backup archive data files and removed all filed older than 4 weeks. Rebooted analyser and tested. | B | SWO-90568 | No | Arrived on site and attended machine to confirm customer reported fault. Pluuged in keyboard to analyser and navigated to C: and D: drive of the machine. Found backup archive data files and removed all filed older than 4 weeks. Rebooted analyser and tested. |  | Setup | HW | CPU | HDD
ABL90 FLEX PLUS |  | B | D08417729033F0B3 | No |  |  | Assisted Replacement | SC | SC | SC | PER?
ABL90 FLEX PLUS | Nieuwe inletnaald, wastetubing, pinchvalvetubing en connectorgasket gemonteerd. Na opstarten en controle metingen een bloedmonster bepaald. OK. | Vervolgbezoek brengen om de hemolyzer te vervangen. | SWO-90454 | No | New Inletneedle, wastetubing, pinchvalvetubing and Connectorgasket mounted. After start-up and control measurements a blood sample is determined. Ok. | Follow-up visit to replace the Hemolyzer. | Defect | HW | PM Kit | PM Kit
ABL90 FLEX PLUS | Ran all service test passed ok. Pack ok. Ran sample through unit to add lubricant as unit is new and has no samples through. Unit has started up ok. Cal QC ok | Ran all service test passed ok. Pack ok. Ran sample through unit to add lubricant as unit is new and has no samples through. Unit has started up ok. Cal QC ok | SWO-90542 | No | Ran all service test passed ok. Pack ok. Ran sample through unit to add lubricant as unit is new and has no samples through. Unit has started up ok. Cal QC ok | Ran all service test passed ok. Pack ok. Ran sample through unit to add lubricant as unit is new and has no samples through. Unit has started up ok. Cal QC ok | Assisted Replacement | SC | SC | Wetability
ABL90 FLEX | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-90479 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL90 FLEX |  | B | SWO-90352 | No |  |  | Unknown | SC | Glu interference | Glu interference
ABL90 | Substiuição da cassete de electrodos, estabilização e passagem dos QCs. | B | D08352CC9004DDF0 | No | Substiution of the cassette of electrodes, stabilization and passage of the QCs. |  | Assisted Replacement | SC | SC | Reference
AQT90 FLEX | 更换了光学单元903-364后修复 | 更换了光学单元903-364后修复 | SWO-90248 | No | Repair after replacement of optical unit 903-364 | Repair after replacement of optical unit 903-364 | Defect | HW | Process Module | Optical Unit | 131
ABL827 FLEX | Reseated the EL MET pump tube and ran 2 point calibration. This passed without any rinse errors. | Reseated the EL MET pump tube and ran 2 point calibration. This passed without any rinse errors. | SWO-90528 | No | Reseated the EL MET pump tube and ran 2 point calibration. This passed without any rinse errors. | Reseated the EL MET pump tube and ran 2 point calibration. This passed without any rinse errors. | Defect | HW | El/Met Module | Pump tubing
ABL835 FLEX | Skiftet Inlet probe og studs ved kapillær rens | Skiftet Inlet probe og studs ved kapillær rens | SWO-90408 | No | Switched Inlet probe and studs at the capillary's | Switched Inlet probe and studs at the capillary's | Defect | HW | Inlet module | Nipple kit | Nipple kit
ABL825 FLEX | Verfied high temp - found auto check fan not working. | Replaced the fan and tested. | D083E0C4902CC181 | No | Verfied high temp - found auto check fan not working. | Replaced the fan and tested. | Defect | HW | AutoCheck Module | FAN
ABL835 FLEX | 0316 and 0258 neon lamp errors in log. Neon position number is .099 to .0101 | A new fiber optic and lamp unit were no help. Replaced spectrometer unit and aligned oximeter. | D083FCB8900771A1 | No | 0316 and 0258 neon lamp errors in log. Neon position number is .099 to .0101 | A new fiber optic and lamp unit were no help. Replaced spectrometer unit and aligned oximeter. | Defect | HW | Oxi module | Spectrometer
ABL825 FLEX | Modul gereinigt und gangbar gemacht Kalibration und QC durchgeführt | Modul gereinigt und gangbar gemacht Kalibration und QC durchgeführt | SWO-90465 | No | Module cleaned and made workable calibration and QC performed | Module cleaned and made workable calibration and QC performed | Adjustment | HW | AutoCheck Module | Carroussel
ABL800 BASIC | Sampler tray position PCB erneuert, FlexQ-Modul im Service justiert und überprüft --> ohne Erfolg --> PCB zurück getauscht, FlexQ-Control PCB erneuert, AUinstall (da inkonsistente Software) durchgeführt, System im Service justiert und überprüft --> Testmessungen durchgeführt, System und Funktion i.O. | Sampler tray position PCB erneuert, FlexQ-Modul im Service justiert und überprüft --> ohne Erfolg --> PCB zurück getauscht, FlexQ-Control PCB erneuert, AUinstall (da inkonsistente Software) durchgeführt, System im Service justiert und überprüft --> Testmessungen durchgeführt, System und Funktion i.O. | SWO-90515 | No | Sampler tray position PCB renewed, FlexQ module in service adjusted and checked --> without success --> PCB replaced, FlexQ-Control PCB renewed, AUinstall (because inconsistent software) performed, system in service adjusted and checked --> test measurements system and function in the | Sampler tray position PCB renewed, FlexQ module in service adjusted and checked --> without success --> PCB replaced, FlexQ-Control PCB renewed, AUinstall (because inconsistent software) performed, system in service adjusted and checked --> test measurements system and function in the | Defect | HW | FlexQ Module | PCB
ABL835 FLEX | Ran inlet calibration in service, failed multiple tests Replaced inlet gasket and inlet probe-No change | Replaced friction detector Test in service-OK | D083F4B3905EA796 | No | Ran inlet calibration in service; failed multiple tests Replaced inlet gasket and inlet probe-No change | Replaced friction detector Test in service-OK | Defect | HW | Inlet module | Friction detector unit
ABL837 FLEX | Jeg har skiftet inlet modulet. Jeg har kørt flere test af inlet modulet og alt ser fint ud nu. Alt er OK. | Jeg har skiftet inlet modulet. Jeg har kørt flere test af inlet modulet og alt ser fint ud nu. Alt er OK. | SWO-90689 | No | I changed the inlet module. I have been running several tests of the inlet module and everything looks fine now. Everything is OK. | I changed the inlet module. I have been running several tests of the inlet module and everything looks fine now. Everything is OK. | Defect | HW | Inlet module | Probe cal.
ABL815 FLEX | Halogenlamp justiert, Probelauf, Kal und QC ok | Halogenlamp justiert, Probelauf, Kal und QC ok | SWO-90491 | No | Halogen lamp adjusted, test run, Kal and QC ok | Halogen lamp adjusted, test run, Kal and QC ok | Adjustment | HW | Oxi module | Lamp unit
AQT90 FLEX | Cleaned needle assembly, calibrated height and in to wash block. Repositioned the wash block to inlet PCB cable. Ran system clean and 5 LQC's. | Cleaned needle assembly, calibrated height and in to wash block. Repositioned the wash block to inlet PCB cable. Ran system clean and 5 LQC's. | SWO-90617 | Yes | Cleaned needle assembly, calibrated height and in to wash block. Repositioned the wash block to inlet PCB cable. Ran system clean and 5 LQC's. | Cleaned needle assembly, calibrated height and in to wash block. Repositioned the wash block to inlet PCB cable. Ran system clean and 5 LQC's. | Re classified RMED | Not in PPM
AQT90 FLEX | New Optical Unit fitted. | New Optical Unit fitted. | SWO-90527 | Yes | New Optical Unit fitted. | New Optical Unit fitted. | Re classified RMED | Not in PPM
AQT90 FLEX | Keine weitere Maßnahmen notwendig. Anwender haben angefangene Testkassetten vom alten AQT in neuen eingelegt. Ausgestoßene Testcups werden aber nicht erkannt. AQT hat Keine Testcups erkannt und ging in Fehler mit Meldung 3553. Angefangene Testkassetten dürfen nicht in einen anderen AQT eingelegt werden. Keine weitere Beanstandung beim AQT vorhanden. Systemüberprüfung und Qualitätskontrolle o.K. | Keine weitere Maßnahmen notwendig. Anwender haben angefangene Testkassetten vom alten AQT in neuen eingelegt. Ausgestoßene Testcups werden aber nicht erkannt. AQT hat Keine Testcups erkannt und ging in Fehler mit Meldung 3553. Angefangene Testkassetten dürfen nicht in einen anderen AQT eingelegt werden. Keine weitere Beanstandung beim AQT vorhanden. Systemüberprüfung und Qualitätskontrolle o.K. | SWO-90667 | No | No further action is necessary. Users have inserted started test cassettes from the old AQT into new ones. However, outcast test cups are not recognized. AQT did not recognize any test cups and went into error with message 3553. Started test cassettes may not be inserted into another AQT. No further complaint exists with the AQT. System control and quality control o.K. | No further action is necessary. Users have inserted started test cassettes from the old AQT into new ones. However, outcast test cups are not recognized. AQT did not recognize any test cups and went into error with message 3553. Started test cassettes may not be inserted into another AQT. No further complaint exists with the AQT. System control and quality control o.K. | User error | Consumables | Test Kit | Cartridge re-use
TCM-COMBIM | Check sensor module connections, reboot monitor several times could not reproduce the error. | Performed CombiM monitor performance test and sensor simulator test - Passed | D083F002900DCF21 | No | Check sensor module connections, reboot monitor several times could not reproduce the error. | Performed CombiM monitor performance test and sensor simulator test - Passed | No fault found | No fault found | No fault found | No fault found
TCM/Tosca | replicate error. Opened TCM and cf card was dislodged, | Opened TCM and placed cf back in place. replaced broken handle | D0842F559036C888 | No | replicate error. Opened TCM and cf card was dislodged, | Opened TCM and placed cf back in place. replaced broken handle | Defect | HW | Handle
ABL80 FLEX CO-OX | 清洁血氧模块正常。 | 清洁血氧模块正常。 | SWO-90911 | No | Clean the blood oxygen module normally. | Clean the blood oxygen module normally.
ABL80 FLEX BASIC | 更换测试卡正常。 | 更换测试卡正常。 | SWO-90832 | No | The replacement test card is OK. | The replacement test card is OK.
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-92570 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL80 FLEX BASIC | SC接続不良、再取付け、接続確認。センサーボードチェック | SC接続不良、再取付け、接続確認。センサーボードチェック | SWO-91099 | No | SC connection failure, re-installation, connection confirmation. Sensor board check | SC connection failure, re-installation, connection confirmation. Sensor board check
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-91969 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL90 FLEX PLUS | Kupplung 886-913 schleift, neu justiert und befestigt | Kupplung 886-913 schleift, neu justiert und befestigt | SWO-90879 | No | Coupling 886-913 grinds, re-adjusted and fastened | Coupling 886-913 grinds, re-adjusted and fastened | Adjustment | HW | Auto Inlet module | Inlet Coupling
ABL90 FLEX | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-91975 | No | Repair after replacing a new test card | Repair after replacing a new test card | Assisted Replacement | SC | SC | Glu
AQT90 FLEX |  | B | SWO-90590 | No |  |  | Defect | HW | Process Module | Optical Unit | 129
ABL827 FLEX | Cleaned the valve block and replaced the pH BG module air valve membrane. Also ensured the smooth operation of both module valves. | Cleaned the valve block and replaced the pH BG module air valve membrane. Also ensured the smooth operation of both module valves. | SWO-90806 | No | Cleaned the valve block and replaced the pH BG module air valve membrane. Also ensured the smooth operation of both module valves. | Cleaned the valve block and replaced the pH BG module air valve membrane. Also ensured the smooth operation of both module valves. | Defect | HW | pH/BG Module | Valve, air | Membrane
ABL825 FLEX | Gerinnsel vor Ca und Ablagerungen vor pH entfernt. Alle Messmodul-Ventile ausgebaut und kontrolliert, keine Auffälligkeiten. Sichtkontrolle aller Komponenten, OK. Test aller Fluidicmodul-Ventile im Servicemenü, OK. Isolationstest, OK. SW Update auf V6.16 und Anpassung der pCO2 Zielwerte durch Installation der aktualisierten Beipackzettel der verwendeten QC Lose mit Anpassung der Zielwerte in der Radiance Konfiguration. Austausch des kompletten pH BG Moduls mit Einsatz der Original-Elektroden. Cl Elektrode (Parameter nicht aktiviert) durch Dummy ersetzt. Neustart, mehrere Kalibrationen, alle ohne Auffälligkeiten, Endkontrolle: 3 Level AC5+ gemessen => Alle Werte OK | Gerinnsel vor Ca und Ablagerungen vor pH entfernt. Alle Messmodul-Ventile ausgebaut und kontrolliert, keine Auffälligkeiten. Sichtkontrolle aller Komponenten, OK. Test aller Fluidicmodul-Ventile im Servicemenü, OK. Isolationstest, OK. SW Update auf V6.16 und Anpassung der pCO2 Zielwerte durch Installation der aktualisierten Beipackzettel der verwendeten QC Lose mit Anpassung der Zielwerte in der Radiance Konfiguration. Austausch des kompletten pH BG Moduls mit Einsatz der Original-Elektroden. Cl Elektrode (Parameter nicht aktiviert) durch Dummy ersetzt. Neustart, mehrere Kalibrationen, alle ohne Auffälligkeiten, Endkontrolle: 3 Level AC5+ gemessen => Alle Werte OK | SWO-90794 | No | Removed clots from Ca and deposits before pH. All measuring module valves are removed and controlled, no abnormalities. Visual inspection of all components, OK. Test of all fluidic module valves in the service menu, OK. Isolation test, OK. SW update to V6.16 and adjust the pCO2 target values by installing the updated package leaflets of the used QC lots with adjustment of the target values in the radiance configuration. Replacement of the complete pH BG module with the use of the original electrodes. Cl Electrode (parameter not activated) replaced by dummy. Restart, multiple calibrations, all without abnormalities, final control: 3 levels AC5+ measured => All values OK | Removed clots from Ca and deposits before pH. All measuring module valves are removed and controlled, no abnormalities. Visual inspection of all components, OK. Test of all fluidic module valves in the service menu, OK. Isolation test, OK. SW update to V6.16 and adjust the pCO2 target values by installing the updated package leaflets of the used QC lots with adjustment of the target values in the radiance configuration. Replacement of the complete pH BG module with the use of the original electrodes. Cl Electrode (parameter not activated) replaced by dummy. Restart, multiple calibrations, all without abnormalities, final control: 3 levels AC5+ measured => All values OK | Defect | HW | pH/BG Module | ph/BG Module
ABL825 FLEX | 关机重启后正常。 | 关机重启后正常。 | SWO-90834 | No | Normal after the shutdown restarts. | Normal after the shutdown restarts. | Unknown | HW | DMS_PC | Touch Screen | Reset analyzer
ABL837 FLEX | Skiftet GAS1 Kal. + QC | Skiftet GAS1 Kal. + QC | SWO-90764 | No | Skiftet GAS1 Kal. + QC | Skiftet GAS1 Kal. + QC | Assisted Replacement | Gas Cylinder | Gas Cylinder | Gas 1
ABL825 FLEX | Gerät geprüft. Elektroden überprüft. Ventile im PH BG Modul geprüft. - kein Fehler. Barowert geprüft und eingegeben. - 743mmHg Diverse Kal und QC durchgeführt. | Gerät geprüft. Elektroden überprüft. Ventile im PH BG Modul geprüft. - kein Fehler. Barowert geprüft und eingegeben. - 743mmHg Diverse Kal und QC durchgeführt. | SWO-91191 | No | device checked. electrodes. valves in the PH BG module. - no error. Baro value checked and entered. - 743mmHg Various Kal and QC performed. | device checked. electrodes. valves in the PH BG module. - no error. Baro value checked and entered. - 743mmHg Various Kal and QC performed. | Adjustment | HW | Analyzer Control Module | Barometer
ABL810 FLEX | Die betroffenen Parameter Membraniert. | Die betroffenen Parameter Membraniert. | SWO-90813 | No | The affected parameters are membraneed. | The affected parameters are membraneed. | Assisted Replacement | Electrode | Membrane | pCO2
ABL815 FLEX | Verschlauchung überprüft und überarbeitet. Einlassnadel und Pumpenschläuche der Elektrodenmodule ersetzt. Pumpschläuche aus Bestand Überhang Wartungskits. | Verschlauchung überprüft und überarbeitet. Einlassnadel und Pumpenschläuche der Elektrodenmodule ersetzt. Pumpschläuche aus Bestand Überhang Wartungskits. | SWO-90697 | Yes | Hose checked and revised. inlet needle and pump hoses of the electrode modules. Pump hoses made of stock overhang maintenance kits. | Hose checked and revised. inlet needle and pump hoses of the electrode modules. Pump hoses made of stock overhang maintenance kits. | Re classified RMED | Not customer (DEMO/OOB)
AQT90 FLEX | Fehlersuche: Reagenzröhrchen falsch herum eingesetzt, daher Nadel verbogen! Nadel erneuert, QC i.O. | Fehlersuche: Reagenzröhrchen falsch herum eingesetzt, daher Nadel verbogen! Nadel erneuert, QC i.O. | SWO-90841 | No | Troubleshooting: Reagent tubes inserted incorrectly, therefore needle bent! Needle renewed, QC i.O. | Troubleshooting: Reagent tubes inserted incorrectly, therefore needle bent! Needle renewed, QC i.O. | User error | HW | Aspiration Module | Bend needle
AQT90 FLEX | Need to replace needle | B | SWO-90857 | No | Need to replace needle |  | Defect | HW | Aspiration Module | Bend needle
ABL80 FLEX BASIC | 更换两张测试卡正常。 | 更换两张测试卡正常。 | SWO-90936 | No | Replacing the two test cards is ok. | Replacing the two test cards is ok.
ABL80 CO-OX | Troubleshot and found a clot in the fluid path. | Flushed the fluid path and this resolved the problem. | D08673897C19A633 | No | Troubleshot and found a clot in the fluid path. | Flushed the fluid path and this resolved the problem.
ABL90 FLEX | Zeitserver Einstellung vorgenommen i.o Nach Neustart Zeit i.o | Zeitserver Einstellung vorgenommen i.o Nach Neustart Zeit i.o | SWO-90888 | No | Time server setting made i.o After reboot time i.o | Time server setting made i.o After reboot time i.o | Setup | SW | Time setting | Time setting
ABL90 FLEX | 更换测试卡正常。 | 更换测试卡正常。 | SWO-90921 | No | The replacement test card is OK. | The replacement test card is OK. | Assisted Replacement | SC | SC | Reference
AQT90 FLEX | 从公司找了一张定标卡，重新定标正常。 | 从公司找了一张定标卡，重新定标正常。 | SWO-90659 | No | From the company to find a calibration card, re-calibration normal. | From the company to find a calibration card, re-calibration normal. | Assisted Replacement | Consumables | Test Kit | CKMB
ABL835 FLEX | Zamjena igle i crijeva između igle i elektroda. Čišćenje linija otpada. Dekontaminacija instrumenta. Preporuka: koristiti Clot catcher (Kat.br. 906-020) | Zamjena igle i crijeva između igle i elektroda. Čišćenje linija otpada. Dekontaminacija instrumenta. Preporuka: koristiti Clot catcher (Kat.br. 906-020) | SWO-91069 | No | Replacing the needle and hose between the needle and the Electris. Cleaning the lines of debris. Decontamination of the instrument. Recommendation: Use Clot catcher (Kat.br. 906-020) | Replacing the needle and hose between the needle and the Electris. Cleaning the lines of debris. Decontamination of the instrument. Recommendation: Use Clot catcher (Kat.br. 906-020) | Cleaned | Fluidic pathway | Inlet module | Clot
ABL825 FLEX |  | B | D08695AE7C412786 | Yes |  |  | Re classified RMED | 6.16 SWU
ABL835 FLEX | Fehlersuche, Empf. und Driftfehler trotz mehrmaliger Membranierung, Messkammern nochmals gründlich gereinigt (El Met Modul), fehlenden unteren Dichtungsring angebracht, Elektroden- und Messkappenkontakte mit Glasfaserpinsel gereinigt, Reinigungsprogramm, Kal.1 Lösung neu, Pumpenkal. und LS Justage, Empf. jetzt ok bei 484pA mM System-und Funktionskontrolle i.O. | Fehlersuche, Empf. und Driftfehler trotz mehrmaliger Membranierung, Messkammern nochmals gründlich gereinigt (El Met Modul), fehlenden unteren Dichtungsring angebracht, Elektroden- und Messkappenkontakte mit Glasfaserpinsel gereinigt, Reinigungsprogramm, Kal.1 Lösung neu, Pumpenkal. und LS Justage, Empf. jetzt ok bei 484pA mM System-und Funktionskontrolle i.O. | SWO-90959 | No | Troubleshooting, Empf. and drift fault sizing despite several times membraneing, measuring chambers again thoroughly cleaned (El Met module), missing lower sealing ring attached, electrode and measuring cap contacts cleaned with fiberglass brush, cleaning program, Kal.1 solution new, pump scale. and LS Adjustment, Empf. now ok at 484pA mM system and function control i.O. | Troubleshooting, Empf. and drift fault sizing despite several times membraneing, measuring chambers again thoroughly cleaned (El Met module), missing lower sealing ring attached, electrode and measuring cap contacts cleaned with fiberglass brush, cleaning program, Kal.1 solution new, pump scale. and LS Adjustment, Empf. now ok at 484pA mM system and function control i.O. | Assisted Replacement | Electrode | Electrode | Lac | Sealing ring
ABL825 FLEX | Fehlersuche, LS-Justage im Service - i.O., in Absprache mit Kunden pH-Gerinnsel-Erkennung eingeschaltet, Kal u. QC i.O. | Fehlersuche, LS-Justage im Service - i.O., in Absprache mit Kunden pH-Gerinnsel-Erkennung eingeschaltet, Kal u. QC i.O. | SWO-90902 | No | Troubleshooting, LS-Adjustment in service - i.O., switched on in consultation with customer pH clot detection, Kal & QC i.O. | Troubleshooting, LS-Adjustment in service - i.O., switched on in consultation with customer pH clot detection, Kal & QC i.O. | Adjustment | HW | pH/BG Module | LS-adjustment
ABL837 FLEX | the monitor was not functioning properly | The monitor was replaced and the tBil issues with patients was discussed. A complete OXI test was performed and the QC was checked was checked with nothing out of the ordinary to report. A decontaminatin procedure was discussed as well to perform on a monthly routine. Need appropriate # of dummy electrodes to run the decontamination procedure. | D086769A7C1AA11A | No | the monitor was not functioning properly | The monitor was replaced and the tBil issues with patients was discussed. A complete OXI test was performed and the QC was checked was checked with nothing out of the ordinary to report. A decontaminatin procedure was discussed as well to perform on a monthly routine. Need appropriate # of dummy electrodes to run the decontamination procedure. | Defect | HW | DMS_PC | Touch Screen
AQT90 FLEX | Adjusted scanner alignment and tested. Cleaned scanner encapsulation casing. | Adjusted scanner alignment and tested. Cleaned scanner encapsulation casing. | SWO-90950 | Yes | Adjusted scanner alignment and tested. Cleaned scanner encapsulation casing. | Adjusted scanner alignment and tested. Cleaned scanner encapsulation casing. | Re classified RMED | Not in PPM
TCM400 |  | B | SWO-90869 | No
ABL80 FLEX BASIC | Reinstalled the SC but Lac was showing error with Drift error, flushed the hot water but problem is remaining same, replace with new SC found it was working fine. | Reinstalled the SC but Lac was showing error with Drift error, flushed the hot water but problem is remaining same, replace with new SC found it was working fine. | SWO-91160 | No | Reinstalled the SC but Lac was showing error with Drift error, flushed the hot water but problem is remaining same, replace with new SC found it was working fine. | Reinstalled the SC but Lac was showing error with Drift error, flushed the hot water but problem is remaining same, replace with new SC found it was working fine.
ABL80 OSM | Unable to perform any analysis. Unit lock while customer tried to push users from Aqure | Unable to resolve thru normal means. Issue was resolved by performing a system reset . Performed reset, installed customer setup. Installed solution pack from customer inventory, which will be replaced. | D087CECAF2191025 | No | Unable to perform any analysis. Unit lock while customer tried to push users from Aqure | Unable to resolve thru normal means. Issue was resolved by performing a system reset . Performed reset, installed customer setup. Installed solution pack from customer inventory, which will be replaced.
ABL90 FLEX | 回路徹底洗浄、ソフトウェアバージョンアップ(3.1MR3→MR7)実施。サービス画面にて各ファクター確認。Cal QC問題なし。 | 回路徹底洗浄、ソフトウェアバージョンアップ(3.1MR3→MR7)実施。サービス画面にて各ファクター確認。Cal QC問題なし。 | SWO-91184 | No | Thorough circuit cleaning and software upgrade (3.1MR3→MR7). Check each factor on the service screen. Cal QC no problem. | Thorough circuit cleaning and software upgrade (3.1MR3→MR7). Check each factor on the service screen. Cal QC no problem. | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway | Wetability ?
ABL90 |  | B | D08B80D65419D286 | No |  |  | Assisted Replacement | SC | SC | Lac
ABL90 FLEX | Gerät überprüft Blutmessungen durchgeführt, Übertragung überprüft - ok es werden die nicht temperaturkorrigierten Werte vom Roche Cobas System übernommen, bei diesen Parametern erscheint auch auf dem Ausdruck kein Wert! kein Gerätefehler | Gerät überprüft Blutmessungen durchgeführt, Übertragung überprüft - ok es werden die nicht temperaturkorrigierten Werte vom Roche Cobas System übernommen, bei diesen Parametern erscheint auch auf dem Ausdruck kein Wert! kein Gerätefehler | SWO-91205 | No | Device checks blood measurements performed, transmission checked - ok the non-temperature-corrected values are taken over by the Roche Cobas system, with these parameters also no value appears on the expression! No device error | Device checks blood measurements performed, transmission checked - ok the non-temperature-corrected values are taken over by the Roche Cobas system, with these parameters also no value appears on the expression! No device error | No fault found | Temp corr. | pO2 | pO2
ABL90 FLEX PLUS | Timesettings nochmals durchgeführt. Zeitgeist war deaktiviert. Kalibration und QC's vorgenommen. | Timesettings nochmals durchgeführt. Zeitgeist war deaktiviert. Kalibration und QC's vorgenommen. | SWO-91222 | No | timesettings again. Zeitgeist was disabled. calibration and QC's. | timesettings again. Zeitgeist was disabled. calibration and QC's. | Setup | SW | Time setting | Time setting
AQT90 FLEX | 更换了孵育盘PCB 903-138后修复 | 更换了孵育盘PCB 903-138后修复 | SWO-90854 | No | Repairafter after replacement of incubation disk PCB 903-138 | Repairafter after replacement of incubation disk PCB 903-138 | Defect | HW | Incubator module | PCB, Incubator | 161
ABL815 FLEX | Membranventile der Module kontrolliert, Referenzelektrode kontrolliert, Messkammer pH & Ref gereinigt, Pumpenschläuche der Messmodule ersetzt, PumpenKal, korrekte Kal abgewartet, QK ok | B | SWO-91267 | No | Diaphragm valves of the modules checked, reference electrode checked, measuring chamber pH & Ref cleaned, pump hoses of the measuring modules replaced, pumpsKal, correct scale awaiting, QK ok |  | Defect | HW | pH/BG Module | Pump tubes
ABL800 BASIC | オキシPCB交換、オートチェックモジュール入れ替え（実験室納品機のもの）。原因不明のため現在持ち帰り中。 | B | SWO-91104 | No | Oxy PCB replacement, auto check module replacement (laboratory delivery machine). It is being taken home now for unknown reason. |  | Defect | HW | Oxi module | PCB
ABL820 FLEX | March Stats compared to analyzer 2323N 323: 97.9 - 92, 68.9 - 64.7, 41.3 - 38.1, 21.9 - 20.7. Low on all pCO2. Replaced pCO2 electrode and gas tanks. Exchanged BG module. | Results have increased slightly to 21, 39.2, 65.7. will monitor for a few days and return to further troubleshoot. | D088E97868140463 | No | March Stats compared to analyzer 2323N 323: 97.9 - 92, 68.9 - 64.7, 41.3 - 38.1, 21.9 - 20.7. Low on all pCO2. Replaced pCO2 electrode and gas tanks. Exchanged BG module. | Results have increased slightly to 21, 39.2, 65.7. will monitor for a few days and return to further troubleshoot. | Assisted Replacement | Electrode | Electrode | pCO2
ABL825 FLEX | Allgemeine Kontrolle ok. CPU Unit ersetzt ok. Sybase geladen System i.o. Kal i.o | Allgemeine Kontrolle ok. CPU Unit ersetzt ok. Sybase geladen System i.o. Kal i.o | SWO-91174 | No | General control ok. CPU Unit replaces ok. Sybase loaded system i.o. Kal i.o | General control ok. CPU Unit replaces ok. Sybase loaded system i.o. Kal i.o | Defect | HW | DMS_PC | CPU
AQT90 FLEX | Systemreinigung ausgelöst Systemüberprüfung durchgeführt Neue QC Lösung TNT Level 2 gestartet => Ergebnis i.O. Reinigungslösung auf Station gelassen | Systemreinigung ausgelöst Systemüberprüfung durchgeführt Neue QC Lösung TNT Level 2 gestartet => Ergebnis i.O. Reinigungslösung auf Station gelassen | SWO-91226 | No | System cleaning triggered System check performed New QC solution TNT Level 2 started => Result i.O. cleaning solution left on station | System cleaning triggered System check performed New QC solution TNT Level 2 started => Result i.O. cleaning solution left on station | User error | Consumables | Cleaning Solution | Tubes
ABL80 FLEX BASIC | 更换测试卡 | 更换测试卡 | SWO-91342 | No | Replace the test card | Replace the test card
ABL825 FLEX | Leihmodul (838R1547N009), tHb Kalibration, Funktionskontrolle. | Leihmodul (838R1547N009), tHb Kalibration, Funktionskontrolle. | SWO-91336 | No | Leihmodul (838R1547N009), tHb Kalibration, Funktionskontrolle. | Leihmodul (838R1547N009), tHb Kalibration, Funktionskontrolle. | Defect | HW | Oxi module | Neon Lamp, low
ABL80 FLEX BASIC | Found that Lac parameter was showing low, flushed with warm water 4 times but problem was remaining same, waited for New SC, received the Installed the New SC found Lac parameter was functional. | Found that Lac parameter was showing low, flushed with warm water 4 times but problem was remaining same, waited for New SC, received the Installed the New SC found Lac parameter was functional. | SWO-92283 | No | Found that Lac parameter was showing low, flushed with warm water 4 times but problem was remaining same, waited for New SC, received the Installed the New SC found Lac parameter was functional. | Found that Lac parameter was showing low, flushed with warm water 4 times but problem was remaining same, waited for New SC, received the Installed the New SC found Lac parameter was functional.
ABL90 FLEX | Gerät überprüft, Meßsystem durchgespült. Meßkassette ausgetauschtSC 90 1200 946-060 R0014 eingebaut, tHb Kal o.k. | Gerät überprüft, Meßsystem durchgespült. Meßkassette ausgetauschtSC 90 1200 946-060 R0014 eingebaut, tHb Kal o.k. | SWO-91713 | No | device checked, measuring system flushed through. Measuring cassette replacedSC 90 1200 946-060 R0014 built-in, tHb Kal ok. | device checked, measuring system flushed through. Measuring cassette replacedSC 90 1200 946-060 R0014 built-in, tHb Kal ok. | User error | HW | Inlet Open | Inlet Open
ABL90 |  | B | D08B8BB754343A83 | No |  |  | Assisted Replacement | SC | SC | Glu
ABL90 FLEX | tHb-Kal., Kal. + QC i.O. | tHb-Kal.; Kal. + QC i.O. | SWO-91595 | No | tHb-Kal.; Kal. + QC i.O. | tHb-Kal.; Kal. + QC i.O. | Adjustment | HW | Oxi module | tHb Cal
ABL90 FLEX PLUS |  | B | D08B5F2354413F51 | No |  |  | Assisted Replacement | SC | SC | Reference | PER?
ABL827 FLEX | Even though the El Met module valves looked OK they were old so new ones were fitted. While the analyser was starting up a hole was found in the tube between the pre-heater & the measuring chambers on the El Met module so a new one was fitted. | Even though the El Met module valves looked OK they were old so new ones were fitted. While the analyser was starting up a hole was found in the tube between the pre-heater & the measuring chambers on the El Met module so a new one was fitted. | SWO-91525 | No | Even though the El Met module valves looked OK they were old so new ones were fitted. While the analyser was starting up a hole was found in the tube between the pre-heater & the measuring chambers on the El Met module so a new one was fitted. | Even though the El Met module valves looked OK they were old so new ones were fitted. While the analyser was starting up a hole was found in the tube between the pre-heater & the measuring chambers on the El Met module so a new one was fitted. | Defect | HW | pH/BG Module | Tube
ABL835 FLEX | Se limpia espejo de módulo de autocheck al apreciarse mucho polvo. Se comprueba funcionamiento, se reconocen correctamente los controles en el carrusel. Se observan fallos de calibración del lactato. Se sustituye membrana y se calibra equipo. Se comprueba funcionamiento mediante QC. | B | D09F6F6BB40BAE5F | No | Self-checking module mirror is cleaned when a lot of dust is appreciated. It checks operation, correctly recognizes the controls on the carousel. Calibration failures of the lactate are observed. Membrane is replaced and equipment calibrated. Operation is checked by QC. |  | Cleaned | HW | Fluidic pathway | Fluidic pathway
ABL825 FLEX | Gerät überpüft, Fluidic channel, V7 und V 8 erneuert | Gerät überpüft, Fluidic channel, V7 und V 8 erneuert | SWO-91711 | No | Device superseded, Fluidic channel, V7 and V 8 renewed | Device superseded, Fluidic channel, V7 and V 8 renewed | Defect | HW | Fluidic manifold | Valve
ABL827 FLEX | Einbau des Ersatz Oxymetrie Moduls in den Analysator. Einspielung der Analysator Software ins Modul. Kalibration und tHb Kalibration im Service und im Anwender Bereich durchgeführt. Kalibrationen und Qualitätskontrolle o. K. | Tausch des Oxymetrie Moduls. Einstellung des Moduls. tHb Kalibration durchgefallen. Fehlermeldung: Halogen lamp voltage too high . Fahrt zur Geschäftsstelle und Abholen eines Ersatz-Moduls. Oxymetrie Werte nicht messbar. | SWO-91610 | No | Installation of the replacement oxymetry module in the analyzer. Inserting the analyzer software into the module. calibration and tHb calibration in the service and in the user area. Calibrations and quality control o.K. | Replacement of the oxymetry module. Setting the module. tHb calibration. Error: Halogen lamp voltage too high . Drive to the office and pick up a replacement module. Oxymetry values cannot be measured. | Defect | HW | Oxi module | Halogen Lamp
ABL827 FLEX | Verschlissene Einlaßsonde sowie Einlaßdichtung ersetzt | B | SWO-92665 | No | Replaced worn inlet probe and inlet seal |  | Replaced | HW | Inlet module | Inlet gasket
ABL827 FLEX | Customer performed protein removal | Replaced hemolyzer. Adjusted lamp voltage. Ran tHB cal in service and customer mode | D08B913D544EFB62 | No | Customer performed protein removal | Replaced hemolyzer. Adjusted lamp voltage. Ran tHB cal in service and customer mode | Defect | HW | Oxi module | Hemolyzer
ABL825 FLEX | Revisit to replace the Oximetry PCB board and, finish installation. The complete oxdimetry test Neon position is 1.901 and the service ctHb calibration fails, status 31. | Replaced all listed part to resolve error 315 Neon number too high. Adjusted the lamp and, performed a complete oximetry test, all passing within specifications. | D08BC09C545F6950 | No | Revisit to replace the Oximetry PCB board and, finish installation. The complete oxdimetry test Neon position is 1.901 and the service ctHb calibration fails, status 31. | Replaced all listed part to resolve error 315 Neon number too high. Adjusted the lamp and, performed a complete oximetry test, all passing within specifications. | Defect | HW | Oxi module | PCB, Oxi
ABL835 FLEX |  | B | D08BD52A54223A87 | No |  |  | Defect | HW | Inlet module | Probe bend | bowed needle aspiration
AQT90 FLEX |  | B | D08B54F05460DB28 | No |  |  | Defect | SW | Analyzer Control Module | Re installed SW
ABL80 FLEX BASIC | 更换测试卡。 | 更换测试卡。 | SWO-91705 | No | Replace the test card. | Replace the test card.
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-92413 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL80 FLEX BASIC | 更换两张测试卡正常。 | 更换两张测试卡正常。 | SWO-92968 | No | Replacing the two test cards is ok. | Replacing the two test cards is ok.
ABL80 CO-OX | Will replace Sensor Conn PCB and do an annual PM. | REplaced the Sensor Conn PCB and did the annual PM. The system is still not initialized. Ordered analog PCB. 4 6 2016 Returned with parts and tried a new SP first and that resolved the issue, replaced SP with client stock. Replaced client stock on SC and turned in credit form. | D08D3807CA5D46B9 | No | Will replace Sensor Conn PCB and do an annual PM. | REplaced the Sensor Conn PCB and did the annual PM. The system is still not initialized. Ordered analog PCB. 4 6 2016 Returned with parts and tried a new SP first and that resolved the issue, replaced SP with client stock. Replaced client stock on SC and turned in credit form.
ABL80 FLEX BASIC | 更换测试卡修复仪器 | 更换测试卡修复仪器 | SWO-92659 | No | Replace the test card repair instrument | Replace the test card repair instrument
ABL90 FLEX | 更换测试卡正常 | 更换测试卡正常 | SWO-91984 | No | Replace the test card ok | Replace the test card ok | Assisted Replacement | SC | SC | Reference | 1361
ABL800 FLEX | As suggested by Staff, Units and reference ranges have been changed. Anlyser solutions replaced. Analyser serviced and calibrated properly, handed over in a fine working condition. | As suggested by Staff, Units and reference ranges have been changed. Anlyser solutions replaced. Analyser serviced and calibrated properly, handed over in a fine working condition. | SWO-92129 | Yes | As suggested by Staff, Units and reference ranges have been changed. Anlyser solutions replaced. Analyser serviced and calibrated properly, handed over in a fine working condition. | As suggested by Staff, Units and reference ranges have been changed. Anlyser solutions replaced. Analyser serviced and calibrated properly, handed over in a fine working condition. | Re classified RMED | Setup
ABL825 FLEX | Se revisan y ajustan tubos de módulos y bombas. Se hacen calibraciones de sensores de líquido, bombas y entrada. Se hacen calibraciones a dos puntos. Se pasan controles de calidad. Se rellena carrusel de QC. | B | D09F7283B456AE63 | No | Module tubes and pumps are checked and adjusted. Calibrations of liquid sensors, pumps and input are made. Two-point calibrations are made. Quality controls are passed. QC carousel is filled. |  | Cleaned | HW | Fluidic pathway | Fluidic pathway
ABL825 FLEX | Gasflaschendruck überprüft. o.k. Flowregulator 2 und Drucksensor erneuert. Einlaufverfahren und Servicetests durchgeführt. Nach Neustart A D converter failed. Neustart durchgeführt. Alle Funktionen wieder o.k. | Gasflaschendruck überprüft. o.k. Flowregulator 2 und Drucksensor erneuert. Einlaufverfahren und Servicetests durchgeführt. Nach Neustart A D converter failed. Neustart durchgeführt. Alle Funktionen wieder o.k. | SWO-91953 | No | gas cylinder pressure checked. o.k. Flowregulator 2 and pressure sensor renewed. inlet procedures and service tests. After rebootIng A D converter failed. Restart. All functions again ok. | gas cylinder pressure checked. o.k. Flowregulator 2 and pressure sensor renewed. inlet procedures and service tests. After rebootIng A D converter failed. Restart. All functions again ok. | Defect | HW | Gas module | Flow regulator 2
ABL837 FLEX | found that there were demo keys that were not disabled | disabled all demo keys on analyzer. | D08CC234CA19ACC1 | No | found that there were demo keys that were not disabled | disabled all demo keys on analyzer. | Defect | SW | Demo Key Exp | Demo Key Exp
ABL835 FLEX | Looked in datalogs, Found that Fariba changed Reference membrane on April 1 2016 with lot number 1923 and all 3 levels shifted down to low end. Recommend software upgrade. Downloaded all data and setup. First attempt to upgrade had corrupt files. | Had to Reload Windows and old software to 6.15 and updated data and setup then Upgraded to 6.16 | D08CE9C4CA50D469 | No | Looked in datalogs, Found that Fariba changed Reference membrane on April 1 2016 with lot number 1923 and all 3 levels shifted down to low end. Recommend software upgrade. Downloaded all data and setup. First attempt to upgrade had corrupt files. | Had to Reload Windows and old software to 6.15 and updated data and setup then Upgraded to 6.16 | Assisted Replacement | Electrode | Electrode | pCO2
ABL837 FLEX | Replaced the Reed sensor. pH BG assymetry error noticed intermittently.Error 538 and 793. Checked all logs. The cover is still having issues. If a new solution bottle is installed and the cover is put back on then the cover and the window over the electrodes is detected. | Replaced Reed Sensor. Inspected all tubing. Replaced black tubing from multi connector to ph bg dual valve housing. replaced pump tubing(customer stock) for ph bg and Ei met modules. Checked all pump rotors. Ran pump calibrations, calibrations and LS adjust. | D08CF2AECA4F97D4 | No | Replaced the Reed sensor. pH BG assymetry error noticed intermittently. Error 538 and 793. Checked all logs. The cover is still having issues. If a new solution bottle is installed and the cover is put back on then the cover and the window over the electrodes is detected. | Replaced Reed Sensor. Inspected all tubing. Replaced black tubing from multi connector to ph bg dual valve housing. replaced pump tubing(customer stock) for ph bg and Ei met modules. Checked all pump rotors. Ran pump calibrations, calibrations and LS adjust. | Defect | HW | Basic unit | Cover Sensor
ABL825 FLEX | Gerät geprüft. Ventile im Service überprüft. Airventil PH BG Modul ztw. ohne Funktion- Ventil erneuert. LS Justage und Pumpenkal im Service gestartet, Asymmetriefehler bleibt. Pumpenschläuche erneuert. Ref.Membrane erneuert. Diverse Kal und QC durchgeführt. | Gerät geprüft. Ventile im Service überprüft. Airventil PH BG Modul ztw. ohne Funktion- Ventil erneuert. LS Justage und Pumpenkal im Service gestartet, Asymmetriefehler bleibt. Pumpenschläuche erneuert. Ref.Membrane erneuert. Diverse Kal und QC durchgeführt. | SWO-91830 | No | device checked. valves in the service. Airvalve PH BG module ztw. without function valve. LS adjustment and pump scale started in service, asymmetry error remains. Pump hoses renewed. Ref.Membrane renewed. Various Kal and QC performed. | device checked. valves in the service. Airvalve PH BG module ztw. without function valve. LS adjustment and pump scale started in service, asymmetry error remains. Pump hoses renewed. Ref.Membrane renewed. Various Kal and QC performed. | Defect | HW | pH/BG Module | Valve, air
ABL80 FLEX BASIC | 更换两张测试卡正常。 | 更换两张测试卡正常。 | SWO-91977 | No | Replacing the two test cards is ok. | Replacing the two test cards is ok.
ABL80 FLEX BASIC | 更换测试卡修复仪器 | 更换测试卡修复仪器 | SWO-92346 | No | Replace the test card repair instrument | Replace the test card repair instrument
ABL80 FLEX CO-OX | Kompletten Analysator demontiert und intensiv gereinigt, montiert, neuen LP vom Kunden eingesetzt (Lot 220628), vermutlich fehlerhaften undichten LP aufgrund zu großer Toleranz der Stutzen (Lot219850) reklamiert und entsorgt, Kal, QC | B | SWO-92339 | No | Complete analyzer dismantled and intensively cleaned, assembled, new LP used by the customer (lot 220628), probably faulty leaking LP due to too high tolerance of the nozzles (Lot219850) complained and disposed of, Kal, QC
ABL80 FLEX BASIC |  | B | SWO-92007 | No
ABL80 CO-OX | Performed visual inspection of fluidics, no issues observed or noted. No blockage in analzyer at this time. Confirmed pH & pCO2 I A and pO2 internal QC Sol ID C8302 elevated. | Flushed waste fluid path, cleaned waste port, Reinstalled previously used SC replaced during troubleshooting of blockage, reactivated pH & pCO2. All parameters now within specifications although C8302 continues to trend high but in range. Suspect an issue with the SP internal QC. | D08E49A4405B0B07 | No | Performed visual inspection of fluidics, no issues observed or noted. No blockage in analzyer at this time. Confirmed pH & pCO2 I A and pO2 internal QC Sol ID C8302 elevated. | Flushed waste fluid path, cleaned waste port, Reinstalled previously used SC replaced during troubleshooting of blockage, reactivated pH & pCO2. All parameters now within specifications although C8302 continues to trend high but in range. Suspect an issue with the SP internal QC.
ABL80 FLEX BASIC | 納品して1週間足らずでのエラーのためCFカード、パワーサプライ、CPUボード交換実施。 | B | SWO-97278 | No | Cf card, power supply, CPU board replacement was implemented due to errors in less than a week.
ABL90 FLEX | a készülék a helyszínen nem javítható, szervizbe szállítása szükséges. A javítás idejére csere készülék üzembe helyezése. | B | SWO-92201 | No | The machine is not repairable to be repaired on-site. Installation of replacement device during repair. |  | Unknown | Removed by RSUB | SC Door | SC Door
ABL90 FLEX PLUS | Inlet Modul abgenommen, LP90 entnommen, Einlassdichtung gereinigt, Einlassnadel sah verstopft auf, Einlassnadel gespült - ist wieder frei, Fließwege gespült, LP90 und Inlet wieder installiert, ABL befindet sich in Startphase und Zeit läuft runter, tHb Kal. durchgeführt, System- und Funktionskontrolle in Ordnung. | Inlet Modul abgenommen, LP90 entnommen, Einlassdichtung gereinigt, Einlassnadel sah verstopft auf, Einlassnadel gespült - ist wieder frei, Fließwege gespült, LP90 und Inlet wieder installiert, ABL befindet sich in Startphase und Zeit läuft runter, tHb Kal. durchgeführt, System- und Funktionskontrolle in Ordnung. | SWO-92162 | No | Inlet module removed, LP90 removed, inlet seal cleaned, inlet needle looked clogged, inlet needle flushed again - is free again, flow paths rinsed, LP90 and Inlet reinstalled, ABL is in start phase and time runs down, tHb Kal. system and function control in order. | Inlet module removed, LP90 removed, inlet seal cleaned, inlet needle looked clogged, inlet needle flushed again - is free again, flow paths rinsed, LP90 and Inlet reinstalled, ABL is in start phase and time runs down, tHb Kal. system and function control in order. | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
ABL90 | Actualização do sw para versão 3.1MR6.PT, limpeza do ciruito da amostra, verificação do funcionamento da válvula de oxi, substituição da sonda , da entrada sem sucesso, substituição do pack de soluções, testes no modo de serviço, establização e passagem do QC. | B | D08E5D83400C09C0 | No | Update of SW to version 3.1MR6.PT, cleaning of the sample circuit, check the operation of the Oxy valve, replacement of the probe, entry without success, replacement of the solution pack, tests in the mode of service, establishment and passage of QC. |  | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
ABL90 | Limpeza do circuito, substituição da entrada e da sonda sem sucesso, substituição do pack de soluções, establização e passagem do QC. Tudo ok. | B | D08E388E403BCD37 | No | Cleaning the circuit, replacing the input and the probe without success, replacing the solution pack, establishing and passing the QC. Okay, all right. |  | Assisted Replacement | SP | SP | FT
AQT90 FLEX |  | B | SWO-92067 | No |  |  | Defect | HW | Storage Module | Barcode reader, Storage wheel | BR1
AQT90 FLEX |  | B | SWO-92146 | No |  |  | Defect | HW | Process Module | Process Unit motor | 5784
AQT90 FLEX |  | B | SWO-92014 | No |  |  | Defect | HW | Process Module | Process Unit sensor | 5784
ABL835 FLEX | Nieuwe blauwe o-ring op de electrode gemonteerd. Meetkamer schoongemaakt. Na twee kalibraties is de analyzer weer meetklaar. | Nieuwe blauwe o-ring op de electrode gemonteerd. Meetkamer schoongemaakt. Na twee kalibraties is de analyzer weer meetklaar. | SWO-92117 | No | New Blue O-ring mounted on the electrode. Room cleaned. After two calibrations, the analyzer is ready to be measured again. | New Blue O-ring mounted on the electrode. Room cleaned. After two calibrations, the analyzer is ready to be measured again. | Assisted Replacement | Electrode | Electrode | pH | Sealing ring
ABL835 FLEX | New FPC cable fitted | New FPC cable fitted | SWO-92258 | No | New FPC cable fitted | New FPC cable fitted | Defect | HW | AutoCheck Module | FPC Cable
ABL825 FLEX |  | B | D08DF3AD403DDB7B | No |  |  | Cleaned | Fluidic pathway | Fluidic pathway | Lac
ABL825 FLEX | an ongoing issue that tends to come and go , lactate , glucose, Ica, chloride. | sending new lot of membranes | D08E20294005B397 | No | an ongoing issue that tends to come and go , lactate , glucose, Ica, chloride. | sending new lot of membranes | Assisted Replacement | Electrode | Membrane | Lac | QC Low
ABL827 FLEX | Found in the log files that the Rinse solution was expired. | Replaced the Met II bottle, performed a HI Crea tests and verified that the system was still conntected to Radiance. Customer performed a blood sample and all information transfered with no proble. Will follow up with the customer after the automatic Hi Crea check runs again at 11:30PM to insure that the systems is connected. | D08E8E1B40130BF0 | No | Found in the log files that the Rinse solution was expired. | Replaced the Met II bottle, performed a HI Crea tests and verified that the system was still conntected to Radiance. Customer performed a blood sample and all information transfered with no proble. Will follow up with the customer after the automatic Hi Crea check runs again at 11:30PM to insure that the systems is connected. | Setup | SW | His/lis | His/Lis
ABL820 FLEX | If the syringe flap is opened and closed before the user has logged in, the start button will freeze. However, this does not happen on the customers other two analyzers. | Performed software upgrade to 6.16, and this resolved the problem. | D08E4BBC403D8131 | No | If the syringe flap is opened and closed before the user has logged in, the start button will freeze. However, this does not happen on the customers other two analyzers. | Performed software upgrade to 6.16, and this resolved the problem. | Defect | SW | DMS_PC | Touch Screen
ABL837 FLEX | Jeg har lagt ABL software 6.15 på igen. Jeg har lagt backup setup ind igen. Alt er OK nu og der er igen forbindelse til Radiance. | Jeg har lagt ABL software 6.15 på igen. Jeg har lagt backup setup ind igen. Alt er OK nu og der er igen forbindelse til Radiance. | SWO-92219 | No | I have put the ABL software 6.15 back on. I have put backup setup back in. Everything is OK now and there is again connection to Radiance. | I have put the ABL software 6.15 back on. I have put backup setup back in. Everything is OK now and there is again connection to Radiance. | Defect | SW | Ver 6.16 | Ver 6.16
ABL80 FLEX BASIC | 更换测试卡修复仪器 | 更换测试卡修复仪器 | SWO-94405 | No | Replace the test card repair instrument | Replace the test card repair instrument
ABL80 CO-OX | Completed visual inspection. observed dried, powdery material on a portion of the Sensor Cassette connector indicating fluid had been on the connector. | Installed maintenance kit as precautionary measure. Replace Sensor Cassette connector PCB. Performed applicable service diagnostics test and adjustments. Replaced with SP with lot# 221065 to resolve internal QC issues. Root cause of Internal QC recovery issue was caused by an issue with SP lot#220240 X 2. | D08F5755B62FC9A6 | No | Completed visual inspection. observed dried, powdery material on a portion of the Sensor Cassette connector indicating fluid had been on the connector. | Installed maintenance kit as precautionary measure. Replace Sensor Cassette connector PCB. Performed applicable service diagnostics test and adjustments. Replaced with SP with lot# 221065 to resolve internal QC issues. Root cause of Internal QC recovery issue was caused by an issue with SP lot#220240 X 2.
ABL80 FLEX CO-OX | SC erneut initialisiert, Fehler tritt erneut auf, QCs durchgeführt. Kunde hat keine weitere SC. für den Kunden eine neu SC zu morgen vor 8 Uhr bestellt. | SC erneut initialisiert, Fehler tritt erneut auf, QCs durchgeführt. Kunde hat keine weitere SC. für den Kunden eine neu SC zu morgen vor 8 Uhr bestellt. | SWO-92392 | No | SC reinitialized, error reoccurs, QCs performed. Customer has not ordered another SC. for the customer a new SC tomorrow before 8 o'clock. | SC reinitialized, error reoccurs, QCs performed. Customer has not ordered another SC. for the customer a new SC tomorrow before 8 o'clock.
ABL80 FLEX BASIC | reinstalled the sc80 but fail to calibration. replace the SP & tried but issue remain same. replace the SC 80 | B | SWO-92561 | No | reinstalled the sc80 but fail to calibration. replace the SP & tried but issue remain same. replace the SC 80
ABL90 FLEX | Fandt at opstæning for remote adgang på ABLen var sat forkert. Dette er nu rettet. Under fejlfinding er der også lagt software version 3.1 MR8 på ABLen. Alt er OK igen. | Fandt at opstæning for remote adgang på ABLen var sat forkert. Dette er nu rettet. Under fejlfinding er der også lagt software version 3.1 MR8 på ABLen. Alt er OK igen. | SWO-92930 | No | Found that the upturn for remote access on the ABLen was set incorrectly. This is now corrected. During troubleshooting, software version 3.1 MR8 is also added to the ABLen. Everything is OK again. | Found that the upturn for remote access on the ABLen was set incorrectly. This is now corrected. During troubleshooting, software version 3.1 MR8 is also added to the ABLen. Everything is OK again. | Setup | SW | Roma.ini | NetOP
ABL90 | Found no errors present on analyzer in last 12 hours. All cals check OK, QC's check OK. Activity log shows no errors since midnight sequence. Some flow issues have been occurring 4 5 and 4 6. Inhomogenous sample error occurred. Customer has completed multiple back flushes. SC and SP were replaced. Radiometer service recommended SC replacement on 4 6, but calibrations began to be acceptable. One additional error then occurred on a sample. | Service observed three test samples with no errors. (No others were available) Bay area checked - no leaks visible. Given error history, new inlet probe was installed. Customer advised to run samples normally, or to at least run extra blood available to verify no further errors occur. If any cal, patient, or flow errors occur, then a change of SC should be made. | D09068992C537B4B | No | Found no errors present on analyzer in last 12 hours. All cals check OK, QC's check OK. Activity log shows no errors since midnight sequence. Some flow issues have been occurring 4 5 and 4 6. Inhomogenous sample error occurred. Customer has completed multiple back flushes. SC and SP were replaced. Radiometer service recommended SC replacement on 4 6, but calibrations began to be acceptable. One additional error then occurred on a sample. | Service observed three test samples with no errors. (No others were available) Bay area checked - no leaks visible. Given error history, new inlet probe was installed. Customer advised to run samples normally, or to at least run extra blood available to verify no further errors occur. If any cal, patient, or flow errors occur, then a change of SC should be made. | No fault found | No fault found | Fluidic pathway | FT | Good
ABL90 FLEX | Messstrecke gespült, Serviceprogramme überprüft, selbigen LP installiert, Kal, QC | Messstrecke gespült, Serviceprogramme überprüft, selbigen LP installiert, Kal, QC | SWO-92562 | No | Measuring track flushed, service programs checked, same LP installed, Kal, QC | Measuring track flushed, service programs checked, same LP installed, Kal, QC | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
ABL90 FLEX PLUS |  | B | D08F38E9B64CEDF0 | Yes |  |  | Re classified RMED | PER
ABL90 FLEX |  | B | SWO-92303 | Yes |  |  | Re classified RMED | AQT SWO
ABL90 FLEX |  | B | SWO-92364 | No |  |  | Defect | HW | Basic Unit | Sensor Interface PCB | pO2 (Return Request - 16.5-16)
ABL90 FLEX | 更换新测试卡后修复 | pO2 complete test中，zpO2_DarkNo不通过，按照维修手册更换了903-183后修复，903-165未更换 | SWO-92740 | No | Repair after replacing a new test card | pO2 complete test, zpO2-DarkNo does not pass, according to the repair manual replaced 903-183 after repair, 903-165 not replaced
ABL825 FLEX | 更换测试卡修复仪器 | 更换测试卡修复仪器 | SWO-92389 | No | Replace the test card repair instrument | Replace the test card repair instrument
ABL800 FLEX | エアライン清掃・タンパク除去・LS調整・ポンプCALを行うがCALが通ったり通らなかったり コンプリートテスト・tHbCAL OK だがQC測定でOXI Spectrum mismatchが出たり出なかったり OXIﾓｼﾞｭｰﾙ入替・コンプリートテスト・CAL・tHbCAL・QC測定データ確認 | B | SWO-94384 | No | Airline cleaning, protein removal, LS adjustment, pump CAL, cal does not pass or pass, complete test tHbCAL OK, but OXI Spectrum mismatch does not come out in QC measurement OXI module replacement, complete test, CAL, tHbCAL, QC measurement data confirmation |  | Cleaned | Fluidic pathway | Fluidic manifold | tHb
ABL800 BASIC | Cleaned the inlet casket, after a successful 2 point calibration machine is working fine. | Cleaned the inlet casket, after a successful 2 point calibration machine is working fine. | SWO-92786 | No | Cleaned the inlet casket, after a successful 2 point calibration machine is working fine. | Cleaned the inlet casket, after a successful 2 point calibration machine is working fine. | User error | Setup | Cal Schedule | Cal stopped
ABL827 FLEX | Customer re-membraned electrode | Replaced El Met air valve. Replaced all pump tubing | D08F5904B64D73E5 | No | Customer re-membraned electrode | Replaced El Met air valve. Replaced all pump tubing | Assisted Replacement | Electrode | Membrane | Lac | Calibration
ABL827 FLEX | Inlet Nippel erneuert, Funktionskontrolle. | Inlet Nippel erneuert, Funktionskontrolle. | SWO-92422 | No | Inlet nipple renewed, function control. | Inlet nipple renewed, function control. | Defect | HW | Inlet module | Leaking | Nipple kit
ABL800Basic |  | B | D08F5EB2B62BB700 | Yes |  |  | Re classified RMED | Moved analyzer
AQT90 FLEX | Systemreinigung durchgeführt Mitgebrachte QC Lot 11234 gemessen Mitgebrachte QC auf Station gelassen Fehlerhafte QC Lösungen mitgenommen | Systemreinigung durchgeführt Mitgebrachte QC Lot 11234 gemessen Mitgebrachte QC auf Station gelassen Fehlerhafte QC Lösungen mitgenommen | SWO-92503 | No | System cleaning carried out brought with it QC Lot 11234 measured brought to station left Faulty QC solutions taken | System cleaning carried out brought with it QC Lot 11234 measured brought to station left Faulty QC solutions taken | Cleaned | HW | Process Module | System cleaning
AQT90 FLEX | Cup dryer PCB vervangen. | Cup dryer PCB vervangen. | SWO-92404 | No | Cup dryer PCB vervangen. | Cup dryer PCB vervangen. | Defect | HW | Process Module | PCB, Cup dryer | 168
AQT90 FLEX |  | B | D08F4965B6247637 | No |  |  | Not Cat | Not cat. | Not cat. | Not cat.
ABL90 FLEX | Fließwege gereinigt, neue Schläuche eingebaut. Einlass Sonde ersetzt. | B | SWO-92995 | No | Flow paths cleaned, new hoses installed. Inlet probe replaced. |  | Defect | HW | PM Kit | PM Kit
ABL90 FLEX | Replaced faulty inlet gasket with holder. Replaced faulty Sensor Cassette 946-008R0373N060. | Replaced faulty inlet gasket with holder. Replaced faulty Sensor Cassette 946-008R0373N060. | SWO-92812 | No | Replaced faulty inlet gasket with holder. Replaced faulty Sensor Cassette 946-008R0373N060. | Replaced faulty inlet gasket with holder. Replaced faulty Sensor Cassette 946-008R0373N060. | Assisted Replacement | SC | SC | pO2
ABL90 FLEX | 更换新试剂包后修复 | 更换新试剂包后修复 | SWO-92945 | No | Repair after replacing a new kit | Repair after replacing a new kit | Assisted Replacement | SP | SP | Chip Read
AQT90 FLEX |  | B | SWO-92617 | No |  |  | Defect | HW | Aspiration Module | Needle Wash Unit | 1267
AQT90 FLEX | 更换了Hct模块903-329后修复(by andrew) | 更换了Hct模块903-329后修复(by andrew) | SWO-92360 | No | Repair after replacing Hct module 903-329 (by andrew) | Repair after replacing Hct module 903-329 (by andrew) | Defect | HW | Aspiration Module | Needle Wash Unit | 3414
ABL827 FLEX | Soaked the El Met chambers with protein remover and removed deposits from the Lactate chamber. Electrode connector replaced as it was a recommended countermeasure for problems with low sensitivity. Lactate membrane replaced because of low sensitivity. New Ca membrane fitted because it was failing QC. | Soaked the El Met chambers with protein remover and removed deposits from the Lactate chamber. Electrode connector replaced as it was a recommended countermeasure for problems with low sensitivity. Lactate membrane replaced because of low sensitivity. New Ca membrane fitted because it was failing QC. | SWO-92827 | No | Soaked the El Met chambers with protein remover and removed deposits from the Lactate chamber. Electrode connector replaced as it was a recommended countermeasure for problems with low sensitivity. Lactate membrane replaced because of low sensitivity. New Ca membrane fitted because it was failing QC. | Soaked the El Met chambers with protein remover and removed deposits from the Lactate chamber. Electrode connector replaced as it was a recommended countermeasure for problems with low sensitivity. Lactate membrane replaced because of low sensitivity. New Ca membrane fitted because it was failing QC. | Defect | HW | El/Met Module | Electrode Connector | Lac
ABL827 FLEX | Measuring chambers removed and soaked with protein remover. New sealing ring fitted to the Glu electrode. New Lactate membrane fitted. New Ca membrane fitted because it was failing calibrations. | Measuring chambers removed and soaked with protein remover. New sealing ring fitted to the Glu electrode. New Lactate membrane fitted. New Ca membrane fitted because it was failing calibrations. | SWO-92828 | No | Measuring chambers removed and soaked with protein remover. New sealing ring fitted to the Glu electrode. New Lactate membrane fitted. New Ca membrane fitted because it was failing calibrations. | Measuring chambers removed and soaked with protein remover. New sealing ring fitted to the Glu electrode. New Lactate membrane fitted. New Ca membrane fitted because it was failing calibrations. | Assisted Replacement | Electrode | Membrane | Glu | Sealing ring
ABL827 FLEX |  | B | D09436588E47B707 | No |  |  | Defect | HW | Oxi module | Tubings
ABL825 FLEX | Comprobación del estado del analizador. Se realiza un Oxi complete test, se observa valores de lámpara de Neon elevados. Se procede a la comprobación de conexiones y de lámapara. Se pide nueva targeta PCB. | Se pide targeta OXI PCB | D090A5952C5586F0 | No | Checking the status of the scanner. An Oxi complete test is performed, high Neon lamp values are observed. Connections and lamapara check are carried out. New PCB targeta is requested. | Targeta OXI PCB is ordered | Revisit required | HW | Oxi module | PCB, Oxi
ABL825 FLEX | 更换了硬盘910-319，Au板902-847，血氧单元902-833后修复(by andrew) | 更换了硬盘910-319，Au板902-847，血氧单元902-833后修复(by andrew) | SWO-92388 | No | Replaced hard drive 910-319, Au plate 902-847, oxygen unit 902-833 after repair (by andrew) | Replaced hard drive 910-319, Au plate 902-847, oxygen unit 902-833 after repair (by andrew) | Defect | HW | Basic unit | PCB, Analyzer Control
ABL827 FLEX | Gerät überprüft - Ref.-Membran - rötlich Laut dem Gerätespeicher wurde die Ref.-Membran am 15.03 um 01.53 erneuert. In der anschließenden Kal ist kein Drift vorhanden, daher liegt der Verdacht sehr nahe, dass die Membran fälschlicher weise registriert wurde. Davor wurde die Ref.-Membran am 14.02 bezogen, daher wurde die Membran außerhalb der Radiometer-Vorgaben betrieben. Ref.-Elektrode neu bezogen, Gerät kalibriert, QC´s durchgef. blutige Vergleichsmessungen zum AWR durchgeführt - der Na Wert ist genau identisch. Das Problem mit Frau Albrecht, Dr. Langer und Frau Planz besprochen Es lag kein gerätetechnisches Problem vor | B | SWO-92738 | No | Device checked - Ref.-membrane - reddish According to the device memory, the ref.-membrane was renewed on 15.03 at 01.53. There is no drift in the subsequent cal, so it is very obvious that the membrane was falsely registered. Before that, the ref. membrane was purchased on 14.02, therefore the membrane was operated outside the radiometer specifications. Ref.-electrode re-covered, device calibrated, QC's through. bloody comparative measurements to the AWR - the Na value is exactly the same. The problem discussed with Ms. Albrecht, Dr. Langer and Ms. Planz There was no equipment problem |  | User error | Electrode | Membrane | Reference | Exp time doubled
ABL835 FLEX | Cal 1 needed to be replaced - empty. Inlet Module Liquid sensor is suspect. | Replaced Cal 1 Sol. Replaced Inlet Liquid Sensor.Cleaned tip of pH Electrode and cleaned pH Measuring Chamber. | D09079672C2E72EC | No | Cal 1 needed to be replaced - empty. Inlet Module Liquid sensor is suspect. | Replaced Cal 1 Sol. Replaced Inlet Liquid Sensor.Cleaned tip of pH Electrode and cleaned pH Measuring Chamber. | Defect | HW | Inlet module | LS
ABL837 FLEX | pCO2電極交換×2個、pCO2ﾒﾝﾌﾞﾗﾝﾛｯﾄ変更ﾁｪｯｸ×3ロット、Ref・Clﾒﾝﾌﾞﾗﾝ交換ﾁｪｯｸ×各1回、電極ﾓｼﾞｭｰﾙﾎﾟﾝﾌﾟﾁｭｰﾌﾞ交換×1回。で直らず、ﾓｼﾞｭｰﾙ交換。 | B | SWO-92585 | No | PCO2 electrode replacement × 2 pieces, pCO 2 membrane lot change check × 3 lot, Ref clmenbran exchange check × once each, electrode module pump tube replacement × 1 times. module replacement without fixing it. |  | Assisted Replacement | Electrode | Electrode | pCO2
ABL810 FLEX |  | B | D094A9688E2C057E | No |  |  | Adjustment | HW | Oxi module | Lamp unit
AQT90 FLEX | cleaned all the sensors as well as removed the broken cartridge. still having problem with cartridge slider. | cleaned all the sensors as well as removed the broken cartridge. still having problem with cartridge slider. | SWO-92779 | No | cleaned all the sensors as well as removed the broken cartridge. still having problem with cartridge slider. | cleaned all the sensors as well as removed the broken cartridge. still having problem with cartridge slider. | Defect | Consumables | Test Kit | Cartridge | 5784 (Broken or Label)
TCM CombiM | CPUﾎﾞｰﾄﾞ交換 | B | SWO-92766 | No | CPU board replacement |  | Defect | HW | PCB, Main CPU | PCB, Main CPU
ABL80 FLEX BASIC | 将公司备用机给医院使用，仪器返回公司库房继续检修。 | 将公司备用机给医院使用，仪器返回公司库房继续检修。 | SWO-92924 | No | The company spare machine to the hospital to use, the instrument returned to the company warehouse to continue maintenance. | The company spare machine to the hospital to use, the instrument returned to the company warehouse to continue maintenance.
ABL800 BASIC |  | B | SWO-92829 | No |  |  | Defect | HW | DMS_PC | Power Supply
ABL80 FLEX BASIC | 更换新测试卡修复仪器 | 更换新测试卡修复仪器 | SWO-96116 | No | Replace the new test card repair instrument | Replace the new test card repair instrument
ABL80 FLEX BASIC |  | 更换了插线636-659后修复 | SWO-92998 | No |  | Repairafter after replacing plug 636-659
ABL80 FLEX CO-OX | Gerät sendet weiterhin falsch. | Gerät sendet weiterhin falsch. | SWO-93160 | No | Device continues to send incorrectly. | Device continues to send incorrectly.
ABL90 |  | B | D09447698E59B70B | No |  |  | Defect | HW | Oxi module | Hemolyzer
ABL90 FLEX | Datei Archiv und letzte Problem-Messungen überprüft,fällige tHb Kal durchgeführt, Haken bei Ableitung Ästimierung abgeleiteter Parameter gesetzt, mehrere Testmessungen durchgeführt - alles in Ordnung, System- und Funktionskontrolle in Ordnung. | Datei Archiv und letzte Problem-Messungen überprüft,fällige tHb Kal durchgeführt, Haken bei Ableitung Ästimierung abgeleiteter Parameter gesetzt, mehrere Testmessungen durchgeführt - alles in Ordnung, System- und Funktionskontrolle in Ordnung. | SWO-93068 | No | File archive and last problem measurements checked,due tHb Kal performed, hook set when deriving aeitating derived parameters, several test measurements performed - all in order, system and function control in order. | File archive and last problem measurements checked,due tHb Kal performed, hook set when deriving aeitating derived parameters, several test measurements performed - all in order, system and function control in order. | Setup | SW | Reported derived parameters | Reported derived parameters
ABL90 | Checked all logs for errors. | Peform several back flushes. A new Sc was replaced by the user. | D094A7D28E405699 | No | Checked all logs for errors. | Peform several back flushes. A new Sc was replaced by the user. | Assisted Replacement | SC | SC | pCO2
ABL90 FLEX | Hemolyzer erneuert, Gerät kalibriert und QC durchgeführt. | Gerät signalisiert weit der Auslieferung immer wieder QC Level 2 Probleme bei der Oximetrie. Gerät mehrfach mit warmen Wasser durchgespült, Schlauchwege und Dichtungen aufgrund inhomogener Flüssigkeit erneuert. Anschließend QC erfolgreich durchgeführt. | SWO-93618 | No | Hemolyzer renewed, device calibrated and QC performed. | Device signals far the delivery repeatedly QC Level 2 problems with oximetry. Device repeatedly rinsed with warm water, hoseways and seals renewed due to inhomogeneous liquid. Subsequently QC successfully performed. | Defect | HW | Oxi module | Hemolyzer
ABL90 FLEX | Vervanging van Sensor Interface PCB Opstart, calibratie en controle : OK | Vervanging van Sensor Interface PCB Opstart, calibratie en controle : OK | SWO-93083 | No | Replacement of Sensor Interface PCB boot, calibration and control: OK | Replacement of Sensor Interface PCB boot, calibration and control: OK | Defect | HW | Basic Unit | Sensor Interface PCB | Temp (Return Request - 16.5-16)
ABL90 FLEX | Remplacé le kit de maintenance périodique e.o Nettoyage du système fluidique e.o Calibration e.o CQ e.o | Remplacé le kit de maintenance périodique e.o Nettoyage du système fluidique e.o Calibration e.o CQ e.o | SWO-93075 | No | Replaced the periodic maintenance kit e.o Fluid system cleaning e.o Calibration e.o CQ e.o | Replaced the periodic maintenance kit e.o Fluid system cleaning e.o Calibration e.o CQ e.o | Defect | HW | PM Kit | PM Kit
ABL90 FLEX |  | B | SWO-92853 | No |  |  | Unknown | Unclear | SC | SC
AQT90 FLEX | 更换了扫描器910-304和902-924后修复 | 更换了扫描器910-304和902-924后修复 | SWO-92982 | No | Replaced scanners 910-304 and 902-924 after repair | Replaced scanners 910-304 and 902-924 after repair | Defect | HW | Storage Module | Barcode reader, Storage wheel | BR1
ABL815 FLEX | Gerinnsel vor Ca entfernt, Elektrode erneuert. | Gerinnsel vor Ca entfernt, Elektrode erneuert. | SWO-93113 | No | Clots removed from Ca, electrode renewed. | Clots removed from Ca, electrode renewed. | Assisted Replacement | Electrode | Electrode | Ca
ABL800 BASIC | Fehlersuche, Laktat und Glukose wurden bereits membraniert, Messkammer, Elektroden- und Messkappenkontakte nochmals gründlich gereinigt, an Laktatelektrode fehlte der untere Dichtungsring->angebracht, Laktatelektrode und Elektrolytelektroden neu membraniert, Fluidic- und Measuringmodule Ventile überprüft ok, LS Justage und Kal.1+Kal.2 Lösung erneuert, mehrere Kal. ok System-und Funktionskontrolle i.O. | Fehlersuche, Laktat und Glukose wurden bereits membraniert, Messkammer, Elektroden- und Messkappenkontakte nochmals gründlich gereinigt, an Laktatelektrode fehlte der untere Dichtungsring->angebracht, Laktatelektrode und Elektrolytelektroden neu membraniert, Fluidic- und Measuringmodule Ventile überprüft ok, LS Justage und Kal.1+Kal.2 Lösung erneuert, mehrere Kal. ok System-und Funktionskontrolle i.O. | SWO-93009 | No | Troubleshooting, lactate and glucose have already been membraned, measuring chamber, electrode and measuring cap contacts have been thoroughly cleaned again, lactate electrode lacked the lower sealing ring->attached, lactate electrodes and electrolyte electrodes re-membrane, Fluidic and Measuring module valves checked ok, LS adjustment and Kal.1+Kal.2 solution renewed, several cal. ok System and function control i.O. | Troubleshooting, lactate and glucose have already been membraned, measuring chamber, electrode and measuring cap contacts have been thoroughly cleaned again, lactate electrode lacked the lower sealing ring->attached, lactate electrodes and electrolyte electrodes re-membrane, Fluidic and Measuring module valves checked ok, LS adjustment and Kal.1+Kal.2 solution renewed, several cal. ok System and function control i.O. | Assisted Replacement | Electrode | Electrode | Lac | Sealing ring
ABL825 FLEX | customer has been having issues with drift issues on cals Met and Electrolytes and erratic qc values As per ASM i was requested to troubleshoot operation of analyzer nstall permeant Keys on to system | Ph BG and Ei met modual disassembled and inspected Fluidic modual disassembled and inspected Solution rotor was loose when attached to drive motor replaced performed 6 month Pm tested operation | D09491558E3933A6 | No | customer has been having issues with drift issues on cals Met and Electrolytes and erratic qc values As per ASM i was requested to troubleshoot operation of analyzer nstall permeant Keys on to system | Ph BG and Ei met modual disassembled and inspected Fluidic modual disassembled and inspected Solution rotor was loose when attached to drive motor replaced performed 6 month Pm tested operation | Adjustment | HW | Fluidic manifold | Pump rotor | fixed
ABL800 BASIC | プローブ交換 | B | SWO-92952 | No | Probe exchange |  | Defect | HW | Inlet module | Probe bend | sample transfer error // probe bend
ABL825 FLEX | checked logs for errors checked system for errors ran a vk kit to verify parameters ran complete oxi test | lvl 3 on verification was on the high side but still in recomended to change membrane - after membrane change vale was right in the middle of the spectrum - unit is running within manufactures spec | D0943D278E62D469 | No | checked logs for errors checked system for errors ran a vk kit to verify parameters ran complete oxi test | lvl 3 on verification was on the high side but still in recomended to change membrane - after membrane change vale was right in the middle of the spectrum - unit is running within manufactures spec | Assisted Replacement | Electrode | Membrane | Ca
ABL835 FLEX | Lampe justeret (meget ustabil), kalibreret, slutafprøvet. | Lampe justeret (meget ustabil), kalibreret, slutafprøvet. | SWO-93092 | No | Lamp adjusted (very unstable), calibrated, final tested. | Lamp adjusted (very unstable), calibrated, final tested. | Adjustment | HW | Oxi module | Lamp unit
ABL825 FLEX | Neue Festplatte bestellt. | Konfiguration übertragen. | SWO-93002 | No | New hard drive ordered. | configuration. | Revisit required | HW | DMS_PC | HDD
ABL825 FLEX | Konfiguration übertragen. | Konfiguration übertragen. | SWO-93040 | Yes | configuration. | configuration. | Re classified RMED | Same day upSWO
AQT90 FLEX |  | B | D0944E288E1FA652 | No |  |  | Defect | HW | Storage Module | Cartridge guide | 5784
ABL80 |  | B | D095B6E60455FA39 | No
ABL90 FLEX | Einlasssonde und Dichtung erneuert. Kunde hat eine neue Einlasssonde aus dem Kreißsaal organisiert. | Einlasssonde und Dichtung erneuert. Kunde hat eine neue Einlasssonde aus dem Kreißsaal organisiert. | SWO-93229 | No | Inlet probe and seal renewed. Customer has organized a new entrance probe from the Kreißsaal. | Inlet probe and seal renewed. Customer has organized a new entrance probe from the Kreißsaal. | Defect | HW | Inlet Gasket | Inlet gasket
ABL90 | Inspected analzyer, no issues found, service diagnostics detected no errors or failures. Replaced SC lot# R0750 with same lot#, resolved pH sensitivity issue, introduced issue with pCO2 sensitivity low, out of range. | Installed second SC lot# R0381, all issues resolved. Ordered 2 replacements for customers via credit form process. | D095F12B04191801 | No | Inspected analzyer, no issues found, service diagnostics detected no errors or failures. Replaced SC lot# R0750 with same lot#, resolved pH sensitivity issue, introduced issue with pCO2 sensitivity low, out of range. | Installed second SC lot# R0381, all issues resolved. Ordered 2 replacements for customers via credit form process. | Assisted Replacement | SC | SC | pH
ABL90 FLEX | Gerät überprüft, Software Upgrade und tHb Kal durchgeführt Ca nicht relevant | Kunde hat Analysator bereits mehrfach neu gestartet ! Fehler besteht weiterhin ! Kunde wünscht Einsatz eines Technikers ! | SWO-93250 | No | Device checked, software upgrade and tHb Kal performed Ca not relevant | Customer has already restarted Analyzer several times ! Error persists ! Customer wants the use of a technician ! | Defect | SW | Windows | Windows
ABL90 | Upon arrival the analyzer was requesting the roplacement of the solution pace and inlet gasket. Found in the data logs that the system was complaing about inhomogenous rinse solution and some vaccum problems also. | Installed part number 905-899 maintenance kit and the problem went away. However, needed to replace the SC during the troubleshooting process due to PCO2 failing sensiitivity iduring calibrations> | D095A54F04585FD1 | No | Upon arrival the analyzer was requesting the roplacement of the solution pace and inlet gasket. Found in the data logs that the system was complaing about inhomogenous rinse solution and some vaccum problems also. | Installed part number 905-899 maintenance kit and the problem went away. However, needed to replace the SC during the troubleshooting process due to PCO2 failing sensiitivity iduring calibrations> | Defect | HW | PM Kit | PM Kit
ABL90 FLEX PLUS |  | B | D096A45F7A3BE60F | No |  |  | Assisted Replacement | SC | SC | Reference
ABL90 FLEX PLUS |  | B | D0959270041EDB04 | Yes |  |  | Re classified RMED | SWU
ABL90 FLEX PLUS |  | B | D095905D0454DB03 | Yes |  |  | Re classified RMED | SWU
ABL90 FLEX |  | B | SWO-93259 | No |  |  | Defect | HW | Basic Unit | Sensor Interface PCB | Temp (Return Request - 16.5-16)
AQT90 FLEX |  | B | SWO-93243 | No |  |  | Defect | HW | Process Module | Optical Unit | 135
AQT90 FLEX | 更换了903-236后修复( by andrew) | 更换了903-236后修复( by andrew) | SWO-92893 | No | Repair after replacement of 903-236 (by andrew) | Repair after replacement of 903-236 (by andrew) | Defect | HW | Process Module | Process wheel | 129
AQT90 FLEX | 更换了636-550后修复(by andrew) | 更换了636-550后修复(by andrew) | SWO-92965 | No | Repair after replacement of 636-550 (by andrew) | Repair after replacement of 636-550 (by andrew) | Defect | HW | Process Module | Process Unit sensor | 5784
AQT90 FLEX |  | B | SWO-93200 | No |  |  | Defect | HW | Incubator module | PCB, Incubator | 2168
ABL825 FLEX |  | B | D095FB5204494BC4 | No |  |  | Defect | HW | AutoCheck Module | AutoCheck Module
ABL825 FLEX | 更换LAC电极膜正常。 | 更换LAC电极膜正常。 | SWO-93536 | No | Replacing the LAC electrode film is normal. | Replacing the LAC electrode film is normal. | Assisted Replacement | Electrode | Membrane | Lac | Calibration
ABL835 FLEX | Lacメンブラン内気泡除去、EL Metモジュールのフラットメンブランバルブ整備。PHシーリングリング交換。 | Lacメンブラン内気泡除去、EL Metモジュールのフラットメンブランバルブ整備。PHシーリングリング交換。 | SWO-93438 | No | Lac membrane bubble removal and flat membrane valve maintenance of EL Met module. PH sealing ring replacement. | Lac membrane bubble removal and flat membrane valve maintenance of EL Met module. PH sealing ring replacement. | Assisted Replacement | Electrode | Membrane | Lac | Response error
ABL810 FLEX | Elektrode erneuert und mehrfach kalibriert. | Elektrode erneuert und mehrfach kalibriert. | SWO-93294 | No | electrode renewed and calibrated several times. | electrode renewed and calibrated several times. | Assisted Replacement | Electrode | Electrode | pCO2
ABL825 FLEX |  | B | SWO-93220 | No |  |  | Defect | HW | Oxi module | Oxi module | 316
ABL817 FLEX | Verified. | Replaced the Inlet Gasket 902-668. Also, the Crea membrane replace is due. Replaced the A and B membranes from the customers stock as a courtesy. No leaks detected after all calibrations and QC's were performed. | D095D56104375FD6 | No | Verified. | Replaced the Inlet Gasket 902-668. Also, the Crea membrane replace is due. Replaced the A and B membranes from the customers stock as a courtesy. No leaks detected after all calibrations and QC's were performed. | Defect | HW | Inlet module | Inlet gasket
TCM-COMBIM | After 12 days the unit would not boot up. | Replaced the CPU pcb and the CF Card. | D095678B044BEEDD | No | After 12 days the unit would not boot up. | Replaced the CPU pcb and the CF Card. | Defect | HW | PCB, Main CPU | PCB, Main CPU
TCM/Tosca | Removed unit and returned to the workshop for repair. Upon further investigation it was the fuse that had blown. Replaced fuse and monitor is now functioning. Returned the monitor to customer and tested. | Service call required. | SWO-93323 | No | Removed unit and returned to the workshop for repair. Upon further investigation it was the fuse that had blown. Replaced fuse and monitor is now functioning. Returned the monitor to customer and tested. | Service call required. | Defect | HW | Power Supply | Power Supply | Fuse
ABL80 FLEX BASIC | 更换新测试卡修复仪器 | 更换新测试卡修复仪器 | SWO-93525 | No | Replace the new test card repair instrument | Replace the new test card repair instrument
ABL80 FLEX BASIC | 更换552-039一个正常。(by zhenjie) | 更换552-039一个正常。(by zhenjie) | SWO-92929 | No | Replace 552-039 one normal. (by zhenjie) | Replace 552-039 one normal. (by zhenjie)
ABL1 | Replaced Sensor connector, zeroing done, installed new SC. | Replaced Sensor connector, zeroing done, installed new SC. | SWO-94320 | No | Replaced Sensor connector, zeroing done, installed new SC. | Replaced Sensor connector, zeroing done, installed new SC.
ABL90 |  | B | D096CECB7A5D8D0F | No |  |  | Defect | HW | Oxi module | Hemolyzer
ABL90 |  | B | D0971CFD7A51A866 | No |  |  | Assisted Replacement | SP | SP | Glu/Lac
ABL90 FLEX PLUS | Kalibrierintervall 15min verschoben und QC Plan mit Kal verknüpft, ein Gerät ist somit immer verfügbar, Netzwerkdosen nicht verfügbar müssen erst bereitgestellt werden, Frage ob beide Geräte an einer Dose via Hub Switch betrieben werden können - ja | Kalibrierintervall 15min verschoben und QC Plan mit Kal verknüpft, ein Gerät ist somit immer verfügbar, Netzwerkdosen nicht verfügbar müssen erst bereitgestellt werden, Frage ob beide Geräte an einer Dose via Hub Switch betrieben werden können - ja | SWO-93959 | No | Calibration interval 15min shifted and QC plan linked to Kal, a device is always available, network boxes not available must first be provided, question whether both devices can be operated on a can via hub switch - yes | Calibration interval 15min shifted and QC plan linked to Kal, a device is always available, network boxes not available must first be provided, question whether both devices can be operated on a can via hub switch - yes | Setup | Cal Schedule | Cal Schedule | Cal Schedule
ABL90 FLEX PLUS |  | B | D0970C2C7A4AD5F5 | No |  |  | Assisted Replacement | SP | SP | FT
AQT90 FLEX |  | B | SWO-93439 | No |  |  | Defect | HW | Aspiration Module | Needle Wash Unit | 1267
AQT90 FLEX |  | B | SWO-93503 | No |  |  | Assisted Replacement | Consumables | Test Kit | CKMB CAL | 124
ABL837 FLEX | Customer advised Hct reading giving error 791: Parameter could not be measured. | FSE found no issue running QC, blood sample run and Hct detected properly. FSE suspects machine needed time to condition after adding new parameter. | D096AFED7A3FADEC | No | Customer advised Hct reading giving error 791: Parameter could not be measured. | FSE found no issue running QC, blood sample run and Hct detected properly. FSE suspects machine needed time to condition after adding new parameter. | No fault found | No fault found | No fault found | No fault found
ABL825 FLEX | Customer tried to reboot | Replaced OXI PCB. Adjusted lamp. Ran tHB cal in service and customer modes. Ran linearity for customer | D096D5AF7A124236 | No | Customer tried to reboot | Replaced OXI PCB. Adjusted lamp. Ran tHB cal in service and customer modes. Ran linearity for customer | Defect | HW | Oxi module | PCB, Oxi
ABL825 FLEX | 比較電極交換にて様子を見ます。 病院の比較電極は社内でチェック。 | 比較電極交換にて様子を見ます。 病院の比較電極は社内でチェック。 | SWO-93751 | No | I see the appearance by the comparison electrode exchange. Check the comparison electrodes of the hospital in-house. | I see the appearance by the comparison electrode exchange. Check the comparison electrodes of the hospital in-house. | Assisted Replacement | Electrode | Electrode | pH
ABL827 FLEX | Daten gesichert, Software V6.15 neu aufgespielt, Daten zurück gespielt, Gerät kalibriert, QC´s durchgef. | in der Nacht war ein Stromausfall, seitdem besteht das Problem; der Anwender wünscht einen Vorort-Service zur Beseitigung der Geräte-Störung; | SWO-93842 | No | Data backed up, software V6.15 replayed, data played back, device calibrated, QC's through. | at night was a power outage, since then the problem exists; the user wants a suburban service to eliminate the device malfunction; | Defect | SW | Windows | Off/On
ABL837 FLEX | Spildslangen sat på plads under Auto QC modulet | Spildslangen sat på plads under Auto QC modulet | SWO-93505 | No | The waste hose put in place under the Auto QC module | The waste hose put in place under the Auto QC module | Adjustment | HW | Basic unit | Waste tube
ABL825 FLEX | Kunden Oxi Modul nach erfolgter Reparatur zurückgetauscht. Nach Neustart ist das FlexQ Modul wieder aktiv. tHb Kal. | Kunden Oxi Modul nach erfolgter Reparatur zurückgetauscht. Nach Neustart ist das FlexQ Modul wieder aktiv. tHb Kal. | SWO-93693 | No | Customer Oxi module exchanged after repair. After rebooting, the FlexQ module is active again. tHb Kal. | Customer Oxi module exchanged after repair. After rebooting, the FlexQ module is active again. tHb Kal. | Defect | HW | FlexQ Module | FlexQ Module
ABL80 |  | B | D096D3527A401925 | No
ABL80 FLEX BASIC | 更换测试卡修复仪器 | 更换测试卡修复仪器 | SWO-94400 | No | Replace the test card repair instrument | Replace the test card repair instrument
ABL80 CO-OX | will replace CPU, barcode reader, and cable | Replaced CPU, barcode reader, and cable. All working normal now. | D0988604F01DE843 | No | will replace CPU, barcode reader, and cable | Replaced CPU, barcode reader, and cable. All working normal now.
ABL80 FLEX BASIC | ＳＣ交換実施にて改善。 | B | SWO-94069 | No | Improved by SC exchange.
ABL90 | Verified error, unable to reset or reinitialize battery. | Replaced battery. | D0983C55F0363906 | No | Verified error, unable to reset or reinitialize battery. | Replaced battery. | Defect | HW | Battery module | Battery
ABL90 FLEX | 更换新测试卡修复仪器 | 更换新测试卡修复仪器 | SWO-93862 | No | Replace the new test card repair instrument | Replace the new test card repair instrument
ABL800 BASIC | Einlassnippel ersetzt, Gerät gereinigt, Luftkanal gespült, Kal und QC durchgeführt | Einlassnippel ersetzt, Gerät gereinigt, Luftkanal gespült, Kal und QC durchgeführt | SWO-94009 | No | Replacement inlet nipples, device cleaned, duct flushed, Kal and QC performed | Replacement inlet nipples, device cleaned, duct flushed, Kal and QC performed | Defect | HW | Inlet module | Leaking | Nipple kit
ABL825 FLEX | Ph bg and Ei Met moduals have been disassembled all valves fluidic blocks and measuring chambers have been ispected. Fluidic modual has been dissasembled all valves have been inspected and tested 6 month pm was performed Tested voltages on power supply passed no issues with Au connectors grounding pads and groung connections measured no consernable ohm readings | Found issue with Solution rotor replaced Ph oring replaced Ref electrode and new lot 1917 of ref membranes installed unit had performed well over night and will be monitored for operatio | D0982701F012AE19 | No | Ph bg and Ei Met moduals have been disassembled all valves fluidic blocks and measuring chambers have been ispected. Fluidic modual has been dissasembled all valves have been inspected and tested 6 month pm was performed Tested voltages on power supply passed no issues with Au connectors grounding pads and groung connections measured no consernable ohm readings | Found issue with Solution rotor replaced Ph oring replaced Ref electrode and new lot 1917 of ref membranes installed unit had performed well over night and will be monitored for operatio | Assisted Replacement | Electrode | Membrane | Reference
ABL835 FLEX | N A | Gave customer part number for tubing. | D0982C04F013AB3B | Yes | N A | Gave customer part number for tubing. | Re classified RMED
ABL837 FLEX |  | B | D09807FBF020192F | Yes |  |  | Re classified RMED | Training of customer
ABL800 BASIC | 更换4个LAC电极膜正常。 | 更换4个LAC电极膜正常。 | SWO-94113 | No | Replacing 4 LAC electrode membranes is normal. | Replacing 4 LAC electrode membranes is normal. | Assisted Replacement | Electrode | Membrane | Lac | Calibration
ABL825 FLEX | neue Elektrode eingesetzt - starke Schwankungen- ElMet Elektroden komplett membraniert, Modulpumpenschläuche ersetzt, Neustart, mehrere Kal, QC | neue Elektrode eingesetzt - starke Schwankungen- ElMet Elektroden komplett membraniert, Modulpumpenschläuche ersetzt, Neustart, mehrere Kal, QC | SWO-94093 | No | new electrode inserted - strong fluctuations- ElMet electrodes completely membraneed, module pump hoses replaced, restart, several Kal, QC | new electrode inserted - strong fluctuations- ElMet electrodes completely membraneed, module pump hoses replaced, restart, several Kal, QC | Assisted Replacement | Electrode | Membrane | Lac | Sensitivity
AQT90 FLEX | Inlet insert parts re-assembled, tested as ok. Inlet insert refitted to analyser. | Inlet insert parts re-assembled, tested as ok. Inlet insert refitted to analyser. | SWO-94040 | No | Inlet insert parts re-assembled, tested as ok. Inlet insert refitted to analyser. | Inlet insert parts re-assembled, tested as ok. Inlet insert refitted to analyser. | Defect | HW | Aspiration Module | Tube Holder
AQT90 FLEX | Found inlet door was not opening, instrument was showing error, found removed the door assemble and replaced the drive belt for inlet door, restarted and checked in the service program but problem was remaining same, need to replace with new inlet door assemble and check awaiting for spare part. | Found inlet door was not opening, instrument was showing error, found removed the door assemble and replaced the drive belt for inlet door, restarted and checked in the service program but problem was remaining same, need to replace with new inlet door assemble and check awaiting for spare part. | SWO-94316 | No | Found inlet door was not opening, instrument was showing error, found removed the door assemble and replaced the drive belt for inlet door, restarted and checked in the service program but problem was remaining same, need to replace with new inlet door assemble and check awaiting for spare part. | Found inlet door was not opening, instrument was showing error, found removed the door assemble and replaced the drive belt for inlet door, restarted and checked in the service program but problem was remaining same, need to replace with new inlet door assemble and check awaiting for spare part. | Defect | HW | Cover and Chassis | Inlet Door
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-94163 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL80 CO-OX | Error occurring mostly with level 3 external control. | Checked line for obstructions. Performed co-ox test and liquid sensor adjust in diagnostics. | D0995602665A5EEA | No | Error occurring mostly with level 3 external control. | Checked line for obstructions. Performed co-ox test and liquid sensor adjust in diagnostics.
ABL90 |  | B | D0995965663B50C9 | No |  |  | Defect | HW | CF card | CF card
ABL90 FLEX | Flushed unit swapped SC | Flushed unit swapped SC | SWO-94296 | No | Flushed unit swapped SC | Flushed unit swapped SC | Assisted Replacement | SC | SC | pO2
ABL90 FLEX | Replaced SP flushed unit but still failing QC 3 . Replaced SC ran calibration all passed ok ran all QC passed ok | Replaced SP flushed unit but still failing QC 3 . Replaced SC ran calibration all passed ok ran all QC passed ok | SWO-94282 | No | Replaced SP flushed unit but still failing QC 3 . Replaced SC ran calibration all passed ok ran all QC passed ok | Replaced SP flushed unit but still failing QC 3 . Replaced SC ran calibration all passed ok ran all QC passed ok | Assisted Replacement | SC | SC | pO2
ABL90 FLEX | Flushed unit ran test all passed ok. Still failing QC replaced SP calibration and QC ok | Flushed unit ran test all passed ok. Still failing QC replaced SP calibration and QC ok | SWO-94290 | Yes | Flushed unit ran test all passed ok. Still failing QC replaced SP calibration and QC ok | Flushed unit ran test all passed ok. Still failing QC replaced SP calibration and QC ok | Re classified RMED | SP | pO2 | pO2
ABL90 |  | B | D0997F9E66301937 | No |  |  | Defect | HW | Inlet Gasket | Inlet gasket
ABL90 FLEX | Lösungspack ausgetauscht. Einlassnadel und Einlassdichtung gespült. Sensorkasette ausgetauscht. Kalibriert und QC's durchgeführt. | Lösungspack ausgetauscht. Einlassnadel und Einlassdichtung gespült. Sensorkasette ausgetauscht. Kalibriert und QC's durchgeführt. | SWO-94268 | No | solution pack. Inlet needle and inlet seal rinsed. Sensor catheter replaced. Calibrated and QC's performed. | solution pack. Inlet needle and inlet seal rinsed. Sensor catheter replaced. Calibrated and QC's performed. | Assisted Replacement | SC | SC | pCO2
ABL90 FLEX | 更换新测试卡修复仪器 | 更换新测试卡修复仪器 | SWO-94478 | No | Replace the new test card repair instrument | Replace the new test card repair instrument
ABL90 FLEX | 更换测试卡正常。 | 更换测试卡正常。 | SWO-94126 | No | The replacement test card is OK. | The replacement test card is OK. | Assisted Replacement | SC | SC | pO2
ABL820 FLEX | 需要借配件 printer做更換 | B | SWO-94153 | No | Need to borrow accessories printer for replacement |  | Defect | HW | Printer module | Printer head
ABL800 BASIC | presently cleaned the chamber and replaced the potassium membrane it need to replace with new measuring chamber . | presently cleaned the chamber and replaced the potassium membrane it need to replace with new measuring chamber . | SWO-94355 | No | presently cleaned the chamber and replaced the potassium membrane it need to replace with new measuring chamber . | presently cleaned the chamber and replaced the potassium membrane it need to replace with new measuring chamber . | Cleaned | Fluidic pathway | El/Met Module | K+
ABL825 FLEX | Installation Option Keys lt. Vertrag | Installation Option Keys lt. Vertrag | SWO-94256 | No | Installation Option Keys according to Contract | Installation Option Keys according to Contract | Setup | SW | Demo Key Exp | Demo Key Exp
ABL837 FLEX |  | B | D09D326BC81E81ED | Yes |  |  | Re classified RMED | SWU
ABL825 FLEX | Oxi Modul erneuert. Service tHb und Anwender tHb Kal durchgeführt. QC wiederholt. | Oxi Modul erneuert. Service tHb und Anwender tHb Kal durchgeführt. QC wiederholt. | SWO-94293 | No | Oxi module renewed. service tHb and user tHb Kal. QC. | Oxi module renewed. service tHb and user tHb Kal. QC. | Defect | HW | Oxi module | Oxi module
ABL815 FLEX | Datensicherung, Applikationssoftware v6.16 deinstalliert und alte Dateien gelöscht, Neuinstallation der Applikationssoftware v6.16, Datenrücksicherung, Probelauf, Patientenbericht (patres.rpt) aktualisiert | Datensicherung, Applikationssoftware v6.16 deinstalliert und alte Dateien gelöscht, Neuinstallation der Applikationssoftware v6.16, Datenrücksicherung, Probelauf, Patientenbericht (patres.rpt) aktualisiert | SWO-94205 | No | Data backup, application software v6.16 uninstalled and old files deleted, reinstallation of application software v6.16, data backup, trial run, patient report (patres.rpt) updated | Data backup, application software v6.16 uninstalled and old files deleted, reinstallation of application software v6.16, data backup, trial run, patient report (patres.rpt) updated | Defect | SW | Printer module | Print out
ABL810 FLEX | Spannungswert des Gas1 Ventils zu niedrig, Gas2 Ventil o. K. Evtl. Tausch des Gasventils und des Gasdruck-Minderers. Tausch der pCO2 Elektrode. Qualitätskontroll-Wert bleibt außer Grenzwert. Eigene pCO2 Elektrode wieder eingelegt. Gasventil1 und Gasdruckminderer bestellt. Kalibrationen und Qualitätskontrolle o. K. auch bei pCO2, vmtl. nut temporär. | Spannungswert des Gas1 Ventils zu niedrig, Gas2 Ventil o. K. Evtl. Tausch des Gasventils und des Gasdruck-Minderers. Tausch der pCO2 Elektrode. Qualitätskontroll-Wert bleibt außer Grenzwert. Eigene pCO2 Elektrode wieder eingelegt. Gasventil1 und Gasdruckminderer bestellt. Kalibrationen und Qualitätskontrolle o. K. auch bei pCO2, vmtl. nut temporär. | SWO-94263 | No | Voltage value of the Gas1 valve too low, Gas2 valve o.K. Possible replacement of the gas valve and the gas pressure reducer. Replacement of the pCO2 electrode. Quality control value remains out of value. Own pCO2 electrode inserted again. gas valve1 and gas pressure reducer. Calibrations and quality control o. K. also at pCO2, vmtl. nut temporary. | Voltage value of the Gas1 valve too low, Gas2 valve o.K. Possible replacement of the gas valve and the gas pressure reducer. Replacement of the pCO2 electrode. Quality control value remains out of value. Own pCO2 electrode inserted again. gas valve1 and gas pressure reducer. Calibrations and quality control o. K. also at pCO2, vmtl. nut temporary. | Assisted Replacement | Electrode | Electrode | pCO2
ABL810 FLEX | Tausch des Autocheck Moduls notwendig, bestellt. Analysator messbereit. | Tausch des Autocheck Moduls notwendig, bestellt. Analysator messbereit. | SWO-94264 | No | Replacement of the Autocheck module necessary, ordered. Analyzer ready to measure. | Replacement of the Autocheck module necessary, ordered. Analyzer ready to measure. | Revisit required | HW | AutoCheck Module | AutoCheck Module
ABL80 FLEX BASIC | 更换测试卡正常。 | 更换测试卡正常。 | SWO-95739 | No | The replacement test card is OK. | The replacement test card is OK.
ABL80 FLEX BASIC | 更换新测试卡修复仪器 | 更换新测试卡修复仪器 | SWO-94365 | No | Replace the new test card repair instrument | Replace the new test card repair instrument
ABL90 FLEX | 更换903-183修复仪器 | 更换903-183修复仪器 | SWO-93796 | No | Replace the 903-183 repair instrument | Replace the 903-183 repair instrument | Defect | HW | Basic Unit | Sensor Interface PCB | (Return Request - 16.5-16)
AQT90 FLEX | 更换配件903-329和636-561修复仪器 | 更换配件903-329和636-561修复仪器 | SWO-94338 | No | Replacement accessories 903-329 and 636-561 repair instruments | Replacement accessories 903-329 and 636-561 repair instruments | Defect | HW | Aspiration Module | Needle Wash Unit | 3415
ABL90 FLEX | 手动掰直进样针 | 手动掰直进样针 | SWO-94366 | No | Manual straight into the sample needle | Manual straight into the sample needle | Unknown | Unknown | Unknown | Unknown
AQT90 FLEX | 更换903-210修复仪器 | 更换903-210修复仪器 | SWO-94088 | No | Replace the 903-210 repair instrument | Replace the 903-210 repair instrument | Defect | HW | Process Module | PCB, Cup dryer | 168
ABL80 FLEX BASIC | 更换测试卡正常。 | 更换测试卡正常。 | SWO-95719 | No | The replacement test card is OK. | The replacement test card is OK.
ABL80 FLEX BASIC | 更换测试卡正常。 | 更换测试卡正常。 | SWO-96314 | No | The replacement test card is OK. | The replacement test card is OK.
ABL80 FLEX CO-OX | Kein Fehler feststellbar. Gesamte Verschlauchung gereinigt.Wahrscheinlich handelt es sich um ein Bedienungsfehler. | Kein Fehler feststellbar. Gesamte Verschlauchung gereinigt.Wahrscheinlich handelt es sich um ein Bedienungsfehler. | SWO-94437 | No | No error detectable. Total hose cleaning. It is probably an operating error. | No error detectable. Total hose cleaning. It is probably an operating error.
ABL80 CO-OX | Two SC had previously been installed. Customer replaced SP from -R- instructions. | Removed reagents and cleaned all tubing. Re-installed SC and SP, system cycle checking OK. Patient sample checks OK. | D09D5D87C8117DF3 | No | Two SC had previously been installed. Customer replaced SP from -R- instructions. | Removed reagents and cleaned all tubing. Re-installed SC and SP, system cycle checking OK. Patient sample checks OK.
ABL90 FLEX | Tausch des Displays. Nach Neustart ist Funktion des Bildschirms wieder o. K. Reaktion wieder Punktgenau. Kalibrationen und Qualitätskontrolle o. K. | Tausch des Displays. Nach Neustart ist Funktion des Bildschirms wieder o. K. Reaktion wieder Punktgenau. Kalibrationen und Qualitätskontrolle o. K. | SWO-94633 | No | Replacing the display. After rebooting, the function of the screen is again o.K. reaction again point-accurate. Calibrations and quality control o.K. | Replacing the display. After rebooting, the function of the screen is again o.K. reaction again point-accurate. Calibrations and quality control o.K. | Defect | HW | DMS_PC | Touch screen
ABL90 FLEX PLUS | Replaced inlet gasket with holder. | Replaced inlet gasket with holder. | SWO-94652 | No | Replaced inlet gasket with holder. | Replaced inlet gasket with holder. | Defect | HW | Auto Inlet module | Inlet Gasket
ABL90 FLEX PLUS |  | B | D09E2BF73E17F778 | No |  |  | Adjustment | HW | Oxi module | tHb Cal | Please help me by clarifying the following : 1. According to the sold date, this analyzer is only 6 days old at upSWO visit. 2. Why do a new installation like this drive upSWO. a. Isn’t tHb Calibration a part of the installation procedure?
AQT90 FLEX | 更换636-550修复仪器(by zhenjie) | 更换636-550修复仪器(by zhenjie) | SWO-94497 | No | Replace the 636-550 repair instrument (by zhenjie) | Replace the 636-550 repair instrument (by zhenjie) | Defect | HW | Process Module | Process Unit sensor | 5784
ABL835 FLEX | Limpieza del sistema completo de la unidad cambio de la valvula de solución del modulo de PH BG así como sustitución de las juntas de las valvulas de aire y valvulas de soluciónes de lavado salida de aire y salida de desperdicios del modulo de PH BG y modulo de fluidica, calibración y test controles de calidad funcionanndo correctamente. | Cambio de valvula de solución de lavado modulo PH BG | D09D2388C84B1C57 | No | Cleaning the entire unit system by changing the solution valve of the PH BG module as well as replacing the joints of the air valves and valves of air outlet washing and waste outlet solutions of the PH BG module and fluidic module , calibration and test quality controls operating properly. | Change of valve washing solution module PH BG | Defect | HW | pH/BG Module | Valve
ABL835 FLEX |  | B | D09D4BB4C846284F | Yes |  |  | Re classified RMED | PM
ABL835 FLEX | Gerät überprüft, Meßkammer und Luftauslass gesäubert, E777 (RV02) erneuert, Kal und QC o.k. | Gerät überprüft, Meßkammer und Luftauslass gesäubert, E777 (RV02) erneuert, Kal und QC o.k. | SWO-95611 | No | Device checked, measuring chamber and air outlet cleaned, E777 (RV02) renewed, Kal and QC ok | Device checked, measuring chamber and air outlet cleaned, E777 (RV02) renewed, Kal and QC ok | Assisted Replacement | Electrode | Electrode | pH
ABL800 BASIC | 更换945-619修复仪器(by zhenjie) | 更换945-619修复仪器(by zhenjie) | SWO-94330 | No | Replace the 945-619 repair instrument (by zhenjie) | Replace the 945-619 repair instrument (by zhenjie) | Assisted Replacement | Electrode | Electrode | Lac | Calibration
ABL825 FLEX | Messkammer und Schlauchsystem geprüft, Modulventile V1-V4 demontiert und gereinigt, Nullpunkt Abgleich durchgeführt, Lac. und Ca. Elektrode neu membraniert, Funktionskontrolle ok | B | SWO-94532 | No | Measuring chamber and hose system tested, module valves V1-V4 dismantled and cleaned, zero point matching performed, Lac. and approx. electrode re-membraneed, function control ok |  | Cleaned | Fluidic pathway | Fluidic manifold | Lac
ABL825 FLEX | Messkammer und Schlauchsystem geprüft, Modulventile V1-V4 demontiert und gereinigt, Nullpunkt Abgleich durchgeführt, Lactat Elektrode neu membraniert (Lot 3002), Funktionskontrolle ok | B | SWO-94533 | No | Measuring chamber and hose system tested, module valves V1-V4 dismantled and cleaned, zero point matching performed, lactate electrode re-membraneed (lot 3002), function control ok |  | Cleaned | Fluidic pathway | Fluidic manifold | Lac
ABL800 BASIC | 更换电极膜修复仪器 | 更换电极膜修复仪器 | SWO-94509 | No | Replace electrode membrane repair instrument | Replace electrode membrane repair instrument | Assisted Replacement | Electrode | Membrane | Lac | Calibration
ABL820 FLEX | bad inlet preheater, and pump tubes are flat. | Replaced inlet preheater and all pump tubes | D09D1985C813E853 | No | bad inlet preheater, and pump tubes are flat. | Replaced inlet preheater and all pump tubes | Defect | HW | Inlet module | Preheater
ABL80 FLEX BASIC | Kept SC in fridge and then reinstall. CL was inactive. Done active parameters and install again, Now unit is working fine. | Kept SC in fridge and then reinstall. CL was inactive. Done active parameters and install again, Now unit is working fine. | SWO-96431 | No | Kept SC in fridge and then reinstall. CL was inactive. Done active parameters and install again, Now unit is working fine. | Kept SC in fridge and then reinstall. CL was inactive. Done active parameters and install again, Now unit is working fine.
ABL80 FLEX BASIC | 更换测试卡修复仪器 | 更换测试卡修复仪器 | SWO-95442 | No | Replace the test card repair instrument | Replace the test card repair instrument
ABL80 FLEX BASIC |  | 更换测试卡修复仪器 | SWO-95443 | No |  | Replace the test card repair instrument
ABL80 FLEX BASIC |  | B | SWO-94681 | No
ABL90 FLEX | Gerät überprüft, JW durchgeführt, Meßkassette ausgetauscht | Gerät überprüft, JW durchgeführt, Meßkassette ausgetauscht | SWO-95615 | No | Device checked, JW performed, measuring cassette replaced | Device checked, JW performed, measuring cassette replaced | Assisted Replacement | SC | SC | Lac
ABL90 | via phone | shipped sc and sp from my stock via air canada same day. | D09E7B7B3E1E6082 | Yes | via phone | shipped sc and sp from my stock via air canada same day. | Re classified RMED | support
ABL90 FLEX | Fehlersuche -> Schlauch zwischen MK u. Hemolyzer def., Gerät aus, Schlauch gew., Gerät ein, Kal. i.O. | Fehlersuche -> Schlauch zwischen MK u. Hemolyzer def.; Gerät aus; Schlauch gew.; Gerät ein; Kal. i.O. | SWO-95532 | No | Troubleshooting -> Hose between MK and Hemolyzer def.; device off; Hose. device; Kal. Ok. | Troubleshooting -> Hose between MK and Hemolyzer def.; device off; Hose. device; Kal. Ok. | Defect | HW | Oxi module | Pinch Valve tube
ABL90 FLEX | Gerät in betrieb genommen, SC 90 und LP 90 ersetzt, Konfiguration eingegeben, Datenverbindung zum Netzwerk hergestellt, Funktionskontrolle ok | B | SWO-95752 | Yes | Device put into operation, SC 90 and LP 90 replaced, configuration entered, data connection to the network established, function control ok |  | Re classified RMED | Re location / Installation
ABL90 | Glucose low out of limits on S7050 QC. Results dropped rapidly from 250 0n 4 11 to 208 today. Unit going through 20 day QC study. | Tried running blood samples to see if sensor would rebound. It did not. Swapped sensors with a unit that does not have glucose enabled. | D09E54C93E0D1410 | No | Glucose low out of limits on S7050 QC. Results dropped rapidly from 250 0n 4 11 to 208 today. Unit going through 20 day QC study. | Tried running blood samples to see if sensor would rebound. It did not. Swapped sensors with a unit that does not have glucose enabled. | Assisted Replacement | SC | SC | Glu
ABL90 FLEX PLUS | Inletmodule vervangen. | Flatcable vervangen tussen Inlet detector PCB en Wet Section Control PCB. | SWO-95529 | No | Replace Inletmodule. | Flatcable vervangen tussen Inlet detector PCB en Wet Section Control PCB. | Defect | HW | Auto Inlet module | Auto Inlet module | being returned / Luc
TCM4 |  | B | SWO-95518 | No
AQT90 FLEX |  | B | SWO-94663 | No |  |  | Defect | HW | Process Module | Optical Unit | 131
ABL825 FLEX | PC Modul ausgebaut CPU Unit erneuert (neue MAC Adresse: 00-04-5F-90-15-96), Modul eingebaut Gerät überprüft | PC Modul ausgebaut CPU Unit erneuert (neue MAC Adresse: 00-04-5F-90-15-96), Modul eingebaut Gerät überprüft | SWO-95501 | No | PC module removed CPU unit renewed (new MAC address: 00-04-5F-90-15-96), module built-in device checked | PC module removed CPU unit renewed (new MAC address: 00-04-5F-90-15-96), module built-in device checked | Defect | HW | DMS_PC | CPU
ABL825 FLEX |  | B | D0A0EF6D2A53C729 | No |  |  | Defect | HW | Inlet module | Leak
ABL837 FLEX | チャンバー内洗浄で改善しなかったため再度Refメンブラン交換 | B | SWO-97065 | No | Ref membrane replacement again because it did not improve by cleaning in the chamber |  | Assisted Replacement | Electrode | Membrane | pO2
ABL825 FLEX | 更换PH电极修复仪器（by zhenjie) | 更换PH电极修复仪器（by zhenjie) | SWO-96025 | No | Replace the PH electrode repair instrument (by zhenjie) | Replace the PH electrode repair instrument (by zhenjie) | Assisted Replacement | Electrode | Electrode | pH
AQT90 FLEX |  | B | D09E92293E313691 | No |  |  | Defect | HW | Analyzer Control Module | CPU
ABL80 FLEX Basic version | Passagem do QC com valor baixo no nivel 4, substituição do pack e foi gerando um erro de C V, substituição da SC, establização e passagem do QC. | B | D09FD005B4243CFB | No | QC passage with low value at level 4, replacement of the pack and was generating a C V error, replacement of SC, establishment and passage of QC.
ABL80 FLEX BASIC | 更换测试卡正常。 | 更换测试卡正常。 | SWO-95740 | No | The replacement test card is OK. | The replacement test card is OK.
ABL80 CO-OX | Found fluid leaking from sample port of sensor cassette indicating there was a blockage in waste tubing and might be pinched. | Adjusted waste tubing. Verified instrument operation. | D0A0A6ED2A0DDB87 | No | Found fluid leaking from sample port of sensor cassette indicating there was a blockage in waste tubing and might be pinched. | Adjusted waste tubing. Verified instrument operation.
ABL80 FLEX BASIC | 更换新测试卡修复仪器 | 更换新测试卡修复仪器 | SWO-96117 | No | Replace the new test card repair instrument | Replace the new test card repair instrument
ABL90 | Error 1277, fluid transport issues. Ran service checks. Error 1179,13501, 1351, and 1352. | Replaced inlet gasket, connector gasket and pump rotor. | D0A5F2AE02123E59 | No | Error 1277, fluid transport issues. Ran service checks. Error 1179,13501, 1351, and 1352. | Replaced inlet gasket, connector gasket and pump rotor. | Defect | HW | Pump Rotor | Pump Rotor
ABL90 |  | B | D09FD886B44260BE | No |  |  | Adjustment | HW | Sensor Flow selector | Sensor Flow selector
ABL90 FLEX PLUS | Einlassdichtungshalter erneuert und SP erneut installiert | Einlassdichtungshalter erneuert und SP erneut installiert | SWO-95884 | No | Inlet seal holder renewed and SP reinstalled | Inlet seal holder renewed and SP reinstalled | Defect | HW | Auto Inlet module | Inlet Gasket
TCM CombiM |  | B | SWO-95704 | No
TCM CombiM | 須借貨 903-365 Main CPU PCB 做更換 | B | SWO-95727 | No | Borrow 903-365 Main CPU PCB for replacement |  | Defect | HW | PCB, Main CPU | PCB, Main CPU
TCM4 |  | B | D09FA7DBB4474BB5 | No |  |  | Defect | HW | Power Supply | Power Supply
ABL90 FLEX | 更换测试卡正常。 | 更换测试卡正常。 | SWO-95742 | No | The replacement test card is OK. | The replacement test card is OK. | Assisted Replacement | SC | SC | Reference
ABL90 FLEX |  | B | SWO-95705 | No |  |  | Defect | HW | DMS_PC | Barcode reader
ABL825 FLEX |  | B | D09F7E48B40A4C9D | No |  |  | Defect | HW | Fluidic manifold | Pump tubes
ABL820 FLEX | Analysator geöffnet, alle Stecker von der AU gezogen, keine Änderung, ABC LEDs leuchten durchgehend rot. AU Software Installation kommt nicht über den Punkt waiting for Bootstrap Code hinaus. Analysator zur Reparatur abgezogen. | Analysator geöffnet, alle Stecker von der AU gezogen, keine Änderung, ABC LEDs leuchten durchgehend rot. AU Software Installation kommt nicht über den Punkt waiting for Bootstrap Code hinaus. Analysator zur Reparatur abgezogen. | SWO-96237 | No | Analyzer open, all plugs pulled from the AU, no change, ABC LEDs glow throughout red. AU Software installation does not go beyond waiting for bootstrap code. analyzer for repair. | Analyzer open, all plugs pulled from the AU, no change, ABC LEDs glow throughout red. AU Software installation does not go beyond waiting for bootstrap code. analyzer for repair. | Defect | HW | Basic unit | Analyzer Control PCB
ABL820 FLEX | found that the ABLUser had been locked. | rebooted the analyzer which corrected that issue. | D09F8ED5B4088421 | No | found that the ABLUser had been locked. | rebooted the analyzer which corrected that issue. | Defect | SW | DMS_PC | Windows
ABL810 FLEX | Tausch des Gas Moduls 1. Reinigung der PCO2 Messkammer und entfernen der Luftbläschen aus dem Elektrolyt. pCO2 Werte bei allen Analysatoren in Mosbach und Buchen grenzwertig zum unteren Grenzwert. Daten gezogen zur Auswertung und Eröffnung einer Techline. Qualitätskontrolle fehlerhaft. Messkammer undicht. Gereinigt und unterer Elektrodendichtring bei der pH Elektrode erneuert. Weiterhin undichtigkeit, unterer Elektrodendichtring der Dumy Electrode in Cl Messkammer erneuert. Weiterhin Undichtigkeit, LS Justage und Pumpen Kal., kein Erfolg. LS Fehler, unterer LS Sensor in pH BG Modul defekt, erneuert. Einlassdichtung erneuert. Wieterhin Qualitätskontroll Grenzverletzung bei Kevel rot. Kalibration mit neuer Membran Charge, kein Erfolg. Tausch pCO2 Elektrode. Kalibrationen o. K., und Qualitätskontrolle bei pCO2 nicht o. K. | Tausch des pH BG Moduls. LS Justage, Pumpenkal und Nullpunktkalibrierung durchgeführt. Einkalibriert, Qualitätskontroll Werte zielwertig. Service Dump gezogen zur Untersuchung im Technischen Support. Neustart des Analysators, anschließend bleibt Analysator im Laden der Anwendersoftware stehen mit Meldung; Data upgrade failed- status code 2 . Neueinspielung der Anwendersoftware. Konfiguration wird nicht sngenommen. Neukonfiguration des Analysators. Kalibrationen und Qualitätskontrolle o. K. | SWO-95968 | No | Exchange of the gas module 1. Cleaning of the PCO2 measuring chamber and removing the air bubbles from the electrolyte. pCO2 values for all analyzers in Mosbach and Buchen borderline to the lower limit value. Data drawn for evaluation and opening of a techline. Quality control faulty. measuring chamber leaks. Cleaned and lower electrode sealing ring renewed at the pH electrode. Further leakage, lower electrode sealing ring of the Dumy Electrode in Cl measuring chamber renewed. Continued leakage, LS adjustment and pumps Cal., no success. LS error, lower LS sensor in pH BG module defective, renewed. Inlet seal renewed. Wieterhin Quality Control Boundary Violation at Kevel Red. Calibration with new membrane batch, no success. Replacement pCO2 electrode. Calibrations o. K., and quality control at pCO2 not o. K. | Replacement of the pH BG module. LS adjustment, pump scale and zero-point calibration. Calibrated, quality control values targeted. Service Dump pulled for investigation in Technical Support. Restart of the analyzer, then analyser remains in the loading of the user software with message; Data upgrade failed- status code 2 . Re-recording of the user software. Configuration is not taken sn. Reconfiguration of the analyzer. Calibrations and quality control o.K. | Defect | HW | pH/BG Module | pCO2
ABL825 FLEX | Lactat Elektrode nach SI 16.004 Anleitung neu membraniert, Funktionskontrolle ok | Lactat Elektrode nach SI 16.004 Anleitung neu membraniert, Funktionskontrolle ok | SWO-95869 | No | Lactate electrode according to SI 16.004 instructions re-membraneed, function control ok | Lactate electrode according to SI 16.004 instructions re-membraneed, function control ok | Assisted Replacement | Electrode | Membrane | Lac | Sensitivity
ABL825 FLEX | Lactat Elektrode nach Anleitung SI 16.004 neu membraniert, Funktionskontrolle ok | Lactat Elektrode nach Anleitung SI 16.004 neu membraniert, Funktionskontrolle ok | SWO-95870 | No | Lactate electrode according to instructions SI 16.004 re-membraneed, function control ok | Lactate electrode according to instructions SI 16.004 re-membraneed, function control ok | Assisted Replacement | Electrode | Membrane | Lac | Sensitivity
ABL810 FLEX | Tausch des AutoCheck Moduls. Reinigung der pCO2 Messkammer, pCO2 Elektrode Luftbläschen aus Elektrolyt herausgeklopft. Qualitätskontroll Messwerte bei Oxymetrie erhöht, Proteinentfernung durchgeführt und tHb Kalibration erneut durchgeführt. Kalibrationen und Qualitätskontrolle o. K. | Tausch des AutoCheck Moduls. Reinigung der pCO2 Messkammer, pCO2 Elektrode Luftbläschen aus Elektrolyt herausgeklopft. Qualitätskontroll Messwerte bei Oxymetrie erhöht, Proteinentfernung durchgeführt und tHb Kalibration erneut durchgeführt. Kalibrationen und Qualitätskontrolle o. K. | SWO-95966 | No | Replacing the AutoCheck module. Cleaning of the pCO2 measuring chamber, pCO2 electrode air bubbles knocked out of electrolyte. Quality control measured values in oxymetry increased, protein removal performed and tHb calibration performed again. Calibrations and quality control o.K. | Replacing the AutoCheck module. Cleaning of the pCO2 measuring chamber, pCO2 electrode air bubbles knocked out of electrolyte. Quality control measured values in oxymetry increased, protein removal performed and tHb calibration performed again. Calibrations and quality control o.K. | Defect | HW | AutoCheck Module | Ampoule breaker
ABL810 FLEX |  | B | D0A0EE272A6293F5 | No |  |  | Assisted Replacement | Electrode | Membrane | pCO2
ABL820 FLEX | Found on of the nipple fittings for inlet gasket was defective had a piece of plastic bent over blocking flow | replaced the nipple kit assy, and had paperwork signed for installations. | D09FB32CB439E869 | No | Found on of the nipple fittings for inlet gasket was defective had a piece of plastic bent over blocking flow | replaced the nipple kit assy, and had paperwork signed for installations. | Defect | HW | Inlet module | Leaking | Nipple kit
ABL80 FLEX BASIC | 更换新测试卡修复仪器 | 更换新测试卡修复仪器 | SWO-96052 | No | Replace the new test card repair instrument | Replace the new test card repair instrument
ABL80 | The problem originated as a pO2 issue and the customer repalced a solution pack and sensor cassete. Then the issue became a pCO2 problem, and another sensor cassette was replaced, but this did not resolve the problem. I found that none of the QC was scanned in, so I instructed the customer on how to scan in the QC. I performed a zero calibration, and replaced another solution pack, but this also did not resolve the problem. I ordered another sensor cassette overnight. | Replaced the sensor cassette and this resolved the problem. | D0A0E9932A021A7C | No | The problem originated as a pO2 issue and the customer repalced a solution pack and sensor cassete. Then the issue became a pCO2 problem, and another sensor cassette was replaced, but this did not resolve the problem. I found that none of the QC was scanned in, so I instructed the customer on how to scan in the QC. I performed a zero calibration, and replaced another solution pack, but this also did not resolve the problem. I ordered another sensor cassette overnight. | Replaced the sensor cassette and this resolved the problem.
ABL90 FLEX | Mehrere Probenmessungen durchgeführt, bei Start der Probenreihe war das Phänomen bei der ersten Messung ebenfalls zu erkennen, weitere Messungen (9 Stk.) warenerfolgreich und wie gewohnt durchgeführt, verschiedene Varianten mit oder ohne Eingabe waren ebenfalls erfolgreich, Kundin gebeten das Problem zu beobachten, Kundin würde sich bei weiteren Vorkommnissen wieder melden, Kundenkontakt von seitens Radiometer wird gehalten. | Software V3.1 MR8 erneut aufgespielt, Konfiguration vorher gesichert und ausgedruckt, Software Update lief ohne Probleme durch, Warmstart nach Software Update erfolgreich, manuelle Kal. durchgeführt und bestanden, QC Messungen durchgeführt, mehrere Probenmessungen durchgeführt, welche alle erfolgreich waren, System- und Funktionskontrolle in Ordnung. | SWO-96346 | No | Several sample measurements were carried out, at the start of the sample series the phenomenon was also visible at the first measurement, further measurements (9 pcs.) were successful and carried out as usual, different variants with or without input were also successful, Customer asked to observe the problem, customer would report again in case of further incidents, customer contact from radiometers is maintained. | Software V3.1 MR8 replayed, configuration previously backed up and printed, software update went through without any problems, warm start after software update successful, manual cal. performed and passed, QC measurements performed, several sample measurements were carried out, all of which were successful, system and function control in order. | Unknown | patient result | patient result | patient result
ABL90 FLEX | szerviz pO2 teszt, oldat csere, kalibrációk, QC mérés | B | SWO-95992 | No | Service pO2 test, Solution replacement, calibrations, QC measurement |  | Assisted Replacement | SP | SP | pO2
ABL90 |  | B | D0A0D8542A5B1943 | Yes |  |  | Re classified RMED | SWU
ABL90 FLEX PLUS |  | B | D0A116462A24C72C | Yes |  |  | Re classified RMED | PER
ABL90 FLEX | 更换配件903-158修复仪器(by zhenjie) | 更换配件903-158修复仪器(by zhenjie) | SWO-96069 | No | Replacement accessory 903-158 repair instrument (by zhenjie) | Replacement accessory 903-158 repair instrument (by zhenjie) | Defect | HW | SP PCB | SP PCB
AQT90 FLEX | 更换903-364修复仪器(by zhenjie) | 更换903-364修复仪器(by zhenjie) | SWO-95511 | No | Replace the 903-364 repair instrument (by zhenjie) | Replace the 903-364 repair instrument (by zhenjie) | Defect | HW | Process Module | Optical Unit | 131
ABL800 BASIC | check the log. reset the oximetery section. shutdown system around 30-40 min . reset the PCB. restart the system & calibrate the system. | B | SWO-96268 | No | check the log. reset the oximetery section. shutdown system around 30-40 min . reset the PCB. restart the system & calibrate the system. |  | Defect | HW | Oxi module | Neon Lamp, low
ABL837 FLEX |  | B | D0A148272A3C194C | Yes |  |  | Re classified RMED | Setup, customer request
ABL837 FLEX | メンブラン内の気泡抜き メンブラン取り付け直し エアライン・測定チャンバー清掃 タンパク除去・LS調整 CAL・QC測定データ確認 | メンブラン内の気泡抜き メンブラン取り付け直し エアライン・測定チャンバー清掃 タンパク除去・LS調整 CAL・QC測定データ確認 | SWO-96300 | No | Re-installing membranes without bubbles in membranes, removing proteins, AND confirming LS adjustment CAL and QC measurement data | Re-installing membranes without bubbles in membranes, removing proteins, AND confirming LS adjustment CAL and QC measurement data | Cleaned | HW | Fluidic pathway | Crea
ABL800 BASIC | Kontakte der Ref-Elektrode und Kontaktkappen nochmals gereinigt. Referenz-Messkammer gereinigt. Kontakte der Ref-Elektrodenkappe justiert. Kalibrationen und QC's vorgenommen. | Kontakte der Ref-Elektrode und Kontaktkappen nochmals gereinigt. Referenz-Messkammer gereinigt. Kontakte der Ref-Elektrodenkappe justiert. Kalibrationen und QC's vorgenommen. | SWO-96241 | No | Contacts of the ref electrode and contact caps cleaned again. Reference measuring chamber cleaned. contacts of the ref electrode cap. calibrations and QC's. | Contacts of the ref electrode and contact caps cleaned again. Reference measuring chamber cleaned. contacts of the ref electrode cap. calibrations and QC's. | Cleaned | HW | pH/BG Module | Reference
ABL820 FLEX | Kal., QCs, Datenverbindung und NetOp i.O. Austausch des Gebrauchtteils bei der nächsten Wartung geplant. | Analysator geöffnet, alle Stecker von der AU gezogen, keine Änderung, ABC LEDs leuchten durchgehend rot. AU Software Installation kommt nicht über den Punkt waiting for Bootstrap Code hinaus. Analysator zur Reparatur abgezogen. | SWO-96234 | No | Cal., QCs, data connection and NetOp i.O. replacement of the used part planned at the next maintenance. | Analyzer open, all plugs pulled from the AU, no change, ABC LEDs glow throughout red. AU Software installation does not go beyond waiting for bootstrap code. analyzer for repair. | Defect | HW | Basic unit | Analyzer Control PCB
ABL827 FLEX | Reviewed March WDC report also found O2Hb running high on all levels of QC. | Performed service tHb calibration & verification. | D0A0E4B92A56FAC8 | No | Reviewed March WDC report also found O2Hb running high on all levels of QC. | Performed service tHb calibration & verification. | Adjustment | HW | Oxi module | tHb cal
ABL837 FLEX | Spild studs skiftet SW opdateret til ver. 6.16 | Spild studs skiftet SW opdateret til ver. 6.16 | SWO-96682 | No | Scrap Studs switched SW updated to ver. 6.16 | Scrap Studs switched SW updated to ver. 6.16 | Defect | HW | Fluidic manifold | Y piece
ABL825 FLEX | zur Fehlersuche Display ersetzt und Kalibriert. zur weiteren Beobachtung | Neue PC Unit eingesetzt. Kal und QC nach Einlaufphase ok. Zur weiteren Beobachtung | SWO-96195 | No | for troubleshooting display is replaced and calibrated. for further observation | New PC Unit. Kal and QC after run-in phase ok. For further observation | Defect | HW | DMS_PC | Touch screen
ABL835 FLEX | Customer had replaced membrane using lot R0736 but could not get qc to inprove | tested gas valves passsed cleaned electrode and replaced membrane using lot r0636 | D0A0D5942A23CED9 | No | Customer had replaced membrane using lot R0736 but could not get qc to inprove | tested gas valves passsed cleaned electrode and replaced membrane using lot r0636 | Assisted Replacement | Electrode | Membrane | pCO2
ABL835 FLEX |  | B | D0A132562A071949 | No |  |  | Assisted Replacement | Electrode | Membrane | pCO2
ABL835 FLEX | Applications Software erneut installiert und alle Daten wieder hergestellt. | Applications Software erneut installiert und alle Daten wieder hergestellt. | SWO-96259 | No | Applications software reinstalled and all data restored. | Applications software reinstalled and all data restored. | Defect | SW | DMS_PC | Missing button
AQT90 FLEX | Replaced optical unit set up and test ok. System check passed ok. Recalibrate D-dim ok Ran QC passed ok | Advised customer to run system test this failed. Engineer required | SWO-96279 | No | Replaced optical unit set up and test ok. System check passed ok. Recalibrate D-dim ok Ran QC passed ok | Advised customer to run system test this failed. Engineer required | Defect | HW | Process Module | Optical Unit | 135
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-97822 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL80 FLEX BASIC | 更换902-972一个正常。(by zhenjie) | 更换902-972一个正常。(by zhenjie) | SWO-95985 | No | Replace 902-972 one normal. (by zhenjie) | Replace 902-972 one normal. (by zhenjie)
ABL90 FLEX | Software V3.1 MR8 erneut aufgespielt, Konfiguration vorher gesichert und ausgedruckt, Software Update lief ohne Probleme durch, Warmstart nach Software Update erfolgreich, manuelle Kal. durchgeführt und bestanden, QC Messungen durchgeführt, mehrere Probenmessungen durchgeführt, welche alle erfolgreich waren, System- und Funktionskontrolle in Ordnung. | Software V3.1 MR8 erneut aufgespielt, Konfiguration vorher gesichert und ausgedruckt, Software Update lief ohne Probleme durch, Warmstart nach Software Update erfolgreich, manuelle Kal. durchgeführt und bestanden, QC Messungen durchgeführt, mehrere Probenmessungen durchgeführt, welche alle erfolgreich waren, System- und Funktionskontrolle in Ordnung. | SWO-96502 | Yes | Software V3.1 MR8 replayed, configuration previously backed up and printed, software update went through without any problems, warm start after software update successful, manual cal. performed and passed, QC measurements performed, several sample measurements were carried out, all of which were successful, system and function control in order. | Software V3.1 MR8 replayed, configuration previously backed up and printed, software update went through without any problems, warm start after software update successful, manual cal. performed and passed, QC measurements performed, several sample measurements were carried out, all of which were successful, system and function control in order. | Re classified RMED | Continued Work
ABL90 FLEX | Backflush Prcedure ohne Erfolg. Hemolyzer Unit und Ventilschlauch erneuert. Oxi Complete Test ok. tHb Kalibration durchgeführt (100,4%). Kalibrationen und QC's vorgenommen. | Backflush Prcedure ohne Erfolg. Hemolyzer Unit und Ventilschlauch erneuert. Oxi Complete Test ok. tHb Kalibration durchgeführt (100,4%). Kalibrationen und QC's vorgenommen. | SWO-96511 | No | Backflush Prcedure without success. Hemolyzer unit and valve hose. Oxi Complete Test ok. tHb calibration (100.4%). calibrations and QC's. | Backflush Prcedure without success. Hemolyzer unit and valve hose. Oxi Complete Test ok. tHb calibration (100.4%). calibrations and QC's. | Defect | HW | Oxi module | Hemolyzer
ABL90 FLEX | Daten gesichert und SC getauscht. | Daten gesichert und SC getauscht. | SWO-96397 | No | Data backed up and SC exchanged. | Data backed up and SC exchanged. | Assisted Replacement | SC | SC | Na
ABL90 FLEX | Fehlerhafter SP wieder eingesetzt.Initialisierung 1 Std abgewartet. Gerät Messbereit | Fehlerhafter SP wieder eingesetzt.Initialisierung 1 Std abgewartet. Gerät Messbereit | SWO-96560 | No | Faulty SP reinused. Initialization 1 hr waited. Device Ready to measure | Faulty SP reinused. Initialization 1 hr waited. Device Ready to measure | Adjustment | HW | SP PCB | SP PCB, contact
ABL90 FLEX PLUS | Gerät überprüft: LP nicht richtig aktiviert ( nicht korrekt nach unten gedrückt) -> LP korrekt aktiviert, Startphase abgewartet, alle fehlerhafte QC wiederholt | Gerät überprüft: LP nicht richtig aktiviert ( nicht korrekt nach unten gedrückt) -> LP korrekt aktiviert, Startphase abgewartet, alle fehlerhafte QC wiederholt | SWO-96590 | No | Device checked: LP not properly activated (not properly pressed down) -> LP correctly activated, start phase waited, all faulty QC repeated | Device checked: LP not properly activated (not properly pressed down) -> LP correctly activated, start phase waited, all faulty QC repeated | User error | SP Activation | SP Activation | SP Activation
ABL90 FLEX PLUS | auto inlet PCB is vervangen. | auto inlet PCB is vervangen. | SWO-96428 | No | auto inlet PCB is vervangen. | auto inlet PCB is vervangen. | Defect | HW | Auto Inlet module | Auto inlet PCB
ABL800 BASIC | Found Printer head having problem, cleaned the head but problem is remaining same, replaced the new printer head from our existing stock restarted instrument found instrument is functional. | Found Printer head having problem, cleaned the head but problem is remaining same, replaced the new printer head from our existing stock restarted instrument found instrument is functional. | SWO-97948 | No | Found Printer head having problem, cleaned the head but problem is remaining same, replaced the new printer head from our existing stock restarted instrument found instrument is functional. | Found Printer head having problem, cleaned the head but problem is remaining same, replaced the new printer head from our existing stock restarted instrument found instrument is functional. | Defect | HW | Printer module | Printer head
ABL825 FLEX | Mangels Ersatzteile gebrauchtes Modul verbaut. Original Modul wird in der Werkstatt repariert. | Neon Lamp ersetzt, Oxy complete Test, Kalibrationen, QC-Kontrolle Rücktausch und Test vor Ort | SWO-96458 | No | In the absence of spare parts used module installed. Original module is repaired in the workshop. | Neon lamp replaced, oxy complete test, calibrations, QC control exchange and on-site test | Defect | HW | Oxi module | Neon Lamp, low
ABL810 FLEX | Tausch des pH BG Moduls. LS Justage, Pumpenkal und Nullpunktkalibrierung durchgeführt. Einkalibriert, Qualitätskontroll Werte zielwertig. Service Dump gezogen zur Untersuchung im Technischen Support. Neustart des Analysators, anschließend bleibt Analysator im Laden der Anwendersoftware stehen mit Meldung, Data upgrade failed- status code 2 . Neueinspielung der Anwendersoftware. Konfiguration wird nicht sngenommen. Neukonfiguration des Analysators. Kalibrationen und Qualitätskontrolle o. K. | Tausch des pH BG Moduls. LS Justage, Pumpenkal und Nullpunktkalibrierung durchgeführt. Einkalibriert, Qualitätskontroll Werte zielwertig. Service Dump gezogen zur Untersuchung im Technischen Support. Neustart des Analysators, anschließend bleibt Analysator im Laden der Anwendersoftware stehen mit Meldung; Data upgrade failed- status code 2 . Neueinspielung der Anwendersoftware. Konfiguration wird nicht sngenommen. Neukonfiguration des Analysators. Kalibrationen und Qualitätskontrolle o. K. | SWO-96533 | No | Replacement of the pH BG module. LS adjustment, pump scale and zero-point calibration. Calibrated, quality control values targeted. Service Dump pulled for investigation in Technical Support. Restart of the analyzer, then analyser remains in the loading of the user software with message; Data upgrade failed- status code 2 . Re-recording of the user software. Configuration is not taken sn. Reconfiguration of the analyzer. Calibrations and quality control o.K. | Replacement of the pH BG module. LS adjustment, pump scale and zero-point calibration. Calibrated, quality control values targeted. Service Dump pulled for investigation in Technical Support. Restart of the analyzer, then analyser remains in the loading of the user software with message; Data upgrade failed- status code 2 . Re-recording of the user software. Configuration is not taken sn. Reconfiguration of the analyzer. Calibrations and quality control o.K. | Defect | HW | pH/BG Module | pCO2
ABL810 FLEX | Tausch des pH BG Moduls. LS Justage, Pumpenkal und Nullpunktkalibrierung durchgeführt. Einkalibriert. Kalibrationen und Qualitätskontrolle o. K., Werte zielwertig. | Tausch des pH BG Moduls. LS Justage, Pumpenkal und Nullpunktkalibrierung durchgeführt. Einkalibriert. Kalibrationen und Qualitätskontrolle o. K., Werte zielwertig. | SWO-96534 | No | Replacement of the pH BG module. LS adjustment, pump scale and zero-point calibration. Calibrated. Calibrations and quality control o.K., values target. | Replacement of the pH BG module. LS adjustment, pump scale and zero-point calibration. Calibrated. Calibrations and quality control o.K., values target. | Defect | HW | pH/BG Module | pCO2
AQT90 FLEX | Testkassette entfernt und Lichtschranke gereinigt | Defekte TNT-Parameterkassette erneuert. | SWO-96464 | No | Test cassette removed and light barrier cleaned | Defective TNT parameter cartridge renewed. | Defect | Consumables | Test Kit | Cartridge, TnT | TNT - 5784
ABL90 FLEX | New inlet gasket with holder fitted. Cycled through the fluid transport tests, all tests passed. Analyser performed start up procedure. | New inlet gasket with holder fitted. Cycled through the fluid transport tests, all tests passed. Analyser performed start up procedure. | SWO-96606 | No | New inlet gasket with holder fitted. Cycled through the fluid transport tests, all tests passed. Analyser performed start up procedure. | New inlet gasket with holder fitted. Cycled through the fluid transport tests, all tests passed. Analyser performed start up procedure. | User error | SP Activation | SP Activation | SP Activation
AQT90 FLEX | 清理存储盘的破碎塑料片 | 清理存储盘的破碎塑料片 | SWO-96597 | No | Clean up the broken plastic sheet of the storage tray | Clean up the broken plastic sheet of the storage tray | Defect | Consumables | Test Kit | Cartridge | 5784 (Broken or Label)
AQT90 FLEX | Has to change the optical unit and process wheel. | B | SWO-96987 | No | Has to change the optical unit and process wheel. |  | Defect | HW | Process Module | Optical Unit | 131
ABL80 FLEX BASIC | 更换测试卡正常。 | 更换测试卡正常。 | SWO-96638 | No | The replacement test card is OK. | The replacement test card is OK.
AQT90 FLEX | 更换903-031，903-030各一个正常。AU板大电流冲击到了测试卡传送板，造成损坏。---Jack.chu | 更换903-031，903-030各一个正常。AU板大电流冲击到了测试卡传送板，造成损坏。---Jack.chu | SWO-96584 | No | Replace 903-031, 903-030 each normal. The AU plate has a high current that hits the test card conveyor, causing damage. ---Jack.chu | Replace 903-031, 903-030 each normal. The AU plate has a high current that hits the test card conveyor, causing damage. ---Jack.chu | Defect | HW | Process Module | PCB, Cup transport - Rear | 1999
ABL80 FLEX BASIC | 更换测试卡修复仪器 | 更换测试卡修复仪器 | SWO-97641 | No | Replace the test card repair instrument | Replace the test card repair instrument
ABL80 FLEX BASIC | 更换新测试卡修复仪器 | 更换新测试卡修复仪器 | SWO-96992 | No | Replace the new test card repair instrument | Replace the new test card repair instrument
ABL90 FLEX | Gleiche SP und SC wiederverwendet.Spülung der SC unterbrochen.PCO2 und Glu..Lac Kalfehler Neue SC eingebaut..Lac zeigt keine Empf. nach 3 Kal Erneuter Wechsel SC Kal und QC ok | Gleiche SP und SC wiederverwendet.Spülung der SC unterbrochen.PCO2 und Glu..Lac Kalfehler Neue SC eingebaut..Lac zeigt keine Empf. nach 3 Kal Erneuter Wechsel SC Kal und QC ok | SWO-96869 | No | Same SP and SC reused. rinse of the SC interrupted. PCO2 and Glu.. Lac Calf Error New SC built-in. Lac shows no empf. after 3 Kal New change SC Kal and QC ok | Same SP and SC reused. rinse of the SC interrupted. PCO2 and Glu.. Lac Calf Error New SC built-in. Lac shows no empf. after 3 Kal New change SC Kal and QC ok | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
ABL90 |  | B | D0A5E88C0234194D | No |  |  | Defect | HW | Inlet Gasket | Inlet gasket
ABL90 FLEX | Netzwerkdaten eingegeben (IP 10.52.202.69), CobasIT-Verbindung eingegeben (10.32.1.16:55578), Verbindung überprüft - Ping i.O., Datennachsendung, Radiance-Verbindung ausgeschalten u. Warteschlange gelöscht | Netzwerkdaten eingegeben (IP 10.52.202.69), CobasIT-Verbindung eingegeben (10.32.1.16:55578), Verbindung überprüft - Ping i.O., Datennachsendung, Radiance-Verbindung ausgeschalten u. Warteschlange gelöscht |  | No | Network data entered (IP 10.52.202.69), CobasIT connection entered (10.32.1.16:55578), connection checked - ping i.O., data mailing, radiance connection turned off and queue deleted | Network data entered (IP 10.52.202.69), CobasIT connection entered (10.32.1.16:55578), connection checked - ping i.O., data mailing, radiance connection turned off and queue deleted | #N/A | #N/A | #N/A | #N/A | #N/A
ABL90 FLEX | Netzwerkdaten eingegeben (10.52.202.70), CobasIT-Verbindung eingegeben (10.32.1.16:55578), Radiance-Verbindung ausgeschalten u. Warteschlange gelöscht, Daten nachgesendet, Ping i.O. Name: 090R0770N006 | Netzwerkdaten eingegeben (10.52.202.70), CobasIT-Verbindung eingegeben (10.32.1.16:55578), Radiance-Verbindung ausgeschalten u. Warteschlange gelöscht, Daten nachgesendet, Ping i.O. Name: 090R0770N006 | SWO-96777 | No | Network data entered (10.52.202.70), CobasIT connection entered (10.32.1.16:55578), radiance connection turned off and queue deleted, data sent back, ping i.o. Name: 090R07070N006 | Network data entered (10.52.202.70), CobasIT connection entered (10.32.1.16:55578), radiance connection turned off and queue deleted, data sent back, ping i.o. Name: 090R07070N006 | Setup | SW | His/lis | His/Lis | Corbas IT
ABL90 FLEX | Netzwerkeinstellungen überprüft u. gespeichert, CobasIT-Verbindung eingegeben, Ping i.O. | B | SWO-96967 | No | Network settings checked & saved, CobasIT connection entered, ping i.O. |  | Setup | SW | His/lis | His/Lis | Corbas IT
ABL90 FLEX | 更换测试卡修复仪器 | 更换测试卡修复仪器 | SWO-97818 | No | Replace the test card repair instrument | Replace the test card repair instrument
AQT90 FLEX |  | B | SWO-96719 | No |  |  | Defect | HW | Aspiration Module | Cable | 5669
AQT90 FLEX | Checked and aligned the barcode reader. | Checked and aligned the barcode reader. | SWO-96713 | No | Checked and aligned the barcode reader. | Checked and aligned the barcode reader. | Adjustment | HW | Storage Module | Barcode reader, Storage wheel
ABL805 FLEX | Comprobaciónd el estado del analizador. Se procede a la limpieza del electrodo con electrode cleaning solution así con agua destilada. Se realizan varias Cal para estabilizar y Qc´s para comprobar funcionamiento, tambien se pasan muestras. pH parece recuperado. | Limpieza electrodo de pH | D0A5CB680236C1F5 | No | Check the status of the scanner. The electrode is cleaned with electrode cleaning solution as well as distilled water. Several Cal are performed to stabilize and Qc's to check performance, samples are also passed. pH seems recovered. | Cleaning pH electrode | Cleaned | HW | pH/BG Module | pH
ABL800 BASIC | Found analyzer software error, Installed the windows and analyzer software and restarted instrument done setup found instrument is functional. | Found analyzer software error, Installed the windows and analyzer software and restarted instrument done setup found instrument is functional. | SWO-97951 | No | Found analyzer software error, Installed the windows and analyzer software and restarted instrument done setup found instrument is functional. | Found analyzer software error, Installed the windows and analyzer software and restarted instrument done setup found instrument is functional. | Defect | SW | DMS_PC | Windows
ABL835 FLEX | Replace faulty power supply. | Replace faulty power supply. | SWO-96949 | No | Replace faulty power supply. | Replace faulty power supply. | Defect | HW | Basic unit | Power Supply
ABL827 FLEX | Remove and strip down FlexQ module, clean extensive blood spillage from chassis and other parts, re orientate guide shaft by 180 degrees as current side in use was worn and slipping, re assemble. | Remove and strip down FlexQ module, clean extensive blood spillage from chassis and other parts, re orientate guide shaft by 180 degrees as current side in use was worn and slipping, re assemble. | SWO-96912 | No | Remove and strip down FlexQ module, clean extensive blood spillage from chassis and other parts, re orientate guide shaft by 180 degrees as current side in use was worn and slipping, re assemble. | Remove and strip down FlexQ module, clean extensive blood spillage from chassis and other parts, re orientate guide shaft by 180 degrees as current side in use was worn and slipping, re assemble. | Adjustment | HW | AutoCheck Module | Guide shaft
ABL835 FLEX | pCO2電極・メンブラン交換 pHシーリング・ダミー電極洗浄・パク除・LS調整 CAL3回入れQC Lv2は少し上がったが、Lv4は範囲外 pHシーリングリング交換、チャンバー押さえつけ直し、LS調整・ポンプCAL CAL・QC Lv2・Lv4データ確認 | pCO2電極・メンブラン交換 pHシーリング・ダミー電極洗浄・パク除・LS調整 CAL3回入れQC Lv2は少し上がったが、Lv4は範囲外 pHシーリングリング交換、チャンバー押さえつけ直し、LS調整・ポンプCAL CAL・QC Lv2・Lv4データ確認 | SWO-96818 | No | pCO2 electrode, membrane exchange pH sealing, dummy electrode cleaning, pak removal, LS adjustment CAL3 times put QC Lv2 went up a little, Lv4 is out of range pH sealing ring replacement, chamber pressing re-press, LS adjustment / pump CAL CAL / QC Lv2 and Lv4 data confirmation | pCO2 electrode, membrane exchange pH sealing, dummy electrode cleaning, pak removal, LS adjustment CAL3 times put QC Lv2 went up a little, Lv4 is out of range pH sealing ring replacement, chamber pressing re-press, LS adjustment / pump CAL CAL / QC Lv2 and Lv4 data confirmation | Assisted Replacement | Electrode | Electrode | pCO2
ABL835 FLEX | Windows erreur ne demarre pas | Remplac disque dur et reload windows et DMS sotware et reload configuration client | D0A60BC1021AD9CA | No | Windows error doesn't start | Remplac disque dur et reload windows et DMS sotware et reload configuration client | Defect | HW | DMS_PC | HDD
ABL825 FLEX | Neue PC Unit eingesetzt. Kal und QC nach Einlaufphase ok. Zur weiteren Beobachtung | Neue PC Unit eingesetzt. Kal und QC nach Einlaufphase ok. Zur weiteren Beobachtung | SWO-96871 | No | New PC Unit. Kal and QC after run-in phase ok. For further observation | New PC Unit. Kal and QC after run-in phase ok. For further observation | Defect | HW | DMS_PC | CPU
ABL825 FLEX |  | B | D0A5EC0F020120D7 | No |  |  | Defect | HW | Inlet module | Leaking | Nipple kit
ABL825 FLEX | Gerät geprüft. Aus El Met Modul läuft Flüssigkeit aus. Elektroden geprüft und Messkammer gesäubert. Einlassdichtung und Einlassschlauch erneuert. Lufteinlass gespült,LS Justage und Pumpenkal gestartet. Diverse Kal und QC durchgeführt. | Gerät geprüft. Aus El Met Modul läuft Flüssigkeit aus. Elektroden geprüft und Messkammer gesäubert. Einlassdichtung und Einlassschlauch erneuert. Lufteinlass gespült,LS Justage und Pumpenkal gestartet. Diverse Kal und QC durchgeführt. | SWO-96900 | No | device checked. Liquid leaks from El Met module. electrodes tested and measuring chamber cleaned. Inlet seal and inlet hose renewed. Air inlet rinsed,LS adjustment and pump scale started. Various Kal and QC performed. | device checked. Liquid leaks from El Met module. electrodes tested and measuring chamber cleaned. Inlet seal and inlet hose renewed. Air inlet rinsed,LS adjustment and pump scale started. Various Kal and QC performed. | Defect | HW | El/Met Module | Tubings
ABL825 FLEX | Had customer look in data logs and found error codes 0251 and 0579, halogen lamp to high. | Adjusted lamp, did complete oximetry test, did THB is service and for user and as sample. | D0A605C6022D72D9 | No | Had customer look in data logs and found error codes 0251 and 0579, halogen lamp to high. | Adjusted lamp, did complete oximetry test, did THB is service and for user and as sample. | Adjustment | HW | Oxi module | Lamp unit
ABL815 FLEX | Halogenlampe nachjustiert. Complete Oximetrietest durchgeführt. Halogenintensität bei 3256. Service-tHb-Kalibration durchgeführt. Kalibrationen und Qualitätskontrollen in Ordnung. | Halogenlampe nachjustiert. Complete Oximetrietest durchgeführt. Halogenintensität bei 3256. Service-tHb-Kalibration durchgeführt. Kalibrationen und Qualitätskontrollen in Ordnung. | SWO-96940 | No | Halogen lamp readjusted. Complete oximetry test. halogen intensity at 3256. Service tHb calibration. calibrations and quality controls in order. | Halogen lamp readjusted. Complete oximetry test. halogen intensity at 3256. Service tHb calibration. calibrations and quality controls in order. | Adjustment | HW | Oxi module | Lamp unit
AQT90 FLEX | Fehlersuche, Austausch touchpad, Funktionskontrolle | Fehlersuche, Austausch touchpad, Funktionskontrolle | SWO-96768 | No | Troubleshooting, replacement touchpad, function control | Troubleshooting, replacement touchpad, function control | Defect | HW | Analyzer Control Module | Touch screen
AQT90 FLEX | Replaced needle. carried out needle calibration. Analyser giving flow error. Found fault to be the needle wash unit had water damage. return visit required to replace needle wash unit and re calibrate | Replaced Needle wash unit and optical sensor. carried out calibration of needle and full HCT calibration. Ran QC All now working ok | SWO-97564 | No | Replaced needle. carried out needle calibration. Analyser giving flow error. Found fault to be the needle wash unit had water damage. return visit required to replace needle wash unit and re calibrate | Replaced Needle wash unit and optical sensor. carried out calibration of needle and full HCT calibration. Ran QC All now working ok | Defect | HW | Aspiration Module | Bend needle | 38
AQT90 FLEX | Analysator geöffnet, Safe Positions durchgeführt. CRP Testkassette LOT:11429 mit fehlendem Aufkleber entfernt. Anwender über Sachverhalt und Fehlervermeidung (Aufkleber vor Installation fest andrücken) aufgeklärt. | Analysator geöffnet, Safe Positions durchgeführt. CRP Testkassette LOT:11429 mit fehlendem Aufkleber entfernt. Anwender über Sachverhalt und Fehlervermeidung (Aufkleber vor Installation fest andrücken) aufgeklärt. | SWO-96974 | No | Analyzer open, Safe Positions performed. CRP test cartridge LOT:11429 with missing sticker removed. Users informed about facts and error avoidance (putting stickers firmly in place before installation). | Analyzer open, Safe Positions performed. CRP test cartridge LOT:11429 with missing sticker removed. Users informed about facts and error avoidance (putting stickers firmly in place before installation). | Defect | Consumables | Test Kit | Barcode Label | 5784
AQT90 FLEX | Defekte TNT-Parameterkassette erneuert. | Defekte TNT-Parameterkassette erneuert. | SWO-96797 | No | Defective TNT parameter cartridge renewed. | Defective TNT parameter cartridge renewed.
ABL80 FLEX BASIC | 更换测试卡修复仪器 | 更换测试卡修复仪器 | SWO-97645 | No | Replace the test card repair instrument | Replace the test card repair instrument
ABL80 |  | B | D0A6E9117845BE3A | No
ABL80 FLEX BASIC | 更换舱门903-363后修复---jack.chu | 更换舱门903-363后修复---jack.chu | SWO-96721 | No | Repair ---jack.chu after replacing hatch 903-363 | Repair ---jack.chu after replacing hatch 903-363
ABL80 FLEX BASIC | O型圈一个折了，尾部的拿了东西缠上了。---jack.chu | O型圈一个折了，尾部的拿了东西缠上了。---jack.chu | SWO-96650 | No | The O-ring is folded, and the tail is entangled with something. ---jack.chu | The O-ring is folded, and the tail is entangled with something. ---jack.chu
ABL80 FLEX BASIC |  | B | SWO-97029 | No
ABL80 FLEX BASIC | 更换902-972一个正常。---jack.chu | 更换902-972一个正常。---jack.chu | SWO-96799 | No | Replace 902-972 one normal. ---jack.chu | Replace 902-972 one normal. ---jack.chu
ABL800 FLEX | 修理完了モジュール入替 OXIコンプリートテスト確認 CAL・tHb CAL・StHb CAL・QC測定データ確認 | 修理完了モジュール入替 OXIコンプリートテスト確認 CAL・tHb CAL・StHb CAL・QC測定データ確認 | SWO-97276 | No | Repair Completion Module Replacement OXI Complete Test Confirmation CAL, tHb CAL, StHb CAL, QC Measurement Data Confirmation | Repair Completion Module Replacement OXI Complete Test Confirmation CAL, tHb CAL, StHb CAL, QC Measurement Data Confirmation | Defect | HW | Oxi module | tHb cal
ABL800 BASIC | So clean ph electrode and measuring chamber and again perform the samples but still it shows the error. Check and replace the inlet gasket and pump tubes. Then perform the Samples and it gives the result without any error. After replace the tubes and gasket unit is working properly. | So clean ph electrode and measuring chamber and again perform the samples but still it shows the error. Check and replace the inlet gasket and pump tubes. Then perform the Samples and it gives the result without any error. After replace the tubes and gasket unit is working properly. | SWO-97806 | No | So clean ph electrode and measuring chamber and again perform the samples but still it shows the error. Check and replace the inlet gasket and pump tubes. Then perform the Samples and it gives the result without any error. After replace the tubes and gasket unit is working properly. | So clean ph electrode and measuring chamber and again perform the samples but still it shows the error. Check and replace the inlet gasket and pump tubes. Then perform the Samples and it gives the result without any error. After replace the tubes and gasket unit is working properly. | Defect | HW | pH/BG Module | Tubings
ABL800 BASIC | Found Sample error replaced the inlet gasket, restarted instrument performed the sample found results are ok instrument is functional. | Found Sample error replaced the inlet gasket, restarted instrument performed the sample found results are ok instrument is functional. | SWO-97903 | No | Found Sample error replaced the inlet gasket, restarted instrument performed the sample found results are ok instrument is functional. | Found Sample error replaced the inlet gasket, restarted instrument performed the sample found results are ok instrument is functional. | Defect | HW | Inlet module | Inlet gasket
ABL815 FLEX | Reinigungsverfahren für Lac. Elektrode durchgeführt, den Anwendern erklärt, Ausdruck überreicht. Lac.neu bespannt. V1-V4 gereinigt, Öl aus den Ventilen entfernt. | Reinigungsverfahren für Lac. Elektrode durchgeführt, den Anwendern erklärt, Ausdruck überreicht. Lac.neu bespannt. V1-V4 gereinigt, Öl aus den Ventilen entfernt. | SWO-97919 | No | Cleaning process for Lac. electrode performed, explained to the users, expression handed over. Lac.new. V1-V4 cleaned, oil removed from the valves. | Cleaning process for Lac. electrode performed, explained to the users, expression handed over. Lac.new. V1-V4 cleaned, oil removed from the valves. | Cleaned | Electrode | Electrode | Lac
ABL827 FLEX |  | B | D0A7145C780EB749 | Yes |  |  | Re classified RMED | Demo setup
ABL800 BASIC | 更换了Lac电极后解决 | 更换了Lac电极后解决 | SWO-97860 | No | Resolved after replacing Lac electrode | Resolved after replacing Lac electrode | Assisted Replacement | Electrode | Electrode | Lac | Calibration
ABL820 FLEX | on samples the sample only is asperated to the ref chamber, the whole chamber does not fill to the top LS. checked upper and lower valves, all check out replaced #6 and both upper valves at dual valve housing | Replaced valves 6 and upper valves and swapped #7 & 8. also replaced lower sealing ring due to fluid leaking from BG gutter at the end of each cycle. Suspect bad fluidic PCB. ordered part will return 4 28 2016 Friday | D0A7979C784B49FE | No | on samples the sample only is asperated to the ref chamber, the whole chamber does not fill to the top LS. checked upper and lower valves, all check out replaced #6 and both upper valves at dual valve housing | Replaced valves 6 and upper valves and swapped #7 & 8. also replaced lower sealing ring due to fluid leaking from BG gutter at the end of each cycle. Suspect bad fluidic PCB. ordered part will return 4 28 2016 Friday | Defect | HW | pH/BG Module | Valve
AQT90 FLEX | Replaced Needle wash unit and optical sensor. carried out calibration of needle and full HCT calibration. Ran QC All now working ok | Replaced Needle wash unit and optical sensor. carried out calibration of needle and full HCT calibration. Ran QC All now working ok | SWO-97566 | No | Replaced Needle wash unit and optical sensor. carried out calibration of needle and full HCT calibration. Ran QC All now working ok | Replaced Needle wash unit and optical sensor. carried out calibration of needle and full HCT calibration. Ran QC All now working ok | Defect | HW | Aspiration Module | Sensor, needle | 38
AQT90 FLEX | Mit Herrn Kozelj telefonisch das Problem lokalisiert und gelöst. Zwischendrin mit Herrn Dr Lacha beim konfigurieren eines frei konfigurierbaren Feldes am ABL 800 auf der Intensivstation geholfen Herrn Dr von Freiburg demonstriert und Herrn Dr Lacha Bescheid gegeben, dass nun die Abfrage funktioniert. AQT ist Messbereit. | Mit Herrn Kozelj telefonisch das Problem lokalisiert und gelöst. Zwischendrin mit Herrn Dr Lacha beim konfigurieren eines frei konfigurierbaren Feldes am ABL 800 auf der Intensivstation geholfen Herrn Dr von Freiburg demonstriert und Herrn Dr Lacha Bescheid gegeben, dass nun die Abfrage funktioniert. AQT ist Messbereit. | SWO-97425 | Yes | With Mr. Kozelj by telephone, the problem is located and solved. In between with Dr. Lacha configuring a freely configurable field on the ABL 800 in the intensive care unit, Dr von Freiburg demonstrated and informed Dr Lacha that the query is now working. AQT is ready to measure. | With Mr. Kozelj by telephone, the problem is located and solved. In between with Dr. Lacha configuring a freely configurable field on the ABL 800 in the intensive care unit, Dr von Freiburg demonstrated and informed Dr Lacha that the query is now working. AQT is ready to measure. | Re classified RMED | Telph Support
AQT90 FLEX | AQT mit CRP Lot 11124 neu kalibriert. Funktionskontrolle iO. | AQT mit CRP Lot 11124 neu kalibriert. Funktionskontrolle iO. | SWO-97129 | No | AQT recalibrated with CRP Lot 11124. Functional control iO. | AQT recalibrated with CRP Lot 11124. Functional control iO. | Defect | Consumables | Test Kit | Barcode Label | 5784
ABL80 FLEX BASIC | 更换了打印机主板后修复---jack.chu | 更换了打印机主板后修复---jack.chu | SWO-97962 | No | Repair ---jack.chu after replacing printer motherboard | Repair ---jack.chu after replacing printer motherboard
ABL80 FLEX BASIC |  | B | SWO-97366 | No
ABL80 FLEX BASIC |  | B | SWO-97338 | No
ABL90 | Checked SP | Replaced SP and Ran all SVC checks. | D0B3A75C14060550 | No | Checked SP | Replaced SP and Ran all SVC checks. | Assisted Replacement | SP | SP | FT
ABL90 |  | B | D0A8CE9EEE4681C8 | No |  |  | Assisted Replacement | SC | SC | pH
ABL90 | Upon arrival verified wireless usb adapter not working correctly. | Replaced wireless usb adapter and installed the necessary drivers to make it work correctly. | D0A8B055EE62F0DC | No | Upon arrival verified wireless usb adapter not working correctly. | Replaced wireless usb adapter and installed the necessary drivers to make it work correctly. | Setup | HW | USB Wifi | USB Wifi
ABL90 FLEX PLUS |  | B | D0A86F6DEE08ABA9 | Yes |  |  | Re classified RMED | PER
ABL90 FLEX PLUS |  | B | D0A8D497EE0D84E3 | Yes |  |  | Re classified RMED | SWU
TCM TOSCA | 更换了电池431-018C后修复---jack.chu | 更换了电池431-018C后修复---jack.chu | SWO-96783 | No | Repaired after replacing battery 431-018C---jack.chu | Repaired after replacing battery 431-018C---jack.chu
ABL90 FLEX | 更换试剂包修复仪器 | 更换试剂包修复仪器 | SWO-97580 | No | Replace reagent kit repair instruments | Replace reagent kit repair instruments
ABL90 FLEX | 更换新测试卡修复仪器 | 更换新测试卡修复仪器 | SWO-97611 | No | Replace the new test card repair instrument | Replace the new test card repair instrument | Assisted Replacement | SC | SC | Ref
ABL90 | Substituição do Spack sem sucesso, substituição da entrada, testes no modo de serviço técnico, establização e passagem do QC. | B | D0A8B5DBEE3FC82F | No | Replacing the Spack without success, replacing the input, testing in the mode of technical service, establishment and passage of QC. |  | Defect | HW | Inlet Gasket | Inlet gasket
ABL815 FLEX |  | B | D0A8D6D6EE3784EF | Yes |  |  | Re classified RMED | SWU
ABL800 BASIC | calibrate the inlet. but needle movement error random. machine under observation. | B | SWO-98099 | No | calibrate the inlet. but needle movement error random. machine under observation. |  | Defect | HW | Inlet module | Inlet module
ABL800 BASIC | Replace the membrane with new one. Again performed 2 point calibration. After replace the membrane unit is working properly. | Replace the membrane with new one. Again performed 2 point calibration. After replace the membrane unit is working properly. | SWO-98089 | No | Replace the membrane with new one. Again performed 2 point calibration. After replace the membrane unit is working properly. | Replace the membrane with new one. Again performed 2 point calibration. After replace the membrane unit is working properly. | Assisted Replacement | Electrode | Membrane | Na+
ABL835 FLEX |  | B | D0A88EC3EE336218 | Yes |  |  | Re classified RMED | SWU
AQT90 FLEX |  | B | D0A83D6CEE023937 | No |  |  | Cleaned | HW | Process Module | Tubes
ABL80 FLEX BASIC | 更换903-334一个 | 更换903-334一个 | SWO-97602 | No | Replace 903-334 one | Replace 903-334 one
ABL80 FLEX BASIC | 更换新测试卡修复仪器 | 更换新测试卡修复仪器 | SWO-97608 | No | Replace the new test card repair instrument | Replace the new test card repair instrument
ABL80 FLEX BASIC | Remove Sc and cleaned it with Distilled water.Than kept in fridge for 30 minutes. Then reinstalled Sc and checked with QC. Results are coming ok. | Remove Sc and cleaned it with Distilled water.Than kept in fridge for 30 minutes. Then reinstalled Sc and checked with QC. Results are coming ok. | SWO-98037 | No | Remove Sc and cleaned it with Distilled water. Than kept in fridge for 30 minutes. Then reinstalled Sc and checked with QC. Results are coming ok. | Remove Sc and cleaned it with Distilled water. Than kept in fridge for 30 minutes. Then reinstalled Sc and checked with QC. Results are coming ok.
ABL80 CO-OX | CONFIRMED ISSUE WHEN ONSITE. | RESOLUTION OF BARCODE PRINT OUT ON PAPER WAS TOO LOW. BCR WAS ABLE TO READ QC AND OTHER MISCELANEOUS BARCODE LABELS. ADVISED CUSTOMER TO IMPROVE BARCODE RESOLUTION. ALSO REPLACED HARD DRIVE FLASH CARD TO PREPARE SYSTEM FOR REMOTE SUPPORT WITH SP2 LOADED. | D0AA098E64364F47 | No | CONFIRMED ISSUE WHEN ONSITE. | RESOLUTION OF BARCODE PRINT OUT ON PAPER WAS TOO LOW. BCR WAS ABLE TO READ QC AND OTHER MISCELANEOUS BARCODE LABELS. ADVISED CUSTOMER TO IMPROVE BARCODE RESOLUTION. ALSO REPLACED HARD DRIVE FLASH CARD TO PREPARE SYSTEM FOR REMOTE SUPPORT WITH SP2 LOADED.
ABL90 FLEX | Fehlersuche, im Serviceprogr.div. Einstellungen vorgenommen, System gereinigt, Kal. + QC i.O. | Fehlersuche; im Serviceprogr.div. Einstellungen vorgenommen; System gereinigt; Kal. + QC i.O. | SWO-97762 | No | Troubleshooting; in Serviceprogr.div. settings made; System cleaned; Kal. + QC in the first place | Troubleshooting; in Serviceprogr.div. settings made; System cleaned; Kal. + QC in the first place | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
ABL90 FLEX | Einlass überprüft, Einlasssonde u. -halter gereinigt, Pressurtest u. Gerinnselentfernung durchgeführt, Kal u. QC i.O. | B | SWO-97738 | No | Inlet checked, inlet probe and holder cleaned, pressurtest and clot removal performed, Kal & QC i.O. |  | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
ABL90 |  | B | D0A99CCA641EBE6F | Yes |  |  | Re classified RMED | SWU
ABL90 | Error 1228 present for failing calibrations. | Replaced Sensor Interface Bd. | D0A9CD2B641F6FE4 | No | Error 1228 present for failing calibrations. | Replaced Sensor Interface Bd. | Defect | HW | Basic Unit | Sensor Interface PCB | (Return Request - 16.5-16) Received 216-5-23
ABL825 FLEX | No investigation necessary | Performed software upgrade to version 6.16. Performed all necessary calibrations. Installed new package inserts for new ranges. | D0A9C385645AC42F | Yes | No investigation necessary | Performed software upgrade to version 6.16. Performed all necessary calibrations. Installed new package inserts for new ranges. | Re classified RMED | SWU
ABL825 FLEX | Gerät überprüft, Elektrode und Kontakte kontrolliert und gereinigt, Kal Lösungen ausgetauscht, E777 (RP01) erneuert | Gerät überprüft, Elektrode und Kontakte kontrolliert und gereinigt, Kal Lösungen ausgetauscht, E777 (RP01) erneuert | SWO-97791 | No | Device checked, electrode and contacts checked and cleaned, Kal solutions replaced, E777 (RP01) renewed | Device checked, electrode and contacts checked and cleaned, Kal solutions replaced, E777 (RP01) renewed | Assisted Replacement | Electrode | Electrode | pH
ABL800 BASIC | replace the gas 2 . calibration sensitivity low. replace the membrane. calibrate multiple time.calibration pass & qc pass. | B | SWO-98095 | No | replace the gas 2 . calibration sensitivity low. replace the membrane. calibrate multiple time.calibration pass & qc pass. |  | Assisted Replacement | Electrode | Membrane | pCO2
ABL810 FLEX | Tausch des pH BG Moduls. LS Justage, Pumpenkal und Nullpunktkalibrierung durchgeführt. Kalibrationen und Qualitätskontrolle o. K. Gemessene Werte sind nicht zielwertig, aber innerhalb der Grenzwerte. | Tausch des pH BG Moduls. LS Justage, Pumpenkal und Nullpunktkalibrierung durchgeführt. Kalibrationen und Qualitätskontrolle o. K. Gemessene Werte sind nicht zielwertig, aber innerhalb der Grenzwerte. | SWO-97782 | No | Replacement of the pH BG module. LS adjustment, pump scale and zero-point calibration. Calibrations and quality control o.K. Measured values are not target values, but within the limits. | Replacement of the pH BG module. LS adjustment, pump scale and zero-point calibration. Calibrations and quality control o.K. Measured values are not target values, but within the limits. | Defect | HW | pH/BG Module | pCO2
ABL800 BASIC | Found calibration error for Lac parameter, performed the Lac electrode cleaning as per cleaning procedure replaced the membrane and calibrated found Lac parameter was functional. | Found calibration error for Lac parameter, performed the Lac electrode cleaning as per cleaning procedure replaced the membrane and calibrated found Lac parameter was functional. | SWO-98101 | No | Found calibration error for Lac parameter, performed the Lac electrode cleaning as per cleaning procedure replaced the membrane and calibrated found Lac parameter was functional. | Found calibration error for Lac parameter, performed the Lac electrode cleaning as per cleaning procedure replaced the membrane and calibrated found Lac parameter was functional. | Assisted Replacement | Electrode | Membrane | Lac | Calibration
AQT90 FLEX | Laser beam in het service menu gecontroleerd en daaruit blijkt dat het niet correct is afgesteld. Zodoende laser beam opnieuw heb afgesteld en gecontroleerd. | Laser beam in het service menu gecontroleerd en daaruit blijkt dat het niet correct is afgesteld. Zodoende laser beam opnieuw heb afgesteld en gecontroleerd. | SWO-97807 | No | Laser Beam in the service menu and it shows that it has not been adjusted correctly. So laser beam has been re-tuned and checked. | Laser Beam in the service menu and it shows that it has not been adjusted correctly. So laser beam has been re-tuned and checked. | Adjustment | HW | Storage Module | Barcode reader, Storage wheel
ABL80 FLEX CO-OX | JW durchgeführt, SC reinitialisiert , keine Verbesserung, Datensicherung u. RtP erstellt an TS geschickt, Info an Kunden | JW durchgeführt, SC reinitialisiert , keine Verbesserung, Datensicherung u. RtP erstellt an TS geschickt, Info an Kunden | SWO-97946 | No | JW performed, SC reinitialized, no improvement, data backup and RtP created sent to TS, info sent to customers | JW performed, SC reinitialized, no improvement, data backup and RtP created sent to TS, info sent to customers
ABL80 OSM | Unable to perform any system activity. Unit locked up while pushing users to Aqure | Performed system reset . Installed setup. Performed software up grade to version 3.14. | D0AAF75EDA514DCF | No | Unable to perform any system activity. Unit locked up while pushing users to Aqure | Performed system reset . Installed setup. Performed software up grade to version 3.14.
ABL80 FLEX BASIC | Cleaned SC with Distilled water and re hydrated again. Then re install SC. Now unit is working fine. | Cleaned SC with Distilled water and re hydrated again. Then re install SC. Now unit is working fine. | SWO-98043 | No | Cleaned SC with Distilled water and re hydrated again. Then re install SC. Now unit is working fine. | Cleaned SC with Distilled water and re hydrated again. Then re install SC. Now unit is working fine.
ABL90 FLEX | Gerät überprüft, Sensorkassette ausgetauscht | Gerät überprüft, Sensorkassette ausgetauscht | SWO-98169 | No | Device checked, sensor cartridge replaced | Device checked, sensor cartridge replaced | Assisted Replacement | SC | SC | Glu
ABL90 FLEX | Beim öffnen der Sensor Bay festgestellt, dass Kontakte für SC90 (von Sensorkassetten Interfaceboatd) schwarz, verschmort, unterschiedlich lang sind, Gerät abgezogen und in GS verbracht, Leihgerät ABL90 Plus platziert, diensthabende Ärzte eine kurze Einweisung gegegeben, Leihgerät arbeitet ohne Probleme. | Der Kunde bittet um einen Service und Rückruf. | SWO-98038 | No | When opening the Sensor Bay it was found that contacts for SC90 (from sensor cassettes Interfaceboatd) are black, shaded, different length, device removed and moved into GS, loaner ABL90 Plus placed, on-duty doctors given a brief briefing, Rental device works without problems. | The customer asks for a service and callback. | Revisit required | Removed by RSUB | Removed Analyzer | Removed Analyzer
ABL90 FLEX PLUS | Contrôle d'appareil dans le menu service : tout les test sont OK Débouchage du système préventive et nettoyage du joint d'intro . Calibrations et CQ : OK | Contrôle d'appareil dans le menu service : tout les test sont OK Débouchage du système préventive et nettoyage du joint d'intro . Calibrations et CQ : OK | SWO-97972 | No | Device control in the service menu: all tests are OK Unblocking the preventative system and cleaning the intro joint. Calibrations and CQ: OK | Device control in the service menu: all tests are OK Unblocking the preventative system and cleaning the intro joint. Calibrations and CQ: OK | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
ABL90 FLEX | 更换了电源模块910-285后修复---jack.chu | 更换了电源模块910-285后修复---jack.chu | SWO-97586 | No | Repair edified after replacing power module 910-285---jack.chu | Repair edified after replacing power module 910-285---jack.chu | Defect | HW | Power Supply | Power Supply
AQT90 FLEX | 更换了孵育盘传感器903-138后修复---jack.chu | 更换了孵育盘传感器903-138后修复---jack.chu | SWO-97839 | No | Repaired after replacing the incubation disk sensor 903-138---jack.chu | Repaired after replacing the incubation disk sensor 903-138---jack.chu | Defect | Hw | Incubator module | PCB, Incubator | 5784
ABL827 FLEX | Replaced the El Met module rinse valve. Lactate membrane also replaced because the calibration sensitivity was below the lower limit. | Replaced the El Met module rinse valve. Lactate membrane also replaced because the calibration sensitivity was below the lower limit. | SWO-98072 | No | Replaced the El Met module rinse valve. Lactate membrane also replaced because the calibration sensitivity was below the lower limit. | Replaced the El Met module rinse valve. Lactate membrane also replaced because the calibration sensitivity was below the lower limit. | Defect | HW | Crea Module | Valve
ABL800 BASIC | Found all sample measurements was unstable for pH, replaced the reference membrane and cleaned the pH electrode and calibrated performed the sample measurement found pH values are ok, found instrument is functional. | Found all sample measurements was unstable for pH, replaced the reference membrane and cleaned the pH electrode and calibrated performed the sample measurement found pH values are ok, found instrument is functional. | SWO-98103 | No | Found all sample measurements was unstable for pH, replaced the reference membrane and cleaned the pH electrode and calibrated performed the sample measurement found pH values are ok, found instrument is functional. | Found all sample measurements was unstable for pH, replaced the reference membrane and cleaned the pH electrode and calibrated performed the sample measurement found pH values are ok, found instrument is functional. | Assisted Replacement | Electrode | Membrane | Reference
ABL800 BASIC | Taken the membranes from the distributor place and replaced the membranes, done the calibration. Run a sample, machine is working fine. | Taken the membranes from the distributor place and replaced the membranes, done the calibration. Run a sample, machine is working fine. | SWO-98083 | No | Taken the membranes from the distributor place and replaced the membranes, done the calibration. Run a sample, machine is working fine. | Taken the membranes from the distributor place and replaced the membranes, done the calibration. Run a sample, machine is working fine. | Assisted Replacement | Electrode | Membrane | pCO2
ABL825 FLEX | Performed qc and followed flow of sample looking for air bubbles at ph slot. | Replaced ph bg pump tubing and tubing between ph bg and ei met module. | D0AAE223DA2EEC9C | No | Performed qc and followed flow of sample looking for air bubbles at ph slot. | Replaced ph bg pump tubing and tubing between ph bg and ei met module. | Defect | HW | pH/BG Module | Pump tubes
AQT90 FLEX | Reagenzienkassette manuell ausgeworfen Cups werden nicht in den Pack abtransportiert, Cups entfernt neuen Kundenreagenzienpack installiert Systemüberprüfung und Reinigung durchgeführt | Reagenzienkassette manuell ausgeworfen Cups werden nicht in den Pack abtransportiert, Cups entfernt neuen Kundenreagenzienpack installiert Systemüberprüfung und Reinigung durchgeführt | SWO-98059 | No | Reagent cartridge manually ejected cups are not transported into the pack, cups removed new customer reagent pack installed system review and cleaning performed | Reagent cartridge manually ejected cups are not transported into the pack, cups removed new customer reagent pack installed system review and cleaning performed | Not Cat | Not cat. | Not cat. | Not cat. | 5784
ABL90 FLEX | Oxi teszt hibás, cserekészülék kihelyezése (S N: 754R2143N001). Műszer elszállítva további tesztekre, javításra | B | SWO-98118 | No | Oxy test Defective, replace replacement (S N: 754R2143N001). Instrument shipped for further testing, repair |  | Defect | HW | Oxi module | Hemolyzer
ABL90 FLEX | 更换测试卡正常。 | 更换试剂包正常。 | SWO-100019 | No | The replacement test card is OK. | The replacement kit is OK. | Unknown | Revisit required | Revisit required | Revisit required
ABL827 FLEX | Customer replaced membrane | Checked and cleaned measuring chamber. Checker electrode connector.Performed zero current procedure in service Replaced glucose electrode | D0AC4D41503F27B1 | No | Customer replaced membrane | Checked and cleaned measuring chamber. Checker electrode connector. Performed zero current procedure in service Replaced glucose electrode | Assisted Replacement | Electrode | Membrane | Glu
ABL800 BASIC | Checked and found Touch Screen intermediately working, need to order for New Touch Screen. | B | SWO-98591 | No | Checked and found Touch Screen intermediately working, need to order for New Touch Screen. |  | Revisit required | HW | DMS_PC | Touch screen
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-98134 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL90 FLEX | 更换试剂包正常。 | 更换试剂包正常。 | SWO-100020 | No | The replacement kit is OK. | The replacement kit is OK. | Assisted Replacement | SP | SP | Oxi QC
ABL825 FLEX | 更换电极945-612后解决---jack.chu | 更换电极945-612后解决---jack.chu | SWO-97947 | No | After replacing the electrode 945-612, solve the ---jack.chu | After replacing the electrode 945-612, solve the ---jack.chu | Assisted Replacement | Electrode | Electrode | pCO2 | 375
ABL80 FLEX CO-OX | Remplacé la valve de l'oxymétrie e.o garantie Remplacé le kit de maintenance e.o garantie Cycle de qualité e.o | Remplacé la valve de l'oxymétrie e.o garantie Remplacé le kit de maintenance e.o garantie Cycle de qualité e.o | SWO-98352 | No | Replaced the oxymetry valve e.o warranty Replaced the maintenance kit e.o guaranteed Cycle of quality e.o | Replaced the oxymetry valve e.o warranty Replaced the maintenance kit e.o guaranteed Cycle of quality e.o
ABL80 FLEX Basic version |  | B | D0AECCB83C2D1AC0 | No
ABL90 | Verified. | Replaced the inlet gasket from the customers stock per the service manual. | D0AE97353C0C9243 | No | Verified. | Replaced the inlet gasket from the customers stock per the service manual. | Defect | HW | Inlet Gasket | Inlet gasket
ABL90 FLEX PLUS |  | B | D0AEF7273C0EF6E2 | No |  |  | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
ABL90 FLEX PLUS |  | B | D0AEF97E3C2BF6E5 | No |  |  | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
ABL835 FLEX | 代替モジュール設置 持つ帰りチェック | 代替モジュール設置 持つ帰りチェック | SWO-99082 | No | Alternative module installation with return check | Alternative module installation with return check | Revisit required | HW | FlexQ Module | FlexQ Module
ABL825 FLEX | Airventil gespült - Einlassbereich demonstriert, überprüft und kalibriert - Kalibrationen - Funktionsüberprüfung etc. | Airventil gespült - Einlassbereich demonstriert, überprüft und kalibriert - Kalibrationen - Funktionsüberprüfung etc. | SWO-98327 | No | Airvalve rinsed - inlet area demonstrated, checked and calibrated - calibrations - function check, etc. | Airvalve rinsed - inlet area demonstrated, checked and calibrated - calibrations - function check, etc. | Cleaned | HW | Fluidic manifold | Air vent
ABL835 FLEX | corrélatin sur HBCO ne fonctionne pas | Effectué calibration hemolyser et thb en mode service | D0AEBDF93C46EA34 | No | correlated on HBCO does not work | Performed hemolyser and thb calibration in service mode | Adjustment | HW | Oxi module | tHb cal
ABL835 FLEX | correlation ne fonctiomme pas Sur HBCO | Calibr hemolyser service THb calibration | D0AEBCA43C45EA31 | No | correlation does not fonctiomme On HBCO | Calibr hemolyser service THb calibration | Adjustment | HW | Oxi module | tHb cal
ABL827 FLEX | Kevin Elliott and I worked with them over the phone and asked them to membrane the REF. The PH and electrolyte sensitivities are still out. Tried REF electrode and connector. Tried Cal 1 and 2 solutions. I went to replace the chamber and found someone had pushed so hard on the pO2 electrode it was crammed up in the pO2 connector.Tried REF electrode and connector and pH BG chamber and still high. Will order parts and return. Amy said they may not want it repaired. I told her I would call be by 4:30. I called Amy and she said they had decided not to repair the analyzer. They are going to connect the replacement 827 and put it in service. She will let me know when to pick this one up. | Pulled pO2 electrode out of the connector and the pH and electrolytes came right back in. Replaced REF connector and pH BG chamber. | D0AE9E9E3C0937C8 | No | Kevin Elliott and I worked with them over the phone and asked them to membrane the REF. The PH and electrolyte sensitivities are still out. Tried REF electrode and connector. Tried Cal 1 and 2 solutions. I went to replace the chamber and found someone had pushed so hard on the pO2 electrode it was crammed up in the pO2 connector. Tried REF electrode and connector and pH BG chamber and still high. Will order parts and return. Amy said they may not want it repaired. I told her I would call be by 4:30. I called Amy and she said they had decided not to repair the analyzer. They are going to connect the replacement 827 and put it in service. She will let me know when to pick this one up. | Pulled pO2 electrode out of the connector and the pH and electrolytes came right back in. Replaced REF connector and pH BG chamber. | Revisit required | HW | pH/BG Module | Electrode Connector
ABL810 FLEX | PCO2 Elektrode ersetzt. PCO2 Membrane mit andere Lot aufgezogen, Elektroden Kontakte gereinigt. pO 2 und Ref. Membraniert. Gerät Einkalibriert. | PCO2 Elektrode ersetzt. PCO2 Membrane mit andere Lot aufgezogen, Elektroden Kontakte gereinigt. pO 2 und Ref. Membraniert. Gerät Einkalibriert. | SWO-98377 | No | PCO2 electrode replaced. PCO2 membrane with other solder mounted, electrode contacts cleaned. pO 2 and Ref. Membraneed. Device calibrated. | PCO2 electrode replaced. PCO2 membrane with other solder mounted, electrode contacts cleaned. pO 2 and Ref. Membraneed. Device calibrated. | Assisted Replacement | Electrode | Electrode | pCO2
ABL815 FLEX | Meßkammer überprüft, Ref. Elektrode neu membraniert, Kal1 und 2 Lösungen ersetzt, Kal, QC, Key`s eingespielt, Bed.Hb geliefert | Meßkammer überprüft, Ref. Elektrode neu membraniert, Kal1 und 2 Lösungen ersetzt, Kal, QC, Key`s eingespielt, Bed.Hb geliefert | SWO-98298 | No | Measuring chamber checked, ref. Electrode re-membraneed, Kal1 and 2 solutions replaced, Kal, QC, Key's recorded, Bed.Hb supplied | Measuring chamber checked, ref. Electrode re-membraneed, Kal1 and 2 solutions replaced, Kal, QC, Key's recorded, Bed.Hb supplied | Setup | SW | Demo Key Exp | Demo Key Exp
ABL815 FLEX | PC-Modul sperrt an vorderer Kante, mit Hebel herausgezogen. Abdeckung DMS-Cover mit Schrauben aus Altgerät befestigt. Stromvesorgungskabel AU-Platine korrekt befestigt. Einlassnadel erneuert. | PC-Modul sperrt an vorderer Kante, mit Hebel herausgezogen. Abdeckung DMS-Cover mit Schrauben aus Altgerät befestigt. Stromvesorgungskabel AU-Platine korrekt befestigt. Einlassnadel erneuert. | SWO-98266 | Yes | PC module locks at the front edge, pulled out with lever. Cover DMS-Cover fastened with screws made of old equipment. Power supply cable AU board correctly attached. Inlet needle renewed. | PC module locks at the front edge, pulled out with lever. Cover DMS-Cover fastened with screws made of old equipment. Power supply cable AU board correctly attached. Inlet needle renewed. | Re classified RMED | OOB
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-99182 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL80 FLEX BASIC | 更换测试卡正常。 | 更换测试卡正常。 | SWO-98688 | No | The replacement test card is OK. | The replacement test card is OK.
ABL80 FLEX BASIC | 更换新测试卡修复仪器 | 更换新测试卡修复仪器 | SWO-99122 | No | Replace the new test card repair instrument | Replace the new test card repair instrument
ABL90 |  | B | D0B7732B76097271 | No |  |  | Assisted Replacement | SC | SC | Reference
ABL90 | I went to fedex, but part would not arrive until 11am I tried to install a new CF card and reprogram, but it did not work Had to go back to Fedex and pick up the new CPU. | Installed the new CPU Downgrade to SYS8 Installed Windows Reprogramed with 3.1 MR8. | D0B003B9B2214931 | No | I went to fedex, but part would not arrive until 11am I tried to install a new CF card and reprogram, but it did not work Had to go back to Fedex and pick up the new CPU. | Installed the new CPU Downgrade to SYS8 Installed Windows Reprogramed with 3.1 MR8. | Defect | HW | CPU | Reboot
ABL90 FLEX | 做了血红蛋白定标。 | 做了血红蛋白定标。 | SWO-98415 | No | Hemoglobin is calibrated. | Hemoglobin is calibrated. | Adjustment | HW | Oxi module | tHb Cal
ABL825 FLEX | Testweise das mainboard getauscht. Keine Besserung Festplatte kann nicht gestartet werden. Netzteil getauscht. Danach wurde Festplatte wieder ordnungsgemäß hochgefahren und Analysator könnte starten. Kalibrierungen durchgeführt. QCs für alle Level durchgeführt. --> i.O. ABL wieder Messbereit. | Testweise das mainboard getauscht. Keine Besserung Festplatte kann nicht gestartet werden. Netzteil getauscht. Danach wurde Festplatte wieder ordnungsgemäß hochgefahren und Analysator könnte starten. Kalibrierungen durchgeführt. QCs für alle Level durchgeführt. --> i.O. ABL wieder Messbereit. | SWO-98558 | No | Test-wise the mainboard was replaced. No improvement hard drive can not be started. power supply is replaced. After that, hard drive was booted back up properly and analyzer could start. calibrations. QCs for all levels. --> i.O. ABL ready to measure again. | Test-wise the mainboard was replaced. No improvement hard drive can not be started. power supply is replaced. After that, hard drive was booted back up properly and analyzer could start. calibrations. QCs for all levels. --> i.O. ABL ready to measure again. | Defect | HW | Power Supply | HDD down
ABL815 FLEX | Leaking at inlet area. | Replaced inlet probe Gasket | D0B04D87B230A635 | No | Leaking at inlet area. | Replaced inlet probe Gasket | Defect | HW | Inlet module | Probe cal.
ABL825 FLEX | Error code 963-Leakage currents detected. Inlet liquid sensor error - 529 538.539.and 541 -Pump errors. | B | D0B0509BB24FA63A | No | Error code 963-Leakage currents detected. Inlet liquid sensor error - 529 538.539.and 541 -Pump errors. |  | Unknown | HW | Unknown | Leak Current
ABL815 FLEX | Lampe neu justiert. | Lampe neu justiert. | SWO-98583 | No | Lamp re-adjusted. | Lamp re-adjusted. | Adjustment | HW | Oxi module | Lamp unit | High
TCM-COMBIM |  | B | D0AFF362B2211AC4 | No |  |  | Defect | HW | PCB, Main CPU | PCB, Main CPU
ABL80 FLEX BASIC | 更换新测试卡修复仪器 | 更换新测试卡修复仪器 | SWO-98995 | No | Replace the new test card repair instrument | Replace the new test card repair instrument
ABL80 FLEX BASIC | 更换两张测试卡正常。 | 更换两张测试卡正常。 | SWO-99231 | No | Replacing the two test cards is ok. | Replacing the two test cards is ok.
ABL80 FLEX BASIC | 更换测试卡正常。 | 更换测试卡正常。 | SWO-98709 | No | The replacement test card is OK. | The replacement test card is OK.
ABL90 | Needs software patch installed. | Installed Ralls software patch. | D0B149F92860D219 | No | Needs software patch installed. | Installed Ralls software patch. | Setup | SW | His/lis | His/Lis
ABL90 | Needs Ralls patch installed. | Installed Ralls patch | D0B25C569E196D67 | No | Needs Ralls patch installed. | Installed Ralls patch | Setup | SW | His/lis | His/Lis
ABL90 |  | B | D0B11AB0283BC677 | No |  |  | Setup | SW | DMS_PC | Barcode Reader
ABL90 |  | B | D0B11E47285F1376 | Yes |  |  | Re classified RMED | Installation - Demo
ABL90 FLEX PLUS | Gewünschten unteren pH-Messbereich von 6,75 auf 6,50 eingestellt | Gewünschten unteren pH-Messbereich von 6,75 auf 6,50 eingestellt | SWO-98654 | No | Desired lower pH measuring range set from 6.75 to 6.50 | Desired lower pH measuring range set from 6.75 to 6.50 | Setup | SW | Reportable range | pH
ABL90 FLEX | 更换新测试卡修复仪器 | 更换新测试卡修复仪器 | SWO-99029 | No | Replace the new test card repair instrument | Replace the new test card repair instrument | Assisted Replacement | SC | SC | pCO2 | 378
ABL90 FLEX | 更换测试卡正常。 | 更换测试卡正常。 | SWO-98750 | No | The replacement test card is OK. | The replacement test card is OK. | Assisted Replacement | SC | SC | Reference | 1361
ABL90 FLEX | 更换试剂包正常。 | 更换试剂包正常。 | SWO-99388 | No | The replacement kit is OK. | The replacement kit is OK. | User error | SP Activation | SP Activation | SP Activation
AQT90 FLEX | 更换903-210一个正常。 | 更换903-210一个正常。 | SWO-98476 | No | Replace the 903-210 one normal. | Replace the 903-210 one normal. | Defect | HW | Process Module | PCB, Cup dryer | 17
AQT90 FLEX | REVISIÓN DE HISTÓRICOS. DESMONTAJE Y LIMPIEZA DEL LECTOR DE CÓDIGO DE BARAS DE LA RUEDA DE REACTIVOS. AJUSTES DEL LECTOR. LIMPIEZA GENERAL DEL EQUIPO: SENSORES ÓPTICOS, YMECANISMOS. TESTS DE MENU DE SERVICE. RELIZAR LIMPIEZA CON CLEANING SOLUTION. TEST DE LECTURA DEL LECTOR DE PAQUETES DE REACTIVOS. CONTROLES | REVISIÓN DE HISTÓRICOS. DESMONTAJE Y LIMPIEZA DEL LECTOR DE CÓDIGO DE BARAS DE LA RUEDA DE REACTIVOS. AJUSTES DEL LECTOR. LIMPIEZA GENERAL DEL EQUIPO: SENSORES ÓPTICOS, YMECANISMOS. TESTS DE MENU DE SERVICE. RELIZAR LIMPIEZA CON CLEANING SOLUTION. TEST DE LECTURA DEL LECTOR DE PAQUETES DE REACTIVOS. | D0B2AEFB9E5190E3 | No | HISTORICAL REVIEW. DISASSEMBLY AND CLEANING OF THE CODE READER OF THE REACTIVE WHEEL. READER SETTINGS. GENERAL CLEANING OF THE EQUIPMENT: OPTICAL SENSORS, YMECHANISMS. SERVICE MENU TESTS. RELIZE CLEANING WITH CLEANING SOLUTION. REAGENT PACKET READER READING TEST. Controls | HISTORICAL REVIEW. DISASSEMBLY AND CLEANING OF THE CODE READER OF THE REACTIVE WHEEL. READER SETTINGS. GENERAL CLEANING OF THE EQUIPMENT: OPTICAL SENSORS, YMECHANISMS. SERVICE MENU TESTS. RELIZE CLEANING WITH CLEANING SOLUTION. REAGENT PACKET READER READING TEST. | Cleaned | HW | Storage Module | Barcode reader, Storage wheel
ABL825 FLEX |  | B | D0B139EA2836364C | No |  |  | Assisted Replacement | Electrode | Electrode | pO2
ABL835 FLEX | On inspecting logs found that since problem was reported 3 5 16 that 14 samples had gone through ok and given pH results. Carry out a further 2 protein removals. Lactate needs re membraning, no membranes available, informed the lab that it needed doing. | On inspecting logs found that since problem was reported 3 5 16 that 14 samples had gone through ok and given pH results. Carry out a further 2 protein removals. Lactate needs re membraning, no membranes available, informed the lab that it needed doing. | SWO-98961 | No | On inspecting logs found that since problem was reported 3 5 16 that 14 samples had gone through ok and given pH results. Carry out a further 2 protein removals. Lactate needs re membraning, no membranes available, informed the lab that it needed doing. | On inspecting logs found that since problem was reported 3 5 16 that 14 samples had gone through ok and given pH results. Carry out a further 2 protein removals. Lactate needs re membraning, no membranes available, informed the lab that it needed doing. | Cleaned | HW | pH/BG Module | pH
ABL825 FLEX | Display ersetzt. | B | SWO-99600 | No | Display replaced. |  | Defect | HW | DMS_PC | Touch screen
ABL837 FLEX |  | B | D0B1483F28402567 | No |  |  | Setup | SW | Settings | Printer
ABL825 FLEX | Im Servicemenu verschiedene Tests durchgeführt. Pumpenrotoren überprüft und gereinigt. Beide Messkammern ersetzt. Mehrere Kapillarmessungen mit Flexmodus durchgeführt, alle i.o. Kal., QC und System i.o. Funktionskontrolle i.o. | Im Servicemenu verschiedene Tests durchgeführt. Pumpenrotoren überprüft und gereinigt. Beide Messkammern ersetzt. Mehrere Kapillarmessungen mit Flexmodus durchgeführt, alle i.o. Kal., QC und System i.o. Funktionskontrolle i.o. | SWO-99247 | No | Various tests are carried out in the service menu. Pump rotors checked and cleaned. Replaced both measuring chambers. Several capillary measurements performed with flex mode, all i.o. Cal., QC and System i.o. function control i.o. | Various tests are carried out in the service menu. Pump rotors checked and cleaned. Replaced both measuring chambers. Several capillary measurements performed with flex mode, all i.o. Cal., QC and System i.o. function control i.o. | Defect | SW | Flex Mode | Response Error | Electrolyte
ABL835 FLEX |  | B | D0B1448828622566 | No |  |  | Defect | HW | Inlet module | Probe cal.
AQT90 FLEX | Fehlersuche, Systemreinigung, Test aller Komponenten, Analysatorabsturz während einer Kalibrierung, Reparaturersatzgerät aufgestellt | Vor Ort Service wird gewünscht | SWO-98726 | No | Troubleshooting, system cleaning, testing of all components, analyzer crash during calibration, repair replacement device set up | On-site service is desired | Defect | SW | Analyzer Control Module | System Cleaning | 3264
AQT90 FLEX | PISA hat sich im Serviceprogramm selbsttätig herunter gefahren, während des Neustart war im Bereich Provess-Wheel Punsh-Down zum Solution Pack ein Knack -Geräusch zu hören gewesen. Nach dem Neustart kein Fehler mehr. | PISA hat sich im Serviceprogramm selbsttätig herunter gefahren, während des Neustart war im Bereich Provess-Wheel Punsh-Down zum Solution Pack ein Knack -Geräusch zu hören gewesen. Nach dem Neustart kein Fehler mehr. | SWO-98862 | No | PISA has automatically shut down in the service program, during the reboot there was a cracking noise in the Provess Wheel Punsh-Down area to the Solution Pack. No more errors after the restart. | PISA has automatically shut down in the service program, during the reboot there was a cracking noise in the Provess Wheel Punsh-Down area to the Solution Pack. No more errors after the restart. | Defect | Consumables | SP | SP | 5684
AQT90 FLEX | REVISIÓN DE HISTÓRICOS. LIMPIEZA Y ACONDICIONAMIENTO DE LA RUEDA DE INCUBACIÓN Y RACK DE SENSORES ACOSIADOS. SUSTITUIR TUBOS, FILTROS Y CABEZALES DE BOMBAS PERISTÁLTICAS. LIMPIEZA DE BOMBAS, SENSORES ÓPTICOS Y LECTORES DE CÓDIGOS DE BARRAS. TEST DE MENU SERVICE. LIMPIEZA CON CLEANING SOLUTION. CALIBRACIONES Y CONTROLES | REVISIÓN DE HISTÓRICOS. LIMPIEZA Y ACONDICIONAMIENTO DE LA RUEDA DE INCUBACIÓN Y RACK DE SENSORES ACOSIADOS. SUSTITUIR TUBOS, FILTROS Y CABEZALES DE BOMBAS PERISTÁLTICAS. LIMPIEZA DE BOMBAS, SENSORES ÓPTICOS Y LECTORES DE CÓDIGOS DE BARRAS. TEST DE MENU SERVICE. LIMPIEZA CON CLEANING SOLUTION. | D0B2651A9E52AC7A | No | HISTORICAL REVIEW. CLEANING AND CONDITIONING OF THE INCUBATION WHEEL AND RACK OF ACOSIADOS SENSORS. SUBSTITUTE TUBES, FILTERS AND HEADS OF PERISTALTIC PUMPS. CLEANING OF PUMPS, OPTICAL SENSORS AND BARS CODE READERS. MENU SERVICE TEST. CLEANING SOLUTION CLEANING. CALIBRATIONS AND CONTROLS | HISTORICAL REVIEW. CLEANING AND CONDITIONING OF THE INCUBATION WHEEL AND RACK OF ACOSIADOS SENSORS. SUBSTITUTE TUBES, FILTERS AND HEADS OF PERISTALTIC PUMPS. CLEANING OF PUMPS, OPTICAL SENSORS AND BARS CODE READERS. MENU SERVICE TEST. CLEANING SOLUTION CLEANING. | Cleaned | HW | Incubator module | Sensor, Incubator
AQT90 FLEX | Gerät mit Kundin über Telefon neu gestartet, nach Rückruf zusammen mit Kundin versucht Testkassetten einzulegen, Meldungen: Strichcode nicht erkannt, nach Ankunft am Gerät mehrere Testkassetten herausgenommen und wieder eingelegt, auch Testkassetten, welche vorher nicht erkannt wurden, hierbei gab es keinerlei Probleme, AQT zeigt jetzt wieder eine grüne Ampel und alle Bestände sind im Normbereich, Übertragung der QC Ergebnisse im Labor überprüft - Übertragung funktioniert, System- und Funktionskontrolle in Ordnung. | der Anwender hat das Einlegen neuer Testkassetten ins Storage Wheel mehrfach bversucht, allerdings ohne Erfolg; Scanner-Funktion Strichcode-Etiketten; der Anwender wünscht einen Vorort-Service zur Beseitigung der Geräte-Störung; | SWO-99154 | Yes | Device with customer restarted via telephone, after callback together with customer tries to insert test cassettes, messages: bar code not recognized, after arrival at the device several test cassettes removed and inserted again, also test cassettes, which were not previously There were no problems, AQT now shows a green traffic light again and all stocks are in the normal range, transmission of QC results checked in the laboratory - transmission works, system and function control fine. | the user has tried to insert new test cartridges into the storage wheel several times, but to no avail; Scanner function bar code labels; the user wants a suburban service to eliminate the device malfunction; | Re classified RMED | Phone support
TCM TOSCA | Tosca modul udskiftet. Det nye hvæser også! Techlinesag! | B | SWO-98881 | No | Tosca module replaced. The new hisses too! Techline Case! |  | Defect | HW | Tosca module |  | Techline
TCM TOSCA | Tosca modul udskiftet, slutafprøvet. | Tosca modul udskiftet, slutafprøvet. | SWO-98877 | No | Tosca module replaced, final tested. | Tosca module replaced, final tested. | Defect | HW | Tosca module |  | Techline
ABL80 FLEX BASIC | 更换测试卡修复仪器 | 更换测试卡修复仪器 | SWO-99148 | No | Replace the test card repair instrument | Replace the test card repair instrument
ABL80 FLEX BASIC | 更换新测试卡修复仪器 | 更换新测试卡修复仪器 | SWO-99208 | No | Replace the new test card repair instrument | Replace the new test card repair instrument
ABL80 FLEX BASIC | 检查902-972，发现有损坏的痕迹，更换902-927修复仪器 by jesson | 检查902-972，发现有损坏的痕迹，更换902-927修复仪器 by jesson | SWO-98740 | No | Check 902-972 and find signs of damage, replace the 902-927 repair instrument by jesson | Check 902-972 and find signs of damage, replace the 902-927 repair instrument by jesson
ABL80 FLEX BASIC | 更换测试卡正常。 | 更换测试卡正常。 | SWO-99330 | No | The replacement test card is OK. | The replacement test card is OK.
ABL80 FLEX BASIC | 更换测试卡正常 | 更换测试卡正常 | SWO-99329 | No | Replace the test card ok | Replace the test card ok
ABL80 FLEX BASIC | Cleaned the waste connector & checked equipment is functional | Cleaned the waste connector & checked equipment is functional | SWO-99376 | No | Cleaned the waste connector & checked equipment is functional | Cleaned the waste connector & checked equipment is functional
ABL827 FLEX | Fitted another EL MET module to see if the lactate calibrations would be any better. New measuring chambers fitted to this module because the ones on there were broken. Swapped over the pump tube from the original module as this was newer. Calibrations & QCs would not pass with this module so the original one was refitted. Another module will be fitted to a working analyser and tested over a few days to ensure it is ok and then will be fitted to this analyser. | Fitted another EL MET module to see if the lactate calibrations would be any better. New measuring chambers fitted to this module because the ones on there were broken. Swapped over the pump tube from the original module as this was newer. Calibrations & QCs would not pass with this module so the original one was refitted. Another module will be fitted to a working analyser and tested over a few days to ensure it is ok and then will be fitted to this analyser. | SWO-99063 | No | Fitted another EL MET module to see if the lactate calibrations would be any better. New measuring chambers fitted to this module because the ones on there were broken. Swapped over the pump tube from the original module as this was newer. Calibrations & QCs would not pass with this module so the original one was refitted. Another module will be fitted to a working analyser and tested over a few days to ensure it is ok and then will be fitted to this analyser. | Fitted another EL MET module to see if the lactate calibrations would be any better. New measuring chambers fitted to this module because the ones on there were broken. Swapped over the pump tube from the original module as this was newer. Calibrations & QCs would not pass with this module so the original one was refitted. Another module will be fitted to a working analyser and tested over a few days to ensure it is ok and then will be fitted to this analyser. | Assisted Replacement | Electrode | Membrane | Lac | Calibration
ABL827 FLEX | Fit new reference electrode. | Fit new reference electrode. | SWO-99058 | No | Fit new reference electrode. | Fit new reference electrode. | Assisted Replacement | Electrode | Electrode | Reference
ABL837 FLEX | Fit new reference electrode | Fit new reference electrode | SWO-99061 | No | Fit new reference electrode | Fit new reference electrode | Assisted Replacement | Electrode | Electrode | Reference
ABL835 FLEX | Fit new reference electrode | Fit new reference electrode | SWO-99062 | No | Fit new reference electrode | Fit new reference electrode | Assisted Replacement | Electrode | Electrode | Reference
ABL837 FLEX | インレットプローブ交換 CAL確認 | インレットプローブ交換 CAL確認 | SWO-99079 | No | Inlet Probe Replacement CAL Confirmation | Inlet Probe Replacement CAL Confirmation | Defect | HW | Inlet module | Probe bend | State inlet probe is the hold state hold bend on the lower side, the probe is flying out not completely return gasket // the inlet probe is bent
ABL825 FLEX | Ran patient sample. | Installed a patch in the roma.ini file. | D0B25D179E526D6A | No | Ran patient sample. | Installed a patch in the roma.ini file. | Setup | SW | His/lis | His/Lis
ABL827 FLEX | Switched carousels-No chnage | Replaced auto check flex cable | D0B79EA97663776C | No | Switched carousels-No chnage | Replaced auto check flex cable | Defect | HW | AutoCheck Module | FPC Cable
ABL800 BASIC | Shut down the machine for half an hour, restarted the machine, all the problems got resolved. Meet Dr. Khedkar to discuss the issue, someone from the staff switched OFF the main supply, because of which all the oxymetry parameters failed. | Shut down the machine for half an hour, restarted the machine, all the problems got resolved. Meet Dr. Khedkar to discuss the issue, someone from the staff switched OFF the main supply, because of which all the oxymetry parameters failed. | SWO-99341 | No | Shut down the machine for half an hour, restarted the machine, all the problems got resolved. Meet Dr. Khedkar to discuss the issue, someone from the staff switched OFF the main supply, because of which all the oxymetry parameters failed. | Shut down the machine for half an hour, restarted the machine, all the problems got resolved. Meet Dr. Khedkar to discuss the issue, someone from the staff switched OFF the main supply, because of which all the oxymetry parameters failed. | Unknown | HW | Oxi module | Neon Lamp, low
ABL80 FLEX BASIC | 更换测试卡修复仪器 | 更换测试卡修复仪器 | SWO-99185 | No | Replace the test card repair instrument | Replace the test card repair instrument
ABL80 CO-OX | user keeps getting fluid blockage error. Found that no fluid can move through the hemolyzer. | Replaced hemolyzers and all tubing to make sure it is all clear. | D0B3A446145165AE | No | user keeps getting fluid blockage error. Found that no fluid can move through the hemolyzer. | Replaced hemolyzers and all tubing to make sure it is all clear.
ABL80 FLEX BASIC | 故障代码1005 ，机器不能连接I-button，更换903-228一个正常。By Jesson | 故障代码1005 ，机器不能连接I-button，更换903-228一个正常。By Jesson | SWO-98999 | No | Fault code 1005, the machine can not connect to I-button, replace 903-228 a normal. By Jesson | Fault code 1005, the machine can not connect to I-button, replace 903-228 a normal. By Jesson
ABL90 FLEX | 内部洗浄。 | 内部洗浄。 | SWO-99352 | No | Internal cleaning. | Internal cleaning. | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
ABL90 FLEX | Műszer visszaszállítása és üzembehelyezése, Kalibrációk, QC mérések | B | SWO-99207 | Yes | Instrument return and commissioning, calibration, QC measurements |  | Re classified RMED | LEH Mail
ABL90 FLEX PLUS |  | B | D0B3D5C51411E891 | Yes |  |  | Re classified RMED | double entry
ABL90 FLEX | 连续更换四张测试卡后修复 | 连续更换四张测试卡后修复 | SWO-99190 | No | Repair after four consecutive test cards | Repair after four consecutive test cards | Assisted Replacement | SC | SC | Na
AQT90 FLEX | 进样门没有正常关闭，检查进样门 机械部分，发现进样门没有接触到微动开关，重新安装进样传感器修复仪器 by jesson | 进样门没有正常关闭，检查进样门 机械部分，发现进样门没有接触到微动开关，重新安装进样传感器修复仪器 by jesson | SWO-99131 | No | The sample door is not closed properly, check the mechanical part of the sample door, find that the sample door is not in contact with the microswitch, reinstall the sample sensor repair instrument by jesson | The sample door is not closed properly, check the mechanical part of the sample door, find that the sample door is not in contact with the microswitch, reinstall the sample sensor repair instrument by jesson | Adjustment | HW | Cover and Chassis | Inlet Door | 3627
ABL80 | Chequeo de equipo observando que está todo en verde y que todas las calibraciones y ciclos de calidad estan OK. chequeo de histórico de eventos viendo que en la instalación del SC tuvieron errores multiparamétricos en la inicialización de éste y que tras varias cal y ciclos de calidad al cabo de 2 horas se puso todo OK. Chequeo de historico de muestras viendo que 4 o 5 muestras en las semanas 16 17 salieron con !!! y mostrando el error posible aire en la muestra. Se observa que el SC que retirarón gastado tenía la sonda ligeramente doblada. | chequeo de reactivos, historico de datos, realización de calibración, ciclos de calidad | D0B3BB831416D86D | No | Checking equipment to see that everything is green and that all calibrations and quality cycles are OK. checking of historical events showing that in the installation of the SC they had multiparametric errors in the initialization of this and that after several lime and quality cycles after 2 hours everything was set OK. Checking out historical samples showing that 4 or 5 samples in weeks 16 17 came out with !!! and showing the possible air error in the sample. It is observed that the spent SC had the probe slightly bent. | reagent checking, data history, calibration, quality cycles
ABL835 FLEX | インレットプローブ曲がりの為、交換。 FLEXQポジション調整、LS調整実施。 | インレットプローブ曲がりの為、交換。 FLEXQポジション調整、LS調整実施。 | SWO-99107 | No | Replace the inlet probe for bending. FlexQ position adjustment and LS adjustment. | Replace the inlet probe for bending. FlexQ position adjustment and LS adjustment. | Adjustment | HW | FlexQ Module | Position Calibration
ABL825 FLEX | Tried to power back on with no help, checked power supply voltages and they were alll good. | Installed new CPU. | D0B50AB18A154B9B | No | Tried to power back on with no help, checked power supply voltages and they were alll good. | Installed new CPU. | Defect | HW | DMS_PC | CPU
ABL800 BASIC | Last week we have performed the Lac cleaning with protein removal but problem was remaining same, found sensitivity error and drift error, replaced the standby Lac electrode and with new membrane calibrated found calibration was OK, it was under observation. | Last week we have performed the Lac cleaning with protein removal but problem was remaining same, found sensitivity error and drift error, replaced the standby Lac electrode and with new membrane calibrated found calibration was OK, it was under observation. | SWO-99317 | No | Last week we have performed the Lac cleaning with protein removal but problem was remaining same, found sensitivity error and drift error, replaced the standby Lac electrode and with new membrane calibrated found calibration was OK, it was under observation. | Last week we have performed the Lac cleaning with protein removal but problem was remaining same, found sensitivity error and drift error, replaced the standby Lac electrode and with new membrane calibrated found calibration was OK, it was under observation. | Assisted Replacement | Electrode | Membrane | Lac | Calibration
AQT90 FLEX | Vervollständigung der Konfiguration Kalibrierung, Funktionskontrolle | Vor Ort Service wird gewünscht | SWO-99136 | Yes | Completion of configuration calibration, function control | On-site service is desired | Re classified RMED | Phone support
AQT90 FLEX | 开机报错，软件不匹配，更换903-198后，重新安装windows ，应用软件后，需要重新定标。 By jesson | 重新定标正常。 | SWO-98414 | No | Boot error, software does not match, replace 903-198, reinstall windows, application software, need to re-label. By Jesson | Recalibration is normal. | Defect | HW | Analyzer Control Module | PCB, Main board
AQT90 FLEX |  | B | SWO-99384 | No |  |  | Defect | HW | Incubator module | PCB, Incubator | 2168
AQT90 FLEX | 重新定标正常。 | 重新定标正常。 | SWO-99346 | No | Recalibration is normal. | Recalibration is normal. | Revisit required | Consumables | Test Kit | Test Kit
ABL80 FLEX BASIC |  | B | SWO-99420 | No
ABL80 FLEX CO-OX | Seit letztem Service Na QC Werte knapp über ZW im Level2, K QC In Level 1 aber an oberen Grenze, so2 QC Werte grenzwertig, Zero Calibrierung lt. RDEU Vorgaben durchgeführt, alle Kanäle auf 0 gesetzt, danach LP leer-> ersetzt, danach sO2 QC Fehler im Level 2, thb Kal durchgeführt,danach sO2 QC wieder i.O., K QC im Level 1 immer noch im oberen Bereichen -> Info an Anwender | Seit letztem Service Na QC Werte knapp über ZW im Level2, K QC In Level 1 aber an oberen Grenze, so2 QC Werte grenzwertig, Zero Calibrierung lt. RDEU Vorgaben durchgeführt, alle Kanäle auf 0 gesetzt, danach LP leer-> ersetzt, danach sO2 QC Fehler im Level 2, thb Kal durchgeführt,danach sO2 QC wieder i.O., K QC im Level 1 immer noch im oberen Bereichen -> Info an Anwender | SWO-99507 | No | Since last service Na QC values just above ZW in level2, K QC In level 1 but at upper limit, so2 QC values borderline, zero calibration carried out according to RDEU specifications, all channels set to 0, then LP empty-> replaced, then sO2 QC error in level 2 , thb Kal performed, then sO2 QC again i.O., K QC in level 1 still in the upper range -> Info to users | Since last service Na QC values just above ZW in level2, K QC In level 1 but at upper limit, so2 QC values borderline, zero calibration carried out according to RDEU specifications, all channels set to 0, then LP empty-> replaced, then sO2 QC error in level 2 , thb Kal performed, then sO2 QC again i.O., K QC in level 1 still in the upper range -> Info to users
ABL80 OSM | Veified the problem by talking with thecustomer. Cusyomer was able to push the information to the analyzer through Aqure. | Replaced the CF card and solution pack. | D0B785D1765AEE6B | No | Veified the problem by talking with thecustomer. Cusyomer was able to push the information to the analyzer through Aqure. | Replaced the CF card and solution pack.
ABL80 FLEX BASIC | 更换测试卡正常。 | 更换测试卡正常。 | SWO-101790 | No | The replacement test card is OK. | The replacement test card is OK.
ABL90 FLEX | Folyadékrendszer többszöri öblítése,-tisztítása, Tűtartó csere, OXI tesztek, Szervíz tesztek, SC csere (790-93), Kalibrálás, QC mérés | B | SWO-99559 | No | Repeated rinsing, cleaning of the liquid system, needle holder replacement, OXY tests, service tests, SC replacement (790-93), calibration, QC measurement |  | Replaced | HW | Inlet | Probe
ABL90 FLEX | Reinstalled Windows xp to analyser and then reinstalled application software version mr8. Ran all tests and checks once completed and recalibrated touch screen. Customer to monitor | Reinstalled Windows xp to analyser and then reinstalled application software version mr8. Ran all tests and checks once completed and recalibrated touch screen. Customer to monitor | SWO-99662 | No | Reinstalled Windows xp to analyser and then reinstalled application software version mr8. Ran all tests and checks once completed and recalibrated touch screen. Customer to monitor | Reinstalled Windows xp to analyser and then reinstalled application software version mr8. Ran all tests and checks once completed and recalibrated touch screen. Customer to monitor | Defect | SW | win XP | Low on disc Space
ABL90 FLEX | Gebrauchtes LP reaktiviert, Funktion OK. | Gebrauchtes LP reaktiviert, Funktion OK. | SWO-99711 | No | Used LP reactivated, function OK. | Used LP reactivated, function OK. | User error | SP Activation | SP Activation | SP Activation
ABL90 | Ran THB calibrations, 2pt calibration, qc and blood samples and all looks good. | There was no correction needed, just testing. | D0B763A57619F2F2 | No | Ran THB calibrations, 2pt calibration, qc and blood samples and all looks good. | There was no correction needed, just testing. | Adjustment | HW | Oxi module | tHb Cal
ABL90 FLEX PLUS |  | B | D0BA08AC623B7F80 | Yes |  |  | Re classified RMED | PER
ABL90 FLEX PLUS |  | B | D0B79AC5760F2B26 | Yes |  |  | Re classified RMED | Logistical PER
ABL90 FLEX | 更换测试卡修复仪器 | 更换测试卡修复仪器 | SWO-100870 | No | Replace the test card repair instrument | Replace the test card repair instrument
ABL90 FLEX |  | B | SWO-99419 | No |  |  | Defect | HW | Sensor Interface PCB | Temperature | (Return Request - 16.5-16)
AQT90 FLEX |  | B | SWO-99491 | No |  |  | Defect | HW | Storage Module | Barcode reader, Storage wheel
AQT90 FLEX | 更换903-329正常（by zhenjie) | 更换903-329正常（by zhenjie) | SWO-99336 | No | Replace 903-329 Normal (by zhenjie) | Replace 903-329 Normal (by zhenjie) | Defect | HW | Aspiration Module | Needle Wash Unit | 1223
ABL825 FLEX | CPU gele tres souvent et maintenance préventive due | Effecué la maintenance bi-annuel selon les normes remplacé les tubes de pompe nettoyé pompes et flush v7.v8 verifier electrode o'ringfull oxi test et calibration service THB calibration nettoyé les pompes vérifié electronique calibration zero po2 ,lactate.glucose . nettoyé autocheck laser. install nouveau cpu et installé clé permanante | D0B7784D765E8A9C | No | CPU very often geles and preventative maintenance due | Effecué bi-annual maintenance according to standards replaced pump tubes cleaned pumps and flush v7.v8 check electronico'ringfull oxi test and calibration service THB calibration cleaned pumps verified electronic calibration zero po2, lactate.glucose . cleaned laser autocheck. installs new cpu and installed permanent key | Defect | HW | DMS_PC | CPU
ABL825 FLEX | Needs critical limits patch installed. | Installed critical limits patch. | D0B78701761DEB55 | Yes | Needs critical limits patch installed. | Installed critical limits patch. | Re classified RMED | Planned activity
ABL825 FLEX | Needs critical limits patch installed. | Installed critical limits patch. | D0B78802762CEB58 | Yes | Needs critical limits patch installed. | Installed critical limits patch. | Re classified RMED | Planned activity
ABL825 FLEX | nach erneutem Auftreten d .Fehlers Display wieder zurückgetauscht. Display Kabel ersetzt. z.Test. | Neue PC Unit eingesetzt. Kal und QC nach Einlaufphase ok. Zur weiteren Beobachtung | SWO-99715 | No | after recurrence d . Error display replaced again. Display cable replaced. e.g.Test. | New PC Unit. Kal and QC after run-in phase ok. For further observation | Defect | HW | DMS_PC | CPU | Touch screen
ABL810 FLEX | BG Messkammer gereinigt, Membraniert. | Tausch des pH BG Moduls. LS Justage, Pumpenkal und Nullpunktkalibrierung durchgeführt. Kalibrationen und Qualitätskontrolle o. K. Gemessene Werte sind nicht zielwertig, aber innerhalb der Grenzwerte. | SWO-99601 | No | BG measuring chamber cleaned, membraneed. | Replacement of the pH BG module. LS adjustment, pump scale and zero-point calibration. Calibrations and quality control o.K. Measured values are not target values, but within the limits. | Defect | HW | pH/BG Module | pCO2 & pO2
ABL810 FLEX | Neue IP Adressen eingegeben. Jeweils eine AQC Messung und Patientenmessung durchgeführt, Hr. Corte hat mir die Funktion bestätigt. | Neue IP Adressen eingegeben. Jeweils eine AQC Messung und Patientenmessung durchgeführt, Hr. Corte hat mir die Funktion bestätigt. | SWO-99599 | No | New IP addresses entered. An AQC measurement and patient measurement was performed, Mr. Corte confirmed the function to me. | New IP addresses entered. An AQC measurement and patient measurement was performed, Mr. Corte confirmed the function to me. | Setup | Sw | His/lis | Conworx
ABL810 FLEX | Die Conworx Konfiguration eingegeben. Zur Kontrolle ein AQC und Patientenmessung durchgeführt. Hr.Corte hat mir die Funktion telefonisch bestätigt. | Die Conworx Konfiguration eingegeben. Zur Kontrolle ein AQC und Patientenmessung durchgeführt. Hr.Corte hat mir die Funktion telefonisch bestätigt. | SWO-99598 | No | The Conworx configuration is entered. An AQC and patient measurement are performed for control. Mr.Corte confirmed the function to me by phone. | The Conworx configuration is entered. An AQC and patient measurement are performed for control. Mr.Corte confirmed the function to me by phone. | Setup | Sw | His/lis | Conworx
AQT90 FLEX | 引上げ修理（ﾀｰﾐﾅﾙ） | B | SWO-99526 | No | Raiserepair (Terminal) |  | Revisit required | HW | Cover and Chassis | Storage Door
ABL80 FLEX Basic version |  | B | D0B8BFF7EC115D32 | No
ABL80 FLEX BASIC |  | B | SWO-99759 | No
ABL80 FLEX BASIC |  | B | SWO-99718 | No
ABL80 FLEX |  | B | SWO-100022 | No
ABL90 FLEX | Netzwerkeinstellungen sind nicht möglich. Es fehlt an Ansprechpartner bei der EDV | Netzwerkeinstellungen sind nicht möglich. Es fehlt an Ansprechpartner bei der EDV | SWO-99900 | No | Network settings are not possible. There is a lack of contact persons at the IT | Network settings are not possible. There is a lack of contact persons at the IT | Setup | SW | DMS_PC | Barcode Reader
ABL90 FLEX PLUS | ｻｰﾋﾞｽﾁｪｯｸ：CavitationAbsorb 0.12[5.00-30.00]で範囲外。 tHbｷｬﾙOK、QCOK。 バルブ動作OK。 バルブチューブ交換後、CaviotationAbsorb 14.78で改善。 | B | SWO-99844 | No | Service check: CavitationAbsorb 0.12 [5.00-30.00] out of range. tHb cal OK, QCOK. Valve operation OK. Improved at CaviotationAbsorb 14.78 after valve tube replacement. |  | Defect | HW | Oxi module | Pinch valve tube
ABL90 FLEX PLUS |  | B | D0B8A15FEC3E40BC | Yes |  |  | Re classified RMED | SWU
ABL90 FLEX PLUS | Defekte CPU Einheit ersetzt, Anwendersoftware MR8 eingespielt, Konfiguration von St. 62 gespeichert und eingespielt, sämtliche notwendige Einstellungen auf Systemebene (feste Zielwerte, Zeit, Datum etc.) vorgenommen und mit Befehl auf der CF Card gespeichert, Kal, Interne QC | B | SWO-99969 | No | Defective CPU unit replaced, user software MR8 recorded, configuration of st. 62 saved and recorded, all necessary settings made at system level (fixed target values, time, date, etc.) and saved with command on the CF Card, Kal, Internal QC |  | Defect | HW | CPU | CPU
ABL827 FLEX | Fitted another EL MET module which was working on an analyser in our workshop. The EL MET module removed will be fitted to the workshop analyser to see how it performs on that. New membrane fitted to the Lactate electrode because the sensitivity was below the lower limit. Cleaning cycle run as currently recommended to do after changing the Lactate membrane. | Fitted another EL MET module which was working on an analyser in our workshop. The EL MET module removed will be fitted to the workshop analyser to see how it performs on that. New membrane fitted to the Lactate electrode because the sensitivity was below the lower limit. Cleaning cycle run as currently recommended to do after changing the Lactate membrane. | SWO-99967 | No | Fitted another EL MET module which was working on an analyser in our workshop. The EL MET module removed will be fitted to the workshop analyser to see how it performs on that. New membrane fitted to the Lactate electrode because the sensitivity was below the lower limit. Cleaning cycle run as currently recommended to do after changing the Lactate membrane. | Fitted another EL MET module which was working on an analyser in our workshop. The EL MET module removed will be fitted to the workshop analyser to see how it performs on that. New membrane fitted to the Lactate electrode because the sensitivity was below the lower limit. Cleaning cycle run as currently recommended to do after changing the Lactate membrane. | Assisted Replacement | Electrode | Membrane | Lac | Calibration
ABL837 FLEX | Removed the auto check module cleaned the module and scanner restarted analyzer found auto check module is working. | Removed the auto check module cleaned the module and scanner restarted analyzer found auto check module is working. | SWO-101641 | No | Removed the auto check module cleaned the module and scanner restarted analyzer found auto check module is working. | Removed the auto check module cleaned the module and scanner restarted analyzer found auto check module is working. | Cleaned | HW | AutoCheck Module | Barcode scanner
ABL825 FLEX | Herr Rahner hat mir mitgeteilt das der Scanner am ABL nicht funktioniert. Scanner ausgebaut und gereinigt. Danach funktionierte der Scanner wieder. Mehrere Kalibrierungen durchgeführt. Nach mehrmaligem Kalibrieren -> Flüssigkeitsaustritt da PH Elektrode undicht. Dichtring erneuert. Nochmals mehrmals kalibriert. QCs für alle Level durchgeführt. --> i.O. ABL wieder Messbereit. | Herr Rahner hat mir mitgeteilt das der Scanner am ABL nicht funktioniert. Scanner ausgebaut und gereinigt. Danach funktionierte der Scanner wieder. Mehrere Kalibrierungen durchgeführt. Nach mehrmaligem Kalibrieren -> Flüssigkeitsaustritt da PH Elektrode undicht. Dichtring erneuert. Nochmals mehrmals kalibriert. QCs für alle Level durchgeführt. --> i.O. ABL wieder Messbereit. | SWO-99821 | No | Mr Rahner told me that the scanner on the ABL was not working. Scanner removed and cleaned. After that, the scanner worked again. Multiple calibrations performed. After several calibrations -> Liquid leakage since PH electrode leaks. Sealing ring renewed. Calibrated again several times. QCs for all levels. --> i.O. ABL ready to measure again. | Mr Rahner told me that the scanner on the ABL was not working. Scanner removed and cleaned. After that, the scanner worked again. Multiple calibrations performed. After several calibrations -> Liquid leakage since PH electrode leaks. Sealing ring renewed. Calibrated again several times. QCs for all levels. --> i.O. ABL ready to measure again. | Cleaned | HW | Basic unit | Barcode reader
ABL800 BASIC | GAS 2 HAS BEEN REPLACED NOW EQUIPMENT IS WORKIN G WELL | GAS 2 HAS BEEN REPLACED NOW EQUIPMENT IS WORKIN G WELL | SWO-101026 | No | GAS 2 HAS BEEN REPLACED NOW EQUIPMENT IS WORKIN G WELL | GAS 2 HAS BEEN REPLACED NOW EQUIPMENT IS WORKIN G WELL | Assisted Replacement | Gas | Gas2 | Gas 2
ABL800 BASIC | Conducted user training for the entire technician’s, explained the sample measurement, replacement of solutions paper, QC replacement and measurement’s and basic troubleshoot. | Conducted user training for the entire technician’s, explained the sample measurement, replacement of solutions paper, QC replacement and measurement’s and basic troubleshoot. | SWO-101644 | Yes | Conducted user training for the entire technician’s, explained the sample measurement, replacement of solutions paper, QC replacement and measurement’s and basic troubleshoot. | Conducted user training for the entire technician’s, explained the sample measurement, replacement of solutions paper, QC replacement and measurement’s and basic troubleshoot. | Re classified RMED | Training of customer
ABL810 FLEX | Anbindung an Conworx umgesetzt. Durch Hr. Corte die Funktion bestätig bekommen. | B | SWO-100055 | No | connection to Conworx. By Mr. Corte get the function confirmed. |  | Setup | Sw | His/lis | Conworx
ABL835 FLEX |  | B | D0B8C325EC2F0003 | No |  |  | Defect | HW | Inlet module | Inlet module
ABL825 FLEX | Fehlersuche, Reinigung Elektrodenkonnektor und Elektrode, Funktionskontrolle | Kd. bittet um Fehlerbehebung bei Lac und Überprüfung der Auto QC Konfiguration. | SWO-99740 | No | Troubleshooting, cleaning electrode connector and electrode, function control | Kd. asks for troubleshooting of Lac and checking the Auto QC configuration. | Cleaned | HW | El/Met Module | Electrode Connector | Lac
ABL810 FLEX | Die Konfiguration für Conworx Anbindung eingestellt, die Funktion wurde durch Hr.Corte bestätigt. | B | SWO-100056 | No | The configuration for Conworx connection is set, the function has been confirmed by Mr.Corte. |  | Setup | Sw | His/lis | Conworx
AQT90 FLEX | Tray door is broken. New cover to be shipped. Blank samples completed. | 5 10 16 Installed new front cover. | D0B94774EC29368F | No | Tray door is broken. New cover to be shipped. Blank samples completed. | 5 10 16 Installed new front cover. | Defect | HW | Cover and Chassis | Storage Door
ABL80 FLEX BASIC |  | B | SWO-100008 | No
ABL80 FLEX BASIC |  | B | SWO-100042 | No
ABL80 CO-OX | Inspected all data logs, event logs, qc logs. | Found no problems in all logs. Manual qc was ran alot due to verification kit being ran. There was no problem that I could visually see. | D0B9EB6B6214D7A8 | No | Inspected all data logs, event logs, qc logs. | Found no problems in all logs. Manual qc was ran alot due to verification kit being ran. There was no problem that I could visually see.
ABL80 FLEX CO-OX | Waste Manifold ausgetauscht da undicht (BS) Nacht tausch kein Flüssigkeitsaustritt mehr | Waste Manifold ausgetauscht da undicht (BS) Nacht tausch kein Flüssigkeitsaustritt mehr | SWO-100947 | No | Waste Manifold replaced since leaking (BS) Night exchange no more liquid leakage | Waste Manifold replaced since leaking (BS) Night exchange no more liquid leakage
ABL90 FLEX | Der var ikke forbindelse til netværksstikket | Der var ikke forbindelse til netværksstikket | SWO-100857 | No | There was no connection to the network connector | There was no connection to the network connector | Setup | HW | Cable | Connector
ABL90 FLEX PLUS |  | B | D0BA043C621C458B | No |  |  | Adjustment | HW | Auto Inlet module | Inlet Position Cal.
ABL90 FLEX PLUS | inspected entire inlet module and found no fault, entered service mode and lowered inlet probe, restarted analyzer and gave pump cal error, done back flush and was ok. all cals ok and Qc's ok | inspected entire inlet module and found no fault, entered service mode and lowered inlet probe, restarted analyzer and gave pump cal error, done back flush and was ok. all cals ok and Qc's ok | SWO-101192 | No | inspected entire inlet module and found no fault, entered service mode and lowered inlet probe, restarted analyzer and gave pump cal error, done back flush and was ok. all cals ok and Qc's ok | inspected entire inlet module and found no fault, entered service mode and lowered inlet probe, restarted analyzer and gave pump cal error, done back flush and was ok. all cals ok and Qc's ok | Adjustment | HW | Auto Inlet module | Inlet Position Cal.
ABL90 FLEX | 更换新测试卡修复仪器 | 更换新测试卡修复仪器 | SWO-100885 | No | Replace the new test card repair instrument | Replace the new test card repair instrument
AQT90 FLEX |  | B | SWO-100057 | No |  |  | Defect | HW | Process Module | PCB, Cup dryer | 168
ABL800 FLEX | install new window and application software | install new window and application software | SWO-101627 | No | install new window and application software | install new window and application software | Defect | SW | Windows | Re-installed
ABL800 BASIC | Clean the chamber ,run cleaning and recalibrate 2-3 times and now senstivity is OK and problem has been resolved. | Replace Lac membrane by following procedure and done cleaning then calibrate which is OK. | SWO-102451 | No | Clean the chamber ,run cleaning and recalibrate 2-3 times and now senstivity is OK and problem has been resolved. | Replace Lac membrane by following procedure and done cleaning then calibrate which is OK. | Assisted Replacement | Electrode | Membrane | Lac | Sensitivity
ABL830 FLEX | ｼｬｯﾄﾀﾞｳﾝ後Oxiﾓｼﾞｭｰﾙ交換。 再起動後はｺﾝﾌﾟﾘｰﾄﾃｽﾄ、ｻｰﾋﾞｽtHb実施。 | B | SWO-101702 | No | Oxi module replacement after shutdown. Complete test after reboot, service tHb implementation. |  | Defect | HW | Oxi module | Neon Lamp, low
ABL800 BASIC | Kept the machine OFF for half an hour. Restarted the machine. Taken the Gas 1 Cylinder from the distributor office, installed it, also changed the Rinse solution since the level was low. Done the calibration. Machine is working fine. | B | SWO-101654 | No | Kept the machine OFF for half an hour. Restarted the machine. Taken the Gas 1 Cylinder from the distributor office, installed it, also changed the Rinse solution since the level was low. Done the calibration. Machine is working fine. |  | Adjustment | HW | Oxi module | tHb cal
ABL825 FLEX | Oxi spectrum mismatch on qc. Suspect is malfunctioning hemolyzer. | Replaced hemolyzer. Performed complete oxi test - PASS. Performed thb cal. Verified instrument operation. | D0BA3CEA6212FAFB | No | Oxi spectrum mismatch on qc. Suspect is malfunctioning hemolyzer. | Replaced hemolyzer. Performed complete oxi test - PASS. Performed thb cal. Verified instrument operation. | Defect | HW | Oxi module | Hemolyzer
ABL825 FLEX | tried to scan QC and it would not scan, found lense lose-- extensive travel to and from Davis Ca--customer recieved double QC with different lots. picked up QC and sent back to westlake on RGAy16-122413c | replaced bar code scanner; | D0BA35D5623466BA | No | tried to scan QC and it would not scan, found lense lose-- extensive travel to and from Davis Ca--customer recieved double QC with different lots. picked up QC and sent back to westlake on RGAy16-122413c | replaced bar code scanner; | Defect | HW | Basic unit | Barcode Reader
ABL825 FLEX | Tausch Lactatelektrode, Funktionskontrolle | Tausch Lactatelektrode, Funktionskontrolle | SWO-100926 | No | Replacement of lactate electrode, function control | Replacement of lactate electrode, function control | Assisted Replacement | Electrode | Electrode | Lac | Calibration
AQT90 FLEX | Cleaned process wheel and reference.Checked with patient sample.Still showing the same error. Performed system check.that also failed.Showing flash lamp intensity out of range.Its value is -11.7 | B | SWO-101023 | No | Cleaned process wheel and reference. Checked with patient sample. Still showing the same error. Performed system check.that also failed. Showing flash lamp intensity out of range. Its value is -11.7 |  | Defect | HW | Process Module | Optical Unit | 131
AQT90 FLEX | Cleaned the instrument with vacuum cleaner, replaced the inlet door unit and restarted instrument found inlet door is functional. | Cleaned the instrument with vacuum cleaner, replaced the inlet door unit and restarted instrument found inlet door is functional. | SWO-101647 | No | Cleaned the instrument with vacuum cleaner, replaced the inlet door unit and restarted instrument found inlet door is functional. | Cleaned the instrument with vacuum cleaner, replaced the inlet door unit and restarted instrument found inlet door is functional. | Cleaned | HW | Cover and Chassis | Inlet Door
AQT90 FLEX | AQT zerlegt und zwei verkantete Cups aus dem Process Wheel entfernt. Dann musste Ventiel 1 wieder stabilisiert werden es saß lose auf der Weetsection und hat mehrere Flüüsigkeitstransportfehlrer verursacht weitere Produktionsfehler sind nicht auszuschließen. Die Cups konnten nicht Abtransportiert werden da die Verriegelung vom LP auch defekt war. | AQT zerlegt und zwei verkantete Cups aus dem Process Wheel entfernt. Dann musste Ventiel 1 wieder stabilisiert werden es saß lose auf der Weetsection und hat mehrere Flüüsigkeitstransportfehlrer verursacht weitere Produktionsfehler sind nicht auszuschließen. Die Cups konnten nicht Abtransportiert werden da die Verriegelung vom LP auch defekt war. | SWO-101010 | No | AQT disassembles and removes two edged cups from the Process Wheel. Then Ventiel 1 had to be stabilized again it sat loosely on the Weetsection and has caused several liquid transport failures further production errors are not to be ruled out. The cups could not be removed because the locking of the LP was also defective. | AQT disassembles and removes two edged cups from the Process Wheel. Then Ventiel 1 had to be stabilized again it sat loosely on the Weetsection and has caused several liquid transport failures further production errors are not to be ruled out. The cups could not be removed because the locking of the LP was also defective. | Defect | Consumables | SP | SP | 5784
AQT90 FLEX | Durchführung einer Wartungsmassnahme (Jahreswartung) mit allen notwendigen Reinigungsmaßnahmen von Bsugruppen. Überprüfung des Flows, vor Wartungsbeginns, ergabt normale Werte zwischen 1350 und 1450 cpunts, o. K. Durchführung von Blankmessungen, Proteinentfernung, Reinigung und Systemüberprüfung durchgeführt o. K. Service Dump vor Wartungsbeginn und danach gezogen. Flow Werte nach Wartungsmassnahme unter 1200 counts. Qualitätskontrolle o. K. Fragebogen zurFalsch positiven Messung mit Anwender ausgefüllt. | Durchführung einer Wartungsmassnahme (Jahreswartung) mit allen notwendigen Reinigungsmaßnahmen von Bsugruppen. Überprüfung des Flows, vor Wartungsbeginns, ergabt normale Werte zwischen 1350 und 1450 cpunts, o. K. Durchführung von Blankmessungen, Proteinentfernung, Reinigung und Systemüberprüfung durchgeführt o. K. Service Dump vor Wartungsbeginn und danach gezogen. Flow Werte nach Wartungsmassnahme unter 1200 counts. Qualitätskontrolle o. K. Fragebogen zurFalsch positiven Messung mit Anwender ausgefüllt. | SWO-101016 | No | Execution of a maintenance measure (annual maintenance) with all necessary cleaning measures of blot groups. Checking the flow, before the start of maintenance, yields normal values between 1350 and 1450 cpunts, o.k. Execution of blank measurements, protein removal, cleaning and system verification performed o. K. Service Dump before maintenance start and then drawn. Flow values after maintenance measure below 1200 counts. Quality control o.K. Questionnaire forfalse positive measurement filled out with user. | Execution of a maintenance measure (annual maintenance) with all necessary cleaning measures of blot groups. Checking the flow, before the start of maintenance, yields normal values between 1350 and 1450 cpunts, o.k. Execution of blank measurements, protein removal, cleaning and system verification performed o. K. Service Dump before maintenance start and then drawn. Flow values after maintenance measure below 1200 counts. Quality control o.K. Questionnaire forfalse positive measurement filled out with user. | Performance | Discrepant Result | Process Module | Cup wash unit
ABL80 FLEX BASIC | 更换测试卡正常。 | 更换测试卡正常。 | SWO-101400 | No | The replacement test card is OK. | The replacement test card is OK.
ABL80 FLEX BASIC | 更换903-334和902-972后修复(by zhenjie) | 更换903-334和902-972后修复(by zhenjie) | SWO-100228 | No | Repair after replacing 903-334 and 902-972 (by-zhenjie) | Repair after replacing 903-334 and 902-972 (by-zhenjie)
ABL80 FLEX BASIC | 更换902-972修复仪器(by zhenjie) | 更换902-972修复仪器(by zhenjie) | SWO-100886 | No | Replace the 902-972 repair instrument (by zhenjie) | Replace the 902-972 repair instrument (by zhenjie)
ABL80 FLEX BASIC | Cleaned Sc with distilled water and kept in fridge. Then reinstalled SC. Checked with Quality controls. Instrument is working fine. | Cleaned Sc with distilled water and kept in fridge. Then reinstalled SC. Checked with Quality controls. Instrument is working fine. | SWO-101604 | No | Cleaned Sc with distilled water and kept in fridge. Then reinstalled SC. Checked with Quality controls. Instrument is working fine. | Cleaned Sc with distilled water and kept in fridge. Then reinstalled SC. Checked with Quality controls. Instrument is working fine.
ABL80 FLEX CO-OX | Tool in Grundeinstellung auf PC installiert, Pfade angepasst, Datensicherung - überspielen und erstellen der Monatsberichte geschult, Bewertung und Dokumentation besprochen | Tool in Grundeinstellung auf PC installiert, Pfade angepasst, Datensicherung - überspielen und erstellen der Monatsberichte geschult, Bewertung und Dokumentation besprochen | SWO-101847 | No | Tool installed on PC as a default, paths adapted, data backup - overplay and create the monthly reports trained, evaluation and documentation discussed | Tool installed on PC as a default, paths adapted, data backup - overplay and create the monthly reports trained, evaluation and documentation discussed
ABL90 FLEX | System gespült, Durchfluss überprüft, Anschlussdichtung ersetzt u. LP gereinigt, Durchflussüberprüfung i.O., LP er-initialisiert, Startphase abgewartet, Kal u. QC i.O. | System gespült, Durchfluss überprüft, Anschlussdichtung ersetzt u. LP gereinigt, Durchflussüberprüfung i.O., LP er-initialisiert, Startphase abgewartet, Kal u. QC i.O. | SWO-101129 | No | System flushed, flow checked, connection seal replaced and LP cleaned, flow check i.O., LP er-initialized, start phase awaited, Kal & QC i.O. | System flushed, flow checked, connection seal replaced and LP cleaned, flow check i.O., LP er-initialized, start phase awaited, Kal & QC i.O. | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
ABL90 |  | B | D0BCD8C04E461E4B | Yes |  |  | Re classified RMED | Telph Support
ABL90 FLEX PLUS | System über Kurzabschalttaster auf der Rückseite neu gestartet, im Service Einlass geprüft und wieder richtig zusammen gebaut, Startphase, mit Kunden den Einlassdichtungstausch durchgesprochen | Fr. Tespedes bittet dringend um einen Service Einsatz. Bitte kurz Bescheid geben, wann der Service erfolgt. | SWO-101126 | No | System restarted via short switch-off button on the back, checked in the service inlet and reassembled properly together, start phase, with customers the inlet seal exchange discussed | Ms. Tespedes is urgently asking for a service. Please inform you shortly when the service will take place. | User error | HW | Auto Inlet module | Inlet Gasket
ABL90 FLEX PLUS | Cleaned pin contact reinstall SC ran calibration status and sensitivity ok. Ran QC all passed ok | B | SWO-101202 | No | Cleaned pin contact reinstall SC ran calibration status and sensitivity ok. Ran QC all passed ok |  | Cleaned | HW | SC contact pin | SC contact pin
ABL90 FLEX | 更换测试卡修复仪器 | 更换测试卡修复仪器 | SWO-101131 | No | Replace the test card repair instrument | Replace the test card repair instrument
ABL90 FLEX | 更换测试卡正常。 | 更换测试卡正常。 | SWO-101373 | No | The replacement test card is OK. | The replacement test card is OK. | Assisted Replacement | SC | SC | Lac
ABL90 FLEX |  | B | SWO-101160 | No |  |  | Defect | HW | Basic Unit | Sensor Interface PCB | Return request - 216-5-18
AQT90 FLEX |  | B | SWO-101136 | No |  |  | Defect | HW | Process Module | Optical Unit | 131
ABL815 FLEX | Cal 2 érzékelő csere. | B | SWO-101081 | No | Cal 2 sensor replacement. |  | Defect | HW | Basic unit (refurb) | Bottle detector PCB
ABL835 FLEX | Waste érzékelő csere. | B | SWO-101080 | No | Waste sensor replacement. |  | Defect | HW | Basic unit (refurb) | Bottle detector PCB
ABL800 BASIC | Replace the Lac membrane by following proper replacement procedure and done 2-point calibration and resolve the problem | Replace Lac membrane by following procedure and done cleaning then calibrate which is OK. | SWO-102923 | No | Replace the Lac membrane by following proper replacement procedure and done 2-point calibration and resolve the problem | Replace Lac membrane by following procedure and done cleaning then calibrate which is OK. | Assisted Replacement | Electrode | Membrane | Lac | Sensitivity
ABL800 BASIC | Replaced the waste tube & checked equipment is functional | Replaced the waste tube & checked equipment is functional | SWO-101671 | No | Replaced the waste tube & checked equipment is functional | Replaced the waste tube & checked equipment is functional | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
ABL835 FLEX | Lacメンブラン不良の為、交換。 | Lacメンブラン不良の為、交換。 | SWO-101575 | No | Lac membrane bad, replacement. | Lac membrane bad, replacement. | Assisted Replacement | Electrode | Membrane | Lac | Sensitivity
ABL835 FLEX | WDC results inconsistent over time. | Recleaned contacts, checked valves and pumps. Replaced reference and calcium electrodes. | D0BB6B09D840AA77 | No | WDC results inconsistent over time. | Recleaned contacts, checked valves and pumps. Replaced reference and calcium electrodes. | Assisted Replacement | Electrode | Electrode | Ca
ABL825 FLEX | During pH BG cal unit performs a cal 1 due to drift | Replaced pH electrode. | D0BB7325D858D640 | No | During pH BG cal unit performs a cal 1 due to drift | Replaced pH electrode. | Assisted Replacement | Electrode | Electrode | pH
ABL835 FLEX | Inlet coule et probleme trouvé rouleau de des déchets pas engagé et tube inlet pincé | Remplacé tube pompe BG et Met inlet tube | D0BB17D9D81F46F9 | No | Inlet sinks and problem found waste roll not engaged and tube inlet pinched | Replaced pump tube BG and Met inlet tube | Defect | HW | Fluidic manifold | Waste Pump
ABL825 FLEX | Power Supply erneuert. Kal und QC abgewartet. | Neue PC Unit eingesetzt. Kal und QC nach Einlaufphase ok. Zur weiteren Beobachtung | SWO-101269 | No | Power Supply renewed. Kal and QC waited. | New PC Unit. Kal and QC after run-in phase ok. For further observation | Defect | HW | DMS_PC | CPU
ABL800 BASIC | cleaned the touch screen and done calibration for touch screen.its worked for one hours and again giving problem.it is need to replace with new touch screen. | cleaned the touch screen and done calibration for touch screen.its worked for one hours and again giving problem.it is need to replace with new touch screen. | SWO-101776 | No | cleaned the touch screen and done calibration for touch screen.its worked for one hours and again giving problem.it is need to replace with new touch screen. | cleaned the touch screen and done calibration for touch screen.its worked for one hours and again giving problem.it is need to replace with new touch screen. | Revisit required | HW | DMS_PC | Touch screen
ABL80 FLEX BASIC | Try To install new SC but same problem occur.Perform Zeroing process Po2 parameter fail in zeroing we have to replace Sensor board | Try To install new SC but same problem occur.Perform Zeroing process Po2 parameter fail in zeroing we have to replace Sensor board | SWO-101630 | No | Try To install new SC but same problem occur. Perform Zeroing process Po2 parameter fail in zeroing we have to replace Sensor board | Try To install new SC but same problem occur. Perform Zeroing process Po2 parameter fail in zeroing we have to replace Sensor board
ABL90 | Numerous errors - 1405,1063, 0751, 1087, 1065. | Back flush completed, SP changed twice, inlet gasket and probe change with some improvement. Replaced SC, errors have subsided. | D0BCAC124E4C3E1D | No | Numerous errors - 1405,1063, 0751, 1087, 1065. | Back flush completed, SP changed twice, inlet gasket and probe change with some improvement. Replaced SC, errors have subsided. | Assisted Replacement | SC | SC | Temp
ABL90 FLEX PLUS |  | B | D0C03664B0297F98 | No |  |  | Defect | HW | Auto Inlet module | Motor assembly | Return request (216-5-18)
ABL90 FLEX PLUS |  | B | D0C02D0CB014A36A | Yes |  |  | Re classified RMED | PER upSWO
AQT90 FLEX |  | 更换了借用的903-118后未修复仪器，继续出现相同故障，需要再增加借件 | SWO-101367 | No |  | Replaced the borrowed 903-118 after the unrepaired instrument, continue to have the same fault, need to add the loan | Defect | HW | Process Module | Cup dryer, Air filter | 168
ABL815 FLEX | nyomtató mechanika cseréje | B | SWO-101493 | No | changing printer mechanics |  | Defect | HW | Printer module | Printer Module
ABL815 FLEX |  | B | D0BCAC094E00E0B8 | No |  |  | Assisted Replacement | Electrode | Electrode | Reference
ABL827 FLEX |  | B | D0C0355FB03F7F93 | No |  |  | Defect | HW | Oxi module | Hemolyzer
ABL835 FLEX | Tried new reference membrane and switch pCO electrodes with another machine-No change Back up all data | Installed software version 6.16 Installed new reference membrane lot 1926 | D0BC8EC54E5F2705 | No | Tried new reference membrane and switch pCO electrodes with another machine-No change Back up all data | Installed software version 6.16 Installed new reference membrane lot 1926 | Assisted Replacement | Electrode | Membrane | Reference
ABL825 FLEX |  | B | D0BC8A484E34AC2F | No |  |  | Defect | HW | Inlet module | Probe
ABL815 FLEX | Dateien überprüft, Lac-Elektrode ersetzt (alt: UP06, neu: UP04), Weiterhin Kal-Empfindlichkeitsfehler, Rücktausch Elektrode, Membranierung mit anderem Lot (R2982), Kal u. QC i.O. | B | SWO-101599 | No | Files checked, Lac electrode replaced (old: UP06, new: UP04), Further cal sensitivity error, exchange electrode, membraneing with other solder (R2982), Kal & QC i.O. |  | Assisted Replacement | Electrode | Membrane | Lac | Sensitivity
AQT90 FLEX | Opened the front panel, cleaned the dust on the barcode scanner. Restarted the machine. As per request by Pascal Sir, taken the troubleshooting training for the technicians handling the machine. Check the stock of consumables. Machine is working fine. | Opened the front panel, cleaned the dust on the barcode scanner. Restarted the machine. As per request by Pascal Sir, taken the troubleshooting training for the technicians handling the machine. Check the stock of consumables. Machine is working fine. | SWO-101653 | No | Opened the front panel, cleaned the dust on the barcode scanner. Restarted the machine. As per request by Pascal Sir, taken the troubleshooting training for the technicians handling the machine. Check the stock of consumables. Machine is working fine. | Opened the front panel, cleaned the dust on the barcode scanner. Restarted the machine. As per request by Pascal Sir, taken the troubleshooting training for the technicians handling the machine. Check the stock of consumables. Machine is working fine. | Cleaned | HW | Storage Module | Barcode reader, Storage wheel
TCM/Tosca | Remembrain sensor checked for damage ok. Test calibration time ok. Test sensor and unit on engineer . pco2 readings 4.9 kPa spo2 97% PR 83 All reads well within specification. Note customer must use the contact gel when connecting attachment clips. Using the wrong solution will have a big impact on the readings. | Remembrain sensor checked for damage ok. Test calibration time ok. Test sensor and unit on engineer . pco2 readings 4.9 kPa spo2 97% PR 83 All reads well within specification. Note customer must use the contact gel when connecting attachment clips. Using the wrong solution will have a big impact on the readings. | SWO-101615 | No | Remembrain sensor checked for damage ok. Test calibration time ok. Test sensor and unit on engineer . pco2 readings 4.9 kPa spo2 97% PR 83 All reads well within specification. Note customer must use the contact gel when connecting attachment clips. Using the wrong solution will have a big impact on the readings. | Remembrain sensor checked for damage ok. Test calibration time ok. Test sensor and unit on engineer . pco2 readings 4.9 kPa spo2 97% PR 83 All reads well within specification. Note customer must use the contact gel when connecting attachment clips. Using the wrong solution will have a big impact on the readings.
ABL90 |  | B | D0BE0976C41F0019 | No |  |  | User error | SP Installation | SP Installation | SP Installation
ABL80 FLEX BASIC |  | B | SWO-101695 | No
ABL80 FLEX BASIC | 更换测试卡正常。 | 更换测试卡正常。 | SWO-101785 | No | The replacement test card is OK. | The replacement test card is OK.
ABL80 FLEX BASIC | 更换新试剂包修复仪器 | 更换新试剂包修复仪器 | SWO-101832 | No | Replace the new reagent kit repair instrument | Replace the new reagent kit repair instrument
ABL80 CO-OX | Analyzer has been failing system cycle. Unit checking OK since late AM. Failure has been level 2 internal QC (C8302 lot 220849)value around 74.0 | Value now running 75.9 and OK. Check valve operation and for leaks. All check ok. If pCO2 drops low again, SP change needed. | D0C0C344B01A89C8 | No | Analyzer has been failing system cycle. Unit checking OK since late AM. Failure has been level 2 internal QC (C8302 lot 220849)value around 74.0 | Value now running 75.9 and OK. Check valve operation and for leaks. All check ok. If pCO2 drops low again, SP change needed.
ABL90 | Analyzer has been out of service since 5 6. New SP was installed, top was not fully activated down. Failing multiple fluid flow tests, poewered on off numerous times resulting in no power to analyzer. | All fuses check OK, power supply checks ok, Analog PCB checks OK, replaced CPU PCB. Re-installed software, SP, and SC. All checking OK. | D0C08C49B0181363 | No | Analyzer has been out of service since 5 6. New SP was installed, top was not fully activated down. Failing multiple fluid flow tests, poewered on off numerous times resulting in no power to analyzer. | All fuses check OK, power supply checks ok, Analog PCB checks OK, replaced CPU PCB. Re-installed software, SP, and SC. All checking OK. | Defect | HW | CPU | CPU
ABL90 FLEX PLUS | Replaced sensor interface PCB. Calibration and QC passing ok. | Cleaned pin contact reinstall SC ran calibration status and sensitivity ok. Ran QC all passed ok | SWO-101761 | No | Replaced sensor interface PCB. Calibration and QC passing ok. | Cleaned pin contact reinstall SC ran calibration status and sensitivity ok. Ran QC all passed ok | Defect | HW | Basic Unit | Sensor Interface PCB | Return request - 216-5-18
AQT90 FLEX | 将卡盘的杯子取出。(by andrew) | 将卡盘的杯子取出。(by andrew) | SWO-101705 | No | Remove the cup from the chuck. (by andrew) | Remove the cup from the chuck. (by andrew) | Defect | Consumables | Test Kit | Cartridge | 5784 (Punch out)
ABL800 BASIC | Clean the chamber,Interchange the Glu & Lac to confirm the electrode and finally found the problem in the membrane and replace the membrane . | Replace Lac membrane by following procedure and done cleaning then calibrate which is OK. | SWO-102938 | No | Clean the chamber,Interchange the Glu & Lac to confirm the electrode and finally found the problem in the membrane and replace the membrane . | Replace Lac membrane by following procedure and done cleaning then calibrate which is OK. | Assisted Replacement | Electrode | Membrane | Lac | Sensitivity
ABL837 FLEX | Zamjena pisača. | B | SWO-101746 | No | Replacing the printer. |  | Defect | HW | Printer module | Printer Module
ABL825 FLEX |  | B | D0C0338AB01B771B | Yes |  |  | Re classified RMED | SWU
ABL825 FLEX | Looked in data logs and found a few oxi spectrum mismatch on blood samples. Did complete oxi test in Service. Found lamp a little high. | Adjusted lamp, Did complete Oxi test, Did THB in service and as user and as sample. | D0C06E36B016A16C | No | Looked in data logs and found a few oxi spectrum mismatch on blood samples. Did complete oxi test in Service. Found lamp a little high. | Adjusted lamp, Did complete Oxi test, Did THB in service and as user and as sample. | Adjustment | HW | Oxi module | Lamp unit
ABL825 FLEX | Looked in datalogs and found received 0581, oxi spectrum mismatch. Did complete oximetry test in service. Lamp failed for Neon Intensity to low, was only 33% when it should be greater than 50%. | Replaced lamp unit, Adjusted the lamp, Did THB in service and as user and as sample. | D0C057D9B02FA167 | No | Looked in datalogs and found received 0581, oxi spectrum mismatch. Did complete oximetry test in service. Lamp failed for Neon Intensity to low, was only 33% when it should be greater than 50%. | Replaced lamp unit, Adjusted the lamp, Did THB in service and as user and as sample. | Defect | HW | Oxi module | Lamp unit
ABL800 BASIC | Replaced Cal 1 bottle with a new one. After successful calibration, machine is working fine. Also trained the technicians on how to resolve this problem. | Replaced Cal 1 bottle with a new one. After successful calibration, machine is working fine. Also trained the technicians on how to resolve this problem. | SWO-103063 | No | Replaced Cal 1 bottle with a new one. After successful calibration, machine is working fine. Also trained the technicians on how to resolve this problem. | Replaced Cal 1 bottle with a new one. After successful calibration, machine is working fine. Also trained the technicians on how to resolve this problem. | Assisted Replacement | Consumables | Solution Bottle | Cal 1
AQT90 FLEX | Cleaned the Analyzer with Vacuum cleaner, performed the System clean, calibrated analyzer for Tnt and CRP performed the test found instrument is functional. | Cleaned the Analyzer with Vacuum cleaner, performed the System clean, calibrated analyzer for Tnt and CRP performed the test found instrument is functional. | SWO-102825 | Yes | Cleaned the Analyzer with Vacuum cleaner, performed the System clean, calibrated analyzer for Tnt and CRP performed the test found instrument is functional. | Cleaned the Analyzer with Vacuum cleaner, performed the System clean, calibrated analyzer for Tnt and CRP performed the test found instrument is functional. | Re classified RMED
ABL80 FLEX BASIC | 更换测试卡正常。 | 更换测试卡正常。 | SWO-102515 | No | The replacement test card is OK. | The replacement test card is OK.
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-102117 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL80 FLEX CO-OX | Fehlersuche, Gerät aus, Hemolyzer gew., Gerät ein, div. Kal. durchgef., Kal. + QC i.O. | Fehlersuche; Gerät aus; Hemolyzer gew.; Gerät ein; div. Kal. durchgef.; Kal. + QC i.O. | SWO-101882 | No | Troubleshooting; device off; Hemolyzer. device; div. Kal. through it; Kal. + QC in the first place | Troubleshooting; device off; Hemolyzer. device; div. Kal. through it; Kal. + QC in the first place
ABL80 FLEX BASIC | 更换新试剂包后修复 | 更换新试剂包后修复 | SWO-102161 | No | Repair after replacing a new kit | Repair after replacing a new kit
ABL80 FLEX Basic version |  | B | D0C29F039C370505 | No
ABL90 FLEX | Replaced probe and inlet gasket with holder. | B | SWO-102310 | No | Replaced probe and inlet gasket with holder. |  | Replaced | HW | Probe bend | Probe bend
ABL90 |  | B | D0C2AFF29C10C9C3 | Yes |  |  | Re classified RMED | Installation
ABL90 FLEX PLUS |  | B | D0C1884F260E2C75 | Yes |  |  | Re classified RMED | Logistical upSWO
ABL815 FLEX | An beiden Elektroden fehlten die Dichtungsringe, Elektroden richtig bezogen Gerät kalibriert, | An beiden Elektroden fehlten die Dichtungsringe, Elektroden richtig bezogen Gerät kalibriert, | SWO-102225 | No | Both electrodes lacked the sealing rings, electrodes properly covered device calibrated, | Both electrodes lacked the sealing rings, electrodes properly covered device calibrated, | Assisted Replacement | Electrode | Electrode | Sealing Ring | Lac
ABL825 FLEX |  | B | D0C1697D26347916 | Yes |  |  | Re classified RMED | SWU
ABL825 FLEX |  | B | D0C16A9F26627919 | Yes |  |  | Re classified RMED | SWU
ABL825 FLEX | Checked voltage to computer-low | Replaced power supply Computer boots up fine | D0C1947A2627D2EE | No | Checked voltage to computer-low | Replaced power supply Computer boots up fine | Defect | HW | Basic Unit | Power Supply
ABL800 BASIC | ｶﾞｽﾄﾗﾝｽﾃﾞｭｰｻ交換では改善しないため、代替ｶﾞｽﾓｼﾞｭｰﾙを設置。 | B | SWO-102810 | No | Alternative gas modules are installed because gas transducer replacement does not improve. |  | Defect | HW | Gas module | Pressure Transducer
ABL800 BASIC | clean the electrode. clean the o-ring. clean the contact. drift in calibration. if problem repeat need to check the Lac electrode. | clean the electrode. clean the o-ring. clean the contact. drift in calibration. if problem repeat need to check the Lac electrode. | SWO-102088 | No | clean the electrode. clean the o-ring. clean the contact. drift in calibration. if problem repeat need to check the Lac electrode. | clean the electrode. clean the o-ring. clean the contact. drift in calibration. if problem repeat need to check the Lac electrode. | Assisted Replacement | Electrode | Membrane | Lac | Zero Current
ABL800 BASIC | エアライン洗浄 チューブ交換 液センサー調整 | B | SWO-104568 | No | Airline Cleaning Tube Replacement Liquid Sensor Adjustment |  | Cleaned | HW | Fluidic manifold | Air vent
ABL815 FLEX | Gerät aus und wieder eingeschaltet, kalibriert und Parameter entsperrt, | Gerät aus und wieder eingeschaltet, kalibriert und Parameter entsperrt, | SWO-101969 | No | device off and back on, calibrated and parameters unlocked, | device off and back on, calibrated and parameters unlocked, | Defect | HW | Oxi module | Spectrum error
ABL827 FLEX | GETTING JUST DOTS AND NO NUMBERS FOR PO2 AND PCO2 FAILING DRIFT ON PH AND K+ EVEN THOUGH DRIFT AT .01.LOOKED IN CALIBRATION DRIFT TOLERANCES AND FOUND INCORRECT VERY HIGH VALUES.ALSO BAROMETRIC PRESSURE VALUES MISSING.ALSO FOUND IN ACTIVITY LOGS THAT SETTINGS WERE TRANSFERRED FROM A USB BUT FAILED WHICH CORRUPTED THE SOFTWARE | REMOVE KIRA AND REINSTALLED VERSION 6.16 SOFTWARE.SWAPPED OUT THE ANALYZER CONTROL BOARD FOR THE MISSING BAROMETER VALUES.REINSTALLED QC AND PERFORMED MULTIPLE CALIBRATIONS INCLUDING THB CALIBRATION | D0C1B66326631F23 | No | GETTING JUST DOTS AND NO NUMBERS FOR PO2 AND PCO2 FAILING DRIFT ON PH AND K+ EVEN THOUGH DRIFT AT .01.LOOKED IN CALIBRATION DRIFT TOLERANCES AND FOUND INCORRECT VERY HIGH VALUES. ALSO BAROMETRIC PRESSURE VALUES MISSING. ALSO FOUND IN ACTIVITY LOGS THAT SETTINGS WERE TRANSFERRED FROM A USB BUT FAILED WHICH CORRUPTED THE SOFTWARE | REMOVE KIRA AND REINSTALLED VERSION 6.16 SOFTWARE. SWAPPED OUT THE ANALYZER CONTROL BOARD FOR THE MISSING BAROMETER VALUES. REINSTALLED QC AND PERFORMED MULTIPLE CALIBRATIONS INCLUDING THB CALIBRATION | Defect | HW | Basic Unit | Analyzer Control PCB | Barometer
AQT90 FLEX | check the system intensity. intensity low. need to replace the optical unit. | check the system intensity. intensity low. need to replace the optical unit. | SWO-102086 | No | check the system intensity. intensity low. need to replace the optical unit. | check the system intensity. intensity low. need to replace the optical unit. | Defect | HW | Process Module | Optical Unit
AQT90 FLEX | Needle wash cleaning procedure en Fluid path gereinigd. Naald afgeregeld. | B | SWO-101978 | No | Needle Wash Cleaning procedure and Fluid path cleaned. Needle is adjusted. |  | Cleaned | HW | Aspiration Module | Needle Wash Unit
AQT90 FLEX | プリンターユニット交換後も印字しないので角田氏相談、いろいろ実施するも同様なので、引上げ修理予定 | B | SWO-102109 | No | Mr. Kakuda consultation because it does not print even after the printer unit replacement, so it is the same as implementing various things, so it will be raised and repaired |  | Not Cat | Not cat. | Not cat. | Not cat.
ABL80 FLEX BASIC | flushed the inlet tube and waste pump tube. ABL 80 is workinf fine now. | flushed the inlet tube and waste pump tube. ABL 80 is workinf fine now. | SWO-102121 | No | flushed the inlet tube and waste pump tube. ABL 80 is workinf fine now. | flushed the inlet tube and waste pump tube. ABL 80 is workinf fine now.
ABL80 CO-OX | The analyzer was showing fluid blockage | I replaced all the tubings and ran cal's | D0C2DED39C503182 | No | The analyzer was showing fluid blockage | I replaced all the tubings and ran cal's
ABL90 | sO2 is blank on patient results when the value is above the reportable range. | In roma.imi I checked and, confirmed that the analyzer is set-up to show values above the reportable range, confirmed that iti is enabled. Checked and, confirmed that 'Repression' is turned off and, that 'Suppression' is turned 'On'. pO2 is failing on qc level 1 message 521 inhomogeneous sample error. Back-flush the system, replaced a new SC. Analyzer does not have the inlet clip, replace the inlet clip. Gave the customer a clip for her other ABL90. | D0C2C00C9C2CA2CC | No | sO2 is blank on patient results when the value is above the reportable range. | In roma.imi I checked and, confirmed that the analyzer is set-up to show values above the reportable range, confirmed that iti is enabled. Checked and, confirmed that 'Repression' is turned off and, that 'Suppression' is turned 'On'. pO2 is failing on qc level 1 message 521 inhomogeneous sample error. Back-flush the system, replaced a new SC. Analyzer does not have the inlet clip, replace the inlet clip. Gave the customer a clip for her other ABL90. | Assisted Replacement | SC | SC | Wetability
ABL90 | Unit would connect for short period of time, then drop off. Different amount of time between connection and loss of connection. | Replaced Belkin Adaptor.Signal strength a little stronger. | D0CB6892D6029A32 | No | Unit would connect for short period of time, then drop off. Different amount of time between connection and loss of connection. | Replaced Belkin Adaptor.Signal strength a little stronger. | Setup | Wifi | His/lis | His/Lis
ABL90 FLEX | Gerät manuell gespült, Einlassdichtung erneuert, Kaltstartphase abgewartet, Kal und QCs durchgeführt. | Kunde bittet um einen Einsatz des zuständigen Technikers, bitte um Kontaktaufnahme. | SWO-102325 | No | Device manually rinsed, inlet seal renewed, cold start phase awaited, Kal and QCs performed. | Customer asks for the use of the responsible technician, please contact us. | Defect | HW | Inlet Gasket | Inlet gasket
ABL90 |  | B | D0C2D7E89C0FA4E6 | No |  |  | No fault found | No fault found | No fault found | No fault found
ABL827 FLEX |  | B | D0C2E3349C019408 | No |  |  | Assisted Replacement | Electrode | Electrode | pH
ABL825 FLEX | ABL Inbetriebnahme Austausch gegen ABL 754R0454N012 | ABL Inbetriebnahme Austausch gegen ABL 754R0454N012 | SWO-102313 | Yes | ABL Commissioning Replacement for ABL 754R0454N012 | ABL Commissioning Replacement for ABL 754R0454N012 | Re classified RMED | Demo setup
ABL815 FLEX | GARANTIE! Schläuche im Umfang einer Jahreswartung erneuert. | GARANTIE! Schläuche im Umfang einer Jahreswartung erneuert. | SWO-102352 | No | Guarantee! Hoses to the extent of an annual maintenance. | Guarantee! Hoses to the extent of an annual maintenance. | Defect | HW | Basic unit | Tubings
ABL820 FLEX | Viewed logs to confirm error. Inspected auto check. Carousel sticking in the auto check. | Ran service program and cleaned all sensors and mirrors. Flex cable slightly skewed - reseated and secured. | D0C2D5409C58299A | No | Viewed logs to confirm error. Inspected auto check. Carousel sticking in the auto check. | Ran service program and cleaned all sensors and mirrors. Flex cable slightly skewed - reseated and secured. | Defect | HW | AutoCheck Module | FPC Cable
ABL825 FLEX | OXI 10 Key bestellt und installiert. | OXI 10 Key bestellt und installiert. | SWO-102360 | No | OXI 10 Key ordered and installed. | OXI 10 Key ordered and installed. | Setup | SW | Demo Key Exp | Demo Key Exp
ABL825 FLEX | OXI 10 Key bestellt und installiert. | OXI 10 Key bestellt und installiert. | SWO-102357 | No | OXI 10 Key ordered and installed. | OXI 10 Key ordered and installed. | Setup | SW | SW key Expired | SW key Expired
ABL800 BASIC | ＥＬ Ｍｅｔモジュール持ち帰りチェック | ＥＬ Ｍｅｔモジュール持ち帰りチェック | SWO-103175 | No | EL Met Module Takeaway Check | EL Met Module Takeaway Check | Unknown | HW | El/Met Module | Glu
TCM TOSCA | Electrode test cable recieved. Try to test electrode but TOSCA Electrode can not be tested by this cable only CombiM electrode can be tested. Need to cross check with another electrode- 2. 5621000- TOSCA Sensor 92, tcPCO2 SpO2 PR. Sending TO request. Ask Mahendra to visit next week for Application related quaries and Dr satisfaction. | Electrode test cable recieved. Try to test electrode but TOSCA Electrode can not be tested by this cable only CombiM electrode can be tested. Need to cross check with another electrode- 2. 5621000- TOSCA Sensor 92, tcPCO2 SpO2 PR. Sending TO request. Ask Mahendra to visit next week for Application related quaries and Dr satisfaction. | SWO-102309 | No | Electrode test cable recieved. Try to test electrode but TOSCA Electrode can not be tested by this cable only CombiM electrode can be tested. Need to cross check with another electrode- 2. 5621000- TOSCA Sensor 92, tcPCO2 SpO2 PR. Sending TO request. Ask Mahendra to visit next week for Application related quaries and Dr satisfaction. | Electrode test cable recieved. Try to test electrode but TOSCA Electrode can not be tested by this cable only CombiM electrode can be tested. Need to cross check with another electrode- 2. 5621000- TOSCA Sensor 92, tcPCO2 SpO2 PR. Sending TO request. Ask Mahendra to visit next week for Application related quaries and Dr satisfaction.
ABL80 FLEX BASIC | Replaced Sensor board & after doing zeroing, installed sensor cassette & solution pack. Analyzer ready. | Replaced Sensor board & after doing zeroing, installed sensor cassette & solution pack. Analyzer ready. | SWO-102762 | No | Replaced Sensor board & after doing zeroing, installed sensor cassette & solution pack. Analyzer ready. | Replaced Sensor board & after doing zeroing, installed sensor cassette & solution pack. Analyzer ready.
ABL80 FLEX CO-OX OSM | ＜代替設置＞ ヘモライザー内部にて詰まりが発生しており水道水や次亜塩でも開通せず、代替えヘモライザー設置。 | B | SWO-102861 | No | <代替設置> clogging occurs inside the hemolizer, and it does not open even with tap water or sub-salt, and an alternative hemolyzer is installed.
ABL80 FLEX BASIC | 更换新测试卡修复仪器 | 更换新测试卡修复仪器 | SWO-104208 | No | Replace the new test card repair instrument | Replace the new test card repair instrument
ABL80 FLEX Basic version |  | B | D0C4097E122F0511 | No
ABL90 |  | B | D0C4063712616353 | No |  |  | Defect | HW | Oxi module | Pinch valve tube
ABL90 |  | B | D0C405C812086352 | Yes |  |  | Re classified RMED | See DC463712616353
ABL90 | No problems found with analyzer. Noted many 1178 errors in data log. | Perform all service tests. Visually check analyzer. | D0C423F1121CF999 | No | No problems found with analyzer. Noted many 1178 errors in data log. | Perform all service tests. Visually check analyzer. | No fault found | No fault found | No fault found | No fault found
ABL90 FLEX | Nettoyage du la section fluidique e.o Remplacé les conduites souples e.o Test du système à l'aide du programme de service e.o Remplacé le SC 90 ( tjrs instable après plusieurs heures d'attentes : sensibilité PCO2 trop bas ) Lot de SC (commande du 07.03.2016) probablement défectueux Calibration et CQ ok ( hormis PCO2 ) | Nettoyage du la section fluidique e.o Remplacé les conduites souples e.o Test du système à l'aide du programme de service e.o Remplacé le SC 90 ( tjrs instable après plusieurs heures d'attentes : sensibilité PCO2 trop bas ) Lot de SC (commande du 07.03.2016) probablement défectueux Calibration et CQ ok ( hormis PCO2 ) | SWO-102708 | No | Cleaning the fluid section e.o Replaced the flexible pipes e.o System testing using the e.o service program Replaced the SC 90 (unstable tjrs after several hours of waiting: PCO2 sensitivity too low) SC lot (order of 07.03.2016) probably faulty Calibration and CQ ok (except PCO2) | Cleaning the fluid section e.o Replaced the flexible pipes e.o System testing using the e.o service program Replaced the SC 90 (unstable tjrs after several hours of waiting: PCO2 sensitivity too low) SC lot (order of 07.03.2016) probably faulty Calibration and CQ ok (except PCO2) | Assisted Replacement | SC | SC | pCO2
ABL90 FLEX | 更换测试卡正常。 | 更换测试卡正常。 | SWO-102813 | No | The replacement test card is OK. | The replacement test card is OK. | Assisted Replacement | SC | SC | pCO2
AQT90 FLEX | 更换903-364 重新定标 | 更换903-364 重新定标 | SWO-102579 | No | Replace 903-364 Re-calibration | Replace 903-364 Re-calibration | Defect | HW | Process Module | Optical Unit | 124
ABL800 BASIC | Found all the samples results Na values are low, checked and found Ref membrane problem, replaced the Ref membrane calibrated found instrument is functional, performed the sample results are ok for Na parameter. | Found all the samples results Na values are low, checked and found Ref membrane problem, replaced the Ref membrane calibrated found instrument is functional, performed the sample results are ok for Na parameter. | SWO-102827 | No | Found all the samples results Na values are low, checked and found Ref membrane problem, replaced the Ref membrane calibrated found instrument is functional, performed the sample results are ok for Na parameter. | Found all the samples results Na values are low, checked and found Ref membrane problem, replaced the Ref membrane calibrated found instrument is functional, performed the sample results are ok for Na parameter. | Assisted Replacement | Electrode | Membrane | Reference
ABL835 FLEX | Engineer to rest tend site and replace the spectrometer to see if this improves the issue. Engineer to arrange a replacement pH electrode to be sent to the customer FAO Rovinder Madhar | Engineer to rest tend site and replace the spectrometer to see if this improves the issue. Engineer to arrange a replacement pH electrode to be sent to the customer FAO Rovinder Madhar | SWO-102652 | No | Engineer to rest tend site and replace the spectrometer to see if this improves the issue. Engineer to arrange a replacement pH electrode to be sent to the customer FAO Rovinder Madhar | Engineer to rest tend site and replace the spectrometer to see if this improves the issue. Engineer to arrange a replacement pH electrode to be sent to the customer FAO Rovinder Madhar | Revisit required | HW | Oxi module | cBil
ABL825 FLEX | Analysator heruntergefahren - Erneuerung CPU - Allgemeine Reinigung - Funktionsüberprüfung etc. | Analysator heruntergefahren - Erneuerung CPU - Allgemeine Reinigung - Funktionsüberprüfung etc. | SWO-102705 | No | Analyzer shut down - Renewal CPU - General cleaning - Function check, etc. | Analyzer shut down - Renewal CPU - General cleaning - Function check, etc. | Defect | HW | DMS_PC | Touch screen
ABL825 FLEX | Datensicherung - Software Upgrade erneut gestartet (keine Fehlerbeseitigung) - Datenwiederherstellung - augenblicklich trotz mehrerer Telefonaten keine Problemlösung (Meldung erfolgt) - Remote Key bis auf weiteres ausgeschaltet ! | Datensicherung - Software Upgrade erneut gestartet (keine Fehlerbeseitigung) - Datenwiederherstellung - augenblicklich trotz mehrerer Telefonaten keine Problemlösung (Meldung erfolgt) - Remote Key bis auf weiteres ausgeschaltet ! | SWO-102706 | No | Backup - Software upgrade restarted (no bug fix) - Data recovery - no problem solving (message is made) immediately despite several phone calls - Remote Key turned off until further notice ! | Backup - Software upgrade restarted (no bug fix) - Data recovery - no problem solving (message is made) immediately despite several phone calls - Remote Key turned off until further notice ! | Unknown | SW | His/lis | NetOp
ABL825 FLEX | Checked the data logs for current failures. | Performed a inlet calibration and ran several test samples to verify the problem was solved. | D0C48A1F12566CCB | No | Checked the data logs for current failures. | Performed a inlet calibration and ran several test samples to verify the problem was solved. | Unknown | HW | Inlet module | Probe cal.
ABL825 FLEX | Oximetrie ausgebaut und Steckverbindung überprüft, Neonlampe ausgebaut und Kabel überprüft, wieder zusammengesetzt, Oxicomplete Test weiterhin fehlerhaft, Rep abgebrochen um Material aus Büro(Berlin)zu holen, nach Wiederankunft Leihmodul Oxi in das Gerät verbaut und justier, Neustart, Oxi kalibriert und tHb Kal durchgeführt, FlexQ Mischer kalibriert, Kal und QC | Oximetrie ausgebaut und Steckverbindung überprüft, Neonlampe ausgebaut und Kabel überprüft, wieder zusammengesetzt, Oxicomplete Test weiterhin fehlerhaft, Rep abgebrochen um Material aus Büro(Berlin)zu holen, nach Wiederankunft Leihmodul Oxi in das Gerät verbaut und justier, Neustart, Oxi kalibriert und tHb Kal durchgeführt, FlexQ Mischer kalibriert, Kal und QC | SWO-102746 | No | Oximetry removed and plug-in connection checked, neon lamp removed and cable checked, reassembled, Oxicomplete test still faulty, Rep aborted to fetch material from office(Berlin), after re-arrival loan module Oxi installed in the device and adjustment, restart, Oxi calibrated and tHb Kal performed, FlexQ mixer calibrated, Kal and QC | Oximetry removed and plug-in connection checked, neon lamp removed and cable checked, reassembled, Oxicomplete test still faulty, Rep aborted to fetch material from office(Berlin), after re-arrival loan module Oxi installed in the device and adjustment, restart, Oxi calibrated and tHb Kal performed, FlexQ mixer calibrated, Kal and QC | Defect | HW | Oxi module | Oxi module
ABL800 BASIC | 更换配件910-320 修复仪器（by andrew） | 更换配件910-320 修复仪器（by andrew） | SWO-102826 | No | Replace accessories 910-320 repair instrument (by andrew) | Replace accessories 910-320 repair instrument (by andrew) | Defect | HW | DMS_PC | Touch screen
ABL815 FLEX | Reinigungsprozedur gem. Anweisung RMED bei neu installierter Lac-Elektrode (neu: UP-04, alt: UP-06), Kal u. QC i.O. | Dateien überprüft, Lac-Elektrode ersetzt (alt: UP06, neu: UP04), Weiterhin Kal-Empfindlichkeitsfehler, Rücktausch Elektrode, Membranierung mit anderem Lot (R2982), Kal u. QC i.O. | SWO-102654 | No | Cleaning procedure according to instruction RMED for newly installed Lac electrode (new: UP-04, old: UP-06), Kal & QC i.O. | Files checked, Lac electrode replaced (old: UP06, new: UP04), Further cal sensitivity error, exchange electrode, membraneing with other solder (R2982), Kal & QC i.O. | Assisted Replacement | Electrode | Membrane | Lac | Sensitivity
AQT90 FLEX | Removed all empty cartridges and installed a new full one. | Removed all empty cartridges and installed a new full one. | SWO-102664 | No | Removed all empty cartridges and installed a new full one. | Removed all empty cartridges and installed a new full one. | Defect | Consumables | Test Kit | Cartridge | 3553 (Punch out)
TCM400 | Reset the electrode. Reset the modules. membrane replace. observe the calibration & reading. | Reset the electrode. Reset the modules. membrane replace. observe the calibration & reading. | SWO-103007 | No | Reset the electrode. Reset the modules. membrane replace. observe the calibration & reading. | Reset the electrode. Reset the modules. membrane replace. observe the calibration & reading.
ABL80 CO-OX | Delivered new SC and SP which has been calibrating with pCO2 values in middle of range. | Exchanged new SP and SC with those in customers analyzer, pCO2 problem transfers with SC SP. Analyzer QC values check in middle of range with the new SC and SP. Externl QC checks OK. Radiometer withh check with customer after the weekend for further errors and QC results. | D0C5470B881D5487 | No | Delivered new SC and SP which has been calibrating with pCO2 values in middle of range. | Exchanged new SP and SC with those in customers analyzer, pCO2 problem transfers with SC SP. Analyzer QC values check in middle of range with the new SC and SP. Externl QC checks OK. Radiometer withh check with customer after the weekend for further errors and QC results.
ABL80 CO-OX | The analyzer would not run a System cycle | I checked all tubings and replaced the hemolyzer | D0C54ADA881E6642 | No | The analyzer would not run a System cycle | I checked all tubings and replaced the hemolyzer
ABL90 FLEX | Software komplett neu installiert tHb Kalibriert, Gerät konfiguriert Gerät kalibriert und überprüft | Software komplett neu installiert tHb Kalibriert, Gerät konfiguriert Gerät kalibriert und überprüft | SWO-103075 | No | Software completely reinstalled tHb calibrated, device configured device calibrated and checked | Software completely reinstalled tHb calibrated, device configured device calibrated and checked | Adjustment | HW | Oxi module | tHb Cal
ABL90 | OBF - Initial Sensor Cassette not reading pCO2, and Glu. Issue remained with Glu sporadic Sensitivity issues. | First Sensor Cassette replaced to resolve pCO2 and Glu issues. -pCO2 and Glu resolved for testing purposes. Second Sensor Cassette replaced as Glu continued to spike in Sensitivity. - Verified pin contacts and ran verification with Sensory Simulator, pO2 sensor readings failed. Third Sensor Cassette replaced due to leakage from top of Cassette. Replaced Sensory Interface board, ran sensor simulator, all values returned as pass. Reran VK kit to all levels passed, no Glu issues. | D0C52E748811CADA | No | OBF - Initial Sensor Cassette not reading pCO2, and Glu. Issue remained with Glu sporadic Sensitivity issues. | First Sensor Cassette replaced to resolve pCO2 and Glu issues. -pCO2 and Glu resolved for testing purposes. Second Sensor Cassette replaced as Glu continued to spike in Sensitivity. - Verified pin contacts and ran verification with Sensory Simulator, pO2 sensor readings failed. Third Sensor Cassette replaced due to leakage from top of Cassette. Replaced Sensory Interface board, ran sensor simulator, all values returned as pass. Reran VK kit to all levels passed, no Glu issues. | Defect | HW | Basic Unit | Sensor Interface PCB
ABL90 FLEX | 更换新试剂包后解决 | 更换新试剂包后解决 | SWO-103218 | No | Resolve after replacing a new reagent kit | Resolve after replacing a new reagent kit
AQT90 FLEX | 清洁处理盘和光学单元模块灯 | 清洁处理盘和光学单元模块灯 | SWO-102470 | No | Cleaning the processing disc and optical unit module lamp | Cleaning the processing disc and optical unit module lamp | Defect | HW | Process Module | Optical Unit | 131
AQT90 FLEX | 更换910-330一个正常。(by andrew) | 更换910-330一个正常。(by andrew) | SWO-102180 | No | Replace the 910-330 one normal. (by andrew) | Replace the 910-330 one normal. (by andrew) | Defect | HW | Analyzer Control Module | CPU
ABL835 FLEX | Revisión normal del equipo con limpieza de la cámara de medida y salida del desperdicio, comprobación del módulo oxi, realización de operaciones de mantenimiento y prueba con controles. | B | D0C55D258853E1DC | Yes | Normal equipment check with measuring chamber cleaning and waste output, oxy module check, maintenance operations and testing with controls. |  | Re classified RMED | PM
ABL800 BASIC | Replace Lac membrane by following procedure and done cleaning then calibrate which is OK. | Replace Lac membrane by following procedure and done cleaning then calibrate which is OK. | SWO-102993 | No | Replace Lac membrane by following procedure and done cleaning then calibrate which is OK. | Replace Lac membrane by following procedure and done cleaning then calibrate which is OK. | Assisted Replacement | Electrode | Membrane | Lac | Sensitivity
ABL825 FLEX | Fit new inlet nipple on syringe side, new inlet gasket, replace clear tube from waste nipple to tubing organiser, soak clear tubing to waste. | B | SWO-103102 | No | Fit new inlet nipple on syringe side, new inlet gasket, replace clear tube from waste nipple to tubing organiser, soak clear tubing to waste. |  | Defect | HW | Inlet module | Leaking | Nipple kit
ABL800 BASIC | 装置メンテナンス必要。キャリブレーション中にリンス溶液、cal1溶液がありませんでした。リンス溶液、cal1,cal2,クリーニング溶液交換実施。QC測定実施。 | 装置メンテナンス必要。キャリブレーション中にリンス溶液、cal1溶液がありませんでした。リンス溶液、cal1,cal2,クリーニング溶液交換実施。QC測定実施。 | SWO-103338 | No | Equipment maintenance required. There was no rinse solution during calibration, cal1 solution. Rinsing solution, cal1, cal2, cleaning solution exchange. QC measurement. | Equipment maintenance required. There was no rinse solution during calibration, cal1 solution. Rinsing solution, cal1, cal2, cleaning solution exchange. QC measurement. | Assisted Replacement | Consumables | Solution Bottle | Cal 1
ABL835 FLEX |  | B | D0C73D2D743833BC | No |  |  | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
ABL825 FLEX | Error codes 538, 539 and 541. | Replaced pH BG module pcb.Reloaded software. Replaced inlet gasket. Replaced tubings. | D0C5ACE6883D4C6F | No | Error codes 538, 539 and 541. | Replaced pH BG module pcb. Reloaded software. Replaced inlet gasket. Replaced tubings. | Defect | HW | pH/BG Module | PCB
AQT90 FLEX | Perform TnT calibration. now machine is running ok | Perform TnT calibration. now machine is running ok | SWO-104337 | No | Perform TnT calibration. now machine is running ok | Perform TnT calibration. now machine is running ok | Assisted Replacement | Consumables | Test Kit | CAL | TnT
ABL90 FLEX | 更换了新试剂包后又更换了新测试卡解决故障 | 更换了新试剂包后又更换了新测试卡解决故障 | SWO-103181 | No | New test card was replaced and new test card was replaced | New test card was replaced and new test card was replaced
ABL90 FLEX | 更换新试剂包修复仪器 | 更换新试剂包修复仪器 | SWO-103194 | No | Replace the new reagent kit repair instrument | Replace the new reagent kit repair instrument | Assisted Replacement | SP | SP | LS error | 1176
ABL80 FLEX BASIC | 更换测试卡解决 | 更换测试卡解决 | SWO-104032 | No | Replace the test card to resolve | Replace the test card to resolve
ABL80 FLEX CO-OX | 更换新测试卡修复仪器 | 更换新测试卡修复仪器 | SWO-103496 | No | Replace the new test card repair instrument | Replace the new test card repair instrument
ABL90 FLEX | Gerät wie folgt vorgefunden: LP ist nicht korrekt installiert - es ist nicht eingerastet, daher kann das LP nicht eingelesen werden. Gerät neu gestartet - LP korrekt eingelegt - LP wird erkannt LP wieder entfernt und von Frau Lücke einlegen lassen - LP wird nicht richtig eingerastet - daher erscheint wieder ein LP Fehler kurze Einweisung durchgeführt Frau Lücke erneuert das LP erneut - LP wird korrekt gelesen - Test wiederholt - LP wird gelesen Kal und QC´s durchgef. - Gerät OK Es lag kein Gerätefehler vor | Servicetechniker wird benötigt | SWO-103291 | No | Device found as follows: LP is not installed correctly - it is not locked, so the LP cannot be read in. Device restarted - LP inserted correctly - LP is detected LP removed again and let by woman insert gap - LP is not snapped properly - therefore appears again an LP error short briefing performed Mrs. Gap renews the LP again - LP is read correctly - Test wi ederholt - LP is read Kal and QC's. - Device OK There was no device error | Service technician is needed | User error | Installation | SP Activation | SP Activation
ABL90 |  | B | D0C9258DEA198FE0 | No |  |  | Assisted Replacement | SP | SP | Pump cal error
AQT90 FLEX | 清洁光单 | 清洁光单 | SWO-103292 | No | Cleaning the light sheet | Cleaning the light sheet | Cleaned | HW | Process Module | Optical Unit | 124
AQT90 FLEX |  | B | SWO-103231 | No |  |  | Cleaned | HW | Analyzer Control Module | Barcode reader, display
ABL827 FLEX | Installed the analyser and copied the set-up from the existing analyser. Samples will be put through this replacement analyser while the existing analyser will be left running to see what difference this makes to the Lactate. | Installed the analyser and copied the set-up from the existing analyser. Samples will be put through this replacement analyser while the existing analyser will be left running to see what difference this makes to the Lactate. | SWO-103490 | No | Installed the analyser and copied the set-up from the existing analyser. Samples will be put through this replacement analyser while the existing analyser will be left running to see what difference this makes to the Lactate. | Installed the analyser and copied the set-up from the existing analyser. Samples will be put through this replacement analyser while the existing analyser will be left running to see what difference this makes to the Lactate. | Assisted Replacement | Electrode | Electrode | Lac | Calibration
ABL825 FLEX | Scanner ersetzt, Funktionskontrolle | Scanner ersetzt, Funktionskontrolle | SWO-103430 | No | Scanner replaced, function control | Scanner replaced, function control | Defect | HW | Basic unit | Barcode scanner
ABL835 FLEX | tube from preheater to ph module had hole in it. user also stated that cannot aspirate cap samples | replaced rubber tube from customer stock everything worked fine after that. | D0C921FAEA629B47 | No | tube from preheater to ph module had hole in it. user also stated that cannot aspirate cap samples | replaced rubber tube from customer stock everything worked fine after that. | Defect | HW | Inlet module | Tubings
ABL825 FLEX | Customer cleaned measuring chamber and remembraned the pCO2 electrode. Checked gas flow and pCO2 electrode channel. | Replaced the pH electrode seal ring and installed software version 6.16. | D0C92945EA5D5EC1 | No | Customer cleaned measuring chamber and remembraned the pCO2 electrode. Checked gas flow and pCO2 electrode channel. | Replaced the pH electrode seal ring and installed software version 6.16. | Assisted Replacement | Electrode | Membrane | pCO2
ABL800 BASIC | replaced with new display unit and calibrated touch screen now touch screen is working. adjusted the halogen intensity and done 2 point calibration and patient sample all results found with in the range. | replaced with new display unit and calibrated touch screen now touch screen is working. adjusted the halogen intensity and done 2 point calibration and patient sample all results found with in the range. | SWO-103262 | No | replaced with new display unit and calibrated touch screen now touch screen is working. adjusted the halogen intensity and done 2 point calibration and patient sample all results found with in the range. | replaced with new display unit and calibrated touch screen now touch screen is working. adjusted the halogen intensity and done 2 point calibration and patient sample all results found with in the range. | Defect | HW | DMS_PC | Touch screen
ABL815 FLEX | Workaround gem. RMED durchgeführt, Kalu. QC i.O. | B | SWO-103293 | No | Workaround gem. RMED durchgeführt, Kalu. QC i.O. |  | Cleaned | Electrode | Electrode | Lac
ABL815 FLEX | Workaround RMED durchgeführt, akz. Kal abgewartet, Kal u. QC i.O. | B | SWO-103294 | No | Workaround RMED, anc. Kal waited, Kal & QC i.O. |  | Cleaned | Electrode | Electrode | Lac
AQT90 FLEX | Analysator geöffnet. Driverbellts erneuert. Ein Sicherungsring der Motorwelle fehlt, ersetzt. Initialisierung. Analysator messbereit. | FSE-Einsatz erforderlich | SWO-103435 | No | Analyzer open. Driverbellts renewed. A fuse ring of the motor shaft is missing, replaced. Initialization. Analyzer ready to measure. | FSE deployment required | Defect | HW | Cover and Chassis | Inlet Door | 5784
TCM/Tosca | Arrived on site after customer had the machine returned to the hospital from patient home. Switched analyser on and allowed to calibrate. Calibration took a long time and then failed asking to be remembraned. Removed sensor from calibration bay cleaned membrane with paper towel dampened with warm water to remove and dust and excess gel. Cleaned the sensor bay and then reseated. Tried recalibrating and this time it passed after a long time. Mems had a pack with contact gel, electrolyte solution and membranes in it. Removed membrane and tried to. Fit new one but found the membranes they had were for sensor 82 and not the sensor 92. Cleaned the sensor and tried to refit the membrane I had removed but the calibration will not pass now. Suggest remembraning and retesting when patient receives. Could not do this here as no consumables available. | Arrived on site after customer had the machine returned to the hospital from patient home. Switched analyser on and allowed to calibrate. Calibration took a long time and then failed asking to be remembraned. Removed sensor from calibration bay cleaned membrane with paper towel dampened with warm water to remove and dust and excess gel. Cleaned the sensor bay and then reseated. Tried recalibrating and this time it passed after a long time. Mems had a pack with contact gel, electrolyte solution and membranes in it. Removed membrane and tried to. Fit new one but found the membranes they had were for sensor 82 and not the sensor 92. Cleaned the sensor and tried to refit the membrane I had removed but the calibration will not pass now. Suggest remembraning and retesting when patient receives. Could not do this here as no consumables available. | SWO-103346 | No | Arrived on site after customer had the machine returned to the hospital from patient home. Switched analyser on and allowed to calibrate. Calibration took a long time and then failed asking to be remembraned. Removed sensor from calibration bay cleaned membrane with paper towel dampened with warm water to remove and dust and excess gel. Cleaned the sensor bay and then reseated. Tried recalibrating and this time it passed after a long time. Mems had a pack with contact gel, electrolyte solution and membranes in it. Removed membrane and tried to. Fit new one but found the membranes they had were for sensor 82 and not the sensor 92. Cleaned the sensor and tried to refit the membrane I had removed but the calibration will not pass now. Suggest remembraning and retesting when patient receives. Could not do this here as no consumables available. | Arrived on site after customer had the machine returned to the hospital from patient home. Switched analyser on and allowed to calibrate. Calibration took a long time and then failed asking to be remembraned. Removed sensor from calibration bay cleaned membrane with paper towel dampened with warm water to remove and dust and excess gel. Cleaned the sensor bay and then reseated. Tried recalibrating and this time it passed after a long time. Mems had a pack with contact gel, electrolyte solution and membranes in it. Removed membrane and tried to. Fit new one but found the membranes they had were for sensor 82 and not the sensor 92. Cleaned the sensor and tried to refit the membrane I had removed but the calibration will not pass now. Suggest remembraning and retesting when patient receives. Could not do this here as no consumables available.
TCM400 | reset the cable. reset the modules. reset system. recalibrate the membrane. observe the system. | reset the cable. reset the modules. reset system. recalibrate the membrane. observe the system. | SWO-104191 | No | reset the cable. reset the modules. reset system. recalibrate the membrane. observe the system. | reset the cable. reset the modules. reset system. recalibrate the membrane. observe the system.
ABL80 FLEX BASIC | 更换测试卡正常。 | 更换测试卡正常。 | SWO-103736 | No | The replacement test card is OK. | The replacement test card is OK.
ABL80 FLEX BASIC | 更换新测试卡修复仪器 | 更换新测试卡修复仪器 | SWO-104223 | No | Replace the new test card repair instrument | Replace the new test card repair instrument
ABL80 FLEX BASIC | 更换新测试卡修复仪器 | 更换新测试卡修复仪器 | SWO-104225 | No | Replace the new test card repair instrument | Replace the new test card repair instrument
ABL90 FLEX | Fehlersuche, Fließwege gespült, div. Serviceprogr. durchgef., Anwender im Umgang mit den Gerinselfängern geschult (knicken bei zu starkem Drücken ab), Servicebox hinterlegt, Kal. + QC i.O. | Fehlersuche; Fließwege gespült; div. Serviceprogr. durchgef.; Anwender im Umgang mit den Gerinselfängern geschult (knicken bei zu starkem Drücken ab); Servicebox hinterlegt; Kal. + QC i.O. | SWO-104129 | No | Troubleshooting; Flow paths rinsed; div. Serviceprogr. through it; Users trained in handling the Gerinsel catchers (buckling off when pushing too much); Service box deposited; Kal. + QC in the first place | Troubleshooting; Flow paths rinsed; div. Serviceprogr. through it; Users trained in handling the Gerinsel catchers (buckling off when pushing too much); Service box deposited; Kal. + QC in the first place | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
ABL90 FLEX | Lösungspack erneuert. Fehler behoben. Einlassdichtung mit starkem Verschleiß erneuert, und 1x als Ersatz bereitgestellt. Einlasssonde als Ersatz bereitgestellt. Einweisung in das Reklamationsverfahren gegeben, Anleitung 1 und 2 bereitgestellt. | Lösungspack erneuert. Fehler behoben. Einlassdichtung mit starkem Verschleiß erneuert, und 1x als Ersatz bereitgestellt. Einlasssonde als Ersatz bereitgestellt. Einweisung in das Reklamationsverfahren gegeben, Anleitung 1 und 2 bereitgestellt. | SWO-103700 | No | Solution pack renewed. Fixed a bug. Inlet seal renewed with heavy wear, and provided 1x as a replacement. inlet probe as a replacement. instruction in the complaint procedure, instructions 1 and 2 provided. | Solution pack renewed. Fixed a bug. Inlet seal renewed with heavy wear, and provided 1x as a replacement. inlet probe as a replacement. instruction in the complaint procedure, instructions 1 and 2 provided. | Assisted Replacement | SP | SP | pO2
ABL90 | No investigation needed. | Installed Rals patch for qc ranges. | D0CA42DB600BE8CD | No | No investigation needed. | Installed Rals patch for qc ranges. | Setup | His/Lis | His/lis | His/Lis
ABL90 | No investigation needed. | Installed Rals patch for qc ranges. | D0CA43B5600DE8D0 | No | No investigation needed. | Installed Rals patch for qc ranges. | Setup | His/Lis | His/lis | His/Lis
ABL90 | Error 1352. Tried back flush procedure. No blockage detected. Blockage may have eventually worked it's way out of system. | Perform all service programs in Service mode. All passed. | D0CA7AAB6022DC79 | No | Error 1352. Tried back flush procedure. No blockage detected. Blockage may have eventually worked it's way out of system. | Perform all service programs in Service mode. All passed. | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
ABL90 FLEX PLUS | Arrived on site and attended analyser. program. Looked at locking hook inside analyser and this was in a raised position. Rebooted analyser and entered service performed eject, prepare and reinsertion of pack. Exited service program and performed the replacement procedure for solution pack. This worked correctly and accepted pack. Analyser started a 55 minute startup. Analyser went to intervention for inlet gasket. I performed gasket clean and refit as no other gaskets were available. Next the solution pack intervention screen which I completed. Then the analyser flush sequence. I performed this and as a precaution performed the sensor cassette procedure. I clicked on test again once all of these were done. Analyser started 55 minute startup. Analyser came to solution pack change screen so I went to collect new sp. when I arrived back I activated the sp and then attempted the replacement procedure. Button to continue was not working. Rebooted analyser and when it powered back up, the startup worked correctly with current pack installed. Left this conditioning and new activated pack next to analyser as current pack only has 39 actions left and will need replacing today. All parameters green except pco2. Customer to monitor. | Arrived on site and attended analyser. program. Looked at locking hook inside analyser and this was in a raised position. Rebooted analyser and entered service performed eject, prepare and reinsertion of pack. Exited service program and performed the replacement procedure for solution pack. This worked correctly and accepted pack. Analyser started a 55 minute startup. Analyser went to intervention for inlet gasket. I performed gasket clean and refit as no other gaskets were available. Next the solution pack intervention screen which I completed. Then the analyser flush sequence. I performed this and as a precaution performed the sensor cassette procedure. I clicked on test again once all of these were done. Analyser started 55 minute startup. Analyser came to solution pack change screen so I went to collect new sp. when I arrived back I activated the sp and then attempted the replacement procedure. Button to continue was not working. Rebooted analyser and when it powered back up, the startup worked correctly with current pack installed. Left this conditioning and new activated pack next to analyser as current pack only has 39 actions left and will need replacing today. All parameters green except pco2. Customer to monitor. | SWO-103584 | No | Arrived on site and attended analyser. program. Looked at locking hook inside analyser and this was in a raised position. Rebooted analyser and entered service performed eject, prepare and reinsertion of pack. Exited service program and performed the replacement procedure for solution pack. This worked correctly and accepted pack. Analyser started a 55 minute startup. Analyser went to intervention for inlet gasket. I performed gasket clean and refit as no other gaskets were available. Next the solution pack intervention screen which I completed. Then the analyser flush sequence. I performed this and as a precaution performed the sensor cassette procedure. I clicked on test again once all of these were done. Analyser started 55 minute startup. Analyser came to solution pack change screen so I went to collect new sp. when I arrived back I activated the sp and then attempted the replacement procedure. Button to continue was not working. Rebooted analyser and when it powered back up, the startup worked correctly with current pack installed. Left this conditioning and new activated pack next to analyser as current pack only has 39 actions left and will need replacing today. All parameters green except pco2. Customer to monitor. | Arrived on site and attended analyser. program. Looked at locking hook inside analyser and this was in a raised position. Rebooted analyser and entered service performed eject, prepare and reinsertion of pack. Exited service program and performed the replacement procedure for solution pack. This worked correctly and accepted pack. Analyser started a 55 minute startup. Analyser went to intervention for inlet gasket. I performed gasket clean and refit as no other gaskets were available. Next the solution pack intervention screen which I completed. Then the analyser flush sequence. I performed this and as a precaution performed the sensor cassette procedure. I clicked on test again once all of these were done. Analyser started 55 minute startup. Analyser came to solution pack change screen so I went to collect new sp. when I arrived back I activated the sp and then attempted the replacement procedure. Button to continue was not working. Rebooted analyser and when it powered back up, the startup worked correctly with current pack installed. Left this conditioning and new activated pack next to analyser as current pack only has 39 actions left and will need replacing today. All parameters green except pco2. Customer to monitor. | Adjustment | HW | SP Locking mechanism | SP Locking mechanism
ABL825 FLEX |  | B | D0CA62086014AF5C | No |  |  | Defect | HW | Gas module | Pressure Transducer
ABL825 FLEX | Reinigung Kalibrierung | Reinigung Kalibrierung | SWO-103655 | No | Cleaning Calibration | Cleaning Calibration | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
ABL835 FLEX | ＦｌｅｘＱ返却 | ＦｌｅｘＱ返却 | SWO-103999 | No | FlexQ Return | FlexQ Return | Defect | HW | FlexQ Module | FlexQ Module | Replaced
ABL835 FLEX | Ｅｌ Metモジュール返却 | Ｅｌ Metモジュール返却 | SWO-104002 | No | El Met Module Return | El Met Module Return | Assisted Replacement | Electrode | Electrode | Lac | Response error
ABL835 FLEX | En o-ring fra ref. elektroden var ved et uheld landet ved elektrode kontakterne | En o-ring fra ref. elektroden var ved et uheld landet ved elektrode kontakterne | SWO-103494 | No | An O-ring from ref. The electrode was accidentally landed by the electrode contacts | An O-ring from ref. The electrode was accidentally landed by the electrode contacts | User error | Electrode | Electrode | Reference | Sealing Ring
ABL810 FLEX | V1-V3 ausgebaut und gereinigt Abfallpumpenschläuche gewechselt Einlaßkal, Verbindungsschlauch pH Modul Einlaßmodul und Einlaßschlauch gewechselt | V1-V3 ausgebaut und gereinigt Abfallpumpenschläuche gewechselt Einlaßkal, Verbindungsschlauch pH Modul Einlaßmodul und Einlaßschlauch gewechselt | SWO-103704 | No | V1-V3 removed and cleaned Waste pump hoses changed inlet scale, connecting hose pH module inlet module and inlet hose changed | V1-V3 removed and cleaned Waste pump hoses changed inlet scale, connecting hose pH module inlet module and inlet hose changed | Defect | HW | Fluidic manifold | Waste tube
ABL80 FLEX BASIC | 更换测试卡修复仪器 | 更换测试卡修复仪器 | SWO-103837 | No | Replace the test card repair instrument | Replace the test card repair instrument
ABL80 FLEX BASIC |  | B | SWO-103802 | No
ABL80 CO-OX | confirmed problem. | Replaced the barcode scanner. It is still not working. The analyzer needs to be returned to Radiometer for Depot repair. | D0CBAE20D60F53A5 | No | confirmed problem. | Replaced the barcode scanner. It is still not working. The analyzer needs to be returned to Radiometer for Depot repair.
ABL80 FLEX BASIC |  | B | SWO-103799 | No
ABL90 FLEX | Einlassdichtung aus KDB-Bestand ersetzt | Einlassdichtung aus KDB-Bestand ersetzt | SWO-103892 | No | Inlet seal replaced by KDB stock | Inlet seal replaced by KDB stock | Assisted Replacement | SC | SC | Lac
ABL90 FLEX | SP csere | SP csere | SWO-103813 | Yes | SP replacement | SP replacement | Re classified RMED | PER/Phone
ABL90 FLEX PLUS |  | B | D0CBD1CFD61C1265 | Yes |  |  | Re classified RMED | PER
AQT90 FLEX |  | B | SWO-103747 | No |  |  | Cleaned | HW | Storage Module | Barcode reader, Storage wheel
AQT90 FLEX | 更换新处理包安装后解决 | 更换新处理包安装后解决 | SWO-103740 | No | Resolve after replacing a new processing pack is installed | Resolve after replacing a new processing pack is installed | Defect | Consumables | SP | SP
AQT90 FLEX |  | B | SWO-103733 | No |  |  | Not Cat | Not cat. | Not cat. | Not cat.
AQT90 FLEX | Rerun QCs. | Rerun QCs. | SWO-103955 | No | Rerun QCs. | Rerun QCs. | No fault found | Performance | LQC Kit | LQC
ABL800 BASIC | So perform protein removal and clean the electrode. Again perform 2 point calibration for 2 times and values are within the range. Now unit is working properly. | So perform protein removal and clean the electrode. Again perform 2 point calibration for 2 times and values are within the range. Now unit is working properly. | SWO-104459 | No | So perform protein removal and clean the electrode. Again perform 2 point calibration for 2 times and values are within the range. Now unit is working properly. | So perform protein removal and clean the electrode. Again perform 2 point calibration for 2 times and values are within the range. Now unit is working properly. | Cleaned | Electrode | Electrode | Ca
ABL837 FLEX | Entlüftungsanschluss gespült. Gas2-Ventil erneuert. | Entlüftungsanschluss gespült. Gas2-Ventil erneuert. | SWO-103931 | No | Vent connection flushed. Gas2 valve renewed. | Vent connection flushed. Gas2 valve renewed. | Defect | HW | Gas module | Flow regulator 2
ABL837 FLEX | AU-Software über auinstall installiert. Erneuter Abbruch mit Fehlermldung: Download D: ABL700 AUEngine sw appbase.out appbase.out failed . AUE->AUS 461 14 RO_RTC_REQUEST 0002 Logoff Installation Utility 3 0002 Accepting Installation Utility 3 Neustart DMS - erneute Installation. Nach Beendigung der Installation wurde Verbindung zur DMS nicht aufgebaut, ftp geöffnet. erneuter Neustart DMS nach Wartezeit von Ca 10min. Nach erneutem Booten wurde Verbindung zur AU korrekt aufgebaut. | AU-Software über auinstall installiert. Erneuter Abbruch mit Fehlermldung: Download D: ABL700 AUEngine sw appbase.out appbase.out failed . AUE->AUS 461 14 RO_RTC_REQUEST 0002 Logoff Installation Utility 3 0002 Accepting Installation Utility 3 Neustart DMS - erneute Installation. Nach Beendigung der Installation wurde Verbindung zur DMS nicht aufgebaut, ftp geöffnet. erneuter Neustart DMS nach Wartezeit von Ca 10min. Nach erneutem Booten wurde Verbindung zur AU korrekt aufgebaut. | SWO-103929 | No | AU software installed via auinstall. Re-cancellation with error: Download D: ABL700 AUEngine sw appbase.out appbase.out failed . AUE->FROM 461 14 RO_RTC_REQUEST 0002 Logoff Installation Utility 3 0002 Accepting Installation Utility 3 Reboot DMS - Reinstall. After the installation was completed, the connection to the DMS was not established, ftp was opened. restart DMS after waiting time of Approx 10min. After rebooting, the connection to the AU was established correctly. | AU software installed via auinstall. Re-cancellation with error: Download D: ABL700 AUEngine sw appbase.out appbase.out failed . AUE->FROM 461 14 RO_RTC_REQUEST 0002 Logoff Installation Utility 3 0002 Accepting Installation Utility 3 Reboot DMS - Reinstall. After the installation was completed, the connection to the DMS was not established, ftp was opened. restart DMS after waiting time of Approx 10min. After rebooting, the connection to the AU was established correctly. | Unknown | SW | Upgrade | Upgrade
ABL800 BASIC | gas 1 empty .causing calibration error. guide the customer for gas replacement . | gas 1 empty .causing calibration error. guide the customer for gas replacement . | SWO-104187 | No | gas 1 empty .causing calibration error. guide the customer for gas replacement . | gas 1 empty .causing calibration error. guide the customer for gas replacement . | Assisted Replacement | Gas 1 | Gas 1 | Gas 1
ABL810 FLEX | AU PCB inklusive Gasmodul aus GS Bestand ersetzt. Abweichenden Zählerstand beachten!!! | Da die Verantwortlichen vom Labor dabei sein wollen wird der Einsatz auf kommenden Mittwoch gelegt. Es wird dann ein anderes PH BG Modul eingebaut und die Ventiele des Gasmoduls gewechselt um den Fehler endlich in den Griff zu bekommen. | SWO-103811 | No | AU PCB including gas module from GS stock replaced. Observe the different meter reading!!! | As the lab officials want to be there, the deployment will be put on hold next Wednesday. Another PH BG module is then installed and the valves of the gas module are changed to finally get a grip on the error. | Revisit required | Electrode | Electrode | pCO2
ABL820 FLEX | Found PH electrode to be faulty. | Ordered and replaced Ph electrode. Performed necessary calibrations | D0CBFF55D65684A9 | No | Found PH electrode to be faulty. | Ordered and replaced Ph electrode. Performed necessary calibrations | Assisted Replacement | Electrode | Electrode | pH
ABL80 FLEX BASIC | Performed zeroing of all channels with zeroing cassette | Performed zeroing of all channels with zeroing cassette | SWO-104026 | No | Performed zeroing of all channels with zeroing cassette | Performed zeroing of all channels with zeroing cassette
ABL80 FLEX BASIC | Fehlersuche, neue SC installiert, Initialisierung -> Flüssigkeitstransportfehler, LP gew., alles i.O. | Fehlersuche; neue SC installiert; Initialisierung -> Flüssigkeitstransportfehler; LP gew.; alles i.O. | SWO-104151 | No | Troubleshooting; new SC installed; Initialization -> Liquid transport error; LP. everything in the | Troubleshooting; new SC installed; Initialization -> Liquid transport error; LP. everything in the
ABL90 FLEX | Kundengespräch, Fehlersuche, Justagen, Kalibrationen und Qualitätskontrollen durchgeführt. | Kundengespräch, Fehlersuche, Justagen, Kalibrationen und Qualitätskontrollen durchgeführt. | SWO-104066 | No | Customer consultation, troubleshooting, adjustments, calibrations and quality controls. | Customer consultation, troubleshooting, adjustments, calibrations and quality controls. | Assisted Replacement | SC | SC | Glu
ABL90 | Copied all data logs, did a service dump. | No correction yet. Copied logs for CQC case. | D0CCB91C4C613037 | No | Copied all data logs, did a service dump. | No correction yet. Copied logs for CQC case. | Unknown | Revisit required | tHb Performance | tHb Performance
ABL90 FLEX | Einladichtung erneuert Kalibration und QC durchgeführt | Sensorkassette erneuert (Kundenmaterial) | SWO-104067 | No | Single seal renews calibration and QC performed | Sensor cartridge renewed (customer material) | Defect | HW | Inlet module | Inlet gasket
ABL90 FLEX | Sensorkassette erneuert (Kundenmaterial) | Sensorkassette erneuert (Kundenmaterial) | SWO-104170 | No | Sensor cartridge renewed (customer material) | Sensor cartridge renewed (customer material) | Assisted Replacement | SP | SP | FT
ABL90 FLEX | Kal wiederholt - Fehler bleibt, Spülen Fließwege, Kal wiederholt - Kal i.O., QC i.O. | B | SWO-104143 | No | Kal repeated - error remains, flushing flow paths, Kal repeated - Kal i.O., QC i.O. |  | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
ABL90 | REPLACED CASSETTE AS COURTESY AND NOTIFIED GINA (CAS) AS POSIBLE TRAINING ISSUE | B | D0CCD9134C086C78 | No | REPLACED CASSETTE AS COURTESY AND NOTIFIED GINA (CAS) AS POSIBLE TRAINING ISSUE |  | Assisted Replacement | SC | SC | Wetability
ABL90 FLEX PLUS | Arrived on site and checked the patient results log to confirm the reported errors. Found 1070 sensor response error and also 1228 correction for lactate contains errors from K+, Na+, Ca2+. Entered service programs and removed the solution pack. Powered down the analyser and took out the wet section of the analyser for investigation. Manually opened the sensor bay door and removed the sensor cassette. Checked for signs of damage or spillage spray around the sensor Bay Area paying particular attention to the sensor pins above the sensor cassette. There was no sign of liquid damage or physical damage. All connector pins were straight and undamaged. Tried to push each pin backwards feeling for any resistance or clicking. None found. Checked that they were all straight and not loose. Refitted sensor cassette and removed the top sensor bay door before manually winding the bay closed again. This was to check alignment of the pins with the connection points on the sensor cassette. These were all correctly positioned and were not intruding on the plastic parts of the cassette. Also checked that the pins were in fact receding as they came into contact with the connection points when the bay closed. This was also functioning correctly. Cleaned the connector pins with a cotton bud slightly dampened with di water and then dried again with Joseph paper. Refitted all parts into analyser and powered up. Ran calibrations and qcs after the warm start of 10 minutes had finished. Calibration and qcs passed without problem and the qc plots looked well. No patient sample available to be run and checked. Customer to monitor. | Arrived on site and checked the patient results log to confirm the reported errors. Found 1070 sensor response error and also 1228 correction for lactate contains errors from K+, Na+, Ca2+. Entered service programs and removed the solution pack. Powered down the analyser and took out the wet section of the analyser for investigation. Manually opened the sensor bay door and removed the sensor cassette. Checked for signs of damage or spillage spray around the sensor Bay Area paying particular attention to the sensor pins above the sensor cassette. There was no sign of liquid damage or physical damage. All connector pins were straight and undamaged. Tried to push each pin backwards feeling for any resistance or clicking. None found. Checked that they were all straight and not loose. Refitted sensor cassette and removed the top sensor bay door before manually winding the bay closed again. This was to check alignment of the pins with the connection points on the sensor cassette. These were all correctly positioned and were not intruding on the plastic parts of the cassette. Also checked that the pins were in fact receding as they came into contact with the connection points when the bay closed. This was also functioning correctly. Cleaned the connector pins with a cotton bud slightly dampened with di water and then dried again with Joseph paper. Refitted all parts into analyser and powered up. Ran calibrations and qcs after the warm start of 10 minutes had finished. Calibration and qcs passed without problem and the qc plots looked well. No patient sample available to be run and checked. Customer to monitor. | SWO-104115 | No | Arrived on site and checked the patient results log to confirm the reported errors. Found 1070 sensor response error and also 1228 correction for lactate contains errors from K+, Na+, Ca2+. Entered service programs and removed the solution pack. Powered down the analyser and took out the wet section of the analyser for investigation. Manually opened the sensor bay door and removed the sensor cassette. Checked for signs of damage or spillage spray around the sensor Bay Area paying particular attention to the sensor pins above the sensor cassette. There was no sign of liquid damage or physical damage. All connector pins were straight and undamaged. Tried to push each pin backwards feeling for any resistance or clicking. None found. Checked that they were all straight and not loose. Refitted sensor cassette and removed the top sensor bay door before manually winding the bay closed again. This was to check alignment of the pins with the connection points on the sensor cassette. These were all correctly positioned and were not intruding on the plastic parts of the cassette. Also checked that the pins were in fact receding as they came into contact with the connection points when the bay closed. This was also functioning correctly. Cleaned the connector pins with a cotton bud slightly dampened with di water and then dried again with Joseph paper. Refitted all parts into analyser and powered up. Ran calibrations and qcs after the warm start of 10 minutes had finished. Calibration and qcs passed without problem and the qc plots looked well. No patient sample available to be run and checked. Customer to monitor. | Arrived on site and checked the patient results log to confirm the reported errors. Found 1070 sensor response error and also 1228 correction for lactate contains errors from K+, Na+, Ca2+. Entered service programs and removed the solution pack. Powered down the analyser and took out the wet section of the analyser for investigation. Manually opened the sensor bay door and removed the sensor cassette. Checked for signs of damage or spillage spray around the sensor Bay Area paying particular attention to the sensor pins above the sensor cassette. There was no sign of liquid damage or physical damage. All connector pins were straight and undamaged. Tried to push each pin backwards feeling for any resistance or clicking. None found. Checked that they were all straight and not loose. Refitted sensor cassette and removed the top sensor bay door before manually winding the bay closed again. This was to check alignment of the pins with the connection points on the sensor cassette. These were all correctly positioned and were not intruding on the plastic parts of the cassette. Also checked that the pins were in fact receding as they came into contact with the connection points when the bay closed. This was also functioning correctly. Cleaned the connector pins with a cotton bud slightly dampened with di water and then dried again with Joseph paper. Refitted all parts into analyser and powered up. Ran calibrations and qcs after the warm start of 10 minutes had finished. Calibration and qcs passed without problem and the qc plots looked well. No patient sample available to be run and checked. Customer to monitor. | Adjustment | HW | Sensor Cassette Contacts | Sensor Cassette Contacts
ABL90 FLEX PLUS |  | B | D0CCE8D14C1D31A3 | Yes |  |  | Re classified RMED | Training of customer
TCM CombiM |  | B | SWO-104038 | No
ABL90 FLEX | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-104196 | No | Repair after replacing a new test card | Repair after replacing a new test card | Assisted Replacement | SC | SC | Reference
AQT90 FLEX |  | B | SWO-104039 | No |  |  | Defect | HW | Process Module | Optical Unit | 135
ABL825 FLEX | Verified the problem. | Replaced the nipple. | D0CCC42B4C42056B | No | Verified the problem. | Replaced the nipple. | Defect | HW | Inlet module | Nipple kit | Nipple kit
ABL835 FLEX | Jeg har vist Lotte hvordan man skifter denne spildstuds. | Jeg har vist Lotte hvordan man skifter denne spildstuds. | SWO-104110 | No | I have shown Lotte how to change this waste studs. | I have shown Lotte how to change this waste studs. | Defect | HW | Inlet module | Nipple kit | Nipple kit
ABL825 FLEX |  | B | D0CCC3704C1D34E3 | No |  |  | Defect | HW | Inlet module | Friction detector unit | Known issue
ABL80 FLEX BASIC |  | B | SWO-104204 | No
ABL80 FLEX BASIC | 更换902-972和903-334修复仪器---jack.chu | 更换902-972和903-334修复仪器---jack.chu | SWO-104073 | No | Replace 902-972 and 903-334 repair instruments---jack.chu | Replace 902-972 and 903-334 repair instruments---jack.chu
ABL80 FLEX CO-OX | ｾﾝｻｶｾｯﾄ持参し交換するが同じｴﾗｰで改善せず、現場にもう1個在庫があり、そのｶｾｯﾄに交換してOK。 | B | SWO-104398 | No | It does not improve by the same error though the sensor cassette is brought and exchanged, there is another stock in the site, and it exchanges it with the cassette and it is OK.
ABL90 FLEX | Fitted new Hemolyzer set up and test ok. Ran QC and the Thb is still failing. Spectrometer is faulty. | Customer has flushed unit and ran Thb calibration this has passed but still failing QC . Engineer required . | SWO-104391 | No | Fitted new Hemolyzer set up and test ok. Ran QC and the Thb is still failing. Spectrometer is faulty. | Customer has flushed unit and ran Thb calibration this has passed but still failing QC . Engineer required . | Unknown | Revisit required | Oxi module | Spectrometer
AQT90 FLEX |  | B | SWO-104181 | No |  |  | Defect | HW | Aspiration Module | Needle Wash Unit | 3412
ABL825 FLEX | Trouble shot and found a bent probe. | Replaced inlet probe. | D0CE4048C2268594 | No | Trouble shot and found a bent probe. | Replaced inlet probe. | Defect | HW | Inlet module | Probe bend | Leak from inlet. // Trouble shot and found a bent probe. // Replaced inlet probe.
ABL827 FLEX | There was a leak in analyzer between pH Bg and EL Met module from grey tubing at pH module. Also same type of leak between EL Met and blank module at grey tubing attached to EL Met module. Was not leaking when arrived. | Replaced grey module connector tubings. Perform zero current calibration in service mode. Ran one level of VK material and got 1032. | D0CE4910C247764C | No | There was a leak in analyzer between pH Bg and EL Met module from grey tubing at pH module. Also same type of leak between EL Met and blank module at grey tubing attached to EL Met module. Was not leaking when arrived. | Replaced grey module connector tubings. Perform zero current calibration in service mode. Ran one level of VK material and got 1032. | Defect | HW | Basic Unit | Tubings
AQT90 FLEX | replaced the needle Assembly & needle sensor | replaced the needle Assembly & needle sensor | SWO-104354 | No | replaced the needle Assembly & needle sensor | replaced the needle Assembly & needle sensor | Defect | HW | Aspiration Module | Bend needle | Pipette 38
ABL80 CO-OX | Found evidence of fluid on the sensor connect board. | Replaced the sensor connect board. | D0CFC2753850C438 | No | Found evidence of fluid on the sensor connect board. | Replaced the sensor connect board.
AQT90 FLEX | 取出断角测试卡后仪器恢复使用 | 取出断角测试卡后仪器恢复使用 | SWO-104472 | No | The instrument is back in use after removing the break-out test card | The instrument is back in use after removing the break-out test card | Defect | Consumables | Test Kit | Cartridge | 5784 (Broken or Label)
AQT90 FLEX | 更换光学单元903-364后修复---jack.chu | 更换光学单元903-364后修复---jack.chu | SWO-104284 | No | Repair ---jack.chu after replacing optical unit 903-364 | Repair ---jack.chu after replacing optical unit 903-364 | Defect | HW | Process Module | Optical Unit | 131
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-105949 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL90 FLEX PLUS |  | B | D0D0A2A5AE13A35A | No |  |  | Assisted Replacement | SC | SC | Glu/Lac/Na
ABL90 FLEX | 更换新试剂包后修复 | 更换新试剂包后修复 | SWO-104522 | No | Repair after replacing a new kit | Repair after replacing a new kit | Assisted Replacement | SP | SP | pCO2
AQT90 FLEX | 取出损坏测试卡，将存储盘复位后修复 | 取出损坏测试卡，将存储盘复位后修复 | SWO-104722 | No | Remove the damaged test card and repair it after the storage tray is reset | Remove the damaged test card and repair it after the storage tray is reset | Defect | Consumables | Test Kit | Cartridge | 5689 (Broken or Label)
AQT90 FLEX | 更换903-028 Cabinet Top PCB | 更换903-028 Cabinet Top PCB | SWO-104308 | No | 更换903-028 Cabinet Top PCB | 更换903-028 Cabinet Top PCB | Defect | HW | Cover and Chassis | PCB, Cabinet | 5784
ABL80 FLEX BASIC |  | 更换了信号板903-334后修复---jack.chu | SWO-104523 | No |  | Repair ---jack.chu after replacing signal board 903-334
ABL90 FLEX | Einlassnadel verbogen, Einlassnadel aus Kundenbestand erneuert. Flüssigkeitstransport im Servicemenü überprüft, in Ordnung. | Einlassnadel verbogen, Einlassnadel aus Kundenbestand erneuert. Flüssigkeitstransport im Servicemenü überprüft, in Ordnung. | SWO-104848 | No | Inlet needle bent, inlet needle renewed from customer base. Liquid transport checked in the service menu, all right. | Inlet needle bent, inlet needle renewed from customer base. Liquid transport checked in the service menu, all right. | Replaced | HW | Inlet Module | Probe bend
ABL90 FLEX | Fehlersuche - WS funktioniert, DMS nicht, ET bestellt und ABL90 abgezogen. | Fehlersuche - WS funktioniert, DMS nicht, ET bestellt und ABL90 abgezogen. | SWO-105406 | No | Troubleshooting - WS works, DMS does not, ET ordered and ABL90 subtracted. | Troubleshooting - WS works, DMS does not, ET ordered and ABL90 subtracted. | Unknown | Removed by RSUB | Analyzer removed | Analyzer removed
ABL90 FLEX PLUS | Techline. RBNL 749718508 ABL90 Flex Plus | Techline. RBNL 749718508 ABL90 Flex Plus | SWO-104849 | Yes | Mr. Techline. RBNL 749718508 ABL90 Flex Plus | Mr. Techline. RBNL 749718508 ABL90 Flex Plus | Re classified RMED | PER | OOB/TL
ABL90 FLEX PLUS | Heb de sensor contacten handmatig een aantal malen ingeduwd met een naalddoosje tegen de probes aan. Hierna afgepoets met alkohol. De sensor doorgespoeld met bloed en geïnstalleerd. Opgestart en gekalibreerd (55 min) zonder problemen. Twee bloedmetingen uitgevoerd. | Heb de sensor contacten handmatig een aantal malen ingeduwd met een naalddoosje tegen de probes aan. Hierna afgepoets met alkohol. De sensor doorgespoeld met bloed en geïnstalleerd. Opgestart en gekalibreerd (55 min) zonder problemen. Twee bloedmetingen uitgevoerd. | SWO-104846 | No | Have manually pushed the sensor contacts a number of times with a needle box against the probes. Hereafter, with Alkohol. The sensor is washed through with blood and installed. Booted and calibrated (55 min) without any problems. Two blood measurements performed. | Have manually pushed the sensor contacts a number of times with a needle box against the probes. Hereafter, with Alkohol. The sensor is washed through with blood and installed. Booted and calibrated (55 min) without any problems. Two blood measurements performed. | Adjustment | HW | SC contact pin | SC contact pin
ABL90 FLEX | 更换测试卡 | 更换测试卡 | SWO-104550 | No | Replace the test card | Replace the test card | Assisted Replacement | SC | SC | Lac - Cal
AQT90 FLEX |  | B | SWO-104570 | No |  |  | Defect | HW | Analyzer Control Module | Barcode reader, display
ABL825 FLEX | Austausch kompl. PC Modul aus ABL825 FLEX 754R1221N005 ( ZK Bad Berka IC 3) Radianceanbindung überprüft, muß noch angebunden werden Einlaufverfahren | Austausch kompl. PC Modul aus ABL825 FLEX 754R1221N005 ( ZK Bad Berka IC 3) Radianceanbindung überprüft, muß noch angebunden werden Einlaufverfahren | SWO-104788 | No | Replacement compl. PC module made of ABL825 FLEX 754R1221N005 ( ZK Bad Berka IC 3) Radiance connection checked, has yet to be connected inlet procedure | Replacement compl. PC module made of ABL825 FLEX 754R1221N005 ( ZK Bad Berka IC 3) Radiance connection checked, has yet to be connected inlet procedure | Defect | HW | DMS_PC | DMS_PC | Barcode - USB
ABL800 BASIC | replace the cylinder.calibrate system. calibration pass. guided customer for replacement. | replace the cylinder.calibrate system. calibration pass. guided customer for replacement. | SWO-106205 | No | replace the cylinder.calibrate system. calibration pass. guided customer for replacement. | replace the cylinder.calibrate system. calibration pass. guided customer for replacement. | Assisted Replacement | Gas 1 | Gas1 | Gas 1 | Custemer should change this by themself
ABL815 FLEX |  | B | D0D4E10E106060FF | No |  |  | Defect | HW | Inlet module | Inlet gasket | Difficult to understand
ABL815 FLEX |  | B | D0D4E28310256102 | No |  |  | Defect | HW | Gas module | Gas 2
ABL815 FLEX | Option Key eingeschaltet - Funktionsüberprüfung etc. | Option Key eingeschaltet - Funktionsüberprüfung etc. | SWO-104987 | No | Option Key switched on - function check etc. | Option Key switched on - function check etc. | Setup | SW | Option key | Option key
AQT90 FLEX | Inlet Wheel demontiert, Lichtschranke gereinigt. | Inlet Wheel demontiert, Lichtschranke gereinigt. | SWO-104755 | No | Inlet Wheel dismantled, light barrier cleaned. | Inlet Wheel dismantled, light barrier cleaned. | Cleaned | HW | Aspiration Module | Inlet Wheel Sensor
ABL80 FLEX Basic version |  | B | D0D3298D9A3A2FA4 | No
ABL80 FLEX Basic version |  | B | D0D3B31B9A602DA2 | No
ABL80 | checked tubing and waste pump. Found it hard to inject water into the solution pack waste port. | Flushed the waste lines and replaced the solution pack. | D0D332359A633D10 | No | checked tubing and waste pump. Found it hard to inject water into the solution pack waste port. | Flushed the waste lines and replaced the solution pack.
ABL90 | After replacement of the sensor cassette (SC), pO2 values are above the acceptable QC range. Lac and glucose have also been activated. | Incorrect SC was installed on the analyzer. SC 946-008, which includes metabolites, was installed. Customer typically uses SC 946-013, with electrolytes and oximetry only. Replaced the SC and verified proper parameters. Also upgraded system software to current version, 3.1MR8. Analyzer is within acceptable operating specifications. | D0D33C249A514486 | No | After replacement of the sensor cassette (SC), pO2 values are above the acceptable QC range. Lac and glucose have also been activated. | Incorrect SC was installed on the analyzer. SC 946-008, which includes metabolites, was installed. Customer typically uses SC 946-013, with electrolytes and oximetry only. Replaced the SC and verified proper parameters. Also upgraded system software to current version, 3.1MR8. Analyzer is within acceptable operating specifications. | Assisted Replacement | SC | SC | pO2
ABL90 FLEX | Fehlersuche, LP u. Einlaßhalter gew., Initialisierung abgewartet, div. Serviceprogr. angewendet, Kal. + QC i.O. | Fehlersuche; LP u. Einlaßhalter gew.; Initialisierung abgewartet; div. Serviceprogr. angewendet; Kal. + QC i.O. | SWO-105163 | No | Troubleshooting; LP and inlet holder; initialization awaited; div. Serviceprogr. applied; Kal. + QC in the first place | Troubleshooting; LP and inlet holder; initialization awaited; div. Serviceprogr. applied; Kal. + QC in the first place | Assisted Replacement | SP | SP | pCO2
ABL90 | Test patients were QC material transferred into a syringe and due to not being whole blood and too many values out of reportable range, some parameters were not displayed due to ranges settings. Patient report file appears to be corrupted. | Walked customer through best way to test system reports and how to edit them since he had a few questions of options available. Per Jim Gibbons, reloaded report layout files for S W version 3.1 MR8. Copied setup from other 90 new analyzer & Rebooted system. | D0D32C8A9A0E6B2C | No | Test patients were QC material transferred into a syringe and due to not being whole blood and too many values out of reportable range, some parameters were not displayed due to ranges settings. Patient report file appears to be corrupted. | Walked customer through best way to test system reports and how to edit them since he had a few questions of options available. Per Jim Gibbons, reloaded report layout files for S W version 3.1 MR8. Copied setup from other 90 new analyzer & Rebooted system. | Setup | Patient report layout | Patient report layout | Patient report layout
ABL90 FLEX | Corrected tHb parameter setup, removing offset of 6 g L and resetting this to zero. | Corrected tHb parameter setup, removing offset of 6 g L and resetting this to zero. | SWO-105265 | No | Corrected tHb parameter setup, removing offset of 6 g L and resetting this to zero. | Corrected tHb parameter setup, removing offset of 6 g L and resetting this to zero. | User error | Setup | Oxi module | tHb - offset
ABL90 FLEX | Gehäuse getauscht Gerät komplettiert alle beschädigten Teile ersetzt SC und LP erneuert anschließender Testlauf | Gehäuse getauscht Gerät komplettiert alle beschädigten Teile ersetzt SC und LP erneuert anschließender Testlauf | SWO-105121 | Yes | Housing swapped device completes all damaged parts replaces SC and LP renews subsequent test run | Housing swapped device completes all damaged parts replaces SC and LP renews subsequent test run | Re classified RMED | Demo (OOB)
ABL90 | Performed Back Flush Ran diagnostics. | Replaced SP. | D0DC25E2D45BB8EE | No | Performed Back Flush Ran diagnostics. | Replaced SP. | Assisted Replacement | SP | SP | FT
ABL90 FLEX PLUS | Analysator überprüft - Erneuerung SP (auch LOT UR03) - Erneuerung Einlassdichtung - Funktionsüberprüfung ohne Beanstandung | Analysator überprüft - Erneuerung SP (auch LOT UR03) - Erneuerung Einlassdichtung - Funktionsüberprüfung ohne Beanstandung | SWO-105106 | No | Analyzer checked - Renewal SP (also LOT UR03) - Renewal inlet seal - Function check without complaint | Analyzer checked - Renewal SP (also LOT UR03) - Renewal inlet seal - Function check without complaint | Assisted Replacement | SP | SP | pCO2 | QC
AQT90 FLEX | 更换配件910-304修复仪器 | 更换配件910-304修复仪器 | SWO-105249 | No | Replacement accessories 910-304 repair instrument | Replacement accessories 910-304 repair instrument | Defect | HW | Storage Module | Barcode reader, Storage wheel
AQT90 FLEX | 更换903-329一个正常。(by andrew) | 更换903-329一个正常。(by andrew) | SWO-104940 | No | Replace the 903-329 one normal. (by andrew) | Replace the 903-329 one normal. (by andrew) | Defect | HW | Aspiration Module | Needle Wash Unit | 1258
AQT90 FLEX | 清洁Hct模块，正常，需观察 | 清洁Hct模块，正常，需观察 | SWO-104710 | No | Clean Hct module, normal, to be observed | Clean Hct module, normal, to be observed | Cleaned | HW | Aspiration Module | Needle Wash Unit | 1257
ABL825 FLEX | Kein Fehler mehr feststellbar ! In Network Connections Link Speed von Auto auf 100 M Bit gestellt - Funktionsüberprüfung etc | Kein Fehler mehr feststellbar ! In Network Connections Link Speed von Auto auf 100 M Bit gestellt - Funktionsüberprüfung etc | SWO-105192 | No | No more error detectable ! In Network Connections Link Speed from Auto set to 100 M bit - function check etc | No more error detectable ! In Network Connections Link Speed from Auto set to 100 M bit - function check etc | No fault found | SW | NetOP | NetOP
ABL825 FLEX | Erneuerung Referenz Elektrode - mehrfache Kalibrationen etc. - Funktionsüberprüfung | Erneuerung Referenz Elektrode - mehrfache Kalibrationen etc. - Funktionsüberprüfung | SWO-105047 | No | Renewal reference electrode - multiple calibrations etc. - function check | Renewal reference electrode - multiple calibrations etc. - function check | Assisted Replacement | Electrode | Electrode | Reference
ABL800 BASIC |  | B | SWO-106190 | No |  |  | Defect | HW | FlexQ Module | FlexQ Module
ABL825 FLEX | Fehlersuche, Einlasssonde u. -Dichtung ersetzt, Pumpenschlauch pH BG-Modul ersetzt, ABL aus- u. eingeschaltet, Kal i.O., QC i.O. Kunde überwacht Kal, um zu sehen, ob Fehler beseitigt ist. | B | SWO-104958 | No | Troubleshooting, inlet probe and seal replaced, pump hose pH BG module replaced, ABL off and switched on, Kal i.O., QC i.O. customer monitors Kal to see if errors are fixed. |  | Assisted Replacement | Electrode | Electrode | Reference
ABL815 FLEX | pCO2 errors on quality controls. Error code 0090(low) | Replaced membranes with different lot numbers after several replacement of previous lot failed. | D0D383B09A4B391F | No | pCO2 errors on quality controls. Error code 0090(low) | Replaced membranes with different lot numbers after several replacement of previous lot failed. | Assisted Replacement | Electrode | Membrane | pCO2 | Ver 6.15
ABL825 FLEX | Intensive leaking. Solutions Pump error 923 | Replaced Solutions pump tubing (error cleared), but massve leaking. Inspecetd all tubings, pump rotors, membrane valves, inlet probe, gasket and inlet module. Isolated leaking to inlet preheater. Ordered inlet preheater module(backordered). | D0D380B39A133917 | No | Intensive leaking. Solutions Pump error 923 | Replaced Solutions pump tubing (error cleared), but massve leaking. Inspecetd all tubings, pump rotors, membrane valves, inlet probe, gasket and inlet module. Isolated leaking to inlet preheater. Ordered inlet preheater module(backordered). | Revisit required | HW | Inlet module | Preheater
ABL815 FLEX | Unable to scan. Determined bar code scanner is defective. Ordered bar code scanner (backordered). | B | D0D3831A9A24391D | No | Unable to scan. Determined bar code scanner is defective. Ordered bar code scanner (backordered). |  | Revisit required | HW | Basic Unit | Barcode reader
ABL825 FLEX |  | B | D0D34F6D9A2C97C3 | No |  |  | Defect | HW | Inlet module | Probe bend | Needle twisted repeatedly
ABL835 FLEX | Bildschirm mehrmals kalibriert. --> keine Besserung. Bildschirm und Verbindungskabel erneuert. Bildschirm erneut kalibriert. Bildschirm lässt sich wieder bedienen. Nullpunkt Kalibrierung duchgeführt. Lactatelectrode neu membraniert. Mehrere Kalibrierungen durchgeführt. ABL verbindet sich dauernd mit Roche System und verliert die Verbindung wieder Daten werden nicht ins Labor gesendet. Problem war vor einigen Tagen schoneinmal und dann wieder weg. Um auszuschließen das der Analysator nicht sendet wurde Analysator via Crossover kabel an Rechner von Techniker angeschlossen (beide im gleichen Netz) und per Astm.net überprüft. Analysator sendet Daten einwandfrei heraus. --> Analysator i.O. Herrn Lattoch (it) anhand des Programmes gezeigt das Analysator sendet und der Fehler nicht am Analysator liegen kann. Frau Knebel aus dem Labor mitgeteilt das sie das Roche System neu starten soll. Nach Neustart wurden Daten wieder einwandfrei gesendet und Verbindung wird gehalten. QCs und Patientenberichte werden wieder ordnungsgemäß gesendet. Frau Lutteri über die gesamte Situation informiert. QCs für alle Level durchgeführt. Überprüft ob QCs ins Labor übertragen werden | Bildschirm mehrmals kalibriert. --> keine Besserung. Bildschirm und Verbindungskabel erneuert. Bildschirm erneut kalibriert. Bildschirm lässt sich wieder bedienen. Nullpunkt Kalibrierung duchgeführt. Lactatelectrode neu membraniert. Mehrere Kalibrierungen durchgeführt. ABL verbindet sich dauernd mit Roche System und verliert die Verbindung wieder Daten werden nicht ins Labor gesendet. Problem war vor einigen Tagen schoneinmal und dann wieder weg. Um auszuschließen das der Analysator nicht sendet wurde Analysator via Crossover kabel an Rechner von Techniker angeschlossen (beide im gleichen Netz) und per Astm.net überprüft. Analysator sendet Daten einwandfrei heraus. --> Analysator i.O. Herrn Lattoch (it) anhand des Programmes gezeigt das Analysator sendet und der Fehler nicht am Analysator liegen kann. Frau Knebel aus dem Labor mitgeteilt das sie das Roche System neu starten soll. Nach Neustart wurden Daten wieder einwandfrei gesendet und Verbindung wird gehalten. QCs und Patientenberichte werden wieder ordnungsgemäß gesendet. Frau Lutteri über die gesamte Situation informiert. QCs für alle Level durchgeführt. Überprüft ob QCs ins Labor übertragen werden | SWO-105204 | No | screen calibrated several times. --> no improvement. screen and connection cable. screen recalibrated. screen can be operated again. Zero-point calibration guided by. Lactate electrode. Multiple calibrations performed. ABL connects to Roche system all the time and loses the connection again data is not sent to the lab. Problem was a few days ago already once and then away again. To rule out that the analyzer does not send, the analyzer was connected via crossover cable to computer of technicians (both in the same network) and checked by Astm.net. Analyzer sends out data correctly. --> Analyzer usually Mr. Lattoch (it) shown on the program the analyzer sends and the error cannot be on the analyzer. Ms. Knebel from the lab announced that she should restart the Roche system. After rebooting, data has been sent back correctly and the connection is maintained. QCs and patient reports are sent back correctly. Mrs Lutteri is aware of the whole situation. QCs for all levels. Checks whether QCs are transferred to the laboratory | screen calibrated several times. --> no improvement. screen and connection cable. screen recalibrated. screen can be operated again. Zero-point calibration guided by. Lactate electrode. Multiple calibrations performed. ABL connects to Roche system all the time and loses the connection again data is not sent to the lab. Problem was a few days ago already once and then away again. To rule out that the analyzer does not send, the analyzer was connected via crossover cable to computer of technicians (both in the same network) and checked by Astm.net. Analyzer sends out data correctly. --> Analyzer usually Mr. Lattoch (it) shown on the program the analyzer sends and the error cannot be on the analyzer. Ms. Knebel from the lab announced that she should restart the Roche system. After rebooting, data has been sent back correctly and the connection is maintained. QCs and patient reports are sent back correctly. Mrs Lutteri is aware of the whole situation. QCs for all levels. Checks whether QCs are transferred to the laboratory | Defect | HW | DMS_PC | Touch screen
ABL835 FLEX |  | B | D0D345059A275B5E | No |  |  | Defect | HW | Inlet module | Probe cal.
ABL800 BASIC | 更换LAC电膜膜正常。 | 更换LAC电膜膜正常。 | SWO-105604 | No | Replace the LAC electric membrane is normal. | Replace the LAC electric membrane is normal. | Assisted Replacement | Electrode | Membrane | Lac | Sensitivity
ABL835 FLEX | Applikationssoftware deinstalliert, Ordner ABL gelöscht, Neuinstallation der ABL-Software, Neustart, Optionkeys installiert (BG-Key fehlt -> ist nachbestellt durch USi), Konfiguration des ABL, Firewall konfiguriert, Timeserver konfiguriert, Netzwerkanbindung, Benutzer von Radiance hochgeladen, KAL und QK ok | Applikationssoftware deinstalliert, Ordner ABL gelöscht, Neuinstallation der ABL-Software, Neustart, Optionkeys installiert (BG-Key fehlt -> ist nachbestellt durch USi), Konfiguration des ABL, Firewall konfiguriert, Timeserver konfiguriert, Netzwerkanbindung, Benutzer von Radiance hochgeladen, KAL und QK ok | SWO-105283 | Yes | Application software uninstalled, folder ABL deleted, reinstallation of the ABL software, restart, option keys installed (BG key is missing -> is reordered by USi), configuration of the ABL, firewall configured, time server configured, network connection, Users of Radiance uploaded, KAL and QK ok | Application software uninstalled, folder ABL deleted, reinstallation of the ABL software, restart, option keys installed (BG key is missing -> is reordered by USi), configuration of the ABL, firewall configured, time server configured, network connection, Users of Radiance uploaded, KAL and QK ok | Re classified RMED | SWU
AQT90 FLEX | Gerät demontiert und geprüft. Optische Einheit abgebaut, diverse Reinigungsarbeiten durchgeführt. Proceswheel gereinigt. Gerät wieder zusammengebaut. Neustart und Systemüberprüfung wiederholt, Fehler bleibt. Neue Optische Einheit bestellt. QC durchgeführt, ok. | Anwenderin startetet das Gerät neu. Fehler bleibt bestehen. | SWO-105633 | No | device is dismantled and tested. Optical unit dismantled, various cleaning work carried out. Proceswheel cleaned. reassembled. Restart and system check repeated, error remains. New optical unit ordered. QC performed, ok. | User restarts the device. Errors remain. | Revisit required | HW | Process Module | Optical Unit | 131
AQT90 FLEX |  | B | D0D388B49A341308 | Yes |  |  | Re classified RMED | SWU
AQT90 FLEX | Servicedump gezogen. Fließwegwerte in Ordnung. Safe Positions im Service angefahren - ok. Optische Einheit mehrmals gefahren - ok. Motorspindel der optischen Einheit eingefettet. 5 Testmessungen mit Pufferflüssigkeit durchgeführt - ok. Systemüberprüfung (ok) und Systemreinigung durchgeführt. LQC's durchgeführt. Eine Patientenmessung (CRP) vorgenommen - wurde ohne Abbruch gemessen (von Hr. Bernecker durchgeführt). | Servicedump gezogen. Fließwegwerte in Ordnung. Safe Positions im Service angefahren - ok. Optische Einheit mehrmals gefahren - ok. Motorspindel der optischen Einheit eingefettet. 5 Testmessungen mit Pufferflüssigkeit durchgeführt - ok. Systemüberprüfung (ok) und Systemreinigung durchgeführt. LQC's durchgeführt. Eine Patientenmessung (CRP) vorgenommen - wurde ohne Abbruch gemessen (von Hr. Bernecker durchgeführt). | SWO-105217 | No | Service dump. Flow path values in order. Safe positions in service approached - ok. Optical unit driven several times - ok. motor spindle of the optical unit. 5 test measurements carried out with buffer fluid - ok. system check (ok) and system cleaning. LQC's. A patient measurement (CRP) was carried out - was measured without discontinuation (performed by Mr. Bernecker). | Service dump. Flow path values in order. Safe positions in service approached - ok. Optical unit driven several times - ok. motor spindle of the optical unit. 5 test measurements carried out with buffer fluid - ok. system check (ok) and system cleaning. LQC's. A patient measurement (CRP) was carried out - was measured without discontinuation (performed by Mr. Bernecker). | Power off/on | HW | Process Module | Optical Unit | 541
ABL80 FLEX BASIC | Installed new solution pack and SC. Still problem is there. Need to cross check with sensor connector board and signal data board. | Installed new solution pack and SC. Still problem is there. Need to cross check with sensor connector board and signal data board. | SWO-106131 | No | Installed new solution pack and SC. Still problem is there. Need to cross check with sensor connector board and signal data board. | Installed new solution pack and SC. Still problem is there. Need to cross check with sensor connector board and signal data board.
ABL80 FLEX BASIC | Found Sensor cassette got clotted with blood, Suggested to customer replaced new Sensor cassette, waited for Consumables but it was not available with our distributor. Postponed to next week. | Found Sensor cassette got clotted with blood, Suggested to customer replaced new Sensor cassette, waited for Consumables but it was not available with our distributor. Postponed to next week. | SWO-106188 | No | Found Sensor cassette got clotted with blood, Suggested to customer replaced new Sensor cassette, waited for Consumables but it was not available with our distributor. Postponed to next week. | Found Sensor cassette got clotted with blood, Suggested to customer replaced new Sensor cassette, waited for Consumables but it was not available with our distributor. Postponed to next week.
ABL90 | No investigation needed. | Installed patch. | D0D46C54101B4413 | No | No investigation needed. | Installed patch. | Setup | His/Lis | His/lis | Rall patch
ABL90 | No investigation needed. | Installed Rals patch. | D0D46D7E10614416 | No | No investigation needed. | Installed Rals patch. | Setup | His/Lis | His/lis | Rall patch
AQT90 FLEX | 更换902-924一个正常。 by jesson | 更换902-924一个正常。 by jesson | SWO-104708 | No | Replace the 902-924 one normal. by jesson | Replace the 902-924 one normal. by jesson | Defect | HW | Analyzer Control Module | Barcode reader, display | BR2
AQT90 FLEX | 重新安装系统，使用定标卡完成定标后解决 By Jesson | 重新安装系统，使用定标卡完成定标后解决 By Jesson | SWO-105063 | No | Reinstall the system and resolve By Jesson after calibration is completed with the calibration card | Reinstall the system and resolve By Jesson after calibration is completed with the calibration card | Defect | SW | Analyzer Control Module | Re installed SW
ABL835 FLEX | Waste ellenőrzése, waste detector és jelenlét érzékelő panel csere. Kalibrálás, QC mérés | B | SWO-105173 | No | Waste control, waste detector and presence sensing panel replacement. Calibration, QC measurement |  | Defect | HW | Basic Unit | Waste detector
ABL825 FLEX |  | B | D0D4A3AF104E30FB | Yes |  |  | Re classified RMED | Half year service
ABL800 BASIC | Chamber replaced from other machine to rectify the error | Chamber replaced from other machine to rectify the error | SWO-105553 | No | Chamber replaced from other machine to rectify the error | Chamber replaced from other machine to rectify the error | Defect | HW | El/Met Module | Measuring chamber | INVESTIGATION: bend inlet probe modification because due to the inlet probe.
ABL835 FLEX | ｲﾝﾚｯﾄﾌﾟﾛｰﾌﾞの曲りによる為インレットプローブ修正。 | ｲﾝﾚｯﾄﾌﾟﾛｰﾌﾞの曲りによる為インレットプローブ修正。 | SWO-108097 | No | Inlet probe correction due to bending of inlet probe. | Inlet probe correction due to bending of inlet probe. | Defect | HW | Inlet module | Probe bend | The membrane exchange after the leak detection can not be restarted.
ABL827 FLEX | FlexQ-Modul (R0581N002) ausgebaut und (R0660N005) eingebaut, FlexQ im Service justiert, System und Funktion i.O. | FlexQ-Modul (R0581N002) ausgebaut und (R0660N005) eingebaut, FlexQ im Service justiert, System und Funktion i.O. | SWO-105319 | No | FlexQ module (R0581N002) removed and installed (R0660N005), FlexQ adjusted in service, system and function in the following | FlexQ module (R0581N002) removed and installed (R0660N005), FlexQ adjusted in service, system and function in the following | Defect | HW | FlexQ Module | Position error
ABL825 FLEX |  | B | D0DBFB0DD40D30CE | No |  |  | Assisted Replacement | Electrode | Electrode | Lac
ABL820 FLEX | Found that the users inlet probe was bent | replaced inlet probe from customer stock which corrected the issue. | D0D47AD71060DB71 | No | Found that the users inlet probe was bent | replaced inlet probe from customer stock which corrected the issue. | Defect | HW | Inlet module | Probe bend | User called stating that they are getting leak error // Found that the users inlet probe was bent // Replaced inlet probe from customer stock which corrected the issue.
ABL825 FLEX | Checked the cal verification data and, confirmed the problem. | Removed the bad calibration verification data by reloading windows and the analyzer software. | D0D4AA831001C383 | No | Checked the cal verification data and, confirmed the problem. | Removed the bad calibration verification data by reloading windows and the analyzer software. | Defect | SW | SW | Data base | 6.16
ABL825 FLEX |  | B | D0D6DCCAFC4E6B84 | No |  |  | Assisted Replacement | Electrode | Electrode | Ca
AQT90 FLEX | Gerät demontiert. Optische Einheit erneuert und kalibriert. Systemüberprüfung gestartet, ok. 2 Parameter kalibriert. QC durchgeführt. | Anwenderin startetet das Gerät neu. Fehler bleibt bestehen. | SWO-105631 | No | device dismantled. Optical unit renewed and calibrated. System check started, ok. 2 parameters calibrated. QC. | User restarts the device. Errors remain. | Defect | HW | Process Module | Optical Unit | 131
AQT90 FLEX |  | B | D0D48CA2105657AD | Yes |  |  | Re classified RMED | Planned activity | to get Service dump
AQT90 FLEX | Barcodescanner war werksseitig nicht mittig justiert. Scanner neu abgeglichen. Gerät überprüft | Barcodescanner war werksseitig nicht mittig justiert. Scanner neu abgeglichen. Gerät überprüft | SWO-105518 | No | Barcode scanner was not factory adjusted in the middle. Scanner re-matched. Device checked | Barcode scanner was not factory adjusted in the middle. Scanner re-matched. Device checked | Adjustment | HW | Storage Module | Barcode reader, Storage wheel | SYMPTOM: leak detection
ABL80 CO-OX |  | B | D0DC5807D404BCC3 | No
ABL80 FLEX CO-OX | pO2 Kanal eingestellt, SC ersetzt, Gerät kalibriert, Fehler bleibt, LP ersetzt, Gerät kalibriert, | pO2 Kanal eingestellt, SC ersetzt, Gerät kalibriert, Fehler bleibt, LP ersetzt, Gerät kalibriert, | SWO-105973 | No | pO2 channel set, SC replaced, device calibrated, error remains, LP replaced, device calibrated, | pO2 channel set, SC replaced, device calibrated, error remains, LP replaced, device calibrated,
ABL80 FLEX BASIC | 更换新的前面板解决故障 by jesson | 更换新的前面板解决故障 by jesson | SWO-105662 | No | Replace the new front panel to fix the fault by jesson | Replace the new front panel to fix the fault by jesson
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-105938 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL90 |  | B | D0D5A08D8623F43C | Yes |  |  | Re classified RMED | SWU
ABL90 FLEX |  | B | SWO-105791 | Yes |  |  | Re classified RMED | Phone call
ABL90 FLEX | Changed the inlet probe. Entered service. Ran all fluid service tests. All passed. | Changed the inlet probe. Entered service. Ran all fluid service tests. All passed. | SWO-105928 | No | Changed the inlet probe. Entered service. Ran all fluid service tests. All passed. | Changed the inlet probe. Entered service. Ran all fluid service tests. All passed. | Replaced | HW | Inlet Module | Probe bend
ABL90 FLEX | LP über Servicemenü eingesetzt. Chipdaten werden gelesen. Kontakte Solution Pack ID PCB in Ordnung.Nach Neuinitialisierung Pack bereit. Nach mehreren Kalibrationen Empfindlichkeit pCO2 bei 12%. Sensorkassette erneuert. CCF eingeschickt. | LP über Servicemenü eingesetzt. Chipdaten werden gelesen. Kontakte Solution Pack ID PCB in Ordnung.Nach Neuinitialisierung Pack bereit. Nach mehreren Kalibrationen Empfindlichkeit pCO2 bei 12%. Sensorkassette erneuert. CCF eingeschickt. | SWO-105903 | No | LP via service menu. Chip data is read. Contacts Solution Pack ID PCB fine.After reinitialization pack ready. After several calibrations sensitivity pCO2 at 12%. Sensor cartridge is renewed. CCF. | LP via service menu. Chip data is read. Contacts Solution Pack ID PCB fine.After reinitialization pack ready. After several calibrations sensitivity pCO2 at 12%. Sensor cartridge is renewed. CCF. | User error | SP Installation | SP Installation | SP Installation
ABL90 |  | B | D0D5AAE6864372FE | No |  |  | Unknown | Revisit required | Revisit required | Vacum error
ABL90 FLEX PLUS | Removed solution pack and sensor cassette for inspection. Solution pack was not activated correctly as there were still some pouches that could be heard being pierced when I pushed firmly on the top of the pack. Cleaned the inlet gasket free of a small clot on the inside of it and reseated. Ran fluid transport checks and is now passing. | Removed solution pack and sensor cassette for inspection. Solution pack was not activated correctly as there were still some pouches that could be heard being pierced when I pushed firmly on the top of the pack. Cleaned the inlet gasket free of a small clot on the inside of it and reseated. Ran fluid transport checks and is now passing. | SWO-105930 | No | Removed solution pack and sensor cassette for inspection. Solution pack was not activated correctly as there were still some pouches that could be heard being pierced when I pushed firmly on the top of the pack. Cleaned the inlet gasket free of a small clot on the inside of it and reseated. Ran fluid transport checks and is now passing. | Removed solution pack and sensor cassette for inspection. Solution pack was not activated correctly as there were still some pouches that could be heard being pierced when I pushed firmly on the top of the pack. Cleaned the inlet gasket free of a small clot on the inside of it and reseated. Ran fluid transport checks and is now passing. | User error | SP Activation | SP Activation | SP Activation
AQT90 FLEX | 更换903-364修复仪器By jesson | 更换903-364修复仪器By jesson | SWO-104926 | No | Replace the 903-364 repair instrument By Jesson | Replace the 903-364 repair instrument By Jesson | Defect | HW | Process Module | Optical Unit | 135
ABL837 FLEX | Removed Creatinine Electrodes and dummy's. Cleaned and inspected chambers. Checked solutions and found Cal 1 had not had additive added to it. Mixed in additive and repeated calibration. Some improvement has been made, but it seems all bottles had not been mixed correctly. Replaced all solutions and membranes. | Clean the chambers including the dummy Electrode spaces and replace the membranes with a different lot. Customer to call back if problem has not been resolved. 01 06 - customer has called back as the probLem has not been resolved. | SWO-105773 | No | Removed Creatinine Electrodes and dummy's. Cleaned and inspected chambers. Checked solutions and found Cal 1 had not had additive added to it. Mixed in additive and repeated calibration. Some improvement has been made, but it seems all bottles had not been mixed correctly. Replaced all solutions and membranes. | Clean the chambers including the dummy Electrode spaces and replace the membranes with a different lot. Customer to call back if problem has not been resolved. 01 06 - customer has called back as the probLem has not been resolved. | User error | Consumables | Solution bottle | Solutions bottle
ABL800 BASIC | pH Elektrode E777 Lot SK03 ausgebaut und ersetzt, ABL mehrfach Kalibriert, Funktion OK. | Kundin bittet morgen um einen Rückruf und Service. | SWO-105804 | No | pH electrode E777 Lot SK03 removed and replaced, ABL multiple calibrated, function OK. | Customer asks for a callback and service tomorrow. | Assisted Replacement | Electrode | Electrode | pH
ABL800 BASIC | 測定チャンバー(El Met側)交換 LS調整・ポンプCAL タンパク除去 CAL・QC測定データ確認 | B | SWO-108649 | No | Measurement chamber (El Met side) replacement LS adjustment pump CAL protein removal CAL/QC measurement data confirmation |  | Defect | HW | El/Met Module | Measuring chamber | Lac
ABL820 FLEX | Checked QC calibration logs for errors. QC error 374. Customer replaced inlet probe, inlet nipples and inlet gasket | Replaced inlet probe, waste tubes and lower sensor inlet tube. Run run QCs. | D0D5E726862F4B94 | No | Checked QC calibration logs for errors. QC error 374. Customer replaced inlet probe, inlet nipples and inlet gasket | Replaced inlet probe, waste tubes and lower sensor inlet tube. Run run QCs. | Defect | HW | Inlet module | Waste tube
ABL835 FLEX | Renset inletet. Forklaret kunden, at fejl 328 skyldes, at ABL'en ikke kan flytte prøvematerialet ind i ABL'en(koaguleret clot koagel). ADG SUSØ følger op. | Renset inletet. Forklaret kunden, at fejl 328 skyldes, at ABL'en ikke kan flytte prøvematerialet ind i ABL'en(koaguleret clot koagel). ADG SUSØ følger op. | SWO-105784 | No | Purified Inletet. Explained to the customer that error 328 is due to the fact that the ABL cannot move the test material into the ABL (coagulated clot Koagel). ADG SUSØ follows up. | Purified Inletet. Explained to the customer that error 328 is due to the fact that the ABL cannot move the test material into the ABL (coagulated clot Koagel). ADG SUSØ follows up. | Cleaned | HW | Fluidic pathway | Inlet module
ABL825 FLEX | Confirmed problem. | Replaced the pH electrode. | D0D5E9F986486F8D | No | Confirmed problem. | Replaced the pH electrode. | Assisted Replacement | Electrode | Electrode | pH
ABL800 BASIC | TRY TO FIX GEAR ASSEMBLY BUT IT IS STILL NOT WORKING NEED TO CHANGE PRINTER HEAD | TRY TO FIX GEAR ASSEMBLY BUT IT IS STILL NOT WORKING NEED TO CHANGE PRINTER HEAD | SWO-105985 | No | TRY TO FIX GEAR ASSEMBLY BUT IT IS STILL NOT WORKING NEED TO CHANGE PRINTER HEAD | TRY TO FIX GEAR ASSEMBLY BUT IT IS STILL NOT WORKING NEED TO CHANGE PRINTER HEAD | Defect | HW | Printer module | Printer Head
ABL825 FLEX | Leihmodul ausgebaut und Kundenoximodul wieder eingesetzt, justiert und kalibriert, Kal und QC | Leihmodul ausgebaut und Kundenoximodul wieder eingesetzt, justiert und kalibriert, Kal und QC | SWO-106057 | No | Rental module removed and customer oximodule re-inserted, adjusted and calibrated, Kal and QC | Rental module removed and customer oximodule re-inserted, adjusted and calibrated, Kal and QC | Revisit required | HW | Oxi module | Oxi module | Customer module reinstalled
ABL815 FLEX | Kunde Wartet auf Lac. Membran Lieferung, am 25.05. bestellt. Lac. Membranen mit gebracht, Reinigungsverfahren für Lac. Membranen durchgeführt und Anwender eingewiesen. FlexQ Modul demontiert und gereinigt. | Kunde Wartet auf Lac. Membran Lieferung, am 25.05. bestellt. Lac. Membranen mit gebracht, Reinigungsverfahren für Lac. Membranen durchgeführt und Anwender eingewiesen. FlexQ Modul demontiert und gereinigt. | SWO-106365 | No | Customer Waiting for Lac. Membrane delivery, ordered on 25.05. Lac. membranes, cleaning process for Lac. membranes and instructed users. FlexQ module dismantled and cleaned. | Customer Waiting for Lac. Membrane delivery, ordered on 25.05. Lac. membranes, cleaning process for Lac. membranes and instructed users. FlexQ module dismantled and cleaned. | Assisted Replacement | Electrode | Membrane | Lac
ABL825 FLEX |  | B | SWO-105689 | No |  |  | Defect | HW | Oxi module | Oxi module | 251
ABL815 FLEX | PH-Elektrode gereinigt, Messkammersitz der PH und Ref.-Elektrode gereinigt, Kal, weiterhin Empfindlichkeit zu niedrig, Wechsel der PH-Elektrode (Garantieelektrode UP01), Kal, QC | PH-Elektrode gereinigt, Messkammersitz der PH und Ref.-Elektrode gereinigt, Kal, weiterhin Empfindlichkeit zu niedrig, Wechsel der PH-Elektrode (Garantieelektrode UP01), Kal, QC | SWO-105935 | No | PH electrode cleaned, measuring chamber seat of PH and ref electrode cleaned, Kal, still sensitivity too low, change of PH electrode (guarantee electrode UP01), Kal, QC | PH electrode cleaned, measuring chamber seat of PH and ref electrode cleaned, Kal, still sensitivity too low, change of PH electrode (guarantee electrode UP01), Kal, QC | Assisted Replacement | Electrode | Electrode | pH
AQT90 FLEX | Gerät geöffnet und reichlich Staub vom Scannergehäuse entfernt. Aktive Kalibrationen ausgelesen. Uhrzeit über eine Stunde zurück eingestellt. Im Servicemenü und während Neustart die Scanfunktion getestet, OK | Gerät geöffnet und reichlich Staub vom Scannergehäuse entfernt. Aktive Kalibrationen ausgelesen. Uhrzeit über eine Stunde zurück eingestellt. Im Servicemenü und während Neustart die Scanfunktion getestet, OK | SWO-106150 | No | device and removes plenty of dust from the scanner housing. Active calibrations. time over one hour. Tested in the service menu and during restart the scan function, OK | device and removes plenty of dust from the scanner housing. Active calibrations. time over one hour. Tested in the service menu and during restart the scan function, OK | Cleaned | HW | Storage Module | Barcode reader, Storage wheel
AQT90 FLEX |  | B | D0D5D9F2860E7239 | No |  |  | Defect | HW | Analyzer Control Module | CF Card | 1999
AQT90 FLEX |  | B | D0D5FCAD8602966E | Yes |  |  | Re classified RMED | Service Dump
AQT90 FLEX | Dismantled the machine Inlet door panel and reassembled and still it remains the Same. Need to replace the front panel completely. | B | SWO-105946 | No | Dismantled the machine Inlet door panel and reassembled and still it remains the Same. Need to replace the front panel completely. |  | Revisit required | HW | Cover and Chassis | Inlet Door
AQT90 FLEX | done the system clean process and noted the counter details. | done the system clean process and noted the counter details. | SWO-106196 | No | done the system clean process and noted the counter details. | done the system clean process and noted the counter details. | User error | SW | Cleaning Solution | Cleaning Solution
ABL80 FLEX BASIC | 更换测试卡正常。 | 更换测试卡正常。 | SWO-106026 | No | The replacement test card is OK. | The replacement test card is OK.
ABL80 FLEX BASIC | Cleaned Sc and kept for re hydration. Then reinstalled Sc but still problem is there. Need to replace SC. | Cleaned Sc and kept for re hydration. Then reinstalled Sc but still problem is there. Need to replace SC. | SWO-106100 | No | Cleaned Sc and kept for re hydration. Then reinstalled Sc but still problem is there. Need to replace SC. | Cleaned Sc and kept for re hydration. Then reinstalled Sc but still problem is there. Need to replace SC.
ABL80 FLEX BASIC | 機器引き上げ | B | SWO-105947 | No | Equipment pulling up
ABL80 CO-OX | Sometimes fails all or some QC OXI parameters and are marked I A . System cycle repeated and all green. Then get some or all OXI external QC parameters red and then I A. Found short, large clear tubing behind front panel bent in two. one end of this goes to black port on manifold. | Replaced part from PM kit. Same issue after System Cycle. Ran a tHb Calibration and System Cycle. Same issue. Replaced Hemolyzer and align. Same issue. Order parts and will return Monday. | D0D6E901FC0778C0 | No | Sometimes fails all or some QC OXI parameters and are marked I A . System cycle repeated and all green. Then get some or all OXI external QC parameters red and then I A. Found short, large clear tubing behind front panel bent in two. one end of this goes to black port on manifold. | Replaced part from PM kit. Same issue after System Cycle. Ran a tHb Calibration and System Cycle. Same issue. Replaced Hemolyzer and align. Same issue. Order parts and will return Monday.
ABL90 |  | B | D0D717A8FC571D7E | Yes |  |  | Re classified RMED | Installation
ABL90 |  | B | D0D6E1A6FC026E6E | No |  |  | Adjustment | HW | Oxi module | tHb Cal
ABL90 FLEX PLUS | Fehlersuche, Datensicherung, Gerät aus, neue CF-Card eingesetzt, Gerät ein, kompl. SW-Installation durchgef., Konfig. neu geladen, Versuch Netzwerkverbindung herzustellen gescheitert, div. Telefonate mit IT Vivantes u. RMed = IP-Adresse 2x vergeben -> Klärung am Montag | Fehlersuche; Datensicherung; Gerät aus; neue CF-Card eingesetzt; Gerät ein; kompl. SW-Installation durchgef.; Konfig. neu geladen; Versuch Netzwerkverbindung herzustellen gescheitert; div. Telefonate mit IT Vivantes u. RMed = IP-Adresse 2x vergeben -> Klärung am Montag | SWO-106042 | No | Troubleshooting; data backup; device off; new CF-Card; device; Complete. SW installation; Config. reloaded; Attempt to establish network connection failed; div. Phone calls with IT Vivantes and RMed = IP address 2x assigned -> Clarification on Monday | Troubleshooting; data backup; device off; new CF-Card; device; Complete. SW installation; Config. reloaded; Attempt to establish network connection failed; div. Phone calls with IT Vivantes and RMed = IP address 2x assigned -> Clarification on Monday | Defect | HW | CF Card | CF Card
ABL90 FLEX PLUS | correctly in serted solution pack but started failing p02 on Qc's and had to replace solution pack | correctly in serted solution pack but started failing p02 on Qc's and had to replace solution pack | SWO-106551 | No | correctly in serted solution pack but started failing p02 on Qc's and had to replace solution pack | correctly in serted solution pack but started failing p02 on Qc's and had to replace solution pack | User error | SP Activation | SP Activation | SP Activation
ABL80 FLEX BASIC | Try to install with new SP and by reset but same problem. Replaced SC- new SC given. Per made for failed SC. | Try to install with new SP and by reset but same problem. Replaced SC- new SC given. Per made for failed SC. | SWO-106214 | No | Try to install with new SP and by reset but same problem. Replaced SC- new SC given. Per made for failed SC. | Try to install with new SP and by reset but same problem. Replaced SC- new SC given. Per made for failed SC.
ABL800 BASIC | clean the gasket. run protein removal.clean the fluidic path. leakage in gasket causing sample error.machine under observation. | clean the gasket. run protein removal.clean the fluidic path. leakage in gasket causing sample error.machine under observation. | SWO-106179 | No | clean the gasket. run protein removal.clean the fluidic path. leakage in gasket causing sample error.machine under observation. | clean the gasket. run protein removal.clean the fluidic path. leakage in gasket causing sample error.machine under observation. | Cleaned | Fluidic pathway | Inlet module | Inlet module
ABL835 FLEX | Comprobar electrodo de PCO2 , activar el electrodo, calibrar y probar con controles. Actualizar el software a la versión 6.16. ES. | Activar el electrodo de PCO2. | D0D6D163FC5868CC | No | Check PCO2 electrode, activate the electrode, calibrate and test with controls. Update the software to version 6.16. Is. | Activate the PCO2 electrode. | Setup | SW | Installation | 6.16
ABL827 FLEX | pCO2 Elektrode erneuert. Funktionskontrolle iO. | pCO2 Elektrode erneuert. Funktionskontrolle iO. | SWO-106044 | No | pCO2 electrode. Functional control iO. | pCO2 electrode. Functional control iO. | Assisted Replacement | Electrode | Electrode | pCO2 | 6.16?
ABL835 FLEX | confirmed problem. | Removed the measuring modules and the fluidic block and cleaned and reseated the valves. | D0D73A1AFC370B6F | No | confirmed problem. | Removed the measuring modules and the fluidic block and cleaned and reseated the valves. | Cleaned | Fluidic pathway | El/Met Module | El/Met Module
ABL825 FLEX | El Met und PH BG- Pumpenschläuche ausgetauscht. Pumpenkal und LS Justage wiederholt. | El Met und PH BG- Pumpenschläuche ausgetauscht. Pumpenkal und LS Justage wiederholt. | SWO-106072 | No | El Met and PH BG pump hoses. pump scale and LS adjustment. | El Met and PH BG pump hoses. pump scale and LS adjustment. | Defect | HW | pH/BG Module | Pump tubes
ABL825 FLEX | Fuite liquide sur le module d'introduction e.o Remplacé le kit de maintenance semi-annuel e.o Calibration e.o contrôle de qualité e.o | Fuite liquide sur le module d'introduction e.o Remplacé le kit de maintenance semi-annuel e.o Calibration e.o contrôle de qualité e.o | SWO-106049 | No | Liquid leak on the introductory module e.o Replaced the semi-annual maintenance kit e.o Calibration e.o quality control e.o | Liquid leak on the introductory module e.o Replaced the semi-annual maintenance kit e.o Calibration e.o quality control e.o | Defect | HW | Inlet module | Inlet module | Leakage
ABL80 FLEX BASIC | 更换测试卡 | 更换测试卡 | SWO-106647 | No | Replace the test card | Replace the test card
ABL80 FLEX BASIC | 观察定标过程，发现定标时，测试卡内没有气泡，分别更换测试卡和试剂包，故障没有解决。检查机器，发现902-272上有腐蚀痕迹，拆下902-272检查发现背面还残留有液体。更换902-272后故障没有解决，更换903-334后解决故障 by jesson | 观察定标过程，发现定标时，测试卡内没有气泡，分别更换测试卡和试剂包，故障没有解决。检查机器，发现902-272上有腐蚀痕迹，拆下902-272检查发现背面还残留有液体。更换902-272后故障没有解决，更换903-334后解决故障 by jesson | SWO-106200 | No | Observe the calibration process, found that the calibration time, the test card is not bubbles, respectively, replace the test card and reagent package, the fault is not resolved. Check the machine and find traces of corrosion on the 902-272, and remove the 902-272 to check that there is still liquid on the back. Replacement 902-272 after the fault is not resolved, replace 903-334 after the fault by jesson | Observe the calibration process, found that the calibration time, the test card is not bubbles, respectively, replace the test card and reagent package, the fault is not resolved. Check the machine and find traces of corrosion on the 902-272, and remove the 902-272 to check that there is still liquid on the back. Replacement 902-272 after the fault is not resolved, replace 903-334 after the fault by jesson
AQT90 FLEX | 拆 库存机更换903-186，910-339，636-666各一个正常。By jesson | 拆 库存机更换903-186，910-339，636-666各一个正常。By jesson | SWO-106007 | No | Dismantling the stock machine replacement 903-186, 910-339, 636-666 each normal. By Jesson | Dismantling the stock machine replacement 903-186, 910-339, 636-666 each normal. By Jesson | Defect | HW | Aspiration Module | Barcode reader, Inlet
ABL800 BASIC | Clean all chambers,change Lac membrane and done 2-point calibration and resolve the problem | Clean all chambers,change Lac membrane and done 2-point calibration and resolve the problem | SWO-106250 | No | Clean all chambers,change Lac membrane and done 2-point calibration and resolve the problem | Clean all chambers,change Lac membrane and done 2-point calibration and resolve the problem | Assisted Replacement | Electrode | Membrane | Lac | Sensitivity
AQT90 FLEX | Remove out the stacked TnT Cartridge from storage wheel. Clean the Storage wheel and optical sensor. Connect all the module again and start the machine . | Remove out the stacked TnT Cartridge from storage wheel. Clean the Storage wheel and optical sensor. Connect all the module again and start the machine . | SWO-109056 | No | Remove out the stacked TnT Cartridge from storage wheel. Clean the Storage wheel and optical sensor. Connect all the module again and start the machine . | Remove out the stacked TnT Cartridge from storage wheel. Clean the Storage wheel and optical sensor. Connect all the module again and start the machine . | Defect | Consumables | Test Kit | Cartridge | 5689 (Broken or Label)
ABL80 FLEX BASIC | Replaced new Sc and Sp. Unit is working fine. | Replaced new Sc and Sp. Unit is working fine. | SWO-108592 | No | Replaced new Sc and Sp. Unit is working fine. | Replaced new Sc and Sp. Unit is working fine.
ABL80 FLEX BASIC | Introduce the PICO 70 Sampler to reduce the humanistic error during pricking of arterial blood from neonates. Customer was agree and asking for quotation for procure and further financially negotiation. | After checking all technical point, Ran the RL as a sample and compare the target values. which are as below- Target Analyzed Value of RL Na+ 131 137 K+ 5 5.2 Both values within 1'SD. which indicate the most accurate reporting. Convinced and aware the customer about the Radiometer QC for more authenticate clarity. Customer Ran the sample for his satisfaction again. Now issue has been resolved. Note-It is advised to procure QC for compare of quality of results related issue. | SWO-109063 | No | Introduce the PICO 70 Sampler to reduce the humanistic error during pricking of arterial blood from neonates. Customer was agree and asking for quotation for procure and further financially negotiation. | After checking all technical point, Ran the RL as a sample and compare the target values. which are as below- Target Analyzed Value of RL Na+ 131 137 K+ 5 5.2 Both values within 1'SD. which indicate the most accurate reporting. Convinced and aware the customer about the Radiometer QC for more authenticate clarity. Customer Ran the sample for his satisfaction again. Now issue has been resolved. Note-It is advised to procure QC for compare of quality of results related issue.
ABL80 FLEX BASIC | After checking all technical point, Ran the RL as a sample and compare the target values. which are as below- Target Analyzed Value of RL Na+ 131 137 K+ 5 5.2 Both values within 1'SD. which indicate the most accurate reporting. Convinced and aware the customer about the Radiometer QC for more authenticate clarity. Customer Ran the sample for his satisfaction again. Now issue has been resolved. Note-It is advised to procure QC for compare of quality of results related issue. | After checking all technical point, Ran the RL as a sample and compare the target values. which are as below- Target Analyzed Value of RL Na+ 131 137 K+ 5 5.2 Both values within 1'SD. which indicate the most accurate reporting. Convinced and aware the customer about the Radiometer QC for more authenticate clarity. Customer Ran the sample for his satisfaction again. Now issue has been resolved. Note-It is advised to procure QC for compare of quality of results related issue. | SWO-109057 | No | After checking all technical point, Ran the RL as a sample and compare the target values. which are as below- Target Analyzed Value of RL Na+ 131 137 K+ 5 5.2 Both values within 1'SD. which indicate the most accurate reporting. Convinced and aware the customer about the Radiometer QC for more authenticate clarity. Customer Ran the sample for his satisfaction again. Now issue has been resolved. Note-It is advised to procure QC for compare of quality of results related issue. | After checking all technical point, Ran the RL as a sample and compare the target values. which are as below- Target Analyzed Value of RL Na+ 131 137 K+ 5 5.2 Both values within 1'SD. which indicate the most accurate reporting. Convinced and aware the customer about the Radiometer QC for more authenticate clarity. Customer Ran the sample for his satisfaction again. Now issue has been resolved. Note-It is advised to procure QC for compare of quality of results related issue.
ABL80 FLEX BASIC |  | B | SWO-106329 | No
ABL90 FLEX PLUS | Analyzer uitgezet en manueel de deurtjes gesloten. Na herstarten moeten de reagenspack en de sensorcassette opnieuw geïnstalleerd worden maar dat verloopt nu normaal. De deurtjes gaan keurig open. Na opstarten (55 minuten) is de gevoeligheid van de glu en lac te veel te laag, iets meer dan 40pA mmol. Nieuwe sensor geïnstalleerd (R0304). Per rapport van defecte (R0295) gemaakt. | Analyzer uitgezet en manueel de deurtjes gesloten. Na herstarten moeten de reagenspack en de sensorcassette opnieuw geïnstalleerd worden maar dat verloopt nu normaal. De deurtjes gaan keurig open. Na opstarten (55 minuten) is de gevoeligheid van de glu en lac te veel te laag, iets meer dan 40pA mmol. Nieuwe sensor geïnstalleerd (R0304). Per rapport van defecte (R0295) gemaakt. | SWO-106608 | No | Analyzer is turned off and the doors are closed manually. After restarting, the reagent pack and the sensor cassette must be reinstalled, but that is now normal. The doors are nicely open. After booting (55 minutes) The sensitivity of the Glu and Lac is too low, just over 40pA mmol. New sensor installed (R0304). Per Report of Defective (R0295) made. | Analyzer is turned off and the doors are closed manually. After restarting, the reagent pack and the sensor cassette must be reinstalled, but that is now normal. The doors are nicely open. After booting (55 minutes) The sensitivity of the Glu and Lac is too low, just over 40pA mmol. New sensor installed (R0304). Per Report of Defective (R0295) made. | Adjustment | SC Door | SC Door | Off/On
ABL90 FLEX PLUS | Shutdown analyses and manual (torx- screwdriver) closed the doors. After restart analyzer ask to replace reagentpack and sensorcassette. This can be done in normal way. Bay normal open and close. After startup the sensitivity glucose and lactate is to low, so to instal a new sensor cassette. | Shutdown analyses and manual (torx- screwdriver) closed the doors. After restart analyzer ask to replace reagentpack and sensorcassette. This can be done in normal way. Bay normal open and close. After startup the sensitivity glucose and lactate is to low, so to instal a new sensor cassette. | SWO-108778 | Yes | Shutdown analyses and manual (torx- screwdriver) closed the doors. After restart analyzer ask to replace reagentpack and sensorcassette. This can be done in normal way. Bay normal open and close. After startup the sensitivity glucose and lactate is to low, so to instal a new sensor cassette. | Shutdown analyses and manual (torx- screwdriver) closed the doors. After restart analyzer ask to replace reagentpack and sensorcassette. This can be done in normal way. Bay normal open and close. After startup the sensitivity glucose and lactate is to low, so to instal a new sensor cassette. | Re classified RMED | Double entry |  |  | LEH
ABL90 FLEX PLUS | SP und SC erneuert | SP und SC erneuert | SWO-106542 | No | SP and SC renewed | SP and SC renewed | Assisted Replacement | SC | SC | Wetability
ABL90 FLEX | 更换测试卡，试剂包正常。 | 更换测试卡，试剂包正常。 | SWO-107693 | No | Replace the test card and the reagent package is OK. | Replace the test card and the reagent package is OK.
ABL90 FLEX | 更换测试卡正常。 | 更换测试卡正常。 | SWO-106330 | No | The replacement test card is OK. | The replacement test card is OK. | Assisted Replacement | SC | SC | Reference | 1361
ABL825 FLEX | Examined waste tubing area. Found defective waste tube. | Replaced waste tubes, with customer stock. | D0DA9AE75E4F503F | No | Examined waste tubing area. Found defective waste tube. | Replaced waste tubes, with customer stock. | Defect | HW | Basic Unit | Waste tube
ABL835 FLEX | could not find any reason for it not passing. | Checked and cleaned all the wells that the electrodes sit in. preformed pump cal and ls adjust on analyzer. | D0DBF522D411267E | No | could not find any reason for it not passing. | Checked and cleaned all the wells that the electrodes sit in. preformed pump cal and ls adjust on analyzer. | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
ABL825 FLEX | I inspected tray for anything rubbing it and did not find anything. | Performed a completed sampler tray calibration. | D0DAA52A5E3CD931 | No | I inspected tray for anything rubbing it and did not find anything. | Performed a completed sampler tray calibration. | Adjustment | HW | FlexQ Module | FlexQ Module
ABL825 FLEX | Defekte neue Einlassdichtung und Einlassnadel ersetzt, div. Dichtungen an Elektroden erneuert. Gerät mehrfach kalibriert und mit QCs überprüft, OK. | Defekte neue Einlassdichtung und Einlassnadel ersetzt, div. Dichtungen an Elektroden erneuert. Gerät mehrfach kalibriert und mit QCs überprüft, OK. | SWO-106667 | No | Defective new inlet seal and inlet needle replaced, various seals on electrodes renewed. Device calibrated multiple times and checked with QCs, OK. | Defective new inlet seal and inlet needle replaced, various seals on electrodes renewed. Device calibrated multiple times and checked with QCs, OK. | Defect | HW | Probe | Probe | Clot?
ABL825 FLEX | AC's Lot 7745 611 abgezogen, Level2 Lot 649 installiert, Funktionskontrolle ok | AC's Lot 7745 611 abgezogen, Level2 Lot 649 installiert, Funktionskontrolle ok | SWO-106582 | No | AC's Lot 7745 611 subtracted, Level2 Lot 649 installed, function control ok | AC's Lot 7745 611 subtracted, Level2 Lot 649 installed, function control ok | Assisted Replacement | QC | Level 2 | Level 2 | Strange
ABL800 BASIC | 更换新Lac电极膜后解决 | 更换新Lac电极膜后解决 | SWO-106314 | No | Solution after replacing the new Lac electrode film | Solution after replacing the new Lac electrode film | Assisted Replacement | Electrode | Membrane | Lac | Sensitivity
ABL815 FLEX | Erneuerung Referenz Elektrode - Funktionsüberprüfung etc. | Erneuerung Referenz Elektrode - Funktionsüberprüfung etc. | SWO-106612 | No | Renewal Reference Electrode - Function Verification etc. | Renewal Reference Electrode - Function Verification etc. | Assisted Replacement | Electrode | Electrode | Reference
ABL827 FLEX |  | B | D0E5F1A3840542EB | Yes |  |  | Re classified RMED | Installation
AQT90 FLEX | replaced the new optical unit its failed the system check optical intensity failure we need to replace new optical unit again | B | SWO-107581 | No | replaced the new optical unit its failed the system check optical intensity failure we need to replace new optical unit again |  | Defect | HW | Process Module | Optical Unit | 124
AQT90 FLEX | Performed TnI calibration with a fresh calibration cartridge, now TnI Test cartridge is being detected in storage wheel | Performed TnI calibration with a fresh calibration cartridge, now TnI Test cartridge is being detected in storage wheel | SWO-108330 | No | Performed TnI calibration with a fresh calibration cartridge, now TnI Test cartridge is being detected in storage wheel | Performed TnI calibration with a fresh calibration cartridge, now TnI Test cartridge is being detected in storage wheel | User error | Consumables | Test Kit | Cartridge | Installation
ABL80 FLEX BASIC | 更换测试卡修复仪器 | 更换测试卡修复仪器 | SWO-107415 | No | Replace the test card repair instrument | Replace the test card repair instrument
ABL80 FLEX BASIC | カセット交換 | B | SWO-112182 | No | Cassette replacement
ABL90 FLEX | Lösungspack korrekt installiert. Gerät durchgespült, Kal und QC durchgeführt. Gerät OK | Lösungspack korrekt installiert. Gerät durchgespült, Kal und QC durchgeführt. Gerät OK | SWO-107727 | No | solution pack installed correctly. device, Kal and QC performed. Device OK | solution pack installed correctly. device, Kal and QC performed. Device OK | User error | SP Activation | SP Activation | SP Activation
ABL90 FLEX | Fehlersuche, div. Test`s im Service durchgef., Einlaßhalter gew., SC gew., LP neu initialisiert -> weiter Fehler, alte SC eingesetzt, Initialisierung abgewartet, weiter Cl-Fehler, Abbruch d. Arbeiten, um Material zu organisieren, Cl abgeschaltet, Kunde informiert, Kal. + QC ohne Cl i.O. | Fehlersuche; div. Test`s im Service durchgef.; Einlaßhalter gew.; SC gew.; LP neu initialisiert -> weiter Fehler; alte SC eingesetzt; Initialisierung abgewartet; weiter Cl-Fehler; Abbruch d. Arbeiten, um Material zu organisieren; Cl abgeschaltet; Kunde informiert; Kal. + QC ohne Cl i.O. | SWO-107434 | No | Troubleshooting; div. Test it in the service. Inlet holder. SC. LP reinitialized -> further error; old SC used; initialization awaited; further Cl error; Demolition of the works to organize material; Cl switched off; customer informed; Kal. + QC without Cl i.O. | Troubleshooting; div. Test it in the service. Inlet holder. SC. LP reinitialized -> further error; old SC used; initialization awaited; further Cl error; Demolition of the works to organize material; Cl switched off; customer informed; Kal. + QC without Cl i.O. | Unknown | Revisit required | - | CL
ABL90 |  | B | D0DD54E34A3575E7 | No |  |  | Defect | HW | Sensor Bay | Sensor Bay
ABL90 |  | B | D0DC21DDD45C532F | Yes |  |  | Re classified RMED
ABL90 |  | B | D0DC139FD46332E4 | Yes |  |  | Re classified RMED | RE installation
TCM CombiM |  | B | SWO-107413 | No |  |  | Defect | HW | Power Supply | Power Supply
AQT90 FLEX |  | B | SWO-107395 | No |  |  | Defect | HW | Process Module | Optical Unit | 131
ABL800 BASIC | reinstalled the windows xp and application software. now instrument working fine. | reinstalled the windows xp and application software. now instrument working fine. | SWO-108708 | No | reinstalled the windows xp and application software. now instrument working fine. | reinstalled the windows xp and application software. now instrument working fine. | Defect | SW | Roma.ini | Roma.ini
ABL835 FLEX | CPUユニット交換 CAL・QC測定データ確認 | B | SWO-108797 | No | CPU unit replacement CAL/QC measurement data confirmation |  | Defect | HW | DMS_PC | CPU
ABL800 BASIC | Has to change the printer head. Part Number: 910-318 printer head 636-480 : Cable for Flap detector PCB | B | SWO-108287 | No | Has to change the printer head. Part Number: 910-318 printer head 636-480 : Cable for Flap detector PCB |  | Defect | HW | Printer module | Printer Head
ABL825 FLEX | Herr Jakob hat mir mitgeteilt das er schon mehrmals in den letzten Wochen membraniert hat. Erst alles wieder i.O. Sei und nach kurzer Zeit die pCO2 Empfindlichkeit in der Kalibrierung wieder absinkt. pCO2 Elektrode getauscht und neu membraniert. Mehrere Kalibrierungen durchgeführt. QCs für alle Level durchgeführt. ABL wieder Messbereit. Mehrere Patientenmessungen durchführen lassen. | Herr Jakob hat mir mitgeteilt das er schon mehrmals in den letzten Wochen membraniert hat. Erst alles wieder i.O. Sei und nach kurzer Zeit die pCO2 Empfindlichkeit in der Kalibrierung wieder absinkt. pCO2 Elektrode getauscht und neu membraniert. Mehrere Kalibrierungen durchgeführt. QCs für alle Level durchgeführt. ABL wieder Messbereit. Mehrere Patientenmessungen durchführen lassen. | SWO-107652 | No | Mr Jakob has told me that he has been membraneing several times in recent weeks. First everything is back in the o.o. and after a short time the pCO2 sensitivity in the calibration drops again. pCO2 electrode is exchanged and re-membraneed. Multiple calibrations performed. QCs for all levels. ABL ready to measure again. Have multiple patient measurements carried out. | Mr Jakob has told me that he has been membraneing several times in recent weeks. First everything is back in the o.o. and after a short time the pCO2 sensitivity in the calibration drops again. pCO2 electrode is exchanged and re-membraneed. Multiple calibrations performed. QCs for all levels. ABL ready to measure again. Have multiple patient measurements carried out. | Assisted Replacement | Electrode | Electrode | pCO2
ABL825 FLEX | Glucose had no sensitivity reading which could indicate the electrode could have been damaged. Upon closer look of the calibration result message, the Glucose parameter was not installed and the Option Key has expired. New option key will have to be created. | Installed new demo option key for Glucose. Verified instrument operation. | D0DBF6E2D45531EF | No | Glucose had no sensitivity reading which could indicate the electrode could have been damaged. Upon closer look of the calibration result message, the Glucose parameter was not installed and the Option Key has expired. New option key will have to be created. | Installed new demo option key for Glucose. Verified instrument operation. | Setup | SW | Demo Key Exp | Demo Key Exp
ABL825 FLEX | It was stuck at Radiometer boot up screen. | Reloaded Radiometer software. Reloaded setup. | D0DBFBF3D432CF0E | No | It was stuck at Radiometer boot up screen. | Reloaded Radiometer software. Reloaded setup. | Defect | SW | Reload SW | Reload SW
ABL820 FLEX | Defective electrode ceramic jacket fell off | Replaced Electrode. also checked windows setting to verify all sound settings are turned up to 100% but no help on Analtzer sound. Not chiming. will follow up | D0DC6D8DD4516EA1 | No | Defective electrode ceramic jacket fell off | Replaced Electrode. also checked windows setting to verify all sound settings are turned up to 100% but no help on Analtzer sound. Not chiming. will follow up | Assisted Replacement | Electrode | Electrode | pCO2
ABL810 FLEX | Bedienung mit angeschlossener Tastatur normal - Touchpanel scheint defekt zu sein - muss erst bestellt werden, externe Maus angeschlossen um Gerät bedienbar zu machen | B | SWO-108357 | No | Operation with connected keyboard normal - touch panel seems to be defective - must first be ordered, external mouse connected to make device operable |  | Revisit required | HW | DMS_PC | Touch screen
ABL800 BASIC | replaced the new gas2 and performed the gas check now machine working fine | replaced the new gas2 and performed the gas check now machine working fine | SWO-107603 | No | replaced the new gas2 and performed the gas check now machine working fine | replaced the new gas2 and performed the gas check now machine working fine | Assisted Replacement | Gas 2 | Gas 2 | Gas 2
ABL80 FLEX BASIC | 更换新测试卡修复仪器 | 更换新测试卡修复仪器 | SWO-107740 | No | Replace the new test card repair instrument | Replace the new test card repair instrument
ABL80 FLEX BASIC | 更换902-975一个正常。(by zhenjie) | 更换902-975一个正常。(by zhenjie) | SWO-106482 | No | Replace the 902-975 one normal. (by zhenjie) | Replace the 902-975 one normal. (by zhenjie)
ABL80 FLEX BASIC | 更换新测试卡修复仪器 | 更换新测试卡修复仪器 | SWO-107763 | No | Replace the new test card repair instrument | Replace the new test card repair instrument
ABL90 |  | B | D0DD56D24A62B3ED | No |  |  | Defect | HW | Inlet Module | Inlet Module
ABL90 |  | B | D0DEF25DC0100EF9 | Yes |  |  | Re classified RMED | TU
ABL90 FLEX | Parameter wieder aktiviert, 3x Kal., alles i.O. | Parameter wieder aktiviert; 3x Kal.; alles i.O. | SWO-107908 | No | Parameter reactivated; 3x Kal.; everything in the | Parameter reactivated; 3x Kal.; everything in the | No fault found | SC | CL | CL
ABL90 FLEX PLUS |  | B | D0DD05BF4A30AFD8 | Yes |  |  | Re classified RMED | PER
AQT90 FLEX | 更换903-364修复仪器(by zhenjie) | 更换903-364修复仪器(by zhenjie) | SWO-107459 | No | Replace the 903-364 repair instrument (by zhenjie) | Replace the 903-364 repair instrument (by zhenjie) | Defect | HW | Process Module | Optical Unit | 129
ABL825 FLEX |  | B | D0DD4D824A56F599 | No |  |  | Unknown | HW | pH/BG Module | pO2 | Strange
ABL800 BASIC | Cleaned auto check scanner. Rescanned QC material. | Cleaned auto check scanner. Rescanned QC material. | SWO-108882 | No | Cleaned auto check scanner. Rescanned QC material. | Cleaned auto check scanner. Rescanned QC material. | Cleaned | HW | AutoCheck Module | Barcode reader
ABL835 FLEX |  | B | D0DD331D4A0351D1 | No |  |  | Defect | HW | Inlet module | Probe Holder
ABL817 FLEX | Ausgewerteten Daten ergaben auch im tHb und K eine Veränderung innerhalb QC Grenzen. Isolation im Service überprüft, OK. E1001 membraniert. QC OK. | Fehler tritt sporadisch auf, Modul ausgebaut Ventile gereinigt, QC Daten zur weiteren Auswertung mitgenommen. | SWO-107933 | No | Evaluated data also showed a change within QC limits in tHb and K. Isolation in service checked, OK. E1001 membrane. QC OK. | Error occurs sporadically, module removed valves cleaned, QC data taken for further evaluation. | Assisted Replacement | Electrode | Membrane | Reference
ABL817 FLEX | Fehler tritt sporadisch auf, Modul ausgebaut Ventile gereinigt, QC Daten zur weiteren Auswertung mitgenommen. | Fehler tritt sporadisch auf, Modul ausgebaut Ventile gereinigt, QC Daten zur weiteren Auswertung mitgenommen. | SWO-107934 | No | Error occurs sporadically, module removed valves cleaned, QC data taken for further evaluation. | Error occurs sporadically, module removed valves cleaned, QC data taken for further evaluation. | Cleaned | Fluidic pathway | Fluidic pathway | Na+
ABL825 FLEX | Checked lamp voltage. | Adjusted lamp voltage to specifications. | D0DD56584A4CD229 | No | Checked lamp voltage. | Adjusted lamp voltage to specifications. | Adjustment | HW | Oxi module | Lamp unit
ABL815 FLEX | pCO2 low on level 2 QC | Cleaned electrode and chamber. | D0DD933F4A374720 | No | pCO2 low on level 2 QC | Cleaned electrode and chamber. | Cleaned | Fluidic pathway | Fluidic pathway | pCO2
ABL800 BASIC | Performed complete oximetry test, where neon intensity and neon number was found out of range. So done the tHb calibration after which kept the machine OFF for 30 minutes. After restarting the analyzer and successful calibration, machine is working fine. Kept the machine under observation. | Performed complete oximetry test, where neon intensity and neon number was found out of range. So done the tHb calibration after which kept the machine OFF for 30 minutes. After restarting the analyzer and successful calibration, machine is working fine. Kept the machine under observation. | SWO-108622 | No | Performed complete oximetry test, where neon intensity and neon number was found out of range. So done the tHb calibration after which kept the machine OFF for 30 minutes. After restarting the analyzer and successful calibration, machine is working fine. Kept the machine under observation. | Performed complete oximetry test, where neon intensity and neon number was found out of range. So done the tHb calibration after which kept the machine OFF for 30 minutes. After restarting the analyzer and successful calibration, machine is working fine. Kept the machine under observation. | Adjustment | HW | Oxi module | tHb cal
ABL815 FLEX | Reinigung - Spülung Abfallverschlauchung - Proteinentfernung doppelt - Funktionsüberprüfung etc. Bem.: Gerinselgarantie (Kunde verwendet Pico 50 und Safe Pico) | Reinigung - Spülung Abfallverschlauchung - Proteinentfernung doppelt - Funktionsüberprüfung etc. Bem.: Gerinselgarantie (Kunde verwendet Pico 50 und Safe Pico) | SWO-107963 | No | Cleaning - Flushing waste hose - Protein removal double - function check, etc. Bem.: Gerinsel guarantee (customer uses Pico 50 and Safe Pico) | Cleaning - Flushing waste hose - Protein removal double - function check, etc. Bem.: Gerinsel guarantee (customer uses Pico 50 and Safe Pico) | Defect | HW | Fluidic manifold | Waste tube
ABL825 FLEX | Replaced Lactate electrode. | Replaced Lactate electrode. | SWO-108005 | No | Replaced Lactate electrode. | Replaced Lactate electrode. | Assisted Replacement | Electrode | Electrode | Lac
ABL800 BASIC | Cleaned inlet rail surface. Checked the inlet probe and inlet gasket. Problem couldn't be rectified. | B | SWO-108278 | No | Cleaned inlet rail surface. Checked the inlet probe and inlet gasket. Problem couldn't be rectified. |  | Revisit required | HW | Inlet module | Probe
ABL800 BASIC | cleaned the screen and calibrated the touch screen then also its not working properly we need to replace new touch screen. | B | SWO-108653 | No | cleaned the screen and calibrated the touch screen then also its not working properly we need to replace new touch screen. |  | Revisit required | HW | DMS_PC | Touch screen
ABL810 FLEX | Display komplett ersetzt, kalibriert und Systemeinstellungen vorgenommen, Test i. O. | B | SWO-108356 | No | Display completely replaced, calibrated and system settings made, test in the way |  | Defect | HW | DMS_PC | Touch screen
AQT90 FLEX | Cleaned dust inside AQT90, restarted AQT90 successfully | Cleaned dust inside AQT90, restarted AQT90 successfully | SWO-108350 | No | Cleaned dust inside AQT90, restarted AQT90 successfully | Cleaned dust inside AQT90, restarted AQT90 successfully | Cleaned | HW | Storage Module | Storage wheel, Sensor | Rotatestoragewheel
AQURE | Arrived on site and attended analyser to investigate customer reported fault. checked through the analyser radiance aqure connection settings. these were all set correctly and the handshake connection status works correctly. I then checked the netop remote support option in communications menu. unticked this and reapplied it. this connected properly also. I plugged in a usb keyboard and went into the remote support software settings and made sure that the encryption settings were set correctly. Netop was not ticked in the encryption menu. I set this and applied the changes. I than asked the customer to retry the connection. this did not resolve the issue. I checked the firewall setup and made sure that netop was set in the exceptions list. it was in the list. I then went to the lab and tried the connection using bot IE and Chrome browser. this did not make a difference. the actual issue is with aqure timing out when trying to load the data for the ITU analyser. Paul Northall advised that the Timeout setting may need to be lengthend to accommodate for the low resourses of the demo AQURE server. IT will log in and adjust this. | Arrived on site and attended analyser to investigate customer reported fault. checked through the analyser radiance aqure connection settings. these were all set correctly and the handshake connection status works correctly. I then checked the netop remote support option in communications menu. unticked this and reapplied it. this connected properly also. I plugged in a usb keyboard and went into the remote support software settings and made sure that the encryption settings were set correctly. Netop was not ticked in the encryption menu. I set this and applied the changes. I than asked the customer to retry the connection. this did not resolve the issue. I checked the firewall setup and made sure that netop was set in the exceptions list. it was in the list. I then went to the lab and tried the connection using bot IE and Chrome browser. this did not make a difference. the actual issue is with aqure timing out when trying to load the data for the ITU analyser. Paul Northall advised that the Timeout setting may need to be lengthend to accommodate for the low resourses of the demo AQURE server. IT will log in and adjust this. | SWO-107984 | Yes | Arrived on site and attended analyser to investigate customer reported fault. checked through the analyser radiance aqure connection settings. these were all set correctly and the handshake connection status works correctly. I then checked the netop remote support option in communications menu. unticked this and reapplied it. this connected properly also. I plugged in a usb keyboard and went into the remote support software settings and made sure that the encryption settings were set correctly. Netop was not ticked in the encryption menu. I set this and applied the changes. I than asked the customer to retry the connection. this did not resolve the issue. I checked the firewall setup and made sure that netop was set in the exceptions list. it was in the list. I then went to the lab and tried the connection using bot IE and Chrome browser. this did not make a difference. the actual issue is with aqure timing out when trying to load the data for the ITU analyser. Paul Northall advised that the Timeout setting may need to be lengthend to accommodate for the low resourses of the demo AQURE server. IT will log in and adjust this. | Arrived on site and attended analyser to investigate customer reported fault. checked through the analyser radiance aqure connection settings. these were all set correctly and the handshake connection status works correctly. I then checked the netop remote support option in communications menu. unticked this and reapplied it. this connected properly also. I plugged in a usb keyboard and went into the remote support software settings and made sure that the encryption settings were set correctly. Netop was not ticked in the encryption menu. I set this and applied the changes. I than asked the customer to retry the connection. this did not resolve the issue. I checked the firewall setup and made sure that netop was set in the exceptions list. it was in the list. I then went to the lab and tried the connection using bot IE and Chrome browser. this did not make a difference. the actual issue is with aqure timing out when trying to load the data for the ITU analyser. Paul Northall advised that the Timeout setting may need to be lengthend to accommodate for the low resourses of the demo AQURE server. IT will log in and adjust this.
ABL80 FLEX CO-OX OSM | センサーカセットにて現象発生。 カセット交換にて復旧致しました。 | センサーカセットにて現象発生。 カセット交換にて復旧致しました。 | SWO-108052 | No | A phenomenon occurs in the sensor cassette. We have recovered by cassette exchange. | A phenomenon occurs in the sensor cassette. We have recovered by cassette exchange.
ABL80 FLEX CO-OX OSM | Fehlersuche Fließwege überprüft Testmessungen i.O. | Fehlersuche Fließwege überprüft Testmessungen i.O. | SWO-108090 | No | Troubleshooting Flow paths checks test measurements in the | Troubleshooting Flow paths checks test measurements in the
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-109447 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL80 FLEX BASIC | So SC need replacement. Replace the SC with new one. After replace the SC unit is working properly. | So SC need replacement. Replace the SC with new one. After replace the SC unit is working properly. | SWO-108366 | No | So SC need replacement. Replace the SC with new one. After replace the SC unit is working properly. | So SC need replacement. Replace the SC with new one. After replace the SC unit is working properly.
ABL90 |  | B | D0E02052364CB410 | Yes |  |  | Re classified RMED | Customer training
ABL90 FLEX | ctHb Berichtbarer Bereich überprüft -0.20 bis 27 g dl. I.O. Oxi. Koplit Test I.O. Abl Kalibriert und tHb Kalibration durchgeführt I.O. Drei Blutproben gemessen alle Werte wurden angezeigt. | ctHb Berichtbarer Bereich überprüft -0.20 bis 27 g dl. I.O. Oxi. Koplit Test I.O. Abl Kalibriert und tHb Kalibration durchgeführt I.O. Drei Blutproben gemessen alle Werte wurden angezeigt. | SWO-108211 | No | ctHb Reportable range checked -0.20 to 27 g dl. I.O. Oxi. Koplit Test I.O. Abl Calibrated and tHb Calibration Performed I.O. Three blood samples measured all values were displayed. | ctHb Reportable range checked -0.20 to 27 g dl. I.O. Oxi. Koplit Test I.O. Abl Calibrated and tHb Calibration Performed I.O. Three blood samples measured all values were displayed. | Adjustment | HW | Oxi module | tHb Cal
ABL90 FLEX | Gerät überprüft, Mehrmals die Software überspielt, Konfiguration eingespielt, Konfiguration von Hand eingegeben | Konfiguration, Netzwerkanbindung angepasst, Sensorkassette installiert, Abklärung offener Fragen Benutzerverwaltung wird am Montag mit Herrn Janning (EDV) geklärt | SWO-108564 | No | Device checked, multiple times overplayed the software, configuration played in, configuration entered manually | Configuration, network connection adapted, sensor cartridge installed, clarification of open questions User administration will be clarified on Monday with Mr. Janning (EDV) | Defect | SW | Error code 2 | Error code 2
ABL90 FLEX | Software mehrfach überspielt, Konfiguration eingelesen, Lösungspack ersetzt, tHb Kal durchgeführt Folgebesuch | Konfiguration, Netzwerkanbindung angepasst, Sensorkassette installiert, Abklärung offener Fragen Benutzerverwaltung wird am Montag mit Herrn Janning (EDV) geklärt | SWO-108562 | Yes | Software repeatedly overplayed, configuration read in, solution pack replaced, tHb Kal carried out follow-up visit | Configuration, network connection adapted, sensor cartridge installed, clarification of open questions User administration will be clarified on Monday with Mr. Janning (EDV) | Re classified RMED | Double entry
ABL90 FLEX PLUS | Replaced solution pack and sensor cassette. | Replaced solution pack and sensor cassette. | SWO-109182 | No | Replaced solution pack and sensor cassette. | Replaced solution pack and sensor cassette. | Assisted Replacement | SC | SC | Lac
ABL90 | Verificação dos registros, limpeza do sensor sem resultado, desmontagem do analisador, lubrificação do motor e limpeza do sensor, novamente sem resultado | B | D0DF0C6BC01539DA | No | Checking records, cleaning the sensor without result, disassembly of the analyzer, engine lubrication and sensor cleaning, again without result |  | Unknown | Revisit required | Revisit required | Revisit required
ABL825 FLEX | Einsatz einer neuen Lactat Elektrode. Membranmaterial Lot Nr. R3057. Empfindlichkeit steigt auf 320 pA, besser wie alte Elektrode. Wert höher erwartet. Test mit Elektrode, falls Empfindlichkeit weiter schnell absinkt, Tausch des EL MeET Moduls empfohlen. Kalibrationen und Qualitätskontrolle o. K. | Einsatz einer neuen Lactat Elektrode. Membranmaterial Lot Nr. R3057. Empfindlichkeit steigt auf 320 pA, besser wie alte Elektrode. Wert höher erwartet. Test mit Elektrode, falls Empfindlichkeit weiter schnell absinkt, Tausch des EL MeET Moduls empfohlen. Kalibrationen und Qualitätskontrolle o. K. | SWO-108273 | No | Use of a new lactate electrode. Membrane material Lot No. R3057. Sensitivity rises to 320 pA, better than old electrode. Value higher expected. Test with electrode, if sensitivity continues to drop rapidly, replacement of the EL MeET module recommended. Calibrations and quality control o.K. | Use of a new lactate electrode. Membrane material Lot No. R3057. Sensitivity rises to 320 pA, better than old electrode. Value higher expected. Test with electrode, if sensitivity continues to drop rapidly, replacement of the EL MeET module recommended. Calibrations and quality control o.K. | Assisted Replacement | Electrode | Electrode | Lac
ABL825 FLEX | Reinigung und polieren des Rahmens. Analysator reagiert wieder normal und schnell auf Eingabe auf Display. Tausch Hardware nicht notwendig. Einweisung Mitarbeiter Kollege. Kalibrationen und Qualitätskontrolle o. K. | Reinigung und polieren des Rahmens. Analysator reagiert wieder normal und schnell auf Eingabe auf Display. Tausch Hardware nicht notwendig. Einweisung Mitarbeiter Kollege. Kalibrationen und Qualitätskontrolle o. K. | SWO-108274 | No | Cleaning and polishing of the frame. Analyzer responds normally and quickly to input to display. Swap hardware not necessary. Instruction employee colleague. Calibrations and quality control o.K. | Cleaning and polishing of the frame. Analyzer responds normally and quickly to input to display. Swap hardware not necessary. Instruction employee colleague. Calibrations and quality control o.K. | Defect | HW | DMS_PC | Touch screen
ABL810 FLEX | PH Dichtring erneuert. Cl Dummyelektrode erneuert. PCO2 Elektrode erneuert. Alle Schläuche aus dem JW Kit erneuert. Mehrfach kalibriert und QC Kontrollen durchgeführt. Pumpenrotor und Ventile des PH BG Moduls erneuert. Erneut mehrfach kalibriert und mehrere QC Kontrollen gestartet. Werte haben sich deutlich verbessert. | PH Dichtring erneuert. Cl Dummyelektrode erneuert. PCO2 Elektrode erneuert. Alle Schläuche aus dem JW Kit erneuert. Mehrfach kalibriert und QC Kontrollen durchgeführt. Pumpenrotor und Ventile des PH BG Moduls erneuert. Erneut mehrfach kalibriert und mehrere QC Kontrollen gestartet. Werte haben sich deutlich verbessert. | SWO-108143 | No | PH sealing ring renewed. Cl dummy electrode renewed. PCO2 electrode renewed. All hoses from the JW kit renewed. Multiple calibrated and QC checks performed. pump rotor and valves of the PH BG module. Again calibrated multiple times and several QC checks started. Values have improved significantly. | PH sealing ring renewed. Cl dummy electrode renewed. PCO2 electrode renewed. All hoses from the JW kit renewed. Multiple calibrated and QC checks performed. pump rotor and valves of the PH BG module. Again calibrated multiple times and several QC checks started. Values have improved significantly. | Setup | SW | Configuration | Configuration
ABL800 BASIC | ﾗﾝﾌﾟﾕﾆｯﾄ交換にて対応。 | ﾗﾝﾌﾟﾕﾆｯﾄ交換にて対応。 | SWO-109861 | No | It corresponds by the lamp unit exchange. | It corresponds by the lamp unit exchange. | Defect | HW | Oxi module | Lamp unit
ABL800 BASIC | ｶﾞｽﾄﾗﾝｽﾃﾞｭｰｻ、ﾌﾛｰﾚｷﾞｭﾚｰﾀ(Gas1)交換 | ｶﾞｽﾄﾗﾝｽﾃﾞｭｰｻ、ﾌﾛｰﾚｷﾞｭﾚｰﾀ(Gas1)交換 | SWO-109919 | No | Gas transducers, flow regulator (Gas1) replacement | Gas transducers, flow regulator (Gas1) replacement | Defect | HW | Gas module | Pressure Transducer
ABL800 BASIC | Replaced the new Display unit for DMS910, ABL800 FLEX (910-320), restarted instrument found touchscreen is working, found instrument is functional. | Replaced the new Display unit for DMS910, ABL800 FLEX (910-320), restarted instrument found touchscreen is working, found instrument is functional. | SWO-108072 | No | Replaced the new Display unit for DMS910, ABL800 FLEX (910-320), restarted instrument found touchscreen is working, found instrument is functional. | Replaced the new Display unit for DMS910, ABL800 FLEX (910-320), restarted instrument found touchscreen is working, found instrument is functional. | Defect | HW | DMS_PC | Touch screen
ABL810 FLEX |  | B | D0E34E2898510F9F | No |  |  | Defect | HW | DMS_PC | Touch screen
TCM TOSCA | Mam-brain replaced by said procedure. Study done for 2-3 hours.Dr satisfied with TCM redings and want test for next three studies and than give feedback. | Mam-brain replaced by said procedure. Study done for 2-3 hours.Dr satisfied with TCM redings and want test for next three studies and than give feedback. | SWO-108489 | No | Mam-brain replaced by said procedure. Study done for 2-3 hours. Dr satisfied with TCM redings and want test for next three studies and than give feedback. | Mam-brain replaced by said procedure. Study done for 2-3 hours. Dr satisfied with TCM redings and want test for next three studies and than give feedback.
ABL80 FLEX CO-OX | Meßkassette gespült, tHb Kal durchgeführt, Konfiguration angepasst | Meßkassette gespült, tHb Kal durchgeführt, Konfiguration angepasst | SWO-108617 | No | Measuring cassette flushed, tHb Kal performed, configuration adapted | Measuring cassette flushed, tHb Kal performed, configuration adapted
ABL80 CO-OX | Unit leaking from inlet | Checked waste tubes for blockage, Replaced waste manifold and stained door hinge. Found problem to be defective solution pack. Replaced pack | D0E0FCB9AC344161 | No | Unit leaking from inlet | Checked waste tubes for blockage, Replaced waste manifold and stained door hinge. Found problem to be defective solution pack. Replaced pack
ABL90 |  | B | D0E0229A361AB412 | Yes |  |  | Re classified RMED | SWU
ABL90 |  | B | D0E0262F361BB415 | No |  |  | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
ABL90 FLEX | LP getauscht alle QC IO | LP getauscht alle QC IO | SWO-108756 | No | LP swapped all QC IO | LP swapped all QC IO | Assisted Replacement | SP | SP | pCO2
ABL90 FLEX | Konfiguration, Netzwerkanbindung angepasst, Sensorkassette installiert, Abklärung offener Fragen Benutzerverwaltung wird am Montag mit Herrn Janning (EDV) geklärt | Konfiguration, Netzwerkanbindung angepasst, Sensorkassette installiert, Abklärung offener Fragen Benutzerverwaltung wird am Montag mit Herrn Janning (EDV) geklärt | SWO-108566 | No | Configuration, network connection adapted, sensor cartridge installed, clarification of open questions User administration will be clarified on Monday with Mr. Janning (EDV) | Configuration, network connection adapted, sensor cartridge installed, clarification of open questions User administration will be clarified on Monday with Mr. Janning (EDV) | Setup | His/Lis | network | network | LEH Mail (3X)
ABL90 FLEX | Code39 spezifische Einstellungen hinzugefügt, Neustart abgewartet, Barcode getestet, i.O. | B | SWO-108765 | No | Added Code39 specific settings, waited for restart, barcode tested, i.O. |  | Setup | SW | DMS_PC | Barcode Reader
ABL825 FLEX | Locked in datalogs and found that lactate and glucose has excessive drift, lactate drift .4, .3, and .2 and glucose was 16 and 18. Inspected measuring chamber and glucose and lactate connectors. found measuring chamber worn and both connectors worn. Also noticed annual maintenance is due next month. | Replaced measuring chamber and both lactate and glucose connectors and did cleaning of tip of lactate electrode and re-membraned and completed annual maintenance. All tubing and parts replaced from annual kit. | D0E007693604B910 | No | Locked in datalogs and found that lactate and glucose has excessive drift; lactate drift .4, .3, and .2 and glucose was 16 and 18. Inspected measuring chamber and glucose and lactate connectors. found measuring chamber worn and both connectors worn. Also noticed annual maintenance is due next month. | Replaced measuring chamber and both lactate and glucose connectors and did cleaning of tip of lactate electrode and re-membraned and completed annual maintenance. All tubing and parts replaced from annual kit. | Assisted Replacement | Electrode | Membrane | Lac
ABL835 FLEX | Calibration d sur Ph et electolyte et Qc hors limite sur sodium et o'ring Ph | Remplacé electrode de reference et electode de sodium | D0DF9545363D5662 | No | D calibration on Ph and electolyte and Qc off limit on sodium and o'ring Ph | Replaced reference electrode and sodium electode | Assisted Replacement | Electrode | Electrode | Reference
ABL835 FLEX | Hbユニットの直接洗浄実施、tHb CAL、QC測定 | B | SWO-108334 | No | Direct cleaning of Hb units, tHb CAL and QC measurements |  | Adjustment | HW | Oxi module | tHb cal
ABL827 FLEX | Kappe der QC Ampullen war zu Hoch positioniert. Nach korrekter Position der Ampullen Kappen funktionierte das AQC Modul einwandfrei. Absprache mit Frau Führbach getroffen falls der Fehler wieder Auftritt wird sie sich nochmals melden. | Kappe der QC Ampullen war zu Hoch positioniert. Nach korrekter Position der Ampullen Kappen funktionierte das AQC Modul einwandfrei. Absprache mit Frau Führbach getroffen falls der Fehler wieder Auftritt wird sie sich nochmals melden. | SWO-108754 | No | Cap of the QC ampoules was positioned too high. After the correct position of the ampoules caps, the AQC module worked perfectly. Agreement with Ms. Führbach if the error again she will contact again. | Cap of the QC ampoules was positioned too high. After the correct position of the ampoules caps, the AQC module worked perfectly. Agreement with Ms. Führbach if the error again she will contact again. | Defect | Consumables | AutoCheck Module | Autocheck Ampoule
AQT90 FLEX |  | B | D0DFA2BB360C7996 | No |  |  | Defect | HW | Analyzer Control Module | Barcode reader, display
ABL80 FLEX BASIC | Placed order for new printer head | B | SWO-109422 | No | Placed order for new printer head
ABL90 FLEX PLUS | Fluidic tests ok and replaced SC. | Fluidic tests ok and replaced SC. | SWO-110673 | No | Fluidic tests ok and replaced SC. | Fluidic tests ok and replaced SC. | Assisted Replacement | SC | SC | K+ and Ca++ | 134
ABL80 FLEX BASIC |  | B | SWO-108680 | No
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-109448 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL90 FLEX | 更换903-183修复仪器(by zhenjie) | 更换903-183修复仪器(by zhenjie) | SWO-108359 | No | Replace the 903-183 repair instrument (by zhenjie) | Replace the 903-183 repair instrument (by zhenjie) | Defect | HW | Basic Unit | Sensor Interface PCB
AQT90 FLEX | 清洁仪器，重新定标 | 清洁仪器，重新定标 | SWO-108666 | No | Clean the instrument and re-label it | Clean the instrument and re-label it | Cleaned | HW | Process Module | Process Unit sensor | 5257
ABL825 FLEX | Überprüfung und Analysator teilweise demontiert - Erneuerung CPU Einheit - Rückmontage - Erneuerung pH Dichtring (getrocknete Blutreste Auslauf pH BG Modul) - Funktionsüberprüfung etc. | Überprüfung und Analysator teilweise demontiert - Erneuerung CPU Einheit - Rückmontage - Erneuerung pH Dichtring (getrocknete Blutreste Auslauf pH BG Modul) - Funktionsüberprüfung etc. | SWO-108696 | No | Checking and analyzer partially dismantled - renewal CPU unit - back assembly - renewal pH sealing ring (dried blood residues outlet pH BG module) - function check etc. | Checking and analyzer partially dismantled - renewal CPU unit - back assembly - renewal pH sealing ring (dried blood residues outlet pH BG module) - function check etc. | Defect | HW | DMS_PC | CPU
AQT90 FLEX | UNDEFINED TIMIMG ERROR | RESET SYSTEM AND CLEARED ALL SYSTEMS IN SERVICE AND REMOVED SAMPLES | D0E3D0E798411DA6 | No | UNDEFINED TIMIMG ERROR | RESET SYSTEM AND CLEARED ALL SYSTEMS IN SERVICE AND REMOVED SAMPLES | Defect | SW | Analyzer Control Module | Off/On | 5357
ABL80 FLEX CO-OX | Flüssigkeitswege kontrolliert, Schlauchkit erneuert, Hemolysator durchgespült, CO-OX Test durchgeführt (Neon Position Fehler). Oximetrie deaktiviert | Flüssigkeitswege kontrolliert, Schlauchkit erneuert, Hemolysator durchgespült, CO-OX Test durchgeführt (Neon Position Fehler). Oximetrie deaktiviert | SWO-109106 | No | Fluid pathways checked, hose kit renewed, hemolyser flushed through, CO-OX test performed (neon position error). Oximetry disabled | Fluid pathways checked, hose kit renewed, hemolyser flushed through, CO-OX test performed (neon position error). Oximetry disabled
ABL80 FLEX BASIC | 更换两张新测试卡修复仪器 | 更换两张新测试卡修复仪器 | SWO-108852 | No | Replace two new test card repair instruments | Replace two new test card repair instruments
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-109446 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-109477 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL80 FLEX BASIC | Cleaned Sc and calibrated again. calibration is coming ok. Checked with samples and found that in normal patient PO2 and SO2 is coming low.Need to check with QC. | Checked with level 1 and level 3 Quality control. Found that PO2 results and Ph results are coming low that acceptable range. Need to replace sensor cassette. | SWO-110430 | No | Cleaned Sc and calibrated again. calibration is coming ok. Checked with samples and found that in normal patient PO2 and SO2 is coming low. Need to check with QC. | Checked with level 1 and level 3 Quality control. Found that PO2 results and Ph results are coming low that acceptable range. Need to replace sensor cassette.
ABL90 FLEX | Gerät aus, Sensor Platine gew., Gerät ein, Startphase abgewartet, Kal. + QC i.O. | Gerät aus; Sensor Platine gew.; Gerät ein; Startphase abgewartet; Kal. + QC i.O. | SWO-108919 | No | device off; Sensor board. device; start-up phase awaited; Kal. + QC in the first place | device off; Sensor board. device; start-up phase awaited; Kal. + QC in the first place | Defect | HW | Basic Unit | Sensor Interface PCB
ABL90 |  | B | D0E375E4981EB418 | No |  |  | Adjustment | HW | Oxi module | tHb Cal
TCM4 | 更换了气体模块后修复---jack.chu | 更换了气体模块后修复---jack.chu | SWO-110882 | No | Repair ---jack.chu after replacing gas module | Repair ---jack.chu after replacing gas module
TCM CombiM |  | B | SWO-108789 | No
AQT90 FLEX |  | B | SWO-108806 | No |  |  | Defect | HW | Storage Module | Wheel lock, Sensor | 3998
ABL835 FLEX | 大体Oxiﾓｼﾞｭｰﾙと入れ替え。 | 大体Oxiﾓｼﾞｭｰﾙと入れ替え。 | SWO-109788 | No | It is roughly replaced with the Oxi module. | It is roughly replaced with the Oxi module. | Defect | HW | Oxi module | Oxi module
ABL825 FLEX | pCo2 reading at or above package insert. | Replaced pCO2 electrode. | D0E39EFB985AE845 | No | pCo2 reading at or above package insert. | Replaced pCO2 electrode. | Assisted Replacement | Electrode | Electrode | pCO2
ABL800 BASIC | touch screen calibration done. system under observation. if problem repeated necessary action will be taken. | touch screen calibration done. system under observation. if problem repeated necessary action will be taken. | SWO-110101 | No | touch screen calibration done. system under observation. if problem repeated necessary action will be taken. | touch screen calibration done. system under observation. if problem repeated necessary action will be taken. | Adjustment | HW | DMS_PC | Touch screen
ABL825 FLEX | PC Modul aus GS Gerät eingebaut. Konfig. Anbindung ok. zu Probelauf. | PC Modul aus GS Gerät eingebaut. Konfig. Anbindung ok. zu Probelauf. | SWO-109039 | No | PC module built-in from GS device. Config. Connection ok. to test run. | PC module built-in from GS device. Config. Connection ok. to test run. | Revisit required | HW | DMS_PC | DMS_PC
ABL825 FLEX | Driftgrenzen leicht erweitert. | B | SWO-109092 | No | Drift limits slightly extended. |  | Assisted Replacement | Electrode | Electrode | Reference
ABL827 FLEX | DATA LOGS SHOW HALOGEN INTENSITY TOO HIGH.ALSO 247 SPECTRUM A TO D ERRORS | UNABLE TO GET THE LAMP TO STAY STEADY.REPLACED LAMP UNIT AND OXI MODULE BOARD.ADJUST LAMP TO 3385.PERFORMED THB CAL IN CUSTOMER AND SERVICE MODES | D0E37158981F891B | No | DATA LOGS SHOW HALOGEN INTENSITY TOO HIGH. ALSO 247 SPECTRUM A TO D ERRORS | UNABLE TO GET THE LAMP TO STAY STEADY. REPLACED LAMP UNIT AND OXI MODULE BOARD. ADJUST LAMP TO 3385.PERFORMED THB CAL IN CUSTOMER AND SERVICE MODES | Defect | HW | Oxi module | Lamp unit
AQT90 FLEX | DISCUSSED WITH TECH SUPPORT BUT WHEN ARRIVED ALL ERRORS GONE | NONE NECESSARY SENT A SERVICE DUMP TO BRIAN VAZQUEZ | D0E3D4FF98551DAB | Yes | DISCUSSED WITH TECH SUPPORT BUT WHEN ARRIVED ALL ERRORS GONE | NONE NECESSARY SENT A SERVICE DUMP TO BRIAN VAZQUEZ | Re classified RMED | To collect Data
AQT90 FLEX | Removed the broken Piece and checked, all working good. Planned to calibrate the new lot. | Removed the broken Piece and checked, all working good. Planned to calibrate the new lot. | SWO-109746 | No | Removed the broken Piece and checked, all working good. Planned to calibrate the new lot. | Removed the broken Piece and checked, all working good. Planned to calibrate the new lot. | Defect | Consumables | Test Kit | Cartridge | Broken
TCM TOSCA | Gas low found only 10%. Unable to calibrate now. need to order new gas cylinder. Ask Mr C S Singh to order the same. Revised pricing required for order. old order copy asked for part no and old pricing. | Gas low found only 10%. Unable to calibrate now. need to order new gas cylinder. Ask Mr C S Singh to order the same. Revised pricing required for order. old order copy asked for part no and old pricing. | SWO-109412 | No | Gas low found only 10%. Unable to calibrate now. need to order new gas cylinder. Ask Mr C S Singh to order the same. Revised pricing required for order. old order copy asked for part no and old pricing. | Gas low found only 10%. Unable to calibrate now. need to order new gas cylinder. Ask Mr C S Singh to order the same. Revised pricing required for order. old order copy asked for part no and old pricing.
ABL80 FLEX CO-OX | Gerät überprüft, Hemolyser erneuert, tHb Kal o.k. | Gerät überprüft, Hemolyser erneuert, tHb Kal o.k. | SWO-109492 | No | Device checked, Hemolyser renewed, tHb Kal o.k. | Device checked, Hemolyser renewed, tHb Kal o.k.
ABL80 FLEX BASIC | 收到的配件是坏的(by andrew) | 更换了借用的电机。还要借用打印机模块，第一次借用时收到的备件有质量问题，第二次借用时仓库缺货未能借到。暂时先不处理，等仓库到货再借用。（by andrew） | SWO-109118 | No | The accessories received are bad (by andrew) | The borrowed motor has been replaced. Also borrow the printer module, the first time to borrow the spare parts received quality problems, the second loan warehouse out of stock failed to borrow. For the time being, do not deal with, wait for the warehouse to arrive before borrowing. (by andrew)
ABL80 FLEX BASIC | Lüfter erneuert | Lüfter erneuert | SWO-109262 | No | Fan renewed | Fan renewed
ABL80 FLEX BASIC | 更换测试卡修复仪器 | 更换测试卡修复仪器 | SWO-109154 | No | Replace the test card repair instrument | Replace the test card repair instrument
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-109485 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL80 FLEX BASIC | 更换新测试卡后解决 | 更换新测试卡后解决 | SWO-109791 | No | Resolve after replacing a new test card | Resolve after replacing a new test card
ABL90 |  | B | D0E4F9490E2B1999 | No |  |  | Setup | His/Lis | His/lis | His/Lis | Could be reclass
ABL90 | COMPLETED COMPLETE OXI TEST AND RETESTED SOLUTION PACK. RAN 14 SAMPLES TO REPRODUCE ERROR. COULD NOT REPRODUCE ERROR | NONE NECESSARY | D0E4C2F90E1A40B0 | No | COMPLETED COMPLETE OXI TEST AND RETESTED SOLUTION PACK. RAN 14 SAMPLES TO REPRODUCE ERROR. COULD NOT REPRODUCE ERROR | NONE NECESSARY | Unknown | HW | Oxi module | Hemolyzer
ABL90 FLEX PLUS | Gerät überprüft. SC zeigt im Verlauf steigende Lac Werte Level C seit 2.6.16 Reklamation nach Willich. SC eingebaut keine Messung PCo2 Lac möglich. Am 15.6.16 erneuter Besuch QC nun ok. | Gerät überprüft. SC zeigt im Verlauf steigende Lac Werte Level C seit 2.6.16 Reklamation nach Willich. SC eingebaut keine Messung PCo2 Lac möglich. Am 15.6.16 erneuter Besuch QC nun ok. | SWO-109403 | No | device checked. SC shows in the course of increasing Lac values level C since 2.6.16 complaint to Willich. SC built in no measurement PCo2 Lac possible. On 15.6.16 again visit QC now ok. | device checked. SC shows in the course of increasing Lac values level C since 2.6.16 complaint to Willich. SC built in no measurement PCo2 Lac possible. On 15.6.16 again visit QC now ok. | Assisted Replacement | SC | SC | Lac
ABL90 | Troca da entrada e inspeção da sonda, limpeza do circuito, estabilização e passagem dos QCs | B | D0E50C9F0E221838 | No | Exchange of input and inspection of the probe, cleaning the circuit, stabilization and passage of the QCs |  | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
ABL837 FLEX | Found problem with Internal Barcode reader, removed the unit and cleaned the Barcode reader but problem was remaining same, need to replace with new. | Found problem with Internal Barcode reader, removed the unit and cleaned the Barcode reader but problem was remaining same, need to replace with new. | SWO-109310 | No | Found problem with Internal Barcode reader, removed the unit and cleaned the Barcode reader but problem was remaining same, need to replace with new. | Found problem with Internal Barcode reader, removed the unit and cleaned the Barcode reader but problem was remaining same, need to replace with new. | Revisit required | HW | Basic Unit | Barcode reader
ABL835 FLEX | Lac電極交換実施 | B | SWO-112191 | No | Lac electrode replacement |  | Assisted Replacement | Electrode | Electrode | Lac | Sensitivity
ABL827 FLEX | Einlaßnadel vom Kundenmaterial ersetzt, Einlaßkalibrierung, Kal &QK ok, Reinigungslösung ersetzt | Einlaßnadel vom Kundenmaterial ersetzt, Einlaßkalibrierung, Kal &QK ok, Reinigungslösung ersetzt | SWO-109197 | No | Inlet needle replaced by customer material, inlet calibration, Kal &QK ok, cleaning solution replaced | Inlet needle replaced by customer material, inlet calibration, Kal &QK ok, cleaning solution replaced | Defect | HW | Inlet module | Probe bend | inlet needle bent
ABL815 FLEX |  | B | D0E89480704F54C0 | No |  |  | Defect | HW | Basic Unit | Analyzer Control PCB
ABL825 FLEX | Fehlersuche, zur Zeit kein Fehler feststellbar. Daten lassen einen sporadischen Ausfall des Friction detectors vermuten. Ersatzteil bestellt. | Fehlersuche, zur Zeit kein Fehler feststellbar. Daten lassen einen sporadischen Ausfall des Friction detectors vermuten. Ersatzteil bestellt. | SWO-109765 | No | Troubleshooting, no error can be detected at the moment. Data suggest a sporadic failure of the Friction detector. spare part ordered. | Troubleshooting, no error can be detected at the moment. Data suggest a sporadic failure of the Friction detector. spare part ordered. | Defect | HW | Inlet module | Probe cal.
ABL825 FLEX | Gerät demontiert und gereinigt, Verstopfung beseitigt, Funktionskontrolle ok | Gerät demontiert und gereinigt, Verstopfung beseitigt, Funktionskontrolle ok | SWO-109560 | No | Device dismantled and cleaned, clogging eliminated, function control ok | Device dismantled and cleaned, clogging eliminated, function control ok | Cleaned | Fluidic pathway | Fluidic pathway | Waste
ABL80 FLEX CO-OX | Jahreswartung durchgeführt i.o Allgemeine Kontrolle i.o Zero unf therm Kal. duchgeführt i.o Thb Kal. i.o Manueller QA Zyklus i.o Datenbank bereinigt i.o | Jahreswartung durchgeführt i.o Allgemeine Kontrolle i.o Zero unf therm Kal. duchgeführt i.o Thb Kal. i.o Manueller QA Zyklus i.o Datenbank bereinigt i.o | SWO-109601 | No | Annual maintenance carried out i.o General control i.o Zero unf therm Kal. led by i.o Thb Kal. i.o Manual QA cycle i.o database cleaned i.o | Annual maintenance carried out i.o General control i.o Zero unf therm Kal. led by i.o Thb Kal. i.o Manual QA cycle i.o database cleaned i.o
ABL80 FLEX CO-OX | Jahreswartung durchgeführt i.o Allgemeine Kontrolle i.o Zero und therm Kal. durchgeführt i.o tHb Kal. i.o Manueller QA Zyklus i.o Datenbank bereinigt i.o | Jahreswartung durchgeführt i.o Allgemeine Kontrolle i.o Zero und therm Kal. durchgeführt i.o tHb Kal. i.o Manueller QA Zyklus i.o Datenbank bereinigt i.o | SWO-109602 | No | Annual maintenance carried out i.o General control i.o Zero and therm Kal. i.o tHb Kal. i.o Manual QA cycle i.o database cleaned i.o | Annual maintenance carried out i.o General control i.o Zero and therm Kal. i.o tHb Kal. i.o Manual QA cycle i.o database cleaned i.o
ABL80 | se procedio a la sustitución de la touch panel defectuosa. tras esto, se observo neumática obstruida. se realizo kit de tubos anual. tras esto, se monto equipo y se calibro. se comprobó correcto funcionamiento mediante CC. | B | D0EDB40A4852F741 | No | proceeded to replace the faulty touch panel. after this, there was a blocked pneumatic observe. pipe kit was done annually. after this, team is mounted and heated. DC operation was proven to be correct.
ABL80 FLEX CO-OX | Konfiguration angepasst, Sensorkassette und Lösungspack installiert | Technikereinsatz erbeten | SWO-109889 | No | Configuration adapted, sensor cartridge and solution pack installed | Technicians requested
ABL80 FLEX BASIC | Checked with level 1 and level 3 Quality control. Found that PO2 results and Ph results are coming low that acceptable range. Need to replace sensor cassette. | Checked with level 1 and level 3 Quality control. Found that PO2 results and Ph results are coming low that acceptable range. Need to replace sensor cassette. | SWO-110431 | No | Checked with level 1 and level 3 Quality control. Found that PO2 results and Ph results are coming low that acceptable range. Need to replace sensor cassette. | Checked with level 1 and level 3 Quality control. Found that PO2 results and Ph results are coming low that acceptable range. Need to replace sensor cassette.
ABL80 FLEX BASIC | 修理完了機返却、セットアップ。 | 修理完了機返却、セットアップ。 | SWO-109778 | No | Return machine repair completion machine, setup. | Return machine repair completion machine, setup.
ABL90 | Unit passed pO2 diagnostics, several SCs' were used prior to replacing Pump Rotor. | Replace suspect Pump Rotor. | D0E64E51843202C9 | No | Unit passed pO2 diagnostics, several SCs' were used prior to replacing Pump Rotor. | Replace suspect Pump Rotor. | Defect | HW | Pump rotor | pO2 | Ret.request 23/3-16
ABL90 |  | B | D0E623BB842A19A3 | No |  |  | Defect | Ext. HW | Ext. HW | His/Lis Cable
ABL90 | Customer tried performing back flush procedure with difficulty. Onsite troubleshooting traced blockage to the inlet side of the cassette bay unit. Not able to flush blockage from bay unit. | Replace sensor cassette bay unit. Perform all service fluid transport service programs. | D0ED8F7B4826416E | No | Customer tried performing back flush procedure with difficulty. Onsite troubleshooting traced blockage to the inlet side of the cassette bay unit. Not able to flush blockage from bay unit. | Replace sensor cassette bay unit. Perform all service fluid transport service programs. | Cleaned | Fluidic pathway | Fluidic pathway | Cassette retainer
ABL827 FLEX | Replaced pressure transducer unit with chassis. Tested and checked all functions. All working ok | Replaced pressure transducer unit with chassis. Tested and checked all functions. All working ok | SWO-109705 | No | Replaced pressure transducer unit with chassis. Tested and checked all functions. All working ok | Replaced pressure transducer unit with chassis. Tested and checked all functions. All working ok | Defect | HW | Gas module | Gas module
ABL825 FLEX |  | B | D0E5FB108433CDCE | No |  |  | Defect | HW | Oxi module | Oxi module
ABL825 FLEX | A) Exchanged the display and cable and the display still did not work. B) The analyzer is shown as the incorrect type. C) The FlexQ samples are often inhomogeneous. D) User has questions on Remote Support. | A) Replaced the CPU module and the display now works as it should. B) Disabled the unused parameters and performed the update system type program. C) Performed the Sample Tray alignment program and aligned the sample holders to the inlet probe. D) Gave the user documentation regarding remote support and explained the process. | D0F550310C3286A7 | No | A) Exchanged the display and cable and the display still did not work. B) The analyzer is shown as the incorrect type. C) The FlexQ samples are often inhomogeneous. D) User has questions on Remote Support. | A) Replaced the CPU module and the display now works as it should. B) Disabled the unused parameters and performed the update system type program. C) Performed the Sample Tray alignment program and aligned the sample holders to the inlet probe. D) Gave the user documentation regarding remote support and explained the process. | Defect | HW | DMS_PC | CPU | Touch screen
ABL817 FLEX | Verified. | Remembrained the A and B electrodes. Customer gas 2 was out. Waiting for customer to locate. Contacted the customer and they installed a new gas 2 tank. Analyzer is now back in service. Customer performed all calibrations and QC's. | D0E68E59840DB712 | No | Verified. | Remembrained the A and B electrodes. Customer gas 2 was out. Waiting for customer to locate. Contacted the customer and they installed a new gas 2 tank. Analyzer is now back in service. Customer performed all calibrations and QC's. | Assisted Replacement | Electrode | Membrane | Crea
ABL80 FLEX BASIC | 更换了借用的电机。还要借用打印机模块，第一次借用时收到的备件有质量问题，第二次借用时仓库缺货未能借到。暂时先不处理，等仓库到货再借用。（by andrew） | 更换了借用的电机。还要借用打印机模块，第一次借用时收到的备件有质量问题，第二次借用时仓库缺货未能借到。暂时先不处理，等仓库到货再借用。（by andrew） | SWO-109796 | No | The borrowed motor has been replaced. Also borrow the printer module, the first time to borrow the spare parts received quality problems, the second loan warehouse out of stock failed to borrow. For the time being, do not deal with, wait for the warehouse to arrive before borrowing. (by andrew) | The borrowed motor has been replaced. Also borrow the printer module, the first time to borrow the spare parts received quality problems, the second loan warehouse out of stock failed to borrow. For the time being, do not deal with, wait for the warehouse to arrive before borrowing. (by andrew)
ABL80 FLEX BASIC | 更换新测试卡修复仪器 | 更换新测试卡修复仪器 | SWO-111934 | No | Replace the new test card repair instrument | Replace the new test card repair instrument
ABL80 FLEX BASIC | 更换新测试卡修复仪器 | 更换新测试卡修复仪器 | SWO-109982 | No | Replace the new test card repair instrument | Replace the new test card repair instrument
ABL90 FLEX | Flushed system and checked for leaks ok. Ran service test passed ok. Replaced SP . 1 hour start up left to settle down. | Flushed system and checked for leaks ok. Ran service test passed ok. Replaced SP . 1 hour start up left to settle down. | SWO-110235 | No | Flushed system and checked for leaks ok. Ran service test passed ok. Replaced SP . 1 hour start up left to settle down. | Flushed system and checked for leaks ok. Ran service test passed ok. Replaced SP . 1 hour start up left to settle down. | Assisted Replacement | SP | SP | pO2
ABL90 | Several SCs' have been used to troubleshoot problem with no results. SPs' were also replaced. pO2 passing diagnostics.No errors other than failing Study. | Replace Pump Rotor. | D0E789ADFA626A25 | No | Several SCs' have been used to troubleshoot problem with no results. SPs' were also replaced. pO2 passing diagnostics. No errors other than failing Study. | Replace Pump Rotor. | Defect | HW | Pump rotor | pO2 | Ret.request 23/3-16
ABL90 FLEX PLUS |  | B | D0E73FF6FA5C9E4B | Yes |  |  | Re classified RMED | SWU
ABL90 FLEX PLUS | Complete display vervangen. Geplaatst: 910-353P1503N0573 Eruit: 910-353P1535N1775 | Complete display vervangen. Geplaatst: 910-353P1503N0573 Eruit: 910-353P1535N1775 | SWO-110515 | No | Replace the Complete display. Posted: 910-353P1503N0573 look like: 910-353P1535N1775 | Replace the Complete display. Posted: 910-353P1503N0573 look like: 910-353P1535N1775 | Defect | HW | DMS_PC | Display
ABL90 FLEX PLUS |  | B | D0E743A1FA149337 | Yes |  |  | Re classified RMED | SWU
ABL90 FLEX PLUS | Remove and inspect inlet within service programs, clean inlet gasket, check inlet probe was straight, calibrate inlet mechanism. Also visit other two new analysers on site on CICU and Ward 15 to compare SP usage and error logs. | Remove and inspect inlet within service programs, clean inlet gasket, check inlet probe was straight, calibrate inlet mechanism. Also visit other two new analysers on site on CICU and Ward 15 to compare SP usage and error logs. | SWO-110413 | No | Remove and inspect inlet within service programs, clean inlet gasket, check inlet probe was straight, calibrate inlet mechanism. Also visit other two new analysers on site on CICU and Ward 15 to compare SP usage and error logs. | Remove and inspect inlet within service programs, clean inlet gasket, check inlet probe was straight, calibrate inlet mechanism. Also visit other two new analysers on site on CICU and Ward 15 to compare SP usage and error logs. | Adjustment | HW | Auto Inlet module | Inlet Position Cal.
ABL90 FLEX | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-109930 | No | Repair after replacing a new test card | Repair after replacing a new test card
AQT90 FLEX | 更换了903-364后修复(by andrew) | 更换了903-364后修复(by andrew) | SWO-109568 | No | Repair after replacement of 903-364 (by andrew) | Repair after replacement of 903-364 (by andrew) | Defect | HW | Process Module | Optical Unit | 131
ABL825 FLEX | Nieuw speakertje gemonteerd. Nu weer geluid, b.v. bij intoetsen. | Nieuw speakertje gemonteerd. Nu weer geluid, b.v. bij intoetsen. | SWO-109991 | No | New Speaker mounted. Now sound again, e.g. when writing. | New Speaker mounted. Now sound again, e.g. when writing. | Defect | HW | DMS_PC | Loud Speaker
ABL835 FLEX | Spectrometer replaced. Halogen lamp adjusted to 3364. Service tHb calibration ran and passed. The customer is to report back to the local engineer if the results of the next EQA. | Spectrometer replaced. Halogen lamp adjusted to 3364. Service tHb calibration ran and passed. The customer is to report back to the local engineer if the results of the next EQA. | SWO-109988 | No | Spectrometer replaced. Halogen lamp adjusted to 3364. Service tHb calibration ran and passed. The customer is to report back to the local engineer if the results of the next EQA. | Spectrometer replaced. Halogen lamp adjusted to 3364. Service tHb calibration ran and passed. The customer is to report back to the local engineer if the results of the next EQA. | Defect | HW | Oxi module | Spectrometer | WEQAS
ABL800 BASIC | オートサンプラー モジュール引き上げ検証、 | B | SWO-109993 | No | Autosampler module pull verification, |  | Adjustment | HW | FlexQ Module | FlexQ Module
ABL817 FLEX | Ventile vom pH BG, EL MET und Crea Modul gereinigt und zum Teil die Membranen ern., HJW Kit eingebaut, Luftweg gereinigt, Ref.-Connector ern., Messkammer gereinigt, div. Elektroden neu bezogen, Gerät kalibriert, QC´s durchgef. | Fehler tritt sporadisch auf, Modul ausgebaut Ventile gereinigt, QC Daten zur weiteren Auswertung mitgenommen. | SWO-109966 | No | Valves cleaned by pH BG, EL MET and Crea module and partly the membranes ern., HJW kit installed, airway cleaned, ref.-connector ern., measuring chamber cleaned, various electrodes re-obtained, device calibrated, QC's through. | Error occurs sporadically, module removed valves cleaned, QC data taken for further evaluation. | Cleaned | Fluidic pathway | Fluidic pathway | Na+
ABL825 FLEX | FluidicChannelUnit ausgebaut | B | SWO-110253 | No | FluidicChannelUnit expanded |  | Defect | HW | Fluidic manifold | Fluidic manifold (leak current)
ABL815 FLEX | FluidicChannelUnit ausgebaut und Abfallventil ausgebaut - Ventil nicht fest genug verschraubt -komplett mit Lösung vollgelaufen - ersetzt, alle anderen Ventile überprüft - ebenfalls alle nicht fest genug und feucht - alle Ventile ausgebaut und getrocknet, alles wieder zusammengesetzt, Abfallschlauch am Anschluß 2 war im Durchführungsstück verklemmt - korrigiert - dadurch Überdruck bei den Ventilen und dadurch sehr wahrscheinlich Ventile undicht und damit feucht --> Isolationsfehler, Neustart, Kal und QC | FluidicChannelUnit ausgebaut | SWO-113658 | No | FluidicChannelUnit removed and waste valve removed - valve not screwed tightly enough - completely full with solution - replaced, all other valves checked - also not firm enough and moist - all valves removed and dried, all again compound, waste hose at connection 2 was jammed in the feed-in piece - corrected - thus overpressure at the valves and thus very likely valves leakand and thus damp --> Insulation error, restart, cal and QC | FluidicChannelUnit expanded | Defect | HW | Fluidic manifold | Valve
ABL80 FLEX CO-OX | Layout neu eingestellt und getestet | Layout neu eingestellt und getestet | SWO-110364 | No | Layout re-adjusted and tested | Layout re-adjusted and tested
ABL80 FLEX BASIC | 更换了前面板924-738后修复(by andrew) | 更换了前面板924-738后修复(by andrew) | SWO-109827 | No | Replacement of front panel 924-738 after repair (by andrew) | Replacement of front panel 924-738 after repair (by andrew)
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-110361 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-112181 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL80 | Telefonicamente se les indica que instalen nuevo SC pero no tienen. Se pide material por defectuoso. Al instalar nuevo SC no lo detecta. Se observa sensor board con los pines en muy mal estado. Traslado equipo a taller Madrid. Cambio sensor board. Tests Ok. Ahora detecta SC. Ciclo de calidad OK. Traslado de equipo de nuevo a su ubicación. | Cambio sensor board | D0E8A8757023A9FF | No | They are prompted to install new SCs but do not have them. Material is requested for defective. Installing new SC does not detect it. Sensor board is observed with the pins in very bad condition. Transfer equipment to workshop Madrid. Sensor board change. Tests Ok. Now it detects SC. Quality cycle OK. Moving equipment back to your location. | Change sensor board
ABL90 FLEX | Oximetrie im Servicemenü überprüft. Alles in Ordnung. Einweisung Herrn Hähnel in Servicemenü. Kalibrationen und Qualitätskontrollen in Ordnung. | Oximetrie im Servicemenü überprüft. Alles in Ordnung. Einweisung Herrn Hähnel in Servicemenü. Kalibrationen und Qualitätskontrollen in Ordnung. | SWO-110210 | No | Oximetry in the service menu. all right. Instruction Mr. Hähnel in the service menu. calibrations and quality controls in order. | Oximetry in the service menu. all right. Instruction Mr. Hähnel in the service menu. calibrations and quality controls in order. | Unknown | HW | Oxi module | Spectrum error
ABL90 FLEX PLUS | Fehlersuche, Scanner mit anderem quer getauscht -> gl. Effekt, nach Einschalten d. Bildschirmschoners u. anschließendem berühren d. Bildschirm`s, alles i.O., Lösung: - durch Antippen d. Bildschirmschoners, oder durch Antippen d. Feldes Name (Personal ID) sofortige Scannerfunktion | Fehlersuche; Scanner mit anderem quer getauscht -> gl. Effekt; nach Einschalten d. Bildschirmschoners u. anschließendem berühren d. Bildschirm`s, alles i.O.; Lösung: - durch Antippen d. Bildschirmschoners, oder durch Antippen d. Feldes Name (Personal ID) sofortige Scannerfunktion | SWO-110302 | No | Troubleshooting; Scanner with other cross-exchanged -> gl. Effect; after turning on the screen saver and then touching the screen, all in the picture; Solution: - by tapping the screen saver, or by tapping the field Name (Personal ID) instant scanner function | Troubleshooting; Scanner with other cross-exchanged -> gl. Effect; after turning on the screen saver and then touching the screen, all in the picture; Solution: - by tapping the screen saver, or by tapping the field Name (Personal ID) instant scanner function | Defect | SW | DMS_PC | Touch screen
AQT90 FLEX | 更换配件903-329修复仪器(by andrew) | 更换配件903-329修复仪器(by andrew) | SWO-109814 | No | Replacement accessory 903-329 repair instrument (by andrew) | Replacement accessory 903-329 repair instrument (by andrew) | Defect | HW | Aspiration Module | Needle Wash Unit | 1264, 1265
ABL825 FLEX | 更换了905-847后修复(by andrew) | 更换了905-847后修复(by andrew) | SWO-109801 | No | Repair after replacement of 905-847 (by andrew) | Repair after replacement of 905-847 (by andrew) | Defect | HW | Inlet module | Nipple kit | Nipple kit
ABL800 BASIC | Reinigung der Messkammern und bürsten der Elektrodenkontakte. Proteinentfernung durchgeführt. pH Elektrode gereinigt und unteren Elektrodendichring erneuert. tHb Kalibration durchgeführt mit eigenem Kalibrator. Metabolitelektroden mussten neu membraniert werden. Kalibrationen und Qualitätskontrolle o. K. | Reinigung der Messkammern und bürsten der Elektrodenkontakte. Proteinentfernung durchgeführt. pH Elektrode gereinigt und unteren Elektrodendichring erneuert. tHb Kalibration durchgeführt mit eigenem Kalibrator. Metabolitelektroden mussten neu membraniert werden. Kalibrationen und Qualitätskontrolle o. K. | SWO-110395 | No | Cleaning of the measuring chambers and brushing the electrode contacts. protein removal. pH electrode cleaned and lower electrode tight ring renewed. tHb calibration carried out with its own calibrator. Metabolite electrodes had to be re-membraneed. Calibrations and quality control o.K. | Cleaning of the measuring chambers and brushing the electrode contacts. protein removal. pH electrode cleaned and lower electrode tight ring renewed. tHb calibration carried out with its own calibrator. Metabolite electrodes had to be re-membraneed. Calibrations and quality control o.K. | Cleaned | Fluidic pathway | pH/BG Module | pH
ABL817 FLEX | Fließwege kontrolliert, OXI Test durchgeführt - OK, Netzteil ern., Gerät kalibriert, QC´s durchgef. | Fehler tritt sporadisch auf, Modul ausgebaut Ventile gereinigt, QC Daten zur weiteren Auswertung mitgenommen. | SWO-110315 | No | Flow paths controlled, OXI test performed - OK, power supply ern., device calibrated, QC's through. | Error occurs sporadically, module removed valves cleaned, QC data taken for further evaluation. | Unknown | HW | Basic Unit | Power Supply
ABL825 FLEX |  | B | D0E8B08670577DC8 | No |  |  | Assisted Replacement | Electrode | Electrode | Reference
ABL825 FLEX |  | B | D0E8B20D70497DCF | Yes |  |  | Re classified RMED |  |  |  | Might give pCO2 problem
ABL825 FLEX | 更换了905-847后修复 | B | SWO-109918 | No | Repairafter after replacement of 905-847 |  | Defect | HW | Inlet module | Nipple kit | Nipple kit
ABL825 FLEX | Flushed waste tubing and checked waste pump. | Replaced tubing connectors below inlet gasket and waste tubing from inlet gasket to waste pump. | D0E8725570182858 | No | Flushed waste tubing and checked waste pump. | Replaced tubing connectors below inlet gasket and waste tubing from inlet gasket to waste pump. | Defect | HW | Inlet module | Leaking | Nipple kit
ABL825 FLEX |  | B | D0EC6E0DD249FADF | No |  |  | Unknown | Unknown | Unknown | pO2
AQT90 FLEX | Fehlersuche, Reinigung Spiegel, Funktionskontrolle , Fehler nicht behebbar | Vorort Service gewünscht | SWO-110548 | No | Troubleshooting, cleaning mirror, function control, error not removable | Suburb service desired | Revisit required | HW | Storage Module | Barcode reader, Storage wheel
AQT90 FLEX |  | B | D0E8B5FE70639C9C | No |  |  | Defect | Consumables | Test Kit | Cartridge | 5769 (Punchout)
AQURE | Reset and checked all remote support settings ok. Rebooted analyser ok. Katy has tested quick access working ok. | Reset and checked all remote support settings ok. Rebooted analyser ok. Katy has tested quick access working ok. | SWO-110376 | Yes | Reset and checked all remote support settings ok. Rebooted analyser ok. Katy has tested quick access working ok. | Reset and checked all remote support settings ok. Rebooted analyser ok. Katy has tested quick access working ok.
ABL800 FLEX | Cleaning and protein removal performed. Performed the suggested electrode cleaning procedure and replaced the membranes. Still showing the same error. Has to replace the electrodes | B | SWO-110640 | No | Cleaning and protein removal performed. Performed the suggested electrode cleaning procedure and replaced the membranes. Still showing the same error. Has to replace the electrodes |  | Assisted Replacement | Electrode | Electrode | Lac | Calibration
ABL815 FLEX | Gerinnsel im Kanal vor der PH Elektrode, und Blut im Einlassmodul, Gerät gespült Blutreste entfernt, Kal Empf weiter nur ... Status bei 5,4, PH Elektrode Lot UB 02 Glasspitze mechanisch def Glas gerissen, PH Elektrode Lot LB 02 zur Überbrückung eingesetzt, Kal und QC durchgeführt E-Teil wird bestellt. | B | SWO-110535 | No | Clots in the canal in front of the PH electrode, and blood in the inlet module, device flushed blood residues removed, kale empf continues only ... Status at 5.4, PH electrode lot UB 02 glass tip mechanicaldef glass cracked, PH electrode solder LB 02 used for bridging, Kal and QC performed E part is ordered. |  | Assisted Replacement | Electrode | Electrode | pH
ABL80 FLEX BASIC | 更换测试卡解决 | 更换测试卡解决 | SWO-111570 | No | Replace the test card to resolve | Replace the test card to resolve
ABL80 FLEX BASIC | 测试卡修复仪器 | 测试卡修复仪器 | SWO-111572 | No | Test card repair instrument | Test card repair instrument
ABL80 FLEX BASIC |  | B | SWO-110586 | No
ABL80 OSM | Performed thb cal with not help, ran service test and passed. | Replaced hemolyzer. | D0EC2EA4D24DC030 | No | Performed thb cal with not help, ran service test and passed. | Replaced hemolyzer.
ABL90 FLEX | Wet Section ausgebaut, Sensorfahne ausgerichtet | Kd. bittet um Besuch eines Servicetechnikers. | SWO-110880 | No | Wet Section expanded, sensor flag aligned | Kd. asks to visit a service technician. | Adjustment | HW | Sensor Cassette Door | Sensor plate
ABL90 FLEX PLUS | Fehlersuche, LP-Wechsel nicht mgl., da kein Material vor Ort, LP quer zum Zweitgerät getauscht -> gl. Fehler = LP def., LP wieder zurückgetauscht, Fehler weiterhin, Abbruch d. Arbeiten, Anwender informiert -> kümmert sich um Bestellung -> meldet sich bei Bedarf, wenn neues Material eingetroffen ist | Fehlersuche; LP-Wechsel nicht mgl., da kein Material vor Ort; LP quer zum Zweitgerät getauscht -> gl. Fehler = LP def.; LP wieder zurückgetauscht; Fehler weiterhin; Abbruch d. Arbeiten; Anwender informiert -> kümmert sich um Bestellung -> meldet sich bei Bedarf, wenn neues Material eingetroffen ist | SWO-110703 | No | Troubleshooting; LP change not with me, since no material on site; LP crossto the second unit -> gl. Error = LP def.; LP exchanged again; Errors continue; Demolition of the works; User informs -> takes care of ordering -> reports when new material arrives if necessary | Troubleshooting; LP change not with me, since no material on site; LP crossto the second unit -> gl. Error = LP def.; LP exchanged again; Errors continue; Demolition of the works; User informs -> takes care of ordering -> reports when new material arrives if necessary | Assisted Replacement | SP | SP | pCO2 Level C
TCM CombiM | 6.13已经报修过了，这单需要取消。 | 6.13已经报修过了，这单需要取消。 | SWO-110573 | No | 6.13 has been reported for repair and this order needs to be cancelled. | 6.13 has been reported for repair and this order needs to be cancelled. | Defect | HW | Power Supply | Power Supply
ABL90 FLEX |  | B | SWO-110556 | No |  |  | Defect | HW | Basic Unit | Sensor Interface PCB
ABL90 FLEX |  | B | SWO-110618 | No |  |  | Defect | HW | ABL9 USER INTERFACE MODULE | ABL9 USER INTERFACE MODULE
AQT90 FLEX |  | B | SWO-110634 | No |  |  | Defect | HW | Cover and Chassis | Inlet Door | 5784
ABL800 BASIC | touch screen calibration. system under observation. | touch screen calibration. system under observation. | SWO-111814 | No | touch screen calibration. system under observation. | touch screen calibration. system under observation. | Adjustment | HW | DMS_PC | Touch screen
ABL825 FLEX |  | B | D0EC52C6D221CC6D | No |  |  | Assisted Replacement | Electrode | Electrode | Lac | Lifetime
ABL800 BASIC | replaced the pCO2 membrane done 2 point calibration found all results found with in the range. | replaced the pCO2 membrane done 2 point calibration found all results found with in the range. | SWO-112134 | No | replaced the pCO2 membrane done 2 point calibration found all results found with in the range. | replaced the pCO2 membrane done 2 point calibration found all results found with in the range. | Assisted Replacement | Electrode | Membrane | pCO2
ABL825 FLEX | Checked all logs to verify past performance and found several 791 errors in a row. Visual and operational inspection found V7 valve membrane sticking in fluidic housing. | Replaced the V7 valve and membrane. Ran 10 blood samples with no 791 errors. | D0EC983ED24AFC9A | No | Checked all logs to verify past performance and found several 791 errors in a row. Visual and operational inspection found V7 valve membrane sticking in fluidic housing. | Replaced the V7 valve and membrane. Ran 10 blood samples with no 791 errors. | Defect | HW | Fluidic manifold | Vent, air | V7
ABL800 BASIC | REPLACE PRINTER HEAD | REPLACE PRINTER HEAD | SWO-110723 | No | REPLACE PRINTER HEAD | REPLACE PRINTER HEAD | Revisit required | HW | Printer module | Printer Head
ABL825 FLEX | Some leaking reported during comparison testing. | Replaced 2 pump rotors (pH BG and waste). | D0ECA88BD21FAC74 | No | Some leaking reported during comparison testing. | Replaced 2 pump rotors (pH BG and waste). | Defect | HW | Fluidic manifold | Pump Rotor | Waste
ABL825 FLEX |  | B | D0EC80ACD22CDDAB | No |  |  | Defect | HW | Inlet module | Leaking
ABL825 FLEX |  | B | D0ED924348197861 | No |  |  | Revisit required | Electrode | Electrode | pO2
ABL80 FLEX BASIC |  | B | SWO-110890 | No
ABL80 FLEX BASIC | 更换新测试卡修复仪器 | 更换新测试卡修复仪器 | SWO-111924 | No | Replace the new test card repair instrument | Replace the new test card repair instrument
ABL80 FLEX BASIC | 更换新测试卡修复仪器 | 更换新测试卡修复仪器 | SWO-111931 | No | Replace the new test card repair instrument | Replace the new test card repair instrument
ABL80 FLEX BASIC |  | B | SWO-110846 | No
ABL80 FLEX BASIC | performed zeroing and replaced sensor cassette problem persists. need to replace sensor board. | performed zeroing and replaced sensor cassette problem persists. need to replace sensor board. | SWO-112694 | No | performed zeroing and replaced sensor cassette problem persists. need to replace sensor board. | performed zeroing and replaced sensor cassette problem persists. need to replace sensor board.
ABL90 FLEX | Fehlersuche vor Ort. Analysator scheint verstopft. Komplettes Schlauchsystem gespült. Keine Besserung. Komplettes Schlauchsystem getauscht. ( in diesem Zuge die Jahreswertung mit durchgeführt. Siehe extra Bericht. Sensorkassette aus Kundenbestand getauscht. Danach war Analysator wieder Messbereit. Bei eintreffen hatte Analysator keine Radiance Verbindung. Deshalb mit EDV Herrn Knappe, Medizintechnik Herrn Janßen sowie mit Herrn Kozelj telefoniert um Problem zu beheben. Nach erneuter Rücksprache mit Herrn Kozelj war das Gerät dann wieder mit Radiance verbunden. Herrn Janßen und Herrn knappe darüber informiert. QCs für alle Level durchgeführt. ABL wieder Messbereit. | Fehlersuche vor Ort. Analysator scheint verstopft. Komplettes Schlauchsystem gespült. Keine Besserung. Komplettes Schlauchsystem getauscht. ( in diesem Zuge die Jahreswertung mit durchgeführt. Siehe extra Bericht. Sensorkassette aus Kundenbestand getauscht. Danach war Analysator wieder Messbereit. Bei eintreffen hatte Analysator keine Radiance Verbindung. Deshalb mit EDV Herrn Knappe, Medizintechnik Herrn Janßen sowie mit Herrn Kozelj telefoniert um Problem zu beheben. Nach erneuter Rücksprache mit Herrn Kozelj war das Gerät dann wieder mit Radiance verbunden. Herrn Janßen und Herrn knappe darüber informiert. QCs für alle Level durchgeführt. ABL wieder Messbereit. | SWO-111149 | No | Troubleshooting on site. Analyzer seems clogged. Complete hose system flushed. No improvement. Complete hose system replaced. ( in this context, the annual evaluation was carried out. See extra report. Sensor cartridge replaced from customer stock. After that, the analyzer was ready to measure again. When analyser arrived, no radiance connection. That is why Mr Knappe, Medical Technology Mr Janßen and Mr Kozelj are on the phone with MR Knappe in order to fix the problem. After further consultation with Mr Kozelj, the device was then reconnected to Radiance. Mr Janßen and Mr. Knappe informed us of this. QCs for all levels. ABL ready to measure again. | Troubleshooting on site. Analyzer seems clogged. Complete hose system flushed. No improvement. Complete hose system replaced. ( in this context, the annual evaluation was carried out. See extra report. Sensor cartridge replaced from customer stock. After that, the analyzer was ready to measure again. When analyser arrived, no radiance connection. That is why Mr Knappe, Medical Technology Mr Janßen and Mr Kozelj are on the phone with MR Knappe in order to fix the problem. After further consultation with Mr Kozelj, the device was then reconnected to Radiance. Mr Janßen and Mr. Knappe informed us of this. QCs for all levels. ABL ready to measure again. | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
ABL90 FLEX | CPU交換 | B | SWO-111169 | No | CPU switch |  | Defect | HW | DMS_PC | CPU
ABL90 FLEX | No action required today, data logs taken for analysis. | No action required today, data logs taken for analysis. | SWO-111151 | No | No action required today, data logs taken for analysis. | No action required today, data logs taken for analysis. | No fault found | - | - | pO2
ABL90 | sO2 QC low out of insert range on A7740. | Replaced and aligned hemolyzer unit. | D0EDAAB448052603 | No | sO2 QC low out of insert range on A7740. | Replaced and aligned hemolyzer unit. | Defect | HW | Oxi module | Hemolyzer
ABL90 FLEX PLUS |  | B | D0EDC913480CEA03 | No |  |  | Defect | SW | Roma Loader | Off/On
TCM4 | 更换电极修复仪器---jack.chu | B | SWO-110972 | No | Replacement electrode repair instrument---jack.chu
TCM4 | 更换新电极修复仪器---jack.chu | 更换新电极修复仪器---jack.chu | SWO-111210 | No | Replace the new electrode repair instrument---jack.chu | Replace the new electrode repair instrument---jack.chu
ABL90 FLEX | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-111214 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL90 FLEX |  | B | SWO-110952 | No |  |  | Defect | HW | DMS_PC | CPU
AQT90 FLEX | 更换新条码扫描器修复仪器---jack.chu | 更换新条码扫描器修复仪器---jack.chu | SWO-110857 | No | Replace the new barcode scanner repair instrument---jack.chu | Replace the new barcode scanner repair instrument---jack.chu | Defect | HW | Analyzer Control Module | Barcode reader, display
AQT90 FLEX |  | B | SWO-110962 | No |  |  | Assisted Replacement | Consumables | Test Kit | NT-PRO BNP, Cal
ABL800 BASIC | replace the ref membrane. replace the Na membrane. calibration pass. qc pass | replace the ref membrane. replace the Na membrane. calibration pass. qc pass | SWO-111813 | No | replace the ref membrane. replace the Na membrane. calibration pass. qc pass | replace the ref membrane. replace the Na membrane. calibration pass. qc pass | Assisted Replacement | Electrode | Membrane | Na+
ABL800 BASIC | Clean auto check module but problem is not resolving, need to replace auto check scanner | B | SWO-112145 | No | Clean auto check module but problem is not resolving, need to replace auto check scanner |  | Revisit required | HW | AutoCheck Module | Barcode reader
ABL825 FLEX | Messkammer gereinigt, Konnektoren gesäubert und pH Elektrode gereinigt und Dichtungsring ersetzt | Messkammer gereinigt, Konnektoren gesäubert und pH Elektrode gereinigt und Dichtungsring ersetzt | SWO-111128 | No | Measuring chamber cleaned, connectors cleaned and pH electrode cleaned and sealing ring replaced | Measuring chamber cleaned, connectors cleaned and pH electrode cleaned and sealing ring replaced | Cleaned | Fluidic pathway | Fluidic pathway | Electrolytes
ABL825 FLEX | Messkammer gereinigt, Konnektoren und Elektroden überprüft und gereinigt, pH Dichtungsring ersetzt, ABL mehrfach kalibriert, keine Driften mehr feststellbar | Messkammer gereinigt, Konnektoren und Elektroden überprüft und gereinigt, pH Dichtungsring ersetzt, ABL mehrfach kalibriert, keine Driften mehr feststellbar | SWO-111129 | No | Measuring chamber cleaned, connectors and electrodes checked and cleaned, pH sealing ring replaced, ABL calibrated several times, no more drifts can be detected | Measuring chamber cleaned, connectors and electrodes checked and cleaned, pH sealing ring replaced, ABL calibrated several times, no more drifts can be detected | Cleaned | Fluidic pathway | Fluidic pathway | pH
ABL800 BASIC | feed back for touch screen issue. machine working fine after calibration. machine under observation. | feed back for touch screen issue. machine working fine after calibration. machine under observation. | SWO-111811 | Yes | feed back for touch screen issue. machine working fine after calibration. machine under observation. | feed back for touch screen issue. machine working fine after calibration. machine under observation. | Re classified RMED | Courtesy
ABL827 FLEX | Tried cleaning electrode and different lot of membranes | Replaced lactate electrode | D0EDA90E483D2F9C | No | Tried cleaning electrode and different lot of membranes | Replaced lactate electrode | Assisted Replacement | Electrode | Electrode | Lac | Sensitivity
ABL825 FLEX | Level 3 and one running intermittently high then low, very unstable looked through data started in May after customer received new membranes customer also replace electrode | replaced electrode with new one under warranty performed a zero current cal; I am going to let the analyzer stablilize for 24 hours and check on the machine in the morning | D0ED8DF6480150FE | No | Level 3 and one running intermittently high then low; very unstable looked through data started in May after customer received new membranes customer also replace electrode | replaced electrode with new one under warranty performed a zero current cal; I am going to let the analyzer stablilize for 24 hours and check on the machine in the morning | Assisted Replacement | Electrode | Electrode | Lac | -
ABL825 FLEX | looked in data logs and got error code 735 and 736. Looked in Autocheck and found broken fan. | Replaced auto check fan. | D0ED99A04844B44B | No | looked in data logs and got error code 735 and 736. Looked in Autocheck and found broken fan. | Replaced auto check fan. | Defect | HW | Auto Inlet module | FAN broken
ABL800 BASIC | 更换新Lac电极膜后解决 | 更换新Lac电极膜后解决 | SWO-110883 | No | Solution after replacing the new Lac electrode film | Solution after replacing the new Lac electrode film | Assisted Replacement | Electrode | Membrane | Lac | -
ABL800 BASIC | Clean chamber ,Run cleaning and 2-point calibration given and resolve the problem.Run QC which is also OK | B | SWO-112545 | No | Clean chamber ,Run cleaning and 2-point calibration given and resolve the problem. Run QC which is also OK |  | Cleaned | Electrode | Electrode | Lac | Sensitivity
ABL837 FLEX | CPU PCB udskiftet, slutafprøvet. | CPU PCB udskiftet, slutafprøvet. | SWO-111088 | No | CPU PCB udskiftet, slutafprøvet. | CPU PCB udskiftet, slutafprøvet. | Defect | HW | DMS_PC | CPU
AQT90 FLEX | Cleaned storage door and front cover, now door is operating smoothly | Cleaned storage door and front cover, now door is operating smoothly | SWO-111471 | No | Cleaned storage door and front cover, now door is operating smoothly | Cleaned storage door and front cover, now door is operating smoothly | Cleaned | HW | Cover and Chassis | Storage Door
ABL80 FLEX BASIC |  | B | SWO-111311 | No
ABL80 FLEX BASIC | 连续更换两张测试卡后修复 | 连续更换两张测试卡后修复 | SWO-111260 | No | Repair after replacing two test cards in a row | Repair after replacing two test cards in a row
ABL80 FLEX BASIC |  | B | SWO-111290 | No
ABL90 FLEX | Analysator demontiert - Monitorkontakte überprüft und neu fixiert (keine Fehlerbeseitigung) - Erneuerung Display Unit - Zeiteinstellung korrigiert - Funktionsüberprüfung etc. ohne Beanstandung! | Analysator demontiert - Monitorkontakte überprüft und neu fixiert (keine Fehlerbeseitigung) - Erneuerung Display Unit - Zeiteinstellung korrigiert - Funktionsüberprüfung etc. ohne Beanstandung! | SWO-111316 | No | Analyzer disassembled - Monitor contacts checked and re-fixed (no troubleshooting) - Renewal Display Unit - Time setting corrected - Function check etc. without complaint! | Analyzer disassembled - Monitor contacts checked and re-fixed (no troubleshooting) - Renewal Display Unit - Time setting corrected - Function check etc. without complaint! | Defect | HW | DMS_PC | Display
ABL90 | found that | B | D0EEEBCABE2B7284 | Yes | found that |  | Re classified RMED | Insufficient text
ABL90 FLEX PLUS | Replaced the blocked hemolyzer. | B | SWO-113704 | No | Replaced the blocked hemolyzer. |  | Defect | HW | Oxi module | Hemolyzer
ABL90 FLEX |  | B | SWO-111263 | No |  |  | Revisit required | - | - | Pump Rotor
ABL827 FLEX | As the zero current calibration was recently performed due to zero current errors the El Met PCB was replaced and the analyser put back into use. The Ca membrane was replaced because the sensitivity was too low and the Lac was replaced because the sensitivity had dropped close to the lower limit. Once this analyser was back up and running the loan analyser was run down and removed from site. | As the zero current calibration was recently performed due to zero current errors the El Met PCB was replaced and the analyser put back into use. The Ca membrane was replaced because the sensitivity was too low and the Lac was replaced because the sensitivity had dropped close to the lower limit. Once this analyser was back up and running the loan analyser was run down and removed from site. | SWO-111438 | No | As the zero current calibration was recently performed due to zero current errors the El Met PCB was replaced and the analyser put back into use. The Ca membrane was replaced because the sensitivity was too low and the Lac was replaced because the sensitivity had dropped close to the lower limit. Once this analyser was back up and running the loan analyser was run down and removed from site. | As the zero current calibration was recently performed due to zero current errors the El Met PCB was replaced and the analyser put back into use. The Ca membrane was replaced because the sensitivity was too low and the Lac was replaced because the sensitivity had dropped close to the lower limit. Once this analyser was back up and running the loan analyser was run down and removed from site. | Assisted Replacement | Electrode | Electrode | Lac | Sensitivity
ABL827 FLEX | Replaced the Gas 2 flow regulator with one removed from a decommissioned analyser (it had only been installed on that analyser for 1 month) and now it is showing 4.9V for a full bottle on Gas 2. Checked the Gas 1 flow regulator with the same bottle and this was reading 4.5V so that is ok. Protein removal was pending so it was run. | Replaced the Gas 2 flow regulator with one removed from a decommissioned analyser (it had only been installed on that analyser for 1 month) and now it is showing 4.9V for a full bottle on Gas 2. Checked the Gas 1 flow regulator with the same bottle and this was reading 4.5V so that is ok. Protein removal was pending so it was run. | SWO-111439 | No | Replaced the Gas 2 flow regulator with one removed from a decommissioned analyser (it had only been installed on that analyser for 1 month) and now it is showing 4.9V for a full bottle on Gas 2. Checked the Gas 1 flow regulator with the same bottle and this was reading 4.5V so that is ok. Protein removal was pending so it was run. | Replaced the Gas 2 flow regulator with one removed from a decommissioned analyser (it had only been installed on that analyser for 1 month) and now it is showing 4.9V for a full bottle on Gas 2. Checked the Gas 1 flow regulator with the same bottle and this was reading 4.5V so that is ok. Protein removal was pending so it was run. | Defect | HW | Gas module | Flow regulator 2
ABL825 FLEX | Inspected tray for damage of something bumping it and none was found. | Replaced sampler tray and performed calibrations. | D0EEC2DEBE2BCEE7 | No | Inspected tray for damage of something bumping it and none was found. | Replaced sampler tray and performed calibrations. | Defect | HW | FlexQ Module | Tray
ABL825 FLEX | Visual and operational inspection found a bent probe. PCO2 trending low on all three levels of QC. | Replaced the probe and inlet gasket. flushed the waste nipples and waste lines. Replaced the REF membrane. Performed several rinse cycles and cleaning cycles to ensure no leaks. | D0EEFD65BE397C75 | No | Visual and operational inspection found a bent probe. PCO2 trending low on all three levels of QC. | Replaced the probe and inlet gasket. flushed the waste nipples and waste lines. Replaced the REF membrane. Performed several rinse cycles and cleaning cycles to ensure no leaks. | Defect | HW | Inlet module | Probe bend | Customer called with analyzer leaking from the front. // Visual and operational inspection found a bent probe. PCO2 trending low on all three levels of QC.
ABL825 FLEX | Visual and operational inspection found a bent probe. | Replaced the probe and inlet gasket. Preformed several flushes, cleaning, calibrations and qc with no leaks present. | D0EEF9FEBE267C70 | No | Visual and operational inspection found a bent probe. | Replaced the probe and inlet gasket. Preformed several flushes, cleaning, calibrations and qc with no leaks present. | Defect | HW | Inlet module | Probe bend | Customer called with analyzer leaking from the front. // Visual and operational inspection found a bent probe.
ABL800 BASIC | CHANGE PRINTERHEAD NOW MACHINE IS OK. | CHANGE PRINTERHEAD NOW MACHINE IS OK. | SWO-112000 | No | CHANGE PRINTERHEAD NOW MACHINE IS OK. | CHANGE PRINTERHEAD NOW MACHINE IS OK. | Defect | HW | Printer module | Printer Head
ABL815 FLEX |  | B | D0F2E8AE2056205A | No |  |  | Defect | HW | Fluidic pathway | Tubing
AQT90 FLEX | Überprüfung und Reinigung von Waschblock, Nadel und Filter da mit Blutrückständen verschmutzt Cup Wascheinheit und Process Wheel sauber Spülung über Pumpe 2 mit HypoChlorid Startphase i.O. Durchführung Reinigung und Überprüfung vom Reinigungsintervall => eingestellt bei 100 Messungen Überprüfung QC Material => QC Röhrchen waren zum Teil schon mal ungekühlt => QC Probenmaterial im Röhrchen verlaufen und wieder eingefroren (TnT Level 1 Lot 11463, TnT Level 2 Lot 11464, CRP Level 1 Lot 11537 und CRP Level 2 Lot 11538) Durchführung Kunden LQC mit Lot 11463 => weiterhin fehlerhaft Mitgebrachtes LQC Lot 11463 gemessen => bestanden => LQC Material beim Kunden gelassen Mit Kunden über Temperaturproblematik gesprochen => Info auch an VBM | Überprüfung und Reinigung von Waschblock, Nadel und Filter da mit Blutrückständen verschmutzt Cup Wascheinheit und Process Wheel sauber Spülung über Pumpe 2 mit HypoChlorid Startphase i.O. Durchführung Reinigung und Überprüfung vom Reinigungsintervall => eingestellt bei 100 Messungen Überprüfung QC Material => QC Röhrchen waren zum Teil schon mal ungekühlt => QC Probenmaterial im Röhrchen verlaufen und wieder eingefroren (TnT Level 1 Lot 11463, TnT Level 2 Lot 11464, CRP Level 1 Lot 11537 und CRP Level 2 Lot 11538) Durchführung Kunden LQC mit Lot 11463 => weiterhin fehlerhaft Mitgebrachtes LQC Lot 11463 gemessen => bestanden => LQC Material beim Kunden gelassen Mit Kunden über Temperaturproblematik gesprochen => Info auch an VBM | SWO-111483 | No | Checking and cleaning of washing block, needle and filter since contaminated with blood residue cup washing unit and Process Wheel clean rinsing via pump 2 with hypochloride start phase i.O. Carrying out cleaning and verification of the cleaning interval => set at 100 Measurements Review QC Material => QC Tubes were sometimes uncooled => QC Sample material in the tube and frozen again (TnT Level 1 Lot 11463, TnT Level 2 Lot 11464, CRP Level 1 Lot 11537 and CRP Level 2 Lot 11538) Performing Customer LQC with Lot 11463 => wei terhin faulty LQC Lot 11463 measured => passed => LQC material left at the customer With customer talked about temperature problems => Info also to VBM | Checking and cleaning of washing block, needle and filter since contaminated with blood residue cup washing unit and Process Wheel clean rinsing via pump 2 with hypochloride start phase i.O. Carrying out cleaning and verification of the cleaning interval => set at 100 Measurements Review QC Material => QC Tubes were sometimes uncooled => QC Sample material in the tube and frozen again (TnT Level 1 Lot 11463, TnT Level 2 Lot 11464, CRP Level 1 Lot 11537 and CRP Level 2 Lot 11538) Performing Customer LQC with Lot 11463 => wei terhin faulty LQC Lot 11463 measured => passed => LQC material left at the customer With customer talked about temperature problems => Info also to VBM | Cleaned | HW | Process Module | Cup wash unit
ABL80 FLEX CO-OX | Fehlersuche, Gerät aus, div. Schlauchmaterial gew., Gerät ein, LP (alt: 222497)u. SC (alt: 221624) erneuert, Initialisierung abgewartet -> pO2-Fehler QC 1, div. Kal., Kal. i.O., QC weiter Fehler, erneut LP (alt: 222324, neu: 223026) gew., erneute Initialisierung, Kal. + QC i.O. | Fehlersuche; Gerät aus; div. Schlauchmaterial gew.; Gerät ein; LP (alt: 222497)u. SC (alt: 221624) erneuert; Initialisierung abgewartet -> pO2-Fehler QC 1; div. Kal.; Kal. i.O.; QC weiter Fehler; erneut LP (alt: 222324, neu: 223026) gew.; erneute Initialisierung; Kal. + QC i.O. | SWO-111541 | No | Troubleshooting; device off; div. hose material. device; LP (old: 222497)and SC (old: 221624) renewed; Initialization awaited -> pO2 error QC 1; div. Kal.; Kal. i.O.; QC continues errors; AGAIN LP (old: 222324, new: 223026); re-initialization; Kal. + QC in the first place | Troubleshooting; device off; div. hose material. device; LP (old: 222497)and SC (old: 221624) renewed; Initialization awaited -> pO2 error QC 1; div. Kal.; Kal. i.O.; QC continues errors; AGAIN LP (old: 222324, new: 223026); re-initialization; Kal. + QC in the first place
ABL80 FLEX BASIC | 更换测试卡修复仪器 | 更换测试卡修复仪器 | SWO-111614 | No | Replace the test card repair instrument | Replace the test card repair instrument
ABL90 |  | B | D0FE5D47463C1B38 | No |  |  | Defect | HW | Basic Unit | Sensor Interface PCB
ABL90 | To get unit to reconnect to customers system analyzer must be moved around. | Replaced wireless adapter. | D0F04065341A6036 | No | To get unit to reconnect to customers system analyzer must be moved around. | Replaced wireless adapter. | Defect | HW | Wireless adaptor | Wireless adaptor
ABL90 FLEX PLUS |  | B | D0F5087C0C4550F8 | Yes |  |  | Re classified RMED | PER
ABL90 FLEX PLUS |  | B | D0F02DBB344E5E39 | No |  |  | Replaced | HW | Auto Inlet module | Probe bend
ABL90 FLEX | 更换配件636-568和910-342修复仪器---jack.chu | 更换配件636-568和910-342修复仪器---jack.chu | SWO-111314 | No | Replacement accessories 636-568 and 910-342 repair instruments ---jack.chu | Replacement accessories 636-568 and 910-342 repair instruments ---jack.chu | Defect | HW | DMS_PC | User interface module
AQT90 FLEX | 更换636-550修复仪器---jack.chu | 更换636-550修复仪器---jack.chu | SWO-111305 | No | Replace 636-550 repair instrument---jack.chu | Replace 636-550 repair instrument---jack.chu | Defect | HW | Process Module | Process Unit sensor | 5257
ABL835 FLEX | 回路消毒作業実施。 全バルブ分解点検、外観消毒。 全メンブラン、全溶液回収 | B | SWO-112636 | No | Circuit disinfection work implemented. All valves disassembly inspection, appearance disinfection. All membranes, all solution recovery |  | Assisted Replacement | Electrode | Membrane | Lac | Sensitivity
ABL800 BASIC | 更换902-833一个正常。---jack.chu | 更换902-833一个正常。---jack.chu | SWO-111192 | No | Replace 902-833 one normal. ---jack.chu | Replace 902-833 one normal. ---jack.chu | Defect | HW | Oxi module | Oxi module
ABL835 FLEX |  | B | D0F05CC134166491 | Yes |  |  | Re classified RMED
ABL825 FLEX |  | B | SWO-111511 | No |  |  | Defect | HW | DMS_PC | Touch screen
ABL825 FLEX | Found hardware error codes pointed to Fuse Failure. Machine supposed to be shipped at software version 6.16. It was shipped at 6.15 | Ordered new fuse for overnight delivery. Replaced fuse. Performed software upgrade to version 6.16. Performed startup calibrations. | D0EFFE1A34130E39 | No | Found hardware error codes pointed to Fuse Failure. Machine supposed to be shipped at software version 6.16. It was shipped at 6.15 | Ordered new fuse for overnight delivery. Replaced fuse. Performed software upgrade to version 6.16. Performed startup calibrations. | Defect | HW | Basic Unit | Analyzer Control PCB | Fuse
ABL825 FLEX |  | B | D0FA52BBE4181764 | No |  |  | Unknown | HW | Inlet module | Friction detector unit | LEH asked RPAC
ABL825 FLEX | Found faulty fuse on AU board. | Replaced Faulty fuse. Restarted analyzer. Performed necessary adjustments and calibrations. | D0F008CB34310E42 | No | Found faulty fuse on AU board. | Replaced Faulty fuse. Restarted analyzer. Performed necessary adjustments and calibrations. | Defect | HW | Basic Unit | Analyzer Control PCB | Fuse
ABL825 FLEX | Lac wurde bereits membraniert- beobachten, Hakenventil vor Oximetrie schließt nicht sauber - ausgebaut und gängig gemacht, eingebaut und mehrere Tests durchgeführt, tHb, Kal und QC, nach einigen Messungen erneut der gleiche Fehler, Hakenventil ausgbaut und ersetzt, erneut mehrere Messungen durchgeführt - fehlerfrei | B | SWO-111771 | No | Lac has already been diaphragm-observed, hook valve before oximetry does not close clean - removed and commonly made, installed and carried out several tests, tHb, Kal and QC, after some measurements again the same error, hook valve expanded and replaced, again several measurements taken - error-free |  | Defect | HW | Oxi module | Pinch Valve
AQURE | The nguestx.ocx ActiveX component needed to be installed. We got a member of IT to log on with admin rights and to open the Quick Access window which allowed nguestx.ocx to be installed. Logged back on to the PC with the normal log in and checked that the Quick Access window still worked, which it did. | The nguestx.ocx ActiveX component needed to be installed. We got a member of IT to log on with admin rights and to open the Quick Access window which allowed nguestx.ocx to be installed. Logged back on to the PC with the normal log in and checked that the Quick Access window still worked, which it did. | SWO-111805 | Yes | The nguestx.ocx ActiveX component needed to be installed. We got a member of IT to log on with admin rights and to open the Quick Access window which allowed nguestx.ocx to be installed. Logged back on to the PC with the normal log in and checked that the Quick Access window still worked, which it did. | The nguestx.ocx ActiveX component needed to be installed. We got a member of IT to log on with admin rights and to open the Quick Access window which allowed nguestx.ocx to be installed. Logged back on to the PC with the normal log in and checked that the Quick Access window still worked, which it did.
ABL80 FLEX BASIC | 更换903-334,903-099各一个正常---jack.chu | 更换903-334,903-099各一个正常---jack.chu | SWO-111267 | No | Replace 903-334,903-099 each a normal ---jack.chu | Replace 903-334,903-099 each a normal ---jack.chu
ABL80 FLEX BASIC | done re hydration of Sc and kept in fridge. Then reinstalled SC. Instrument is working fine. | done re hydration of Sc and kept in fridge. Then reinstalled SC. Instrument is working fine. | SWO-112081 | No | done re hydration of Sc and kept in fridge. Then reinstalled SC. Instrument is working fine. | done re hydration of Sc and kept in fridge. Then reinstalled SC. Instrument is working fine.
ABL80 FLEX | Calibrations and QC all correct, advise to run sample again, show message on printout, Setup QC and Access codes according to customer needs | B | SWO-113017 | No | Calibrations and QC all correct, advise to run sample again, show message on printout, Setup QC and Access codes according to customer needs
ABL90 FLEX | Lösungspack-ID-PCB erneuert. | Lösungspack-ID-PCB erneuert. | SWO-111910 | No | Lösungspack-ID-PCB erneuert. | Lösungspack-ID-PCB erneuert. | Defect | HW | Solution Pack PCB Id | Solution Pack PCB Id
ABL90 FLEX | Fließwege gespült, Einlaufverfahren Neue Bedienungsanleitung abgegeben | Fließwege gespült, Einlaufverfahren Neue Bedienungsanleitung abgegeben | SWO-111992 | No | Flow paths rinsed, inlet procedure New operating instructions issued | Flow paths rinsed, inlet procedure New operating instructions issued | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway | 1342
ABL90 FLEX PLUS | Fehlersuche, Gerät aus, CPU gew., Gerät ein, Startphase abgewartet, Hr. Hanke der IT neue Mac-Adresse mitgeteilt, Kal. -> Fehler SC, vorhandene SC neu eingelesen, Initialisierung abgewartet -> weiter Fehler, neue SC initialisiert, Kal. i.O., QC tHb-Fehler, tHb-Kal. durchgef., Kal. + QC i.O. | Fehlersuche; Gerät aus; CPU gew.; Gerät ein; Startphase abgewartet; Hr. Hanke der IT neue Mac-Adresse mitgeteilt; Kal. -> Fehler SC; vorhandene SC neu eingelesen; Initialisierung abgewartet -> weiter Fehler; neue SC initialisiert; Kal. i.O.; QC tHb-Fehler; tHb-Kal. durchgef.; Kal. + QC i.O. | SWO-111958 | No | Troubleshooting; device off; CPU. device; start-up phase awaited; Mr. Hanke informed the IT new Mac address; Kal. -> Error SC; re-read existing SC; Initialization awaited -> further error; new SC initialized; Kal. i.O.; QC tHb error; tHb-Kal. through it; Kal. + QC in the first place | Troubleshooting; device off; CPU. device; start-up phase awaited; Mr. Hanke informed the IT new Mac address; Kal. -> Error SC; re-read existing SC; Initialization awaited -> further error; new SC initialized; Kal. i.O.; QC tHb error; tHb-Kal. through it; Kal. + QC in the first place | Defect | HW | DMS_PC | CPU
ABL90 FLEX | 更换新测试卡修复仪器 | 更换新测试卡修复仪器 | SWO-111918 | No | Replace the new test card repair instrument | Replace the new test card repair instrument
ABL90 FLEX | 更换测试卡正常。 | 更换测试卡正常。 | SWO-111947 | No | The replacement test card is OK. | The replacement test card is OK. | Cleaned | Fluidic pathway | Fluidic pathway | SC
AQT90 FLEX |  | B | SWO-111827 | No |  |  | Defect | HW | Cover and Chassis | Storage Door
ABL837 FLEX | Cleaned the insturment and repalced the internal Barcode reader restarted the insturment found barcode reader was functional. | Cleaned the insturment and repalced the internal Barcode reader restarted the insturment found barcode reader was functional. | SWO-112114 | No | Cleaned the insturment and repalced the internal Barcode reader restarted the insturment found barcode reader was functional. | Cleaned the insturment and repalced the internal Barcode reader restarted the insturment found barcode reader was functional. | Defect | HW | Basic Unit | Barcode reader
ABL825 FLEX | Fehlersuche Wartung Ref. elektrode,( Kunde. hat Elekt. bereits membr.) Meßkammer gereinigt, Schlauch zw.Modulen erneuert nochmal neu Ref. membraniert | Fehlersuche Wartung Ref. elektrode,( Kunde. hat Elekt. bereits membr.) Meßkammer gereinigt, Schlauch zw.Modulen erneuert nochmal neu Ref. membraniert | SWO-112035 | No | Troubleshooting Maintenance Ref. electrode,( customer. has elekt. already membr.) Measuring chamber cleaned, hose between Modules renews new Ref. membraneed | Troubleshooting Maintenance Ref. electrode,( customer. has elekt. already membr.) Measuring chamber cleaned, hose between Modules renews new Ref. membraneed | Assisted Replacement | Electrode | Electrode | Reference
ABL835 FLEX | . pH sealing ring replaced. A new membrane valve was fitted to the air valve in the pH BG module. Level 4 QC was rescanned into the analyser for software V6.16 and the ranges are now correct. | . pH sealing ring replaced. A new membrane valve was fitted to the air valve in the pH BG module. Level 4 QC was rescanned into the analyser for software V6.16 and the ranges are now correct. | SWO-111882 | No | . pH sealing ring replaced. A new membrane valve was fitted to the air valve in the pH BG module. Level 4 QC was rescanned into the analyser for software V6.16 and the ranges are now correct. | . pH sealing ring replaced. A new membrane valve was fitted to the air valve in the pH BG module. Level 4 QC was rescanned into the analyser for software V6.16 and the ranges are now correct. | Setup | QC Ranges | QC Ranges | QC Ranges | 6.16
ABL815 FLEX | While inspectig all 7 new analyzers, I spotted a leak. Error code 541. I tried tubing replacements. | Replaced pump rotors (pH BG and Waste). Replaced solutions pump tubing. Ran pump and liquid sensor calibrations. | D0F18729AA04109B | No | While inspectig all 7 new analyzers, I spotted a leak. Error code 541. I tried tubing replacements. | Replaced pump rotors (pH BG and Waste). Replaced solutions pump tubing. Ran pump and liquid sensor calibrations. | Defect | HW | Inlet module | Leaking
ABL800 BASIC | INLET PROBE NEED TO BE REPLACE WITH NEW ONE | B | SWO-112115 | No | INLET PROBE NEED TO BE REPLACE WITH NEW ONE |  | Defect | HW | Inlet module | Probe bend | HEADLINE: INLET PROBE PROBLEM // COMPLAINT : MACHINE SHOWING RINSE ERROR. // SYMPTOM : RINSE ERROR // INVESTIGATION: INLET PROBE NEED TO BE REPLACE WITH NEW ONE
ABL825 FLEX | Einlassmodul leihweise ausgetauscht, div. Schläuche und O-Ringe erneuert, sw installiert, Gerät mehrfach kalibriert und mit QC überprüft, OK. altes Inletmodul: 902-783R2404N007 Neues: 902-783R0396N007 | Einlassmodul leihweise ausgetauscht, div. Schläuche und O-Ringe erneuert, sw installiert, Gerät mehrfach kalibriert und mit QC überprüft, OK. altes Inletmodul: 902-783R2404N007 Neues: 902-783R0396N007 | SWO-112061 | No | Inlet module replaced on loan, various hoses and O-rings renewed, sw installed, device calibrated several times and checked with QC, OK. old inlet module: 902-783R2404N007 New: 902-783R0396N007 | Inlet module replaced on loan, various hoses and O-rings renewed, sw installed, device calibrated several times and checked with QC, OK. old inlet module: 902-783R2404N007 New: 902-783R0396N007 | Defect | HW | Inlet module | Leaking
ABL827 FLEX | VISUAL INSPECTION SHOWS THE WASTE LINE BEHIND THE WASTE BOTTLE WAS NOT CONNECTED | RECONNECTED THE WASTE LINE AND TRAINED THE CUSTOMER ON PROPER FLUSHING TECHNIQUES FOR CLOTS ON THE ANALYZER.PERFORMED CALIBRATIONS | D0F1424AAA5171A3 | No | VISUAL INSPECTION SHOWS THE WASTE LINE BEHIND THE WASTE BOTTLE WAS NOT CONNECTED | RECONNECTED THE WASTE LINE AND TRAINED THE CUSTOMER ON PROPER FLUSHING TECHNIQUES FOR CLOTS ON THE ANALYZER. PERFORMED CALIBRATIONS | Defect | HW | Basic Unit | Waste tube
AQT90 FLEX | Software opnieuw geïnstalleerd. | Software opnieuw geïnstalleerd. | SWO-111990 | No | Software reinstalled. | Software reinstalled. | Defect | SW | Analyzer Control Module | Re installed SW
AQT90 FLEX | Performed quality control with multicheck for TnI and Nt-proBNP parameters, both parameters were reported within QC range | Performed quality control with multicheck for TnI and Nt-proBNP parameters, both parameters were reported within QC range | SWO-111945 | No | Performed quality control with multicheck for TnI and Nt-proBNP parameters, both parameters were reported within QC range | Performed quality control with multicheck for TnI and Nt-proBNP parameters, both parameters were reported within QC range | User | Consumables | LQC Kit | NT-proBNP
ABL80 FLEX BASIC | シグナルデータボード交換 | B | SWO-112117 | No | Signal data board exchange
AQT90 FLEX |  | B | SWO-112105 | No |  |  | Defect | HW | Analyzer Control Module | Barcode reader, display
AQT90 FLEX | 更换试剂包修复仪器 | 更换试剂包修复仪器 | SWO-113619 | No | Replace reagent kit repair instruments | Replace reagent kit repair instruments
ABL80 FLEX BASIC | 更换测试卡修复仪器 | 更换测试卡修复仪器 | SWO-112316 | No | Replace the test card repair instrument | Replace the test card repair instrument
ABL90 FLEX | Netzteil erneuert Kalibration und QC durchgeführt | Netzteil erneuert Kalibration und QC durchgeführt | SWO-112167 | No | Power supply renewed calibration and QC performed | Power supply renewed calibration and QC performed | Defect | HW | Power Supply | Power Supply
ABL90 FLEX |  | B | SWO-112124 | No |  |  | Defect | HW | Oxi module | Hemolyzer | 89
AQT90 FLEX |  | B | SWO-112136 | No |  |  | Defect | HW | Process Module | Optical Unit | 135
ABL80 CO-OX |  | B | D0FEE4EEBC5FFBA9 | No
ABL80 FLEX CO-OX | Schlauchverbindungen und Ventile geprüft -> i.O. SC ersetzt, danach wurde QC bestanden | Schlauchverbindungen und Ventile geprüft -> i.O. SC ersetzt, danach wurde QC bestanden | SWO-112511 | No | Hose connections and valves tested -> i.O. SC replaced, after which QC was passed | Hose connections and valves tested -> i.O. SC replaced, after which QC was passed
ABL80 FLEX BASIC | 更换测试卡正常。 | 更换测试卡正常。 | SWO-112567 | No | The replacement test card is OK. | The replacement test card is OK.
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-112337 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL80 FLEX BASIC |  | B | SWO-112267 | No
ABL80 FLEX BASIC | 零点校准，更换测试卡 | 零点校准，更换测试卡 | SWO-112919 | No | Zero-point calibration, replacement of test card | Zero-point calibration, replacement of test card
ABL80 FLEX BASIC | 更换测试卡正常。 | 更换测试卡正常。 | SWO-112561 | No | The replacement test card is OK. | The replacement test card is OK.
ABL80 FLEX BASIC | replace the sc80. replace sp 80 . calibrate system. calibration pass.qc pass. per in process. | replace the sc80. replace sp 80 . calibrate system. calibration pass.qc pass. per in process. | SWO-113835 | No | replace the sc80. replace sp 80 . calibrate system. calibration pass.qc pass. per in process. | replace the sc80. replace sp 80 . calibrate system. calibration pass.qc pass. per in process.
ABL80 FLEX CO-OX | LP wurde getauscht und die Einstellungen am Patientenausdruck geändert. | LP wurde getauscht und die Einstellungen am Patientenausdruck geändert. | SWO-112618 | No | LP was changed and the settings on the patient printout were changed. | LP was changed and the settings on the patient printout were changed.
ABL80 FLEX BASIC |  | B | SWO-112281 | No
ABL90 FLEX | Einlassnadel ist zu Einlassdichtung nicht richtig positioniert, korrigiert, Selbiges LP installiert, alle Serviceprogramme überprüft, i.O., Kal, QC | Einlassnadel ist zu Einlassdichtung nicht richtig positioniert, korrigiert, Selbiges LP installiert, alle Serviceprogramme überprüft, i.O., Kal, QC | SWO-112544 | No | Inlet pin is not positioned correctly to inlet seal, corrected, same LP installed, all service programs checked, i.O., Kal, QC | Inlet pin is not positioned correctly to inlet seal, corrected, same LP installed, all service programs checked, i.O., Kal, QC | User error | HW | Inlet module | Inlet gasket
ABL90 FLEX | SC90 erneuert. | SC90 erneuert. | SWO-112361 | No | SC90. | SC90. | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
ABL90 |  | B | D0F50A140C3CA533 | No |  |  | Cleaned | Fluidic pathway | Fluidic pathway | Inlet gasket
ABL90 | David Topaloff had tried a display and CPU. He received the wrong CPU. I got a CPU from Kevin Elliott. He put it in an analyzer and thought it solved the problem but didn't. He put it on the SWO (B9YW3R3WG) so it had been claimed. He said it was new but it had already been claimed so I did not claim it again. The CPU solved the issue. The analyzer had been sitting for about 20 hours and the SC had to be replaced. I replaced the SC and the customer said they would get it up. I didn't need to sit around and wait for it. | Replaced the CPU and loaded Sybase 16. | D0F53CDC0C0E750C | No | David Topaloff had tried a display and CPU. He received the wrong CPU. I got a CPU from Kevin Elliott. He put it in an analyzer and thought it solved the problem but didn't. He put it on the SWO (B9YW3R3WG) so it had been claimed. He said it was new but it had already been claimed so I did not claim it again. The CPU solved the issue. The analyzer had been sitting for about 20 hours and the SC had to be replaced. I replaced the SC and the customer said they would get it up. I didn't need to sit around and wait for it. | Replaced the CPU and loaded Sybase 16. | Defect | HW | DMS_PC | CPU
ABL90 FLEX | Touchscreen reagiert nicht, erneuert | Touchscreen reagiert nicht, erneuert | SWO-112385 | No | Touchscreen does not react, renews | Touchscreen does not react, renews | Defect | HW | DMS_PC | Touch Screen
ABL90 | Inspected inlet for clots and none was found. | Customer performed flush | D0F4F5920C16EDC0 | No | Inspected inlet for clots and none was found. | Customer performed flush | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
ABL90 | Analyzer stuck in loop. Could resolve in service fluidics. Waste port seemed to be leaking around black waste tubing. Inspection of system is needed. | Replace parts from PM kit. Inspect system while apart. Replaced Inlet Gasket W Connector. Flush out waste port. Restart and run fluidics tests. | D0F538220C489D65 | No | Analyzer stuck in loop. Could resolve in service fluidics. Waste port seemed to be leaking around black waste tubing. Inspection of system is needed. | Replace parts from PM kit. Inspect system while apart. Replaced Inlet Gasket W Connector. Flush out waste port. Restart and run fluidics tests. | Defect | HW | Waste tube | Waste tube
ABL90 FLEX | 更换两张 新测试卡修复仪器 | 更换两张 新测试卡修复仪器 | SWO-112233 | No | Replace two new test card repair instruments | Replace two new test card repair instruments | Assisted Replacement | SC | SC | K+ | 131
ABL90 FLEX |  | B | SWO-112319 | No |  |  | Defect | HW | SC ID PCB | SC Interface PCB
AQT90 FLEX | 定标完成后，定标没有通过，提示故障代码1031，更换了903-364后修复 by jesson | 定标完成后，定标没有通过，提示故障代码1031，更换了903-364后修复 by jesson | SWO-112099 | No | After the calibration is completed, the bid did not pass, prompt fault code 1031, replaced 903-364 after repair by jesson | After the calibration is completed, the bid did not pass, prompt fault code 1031, replaced 903-364 after repair by jesson | Defect | HW | Process Module | Optical Unit | 131
ABL827 FLEX |  | B | D0F6419F825B54D3 | No |  |  | Defect | HW | AutoCheck Module | FPC Cable
ABL835 FLEX | Borrow cal gas from Gsh. Order inlet needle, straighten in order for them to use. Replace inlet gasket borrowed from GSH. I have ordered a inlet gasket and inlet needle , jobcard 4265 | B | SWO-112713 | Yes | Borrow cal gas from Gsh. Order inlet needle, straighten in order for them to use. Replace inlet gasket borrowed from GSH. I have ordered a inlet gasket and inlet needle , jobcard 4265 |  | Re classified RMED | Logistical upSWO
ABL800 BASIC | check for gas pressure test . change gas 2 .calibrate system. calibration pass. | check for gas pressure test . change gas 2 .calibrate system. calibration pass. | SWO-113832 | No | check for gas pressure test . change gas 2 .calibrate system. calibration pass. | check for gas pressure test . change gas 2 .calibrate system. calibration pass. | Assisted Replacement | Gas Cylinder | Gas Cylinder | Gas 2
ABL827 FLEX | Crea-Elektroden und Crea-Konnektorkontakte gereinigt, Messkammer gereinigt, E8088 (EV01) und E8089 (JT01) eingebaut und E8088 (SK03) und E8089 (SK03) ausgebaut, System und Funktion i.O. | B | SWO-112369 | No | Crea electrodes and Crea connector contacts cleaned, measuring chamber cleaned, E8088 (EV01) and E8089 (JT01) installed and E8088 (SK03) and E8089 (SK03) removed, system and function i.O. |  | Assisted Replacement | Electrode | Electrode | Crea
ABL835 FLEX |  | B | D0F553910C52F907 | Yes |  |  | Re classified RMED | Courtesy
ABL825 FLEX | Checked valves, tubing and pumps. Found pH BG module gas valve not opening properly. | Replaced the pH BG module gas valve. | D0F536280C5A5228 | No | Checked valves, tubing and pumps. Found pH BG module gas valve not opening properly. | Replaced the pH BG module gas valve. | Defect | HW | pH/BG Module | Valve, air
ABL825 FLEX |  | B | D0F58B5A0C280B96 | No |  |  | Cleaned | Fluidic pathway | Inlet module | Inlet gasket
ABL815 FLEX | error code 541 | Replaced broken Probe Holder assy. Replace inlet probe. | D0F57D2C0C4A9A46 | No | error code 541 | Replaced broken Probe Holder assy. Replace inlet probe. | Defect | HW | Inlet module | Probe holder | 541
ABL825 FLEX | Upon arrival, no leaking. Ran two point calbrations. Ran blood samples | Upgraded software to 6.16 version (possible changes to fluid flow). Replaced inlet gasket (previously replaced). | D0F579740C269A3F | No | Upon arrival, no leaking. Ran two point calbrations. Ran blood samples | Upgraded software to 6.16 version (possible changes to fluid flow). Replaced inlet gasket (previously replaced). | Defect | HW | Inlet module | Inlet gasket
ABL810 FLEX | Ortshöhenkorrektur zurückgesetzt. Level 1 Kontrollgrenzen für Software > 6.16 eingegeben. Nach Korrektur exakt Zielwert. Halbjahreswartungskit eingesetzt. | Ortshöhenkorrektur zurückgesetzt. Level 1 Kontrollgrenzen für Software > 6.16 eingegeben. Nach Korrektur exakt Zielwert. Halbjahreswartungskit eingesetzt. | SWO-112603 | No | Place elevation correction reset. Level 1 Control limits for software > 6.16 entered. Exactly target value after correction. half-year maintenance kit. | Place elevation correction reset. Level 1 Control limits for software > 6.16 entered. Exactly target value after correction. half-year maintenance kit. | Setup | SW | Altitude corr. Reset | Altitude corr. Reset
ABL800 BASIC | 6.21号更换电极模，6.27号灵敏度就降到140，更换Lac电极膜修复仪器 by jesson | 6.21号更换电极模，6.27号灵敏度就降到140，更换Lac电极膜修复仪器 by jesson | SWO-112290 | No | 6.21 replacement electrode mold, 6.27 sensitivity is reduced to 140, replace Lac electrode membrane repair instrument by jesson | 6.21 replacement electrode mold, 6.27 sensitivity is reduced to 140, replace Lac electrode membrane repair instrument by jesson | Assisted Replacement | Electrode | Membrane | Lac
AQT90 FLEX | ｵﾌﾟﾃｨｶﾙﾕﾆｯﾄ交換。 | ｵﾌﾟﾃｨｶﾙﾕﾆｯﾄ交換。 | SWO-112562 | No | Optical unit replacement. | Optical unit replacement. | Defect | HW | Process Module | Optical Unit | 129
AQT90 FLEX |  | B | D0F532640C2DC3D0 | No |  |  | Defect | HW | Analyzer Control Module | CF Card
ABL80 FLEX BASIC |  | B | SWO-112584 | No
ABL80 FLEX BASIC |  | B | SWO-112643 | No
ABL80 OSM | Inspected for leaks and clots and none was found. | Replaced the manifold unit. Also will retube unit. | D0F6482E82495D0D | No | Inspected for leaks and clots and none was found. | Replaced the manifold unit. Also will retube unit.
ABL90 FLEX | Dienst RMedDstManagerService gestoppt. Daten auf CF-Karte gesichert, Gerät neu gestartet. Mit QC getestet. Zeitintegrität gewährleistet. | Dienst RMedDstManagerService gestoppt. Daten auf CF-Karte gesichert, Gerät neu gestartet. Mit QC getestet. Zeitintegrität gewährleistet. | SWO-112730 | No | Service RMedDstManagerService stopped. Data backed up to CF card, device restarted. Tested with QC. time integrity. | Service RMedDstManagerService stopped. Data backed up to CF card, device restarted. Tested with QC. time integrity. | Setup | SW | Time setting | Time setting | See previous
ABL90 FLEX | Remplacé les conduites souples de la section fluidique e.o Nettoyage du système fluidique e.o Calibration e.o contrôle de qualité e.o | Remplacé les conduites souples de la section fluidique e.o Nettoyage du système fluidique e.o Calibration e.o contrôle de qualité e.o | SWO-112832 | No | Replaced the flexible pipes of the fluidic section e.o Fluid system cleaning e.o Calibration e.o quality control e.o | Replaced the flexible pipes of the fluidic section e.o Fluid system cleaning e.o Calibration e.o quality control e.o | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway | PM Kit
ABL90 FLEX | QC Tabelle kann Fehler bei Oximetrie nicht verarbeiten und stürzt ab - Wert gelöscht danach i.O., automatische Lieferungen müssen überprüft werden da teils Verfallsdaten schon fast erreicht und Material nicht mehr verwendbar ist, | Zusammen am Telefon die korrekte Uhrzeit eingestellt. Wollte mit ihr noch schauen, ob irgendwelche Daten in den Archiven zu finden sind. Sie hat jetzt dafür keine Zeit, da der Praxisbetrieb läuft. Sie bittet morgen um einen Rückruf zwischen 12.30 und 14.30 Uhr. | SWO-112929 | Yes | QC table can not process errors in oximetry and crashes - value deleted afterwards i.O., automatic deliveries must be checked because partly expiry data is almost reached and material is no longer usable, | Together on the phone set the correct time. Wanted to see with her if any data can be found in the archives. She doesn't have time for that now, because the practice is ongoing. She will ask for a callback between 12.30pm and 2.30pm tomorrow. | Re classified RMED | phone
ABL90 |  | B | D0F6C465822D7CF1 | No |  |  | Adjustment | HW | Oxi module | tHb Cal
ABL90 FLEX | Flüssigkeitssystem mit Reinigungslösung gespült, Serviceprogramme geprüft, tHb-Kal ok, QK ok | Flüssigkeitssystem mit Reinigungslösung gespült, Serviceprogramme geprüft, tHb-Kal ok, QK ok | SWO-112659 | No | Liquid system with cleaning solution rinsed, service programs tested, tHb-Kal ok, QK ok | Liquid system with cleaning solution rinsed, service programs tested, tHb-Kal ok, QK ok | Adjustment | HW | Oxi module | tHb Cal
ABL90 | checked hemolyzer checked valve checked tubing ran multiple samples all failing | replaced hemolyzer ran multiple samples all passed | D0F6AA69824E6DE3 | No | checked hemolyzer checked valve checked tubing ran multiple samples all failing | replaced hemolyzer ran multiple samples all passed | Defect | HW | Oxi module | Hemolyzer | Spectrum mismatch
TCM CombiM |  | B | SWO-112650 | No
ABL90 FLEX | 更换测试卡正常。 | 更换测试卡正常。 | SWO-112901 | No | The replacement test card is OK. | The replacement test card is OK. | Assisted Replacement | SC | SC | K+ | Calibration
ABL90 FLEX | 更换测试卡正常。 | 更换测试卡正常。 | SWO-113186 | No | The replacement test card is OK. | The replacement test card is OK.
TCM5 FLEX |  | B | D0F640D5821A76DF | No
ABL827 FLEX | The 2nd hand Gas 2 flow regulator was replaced with a new one now that I had one. After fitting the new flow regulator I checked around the connection with the gas line and found that there was a hole in the tube used to secure the gas line to the flow regulator and it was leaking from there when the regulator was opened. A new piece of tube was fitted and there was no sign of a leak. | The 2nd hand Gas 2 flow regulator was replaced with a new one now that I had one. After fitting the new flow regulator I checked around the connection with the gas line and found that there was a hole in the tube used to secure the gas line to the flow regulator and it was leaking from there when the regulator was opened. A new piece of tube was fitted and there was no sign of a leak. | SWO-112602 | No | The 2nd hand Gas 2 flow regulator was replaced with a new one now that I had one. After fitting the new flow regulator I checked around the connection with the gas line and found that there was a hole in the tube used to secure the gas line to the flow regulator and it was leaking from there when the regulator was opened. A new piece of tube was fitted and there was no sign of a leak. | The 2nd hand Gas 2 flow regulator was replaced with a new one now that I had one. After fitting the new flow regulator I checked around the connection with the gas line and found that there was a hole in the tube used to secure the gas line to the flow regulator and it was leaking from there when the regulator was opened. A new piece of tube was fitted and there was no sign of a leak. | Defect | HW | Gas module | Tubing
ABL827 FLEX | Reinigung Einlassdichtung, verbogene Einlassnadel gerichtet, korrekte Kal abgewartet, QK ok, Testmessungen an Blut ok | Reinigung Einlassdichtung, verbogene Einlassnadel gerichtet, korrekte Kal abgewartet, QK ok, Testmessungen an Blut ok | SWO-112816 | No | Cleaning inlet seal, bent inlet needle directed, correct scale awaiting, QK ok, test measurements on blood ok | Cleaning inlet seal, bent inlet needle directed, correct scale awaiting, QK ok, test measurements on blood ok | Defect | HW | Inlet module | Probe bend
ABL800 BASIC | replaced the all membranes and done 2point calibration and QC all results found with in the range. | replaced the all membranes and done 2point calibration and QC all results found with in the range. | SWO-113860 | No | replaced the all membranes and done 2point calibration and QC all results found with in the range. | replaced the all membranes and done 2point calibration and QC all results found with in the range. | Assisted Replacement | Electrode | Membrane | Electrolytes
ABL810 FLEX | Ortshöhenkorrektur zurückgesetzt. | Ortshöhenkorrektur zurückgesetzt. | SWO-112605 | No | Place elevation correction reset. | Place elevation correction reset. | Setup | SW | Altitude corr. Reset | Altitude corr. Reset
ABL810 FLEX | Gerät abgeschalten und Neonlampe ausgebaut - ersetzt, erneut Complete Test durchgeführt, jetzt Halogenlampenfehler und Hemolyzer fehlerhaft, abgebrochen - neues Oximodul wird bestellt | Gerät abgeschalten und Neonlampe ausgebaut - ersetzt, erneut Complete Test durchgeführt, jetzt Halogenlampenfehler und Hemolyzer fehlerhaft, abgebrochen - neues Oximodul wird bestellt | SWO-112682 | No | Switch off the device and remove the neon lamp - replaced, completed again, now halogen lamp fault and hemolyzer faulty, aborted - new oximodul is ordered | Switch off the device and remove the neon lamp - replaced, completed again, now halogen lamp fault and hemolyzer faulty, aborted - new oximodul is ordered | Revisit required | HW | Oxi module | Oxi module
ABL825 FLEX | Einlassnadel verbogen und Einlassdichtung verstopft, aus Kundenbestand ersetzt, Pumpenkal und LS Kalibrierung durchgeführt, Kal und QC durchgeführt | Einlassnadel verbogen und Einlassdichtung verstopft, aus Kundenbestand ersetzt, Pumpenkal und LS Kalibrierung durchgeführt, Kal und QC durchgeführt | SWO-112941 | No | Inlet needle bent and inlet seal clogged, replaced from customer stock, pump scale and LS calibration performed, Kal and QC performed | Inlet needle bent and inlet seal clogged, replaced from customer stock, pump scale and LS calibration performed, Kal and QC performed | Defect | HW | Inlet module | Probe bend
TCM/Tosca | Confirmed the reported issues. | Isolated the problem to the customer's tc sensor which has a damaged cable. | D0F67E98821E30A8 | No | Confirmed the reported issues. | Isolated the problem to the customer's tc sensor which has a damaged cable. | Defect | HW | Electrode | Cable
ABL80 FLEX BASIC | 更换测试卡修复仪器 | 更换测试卡修复仪器 | SWO-113652 | No | Replace the test card repair instrument | Replace the test card repair instrument
ABL80 FLEX BASIC | 更换测试卡修复仪器 | 更换测试卡修复仪器 | SWO-112957 | No | Replace the test card repair instrument | Replace the test card repair instrument
ABL80 FLEX BASIC |  | B | SWO-112975 | No
ABL90 FLEX | Anschlussdichtung erneuert. | Anschlussdichtung erneuert. | SWO-113138 | No | Connection seal renewed. | Connection seal renewed. | User error | Cal/QC Trouble shooting | Cal/QC Trouble shooting | Cal/QC Trouble shooting
ABL90 FLEX | pO2- und Oxi Complete Test zur Kalibrierung durchgeführt. Kalibrationen und Qualitätskontrollen in Ordnung. | pO2- und Oxi Complete Test zur Kalibrierung durchgeführt. Kalibrationen und Qualitätskontrollen in Ordnung. | SWO-113005 | No | pO2 and Oxi Complete test for calibration. calibrations and quality controls in order. | pO2 and Oxi Complete test for calibration. calibrations and quality controls in order. | Defect | HW | Inlet module | Inlet gasket
ABL90 FLEX | Geräteüberprüfung (kein Fehler feststellbar) - Erneuerung Sensorkassette - Analysator ohne Beanstandung! Bemerkung: Kunde muss Kühlkette überprüfen ! (Lieferweg im Hause) | Geräteüberprüfung (kein Fehler feststellbar) - Erneuerung Sensorkassette - Analysator ohne Beanstandung! Bemerkung: Kunde muss Kühlkette überprüfen ! (Lieferweg im Hause) | SWO-113118 | No | Device verification (no error detectable) - Renewal sensor cassette - Analyzer without complaint! Note: Customer must check cold chain ! (delivery route in-house) | Device verification (no error detectable) - Renewal sensor cassette - Analyzer without complaint! Note: Customer must check cold chain ! (delivery route in-house) | Assisted Replacement | SC | SC | K+
ABL90 FLEX | 29-06-2016. ---------------------- Jeg har udskiftet OXI ventil og den er OK nu. Jeg har kørt alle test i service under Fluid treansport. Der er fejl på pCO2, Glu og Lac. Der er skiftet sensorkassette og jeg afventer Merete om dette afhjælper fejlen. Merete har ringet og der er stadig fejl på Lac. Merete skifter væskepakken. Jeg skal kontakt Ørn i morgen og høre hvad status er på ABL90. 30-06-2016 --------------- Jeg har talt med Ørn og alt er grønt og ABL90 virker fint nu. | 29-06-2016. ---------------------- Jeg har udskiftet OXI ventil og den er OK nu. Jeg har kørt alle test i service under Fluid treansport. Der er fejl på pCO2, Glu og Lac. Der er skiftet sensorkassette og jeg afventer Merete om dette afhjælper fejlen. Merete har ringet og der er stadig fejl på Lac. Merete skifter væskepakken. Jeg skal kontakt Ørn i morgen og høre hvad status er på ABL90. 30-06-2016 --------------- Jeg har talt med Ørn og alt er grønt og ABL90 virker fint nu. | SWO-113295 | No | 29-06-2016.----------------------I have replaced the OXI valve and it is OK now. I have run all tests in service during Fluid Treansport. There are errors at pCO2, Glu and Lac. There is changed sensor cassette and I await Merete about this corrects the error. Merete has called and there are still flaws on the LAC. Merete changes the liquid package. I'm going to contact Eagle tomorrow and hear what status is on ABL90. 30-06-2016---------------I have spoken with Eagle and everything is green and ABL90 works fine now. | 29-06-2016.----------------------I have replaced the OXI valve and it is OK now. I have run all tests in service during Fluid Treansport. There are errors at pCO2, Glu and Lac. There is changed sensor cassette and I await Merete about this corrects the error. Merete has called and there are still flaws on the LAC. Merete changes the liquid package. I'm going to contact Eagle tomorrow and hear what status is on ABL90. 30-06-2016---------------I have spoken with Eagle and everything is green and ABL90 works fine now. | Defect | HW | Oxi module | Pinch Valve
ABL90 | checked data logs, observed errors. Suspect possible clot in fluid path. | Performed backflush with inliet gasket off and then again with it on. Performed multiple service program tests. Pressure test, LS adjust, pump calibration, clot detect removal, visual flow check. All passed. | D0F7A952F813344F | No | checked data logs, observed errors. Suspect possible clot in fluid path. | Performed backflush with inliet gasket off and then again with it on. Performed multiple service program tests. Pressure test, LS adjust, pump calibration, clot detect removal, visual flow check. All passed. | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
ABL90 FLEX | Fließweg gespült, Sensorkasette erneuert, nach dem Wechsel der Sensorkasette ist Lac in der QC C unterhalb der Grenze, mehrfach kalibriert aber dies brachte keine Verbesserung, Lösungspack ausgetauscht. Kalibriert und QC's durchgeführt. Reklamation der SC und LP veranlasst. | Fließweg gespült, Sensorkasette erneuert, nach dem Wechsel der Sensorkasette ist Lac in der QC C unterhalb der Grenze, mehrfach kalibriert aber dies brachte keine Verbesserung, Lösungspack ausgetauscht. Kalibriert und QC's durchgeführt. Reklamation der SC und LP veranlasst. | SWO-113141 | No | Flow path flushed, sensor casette renewed, after changing the sensor casette, Lac is in the QC C below the limit, repeatedly calibrated but this brought no improvement, solution pack replaced. Calibrated and QC's performed. complaint of the SC and LP. | Flow path flushed, sensor casette renewed, after changing the sensor casette, Lac is in the QC C below the limit, repeatedly calibrated but this brought no improvement, solution pack replaced. Calibrated and QC's performed. complaint of the SC and LP. | Assisted Replacement | SC | SC | Glu
ABL90 FLEX PLUS |  | B | D0F7C989F863EC27 | No |  |  | Revisit required | SC | SC | Wetability
ABL90 FLEX PLUS | Replaced Hemolyzer set up and test ok. Ran tHb cal passed ok. Ran QC all passed ok. | Replaced Hemolyzer set up and test ok. Ran tHb cal passed ok. Ran QC all passed ok. | SWO-113144 | No | Replaced Hemolyzer set up and test ok. Ran tHb cal passed ok. Ran QC all passed ok. | Replaced Hemolyzer set up and test ok. Ran tHb cal passed ok. Ran QC all passed ok. | Defect | HW | Oxi module | Hemolyzer
ABL90 FLEX PLUS | Reloaded AU software(V3.1 MR8). Checked all inlet hardware. Reinitialized system. | Reloaded AU software(V3.1 MR8). Checked all inlet hardware. Reinitialized system. | SWO-113107 | No | Reloaded AU software(V3.1 MR8). Checked all inlet hardware. Reinitialized system. | Reloaded AU software(V3.1 MR8). Checked all inlet hardware. Reinitialized system. | Adjustment | HW | Auto Inlet module | Off/On
ABL90 FLEX | 更换试剂包修复仪器 | 更换试剂包修复仪器 | SWO-113625 | No | Replace reagent kit repair instruments | Replace reagent kit repair instruments
AQT90 FLEX | 更换处理包。 | 更换处理包。 | SWO-112931 | No | Replace the handling package. | Replace the handling package. | Assisted Replacement | Consumables | SP | SP | 1999
AQT90 FLEX |  | B | SWO-112904 | No |  |  | Defect | HW | Process Module | Optical Unit | 131
AQT90 FLEX | 做样本，故障提示1024，1028，做系统检查，其中暗视野超范围，按照流程处理，没有效果，更换了903-364后修复机器 by jesson | 做样本，故障提示1024，1028，做系统检查，其中暗视野超范围，按照流程处理，没有效果，更换了903-364后修复机器 by jesson | SWO-112704 | No | Do the sample, fault tip 1024, 1028, do the system check, wherein the dark field of view is over range, according to the process, no effect, replaced 903-364 after repair machine by jesson | Do the sample, fault tip 1024, 1028, do the system check, wherein the dark field of view is over range, according to the process, no effect, replaced 903-364 after repair machine by jesson | Defect | HW | Process Module | Optical Unit | 128
AQT90 FLEX |  | B | SWO-112920 | No |  |  | Defect | HW | Storage Module | PCB, cup transport | 5689
AQT90 FLEX |  | B | SWO-112926 | No |  |  | Defect | HW | Incubator module | PCB, Incubator | 12
ABL800 FLEX | Replaced Lac and Glucose Electrodes. Replaced membranes. Performed two point calibration and cleaning. Checked Patient samples and QC. Now working fine. | Replaced Lac and Glucose Electrodes. Replaced membranes. Performed two point calibration and cleaning. Checked Patient samples and QC. Now working fine. | SWO-113766 | No | Replaced Lac and Glucose Electrodes. Replaced membranes. Performed two point calibration and cleaning. Checked Patient samples and QC. Now working fine. | Replaced Lac and Glucose Electrodes. Replaced membranes. Performed two point calibration and cleaning. Checked Patient samples and QC. Now working fine. | Assisted Replacement | Electrode | Electrode | Glu/Lac | Response
ABL835 FLEX | Spildslanger skiftet Inlet kal. | Spildslanger skiftet Inlet kal. | SWO-113108 | No | Spildslanger the turn Inlet shell. | Spildslanger the turn Inlet shell. | Defect | HW | Basic Unit | Waste tube
ABL810 FLEX | Tausch des unteren LS Sensors im pH BG Modul, nicht besser. Erneuerung des LS Sensoes im Einlassmodul, LS Werte wieder besser, Fehler behoben. Einlassschlauch erneuert. Beim Einsaugen der Qualitätskontroll Lösung, kann sie nicht ganz hochgesogen werden im pH BG Modul und in das Oximetrie Modul. Rollenkränze o. K., V7 durchgespült, o. K. Vmtl. ist Friction detection im Einlassmodul defekt. Ersatzteil bestellt. Analysator kann nicht verwendet werden. | Einbau Ersatz Einlassmodul. (Ersatz-Modul 902-783R0109N001; Eigenes Modul 902-783R2406N002) Einspielung der aktuellen Software ins Modul. LS Justage, Pumpenkal durchgeführt. Proteinentfernung fällig, durchgeführt. Keine Flüssigkeitsaustritt mehr und Druckaufbau in Inletmodul mehr, Funktion o. K. Kalibrationen und Qualitätskontrolle o. K. | SWO-113136 | No | Replacement of the lower LS sensor in the pH BG module, no better. Renewal of the LS Sensoes in the inlet module, LS values better again, error fixed. Inlet hose renewed. When sucking in the quality control solution, it cannot be completely sucked high in the pH BG module and in the oximetry module. Roller wreaths o. K., V7 rinsed through, o. K. Vmtl. Friction detection in the inlet module is defective. spare part ordered. Unable to use analyzer. | Installation of replacement inlet module. (Replacement module 902-783R0109N001; Own module 902-783R2406N002) Inserting the current software into the module. LS Adjustment, pump scale. protein removal due. No more liquid leakage and pressure build-up in inlet module, function o.k. calibrations and quality control o.K. | Defect | HW | Inlet module | LS
ABL827 FLEX |  | B | D0F79366F837EA62 | Yes |  |  | Re classified RMED
ABL815 FLEX | Kontrolle Roma.ini, Korrektur Roma.ini, Warmstart, Probelauf Scanner ok | Kontrolle Roma.ini, Korrektur Roma.ini, Warmstart, Probelauf Scanner ok | SWO-113037 | No | Control Roma.ini, correction Roma.ini, warm start, test run scanner ok | Control Roma.ini, correction Roma.ini, warm start, test run scanner ok | Setup | Sw | Roma.ini | Roma.ini
AQT90 FLEX |  | B | D0FEE598BC5EFBAB | No |  |  | Defect | HW | Incubator module | Punch up
ABL80 FLEX BASIC | 更换测试卡。 | 更换测试卡。 | SWO-113331 | No | Replace the test card. | Replace the test card.
ABL80 FLEX BASIC |  | B | SWO-113291 | No
ABL80 FLEX BASIC | 用零点测试卡检查，需要更换903-334，更换后，机器正常by jesson | B | SWO-113166 | No | Check with zero test card, need to replace 903-334, after replacement, the machine normal by jesson
ABL80 CO-OX | Level 3 QC reading high at or just below high end of package insert range. 2 different lots of QC yield same results. Multiple SP's and SC's have been replaced. | A Techline case is needed here for escalation purposes. Download all log files and DB. Will create case later today. | D0F8D2186E1A907F | No | Level 3 QC reading high at or just below high end of package insert range. 2 different lots of QC yield same results. Multiple SP's and SC's have been replaced. | A Techline case is needed here for escalation purposes. Download all log files and DB. Will create case later today.
ABL90 FLEX | Windows operációs rendszer és analizátor szoftverek újratelepítése, Fogyóanyagok telepítése Inicializálás,Riportok beállítása, Kalibrációk (SC, tHb),QC mérések | B | SWO-113455 | No | Windows operating system and analyzer software reinstalling, consumables installation initialization, setting reports, calibration (SC, tHb), QC measurements |  | Defect | SW | SW | Re installed SW
ABL90 | pCO2 failures are present after replacement of sensor cassette (SC). pCO2 is failing low. Flushing of the SC has not resolved the issue. | Replaced the SC and verified successful pCO2 calibration. Printed and submitted the credit claim form on behalf of the customer. Analyzer is within acceptable operating specifications. | D0F906556E09E212 | No | pCO2 failures are present after replacement of sensor cassette (SC). pCO2 is failing low. Flushing of the SC has not resolved the issue. | Replaced the SC and verified successful pCO2 calibration. Printed and submitted the credit claim form on behalf of the customer. Analyzer is within acceptable operating specifications. | Assisted Replacement | SC | SC | pCO2
ABL90 FLEX | Batterij defect Deze toont dat ze voor 100% is opgeladen maar wanneer de netspanning wordt afgesloten valt het toestel toch uit. Defecte batterij (P0059N024) vervangen door nieuwe (P0112N011). Test OK Alle toestellen werden nagekeken : IZ13 en PACU dienen eveneens een nieuwe batterij te krijgen | Batterij defect Deze toont dat ze voor 100% is opgeladen maar wanneer de netspanning wordt afgesloten valt het toestel toch uit. Defecte batterij (P0059N024) vervangen door nieuwe (P0112N011). Test OK Alle toestellen werden nagekeken : IZ13 en PACU dienen eveneens een nieuwe batterij te krijgen | SWO-113628 | No | Battery failure It shows that it has been charged for 100% but when the mains voltage is shut down the unit still shuts off. Defective battery (P0059N024) replaced with new (P0112N011). Test OK All devices were checked: IZ13 and PACU also need to get a new battery | Battery failure It shows that it has been charged for 100% but when the mains voltage is shut down the unit still shuts off. Defective battery (P0059N024) replaced with new (P0112N011). Test OK All devices were checked: IZ13 and PACU also need to get a new battery | Defect | HW | Battery module | Battery module | To be returned
ABL90 | Customer replaced the sensosr cassette. Analyzer is in the intervention screen and requesting sensor cassette replacement. Performed fluid transport procedures in service mode, flushed both the analyzer and sensor cassette and re-initialized the sensor cassette. Errors are still present and cassette is in need of replacement. | Replaced sensor cassette. All parameters are now properly calibrated. Analyzer is within acceptable operating specifications. | D0F92F466E1D264C | No | Customer replaced the sensosr cassette. Analyzer is in the intervention screen and requesting sensor cassette replacement. Performed fluid transport procedures in service mode, flushed both the analyzer and sensor cassette and re-initialized the sensor cassette. Errors are still present and cassette is in need of replacement. | Replaced sensor cassette. All parameters are now properly calibrated. Analyzer is within acceptable operating specifications. | Assisted Replacement | SC | SC | Multi
ABL90 FLEX PLUS |  | B | D0F8E73A6E552672 | No |  |  | Defect | Hw | Sensor cassette bay unit | SC bay unit
ABL827 FLEX | Replaced the EL MET Module measuring chamber. New Lactate electrode fitted and it was cleaned with fresh hypochlorite prior to installation as per the procedure from Copenhagen. New membrane fitted to the electrode and a cleaning cycle run immediately after installing it. Ca membrane replaced because of unstable calibrations after changing the measuring chamber. Crea membranes changed because of a 0954 error on level 2 QC (recommended action for this error on membranes over 24Hrs old) | As the zero current calibration was recently performed due to zero current errors the El Met PCB was replaced and the analyser put back into use. The Ca membrane was replaced because the sensitivity was too low and the Lac was replaced because the sensitivity had dropped close to the lower limit. Once this analyser was back up and running the loan analyser was run down and removed from site. | SWO-113586 | No | Replaced the EL MET Module measuring chamber. New Lactate electrode fitted and it was cleaned with fresh hypochlorite prior to installation as per the procedure from Copenhagen. New membrane fitted to the electrode and a cleaning cycle run immediately after installing it. Ca membrane replaced because of unstable calibrations after changing the measuring chamber. Crea membranes changed because of a 0954 error on level 2 QC (recommended action for this error on membranes over 24Hrs old) | As the zero current calibration was recently performed due to zero current errors the El Met PCB was replaced and the analyser put back into use. The Ca membrane was replaced because the sensitivity was too low and the Lac was replaced because the sensitivity had dropped close to the lower limit. Once this analyser was back up and running the loan analyser was run down and removed from site. | Assisted Replacement | Electrode | Electrode | Lac
ABL827 FLEX | Replaced the EL MET Module measuring chamber. New Lactate electrode fitted and it was cleaned with fresh hypochlorite prior to installation as per the procedure from Copenhagen. New membrane fitted to the electrode and a cleaning cycle run immediately after installing it. Glu, Ref & Ca membranes replaced because they were all due with the next 7 days. | Measuring chambers removed and soaked with protein remover. New sealing ring fitted to the Glu electrode. New Lactate membrane fitted. New Ca membrane fitted because it was failing calibrations. | SWO-113587 | No | Replaced the EL MET Module measuring chamber. New Lactate electrode fitted and it was cleaned with fresh hypochlorite prior to installation as per the procedure from Copenhagen. New membrane fitted to the electrode and a cleaning cycle run immediately after installing it. Glu, Ref & Ca membranes replaced because they were all due with the next 7 days. | Measuring chambers removed and soaked with protein remover. New sealing ring fitted to the Glu electrode. New Lactate membrane fitted. New Ca membrane fitted because it was failing calibrations. | Assisted Replacement | Electrode | Electrode | Lac
ABL825 FLEX | Geräte- und Datenüberprüfung Gas 1 erneuert pCO 2 membraniert ( Kunde hatte bereits mehrfach in den letzten Tagen membr.) Überprüfung Ventile am BG Modul, pH Elektrode u. Meßkammer Dichtring erneuert Einlaß , incl. Nadel gereinigt | Geräte- und Datenüberprüfung Gas 1 erneuert pCO 2 membraniert ( Kunde hatte bereits mehrfach in den letzten Tagen membr.) Überprüfung Ventile am BG Modul; pH Elektrode u. Meßkammer Dichtring erneuert Einlaß , incl. Nadel gereinigt | SWO-113394 | No | Device and data verification Gas 1 renewed pCO 2 membraneed (customer had already met several times in the last days) Checking valves on the BG module; pH electrode and measuring chamber sealing ring renewed inlet, incl. needle cleaned | Device and data verification Gas 1 renewed pCO 2 membraneed (customer had already met several times in the last days) Checking valves on the BG module; pH electrode and measuring chamber sealing ring renewed inlet, incl. needle cleaned | Assisted Replacement | Electrode | Membrane | pCO2
ABL825 FLEX | check the peer ranges and Lac is running on the high end and intermittently outside the peer 2 range , looked through logs this happened after the new lot of QC was put in, order a different lot of QC to try Level 1 673, level 2 629, Level 3 609, Will open a tech line case both analyzers are running on the high end and intermittently out on Level 3 QC | Ran different lot of QC and still showing higher than the peer 2--Level 3 10.8; performed a servcie disinfection procedure per the service manuall; replaced all membranes and pump tubings; replace pH electrode -- still runing on the high end of the peer 2 range-- intermittently is high outside the peer 2 | D0F8D00D6E2294DC | No | check the peer ranges and Lac is running on the high end and intermittently outside the peer 2 range ; looked through logs this happened after the new lot of QC was put in; order a different lot of QC to try Level 1 673; level 2 629; Level 3 609; Will open a tech line case both analyzers are running on the high end and intermittently out on Level 3 QC | Ran different lot of QC and still showing higher than the peer 2--Level 3 10.8; performed a servcie disinfection procedure per the service manuall; replaced all membranes and pump tubings; replace pH electrode -- still runing on the high end of the peer 2 range-- intermittently is high outside the peer 2 | Assisted Replacement | Electrode | Membrane | Lac | Techline
ABL825 FLEX | Glucoseelektrode erneuert, Gerät kalibriert und mit QC überprüft E7066 defekt: UK-02 | Glucoseelektrode erneuert, Gerät kalibriert und mit QC überprüft E7066 defekt: UK-02 | SWO-113497 | No | Glucose electrode renewed, device calibrated and checked with QC E7066 defective: UK-02 | Glucose electrode renewed, device calibrated and checked with QC E7066 defective: UK-02 | Assisted Replacement | Electrode | Electrode | Glu
ABL810 FLEX | Einbau Ersatz Einlassmodul. (Ersatz-Modul 902-783R0109N001, Eigenes Modul 902-783R2406N002) Einspielung der aktuellen Software ins Modul. LS Justage, Pumpenkal durchgeführt. Proteinentfernung fällig, durchgeführt. Keine Flüssigkeitsaustritt mehr und Druckaufbau in Inletmodul mehr, Funktion o. K. Kalibrationen und Qualitätskontrolle o. K. | Einbau Ersatz Einlassmodul. (Ersatz-Modul 902-783R0109N001; Eigenes Modul 902-783R2406N002) Einspielung der aktuellen Software ins Modul. LS Justage, Pumpenkal durchgeführt. Proteinentfernung fällig, durchgeführt. Keine Flüssigkeitsaustritt mehr und Druckaufbau in Inletmodul mehr, Funktion o. K. Kalibrationen und Qualitätskontrolle o. K. | SWO-113461 | No | Installation of replacement inlet module. (Replacement module 902-783R0109N001; Own module 902-783R2406N002) Inserting the current software into the module. LS Adjustment, pump scale. protein removal due. No more liquid leakage and pressure build-up in inlet module, function o.k. calibrations and quality control o.K. | Installation of replacement inlet module. (Replacement module 902-783R0109N001; Own module 902-783R2406N002) Inserting the current software into the module. LS Adjustment, pump scale. protein removal due. No more liquid leakage and pressure build-up in inlet module, function o.k. calibrations and quality control o.K. | Defect | HW | Inlet module | Inlet module
ABL800 BASIC |  | B | SWO-113293 | No |  |  | Assisted Replacement | Electrode | Membrane | Lac | 378
ABL815 FLEX |  | Technikereinsatz erforderlich | SWO-113551 | Yes |  | Technician stake required | Re classified RMED
ABL815 FLEX | Oximetriemodul komplett ersetzt - justiert und kalibriert, Service tHb, Kal, tHb Kal, QC | Oximetriemodul komplett ersetzt - justiert und kalibriert, Service tHb, Kal, tHb Kal, QC | SWO-113682 | No | Oximetry module completely replaced - adjusted and calibrated, service tHb, Kal, tHb Kal, QC | Oximetry module completely replaced - adjusted and calibrated, service tHb, Kal, tHb Kal, QC | Defect | HW | Oxi module | Oxi module
AQT90 FLEX | Removed bar code scanner and cover and realigned scanner beam and test ok. Refitted cover and retest scanning ok. Ran system clean ok. Ran QC ok. | Removed bar code scanner and cover and realigned scanner beam and test ok. Refitted cover and retest scanning ok. Ran system clean ok. Ran QC ok. | SWO-113379 | No | Removed bar code scanner and cover and realigned scanner beam and test ok. Refitted cover and retest scanning ok. Ran system clean ok. Ran QC ok. | Removed bar code scanner and cover and realigned scanner beam and test ok. Refitted cover and retest scanning ok. Ran system clean ok. Ran QC ok. | Cleaned | HW | Storage Module | Barcode reader, Storage wheel
ABL80 CO-OX | Found no leaks, but the hemolyzer does not appear to be working properly. | Replaced the hemolyzer assembly. | D0FA18E7E46296AA | No | Found no leaks, but the hemolyzer does not appear to be working properly. | Replaced the hemolyzer assembly.
ABL80 FLEX BASIC | 检查902-972接口板，未发现异常，用零点校准卡校准后，按照提示需要更换903-334，更换903-334一个正常By Jesson | 检查902-972接口板，未发现异常，用零点校准卡校准后，按照提示需要更换903-334，更换903-334一个正常By Jesson | SWO-113675 | No | Check the 902-972 interface board, no abnormalities found, after calibration with zero-point calibration card, follow the prompt need to replace 903-334, replace 903-334 a normal By Jesson | Check the 902-972 interface board, no abnormalities found, after calibration with zero-point calibration card, follow the prompt need to replace 903-334, replace 903-334 a normal By Jesson
ABL80 CO-OX | Techline case needs to be created. | Create Techline case and submit logs and DB to support email address. | D0FA05C5E404D0E0 | No | Techline case needs to be created. | Create Techline case and submit logs and DB to support email address.
ABL80 FLEX BASIC | 分解してメンテナンスキット交換 | 分解してメンテナンスキット交換 | SWO-113858 | No | Disassemble and replace maintenance kit | Disassemble and replace maintenance kit
ABL90 | As above. Customer did several flush procedures and did see some debris exit during flush. | Perofrm several back flush procedures with hot tap water. | D0FA1F25E4355E74 | No | As above. Customer did several flush procedures and did see some debris exit during flush. | Perofrm several back flush procedures with hot tap water. | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
ABL90 FLEX | Erneuerung SD Karte - Installation Windows und Applikationssoftware - Konfigurationseinstellungen - Installation SP und SC - Funktionsüberprüfung etc. | Erneuerung SD Karte - Installation Windows und Applikationssoftware - Konfigurationseinstellungen - Installation SP und SC - Funktionsüberprüfung etc. | SWO-113767 | No | Renewal SD card - Installation Windows and application software - Configuration settings - Installation SP and SC - Function verification etc. | Renewal SD card - Installation Windows and application software - Configuration settings - Installation SP and SC - Function verification etc. | Defect | HW | CF Card | CF Card
AQT90 FLEX | 做样本提示干燥温度超范围，更换了903-210后，重新标定后正常 by jesson | 做样本提示干燥温度超范围，更换了903-210后，重新标定后正常 by jesson | SWO-112991 | No | Do the sample prompt drying temperature over range, replace 903-210, recalibrate after normal by jesson | Do the sample prompt drying temperature over range, replace 903-210, recalibrate after normal by jesson | Defect | HW | Process Module | PCB, Cup dryer | 169
AQT90 FLEX |  | B | SWO-113655 | No |  |  | Defect | HW | Process Module | Process Unit sensor | 5784
ABL800 BASIC | Found all the sample measurements Na values are high, found calibration results are low, replaced the pCO2, Ref & Na membranes, calibrated found results are within the range, performed the QC level2 found instrument is functional. | Found all the sample measurements Na values are high, found calibration results are low, replaced the pCO2, Ref & Na membranes, calibrated found results are within the range, performed the QC level2 found instrument is functional. | SWO-113855 | No | Found all the sample measurements Na values are high, found calibration results are low, replaced the pCO2, Ref & Na membranes, calibrated found results are within the range, performed the QC level2 found instrument is functional. | Found all the sample measurements Na values are high, found calibration results are low, replaced the pCO2, Ref & Na membranes, calibrated found results are within the range, performed the QC level2 found instrument is functional. | Assisted Replacement | Electrode | Membrane | Na+
ABL825 FLEX | part received, instrumnet failed calibration checks in service | i replaced the inlet module pcb board,failed lac calibration replaced lactate membrane | D0FA5130E44826ED | No | part received, instrumnet failed calibration checks in service | i replaced the inlet module pcb board,failed lac calibration replaced lactate membrane | Defect | HW | Inlet module | Probe cal.
ABL810 FLEX | Keys eingespielt, Lampe justiert, CompleteOxiTest,, Service tHb, Kal,tHb Kal, QC | Keys eingespielt, Lampe justiert, CompleteOxiTest,, Service tHb, Kal,tHb Kal, QC | SWO-113686 | No | Keys eingespielt, Lampe justiert, CompleteOxiTest,, Service tHb, Kal,tHb Kal, QC | Keys eingespielt, Lampe justiert, CompleteOxiTest,, Service tHb, Kal,tHb Kal, QC | Adjustment | HW | Oxi module | Lamp unit
ABL810 FLEX | neu membraniert, pCO2 Elektrode ersetzt, Inserts korrigiert, Ventilmembran ersetzt, Schlauch ersetzt, mehrere Kal, QC nach wie vor zu niedrig - abgebrochen | Kunde wird später die QC-Messung Level 2 wiederholen Elektrode wurde u.a. vor ca. 1 h neu membraniert | SWO-113687 | No | new membrane, pCO2 electrode replaced, inserts corrected, valve membrane replaced, hose replaced, several kal, QC still too low - aborted | Customer will later repeat the QC measurement level 2 electrode was re-membraneed, among other things, about 1 h ago | Assisted Replacement | Electrode | Electrode | pCO2
ABL1 | Reinstalled Sc, Problem remains, Replaced cassette with new one, PER sent for investigation and replacement. | Reinstalled Sc, Problem remains, Replaced cassette with new one, PER sent for investigation and replacement. | SWO-115349 | No | Reinstalled Sc, Problem remains, Replaced cassette with new one, PER sent for investigation and replacement. | Reinstalled Sc, Problem remains, Replaced cassette with new one, PER sent for investigation and replacement.
AQT90 FLEX | 按照库存机检查光单流程，需要更换光单，更换了光学单元903-364后修复 by jesson | 按照库存机检查光单流程，需要更换光单，更换了光学单元903-364后修复 by jesson | SWO-113647 | No | According to the inventory machine inspection light sheet process, need to replace the light sheet, replaced the optical unit 903-364 after repair by jesson | According to the inventory machine inspection light sheet process, need to replace the light sheet, replaced the optical unit 903-364 after repair by jesson | Defect | HW | Process Module | Optical Unit | 131
AQT90 FLEX | 拆仪器屏幕，清洁条码扫描器 | 拆仪器屏幕，清洁条码扫描器 | SWO-113886 | No | Remove the instrument screen, clean the barcode scanner | Remove the instrument screen, clean the barcode scanner | Cleaned | HW | Analyzer Control Module | Barcode reader, display
ABL825 FLEX |  | B | SWO-113849 | No |  |  | Assisted Replacement | Electrode | Membrane | Lac
ABL80 FLEX BASIC |  | B | SWO-113906 | No
ABL80 FLEX BASIC | replaced solution pack and activate ca and lac and given two point calibration. Analyser is working fine now. | replaced solution pack and activate ca and lac and given two point calibration. Analyser is working fine now. | SWO-114818 | No | replaced solution pack and activate ca and lac and given two point calibration. Analyser is working fine now. | replaced solution pack and activate ca and lac and given two point calibration. Analyser is working fine now.
AQT90 FLEX |  | B | SWO-113904 | No |  |  | Defect | HW | Storage Module | Barcode reader, Storage wheel
ABL800 BASIC | CALIBRATED THE TOUCH SCREEN NOW ITS WORKING BUT IT NEED TO REPLACE. | B | SWO-114819 | No | CALIBRATED THE TOUCH SCREEN NOW ITS WORKING BUT IT NEED TO REPLACE. |  | Defect | HW | DMS_PC | Touch Screen
ABL80 FLEX BASIC | 更换新测试卡修复仪器 | 更换新测试卡修复仪器 | SWO-115568 | No | Replace the new test card repair instrument | Replace the new test card repair instrument
ABL80 FLEX CO-OX | Klappe des Lösungspacks erneuert. Parameter wieder aktiviert und kalibriert. Weiterhin Kalibrationsfehler im pH und pCo2. Oximetrie überprüft. Dabei wurde festgestellt, dass der Hemolysator undicht ist. Ersatzteil bestellt. | Hemolyzer ausgetauscht. Gerät kalibriert und überprüft. Wieder zurück zum Kunden gebracht. | SWO-114210 | No | flap of the solution pack. parameter is reactivated and calibrated. Further calibration errors in pH and pCo2. Oximetry. It was determined that the hemolyser is leaking. spare part ordered. | Hemolyzer. calibrated and checked. Brought back to the customer.
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-115850 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL80 FLEX BASIC | reset the solution pack .fail to installation. sp 80 faulty. installed new solution pack problem resolved. | B | SWO-114223 | No | reset the solution pack .fail to installation. sp 80 faulty. installed new solution pack problem resolved.
ABL90 |  | B | D1182D65F4336618 | No |  |  | Unknown | Unknown | Unknown | Unknown
ABL90 FLEX | Fehlersuche, div. Serviceprogr. durchgef., Einlaßdichtung gew., Startphase abgewartet, LP gew., Kal. + QC i.O. | Fehlersuche; div. Serviceprogr. durchgef.; Einlaßdichtung gew.; Startphase abgewartet; LP gew.; Kal. + QC i.O. | SWO-114226 | No | Troubleshooting; div. Serviceprogr. through it; inlet seal. start-up phase awaited; LP. Kal. + QC in the first place | Troubleshooting; div. Serviceprogr. through it; inlet seal. start-up phase awaited; LP. Kal. + QC in the first place | Defect | HW | Inlet module | Inlet gasket
ABL90 FLEX | Avoid use of BD syringes or any other make with silicon grease within the syringe. | Avoid use of BD syringes or any other make with silicon grease within the syringe. | SWO-114117 | No | Avoid use of BD syringes or any other make with silicon grease within the syringe. | Avoid use of BD syringes or any other make with silicon grease within the syringe. | Assisted Replacement | SC | SC | pO2 | silicone
ABL90 FLEX | Ersatzgerät zur weiteren Überprüfung aufgestellt. Netzwerkdaten eingestellt. Verbindung zu isoft überprüft. | Ersatzgerät zur weiteren Überprüfung aufgestellt. Netzwerkdaten eingestellt. Verbindung zu isoft überprüft. | SWO-114209 | No | replacement device for further verification. network data. Connection to isoft checked. | replacement device for further verification. network data. Connection to isoft checked. | Defect | HW | Inlet module | Inlet gasket | 163
ABL90 FLEX PLUS |  | B | D0FDC07E462F1073 | No |  |  | Assisted Replacement | SC | SC | Glu/Lac
ABL90 FLEX PLUS |  | B | D0FDC2B046301076 | No |  |  | Assisted Replacement | SC | SC | pCO2
ABL90 FLEX | 更换测试卡正常。 | 更换测试卡正常。 | SWO-114081 | No | The replacement test card is OK. | The replacement test card is OK. | Assisted Replacement | SC | SC | pCO2 | 378
ABL835 FLEX | Limpieza del espejo del módulo de autocheck y prueba de funcionamiento. Se observa problemas en la calibración del lactato. Sustitución de membrana y calibración. Revisión de conexión a Radiance. Se coloca cable en sitio correcto y se corrige IP de equipo. Se comprueba correcto funcionamiento del equipo mediante QC. | B | D11476D692566505 | No | Cleaning the autocheck module mirror and test operation. Problems with lactate calibration are observed. Membrane replacement and calibration. Connection review to Radiance. Cable is placed in the correct place and computer IP is corrected. Proper operation of the equipment is checked using QC. |  | Cleaned | HW | AutoCheck Module | Barcode reader
ABL800 BASIC | upload window and application software. Also clean all valves perform 2-point calibration now machine is ok. | upload window and application software. Also clean all valves perform 2-point calibration now machine is ok. | SWO-115226 | No | upload window and application software. Also clean all valves perform 2-point calibration now machine is ok. | upload window and application software. Also clean all valves perform 2-point calibration now machine is ok. | Defect | SW | Roma.ini | Off/On
ABL825 FLEX |  | B | D0FE2503462A93B5 | No |  |  | User error | Replacement schedule | Replacement schedule | Replacement schedule
ABL825 FLEX | error code 657 top plug never reached pHBG upper LS, 510 EL Met upper LS did not recieve rinse from above, 509 pH BG upper LS did not recieve rinse from above, Lac still running on the high end | inspected all valves; found crystalization build up in the blue Met valve and in the pH valve; replaced both blue air valves in the measuring modules; and replaced the rinse valve in the Met module due to left over liquid found -- after replacing the valves the Lac came down into the mean was running on the high end | D0FDC33846501FE8 | No | error code 657 top plug never reached pHBG upper LS; 510 EL Met upper LS did not recieve rinse from above; 509 pH BG upper LS did not recieve rinse from above; Lac still running on the high end | inspected all valves; found crystalization build up in the blue Met valve and in the pH valve; replaced both blue air valves in the measuring modules; and replaced the rinse valve in the Met module due to left over liquid found -- after replacing the valves the Lac came down into the mean was running on the high end | Defect | HW | pH/BG Module | Valve | 59
ABL800 BASIC | formatted the system and installed the operating system and application software. | formatted the system and installed the operating system and application software. | SWO-115389 | No | formatted the system and installed the operating system and application software. | formatted the system and installed the operating system and application software. | Defect | SW | Application SW | Application SW
ABL827 FLEX |  | B | D1182E68F460661B | No |  |  | Unknown | Unknown | Unknown | 6,16
TCM/Tosca |  | B | D0FE090C4662D85C | No
ABL80 FLEX CO-OX |  | B | SWO-114241 | No
ABL80 FLEX BASIC | 点検作業実施 | B | SWO-116559 | No | Inspection work carried out
ABL80 CO-OX | Had pO2 problem before the sensor cassette was replaced. Customer replaced the solution pack per FSE instructions. Checked tubing, pumps and valves. Performed a service zero point check with no errors. Found no issues with the analyzer. | Replaced the sensor cassette. | D0FF152DBC543A0C | No | Had pO2 problem before the sensor cassette was replaced. Customer replaced the solution pack per FSE instructions. Checked tubing, pumps and valves. Performed a service zero point check with no errors. Found no issues with the analyzer. | Replaced the sensor cassette.
ABL90 |  | B | D0FF6904BC2ABC38 | No |  |  | Assisted Replacement | SC | SC | Lac
ABL90 | 2 pCO2, failed sensitivity, 0378, 0210, QC A, QC B, QC C | 600-100% r r | D0FF417FBC0ABC23 | No | 2 pCO2, failed sensitivity, 0378, 0210, QC A, QC B, QC C | 600-100% r R | Assisted Replacement | SC | SC | pCO2
ABL90 FLEX | Kunde hatte bereits Kassette erneuert, die aber Kalibrierfehler hatte. Neue Kassette geliefert. Gerät überprüft | Kunde hatte bereits Kassette erneuert, die aber Kalibrierfehler hatte. Neue Kassette geliefert. Gerät überprüft | SWO-114492 | No | Customer had already renewed cassette, but had calibration errors. New cassette supplied. Device checked | Customer had already renewed cassette, but had calibration errors. New cassette supplied. Device checked | Assisted Replacement | SC | SC | pCO2
ABL90 FLEX | Kundengespräch, Fehlersuche und Systemneustart durchgeführt. SP90 aus Kundenbestand erneuert. | Kundengespräch, Fehlersuche und Systemneustart durchgeführt. SP90 aus Kundenbestand erneuert. | SWO-114362 | No | customer call, troubleshooting and system restart. SP90 renewed from customer base. | customer call, troubleshooting and system restart. SP90 renewed from customer base. | Assisted Replacement | SP | SP | FT
ABL90 FLEX | Schlauchsystem gespült, Servicetest-Programm, SC muss erneuert werden | Schlauchsystem gespült, Servicetest-Programm, SC muss erneuert werden | SWO-114508 | Yes | Hose system flushed, service test program, SC needs to be renewed | Hose system flushed, service test program, SC needs to be renewed | Re classified RMED | Age , OOB
ABL90 FLEX PLUS |  | B | D0FF0697BC155D68 | No |  |  | Defect | SW | Chrash | Chrash
ABL90 FLEX | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-114536 | No | Repair after replacing a new test card | Repair after replacing a new test card | Assisted Replacement | SC | SC | Cl | 378
AQT90 FLEX | 更换925-016一个和903-117一个正常。（by andrew） | 更换925-016一个和903-117一个正常。（by andrew） | SWO-114010 | No | Replace 925-016 one and 903-117 one normal. (by andrew) | Replace 925-016 one and 903-117 one normal. (by andrew) | Defect | HW | Cover and Chassis | PCB, Cabinet
ABL827 FLEX | I cleaned the barcode readers lens and after that there were no gaps in the laser. Used the barcode reader to log on several times and also removed all the battles and rescanned them without any problems. | I cleaned the barcode readers lens and after that there were no gaps in the laser. Used the barcode reader to log on several times and also removed all the battles and rescanned them without any problems. | SWO-114458 | No | I cleaned the barcode readers lens and after that there were no gaps in the laser. Used the barcode reader to log on several times and also removed all the battles and rescanned them without any problems. | I cleaned the barcode readers lens and after that there were no gaps in the laser. Used the barcode reader to log on several times and also removed all the battles and rescanned them without any problems. | Cleaned | HW | AutoCheck Module | Barcode reader
ABL835 FLEX | Bias sur oximetrie entre les deux appareils | Remplac Hemolyser | D0FF2585BC16D0A0 | No | Bias on oximetrie between the two devices | Remplac Hemolyser | Defect | HW | Oxi module | Hemolyzer
ABL835 FLEX | Bias sur oximetrie entre les deux appareils | remplacé hemolyser | D0FF3164BC0CD0A5 | No | Bias on oximetrie between the two devices | replaced hemolyser | Defect | HW | Oxi module | Hemolyzer
ABL815 FLEX | Glu Elektrode neu membraniert, Datensicherung an allen 3Analysatoren gezogen, Kal, QC | Glu Elektrode neu membraniert, Datensicherung an allen 3Analysatoren gezogen, Kal, QC | SWO-114654 | No | Glu electrode re-membraneed, data backup pulled to all 3 analyzers, Kal, QC | Glu electrode re-membraneed, data backup pulled to all 3 analyzers, Kal, QC | Assisted Replacement | Electrode | Membrane | Glu
ABL800 BASIC | Downgraded ABL800 application to v6.13 from existing v6.15. Touchscreen error was rectified | Downgraded ABL800 application to v6.13 from existing v6.15. Touchscreen error was rectified | SWO-115235 | No | Downgraded ABL800 application to v6.13 from existing v6.15. Touchscreen error was rectified | Downgraded ABL800 application to v6.13 from existing v6.15. Touchscreen error was rectified | Defect | SW | Application SW | Application SW
ABL825 FLEX | Lac Elektrode ausgebaut und neu membraniert, Elektrodenspitze zuvor gereinigt. Kalibrationen, QC Messung, OK. | Lac Elektrode ausgebaut und neu membraniert, Elektrodenspitze zuvor gereinigt. Kalibrationen, QC Messung, OK. | SWO-114359 | No | Lac electrode removed and re-membraneed, electrode tip previously cleaned. Calibrations, QC measurement, OK. | Lac electrode removed and re-membraneed, electrode tip previously cleaned. Calibrations, QC measurement, OK. | Assisted Replacement | Electrode | Electrode | Lac | Sensitivity
AQT90 FLEX | Try to fix sample holder but not working.need to change holder | Try to fix sample holder but not working.need to change holder | SWO-115216 | No | Try to fix sample holder but not working.need to change holder | Try to fix sample holder but not working.need to change holder | Revisit required | HW | Aspiration Module | Sample Holder
ABL80 FLEX BASIC | Gerät geprüft. Schlauchkit erneuert. LP und SC installiert. Messung abgewartet. | Gerät geprüft. Schlauchkit erneuert. LP und SC installiert. Messung abgewartet. | SWO-114839 | No | device checked. Hose kit renewed. LP and SC installed. measurement awaited. | device checked. Hose kit renewed. LP and SC installed. measurement awaited.
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-114584 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL80 FLEX BASIC | 更换配件902-972修复仪器(by andrew) | 更换配件902-972修复仪器(by andrew) | SWO-114013 | No | Replacement accessory 902-972 repair instrument (by andrew) | Replacement accessory 902-972 repair instrument (by andrew)
ABL80 FLEX BASIC | 更换903-334和902-972一个正常。(by andrew) | 更换903-334和902-972一个正常。(by andrew) | SWO-114260 | No | Replace the 903-334 and 902-972 one normal. (by andrew) | Replace the 903-334 and 902-972 one normal. (by andrew)
ABL80 FLEX BASIC | Re- install the SP. After reset and re-install the SC unit is working properly. | Re- install the SP. After reset and re-install the SC unit is working properly. | SWO-115376 | No | Re- install the SP. After reset and re-install the SC unit is working properly. | Re- install the SP. After reset and re-install the SC unit is working properly.
ABL90 FLEX | Gerät geprüft. Serviceprogramme aufgerufen. Fliesswege mehrmals gespült. Einlassnadel und Einlassdichtung geprüft getauscht. Lösungspack und Sensorcassette erneut installiert. Startphase abgewartet und QC durchgeführt. | Gerät geprüft. Serviceprogramme aufgerufen. Fliesswege mehrmals gespült. Einlassnadel und Einlassdichtung geprüft getauscht. Lösungspack und Sensorcassette erneut installiert. Startphase abgewartet und QC durchgeführt. | SWO-114838 | No | device checked. service programs are called. Flow paths rinsed several times. Inlet pin and inlet seal checked. solution pack and sensor cassette reinstalled. start-up phase and QC performed. | device checked. service programs are called. Flow paths rinsed several times. Inlet pin and inlet seal checked. solution pack and sensor cassette reinstalled. start-up phase and QC performed. | Assisted Replacement | SP | SP | Flow selector | 1177
ABL90 | Could not verify. Found that the SIT had connected to the system and corrected the problem. | Problem resolved. No MTR Required. | D10058F3321BBC34 | No | Could not verify. Found that the SIT had connected to the system and corrected the problem. | Problem resolved. No MTR Required. | Setup | SW | Time setting | Time setting
ABL90 |  | B | D100519A3214BC46 | Yes |  |  | Re classified RMED
ABL90 | Found later during the troubleshooting that the analyzer was actually appeared connected to Radiance, however the information was not being processed. | Called the SIT to have the Radiance server reset. Waiting on the on call SIT. Problem resolved server issue. | D1009F1B3224BC38 | No | Found later during the troubleshooting that the analyzer was actually appeared connected to Radiance, however the information was not being processed. | Called the SIT to have the Radiance server reset. Waiting on the on call SIT. Problem resolved server issue. | No fault found | SW | His/lis | Radiance
ABL90 | LS! and LS2 will not pass Liquid Sensor calibration.Errors 531, 537, 529. | Replaced defective Sensor Bay Assy. | D10985B26C5581B8 | No | LS! and LS2 will not pass Liquid Sensor calibration. Errors 531, 537, 529. | Replaced defective Sensor Bay Assy. | Defect | HW | Cassette retainer | Cassette retainer testet
ABL90 |  | B | D101CED3A8213640 | Yes |  |  | Re classified RMED | Installation
ABL90 FLEX PLUS | Einlassdichtung wurde bereits gestern durch Med.technik erneuert, Fehler kam aber heute nach wieder, gebrauchte Einlassdichtung aus S-box eingesetzt, Startphase abgewartet, Kal & QC i.O., aus Reisegepäck zwei Einlassdichtung da gelassen. | Einlassdichtung wurde bereits gestern durch Med.technik erneuert, Fehler kam aber heute nach wieder, gebrauchte Einlassdichtung aus S-box eingesetzt, Startphase abgewartet, Kal & QC i.O.; aus Reisegepäck zwei Einlassdichtung da gelassen. | SWO-114692 | No | Inlet seal was already renewed yesterday by Med.technik, but errors came back today, used inlet seal from S-box used, start phase awaited, Kal & QC i.O.; two inlet seals from luggage. | Inlet seal was already renewed yesterday by Med.technik, but errors came back today, used inlet seal from S-box used, start phase awaited, Kal & QC i.O.; two inlet seals from luggage. | Defect | HW | Auto Inlet module | Inlet Gasket
AQT90 FLEX | 更换910-329解决问题。 | 更换910-329解决问题。 | SWO-114242 | No | Replace the 910-329 to resolve the issue. | Replace the 910-329 to resolve the issue. | Defect | HW | Analyzer Control Module | CF Card
ABL810 FLEX | Descontaminación del equipo. Desmontaje, limpieza y ajuste del lector de barras, limpieza con cleaner electrode solution del electrodo de pH, cambio de la membrana de referecia, así como los contactos de los preamplificadores. Enderezar la sonda de aspiración por problemas de detección de aire durante la aspiración. Limpieza de la cámara de medida. Calibraciones varias para estabilizar las lecturas de pH. Revisión y comprobación del equipo con QC´s. | Limpieza lector código de barras y ajuste del mismo. Limpieza elelctrodo de pH. | D1009CAF32463EF9 | No | Decontamination of equipment. Disassembly, cleaning and adjustment of the bar reader; Cleaning with cleaner electrode solution of the pH electrode, change of the reference membrane, as well as the contacts of the preamps. Straighten the suction probe for air detection problems during suction. Cleaning the measuring chamber. Various calibrations to stabilize pH readings. Review and check the equipment with QC's. | Cleaning barcode reader and adjusting it. Cleaning elctrot of pH. | Cleaned | HW | Basic Unit | Barcode Reader, Sampler
ABL835 FLEX | ｐH電極洗浄 | B | SWO-114812 | No | pH electrode cleaning |  | Cleaned | Electrode | Electrode | pH
ABL825 FLEX | Kalibration und Reinigung FlexQ Modul - Funktionsüberprüfung etc. | Kalibration und Reinigung FlexQ Modul - Funktionsüberprüfung etc. | SWO-114750 | No | Calibration and cleaning FlexQ module - function check etc. | Calibration and cleaning FlexQ module - function check etc. | Cleaned | HW | FlexQ Module | Mixer | 882
ABL825 FLEX | checked power supply, | removed DMS and replaced power supply and hard drive | D100D15832256930 | No | checked power supply, | removed DMS and replaced power supply and hard drive | Defect | HW | DMS_PC | Power supply
ABL825 FLEX | error 751, 657 and 510, Will perform | replacec the pH black valve and performed a annual maintenance; replaced V8 and V6; inspected the valves in the Met module and found air valve with a small rip--leaking out the front replaced inlet nipples | D10083853262FAC3 | No | error 751; 657 and 510; Will perform | replacec the pH black valve and performed a annual maintenance; replaced V8 and V6; inspected the valves in the Met module and found air valve with a small rip--leaking out the front replaced inlet nipples | Cleaned | HW | El/Met Module | Valve
ABL800 BASIC | CALIBRATED THE TOUCH SCREEN NOW ITS WORKING BUT SOME TIMES ITS NOT WORKING ITS NEED TO BE REPLACE WITH NEW ONE. | B | SWO-114821 | No | CALIBRATED THE TOUCH SCREEN NOW ITS WORKING BUT SOME TIMES ITS NOT WORKING ITS NEED TO BE REPLACE WITH NEW ONE. |  | Defect | HW | DMS_PC | Touch Screen
ABL825 FLEX | noted tubing nipple was mared | i replaced inlet tubing nipple | D10056E532079765 | No | noted tubing nipple was mared | i replaced inlet tubing nipple | Defect | HW | Inlet module | Leaking | Nipple kit
ABL825 FLEX | mechanische Verbindungen Lampe Photometer geprüft, Lampeneinstellungen angesehen, Lampe komplett neu justiert, Intensität erhöht, Provokationstest durchgeführt | mechanische Verbindungen Lampe Photometer geprüft, Lampeneinstellungen angesehen, Lampe komplett neu justiert, Intensität erhöht, Provokationstest durchgeführt | SWO-114761 | No | mechanical connections lamp photometer tested, lamp settings viewed, lamp completely re-adjusted, intensity increased, provocation test performed | mechanical connections lamp photometer tested, lamp settings viewed, lamp completely re-adjusted, intensity increased, provocation test performed | Adjustment | HW | Oxi module | Lamp unit
ABL825 FLEX | Flüssigkeit wurde bereits durch Anwender abgewischt. Einlasskalibration im Servicemenü, OK. Fließwege sichtkontrolliert, OK. Spülungen, kein weiterer Fl. Austritt feststellbar. | Flüssigkeit wurde bereits durch Anwender abgewischt. Einlasskalibration im Servicemenü, OK. Fließwege sichtkontrolliert, OK. Spülungen, kein weiterer Fl. Austritt feststellbar. | SWO-114695 | No | Liquid has already been wiped by users. Inlet calibration in the service menu, OK. Flow paths visually controlled, OK. Rinses, no further fl. exit detectable. | Liquid has already been wiped by users. Inlet calibration in the service menu, OK. Flow paths visually controlled, OK. Rinses, no further fl. exit detectable. | Defect | HW | Inlet module | Probe cal.
AQT90 FLEX | Checked and cleaned cup sensor, re-ran cal adjustment. | Checked and cleaned cup sensor, re-ran cal adjustment. | SWO-114944 | No | Checked and cleaned cup sensor, re-ran cal adjustment. | Checked and cleaned cup sensor, re-ran cal adjustment. | Cleaned | HW | Incubator module | Sensor, Cup
ABL80 FLEX CO-OX | Hemolyzer ausgetauscht. Gerät kalibriert und überprüft. Wieder zurück zum Kunden gebracht. | Hemolyzer ausgetauscht. Gerät kalibriert und überprüft. Wieder zurück zum Kunden gebracht. | SWO-114914 | No | Hemolyzer. calibrated and checked. Brought back to the customer. | Hemolyzer. calibrated and checked. Brought back to the customer.
ABL80 FLEX BASIC | 更换配件902-972修复仪器(by andrew) | 更换配件902-972修复仪器(by andrew) | SWO-114257 | No | Replacement accessory 902-972 repair instrument (by andrew) | Replacement accessory 902-972 repair instrument (by andrew)
ABL80 FLEX BASIC | Cleaned SC with distilled water and ten kept in fridge for 30 minutes. Reinstalled Sc but again failed to calibrate. Need to replace Sc. | Cleaned SC with distilled water and ten kept in fridge for 30 minutes. Reinstalled Sc but again failed to calibrate. Need to replace Sc. | SWO-115341 | No | Cleaned SC with distilled water and ten kept in fridge for 30 minutes. Reinstalled Sc but again failed to calibrate. Need to replace Sc. | Cleaned SC with distilled water and ten kept in fridge for 30 minutes. Reinstalled Sc but again failed to calibrate. Need to replace Sc.
ABL80 FLEX BASIC | 更换902-971一个正常。(by andrew) | 更换902-971一个正常。(by andrew) | SWO-114544 | No | Replace the 902-971 one normal. (by andrew) | Replace the 902-971 one normal. (by andrew)
ABL90 FLEX | De batterij (P0064N026) werd vervangen door een nieuwe ( P0120N016) : test OK | B | SWO-114942 | No | The battery (P0064N026) was replaced by a new one (P0120N016): Test OK |  | Defect | HW | Battery module | Battery module
ABL90 FLEX | De batterij laadt niet meer op. Deze werd gereset maar gaf geen beterschap. Hierop de batterij (P0061N012) vervangen door een nieuwe (P0120N014). Test OK | B | SWO-114925 | No | The battery is not charging anymore, it was reset but did not give a well. Replace the battery (P0061N012) with a new one (P0120N014). Test OK |  | Defect | HW | Battery module | Battery module
ABL90 FLEX | Lösungspack Defekt. Reklamationsbogen ausgedruckt und übermittelt. Kalibriert und QC's durchgeführt. | Gerätestatus mit Kunden besprochen. Kunde wünscht einen Vor-Ort Service. | SWO-115124 | No | Solution pack defect. Complaint form printed out and transmitted. Calibrated and QC's performed. | Device status discussed with customers. Customer wants an on-site service. | Assisted Replacement | SP | SP | FT
ABL90 FLEX | IP überprüft und Firewalleinstellungen korrigiert, Einstellungen gesichert, Neustart, QC | B | SWO-114955 | No | IP checked and firewall settings corrected, settings secured, restart, QC |  | Setup | His/lis | His/lis | His/lis
ABL90 | Customer reports that analyzer froze and would not allow logins or replacements. When attempting to replace the SP, the number of cycles showed 0. Upon arrival, the SP was observed to not be fully inserted. | Freed the analyzer for SP acceptance and inserted the SP in service mode. The chip was properly read. Successfully reinstalled SP Analyzer calibration failed. Replaced 2 sensor cassettes and all parameters calibrated properly. Credit forms have been submitted for the customer. Analyzer is within acceptable operating specifications. | D101DB90A837F109 | No | Customer reports that analyzer froze and would not allow logins or replacements. When attempting to replace the SP, the number of cycles showed 0. Upon arrival, the SP was observed to not be fully inserted. | Freed the analyzer for SP acceptance and inserted the SP in service mode. The chip was properly read. Successfully reinstalled SP Analyzer calibration failed. Replaced 2 sensor cassettes and all parameters calibrated properly. Credit forms have been submitted for the customer. Analyzer is within acceptable operating specifications. | User error | SP Installation | SP Installation | SP Installation
ABL90 FLEX | készülék újra indítása, oldat újra telepítése, az installálás után az inicializálás rendben zajlik, az inicializálás után az oldat telepítés megismétlése. Kalibrációk, QC mérés | B | SWO-114948 | No | Rebooting the device, re-installing the solution, after installation The initialization is done in order, after initialization, the solution is repeated. Calibrations, QC measurement |  | Adjustment | SW | SW | Off/On
ABL90 FLEX PLUS |  | B | D1017E5CA81A9861 | No |  |  | Assisted Replacement | SP | SP | Leak current
ABL835 FLEX | A new measuring chamber was fitted. The following tubes were replaced, blue inlet tube, blue tube between the inlet and the pH BG module, both grey tubes linking the measuring modules and the Oxi module, the small grey tube above the pH BG module connecting to the preheated. Lactate chamber and electrode tip cleaned. | A new measuring chamber was fitted. The following tubes were replaced, blue inlet tube, blue tube between the inlet and the pH BG module, both grey tubes linking the measuring modules and the Oxi module, the small grey tube above the pH BG module connecting to the preheated. Lactate chamber and electrode tip cleaned. | SWO-115040 | No | A new measuring chamber was fitted. The following tubes were replaced, blue inlet tube, blue tube between the inlet and the pH BG module, both grey tubes linking the measuring modules and the Oxi module, the small grey tube above the pH BG module connecting to the preheated. Lactate chamber and electrode tip cleaned. | A new measuring chamber was fitted. The following tubes were replaced, blue inlet tube, blue tube between the inlet and the pH BG module, both grey tubes linking the measuring modules and the Oxi module, the small grey tube above the pH BG module connecting to the preheated. Lactate chamber and electrode tip cleaned. | Assisted Replacement | Electrode | Electrode | pCO2
ABL825 FLEX | Unit is in a blue screen condition. Unable to check logs. | Replaced hard drive and CPU. Installed Win 7 and 6.16 software. Transferring data base resulted in Status code 2 error. Corrected error. Unable to load extracted data base. Reinstalled setup and QC. | D101C9C0A83C6558 | No | Unit is in a blue screen condition. Unable to check logs. | Replaced hard drive and CPU. Installed Win 7 and 6.16 software. Transferring data base resulted in Status code 2 error. Corrected error. Unable to load extracted data base. Reinstalled setup and QC. | Defect | HW | DMS_PC | HDD
ABL800 BASIC | need to replace the lactate membrane. after changing the membrane the analyzer working satisfactorily. downloaded the activity and calibration log for the analysis. | need to replace the lactate membrane. after changing the membrane the analyzer working satisfactorily. downloaded the activity and calibration log for the analysis. | SWO-115368 | No | need to replace the lactate membrane. after changing the membrane the analyzer working satisfactorily. downloaded the activity and calibration log for the analysis. | need to replace the lactate membrane. after changing the membrane the analyzer working satisfactorily. downloaded the activity and calibration log for the analysis. | Assisted Replacement | Electrode | Membrane | Lac
ABL810 FLEX | neue Charge Membran verwendet, Luftdruck korrigiert, Dichtung pH unten ersetzt, mehrere Kal, QC | neu membraniert, pCO2 Elektrode ersetzt, Inserts korrigiert, Ventilmembran ersetzt, Schlauch ersetzt, mehrere Kal, QC nach wie vor zu niedrig - abgebrochen | SWO-114957 | No | new batch membrane used, air pressure corrected, seal pH replaced below, several Kal, QC | new membrane, pCO2 electrode replaced, inserts corrected, valve membrane replaced, hose replaced, several kal, QC still too low - aborted | Assisted Replacement | Electrode | Electrode | pCO2
AQT90 FLEX | Reinigung des Barcode Gehäuses und Überpfung des Barcodes. Barcode ist mittig eingestellt. Überprüfung aller Baugruppen, o. K., Funktion o. K. Datalogs gezogen und Service Dump gezogen. | Reinigung des Barcode Gehäuses und Überpfung des Barcodes. Barcode ist mittig eingestellt. Überprüfung aller Baugruppen, o. K., Funktion o. K. Datalogs gezogen und Service Dump gezogen. | SWO-115056 | No | Cleaning the barcode housing and padding the barcode. Barcode is set in the middle. Check all assemblies, o. K., function o. K. Datalogs dragged and service dump dragged. | Cleaning the barcode housing and padding the barcode. Barcode is set in the middle. Check all assemblies, o. K., function o. K. Datalogs dragged and service dump dragged. | Defect | SW | Analyzer Control Module | Off/On
AQT90 FLEX | Mit Firma H+S Kontakt aufgenommen und Korrektur der Schnittstelle vorgenommen. Mehrere Telefongespräche durchgeführt. Korrektur am AQT durchgeführt. Werte werden, wie Anwender gewünscht angezeigt. Anwender kann wie gewünscht Werte aus Auswerte Software freigeben. | Mit Firma H+S Kontakt aufgenommen und Korrektur der Schnittstelle vorgenommen. Mehrere Telefongespräche durchgeführt. Korrektur am AQT durchgeführt. Werte werden, wie Anwender gewünscht angezeigt. Anwender kann wie gewünscht Werte aus Auswerte Software freigeben. | SWO-115055 | No | Contacted company H+S and corrected the interface. Multiple phone calls made. correction performed on the AQT. Values are displayed as desired by the user. User can release values from evaluation software as desired. | Contacted company H+S and corrected the interface. Multiple phone calls made. correction performed on the AQT. Values are displayed as desired by the user. User can release values from evaluation software as desired. | Setup | SW | His/lis | His/Lis
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-115210 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL80 FLEX BASIC |  | B | SWO-115093 | No
ABL90 FLEX | Keine Kalibrierung möglich, Display reagiert ungleichmäßig auf Berührung Display komplett erneuert ( gebraucht ) | Keine Kalibrierung möglich, Display reagiert ungleichmäßig auf Berührung Display komplett erneuert ( gebraucht ) | SWO-115308 | No | No calibration possible, display reacts unevenly to touch display completely renewed (used ) | No calibration possible, display reacts unevenly to touch display completely renewed (used ) | Defect | HW | DMS_PC | touch Screen
ABL90 | USB NOT SHOWING ON DESTINATION DRIVE | CUSTOMER WAS USING A WIRELESS KEYBOARD.IT IS NOT RECOMMENDED TO USE ANY WIRELESS DEVICE AS IT CAN CAUSE EXCESSIVE EMI.REBOOTED ANALYZER SEVERAL TIMES BUT DRIVE ONLY SHOWEDAFTER USING THE SOFT SHUTDOWN BUTTON.IF PROBLEM REAPPEARS WE MAY NEED TO REPLACE THE CPU OR WINDOWS SOFTWARE.S | D10323C31E2A5374 | No | USB NOT SHOWING ON DESTINATION DRIVE | CUSTOMER WAS USING A WIRELESS KEYBOARD.IT IS NOT RECOMMENDED TO USE ANY WIRELESS DEVICE AS IT CAN CAUSE EXCESSIVE EMI. REBOOTED ANALYZER SEVERAL TIMES BUT DRIVE ONLY SHOWEDAFTER USING THE SOFT SHUTDOWN BUTTON. IF PROBLEM REAPPEARS WE MAY NEED TO REPLACE THE CPU OR WINDOWS SOFTWARE. S | Unknown | Ext HW | USB | USB not detected
ABL90 |  | B | D102BD6D1E5F4264 | Yes |  |  | Re classified RMED | Phione support
ABL90 FLEX PLUS | Analysator gespült - tHb KAL - Serviceprogramme aufgerufen - Oximetrie Test ohne Beanstandung - Erneuerung Tubing for Valve (Vakuumfehler) - Servicetests ohne Beanstandung - Funktionsüberprüfung etc. | Analysator gespült - tHb KAL - Serviceprogramme aufgerufen - Oximetrie Test ohne Beanstandung - Erneuerung Tubing for Valve (Vakuumfehler) - Servicetests ohne Beanstandung - Funktionsüberprüfung etc. | SWO-115329 | No | Analyzer flushed - tHb KAL - service programs called - Oximetry test without complaint - Renewal Tubing for Valve (vacuum error) - Service tests without complaint - Function check etc. | Analyzer flushed - tHb KAL - service programs called - Oximetry test without complaint - Renewal Tubing for Valve (vacuum error) - Service tests without complaint - Function check etc. | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
ABL90 FLEX PLUS | Fehlersuche, LP gew., Initialisierung abgewartet, Kal. + QC i.O. | Fehlersuche; LP gew.; Initialisierung abgewartet; Kal. + QC i.O. | SWO-115180 | No | Troubleshooting; LP. initialization awaited; Kal. + QC in the first place | Troubleshooting; LP. initialization awaited; Kal. + QC in the first place | Assisted Replacement | SP | SP | pO2
ABL90 FLEX PLUS | Anwenderfehler -> Einlaßhalter u. Nadel gew., Startphase abgewartet, Einlaß kalibriert, Kal. + QC i.O. | Anwenderfehler -> Einlaßhalter u. Nadel gew.; Startphase abgewartet; Einlaß kalibriert; Kal. + QC i.O. | SWO-115187 | No | User error -> Inlet holder and needle; start-up phase awaited; Inlet calibrated; Kal. + QC in the first place | User error -> Inlet holder and needle; start-up phase awaited; Inlet calibrated; Kal. + QC in the first place | Replaced | HW | Auto Inlet module | Probe bend + Gasket
AQT90 FLEX |  | B | SWO-115083 | No |  |  | Defect | HW | Process Module | Optical Unit | 125
AQT90 FLEX | 更换了进样针后解决(by zhenjie) | 更换了进样针后解决(by zhenjie) | SWO-115509 | No | Post-replaceable feeding needle (by zhenjie) | Post-replaceable feeding needle (by zhenjie) | Defect | HW | Aspiration Module | Bend needle | 4794
ABL800 BASIC | Replace the Lac membrane with following procedure and run cleaning and recalibrate,Now its sensitivity is OK | Replace the Lac membrane with following procedure and run cleaning and recalibrate,Now its sensitivity is OK | SWO-115386 | No | Replace the Lac membrane with following procedure and run cleaning and recalibrate,Now its sensitivity is OK | Replace the Lac membrane with following procedure and run cleaning and recalibrate,Now its sensitivity is OK | Assisted Replacement | Electrode | Membrane | Lac
ABL820 FLEX | モニタ交換 CPUユニット交換 | B | SWO-115429 | No | Monitor replacement CPU unit replacement |  | Defect | HW | DMS_PC | CPU | (Touch Screen)
ABL800 BASIC | Has to change the printer head. Part Number: 910-318 printer head | B | SWO-115357 | No | Has to change the printer head. Part Number: 910-318 printer head |  | Defect | HW | Printer module | Printer Head
ABL835 FLEX |  | B | D102DBA71E285B5E | No |  |  | Assisted Replacement | Electrode | Electrode | pCO2
ABL800 BASIC | PERFORMED COMPLETE OXI TEST AND CONFORMED LAMP UNIT FAILURE.AND PERFORMED tHB CALIBRATION AND CONFORMED NEON POSITION INCORRECT WE NEED TO REPLACE OXI MODULE. | B | SWO-115385 | No | PERFORMED COMPLETE OXI TEST AND CONFORMED LAMP UNIT FAILURE. AND PERFORMED tHB CALIBRATION AND CONFORMED NEON POSITION INCORRECT WE NEED TO REPLACE OXI MODULE. |  | Revisit required | HW | Oxi module | Oxi module
ABL815 FLEX | Med.techniker hat Pos. der Sampler schon kalibriert, geprüft i.O., an FlexQ Tray ist Nase abgebrochen ->ersetzt, Samplersensor kal im Serviceprogramm durchgeführt, Barcodescanner für Sampler im Serviceprog. kalibriert, Patientenprobemessung in Pos1,2,3 durchgeführt fehlerfrei | Med.techniker hat Pos. der Sampler schon kalibriert, geprüft i.O.; an FlexQ Tray ist Nase abgebrochen ->ersetzt, Samplersensor kal im Serviceprogramm durchgeführt, Barcodescanner für Sampler im Serviceprog. kalibriert; Patientenprobemessung in Pos1,2,3 durchgeführt fehlerfrei | SWO-115404 | No | Med.techniker has Pos. the sampler is already calibrated, checked usually; on FlexQ Tray nose is broken ->replaced, sampler sensor kal performed in the service program, barcode scanner for sampler in service program. calibrated; Patient sample measurement performed in Pos1,2,3 error-free | Med.techniker has Pos. the sampler is already calibrated, checked usually; on FlexQ Tray nose is broken ->replaced, sampler sensor kal performed in the service program, barcode scanner for sampler in service program. calibrated; Patient sample measurement performed in Pos1,2,3 error-free | Defect | HW | FlexQ Module | Mixer tray
ABL810 FLEX | Analysator Konfiguration gesichert für Leihgerät. Analysator transportfähig gemacht. Analysator abgezogen und zur Reparatur in Zentrale verschickt. | Analysator Konfiguration gesichert für Leihgerät. Analysator transportfähig gemacht. Analysator abgezogen und zur Reparatur in Zentrale verschickt. | SWO-115337 | No | Analyzer configuration secured for loaner. Analyzer made transportable. analyzer and shipped to headquarters for repair. | Analyzer configuration secured for loaner. Analyzer made transportable. analyzer and shipped to headquarters for repair. | Revisit required | Analyzer replaced | Membrane | pCO2 | 971
ABL800 BASIC | 更换Lac电极膜后修复 | 更换Lac电极膜后修复 | SWO-115141 | No | Repair after replacing Lac electrode film | Repair after replacing Lac electrode film | Assisted Replacement | Electrode | Membrane | Lac | Sensitivity
ABL835 FLEX | Oxi Module Pcb ausgetauscht. tHb Kal in Seviceprogramm. Lampeneinheit justiert. | Oxi Module Pcb ausgetauscht. tHb Kal in Seviceprogramm. Lampeneinheit justiert. | SWO-115317 | No | Oxi Module Pcb. tHb Kal in Sevice program. lamp unit. | Oxi Module Pcb. tHb Kal in Sevice program. lamp unit. | Defect | HW | Oxi module | PCB, Oxi
ABL820 FLEX | Ref,CO2メンブラン交換。チャンバー内掃除、エア回路掃除。圧漏れ確認。 | Ref,CO2メンブラン交換。チャンバー内掃除、エア回路掃除。圧漏れ確認。 | SWO-115322 | No | Ref, CO2 membrane replacement. Chamber cleaning, air circuit cleaning. Press leak confirmation. | Ref, CO2 membrane replacement. Chamber cleaning, air circuit cleaning. Press leak confirmation. | Assisted Replacement | Electrode | Membrane | pCO2
ABL825 FLEX | Die Lac Empf. ist nach wenigen Tagen unterhalb Grenze gerutscht => Elektrode membraniert, ServiceDump und Datenextrakt gezogen und an TS geschickt zwecks Eröffnung einer TechLine. Kalibration, OK. | Die Lac Empf. ist nach wenigen Tagen unterhalb Grenze gerutscht => Elektrode membraniert, ServiceDump und Datenextrakt gezogen und an TS geschickt zwecks Eröffnung einer TechLine. Kalibration, OK. | SWO-115267 | No | The Lac Empf. has slipped below the limit after a few days => electrode membraneed, ServiceDump and data extract pulled and sent to TS to open a TechLine. Calibration, OK. | The Lac Empf. has slipped below the limit after a few days => electrode membraneed, ServiceDump and data extract pulled and sent to TS to open a TechLine. Calibration, OK. | Assisted Replacement | Electrode | Membrane | Lac | 378
ABL825 FLEX | Erneut ist die Lac Empf. unterhalb Grenze gerutscht => Elektrode membraniert, ServiceDump und Datenextrakt gezogen und an TS geschickt zwecks Eröffnung einer TechLine. Kalibration, OK. | Erneut ist die Lac Empf. unterhalb Grenze gerutscht => Elektrode membraniert, ServiceDump und Datenextrakt gezogen und an TS geschickt zwecks Eröffnung einer TechLine. Kalibration, OK. | SWO-115160 | No | Once again, the Lac Empf. slipped below limit => electrode membraneed, ServiceDump and data extract pulled and sent to TS to open a TechLine. Calibration, OK. | Once again, the Lac Empf. slipped below limit => electrode membraneed, ServiceDump and data extract pulled and sent to TS to open a TechLine. Calibration, OK. | Assisted Replacement | Electrode | Membrane | Lac | 378
ABL825 FLEX | Mischerträgerantrieb überprüft, Motor leicht unzentriert (eiert) aber Antrieb ingesamt sehr leichtgängig. Madenschrauben des Linearantriebes nachgezogen, Funktion und Positionierung des Mischerträgers im Servicemenü überprüft, OK. | Mischerträgerantrieb überprüft, Motor leicht unzentriert (eiert) aber Antrieb ingesamt sehr leichtgängig. Madenschrauben des Linearantriebes nachgezogen, Funktion und Positionierung des Mischerträgers im Servicemenü überprüft, OK. | SWO-115154 | No | Mixer carrier drive checked, motor slightly uncentered (etherated) but drive overall very smooth. Maden screws of the linear drive, function and positioning of the mixer carrier checked in the service menu, OK. | Mixer carrier drive checked, motor slightly uncentered (etherated) but drive overall very smooth. Maden screws of the linear drive, function and positioning of the mixer carrier checked in the service menu, OK. | Adjustment | HW | FlexQ Module | Mixer tray
ABL825 FLEX | Die Lac Empf. ist nach wenigen Tagen unterhalb Grenze gerutscht => Elektrode membraniert, ServiceDump und Datenextrakt gezogen und an TS geschickt zwecks Eröffnung einer TechLine. Kalibration, OK. | Die Lac Empf. ist nach wenigen Tagen unterhalb Grenze gerutscht => Elektrode membraniert, ServiceDump und Datenextrakt gezogen und an TS geschickt zwecks Eröffnung einer TechLine. Kalibration, OK. | SWO-115272 | No | The Lac Empf. has slipped below the limit after a few days => electrode membraneed, ServiceDump and data extract pulled and sent to TS to open a TechLine. Calibration, OK. | The Lac Empf. has slipped below the limit after a few days => electrode membraneed, ServiceDump and data extract pulled and sent to TS to open a TechLine. Calibration, OK. | Assisted Replacement | Electrode | Membrane | Lac | 378
ABL825 FLEX | Die Lac Empf. ist nach wenigen Tagen unterhalb Grenze gerutscht => Elektrode membraniert, ServiceDump und Datenextrakt gezogen und an TS geschickt zwecks Eröffnung einer TechLine. Kalibration, OK. | Die Lac Empf. ist nach wenigen Tagen unterhalb Grenze gerutscht => Elektrode membraniert, ServiceDump und Datenextrakt gezogen und an TS geschickt zwecks Eröffnung einer TechLine. Kalibration, OK. | SWO-115264 | No | The Lac Empf. has slipped below the limit after a few days => electrode membraneed, ServiceDump and data extract pulled and sent to TS to open a TechLine. Calibration, OK. | The Lac Empf. has slipped below the limit after a few days => electrode membraneed, ServiceDump and data extract pulled and sent to TS to open a TechLine. Calibration, OK. | Assisted Replacement | Electrode | Membrane | Lac | 378
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-116535 | No | Repair after replacing a new test card | Repair after replacing a new test card
AQT90 FLEX | 更换903-364一个正常。 | 更换903-364一个正常。 | SWO-115174 | No | Replace 903-364 one normal. | Replace 903-364 one normal. | Defect | HW | Process Module | Optical Unit | 131
ABL80 FLEX BASIC | 更换了接口板902-972后修复(by zhenjie) | 更换了接口板902-972后修复(by zhenjie) | SWO-115383 | No | Repair after replacing interface board 902-972 (by zhenjie) | Repair after replacing interface board 902-972 (by zhenjie)
AQT90 FLEX |  | B | SWO-115412 | No |  |  | Defect | HW | Process Module | Optical Unit | 131
AQT90 FLEX | 更换了910-343后修复。903-175未更换。(by zhenjie) | 更换了910-343后修复。903-175未更换。(by zhenjie) | SWO-115380 | No | Repair after replacing 910-343. 903-175 has not been replaced. (by zhenjie) | Repair after replacing 910-343. 903-175 has not been replaced. (by zhenjie) | Defect | HW | Power supply | Power supply
AQT90 FLEX | Performed System Cleaning, Problem Resolve. | Performed System Cleaning, Problem Resolve. | SWO-116757 | No | Performed System Cleaning, Problem Resolve. | Performed System Cleaning, Problem Resolve. | Cleaned | HW | Aspiration Module | HCT Sensor
ABL80 FLEX BASIC | 更换新测试卡修复仪器 | 更换新测试卡修复仪器 | SWO-117783 | No | Replace the new test card repair instrument | Replace the new test card repair instrument
ABL80 FLEX BASIC | 更换新测试卡修复仪器 | 更换新测试卡修复仪器 | SWO-117777 | No | Replace the new test card repair instrument | Replace the new test card repair instrument
ABL80 FLEX BASIC |  | B | SWO-115572 | No
ABL80 FLEX BASIC | replaced the new casstte analyser is working fine. | replaced the new casstte analyser is working fine. | SWO-116773 | No | replaced the new casstte analyser is working fine. | replaced the new casstte analyser is working fine.
ABL90 FLEX PLUS | Ventil im Service überprüft - in Ordnung. Undichtigkeit des Abfallschlauches am Waste Connector festgestellt. Ventilschlauch und Abfallschlauch erneuert. SC erneuert. Kalibrationen und QC's vorgenommen. Kunden eine SC für Lager überlassen, da keine mehr vorhanden war. | Ventil im Service überprüft - in Ordnung. Undichtigkeit des Abfallschlauches am Waste Connector festgestellt. Ventilschlauch und Abfallschlauch erneuert. SC erneuert. Kalibrationen und QC's vorgenommen. Kunden eine SC für Lager überlassen, da keine mehr vorhanden war. | SWO-115983 | No | Valve checked in service - all right. leakage of the waste hose on the Waste Connector. valve hose and waste hose. SC renewed. calibrations and QC's. give customers a SC for warehouses because there was no more. | Valve checked in service - all right. leakage of the waste hose on the Waste Connector. valve hose and waste hose. SC renewed. calibrations and QC's. give customers a SC for warehouses because there was no more. | Defect | HW | Waste tube | Waste tube
ABL80 CO-OX |  | Andy performed PM and Testing | D106A2268058A54A | No |  | Andy performed PM and Testing
ABL90 FLEX | 连续更换两张测试卡后解决 | 连续更换两张测试卡后解决 | SWO-115731 | No | Resolve after replacing two test cards in a row | Resolve after replacing two test cards in a row | Assisted Replacement | SC | SC | pCO2 | 378
AQT90 FLEX |  | B | SWO-115570 | No |  |  | Defect | HW | Process Module | Stepper motor, process wheel
ABL80 CO-OX | ERRORES DE VERIFICACION DE SP Y DE CALIBRACIONES IMPIDIENDO EL FUNCIONAMIENTO DEL MISMO, FUGA DE LIQUIDOS DEBIDO A OBSTRUCCION DEL CIRCUITO DE DESECHOS, SE LIMPIA BIEN TODA LA LINEA DE DESECHO CON DESCONTAMINACION Y, CLAIBRACION SE, DOTA DE NUEVOS REACTIVOS. PRUEBAS DE FUNCIONAMIENTO Y AJUSTE DE CAL htb. CALIBRADO Y VERIFICADO CON Q.C. | B | D106BE33805D5543 | No | SP'S VERIFICATION AND CALIBRATION ERRORS BY IMPARTING THE OPERATION OF THE SAME, LEAKING LIQUIDS DUE TO OBSTRUCTION OF THE DISCLOSURE CIRCUIT, CLEANING ALL THE DISCLOSURE LINE WITH DECONTAMINATION AND, OPERATION TESTS AND ADJUSTMENT OF CAL htb. CALIBRATED AND VERIFIED WITH Q.C.
ABL825 FLEX |  | B | D106DA24804C9D83 | No |  |  | Defect | HW | AutoCheck Module | AutoCheck Module
ABL825 FLEX | Checked halogen lamp voltage and found it to be reading above the upper limit. | Adjusted the halogen lamp voltage. Updated software to version 6.16. | D106B6CD804B7657 | No | Checked halogen lamp voltage and found it to be reading above the upper limit. | Adjusted the halogen lamp voltage. Updated software to version 6.16. | Adjustment | HW | Oxi module | Lamp unit
ABL835 FLEX |  | B | D111ECF2A63AF4FB | No |  |  | Assisted Replacement | Electrode | Membrane | Lac | Sensitivity
ABL800 BASIC | Performed complete oximetry test, where neon intensity and neon number was found out of range after which kept the machine OFF for 30 minutes. After restarting the analyzer, machine is still displaying the same error with all the oximetry parameters disabled. Need to replace the oximetry PCB or Spectrometer. | Performed complete oximetry test, where neon intensity and neon number was found out of range after which kept the machine OFF for 30 minutes. After restarting the analyzer, machine is still displaying the same error with all the oximetry parameters disabled. Need to replace the oximetry PCB or Spectrometer. | SWO-115550 | No | Performed complete oximetry test, where neon intensity and neon number was found out of range after which kept the machine OFF for 30 minutes. After restarting the analyzer, machine is still displaying the same error with all the oximetry parameters disabled. Need to replace the oximetry PCB or Spectrometer. | Performed complete oximetry test, where neon intensity and neon number was found out of range after which kept the machine OFF for 30 minutes. After restarting the analyzer, machine is still displaying the same error with all the oximetry parameters disabled. Need to replace the oximetry PCB or Spectrometer. | Revisit required | HW | Oxi module | Oxi module
ABL800 BASIC | clean inlet gasket and inlet probe.perform 2-point calibration now machine is ok. | clean inlet gasket and inlet probe.perform 2-point calibration now machine is ok. | SWO-116761 | No | clean inlet gasket and inlet probe.perform 2-point calibration now machine is ok. | clean inlet gasket and inlet probe.perform 2-point calibration now machine is ok. | Cleaned | HW | Inlet module | Inlet gasket | Leaking
ABL835 FLEX |  | B | D1094F476C095B69 | Yes |  |  | Re classified RMED | Setup
ABL815 FLEX | Schlauch ersetzt, Abfallpumpenschläuche ersetzt, Start abgewartet, Reinigung, LS Justage, Pumpenkal, Kal | B | SWO-115659 | No | Hose replaced, waste pump hoses replaced, start waited, cleaning, LS adjustment, pump scale, kal |  | Defect | HW | Basic Unit | Waste tube | Clot
ABL825 FLEX | Gerinnsel entfernt, Proteinentfernung mehrfach, Einlassdichtung ersetzt, Messkammern überprüft, LS-Justage Pumpenkal. Nach mehreren Kal immer Instabilität bei Ca => Ca und Referenz membraniert. Kalibrationen, Endkontrolle: QC Messung OK | Gerinnsel entfernt, Proteinentfernung mehrfach, Einlassdichtung ersetzt, Messkammern überprüft, LS-Justage Pumpenkal. Nach mehreren Kal immer Instabilität bei Ca => Ca und Referenz membraniert. Kalibrationen, Endkontrolle: QC Messung OK | SWO-115667 | No | Clots removed, protein removal multiple times, inlet seal replaced, measuring chambers checked, LS-Adjustment pump scale. After several scales always instability at Ca => Ca and reference membraneed. Calibrations, final inspection: QC measurement OK | Clots removed, protein removal multiple times, inlet seal replaced, measuring chambers checked, LS-Adjustment pump scale. After several scales always instability at Ca => Ca and reference membraneed. Calibrations, final inspection: QC measurement OK | Cleaned | Fluidic pathway | El/Met Module | Measuring chamber
ABL825 FLEX | Lac membraniert. ServiceDump und Datenextrakt gezogen zwecks Übersendung an RDEU TS. | Lac membraniert. ServiceDump und Datenextrakt gezogen zwecks Übersendung an RDEU TS. | SWO-115663 | No | Lac membranes. ServiceDump and data extract drawn for sending to RDEU TS. | Lac membranes. ServiceDump and data extract drawn for sending to RDEU TS. | Assisted Replacement | Electrode | Membrane | Lac
ABL80 FLEX BASIC | 更换新测试卡修复仪器 | 更换新测试卡修复仪器 | SWO-116129 | No | Replace the new test card repair instrument | Replace the new test card repair instrument
ABL80 FLEX BASIC | Replace the SP and try again install the SC. Now all the parameters are active. There is problem with the SP. After replace the SP unit is working properly. | Replace the SP and try again install the SC. Now all the parameters are active. There is problem with the SP. After replace the SP unit is working properly. | SWO-116783 | No | Replace the SP and try again install the SC. Now all the parameters are active. There is problem with the SP. After replace the SP unit is working properly. | Replace the SP and try again install the SC. Now all the parameters are active. There is problem with the SP. After replace the SP unit is working properly.
ABL80 CO-OX | PCO2 out high on internal QC C8302. | Replaced SP from my inventory. Ran Calibrations. All internal QC in in. Can perform Linearity Study now. | D10990BE6C49D5B4 | No | PCO2 out high on internal QC C8302. | Replaced SP from my inventory. Ran Calibrations. All internal QC in in. Can perform Linearity Study now.
ABL90 FLEX | Gerät überprüft.SC entfernt.keine Lös erkennbar.Fluidic Wege gespült. Einlassdichtung ersetzt.(Verschleiß) Visual Flowceck mit neuem SP erfolgreich. | Gerät überprüft.SC entfernt.keine Lös erkennbar.Fluidic Wege gespült. Einlassdichtung ersetzt.(Verschleiß) Visual Flowceck mit neuem SP erfolgreich. | SWO-115855 | No | device checked. SC removes.No loes recognizable. Fluidic ways rinsed. replaced inlet seal. (Wear) Visual Flowceck successful with new SP. | device checked. SC removes.No loes recognizable. Fluidic ways rinsed. replaced inlet seal. (Wear) Visual Flowceck successful with new SP. | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
ABL90 | checked data logs, inspected wet section, inspected inlet gasket holder and connector gasket. | Replaced inlet gasket holder and inlet connector gasket (customers supply). Performed pressure test, liquid sensor adjust, pump calibration. | D107F25BF656C1CD | No | checked data logs, inspected wet section, inspected inlet gasket holder and connector gasket. | Replaced inlet gasket holder and inlet connector gasket (customers supply). Performed pressure test, liquid sensor adjust, pump calibration. | Defect | HW | Inlet module | Inlet gasket
ABL90 FLEX | Mit Anwender mehrere Telefongespräche geführt. Vorgehensweise bei pCO2 erklärt. RMED hat neue Sensorkassetten mit verbesserter Empfindlichkeit hergestellt. Anwender wird die neuen SC bestellen. Raum des Analysators ist jetzt im Sommer sehr warm, Fenster auf und nicht klimatisiert. Lösungspack sollte auch gekühlt werden. Kalibrationen und Qualitätskontrolle o. K. | Mit Anwender mehrere Telefongespräche geführt. Vorgehensweise bei pCO2 erklärt. RMED hat neue Sensorkassetten mit verbesserter Empfindlichkeit hergestellt. Anwender wird die neuen SC bestellen. Raum des Analysators ist jetzt im Sommer sehr warm, Fenster auf und nicht klimatisiert. Lösungspack sollte auch gekühlt werden. Kalibrationen und Qualitätskontrolle o. K. | SWO-115924 | No | With user conducted several phone calls. procedure with pCO2. RMED has produced new sensor cartridges with improved sensitivity. User will order the new SC. Room of the analyzer is now very warm in summer, windows on and not air-conditioned. solution pack should also be cooled. Calibrations and quality control o.K. | With user conducted several phone calls. procedure with pCO2. RMED has produced new sensor cartridges with improved sensitivity. User will order the new SC. Room of the analyzer is now very warm in summer, windows on and not air-conditioned. solution pack should also be cooled. Calibrations and quality control o.K. | User error | SP | Storage Temp | Storage Temp | pCO2
ABL90 FLEX PLUS | Gerät abgeschalten und Oximetrie ausgebaut, Hemolyzer ausgebaut und ersetzt, wieder zusammengesetzt, Schlauch Oxi zu Messkammer ersetzt, Neustart abgewartet, mehrere Kal, tHb Kal, QC | Gerät abgeschalten und Oximetrie ausgebaut, Hemolyzer ausgebaut und ersetzt, wieder zusammengesetzt, Schlauch Oxi zu Messkammer ersetzt, Neustart abgewartet, mehrere Kal, tHb Kal, QC | SWO-116240 | No | Switch off the device and removed oximetry, hemolyzer removed and replaced, reassembled, tube Oxi replaced to measuring chamber, restart awaited, several Kal, tHb Kal, QC | Switch off the device and removed oximetry, hemolyzer removed and replaced, reassembled, tube Oxi replaced to measuring chamber, restart awaited, several Kal, tHb Kal, QC | Defect | HW | Oxi module | Hemolyzer
AQT90 FLEX |  | B | SWO-115712 | No |  |  | Defect | HW | Aspiration Module | Inlet Wheel | 436
AQT90 FLEX |  | B | SWO-115795 | No |  |  | Defect | HW | Process Module | Stepper motor, process wheel | 5784
ABL80 CO-OX | Limpieza de la pantalla sustitución de la touch screen, calibración de la nueva panntalla por servicio tecnico, chequeos gennerales de uso funcionando correctamennte. | Cambio de el touch screen. | D107E4B9F64154A9 | No | Cleaning the screen replacing the touch screen, calibration of the new panntalla by technical service, genneral checks of use in good operation. | Changing the touch screen.
ABL825 FLEX |  | B | D1084015F60FDB22 | No |  |  | Setup | SW | His/lis | Aqure
ABL837 FLEX |  | B | D111E7C4A618F4F6 | No |  |  | Cleaned | HW | Inlet module | Inlet gasket
ABL835 FLEX | electrode de remplacement recu Calibration erreur Sensibilit basse mais a l'arrivè la sensibilité 93% | Vérifier gas module entre dans specification vérifier contact et pco2 canal amplification ok | D1081B2FF63FC229 | No | replacement electrode received Calibration error Low Sensitivity but arrived sensitivity 93% | Check gas module enters in specification check contact and pco2 channel amplification ok | Assisted Replacement | Electrode | Electrode | pCO2
ABL825 FLEX | came on site to replace doop sensor, found gear motor slipping | took apart the auto check module and found the pin on the gear motor loose; tighten gear motor pin; Ref membrane due replacement on 7 9 replaced ref membrane cleaned the chamber | D1080395F61CB3D2 | No | came on site to replace doop sensor; found gear motor slipping | took apart the auto check module and found the pin on the gear motor loose; tighten gear motor pin; Ref membrane due replacement on 7 9 replaced ref membrane cleaned the chamber | Adjustment | HW | AutoCheck Module | Motor assembly
ABL835 FLEX | Fehlersuche Analysator neu gestartet Oximetrie überprüft FlexQ kalibriert FlexQ Modul defekt erneuter Kundenbesuch notwendig Einlasssonde getauscht da stark verbogen (Kundenbestand) Einlassdichtung getauscht da stark verschmutzt Analysator kalibriert und überprüft QC Messung durchgeführt | Fehlersuche Analysator neu gestartet Oximetrie überprüft FlexQ kalibriert FlexQ Modul defekt erneuter Kundenbesuch notwendig Einlasssonde getauscht da stark verbogen (Kundenbestand) Einlassdichtung getauscht da stark verschmutzt Analysator kalibriert und überprüft QC Messung durchgeführt | SWO-115879 | No | Troubleshooting analyzer restarted Oximetry checked FlexQ calibrated FlexQ module defective re-entry customer visit necessary inlet probe replaced as heavily bent (customer stock) inlet seal exchanged as heavily polluted analyzer calibrates and checks QC QC Measurement performed | Troubleshooting analyzer restarted Oximetry checked FlexQ calibrated FlexQ module defective re-entry customer visit necessary inlet probe replaced as heavily bent (customer stock) inlet seal exchanged as heavily polluted analyzer calibrates and checks QC QC Measurement performed | Defect | HW | Oxi module | Off/On | A/D Converter
ABL837 FLEX |  | B | D111E151A624F4F0 | Yes |  |  | Re classified RMED | Upgrade with FlexQ
ABL825 FLEX | Used verified good USB thumb drive on this analyzer on all 3 USB ports, analyzer would not recongize USB thumb drive, rebooted analyzer and still would not recognize USB Thumbdrive. | Replaced CPU unit. | D1082EA9F63D843A | No | Used verified good USB thumb drive on this analyzer on all 3 USB ports, analyzer would not recongize USB thumb drive, rebooted analyzer and still would not recognize USB Thumbdrive. | Replaced CPU unit. | Defect | HW | DMS_PC | CPU | USB connection
ABL815 FLEX | Da Lac Membran nur 1 Woche mit vielen Problemen in der Kalibration Tausch der Lac-Elektrode+ anderer Membran (fehlerhafte Lac-Elektrode UP04), Kal, QC | Da Lac Membran nur 1 Woche mit vielen Problemen in der Kalibration Tausch der Lac-Elektrode+ anderer Membran (fehlerhafte Lac-Elektrode UP04), Kal, QC | SWO-115942 | No | Since Lac membrane only 1 week with many problems in the calibration exchange of the Lac electrode+ other membrane (faulty Lac electrode UP04), Kal, QC | Since Lac membrane only 1 week with many problems in the calibration exchange of the Lac electrode+ other membrane (faulty Lac electrode UP04), Kal, QC | Assisted Replacement | Electrode | Membrane | Lac
AQT90 FLEX | 5766 restart analyzer errors. Incubation out of temp plus several other errors. | Replaced main PCB. Attempted to upgrade batch folder but process did not go according to service manual. Collected data dump. | D1080F2DF646657F | No | 5766 restart analyzer errors. Incubation out of temp plus several other errors. | Replaced main PCB. Attempted to upgrade batch folder but process did not go according to service manual. Collected data dump. | Defect | HW | Analyzer Control Module | PCB, Main board | 5766
AQT90 FLEX |  | B | D1183F64F454CBF0 | No |  |  | Assisted Replacement | Consumables | SP | SP
AQT90 FLEX | Replace Cartridge with New one. | Replace Cartridge with New one. | SWO-116763 | No | Replace Cartridge with New one. | Replace Cartridge with New one. | Defect | Consumables | Test Kit | Cartridge | Broken
AQT90 FLEX |  | B | D107E771F610C660 | No |  |  | Assisted Replacement | Consumables | Test Kit | Cartridge | Unknown
ABL80 FLEX BASIC | instaaled new sensor cassette with lactate. failed to install with a message of verification check. | instaaled new sensor cassette with lactate. failed to install with a message of verification check. | SWO-116780 | No | instaaled new sensor cassette with lactate. failed to install with a message of verification check. | instaaled new sensor cassette with lactate. failed to install with a message of verification check.
ABL80 FLEX BASIC | Checked the Calibration and Patient result. Runned the QC and Qc Values are good. | Checked the Calibration and Patient result. Runned the QC and Qc Values are good. | SWO-116770 | No | Checked the Calibration and Patient result. Runned the QC and Qc Values are good. | Checked the Calibration and Patient result. Runned the QC and Qc Values are good.
ABL90 FLEX | Oximetriás modul küvettájának tisztítása, vonalkód olvasó ablak csere. Kalibrálás, QC mérés. | B | SWO-116176 | No | Cleaning the Oximetric module cuveted, Barcode Reader window replacement. Calibration, QC measurement. |  | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
ABL90 FLEX | Removed all archived data, and started a new database. Removed the lis connection as it is not required for the spare analyser. Ran all calibrations and qcs to verify analyser functionality. | Removed all archived data, and started a new database. Removed the lis connection as it is not required for the spare analyser. Ran all calibrations and qcs to verify analyser functionality. | SWO-116227 | No | Removed all archived data, and started a new database. Removed the lis connection as it is not required for the spare analyser. Ran all calibrations and qcs to verify analyser functionality. | Removed all archived data, and started a new database. Removed the lis connection as it is not required for the spare analyser. Ran all calibrations and qcs to verify analyser functionality. | Setup | His/lis | His/lis | His/lis
ABL90 | sO2 QC results erratic and going in and out of range. | Replaced hemolyzer and spectrometer. Added 22=0 and 23=0 to ini file but badges still do not scan. Set back to default. Installed battery option and roll cart. Collected data logs. | D10957CB6C587230 | No | sO2 QC results erratic and going in and out of range. | Replaced hemolyzer and spectrometer. Added 22=0 and 23=0 to ini file but badges still do not scan. Set back to default. Installed battery option and roll cart. Collected data logs. | Defect | HW | Oxi module | Hemolyzer
AQT90 FLEX |  | B | SWO-116102 | No |  |  |  | LEH Mail |  | Not Cat. | 161
ABL800 BASIC | Run decontamination multiple time. Reset the oxi-module & calibrate multiple time. still issue remain same. need to check the hemolyzer. | Run decontamination multiple time. Reset the oxi-module & calibrate multiple time. still issue remain same. need to check the hemolyzer. | SWO-116679 | No | Run decontamination multiple time. Reset the oxi-module & calibrate multiple time. still issue remain same. need to check the hemolyzer. | Run decontamination multiple time. Reset the oxi-module & calibrate multiple time. still issue remain same. need to check the hemolyzer. | Revisit required | HW | Oxi module | Hemolyzer
ABL825 FLEX |  | B | D109646A6C35DB28 | No |  |  | Setup | SW | His/lis | Aqure
ABL825 FLEX | Etat élevée 22 Ph electrode liquid a l'intérieur | Remplacé électrode de reference et nettoyé électrode de Ph et remplacé O'ring | D109104A6C5F4D5F | No | High state 22 Ph electrode liquid inside | Replaced reference electrode and cleaned Ph electrode and replaced O'ring | Assisted Replacement | Electrode | Electrode | Reference
ABL800 BASIC | 引上げ修理 | B | SWO-115978 | No | Raise repair |  | Unknown | Electrode | Electrode | pH
ABL837 FLEX |  | B | D109875F6C385B6C | No |  |  | Defect | SW | Installation | Installation | 6.16?
ABL835 FLEX |  | B | D109518C6C225B6A | Yes |  |  | Re classified RMED | Setup
ABL90 FLEX | Analysator zerlegt und sämtliche Kabel auf Beschädigung und richtigen sitz geprüft. Zu Testzwecken einen anderen Bildschirm angebaut. Kunde soll sich wieder melden falls der Bildschirm wieder ausfällt. | Anwender bittet um Abholung des Gerätes!!! | SWO-116605 | No | Analyzer disassembled and checked all cables for damage and correct fit. A different screen is attached for testing purposes. Customer should report again if the screen fails again. | User asks for the device to be picked up!!! | Defect | HW | DMS_PC | Display
ABL90 FLEX | Go to service mode and run all adjustment in fludic section and retest the instrument and resolve the problem | Go to service mode and run all adjustment in fludic section and retest the instrument and resolve the problem | SWO-118183 | No | Go to service mode and run all adjustment in fludic section and retest the instrument and resolve the problem | Go to service mode and run all adjustment in fludic section and retest the instrument and resolve the problem | Cleaned | Fluidic pathway | Fluidic pathway | Clot
ABL90 | Verified. Customer has no consumabales. Worked with Tech Support and Customer Service to order a sensor cassett. Return visit required. WO left open until problem is resolved. | B | D10ABBBEE22D57CE | No | Verified. Customer has no consumabales. Worked with Tech Support and Customer Service to order a sensor cassett. Return visit required. WO left open until problem is resolved. |  | Revisit required | SC | SC | Lac
ABL90 FLEX | Konfiguration Layout und Pat. Liste ...Testmessung ok. Time Server Konfiguration Uhrzeit ok. | Konfiguration Layout und Pat. Liste ...Testmessung ok. Time Server Konfiguration Uhrzeit ok. | SWO-116513 | No | Configuration layout and Pat. List... Test measurement ok. Time Server configuration time ok. | Configuration layout and Pat. List... Test measurement ok. Time Server configuration time ok. | Setup | SW | Time setting | Time setting
ABL90 FLEX PLUS | Went to replace PCB but the error still remained, Removed the complete inlet module and tested it on Micu's machine and the error carried over to the Micu's machine. Entire inlet module is faulty. | B | SWO-117201 | No | Went to replace PCB but the error still remained, Removed the complete inlet module and tested it on Micu's machine and the error carried over to the Micu's machine. Entire inlet module is faulty. |  | Defect | HW | Auto Inlet module | Auto Inlet module | Email
TCM CombiM | 先借用配件910-246和903-365，更换后未解决。再追加借件910-246。更换电源板910-246后解决(by zhenjie) | 先借用配件910-246和903-365，更换后未解决。再追加借件910-246。更换电源板910-246后解决(by zhenjie) | SWO-116101 | No | First borrow accessories 910-246 and 903-365, after replacement not resolved. An additional 910-246 is added. Post-replacement solution (by-zhenjie) | First borrow accessories 910-246 and 903-365, after replacement not resolved. An additional 910-246 is added. Post-replacement solution (by-zhenjie) | Defect | HW | Power Supply | Power Supply
AQT90 FLEX | 更换处理单元和存储盘条码扫描器 | 更换条码扫描器 | SWO-115813 | No | Replace the processing unit and storage disk barcode scanner | Replace barcode scanner | Defect | HW | Process Module | Optical Unit
ABL825 FLEX | Confirmed problem | Replaced the inlet nipples and waste tubing. Also replaced the inlet probe, inlet tubing and waste pump tubing. | D10A961AE2541B91 | No | Confirmed problem | Replaced the inlet nipples and waste tubing. Also replaced the inlet probe, inlet tubing and waste pump tubing. | Defect | HW | Inlet module | Nipple kit | Leaking
ABL835 FLEX |  | B | D111DA5BA63AF4E6 | No |  |  | Defect | HW | El/Met Module | Ca++
ABL825 FLEX | Modul (902-753R1241N001) ersetzt. Alle Testfunktionen im Servicemenü absolviert, keine Fehler. | Modul (902-753R1241N001) ersetzt. Alle Testfunktionen im Servicemenü absolviert, keine Fehler. | SWO-116453 | No | module (902-753R1241N001). Completed all test functions in the service menu, no errors. | module (902-753R1241N001). Completed all test functions in the service menu, no errors. | Defect | HW | FlexQ Module | FlexQ Module
AQT90 FLEX | Run the system clean procedure 4 times after which run 2 samples. Results are OK. Machine is working absolutely fine. Cleaned the machine internally as lot of dust was found inside the machine. | Run the system clean procedure 4 times after which run 2 samples. Results are OK. Machine is working absolutely fine. Cleaned the machine internally as lot of dust was found inside the machine. | SWO-116790 | No | Run the system clean procedure 4 times after which run 2 samples. Results are OK. Machine is working absolutely fine. Cleaned the machine internally as lot of dust was found inside the machine. | Run the system clean procedure 4 times after which run 2 samples. Results are OK. Machine is working absolutely fine. Cleaned the machine internally as lot of dust was found inside the machine. | Cleaned | HW | Aspiration Module | HCT Sensor
AQT90 FLEX | need the replacement of optical unit. | replaced the BNP sensor. and calibrate the Parameters - TropT, D-Dimer which was successfully calibrated on the analyzer. | SWO-118196 | No | need the replacement of optical unit. | replaced the BNP sensor. and calibrate the Parameters - TropT, D-Dimer which was successfully calibrated on the analyzer. | Defect | HW | Process Module | Optical Unit | 131
AQT90 FLEX | Cup hat sich zwischen Pufferkassette verklemmt. Pufferkassette ersetzt, Cup Transport mit sechs Cups durchgeführt. I.O. Startphase QC Level 1 und 2 gemessen. | Techniker sollte kommen | SWO-116466 | No | Cup has become stuck between buffer cartridges. Buffer cartridge replaced, Cup transport carried out with six cups. I.O. Start phase QC Levels 1 and 2 measured. | Technician should come | Defect | Consumables | SP | SP | 5784
AQT90 FLEX | 代替のフロント部と交換 | B | SWO-117941 | No | Replace with alternate front |  | Defect | HW | Cover and Chassis | Inlet Door
ABL80 FLEX CO-OX | Gerät überprüft, Netzteil erneuert, Kalibration o.k. | Zusammen am Telefon den ABL erneut aus- und eingeschaltet. Display funktioniert wieder. | SWO-116729 | No | Device checked, power supply renewed, calibration ok | Together on the phone, turn the ABL off and on again. Display works again.
ABL80 FLEX CO-OX | SP、機器本体の廃液マニホールド洗浄 | B | SWO-117387 | No | Waste liquid manifold cleaning of SP and equipment body
ABL80 FLEX BASIC | went to active parameters and given two point calibration. analyser is working successfully. deleted database. | went to active parameters and given two point calibration. analyser is working successfully. deleted database. | SWO-116822 | No | went to active parameters and given two point calibration. analyser is working successfully. deleted database. | went to active parameters and given two point calibration. analyser is working successfully. deleted database.
ABL80 FLEX BASIC | Done the printing setup properly and check for printing and give proper training to the technicians | Done the printing setup properly and check for printing and give proper training to the technicians | SWO-118181 | No | Done the printing setup properly and check for printing and give proper training to the technicians | Done the printing setup properly and check for printing and give proper training to the technicians
ABL80 CO-OX | Techline case is present here. Level3 QC Cl is at high end of insert range. | Replaced SP and SC from customer. received NA Linearity error. Replaced SC with one of mine. Same issue. Replaced SP with one of mine. Na error gone. Cl reads 67. Replaced SC with one of mine. CL reads 66 but after a System Cycle, reads 67. Replaced SC with last one of mine, lot 223704. Cl now reads 63, one mmol L above mean. Replaced 1 box of each level of QC as a courtesy so QC study can be completed. | D10BC8A5583A9B0D | No | Techline case is present here. Level3 QC Cl is at high end of insert range. | Replaced SP and SC from customer. received NA Linearity error. Replaced SC with one of mine. Same issue. Replaced SP with one of mine. Na error gone. Cl reads 67. Replaced SC with one of mine. CL reads 66 but after a System Cycle, reads 67. Replaced SC with last one of mine, lot 223704. Cl now reads 63, one mmol L above mean. Replaced 1 box of each level of QC as a courtesy so QC study can be completed.
ABL90 |  | B | D10BD48D5811AF75 | Yes |  |  | Re classified RMED | Planned PM
ABL90 |  | B | D10BACFA580719AF | No |  |  | Defect | HW | Oxi module | Pinch valve tube
ABL90 FLEX | Fehlersuche, Gerät ein u. ausgeschaltet, SC gew., Initialisierung abgewartet, Fehlerprotokolle erstellt, Service dump gezogen, Kal. + QC i.O. | Fehlersuche; Gerät ein u. ausgeschaltet; SC gew.; Initialisierung abgewartet; Fehlerprotokolle erstellt; Service dump gezogen; Kal. + QC i.O. | SWO-116683 | No | Troubleshooting; device turned on and off; SC. initialization awaited; error logs created; Service dump pulled; Kal. + QC in the first place | Troubleshooting; device turned on and off; SC. initialization awaited; error logs created; Service dump pulled; Kal. + QC in the first place | Assisted Replacement | SC | SC | K
ABL90 |  | B | D11092F1303D336D | No |  |  | Setup | Report Lay out | Report Lay out | Report Lay out
ABL90 FLEX | 更换新试剂包安装后解决 | 更换新试剂包安装后解决 | SWO-116832 | No | Post-replacement solution for new reagent kit installation | Post-replacement solution for new reagent kit installation | Assisted Replacement | SP | SP | Chip
ABL815 FLEX | El Met, pH BG modul felső szelepi hibásak, tisztítás, illetve csere után megfelelően működnek. Kalibrálás, QC mérés. | B | SWO-116676 | No | The upper valves of El Met, pH BG module are defective, they function properly after cleaning or replacement. Calibration, QC measurement. |  | Defect | HW | pH/BG Module | Valve
ABL800 BASIC | Found it was problem with inlet gasket, replaced the inlet gasket, calibrated inlet probe found instrument is functional. | Found it was problem with inlet gasket, replaced the inlet gasket, calibrated inlet probe found instrument is functional. | SWO-116746 | No | Found it was problem with inlet gasket, replaced the inlet gasket, calibrated inlet probe found instrument is functional. | Found it was problem with inlet gasket, replaced the inlet gasket, calibrated inlet probe found instrument is functional. | Defect | HW | Inlet module | Inlet gasket
ABL800 FLEX | Replaced Creatine membrane | Replaced Creatine membrane | SWO-116727 | No | Replaced Creatine membrane | Replaced Creatine membrane | Assisted Replacement | Electrode | Membrane | Crea | Sensitivity
ABL800 FLEX | Replaced PO2 & Glu membrane | Replaced PO2 & Glu membrane | SWO-116726 | No | Replaced PO2 & Glu membrane | Replaced PO2 & Glu membrane | Assisted Replacement | Electrode | Membrane | Glu | Sensitivity
ABL835 FLEX | キャピラリー側のチューブニップル破損の為、交換。エアーライン洗浄、LS調整実施。 | キャピラリー側のチューブニップル破損の為、交換。エアーライン洗浄、LS調整実施。 | SWO-117740 | No | Replacement for tube nipple damage on the capillary side. Air line cleaning and LS adjustment. | Replacement for tube nipple damage on the capillary side. Air line cleaning and LS adjustment. | Defect | HW | Inlet module | Leakage
ABL835 FLEX | . The measuring chambers were cleaned with hypochlorite and rinsed out. They were then inspected to ensure no fibres remained behind. The electrodes were inspected to ensure no fibres were stuck on the tips membranes. pO2 no longer retain solution when drained however there is a small amount in the reference and possibly the Cl chambers. The reference membrane is Lot 1942. A new pH sealing ring was fitted to ensure a tight seal. Crystals removed from the inlet preheated | . The measuring chambers were cleaned with hypochlorite and rinsed out. They were then inspected to ensure no fibres remained behind. The electrodes were inspected to ensure no fibres were stuck on the tips membranes. pO2 no longer retain solution when drained however there is a small amount in the reference and possibly the Cl chambers. The reference membrane is Lot 1942. A new pH sealing ring was fitted to ensure a tight seal. Crystals removed from the inlet preheated | SWO-116700 | No | . The measuring chambers were cleaned with hypochlorite and rinsed out. They were then inspected to ensure no fibres remained behind. The electrodes were inspected to ensure no fibres were stuck on the tips membranes. pO2 no longer retain solution when drained however there is a small amount in the reference and possibly the Cl chambers. The reference membrane is Lot 1942. A new pH sealing ring was fitted to ensure a tight seal. Crystals removed from the inlet preheated | . The measuring chambers were cleaned with hypochlorite and rinsed out. They were then inspected to ensure no fibres remained behind. The electrodes were inspected to ensure no fibres were stuck on the tips membranes. pO2 no longer retain solution when drained however there is a small amount in the reference and possibly the Cl chambers. The reference membrane is Lot 1942. A new pH sealing ring was fitted to ensure a tight seal. Crystals removed from the inlet preheated | Setup | SW | Reference ranges | pCO2
ABL827 FLEX | Tried new membranes-No change | Replaced glucose electrode Replaced pH o-ring and cleaned measuring chamber | D10E500944272C39 | No | Tried new membranes-No change | Replaced glucose electrode Replaced pH o-ring and cleaned measuring chamber | Assisted Replacement | Electrode | Electrode | Glu
ABL827 FLEX | Lac Membran erneuert. | Lac Membran erneuert. | SWO-116807 | No | Lac membrane renewed. | Lac membrane renewed. | Assisted Replacement | Electrode | Membrane | Lac | Leakage
ABL835 FLEX | FlexQ Modul getauscht alte s n 902-753R0901n003 FlexQ kalibriert und überprüft Analysator kalibriert und überprüft QC Messung durchgeführt | FlexQ Modul getauscht alte s n 902-753R0901n003 FlexQ kalibriert und überprüft Analysator kalibriert und überprüft QC Messung durchgeführt | SWO-120644 | No | FlexQ module swapped old s n 902-753R0901n003 FlexQ calibrated and verified analyzer calibrated and checked QC measurement performed | FlexQ module swapped old s n 902-753R0901n003 FlexQ calibrated and verified analyzer calibrated and checked QC measurement performed | Defect | HW | FlexQ Module | FlexQ Module
ABL820 FLEX | Error 653 Samples in Ph Bg upper liquid sensor detected prematurely. | I noticed that at midnight after running a Pump calibration the analyzer started failing. Performed another pump calibration Liquid sensork, it all came in after that. | D10BEAC95824E733 | No | Error 653 Samples in Ph Bg upper liquid sensor detected prematurely. | I noticed that at midnight after running a Pump calibration the analyzer started failing. Performed another pump calibration Liquid sensork, it all came in after that. | Adjustment | HW | pH/BG Module | LS
ABL810 FLEX | AC Modul ersetzt(Ersatz:??, Original:902-821R2506N003) SW installiert, Neustart, Kal und beide QC, Datensicherung | B | SWO-116852 | No | AC module replaces(replacement:??, original:902-821R2506N003) SW installed, restart, cal and both QC, backup |  | Defect | HW | AutoCheck Module | AutoCheck Module
ABL810 FLEX | nach Rücksprache technischen Support Willich und RMED Autocheckmodul(Ersatz:902-821R1311N006, Original:902-821R2508N007) ersetzt, Neustart, mehrere Kal, QC | neue Charge Membran verwendet, Luftdruck korrigiert, Dichtung pH unten ersetzt, mehrere Kal, QC | SWO-116853 | No | after consultation technical support Willich and RMED autocheck module (replacement:902-821R1311N006, original:902-821R2508N007) replaced, restart, multiple Kal, QC | new batch membrane used, air pressure corrected, seal pH replaced below, several Kal, QC | Assisted Replacement | Electrode | Membrane | pCO2
ABL825 FLEX | Empfindlichkeit erneut unter 150 => Elektrode ersetzt. | Empfindlichkeit erneut unter 150 => Elektrode ersetzt. | SWO-116734 | No | Sensitivity again under 150 => Electrode replaced. | Sensitivity again under 150 => Electrode replaced. | Assisted Replacement | Electrode | Membrane | Lac | Sensitivity
ABL825 FLEX | Nadel ersetzt. | Nadel ersetzt. | SWO-116740 | No | Needle replaced. | Needle replaced. | Defect | HW | Inlet module | Probe bend
ABL825 FLEX | Touchscreen ohne Funktion, keine Bedienung mehr möglich. Monitor wird bestellt zwecks Austausch. | Touchscreen ohne Funktion, keine Bedienung mehr möglich. Monitor wird bestellt zwecks Austausch. | SWO-116736 | No | Touch screen without function, no more operation possible. Monitor is ordered for replacement. | Touch screen without function, no more operation possible. Monitor is ordered for replacement. | Revisit required | HW | DMS_PC | touch Screen
AQT90 FLEX | Replaced the 24V fuse on the main board of the machine, restarted the machine, after which run all the pending samples. All the results are OK. Machine is working fine. Kept the machine under observation. | Replaced the 24V fuse on the main board of the machine, restarted the machine, after which run all the pending samples. All the results are OK. Machine is working fine. Kept the machine under observation. | SWO-116791 | No | Replaced the 24V fuse on the main board of the machine, restarted the machine, after which run all the pending samples. All the results are OK. Machine is working fine. Kept the machine under observation. | Replaced the 24V fuse on the main board of the machine, restarted the machine, after which run all the pending samples. All the results are OK. Machine is working fine. Kept the machine under observation. | Defect | HW | Main PCB | Fuse (24V)
AQT90 FLEX |  | B | D10BDD8D58015363 | Yes |  |  | Re classified RMED | Planning visit
AQT90 FLEX | Aktuell Probleme mit TnT im Level 2 bei LQC, mehrere Systemreinigung durchgeführt, Reagenzienpack getauscht, Systemüberprüfung alles Ok, eingesetzte TnT Testkassette entnommen und neue, von gleicher LOT eingesetzt, LQC TnT durchgeführt, LQC wurde wieder nicht bestanden - zu niedrig, mit Kunden so verblieben, dass er einen neuen Testkit TnT am Dienstag, 19.07.16 vor 12:00 Uhr zugesandt bekommt, dann wird der Kunde die Kal. für TnT durchführen und im Anschluss direkt die fehlerhafte LQC wiederholen, gibt es hier dann weitere Probleme, wird sich der Kunde wieder melden. | Aktuell Probleme mit TnT im Level 2 bei LQC, mehrere Systemreinigung durchgeführt, Reagenzienpack getauscht, Systemüberprüfung alles Ok, eingesetzte TnT Testkassette entnommen und neue, von gleicher LOT eingesetzt, LQC TnT durchgeführt, LQC wurde wieder nicht bestanden - zu niedrig, mit Kunden so verblieben, dass er einen neuen Testkit TnT am Dienstag, 19.07.16 vor 12:00 Uhr zugesandt bekommt, dann wird der Kunde die Kal. für TnT durchführen und im Anschluss direkt die fehlerhafte LQC wiederholen, gibt es hier dann weitere Probleme, wird sich der Kunde wieder melden. | SWO-116724 | No | Currently problems with TnT in level 2 at LQC, several system cleaning performed, reagent pack exchanged, system check all Ok, inserted TnT test cassette removed and new, used by the same LOT, LQC TnT performed, LQC was again failed - too With customers left so that he will receive a new test kit TnT on Tuesday, 19.07.16 before 12:00, then the customer will be the Kal. for TnT and then repeat the faulty LQC directly, there are further problems here, the customer will contact again. | Currently problems with TnT in level 2 at LQC, several system cleaning performed, reagent pack exchanged, system check all Ok, inserted TnT test cassette removed and new, used by the same LOT, LQC TnT performed, LQC was again failed - too With customers left so that he will receive a new test kit TnT on Tuesday, 19.07.16 before 12:00, then the customer will be the Kal. for TnT and then repeat the faulty LQC directly, there are further problems here, the customer will contact again. | Assisted Replacement | Consumables | LQC Kit | TnT
ABL90 | Verificação da entrada partida, retirar a entrada danificada, substituição por uma nova, verificação do funcionamento da mesma, substituição da cassete de eléctrodos, estabilização e passagem do QC. | B | D10D3973CE3BCE62 | No | Checking the starter, removing the damaged entry, replacing a new one, checking the operation, replacing the electrode cassette, stabilizing and passing the QC. |  | Defect | HW | Inlet module | Inlet gasket
AQT90 FLEX | 更换新进样针后解决---jack.chu | 更换新进样针后解决---jack.chu | SWO-116835 | No | Solve the ---jack.chu after replacing the new sample needle | Solve the ---jack.chu after replacing the new sample needle | Defect | HW | Aspiration Module | Needle assembly | 794
AQT90 FLEX | replaced new solution pack and cartridge. analyser is working fine. | replaced new solution pack and cartridge. analyser is working fine. | SWO-116823 | No | replaced new solution pack and cartridge. analyser is working fine. | replaced new solution pack and cartridge. analyser is working fine. | User error | Consumables | SP | Expired
ABL90 FLEX | 更换测试卡修复仪器 | 更换测试卡修复仪器 | SWO-116872 | No | Replace the test card repair instrument | Replace the test card repair instrument | Assisted Replacement | SC | SC | Reference
ABL90 FLEX |  | B | SWO-116814 | No |  |  | Defect | HW | SC ID PCB | SC Interface PCB
AQT90 FLEX | 更换新试剂包安装后解决 | 更换新试剂包安装后解决 | SWO-116836 | No | Post-replacement solution for new reagent kit installation | Post-replacement solution for new reagent kit installation | Defect | HW | Aspiration Module | Needle Wash Unit | 3414
ABL80 FLEX BASIC | 更换配件903-334和902-972修复仪器---jack.chu | 更换配件903-334和902-972修复仪器---jack.chu | SWO-116795 | No | Replacement accessories 903-334 and 902-972 repair instruments ---jack.chu | Replacement accessories 903-334 and 902-972 repair instruments ---jack.chu
ABL80 FLEX BASIC |  | B | SWO-116845 | No
ABL90 |  | B | D10FAB3BBA62F6ED | Yes |  |  | Re classified RMED | SWU
ABL90 | Desmontagem do equipamento, substituição do material indicado, montagem e verificação do funcionamento, passagem do QC. Tudo ok. | B | D10F7F93BA4BCE66 | No | Disassembly of the equipment, replacement of the indicated material, assembly and verification of the operation, passage of the QC. Okay, all right. |  | Defect | HW | Actuator unit | Actuator unit
ABL90 FLEX |  | B | SWO-116860 | No |  |  | Assisted Replacement | SC | SC | Lac
ABL825 FLEX |  | B | D10F9729BA22D546 | No |  |  | Defect | HW | Printer module | Printer Head
ABL820 FLEX | Troubleshot to a bad valve on the pH BG module. | Replaced both valves on the pH BG module, and this resolved the problem. | D10F937EBA18C055 | No | Troubleshot to a bad valve on the pH BG module. | Replaced both valves on the pH BG module, and this resolved the problem. | Defect | HW | pH/BG Module | Valve
ABL825 FLEX |  | B | D10FB102BA040838 | Yes |  |  | Re classified RMED | Logistical upSWO
ABL827 FLEX | While checking analyzers before going live, spotted leaking at inlet | Replaced inlet gasket. | D10FF945BA558386 | No | While checking analyzers before going live, spotted leaking at inlet | Replaced inlet gasket. | Defect | HW | Inlet module | Leaking
AQT90 FLEX | AQT geöffnet, Barcodescanner ausgebaut und komplett neu eingestellt. Montage und Test des Scanners. Bei erster Messung fiel der Probenhalter samt Probe aus dem Inletwheel => befestigt und alle anderen Komponenten sichtkontrolliert. Testmessung mit Blut und CRP1 Messung angestoßen. IP Adresse sowie Serveradresse etc. nach Vorgabe eingestellt und Übertragung getestet, OK | AQT geöffnet, Barcodescanner ausgebaut und komplett neu eingestellt. Montage und Test des Scanners. Bei erster Messung fiel der Probenhalter samt Probe aus dem Inletwheel => befestigt und alle anderen Komponenten sichtkontrolliert. Testmessung mit Blut und CRP1 Messung angestoßen. IP Adresse sowie Serveradresse etc. nach Vorgabe eingestellt und Übertragung getestet, OK | SWO-117039 | No | AQT opened, barcode scanner removed and completely re-adjusted. Assembly and testing of the scanner. At the first measurement, the sample holder and sample fell from the Inletwheel => attached and all other components were visually controlled. Test measurement with blood and CRP1 measurement initiated. IP address as well as server address etc. set according to default and transmission tested, OK | AQT opened, barcode scanner removed and completely re-adjusted. Assembly and testing of the scanner. At the first measurement, the sample holder and sample fell from the Inletwheel => attached and all other components were visually controlled. Test measurement with blood and CRP1 measurement initiated. IP address as well as server address etc. set according to default and transmission tested, OK | Defect | HW | Aspiration Module | Barcode reader, Inlet
ABL80 FLEX Basic version |  | B | D11C2C7256485170 | No
ABL90 FLEX | Fehler QC pCO2. Lösungspack getauscht. Fehler besteht weiterhin, Sensorkassette erneuert. Gerätekonfiguration überprüft. Datenübertragung ins Labor getestet. Laktat noch nicht messbereit. | Fehler QC pCO2. Lösungspack getauscht. Fehler besteht weiterhin, Sensorkassette erneuert. Gerätekonfiguration überprüft. Datenübertragung ins Labor getestet. Laktat noch nicht messbereit. | SWO-117356 | No | Error QC pCO2. solution pack. Error persists, sensor cartridge renewed. Device configuration checked. data transfer to the laboratory. Lactate not yet ready for measurement. | Error QC pCO2. solution pack. Error persists, sensor cartridge renewed. Device configuration checked. data transfer to the laboratory. Lactate not yet ready for measurement. | Assisted Replacement | SC | SC | pCO2
ABL90 FLEX | Gerät nach vollzogener Reparatur beim Kunden installiert. Kundengespräch durchgeführt. | Gerät nach vollzogener Reparatur beim Kunden installiert. Kundengespräch durchgeführt. | SWO-117340 | Yes | Device installed at the customer after repair has been completed. customer interview. | Device installed at the customer after repair has been completed. customer interview. | Re classified RMED | Re -installed
ABL90 FLEX | Nettoyage du système e.o Calibration e.o contrôle de qualité e.o | Nettoyage du système e.o Calibration e.o contrôle de qualité e.o | SWO-117338 | No | Cleaning the e.o Calibration e.o system quality control e.o | Cleaning the e.o Calibration e.o system quality control e.o | No fault found | No fault found | No fault found | No fault found
ABL90 FLEX | Fließwege gespült, LP und SC aus Kundenbestand ersetzt. Handbuchupdate ausgehändigt. | Sie möchte einen Vor Ort Einsatz. | SWO-117397 | No | Flow paths rinsed, LP and SC replaced from customer stock. manual update. | She wants an on-site operation. | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
ABL90 FLEX | kein Fehler | kein Fehler | SWO-117218 | Yes | no error | no error | Re classified RMED | upSWO related to AQT
ABL90 |  | B | D11AF20FE0438BD7 | Yes |  |  | Re classified RMED | Courtesy
ABL90 | Most failures were due to either Cal errors present during QC. Calibration errors mainly due to Lac, and Glu. pO2 Status values too low.Inspected Fluid path for leaks. | Performed Service pO2 test. Passed. | D1111515300C5EB2 | No | Most failures were due to either Cal errors present during QC. Calibration errors mainly due to Lac, and Glu. pO2 Status values too low. Inspected Fluid path for leaks. | Performed Service pO2 test. Passed. | Unknown | SC | SC | pO2
ABL90 |  | B | D110A71E3029AFBD | No |  |  | Assisted Replacement | SP | SP | pO2
ABL90 | Ran pressure tests all bad. | Replaced allparts in PM kit, pressure tests are good. Lactate will not calibrate. Replaced sensor cassette sensitivity is too low. Shut off lactae so they c an use ABL. Will terturn in AM to check on Lactate. | D1110637302766D8 | No | Ran pressure tests all bad. | Replaced allparts in PM kit, pressure tests are good. Lactate will not calibrate. Replaced sensor cassette sensitivity is too low. Shut off lactae so they c an use ABL. Will terturn in AM to check on Lactate. | Defect | HW | PM kit | PM kit
ABL90 FLEX PLUS |  | B | D110E44D301A5339 | No |  |  | Adjustment | HW | Auto Inlet module | Motor assembly | Email
AQT90 FLEX | 更换了903-364后修复---jack.chu | 更换了903-364后修复---jack.chu | SWO-117122 | No | Repaired after replacing 903-364---jack.chu | Repaired after replacing 903-364---jack.chu | Defect | HW | Process Module | Optical Unit | 135
AQT90 FLEX |  | B | SWO-117172 | No |  |  | Defect | HW | Aspiration Module | Bend needle | 4794
ABL800 BASIC | cleaned the pO2 electrode and replaced the membrane, done zero current calibration in service, done two pint calibration and QC all results found with in the range. | cleaned the pO2 electrode and replaced the membrane, done zero current calibration in service, done two pint calibration and QC all results found with in the range. | SWO-120599 | No | cleaned the pO2 electrode and replaced the membrane, done zero current calibration in service, done two pint calibration and QC all results found with in the range. | cleaned the pO2 electrode and replaced the membrane, done zero current calibration in service, done two pint calibration and QC all results found with in the range. | Assisted Replacement | Electrode | Membrane | pO2
ABL835 FLEX | Had Jerry Kelly assist with problem. He recommended tightening Cal 1 and Cal 2 ports and replace all black tubing. | Tightened Cal1 and Cal2 ports and replaced all black tubing on top of PH BG module and El Met module. Also flushed out blue air nipple on fluidic module. | D110E9523016D84C | No | Had Jerry Kelly assist with problem. He recommended tightening Cal 1 and Cal 2 ports and replace all black tubing. | Tightened Cal1 and Cal2 ports and replaced all black tubing on top of PH BG module and El Met module. Also flushed out blue air nipple on fluidic module. | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
ABL815 FLEX | Probehalber Tausch des EL Met-Moduls 902-824R0810N010, dabei Kundenelektroden im Gerät belassen, Kal, QC Kunden EL MET Modul 902-824R2323N004 | Probehalber Tausch des EL Met-Moduls 902-824R0810N010, dabei Kundenelektroden im Gerät belassen, Kal, QC Kunden EL MET Modul 902-824R2323N004 | SWO-117280 | No | Sample replacement of the EL Met module 902-824R0810N010, leaving customer electrodes in the device, Kal, QC customers EL MET module 902-824R2323N004 | Sample replacement of the EL Met module 902-824R0810N010, leaving customer electrodes in the device, Kal, QC customers EL MET module 902-824R2323N004 | Revisit required | HW | El/Met Module | Na+
ABL820 FLEX | PH is drifting and will not hold sensitivity or status - changed multiple ref membranes and checked ph electrode. ph electrode looks to have leaked some fluid out of the tip trapped air bubbles | replaced ph electrode and membranes that the customer used for trouble shooting | D110A6C6301F70B9 | No | PH is drifting and will not hold sensitivity or status - changed multiple ref membranes and checked ph electrode. ph electrode looks to have leaked some fluid out of the tip trapped air bubbles | replaced ph electrode and membranes that the customer used for trouble shooting | Assisted Replacement | Electrode | Electrode | pH
ABL815 FLEX | Ca not showing on patient report | created new layout and customer says Ca not showing; restore default ; set up patient reports and QC schedule replacement schedule; interface not crossing over; called IT department large animal also not crosssing over; worked with IT to resolve the issue-- IP address for analyzer configured | D110D763300B3C56 | No | Ca not showing on patient report | created new layout and customer says Ca not showing; restore default ; set up patient reports and QC schedule replacement schedule; interface not crossing over; called IT department large animal also not crosssing over; worked with IT to resolve the issue-- IP address for analyzer configured | Setup | His/lis | His/lis | His/lis
AQT90 FLEX | Datensicherung über Windows durchgeführt Software 8.8.16 mittels Update neu aufgespielt Testkassetten eingelegt und verbrauchte entfernt Durchführung Systemüberprüfung => alles i.O. | Datensicherung über Windows durchgeführt Software 8.8.16 mittels Update neu aufgespielt Testkassetten eingelegt und verbrauchte entfernt Durchführung Systemüberprüfung => alles i.O. | SWO-117363 | No | Data backup performed via Windows software 8.8.16 by means of update replayed test cartridges inserted and consumed remotely performing system verification => everything in the order | Data backup performed via Windows software 8.8.16 by means of update replayed test cartridges inserted and consumed remotely performing system verification => everything in the order | Defect | SW | Windows | Windows
AQT90 FLEX | Naald en needle wash unit vervangen. | Naald en needle wash unit vervangen. | SWO-117229 | No | Needle wash unit. | Needle wash unit. | Defect | HW | Aspiration Module | HCT Sensor
AQT90 FLEX | tighten the motor and refix the inlet door, working properly. | tighten the motor and refix the inlet door, working properly. | SWO-118200 | No | tighten the motor and refix the inlet door, working properly. | tighten the motor and refix the inlet door, working properly. | Adjustment | HW | Cover and Chassis | Inlet Door
ABL90 FLEX | Unable to flush due to blocked inlet probe and inlet gasket rubber being incorrectly positioned. Replaced inlet probe and gasket. Ran fluid transport checks and is now performing ok. | B | SWO-119207 | No | Unable to flush due to blocked inlet probe and inlet gasket rubber being incorrectly positioned. Replaced inlet probe and gasket. Ran fluid transport checks and is now performing ok. |  | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
ABL90 | Cl not giving any values Cal or QC. Dotted line. | Replaced Sensor Interface Bd. | D112567DA6609243 | No | Cl not giving any values Cal or QC. Dotted line. | Replaced Sensor Interface Bd. | Defect | HW | Basic Unit | Sensor Interface PCB
ABL90 | Unit indicating that the SP or SC needs to be replaced.Determined customer still using old syringes instead of Safepico. | Customer is to remove old syringes and replace with Safepico. | D1125F57A63E924A | No | Unit indicating that the SP or SC needs to be replaced. Determined customer still using old syringes instead of Safepico. | Customer is to remove old syringes and replace with Safepico. | User error | Sampler/Silicone | Sampler/Silicone | Sampler/Silicone
ABL90 FLEX | Einsetzen einer neuen Sensorkassette und eines neuen Lösungspacks aus Bestand der Geschäftsstelle. Kalibrationen und Qualitätskontrolle o. K. Bemerkung: die Dienstleistung belieferung von Kunden bei betriebsmittelengpass , Srtikel-Nr. 41-007 muss noch in den Artikelliste aufgenommen werden. | Überprüfung - Spülung Analysator - Service Tests durchgeführt - Funktionsüberprüfung etc. | SWO-117920 | No | Insert a new sensor cartridge and a new solution pack from the office's inventory. Calibrations and quality control o.k. Remark: the service supplied to customers in case of equipment shortage, Srtikel-No. 41-007 has yet to be included in the item list. | Verification - Rinse Analyzer - Service Tests Performed - Functional Verification etc. | Assisted Replacement | SC | SC | pCO2
ABL90 FLEX | Gerät überprüft, Wet Section ausgebaut, Sensorkontakte kontrolliert, Keine weitere Überprüfung möglich, da kein Sensorsimulator im Bestand | Gerät überprüft, Wet Section ausgebaut, Sensorkontakte kontrolliert, Keine weitere Überprüfung möglich, da kein Sensorsimulator im Bestand | SWO-117671 | No | Device checked, Wet Section removed, sensor contacts checked, No further verification possible as no sensor simulator in stock | Device checked, Wet Section removed, sensor contacts checked, No further verification possible as no sensor simulator in stock | No fault found | No fault found | No fault found | No fault found
ABL90 FLEX | Gerät überprüft, Lösungspack erneuert | Gerät überprüft, Lösungspack erneuert | SWO-117672 | No | Device checked, solution pack renewed | Device checked, solution pack renewed | Assisted Replacement | SP | SP | pCO2
ABL90 FLEX | Verbindung eingerichtet, Daten gespeichert, Kal i.O., Datensendung eingestellt, QC-Daten gesendet | Verbindung eingerichtet, Daten gespeichert, Kal i.O., Datensendung eingestellt, QC-Daten gesendet | SWO-117699 | No | Connection set up, data stored, calendar, data sent, QC data sent | Connection set up, data stored, calendar, data sent, QC data sent | Setup | His/lis | His/lis | His/lis
ABL90 FLEX PLUS |  | B | D11B0341E0248BD8 | No |  |  | Defect | HW | Battery module | Battery module
AQT90 FLEX |  | B | SWO-117416 | No |  |  | Defect | HW | Process Module | Optical Unit | 131
ABL835 FLEX | Se limpia módulo de autocheck. Se comprueba funcionamiento del módulo. Se pasan controles de calidad. | B | D119B3F06A0352AA | No | Autocheck module is cleaned. Module operation is checked. Quality controls are passed. |  | Cleaned | HW | AutoCheck Module | Barcode reader
ABL800 BASIC | Rescanned consumables. Ran LS adjust and pump calibration. | Rescanned consumables. Ran LS adjust and pump calibration. | SWO-119503 | No | Rescanned consumables. Ran LS adjust and pump calibration. | Rescanned consumables. Ran LS adjust and pump calibration. | User error | Replacement | Consumables | Consumables
ABL817 FLEX | pCO2 Elektrode ern., Gerät kalibriert, QC´s durchgef. | B | SWO-117500 | No | pCO2 electrode ern., device calibrated, QC's through. |  | Assisted Replacement | Electrode | Membrane | pCO2
ABL800 BASIC | 更换LAC电极膜正常。 | 更换LAC电极膜正常。 | SWO-117394 | No | Replacing the LAC electrode film is normal. | Replacing the LAC electrode film is normal. | Assisted Replacement | Electrode | Membrane | Lac | Sensitivity
ABL815 FLEX | pCO2-Elektrode ersetzt. Kalibrationen und Qualitätskontrollen in Ordnung. | Frau Schilde wird pCO2-Elektrode bestellen. | SWO-117430 | No | pCO2 electrode. calibrations and quality controls in order. | Ms. Schilde will order pCO2 electrode. | Assisted Replacement | Electrode | Electrode | pCO2
ABL825 FLEX | Empfindlichkeit trotz Elektrodentausch wieder unterhalb Grenze. Lac Elektrode gegen alte Elektrode (Schulungsgerät) durch E. Dietzel ersetzt und membraniert. | Empfindlichkeit trotz Elektrodentausch wieder unterhalb Grenze. Lac Elektrode gegen alte Elektrode (Schulungsgerät) durch E. Dietzel ersetzt und membraniert. | SWO-117615 | No | Sensitivity despite electrode replacement again below limit. Lac electrode against old electrode (training device) replaced by E. Dietzel and membraneed. | Sensitivity despite electrode replacement again below limit. Lac electrode against old electrode (training device) replaced by E. Dietzel and membraneed. | Assisted Replacement | Electrode | Membrane | Lac | Sensitivity
AQT90 FLEX | Cleaned internal Dust from Analyzer, opend Punch Doun motor and reinnstalled, wheel calibration done. machine working ok. | Cleaned internal Dust from Analyzer, opend Punch Doun motor and reinnstalled, wheel calibration done. machine working ok. | SWO-118206 | No | Cleaned internal Dust from Analyzer, opend Punch Doun motor and reinnstalled, wheel calibration done. machine working ok. | Cleaned internal Dust from Analyzer, opend Punch Doun motor and reinnstalled, wheel calibration done. machine working ok. | Cleaned | HW | Storage Module | Sensor, Cup | Calibrate_13
AQT90 FLEX |  | B | D1124F19A6484716 | No |  |  | Cleaned | HW | Incubator module | Sensor, Incubator | 3553
AQT90 FLEX |  | B | D1122560A60E4713 | No |  |  | Cleaned | HW | Incubator module | Sensor, Incubator | 3553
ABL80 FLEX BASIC | 更换测试卡正常。 | 更换测试卡正常。 | SWO-117965 | No | The replacement test card is OK. | The replacement test card is OK.
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-117642 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL80 CO-OX | Performed Service Test Program and found the Hemolyzer Cavitation Absorbance out of range. | Replaced the hemolyzer unit. Performed Service Test Program and the Hemolyzer failed with a LED light error. Replaced the hemolyzer again and now passes the test. | D11357261C1AF44F | No | Performed Service Test Program and found the Hemolyzer Cavitation Absorbance out of range. | Replaced the hemolyzer unit. Performed Service Test Program and the Hemolyzer failed with a LED light error. Replaced the hemolyzer again and now passes the test.
ABL90 |  | B | D11315621C3E04F9 | No |  |  | Defect | HW | Inlet module | Inlet gasket
ABL90 FLEX | ソフトウェアバージョンアップ | B | SWO-120695 | Yes | Software version upgrade |  | Re classified RMED | SWU
ABL90 FLEX | SC90 geliefert und installiert, Funktion OK. Reklamationsausdruck generiert. | SC90 geliefert und erneuert, es bestehen weiterhin pCO2 Empf. Probleme | SWO-117714 | No | SC90 delivered and installed, function OK. Complaint expression generated. | SC90 delivered and renewed, pCO2 Empf still exist. Problems | Assisted Replacement | SC | SC | pCO2
ABL90 FLEX | SC90 geliefert und erneuert, es bestehen weiterhin pCO2 Empf. Probleme | SC90 geliefert und erneuert, es bestehen weiterhin pCO2 Empf. Probleme | SWO-117713 | No | SC90 delivered and renewed, pCO2 Empf still exist. Problems | SC90 delivered and renewed, pCO2 Empf still exist. Problems | Replaced | HW | Probe bend | Probe bend
ABL90 | Verified once on site. Analyzer failing level 3 internal QC. | Performed a system calibration on two analyzers in the ICU and a retest of the internal QC. Systems are all now back in service. Worked extensively with the customer, Tech Support and Customer to have the customers consumable supplies delivered overnight as the customer has none in stock. NO MTR required. The syringes that the customer is using are Portex 4588P-2, manufactured by Smiths Medical, possible silicone contamination. Still under review. | D112DBF71C0A4724 | No | Verified once on site. Analyzer failing level 3 internal QC. | Performed a system calibration on two analyzers in the ICU and a retest of the internal QC. Systems are all now back in service. Worked extensively with the customer, Tech Support and Customer to have the customers consumable supplies delivered overnight as the customer has none in stock. NO MTR required. The syringes that the customer is using are Portex 4588P-2, manufactured by Smiths Medical, possible silicone contamination. Still under review. | Assisted Replacement | SC | SC | pCO2
ABL90 FLEX | alle Adressen neu eingetragen und getestet, | alle Adressen neu eingetragen und getestet, | SWO-117921 | No | all addresses are re-entered and tested, | all addresses are re-entered and tested, | Setup | His/lis | His/lis | His/lis
ABL90 FLEX | ヘモライザーユニット、スペクトロメーターユニット交換及びバージョンアップ | B | SWO-119430 | No | Hemolyzer unit, spectrometer unit replacement and upgrade |  | Defect | HW | Oxi module | Hemolyzer
ABL90 FLEX | ヘモライザーユニットおよびバージョンアップ実施。 | B | SWO-120690 | No | Hemolizer unit and version upgrade. |  | Defect | HW | Oxi module | Hemolyzer | Spectrum mismatch
ABL90 FLEX | Überprüfung - Spülung Analysator - Service Tests durchgeführt - Funktionsüberprüfung etc. | Überprüfung - Spülung Analysator - Service Tests durchgeführt - Funktionsüberprüfung etc. | SWO-117865 | No | Verification - Rinse Analyzer - Service Tests Performed - Functional Verification etc. | Verification - Rinse Analyzer - Service Tests Performed - Functional Verification etc. | Cleaned | Fluidic pathway | Fluidic pathway | pCO2
ABL90 FLEX | Anwender hatte bereits Lösungspack erneuert, Fehler blieb jedoch bestehen. Auch nach Tausch der Kassette. Anwender wartet nun auf die neue Lieferung des Lösungspacks. | Anwender hatte bereits Lösungspack erneuert, Fehler blieb jedoch bestehen. Auch nach Tausch der Kassette. Anwender wartet nun auf die neue Lieferung des Lösungspacks. | SWO-118072 | No | User had already renewed solution pack, but errors remained. Even after exchanging the cassette. Users are now waiting for the new delivery of the solution pack. | User had already renewed solution pack, but errors remained. Even after exchanging the cassette. Users are now waiting for the new delivery of the solution pack. | Assisted Replacement | SP | SP | pCO2
ABL90 FLEX PLUS | SP90 XB-03 ersetzt, Temperatur am ABL 26 Grad Funktion zur Zeit OK. | SP90 XB-03 ersetzt, Temperatur am ABL 26 Grad Funktion zur Zeit OK. | SWO-119219 | No | SP90 XB-03 replaced, temperature at ABL 26 degree function currently OK. | SP90 XB-03 replaced, temperature at ABL 26 degree function currently OK. | Assisted Replacement | SP | SP | pCO2 | C high
ABL90 FLEX PLUS |  | B | D1135BBB1C30555C | No |  |  | Replaced | HW | Probe bend | Probe bend
ABL90 FLEX PLUS | replaced faulty auto inlet module | Went to replace PCB but the error still remained, Removed the complete inlet module and tested it on Micu's machine and the error carried over to the Micu's machine. Entire inlet module is faulty. | SWO-118028 | Yes | replaced faulty auto inlet module | Went to replace PCB but the error still remained, Removed the complete inlet module and tested it on Micu's machine and the error carried over to the Micu's machine. Entire inlet module is faulty. | Re classified RMED | Double SWO ?? | See July 14.216
AQT90 FLEX | 更换条码扫描器 | 更换条码扫描器 | SWO-117239 | No | Replace barcode scanner | Replace barcode scanner | Defect | HW | Storage Module | Barcode reader, Storage wheel
AQT90 FLEX | 更换新测试卡修复仪器 | 更换新测试卡修复仪器 | SWO-117637 | No | Replace the new test card repair instrument | Replace the new test card repair instrument | Defect | Consumables | Test Kit | Cartridge | 5784
ABL800 BASIC | Upgradation of ABL800 performed to v6.16, it was found that windows hanged in the middle of the process and upgradation could not be completed. | Replaced Hard Disk from stock, SW v6.15 installed successfully, 2 point calibration performed | SWO-118189 | No | Upgradation of ABL800 performed to v6.16, it was found that windows hanged in the middle of the process and upgradation could not be completed. | Replaced Hard Disk from stock, SW v6.15 installed successfully, 2 point calibration performed | Revisit required | SW | SW | 6.16 | Slow
ABL815 FLEX | looked through data logs to see if there is a trend on reports, on site with Dan Hansen | performed a service dump will submit to Radiometer for techline case | D1135E5E1C27723A | No | looked through data logs to see if there is a trend on reports; on site with Dan Hansen | performed a service dump will submit to Radiometer for techline case | Unknown | Electrode | Electrode | Ca
ABL820 FLEX | Matt Rutrough performed the install on this analyzer. The pCO2 failed level 3 QC and the sensitivity stayed low. We tried a different inlet and pH BG module. | Will be replaced. A tCQC case was opened. | D1130B0F1C62B77B | Yes | Matt Rutrough performed the install on this analyzer. The pCO2 failed level 3 QC and the sensitivity stayed low. We tried a different inlet and pH BG module. | Will be replaced. A tCQC case was opened. | Re classified RMED | OOB
ABL825 FLEX |  | B | D119F0E06A0DD099 | No |  |  | Defect | HW | Fluidic manifold | Pump Tubing
ABL800 BASIC | Go to service mode,calibrate the inlet position but it shows error and inlet position error shows and goes to hold position | Check again service mode,move manually the inlet probe,calibrate again which is passedand resolve the problem | SWO-119095 | Yes | Go to service mode,calibrate the inlet position but it shows error and inlet position error shows and goes to hold position | Check again service mode,move manually the inlet probe,calibrate again which is passedand resolve the problem | Re classified RMED | Installation
AQT90 FLEX |  | B | D113DBE61C2643FB | No |  |  | Adjustment | HW | Aspiration Module | Inlet Wheel
AQT90 FLEX | Has to replace Needle Assebmbly. Part Number:903-171 | B | SWO-118208 | No | Has to replace Needle Assebmbly. Part Number:903-171 |  | Defect | HW | Aspiration Module | Bend needle
ABL80 OSM |  | B | D11325421C4D5BB4 | No
ABL80 FLEX BASIC | 更换新测试卡和接口板修复仪器---jack.chu | 更换新测试卡和接口板修复仪器---jack.chu | SWO-117975 | No | Replace the new test card and interface board repair instrument ---jack.chu | Replace the new test card and interface board repair instrument ---jack.chu
ABL80 FLEX CO-OX | SC korrekt installiert, LP erneuert. Karl und QC sind OK. | B | SWO-118219 | No | SC installed correctly, LP renewed. Karl and QC are OK.
ABL90 FLEX | Checked SP and found the top of the pack had not been pushed down fully. Top of pack pushed down until no more clicks here'd and top of pack is level. Refitted and ABL 90 starting up ok. | Checked SP and found the top of the pack had not been pushed down fully. Top of pack pushed down until no more clicks here'd and top of pack is level. Refitted and ABL 90 starting up ok. | SWO-118074 | No | Checked SP and found the top of the pack had not been pushed down fully. Top of pack pushed down until no more clicks here'd and top of pack is level. Refitted and ABL 90 starting up ok. | Checked SP and found the top of the pack had not been pushed down fully. Top of pack pushed down until no more clicks here'd and top of pack is level. Refitted and ABL 90 starting up ok. | User error | SP Activation | SP Activation | SP Activation | 1343
ABL90 FLEX | Sensor Kassette ersetzt Zwei SC angegeben | B | SWO-118112 | No | Sensor cassette replaces two SC specified |  | Assisted Replacement | SC | SC | pO2
ABL90 FLEX | Déboucher le système avec eau javel (avec un vieux cassette). Test en service NOK ( toujours des problèmes pression et transport) Remplacement aiguille , cassette , PS , joint d'intro , tuyaux et joint : NOK Nettoyage du hemolyseur (R1623N0002) avec eau javel NOK Remplacement du hemolyseur (R0417N006) : test OK Calibration tHb et calibration paramètres + CQ OK | B | SWO-118212 | No | Uncork the system with bleach (with an old cassette). NOK in-service test (always pressure and transport problems) Replacement needle, cassette, PS, intro joint, pipes and seal: NOK Hemolyseur Cleaning (R1623N0002) with NOK bleach Replacement of hemolyseur (R0417N006): test OK Calibration tHb and calibration settings - CQ OK |  | Defect | HW | Oxi module | Hemolyzer
ABL90 | The battery pack has only been installed about three weeks. | Replaced the battery pack module. The new battery indicated a full charge within a few minutes. | D114969D923E7881 | No | The battery pack has only been installed about three weeks. | Replaced the battery pack module. The new battery indicated a full charge within a few minutes. | Defect | HW | Battery module | Battery module
ABL90 | The first few days after installation, the battery fully charged and the green light stayed on steadily. After about one week, the light began to blink and has stayed that way dispite the analyzer being plugged in and turned on the entire time. | Replaced the battery pack module. | D114921C9257787B | No | The first few days after installation, the battery fully charged and the green light stayed on steadily. After about one week, the light began to blink and has stayed that way dispite the analyzer being plugged in and turned on the entire time. | Replaced the battery pack module. | Defect | HW | Battery module | Battery module
ABL90 | Several samples were measured and each time this one analyzer measured much lower than the other three analyzers. | Replaced the hemolyzer and the spectrometer. | D114990392107884 | No | Several samples were measured and each time this one analyzer measured much lower than the other three analyzers. | Replaced the hemolyzer and the spectrometer. | Defect | HW | Oxi module | Hemolyzer
ABL90 FLEX PLUS |  | B | D114B2689261FA9E | Yes |  |  | Assisted Replacement | SC | SC | pCO2 | 375
ABL90 FLEX | 更换新测试卡修复仪器 | 更换新测试卡修复仪器 | SWO-118021 | No | Replace the new test card repair instrument | Replace the new test card repair instrument | Assisted Replacement | SC | SC | pCO2
AQT90 FLEX |  | B | SWO-118024 | No |  |  | Defect | HW | Process Module | Wash unit
AQT90 FLEX |  | B | SWO-117980 | No |  |  | Defect | HW | Cover and Chassis | Barcode reader, display
ABL837 FLEX | Black Seal ring missing on Glucose electrode, also Tubing at 6 position Stand-Off crimped. | Replaced O-ring and Straighten Fluidic tubing. Also reseated Waste and Air valves in Fluidic module. | D114BF5B923EACEF | No | Black Seal ring missing on Glucose electrode, also Tubing at 6 position Stand-Off crimped. | Replaced O-ring and Straighten Fluidic tubing. Also reseated Waste and Air valves in Fluidic module. | Defect | Electrode | Electrode | GLu | Sealing ring missing
ABL827 FLEX | Found build-up in pH and reference measuring chambers | Replaced measuring chamber Replaced pump tubes Replaced pH electrode | D1149D7C920D6F22 | No | Found build-up in pH and reference measuring chambers | Replaced measuring chamber Replaced pump tubes Replaced pH electrode | Cleaned | HW | pH/BG Module | Measuring chamber
ABL800 BASIC | Replaced Hard Disk from stock, SW v6.15 installed successfully, 2 point calibration performed | Replaced Hard Disk from stock, SW v6.15 installed successfully, 2 point calibration performed | SWO-118190 | No | Replaced Hard Disk from stock, SW v6.15 installed successfully, 2 point calibration performed | Replaced Hard Disk from stock, SW v6.15 installed successfully, 2 point calibration performed | Defect | SW | SW | 6.16 | Slow
ABL815 FLEX | FlexQ-Modul ersetzt (S N defektes FlexQ-Modul: 902-753R1145N003) Glu-, Ref.-, Lac- und pO2-Elektroden membraniert, da QC-Fehler | FlexQ-Modul ersetzt (S N defektes FlexQ-Modul: 902-753R1145N003) Glu-, Ref.-, Lac- und pO2-Elektroden membraniert, da QC-Fehler | SWO-118152 | No | FlexQ module replaced (S N defective FlexQ module: 902-753R1145N003) Glu, ref., lac and pO2 electrodes membraneed because QC errors | FlexQ module replaced (S N defective FlexQ module: 902-753R1145N003) Glu, ref., lac and pO2 electrodes membraneed because QC errors | Defect | HW | FlexQ Module | Positioning
ABL835 FLEX | Geräte geprüft, Ventilatorfilter gereinigt, mit Herrn OA Dr. Schmidt gesprochen, die Geräte in einen klimatisierten Raum umzustellen, dies wurde abgelehnt, Haustechnik wurde verständigt, Funktionskontrolle ok | Geräte geprüft, Ventilatorfilter gereinigt, mit Herrn OA Dr. Schmidt gesprochen, die Geräte in einen klimatisierten Raum umzustellen, dies wurde abgelehnt, Haustechnik wurde verständigt, Funktionskontrolle ok | SWO-118065 | No | Equipment tested, fan filter cleaned, talked to Mr. OA Dr. Schmidt, the devices converted to an air-conditioned room, this was rejected, building technology was communicated, functional control ok | Equipment tested, fan filter cleaned, talked to Mr. OA Dr. Schmidt, the devices converted to an air-conditioned room, this was rejected, building technology was communicated, functional control ok | Cleaned | HW | Basic Unit | Fan | High temp
ABL800 BASIC | cleaned the screen and it works properly. in queries sample found that it is due to the sample problem. provide the pre-analytical training to the staff. | cleaned the screen and it works properly. in queries sample found that it is due to the sample problem. provide the pre-analytical training to the staff. | SWO-118204 | No | cleaned the screen and it works properly. in queries sample found that it is due to the sample problem. provide the pre-analytical training to the staff. | cleaned the screen and it works properly. in queries sample found that it is due to the sample problem. provide the pre-analytical training to the staff. | Cleaned | HW | DMS_PC | touch Screen
ABL825 FLEX |  | B | D119F1496A25D09A | No |  |  | Defect | HW | pH/BG Module | Valve
ABL805 FLEX | Frische sowie eingetrocknete und eingefärbte Flüssigkeit unterhalb des Einlassbereiches => FrictionDetectionUnit ist bestellt und wird demnächst ersetzt. | Frische sowie eingetrocknete und eingefärbte Flüssigkeit unterhalb des Einlassbereiches => FrictionDetectionUnit ist bestellt und wird demnächst ersetzt. | SWO-118117 | No | Fresh as well as dried and dyed liquid below the inlet area => FrictionDetectionUnit has been ordered and will be replaced soon. | Fresh as well as dried and dyed liquid below the inlet area => FrictionDetectionUnit has been ordered and will be replaced soon. | Revisit required | HW | Inlet module | Leaking
ABL800 BASIC | Check again service mode,move manually the inlet probe,calibrate again which is passedand resolve the problem | Check again service mode,move manually the inlet probe,calibrate again which is passedand resolve the problem | SWO-119099 | Yes | Check again service mode,move manually the inlet probe,calibrate again which is passedand resolve the problem | Check again service mode,move manually the inlet probe,calibrate again which is passedand resolve the problem | Re classified RMED | double SWO
ABL825 FLEX | AC Modul inklusive Align Carousel komplett abgeglichen, Endkontrolle: 2 AC5+ Messungen, beide OK | AC Modul inklusive Align Carousel komplett abgeglichen, Endkontrolle: 2 AC5+ Messungen, beide OK | SWO-118147 | No | AC module including Align Carousel fully matched, final control: 2 AC5+ measurements, both OK | AC module including Align Carousel fully matched, final control: 2 AC5+ measurements, both OK | Adjustment | HW | AutoCheck Module | Carroussel
ABL825 FLEX | Bei erster automatischer Spülung nach Beendigung des Servicemenüs (AC Funktionen geprüft, eingestellt) Ausstoß von Spüllösung unter hohem Druck vorne aus der Einlassdichtung. Austritt von Spüllösung während des gesamten Spülzyklus. Aufruf des Servicemenüs mit Einlasskalibrierung, bestanden. Neustart, Startphase OK, Endkontrolle: AC5+ Messung, OK | Bei erster automatischer Spülung nach Beendigung des Servicemenüs (AC Funktionen geprüft, eingestellt) Ausstoß von Spüllösung unter hohem Druck vorne aus der Einlassdichtung. Austritt von Spüllösung während des gesamten Spülzyklus. Aufruf des Servicemenüs mit Einlasskalibrierung, bestanden. Neustart, Startphase OK, Endkontrolle: AC5+ Messung, OK | SWO-118149 | No | At the first automatic rinse after completion of the service menu (AC functions checked, adjusted) ejection of flushing solution under high pressure at the front of the inlet seal. Discharge of rinsing solution during the entire rinsing cycle. call of the service menu with inlet calibration. Restart, start phase OK, final control: AC5+ measurement, OK | At the first automatic rinse after completion of the service menu (AC functions checked, adjusted) ejection of flushing solution under high pressure at the front of the inlet seal. Discharge of rinsing solution during the entire rinsing cycle. call of the service menu with inlet calibration. Restart, start phase OK, final control: AC5+ measurement, OK | Defect | HW | Inlet module | Leaking
ABL825 FLEX | Oximetrie überprüft, Halogenlampe dauerhaft EIN, ausschalten nicht möglich, Justage oder Überprüfung im Servicemenü nicht möglich => Neustart des Analysators. Abgleich der Lampenspannung, OxiCompleteTest, OK. Endkontrolle: AC5+ L2 Messung, OK | Oximetrie überprüft, Halogenlampe dauerhaft EIN, ausschalten nicht möglich, Justage oder Überprüfung im Servicemenü nicht möglich => Neustart des Analysators. Abgleich der Lampenspannung, OxiCompleteTest, OK. Endkontrolle: AC5+ L2 Messung, OK | SWO-118121 | No | Oximetry checked, halogen lamp permanently ON, switch off not possible, adjustment or verification in the service menu not possible => Restart of the analyzer. Balance of lamp voltage, OxiCompleteTest, OK. Final control: AC5+ L2 measurement, OK | Oximetry checked, halogen lamp permanently ON, switch off not possible, adjustment or verification in the service menu not possible => Restart of the analyzer. Balance of lamp voltage, OxiCompleteTest, OK. Final control: AC5+ L2 measurement, OK | Defect | HW | Oxi module | Missing trigger
ABL825 FLEX | pO2 Nullstrom bei -0,1 => ZeroCurrentCalibration im Servicemenü durchgeführt. Kalibrationen, OK | pO2 Nullstrom bei -0,1 => ZeroCurrentCalibration im Servicemenü durchgeführt. Kalibrationen, OK | SWO-118158 | No | pO2 Zero current at -0.1 => ZeroCurrentCalibration performed in the service menu. Calibrations, OK | pO2 Zero current at -0.1 => ZeroCurrentCalibration performed in the service menu. Calibrations, OK | Adjustment | Electrode | Electrode | Zero current | pO2
AQT90 FLEX | Cleaned the Process wheel and cross checked with the patient sample, all working good. | Cleaned the Process wheel and cross checked with the patient sample, all working good. | SWO-118116 | No | Cleaned the Process wheel and cross checked with the patient sample, all working good. | Cleaned the Process wheel and cross checked with the patient sample, all working good. | Cleaned | HW | Process Module | Process wheel
ABL80 FLEX BASIC | Guided the technician over the phone on how to bypass the battery pack and after doing so, it was observed that the machine is not getting ON. So the machine is bypassing the AC mains supply and trying to run on the available battery power, which means that it is required to replace the power supply. | Guided the technician over the phone on how to bypass the battery pack and after doing so, it was observed that the machine is not getting ON. So the machine is bypassing the AC mains supply and trying to run on the available battery power, which means that it is required to replace the power supply. | SWO-119177 | No | Guided the technician over the phone on how to bypass the battery pack and after doing so, it was observed that the machine is not getting ON. So the machine is bypassing the AC mains supply and trying to run on the available battery power, which means that it is required to replace the power supply. | Guided the technician over the phone on how to bypass the battery pack and after doing so, it was observed that the machine is not getting ON. So the machine is bypassing the AC mains supply and trying to run on the available battery power, which means that it is required to replace the power supply.
ABL80 FLEX | Cleaned opto sensor, tried to reinstalled SP, Problem remains, PER send for replacement. | Cleaned opto sensor, tried to reinstalled SP, Problem remains, PER send for replacement. | SWO-120501 | No | Cleaned opto sensor, tried to reinstalled SP, Problem remains, PER send for replacement. | Cleaned opto sensor, tried to reinstalled SP, Problem remains, PER send for replacement.
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-120714 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL80 CO-OX | Verifed the problem and tried reloading the 3.14 software. This had no efect. Rebooted the analyzer. This also did not work. Called tech support and was informed to send the data logs to go with a techline case. On 7-22-16 tech support sent me a copy of the COOX key for this unit. | Loaded the COOX key for this unt. Replaced the solution pack twice to get the unit to pass. | D1176BD17E506FF7 | No | Verifed the problem and tried reloading the 3.14 software. This had no efect. Rebooted the analyzer. This also did not work. Called tech support and was informed to send the data logs to go with a techline case. On 7-22-16 tech support sent me a copy of the COOX key for this unit. | Loaded the COOX key for this unt. Replaced the solution pack twice to get the unit to pass.
ABL80 FLEX BASIC | 更换新测试卡修复仪器 | 更换新测试卡修复仪器 | SWO-119239 | No | Replace the new test card repair instrument | Replace the new test card repair instrument
ABL80 FLEX BASIC |  | B | SWO-118760 | No
ABL80 FLEX BASIC |  | B | SWO-119163 | No
ABL80 FLEX BASIC | Cleaned Sc with distilled water and kept in fridge for 30 minutes. Then reinstalled SC and done calibration.Calibration done sucessfully. | Cleaned Sc with distilled water and kept in fridge for 30 minutes. Then reinstalled SC and done calibration.Calibration done sucessfully. | SWO-120475 | No | Cleaned Sc with distilled water and kept in fridge for 30 minutes. Then reinstalled SC and done calibration. Calibration done sucessfully. | Cleaned Sc with distilled water and kept in fridge for 30 minutes. Then reinstalled SC and done calibration. Calibration done sucessfully.
ABL80 FLEX CO-OX | Gerät überprüft, fehlende Schrauben ersetzt, Lösungspackverriegelung kontrolliert und ersetzt, Fließwege geprüft, Sevicetest und tHb Kal durchgeführt | Gerät überprüft, fehlende Schrauben ersetzt, Lösungspackverriegelung kontrolliert und ersetzt, Fließwege geprüft, Sevicetest und tHb Kal durchgeführt | SWO-119336 | No | Device checked, missing screws replaced, solution pack lock checked and replaced, flow paths checked, Sevicetest and tHb Kal performed | Device checked, missing screws replaced, solution pack lock checked and replaced, flow paths checked, Sevicetest and tHb Kal performed
ABL80 CO-OX | Error code 2106- possible fluid blockage | Preformed liquid sensor cal, CO-OX test, checked fludic flow, replaced sensor cassette | D1184AB4F40B076A | No | Error code 2106- possible fluid blockage | Preformed liquid sensor cal, CO-OX test, checked fludic flow, replaced sensor cassette
ABL90 FLEX | Hr. kühn hat nach Telefonat den SP getauscht --> Fehler bleibt bestehen--> Hr Kühn informiert, dass es auf der Station sehr warm ist. --> Station besucht --> Problem mit Station besprochen Weitergabe def Information an Hr. Schwarz | Hr. kühn hat nach Telefonat den SP getauscht --> Fehler bleibt bestehen--> Hr Kühn informiert, dass es auf der Station sehr warm ist. --> Station besucht --> Problem mit Station besprochen Weitergabe def Information an Hr. Schwarz | SWO-119304 | No | After a phone call, Mr. Kühn changed the SP --> error remains--> Mr. Kühn informed that it is very warm on the station. --> Station visited --> Problem with station discussed passing def information to Mr. Schwarz | After a phone call, Mr. Kühn changed the SP --> error remains--> Mr. Kühn informed that it is very warm on the station. --> Station visited --> Problem with station discussed passing def information to Mr. Schwarz | No fault found | No fault found | No fault found | pCO2 | Temperature
ABL90 |  | B | D1186FF2F4050025 | No |  |  | Defect | HW | Main PCB | CPU
ABL90 FLEX | Gerät überprüft, Datenbank gelöscht, Software neu installiert, Datensicherung zurückgespielt, Verbrauchsmaterial reinstalliert | Gerät überprüft, Datenbank gelöscht, Software neu installiert, Datensicherung zurückgespielt, Verbrauchsmaterial reinstalliert | SWO-119335 | Yes | Device checked, database deleted, software reinstalled, backup backed up, consumables reinstalled | Device checked, database deleted, software reinstalled, backup backed up, consumables reinstalled | Re classified RMED | SW update
ABL90 FLEX | Einlassdichtung und Einlasssonde wurden vom Anwender erneuert. Der Aufforderung das Lösungspack zu erneuern ist der Anwender nicht nachgekommen. Habe das Lösungspack aus dem Analysator genommen und wieder eingesetzt. Analysator fährt wieder hoch. Danach beginnt der Analysator mit der Konditionierung der SC. Dr.Schneider wird die SC selber tauschen falls der ABL dazu auffordert. | Einlassdichtung und Einlasssonde wurden vom Anwender erneuert. Der Aufforderung das Lösungspack zu erneuern ist der Anwender nicht nachgekommen. Habe das Lösungspack aus dem Analysator genommen und wieder eingesetzt. Analysator fährt wieder hoch. Danach beginnt der Analysator mit der Konditionierung der SC. Dr.Schneider wird die SC selber tauschen falls der ABL dazu auffordert. | SWO-119192 | Yes | Inlet seal and inlet probe have been renewed by the user. The user did not comply with the request to renew the solution pack. Have taken the solution pack from the analyzer and re-inserted it. Analyzer goes up again. After that, the analyzer starts conditioning the SC. Dr.Schneider will replace the SC itself if the ABL requests it. | Inlet seal and inlet probe have been renewed by the user. The user did not comply with the request to renew the solution pack. Have taken the solution pack from the analyzer and re-inserted it. Analyzer goes up again. After that, the analyzer starts conditioning the SC. Dr.Schneider will replace the SC itself if the ABL requests it. | Re classified RMED | Telephone support
ABL90 FLEX | Lösungspack wurde bereits am 22.07.16 vom Anwender getauscht. Danach war zwei Tage alles ok. Nun wieder pCO2 und Glukose-Fehler in der QC- Messung. Nochmals Kassette ausgetauscht. Danach alles im Bereich. | Lösungspack wurde bereits am 22.07.16 vom Anwender getauscht. Danach war zwei Tage alles ok. Nun wieder pCO2 und Glukose-Fehler in der QC- Messung. Nochmals Kassette ausgetauscht. Danach alles im Bereich. | SWO-119443 | No | Solution pack has already been replaced by the user on 22.07.16. After that, two days everything was ok. Now again pCO2 and glucose errors in the QC measurement. Replaced the cassette again. After that, everything in the area. | Solution pack has already been replaced by the user on 22.07.16. After that, two days everything was ok. Now again pCO2 and glucose errors in the QC measurement. Replaced the cassette again. After that, everything in the area. | Assisted Replacement | SC | SC | pCO2
ABL90 FLEX | Fehlersuche, Gerät aus, Hemolyzer gew., Gerät ein, Startphase abgewartet, Kal. i.O., QC weiter Fehler, Hemolyzer wieder zurückgebaut, div. Schlauchmaterial gew., sowie Servicepr. durchgef., LP gew., Kal. + tHb i.O., QC = pCO2-Fehler Level C, fixe Zielwerte eingestellt u. aktiviert, Startphase, Kal. i.O., QC Level a u. C pCO2-Fehler, SC gew., Abbruch d. Arbeiten u. an Kenny übergeben | Fehlersuche; Gerät aus; Hemolyzer gew.; Gerät ein; Startphase abgewartet; Kal. i.O.; QC weiter Fehler; Hemolyzer wieder zurückgebaut; div. Schlauchmaterial gew., sowie Servicepr. durchgef.; LP gew.; Kal. + tHb i.O.; QC = pCO2-Fehler Level C; fixe Zielwerte eingestellt u. aktiviert; Startphase; Kal. i.O.; QC Level a u. C pCO2-Fehler; SC gew.; Abbruch d. Arbeiten u. an Kenny übergeben | SWO-119358 | Yes | Troubleshooting; device off; Hemolyzer. device; start-up phase awaited; Kal. i.O.; QC continues errors; Hemolyzer restored; various hose material, as well as servicepr. through it; LP. Kal. + tHb i.O.; QC = pCO2 error level C; fixed target values set and enabled; Start-up phase; Kal. i.O.; QC Level a and C pCO2 errors; SC. Demolition of works and handed over to Kenny | Troubleshooting; device off; Hemolyzer. device; start-up phase awaited; Kal. i.O.; QC continues errors; Hemolyzer restored; various hose material, as well as servicepr. through it; LP. Kal. + tHb i.O.; QC = pCO2 error level C; fixed target values set and enabled; Start-up phase; Kal. i.O.; QC Level a and C pCO2 errors; SC. Demolition of works and handed over to Kenny | Re classified RMED | OOB
ABL90 FLEX PLUS |  | B | D11840F2F456CBF2 | Yes |  |  | Re classified RMED | Training
ABL90 FLEX | 更换新测试卡后解决 | 更换新测试卡后解决 | SWO-118280 | No | Resolve after replacing a new test card | Resolve after replacing a new test card
ABL825 FLEX |  | B | SWO-119174 | No |  |  | Defect | HW | Oxi module | Oxi module
ABL825 FLEX | Display ersetzt. | Display ersetzt. | SWO-119348 | No | Display replaced. | Display replaced. | Defect | HW | DMS_PC | Touch Screen
ABL815 FLEX | pCO2 out on level 2 QC- error code 0090 Upon arrival on site, pCO2 was back in but at the lower end of range (37.9 -lowest acceptable 37.7). | Replaced dummy electrode (CI). Replaced pH BG tube pump tubing. Ordered new Ref and pCO2 membranes as well. | D11875D4F4226539 | No | pCO2 out on level 2 QC- error code 0090 Upon arrival on site, pCO2 was back in but at the lower end of range (37.9 -lowest acceptable 37.7). | Replaced dummy electrode (CI). Replaced pH BG tube pump tubing. Ordered new Ref and pCO2 membranes as well. | Assisted Replacement | Electrode | Membrane | pCO2
ABL810 FLEX | Datalogs gezogen. Anwender Einweisung vorgeschlagen. Anwender melden sich wegen Einweisungstermin. Kalibrationen und Qualitätskontrolle o. K. | Datalogs gezogen. Anwender Einweisung vorgeschlagen. Anwender melden sich wegen Einweisungstermin. Kalibrationen und Qualitätskontrolle o. K. | SWO-119391 | No | Datalogs. User instruction suggested. Users log in for briefing date. Calibrations and quality control o.K. | Datalogs. User instruction suggested. Users log in for briefing date. Calibrations and quality control o.K. | Revisit required | Electrode | Electrode | pCO2
ABL810 FLEX | Reinigung der Messkammer und der pH Elektrode und Erneuern des unteren Elektrodendichtrings. Anwendern Hanhabung in so einem Fall erläutert, damit sie selber die Aktion durchführen können. Kalibrationen und Qualitätskontrolle o. K. | Reinigung der Messkammer und der pH Elektrode und Erneuern des unteren Elektrodendichtrings. Anwendern Hanhabung in so einem Fall erläutert, damit sie selber die Aktion durchführen können. Kalibrationen und Qualitätskontrolle o. K. | SWO-119392 | No | Cleaning of the measuring chamber and the pH electrode and renewing the lower electrode sealing ring. users are told how to do the action themselves. Calibrations and quality control o.K. | Cleaning of the measuring chamber and the pH electrode and renewing the lower electrode sealing ring. users are told how to do the action themselves. Calibrations and quality control o.K. | Assisted Replacement | Electrode | Electrode | pH | Sealing ring
ABL810 FLEX | Datalogs gezogen und an den Technischen Support von Radiometer geschickt. Kalibrationen und Qualitätskontrolle o. K. | Datalogs gezogen und an den Technischen Support von Radiometer geschickt. Kalibrationen und Qualitätskontrolle o. K. | SWO-119390 | No | datalogs and sent to Radiometer Technical Support. Calibrations and quality control o.K. | datalogs and sent to Radiometer Technical Support. Calibrations and quality control o.K. | Revisit required | Electrode | Electrode | pCO2
ABL810 FLEX | Daten und Service Dump gezogen. Daten an Technischen Support der Firma Radiometer gesandt. Kalibrationen und Qualitätskontrolle o. K. | Daten und Service Dump gezogen. Daten an Technischen Support der Firma Radiometer gesandt. Kalibrationen und Qualitätskontrolle o. K. | SWO-119389 | No | Data and service dump. data sent to technical support of Radiometer. Calibrations and quality control o.K. | Data and service dump. data sent to technical support of Radiometer. Calibrations and quality control o.K. | Revisit required | Electrode | Electrode | pCO2
ABL825 FLEX | Remplacé le détecteur de friction du module aspiration e.o Calibration e.o contrôle de qualité e.o | Remplacé le détecteur de friction du module aspiration e.o Calibration e.o contrôle de qualité e.o | SWO-119170 | No | Replaced the friction sensor of the e.o calibration e.o calibration module e.o quality control e.o | Replaced the friction sensor of the e.o calibration e.o calibration module e.o quality control e.o | Defect | HW | Inlet module | Friction detector unit
ABL800 BASIC | Done 2 Point calibration on ABL800 and checked with QC. Fond that results are within range. | Done 2 Point calibration on ABL800 and checked with QC. Fond that results are within range. | SWO-120476 | No | Done 2 Point calibration on ABL800 and checked with QC. Fond that results are within range. | Done 2 Point calibration on ABL800 and checked with QC. Fond that results are within range. | Cleaned | Fluidic pathway | El/Met Module | Measuring chamber
ABL825 FLEX | Touchscreen reagiert mit Aussetzern, teilweise ist sekunden- bis minutenlang keine Eingabe mehr möglich. Nach Neustart vorerst wieder normale Funktion. Monitor wird bestellt. | Touchscreen reagiert mit Aussetzern; teilweise ist sekunden- bis minutenlang keine Eingabe mehr möglich. Nach Neustart vorerst wieder normale Funktion. Monitor wird bestellt. | SWO-119330 | No | Touchscreen responds with dropouts; sometimes it is not possible to enter for seconds to minutes. After rebooting, normal function resumes for the time being. Monitor is ordered. | Touchscreen responds with dropouts; sometimes it is not possible to enter for seconds to minutes. After rebooting, normal function resumes for the time being. Monitor is ordered. | Defect | HW | DMS_PC | touch Screen
AQT90 FLEX | El equipo tiene bloqueada la tarjeta electronica, se desbloquea y se carga el software completo, se realizan calibraciones y blancos, funcionando correctamente. | Carga completa del programa. | D11D251ECC11A1AC | No | The equipment is locked the electronic card, it unlocks and the entire software is loaded, calibrations are performed and white, working properly. | Full load of the program. | Defect | SW | SW | SW
AQT90 FLEX | Replaced the new Optical unit and process wheel. done the Factory reference calibration all working good. | Cleaned the Process wheel and cross checked with the patient sample, all working good. | SWO-120533 | No | Replaced the new Optical unit and process wheel. done the Factory reference calibration all working good. | Cleaned the Process wheel and cross checked with the patient sample, all working good. | Defect | HW | Process Module | Optical Unit | 131
AQT90 FLEX | Replaced needle assembly. Performed needle calibration. | Has to replace Needle Assebmbly. Part Number:903-171 | SWO-120525 | No | Replaced needle assembly. Performed needle calibration. | Has to replace Needle Assebmbly. Part Number:903-171 | Defect | HW | Aspiration Module | Bend needle
ABL80 FLEX BASIC | 更换测试卡正常。 | 更换测试卡正常。 | SWO-119690 | No | The replacement test card is OK. | The replacement test card is OK.
ABL90 |  | B | D11A08226A31AFC9 | No |  |  | Defect | HW | Inlet module | Inlet gasket
ABL90 FLEX PLUS | Tubing 841-811 vervangen. | Tubing 841-811 vervangen. | SWO-119510 | No | Tubing 841-811 vervangen. | Tubing 841-811 vervangen. | Defect | HW | Oxi module | Pinch valve tube
ABL90 FLEX PLUS |  | B | D119B8F76A4262B9 | Yes |  |  | Re classified RMED | Double SWO | SC | pCO2
ABL90 FLEX PLUS |  | B | D11C5180562BFF1C | No |  |  | Assisted Replacement | SC | SC | pCO2
ABL90 FLEX | 更换测试卡正常。 | 更换测试卡正常。 | SWO-119724 | No | The replacement test card is OK. | The replacement test card is OK. | Assisted Replacement | SC | SC | pCO2 | Sensitivity
ABL835 FLEX | オキシメトリーモジュール交換 | B | SWO-119683 | No | Oxymetry module replacement |  | Defect | HW | Oxi module | Lamp unit
ABL835 FLEX | Testet med forskellige TipCap Inlet modul skiftet | Testet med forskellige TipCap Inlet modul skiftet | SWO-119571 | No | Tested with different TipCap Inlet module shifted | Tested with different TipCap Inlet module shifted | Defect | HW | Inlet module | Cannot penetrate tip cap
ABL825 FLEX | Eine neue Cl. Elektrode eingebaut, Parameter aktiviert und neu gestartet. | Eine neue Cl. Elektrode eingebaut, Parameter aktiviert und neu gestartet. | SWO-119719 | No | A new Cl. electrode installed, parameters activated and restarted. | A new Cl. electrode installed, parameters activated and restarted. | Assisted Replacement | Electrode | Electrode | Cl
ABL815 FLEX | Followup to SWO BA0P4P3C2 | Replaced Ref membranes and pCo2 membranes with new lot. Replaced electrode after levels on QC #2 started dropping to borderline. | D119F1616A21BDD5 | No | Followup to SWO BA0P4P3C2 | Replaced Ref membranes and pCo2 membranes with new lot. Replaced electrode after levels on QC #2 started dropping to borderline. | Assisted Replacement | Electrode | Membrane | pCO2
ABL825 FLEX | Leihweise neues E'Lytemodul installiert, sw installiert, checkout and adjustment durchgeführt, Gerät mehrfach kalibriert und mit QCs überprüft, OK Leihmodul: 902-824R0287N003 | Leihweise neues E'Lytemodul installiert, sw installiert, checkout and adjustment durchgeführt, Gerät mehrfach kalibriert und mit QCs überprüft, OK Leihmodul: 902-824R0287N003 | SWO-119673 | No | On loan new E'Lyte module installed, sw installed, checkout and adjustment carried out, device calibrated multiple times and checked with QCs, OK rental module: 902-824R0287N003 | On loan new E'Lyte module installed, sw installed, checkout and adjustment carried out, device calibrated multiple times and checked with QCs, OK rental module: 902-824R0287N003 | Defect | HW | El/Met Module | Glu
ABL825 FLEX | Messkammer und Schlauchsystem geprüft und gereinigt, Glu. Elektrode neu membraniert, Funktionskontrolle ok | B | SWO-119768 | No | Measuring chamber and hose system tested and cleaned, Glu. Electrode re-membraneed, function control ok |  | Assisted Replacement | Electrode | Membrane | Glu | Leakage
ABL800 BASIC | flushed all the tubing and connect the tube properly, calibration done 2-4 times and run the sample, now working fine. | flushed all the tubing and connect the tube properly, calibration done 2-4 times and run the sample, now working fine. | SWO-120523 | No | flushed all the tubing and connect the tube properly, calibration done 2-4 times and run the sample, now working fine. | flushed all the tubing and connect the tube properly, calibration done 2-4 times and run the sample, now working fine. | Cleaned | Fluidic pathway | Fluidic pathway | Clot
ABL815 FLEX | Beim Eintreffen vor Ort TS ohne Reaktion bei Betätigung, der Fehler sei bereits zuvor mehrfach aufgetreten. Monitorstecker im laufenden Betrieb gezogen und erneut gesteckt - TS arbeitet wieder wie vorgesehen. Neuer Monitor wird bestellt. | Beim Eintreffen vor Ort TS ohne Reaktion bei Betätigung; der Fehler sei bereits zuvor mehrfach aufgetreten. Monitorstecker im laufenden Betrieb gezogen und erneut gesteckt - TS arbeitet wieder wie vorgesehen. Neuer Monitor wird bestellt. | SWO-119512 | No | When arriving on site TS without reaction during actuation; the error had already occurred several times before. Monitor plug pulled and replugged during operation - TS works as intended again. New monitor is ordered. | When arriving on site TS without reaction during actuation; the error had already occurred several times before. Monitor plug pulled and replugged during operation - TS works as intended again. New monitor is ordered. | Revisit required | HW | DMS_PC | touch Screen
ABL825 FLEX | Einlassnadel überprüft, ist nach oben abgebogen geknickt => Nadel ersetzt. Inlet-Calibration im Servicemenü, OK | Einlassnadel überprüft, ist nach oben abgebogen geknickt => Nadel ersetzt. Inlet-Calibration im Servicemenü, OK | SWO-119611 | No | Inlet needle checked, bent upwards => needle replaced. Inlet-Calibration in the service menu, OK | Inlet needle checked, bent upwards => needle replaced. Inlet-Calibration in the service menu, OK | Defect | HW | Inlet module | Probe bend
ABL815 FLEX | Einlassnadel ersetzt. Mehrere Proteinentfernungen in Folge durchgeführt, OxiCompleteTest, OK, fünf x Kal 1 in Folge => OK mit Driften bis max. 6,5 Bei Abschluss-QC Messung erneut Spectrum Mismatch => Oxi Modul ausgebaut, zerlegt und Hemolyzer Unit [908-534 R4258N001]ersetzt und abgeglichen. Neustart, 2 x Kal 1, Service-tHb-Kal, Kalibrationen, Funktionstest, Endkontrolle: 3 Level AC5+ gemessen => Alle Werte OK | Immer wiederkehrend hohe tHb Driften bei Kal => Oximodul aus noch nicht ausgeliefertem ABL 754R2279N004 ausgebaut und hier eingebaut. >Defektes Modul: 902-833R2033N001< Kalibrationen, AC5+ Messung, OK | SWO-119558 | No | Inlet needle replaced. Several protein removals performed in a row, OxiCompleteTest, OK, five x Kal 1 in a row => OK with drifts up to max. 6.5 At completion QC measurement again Spectrum Mismatch => Oxi module removed, disassembled and hemolyzer unit [908-534 R4258N001] replaced and matched. Restart, 2 x Kal 1, Service-tHb-Kal, Calibrations, Function Test, Final Control: 3 Level AC5+ Measured => All Values OK | Constantly high tHb drifting at Kal => Oximodul from ABL 754R2279N004 not yet delivered and installed here. >Defective module: 902-833R2033N001" calibrations, AC5+ measurement, OK | Defect | HW | Oxi module | Hemolyzer
AQT90 FLEX | to replace needle , and Needle sensor as sensor check after replacing needle failed with error code 38 , this might be due to spilled solutions from bent needle | to replace needle , and Needle sensor as sensor check after replacing needle failed with error code 38 , this might be due to spilled solutions from bent needle | SWO-120004 | No | to replace needle , and Needle sensor as sensor check after replacing needle failed with error code 38 , this might be due to spilled solutions from bent needle | to replace needle , and Needle sensor as sensor check after replacing needle failed with error code 38 , this might be due to spilled solutions from bent needle | Defect | HW | Aspiration Module | Bend needle | 4794
AQT90 FLEX | cleaned the analyser with blank cartridge and cleaning solution. | cleaned the analyser with blank cartridge and cleaning solution. | SWO-120712 | No | cleaned the analyser with blank cartridge and cleaning solution. | cleaned the analyser with blank cartridge and cleaning solution. | Cleaned | HW | Aspiration Module | HCT Sensor
AQT90 FLEX | Proben wurde vor kurzem abgebrochen mit der Meldung zu wenig Probe für Messung verfügbar und Dunkelstrommessungsfehler. Deswegen Kundin geraten darauf zu schauen das genügend Probe in dem Röhrchen vorhanden ist. Der Fehler bei der Dunkelstrommessung wurde nur einmal ausgegeben und Herr Brand aus der Technik hat mir mitgeteilt das das an der Temperatur im Raum ( zu warme Umgebung) liegen könnte. Frau Scheipl hat deswegen eine Ventilator aufgestellt und die Rollos herunter gelassen da der Anslysator direkt am Fenster steht. Ich habe vor Ort eine Systemüberprüfung durchgeführt welche i.O. war. Es wurde während meiner Anwesenheit noch einige Patientenmessungen durchgeführt welche auch einwandfrei durchgelaufen sind. Am Nachmittag wird Herr Hein sowieso nochmal vorbei kommen und mit Frau Scheipl reden und QCs durchführen. AQT ist Messbereit. | Proben wurde vor kurzem abgebrochen mit der Meldung zu wenig Probe für Messung verfügbar und Dunkelstrommessungsfehler. Deswegen Kundin geraten darauf zu schauen das genügend Probe in dem Röhrchen vorhanden ist. Der Fehler bei der Dunkelstrommessung wurde nur einmal ausgegeben und Herr Brand aus der Technik hat mir mitgeteilt das das an der Temperatur im Raum ( zu warme Umgebung) liegen könnte. Frau Scheipl hat deswegen eine Ventilator aufgestellt und die Rollos herunter gelassen da der Anslysator direkt am Fenster steht. Ich habe vor Ort eine Systemüberprüfung durchgeführt welche i.O. war. Es wurde während meiner Anwesenheit noch einige Patientenmessungen durchgeführt welche auch einwandfrei durchgelaufen sind. Am Nachmittag wird Herr Hein sowieso nochmal vorbei kommen und mit Frau Scheipl reden und QCs durchführen. AQT ist Messbereit. | SWO-119887 | No | Samples were recently aborted with the message too little sample available for measurement and dark current measurement errors. Therefore, customers advise to look that there is enough sample in the tube. The error in the dark current measurement was only issued once and Mr. Brand from the technical told me that this could be due to the temperature in the room (too warm environment). Ms. Scheipl set up a fan and left the blinds down as the analyzer is right next to the window. I did a system check on site which was the only thing I did. During my presence, some patient measurements were carried out which were also carried out flawlessly. In the afternoon, Mr. Hein will come back anyway and talk to Ms. Scheipl and perform QCs. AQT is ready to measure. | Samples were recently aborted with the message too little sample available for measurement and dark current measurement errors. Therefore, customers advise to look that there is enough sample in the tube. The error in the dark current measurement was only issued once and Mr. Brand from the technical told me that this could be due to the temperature in the room (too warm environment). Ms. Scheipl set up a fan and left the blinds down as the analyzer is right next to the window. I did a system check on site which was the only thing I did. During my presence, some patient measurements were carried out which were also carried out flawlessly. In the afternoon, Mr. Hein will come back anyway and talk to Ms. Scheipl and perform QCs. AQT is ready to measure. | Performance | Temperature | Temperature | Temperature | Dark
ABL80 FLEX BASIC | 更换试剂包正常。 | 更换试剂包正常。 | SWO-119739 | No | The replacement kit is OK. | The replacement kit is OK.
ABL90 FLEX | Kunden einen Lösungspack für ABL90 geliefert, da Lösungspack verfallen, Verfallsdatum 11.07.16, Kunde hat keinen Lösungspack mehr vor Ort, Kunden über anfallende Kosten informiert. | Kunden einen Lösungspack für ABL90 geliefert, da Lösungspack verfallen, Verfallsdatum 11.07.16, Kunde hat keinen Lösungspack mehr vor Ort, Kunden über anfallende Kosten informiert. | SWO-119972 | Yes | Customers delivered a solution pack for ABL90 as solution pack expires, expiry date 11.07.16, customer no longer has a solution pack on site, informs customers about costs incurred. | Customers delivered a solution pack for ABL90 as solution pack expires, expiry date 11.07.16, customer no longer has a solution pack on site, informs customers about costs incurred. | Re classified RMED | Logistical upSWO
ABL90 FLEX | ABL90 zur weiteren Überprüfung in die GS mitgenommen. | Sensor Interface PCB ausgebaut, Kontakte gereinigt und justiert, Funktion im Service überprüft, OK. ABL am 28.7. ausgeliefert und messbereit aufgestellt. | SWO-119981 | No | ABL90 for further review in the GS. | Sensor Interface PCB removed, contacts cleaned and adjusted, function checked in service, OK. ABL delivered on 28.7. and set up ready for measurement. | Defect | HW | Basic Unit | Sensor Interface PCB
ABL90 FLEX | Gerät aufgestellt. Einlaufphase. QC-Fehler pCO2 nach Wechsel Lösungspack. Mehrere Kalibrationen und Qualitätskontrollen. | Gerät aufgestellt. Einlaufphase. QC-Fehler pCO2 nach Wechsel Lösungspack. Mehrere Kalibrationen und Qualitätskontrollen. | SWO-120056 | Yes | device. Run-in phase. QC error pCO2 after change solution pack. Multiple calibrations and quality controls. | device. Run-in phase. QC error pCO2 after change solution pack. Multiple calibrations and quality controls. | Re classified RMED | Re installed |  |  | 163
ABL90 FLEX | Alkatrészek cseréje, Folyadékrendszer áztatása hipóban, Szervíztesztek, Oldatcsere teszt. A oldatpakk nem instalálható megfeleően, flow selector nem tud rendesen kalibrálni. ************************* A készülék javítása a helyszinen nem megoldható, műszer elszállítva a központi szervízbe ************************** | B | SWO-119898 | No | Replacement of parts, liquid system soaking in HiPo, service tests, change of solution test. The solution pack should not be properly calibrated, flow selector can not correctly calibrate. Repair of the appliance is not possible on the site, the instrument is shipped to the central service * * * * * * * * * * * * * * * * * * * * * * * * * * * * * * * * * * * |  | Revisit required | Analyzer removed | Analyzer removed | Analyzer removed
ABL90 | On 7 19 customer was instructed to replaces cassette. problem wasn't resolved. Upon arrival an IVK-R5 was ran and found both parameters repeating, but level 2 and 3 reading high outside package insert. | Flushed unit and reran IVK | D11B3B0EE053378B | No | On 7 19 customer was instructed to replaces cassette. problem wasn't resolved. Upon arrival an IVK-R5 was ran and found both parameters repeating, but level 2 and 3 reading high outside package insert. | Flushed unit and reran IVK | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
ABL90 | Unit does not recognize a USB drive in download mode. Customer is using Interleaved 2 of 5 with a check digit. | Reseated USB connections and unit would detect a USB drive. Downloaded data. Set unit to show results of samples out of the reportable range. Tried different barcode settings available but unit still not scanning ID badges. | D11B180BE0339985 | No | Unit does not recognize a USB drive in download mode. Customer is using Interleaved 2 of 5 with a check digit. | Reseated USB connections and unit would detect a USB drive. Downloaded data. Set unit to show results of samples out of the reportable range. Tried different barcode settings available but unit still not scanning ID badges. | Setup | USB | USB | USB
ABL90 FLEX PLUS | Power Supply erneuert ( Nummer Defektteik P013N454) Sensor Cassette Bay Unit erneuert ( Nummer Altteil R0280N003). Sensor Interface PCB erneuert. tHb Kalibration (100,5%) vorgenommen. Demokassette eingelegt. | Power Supply erneuert ( Nummer Defektteik P013N454) Sensor Cassette Bay Unit erneuert ( Nummer Altteil R0280N003). Sensor Interface PCB erneuert. tHb Kalibration (100,5%) vorgenommen. Demokassette eingelegt. | SWO-119929 | Yes | Power Supply renewed (number Defektteik P013N454) Sensor Cassette Bay Unit renewed (number old part R0280N003). Sensor Interface PCB. tHb calibration (100.5%) Made. Demo cassette inserted. | Power Supply renewed (number Defektteik P013N454) Sensor Cassette Bay Unit renewed (number old part R0280N003). Sensor Interface PCB. tHb calibration (100.5%) Made. Demo cassette inserted. | Re classified RMED | Service training?
ABL90 FLEX PLUS | Arrived on site and attended analyser to confirm customer reported fault. Check gasket did not have any signs of damage or blockage. Checked probe did not have any curves or bends in it which there were none. Ran an inlet calibration in service program but this did not help. Replaced probe with new as a precaution. Tested gasket return in service program and found that it felt like something was grinding as the gasket was pushed back and released. Removed full inlet module and found that there was a dark line ground into the plastic. Checked inlet module and noticed that one of the locating pins was protruding further then the other. Tried to adjust this by screwing it further down into position but the adjustment was not a lot and not enough to resolve the issue. This pin is rubbing on the side of the analyser as the gasket is moved. Compared what we found to the analyser in another ward and both pins were set at the same height on the second analyser and did not touch the side at all. Ordered a new inlet mode for replacement. Second site visit to fit required. | suspect the internal spring is binding. escalated to FSE. | SWO-119815 | No | Arrived on site and attended analyser to confirm customer reported fault. Check gasket did not have any signs of damage or blockage. Checked probe did not have any curves or bends in it which there were none. Ran an inlet calibration in service program but this did not help. Replaced probe with new as a precaution. Tested gasket return in service program and found that it felt like something was grinding as the gasket was pushed back and released. Removed full inlet module and found that there was a dark line ground into the plastic. Checked inlet module and noticed that one of the locating pins was protruding further then the other. Tried to adjust this by screwing it further down into position but the adjustment was not a lot and not enough to resolve the issue. This pin is rubbing on the side of the analyser as the gasket is moved. Compared what we found to the analyser in another ward and both pins were set at the same height on the second analyser and did not touch the side at all. Ordered a new inlet mode for replacement. Second site visit to fit required. | suspect the internal spring is binding. escalated to FSE. | Revisit required | HW | Auto Inlet module | Auto Inlet module | Being returned
ABL90 FLEX | 更换新测试卡修复仪器 | 更换新测试卡修复仪器 | SWO-119773 | No | Replace the new test card repair instrument | Replace the new test card repair instrument
AQT90 FLEX |  | B | SWO-119703 | No |  |  | Defect | HW | Process Module | Optical Unit | 131
ABL827 FLEX | Checked that the electrode connector was making a good contact on the pCO2 electrode and made sure there was no air between the tip of the electrode and the membrane. This made no difference so I replaced the electrode connector with one of the unused ones on the MET2 module which again made no difference. The membrane was replaced with a different lot number and when this didn't work the electrode was replaced. After replacing the electrode the pCO2 status and sensitivity values were OK. Decontamination performed because it was due in a few days and the Cl, Lac, Glu & Crea membranes were replaced. | Checked that the electrode connector was making a good contact on the pCO2 electrode and made sure there was no air between the tip of the electrode and the membrane. This made no difference so I replaced the electrode connector with one of the unused ones on the MET2 module which again made no difference. The membrane was replaced with a different lot number and when this didn't work the electrode was replaced. After replacing the electrode the pCO2 status and sensitivity values were OK. Decontamination performed because it was due in a few days and the Cl, Lac, Glu & Crea membranes were replaced. | SWO-119931 | No | Checked that the electrode connector was making a good contact on the pCO2 electrode and made sure there was no air between the tip of the electrode and the membrane. This made no difference so I replaced the electrode connector with one of the unused ones on the MET2 module which again made no difference. The membrane was replaced with a different lot number and when this didn't work the electrode was replaced. After replacing the electrode the pCO2 status and sensitivity values were OK. Decontamination performed because it was due in a few days and the Cl, Lac, Glu & Crea membranes were replaced. | Checked that the electrode connector was making a good contact on the pCO2 electrode and made sure there was no air between the tip of the electrode and the membrane. This made no difference so I replaced the electrode connector with one of the unused ones on the MET2 module which again made no difference. The membrane was replaced with a different lot number and when this didn't work the electrode was replaced. After replacing the electrode the pCO2 status and sensitivity values were OK. Decontamination performed because it was due in a few days and the Cl, Lac, Glu & Crea membranes were replaced. | Assisted Replacement | Electrode | Electrode | pCO2
ABL800 FLEX | Checked and cleaned electrode chamber. Cleaned Rinse Path. found Calibration solution 2 was empty , replaced it. Execute the 2 point calibration. Checked the Printer also Now ABL800 status is OK. | B | SWO-120190 | No | Checked and cleaned electrode chamber. Cleaned Rinse Path. found Calibration solution 2 was empty , replaced it. Execute the 2 point calibration. Checked the Printer also Now ABL800 status is OK. |  | Assisted Replacement | Consumables | Solution bottle | Solutions bottle
ABL835 FLEX | Have tried reference and Ca++ membranes Also tried referenc and Ca++ electrodes | Found a conbination that reads lower but not at the mean | D11AD648E011D299 | No | Have tried reference and Ca++ membranes Also tried referenc and Ca++ electrodes | Found a conbination that reads lower but not at the mean | Assisted Replacement | Electrode | Electrode | Reference
ABL815 FLEX | this has been an on going problem since install, setup from 6.13 was transfered over to 6.16 after windows and 6.16 was loaded | reloaded windows 7 and 6.16 software; setup all reports and info on analyzer from scratch; set up interface Talked to IT about remote support | D11ACA87E01AE07E | No | this has been an on going problem since install; setup from 6.13 was transfered over to 6.16 after windows and 6.16 was loaded | reloaded windows 7 and 6.16 software; setup all reports and info on analyzer from scratch; set up interface Talked to IT about remote support | Defect | SW | 6.16 | Ca++
ABL825 FLEX | Elektrode erneut (durch funktionsfähige gebrauchte)ersetzt und membraniert. Kalibrationen. | Elektrode erneut (durch funktionsfähige gebrauchte)ersetzt und membraniert. Kalibrationen. | SWO-119909 | No | electrode again (by functional used ones) and membranes. Calibrations. | electrode again (by functional used ones) and membranes. Calibrations. | Assisted Replacement | Electrode | Membrane | Lac
ABL825 FLEX | Diverse Eingabeversuche des Anmeldecodes auf dem Anmelde Bildschirm => Alle vertikalen Berührungen im Bereich der Felder 0,1,4,7 vom oberen bis zum unteren Rand sind OHNE Funktion. Nach Ziehen und Stecken des Monitorkabels Funktion vorerst wieder gegeben, anschließend unwillkürlich ohne äußeren Einfluss ausgelöste unrhythmische Tastentöne alle paar Sekunden in Folge => ABL Kaltstart, Funktion nun vorerst wieder OK. | Besuch durch Rolf Bollmann: Keinerlei Eingaben am TS mehr möglich; Funktion vorübergenehd wiederhergestellt nach zweimaligem Abziehen und Stecken des Monitorkabels. | SWO-119913 | No | Various input attempts of the login code on the login screen => All vertical touches in the range of fields 0,1,4,7 from the top to the bottom are WITHOUT function. After pulling and plugging the monitor cable function is given again for the time being, then involuntarily triggered without external influence unrhythmic key tones every few seconds in a row => ABL cold start, function now OK again for the time being. | Visit by Rolf Bollmann: No more entries at the TS possible; Function bygenehd restored after twice unplugging and plugging the monitor cable. | Defect | HW | DMS_PC | touch Screen
ABL815 FLEX | Immer wiederkehrend hohe tHb Driften bei Kal => Oximodul aus noch nicht ausgeliefertem ABL 754R2279N004 ausgebaut und hier eingebaut. >Defektes Modul: 902-833R2033N001< Kalibrationen, AC5+ Messung, OK | Immer wiederkehrend hohe tHb Driften bei Kal => Oximodul aus noch nicht ausgeliefertem ABL 754R2279N004 ausgebaut und hier eingebaut. >Defektes Modul: 902-833R2033N001< Kalibrationen, AC5+ Messung, OK | SWO-119822 | No | Constantly high tHb drifting at Kal => Oximodul from ABL 754R2279N004 not yet delivered and installed here. >Defective module: 902-833R2033N001" calibrations, AC5+ measurement, OK | Constantly high tHb drifting at Kal => Oximodul from ABL 754R2279N004 not yet delivered and installed here. >Defective module: 902-833R2033N001" calibrations, AC5+ measurement, OK | Defect | HW | Oxi module | Oxi module
AQT90 FLEX |  | escalated to FSE | SWO-119763 | Yes |  | escalated to FSE | Re classified RMED | Phone support |  |  | 3264
AQT90 FLEX | Scanner gereinigt und der Station gezeigt wie sie sich bei diesem Fehler schnell selber helfen können. | Scanner gereinigt und der Station gezeigt wie sie sich bei diesem Fehler schnell selber helfen können. | SWO-120011 | No | scanner and showed the station how to help themselves with this error quickly. | scanner and showed the station how to help themselves with this error quickly. | Cleaned | HW | Storage Module | Barcode reader, Storage wheel
AQT90 FLEX | Done the alignment of motors, now working fine. | Done the alignment of motors, now working fine. | SWO-120521 | No | Done the alignment of motors, now working fine. | Done the alignment of motors, now working fine. | Adjustment | HW | Cover and Chassis | Inlet Door
ABL80 FLEX | Cleaned Internal dust from Analyzer control PCB, reconnect connections of opto sensor of Door latch and Try to Reinstalled, Problem remains . | B | SWO-120265 | No | Cleaned Internal dust from Analyzer control PCB, reconnect connections of opto sensor of Door latch and Try to Reinstalled, Problem remains .
ABL80 FLEX BASIC | 更换测试卡修复仪器 | 更换测试卡修复仪器 | SWO-120002 | No | Replace the test card repair instrument | Replace the test card repair instrument
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-120716 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL80 FLEX BASIC |  | B | SWO-119945 | No
ABL80 FLEX CO-OX | Gerät überprüft, i-Button Kabel erneuert, Lösungspack ersetzt, Sensorkassette und Schlauchsystem durchgespült, Schlauchkit erneuert, Windows Uhrzeit korregiert, Kal und QC o.k, | Gerät überprüft, i-Button Kabel erneuert, Lösungspack ersetzt, Sensorkassette und Schlauchsystem durchgespült, Schlauchkit erneuert, Windows Uhrzeit korregiert, Kal und QC o.k, | SWO-120347 | No | Device checked, i-button cable renewed, solution pack replaced, sensor cassette and hose system flushed through, hose kit renewed, Windows time corrected, Kal and QC ok, | Device checked, i-button cable renewed, solution pack replaced, sensor cassette and hose system flushed through, hose kit renewed, Windows time corrected, Kal and QC ok,
ABL90 FLEX | Nettoyage préventive du circuit de liquide avec eau , javel (vieux cassette) , eau , alcool ( vieux cassette) et eau. Le cassette est nettoyé sépare avec eau. J'ai pas constaté un bouchage dans le système . Calibrations et CQ : OK Test échantillons seringue OK Test échantillons capillaire | B | SWO-120108 | No | Preventive cleaning of the liquid circuit with water, bleach (old cassette), water, alcohol (old cassette) and water. The cassette is cleaned separates with water. I didn't notice a plug in the system. Calibrations and QC: OK Test syringe samples OK Hair Sample Testing |  | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
ABL90 FLEX PLUS | Arrived on site and attended analyser. Checked through logs and found that the analyser had shut down on the 17 07 2016 at 22:52 error code 1349 analyser shutdown due to low battery. Checked for further signs of this fault after and found that there was another occurrence of it on the 22 07 2016. Logs showed that the analyser was showing a low battery warning approximately 15 minutes before shutdown which is to be expected if the analyser had been unplugged and left in standby. Entered service program and tested analyser battery status. Battery showing as connected and on mains. Battery level at 100% when plugged in. Status showed as connected but not on mains when I unplugged the mains cable. I tested battery life over 15 minutes and level dropped to expected levels over this time for a new battery. Plugged mains back in and refreshed status. Analyser showing connected to mains again and battery charging correctly. Battery function looks to be working correctly. Suspect possibility that the analyser has been unplugged and allowed to run out of battery by mistake. Customer to monitor. | Arrived on site and attended analyser. Checked through logs and found that the analyser had shut down on the 17 07 2016 at 22:52 error code 1349 analyser shutdown due to low battery. Checked for further signs of this fault after and found that there was another occurrence of it on the 22 07 2016. Logs showed that the analyser was showing a low battery warning approximately 15 minutes before shutdown which is to be expected if the analyser had been unplugged and left in standby. Entered service program and tested analyser battery status. Battery showing as connected and on mains. Battery level at 100% when plugged in. Status showed as connected but not on mains when I unplugged the mains cable. I tested battery life over 15 minutes and level dropped to expected levels over this time for a new battery. Plugged mains back in and refreshed status. Analyser showing connected to mains again and battery charging correctly. Battery function looks to be working correctly. Suspect possibility that the analyser has been unplugged and allowed to run out of battery by mistake. Customer to monitor. | SWO-120193 | No | Arrived on site and attended analyser. Checked through logs and found that the analyser had shut down on the 17 07 2016 at 22:52 error code 1349 analyser shutdown due to low battery. Checked for further signs of this fault after and found that there was another occurrence of it on the 22 07 2016. Logs showed that the analyser was showing a low battery warning approximately 15 minutes before shutdown which is to be expected if the analyser had been unplugged and left in standby. Entered service program and tested analyser battery status. Battery showing as connected and on mains. Battery level at 100% when plugged in. Status showed as connected but not on mains when I unplugged the mains cable. I tested battery life over 15 minutes and level dropped to expected levels over this time for a new battery. Plugged mains back in and refreshed status. Analyser showing connected to mains again and battery charging correctly. Battery function looks to be working correctly. Suspect possibility that the analyser has been unplugged and allowed to run out of battery by mistake. Customer to monitor. | Arrived on site and attended analyser. Checked through logs and found that the analyser had shut down on the 17 07 2016 at 22:52 error code 1349 analyser shutdown due to low battery. Checked for further signs of this fault after and found that there was another occurrence of it on the 22 07 2016. Logs showed that the analyser was showing a low battery warning approximately 15 minutes before shutdown which is to be expected if the analyser had been unplugged and left in standby. Entered service program and tested analyser battery status. Battery showing as connected and on mains. Battery level at 100% when plugged in. Status showed as connected but not on mains when I unplugged the mains cable. I tested battery life over 15 minutes and level dropped to expected levels over this time for a new battery. Plugged mains back in and refreshed status. Analyser showing connected to mains again and battery charging correctly. Battery function looks to be working correctly. Suspect possibility that the analyser has been unplugged and allowed to run out of battery by mistake. Customer to monitor. | No fault found | No fault found | No fault found | Battery OK
ABL837 FLEX | Cleaned Creatinine module chambers and performed zero current calibration. Replaced all solutions, but sens.1 was too low. Replaced membrane. | 0208 216 4049 Katie Heanie. customer has rescanned QCs and issue is still present escalated too Dave. | SWO-119987 | No | Cleaned Creatinine module chambers and performed zero current calibration. Replaced all solutions, but sens.1 was too low. Replaced membrane. | 0208 216 4049 Katie Heanie. customer has rescanned QCs and issue is still present escalated too Dave. | Assisted Replacement | Electrode | Membrane | Crea
ABL835 FLEX | Luftweg gespült Alle Ventile von ELE Met und PH BG Modul ausgebaut und überprüft sowie gereinigt. Mehrere Kalibrierungen durchgeführt. QCs für alle Level durchgeführt. Testmessungen durchgeführt. ABL wieder Messbereit. | Sr. Heike bittet um einen Rückruf und erneuten Service. | SWO-120263 | No | Airway flushed All valves removed from ELE Met and PH BG module and checked and cleaned. Multiple calibrations performed. QCs for all levels. test measurements. ABL ready to measure again. | Sr. Heike asks for a callback and reservice. | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
ABL800 BASIC |  | B | SWO-120072 | No |  |  | Defect | HW | AutoCheck Module | Barcode reader
ABL825 FLEX |  | B | D11C34475653E6E0 | Yes |  |  | Re classified RMED | PM
ABL800 BASIC | Replaced pCO2 and pO2 membrane. Performed calibration. Run RADIOMETER QC Level 3 and found all parameter within range. Analyzer working satisfactorily. | Replaced pCO2 and pO2 membrane. Performed calibration. Run RADIOMETER QC Level 3 and found all parameter within range. Analyzer working satisfactorily. | SWO-120102 | No | Replaced pCO2 and pO2 membrane. Performed calibration. Run RADIOMETER QC Level 3 and found all parameter within range. Analyzer working satisfactorily. | Replaced pCO2 and pO2 membrane. Performed calibration. Run RADIOMETER QC Level 3 and found all parameter within range. Analyzer working satisfactorily. | Assisted Replacement | Electrode | Membrane | pCO2
ABL825 FLEX |  | B | D11C3333560D62C3 | Yes |  |  | Re classified RMED | PM
ABL825 FLEX | Bei allen drei Analysatoren die richtigen Barcodes für Level rot und gelb eingescannt. Vorab wurden die neue Barcodes vom Intrant heruntergeladen. Unterer Elektrodendichtring der pH Elektrode erneuert, Elektrode gereinigt. Referenz Elektrode neu membraniert mit Los Nr. 1949. Einsatz einer neuen pCO2 Elektrode. Mitarbeiter der-R-IT Abteilung informiert zur Korrektur der pCO2 Werte in Radiance. Kalibrationen und Qualitätskontrolle o. K. | Bei allen drei Analysatoren die richtigen Barcodes für Level rot und gelb eingescannt. Vorab wurden die neue Barcodes vom Intrant heruntergeladen. Unterer Elektrodendichtring der pH Elektrode erneuert, Elektrode gereinigt. Referenz Elektrode neu membraniert mit Los Nr. 1949. Einsatz einer neuen pCO2 Elektrode. Mitarbeiter der-R-IT Abteilung informiert zur Korrektur der pCO2 Werte in Radiance. Kalibrationen und Qualitätskontrolle o. K. | SWO-120198 | No | Scanned for all three analyzers the correct barcodes for level red and yellow. The new barcodes were downloaded from the Intrant in advance. Lower electrode sealing ring of the pH electrode renewed, electrode cleaned. Reference electrode re-membraneed with lot No. 1949. Use of a new pCO2 electrode. Employees of the R-IT department inform about the correction of the pCO2 values in Radiance. Calibrations and quality control o.K. | Scanned for all three analyzers the correct barcodes for level red and yellow. The new barcodes were downloaded from the Intrant in advance. Lower electrode sealing ring of the pH electrode renewed, electrode cleaned. Reference electrode re-membraneed with lot No. 1949. Use of a new pCO2 electrode. Employees of the R-IT department inform about the correction of the pCO2 values in Radiance. Calibrations and quality control o.K. | Assisted Replacement | Electrode | Electrode | pCO2
ABL825 FLEX |  | B | D11D0F9CCC459D2B | No |  |  | Defect | HW | Inlet module | Leaking
ABL815 FLEX | Bildschirm komplett getauscht mit Funktionstest. | Bildschirm komplett getauscht mit Funktionstest. | SWO-120140 | No | Screen completely replaced with functional test. | Screen completely replaced with functional test. | Defect | HW | DMS_PC | touch Screen
ABL825 FLEX | Bildschirm komplett ersetzt mit Funktionstest. | Bildschirm komplett ersetzt mit Funktionstest. | SWO-120136 | No | Screen completely replaced with function test. | Screen completely replaced with function test. | Defect | HW | DMS_PC | touch Screen
ABL825 FLEX | Meldung mit OK weg gedrückt, vorerst keine weiteren Meldungen. | Meldung mit OK weg gedrückt; vorerst keine weiteren Meldungen. | SWO-120141 | No | Message with OK pressed away; no further reports for the time being. | Message with OK pressed away; no further reports for the time being. | Defect | SW | Windows | 6.16
ABL825 FLEX | Bildschirm komplett ersetzt mit Funktionstest. | Bildschirm komplett ersetzt mit Funktionstest. | SWO-120137 | No | Screen completely replaced with function test. | Screen completely replaced with function test. | Defect | HW | DMS_PC | touch Screen
AQT90 FLEX | Gerät überprüft, Diverse Dinge überprüft, Clean install durchgeführt | B | SWO-120346 | Yes | Device checked, Various things checked, Clean install performed |  | Re classified RMED | OOB
ABL80 FLEX | Cleaned SC, Done Zeroing and reinstalled, problem resolve. | Cleaned SC, Done Zeroing and reinstalled, problem resolve. | SWO-120534 | No | Cleaned SC, Done Zeroing and reinstalled, problem resolve. | Cleaned SC, Done Zeroing and reinstalled, problem resolve.
ABL80 CO-OX | pCO2 reading outside package insert and S8302 failing on system cycle | Replaced solution pack from customers aready used packs. | D11D99B7CC3170BF | No | pCO2 reading outside package insert and S8302 failing on system cycle | Replaced solution pack from customers aready used packs.
ABL90 | TRIED MY OWN USB SUCCESSFULLY | INSTRUCTED THE CUSTOMER TO PLACE DRIVE IN THE USB PORT PRIOR TO GOING INTO DISK FUNCTIONS AND TO NEVER OPEN UP THE CVS FILE PRIOR TO UPLOADING TO WDC AS IT WILL CORRUPT THE DATA | D11D41ECCC1545F5 | No | TRIED MY OWN USB SUCCESSFULLY | INSTRUCTED THE CUSTOMER TO PLACE DRIVE IN THE USB PORT PRIOR TO GOING INTO DISK FUNCTIONS AND TO NEVER OPEN UP THE CVS FILE PRIOR TO UPLOADING TO WDC AS IT WILL CORRUPT THE DATA | User error | USB | USB | USB
ABL90 FLEX PLUS | Fehlersuche durchgeführt. Monitor erneuer-ohne Erfolg. Speicherkarte erneuert-ohne Erfolg Netzteil erneuert- ohne Erfolg. CPU einheit bestellt und das Gerät zwecks Reparatur mitgenommen. | Fehlersuche durchgeführt. Monitor erneuer-ohne Erfolg. Speicherkarte erneuert-ohne Erfolg Netzteil erneuert- ohne Erfolg. CPU einheit bestellt und das Gerät zwecks Reparatur mitgenommen. | SWO-120447 | No | Troubleshooting. Monitor renew-without success. Memory card renewed-without success power supply renewed- without success. CPU unit and took the device with you for repair. | Troubleshooting. Monitor renew-without success. Memory card renewed-without success power supply renewed- without success. CPU unit and took the device with you for repair. | Revisit required | Analyzer removed | DMS_PC | CPU
ABL90 FLEX PLUS | Schlauchsystem gespült, Funktionskontrolle ok | Schlauchsystem gespült, Funktionskontrolle ok | SWO-120395 | No | Hose system flushed, function control ok | Hose system flushed, function control ok | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
ABL90 FLEX PLUS |  | B | D11D4DF1CC291BAF | No |  |  | Defect | HW | CF Card | CF Card
ABL90 FLEX PLUS | Arrived on site and attended analyser. When we got there, the analyser was asking for the solution pack to be changed. Changed this and the analyser started a 1 hour start up. Waited for analyser to finish startup and then entered service programs. Lifted inlet to open position, removed inlet module and replaced with new. Ran inlet calibration and tested that the inlet opened and closed to the correct positions. Sensor cassette was due to be replaced tomorrow and the glucose was out of calibration. Customer will change sensor cassette. Analyser running well. | Arrived on site and attended analyser. When we got there, the analyser was asking for the solution pack to be changed. Changed this and the analyser started a 1 hour start up. Waited for analyser to finish startup and then entered service programs. Lifted inlet to open position, removed inlet module and replaced with new. Ran inlet calibration and tested that the inlet opened and closed to the correct positions. Sensor cassette was due to be replaced tomorrow and the glucose was out of calibration. Customer will change sensor cassette. Analyser running well. | SWO-120465 | No | Arrived on site and attended analyser. When we got there, the analyser was asking for the solution pack to be changed. Changed this and the analyser started a 1 hour start up. Waited for analyser to finish startup and then entered service programs. Lifted inlet to open position, removed inlet module and replaced with new. Ran inlet calibration and tested that the inlet opened and closed to the correct positions. Sensor cassette was due to be replaced tomorrow and the glucose was out of calibration. Customer will change sensor cassette. Analyser running well. | Arrived on site and attended analyser. When we got there, the analyser was asking for the solution pack to be changed. Changed this and the analyser started a 1 hour start up. Waited for analyser to finish startup and then entered service programs. Lifted inlet to open position, removed inlet module and replaced with new. Ran inlet calibration and tested that the inlet opened and closed to the correct positions. Sensor cassette was due to be replaced tomorrow and the glucose was out of calibration. Customer will change sensor cassette. Analyser running well. | Defect | HW | Auto Inlet module | Auto Inlet module
AQT90 FLEX |  | B | SWO-120299 | No |  |  | Defect | HW | Storage Module | Punch up | 5297
ABL835 FLEX | Fehlersuche wahren Jahreswartung mit durchgeführt. Nach Jahreswartung und Testmessungen mit Frau Müller erstmal keine Fehler feststellbar. Frau Müller wird alles genau beobachten da sie am Wochenende Dienst hat. ABL ist nach Jahreswartung Messbereit. | Fehlersuche wahren Jahreswartung mit durchgeführt. Nach Jahreswartung und Testmessungen mit Frau Müller erstmal keine Fehler feststellbar. Frau Müller wird alles genau beobachten da sie am Wochenende Dienst hat. ABL ist nach Jahreswartung Messbereit. | SWO-120365 | No | Troubleshooting true annual maintenance performed with. After annual maintenance and test measurements with Ms. Müller, no errors can be detected for the time being. Ms. Müller will be watching everything closely as she is on duty at the weekend. ABL is ready to measure after annual maintenance. | Troubleshooting true annual maintenance performed with. After annual maintenance and test measurements with Ms. Müller, no errors can be detected for the time being. Ms. Müller will be watching everything closely as she is on duty at the weekend. ABL is ready to measure after annual maintenance. | Defect | HW | PM Kit | PM Kit | 719
ABL815 FLEX | Datensicherung, ABL-Software deinstalliert und Ordner ABL gelöscht, Neuinstallation der ABL-Software v6.16, Datensicherung zurückgespielt, Neustart, Konfiguration, Probelauf, Kal u. QK ok, Radianceanbindung ok | B | SWO-120293 | No | Data backup, ABL software uninstalled and folder ABL deleted, reinstallation of ABL software v6.16, data backup played back, restart, configuration, test run, Cal & QK ok, radiance connection ok |  | Defect | SW | 6.16 | 6.16
ABL825 FLEX | adjustment neu durchgeführt, keine Ursache für die Intensitätsänderung erkannt | adjustment neu durchgeführt, keine Ursache für die Intensitätsänderung erkannt | SWO-120624 | No | adjustment re-performed, no cause of the intensity change detected | adjustment re-performed, no cause of the intensity change detected | Adjustment | HW | Oxi module | Lamp unit
AQT90 FLEX | Nadel erneuert. Flüssigkeit im Gerät beseitigt. Cartridge Lasergehäuse gereinigt. Systemüberprüfung ok. Systemreinigung ok. LQC's durchgeführt. | Nadel erneuert. Flüssigkeit im Gerät beseitigt. Cartridge Lasergehäuse gereinigt. Systemüberprüfung ok. Systemreinigung ok. LQC's durchgeführt. | SWO-120403 | No | Needle renewed. liquid in the device. Cartridge laser housing cleaned. System check ok. System cleaning ok. LQC's. | Needle renewed. liquid in the device. Cartridge laser housing cleaned. System check ok. System cleaning ok. LQC's. | Defect | HW | Aspiration Module | Bend needle | 4794
ABL90 FLEX | 溶液パック→カセット交換 | 溶液パック→カセット交換 | SWO-125338 | No | Solution pack → Cassette replacement | Solution pack → Cassette replacement | Assisted Replacement | SC | Lac | Lac
ABL90 FLEX | 更换新测试卡修复仪器 | 更换新测试卡修复仪器 | SWO-121034 | No | Replace the new test card repair instrument | Replace the new test card repair instrument
ABL80 FLEX CO-OX | Fehlersuche, Hemolyzer verstopft->gespült, Verstopfung gelöst, LP Verriegelung defekt->erneuert, nach Neustart kein Flüssigkeitsaustritt mehr vorhanden, LS Justage fehlgeschlagen, Liquid Sensor defekt->erneuert System-und Funktionskontrolle i.O. | Fehlersuche, Hemolyzer verstopft->gespült, Verstopfung gelöst, LP Verriegelung defekt->erneuert, nach Neustart kein Flüssigkeitsaustritt mehr vorhanden, LS Justage fehlgeschlagen, Liquid Sensor defekt->erneuert System-und Funktionskontrolle i.O. | SWO-120923 | No | Troubleshooting, Hemolyzer clogged->flushed->flushed, blockage solved, LP locking defective->renewed, after restart no liquid outlet available, LS Adjustment failed, Liquid Sensor defective->renewed system and function control i.O. | Troubleshooting, Hemolyzer clogged->flushed->flushed, blockage solved, LP locking defective->renewed, after restart no liquid outlet available, LS Adjustment failed, Liquid Sensor defective->renewed system and function control i.O.
ABL80 FLEX CO-OX | Eine neue SC eingesetzt und das System fährt ohne Fehlermeldung hoch. Vermutlich verursacht die alte SC einen Kurzschluss der durch einen Bruch in der Keramik verursacht wurde. Kunde ordert neue SC da keine mehr vor Ort. Der Analysator wird jetzt noch häufiger Kalibrieren. | Eine neue SC eingesetzt und das System fährt ohne Fehlermeldung hoch. Vermutlich verursacht die alte SC einen Kurzschluss der durch einen Bruch in der Keramik verursacht wurde. Kunde ordert neue SC da keine mehr vor Ort. Der Analysator wird jetzt noch häufiger Kalibrieren. | SWO-120852 | No | A new SC is inserted and the system boots up without an error message. Presumably, the old SC causes a short circuit caused by a break in the ceramic. Customer no longer orders new SC on site. The analyzer will now calibrate even more frequently. | A new SC is inserted and the system boots up without an error message. Presumably, the old SC causes a short circuit caused by a break in the ceramic. Customer no longer orders new SC on site. The analyzer will now calibrate even more frequently.
ABL90 FLEX | Fixe Zielwerte waren bereits eingerichtet. Kundin Tabelle mit Fixen Zielwerten ausgehändigt. Kundin wird Zielwerte in Laborsystem eintragen und sich mit Poctopus in Verbindung setzen. | Fixe Zielwerte waren bereits eingerichtet. Kundin Tabelle mit Fixen Zielwerten ausgehändigt. Kundin wird Zielwerte in Laborsystem eintragen und sich mit Poctopus in Verbindung setzen. | SWO-120932 | No | Fixed target values were already set up. Customer table with fixed target values handed out. Customer will enter target values in laboratory system and contact Poctopus. | Fixed target values were already set up. Customer table with fixed target values handed out. Customer will enter target values in laboratory system and contact Poctopus. | No fault found | Fixed assigned values | Allready setup | Values printed for customer
ABL90 FLEX | Fixe Zielwerte waren bereits eingerichtet. Kundin Tabelle mit Fixen Zielwerten ausgehändigt. Kundin wird Zielwerte in Laborsystem eintragen und sich mit Poctopus in Verbindung setzen. tHb Kalibration vorgenommen (101,1%). Sensorkassette und LP erneuert. Kalibration und QC's vorgenommen. Bitte LP und SC bestellen, da keine mehr vorhanden sind. | Frau Selzer angerufen. Frau Selzer wurde von Chefarzt angesprochen. Frau Selzer erklärt, dass bei Ankunft auch keine Personalanmeldung vorhanden war und ich auf Sation und vom Labor nicht angesprochen wurde die Personalanmeldung zu konfigurieren. Es wurden nur die Zielwerte Kontrolliert. Es wurden keine Konfigurationsänderung vorgenommen. Frau Selzer wird dies weiterleiten. Kundin gesagt, dass Konfigurationsänderung kostenpflichtig ist. Sie wird sich bei Bedarf wieder melden. Bin nicht mehr umgedreht, da ich schon 90 km weg war. | SWO-120933 | No | Fixed target values were already set up. Customer table with fixed target values handed out. Customer will enter target values in laboratory system and contact Poctopus. tHb calibration (101.1%). Sensor cassette and LP. calibration and QC's. Please order LP and SC as there are no more. | Ms. Selzer called. Ms Selzer was approached by the chief medical officer. Ms. Selzer explains that there was no staff registration on arrival and I wasn't approached by Sation and the lab to configure the staff registration. Only the target values were checked. No configuration change has been made. Mrs Selzer will pass this on. Customer said that configuration change is chargeable. She will come back if necessary. I'm no longer upside down because I was already 90 km away. | No fault found | Fixed assigned values | Allready setup | Values printed for customer
ABL90 FLEX | Explorer im Hintergrund geschlossen. Zeiteinstellung eingegeben .Speicherbefehl ausgeführt. | Explorer im Hintergrund geschlossen. Zeiteinstellung eingegeben .Speicherbefehl ausgeführt. | SWO-120881 | No | Explorer closed in the background. Time setting entered . Memory command executed. | Explorer closed in the background. Time setting entered . Memory command executed. | Setup | Time | Time | Time
ABL90 FLEX | A megjavított műszer visszaszállítása és üzembehelyezése, Kalibrálás, QC mérések | B | SWO-120888 | No | Redelivery and commissioning of repaired instrument, calibration, QC measurements |  | Defect | HW | Flow selector PCB | Flow selector PCB
ABL90 FLEX | SCを再度交換。サービスプログラムにて各部をチェック | SCを再度交換。サービスプログラムにて各部をチェック | SWO-121010 | No | Replaced the SC again. Check each part in the service program | Replaced the SC again. Check each part in the service program | Assisted Replacement | SC | SC | pCO2
ABL90 FLEX PLUS | Gerät nach Reparatur wieder ausgeliefert. | Gerät nach Reparatur wieder ausgeliefert. | SWO-120893 | No | device after repair. | device after repair. | Unknown | Analyzer removed | Analyzer removed | Analyzer removed
ABL90 FLEX PLUS |  | B | D12151F12E2A9A82 | Yes |  |  | Re classified RMED | SW update | SW update | SW update
ABL90 FLEX PLUS |  | B | D1215C862E3E2B47 | No |  |  | Defect | HW | Printer | Printer
ABL90 FLEX | 换了夹紧阀管修复仪器 | 换了夹紧阀管修复仪器 | SWO-120743 | No | Change of clamping valve repair instrument | Change of clamping valve repair instrument | Defect | HW | Oxi module | Pinch valve tube
ABL90 | CHEQUEO DE NEUMATICA, LIMPIEZA DE CIRCUITOS, SUSTITUCIÓN DE INLET GASKET, CONECTOR GASKET, SONDA, SUSTITUCIÓN DE DOS SC, SUSTITUCIÓN DE SP, EJECUCIÓN DE PROGRAMAS DE SERVICIO EN VARIAS OCASIONES, CALIBRACIONES, PASADO QC. | CHEQUEO DE NEUMATICA, LIMPIEZA DE CIRCUITOS, SUSTITUCIÓN DE INLET GASKET, CONECTOR GASKET, SONDA, SUSTITUCIÓN DE DOS SC, SUSTITUCIÓN DE SP, EJECUCIÓN DE PROGRAMAS DE SERVICIO EN VARIAS OCASIONES, CALIBRACIONES, PASADO QC. | D12115572E0F60FA | No | NEUMATIC CHECK, CIRCUIT CLEANING, GASKET INLET SUBSTITUTE, GASKET CONNECTOR, SONDA, TWO SC SUBSTITUTE, SP SUBSTITUTE, SERVICE PROGRAM EXECUTION IN SEVERAL OccasionS, CALIBRATIONS, PASSED QC. | NEUMATIC CHECK, CIRCUIT CLEANING, GASKET INLET SUBSTITUTE, GASKET CONNECTOR, SONDA, TWO SC SUBSTITUTE, SP SUBSTITUTE, SERVICE PROGRAM EXECUTION IN SEVERAL OccasionS, CALIBRATIONS, PASSED QC. | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway | & Inlet gasket + SC + SP
ABL827 FLEX | The ABL825 that has been running in the workshop is showing much better performance on the Lactate so it was decided to swap the analyser with this one. Analyser s n 754R0796N006 was brought in and put in place of this one. The Crea module from this analyser was fitted to the replacement and demo keys installed so the Crea module could be enabled. The set-up was copied and loaded on to the replacement analyser. The Glu, Lac & Ref membranes were replaced because they were all over 30 days old. The pCO2 membrane was also replace because it perforated when trying to remove an air bubble trapped under the electrode. | The ABL825 that has been running in the workshop is showing much better performance on the Lactate so it was decided to swap the analyser with this one. Analyser s n 754R0796N006 was brought in and put in place of this one. The Crea module from this analyser was fitted to the replacement and demo keys installed so the Crea module could be enabled. The set-up was copied and loaded on to the replacement analyser. The Glu, Lac & Ref membranes were replaced because they were all over 30 days old. The pCO2 membrane was also replace because it perforated when trying to remove an air bubble trapped under the electrode. | SWO-120968 | No | The ABL825 that has been running in the workshop is showing much better performance on the Lactate so it was decided to swap the analyser with this one. Analyser s n 754R0796N006 was brought in and put in place of this one. The Crea module from this analyser was fitted to the replacement and demo keys installed so the Crea module could be enabled. The set-up was copied and loaded on to the replacement analyser. The Glu, Lac & Ref membranes were replaced because they were all over 30 days old. The pCO2 membrane was also replace because it perforated when trying to remove an air bubble trapped under the electrode. | The ABL825 that has been running in the workshop is showing much better performance on the Lactate so it was decided to swap the analyser with this one. Analyser s n 754R0796N006 was brought in and put in place of this one. The Crea module from this analyser was fitted to the replacement and demo keys installed so the Crea module could be enabled. The set-up was copied and loaded on to the replacement analyser. The Glu, Lac & Ref membranes were replaced because they were all over 30 days old. The pCO2 membrane was also replace because it perforated when trying to remove an air bubble trapped under the electrode. | Assisted Replacement | Electrode | Membrane | Lac | Cal
ABL800 BASIC | Ref.-, K-, Glu- und Lac-Elektroden neu membraniert 2-Pkt.-Kalibrierungen | Ref.-, K-, Glu- und Lac-Elektroden neu membraniert 2-Pkt.-Kalibrierungen | SWO-121181 | No | Ref.-, K-, Glu- and Lac-electrodes re-membraneized 2-pcs. calibrations | Ref.-, K-, Glu- and Lac-electrodes re-membraneized 2-pcs. calibrations | Assisted Replacement | Electrode | Membrane | Lac
ABL800 BASIC | Gerätearchiv überprüft, Gerät neu gestartet, Lufteinlass gespült, Einlassnadel überprüft - ist gerade, Einlassdichtung aus Kundenbestand getauscht, Gerät im Service überprüft, Kal. Ok, System- und Funktionskontrolle in Ordnung. | Gerätearchiv überprüft, Gerät neu gestartet, Lufteinlass gespült, Einlassnadel überprüft - ist gerade, Einlassdichtung aus Kundenbestand getauscht, Gerät im Service überprüft, Kal. Ok, System- und Funktionskontrolle in Ordnung. | SWO-121056 | No | Device archive checked, device restarted, air inlet flushed, inlet needle checked - is straight, inlet seal is exchanged from customer stock, device checked in service, Kal. Ok, system and function control all right. | Device archive checked, device restarted, air inlet flushed, inlet needle checked - is straight, inlet seal is exchanged from customer stock, device checked in service, Kal. Ok, system and function control all right. | Defect | HW | Inlet module | Bend needle
ABL825 FLEX | Checked pumps, liquid sensors, valves and tubing. All checked okay. | Replaced solutions pump tube for diagnostic reasons. Updated software to version 6.16. | D12131392E2DFAC7 | No | Checked pumps, liquid sensors, valves and tubing. All checked okay. | Replaced solutions pump tube for diagnostic reasons. Updated software to version 6.16. | Defect | HW | Basic Unit | Pump tubes
ABL800 BASIC | Create New Patient Layout as CK Birla and Select the same as default. Remove the Reportable Ranges. Run the sample for Observation. Now issue has been Resolved. | B | SWO-120832 | No | Create New Patient Layout as CK Birla and Select the same as default. Remove the Reportable Ranges. Run the sample for Observation. Now issue has been Resolved. |  | Setup | Patient Report | Patient Report | Patient Report
ABL800 BASIC | cleaned the nipples and waste line tube, tube got disconnected from the nipple, refix the tube now working fine. | cleaned the nipples and waste line tube, tube got disconnected from the nipple, refix the tube now working fine. | SWO-122176 | No | cleaned the nipples and waste line tube, tube got disconnected from the nipple, refix the tube now working fine. | cleaned the nipples and waste line tube, tube got disconnected from the nipple, refix the tube now working fine. | Cleaned | HW | Fluidic pathway | Waste tube
ABL80 CO-OX | Analyzer would not boot up after corrupt file error. | Replaced CF card and reinstalled customer setup to the best of my ability. Customer will verify normal, reportable and critical ranges are correct. | D12287B6A41E7377 | No | Analyzer would not boot up after corrupt file error. | Replaced CF card and reinstalled customer setup to the best of my ability. Customer will verify normal, reportable and critical ranges are correct.
ABL90 FLEX | Enter service and carry out pO2 test, recalibrate, no change. fit new SP and re calibrate and re run failed QC. | Enter service and carry out pO2 test, recalibrate, no change. fit new SP and re calibrate and re run failed QC. | SWO-121213 | Yes | Enter service and carry out pO2 test, recalibrate, no change. fit new SP and re calibrate and re run failed QC. | Enter service and carry out pO2 test, recalibrate, no change. fit new SP and re calibrate and re run failed QC. | Re classified RMED | Same day upSWO | See SWO-121212
ABL90 FLEX | Fit new Barrery, no change, battery icon still indicating zero percent and not changing after 10 mins with new battery. Return visit required with PSU and or a CPU. | Fit new Barrery, no change, battery icon still indicating zero percent and not changing after 10 mins with new battery. Return visit required with PSU and or a CPU. | SWO-121212 | No | Fit new Barrery, no change, battery icon still indicating zero percent and not changing after 10 mins with new battery. Return visit required with PSU and or a CPU. | Fit new Barrery, no change, battery icon still indicating zero percent and not changing after 10 mins with new battery. Return visit required with PSU and or a CPU. | Defect | HW | Battery module | Battery module
ABL90 FLEX PLUS | Einstellungen der Uhr in der Konfig geändert, SC wurde heute ersetzt, PH und PCo2 Kal Fehler vorhanden,Empf. Zu niedrig bei PH 83,1% und Pco2 bei 23,3%, keine neue SC mehr vor Ort, Med Technik erhält morgen Früh einen Rückruf um den Status und dann evt weitere Maßnahmen zu besprechen. | Rückruf bei Hrn. Becker Empf immer noch zu niedrig, SC ersetzt, Kal und QC durchgeführt, SC reklamiert | SWO-121304 | No | Settings of the watch changed in the config, SC was replaced today, PH and PCo2 Kal error present,Mpf. Too low at PH 83.1% and Pco2 at 23.3%, no new SC on site, Med Technik receives a recall tomorrow morning to discuss status and then evt further measures. | Recall at Mr. Becker Empf still too low, SC replaced, Kal and QC carried out, SC complained | Setup | Time | Time | Time
ABL90 FLEX PLUS | Gerät überprüft, System gespült, Lösungspack und Sensorkassette ausgetauscht | Gerät überprüft, System gespült, Lösungspack und Sensorkassette ausgetauscht | SWO-121302 | No | Device checked, system flushed, solution pack and sensor cartridge replaced | Device checked, system flushed, solution pack and sensor cartridge replaced | Assisted Replacement | SP | SP | SP | & SC, pCO2 Glu Lac
ABL90 FLEX PLUS | Sensorkassette erneuert. Telefonisch nach Einlaufphase Werte abgefragt. Kalibrationen und Qualitätskontrollen in Ordnung. | Sensorkassette erneuert. Telefonisch nach Einlaufphase Werte abgefragt. Kalibrationen und Qualitätskontrollen in Ordnung. | SWO-121235 | No | Sensor cartridge is renewed. Questions values by telephone after the run-in phase. calibrations and quality controls in order. | Sensor cartridge is renewed. Questions values by telephone after the run-in phase. calibrations and quality controls in order. | Assisted Replacement | SC | SC | pCO2 | & Cl
ABL90 FLEX PLUS |  | B | D1229631A41AD4B8 | No |  |  | Assisted Replacement | SC | SC | Reference | 1361
ABL825 FLEX |  | B | D122555DA433717C | No |  |  | Assisted Replacement | Electrode | Membrane | Lac | Cal
ABL835 FLEX | Radiometer to escalate the fault. | Radiometer to escalate the fault. | SWO-121273 | No | Radiometer to escalate the fault. | Radiometer to escalate the fault. | Assisted Replacement | Electrode | Membrane | Reference | pCO2 symptome
ABL800 BASIC | BG側測定室、ｐH部剥離の為、交換 | B | SWO-121279 | No | Bg side measuring chamber, for pH part peeling, replaced |  | Defect | HW | pH/BG Module | pH | Sealing Ring
ABL825 FLEX | Error code 529 (in EI Met) Checked liquid sensors and pump rotors (service mode). Checked membrane valves. | Replaced pump rotor in EI Met module.. Replaced inlet gasket. Replaced dummy electrode in EI Met module. Replaced elecrtode module tubings - both modules. | D122ADD0A407BDD6 | No | Error code 529 (in EI Met) Checked liquid sensors and pump rotors (service mode). Checked membrane valves. | Replaced pump rotor in EI Met module.. Replaced inlet gasket. Replaced dummy electrode in EI Met module. Replaced elecrtode module tubings - both modules. | Unknown | HW | Fluidic pathway | Fluidic pathway
TCM400-4 | Verified said issue. | Replaced CPU board. | D122B668A449FF64 | No | Verified said issue. | Replaced CPU board. | Defect | HW | PCB, Main CPU | PCB, Main CPU
ABL80 FLEX BASIC | 更换了接口板902-972后修复(byzhenjie) | 更换了接口板902-972后修复(byzhenjie) | SWO-121074 | No | Repair after replacing interface board 902-972 (byzhenjie) | Repair after replacing interface board 902-972 (byzhenjie)
ABL80 FLEX BASIC | Analyzer meegenomen voor verder onderzoek met goedkeuring van Mr. Degryse. | Analyzer meegenomen voor verder onderzoek met goedkeuring van Mr. Degryse. | SWO-121561 | No | Analyzer for further research with approval of Mr. Degryse. | Analyzer for further research with approval of Mr. Degryse.
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-121303 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL80 FLEX CO-OX | Gerät überprüft, Leihgerät 304909 aufgestellt und konfiguriert | Gerät überprüft, Leihgerät 304909 aufgestellt und konfiguriert | SWO-121819 | No | Device checked, rental device 304909 set up and configured | Device checked, rental device 304909 set up and configured
ABL80 CO-OX | Solution pack leaking through waste port. Solution pack leak is known issue and will submit credits for customer. Checked waste pump in diagnostics and flushed waste tubing with hypochlorite solution. | Installed new solution pack using customer stock. Verified instrument operation. | D123AD861A0DB055 | No | Solution pack leaking through waste port. Solution pack leak is known issue and will submit credits for customer. Checked waste pump in diagnostics and flushed waste tubing with hypochlorite solution. | Installed new solution pack using customer stock. Verified instrument operation.
ABL80 CO-OX |  | B | D125EEC80637626B | No
ABL90 FLEX | Changed hemolyzer unit. Ran all service tests. Ran tHb calibration. All passed. | Changed hemolyzer unit. Ran all service tests. Ran tHb calibration. All passed. | SWO-121499 | No | Changed hemolyzer unit. Ran all service tests. Ran tHb calibration. All passed. | Changed hemolyzer unit. Ran all service tests. Ran tHb calibration. All passed. | Defect | HW | Oxi module | Hemolyzer
ABL90 FLEX | Gerät geprüft, Kundengespräch, automatisches Drucken deaktiviert | visuelle Kontrolle der Pins vom Sensor Interface Board - kein Fehler feststellbar Gerät bleibt unter Beobachtung falls in den nächsten Tagen wieder SC mit Ref. Elektroden Fehler ausfällt, sollte das Sensor Interface Board getauscht werden | SWO-121430 | No | Device checked, customer call, automatic printing disabled | visual control of the pins from the Sensor Interface Board - no error detectable device remains under observation if in the next few days again SC with Ref. Electrodes error fails, the sensor interface board should be replaced | No fault found | No fault found | No fault found | No fault found | SC 1361
ABL90 FLEX PLUS |  | B | D123CB591A5CD485 | No |  |  | No fault found | Power taken | SC & SP replaced | Power taken
ABL90 FLEX PLUS | Analysator angebunden und Kundengespräch durchgeführt. | Analysator angebunden und Kundengespräch durchgeführt. | SWO-121406 | No | analyzer and customer call carried out. | analyzer and customer call carried out. | Setup | HIS/LIS | HIS/LIS | Aqure
ABL90 FLEX PLUS | Rückruf bei Hrn. Becker Empf immer noch zu niedrig, SC ersetzt, Kal und QC durchgeführt, SC reklamiert | Rückruf bei Hrn. Becker Empf immer noch zu niedrig, SC ersetzt, Kal und QC durchgeführt, SC reklamiert | SWO-121512 | Yes | Recall at Mr. Becker Empf still too low, SC replaced, Kal and QC carried out, SC complained | Recall at Mr. Becker Empf still too low, SC replaced, Kal and QC carried out, SC complained | Re classified RMED | Phone call | SC | pCO2
AQT90 FLEX |  | B | SWO-121354 | No |  |  | Defect | HW | Aspiration Module | Inlet Wheel Sensor
AQT90 FLEX |  | B | SWO-121369 | No |  |  | User error | SW | calibration Kit | Inactivation
ABL825 FLEX | FlexQ Modul teilweise demontiert, gereinigt und kalibriert etc. - Funktionsüberprüfung etc. | FlexQ Modul teilweise demontiert, gereinigt und kalibriert etc. - Funktionsüberprüfung etc. | SWO-121382 | No | FlexQ module partially dismantled, cleaned and calibrated etc. - function check etc. | FlexQ module partially dismantled, cleaned and calibrated etc. - function check etc. | Adjustment | HW | FlexQ Module | FlexQ Module
ABL825 FLEX | Messkammer geprüft und gereinigt, Na. neu membraniert, Elektrodenkappe bestellt | B | SWO-121639 | No | Measuring chamber tested and cleaned, Na. newly membraneed, electrode cap ordered |  | Defect | HW | El/Met Module | Electrode Cap
ABL80 FLEX BASIC | 更换了接口板和新的测试卡后解决(by zhenjie) | 更换了接口板和新的测试卡后解决(by zhenjie) | SWO-121298 | No | Post-fix with replacement of interface board and new test card (by zhenjie) | Post-fix with replacement of interface board and new test card (by zhenjie)
ABL80 FLEX BASIC | check the sp80 smart chip error. replace the sp80. There is leakage in sensor cassette cause the sensor connector board problem. clean the sensor connector board. but fail temp test. need to replace the sensor connector board.sc 80 faulty .replace the sc80 . | B | SWO-121892 | No | check the sp80 smart chip error. replace the sp80. There is leakage in sensor cassette cause the sensor connector board problem. clean the sensor connector board. but fail temp test. need to replace the sensor connector board.sc 80 faulty .replace the sc80 .
ABL80 CO-OX | Customer already replaced SC x2. Replaced SP - analyzer failed to see SP. Replaced with another pack. Failing Hb and pO2. Cleaned and checked all valves and tubing. Completed Oxi test. Failing lamp. | Replaced Hemolyzer unit and broken cable. Oxi test checks OK. Completed total Hb Cal. | D124CDA6905653E0 | No | Customer already replaced SC x2. Replaced SP - analyzer failed to see SP. Replaced with another pack. Failing Hb and pO2. Cleaned and checked all valves and tubing. Completed Oxi test. Failing lamp. | Replaced Hemolyzer unit and broken cable. Oxygen test checks OK. Completed Total Hb Cal.
ABL90 FLEX | visuelle Kontrolle der Pins vom Sensor Interface Board - kein Fehler feststellbar Gerät bleibt unter Beobachtung falls in den nächsten Tagen wieder SC mit Ref. Elektroden Fehler ausfällt, sollte das Sensor Interface Board getauscht werden | visuelle Kontrolle der Pins vom Sensor Interface Board - kein Fehler feststellbar Gerät bleibt unter Beobachtung falls in den nächsten Tagen wieder SC mit Ref. Elektroden Fehler ausfällt, sollte das Sensor Interface Board getauscht werden | SWO-121729 | No | visual control of the pins from the Sensor Interface Board - no error detectable device remains under observation if in the next few days again SC with Ref. Electrodes error fails, the sensor interface board should be replaced | visual control of the pins from the Sensor Interface Board - no error detectable device remains under observation if in the next few days again SC with Ref. Electrodes error fails, the sensor interface board should be replaced | No fault found | No fault found | No fault found | No fault found
ABL90 FLEX | Sensorcassette geliefert und installiert, Kalibrationen, QC Kontrolle | Sensorcassette geliefert und installiert, Kalibrationen, QC Kontrolle | SWO-121859 | Yes | Sensor cassette delivered and installed, calibrations, QC control | Sensor cassette delivered and installed, calibrations, QC control | Re classified RMED | Logistical upSWO | SC delivered | Customer SC expired
ABL90 FLEX | Fehlersuche: Chipdaten werden nicht gelesen, Kontakte gereinigt, im Service überprüft - i.O., LP neu installiert, Kal u. QC i.O. | B | SWO-121877 | No | Troubleshooting: Chip data is not read, contacts are cleaned, checked in the service - i.O., LP reinstalled, Kal & QC i.O. |  | Adjustment | HW | SP PCB | Contacts
ABL90 FLEX PLUS | Keyboard was in Hungarian setting, set to English. Re-entered user barcodes. | Keyboard was in Hungarian setting, set to English. Re-entered user barcodes. | SWO-122745 | No | Keyboard was in Hungarian setting, set to English. Re-entered user barcodes. | Keyboard was in Hungarian setting, set to English. Re-entered user barcodes. | Setup | SW | SW | Barcode &Keyboard
ABL90 FLEX PLUS | Refitted solution pack now ok. | Refitted solution pack now ok. | SWO-121793 | No | Refitted solution pack now ok. | Refitted solution pack now ok. | Assisted Replacement | SP | SP | FT
ABL90 FLEX PLUS | SE SUSTITUYE POR ENTRADA NUEVA. | SE SUSTITUYE POR ENTRADA NUEVA. | D124CC02900F4AB8 | No | REPLACED BY NEW INPUT. | REPLACED BY NEW INPUT. | Defect | HW | Auto Inlet module | Auto Inlet module | requested returned
AQT90 FLEX |  | B | SWO-121694 | No |  |  | Defect | HW | Storage Module | Barcode reader, Storage wheel
ABL825 FLEX | Einlass und Abfallsystem demontiert und gereinigt, Verstopfung beseitigt, Montage, Kal und QK ok | Einlass und Abfallsystem demontiert und gereinigt, Verstopfung beseitigt, Montage, Kal und QK ok | SWO-121834 | No | Inlet and waste system dismantled and cleaned, clogging removed, assembly, cal and QK ok | Inlet and waste system dismantled and cleaned, clogging removed, assembly, cal and QK ok | Cleaned | HW | Inlet module | Inlet
ABL825 FLEX |  | B | D124D9A490087CF1 | No |  |  | Unknown | Electrode | Membrane | Reference
ABL825 FLEX |  | B | D124DA1190057CF4 | No |  |  | Unknown | Electrode | Membrane | Reference
ABL825 FLEX |  | B | D124DA6F90427CF7 | No |  |  | Unknown | Electrode | Membrane | Reference
ABL825 FLEX |  | B | D12517B09012259A | No |  |  | Defect | HW | Inlet module | Friction detector unit
ABL815 FLEX | came in last week and reinstalled windows 7 and software 6.16, I am going to open a techline case for this problem | problem of Ca not being selected on patient layout; use dynamic parameters profile during measurement is selected ; noticed 95ul was selected on the syringe mode | D12508049036FEE1 | No | came in last week and reinstalled windows 7 and software 6.16; I am going to open a techline case for this problem | problem of Ca not being selected on patient layout; use dynamic parameters profile during measurement is selected ; noticed 95ul was selected on the syringe mode | User error | HW | Mode Selection | Mode Selection
ABL80 FLEX BASIC | Done cleaning of SC and removal of air from solution pack. Than kept SC in fridge for 30 minutes. Then reinstalled SC. Unit is working fine. | Done cleaning of SC and removal of air from solution pack. Than kept SC in fridge for 30 minutes. Then reinstalled SC. Unit is working fine. | SWO-122124 | No | Done cleaning of SC and removal of air from solution pack. Than kept SC in fridge for 30 minutes. Then reinstalled SC. Unit is working fine. | Done cleaning of SC and removal of air from solution pack. Than kept SC in fridge for 30 minutes. Then reinstalled SC. Unit is working fine.
ABL80 FLEX BASIC |  | B | SWO-122002 | No
ABL90 FLEX | Arrived on site and attended analyser that was showing the battery 0% problem. Analyser now showing 100% battery. Disconnected the mains cable to test how quickly the battery dissipates. Analyser shut down immediately. Removed rear cover and psu. Fitted new psu and refitted rear cover. Powered up the analyser with mains plugged in. Analyser started a 10 minute warm start as expected. Battery showing 0% after 5 minutes of power on. Suggests that the analyser psu may not be the issue. Entered service programs and checked and checked battery status. Analyser showing as connected to mains supply when mains cable connected but battery showing as initialised but at 0% even after 15 minutes attached to the mains power. Battery light flashing green. Had a new battery tried and a psu both of which did not resolve the issue. New pc module to be ordered and fitted on next visit. Ran calibration and qc after power down to verify analyser functionality. Analyser running well. | Arrived on site and attended analyser that was showing the battery 0% problem. Analyser now showing 100% battery. Disconnected the mains cable to test how quickly the battery dissipates. Analyser shut down immediately. Removed rear cover and psu. Fitted new psu and refitted rear cover. Powered up the analyser with mains plugged in. Analyser started a 10 minute warm start as expected. Battery showing 0% after 5 minutes of power on. Suggests that the analyser psu may not be the issue. Entered service programs and checked and checked battery status. Analyser showing as connected to mains supply when mains cable connected but battery showing as initialised but at 0% even after 15 minutes attached to the mains power. Battery light flashing green. Had a new battery tried and a psu both of which did not resolve the issue. New pc module to be ordered and fitted on next visit. Ran calibration and qc after power down to verify analyser functionality. Analyser running well. | SWO-122144 | No | Arrived on site and attended analyser that was showing the battery 0% problem. Analyser now showing 100% battery. Disconnected the mains cable to test how quickly the battery dissipates. Analyser shut down immediately. Removed rear cover and psu. Fitted new psu and refitted rear cover. Powered up the analyser with mains plugged in. Analyser started a 10 minute warm start as expected. Battery showing 0% after 5 minutes of power on. Suggests that the analyser psu may not be the issue. Entered service programs and checked and checked battery status. Analyser showing as connected to mains supply when mains cable connected but battery showing as initialised but at 0% even after 15 minutes attached to the mains power. Battery light flashing green. Had a new battery tried and a psu both of which did not resolve the issue. New pc module to be ordered and fitted on next visit. Ran calibration and qc after power down to verify analyser functionality. Analyser running well. | Arrived on site and attended analyser that was showing the battery 0% problem. Analyser now showing 100% battery. Disconnected the mains cable to test how quickly the battery dissipates. Analyser shut down immediately. Removed rear cover and psu. Fitted new psu and refitted rear cover. Powered up the analyser with mains plugged in. Analyser started a 10 minute warm start as expected. Battery showing 0% after 5 minutes of power on. Suggests that the analyser psu may not be the issue. Entered service programs and checked and checked battery status. Analyser showing as connected to mains supply when mains cable connected but battery showing as initialised but at 0% even after 15 minutes attached to the mains power. Battery light flashing green. Had a new battery tried and a psu both of which did not resolve the issue. New pc module to be ordered and fitted on next visit. Ran calibration and qc after power down to verify analyser functionality. Analyser running well. | Defect | HW | Battery module | Battery module
ABL90 FLEX | Patientenlayout wie von Hr. Sengl gewünscht abgeändert. | Patientenlayout wie von Hr. Sengl gewünscht abgeändert. | SWO-122087 | Yes | Patient layout as desired by Mr. Sengl. | Patient layout as desired by Mr. Sengl. | Re classified RMED | Courtesy | FSE Called out to other analyzer | See UpSWO SWO-12288
ABL90 FLEX | Arrived on site and attended analyser. Found in solution pack replacement screen. Removed the pack and checked I'd chip for signs of damage or fault but nothing obvious. Inserted the old solution pack that was removed and this then continued the installation and completed once the inlet was closed. Ran the replacement procedure once more and inserted the new pack but it got stuck at the same point and would not recognise the pack once installed. Removed this pack and tried another new pack. This worked correctly. Analyser also required new sensor cassette. Installed new cassette and left analyser to condition. Analyser back up and running again | Arrived on site and attended analyser. Found in solution pack replacement screen. Removed the pack and checked I'd chip for signs of damage or fault but nothing obvious. Inserted the old solution pack that was removed and this then continued the installation and completed once the inlet was closed. Ran the replacement procedure once more and inserted the new pack but it got stuck at the same point and would not recognise the pack once installed. Removed this pack and tried another new pack. This worked correctly. Analyser also required new sensor cassette. Installed new cassette and left analyser to condition. Analyser back up and running again | SWO-122033 | No | Arrived on site and attended analyser. Found in solution pack replacement screen. Removed the pack and checked I'd chip for signs of damage or fault but nothing obvious. Inserted the old solution pack that was removed and this then continued the installation and completed once the inlet was closed. Ran the replacement procedure once more and inserted the new pack but it got stuck at the same point and would not recognise the pack once installed. Removed this pack and tried another new pack. This worked correctly. Analyser also required new sensor cassette. Installed new cassette and left analyser to condition. Analyser back up and running again | Arrived on site and attended analyser. Found in solution pack replacement screen. Removed the pack and checked I'd chip for signs of damage or fault but nothing obvious. Inserted the old solution pack that was removed and this then continued the installation and completed once the inlet was closed. Ran the replacement procedure once more and inserted the new pack but it got stuck at the same point and would not recognise the pack once installed. Removed this pack and tried another new pack. This worked correctly. Analyser also required new sensor cassette. Installed new cassette and left analyser to condition. Analyser back up and running again | Assisted Replacement | SP | SP | Chip
ABL90 FLEX | Hemolyzer Unit erneuert. Oxi Complete Test ok. LP erneuert (QC Fehler Level C). tHb Kalibration (100,5%), Kalibrationen und QC's vorgenommen. Fehler trat nach 3 QC Level B wieder auf. Ersatzteil nicht im Bestand. Nach Rücksprache mit Hr. Sengl Parameter deaktiviert. Wie von Herrn Sengl gewünscht Patientenlayout abgeändert. | Hemolyzer Unit erneuert. Oxi Complete Test ok. LP erneuert (QC Fehler Level C). tHb Kalibration (100,5%), Kalibrationen und QC's vorgenommen. Fehler trat nach 3 QC Level B wieder auf. Ersatzteil nicht im Bestand. Nach Rücksprache mit Hr. Sengl Parameter deaktiviert. Wie von Herrn Sengl gewünscht Patientenlayout abgeändert. | SWO-122088 | No | Hemolyzer Unit. Oxi Complete Test ok. LP renewed (QC Error Level C). tHb calibration (100.5%), calibrations and QC's. Error reappeared after 3 QC Level B. Spare part is not in stock. After consultation with Mr. Sengl parameter deactivated. As requested by Mr. Sengl, patient layout changed. | Hemolyzer Unit. Oxi Complete Test ok. LP renewed (QC Error Level C). tHb calibration (100.5%), calibrations and QC's. Error reappeared after 3 QC Level B. Spare part is not in stock. After consultation with Mr. Sengl parameter deactivated. As requested by Mr. Sengl, patient layout changed. | Defect | HW | Oxi module | Hemolyzer | Potential revisit
ABL90 FLEX | Explained to customer that lower reportable range is 0.4 and he said this may cause an issue, I advised him to take this up with Paul Cawood in the laboratory. Steve has adjusted the lower ref range to 0.4. I have advised Paul Cawood of the above | Explained to customer that lower reportable range is 0.4 and he said this may cause an issue, I advised him to take this up with Paul Cawood in the laboratory. Steve has adjusted the lower ref range to 0.4. I have advised Paul Cawood of the above | SWO-122097 | No | Explained to customer that lower reportable range is 0.4 and he said this may cause an issue, I advised him to take this up with Paul Cawood in the laboratory. Steve has adjusted the lower ref range to 0.4. I have advised Paul Cawood of the above | Explained to customer that lower reportable range is 0.4 and he said this may cause an issue, I advised him to take this up with Paul Cawood in the laboratory. Steve has adjusted the lower ref range to 0.4. I have advised Paul Cawood of the above | Setup | Reportable ranges | Reportable ranges | Reportable ranges
ABL90 FLEX PLUS | Keine Anbindungsdaten im CRM hinterlegt. Kontaktaufnahme mit Herrn Bischof (EDV Gelnhausen). Abfrage der Anbindungsdaten. Kein Ping möglich. Analysator hat IP 169.254.152.86 (Keeper Addi IP, frei vergeben durch abgebrochene Verbindung. Eingabe der Verbindungsdaten auf Windowsebene. Speicherung der Daten in Bios. Neustart. Verbindung hergestellt. Daten wurden automatisch nachgesendet. Analysator messbereit und Angebunden Testmessung, Übertragung ins KIS LIS erfolgreich | Keine Anbindungsdaten im CRM hinterlegt. Kontaktaufnahme mit Herrn Bischof (EDV Gelnhausen). Abfrage der Anbindungsdaten. Kein Ping möglich. Analysator hat IP 169.254.152.86 (Keeper Addi IP; frei vergeben durch abgebrochene Verbindung. Eingabe der Verbindungsdaten auf Windowsebene. Speicherung der Daten in Bios. Neustart. Verbindung hergestellt. Daten wurden automatisch nachgesendet. Analysator messbereit und Angebunden Testmessung, Übertragung ins KIS LIS erfolgreich | SWO-122094 | No | No connection data stored in CRM. Contacting Mr Bishop (EDV Gelnhausen). Querying the connection data. No ping possible. Analyzer has IP 169.254.152.86 (Keeper Addi IP; freely assigned by aborted connection. Enter connection data at the Windows level. Storage of data in bios. Restart. Connection established. Data was automatically sent back. Analyzer ready for measurement and tied test measurement, transmission to The KIS LIS successful | No connection data stored in CRM. Contacting Mr Bishop (EDV Gelnhausen). Querying the connection data. No ping possible. Analyzer has IP 169.254.152.86 (Keeper Addi IP; freely assigned by aborted connection. Enter connection data at the Windows level. Storage of data in bios. Restart. Connection established. Data was automatically sent back. Analyzer ready for measurement and tied test measurement, transmission to The KIS LIS successful | Setup | His/lis | His/lis | Aqure
ABL90 FLEX PLUS | Refitted solution pack and calibrated | Refitted solution pack and calibrated | SWO-122224 | No | Refitted solution pack and calibrated | Refitted solution pack and calibrated | Assisted Replacement | SP | SP | FT
ABL800 FLEX | Cleaned Po2 chamber & checked equipment is funcationl | Cleaned Po2 chamber & checked equipment is funcationl | SWO-123751 | No | Cleaned Po2 chamber & checked equipment is funcationl | Cleaned Po2 chamber & checked equipment is funcationl | Cleaned | Fluidic pathway | pH/BG Module | pO2
ABL800 BASIC |  | B | SWO-121926 | No |  |  | Defect | HW | Inlet module | Nipple kit
ABL835 FLEX |  | B | D12658790616C461 | No |  |  | Defect | HW | AutoCheck Module | Temperature
ABL825 FLEX | Error code Neon lamp number too high | Replaced Oxi module PCB. | D1265F59060F8D06 | No | Error code Neon lamp number too high | Replaced Oxygen Module PCB. | Defect | HW | Oxi module | Neon Lamp | High
ABL835 FLEX | Analysator geprüft, Windows SW neu installiert, SW V6.16 neu aufgespielt, Konfiguration neu eingegeben, Funktionskontrolle ok | Anwender erwartet Rückruf bzw. FSE-Einsatz | SWO-122077 | No | Analyzer checked, Windows SW reinstalled, SW V6.16 replayed, configuration re-entered, function control ok | Users expect recall or FSE use | Defect | SW | Reinstall | Reinstall
ABL825 FLEX | Mit Frau Fassel am Telefon die Lac Elektrode neu membraniert. Mehrer Kalibrierungen von Frau Fassel durchführen lassen. Danach QC wiederholt. Danach war nach Aussage von Frau Fessel der Analysator wieder Messbereit. Als ich dann vor Ort war habe ich nochmals mehrmals kalibriert und alle QCs wiederholt. pCO2 Membran erneuert da Probleme bei QC Level 3 Nach mehreren Kalibrierungen und mehrmaligen membranieren hat sich der pCO2 Fehler bei QC Level 3 nicht beheben lassen. Deswegen Parameter pCO2 gesperrt und Frau Fassel mitgeteilt das am Montag ein Techniker kommt und sich dem Problem annimmt. Die anderen Parameter waren bei den QCs i.O. QC Level 1 neu eingelesen weil das richtige Level nicht im ABL hinterlegt war. Level zwei war zwar eingelesen aber die falschen QCs im Rundel. Rundel richtig bestückt. Alle drei QC Level durchgeführt --> i.O. ABL wieder Messbereit. | Mit Frau Fassel am Telefon die Lac Elektrode neu membraniert. Mehrer Kalibrierungen von Frau Fassel durchführen lassen. Danach QC wiederholt. Danach war nach Aussage von Frau Fessel der Analysator wieder Messbereit. Als ich dann vor Ort war habe ich nochmals mehrmals kalibriert und alle QCs wiederholt. pCO2 Membran erneuert da Probleme bei QC Level 3 Nach mehreren Kalibrierungen und mehrmaligen membranieren hat sich der pCO2 Fehler bei QC Level 3 nicht beheben lassen. Deswegen Parameter pCO2 gesperrt und Frau Fassel mitgeteilt das am Montag ein Techniker kommt und sich dem Problem annimmt. Die anderen Parameter waren bei den QCs i.O. QC Level 1 neu eingelesen weil das richtige Level nicht im ABL hinterlegt war. Level zwei war zwar eingelesen aber die falschen QCs im Rundel. Rundel richtig bestückt. Alle drei QC Level durchgeführt --> i.O. ABL wieder Messbereit. | SWO-122157 | No | With Ms. Fassel on the phone, the Lac electrode is re-membrane. Several calibrations by Ms. Fassel. Then QC repeated. After that, according to Ms. Fessel, the analyzer was ready to measure again. When I was there I calibrated several times and repeated all QCs. pCO2 membrane renews as problems with QC Level 3 After several calibrations and multiple membranes, the pCO2 error with QC Level 3 could not be fixed. That's why pCO2 was blocked and Ms. Fassel told that on Monday a technician came and addressed the problem. The other parameters were reread in the QCs in the qCs in the main QC Level 1 because the correct level was not stored in the ABL. Level two was read in but the wrong QCs in the round. Rundel properly equipped. All three QC levels performed --> i.O. ABL ready to measure again. | With Ms. Fassel on the phone, the Lac electrode is re-membrane. Several calibrations by Ms. Fassel. Then QC repeated. After that, according to Ms. Fessel, the analyzer was ready to measure again. When I was there I calibrated several times and repeated all QCs. pCO2 membrane renews as problems with QC Level 3 After several calibrations and multiple membranes, the pCO2 error with QC Level 3 could not be fixed. That's why pCO2 was blocked and Ms. Fassel told that on Monday a technician came and addressed the problem. The other parameters were reread in the QCs in the qCs in the main QC Level 1 because the correct level was not stored in the ABL. Level two was read in but the wrong QCs in the round. Rundel properly equipped. All three QC levels performed --> i.O. ABL ready to measure again. | User error | HW | Auto Check Module | QC | Wrong QC's
ABL835 FLEX | Mit Frau Fassel am Telefon nochmals die pco2 Elektrode membraniert und kalibriert. Keine Besserung. Bei Eintreffen habe ich die pCO2 Elektrode erneuert und membraniert. Mehrmals kalibriert. Danach war Empf von pCO2 wieder i.O. QCs für alle Level durchgeführt. --> i.O. ABL wieder Messbereit. | Mit Frau Fassel am Telefon nochmals die pco2 Elektrode membraniert und kalibriert. Keine Besserung. Bei Eintreffen habe ich die pCO2 Elektrode erneuert und membraniert. Mehrmals kalibriert. Danach war Empf von pCO2 wieder i.O. QCs für alle Level durchgeführt. --> i.O. ABL wieder Messbereit. | SWO-122158 | No | With Ms. Fassel on the phone again the pco2 electrode membranes and calibrates. No improvement. When I arrived, I renewed and membraneed the pCO2 electrode. Calibrated several times. After that, Empf of pCO2 was again performed i.O. QCs for all levels. --> i.O. ABL ready to measure again. | With Ms. Fassel on the phone again the pco2 electrode membranes and calibrates. No improvement. When I arrived, I renewed and membraneed the pCO2 electrode. Calibrated several times. After that, Empf of pCO2 was again performed i.O. QCs for all levels. --> i.O. ABL ready to measure again. | Assisted Replacement | Electrode | Electrode | pCO2
ABL800 FLEX | Performed complete oximetry test & 2 point calibration done | Performed complete oximetry test & 2 point calibration done | SWO-123800 | No | Performed complete oximetry test & 2 point calibration done | Performed complete oximetry test & 2 point calibration done | Adjustment | HW | Oxi module | Oxi module
ABL90 FLEX | Reinigung Einlassdichtung, Fluid Transport durchgeführt, Kal abgewartet, Kal u. QC i.O. Probenmessung - i.O. | Reinigung Einlassdichtung, Fluid Transport durchgeführt, Kal abgewartet, Kal u. QC i.O. Probenmessung - i.O. | SWO-122626 | No | Cleaning inlet seal, fluid transport carried out, Kal waited, Kal and QC i.O. sample measurement - i.O. | Cleaning inlet seal, fluid transport carried out, Kal waited, Kal and QC i.O. sample measurement - i.O. | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway | & inlet gasket cleaned
ABL90 FLEX | SC und LP erneuert. Nach Erneuerung Kal Fehler pCO2 (44,7%). SC nochmals erneuert. Kalibration und QC's vorgenommen. | SC und LP erneuert. Nach Erneuerung Kal Fehler pCO2 (44,7%). SC nochmals erneuert. Kalibration und QC's vorgenommen. | SWO-122530 | No | SC and LP. After renewal Kal error pCO2 (44.7%). SC renewed. calibration and QC's. | SC and LP. After renewal Kal error pCO2 (44.7%). SC renewed. calibration and QC's. | Assisted Replacement | SC | SC | pCO2
ABL90 | 1343 unable to pump solutions error, 1063 pressure test vacuum error and 1373 solution pack replacement test failed messages are present. The vacuum value of 0.97 is above the acceptable value of 0.25-0.75. | Replaced the inlet gasket and re-ran the pressure test in service. The value is now 0.48. Completed all fluid flow tests successfully. Analyzer is within acceptable operating specifications. | D129E512684D5151 | No | 1343 unable to pump solutions error, 1063 pressure test vacuum error and 1373 solution pack replacement test failed messages are present. The vacuum value of 0.97 is above the acceptable value of 0.25-0.75. | Replaced the inlet gasket and re-ran the pressure test in service. The value is now 0.48. Completed all fluid flow tests successfully. Analyzer is within acceptable operating specifications. | Defect | HW | Inlet module | Inlet gasket
ABL90 FLEX | Spectrometer Unit erneuert. Oxi Complete Test ok. tHB Kalibration (100,0%), Kalibration und QC's vorgenommen. Kunden LP für Lager überlassen. | Spectrometer Unit erneuert. Oxi Complete Test ok. tHB Kalibration (100,0%), Kalibration und QC's vorgenommen. Kunden LP für Lager überlassen. | SWO-122531 | No | Spectrometer Unit. Oxi Complete Test ok. tHB calibration (100.0%), calibration and QC's. customers lp for stock. | Spectrometer Unit. Oxi Complete Test ok. tHB calibration (100.0%), calibration and QC's. customers lp for stock. | Defect | HW | Oxi module | Spectrometer
ABL90 | On arrival analyzer working ok. | Perform back flush procedure. Perform all service checks | D12A1E8F684BBC08 | No | On arrival analyzer working ok. | Perform back flush procedure. Perform all service checks | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway | Na
ABL90 FLEX | Einlaßdichtung gereinigt Intervalle zurückgesetzt tHb-Kal durchgeführt Einweisung Frau Bositz: Fehlersuche, Behebung und CCF | Einlaßdichtung gereinigt Intervalle zurückgesetzt tHb-Kal durchgeführt Einweisung Frau Bositz: Fehlersuche, Behebung und CCF | SWO-122648 | No | Inlet seal cleaned intervals reset tHb-Kal carried instruction Mrs. Bositz: troubleshooting, remediation and CCF | Inlet seal cleaned intervals reset tHb-Kal carried instruction Mrs. Bositz: troubleshooting, remediation and CCF | Setup | maintenance schedule | Inlet module | Inlet gasket & connecter gasket | tHb cal
ABL90 FLEX PLUS | Arrived on site and attended analyser to investigate any further instances of the error 1349. Logged onto analyser and checked activity logs for any further instances since my last visit. Found that there had been another on the 01 08 16 . Analyser showing same symptoms of displaying low battery around 15 minutes before shutdown. Brought new psu for replacement. Powered down the analyser no removed psu from analyser. Fitted new psu and put analyser back together. Powered up the analyser and waited for it to complete its startup procedure. After this I tested the battery performance in the service programs. Unplugged analyser from the mains and the it diverted to better as expected. But when plugged back into mains, the analyser did not return back to running from mains and start charging the battery again. Battery dissipated at a good rate though. Powered down the analyser and replaced psu with another. Powered back on and performed the same tests to verify psu functionality. Entered service program, checked battery status which showed as connected. Disconnected from the mains and analyser diverted to battery power. Plugged analyser back into mains and the analyser did not divert back to mains. Refitted the original psu and powered the analyser back up. Tested this psu in the same way. When unplugged the analyser runs from battery and when plugged back into mains supply, the analyser diverts back to mains power and charges the battery. Performed calibrations and qcs to verify analyser functionality. Analyser running well on this ups. | Arrived on site and attended analyser to investigate any further instances of the error 1349. Logged onto analyser and checked activity logs for any further instances since my last visit. Found that there had been another on the 01 08 16 . Analyser showing same symptoms of displaying low battery around 15 minutes before shutdown. Brought new psu for replacement. Powered down the analyser no removed psu from analyser. Fitted new psu and put analyser back together. Powered up the analyser and waited for it to complete its startup procedure. After this I tested the battery performance in the service programs. Unplugged analyser from the mains and the it diverted to better as expected. But when plugged back into mains, the analyser did not return back to running from mains and start charging the battery again. Battery dissipated at a good rate though. Powered down the analyser and replaced psu with another. Powered back on and performed the same tests to verify psu functionality. Entered service program, checked battery status which showed as connected. Disconnected from the mains and analyser diverted to battery power. Plugged analyser back into mains and the analyser did not divert back to mains. Refitted the original psu and powered the analyser back up. Tested this psu in the same way. When unplugged the analyser runs from battery and when plugged back into mains supply, the analyser diverts back to mains power and charges the battery. Performed calibrations and qcs to verify analyser functionality. Analyser running well on this ups. | SWO-122483 | No | Arrived on site and attended analyser to investigate any further instances of the error 1349. Logged onto analyser and checked activity logs for any further instances since my last visit. Found that there had been another on the 01 08 16 . Analyser showing same symptoms of displaying low battery around 15 minutes before shutdown. Brought new psu for replacement. Powered down the analyser no removed psu from analyser. Fitted new psu and put analyser back together. Powered up the analyser and waited for it to complete its startup procedure. After this I tested the battery performance in the service programs. Unplugged analyser from the mains and the it diverted to better as expected. But when plugged back into mains, the analyser did not return back to running from mains and start charging the battery again. Battery dissipated at a good rate though. Powered down the analyser and replaced psu with another. Powered back on and performed the same tests to verify psu functionality. Entered service program, checked battery status which showed as connected. Disconnected from the mains and analyser diverted to battery power. Plugged analyser back into mains and the analyser did not divert back to mains. Refitted the original psu and powered the analyser back up. Tested this psu in the same way. When unplugged the analyser runs from battery and when plugged back into mains supply, the analyser diverts back to mains power and charges the battery. Performed calibrations and qcs to verify analyser functionality. Analyser running well on this ups. | Arrived on site and attended analyser to investigate any further instances of the error 1349. Logged onto analyser and checked activity logs for any further instances since my last visit. Found that there had been another on the 01 08 16 . Analyser showing same symptoms of displaying low battery around 15 minutes before shutdown. Brought new psu for replacement. Powered down the analyser no removed psu from analyser. Fitted new psu and put analyser back together. Powered up the analyser and waited for it to complete its startup procedure. After this I tested the battery performance in the service programs. Unplugged analyser from the mains and the it diverted to better as expected. But when plugged back into mains, the analyser did not return back to running from mains and start charging the battery again. Battery dissipated at a good rate though. Powered down the analyser and replaced psu with another. Powered back on and performed the same tests to verify psu functionality. Entered service program, checked battery status which showed as connected. Disconnected from the mains and analyser diverted to battery power. Plugged analyser back into mains and the analyser did not divert back to mains. Refitted the original psu and powered the analyser back up. Tested this psu in the same way. When unplugged the analyser runs from battery and when plugged back into mains supply, the analyser diverts back to mains power and charges the battery. Performed calibrations and qcs to verify analyser functionality. Analyser running well on this ups. | No fault found | Original PSU worked | Original PSU worked | Original PSU worked | Battery issue
ABL90 | Verificação do histórico, substituição da SC por uma nova sem sucesso, verificação da SC noutro equipamento e verificou-se o mau funcioanmento da mesma, por fim substituição de outra SC e establização, passagem do QC. Tudo Ok. | B | D12A3030683F195D | No | Checking the history, replacing the SC with a new one without success, checking the SC in another equipment and verified the bad function of the same, finally replacing another SC and establishing, passing the QC. Okay, all right. |  | Assisted Replacement | SC | SC | pCO2
AQT90 FLEX | 做系统检查，提示光学强度超范围，更换了光学单元后修复 by jesson | 做系统检查，提示光学强度超范围，更换了光学单元后修复 by jesson | SWO-122599 | No | Do a system check, prompt the optical strength over range, replace the optical unit after repair by jesson | Do a system check, prompt the optical strength over range, replace the optical unit after repair by jesson | Defect | HW | Process Module | Optical Unit | 131
ABL827 FLEX | Swapped the oxi module with the one from 754R0098N001 as this is not in use at the minute. Ran a complete oxi test which passed. Service tHb calibration run and passed. Crea membranes replaced because it failed level 2 & 4 QC | Swapped the oxi module with the one from 754R0098N001 as this is not in use at the minute. Ran a complete oxi test which passed. Service tHb calibration run and passed. Crea membranes replaced because it failed level 2 & 4 QC | SWO-122573 | No | Swapped the oxi module with the one from 754R0098N001 as this is not in use at the minute. Ran a complete oxi test which passed. Service tHb calibration run and passed. Crea membranes replaced because it failed level 2 & 4 QC | Swapped the oxi module with the one from 754R0098N001 as this is not in use at the minute. Ran a complete oxi test which passed. Service tHb calibration run and passed. Crea membranes replaced because it failed level 2 & 4 QC | Defect | HW | Oxi module | Neon Lamp
ABL825 FLEX | Customer tried protein removal and tHB cal | Replaced hemolyzer and lamp unit. Adjusted lamp voltage. Ran tHb cal in service and customer modes | D129E3246849911B | No | Customer tried protein removal and tHB cal | Replaced hemolyzer and lamp unit. Adjusted lamp voltage. Ran tHb cal in service and customer modes | Defect | HW | Oxi module | Hemolyzer | Sensor Clot
ABL835 FLEX |  | B | D129E8146849BF53 | No |  |  | Defect | HW | Auto Check Module | Fan
ABL825 FLEX | pCO2 Elektrode erneuert. Parameter entsperrt. Kalibrationen und QC's vorgenommen. | pCO2 Elektrode erneuert. Parameter entsperrt. Kalibrationen und QC's vorgenommen. | SWO-122393 | No | pCO2 electrode. Parameter unlocked. calibrations and QC's. | pCO2 electrode. Parameter unlocked. calibrations and QC's. | Assisted Replacement | Electrode | Electrode | pCO2 | 6.13
ABL80 FLEX CO-OX | Sensorkassette demontiert und neu zusammengesetzt, Kal und QC durchgeführt | Sensorkassette demontiert und neu zusammengesetzt, Kal und QC durchgeführt | SWO-122816 | No | Sensor cassette dismantled and reassembled, Kal and QC performed | Sensor cassette dismantled and reassembled, Kal and QC performed
ABL90 FLEX | Neue Patres aufgespielt. Druckbild in Ordnung. Kalibration und QC's vorgenommen. | Neue Patres aufgespielt. Druckbild in Ordnung. Kalibration und QC's vorgenommen. | SWO-122798 | No | New Fathers. Print image in order. calibration and QC's. | New Fathers. Print image in order. calibration and QC's. | Setup | HW | Printer | External printer
ABL90 FLEX | bei eintreffen pCO2 Kal Empf. bei 78%, Kunde verwendet SC 946-009 Lot <500, SC ersetzt, DataLogs gesichert & an TS geschickt, Startphase: pCO2 QC im Bereich aber div. andere Parameter noch mit BackLog Fehler, Kal weiderholt, nur noch pH Kal fehlerhaft, Med.-Technik kalibriert das Gerät weiter & widerholt QC Messung ->meldet sich wenn noch einmal Hilfe nötig | bei eintreffen pCO2 Kal Empf. bei 78%, Kunde verwendet SC 946-009 Lot <500, SC ersetzt, DataLogs gesichert & an TS geschickt, Startphase: pCO2 QC im Bereich aber div. andere Parameter noch mit BackLog Fehler, Kal weiderholt, nur noch pH Kal fehlerhaft, Med.-Technik kalibriert das Gerät weiter & widerholt QC Messung ->meldet sich wenn noch einmal Hilfe nötig | SWO-122749 | No | when pCO2 Kal Empf arrives. at 78%, customer uses SC 946-009 lot •500, SC replaced, DataLogs secured & sent to TS, start phase: pCO2 QC in the range but various other parameters still with BackLog error, Kal weiderholt, only pH Kal faulty, Med.-Technology further calibrates the device & retrieves QC measurement ->reports if help is needed again | when pCO2 Kal Empf arrives. at 78%, customer uses SC 946-009 lot •500, SC replaced, DataLogs secured & sent to TS, start phase: pCO2 QC in the range but various other parameters still with BackLog error, Kal weiderholt, only pH Kal faulty, Med.-Technology further calibrates the device & retrieves QC measurement ->reports if help is needed again | Assisted Replacement | SC | SC | pCO2 | Sensitivity
ABL90 FLEX | Fehlersuche durchgeführt. Schlauch zwischen SC und Hemolyzer war defekt. | Fehlersuche durchgeführt. Schlauch zwischen SC und Hemolyzer war defekt. | SWO-122769 | No | Troubleshooting. Hose between SC and Hemolyzer was defective. | Troubleshooting. Hose between SC and Hemolyzer was defective. | Defect | HW | Oxi module | Pinch valve tube
ABL90 FLEX | Einlassonde und Dichtung dem Kunden ausgeliefert | Einlassonde und Dichtung dem Kunden ausgeliefert | SWO-122708 | No | Inlet probe and seal delivered to the customer | Inlet probe and seal delivered to the customer | Defect | HW | Inlet module | Inlet gasket
ABL90 | CHECKED ALL SETTINGS AND VISUALLY INSPECTED CABLING | DISCOVERED THAT JIM SNESKA SET UP THE ANALYZER FOR SERIAL INTERFACING THROUGH A TERM SERVER AND AT SOME POINT AFTER THAT WAS DONE THE HOSPITA NETWORKING DEPT SET UP A SECOND CONNECTION THROUGH TCP IP AND CHANGED THE SETTING FROM SERIAL TO TCP IP IN THE ANALYZER SETTINGS.REMOVED THE TCP IP CONNECTION AND RESTORED THE STTINGS TO SERIAL. | D12B9ACCDE1EB13F | No | CHECKED ALL SETTINGS AND VISUALLY INSPECTED CABLING | DISCOVERED THAT JIM SNESKA SET UP THE ANALYZER FOR SERIAL INTERFACING THROUGH A TERM SERVER AND AT SOME POINT AFTER THAT WAS DONE THE HOSPITA NETWORKING DEPT SET UP A SECOND CONNECTION THROUGH TCP IP AND CHANGED THE SETTING FROM SERIAL TO TCP IP IN THE ANALYZER SETTINGS. REMOVED THE TCP IP CONNECTION AND RESTORED THE STTINGS TO SERIAL. | Setup | His/lis | His/lis | His/lis
ABL90 | worked with Mark to get analyzers connected to DI through a lantronixs box | IT was trying to connect through TCP IP, it should have been connected through Serial , see Mark's swo | D12C44EF545FADE0 | Yes | worked with Mark to get analyzers connected to DI through a lantronixs box | IT was trying to connect through TCP IP, it should have been connected through Serial , see Mark's swo | Re classified RMED | Same day upSWO | Same day upSWO | see UpSWO D12B9ACCDE1EB13F
ABL90 FLEX PLUS | Lieferung LP | Lieferung LP | SWO-122856 | Yes | Delivery LP | Delivery LP | Re classified RMED | Logistical upSWO | SP | SP
ABL800 FLEX | ｵｷｼﾓｼﾞｭｰﾙ（代替品）へ入替え作業。 | B | SWO-123501 | No | Replacement work for oxy modules (alternatives). |  | Defect | HW | Oxi module | Spectrometer
ABL810 FLEX | blank screen reboot gets no activity from CPU | replaced Power supply and CPU unit is plugged into a UPS that was not effected | D12B4854DE475573 | No | blank screen reboot gets no activity from CPU | replaced Power supply and CPU unit is plugged into a UPS that was not effected | Defect | HW | DMS_PC | PSU
ABL837 FLEX | Unit leaking from pH Chamber and inlet gasket | Replaced pH sealing ring. Inlet gasket waste pump rotor | D12B4554DE5ECE4C | No | Unit leaking from pH Chamber and inlet gasket | Replaced pH sealing ring. Inlet gasket waste pump rotor | Defect | HW | pH/BG Module | pH | Sealing Ring
ABL837 FLEX |  | B | D12B41A1DE1D558E | No |  |  | Assisted Replacement | Electrode | Membrane | Reference
ABL817 FLEX | Verfied once on site. | During troubleshooting, found that the rinse bottle tubing had came disconnected inside the bottle. Replaced the rinse bottle from the customers stock and the errors are no longer present. Also, replaced the A and B Crea membrains. Performed extensive calibrations. Readjusted the dummy electrodes in the chamber no leaks found after the adjustment. Analyzer is now back in service. | D12B963DDE249F0E | No | Verfied once on site. | During troubleshooting, found that the rinse bottle tubing had came disconnected inside the bottle. Replaced the rinse bottle from the customers stock and the errors are no longer present. Also, replaced the A and B Crea membrains. Performed extensive calibrations. Readjusted the dummy electrodes in the chamber no leaks found after the adjustment. Analyzer is now back in service. | Assisted Replacement | Consumables | Rinse solution | Tube
ABL825 FLEX | Elektrolyt Met. Modul ersetzt, SW installiert, Gerät kalibriert und mit QCs überprüft, OK. Defektes Modul: 902-824R2286N007 | Elektrolyt Met. Modul ersetzt, SW installiert, Gerät kalibriert und mit QCs überprüft, OK. Defektes Modul: 902-824R2286N007 | SWO-122952 | No | Electrolyte Met. Module replaced, SW installed, device calibrated and checked with QCs, OK. Defective module: 902-824R2286N007 | Electrolyte Met. Module replaced, SW installed, device calibrated and checked with QCs, OK. Defective module: 902-824R2286N007 | Defect | HW | Elmet | Glu
ABL825 FLEX | Lac. Elektrodenkappe ersetzt, Funktionskontrolle ok | Lac. Elektrodenkappe ersetzt, Funktionskontrolle ok | SWO-123236 | No | Lac. Electrode cap replaced, function control ok | Lac. Electrode cap replaced, function control ok | Defect | HW | Electrode Connector | Lac
ABL815 FLEX | Modul geöffnet und gereinigt. Position kalibriert und Scanner justiert. | Modul geöffnet und gereinigt. Position kalibriert und Scanner justiert. | SWO-122942 | No | Module opened and cleaned. position calibrated and scanner adjusted. | Module opened and cleaned. position calibrated and scanner adjusted. | Adjustment | HW | FlexQ Module | Barcode scanner
ABL825 FLEX | Modul geöffnet und gereinigt. | Modul geöffnet und gereinigt. | SWO-122935 | No | Module opened and cleaned. | Module opened and cleaned. | Adjustment | HW | FlexQ Module | FlexQ Module
ABL815 FLEX | FlexQ Scanner eingestellt. Glu kalibriert. | FlexQ Scanner eingestellt. Glu kalibriert. | SWO-122928 | No | FlexQ Scanner. Glu calibrated. | FlexQ Scanner. Glu calibrated. | Adjustment | HW | FlexQ Module | Barcode scanner
AQT90 FLEX |  | B | D12B41FFDE3AD81A | No |  |  | Defect | Consumables | SP | SP
AQT90 FLEX | Prozesswheel und Lösungspack blockiert. Lösungspack gerüttelt. Safe all positions durchgeführt. Systemreinigung und QC's durchgeführt. | Prozesswheel und Lösungspack blockiert. Lösungspack gerüttelt. Safe all positions durchgeführt. Systemreinigung und QC's durchgeführt. | SWO-122696 | No | Process wheel and solution pack blocked. solution pack. Safe all positions. system cleaning and QC's. | Process wheel and solution pack blocked. solution pack. Safe all positions. system cleaning and QC's. | Defect | Consumables | SP | SP | 5784
AQURE | Made all the SCBU operators Active for the ABL837 & ABL90 and pushed them to the analyser. All the operators were pushed successfully. | B | SWO-122834 | Yes | Made all the SCBU operators Active for the ABL837 & ABL90 and pushed them to the analyser. All the operators were pushed successfully.
ABL80 FLEX BASIC | replace the new sensor connector board. installed new sc 80 & sp80.Calibration pass. | replace the new sensor connector board. installed new sc 80 & sp80.Calibration pass. | SWO-123774 | No | replace the new sensor connector board. installed new sc 80 & sp80. Calibration pass. | replace the new sensor connector board. installed new sc 80 & sp80. Calibration pass.
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-125631 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL80 FLEX BASIC | プレートの調整 | B | SWO-122882 | No | Adjusting the plate
ABL80 FLEX BASIC | So Smartchip was faulty,it need replacement. Repalce the Smartchip with new one. After replace the Smartchip unit is working proeprly. | So Smartchip was faulty,it need replacement. Repalce the Smartchip with new one. After replace the Smartchip unit is working proeprly. | SWO-123788 | No | So Smartchip was faulty,it need replacement. Repalce the Smartchip with new one. After replace the Smartchip unit is working proeprly. | So Smartchip was faulty,it need replacement. Repalce the Smartchip with new one. After replace the Smartchip unit is working proeprly.
ABL90 | Unit failing vacuum test | Replaced inlet gasket | D12C85BD543BECAF | No | Unit failing vacuum test | Replaced inlet gasket | Defect | HW | Inlet module | Inlet gasket
ABL90 FLEX | SC ersetzt, Lot 822 Startphase pCO2 empf. bei ca.24% ABL mehrmals kalibriert die Empf. bleibt bei 24%. >>SC ersetzt Lot. 821 Startphase QC und Kalibration. | Service Einsatz und DATA-LOGS ziehen. | SWO-123404 | No | SC replaced, lot 822 start phase pCO2 empf. at approx.24% ABL several times the empf calibrates. remains at 24%. >>SC replaces lot. 821 Start phase QC and calibration. | Service insert and DATA-LOGS. | Assisted Replacement | SC | SC | pCO2
ABL90 | Upon arrival disocovered unit was idle for 3 months with all consumables expired. | Flushed out unit and replaced inlet connector gasket provided by customer. After the parts were replaced, a new sensor cassette and solution pack were installed. Instructed customer how to replaced consumables and how to do basic troubleshooting as well as adding editing users. | D12C8F8A543CAE19 | No | Upon arrival disocovered unit was idle for 3 months with all consumables expired. | Flushed out unit and replaced inlet connector gasket provided by customer. After the parts were replaced, a new sensor cassette and solution pack were installed. Instructed customer how to replaced consumables and how to do basic troubleshooting as well as adding editing users. | User error | Consumables exp. | Consumables exp. | Consumables exp.
ABL90 FLEX | SP aus Ersatzgerät ohne Fehlermeldungen eingebaut. QC ohne Besserung. Kal ok. Neue SC nach Überprüfung Fließwege..durchspülen eingesetzt. Kal bei Glu..Lac Empf. Fehler.QC bei PCO2 in Grenzen. | SP aus Ersatzgerät ohne Fehlermeldungen eingebaut. QC ohne Besserung. Kal ok. Neue SC nach Überprüfung Fließwege..durchspülen eingesetzt. Kal bei Glu..Lac Empf. Fehler.QC bei PCO2 in Grenzen. | SWO-123259 | No | SP from replacement device without error messages. QC without improvement. Kal ok. New SC after checking flow paths.. rinse. Kal at Glu.. Lac Empf. Error.QC with PCO2 within limits. | SP from replacement device without error messages. QC without improvement. Kal ok. New SC after checking flow paths.. rinse. Kal at Glu.. Lac Empf. Error.QC with PCO2 within limits. | Assisted Replacement | SC | SC | pCO2
ABL90 FLEX PLUS |  | B | D12C7FEB5425AAFF | No |  |  | Replaced | Probe | Probe damaged | Probe damaged
ABL90 FLEX PLUS | Anpassung Überprüfen der Konfiguration, Überprüfen der QC-Kontrollen, SC wechsel da pCO2 Kal. Fehler. Bez. Bilirubin kein Fehler feststellbar. | Mit Chefarzt Dr. Garbe Korrektur von 0,73 eingegeben. Werte werden nun weiter verglichen und Korrekturfaktor ggf. weiter angepasst. Korrektur wird nicht auf QC angewendet und entspr. Info ist im Bildschirm sichtbar. Labor und POCT Managerin informiert. | SWO-123629 | No | Adjustment Check the configuration, check the QC controls, SC switch since pCO2 Cal. Error. Bilirubin no error can be detected. | With Chief Physician Dr. Garbe correction entered of 0.73. Values are now further compared and correction factor is further adjusted if necessary. Correction is not applied to QC and is equivalent to Info is visible on the screen. laboratory and POCT manager. | Setup | SW | SW | Correction factor | Double UpSWO 1/8 & 12/8
ABL90 | TRAS VARIOS CAMBIOS DE REACTIVOS Y NO SIENDO POSIBLE SU SOLUCION, SE PROCEDE A SUSTITUIR LA PLACA SENSOR INTERFACE, SE VERIFICAN CANALES Y SE EFECTUAN CALIBRACIONES, CONTROLES DE CALIDAD ASI COMO MUESTRAS. VERIFICADO. | SUSTITUCION SENSOR INTERFACE | D12C5FAE54401163 | No | AFTER SEVERAL CHANGES OF REACTIVES AND NOT POSSIBLE YOUR SOLUTION, THE INTERFACE SENSOR PLACE IS SUBSTITUTE, CHANNELS ARE CHECKED AND CALIBRATIONS, QUALITY CONTROLS AS SAMPLES ARE EFFECTED. Verified. | SENSOR INTERFACE REPLACEMENT | Defect | HW | SC interface PCB | SC interface PCB
ABL90 FLEX | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-122945 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL90 | Se sustituyo SCi y se comprobó correcto funcionamiento mediante QC. | B | D1409C2BB451B65D | No | SCi was replaced and properly functioned by QC. |  | Assisted Replacement | SC | SC | pCO2
AQT90 FLEX |  | B | SWO-122926 | No |  |  | Defect | HW | Process Module | PCB, Cup dryer | 168
ABL800 BASIC | Replace the rinse valve. Run protein removal multiple time. calibration pass.qc pass. | B | SWO-123773 | No | Replace the rinse valve. Run protein removal multiple time. calibration pass.qc pass. |  | Defect | HW | Fluidic manifold | Valve
ABL800 BASIC | Installed permanent keys | Installed permanent keys | SWO-123174 | Yes | Installed permanent keys | Installed permanent keys | Re classified RMED | Setup | Demo Key
ABL825 FLEX | Defektes Oxi Modul ausgetauscht (alt 902-833R1974N004). Initialisierung. Überprüfung Lichtwerte. Glu musste membraniert werden. Kalibrationen mehrfach durchgeführt. QC in allen Leveln wiederholt | Defektes Oxi Modul ausgetauscht (alt 902-833R1974N004). Initialisierung. Überprüfung Lichtwerte. Glu musste membraniert werden. Kalibrationen mehrfach durchgeführt. QC in allen Leveln wiederholt | SWO-123135 | No | Replaced defective Oxi module (old 902-833R1974N004). Initialization. Checking light values. Glu had to be membraneed. calibrations performed several times. QC repeated at all levels | Replaced defective Oxi module (old 902-833R1974N004). Initialization. Checking light values. Glu had to be membraneed. calibrations performed several times. QC repeated at all levels | Defect | HW | Oxi module | Neon Lamp
ABL815 FLEX | Touch Screen ausgetauscht. Test ok. | Touch Screen ausgetauscht. Test ok. | SWO-123539 | No | Touch screen. Test ok. | Touch screen. Test ok. | Defect | HW | DMS_PC | touch Screen
ABL835 FLEX | The OXI was failing due to the Neon lamp | I replaced the PCB and the lamp with no luck. I talked with mike and tech support and I will order a Power supply, PCB and lamp to put in tomorrow. 8 I replaced the Spectrometer Unit and it came up. | D12C6D0354311A3C | No | The OXI was failing due to the Neon lamp | I replaced the PCB and the lamp with no luck. I talked with mike and tech support and I will order a Power supply, PCB and lamp to put in tomorrow. 8 I replaced the Spectrometer Unit and it came up. | Defect | HW | Oxi module | Neon Lamp | High
ABL825 FLEX | FlexQ Modul aus GS Bestand erneuert. Modul SW Upgrade durchgeführt. Leere Gas1 Flasche ersetzt. | FlexQ Modul aus GS Bestand erneuert. Modul SW Upgrade durchgeführt. Leere Gas1 Flasche ersetzt. | SWO-123030 | No | FlexQ module from GS stock renewed. module SW Upgrade. Empty Gas1 bottle replaced. | FlexQ module from GS stock renewed. module SW Upgrade. Empty Gas1 bottle replaced. | Defect | HW | FlexQ Module | FlexQ Module
AQT90 FLEX | Im Service Menü Fehler Prozess Wheel bei Safe all Position, LP läßt sich im Service Menü nicht entriegeln Fehler bei Solvent Interferenz, Puncher in Punch down Position blockiert, durch starkes rütteln am LP fallen Cups zur Seite danach alle Funktionen ok, Puncher wiederholt betätigt Positionen werden ohne Probleme angefahren LP Lot XB-05 durch LP Lot XB - 14 aus Kundenbestand ersetzt, Systemüberprüfung Systemreinigung und QC durchgeführt, Service dump gezogen | Im Service Menü Fehler Prozess Wheel bei Safe all Position, LP läßt sich im Service Menü nicht entriegeln Fehler bei Solvent Interferenz, Puncher in Punch down Position blockiert, durch starkes rütteln am LP fallen Cups zur Seite danach alle Funktionen ok, Puncher wiederholt betätigt Positionen werden ohne Probleme angefahren LP Lot XB-05 durch LP Lot XB - 14 aus Kundenbestand ersetzt, Systemüberprüfung Systemreinigung und QC durchgeführt, Service dump gezogen | SWO-123152 | No | In the service menu error process wheel at Safe all position, LP can not unlock in the service menu error at solvent interference, puncher blocked in punch down position, by strong shaking at the LP drop cups to the side after all functions ok, puncher repeated actuated positions are driven without problems LP Lot XB-05 by LP Lot XB - 14 replaced from customer inventory, system verification system cleaning and QC performed, service dump pulled | In the service menu error process wheel at Safe all position, LP can not unlock in the service menu error at solvent interference, puncher blocked in punch down position, by strong shaking at the LP drop cups to the side after all functions ok, puncher repeated actuated positions are driven without problems LP Lot XB-05 by LP Lot XB - 14 replaced from customer inventory, system verification system cleaning and QC performed, service dump pulled | Defect | Consumables | SP | SP
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-123588 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL80 FLEX BASIC | 更换了电源模块后修复 by jesson | 更换了电源模块后修复 by jesson | SWO-123295 | No | Fix by jesson after replacing power supply | Fix by jesson after replacing power supply
ABL80 CO-OX | customer reported failing its last 2 EQA tests | retreived data logs and customer data to start a tech line case | D12DB116CA03A6C2 | No | customer reported failing its last 2 EQA tests | retreived data logs and customer data to start a tech line case
ABL80 FLEX BASIC | So put the SC in refrigerator for 2 Hrs. After 2 hrs again try to install new SC. Install the new SC. After install the SC unit is working properly. | So put the SC in refrigerator for 2 Hrs. After 2 hrs again try to install new SC. Install the new SC. After install the SC unit is working properly. | SWO-123786 | No | So put the SC in refrigerator for 2 Hrs. After 2 hrs again try to install new SC. Install the new SC. After install the SC unit is working properly. | So put the SC in refrigerator for 2 Hrs. After 2 hrs again try to install new SC. Install the new SC. After install the SC unit is working properly.
ABL90 FLEX PLUS |  | B | D12DA6F4CA4BF53C | No |  |  | Assisted Replacement | SC | SC | Unknown
ABL90 FLEX PLUS | Replaced solution pack. | Replaced solution pack. | SWO-123985 | No | Replaced solution pack. | Replaced solution pack. | Assisted Replacement | SP | SP | Glu/Lac
ABL90 FLEX PLUS | Fehlende Fixe QC Zielwerte in Roma.ini eingegeben. Lösungspack neu installiert wegen Übernahme der Fixen Zielwerte. Bei Probenmessung festgestellt, dass falsche Gerinselfänger benutzt werden. ( Grund der Verschmutzung ) Bestellformular von Synergy korrigiert, Info an Jona weitergeleitet. Kunde veranlasst Umtausch über Zentrale. ( auch Kreißsaal und Notaufnahme muß korrigiert werden ) Einlass gereinigt und Sensorkassette wegen geringer Anzahl von restlichen Messungen erneuert. Danach, Glu und Lac Fehler im QC Lev. A zusätzlich noch 5 Kal.s durchgeführt , Lev. A weiterhin extrem über 0. ( 22 - 45 ) Kunde wartet bis morgen ab, ob sich SC noch einläuft , und tauscht dann eigenständig die SC. | Fehlende Fixe QC Zielwerte in Roma.ini eingegeben. Lösungspack neu installiert wegen Übernahme der Fixen Zielwerte. Bei Probenmessung festgestellt, dass falsche Gerinselfänger benutzt werden. ( Grund der Verschmutzung ) Bestellformular von Synergy korrigiert, Info an Jona weitergeleitet. Kunde veranlasst Umtausch über Zentrale. ( auch Kreißsaal und Notaufnahme muß korrigiert werden ) Einlass gereinigt und Sensorkassette wegen geringer Anzahl von restlichen Messungen erneuert. Danach, Glu und Lac Fehler im QC Lev. A zusätzlich noch 5 Kal.s durchgeführt , Lev. A weiterhin extrem über 0. ( 22 - 45 ) Kunde wartet bis morgen ab, ob sich SC noch einläuft , und tauscht dann eigenständig die SC. | SWO-123731 | No | Missing Fixe QC target values entered in Roma.ini. Solution pack reinstalled due to adoption of fixed target values. When measuring samples, it is found that incorrect gerinsel catchers are used. ( Cause of pollution ) Corrected order form from Synergy, info forwarded to Jonah. Customer arranges exchange via headquarters. ( also tear room and emergency room must be corrected ) Inlet cleaned and sensor cartridge renewed due to small number of remaining measurements. After that, Glu and Lac error in QC Lev. A additional 5 Kal.s performed, Lev. A still extremely over 0. ( 22 - 45 ) Customer waits until tomorrow to see if SC is still running in, and then exchanges the SC on its own. | Missing Fixe QC target values entered in Roma.ini. Solution pack reinstalled due to adoption of fixed target values. When measuring samples, it is found that incorrect gerinsel catchers are used. ( Cause of pollution ) Corrected order form from Synergy, info forwarded to Jonah. Customer arranges exchange via headquarters. ( also tear room and emergency room must be corrected ) Inlet cleaned and sensor cartridge renewed due to small number of remaining measurements. After that, Glu and Lac error in QC Lev. A additional 5 Kal.s performed, Lev. A still extremely over 0. ( 22 - 45 ) Customer waits until tomorrow to see if SC is still running in, and then exchanges the SC on its own. | Setup | Roma.ini | Roma.ini | Fixed assigned values
ABL800 BASIC | clean the ph chamber & electrode. replace the po2 & ref membrane. run protein removal multiple time. clean the gasket. cleaning multiple time. machine under observation. | clean the ph chamber & electrode. replace the po2 & ref membrane. run protein removal multiple time. clean the gasket. cleaning multiple time. machine under observation. | SWO-123772 | No | clean the ph chamber & electrode. replace the po2 & ref membrane. run protein removal multiple time. clean the gasket. cleaning multiple time. machine under observation. | clean the ph chamber & electrode. replace the po2 & ref membrane. run protein removal multiple time. clean the gasket. cleaning multiple time. machine under observation. | Cleaned | Fluidic pathway | Fluidic pathway | pH
ABL835 FLEX | Cl on level 3 trending high. Changing the membrane brings qc back down to package insert mean. | Replaced Cl electrode | D12D9DCFCA00C28D | No | Cl on level 3 trending high. Changing the membrane brings qc back down to package insert mean. | Replaced Cl electrode | Assisted Replacement | Electrode | Electrode | Cl
ABL800 BASIC | Found printer head broken we need to change Printer Head | Found printer head broken we need to change Printer Head | SWO-123267 | No | Found printer head broken we need to change Printer Head | Found printer head broken we need to change Printer Head | Defect | HW | Printer module | Printer Head
ABL825 FLEX | 代替えモジュールと入れ替え | フリクションデテクターユニット交換 | SWO-123252 | No | Replace with replacement module | Friction detector unit replacement | Defect | HW | Inlet module | Friction detector unit
ABL800 BASIC | I cross checked with another new gas cylinders and still its showing gas low pressure then I cleaned the gas transducer. after cleaning transducer machine showing gas cylinders 100 % then I given 2 point calibration after calibration again machine showing gas low pressure we need to replace gas module . | B | SWO-123350 | No | I cross checked with another new gas cylinders and still its showing gas low pressure then I cleaned the gas transducer. after cleaning transducer machine showing gas cylinders 100 % then I given 2 point calibration after calibration again machine showing gas low pressure we need to replace gas module . |  | Revisit required | HW | Gas Module | Gas Module
ABL80 FLEX BASIC | 更换测试卡修复 | 更换测试卡修复 | SWO-123647 | No | Replacement test card repair | Replacement test card repair
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-123519 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL80 CO-OX | Pc02 showing a positive bias on EQA testing , samples not comparing between the Abl 8xx units and the 80 unit | techline case has been submitted | D12EDFF7401B995B | No | Pc02 showing a positive bias on EQA testing , samples not comparing between the Abl 8xx units and the 80 unit | Techline case has been submitted
ABL80 FLEX CO-OX |  | B | SWO-123594 | No
ABL90 FLEX | Arrived on site and attended analyser with new cpu as a new battery and psu had already been tried but with no success. Powered down analyser and installed new cup but swapped the flash card from the old one into this uni to preserve data. Powered up the analyser and tested. Battery problem persisted. Fitted new battery into analyser and tested. Battery now working. Powered down analyser and fitted the old cup back into analyser with the old flash card transferred back into this unit. Powered up the analyser and tested. Battery still in working condition. Unplugged analyser from the mains and tested battery life. Battery disipatting at the correct rate. Performed calibration and qcs to verify analyser functionality. Analyser running well. | Arrived on site and attended analyser with new cpu as a new battery and psu had already been tried but with no success. Powered down analyser and installed new cup but swapped the flash card from the old one into this uni to preserve data. Powered up the analyser and tested. Battery problem persisted. Fitted new battery into analyser and tested. Battery now working. Powered down analyser and fitted the old cup back into analyser with the old flash card transferred back into this unit. Powered up the analyser and tested. Battery still in working condition. Unplugged analyser from the mains and tested battery life. Battery disipatting at the correct rate. Performed calibration and qcs to verify analyser functionality. Analyser running well. | SWO-123668 | No | Arrived on site and attended analyser with new cpu as a new battery and psu had already been tried but with no success. Powered down analyser and installed new cup but swapped the flash card from the old one into this uni to preserve data. Powered up the analyser and tested. Battery problem persisted. Fitted new battery into analyser and tested. Battery now working. Powered down analyser and fitted the old cup back into analyser with the old flash card transferred back into this unit. Powered up the analyser and tested. Battery still in working condition. Unplugged analyser from the mains and tested battery life. Battery disipatting at the correct rate. Performed calibration and qcs to verify analyser functionality. Analyser running well. | Arrived on site and attended analyser with new cpu as a new battery and psu had already been tried but with no success. Powered down analyser and installed new cup but swapped the flash card from the old one into this uni to preserve data. Powered up the analyser and tested. Battery problem persisted. Fitted new battery into analyser and tested. Battery now working. Powered down analyser and fitted the old cup back into analyser with the old flash card transferred back into this unit. Powered up the analyser and tested. Battery still in working condition. Unplugged analyser from the mains and tested battery life. Battery disipatting at the correct rate. Performed calibration and qcs to verify analyser functionality. Analyser running well. | Defect | HW | Battery module | Battery module
ABL90 FLEX | Remplacement du lecteur barcode + test : Ok | B | SWO-123768 | No | Barcode reader replacement - test: Ok |  | Defect | HW | DMS_PC | Barcode reader
ABL90 FLEX | Replaced with new | Replaced with new | SWO-123569 | No | Replaced with new | Replaced with new | Defect | HW | Inlet module | Inlet gasket
ABL90 FLEX | On arrival analyser had passed all calibrations and QC's and so was in normal operation | On arrival analyser had passed all calibrations and QC's and so was in normal operation | SWO-123660 | No | On arrival analyser had passed all calibrations and QC's and so was in normal operation | On arrival analyser had passed all calibrations and QC's and so was in normal operation | No fault found | No fault found | No fault found | No fault found
ABL90 FLEX | Bestellnummer korrigiert und Lösungspack wegen QC Fehler pCo2 Lev. C erneuert. Fehlerhafter LP wurde Reklamiert. | Bestellnummer korrigiert und Lösungspack wegen QC Fehler pCo2 Lev. C erneuert. Fehlerhafter LP wurde Reklamiert. | SWO-124705 | No | Order number corrected and solution pack due to QC error pCo2 Lev. C renewed. Faulty LP has been complained about. | Order number corrected and solution pack due to QC error pCo2 Lev. C renewed. Faulty LP has been complained about. | Assisted Replacement | SP | SP | pCO2
ABL90 FLEX | Lösungspack wegen QC Fehler erneuert und reklamiert. Bestellformular wegen falscher Bestellnummer Gerinselfänger korrigiert. Konfiguration gesichert. | Lösungspack wegen QC Fehler erneuert und reklamiert. Bestellformular wegen falscher Bestellnummer Gerinselfänger korrigiert. Konfiguration gesichert. | SWO-123728 | No | Solution pack due to QC error renewed and complained. Order form corrected due to incorrect order number Gerinselfänger. Configuration backed up. | Solution pack due to QC error renewed and complained. Order form corrected due to incorrect order number Gerinselfänger. Configuration backed up. | Assisted Replacement | SP | SP | pCO2
ABL90 FLEX PLUS |  | B | D12EDB23402E9C81 | Yes |  |  | Re classified RMED | SW update | SW update | SW update
ABL90 FLEX PLUS | Mit Chefarzt Dr. Garbe Korrektur von 0,73 eingegeben. Werte werden nun weiter verglichen und Korrekturfaktor ggf. weiter angepasst. Korrektur wird nicht auf QC angewendet und entspr. Info ist im Bildschirm sichtbar. Labor und POCT Managerin informiert. | Mit Chefarzt Dr. Garbe Korrektur von 0,73 eingegeben. Werte werden nun weiter verglichen und Korrekturfaktor ggf. weiter angepasst. Korrektur wird nicht auf QC angewendet und entspr. Info ist im Bildschirm sichtbar. Labor und POCT Managerin informiert. | SWO-123633 | Yes | With Chief Physician Dr. Garbe correction entered of 0.73. Values are now further compared and correction factor is further adjusted if necessary. Correction is not applied to QC and is equivalent to Info is visible on the screen. laboratory and POCT manager. | With Chief Physician Dr. Garbe correction entered of 0.73. Values are now further compared and correction factor is further adjusted if necessary. Correction is not applied to QC and is equivalent to Info is visible on the screen. laboratory and POCT manager. | Re classified RMED | Double UpSWO | 1/8 & 12/8
ABL90 FLEX PLUS | Lösungspack und Sensorkassette neu initialisiert. | Lösungspack und Sensorkassette neu initialisiert. | SWO-123730 | No | solution pack and sensor cartridge. | solution pack and sensor cartridge. | Assisted Replacement | SC | SC | Lac | & SP
ABL90 FLEX | 更换测试卡正常。 | 更换测试卡正常。 | SWO-123516 | No | The replacement test card is OK. | The replacement test card is OK. | Assisted Replacement | SC | SC | SC | Ref 1361
ABL800 BASIC | Found Calibration error, for all the pH & electrolytes, replaced the Ref membrane calibrated but problem was remaining same, got the standby reference electrode replaced and calibrated found instrument is functional. Need to replace with new Reference electrode. | Found Calibration error, for all the pH & electrolytes, replaced the Ref membrane calibrated but problem was remaining same, got the standby reference electrode replaced and calibrated found instrument is functional. Need to replace with new Reference electrode. | SWO-123754 | No | Found Calibration error, for all the pH & electrolytes, replaced the Ref membrane calibrated but problem was remaining same, got the standby reference electrode replaced and calibrated found instrument is functional. Need to replace with new Reference electrode. | Found Calibration error, for all the pH & electrolytes, replaced the Ref membrane calibrated but problem was remaining same, got the standby reference electrode replaced and calibrated found instrument is functional. Need to replace with new Reference electrode. | Assisted Replacement | Electrode | Electrode | Reference
ABL815 FLEX |  | B | D131E91E2C15BB62 | No |  |  | Assisted Replacement | Electrode | Electrode | Reference
ABL825 FLEX | Ca++ calcium | Remplacé electrode de referemce lot 5712, installé o'ring lactate et remplacé membrane due | D12EC7AC40540F6B | No | Ca++ calcium | Replaced electrode de referemce lot 5712, installed o'ring lactate and replaced membrane due | Assisted Replacement | Electrode | Electrode | Reference
ABL800 BASIC | TSTへ依頼 | B | SWO-123667 | No | Ask TST |  | Revisit required | Electrode | Electrode | Na+
ABL805 FLEX | Customer replaced waste pump tubing and inlet gasket with no resolve. Replaced inlet wash tubing, solution pump tubing and ph BG module pump tubing. | Flushed waste tubing to waste bottle. Cleaned waste pump rotor. Verified instrument operation. | D132876FA218B09D | No | Customer replaced waste pump tubing and inlet gasket with no resolve. Replaced inlet wash tubing, solution pump tubing and ph BG module pump tubing. | Flushed waste tubing to waste bottle. Cleaned waste pump rotor. Verified instrument operation. | Defect | HW | Inlet module | Tube
ABL800 BASIC | cleaned liquid flow path and checked Cal 2 solution. it was looking contaminated. So changed cal 2 solution and then given calibration. unit is working fine. | cleaned liquid flow path and checked Cal 2 solution. it was looking contaminated. So changed cal 2 solution and then given calibration. unit is working fine. | SWO-123740 | No | cleaned liquid flow path and checked Cal 2 solution. it was looking contaminated. So changed cal 2 solution and then given calibration. unit is working fine. | cleaned liquid flow path and checked Cal 2 solution. it was looking contaminated. So changed cal 2 solution and then given calibration. unit is working fine. | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
AQT90 FLEX | Barcode Scanner ausgebaut, Spiegelflächen gesäubert, zusammengebaut, erneut überprüft, Funktion, OK, Probemessungen durchgeführt, OK, | Barcode Scanner ausgebaut, Spiegelflächen gesäubert; zusammengebaut, erneut überprüft, Funktion, OK; Probemessungen durchgeführt, OK; | SWO-123630 | No | Barcode scanner removed, mirror surfaces cleaned; assembled, re-checked, function, OK; Sample measurements performed, OK; | Barcode scanner removed, mirror surfaces cleaned; assembled, re-checked, function, OK; Sample measurements performed, OK; | Cleaned | HW | Analyzer Control Module | Barcode reader, display
TCM400-6 | paged bruce martz | B | D12EE6F14015608C | No | paged bruce martz
AQURE | Names of ED and HDU analysers changed to match apex | Names of ED and HDU analysers changed to match apex | SWO-123570 | Yes | Names of ED and HDU analysers changed to match apex | Names of ED and HDU analysers changed to match apex
ABL80 FLEX BASIC | 检查机器，发现接口板有腐蚀，更换902-972后，重新校准零点。正常， BY jesson | 检查机器，发现接口板有腐蚀，更换902-972后，重新校准零点。正常， BY jesson | SWO-123533 | No | Check the machine and find corrosion on the interface plate, recalibrate the zero point after replacing the 902-972. Normal, BY jesson | Check the machine and find corrosion on the interface plate, recalibrate the zero point after replacing the 902-972. Normal, BY jesson
AQT90 FLEX | 仪器不启动，检测电源电压正常，更换了910-343后修复， by jesson | 仪器不启动，检测电源电压正常，更换了910-343后修复， by jesson | SWO-123529 | No | The instrument does not start, detects the normal supply voltage, replaces 910-343 after repair, by jesson | The instrument does not start, detects the normal supply voltage, replaces 910-343 after repair, by jesson | Defect | HW | Analyzer removed | PC
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-123822 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-123857 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL80 FLEX BASIC | replace the sc80. calibration pass. | replace the sc80. calibration pass. | SWO-125225 | No | replace the sc80. calibration pass. | replace the sc80. calibration pass.
ABL90 FLEX | Analysator teilweise demontiert - Touch Sreen geöffnet (kein Flüssigkeitseintritt erkennbar) - Erneuerung Touch Screen - Funktionsüberprüfung etc. | Analysator teilweise demontiert - Touch Sreen geöffnet (kein Flüssigkeitseintritt erkennbar) - Erneuerung Touch Screen - Funktionsüberprüfung etc. | SWO-123953 | No | Analyzer partially dismantled - Touch Sreen open (no fluid ingress recognizable) - Renewal Touch Screen - Function check etc. | Analyzer partially dismantled - Touch Sreen open (no fluid ingress recognizable) - Renewal Touch Screen - Function check etc. | Defect | HW | DMS_PC | Touch screen
ABL90 FLEX | Touchscreen neu kalibriert. Problem mit Medizintechnik besprochen --> ABL lässt sich wieder normal bedienen | Touchscreen neu kalibriert. Problem mit Medizintechnik besprochen --> ABL lässt sich wieder normal bedienen | SWO-124037 | No | touchscreen recalibrated. Problem with medical technology discussed --> ABL can be operated normally again | touchscreen recalibrated. Problem with medical technology discussed --> ABL can be operated normally again | Adjustment | HW | DMS_PC | Touch screen
ABL90 FLEX PLUS |  | B | D132CB2BA21B2495 | No |  |  | Assisted Replacement | SC | SC | Glu/Lac
AQT90 FLEX |  | B | SWO-123921 | No |  |  | Defect | HW | Cover and Chassis | Inlet Door | 4558
ABL837 FLEX | Replaced inlet probe. Cleaned clot out of chamber ok. Checked inlet calibration ok. | Replaced inlet probe. Cleaned clot out of chamber ok. Checked inlet calibration ok. | SWO-124063 | No | Replaced inlet probe. Cleaned clot out of chamber ok. Checked inlet calibration ok. | Replaced inlet probe. Cleaned clot out of chamber ok. Checked inlet calibration ok. | Defect | HW | Inlet module | Bend needle
ABL827 FLEX | Printer not working | Replaced printer assembly | D132BF10A231AA29 | No | Printer not working | Replaced printer assembly | Defect | HW | Printer module | Printer
ABL825 FLEX | Inspected cal one bottle which I found empty. | Replaced cal one bottle. | D1329FA4A21330AC | No | Inspected cal one bottle which I found empty. | Replaced cal one bottle. | Assisted Replacement | Solutions | Cal1 | Cal 1
ABL825 FLEX | Funktion überprüft, E788 UB01 hat Glasbruch. E788 erneuert, Funktion OK. | Funktion überprüft, E788 UB01 hat Glasbruch. E788 erneuert, Funktion OK. | SWO-124131 | No | Function checked, E788 UB01 has glass breakage. E788 renewed, function OK. | Function checked, E788 UB01 has glass breakage. E788 renewed, function OK. | Assisted Replacement | Electrode | Electrode | pCO2 | broken
ABL800 BASIC | CPUを新品と交換 | B | SWO-126530 | No | Replacing the CPU with a new one |  | Defect | HW | DMS_PC | CPU
ABL80 FLEX BASIC |  | B | SWO-124159 | No
ABL80 FLEX | Flushed SC with Distilled water, place for cooling for one hour and try to reinstalled, problem remains, replaced SC with new one. | Flushed SC with Distilled water, place for cooling for one hour and try to reinstalled, problem remains, replaced SC with new one. | SWO-125226 | No | Flushed SC with Distilled water, place for cooling for one hour and try to reinstalled, problem remains, replaced SC with new one. | Flushed SC with Distilled water, place for cooling for one hour and try to reinstalled, problem remains, replaced SC with new one.
ABL90 FLEX | Sensor Kassette erneuert | Sensor Kassette erneuert | SWO-124423 | No | Sensor cassette renewed | Sensor cassette renewed | Assisted Replacement | SC | SC | pH | 179
ABL90 FLEX | SC wurde noch nicht ersetzt, nur durch eine verbrauchte, SC 90 946-005 R0815 durch SC 90 Lot R0491 ersetzt, Empf bei 45%, abgewartet und mehrere Kal durchgeführt, LP ersetzt, Fehler bleibt, SC erneut ersetzt mit Lot R0502, Kal und QC durchgeführt | SC wurde noch nicht ersetzt, nur durch eine verbrauchte, SC 90 946-005 R0815 durch SC 90 Lot R0491 ersetzt, Empf bei 45%, abgewartet und mehrere Kal durchgeführt, LP ersetzt, Fehler bleibt, SC erneut ersetzt mit Lot R0502, Kal und QC durchgeführt | SWO-124293 | No | SC has not yet been replaced, only replaced by a used, SC 90 946-005 R0815 by SC 90 Lot R0491, Empf at 45%, waited and several kal performed, LP replaced, error remains, SC replaced with Lot R0502, Kal and QC performed | SC has not yet been replaced, only replaced by a used, SC 90 946-005 R0815 by SC 90 Lot R0491, Empf at 45%, waited and several kal performed, LP replaced, error remains, SC replaced with Lot R0502, Kal and QC performed | Assisted Replacement | SC | SC | pCO2
ABL90 |  | B | D1344F52182AE4AB | No |  |  | Assisted Replacement | SC | SC | Glu | 131
ABL90 FLEX | Schlauchsystem erneuert, Erneuerung Lösungspack, Funktionskontrolle | Schlauchsystem erneuert, Erneuerung Lösungspack, Funktionskontrolle | SWO-124184 | No | Hose system renewed, renewal solution pack, function control | Hose system renewed, renewal solution pack, function control | Defect | HW | PM Kit | PM Kit
ABL90 FLEX PLUS | Powered down and powered back up again, now ok. | Powered down and powered back up again, now ok. | SWO-124550 | No | Powered down and powered back up again, now ok. | Powered down and powered back up again, now ok. | Defect | SW | Sensor Bay | Temp. Reg.
AQT90 FLEX |  | B | SWO-124215 | No |  |  | Defect | HW | Aspiration Module | PCB, Cabinet | 4558
ABL825 FLEX |  | B | D133EF49182D98E4 | No |  |  | Assisted Replacement | Electrode | Electrode | Reference
ABL837 FLEX | Found air valves on all modules stuck in the open position. Leaking current reading 1236. | Cleaned air valves. Removed modules and allow to dry. Checked leak current, reading 67. | D1342418181E7149 | No | Found air valves on all modules stuck in the open position. Leaking current reading 1236. | Cleaned air valves. Removed modules and allow to dry. Checked leak current, reading 67. | Defect | HW | Met II | Valve, air
ABL800 BASIC | Installed permanent keys successfully, 2 point calibration performed, sample measurements performed | Installed permanent keys successfully, 2 point calibration performed, sample measurements performed | SWO-125090 | No | Installed permanent keys successfully, 2 point calibration performed, sample measurements performed | Installed permanent keys successfully, 2 point calibration performed, sample measurements performed | Setup | SW | Demo Key | Expired
ABL815 FLEX | This SWO is to document labor and parts used on 5 31 2016 -SWO B9Z35VXGK (which apparently did not saved to BAS, but there is a PDF copy | Replaced internal Bar code scanner | D1341C59180B75A6 | No | This SWO is to document labor and parts used on 5 31 2016 -SWO B9Z35VXGK (which apparently did not saved to BAS, but there is a PDF copy | Replaced internal Bar code scanner | Defect | HW | Basic Unit | bar code reader
ABL825 FLEX | Keine weiteren QC Verletzungen | B | SWO-124329 | No | No more QC violations |  | assisted Replacement | Electrode | Membrane | pCO2 | Rilibäk
ABL815 FLEX | keine aufälligen Prozesse oder System belastende Anwendung gefunden, Display kalibriert | keine aufälligen Prozesse oder System belastende Anwendung gefunden, Display kalibriert | SWO-124297 | No | no cumbersome processes or system loading application found, display calibrated | no cumbersome processes or system loading application found, display calibrated | Adjustment | HW | Touch screen | calibration
AQT90 FLEX | need to replace the needle. | need to replace the needle. | SWO-124605 | No | need to replace the needle. | need to replace the needle. | Defect | HW | Aspiration Module | Bend needle | 3812
AQT90 FLEX | Checked and found the sample needle has been bend tilt. Need to replace the same. | Clean the Dust fro storage wheel. Shift the AQT-90 Load the TnT and NT Pro BNP along with SP. Run the Sample for Observation. | SWO-125202 | Yes | Checked and found the sample needle has been bend tilt. Need to replace the same. | Clean the Dust fro storage wheel. Shift the AQT-90 Load the TnT and NT Pro BNP along with SP. Run the Sample for Observation. | Re classified RMED | Double UpSWO | See SWO-121212
AQT90 FLEX |  | B | D1343E7D180698E7 | No |  |  | Defect | HW | Aspiration Module | Needle Wash Unit | HCT
ABL80 FLEX BASIC | 更换测试卡正常。 | 更换测试卡正常。 | SWO-125293 | No | The replacement test card is OK. | The replacement test card is OK.
ABL80 FLEX BASIC |  | B | SWO-124427 | No
ABL80 FLEX BASIC | CPU交換。 | CPU交換。 | SWO-124547 | No | CPU switch. | CPU switch.
ABL80 FLEX | Checked waste pump motor, lubricate with zoric acid and reinstalled,Problem remains. Motor Not working, need to replace with new one. | B | SWO-125229 | No | Checked waste pump motor, lubricate with zoric acid and reinstalled,Problem remains. Motor Not working, need to replace with new one.
ABL90 FLEX | Leihgerät 090R0795N001 aufgestellt und konfiguriert, Leihgerät wieder mitgenommen Glukosefehler, Glukose abgeschaltet | Gerät überprüft, Wet Section ausgebaut, Sensorkontakte kontrolliert, Keine weitere Überprüfung möglich, da kein Sensorsimulator im Bestand | SWO-124633 | No | Loaner 090R0795N001 set up and configured, loaner again taken glucose error, glucose switched off | Device checked, Wet Section removed, sensor contacts checked, No further verification possible as no sensor simulator in stock | Unknown | Analyzer removed | Analyzer removed | Analyzer removed
ABL90 FLEX | Gerät überprüft, Sensorkassette 2 mal erneuert, Kontakte kontrolliert | diverse Meßkassetten ausprobiert | SWO-125479 | No | Device checked, sensor cartridge renewed 2 times, contacts checked | tried out various measuring cassettes | Assisted Replacement | SC | SC | Glu | 131
ABL90 FLEX | Sensorkassette und Lösungspack mit laufenden Gerät getauscht, keine Besserung | diverse Meßkassetten ausprobiert | SWO-125483 | Yes | Sensor cassette and solution pack exchanged with running device, no improvement | tried out various measuring cassettes | Re classified RMED | Double UpSWO
ABL90 FLEX PLUS |  | B | D13532E68E058DFC | Yes |  |  | Re classified RMED | Logistical upSWO | Delivery of SC | Delivery of SC
ABL90 FLEX PLUS |  | B | D13507C38E057E7A | No |  |  | Defect | SW | Database crash | Reload SW
ABL90 FLEX PLUS | SW versie 3.1 MR8 geïnstalleerd. | SW versie 3.1 MR8 geïnstalleerd. | SWO-124658 | No | SW version 3.1 MR8 installed. | SW version 3.1 MR8 installed. | Defect | SW | SW | His/lis
ABL90 FLEX PLUS | Bedienungsanleitung passend zum Gerät und Softwareversion geliefert. | Bedienungsanleitung passend zum Gerät und Softwareversion geliefert. | SWO-124580 | Yes | instruction manual to match the device and software version. | instruction manual to match the device and software version. | Re classified RMED | Logistical upSWO |  |  | email
ABL90 FLEX PLUS |  | B | D13530388E388DF9 | Yes |  |  | Re classified RMED | Logistical upSWO | Delivery of SC | Delivery of SC
ABL800 FLEX | Checked all the previous Data and Logs. Take the Feedback from Mr. Umrao and meet with Mr. Vipin. Export the Data backup fro ABL800. ABL800 is working fine. | Checked all the previous Data and Logs. Take the Feedback from Mr. Umrao and meet with Mr. Vipin. Export the Data backup fro ABL800. ABL800 is working fine. | SWO-124532 | Yes | Checked all the previous Data and Logs. Take the Feedback from Mr. Umrao and meet with Mr. Vipin. Export the Data backup fro ABL800. ABL800 is working fine. | Checked all the previous Data and Logs. Take the Feedback from Mr. Umrao and meet with Mr. Vipin. Export the Data backup fro ABL800. ABL800 is working fine. | Re classified RMED | Courtesy
ABL800 FLEX | SE COLOCA SONDA Y SE ENTREGA OTRA DE SUSTITUCION | SE COLOCA SONDA Y SE ENTREGA OTRA DE SUSTITUCION | D135273B8E1B7360 | No | SOUND IS PLACED AND ANOTHER SUBSTITUTE IS DELIVERED | SOUND IS PLACED AND ANOTHER SUBSTITUTE IS DELIVERED | Defect | HW | Inlet module | Probe bend
ABL800 BASIC | Anpassungen entsprechend des Kundenwunsches vorgenommen | Anpassungen entsprechend des Kundenwunsches vorgenommen | SWO-124762 | No | Adjustments made according to the customer's wishes | Adjustments made according to the customer's wishes | Setup | SW | Configuration | Configuration
ABL800 BASIC | ABLのソフトウェアを再インストール（Ver6.16) | ABLのソフトウェアを再インストール（Ver6.16) | SWO-124380 | No | Reinstall abl software (Ver6.16) | Reinstall abl software (Ver6.16) | Defect | SW | Touch screen | SW6.16
ABL815 FLEX | Fließwege überprüft, Verstopfung im Abfallschlauch beseitigt, AC mit korrekter Lot befüllt, Halbjährliche Wartung durchgeführt. | Fließwege überprüft, Verstopfung im Abfallschlauch beseitigt, AC mit korrekter Lot befüllt, Halbjährliche Wartung durchgeführt. | SWO-124620 | No | Flow paths checked, blockage in the waste hose eliminated, AC filled with correct solder, semi-annual maintenance carried out. | Flow paths checked, blockage in the waste hose eliminated, AC filled with correct solder, semi-annual maintenance carried out. | Cleaned | Fluidic pathway | Inlet module | Probe
ABL800 BASIC | Installed permanent keys successfully, 2 point calibration performed, sample measurement performed | Installed permanent keys successfully, 2 point calibration performed, sample measurement performed | SWO-125096 | No | Installed permanent keys successfully, 2 point calibration performed, sample measurement performed | Installed permanent keys successfully, 2 point calibration performed, sample measurement performed | Setup | SW | Demo Key | Expired
ABL825 FLEX |  | B | D13544738E36F60F | No |  |  | Assisted Replacement | Electrode | Membrane | Reference
ABL825 FLEX |  | B | D1354DDF8E53F613 | No |  |  | Assisted Replacement | Electrode | Membrane | Reference
ABL827 FLEX | Lactate calibrations are failing..Calibration error 408 | Replace pump tubing -electrode. Replaced Lactate electrode (warranty). | D1357A548E0C9892 | No | Lactate calibrations are failing.. Calibration error 408 | Replace pump tubing -electrode. Replaced Lactate electrode (warranty). | Assisted Replacement | Electrode | Electrode | Lac | 48
ABL800 BASIC | Cleaned the pH chamber and flushed the probe and other tubing. now working satisfactorily. | Cleaned the pH chamber and flushed the probe and other tubing. now working satisfactorily. | SWO-125219 | No | Cleaned the pH chamber and flushed the probe and other tubing. now working satisfactorily. | Cleaned the pH chamber and flushed the probe and other tubing. now working satisfactorily. | Cleaned | Fluidic pathway | pH/BG Module | Measuring chamber
ABL810 FLEX | verschieden Aktionen durchgeführt, Prozesse im Taskmanager überprüft ob PC stark ausgelastet ist -keine besonderen Ereignisse, Touchpanel reagiert normal, muß weiter beobachtet werden, bei der Gelegenheit bat Frau Dr. Weber um weitere Schulungen am Gerät - in Arbeit | verschieden Aktionen durchgeführt, Prozesse im Taskmanager überprüft ob PC stark ausgelastet ist -keine besonderen Ereignisse, Touchpanel reagiert normal, muß weiter beobachtet werden, bei der Gelegenheit bat Frau Dr. Weber um weitere Schulungen am Gerät - in Arbeit | SWO-124470 | No | various actions performed, processes in the task manager checked whether PC is busy - no special events, touch panel reacts normally, must be monitored further, on the occasion Dr. Weber asked for further training on the device - in progress | various actions performed, processes in the task manager checked whether PC is busy - no special events, touch panel reacts normally, must be monitored further, on the occasion Dr. Weber asked for further training on the device - in progress | Defect | HW | DMS_PC | Touch screen
ABL810 FLEX | pH Dichtring ersetzt. Startphase QC und Kalibration. | pH Dichtring ersetzt. Startphase QC und Kalibration. | SWO-124496 | No | pH sealing ring. Start phase QC and calibration. | pH sealing ring. Start phase QC and calibration. | Defect | HW | Electrode | pH | Sealing Ring
ABL810 FLEX | Ventile im pH BG Modul ausgebaut und leicht geölt, Lufteinlass gereinigt. LS Justage und Pumpen Kalibration. Startphase QC und Kalibration. | Ventile im pH BG Modul ausgebaut und leicht geölt, Lufteinlass gereinigt. LS Justage und Pumpen Kalibration. Startphase QC und Kalibration. | SWO-124711 | No | Valves in the pH BG module removed and lightly oiled, air inlet cleaned. LS adjustment and pump calibration. Start phase QC and calibration. | Valves in the pH BG module removed and lightly oiled, air inlet cleaned. LS adjustment and pump calibration. Start phase QC and calibration. | Adjustment | HW | pH/BG Module | Valve
ABL837 FLEX | Calibrations falling. Oximetry hardware errors.Checked Oximetry system. Lamp unit failing in service checks. | Replaced Lamp unit. Lamp unit adjusted. | D1355BA08E3C3AB4 | No | Calibrations falling. Oximetry hardware errors. Checked Oximetry system. Lamp unit failing in service checks. | Replaced Lamp unit. Lamp unit adjusted. | Defect | HW | Oxi module | Lamp unit
ABL90 | Found that the inlet gasket did not have the rubber piece installed. | Replaced the inlet gasket. Analyzer is now back in service. No MTR Required. Function of the inlet is working as it should. | D136776604093E43 | No | Found that the inlet gasket did not have the rubber piece installed. | Replaced the inlet gasket. Analyzer is now back in service. No MTR Required. Function of the inlet is working as it should. | Defect | HW | Inlet module | Inlet Gasket
ABL90 FLEX | Touchscreen erneut verstellt --> Touchscreen erneuert | Touchscreen neu kalibriert. Problem mit Medizintechnik besprochen --> ABL lässt sich wieder normal bedienen | SWO-124827 | No | Touchscreen re-adjusted --> Touchscreen renewed | touchscreen recalibrated. Problem with medical technology discussed --> ABL can be operated normally again | Defect | HW | DMS_PC | Touch screen
ABL90 |  | B | D1365144040C2E6D | No |  |  | Assisted Replacement | SC | SC | pCO2
ABL90 FLEX PLUS |  | B | D1366ED104463781 | Yes |  |  | Re classified RMED | PER
ABL90 FLEX PLUS |  | B | D136704A040C3786 | Yes |  |  | Re classified RMED | PER
ABL800 BASIC | examine all valves through valve test in service program. examined Ph BG module valves and EL MET valves visually. valves were ok. checked all the tubings. replaced membrane of po2 and pco2. run protien removal. given two point calibration. analyser running sucvcessfully. | examine all valves through valve test in service program. examined Ph BG module valves and EL MET valves visually. valves were ok. checked all the tubings. replaced membrane of po2 and pco2. run protien removal. given two point calibration. analyser running sucvcessfully. | SWO-125223 | No | examine all valves through valve test in service program. examined Ph BG module valves and EL MET valves visually. valves were ok. checked all the tubings. replaced membrane of po2 and pco2. run protien removal. given two point calibration. analyser running sucvcessfully. | examine all valves through valve test in service program. examined Ph BG module valves and EL MET valves visually. valves were ok. checked all the tubings. replaced membrane of po2 and pco2. run protien removal. given two point calibration. analyser running sucvcessfully. | Defect | HW | pH/BG Module | Valve | Revisit required?
ABL825 FLEX | rinse error code 501, 652, 510, 751, 509 -- annual maintenance written up separatley for documentation tracking | flushed the air valve performed an annual maintenance, found pO2 electrode with no oring replaced electrode and remembraned with customer stock--ran pump cal -- pCO2 was coming in high; remembraned and called customer for results; talked to Orlie and then Davis on swing shift for QC results | D136647E045DC344 | No | rinse error code 501, 652, 510, 751, 509 -- annual maintenance written up separatley for documentation tracking | flushed the air valve performed an annual maintenance, found pO2 electrode with no oring replaced electrode and remembraned with customer stock--ran pump cal -- pCO2 was coming in high; remembraned and called customer for results; talked to Orlie and then Davis on swing shift for QC results | Assisted Replacement | Electrode | Electrode | pO2 | Sealing Ring
ABL825 FLEX |  | B | D136A95E040A4723 | Yes |  |  | Re classified RMED | Logistical upSWO
ABL827 FLEX |  | B | D1366EB20446D7A8 | No |  |  | Defect | HW | Inlet module | Tubing
AQT90 FLEX | Replaced the 903-171 NEEDLE ASSEMBLY. | Clean the Dust fro storage wheel. Shift the AQT-90 Load the TnT and NT Pro BNP along with SP. Run the Sample for Observation. | SWO-125204 | No | Replaced the 903-171 NEEDLE ASSEMBLY. | Clean the Dust fro storage wheel. Shift the AQT-90 Load the TnT and NT Pro BNP along with SP. Run the Sample for Observation. | Defect | HW | Aspiration Module | Bend needle | 3812
AQT90 FLEX |  | B | D1366952044B38FC | No |  |  | Adjustment | HW | Cover and Chassis | Inlet Door | SETSAFE 22
ABL800 BASIC | Oxymetry test failure for neon Lamp, Neon lamp needs to replace with new one. | B | SWO-125236 | No | Oxymetry test failure for neon Lamp, Neon lamp needs to replace with new one. |  | Revisit required | HW | Oxi module | Lamp unit | Neon
ABL80 FLEX BASIC | 更换测试卡修复 | 更换测试卡修复 | SWO-125054 | No | Replacement test card repair | Replacement test card repair
ABL80 FLEX BASIC | tested with zeroing casette lac and po2 were out of order. need to replace sensor board. | tested with zeroing casette lac and po2 were out of order. need to replace sensor board. | SWO-125228 | No | tested with zeroing casette lac and po2 were out of order. need to replace sensor board. | tested with zeroing casette lac and po2 were out of order. need to replace sensor board.
ABL90 FLEX | Fließwege mit Proteinentferner und Sensorkassette mit Wasser gespült. Einlassdichtungen und Solution PCB erneuert. Vorhandene gebrauchte SC und LP wieder eingesetzt. Lösungspack, Inlet Probe und Inlet Gasket with holder zur Reserve bereitgestellt. Einlaufverfahren durchgeführt, alle Kal. und QC wieder o.k. thb Kal. durchgeführt und o.k. 5 Blut-Testmessungen durchgeführt o.k. | Fließwege mit Proteinentferner und Sensorkassette mit Wasser gespült. Einlassdichtungen und Solution PCB erneuert. Vorhandene gebrauchte SC und LP wieder eingesetzt. Lösungspack, Inlet Probe und Inlet Gasket with holder zur Reserve bereitgestellt. Einlaufverfahren durchgeführt, alle Kal. und QC wieder o.k. thb Kal. durchgeführt und o.k. 5 Blut-Testmessungen durchgeführt o.k. | SWO-125251 | No | Flow paths with protein remover and sensor cassette rinsed with water. inlet seals and Solution PCB. Existing used SC and LP reinstated. Solution Pack, Inlet Probe and Inlet Gasket with holder provided as a reserve. inlet procedures, all Cal. and QC again o.k. thb Kal. and o.k. 5 blood test measurements carried out ok. | Flow paths with protein remover and sensor cassette rinsed with water. inlet seals and Solution PCB. Existing used SC and LP reinstated. Solution Pack, Inlet Probe and Inlet Gasket with holder provided as a reserve. inlet procedures, all Cal. and QC again o.k. thb Kal. and o.k. 5 blood test measurements carried out ok. | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
ABL90 FLEX | Einlassnadel getauscht und kompletten Jahreswartungskit verbaut. Lösungspack getauscht(Von Techniker mitgebracht). Sensorkassette erneuert. Kalibrierung durchgeführt. QCs für alle Level durchgeführt. ABL wieder Messbereit. Frau Scheer hat gesagt das der Probentyp nicht im Labor in der EDV ankommt. ABL ist an Opus angebunden. Wurde aber laut Frau Scheer nicht von Radiometer Techniker angebunden. Im Patientenprofil ist Probentyp 2 angegeben. Mit Abstimmung von Frau Scheer den Probentyp auf normalen Probentyp umgeändert. Frau Scheer wird beobachten ob der Probentyp nun im richti Labor ankommt. | Einlassnadel getauscht und kompletten Jahreswartungskit verbaut. Lösungspack getauscht(Von Techniker mitgebracht). Sensorkassette erneuert. Kalibrierung durchgeführt. QCs für alle Level durchgeführt. ABL wieder Messbereit. Frau Scheer hat gesagt das der Probentyp nicht im Labor in der EDV ankommt. ABL ist an Opus angebunden. Wurde aber laut Frau Scheer nicht von Radiometer Techniker angebunden. Im Patientenprofil ist Probentyp 2 angegeben. Mit Abstimmung von Frau Scheer den Probentyp auf normalen Probentyp umgeändert. Frau Scheer wird beobachten ob der Probentyp nun im richti Labor ankommt. | SWO-125269 | No | Inlet needle is exchanged and complete annual maintenance kit installed. Solution pack exchanged (brought by technician). Sensor cartridge is renewed. calibration performed. QCs for all levels. ABL ready to measure again. Ms Scheer said the type of sample does not arrive in the computer lab. ABL is connected to Opus. But according to Ms. Scheer, was not connected by Radiometer technicians. Sample type 2 is indicated in the patient profile. With ms. Scheer's vote, the sample type changed to normal sample type. Ms. Scheer will observe whether the sample type now arrives in the richti laboratory. | Inlet needle is exchanged and complete annual maintenance kit installed. Solution pack exchanged (brought by technician). Sensor cartridge is renewed. calibration performed. QCs for all levels. ABL ready to measure again. Ms Scheer said the type of sample does not arrive in the computer lab. ABL is connected to Opus. But according to Ms. Scheer, was not connected by Radiometer technicians. Sample type 2 is indicated in the patient profile. With ms. Scheer's vote, the sample type changed to normal sample type. Ms. Scheer will observe whether the sample type now arrives in the richti laboratory. | Defect | HW | PM Kit | PM Kit | SC/SP incl
ABL90 FLEX | Remplacement de la carte CPU( la disque dur est échangée) et chargement du logiciel. Apres remis en route l'ABL c'est connecté avec aqure et remis en route sans faute, test de barcode : OK | Remplacement de la carte CPU( la disque dur est échangée) et chargement du logiciel. Apres remis en route l'ABL c'est connecté avec aqure et remis en route sans faute, test de barcode : OK | SWO-125158 | No | Replacing the CPU card (the hard drive is exchanged) and loading the software. After restarted the ABL it's connected with aquor and started without fail, barcode test: OK | Replacing the CPU card (the hard drive is exchanged) and loading the software. After restarted the ABL it's connected with aquor and started without fail, barcode test: OK | Defect | HW | DMS_PC | CPU | Barcode issue
ABL90 FLEX | Service Testprogramme , Sensor-Kassette erneuert | Service Testprogramme , Sensor-Kassette erneuert | SWO-125193 | Yes | Service test programs, sensor cassette renewed | Service test programs, sensor cassette renewed | Re classified RMED | Logistical upSWO
ABL90 | Unit passing External QC, but failing Level 3, result value too high.Unit displaying Replace Solution .Replaced SP with same results.Inspected Fluid path for possible air leaks.Ran pO2 test in diagnostics=passed.Recalibrated=OK. | Replaced SC. | D137D25D7A492DC8 | No | Unit passing External QC, but failing Level 3, result value too high. Unit displaying Replace Solution . Replaced SP with same results. Inspected Fluid path for possible air leaks. Ran pO2 test in diagnostics=passed. Recalibrated=OK. | Replaced SC. | Assisted Replacement | SC | SC | pO2
ABL90 FLEX | Replaced gasket with new and straightened probe | Replaced gasket with new and straightened probe | SWO-125131 | No | Replaced gasket with new and straightened probe | Replaced gasket with new and straightened probe | Defect | HW | Inlet module | Inlet Gasket
ABL90 FLEX | Customer happy with analyser on arrival, advised them that analyser will carry out frequent calibrations after changing SC | Customer happy with analyser on arrival, advised them that analyser will carry out frequent calibrations after changing SC | SWO-125183 | No | Customer happy with analyser on arrival, advised them that analyser will carry out frequent calibrations after changing SC | Customer happy with analyser on arrival, advised them that analyser will carry out frequent calibrations after changing SC | No fault found | No fault found | No fault found | No fault found
ABL90 FLEX PLUS |  | B | D137DADA7A0C1A54 | No |  |  | Defect | HW | PM Kit | PM Kit
ABL90 FLEX | 拆进样支架做清洁正常---jack.chu | 拆进样支架做清洁正常---jack.chu | SWO-124968 | No | Remove the sample bracket to do clean normal ---jack.chu | Remove the sample bracket to do clean normal ---jack.chu | Defect | HW | Inlet module | Inlet Gasket
ABL825 FLEX |  | B | D137B7B57A052721 | No |  |  | Assisted Replacement | Electrode | Electrode | Na+
ABL837 FLEX | THE TBIL FOR PATIENT RESULTS WAS REPORTING LOW | THE HEMOLYZER WAS REPLACED | D137ACEF7A5BED8D | No | THE TBIL FOR PATIENT RESULTS WAS REPORTING LOW | THE HEMOLYZER WAS REPLACED | Defect | HW | Oxi module | Hemolyzer
ABL805 FLEX | Found solution was leaking from inlet pre heater on inlet module. | Replaced inlet pre heater. Verified instrument operation. | D138F088F05FB0B1 | No | Found solution was leaking from inlet pre heater on inlet module. | Replaced inlet pre heater. Verified instrument operation. | Defect | HW | Inlet module | Inlet preheater
ABL815 FLEX | Regulierung getestet, Pumpen und Sensoren kalibriert, erneut membraniert mit neuer Charge, mehrere Kal, QC | Regulierung getestet, Pumpen und Sensoren kalibriert, erneut membraniert mit neuer Charge, mehrere Kal, QC | SWO-125342 | No | Regulation tested, pumps and sensors calibrated, re-membraneed with new batch, multiple Kal, QC | Regulation tested, pumps and sensors calibrated, re-membraneed with new batch, multiple Kal, QC | Assisted Replacement | Electrode | Membrane | pCO2
AQT90 FLEX | Clean the Dust fro storage wheel. Shift the AQT-90 Load the TnT and NT Pro BNP along with SP. Run the Sample for Observation. | Clean the Dust fro storage wheel. Shift the AQT-90 Load the TnT and NT Pro BNP along with SP. Run the Sample for Observation. | SWO-125206 | Yes | Clean the Dust fro storage wheel. Shift the AQT-90 Load the TnT and NT Pro BNP along with SP. Run the Sample for Observation. | Clean the Dust fro storage wheel. Shift the AQT-90 Load the TnT and NT Pro BNP along with SP. Run the Sample for Observation. | Re classified RMED | Relocate analyzer | See mail | 216-8-24 / PRARA | 3812
AQT90 FLEX | Touch Screen ersetzt, Touch Screen kalibriert u. gespeichert, Startphase abgewartet, QC i.O. | Touch Screen ersetzt, Touch Screen kalibriert u. gespeichert, Startphase abgewartet, QC i.O. | SWO-125137 | No | Replacement of touch screen, touch screen calibrated and saved, start phase awaited, QC i.O. | Replacement of touch screen, touch screen calibrated and saved, start phase awaited, QC i.O. | Defect | HW | Analyzer Control Module | Display
AQT90 FLEX |  | B | D14085D9B41FB031 | No |  |  | Defect | SW | Analyzer Control Module | Reinstall XP
AQT90 FLEX | 機器清掃 サービスでSafeポジション、システム洗浄。OK。 | 機器清掃 サービスでSafeポジション、システム洗浄。OK。 | SWO-126085 | No | Safe position and system cleaning service with equipment cleaning service. Ok. | Safe position and system cleaning service with equipment cleaning service. Ok. | Defect | SW | Analyzer Control Module | CORRUPTXML | 35
ABL90 FLEX | LP und Schlauchsystem geprüft, Inlet geprüft, Einlassdichtung ersetzt, Funktionskontrolle ok | LP und Schlauchsystem geprüft, Inlet geprüft, Einlassdichtung ersetzt, Funktionskontrolle ok | SWO-125387 | No | LP and hose system tested, Inlet tested, inlet seal replaced, function control ok | LP and hose system tested, Inlet tested, inlet seal replaced, function control ok | Defect | HW | Inlet module | Inlet Gasket
ABL90 FLEX | 回路洗浄。サービスにて各チェック。 | B | SWO-125883 | No | Circuit cleaning. Each check at the service. |  | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
ABL90 FLEX PLUS | ﾍﾓﾗｲｻﾞｰ交換 | B | SWO-125210 | No | Hemolyzer replacement |  | Defect | HW | Oxi module | Hemolyzer
ABL90 FLEX | 更换测试卡正常。 | 更换测试卡正常。 | SWO-125233 | No | The replacement test card is OK. | The replacement test card is OK. | Assisted Replacement | SC | SC | pCO2
AQT90 FLEX |  | B | SWO-125217 | No |  |  | Defect | HW | Process Module | Optical Unit | 131
ABL800 FLEX | Replaced printer head. | Replaced printer head. | SWO-125256 | No | Replaced printer head. | Replaced printer head. | Defect | HW | Printer module | Printer Head
ABL80 FLEX BASIC | 更换了电源模块902-970后修复---jack.chu | 更换了电源模块902-970后修复---jack.chu | SWO-125211 | No | Repair ---jack.chu after replacing power module 902-970 | Repair ---jack.chu after replacing power module 902-970
ABL80 CO-OX |  | B | D13B8091DC36BF45 | No
ABL90 FLEX | Im Rahmen der Wartung alle Schläuche erneuert und Glaskörper des pCO2 sensors gereinigt. Fehler ist allerdings seit einigen Tagen verschwunden. Wahrscheinlich lag ein Fehler im LP vor. | Im Rahmen der Wartung alle Schläuche erneuert und Glaskörper des pCO2 sensors gereinigt. Fehler ist allerdings seit einigen Tagen verschwunden. Wahrscheinlich lag ein Fehler im LP vor. | SWO-125452 | No | As part of maintenance all hoses renewed and vitreous body of the pCO2 sensor cleaned. Mistakes, however, have disappeared for a few days. There was probably an error in the LP. | As part of maintenance all hoses renewed and vitreous body of the pCO2 sensor cleaned. Mistakes, however, have disappeared for a few days. There was probably an error in the LP. | Assisted Replacement | SP | SP | pCO2
ABL90 FLEX PLUS | Removed and replaced OXI tubing. Flushed Analyser and tested using fluid transport programs. Unable to calibrate or QC as solution pack has run out of activities. | B | SWO-125538 | No | Removed and replaced OXI tubing. Flushed Analyser and tested using fluid transport programs. Unable to calibrate or QC as solution pack has run out of activities. |  | Defect | HW | Oxi module | Pinch valve tube
ABL90 FLEX PLUS | informed service Manager who will inform RMED regarding this issue of overlapping info on printouts. | informed service Manager who will inform RMED regarding this issue of overlapping info on printouts. | SWO-125730 | No | informed service Manager who will inform RMED regarding this issue of overlapping info on printouts. | informed service Manager who will inform RMED regarding this issue of overlapping info on printouts. | Defect | SW | Printout | Printout
AQT90 FLEX | 清洁机器内部灰尘，更换定标卡定标后已修复 | 清洁机器内部灰尘，更换定标卡定标后已修复 | SWO-125316 | No | Clean the dust inside the machine and replace the calibration card and fix it | Clean the dust inside the machine and replace the calibration card and fix it | Cleaned | HW | Storage Module | Storage wheel | 161
ABL800 BASIC | 做了去蛋白清洁，擦拭电极舱，仪器恢复 | 做了去蛋白清洁，擦拭电极舱，仪器恢复 | SWO-125370 | No | Did the de-protein clean, wipe the electrode compartment, the instrument to restore | Did the de-protein clean, wipe the electrode compartment, the instrument to restore | Cleaned | Fluidic pathway | Fluidic pathway | Measuring chamber | 476
ABL835 FLEX | Modulet er udskiftet med et ombytter modul. Det ombyttede modul er rengjort for støv. Spejl i modulet og spejl i laser er renset. Modulet er testet og fundet OK igen. Alt er OK. | Modulet er udskiftet med et ombytter modul. Det ombyttede modul er rengjort for støv. Spejl i modulet og spejl i laser er renset. Modulet er testet og fundet OK igen. Alt er OK. | SWO-125981 | No | The module is replaced with a replacement module. The exchanged module is cleaned of dust. Mirror in the module and mirror in the laser is cleaned. The module has been tested and found OK again. Everything is OK. | The module is replaced with a replacement module. The exchanged module is cleaned of dust. Mirror in the module and mirror in the laser is cleaned. The module has been tested and found OK again. Everything is OK. | Cleaned | HW | AutoCheck Module | Barcode reader
ABL827 FLEX | Glucose is too unstable and has been turned off. Error code 0408 Glucose zero error. | Replaced Glucose electrode and membranes. Perfomed Zero current calibration. | D13BB8E3DC09D294 | No | Glucose is too unstable and has been turned off. Error code 0408 Glucose zero error. | Replaced Glucose electrode and membranes. Perfomed Zero current calibration. | Assisted Replacement | Electrode | Electrode | Glu
ABL825 FLEX | Flex Q Modul geprüft, Funktionskontrolle ok | Flex Q Modul geprüft, Funktionskontrolle ok | SWO-125393 | No | Flex Q module checked, function control ok | Flex Q module checked, function control ok | Unknown | HW | FlexQ Module | Barcode reader
ABL80 FLEX BASIC | 更换测试卡正常。 | 更换测试卡正常。 | SWO-125619 | No | The replacement test card is OK. | The replacement test card is OK.